## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault_Class
Z-0902-2015 | 69773 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA Reagent Container, Model No. B1037307 (15 mL), B1034626 (10 mL), and B1034634 (4 mL) |  |  | November 13, 2014 | 2014 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is recalling the ABX PENTRA 400 Reagent Container because they do not always sit properly into the reagent rack. In low reagent level scenarios, this could lead to insufficient reagent volume being pipetted and could lead to incorrect results being reported without an alarm. The defect would result in a false high or false low results for HbA1c and false high results for magnes | Package design/selection | A recall letter was sent on 11/13/14 to all customers who purchased the ABX PENTRA 400 Clinical Chemistry Analyzer. The letter informs the customer of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact their local HORIBA Medical representative. |  | 443 | 443 units | Nationwide in US | Terminated | March 19, 2015 | 126 | 443 | Not_Computer
Z-0880-2015 | 69964 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Alere INRatio PT/INR Test Strips, Alere INRatio |  |  | December 05, 2014 | 2014 | Alere San Diego, Inc. | In certain cases, an INRatio PT/INR Monitor system may provide an INR result that is significantly lower than a result obtained using a reference INR system (laboratory method). This issue can arise if the patient has certain medical conditions. It can also occur if the instructions in the labeling for performing the test are not followed. | Under Investigation by firm | An Urgent: Medical Device Correction, dated December 5, 2014, was sent to Healthcare Professionals. The letter identified the affected product and reason for recall. Customers were instructed to: -Ensure they have read and understand the precautions described in the current product labeling (a complete list of product labeling is provided in Appendix A) and the additional precautions in this notice describing medical conditions that may increase the risk of obtaining a (falsely or erroneously) lower than expected INR result. The INRatio PT/INR Monitor system should NOT be used if the patient has any of the medical conditions described. -Verify that the patient has hematocrit within the range of 30% to 55%. If not, immediately transition the patient to an alternate INR monitoring method.-Perform INR verification testing for the patients using a laboratory INR test method. Immediately transition any patient with a significantly discrepant low result on the INRatioÂżÂż PT/INR Monitor system to an alternative INR monitoring method.-If have forwarded product to another customer, please provide a copy of this letter to them.-Please complete and FAX or e-mail the enclosed Reply Form (Appendix B the last page of this notice) within 10 days to confirm your receipt of this notice. Customers can contact Alere INRatio Recall Hotline by phone at 1-877-929-2579, for questions regarding this notification. Additionally, a website has been established providing information and frequently asked questions. www.inr-care.comPlease FAX or e-mail the completed Reply Form to:Alere San Diego, Inc.Fax: 1-877-929-2580E-mail: Alere4319@stericycle.com. |  | 23894578874 | 77,342 units (12 test strips/box) and 308,947 units (48 test strips/box) | Worldwide Distribution -- US nationwide, Canada, South America, EU, Asia Pacific, Africa, and Middle East. | Open, Classified |  |  | 23895042808 | Not_Computer
Z-0910-2015 | 70144 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | McKesson Cardiology ECG Management |  |  | December 19, 2014 | 2014 | McKesson Israel Ltd. | An ECG procedure time stamp is incorrect, due to a software error in the McKesson Cardiology ECG Management with software versions 13.1 and 13.1.1. | Software design | McKesson sent an Urgent Field Safety Notice dated December 16, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed how to identify the potentially affected procedures and update the procedure date (if relevant) until a software update is made advailable for installation by McKesson. McKesson will proactively identify prior procedures affected by this issue and will notify customers of any issues identified. In addition to the Field Safety Notice, McKesson Israel will provide correction to the affected SW versions and will contact customers for deployment of a software patch that prevents this scenario from occurring. Customers with questions should contact McKesson Support at 1-866-777-0202.For questions regarding this recall call 404-338-3556. |  | 9 | 9 | Worldwide Distribution - USA including MA, MS, NH, NC, TX, WA and Internationally to the United Kingdom. | Terminated | February 25, 2016 | 433 | 9 | Not_Computer
Z-0906-2015 | 69466 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Optetrak Asymmetric Hi-Flex Posterior Stabilized Cemented Femorals, size 3, Catalog #244-02-03, Catalog #244-03-03. |  |  | October 09, 2014 | 2014 | Exactech, Inc. | Mislabeled | Labeling mix-ups | The firm, Exactech, sent an "Urgent Field Safety Notice" letter dated September 30, 2014 to its Consignees/customers by electronic mail on October 9, 2014. The letter described the product, problem and actions to be taken. The customers were instructed to: immediately cease distribution or use of the products; extend information to your accounts that may have this product in their possession; verify whether you have any of the subject femoral components from the specified serial number ranges; complete and return the attached inventory RESPONSE FORM via fax to: 011 61 2 9638 0778 within 48 hrs of receipt of the notice, and contact your Exactech representative to confirm quantities at your location and arrangements for product return and inventory restocking.If you have any questions, contact Manager, Regulatory Affairs at 352-377-1140. |  | 96 | 96 | International Distribution only to countries of: Argentina, Austria, Australia, China, Cyprus, France, Greece, India, Japan, Lebanon, The Netherlands, Spain and Tunisia. | Open, Classified |  |  | 96 | Not_Computer
Z-0909-2015 | 69785 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Cypher MIS Screw System-The system includes screws, various types and sizes of rods, and set screws. Various instruments are also available for use by the surgeon to facilitate implantation of the device. |  |  | October 08, 2014 | 2014 | Ebi, Llc | Inoperative screw head seat splay and/or fracture as a result of positioning of the rod which is a part of the Cypher MIS Screw System. | Device Design | The firm, EBI LLC/Biomet, sent an "URGENT PRODUCT INFORMATION" letter dated 10/6/2014 via Fed Ex to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to follow the steps outlined in the letter. If you have any questions, please feel free to contact Director of Regulatory Affairs at 303-501-8548. |  | 2816 | 2816 approx.units ( 2688 units US and 128 OUS) | Worldwide Distribution: US Distribution to states of: AL, CA, CT, IN, LA, NC, NJ TX &amp; WI; and the country of: The Netherlands. | Terminated | December 21, 2015 | 439 | 2816 | Not_Computer
Z-0908-2015 | 70046 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | I.M. Hole Locator Instrument. Orthopedic manual surgical instrument. |  |  | November 12, 2014 | 2014 | DePuy Orthopaedics, Inc. | Use of excessive force when impacting may lead to intra-operative femoral fracture. | Device Design | The firm, DePuy Synthes, sent an "URGENT INFORMATION - DEVICE CORRECTION NOTICE HP LCS IM Hole Locator" notification dated November 12, 2014 to the affected distributors. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The distributors were instructed to: review the instructions and the information in this notice; provide this notice to your sales consultants, and complete and return the DISTRIBUTOR RESPONSE CARD: Verify Sales Consultant Notifications via fax to (574) 371-4939 or email: kearle@its.jnj.com. For questions, please contact Recall Coordinator, at 574-371-4917 (M-F; 8 a.m. - 5 p.m. EST). |  | 1130 | 1130 | Worldwide Distribution: US (nationwide) including states of: AZ, CA, CO, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MT, NC, NV, NY, OH, PA, TN, TX, VA, WA, and WI; and countries of: Canada, Australia, Austria, Belgium, Czech Republic, France, Germany, India, Israel, Italy, Japan, Malaysia, Russia, South Africa, Switzerland, and United Kingdom. | Terminated | February 24, 2015 | 104 | 1130 | Not_Computer
Z-0907-2015 | 70044 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Ti-Bond Marketing Literature/Labeling, Catalog No. LT10000, MP-14000 |  |  | December 18, 2014 | 2014 | Spinal Elements, Inc | Spinal Elements is recalling the Ti-Bond related brochures that includes LT-10000 and MP-14000 (all revisions) the content regarding the Ti-Bond coating within the literature is unsubstantiated in context and/or misleading for the coating's use in spine applications. | Labeling False and Misleading | Spinal Elements sent an Urgent Field Action letter dated December 22, 2014, to all affected customers. The letter informed the customers to cease all distribution of the literature material. The letter informed the customers that the content regarding the Ti-Bond coating in the literature is unsubstantiated in context and/or misleading for the coating's use in spine applications. In addition to the customer notiification letter, it will include a list of claims related to the field action. The letter also informed customers of the actions to be taken. Customers were instructed to immediately examine their literature inventory, discontinue the use and /or distribution of the affected brochure, destroy all copies or send them back to:Spinal Elements, Inc.3115 Melrose DriveCarlsbad, CA 92010If product was further distributed customers should identify those customers and notify them at once. Customers were also instructed to complete and return the enclosed "Customer Response Form."For questions regarding this recall call 760-607-0121, ext 206.Customers with questions were instructed to contact Customer Service at 1-877-SPINALS, ext 213.Customers with questions were instructed to contact Customer Service at (877) SPINAL5 Ext 213. |  | 1200 | 1200 units | Nationwide Distribution | Terminated | March 10, 2015 | 82 | 1200 | Not_Computer
Z-0905-2015 | 70043 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Femoral Impactor ATTUNE INTUITION Impactor. |  |  | November 10, 2014 | 2014 | DePuy Orthopaedics, Inc. | ATTUNE INTUITION Impactors (cat. nos. 254401003, 254401004, and 254401006) have broken and produced small pieces. Should a fracture occur and not be observed during surgery, there is the potential for these small fractured pieces of the instrument to be left in the patient. | Device Design | DePuy Synthes sent an Urgent Information letter dated November 10, 2014, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to complete the Hospital/User Facility Acknowledgement Car and return to their Deputy Orthopaedics Sales Consultant or fax to 574-371-4939. Customer swere also asked to retain a copy of the complete Acknnowlegement Card in their files along with the notice. Customers with questions were instructed to call 574-371-4917 (M-F; 8 a.m.  5 p.m. EDT).For questions regarding this recall cal 574-371-4917. |  | 4661 | 4661 | Worldwide Distribution: US (nationwide) including P.R. except AK and RI; and Internationally to: Chile, Canada, Australia, Austria, Belgium, China, France, Ireland, India, Israel, Italy, Japan, Malaysia, Netherlands, New Zealand, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Thailand, United Arab Emirates (UAE), and United Kingdom (UK). | Terminated | December 08, 2015 | 393 | 13077 | Not_Computer
Z-0922-2015 | 69927 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | IR remote control JUPITER / TS7500, Material no.1601927, Serial number: to (incl.) 102222333. |  |  | December 03, 2014 | 2014 | Trumpf Medical Systems, Inc. | Complaints have been reported of unintended movement occurring during surgical cases due to the use of damaged JUPITER remotes | Component change control | Trumpf sent an Urgent Medical Device Correction letter on December 3, 2014, to all affected customers. Customers were instructed to complete and return the Confirmation of Receipt form with serial numbers of the affected table(s) and the serial number(s) for any remote controls they have. Replacement remotes will be sent. Customers were instructed to dispose of their recalled remotes. Customers with questions were instructed to call 888-474-9359.For questions regarding this recall call 843-534-0606. |  | 3303 | 3303 total all devices | Worldwide Distribution - USA including AL, AZ, CA, CT, FL, GA, IL, KS, LA, MD, MI, MK, MS, NC, NH, NJ, NY, OH, OK, PA, RI, SC, TX, VA, WA, WI and Internationally to Canada. | Open, Classified |  |  | 3440 | Not_Computer
Z-0913-2015 | 69114 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Symbia Intevo Excel is a non-diagnostic SPECT ICT system with CT support for only attenuation correction and anatomical localization. |  |  | August 12, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Siemens Medical Solutions USA, Inc. has received reports of unintended radial motion of the detectors. This unintended motion, should it occur, may cause the patient compression related injuries. | Device Design | Siemens sent Customer Safety Advisory Notice letters to all customers of record. The letters instructed customers to: 1) continue to use their system(s); 2) place the safety advisory in the system's instructions for use; and, 3) to anticipate that their local service engineer will contact them to perform an inspection and schedule any necessary repairs by December 2014. Customers were also instructed to share the letter with their support staff and, if the equipment was sold or transferred, to forward the safety notice to the new owner of the equipment. Customers with questions can contact their local service representative or Siemens Medical Solutions at 800-888-7436. |  | 6 | 6 systems | Worldwide Distribution - USA (naitonwide) and Internationally to Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lithuania, Malaysia, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Norway, Oman, P.R. China, Pakistan, Peru, Philippines, Poland, Portugal, Qatar, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Open, Classified |  |  | 2213 | Not_Computer
Z-0914-2015 | 69992 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | e.cam -emission computed tomography system |  |  | April 08, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | There is a potential for patient finger injury on the e.cam systems. | Use error | The firm, Siemens, sent a "Customer Safety Advisory Notice" letter dated April 8, 2014 to all direct accounts (customers). The letter described the product, problem and actions to be taken. The letter instructed the customers to follow the instructions in the e.cam operator's manual and the e.cam Patient Bed Safety Addendum of the operator's manual; and to always monitor the patient at all times while they are on the system during initiation of any system movement. The letter further requested that customers forward the letter to the new owner of the equipment in the event that the system is sold or transferred. Customers with questions about this letter can contact Siemens Medical Solutions USA at 800-888-7463 (Monday - Friday, 7:00 AM - 8:00 PM EST). |  | 4462 | 4,462 systems | Worldwide Distribution: US (nationwide) including states of: AK, AR, AZ, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, MA, ME, MI, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WY and Puerto Rico; and countries of: Algeria , Angola , Argentina , Australia , Austria , Azerbaijan , Bahrain , Bangladesh , Belarus , Belgium , Bolivia , Bosnia Herzeg. , Brazil , Bulgaria , Cameroon , Canada , Chile , Colombia , Costa Rica , Croatia , Cuba , Czech Republic , Dem. Rep. Congo , Denmark , Dominican Rep. , Ecuador , Egypt , El Salvador , Finland , France , Georgia , Germany , Ghana , Greece , Guam , Honduras , Hong Kong , India , Indonesia , Iran , Iraq , Ireland , Israel , Italy , Jamaica , Japan , Jordan , Kazakhstan , Kosovo , Kuwait , Latvia , Lebanon , Luxembourg , Macedonia , Malaysia , Mauritius , Mexico , Morocco , Myanmar , Nepal , Netherlands , New Zealand , Niger , Nigeria , Norway , Oman , P.R. China , Pakistan , Palestinian Ter , Peru , Philippines , Poland , Portugal , Qatar , Rep. of Yemen , Republic Korea , Romania , Russian Fed. , Saudi Arabia , Serbia , Singapore , Slovakia , Slovenia , South Africa , Spain , Sri Lanka , Sweden , Switzerland , Syria , Taiwan , Thailand , Trinidad,Tobago , Tunisia , Turkey , Uganda , Ukraine , United Arab Emirates , United Kingdom , Uruguay , Uzbekistan , and Vietnam. | Open, Classified |  |  | 4462 | Not_Computer
Z-0940-2015 | 69902 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | IS3000 Vision Side Cart (VSC) used in conjunction with the da Vinci Si Surgical System IS3000. |  |  | December 05, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive Surgical is initiating a voluntary correction related to da Vinci Si System (IS3000) Vision Side Carts, with threaded-in caster wheels, which can result in the caster wheel bolts gradually unthreading and the wheels loosening. | Device Design | The firm, Intuitive Surgical, sent an "Urgent Medical Device Correction" letter dated December 5, 2014 to all customers on December 5, 2014 by trackable method. The letter described the product, problem and actions to be taken. The customers were instructed to: ensure all affected personnel are fully informed of this notice; identify affected product, and retain a copy of this notice with your da Vinci Si user manual. Please note: the da Vinci Si Vision Side Cart can continue to be used safely. Ensure that all wheels are fully unlocked before attempting to move the Vision Side Cart. Avoid rolling the Vision Side Cart over uneven surfaces (i.e., over cables or a threshold).Intuitive Surgical is working on a newly designed wheel assembly. Uponavailability, Field Service will retrofit the wheel assembly of the affected VisionSide Carts with the new design.If you need further information or support concerning this Medical Device Correction, please contact your Clinical Sales Representative or Intuitive Surgical Customer Service at the numbers listed below:ÂżÂż North America and South America: 8008761310 Option 3 (6 a.m. to 5 p.m. PST)ÂżÂż Japan: 0120565635 or 0355751362 (9 a.m. to 6 p.m. JST)ÂżÂż South Korea: 0232713200 (9 a.m. to 6 p.m. KST)ÂżÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 a.m. to 6 p.m.) or ics@intusurg.com |  | 202 | 202 | Worldwide Distribution: US (nationwide) and countries of: Belgium, France, Germany, Romania, India, Israel, Italy, South Korea, Norway, Slovak Republic, Spain, Sweden, Turkey, and United Kingdom. | Terminated | March 03, 2015 | 88 | 202 | Not_Computer
Z-0934-2015 | 70009 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | AlignRT is a video-based three-dimensional (3D) surface imaging system which is used to image the skin surface of a patient in 3D before and during radiotherapy treatment. |  |  | November 11, 2014 | 2014 | Vision Rt Inc | Potential use error when the external Gate Controller is turned ON after the Vision RT software has been opened. | Under Investigation by firm | VisionRT sent an Urgent Medical Device Correction Notification letter dated November 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers shall:A. Print the instructions in Appendix 3, place them in the front of the product guide and ensure these instructions are read and understood by all those who need to be aware of them within their organisation;B. Maintain awareness of this issue until all actions indicated in this letter have been successfully completed;C. Promptly inform Vision RT if they believe that patient harm occurred due to this issue; ANDD. Complete the acknowledgement in Appendix 1 and return it via email to corrections@visionrt.com.Contact Vision RTShould you have any queries on this letter, please do not hesitate to contact Vision RT by telephone on +44 20 83464300 (866 778-2379 from the US) or as per http://www.visionrt.com/contact/details.For questions regarding this recall call 866-594-5443. |  | 55 | 55 | Nationwide Distribution | Open, Classified |  |  | 55 | Not_Computer
Z-0933-2015 | 69961 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CelluTome Epidermal Harvesting System Harvester (5.0 CM) STERILE EO Rx Only Manufactured for: KCI USA, Inc. Part Number CT-H50 or CT-H25 |  |  | November 19, 2014 | 2014 | KCI USA, INC. | KCI has received reports that, in a small number of cases, the CelluTome Harvester (a component of the CelluTome Epidermal Harvesting System), blades have the potential to drift during shipping and could prematurely protrude into openings of harvester top plate. | Device Design | KCI sent an Urgent - Voluntary Medical Device Field Notification letter to all affected consignees on December 8, 2014. The letter notified the customer of the issue, and provided instructions to inspect their harvester prior to use. If blade drift is identified, instructionswere provided to the customer to either return the device to KCI, or if they wish, how to reset the blade prior to use. The letter instructs the customer to respond to the recall notifcation by returning the product reconciliation form attached the letter. Follow-up notifcations will be send to non-responding consignees.Consignees with questions were instructed to contact their local KCI Representative or Customer Technical Service Center at 1-800-275-4524 and select option 3.For questions regarding this recall call 210-515-4108. |  | 5695 | 1,139/5 Packs | Worldwide Distribution - USA (Nationwide) and Internationally to Australia, Austria, Canada, France, Germany, Ireland, Italy, Isreal, Netherlands, South Africa, Spain, Sweden, Switerland, Turkey, UAE, and United Kingdom. | Terminated | September 18, 2015 | 303 | 5695 | Not_Computer
Z-0937-2015 | 69979 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Stryker Orthobiologics Imbibe Aliquot Needle Bone Cement Needle- Model Numbers: 2110-0505 Beveled Needle, 11 gauge x 4 inch |  |  | December 04, 2014 | 2014 | Orthovita, Inc., dBA Stryker Orthobiologics. | There is the potential for a breach in the inner or outer packaging pouches of all lots of the Stryker Orthobiologics Imbibe Aliquot Needle Bone Cement Needle. | Package design/selection | The firm, Stryker, sent an "URGENT PRODUCT RECALL" letter dated December 4, 2014, to end users/customers. The letter identified the product, problem, and actions to be taken. Customers were instructed to complete and return the URGENT MEDICAL DEVICE RECALL NOTIFICATION ACKNOWLEDGMENT FORM within 5 days via fax to 215-253-5020, contact other hospitals in your territory that have the affected products to arrange return of the product, and return the affected product to: Stryker Orthobiologics, 67 Great Valley Parkway, Malvern, PA 19355. If you have any questions, feel free to contact the Associate Manager, Quality Assurance and Regulatory Compliance at 610-640-1775 x 5299. |  | 2155 | 2155 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | July 28, 2015 | 236 | 91925 | Not_Computer
Z-0939-2015 | 69468 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HomeChoice Automated PD system and HomeChoice Pro Automated PD system Baxter's |  |  | October 06, 2014 | 2014 | Baxter Healthcare Corp. | There are additional and updated warning and cautions that are not in the Patient At-Home Guide for HomeChoice Automated PD Systems and HomeChoice PRO Automated PD Systems. | Labeling design | The firm, Baxter, sent an "IMPORTANT PRODUCT INFORMATION" letter dated October 6, 2014 to all peritoneal dialysis provider direct accounts. Baxter also sent an "IMPORTANT PRODUCT INFORMATION" letter dated October 9, 2014 to all peritoneal dialysis patient direct accounts The letters described the product, problem and actions to be taken. The letters included instructions for the recipients to complete the attached Customer Reply Form and to return it to Baxter via scanned e-mail to: fca@baxter.com or fax : 224-270-5457. In addition, the letter to patients instructed them to contact their healthcare provider to discuss this issue.Customers with clinical questions should contact Baxter's Renal Division Clinic Helpline at 888-736-2573, option 2 (8:00 AM - 5:00 PM CST). Adverse reaction or quality problems experienced with the use of these product can be reported by contacting Baxter Product Surveillance at 800-437-5176 (Monday - Friday, 8:00 AM - 5:30 PM CST) or e-mail to corporate_product_complaints_round_lake@baxter.com. Customers with general questions about the letter should contact The Center for One Baxter at 800-422-9837 (Monday - Friday, 8:00 AM - 5:00 PM CST). |  | 96458 | 96,458 systems | Worldwide Distribution: US (nationwide) Including states of: AK, AL, AR, AZ, CA, CO, CT, DE, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and District of Columbia; and countries of: Algeria, Argentina, Aruba, Australia, Austria, Bahamas, Bahrain, Barbados, Belarus, Belgium, Bosnia, Brazil, Bulgaria, Cayman Islands, Chile, Colombia, Croatia, Cuba, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kuwait, Lebanon, Macedonia, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Poland, Portugal, Quatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | March 28, 2016 | 539 | 96458 | Not_Computer
Z-0936-2015 | 69422 | No | IKK | Isokinetic testing and evaluation system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Cervical Extension Machine |  |  | December 03, 2013 | 2013 | MedX Holdings, Inc. | Review of DHRs revealed that HiPot, Continuity, load and friction re-testing was not completed on these devices during manufacturing. | Unknown/Undetermined by firm | Medx began contacting its consignees December 30, 2013, to schedule testing on the devices at their facility. The technician has been scheduled to complete required testing of these devices in the field.For questions regarding this recall call 322-351-2005. |  | 4 | 4 | Worldwide Distribution - USA including MO, IL, ID, and MN and Internationally to Germany and South Korea. | Open, Classified |  |  | 13 | Not_Computer
Z-0931-2015 | 69753 | No | HTC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Small Electric Drive (SED), part number: 05.001.175: The Small Electric Drive is designed for use in general traumatology, especially hand and food applications involving surgical procedures such as drilling, burring, reaming, pin and wire placement, cutting of bone and hard tissue. |  |  | November 06, 2014 | 2014 | Synthes (USA) Products LLC | The device may operate solely in reverse mode, not operate in reverse mode when intended, not function, or unintentionally start. Unintended start of the hand piece or wrong mode/direction could cause bone/soft tissue, peripheral nerve, and /or vascular damage. Injuries may need surgical or medical intervention. Surgical delay and postoperative local infection are also potential risks. | Device Design | DePuy Synthes sent an Urgent Notice: Medical Device Removal letter dated November 6, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify their consignees using the hospital personnel letter, dated November 6, 2014, as their template for their consignee notification letter, review their inventory and remove the affected lots from stock.Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number. Complete the Verification Section (page 4 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided. Return the Verification Section (page 4 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Send a copy of the completed Verification Section to DePuy Synthes by: Fax: (610) 430-7083 or Scan/email: Fieldaction@synthes.comIf you DO NOT have the identified product, please take the following steps: Complete the attached Verification Section (page 4 of this letter) by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information. Return the documents to DePuy Synthes by: Fax: (610) 430-7083 or Scan/email: Fieldaction@synthes.com Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 4 of the notification.Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 12 | 12 | US Nationwide Distribution to Florida only | Terminated | August 17, 2015 | 284 | 12 | Not_Computer
Z-0932-2015 | 69586 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter Amia Automated PD systems are used in the treatment of adult renal failure patients to perform automated peritoneal dialysis. |  |  | April 24, 2014 | 2014 | Baxter Healthcare Corp. | System error 01779 is produced when the battery cannot be charged due to specific voltage differences between the battery and the charger. | Under Investigation by firm | Baxter sent an IMPORTANT PRODUCT INFORMATION letter dated April 24, 2014 to all affected customers. The peritoneal dialysis provider consignees were instructed to: 1) complete the enclosed customer reply form and return it to Baxter; and, 2) contact Baxter Technical Services at 800-553-6898 should a System Event 01179 occur. Baxter Technical Services is available 24 hours a day, 7 days a week. Clinical questions can be directed to Baxter's Renal Clinical Helpline at 888-736-2543, option 2 (Monday - Friday, 8:00 AM - 5:00 PM CT). General questions regarding the Import Product Information letter can be direct to The Center for One Baxter at 800-422-9837 (Monday - Friday, 8:00 AM - 5:00 PM CT).For questions regarding this recall call 800-422-9837. |  | 55 | 55 units | Nationwide Distribution including GA, MA, NC, NY, PA, and WA. | Terminated | July 21, 2017 | 1184 | 55 | Not_Computer
Z-0941-2015 | 69806 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CSA Medical truFreeze System; Model: CC3-01, a cryogenic surgical device |  |  | November 10, 2014 | 2014 | CSA Medical | An increase in complaint trending for a loss of audible sound associated with the timer has been observed. | Software design | The firm, CSA Medical, sent a letter to its consignees 11/14/2014, via Fedex overnight . The letter explains the event and asks consignees to us a verbal call of time during the cryogen spray period until the root cause is determined and a permanent correction is implemented. Consignees are also asked to complete an acknowledgment form and return it to CSA Medical via, fax, email, telephone, etc. |  | 46 | 46 | US Distribution to states of: AR, CA, CO, CT, DC, FL, GA, IL, IN, LA, MA, MD, MI, NC, NJ, NY, OK, PA, RI, SC, TN, TX, UT, VA, and WV. | Terminated | March 25, 2015 | 135 | 46 | Not_Computer
Z-0930-2015 | 69987 | No | FST | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HarmonyÂż vLED Surgical Lighting System, one unit packaged per box |  |  | November 26, 2014 | 2014 | Steris Corporation | STERIS has learned that groups of LED lights (modules) contained within the vLED light head may illuminate inconsistently. | Software Design Change | STERIS sent an Urgent Voluntary Field Correctionn Notice to all affected consignees via FedEx. The Consignee Notification mailing was initiated the week of December 15, 2014. Customers were instructed that their local STERIS representative would contact their facility to arrange for the installation of updated software. Customers with questions were instructed to call 1-800-548-4873.For questions regarding this recall call 440-392-7231. |  | 140 | 140 units | Worldwide Distribution - USA including Alabama, Alaska, Arkansas, California, District of Columbia, Florida, Illinois, Indiana, Kansas, Louisiana, Mississippi, Nevada, New Jersey, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, Puerto Rico, South Carolina, Tennessee, Texas, Utah, and Virginia and Internationally to Canada, China, Greece, Indonesia, Republic of Korea, Romania, Russian Federation, Taiwan, and Viet Nam. | Terminated | July 11, 2015 | 227 | 140 | Not_Computer
Z-0946-2015 | 69937 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista Myoglobin Calibrator (MYO CAL), Lot 4FD085 |  |  | November 24, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Dimension Vista Myoglobin Calibrator (MYO CAL) may produce a positive shift in MYO QC and Patient test results that exceeds acceptance criteria for this product. A positive shift observed up to 12% at MYO concentrations within and above the reference range of the assay. Depending on quality control limits, this drift may not be detected. | Under Investigation by firm | The firm, Siemens, sent an "Urgent Medical Device Recall"letter dated November 2014, to all customers who received the affected lot. The letter described the product, problem and actions to be taken. The customers were instructed to: Review this letter with your Medical Director; discard your remaining inventory; recalibrate with an alternate in-date lot of MYO CAL, and complete and return the attached Field Correction Effectiveness Check Form and fax the form to (312) 275-7795. Siemens will replace any unused inventory of the affected lot at no charge.If you have any questions, please contact your Siemens Customer Care Center - Technical Solutions at 800-441-9250 or your local Siemens technical support representative. |  | 1230 | 1230 | Worldwide Distribution: US (nationwide) and country of: Canada. | Open, Classified |  |  | 1230 | Not_Computer
Z-0945-2015 | 70081 | No | OFA | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | Enforcement Discretion | Pacemaker Tray (Cardiovascular Surgical Instruments convenience kit) |  |  | October 23, 2014 | 2014 | Centurion Medical Products Corporation | According to the recall notice received from Hospira, the recall was initiated due to a confirmed customer report of particulate in a single unit. Hospira identified the particulate as human hair, embedded in and attached to a pinched area of the stopper. | Nonconforming Material/Component | Centurion sent an Urgent Product Recall Notice dated October 23, 2014, to al affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify and remove all affected inventory that may be in their possession and to forward a copy of thie noctice to any cusotmers to whom the affected product may have been further distributed. Customers were asked to complete the enclosed accountability record whether not they have the product and fax to 517-546-3356. Customers with questions were instructed to call 517-546-5400 Ext. 1135.For questions regarding this recall call 517-546-5400, ext 1135. |  | 8 | 8 kits | Nationwide Distribution including MA, IN, NY, GA, AL, PA, and RI. | Terminated | June 08, 2015 | 228 | 1290 | Not_Computer
Z-0949-2015 | 70066 | Yes | NVZ |  |  | Cardiovascular | PMA | Merlin@home RF Remote Monitoring Transmitter, Model EX1150, with software versions 6.1M or 7.0. |  |  | December 18, 2014 | 2014 | St Jude Medical Cardiac Rhythm Management Division | Transmitter may initiate a software reset resulting in backup operation in some implanted St. Jude Medical Radio Frequency (RF) enabled Implantable Cardioverter Assura, Unify Assura, and Quadra Assura ICDs and Assurity and Allure Pacemakers. | Device Design | The firm, St. Jude Medical, sent an "Important Medical Device Correction" Dear Doctor Letter dated 12/18/14 to customers to inform them that the Merlin@home transmitters may initiate a software reset resulting in backup operation in some implanted St. Jude Medical Radio Frequency (RF) enabled Implantable Cardioverter Defibrillators (ICDs) and Pacemakers. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns are instructed to contact their local St. Jude Medical representative or St. Jude Medical's Technical Services Department at 1-800-722-3774. |  | 257456 | 257,456 units | Worldwide Distribution: US (nationwide) and countries of: EU, Mideast, Canada, Australia/New Zealand, and Asia-Pacific (including Japan). | Terminated | March 30, 2016 | 468 | 257456 | Not_Computer
Z-0948-2015 | 69300 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Abbott m2000sp is an automated system for performing sample preparation for nucleic acid testing. |  |  | November 23, 2013 | 2013 | Abbott Molecular | The x-drive assembly may fail prematurely. This could cause the manipulator or liquid handling arm to contact an object, shake and drip, or improperly pipet. | Nonconforming Material/Component | Abbott Molecular sent a different Field Correction Recall / Urgent Field Safety Notice letters dated October 24, 2014 to each of three sets of customers. Customers whose instruments' X-Drive Assembly belts were perviously inspected by a Field Service Representative and found to not require replacement were advised to contact their local Abbott Molecular Representative or Abbott Molecular Technical Support at 800-553-7042 x 2 should they hear any unusual noises associated with the LiHa or RoMa movement or misalignment of the LiHa or RoMa. Customers whose instruments' X-Drive Assembly belts were previously inspected and replaced by a Field Service Representative were notifieied that this maintenance activity took place. Customers with questions were instructed to contact Abbott Molelcular.For questions regarding this recall call 224-361-7000.Customers whose instruments' X-Drive Assembly belts have not been inspected were notified that a Field Service Representative is scheduled to visit their facility to inspect and if necessary, replace the X-Drive Assembly belt(s). Customers were advised to contact their local Abbott Molecular Representative should they hear any unusual noises associated with the LiHa or RoMa movement or misalignment of the LiHa or RoMa. |  | 176 | 176 devices | Worldwide Distribution - USA including CA, FL, GA, IL, MI, NC, NJ, NY, OH, OK, OR, and TX and Internationally to Australia, Austria, Belgium, Bosnia And Herzegovina, Brazil, Canada, China, Colombia, Costa Rica, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Indonesia, Ireland, Israel, Italy, Kazakhstan, Kenya, Malawi, Mali, Mexico, Morocco, Mozambique, New Zealand, Nigeria, Norway, Pakistan, Poland, Russia, Saudi Arabia, Senegal, South Africa, South Korea, Taiwan, Thailand, Turkey, Ukraine, and United Kingdom. | Open, Classified |  |  | 176 | Not_Computer
Z-0492-2015 | 69787 | No | GAP | Natural nonabsorbable silk surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon PERMA-HAND Silk siliconized black braided 10 strands per packet Ethicon LLC- Made in USA. |  |  | October 17, 2014 | 2014 | Ethicon, Inc. | Ethicon is unable to ensure that PERMA-HAND silk suture (siliconized black braid) meets the expected sterility assurance level for sterile products. | Other | Ethicon sent an Urgent Medical Device Recall Letter/Business Reply Form dated 10/17/2014 via UPS next day mail. |  | 2508 | 2,508 each | Nationwide Distribution | Terminated | August 14, 2015 | 301 | 2508 | Not_Computer
Z-0954-2015 | 70063 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Proteus XR/a with wall stand model number 600-0301. |  |  | December 08, 2014 | 2014 | GE Healthcare | GE Healthcare has recently become aware of a potential safety issue involving the breakage of the two steel cables which support the bucky device in the wall stand of Proteus XR/a X-ray imaging systems.No injuries reported. | Nonconforming Material/Component | Consignees were sent on 12/8/2014 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 17107, dated December 8, 2014. The letter was addressed to Hospital Administrators /Risk Managers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product correction and Contact information (For questions in USA call 800-437-1171, For other countries contact your local GE Healthcare Service Representative. |  | 3096 | 3096 (2486 US, 610 OUS). | USA: All states in continental US including, Guam, DC.OUS; ARAB EMIRATES, AUSTRIA, BOSNIA AND HERZEGOVINA, BRAZIL, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HONDURAS, HONG KONG, HUNGARY, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, KUWAIT, LATVIA, LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PHILIPPINES, POLAND, POLYNESIA, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM. | Terminated | February 03, 2016 | 422 | 3096 | Not_Computer
Z-0953-2015 | 70014 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | The ATTUNE INTUITION Impaction Handle is a re-useable instrument utilized in knee replacement procedures. The instrument interfaces with several related ATTUNE instruments: keel punch, tibial tower, evaluation bullet, fixed bearing tibial impactor, rotating platform tibial impactor, and femoral impactor, for the purpose of impaction, extraction, and insertion applications at steps within the surgical procedure. Due to the frequency of its use, two impaction handles are supplied in each kit. |  |  | November 10, 2014 | 2014 | DePuy Orthopaedics, Inc. | Complaints have indicated that the ATTUNE INTUITION Impaction Handle lever has fractured during impaction. If the lever fractures, there is the potential for fractured pieces of the device, or the inner spring, if released, to be left in the patient if not observed during surgery. | Device Design | On 11/10/2014, URGENT INFORMATION ATTUNE INTUITION IMPACTION HANDLE AND IMPACTORS DEVICE CORRECTION NOTICE notifications were sent to the affected distributors and medical professionals. All distributors were notified via electronic mail, and all medical professionals with affected inventory were also notified by distributor or Sales Consultant via in person delivery or mail or electronic mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, please contact Kim Earle, Recall Coordinator, at 574-371-4917 (M-F; 8 a.m. - 5 p.m. EDT). |  | 7944 | 7,944 | Worldwide Distribution-US (nationwide) except AK and RI, including PR and the countries of Chile, Canada, Australia, Austria, Belgium, France, Hong Kong, Ireland, India, Israel, Italy, Japan, Malaysia, Netherlands, New Zealand, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Thailand, United Arab Emirates (UAE), and United Kingdom (UK | Terminated | December 08, 2015 | 393 | 7944 | Not_Computer
Z-0950-2015 | 69419 | No | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | Labeled in part: I Color Complete / Color Lens, blister packing, 14.5mm. |  |  | September 10, 2014 | 2014 | Intertrade Imports Inc. | A surveillance sample collected by FDA confirmed the presence of Bacillus Cereus and Bacillus Lentus in the contact lenses. I Color Color Complete brand color contact lenses are unapproved medical devices. | Material/Component Contamination | The firm verbally contacted each distributor on September 10, 2014 to notify them about the recall and actions that needs to be taken. |  | 750 | 750 units | Worlldwide Distribution - USA including Texas, California, Florida, and St. Lucia, Barbados and St. Marten. | Open, Classified |  |  | 3600 | Not_Computer
Z-0952-2015 | 70032 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | POWEREASE Instruments Set Screw Breakoff Instrument, REF 2346010, Medtronic Sofamor Danek, USA, Inc. Spinal surgical instrumrent. |  |  | December 03, 2014 | 2014 | Medtronic Sofamor Danek USA Inc | The retaining tabs component of the Set Screw Retaining Drive may be oversized, which will not allow proper mating between the Set Screw Retaining Drive and the set screw. | Process design | Medtronic sent an Urgent - Voluntary Market Removal letter dated December 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Instructions to Customers:1.) Please remove the product from normal storage locations. Do not use this product.2.) As applicable, contact your Medtronic Sales Representative to coordinate the return of affected product. Alternatively, you may call Medtronic's Global Quality Department at (800) 876 -3133, extension 3197 for assistanceIf you need assistance or have any related questions or concerns, please do not hesitate to contact your Medtronic Sales Representative or the Global Quality Department at (800) 876 - 3133, extension 3197. We sincerely apologize for this inconvenience, but this action is necessary to assure that our high standardof quality is maintained. |  | 24 | 24 units | US Distribution to the states of : CA, KS, KY, NC, TX, WA, MI, OH, IA, PA, OR, MS, FL, NY, GA and AR. | Terminated | September 24, 2015 | 295 | 24 | Not_Computer
Z-0947-2015 | 69872 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | Phadia 1000 Instrument (introduced as UniCAP 1000), Article number: 12-3800-01 (All instrument software versions since launch in 2003); Multiple ImmunoCAP assays for Allergen Testing, FDA Cleared for use on Phadia 1000 instrument with IU statement-see attachments. Software utilized in multiple ImmunoCAP assays for Allergen Testing, FDA Cleared for use on Phadia 1000 instrument with IU statement-see attachments. |  |  | November 20, 2014 | 2014 | Phadia US Inc | During an investigation of instrument logs it was determined that In specific circumstances involving multiple steps, a rack sequencing error may occur. This will result in a mismatch between the sample ID and the test result reported for all subsequent sample racks in that run. | Software design | ThermoFisher sent an Urgent Medical Device Product Correction letter to all affected customers. The firm will inform customers of the specific sequence of errors and operator actions that can trigger the event via letter and phone, and will request that they do not respond to the Barcode Read Error and let the sample with the error eject normally from the instrument and be rerun after the barcode read error has been resolved. The firm will issue a mandatory Phadia 1000 Instrument Software (ISW) update that will correct the issue and make efforts to install it within 45 days from date of availability. Customers were asked to return the written response form via fax to: 1-888-243-5214Attention: Customer Support, Phadia 1000 ISW Product Correction Response; EMAIL: Compliance-us.idd@thermofisher.comEmail Subject Line: Phadia 1000 ISW Product Correction Response.Customers with questions were instructed to contact Phadia US Technical Support at 800-346-4364, option 2.For questions regarding this recall call269-492-1957. |  | 146 | 146 | Nationwide Distribution including AZ, CA, CO, CT, FL, GA, HI IL, IN, KS, MD, MI, MN, NC, NJ, NV, NY, OH, OK, OR, PA, TN, TX, UT, VA, and WA. | Terminated | June 30, 2015 | 222 | 146 | Not_Computer
Z-0958-2015 | 69587 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Expression MR200 MRI Patient Monitoring System. |  |  | October 22, 2014 | 2014 | Invivo Corporation | The MR 200 devices failed the initial power up test due to capacitor short and hi-pot testing; possible loss of Patient Monitoring or User/Service Electrical Shock. | Component design/selection | Philips Medical Systems sent an Urgent - Medical Device Correction Letter dated October 22, 2014, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Our records indicate that you have an affected device. The following page provides additional instructions and actions that will be taken to address this problem. If you need any further information or support concerning this issue, please contact your local Philips representative: 877-468-4861 option 1, then option 2. |  | 20 | 20 | CA, FL, IN, KY, MN, NC, NJ, NY, OK, SD and TX Netherlands, Poland and United Kingdom. | Terminated | January 09, 2017 | 810 | 20 | Not_Computer
Z-0959-2015 | 69751 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | IC PROFILER-MR - Model 1123. |  |  | August 08, 2014 | 2014 | Sun Nuclear Corporation | Has a Potential Pinch Hazard to user performing Machine QA in an MRI environment. | Use error | Sun Nuclear Corp. sent a memo dated 08/08/2014 to their 3 customers of the possible manufacturing error. |  | 6 | 6 Units. | Distributed in the states of CA, CO, and OH. | Terminated | March 23, 2015 | 227 | 6 | Not_Computer
Z-0963-2015 | 69747 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator; Model. Numbers: M3535A, M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6 |  |  | November 19, 2014 | 2014 | Philips Medical Systems, Inc. | Device may malfunction, which could cause therapy to be delayed, disabled, or delivered inadvertently. | Software design | On November 19, 2014, Philips notified consignees of the recall via letter. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.A software upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice |  | 74499 | 74,499 devices | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 74499 | Not_Computer
Z-0964-2015 | 70042 | No | MRD |  |  | General &amp; Plastic Surgery | 510(k) | MENTOR MemoryGel Resterilizable Gel Sizer Moderate Plus Profile STERILE. Breast prosthesis sizer. |  |  | December 11, 2014 | 2014 | Mentor Texas, LP. | The IFU provided in the affected units of the MENTOR MemoryGel Resterilizable Sizer is incorrect. Those units include the IFU for the MENTOR MemoryShape Resterilizable Sizer. | Employee error | Mentor sent an Important Labeling Correction letter dated December 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm plans to follow-up with sales representative visits. The recall letter notifies the customer of the issue, but does not instruct them to return the product. Should the customer chose to return the product, all returned product will be tagged to denote no further distribution. The recalled units under firm control have been placed on hold in the inventory system.If you have any questions related to this notice, please contact your local Mentor Sales Representative, or Mentor Customer Service at 1-800-235-5731. |  | 45 | 45 units | US Distribution to the states of : MA, WV, NY, FL, ME, NC, TX, IL, KS, TN, NJ, PA, CA, OR, WI, GA, and WA | Terminated | January 04, 2016 | 389 | 45 | Not_Computer
Z-0915-2015 | 69931 | Yes | NPT |  |  | Cardiovascular | PMA | Lotus TAVR 23mm, |  |  | November 19, 2014 | 2014 | Boston Scientific Corporation | Lotus valve became unlocked during release from the delivery system. This may lead to percutaneous insertion of another valve, resulting in a Valve in Valve (ViV) implantation. If this occurs, the first valve could embolize and it may be necessary to convert the patient to surgery. | Device Design | Boston Scientific notified consignees on November 19, 2014, of the problem via letter through overnight delivery directing consignees to segregate and return affected product to Boston Scientific utilizing the Field Safety Corrective Action (FSCA) Instructions provided. The FSCA details that consignees should: Immediately discontinue use of and segregate product; Complete and return the Account Reply Verification Tracking form and Package/Ship the affected product. |  | 278 | 278 units, all sizes | Distributed only in the countries of Finland, France, Germany, Great Britain, Italy, Norway, Spain, Sweden, and Switzerland. | Terminated | March 26, 2015 | 127 | 278 | Not_Computer
Z-0840-2015 | 69660 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE OEC Brivo 715 Prime Mobile C-Arm X-Ray Product |  |  | September 15, 2014 | 2014 | GE OEC Medical Systems, Inc | mA Accuracy may exceed the design and labeling specification of+ 10%. Per 1020.32(f): Deviation of x-ray tube potential and current from the indicated values shall not exceed the maximum deviation as stated by the manufacturer in accordance with 1020.30(h)(3). | Radiation Control for Health and Safety Act | Planned action(s) to repair defect or to bring product into compliance: GE Healthcare will, without charge, perform a mA calibration on all affected systems in the USA to bring the product back into compliance. The details of which will be included in a subsequent communication to customers or through a GE Healthcare field engineer site visit.Should customers have any questions or concerns regarding these corrections, please do not hesitate to contact the service team for further information at 800-874-7378 option 8. |  | 61 | 61 units installed in US | Worldwide Distribution | Terminated | May 20, 2015 | 247 | 61 | Not_Computer
Z-0965-2015 | 70010 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | MUM-1 (Multiple myeloma oncogene-1), catalog number PRM352 AA; |  |  | December 12, 2014 | 2014 | Biocare Medical Llc | Three lots of MUM-1 may not meet their labeled shelf life, as a drop in staining intensity over time has been observed. | Incorrect or no expiration date | Biocare sent a Medical Device Recall notifications letter to all customers by traceable method. The letters identified the affected product, problem and actions to be taken. Users are asked to immediately discontinue used of the affected lots and dispose of them according to applicable national, state and local laws. Users are requested to complete and return the response form so that proper credit can be issues 800-799-9499. |  | 59 | 59 | Worldwide Distrinution - US Nationwide in the states of TX, WI, MO, WV, SC, MA, WA, TN, AZ, CA, IL, NY, LA and countries of Kuwait, India, Venezuela, New Zealand, Turkey, Uruguay. | Terminated | March 30, 2015 | 108 | 59 | Not_Computer
Z-0960-2015 | 69995 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia T and Symbia T2 computer tomography systems. These systems are used to perform SPECT, CT, and SPECT + CT scans and imaging studies. |  |  | March 05, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | There are confirmed reports of a component coming loose inside the CT enclosure and damaging the system.components. | Component change control | The firm, Siemens, sent an undated "Customer Safety Advisory Notice" letter to all customers. The described the product, problem and actions to be taken. The letters inlcuded instructions for customers to: 1) continue to use their systems while waiting for the field correction to be scheduled and performed; 2) ensure that the safety advisory is placed in the system's instruction for use; and, 3) closely observe the patient during studies. The letter also instructs customers of record to forward a copy of the letter to the new owner if the equipment was sold or transferred. Customers with any questions about this issue can contact Siemens Molecular Imaging at 800-888-7436 (US). |  | 295 | 295 systems | Worldwide Distribution: US Nationwide including Puerto Rico; Austria, Australia, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Czech Republic, Ecuador, Egypt, Finland, France, Germany, Great Britain, Hungary, Ireland, India, Iran, Italy, Japan, Mexico, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Russia, Singapore, Slovakia, Slovenia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, and Tunisia. | Open, Classified |  |  | 295 | Not_Computer
Z-0971-2015 | 69957 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Force; the intended use of computed tomography is to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angels or spiral planes taken at different angles. |  |  | December 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Artifacts are found in acquired imaging on the SOMATOM Force. | Software Manufacturing/Software Deployment | A customer advisory notice, dated December 3, 2014, was sent to end users that identified the product, problem, and action to be taken. Customers were informed that new software would be installed via remote access. |  | 7 | 7 | PA, MA, MN, SC, MD, and TX. | Terminated | April 06, 2015 | 124 | 7 | Not_Computer
Z-0978-2015 | 70060 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Timberline Anchored Lateral Variable Retractable Sleeve LITPCR8600-2101-07, Model No. 86300204 |  |  | August 22, 2013 | 2013 | Biomet Spine, LLC | The inner diameter of the Sleeve shaft is undersized, resulting in interference fit when inserting the Anchored Lateral Straight Retractable Awl or the Anchored Lateral Straight Retractable Drill into the Sleeve shaft. | Component design/selection | The firm, Lanx, sent their Consignees/customers a letter dated October 15, 2013. The letter described the product, problem and actions taken. Lanx and their distributors replaced the affected inventory with the conforming product.Please contact the Chief Compliance Officer at 303-501-8534 or by email: Randy.fesmire@lanx.com. |  | 12 | 12 units | US Distribution to states of: AZ, CA, NY, TN, and TX including PR. | Terminated | March 13, 2015 | 568 | 12 | Not_Computer
Z-0976-2015 | 69091 | No | JSG | Multipurpose culture medium. | Microbiology | Microbiology | 510(K) Exempt | Columbia Agar with 5% Sheep Blood |  |  | June 06, 2014 | 2014 | Becton Dickinson &amp; Co. | A low level of surface and subsurface contamination of Listeria monocytogenes on 9 lots of non-sterile prepared plated media containing sheep blood, typically visible immediately upon removal from the packaging. The contamination is typically observed as 1-3 CFUs per plate when present. | Material/Component Contamination | BD Diagnostic Systems sent a Field Correction Notice dated August 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to continue to follow all surveillance requirements of their laboratory upon confirmation of Listeria monocytogenes. Laboratory staff should notify their respective state health department. Customers with questions should contact BD Technical Services at 800-638-8663.For questions regarding this recall call 410-316-4000. |  | 775 | 775 cases | Worldwide Distribution - US nationwide including Puerto Rico, Belgium, Canada, Philippines, Singapore, and Taiwan. | Terminated | August 12, 2016 | 798 | 9922 | Not_Computer
Z-0970-2015 | 70082 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | EXp Tibial Insert ; Intended Use of Device(s): The Proven Gen-Flex Total Knee System is intended for: a. Total knee replacement due to osteoarthritis, osteonecrosis, rheumatoid arthritis, and/or post-traumatic degenerative problems; and b. Revision of failed previous reconstructions where sufficient bone stock and soft tissue integrity are present. |  |  | December 01, 2014 | 2014 | Stelkast Co | Difficulty locking EXp tibial inserts into tibial trays, caused by a manufacturing condition where a lead in radius on the anterior locking tab is not present on certain insert lots. The absence of this lead in may prevent or make assembly of the insert and tray difficult. | Process design | On December 1, 2014, phone calls were made to all direct accounts to notify them of the recall and to request immediate recovery and return of the SC3425 EXp PS Tibial Tray Inserts and SC3453 EXp CR-HF Inserts from their inventory. |  | 553 | 553 | TX, MD, IN, KS, CA, PA, GA, WA, NY, IL. | Terminated | May 06, 2015 | 156 | 553 | Not_Computer
Z-0973-2015 | 70018 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Malleable C- Retractor, Part Number 397.232, used to provide retraction of the soft tissues during use of trocar instrumentation. Used in the following systems: the Basic Trocar System, Universal Trocar System, Craniofacial Modular Fixation System, Craniofacial (CMF) Distraction System, MatrixORTHOGNATHIC Plating System, and the Matrix MANDIBLE Plating System. |  |  | December 15, 2014 | 2014 | Synthes, Inc. | Potential failure and/or corrosion of the internal spring of the retractor, which may have been manufactured from an incorrect material. Surgical delay may occur. Spring particles could transfer from the instrument to the patient, and may lead to an adverse tissue reaction and/or infection. | Component design/selection | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated December 15, 2014, to its end users and sales consultants. The letter identified the product, problem, and actions to be taken. The End users/sales staff were instructed to immediately remove affected stock from inventory; call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number; complete and return the attached Verification Section to DePuy Synthes via fax: 866-827-3421 or Scan/email: Synthes8645@stericycly.com, even if you do not have any of the affected product, and return any affected devices to Synthes.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes sales consultant. |  | 1762 | 1762 | Worldwide Distribution: US (nationwide) and internationally to Canada. | Open, Classified |  |  | 1762 | Not_Computer
Z-0854-2015 | 69478 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/ zeego systems, model numbers: 10094135, 10094137, 10094139, 10094141, 10280959: dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed include cardiac angiography, neuro angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | September 17, 2014 | 2014 | Siemens Medical Solutions USA, Inc | The Small Focus of a 3 focal spot x-ray tube may fail and further release of radiation will not be possible. The Large Focus and Micro Focus focal spots will still be operational without any restriction. Manual switchover is needed, resulting in a short interruption of the procedure. | Process control | Planned action (s) to repair defect or to bring product into compliance for Siemens Medical Solutions Inc. concerning the corrective action plan (CAP) for the Model(s) Artis Zee/Zeego Stationary Angiographic X-ray Systems.1.Notifying the customer of the defect.2.Implementing a corrective software patch to protect the system from damage inthe event of hardware failure3. Scheduling a service visit from a Siemens representative to perform a hardware replacement.4. Replacing the non-conforming X-ray tube with one manufactured after the nonconformity issue was resolved.For further questions please call (610) 219-6300. |  | 44 | 44 | USA Distribution to the states of MO, MI, NY, OK, CA, LA, CT, IN, VA, MD, WA, HI, ME, TX, NY, OH, PA and CO. | Terminated | August 06, 2015 | 323 | 44 | Not_Computer
Z-0977-2015 | 70065 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | lnline Persuader- PedFuse Pedicle Screw System; Part Number: 11-80164 |  |  | December 19, 2014 | 2014 | SpineFrontier, Inc. | A crack may develop in the handle to the Inline Persuader, which may progress if the user continues to use the affected instrument. | Device Design | Spinefrontier sent an Urgent Voluntary Recall Notification letter dated December 16, 2015, to all affected customers. The notice informed customers of the recall and requests that they initiate the immediate return of the parts listed. In addition, SpineFrontier is requiring confirmation via recall acknowledgement form that consignees have received this notice and that they will initiated the immediate return of the parts. Customers can contact SpineFrontier at 978-232-3990 Monday through Friday 8:30AM to 5:30PM EST with any questions. |  | 4 | 4 devices | Nationwide Distribution including TX, OH, VA, and Washington, DC. | Terminated | January 11, 2016 | 388 | 4 | Not_Computer
Z-0967-2015 | 69830 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 3600 Immunodiagnostic System, Catalog Number 6802783, IVD --- Ortho Clinical Diagnostics. |  |  | November 19, 2014 | 2014 | Ortho-Clinical Diagnostics | Software Anomaly: the firm has identified an anomaly with VITROS System Software Version 3.1 and below, and determined that the software may not properly identify an expired calibration. | Software design | On 11/19/2014, IMPORTANT PRODUCT CORRECTION NOTIFICATION letter (Ref. CL2014-296, dated 11/19/2014) was sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to the US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them of the required actions. Immediately install the most recent ADD using the All Assay Data option. If any of your calibrations are expired, the system will post PZB-005 condition codes. Perform calibrations prior to using those reagents. As expected, calibrations that will expire within one hour will also be flagged. Continue to install the most recent ADD at least once each week using the All Assay Data load option so that the system properly flags expired calibrations. --- Foreign affiliates were informed by e-mail on 11/19/2014 of the issue and instructed to notify their consignees of the issue and required actions. |  | 658 | 658 total: USA - 144, Foreign - 514 | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 2827 | Not_Computer
Z-0972-2015 | 70016 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis Q, and Artis Q.zen ceiling systems; model numbers: 10848281; 10848354: The Artis family is dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed include cardiac angiography, neuro angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | December 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Suboptimal routing of the cable may result in increased wear over time. Without additional measures, damage of a cable may result in limited functionality and in rare cases, failure of a system. An ongoing procedure could be terminated in rare cases. | Device Design | Siemens sent an important customer safety information notice, dated December 3, 2014, to all affected customers. The letter identified the product, problem, and action to be taken. The following hardware modifications will be implemented with the corrective action AX042/14/S:1. A modified cable outlet to provide additional space for the cabling.2. A modified plastic cover at the cable outlet providing optimal cable guidance in the area of the arc. In general all systems will be checked for existing damage prior to the implementation of this action.For questions regarding this recall call 610-219-6300. |  | 50 | 50 | Nationwide Distribution | Terminated | April 06, 2015 | 124 | 50 | Computer
Z-0966-2015 | 70073 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery IGS 730 &amp; Discovery IGS 740. |  |  | December 10, 2014 | 2014 | GE Healthcare | Potential unintentional system motion or rotation due to a positioning user interface issue with Discovery IGS 730 and 740 products. | Software design | Consignees were sent a 12/10/2014 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 12223 dated December 10, 2014. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information ( US call 800-437-1171, other countries contact your local GE Healthcare Service Representative). |  | 24 | 24 ( 7 US, 17 OUS) | Worldwide Distribution-Distributed in the states of CO, CT, MS, NJ, NY, OH, and WA, and the countries of BELGIUM, CANADA, CHINA, FRANCE, GERMANY, JAPAN, MEXICO, PERU, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | June 01, 2015 | 173 | 24 | Not_Computer
Z-0985-2015 | 70074 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Revolution CT., multi-slice (256 detector row) CT scanner. |  |  | December 19, 2014 | 2014 | GE Healthcare | Potential safety issue due to a required torque check on three bolted joints of the Revolution CT scanner.No injuries or illnesses reported. | Process control | Consignees were sent on 12/19/2014 a GE Healthcare "Urgent Medical Device Correction" letter GEHC # 25457 dated December 19, 2014. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (US phone 800-437-1171, Japan 0120-055-919, other countries contact your local GE Healthcare Service Representative). |  | 13 | 13 (7 US, 6 OUS) | US: FL, WA, IL, UT, NY, CA.OUS: Canada, France, Hong Kong, Japan, Switzerland, United Arab Emirates. | Terminated | April 29, 2015 | 131 | 13 | Not_Computer
Z-0983-2015 | 70070 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Timberline Anchored Lateral Fixed Retractable Sleeve Model No. 8630-0205. Timberline Anchored Lateral Fixed Retractable Sleeves are used with the Timberline MPF System to help establish access into the vertebrae for the screws utilized to anchor the plate to the vertebrae. The Timberline Retractable Drill and the Timberline Retractable Awl fit within this Sleeve. |  |  | August 22, 2013 | 2013 | Biomet Spine, LLC | The inner diameter of the Sleeve shaft is undersized, resulting in interference fit when inserting the Anchored Lateral Straight Retractable Awl or the Anchored Lateral Straight Retractable Drill into the Sleeve. | Device Design | The firm, Lanx, sent a letter to their Consignees/customers dated October 15, 2013. The letter describes the product, problem and actions taken. Lanx and their distributors replaced the affected inventory with conforming product. Please contact Chief Compliance Officer at 303-501-8534 or by email: Randy.fesmire@lanx.com if you have any questions regarding this recall. |  | 14 | 14 units | US Distribution to states of: AZ, CA, NY, TN and TX including PR. | Terminated | March 13, 2015 | 568 | 14 | Not_Computer
Z-0990-2015 | 69997 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | Civco, MTSR03 Shoulder Retractor with Wrist Loops Kit with Base, plus two pairs of wrist loops (nylon and sherpa), Non-sterile re-usable wool wrist strap, REF: 20SR03, Rx |  |  | December 11, 2014 | 2014 | Med Tec Inc | The product (wool wrist loop) was manufactured 7cm shorter than product specifications. | Component design/selection | The firm sent e-mailed Field Safety Notice letters and contacted customers via phone on December 11, 2014, informing customers of the issue with the product. Customers were advised they may accept the product as is or request a replacement. A response form was enclosed for consignees to indicate the status of the product.Please contact a CIVCO insides sales representative for additional assistance. You may reach your representative by calling CIVCO Medical Solutions, Radiation Oncology, at 800-842-8688 or 712-737- 8688 between 8:00AM and 5:00PM CST. |  | 2 | 2 | Worldwide Distribution - US Distribution to the states of : AZ, MO, PA and WI., and the countries of : Austria, India and Japan. | Open, Classified |  |  | 13 | Not_Computer
Z-0957-2015 | 69960 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iPlan RT is a radiation treatment planning system that is intended for use in stereotactic, conformal, computer planned, Linac based radiation treatment of cranial, head and neck, and extracranial lesions. |  |  | November 19, 2014 | 2014 | Brainlab AG | iPlan RT Radiation Treatment Planning Software: Potentially incorrect patient positioning when using multiple localized CT image data sets. | Software in the Use Environment | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 19, 2014, to all affected customers. The letter included instructions to: 1) If clinically not required, do not use multiple localized CT scans within one treatment plan (to avoid potentially incorrect Reference and Alignment Set assignments in the first place). 2) If you must use multiple localized CT scans, e.g. for recurrent treatment planning of the same patient, always guarantee that the latest (see note *) CT scan is defined as both Alignment Set and Reference Set during treatment planning. Brainlab will provide a software solution to prevent the described scenario from occurring. Brainlab will actively contact affected customers tentatively starting January 2016 to schedule the update. Brainlab will additionally refine the instructions for use of the iPlan RT treatment planning software in regard to implications of Alignment Set and Reference Set selections and will provide this user guide update to existing affected customers together with the software update.Customers with questions can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 1412 | 1412 systems total | Worldwide Distribution - USA including Alabama, Arizona, Arkansas, California, Colorado, Connecticut, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Vermont, Virginia, Washington, Wisconsin, Wyoming and Internationally to Algeria, Angola, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Latvia, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, Norway, Oman, Panama, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Uruguay, and Venezuela. | Open, Classified |  |  | 1412 | Computer
Z-0982-2015 | 70056 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Anchored Lateral Retractable Drill, Timberline IPF IFU IFU8621-0113, Model 8630-0102. |  |  | August 22, 2013 | 2013 | Biomet Spine, LLC | The outer diameter of the Drill shaft is oversized; thereby, resulting in interference fit when inserting the Drill into the Fixed or Variable Sleeve Assembly. | Component design/selection | Distributors and Direct Sales Reps notified via letter to return product. |  | 15 | 15 | Distributed in PR and the states of NY, AZ, TN, TX, and CA. | Terminated | March 13, 2015 | 568 | 15 | Not_Computer
Z-0986-2015 | 70093 | No | LWH | Surgical sponge scale. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | The SurgiCounter Charging Cradle |  |  | June 16, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | Faulty SurgiCounter Charging Cradle can prevent the scanner from charging or downloading information after a procedure. | Device Design | On 6/16/2013, Improvements to the SurgiCounter" charging base notifications were emailed to the affected users with instructions for requesting a replacement product. The notification included a description of the reason for the recall, affected product, and instructions for responding to the notification. For questions, call the Support team any time at (877) 520-2300 option 1, or support@scmd.com. |  | 3625 | 3625 | US Nationwide Distribution in the states of AK, AL, AR, AZ, CA, CT, FL, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY. | Terminated | August 20, 2015 | 795 | 3625 | Not_Computer
Z-0979-2015 | 69969 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Various Brilliance and Ingenuity Computed Tomography X-ray systems; one unit per package |  |  | December 08, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Improper assembly of the vertical motor/brake system may lead to uncommanded vertical motion. | Process control | Philips sent an Urgent - Field Safety Notice letter dated December 4, 2014 to all affected customers. The US consignees will receive this notice via certified mail and Philips Key Market representatives will distribute the FSN internationally. Additionally, A Field Service Engineer will service the system to resolve the issue. For questions contact your local Philips representative or local Philips Healthcare office. |  | 500 | 500 units | US Nationwide Distribution - AR, AZ, CA, CO, FL, FL, GA, IL, IN, MD, MI, MN, NC, ND, NJ, NY, OH, OK, OR, PA, PA, SC, TN, TX, VA, and WA. The affected product was also distributed to one government consignee. The affected product was also distributed to the following foreign countries: Canada, Albania, Algeria, Australia, Austria, Bangladesh, Belgium, Brazil, Cayman Islands, China, Czech Republic, Denmark, Ecuador, Ethiopia, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Libyan Arab Jamahiriya, Malaysia, Mauritius, Mexico, Myanmar, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Poland, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, United Kingdom, and Vietnam. | Terminated | October 01, 2015 | 297 | 500 | Not_Computer
Z-0987-2015 | 70078 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry XPT interfaced to the CentraLink Data Management System V14x in specific configurations: The CentraLink system software is a network solution provider and multi-system data manager for the instruments and lab automation systems (LAS) within the lab. The CentraLink software consolidates data from all connected instruments so that an operator can review and edit patient and quality control results from a single location. |  |  | December 16, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Enabling sending of preliminary/initial results on the ADVIA Chemistry XPT creates a risk that a critical result could be overwritten by the same result rather than showing the repeated result on CentraLink. A result could be erroneous and critical and appear to be verified upon repeat. | Under Investigation by firm | A Siemens service provider visited each site to verify and/or update customer configuration on the ADVIA Chemistry XPT and CentraLink Systems to ensure proper reporting of results. An Urgent Medical Device Correction letter was issued on December 16, 2014 to all affected customers. Customers were also provided a Customer Bulletin providing instructions on how configuration setting should remain. |  | 7 | 7 | Distributed in the state of WA. | Open, Classified |  |  | 7 | Not_Computer
Z-0981-2015 | 69767 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Drill set long, drill stop compatible for 4.8mm diameter implant, Article Number 040.447S Endosseous dental implant drill set. |  |  | October 27, 2014 | 2014 | Straumann USA, LLC | Drill set may contain incorrect drill. | Process control | Straumann sent a Field Safety Corrective Action by letter dated October 29, 2014 which was delivered via United Parcel Service. The letter identified the product the problem and the action needed to be taken by the customer.Customers are asked to examine their inventory and return the product. In addition, customers are asked to complete a Customer Confirmation Form and return it to Straumann via fax to (978)747-0023. |  | 362 | 362 pieces | Nationwide in the following states: AZ, CA, CO,CT, FL, IA, IL, IN, LA, MA, MD, ME, MN, MO, NC, NJ, NY, OH, OK, OR, PA, RI, TX, VA and WV. Worldwide in the following countries by Institut Straumann AG distribution network: Austria, Australia, Canada, Germany, Spain, Finland, France, United Kingdom, Netherlands, Belgium, Switzerland, Ireland, Italy. | Terminated | April 10, 2015 | 165 | 362 | Not_Computer
Z-0992-2015 | 69739 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Infinite 200PRO |  |  | November 10, 2014 | 2014 | Tecan US, Inc. | Incorrect lumi firmware version installed (E.027 instead of V2.00) | Employee error | Tecan sent an Urgent Field Corrective Action letter dated October2014, to all affected consignees by via FedEx on/or about November 10, 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use te kinetic luminescence measurement workflow requiring movement of th eplate transport out during kinectice measurements until the firmware has been upgraded to current version 2.0. A Tecan service representative will contact you to arrange a time to install the firmware. Additionally, they may want to review prior measurement results obtained with this workflow. Customers with questions were advised to contacte their local Tecan Helpdesk.For questions regarding this recall call 919-361-5200, ext 19524. |  |  |  | Worldwide Distribution - USA including MD and Internationally to Hong Kong, South Korea and the UK. | Terminated | December 21, 2015 | 406 |  | Not_Computer
Z-0980-2015 | 69956 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Linear Accelerator: ARTISTE, ONCOR Avant Garde, ONCOR Impression Plus, ONCOR Expression, ONCOR Impression, PRIUS HI, Mevatron M2/Primus Mid-Energy, MEVATRON PRIMART, MEVATRON KD-2, MEVATRON KDS-2, MEVATRON K2, MEVATRON 6323-2, MEVATRON M2 6740, MEVATRON M2 6700, MEVATRON 6300, MEVATRON MDX-2, MEVATRON MD-2, MEVATRON KD, MEVATRON KDS, MEVATRON 77, MEVATRON K7467, MEVATRON 7767, MEVATRON K8067, MEVATRON M 6300 (MX), MEVATRON M 6730, MEVATRON M 7140, MEVATRON M 7400, MEVATRON MDX, MEVATRON MD (MD, MDX), MEVATRON M 7445 (MXE), MEVATRON M 7400/7440 (MXE), MEVATRON M 6730/6740 (MXE), MEVATRON M 6700 (MX): the intended use of the MEVATRON , PRIMUS , ONCOR and ARTISTE MV System linear accelerator system is to deliver x-ray radiation for therapeutic treatment of cancer. |  |  | December 02, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Use of any unauthorized third-party components on Siemens LINACs may lead to patient or operator injury, damage to the LINAC, and can potentially compromise customer warranty. The dark current dose must be considered in the overall dose calculation, or an overdose and serious injury may occur. | Component design/selection | An urgent field safety notice, dated December 2, 2014, was sent to end users advising them to use only Siemens' authorized components for patient treatments as described in the operator's manual. |  | 320 | 320 | Worldwide Distribution - US Nationwide including Canada and Mexico. | Terminated | May 12, 2015 | 161 | 320 | Not_Computer
Z-0988-2015 | 70024 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Femto LDV Z-Generation Femtosecond Surgical Lasers are used to perform ophthalmic surgeries. |  |  | December 18, 2014 | 2014 | Ziemer Usa Inc | The devices were not shipped with the Starter Kits. | Labeling False and Misleading | Notification letters dated December 18, 2014 were sent to all direct account s(customers). The letters were included with the Starter Kit referred to in the Surgical Procedure Manual. |  | 26 | 26 lasers | US: Nationwide (AL, CA, FL, GA, IL, MN, MO, NC, ND, NH, NJ, NM, NY, PA, TX, VA, WA) | Open, Classified |  |  | 26 | Not_Computer
Z-0994-2015 | 70020 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Ysio Max, Luminos dRF Max and Agile Max systems with software version VE10 and SmartOrtho license. |  |  | December 15, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Potential for composed images to be flipped before being sent to PACS on systems with software version VE10 and SmartOrtho license. Flipped images may be reversed so the associated annotations, e.g. labels (R/L), may be displayed incorrectly, potentially leading to misdiagnosis. | Software design | The firm, Siemens, sent a "Safety Advisory Notice - XP080/14/S" letter dated December 8, 2014, to end users that identified the product, problem, and actions to be taken. Siemens provided instructions to avoid the issue and indicated a software fix would be released as a permanent resolution. The customers were instructed observe this safety notice and comply with the corresponding measure until the update has been fully completed.If you have any questions, contact the Regulatory Technical Specialist at 610-219-2119. |  | 11 | 11 | US Distribution to states of: ID, IL, MI, MN, MO, ND and PA. | Open, Classified |  |  | 11 | Not_Computer
Z-0996-2015 | 69671 | No | GBB | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Powered stretcher chair used during facial and eye surgeries. One chair per package. |  |  | November 14, 2014 | 2014 | TransMotion Medical Inc | The firm is finding that the stretcher chair controllers are moving the chairs involuntarily. | Component design/selection | TransMotion Medical Inc. sent an Urgent Medical Device Recall letter dated November 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed:Discontinue use of the device and contact TransMotion Medical Inc. for a replacement. Suspected faulty pendants must be destroyed by severing the cord and discarding the control. Replacement of the pendant is a long-term solution to the issue due to replacement pendants being not of the same type or design as the failed pendant.Accompanying the recall letter will be an acknowledgment and receipt form that is required to be completed by the customer and returned to TMM via fax or email. This form will confirm that the product resides at the customer site where the unit was originally shipped. If the unit is no longer present at the customer site the customer can indicate this on the form and provide further information as to the whereabouts of the device if known. Ship replacement pendant control to the customer free of charge with installation instructions as to how to replace the pendant on the device. The firm will track non-responders and initiate follow-up contact.For further questions please call Phone: 330-239-4192 ext. 175 or Toll free: 1-866-860-8447 (7:30 AM to 4:00 PM Eastern Time Monday through Friday) |  | 1127 | 1,127 units | Worldwide Distribution - US Distribution and the countries including Canada, Australia, Chile, China, Costa Rica, Germany, Algeria, Israel, Italy, South Korea, Malaysia, Nigeria, Netherlands, Poland, Portugal, Saudi Arabia, Singapore, Thailand, and Venezuela. | Terminated | July 11, 2015 | 239 | 1127 | Not_Computer
Z-0993-2015 | 69560 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | TEGRIS System manufactured by MAQUET GMBH in Germany |  |  | August 21, 2014 | 2014 | Maquet Medical Systems USA | The wrong movement results on the MAGNUS OR table. The button commands on the TEGRIS touchscreen for Lower Leg Up and Upper Leg Down are switched in the software for the integration with the MAGNUS operating table system. | Software design (manufacturing process) | Maquet sent an Urgent Medical Device Field Correction Letter dated August 21, 2014, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following measures until the issue is corrected: The TEGRIS equipped with a software module version up to 2.1.X.X should not be used in conjunction wiht a MAGNUS operating table system. All other MAQUET operating tables are not affected.A MAQUET service representative would contact them and arrange to have their TEGRIS software module updated. Customers with questions were instructed to contact MAQUET Technical Support at 1-888-627-8383 (press option 3, followed by option 1, then option 1 again.For questions regarding this recall call 973-709-7000. |  | 6 | 6 units | US Nationwide Distribution including NE, NY and PA. | Terminated | November 17, 2015 | 453 | 6 | Not_Computer
Z-0997-2015 | 69742 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | Philips IntelliVue TcG10; Monitor, carbon-dioxide, cutaneous. |  |  | November 19, 2014 | 2014 | Philips Medical Systems, Inc. | Philips has discovered that the Instructions for Use (IFU) for the IntelliVue TcG10 Transcutaneous Gas Measurement Module is missing the contraindication that the device is not to be used on patients under gas anesthesia. | Under Investigation by firm | Customers were notified of the recall by letter sent via UPS on November 19, 2014. The Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips asked customers to do the following: Upon receipt of this notification, ensure that the IntelliVue TcG10 addendum to the Instructions for Use is being reviewed and implemented by all members of your staff who operate the device. Review this information with all staff members who are involved in the operation of the IntelliVue TcG10 modules and need to be aware of the contents of this communication. The addendum should be stored with the IntelliVue TcG10 Instructions for Use. If you need any further information or support concerning this issue, please contact your local Philips representative or call us at 1-800-722-9377. |  | 86 | 86 | Distributed in the states of CA, HI, IL, MA, MI, MO, MT, NC, OH, and PA. | Terminated | September 29, 2015 | 314 | 86 | Not_Computer
Z-0995-2015 | 69978 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Ferno Model 35X PROFlexx Stretchers, one unit per package |  |  | December 03, 2014 | 2014 | Ferno-Washington Inc | The wheel castor assemblies may be loosening on the stretchers. | Mixed-up of materials/components | Ferno initiated contact with customers via email and telephone on Wednesday, December 3, 2014. The firm will replace all wheel caster assemblies. |  | 46 | 46 units | The affected product was distributed to the following states: OH, MA, HI, NJ, TN, TX, MO, and GA. The affected product was distributed to the following foreign countries: Canada, Germany and Venezuela. | Terminated | December 01, 2015 | 363 | 46 | Not_Computer
Z-0998-2015 | 70080 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis POWERFLEX PRO PTA Dilatation Catheter |  |  | December 22, 2014 | 2014 | Cordis Corporation | One unit from a demo lot of POWERFLEX PRO PTA Dilatation Catheter was left at a hospital and was not sterilized. | Labeling design | The firm, Cordis, sent an "URGENT MEDICAL DEVICE RECALL (Removal)" letter dated December 29, 2014 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately check their inventory to confirm they do not have any units from the demo lot in possession; identify and set aside any unit from the identified lot; review, complete, sign and return the enclosed Acknowledgement Form directly to Cordis at the fax number on the form; return any affected product to Cordis; share this letter with others in your facility; contact any other facility that has been provided with units of POWERFLEX PRO product; maintain awareness of this notice until all affected product has been returned to Cordis and keep a copy of this notice witn any affected product.For any physician with medical concerns, please contact the Cordis Endovascular Medical Information Line 877-DEVICES Monday through Friday from 9am-5pm. For questions related to the recall that are not adequately addressed in the letter, please contact the Cordis Field Action Coordinator Hal Baden at (786) 313-2365. For questions related to the Acknowledgement Form and its return, please contact Cordis QA at (786) 313-8730. |  | 120 | 120 | Nationwide Distribution: US including states of: CT, NJ, NY, PA, MD, VA, WV, NC, SC, GA, FL, AL, TN, MS, KY, OH, IN, MI, IA, WI, SD, IL, MO, KS, NE, LA, OK, TX, CO, UT, AZ, CA, OR, WA, AR, IN, MN, WY, ND, NV, MA, RI, NH, CT, DE, District of Columbia and Hawaii. | Terminated | July 09, 2015 | 199 | 120 | Not_Computer
Z-1003-2015 | 70136 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Fitness Gear Inversion table (Style STE00059FG or Style STE00118FG), Distributed by ASLI, Wilmington, DE (on carton), Paradigm Health &amp; Wellness, Inc. (in user manuals); The products are inversion tables, which allow the user to secure his or her ankles to the table and invert. The products were intended to be fitness equipment; i.e. consumer products. |  |  | December 29, 2014 | 2014 | Dick's Sporting Goods | Reports of customers falling from the Inversion Table and sustaining injuries. The Inversion Table is a fitness product, not a medical device. DSG removed medical claims from product labels and advertising. | Labeling False and Misleading | DSG sent emails and medical device correction notices, dated December 2014, to consumers who made purchases online and retail store consumers for whom DSG has contact information via its Scorecard customer loyalty program. DSG also posted the recall on their web site and notified their retail stores with instructions to post notices in DICKs stores. DSG plans to remind stores approximately monthly for three months to display the recall poster, beginning one month after the recall announcement. Consumers were asked to return the product to the stores where they were purchased for a full refund. A Press Release, dated January 8, 2015, was also issued about the recall. |  | 32581 | 32581 | Nationwide. | Terminated | September 02, 2015 | 247 | 32581 | Not_Computer
Z-1001-2015 | 70012 | No | LCO | Platelet factor 4 radioimmunoassay. | Hematology | Hematology | 510(k) | PF4 Concentrated Wash (10x), REF 403624, packaged in the following assay kits: 1. ImmucorGamma LIFECODES PF4 IgG (catalog numbers HAT13G, HAT45G) and |  |  | December 15, 2014 | 2014 | Genetic Testing Institute,inc | Some Lots of PF4 Concentrated Wash (10X) are a slightly more dilute concentration. This may lead to slightly increased Optical Density (OD) values. Samples with a higher background reactivity may demonstrate false positive result (borderline reactivity) at or near the cutoff for the assay. | Employee error | Immucor sent an Urgent - Field Safety Notice letter dated December 15, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised to stop using the recalled product and to destroy it. Requested consignees to complete the Customer Response From and return to waukeshatechsupport@immucor.com. Please contact Technical Support should you need additional information at 262.754.1000 or waukeshatechsupport@immucor.com. |  | 6206 | 6206 kits | Worldwide Distribution - Nationwide Distribution to the states of : AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN,TX, UT, VA, WA, WV and WI., and the countries of : Canada, Japan, Belgium, European Union , Australia, Taiwan, Thailand and Scotland. ( some countries not yet identified ) | Terminated | April 30, 2015 | 136 | 6206 | Not_Computer
Z-1002-2015 | 70106 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | Anti-Lambda APC-H7, Catalog Number 656157. |  |  | December 19, 2014 | 2014 | BD Biosciences, Systems &amp; Reagents | Two lots of the Anti-Lambda APC-H7 antibody are contaminated with CD38 antibody. | Vendor change control | BD sent a Recall notification letters dated December, 2014 to affected customers via Certified mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory for affected lots, immediately remove affected reagents from their laboratory, request replacement product and complete and return the enclosed Tracking/Verification form, per the instructions. For questions contact BD Customer Support at 855-236-2772. |  | 15 | 15 units | US Nationwide Distribution in the states of CA, NY, and MI | Terminated | March 11, 2016 | 448 | 15 | Not_Computer
Z-1004-2015 | 69727 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Software Version 3.1 utilized on the following systems: 1) 4600 Chemistry Systems (VITROS 5,1 FS System family member, 2) 5600 Integrated System |  |  | September 29, 2014 | 2014 | Ortho-Clinical Diagnostics | Calibration may not occur when using calibrator barcode labels supplied with VITROS Chemisty Products Calibrator Kit 2. | Software Manufacturing/Software Deployment | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated September 29, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed Calibration may not occur when using calibrator barcode labels supplied with VITROS Chemisty Products Calibrator Kit 2.The resolution to this software anomaly will be in a future version of software. Until the new software is installed on their VITROSÂż System, customers were instructed not to use the calibrator barcode labels for VITROSÂż Calibrator Kit 2. Customers were advised that you manually program the calibration.Customers with questions were instructed to contact Customer Technical Services at 1-800-421-3311. |  | 12 | 12 units | Worldwide Distribution - USA including CO, NC, NJ, NY, PA, MI, WV and Internationally to Canada, Japan and Spain. | Terminated | April 03, 2017 | 917 | 12 | Not_Computer
Z-1000-2015 | 70107 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | CD8 (Leu-2a) PE (In vitro diagnostic), Catalog number 340046. |  |  | December 19, 2014 | 2014 | BD Biosciences, Systems &amp; Reagents | One lot of CD8 PE is contaminated with CD4 FITC and may cause an unexpected staining pattern and incorrect patient data. | Process control | Recall notification letters were sent to all consignees on December 19, 2014 by certified mail. |  | 165 | 165 units | Worldwide distribution. US Nationwide including Puerto Rico, China, Brazil, Japan, Guatemala, India, New Zealand, Australia, Canada, Thailand, Chile and Malaysia. | Terminated | February 23, 2016 | 431 | 165 | Not_Computer
Z-1006-2015 | 70007 | No | BYI | Powered percussor. | Anesthesiology | Anesthesiology | 510(k) | FLUTTER Percussor respiratory device, packaged one unit per clear plastic package. |  |  | November 13, 2014 | 2014 | Aptalis Pharmatech Inc. | Customer notification that the device may be difficult to open or close. | Process control | Aptalis Pharmatech distributed a customer notification letter on November 13, 2014. The letter identified the affected product, problem and actions to be taken. Customers were instructed to forward complaints to Aptalis Pharmatech Customer Service department at 908-927-9600. |  | 49000 | Approximately 49,000 units | Worldwide Distribution - US Nationwide in the states of OH, NM, CO, NC, IL, MD, LA, MN, NJ, SC, AL, VA, MI, MA, and CA. in the countries of: Canada, Australia, Argentina, Guatemala, Serbia, Chile, and Finland. | Terminated | July 11, 2015 | 240 | 49000 | Not_Computer
Z-1007-2015 | 69938 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | The Arrow Edge Hemodialysis Catheterization Set Long-Term Access and NextStep Antegrade Hemodialysis Catheterization Set, Long-Term Access, part numbers: AC-15192-SFX, AC-15232-SFX, AC-15272-SFX, AC-15312-SFX, AC-15422-SFX, CS-15192-SFX, CS-15192-SFXM, CS-15232-SFX, CS-15242-I, CS-15272-SFX, CS-15282-I, CS-15312-SFX, CS-15322-I, CS-15362-I, CS-15422-SFX, CS-15502-SFX, CS-15552-I, CSD-15242-I, CSD-15282-I, and CSD-15322-I. |  |  | February 04, 2013 | 2013 | Arrow International Inc | The tunneling sheath packaged within the kits can crack during placement. If the tunneling sheath cracks, it can causea delay in treatment for the time it takes to retrieve a replacement kit. | Component design/selection | An Urgent Medical Device Recall Notification letter, dated February 4, 2014, was sent to direct consignees that identified the product, problem, and action to be taken. Customers were asked to discontinue use and quarantine affected inventory. The Recall Acknowledgement form should be completed and faxed back to 1-800-343-2935. A customer service representative will contact customers with a Return Goods Authorization (RGA) Number and provide instructions for the return of the product. Customers are asked to contact their local sales representative or Customer Service at 1-800-343-2935 with questions. |  | 2480 | 2,480 | Worldwide Distribution -- US and Canada. | Open, Classified |  |  | 2480 | Not_Computer
Z-1005-2015 | 70017 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Multi-Lumen PI CVC Kit and PSI Kit: 3-L 7 FR X 8 product code: ASK-45703-POSU lot #23F13L0166, exp. 31-Jul-2015: |  |  | December 16, 2014 | 2014 | Arrow International Inc | the Instructions for Use (IFU) were inadvertently removed from the kit. | Labeling False and Misleading | An Urgent Medical Device Corrective Action Notification letter, dated December 16, 2014, was sent to Owens &amp; Minor, 1605 Worldwide Blvd., Hebron, KY 41048-8653, instructing them to immediately discontinue distribution and quarantine any products, use customer letter and Corrective Action Acknowledgement Form templates to communicate this corrective action to any of thier customers that have received product included within the scope of the corrective action, and to have their customers return a completed Corrective Action Acknowledgement Form outlining their remaining inventory of the affected product codes and lot numbers for which replacement IFUs is needed. |  | 420 | 420 | In the state of KY | Terminated | August 10, 2015 | 237 | 420 | Not_Computer
Z-1009-2015 | 69533 | No | MQN | Bone plate. | Dental | Dental | 510(k) | The Synthes Mandible External Fixator; all lots of part nos.: |  |  | October 17, 2014 | 2014 | Synthes, Inc. | Under new testing protocols for MR Environment safety and compatibility, metal devices are no longer designated MR Safe. Some components in the CMF Mandible External Fixator Systems were labeled and/or etched with information indicating they are MR Safe. They can no longer be labeled MR Safe. | Labeling design | Urgent notices of medical device labeling updates, dated October 17, 2014 and October 24, 2014 were sent to end users and sales consultants, which described the product, problem, and action to be taken. These consignees were instructed to inspect inventory and remove out-dated labeling inserts and replace with updated labeling information. |  | 43159 | 43159 | Worldwide Distribution. | Open, Classified |  |  | 43159 | Not_Computer
Z-1011-2015 | 69792 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView model number: 882478 |  |  | November 12, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Unintended detector and gantry movement due to software issues. | Software design | Philips Healthcare sent an Urgent - Field Safety Notice dated November 4, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Philips is advising the customers/users to retain a copy of this Field SafetyBY CUSTOMER I USER Notice with the equipment Instruction for Use, until further notice.Problem #1:Philips is advising the customer/user to perform the following workaround prior tothe implementation of the field safety correction, to avoid contact between thedetector and the head holder as describe above in Problem #1:ONLY if the system is positioned in a Relative-90 configuration with one of thedetectors below the pallet level and the table in in the patient load position:Follow the sequence of operations prior to performing a study utilizing the headholder:- Deploy or store the x-ray panel if necessary;- Run the patient load PPM;- Install the head holder; and- Position the patient on the table and proceed with the studyIn addition, take the following preventative actions:" Monitor the patient during any and all system motions." Know locations of the multiple E-Stop button(s) to halt all system motion.(Emergency Stop: where the system immediately stops all detector andgantry motions by disconnecting power to the motors that control systemmotion) Customers with questions were instructed to contact the Customer Care Solutions Certer at 1-800-722-9377.For questions regarding this recall call 440-483-2119. |  | 1064 | 1064 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Algeria, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, China, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Gabon, Germany, Greece, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, United Kingdom, and Viet Nam. | Open, Classified |  |  | 1064 | Not_Computer
Z-1010-2015 | 70102 | No | OCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Product Code 2C4013, Four Lead TUR Irrigation Set, for use with UROMATIC |  |  | January 14, 2015 | 2015 | Baxter Healthcare Corp. | Potential for failure of the pouch packaging seal at highaltitudes. | Package design/selection | 1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.Baxter Healthcare Center for Service can be reached at 888-229-0001 betweenthe hours of 7:00 AM and 6:00 PM Central Time, Monday through Friday. Pleasehave your Baxter 8-digit ship-to account number, product code, lot number(s) and quantity of product to be returned ready when calling.3. Complete the enclosed Customer Reply Form and return it to Baxter by faxing it to 224-270-5457 or scanning and e-mailing to fca@baxter.com . Returning the Customer Reply Form promptly will prevent you from receiving repeat notices.4. If you distribute this product to other facilities or departments within yourinstitution, please forward a copy of this communication to them.5. If you are a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, please conduct a consumer-level recall of the affected product that you distributed to customers.If product was purchased from a distributor: 1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.The Baxter Healthcare Center for Service can be reached at 888-229-0001between the hours of 7:00 AM and 6:00 PM Central Time, Monday throughFriday.3. Follow your supplier's recall and response process. Please do not return thecustomer reply form to Baxter. |  | 26028 | 26,028 units | US Distribution. | Terminated | April 10, 2017 | 817 | 26028 | Not_Computer
Z-1013-2015 | 70200 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Syngo RT Therapist: The intended use of the SIEMENS branded ARTISTE, ONCOR and PRIMUS family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer |  |  | January 02, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Combination of CTVision with syngo RT Therapist / syngo RT Oncologist 4.3.SP1 automatic registration in Adaptive Targeting might result in wrong offset calculations. Applying this offset can lead to patient mistreatment. Cone Beam imaging is not affected by this problem. | Software Design Change | A customer advisory notice, dated January 2, 2015, was sent to end users which identified the product, problem, and action to be taken. It was recommended that customers use only manual registration with the 6 degrees of freedom option disabled, which provides the correct offset values. A software update will be deployed once a solution is available. |  | 3 | 3 | UT, WI, NY | Terminated | September 10, 2015 | 251 | 3 | Not_Computer
Z-1012-2015 | 70086 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Miethke Shunt System accessories |  |  | December 17, 2014 | 2014 | Aesculap, Inc. | Some Miethke Shunt System accessories may have an incorrect MR safety symbol on the shelf box label. The label should have NO symbol but instead is marked with the MR conditional symbol. The product is MR unsafe. This information is only required to be in the Instructions for Use (IFU). | Labeling False and Misleading | Aesculap sent an Important Correction and Removal and Field Safety Notification letter, dated December 17, 2014, to all affected customers. The letter identiifed the affected product, problem, and actions to be taken. Customers were requested to remove from inventory and contact Aesculap regarding return of the affected devices and acknowledgement form. |  | 772 | 772 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 772 | Not_Computer
Z-0955-2015 | 70133 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Revolution CT is a multi-slice (256 detector row) CT scanner consisting of a gantry, patient table, scanner desktop (operator console), system cabinet, power distribution unit (POU). and interconnecting cables.The system includes image acquisition hardware, image acquisition and reconstruction software, and associated accessories. |  |  | November 12, 2014 | 2014 | GE Healthcare | A required quality control test was not performed during installation associated with the software of the Revolution CT scanner. | Radiation Control for Health and Safety Act | GE Healthcare Planned Action (s) to repair defect or to bring product into Compliance letter dated January 16, 2015 to GE Healthcare.1. A GE Healthcare service representative will perform a noise specification and tolerance calibration to correct all affected systems via Field Action Instruction (FMI 25457).2. A customer notification letter which includes a statement that GE Healthcare will without charge, remedy the defect or bring the product into compliance.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. GE Healthcare Surgery will implement this CAP by May 1, 2015.You may proceed or continue with implementation of the CAP (subject to the conditions noted above).Call Center phone numbers: United States: 800 437-1171, Japan: 0120-055-919Information For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. |  | 13 | 13 units | US Distribution to the states of CA, FL, IL, NY, UT and WA. | Terminated | April 29, 2015 | 168 | 13 | Not_Computer
Z-1019-2015 | 70141 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Zippie Zone Pediatric Rigid Wheelchair Model EIRZ1. |  |  | December 19, 2014 | 2014 | Sunrise Medical (US) LLC | There is the potential for the height adjustable handle to dislodge due to failing locking mechanisms on specific identified chairs. | Device Design | Sunrise Medical sent an Urgent Medical Device Field Correction letters dated December 19, 2014 to the two affected distributors. The letter identified the affected product, problem and actions to be taken. For question contact Sunrise Medical Regulatory Affairs at either (559)294-2840. |  | 3 | 3 | Nationwide Distribution in the states of CO and ME. | Terminated | March 10, 2015 | 81 | 9 | Not_Computer
Z-1014-2015 | 70137 | No | OQY | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Ebb Complete Tamponade System, BD-OTS Obstetric Catheter, REF CTS-1000, Rx Only, Sterile EO |  |  | December 30, 2014 | 2014 | Clinical Innovations, LLC | Potential for a balloon leak in all lots of (CTS-1000) manufactured since August 2014. | Process design | Clinical Innovations sent an "Urgent - Medical Device Recall" letter dated December 31, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They will be instructed to quarantine remaining product. A company representative will visit the hospital and retrieve the product in accordance to instructions detailed on the Representative Recall Reconciliation Instructions. The representative will reconcile the inventory on the Field Reconciliation Form and return the product to Clinical Innovations as indicated on the instructions. The address for the devices to be returned is included on the Representative Recall Reconciliation InstructClinical Innovations is working diligently to resolve this issue. In the meantime ,please quarantine any remaining product in your hospital. A company representative will contact you and make arrangements for the return of the product. Clinical Innovations will provide replacement product as soon as it becomes available. If you have any questions,call Clinical Innovations,at 1-(888)268-6222 or your service representative, as is the name of the Company Contact. |  | 827 | 827 units | Worldwide Distribution - USA (nationwide) and the countries of : Canada, Finland, Estonia, Belgium, Ireland, Sweden, France, Switzerland, Spain, Netherlands, Middle East, Scotland, Germany and Wales. | Terminated | December 28, 2015 | 363 | 827 | Not_Computer
Z-1015-2015 | 70027 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ASSY, POWER SUPPLY 130V 50HZ. This component is used with the following machines: 2008T Hemodialysis Machine, 2008K Hemodialysis Delivery Equipment, 2008 K2 Hemodialysis Delivery Equipment |  |  | December 16, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Unqualified power assemblies shipped to customers. | Under Investigation by firm | Consignees were notified of the recall via letter sent on 12/18/2014 via certified mail. The letter informs consignees of the recall and requests that they take the following steps: Please check your stock for any 50 Hz Power Supplies (P/N 190092) with the name Electro Assemblies, LLC which is visible on the Transformer contained inside the Power Supply and discontinue use. Place the product in a secure, segregated area. Please contact your Fresenius Technical Services at 1-800-227-2572 option 4 then option 1, and reference the Field Action Number above. They will issue you a Return Good Authorization so you can return this product. that a reply form is completed and returned via fax or scanned email copy. In addition, consignees are asked to complete a reply form and return via fax/mail. |  | 10 | 10 | Distributed in OK and FL. | Open, Classified |  |  | 10 | Not_Computer
Z-1016-2015 | 69993 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Airo Mobile Intraoperative CT I Airo; Model #: MobiCT-32. Computed Tomography x-ray system. |  |  | December 08, 2014 | 2014 | Mobius Imaging, LLC | There is a risk that during the transfer of an image and navigation data to the Brainlab Curve Image Guided Surgery Navigation System after a CT scan, an error may occur, causing either no navigation data or incorrect navigation data being transferred to the curve. | Software design | Mobius sent an Urgent Field Safety Notice dated December 8, 2014, to all affected customers. The letter explains the recall and provides advise on action to be taken by the user. The letter also informs consignees that Mobius Imaging is working on a software fix for the software bug/communication error. The final fix is anticipated in early 2015.Advise on action to be taken by the user:The user should be prepared for the potential necessity to perform a manual navigation registration in order to use navigation.In general, do not use navigation data transmitted by the AIRO Mobile CT Scanner of any scan the Brainlab AIR App displayed the "Received Wrong Dataset" error for. To use such scans with the Curve, a manualnavigation registration is required.To avoid the possibility of incorrect data from being transferred to the Curve, Mobius Imaging is recommending that customers only select "NO" when they come to the "Continue with same patient" screen after each scan. They can select the same patient under the "Existing Patients" tab and restart the workflow. Additionally, always adequately verify the accuracy of the navigation registration data as required by the Brainlab Curve" Navigation system software and the AERO Operator's Manual. If you require further clarification, please feel free to contact your local Brain lab Customer SupportRepresentative.Customer Hotline: +49 89 99 15 68 44 or +1800 597 5911 (for US customers) or byE-mail: support@brainlab.com (for US customers: us.support@brainlab.com)Fax Brainlab AG: + 49 89 99 15 68 33Address: Brainlab AG (headquarters), Kapellenstrasse 12, 85622 Feldkirchen, Germany.For questions regarding this recall call 978-615-5025. |  | 315 | 21 devices distributed. 15 of these devices affected by the software v. 1.1.1+ patch | USA including AK, IL, WA, MA, NY, WV, NC, TX, CA, PA and Internationally to Germany, Netherlands, Belgium, Switzerland, Saudi Arabia, Egypt. | Terminated | July 06, 2015 | 210 | 315 | Not_Computer
Z-1026-2015 | 70030 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter, Lactate Assay, Catalog No. OSR6X93 |  |  | December 15, 2014 | 2014 | Beckman Coulter Inc. | N-Acetyl Cysteine (NAC), when administered in therapeutic concentrations for the treatment of acetaminophen overdose, interferes with the assays for these analytes: Cholesterol OSR6X16, Uric Acid OSR6X98, Lactate OSR6X93, and Lipase OSR6X30. NAC interference may lead to falsely low results. | Under Investigation by firm | A customer notification letter was sent to their customers on 12/15/14 to inform them that Beckman Coulter has recently confirmed that N-Acetyl Cysteine (NAC), when administered in therapeutic concentrations (for the treatment of acetaminophen overdose), has been determined to interfere with assays for the following analytes: Cholesterol OSR6X16, Uric Acid OSR6X98, Lactate OSR6X93, and Lipase OSR6X30. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to call (800) 223-0130 in the US and Canada. Customers outside the US and Canada are instructed to contact their local Beckman Coulter Representative. |  | 31799 | 31,799 units total (15,841 units in US) | Nationwide in US and Puerto Rico.World Wide: Albania, Algeria, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia-Herzegovina, Botswana, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyrus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Eritrea, Ethiopia, Finland, France, Germany, Ghana, Gibraltar, Greece, Guadeloupe, Guam, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Lebanon, Lesotho, Liberia, Libyan Arab Jamahiriya, Lithuania, Malawi, Malaysia, Mexico, Morocco, Mozambique, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, Zambia, and Zimbabwe | Terminated | February 19, 2016 | 431 | 31799 | Not_Computer
Z-1027-2015 | 70318 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza: Syngo.plaza is a Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images. It supports the physician in diagnosis and treatment planning. |  |  | January 13, 2015 | 2015 | Siemens Medical Solutions USA, Inc | In case of a system crash, images may not be written to the hard disk and this may result in inconsistencies in the database. In case of a system crash (e.g. blue screen, power outage) images may not be written from cache to the hard disk and might get lost. | Software Manufacturing/Software Deployment | Siemens sent a Safety Advisory Notice dated January 13, 2015, to all affected customers. The Safety Notice identified the product, the problem, and the action to be taken. Siemens recommended checking patient data in the event of a system crash and advised customers that a software update is planned for mid-2015 that will resolve the issue.For questions regarding this recall call 610-219-6300. |  | 50 | 50 | Nationwide Distribution | Open, Classified |  |  | 50 | Not_Computer
Z-1022-2015 | 70068 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | VC150 Vital Signs Monitor (VC150 monitor equipped with Masimo SpO2 technique only); Innokas Medical CARESCAPE VC150 Rx Only Innokas Medical Oy. Catalog numbers 2067980-002, 2067980-006, 2067980-010, and 2067980-014. |  |  | December 03, 2014 | 2014 | INNOKAS MEDICAL OY | If the SpO2 or RRa value is violating the respective alarm limit at the time of switching to monitoring mode, the monitor does not issue visual and audible alarms. | Software design | A notification letter was sent to the distributors on 12/3/2014. The letter requests the devices be returned to the manufacturer for repair. |  | 20 | 20 units | Distributed in the US and Europe. | Terminated | April 24, 2015 | 142 | 20 | Not_Computer
Z-1021-2015 | 70019 | No | HXW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 2.0 Mandible Locking Plate (MLP) Bender / Cutter (part number 329.143) |  |  | December 15, 2014 | 2014 | Synthes, Inc. | Certain lots have the potential for tip breakage while bending 2.0 mm thick plates. If device is used and breaks while in the Operating Room, surgical delay may occur while alternate bending or cutting instruments are located. | Finished device change control | An Urgent Notice: Medical Device Recall, dated December 19, 2014, was sent to end users and sales consultants that described the product, problem, and action to be taken. End users/sales staff were requested to immediately stop use and remove affected product from inventory. Customers were asked to call and obtain a Return Authorization Number, and complete and return the Verification Section of the letter along with the devices to Synthes per the instructions provided. Questions should be directed to 610-719-5450 or your local DePuy Synthes sales consultant. |  | 973 | 973 | Worldwide Distribution -- US and Canada. | Open, Classified |  |  | 973 | Not_Computer
Z-1032-2015 | 70050 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | HEPFLUSH-10 (HEPARIN LOCK FLUSH SOLUTION, USP) 100 USP units per 10 mL (10 USP units per mL), Single Dose Vial, Rx only -- Fresenius Kabi USA, LLC -- NDC 63323-017-10, Product Code 1710 |  |  | December 19, 2014 | 2014 | Fresenius Kabi USA, LLC | The lot failed pH specification. | Component design/selection | Fresenius sent an URGENT DRUG RECALL Letter dated December 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Examine their stock immediately to determine if they have any product from the affected lot. If the consignee is a distributor, immediately notify your customers that have been shipped or may have been shipped this product/lot of this recall and direct them to discontinue distributing or dispensing the affected lot. Please have them prepare to return the product to Fresenius Kabi as instructed in the Urgent Product Recall Response Form. Your customers may retrieve the recall letter and response form at http://www.fresenius-kabi.us /products/pharmaceutical-products/product-updates.html.2. If you have the affected lot available, immediately discontinue distributing or dispensing the lot, and return all units to Fresenius Kabi ,via FedEx Ground, using the enclosed return goods label and packing slip. A FedEx Ground labelcan be obtained by checking the box and noting your mailing address on the enclosed Urgent Product Recall Response Form. It will be mailed to the consignee upon receiving the request. A credit memo will be issued covering the quantity of the return to Fresenius Kabi.3. PLEASE COMPLETE THE ENCLOSED "URGENT PRODUCT RECALL RESPONSE FORM" AND SEND IT BACK TO US IMMEDIATELY VIA EMAIL OR FAX AT 1-708-649-8630. Use the following contact phone numbers as appropriate. Hours of operation: Monday through Friday 8:00 am to 5:00 pm CST: QA Department: (866) 716-2459 for information on how to return product; Vigilance &amp; Medical Affairs: 1(800) 551-7176 for clinical/technical information/Adverse Events (ADE's) reporting. |  | 69725 | Domestic: 69,725 | Nationwide Distribution | Terminated | January 26, 2016 | 403 | 69725 | Not_Computer
Z-1028-2015 | 70034 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Portex Anaerobic Pulsator Arterial Blood Sampling Kit, with Liquid Sodium Heparin for Gases and Electrolytes. |  |  | December 17, 2014 | 2014 | Smiths Medical ASD, Inc. | The outer box label is marked with the correct expiration date of 08/2016. The inner Tyvek Unit Label is marked with 08-2017 which exceeds the actual expiration date by one year. | Employee error | Smiths Medical notified American Medical Depot via telephone on 12 December 2014. Smiths Medical instructed them to put product at their location on hold. American Medical Depots customer will be instructed to return their unused product for replacement. |  | 900 | 900 units | US in the state of FL. | Terminated | June 25, 2015 | 190 | 900 | Not_Computer
Z-1031-2015 | 70232 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HomeChoice/HomeChoice Pro Automated Personal Cycler, for Automated Peritoneal Dialysis. Product codes 5C8310, 5C8310R, R5C8320, and 5C4474. |  |  | October 21, 2014 | 2014 | Baxter Healthcare Corp. | The keypad buttons on HomeChoice devices may be activated without the operator pressing them. | Nonconforming Material/Component | An Urgent Device Correction notification was sent to consignees on October 31, 2014. Within the notification, Baxter Healthcare asked customers to do the following: 1) Contact Baxter Technical Services to arrange for the exchange (swap) of any affected HomeChoice PRO Automated PD System devices. 2) Complete the attached customer reply form and return it to Baxter by either fax or scanned e-mail. For general questions regarding this communication, contact The Center for One Baxter at 900-422-9837, Monday through Friday between the hours of 8:00 AM and 5:00 PM Central Time. |  | 200 | 200 | AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, MI, MO, NC, NE, NJ, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, VT, WA, WI and District of Columbia | Terminated | September 01, 2015 | 315 | 200 | Not_Computer
Z-1029-2015 | 69963 | No | ONN | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Verasense Knee System, Model SYK-TRCR 02. Intelligent disposable tibial insert that measures dynamic loads during orthopedic knee surgery. |  |  | December 05, 2014 | 2014 | Orthosensor, Inc. | Presence of one inoperable (dead) sensor pad. | Under Investigation by firm | OrthoSensor sent an Recall letter dated December 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Required Action for Users" Remove the above-referenced serial numbers from your VERASENSE inventory and quarantine." Complete the enclosed Customer Response Form.Actions to be taken by OrthoSensor" You will be credited for these devices for your full purchase amount. The Returned Materials Authorization (RMA) number for each device in your possession is listed above." OrthoSensor will replenish these devices as soon as they are available.For further questions please call (954) 577-7770. |  | 22 | 22 devices | US Distribution to the states of : IL, MI, NY, CA, VA and NV. | Terminated | March 13, 2015 | 98 | 22 | Not_Computer
Z-1030-2015 | 70313 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow Select Kits (ASK), Maximal Barrier Precautions Kits (CDC) and Arterial Catheterization Kit (NA) with Biopatch Dressing |  |  | January 09, 2015 | 2015 | Arrow International Inc | Certain product codes and lot numbers of Kits with Biopatch Dressing were packaged with the incorrect size Biopatch Dressing. No risk to patients is anticipated. | Packaging process control | Arrow International, Inc. (Arrow) sent notification letters, dated January 9, 2015, to alert their direct account customers that it may be necessary to obtain an alternate Biopatch Dressing if a kit is used that contains an incorrectly sized Biopatch Dressing. Due to the inconvenience, Arrow will reimburse customers $7.50 per kit received. |  | 4484 | 4484 | US (nationwide) and PR. | Open, Classified |  |  | 4484 | Not_Computer
Z-1034-2015 | 69299 | No | OYC |  |  | Clinical Chemistry | PMA | Animas Vibe Insulin Infusion Pump. |  |  | September 06, 2011 | 2011 | Animas Corporation | Calibration factors in the pump overwritten during a programming step. The force sensor could send a lower signal value to the pump processor, with loss of prime warnings, occlusion alarms and the pump unable to detect a cartridge during the prime sequence. Field action initiated 8/29/2011. | Software design (manufacturing process) | The field action was initiated on 8/29/2011 and comprised communication to patients and distributors using email, letters and verbal communication. In addition, notifications were sent to those Health Care Professionals who had patients that were affected by this field action. Health Authorities were notified in those countries where the pump had been distributed (France, Germany, Sweden and United Kingdom) or was in the hands of patients. Replacement product was provided to distributors and end users who were identified as having affected product. |  | 1235 | 1235 | No US distribution, Distributors are located in France, Germany, Sweden and United Kingdom. | Terminated | August 10, 2015 | 1434 | 1235 | Not_Computer
Z-1033-2015 | 70008 | No | MGX |  |  | Toxicology | PMA | DrugCheck Ketamine Dip Test |  |  | November 21, 2014 | 2014 | Express Diagnostics Int'l., Inc. | This recall has been initiated due to potential false positives results when using the test to screen for Ketamine. Use of this product may give incorrect screening results. | Nonconforming Material/Component | Consignees were sent an Express Diagnostic Int. Inc., "Urgent" Letter dated 11//21/2014. The letter described the problem and the product involved in the recall. It advised consignees to quarantine the product. If the product has been further distributed, consignees were advised to notify their customers at once, Consignees were requested to complete and return the "Acknowledgement Form" either via fax to 1-507-526-2252 or e-mail Quality@drugcheck.com. For questions consignees can contact Jackie Gale, Director of Quality Systems and Regulatory Affairs at 1-507-526-3951. |  | 5050 | 5,050 | Worldwide Distribution -- US, including the states of CA, NY, MN, and SC; SWEDEN, CANADA, UNITED KINGDOM, and GUATEMALA. | Open, Classified |  |  | 5050 | Not_Computer
Z-0999-2015 | 70058 | No | OWD |  |  | Molecular Genetics | PMA | cobas 4800 KRAS AMP/DET 24T CE-IVD Mutation Test; CE-IVD 5852170190. |  |  | December 18, 2014 | 2014 | Roche Molecular Systems, Inc. | cobas KRAS Mutation Kit T10786 is generating invalid results for the Mutant Control and/or KRAS Calibrator. | Device Design | RMS issued an Urgent Field Safety Notice on 12/18/2014.A Quality Notification was provided to all users of the of cobasÂż KRAS Mutation Kit CE-IVD, lots T13091, T15435, T15475, T14777, and W00389, informing them of the issue and instructing them to discontinue use and discard any remaining inventory immediately. |  | 660 | 660 pieces | Foreign Distributed. | Open, Classified |  |  | 660 | Not_Computer
Z-1038-2015 | 70021 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab with SA HIS versions 4.0.7.0-4.0.7.1 |  |  | May 20, 2014 | 2014 | SCC Soft Computer | The interface fails to send abnormal flags for Reference Lab test results. | Device Design | The firm sent out 2 CORRECTION COMMUNICATIONS informing consignees of functionality in release of 4.0.7 HIS Add-on interfaces which SCC has designated to be a potential patient Risk-to-Health issue. Further the corrective action indicated HIS releases 4.0.7.2 and higher will be corrected. Hot Fixes have been developed for HIS Add-on patches 4.0.7.0 and 4.0.7.1. Lastly the communication requests that the user "please acknowledge receipt of this task and grant permission to load the mandatory hot fix(es). Thank you." |  | 48 | 48 | Distributed US (nationwide) in the states of OH, GA, NJ, WI, FL, NH, IL, NY, IN, CA, AL, VA, MO, MI, KY, MA, PA, CO, ME, TN, WA, MN, SD, LA, NC, and NV and the country of Canada. | Open, Classified |  |  | 48 | Not_Computer
Z-1039-2015 | 70062 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter Integrated APD Set with Cassette 3-Prong, Sterile, nonpyrogenic fluid path. |  |  | December 08, 2014 | 2014 | Baxter Healthcare Corp. | Baxter is issuing a recall for one lot of Integrated Automated Peritoneal Dialysis set with Cassette 3 prong due to complaints received for leakage. | Device Design | Baxter sent an Urgent Product Recall notification letter dated January 8, 2015 . The letter identified the affected product, problem and action to be taken. The recalling firm has asked the consignees to 1) locate and remove all affected product from their facilities, 2) contact Baxter to arrange for return and credit, and 3) complete the attached customer reply form and return it to Baxter. For questions call 800-422-9637. |  | 8640 | 8,640 units | US Nationwide Distribution in the states of FL, WI, AR, TN, TX, NC, MO, NM, NJ, CT, PA, NV, NY, CA, MI, WI, KS, and OH. | Terminated | October 28, 2015 | 324 | 8640 | Not_Computer
Z-1036-2015 | 69450 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | G- Series Beds; G29 full length bed rails assembly |  |  | December 05, 2014 | 2014 | Invacare Corporation | Invacare identified that the rail mounting bracket on the G29 full length bed rail was out of specification for entrapment safety. | Nonconforming Material/Component | Invacare sent an Urgent Medical Device Recall letter dated December 5, 2014, to all affected customers. The letter identified the product and the problem. Customers with questions should call 440-329-7996.For questions regarding this recall call 440-329-6875. |  | 2499 | 2499 units | US Nationwide Distribution including OH, CO, MO, and NY. | Terminated | June 18, 2015 | 195 | 2499 | Not_Computer
Z-1041-2015 | 70023 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cobas b 221 (versions 5 and 6) Blood Gas and Electrolytes Analyzer |  |  | December 18, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Chemical component of the F3 Fluid pack and environmental bacterial contamination causes a positive bias up to 39% at 4.1 mmol/l for the glucose measurement and 30% at 4.12 mmol/l for lactate. Urea readings are unaffected. | Material/Component Contamination | Roche Diagnostics has confirmed an issue with the calibration and QC stability of MSS parameters for cobas b 221 &lt;6&gt; systems. Internal investigations identified the formation of white deposits in the fluidic system, especially in the SD cartridge tubing, as the cause for the premature failure of the MSS parameters. These white deposits can potentially affect the fluidic pathways of both the Standby solution and the calibration solutions Cal1 to Cal4 (from the S3 Fluid Pack) and may reduce the lifetime of metabolite sensor cartridges.The cause of the white deposit has been identified as a chemical component of the S3 Fluid Pack solutions, as well as environmental bacterial contamination. Since this issue involves the S3 Fluid Pack, only the cobas b 221 &lt;6&gt; system is impacted.This issue may have an effect on the calibration of Glucose and Lactate parameters, which may lead to patient results being affected, as calibration drifts may not be detected by calibration and QC procedures. Glucose and Lactate are the only parameters affected by this issue.As a result, Roche is implementing an interim corrective action to be performed on your systems on a regular basis by Roche Field Engineering Specialist (FESs). Read the Urgent Medical Device Correction (UMDC) in its entirety." Continue to run and maintain your cobas b 221 &lt;6&gt; system as per the Instructions for Use." Complete the attached fax form and fax it to 1-888-656-6385." File this UMDC for future reference. Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-526-2272 if you have questions about the information contained in this UMDC. Public Contact is Todd Siesky, at todd.siesky@roche.com. |  | 103 | 103 | US Nationwide distribution. | Open, Classified |  |  | 103 | Not_Computer
Z-1040-2015 | 70237 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | MobileDiagnost wDR motorized portable diagnostic X-ray systems. |  |  | August 06, 2014 | 2014 | Sedecal S.A. | Reports of unexpected movement. | Other | URGENT - Field Safety Notice letters dated August 6, 2014 were sent to all consignees of record. The letters instructed consignees: 1) The unit is equipped with a safety switch in the drive handle. If the drive handle is released the unit will stop; 2) Do not try to correct unintended movement using the controls on the drive handle; 3) Instead, stop the system immediately by releasing the drive handle entirely; 4) Alternatively, the emergency button can be pressed; 5) Take special care when the system crosses metal strips or plates on the floor or enters an elevator, as these are situations where electrostatic discharge is likely to occur; and, 6) Should the customer feel uncertain regarding this action, he/she can contact Philips Healthcare. |  | 283 | 283 systems | US: Nationwide (AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NM, NY, OH, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI) | Open, Classified |  |  | 283 | Not_Computer
Z-1037-2015 | 70184 | No | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | 16cm Minimally Invasive Attachment. For use with Motor Systems for cutting and shaping bones, including the spine and cranium. Model MIA16-G1 |  |  | December 10, 2014 | 2014 | The Anspach Effort, Inc. | Customer complaint indicating the color bands on the MIA16-G1 were not correct. | Environmental control | Anspach sent an "URGENT MEDICAL DEVICE RECALL" (Removal) letter dated December 10, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm requested the customer to 1) Immediately identify and set aside all product listed below 2) Review, complete, sign and return the attached reply form 3) Share this letter with others in your facility that need to be made aware of this recall 4) If you distribute any of the affected products to other services or facilities, please forward this letter as appropriate 5) Maintain awareness of this notice 7) Keep a copy of this notice. If you have any questions regarding this removal, please contact the Complaint Handling Unit Manager at 561-494-3673 or contact your Anspach Sales Representative. |  | 8 | 8 | US Distribution to the states of: NC and CA., and Internationally to Switzerland. | Open, Classified |  |  | 8 | Not_Computer
Z-1047-2015 | 70288 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Crit Line in a Clip (CLiC) with SW version 2.51 Model Number: CL10041001. A continuous real-time monitor for non-invasive hematocrit, oxygen saturation and percent change in blood volume calculation during hemodialysis treatment. |  |  | December 19, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Potential for misinterpretation of the graphic display of the Blood Volume ( BV) slope | Device Design | FMCRTG, LLC representatives contacted Clinics on 12/19/14 by visit/telephone phone) and removed all CLiC units. |  | 35 | 35 units | CT, NY | Completed |  |  | 35 | Not_Computer
Z-1042-2015 | 70105 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer PERSONA Tibial Articular Surface Inserter; Item # 42-5299-001-00. Manual surgical instrument for knee arthroplasty (orthopedic). |  |  | November 25, 2014 | 2014 | Zimmer, Inc. | Potential for fracture of the tip of the product. The fractured tip is approximately 4mm x 7mm x 3.5mm in size. Retained tip may result in: pain, soft tissue irritation, soft tissue damage, or increased wear due to contact with the fractured tip; Revision surgery to remove the fractured tip | Process control | On 11/25/2014, URGENT MEDICAL DEVICE RECALL LOT SPECIFIC notifications were sent to the affected distributors, hospital Risk Managers and Surgeons with instructions for returning the affected product. All distributors were notified via electronic mail, and all hospital Risk Managers, Surgeons, and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call the customer call center at 1-800-348-2759. |  | 1091 | 1,091 | Worldwide distribution: US nationwide, including Puerto Rico, South Korea, Singapore, Dominican Republic, Canada, Japan, and Germany. | Terminated | December 01, 2015 | 371 | 1091 | Not_Computer
Z-1018-2015 | 70048 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Covidien Trellis-8 Peripheral Infusion System. Models CVT808015, CVT808025, CVT812015, and CVT812025. The following models are not offered for sale in the US: EVT808015, EVT808025, EVT812015, and EVT812025 Sterile EO. |  |  | December 10, 2014 | 2014 | Covidien | A manufacturing error resulted in the risk of incorrect proximal and distal balloon inflation port identification on the units. Units have been identified to have the distal balloon inflation port incorrectly labeled as proximal, and, the proximal balloon port incorrectly labeled as distal. | Mixed-up of materials/components | Consignees were sent an Urgent Product Recall letter dated 12/15/2014. The letter described the issue, identified the affected product, and the required actions. Affected product was to be returned directly to customers' Covidien Sales Rep; or customers were to contact Covidien Service at 1-800-716-6700 to arrange for product return. Questions can be directed to the Covidien Sales Rep or to Covidien Service at 1-800-716-6700. Customers were to complete and return the Verification Form along with the unused product as soon as possible. |  | 1032 | 1032 devices (919 US, 113 OUS) | Worldwide Distribution - US Nationwide and the countries Australia, Canada, Finland, France, Germany, Ireland, Italy, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 31, 2016 | 477 | 1248 | Not_Computer
Z-1043-2015 | 70100 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident PSL HA Solid Back 50 mm |  |  | December 15, 2014 | 2014 | Stryker Howmedica Osteonics Corp. | Five customer complaints were received for the same batch of Trident PSL Shells (part number 540-11-50E lot 49073501) alleging that when the packages were opened there were no dome holes plugs present. | Process control | Stryker sent an Urgent Medical Device letter dated January 7, 2015, with a Recall Notification/Acknowledgement Form via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return the attached Product Recall Acknowledgement Form to so_m_product_field_action_response@stryker.com or fax to 1-855-251-3635.Customers should return the affected product to:Attention: Regulatory ComplianceStryker Orthopaedics325 Corporate DriveMahwah, New Jersey 07430Customers with questions should call 201-831-6365.For questions regarding this recall call 201-831-5000. |  | 7 | seven units | Nationwide Distribution including NY, OH, PA, and TX. | Open, Classified |  |  | 7 | Not_Computer
Z-1048-2015 | 70181 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Si Surgical System, IS3000; - Wall chart for EndoWrist Stapler 45 System. |  |  | January 15, 2015 | 2015 | Intuitive Surgical, Inc. | The Wall chart has been updated because it was noted that Wall Chart (PN 551524-03 Rev A) was inconsistent with the User Manual Adendum (PN 550986-07_B) with regards to neceassary sterilization dry time. | Error in labeling | Medical Device Recall Replacement letters were sent to all customers on January 15, 2015. The letter identified the reason for the recall, the risk to health, affected products, and actions to be taken by the customer or user and the firm. Customers were asked to discard the existing EndoWrist Stapler Documentation and replace it with the new Documentation provided by the firm. Intuitive Surgical will ship replacement instruments at no charge to customers. Then customers were asked to reunite the replacement instruments with the additional components. The firm will retrain customers on the updated instructions for use. A Clinical Sales Representative or Intuitive Surgical Customer Service at 1-800-876-1310, should be contacted if further information or support is needed concerning the recall notification. |  | 1034 | 1,034 | Nationwide Distribution | Terminated | March 03, 2015 | 47 | 1034 | Computer
Z-1050-2015 | 68672 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Infusion System with Hospira MedNet List 20792-04-25, Pole Clamp Assembly Part Number 820-95443-008. A cassette based multi-function infusion system. Each system includes a pumping module and an assortment of disposable IV sets, optional accessories, and an operator's manual. The pole clamp assembly is used to provide a mechanism to tighten and loosen the grip of the pole clamp to remove the infuser from the IV pole. |  |  | June 11, 2014 | 2014 | Hospira Inc. | Plum A+ infusion system pole clamp assemblies were discovered to be cracked. If the knob on the pole clamp assembly is cracked, the infuser cannot be attached to the IV pole since the knob is used to tighten and loosen the grip of the pole clamp to the pole. | Vendor change control | Hospira contacted the the sole customer on June 11, 2014, to inform them of the issue of the potential impact to their Plum A+ infusers list 20792-04-25. Hospira requested that the customer to allow Hospira Field Service Representatives to inspect and replace the impact pole clamp assemblies.On June 12, 2013 Geneva Hospital approved Hospira's request to the inspection and replacement of the impacted pole clamp assemblies. On June 16, 2013 Hospira's Field Service personnel performed the on-site, visual inspection of the pole clamp assemblies and confirmed 15 units were cracked. Hospira determined to replace all 102 pole clamp assemblies that were received by Geneva Hospital. For further questions please call (800) 441- 4100. |  | 102 | 102 infusors | US Distribution to New York. | Completed |  |  | 102 | Not_Computer
Z-1054-2015 | 70185 | No | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | Lubricated Intubating Stylet 6FR that is inserted into an endotracheal tube to facilitate intubation. This product is used in surgical or medical emergency facilities. Primarily used on infants and pediatric patients. |  |  | December 23, 2014 | 2014 | Cardinal Health | A small piece of the stylet sheath has the potential toshear off upon removal of the stylet from the tracheal tube. | Process control | URGENT: PRODUCT RECALL letters dated December 23, 2014 were sent to all customers of record. The letters included instructions for customers to: 1) quarantine the affected product; 2) notify any customers whom you may have distributed the affected product to; 3) contact the appropriate Customer Service group to arrange for return and credit of any affected product in your inventory; and, 4) complete and return the enclosed Acknowledgement Form. Customers with questions about this recall can contact Cardinal Health Quality Systems at 800-292-9332. |  | 104432 | 104,432 stylets | US: Nationwide (AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, WV, WY) + District of Columbia + Guam + US Army Facility in South Korea | Open, Classified |  |  | 104432 | Not_Computer
Z-1049-2015 | 70365 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | FORUM Archive and Viewer, version 3.1, v 3.1.1, (DVD Format) and v 3.2, v 3.2.1.(DVD and USB Flash Drive Format). |  |  | January 22, 2015 | 2015 | Carl Zeiss Meditec AG | Software defect in the FORUM Viewer versions 3.1 and 3.2 which may lead to misinterpretation of the optical coherence tomography (OCT) data. | Software in the Use Environment | Important Field Corrective Action for Forum Viewer letters were sent to all affected US customers on January 22, 2015 by certified mail. |  | 985 | 985 | Nationwide Distribution. | Terminated | March 04, 2016 | 407 | 985 | Not_Computer
Z-1053-2015 | 70260 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set, Model No. 10015862 |  |  | January 13, 2015 | 2015 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion set due to disconnection and leakage issues. | Under Investigation by firm | CareFusion sent an Urgent Medical Device Recall Notification letter dated January 13, 2015, to all affected customers. The letter informs the customers of the problems identified and the actions to be taken. Customers were instructed to complete and return to CareFusion the recall response card. Customers with questions were instructed to CareFusion: Recall related questions, (888) 562-6018; Adverse Event Reports, (888) 812-3266; Technical Questions regarding the Alaris System, (888) 812-3229. |  | 7440 | 7,440 units | Nationwide Distribution | Terminated | March 15, 2016 | 427 | 7440 | Not_Computer
Z-1052-2015 | 70240 | No | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona (TASP) Tibial Articular Surface Provisional Shim, all sizes and Lots Catalog Number: 42-5279-002-00 thru 42-5279-009-04. |  |  | December 11, 2014 | 2014 | Zimmer, Inc. | Potential for delay in surgery and/or ball bearings to be left in the wound. Repeated cycles of sonic cleaning can break down the ball bearing retention feature (swage) on the shim. To prevent this, TASP shims must be manually cleaned and inspected per product insert, 87-6203-991-22. | Device Design | Zimmer sent an URGENT MEDICAL DEVICE CORRECTION notifications letter dated December 11, 2014 to affected distributors via electronic mail, and all hospital Risk Managers and distributors with affected inventory via courier. The letter identified the affected product, problem, instructions for responding to the formal recall notification.and actions to be taken. For questions call the customer call center at 1-800-348-2759. |  | 67963 | 67963 | Worldwide Distribution - USA Nationwide including DC, and PR and the countries of Australia, Austria, Belgium, Canada, Czech Republic, Dominican Republic, France, Germany, India, Iran, Israel, Italy, Japan, Luxembourg, Netherlands, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | October 23, 2015 | 316 | 67963 | Not_Computer
Z-1055-2015 | 70079 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 10 mL Syringe BD Luer-Lok Tip with Blunt Plastic Cannula, Not for Skin Injection ; Catalog #303348, |  |  | December 19, 2014 | 2014 | Becton Dickinson &amp; Company | The stopper of the 10mL BD Luer-Lok syringe is not properly seated in the syringe. | Device Design | BD sent an Urgent Product Advisory Notice letter dated 12/19/2014. The letter identified the affected product, problem and actions to be taken. Customers are advised of the issue with the product and should notify their customers. The attached Product Advisory Response Form should be completed and faxed to 1-866-551-2707. Questions can be directed to 1-866-367-4703. |  | 564000 | 564,000 units | US Nationwide Distribution | Terminated | December 07, 2015 | 353 | 564000 | Not_Computer
Z-1057-2015 | 70414 | No | LYQ | Nonpowered goniometer. | Orthopedic | Orthopedic | 510(k) | KYPHON EXPRESS, Osteo Introducer System, Diamond and Bevel, Size 2, REF T34A, QTY: 1EA, Rx only, STERILE R. For use with spinal fixation devices. |  |  | January 13, 2015 | 2015 | Medtronic Sofamor Danek USA Inc | The product, labeled as containing, one bevel-tipped instrument and one diamond-tipped instrument, may contain either two bevel-tipped instruments or two diamond-tipped instruments. | Process control | On January 12, 2015, the Sales Managers and Sales Force were notified by Post-Market Quality regarding the imminent notification and follow-up activities. The recall notices were delivered to the Risk Managers at impacted consignee sites on January 13, 2015. These recall notices included a Risk Manager Letter informing them on the action, and a consignee-specific questionnaire to indicate their product on-hand. |  | 239 | 239 units | Nationwide and Canada | Terminated | July 13, 2015 | 181 | 239 | Not_Computer
Z-1056-2015 | 70362 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound System, Model 10433816, component 10040596 - UILK2. |  |  | January 12, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | The bolt holding the control panel of the ACUSON SC2000 in a fixed position may fail and the Control panel cannot be locked into position. | Component design/selection | Siemens Medical sent an Customer Safety Advisory Notification to all affected customers on January 12, 2015 , via Certified mail, return receipt requested.The letter identified the product the problem and the action needed to be taken by the customer. At this time there is not mitigation to avoid risk associated with this issue. Local Customer Service Engineers will be contacting each site to replace the console. Until the necessary repairs have been completed, please share the information with all uses of the device.If you have any questions, please contact your local service support person for information regarding timelines and status.Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization who need to be aware of this issue.As always, patient safety issues are a very high priority. To date, no patient injury has been reported. This problem was discovered as part of our ongoing quality process.We sincerely regret any inconvenience this condition may cause in your daily operations. For further questions please call (650) 694-5398. |  | 230 | 230 | Worldwide Distribution - USA and the countries of Australia, Austria, Belarus, Canada, Ecuador, Finland, France, Georgia, Germany, India, Iraq, Ireland, Israel, Italy, Japan, Lithuania, Mexico, Nigeria, P.R. China, Philippines, Portugal, Republic Korea, Russian Fed., Saudi Arabia, Serbia, South Africa, Spain, Suriname, Sweden, Taiwan, Turkey, U.A.E. and United Kingdom. | Terminated | November 12, 2015 | 304 | 230 | Not_Computer
Z-1059-2015 | 70138 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | LINAC Accelerators: ARTISTE, ONCOR Avant Garde, ONCOR Impression Plus, ONCOR Expression, ONCOR Impression, PRIMUS HI, |  |  | December 19, 2014 | 2014 | Siemens Medical Solutions USA, Inc | A hardware update to replace the old PCB Assembly, Head Driver-G41 with PCB with overcurrent protection as a result of reports of the electric board inside the gantry showing signs of overheating, including thermal damage and smoke emission. | Component design/selection | Siemens sent a Customer Advisory Notice Letter TH016/14/S, dated December 19, 2014, to end users that identified the product, problem, and actions to be taken. End users were advised to read the Field Safety Notice and include it with the User's Manual until the hardware update is completed. |  | 310 | 310 | Worldwide Distribution - US Nationwide including Canada and Mexico. | Open, Classified |  |  | 310 | Not_Computer
Z-1058-2015 | 70239 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 980 Ventilator System, Model No. PB980 Ventilator (980xxxxxxxx), |  |  | January 12, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Covidien is issuing a voluntary field action for all Puritan Bennett 980 ventilators due to occasional GUI transient resets that last approximately 30 seconds. | Software design | Covidien sent an Urgent Field Safety Notice on January 12, 2015, to all affected customers. Customers were requested to forward the recall letter to anyone they may have transferred or assigned for use one of the potentially affected ventilators. Covidien is updating software that will modify ventilator alarms and status display messages during a GUI reset. The Puritan Bennett 980 Ventilator Operator Manual will be updated to delete,replace ventilator during a GUI transient reset, and to inform users to monitor the secondary status display until the primary display refreshes. Covidien will communicate with customers when the software and operator manual updates are available.Customers with questions were instructed to contact Technical Support Department at 800-255-6774, option 4, and then option 1.For questions regarding this recall call 303-876-8870. |  | 418 | 418 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | January 06, 2016 | 359 | 418 | Not_Computer
Z-1060-2015 | 70149 | No | BTY | Predictive pulmonary-function value calculator. | Anesthesiology | Anesthesiology | 510(k) | Shape-HF Cardiopulmonary Testing System Disposable Patient Interface (DPI). Part Number 0004-4001. |  |  | December 23, 2014 | 2014 | Shape Medical Systems, Inc | This recall has been initiated because use of the impacted DPIs could result in erroneous testing results potentially leading to incorrect diagnosis and incorrect treatment. The magnitude of error for VO2 or VCO2 parameters could be approximately 20% with use of the impacted product. | Under Investigation by firm | An Urgent Medical Device Recall letter was mailed to consignees on 12/23/2014. The letter identified affected product, described the issue and asked for product to be returned. Shape Technical Service will contact customers to assist in the return of product. A response form was asked to be returned. Customers with questions can call Shape Technical Service at 1-888-906-6266. |  | 564 | 564 devices | Distributed in the states of AZ, FL, GA, MA, MN, NJ. NY, PA, TX, and WI. | Terminated | April 14, 2015 | 112 | 564 | Not_Computer
Z-1008-2015 | 70127 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | MAQUET Servo Humidifier 163; Model No.: 01-06-8125-8, catalog no. XKC01-06-8125-8. |  |  | December 04, 2014 | 2014 | Teleflex, Inc. | Connector mount cracks which may lead to a leak failure during use and necessitate immediate replacement in the breathing circuit. If leak exists and is left untreated, patient may be deprived of adequate anesthetic gases over a period of time and serious adverse health consequences may occur. | Process control | An Urgent - Field Safety Notice, dated December 4, 2014, was sent to end users that identified the product, problem, and action to be taken. Customers were requested to immediately check inventory, discontinue use and distribution, quarantine. Customers were to return the Acknowledgement form and request a return authorization number from their local Maquet representative for affected devices. A credit will be issued upon receipt of the returned affected product. For questions or for support concerning this issue, a local Maquet representative should be contacted at fieldactions@maquet.com. |  | 18 | 18 | Distributed in the states of FL, MI, WV, and KS. | Terminated | March 21, 2017 | 838 | 18 | Not_Computer
Z-1051-2015 | 70257 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | MAQUET, SERVO Humidifier 163, Rx Only, Inmed Mfg. Sdn Bhd. Catalog No: XKC01-06-8125-8, Model No. 01-06-8125-8. |  |  | January 12, 2015 | 2015 | Teleflex Medical | Defects: Cracks that may occur in the center of the corrugated flex tube of the 15 mm and 22 mm ISO female connector. | Process control | Consignees were notified by an Urgent - Field Safety Notice, dated 01/12/2015. The letter identified the affected devices, a description of the problem, and actions to be taken. Customers were asked to check their inventory for any affected product, to cease use and distribution, and immediately quarantine any affected product. Customers are to complete and return the Acknowledgement form and to their local MAQUET representative, who provide a return number. A credit will be issues upon receipt of the affected product. Customer should contact their local MAQUET representative at fieldactions@maquet.com with questions or for support this issue. |  | 100 | 100 units | Distributed to FL, KS, MI and WV. | Terminated | May 10, 2016 | 484 | 100 | Not_Computer
Z-1063-2015 | 70361 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 Automated Breast Volume Scanner (ABVS), model number 10434730. |  |  | January 12, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | There is a low probability the scanner arm will become completely detached from the scanner column. | Vendor change control | Siemens sent a Customer Safety Advisory Notification to all affected sites on January 12, 2015, certified mail, with return receipt. The letter identified the affected product, the issue, and potential risk. Customers were asked to immediately discontinue use of the affected product. A Customer Service Engineer will contact customers to schedule an appointment to make the necessary repairs. Questions should be directed to a local service support person for information regarding timelines and status. Customers are to share the information with all personnel, who need to be aware of the issue. |  | 24 | 24 units | Worldwide Distribution -- US, including the states of TX and MN, and the territory of P.R.; and, the countries of China, Germany, Russian Fed, Saudi Arabia, Taiwan, France, Thailand, and Italy. | Terminated | June 10, 2015 | 149 | 24 | Not_Computer
Z-1062-2015 | 69863 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | CARDIOSAVE Intra-Aortic Balloon Pump and CS300 Intra-Aortic Balloon Pump; and, non-sterile, not for human use, SENSATION and SENSATION PLUS IABC demonstration units. Part Number 0998-00-0800-xx; 0998-00-3023-xx; 0998-UC-0800-xx; 0998-UC-3023-xx |  |  | December 16, 2014 | 2014 | Maquet Datascope Corp - Cardiac Assist Division | Ferrule Cap was not removed on some fiber-optic Sensation and Sensation Plus IABC demonstration units. If fiber-optic connector is inserted into the Maquet CS300 or CARDIOSAVE IABP, cap will prevent connection of unit and may become dislodged and remain in IABPs fiber-optic receptacle. | Device Design | Maquet Inc. sent a Product Recall Medical Device Field Correction letter/Field Correction response form, dated 12/16/2014 via Fed Ex. The letter identified the affected product and the reason for the recall. Users were asked to inspect their inventory for the affected units. If customers have any affected product, they are to follow the instructions provided. Additionally, users are to complete and return the attached Field Correction Response Form to acknowledge receipt of the letter and to indicate completion of the steps outlined. Completed forms should be returned by e-mailing a scanned copy to fieldactions@maquet.com or faxing a copy to 1-973-396-3607. Technical questions should be directed to the Technical Support Department at 1-800-777-4222. |  | 1202 | 838 units-US; 364 units - OUS | Worldwide Distribution | Terminated | June 09, 2017 | 906 | 1202 | Not_Computer
Z-1064-2015 | 69388 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Digital Mobile X-ray Systems used to take diagnostic radiographic exposures or patients. |  |  | September 22, 2014 | 2014 | Virtual Imaging, Inc. | There is a potential for unintended and erratic motion. | Component design/selection | Virtual Imaging sent an FIELD CORRECTION NOTIFICATION letter dated September 22, 2014, to all affected customers. The letters advised direct accounts that a Virtual Imaging Field Engineer with contact them to schedule an appointment for the replacement of the Digital Motion Control Kit. Customers with any concerns about this situation can contact Virtual Imaging's Technical Support at 561-893-8500 or via e-mail at technicalsupport@vifla.com. |  | 316 | 316 (US) 6 (Foreign) | USA (nationwide) to the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV) + US Army facility in South Korea; and the country of Canada. | Open, Classified |  |  | 316 | Not_Computer
Z-1067-2015 | 70157 | No | GWE | Evoked response photic stimulator. | Neurology | Neurology | 510(k) | EPWorks software used in the Protektor 32 |  |  | January 02, 2015 | 2015 | Natus Medical Incorporated | Software error occurs when using remote monitoring; if the remote user tries to stop the free run waveform group, the system will display a message informing the user that they do not have sufficient privilege. | Software design | Natus Medical Inc. DBA Excel-Tech Ltd (Xltek) issued medical device notification letters dated December 22, 2014 to their customers vial email and via courier service on January 2, 2015. Within the notice, customers are instructed are on to removal the previous installation and also on how to install the new program upgrade. Natus Medical Inc. has requested that the Previous Version Removed from Use Confirmation form be completed by each customer and returned. Customers who require assistance can contact Natus Technical service by calling 800-387-7516 or via email Oakville_Technical_Service@xltek.com or ots@natus.com. |  | 248 | 248 units | US Nationwide Distribution | Terminated | March 22, 2017 | 810 | 1176 | Not_Computer
Z-1046-2015 | 70103 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABL90 FLEX analyzer; Model number 393-090. A portable, automated analyzer that measures pH, blood gases, electrolytes, glucose, lactate, bilirubin and oximetry in whole blood. |  |  | December 08, 2014 | 2014 | Radiometer America Inc | The ABL90 analyzer does not always use the most recent calibration data to calculate patient results. This can in some cases lead to a biased patient result. | Software design | Radiometer sent a Field Safety Notice dated December 17, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. What you should do short term: " Please ensure that the operator always closes the inlet when prompted by the analyzer, as per the labeling. " In case the inlet, by mistake, is not closed as per the labeling, please ensure that the following is carried out before measuring a patient sample: " Close the inlet and wait for the analyzer to go to READY (approximately five minutes) " Open the inlet and then close again (to initiate a Rinse) " Wait for the analyzer to go to READY (approximately 1Âż minutes) " Please return the confirmation fax form no. 1 to your RADIOMETER representative when the above actions have been implemented Final Solution provided by RADIOMETER: A mechanism will be provided in an upgraded version of the analyzer software to eliminate the possibility of this error. The new software version will be installed by your local engineer when available.Customers were instructed to ensure that this letter is distributed to the final end-user. Customers with questions were instructed to contact Radiometer America Technical Support at 1-800-736-0600, opt 4. |  | 5002 | 5002 units | Worldwide Distribution -USA including CT, MN, FL, CA, PA, NY, KY, MA, WA, MI, TX, NC, IL, OH, SC, OK, NJ, MD, VA, UT, AL, NM, GA, LA, OR, IN, WI, AZ, ND, CO, MO, WV, ID, TN, VT, NE, IA, NH, MS, ME, AR and Internationally to Canada, Australia, Austria , China, Czech Republic, Denmark, Estonia, Lithuania, Romania, Slovenia, United Arab Emirates, Bolivia, Brazil, Indonesia, Israel, Kuwait, Malaysia, Mexico, Peru, Qatar, Saudi Arabia, Serbia and Montenegro, Syrian Arab Republic, Thailand, Vietnam, Finland, France, Germany, Hong Kong, Hungary, Ireland, Italy, Japan, Netherlands, New Zealand, Norway, Poland, Republic of Korea, Russia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | September 23, 2015 | 289 | 5002 | Not_Computer
Z-1070-2015 | 70426 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Aquarius iNtuition Client Viewer. Findings Workflow module, RECIST 1.1: |  |  | January 21, 2015 | 2015 | TeraRecon, Inc. | Software anomaly related to RECIST1.1 target lesion evaluation criteria in Findings Workflow Module within the Aquarius iNtuition Client Viewer. | Software design | Medical Device Correction letters were sent to all affected users on January 21, 2015 by both e-mail and physical mail - certified mail for us customers and FEDEX for international customers. |  | 91 | 91 | Worldwide Distribution-US (nationwide) and the countries of Canada, Switzerland, Germany, France and Italy. | Terminated | September 30, 2015 | 252 | 91 | Not_Computer
Z-1069-2015 | 70381 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Timberline Cranial Caudal and Posterior Blades, 40 - 180mm Model # 8734-2XXX, 8734-5XXX is a reusable device provided within an aluminum sterilization tray along with other instruments for use in a Timberline procedure |  |  | February 22, 2012 | 2012 | Biomet Spine, LLC | Biomet Spine, LLC is recalling the Timberline Cranial Caudal and Posterior Blades, 40-180mm due to oversize of the set screw hex feature. | Nonconforming Material/Component | Biomet initiated the recall on February 22, 2012, No letters were or will be sent to customers. 100% retrieval was achieved by working directly with the customers.For questions regarding this recall call 303-501-8548. |  | 213 | 213 | Distributed to TX, IN, MO, and NY. No foreign/govt/military/VA. | Terminated | March 23, 2015 | 1125 | 213 | Not_Computer
Z-1066-2015 | 70298 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | System 6 Aseptic Housing. |  |  | January 07, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Stryker is recalling System 6 Aseptic Housing, due to a failed Continuous Bond Line Test from a housing without a continual weld. Customers are instructed to return recalled product to Stryker. | Under Investigation by firm | Stryker sent notice to customers via Fed ex overnight on 1/16/2015, regarding its voluntary recall of Stryker System 6 Aseptic Housings because the laser welder may have failed to produce a continuous bond line for the four specific lots: 13205, 13209, 13210, 13212. Customers are instructed to discontinue use, complete the business reply form, and return all recalled product. Customers may contact Kara Spath at 269-389-4518, or kara.spath@stryker.com. |  | 399 | 399 | Worldwide Distribution-US (nationwide) including Puerto Rico and the states of AR, CA, CO, FL, GA, HI, MD, NC, NY, PA, SC, TX, and VA, and the countries of Australia, Canada, Switzerland, Hong Kong, South Africa, Chile, Korea, Poland, India, Colombia, The Netherlands, and France. | Terminated | March 09, 2016 | 427 | 399 | Not_Computer
Z-1065-2015 | 69837 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftPath Laboratory Information System. Versions 4.3.0.8, 4.3.0.9, 4.3.0.10, 4.3.0.11, 4.3.0.12, 4.3.0.14, 4.3.0.15, and 4.4.0.0 |  |  | November 26, 2014 | 2014 | SCC Soft Computer | Modifications to diagnostic text may be: 1) Saved to the database but not appear on the report sent to the physician; or 2) Documented on the report, but not saved to the database. | Software design | On 11/26/2014 SCC Soft Computer sent a Correction Communication to their consignees: Corrective Action: Versions 4.3.0.8 through 4.4.0.2 will have a mandatory hotfix created. Due to the amount of versions affected, the range of availability dates will be 11/24/2014 through 12/30/2014. Utility: A utility will be made available to be run on your Live system that will identify cases in which there is a disparity between the diagnosis text in the document and the database. The utility output will be provided for client review and validation. |  | 108 | 108 units | Distribution US Nationwide and CANADA. | Open, Classified |  |  | 108 | Not_Computer
Z-1078-2015 | 70248 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCI, Sheridan, Preformed Tracheal Tubes, Cuffed and Uncuffed, Rx Only, Sterile, Teleflex Medical. |  |  | January 09, 2015 | 2015 | Teleflex Medical | Labeling; The affected product packaging is incorrectly labeled as being a cuffed oral ET tube; The correct product labeling is uncuffed nasal ET tube. | Labeling mix-ups | Teleflex Medical sent an " Urgent Medical Device Recall Notification" dated January 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to immediately discontinue use and quarantine any of the affected products, return the product and to complete the enclosed Recall Acknowledgement Form even if they do not have the affected product in stock. Distributors were instructed to communicate the recall to any of their customers who may have received product included within the scope of the recall. Have the customers return any affected product to them, together with a completed Recall Acknowledgement Form.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 300 | 300 units | Worldwide Distribution - US Distribution to the states of FL, GA, IL, KY, MA, MI, MO, MS, OH, TX, WA and WV., and to the countries of Canada and the Netherlands. | Open, Classified |  |  | 300 | Not_Computer
Z-1087-2015 | 70259 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Siemens Symbia T series is intended to aid in detecting, localizing, diagnosing, staging and restaging of lesions, tumors, disease and organ function for the evaluation of diseases and disorders. The images produced by the system can also be used by the physician to aid in radiotherapy treatment planning and interventional radiology procedures. |  |  | October 23, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | There is a potential for the linear bearing car to fail allowing the ball bearings to fall out. | Component design/selection | Consignees were not notified of this issue. |  | 1075 | 1,075 systems (global) | Worldwide Distribution-US (nationwide) including DC and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA, WI, WV, and WY), and the countries of AUSTRALIA , AUSTRIA, BANGLADESH , BELGIUM , BRAZIL , BULGARIA , CANADA , CHILE , CHINA , COLOMBIA , CROATIA , CZECH REPUBLIC , DENMARK , ECUADOR , EGYPT , FINLAND , FRANCE , GERMANY , GREECE , HONG KONG , HUNGARY , INDIA , INDONESIA , IRAN , IRELAND , ISRAEL , ITALY , JAPAN , KUWAIT , LATVIA , LITHUANIA , MEXICO , MOROCCO , NETHERLANDS , NEW ZEALAND , NORWAY , PERU , PHILIPPINES , POLAND , PORTUGAL , QATAR , ROMANIA , RUSSIA , SAUDI ARABIA , SINGAPORE , SLOVAKIA , SOUTH AFRICA , SPAIN , SWEDEN , SWITZERLAND , TAIWAIN , THAILAND , TUNISIA , TURKEY , UNITED ARAB EMIRATES, UNITED KINGDOM , and VENEZUELA. | Completed |  |  | 1794 | Not_Computer
Z-1084-2015 | 70088 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | The MAYFIELDÂż Ultra 360 Base Unit for head positioner, one unit per package. |  |  | December 19, 2014 | 2014 | Integra LifeSciences Corporation | Through an investigation of complaints, Integra found that there is the possibility for the locking handle mechanism of the Base Units to fail during use. | Nonconforming Material/Component | Integra sent an Urgent - Voluntary Medical Device Recall letter dated December 19, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to take the following actions:1. Review your MAYFIELD equipment and determine if you have any MAYFIELD Ultra 360 Base Units2. Once identified we recommend that you discontinue use and remove them from service until they are repairedby an authorized Integra Repair Center3. Complete the attached Recall Acknowledgement and Repair Form4. Upon receipt of the Recall Acknowledgement and Repair Form, Customer Service will contact you and providean RMA number and instructions for returning the productor questions call 609-936-6822.For questions contact Customer Service:USA 800-654-2873Outside USA 609-936-5400 |  | 1053 | 1053 Units | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, WY, and DC. (25 government consignees) and in the foreign countries: CANADA, ALGERIA, ARAB EMIRATES, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CHILE, CHINA, COLUMBIA, CYPRESS, DENMARK, DOMINICAN REPUBLIC, FRANCE, GERMANY, GREAT BRITAN, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN , JORDAN, KOREA, LIBYA, MACEDONIA, MAYLASIA, MEXICO, NETHERLANDS, POLAND, PORTUGAL, PUERTO RICO, ROMANIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, TAIWAN, THAILAND, and TURKEY. | Open, Classified |  |  | 1053 | Not_Computer
Z-1076-2015 | 70051 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The SurgiWrap Bioresorbable Adhesion Barrier Film. SurgiWrap 100x130x0.05mm |  |  | August 26, 2014 | 2014 | MAST Biosurgery USA, Inc | The expiration date of the bar code on the envelope label did not match the expiration date next to the hourglass. The printed expiration date was August of 2017, while the bar code date read as July 31, 2017. The correct expiration date was July 31, 2017. | Labeling Change Control | MAST Biosurgery labeled 49 new envelopes and sent them to the distributor, and subsequently had the distributor remove the contents of the outer envelope (instructions for use and the foil pouches containing the product) and place them into the newly supplied and properly labeled outer envelopes provided by them. The distributor requested to send back the original 49 enveloped.The second batch of new labels was inspected and attached to the new outer envelope to create a properly labeled outer envelope prior to shipping. The firm requested that once the distributor replaced the product and the IFU into the newly labeled envelopes that they provide documentation that the product had been repackaged into the new envelopes. The distributor was also requested to verify that they did not ship any product to customers with the incorrect label. The firm had the distributor return the 49 improperly with the incorrect labels and the first set of "corrected" labels to them for destruction.Further questions call (858) 550-8050. |  | 49 | 49 pieces | Worldwide Distribution to the country of China. | Terminated | February 24, 2015 | 182 | 49 | Not_Computer
Z-1088-2015 | 70279 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Precise Digital Accelerator |  |  | January 07, 2015 | 2015 | Elekta, Inc. | Elekta has identified a batch of securing bolts that fix the Gantry counter frame weight stack may fail. | Other | Elekta sent an Urgent Important Field Safety Notice on January 7, 2015, to all affected customers. Elekta recommended that customers replace the M20 x 50 MM bolts marked 'OF' on the gantry counterbalance frame immediately. The digital linear accelerator must not be used until these 'OF' bolts are replaced. Elekta recommended that customers put the machine in a parked position at 0, and attach the gantry locking straps. A Modification kit will be available as soon as possible which contains the correct bolts. Customers were instructed to complete the Important Field Safety Notice Acknowledgement form and return to their local Elekta Office or Representative. Customers with questions were instructed to contact their local Elekta representative.For questions regarding this recall call 404-993-5886. |  | 46 | 46 | Worldwide Distribution - USA including AZ, CA, FL, KS, LA, MN, MO, NY, NC and Internationally to Austria, Bulgaria, Canada, China, Germany, Hungary, Italy, Malaysia, Malta, Netherlands, Norway, Russia, Spain, Sweden, Taiwan and Ukraine. | Terminated | April 14, 2016 | 463 | 46 | Not_Computer
Z-1083-2015 | 70351 | No | LFJ | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Hemodialysis Catheters and PICCs: Product identification: MCDLTSL24P (SF X 24CM LT HEMO-CATH), MC061422 (14FX15CM SLX DBL LMN ST CATH), |  |  | January 21, 2015 | 2015 | Medical Components, Inc dba MedComp | Drug products contained within the kits may have been rendered ineffective by a duplicate sterilization process. The kits include one or more of the following drugs: 5ml ampule of lidocaine, 3ml Chloraprep one-step, 10ml pre-filled saline syringe, Hygea saline wipes, and 3ml tinted Chloraprep. No injuries or illnesses were reported. | Process control | Medcomp sent a Product Alert/Recall letter dated January 21, 2015 to customers via email. The letter identified the affected product, problem and actions to be taken. Medcomp requested the return of all affected product. Customers were instructed to contact their customer service representative for instructions for returning the product. |  | 243 | 243 | US Nationwide Distribution | Terminated | August 03, 2015 | 194 | 243 | Not_Computer
Z-1082-2015 | 70175 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROWg+ard Blue PLUS Pressure Injectable Two-Lumen CVC Kit. The Multiple-Lumen Catheter permits venous access to the central circulation by way of the femoral jugular, or subclavian veins. |  |  | January 02, 2015 | 2015 | Arrow International Inc | Foreign material found within the kits. The foreign material, white or black particles and/or thread, is visible and can beidentified by the user. | Process control | An urgent Medical Device Recall Notification, dated January 2, 2015, was sent to end users requesting they check their inventory and return affected product per Arrow's instructions. |  | 1195 | 1195 | Nationwide | Open, Classified |  |  | 1195 | Not_Computer
Z-1077-2015 | 70332 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Flyte Togas are components of the Stryker Flyte System and are intended to be worn over any Stryker Flyte Helmet. This toga provides Level 1 and 4 protection according to the Association for the Advancement of Medical Instrumentation (AAMI) Liquid Barrier Performance and Classification System (PB70:2012). |  |  | December 22, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | A single lot (11121171) of Flyte Toga Peel-away X-Large (catalog number 0408-830-100) was labeled with an incorrect expiration date of 2412-12-01. The correct expiration date should have read 2014-12-01. | Error in labeling | Stryker sent an Urgent Medical Device Recall Notification letter dated December 22, 2014, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and responding to the formal recall notification. Customers were instructed to do the following:1. Immediately review this Recall Notification2. Immediately check all stock areas and/or operating room storage and quarantine any affected product found.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify howmany, if any, affected items are currently in their inventory. Complete and return the BRF even if theydo not have any affected product on hand.Note: Your signature on the BRF indicates that you received and understand this Notification andhave followed the instructions in this Notification.4. If customers further distributed this product, they were instructed to forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF. Completed Business Reply Forms should be faxed to 866-521-2762.Upon receipt of the Business Reply Form, Stryker will email a pre-paid shipper. This shipper can be used to return the recalled Flyte Toga(s) to Stryker.Upon receipt of the recalled Flyte Toga(s), credit will be applied to the customer's account.For questions regarding this recall customers were instructed to contact Stryker Instruments: Monday-Friday 8am-5pm (EST) at 269-389-2458.For questions regarding this recall call 269-389-2458. |  | 1530 | 1530 units (170 cases) | Nationwide Distribution including AL, AR, AZ, IA, IL, IN, LA, MA, NC, NE, NM, SC, TN, TX, WI, and WV. | Terminated | May 05, 2015 | 134 | 1530 | Computer
Z-1081-2015 | 70036 | No | LZU |  |  | Ophthalmic | 510(k) | VisiPlugST for the Lacrimal Efficiency Test. It is also sold in the EU as CollaSyn Plugs. Model number 1813, 1814, 1815, 1903, 1904, and 1905. |  |  | November 26, 2014 | 2014 | Lacrimedics Inc | Moisture level may cause the VisiPlugST (or CollaSyn Plugs) to break apart, or crumble upon contact. The firm's data also suggests that a 4 year shelflife may not be sustained. | Nonconforming Material/Component | The firm notified their foreign consignees on 11/26/14. The firm made a decision to conduct a domestic recall on 12/4/14 and will notify US consignees Consignees will be provided with a return label and request any devices that they still have be returned to the firm.Any consignees that have gone out of business will be sent a certified letter to their last known address.CORRECTIVE ACTIONS:The firm has halted sales of the device until the drying method can be improved. Any remaining stock has been placed in quarantine and will be scrapped. |  | 781 | 781 units | Worldwide Distribution-US (nationwide) including the states of AR, CA, FL, ID, IL, KS, MA, MI, NC, NJ, NY, OR, PA, and TX and to the countries of Australia, Brazil, Czech Republic, France, Ireland, Mexico, Peru, Switzerland, and United Kingdom. | Terminated | March 23, 2015 | 117 | 781 | Not_Computer
Z-1080-2015 | 70177 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | ARROW Epidural Needle Component (product number AN-05501). Epidural needles permit access to the epidural space and are used as a conduit for epidural catheter placement. |  |  | January 05, 2015 | 2015 | Arrow International Inc | The product labeling does not specify the 5 year shelf life of the product. | Labeling design | Urgent Medical Device Recall Notifications, dated January 2015, were sent to direct accounts that identified the product, problem, and actions to be taken. Customers were asked to discontinue use and return affected product. |  | 63022 | 63022 | Distribution US Nationwide, Canada and Mexico. | Open, Classified |  |  | 63022 | Not_Computer
Z-1085-2015 | 70258 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | LMA (Laryngeal Mask Airway) Supreme (New Cuff) Size 5, Sterile, Single use, |  |  | January 13, 2015 | 2015 | Teleflex Medical | The labelling of units from the affected lot indicates that it contains a size 5 LMA Supreme (New Cuff), but may contain a size 4 LMA Supreme (New Cuff). | Other | Teleflex sent an Urgent Medical Device Recall Notification letter dated January 13, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and quarantine any products and to return it to Teleflex Medical. Customers were asked to complete the enclosed Recall Acknowledgement Form via fax even if they do not have the affected product in stock. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 2640 | 2640 units | Worldwide Distribution - U.S Nationwide in the states of: AZ, CA, DE, FL, GA, IL, KY, MO, NY, OK, PA, TX, VA &amp; WA and the country of Canada. | Open, Classified |  |  | 2640 | Not_Computer
Z-1089-2015 | 70376 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Telluride Percutaneous Rod Inserter Long, Model #7706-1904. The Inserter is a reusable device provided within an aluminum sterilization case along with other orthopedic manual surgical instruments. The case consists of three trays of instruments, and the middle tray contains a quantity of two Inserters. |  |  | December 19, 2011 | 2011 | Biomet Spine, LLC | The Telluride Percutaneous Rod Inserter Long does not contain a feature that locks the Rod in place during use. There is a potential for the Rod to disengage from the Inserter during insertion, and could lead to harm to the patient. | Device Design | Consignees were informally contacted 12/16/2011 of units in the field requesting retrieval. All units were in Lanx's (Biomet's) possession at the time of the recall initiation on 12/19/2011. |  | 10 | 10 | Distributed to TX, OH, NE and NY. | Terminated | March 23, 2015 | 1190 | 10 | Not_Computer
Z-1090-2015 | 70166 | No | FWP | Chin prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MEDPOR Surgical Implant Contoured Two Piece Chin 5 mm; Part Number 86001; Lot Numbers A1405048, A1404026 |  |  | December 17, 2014 | 2014 | Stryker Craniomaxillofacial Division | Potential for implant damage, implant cracking off/breaking intra-operatively. | Process control | Stryker sent an Urgent Medical Device Recall Notice dated December 23, 2014, to all affected customers with a response form to be completed and returned upon receipt and review. The letter identified the affected product, potential hazards, risk mitigation and the actions to be taken. For questions contact Stryker CMF at 1-800-962-6558, or e-mail at CMF-custserv@stryker.com. |  | 1051 | 1051 | Worldwide Distribution - US Nationwide in the states of PA, KY, VA, CT, UT, MD, AR, FL, TX, AL, DE, IL, IA, MI, CO, OH, GA, CA, DC, MO, LA, WI, ID, SC, NC, NY, MA, NH, CA, OK, MN, KY, IN, OR, TN and the countries of: Argentina, Australia, Canada, Chile, China, Netherlands, France, Spain, South Korea, Colombia, Switzerland, Singapore, and South Africa. | Terminated | November 06, 2015 | 324 | 1051 | Not_Computer
Z-1094-2015 | 70055 | No | MTJ |  |  | General &amp; Plastic Surgery | 510(k) | Hemostatic Bone Putty, for use as a water-soluble implant material and for use in the control of bleeding from bone surfaces. |  |  | January 06, 2012 | 2012 | Synthes, Inc. | Revision A of the Hemostatic Bone Putty Brochure (J10847A) was sent to sales consultants instead of Revision B (J10847B), which corrected an error in the indication statement. | Labeling mix-ups | On/about 12/22/2011, emails containing Revision B brochure, as well as instructions to delete/destroy Revision A, were sent to 38 affected individuals (sales consultants). On/about 1/04/2012, follow up phone calls were made to each member of the sales force who had received Rev. A (110847 A) to confirm that they had replaced the Rev. A with the Rev. B (110847 B). In addition, the members of the sales force were asked whether or not the Rev. A had been disseminated to any HCP's. During the follow up conversation, one member of the sales force stated that he had provided the previous version to a surgeon. On 01/06/2012, the sales consultant spoke to the surgeon in person and the surgeon confirmed that he had disposed of the Rev. A. |  | 38 | 38 | Distributed in the state of WI | Terminated | March 30, 2015 | 1179 | 38 | Not_Computer
Z-1100-2015 | 70280 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | AQUAPAK 340 SW, 340 ML W/040 ADAPTOR, FREN. |  |  | January 14, 2015 | 2015 | Teleflex Medical | The packages may contain foreign material (black particles). | Material/Component Contamination | Teleflex sent and Urgent Medical Device Recall Notification letter dated January 14, 2015 to all affected customers. The letter identified the affect product, problem and actions to be taken. Customers were instructed to immediately discontinue use and quarantine any products with the affected catalog numbers and lot numbers. Customers were asked to complete and return the enclosed Recall Acknowledgement Form. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 15779607 | 15,779,607 units | Worldwide Distribution - U.S. Nationwide and the countries of: Australia, Bahrain, Belgium, Bermuda, Brazil, Canada, China, Ecuador, Germany, Great Britain, Hong Kong, Indonesia, Japan, Jordan, Malaysia, Mexico, New Zealand, Philippines, Singapore, South Korea, Saudi Arabia, Taiwan, Thailand, Uruguay and Zambia. | Open, Classified |  |  | 15779607 | Not_Computer
Z-1114-2015 | 69723 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Arthrosurface HHXL (OVO) 15.6x32mm Taper Post Fixation Component; Catalog Number: 8156-0032, Part Number: 156-0032. HemiCAP Contoured Articular Prosthetic for Humeral head resurfacing. Orthopedic shoulder prosthesis component. |  |  | November 24, 2014 | 2014 | Arthrosurface, Inc. | Fully threaded Taper Post Fixation components not cleared for marketing in the US with current indications. | No Marketing Application | Arthosurface sent an Urgent Medical Device Recall letter on November 24, 2014 via UPS, to all affected customers. The letter informed consignees of the recall and asked them to return two completed forms, 1. Acknowledgement &amp; Receipt Form and 2. Inventory Record via fax, email, or mail. Customers with questions were instcted to contact their local sales representatiave or email RA@arthrosurface.com. For questions regarding this recall call 508-528-4604. |  | 187 | 187 | Worldwide Distribution - USA nationwide. Internationally distributed to Europe, Canada, and Australia.Only products in US affected by this recall. | Terminated | March 27, 2015 | 123 | 786 | Not_Computer
Z-1115-2015 | 70304 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | GemLock Long Hex Driver, Catalog RHL2.5; Also distributed within the Tapered SwissPlus &amp; SwissPlus Implant Systems Kits, Surgical, Complete, Catalog OPCST; Tapered Screw-Vent Implant System Surgical Kit, Complete, Catalog TSVKIT. |  |  | January 22, 2015 | 2015 | Zimmer Dental Inc | The Gemlock Long Hex Driver may not fit into the Fixture Mount Transfer (FMT) or the internal hex of the implant. | Device Design | An Urgent Device Recall Notice was sent on 1/22/15 to their domestic and foreign customers who purchased the Gemlock Long Hex Driver. The letter informs the customers of the problems identified, risks, and the actions to be taken. Customers are instructed to review the notice and check their inventory for the affected units. All affected product is to be immediately quarantined and removed from service. Customers are to contact Simmer Dental Customer Service at 1-800-854-7019 to obtain a Return Authorization Number and set up return of the products. Customers should complete the attached Business Reply Form and fax it to (574) 372-4265 or e-mail to corporatequality.postmarket@zimmer.com. Customers with questions are instructed to contact Zimmer Dental Customer Service at (800) 854-7019, Mon-Fri, 7 am-5pm (Pacific).Zimmer Dental is expanding the recall to include 1 additional lot for catalog no.TSVKIT. Zimmer Dental plans to inform an amended customer notification letter to their US and foreign consignees. The letter will go out to customers on 4/20/15 to inform them of the expanded lot. |  | 627 | 627 units | Worldwide Distribution -- US, Argentina, Australia, Bulgaria, Canada, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, Iran, Israel, Italy, Japan, Lebanon, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Pakistan, Panama, Poland, Romania, Russia, Saudi Arabia, Serbia, Taiwan, Thailand, Tunisia, Turkey, and UAE.Expanded Distribution: Spain | Terminated | August 03, 2015 | 193 | 627 | Not_Computer
Z-1116-2015 | 70492 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Terumo HX2 Temperature Management System provides temperature control of two independent water circuits that directly controls the temperature of patient blood and cardioplegia solution during cardiovascular surgery. |  |  | February 06, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Malfunctioning mixing valve and loss of cooling and/or heating capability. System goes to standby mode and shows error messages EOd (invalid mix valve measured range of motion) or EOE (valve unable to move or complete a motion). Mixing valves need to be replaced at regular intervals. | Device Design | On 2/6/2015, URGENT MEDICAL DEVICE RECALL notifications were mailed to the affected hospitals. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call customer service at 1-800-521-2818. |  | 103 | 103 units | US (nationwide) including the states of AR, CA, CO, CT, FL, IA, IL, KS, KY, MD, MI, MO, MS, NY, OH, PA, TX, UT, WA, and WI and the country of Japan. | Open, Classified |  |  | 103 | Not_Computer
Z-1117-2015 | 70275 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | HARMONIC ACE Curved Shears Instructions for Use, 5MM, IS1200; for use with the da Vinci Standard Surgical System. |  |  | January 21, 2015 | 2015 | Intuitive Surgical, Inc. | Internal labeling review found that the IFU for the Harmonic Ace incorrectly instructs the user on how to manually open the grips. | Error in labeling | A Field Safety Notice entitled, Urgent Medical Device Correction, and dated, January 21, 2015, were sent to customers. The letter identified the problem and the reason for the recall as well as the risk to health. Customers are asked to ensure that all affected personnel are informed of the notice. Additionally, customers are to discard their current Harmonic Ace IFUs. Customers are to refer to the specified manual, reference guide, and instruction card mentioned in the letter. A copy of the notice is to be retained with the da Vinci Standard, S or Si user manual. Intuitive Surgical will remove the Harmonic ACE IFU from all future shipments of the product. Representatives will be available by phone to answer questions related to the Correction. |  | 12467 | 12,467 - both sizes | Worldwide Distribution -- US, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, India, Israel, Italy, Japan, Lebanon, Mexico, Monaco, Netherlands, Pakistan, Panama, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and the United Kingdom. | Terminated | June 01, 2015 | 131 | 12467 | Not_Computer
Z-1121-2015 | 69782 | No | KRB | Thermodilution probe. | Cardiovascular | Cardiovascular | 510(k) | Pulsiocath PiCCO ProAQT Monitoring Kit, single-use cardiac output (CO) sensor, one unit per plastic sterilizable package. |  |  | November 04, 2014 | 2014 | Pulsion Medical Inc | Internal testing and investigation of the packaging revealed cracks and holes in primary packaging | Packaging | Pulsion sent an Urgent Medical Device Recall notification letter on 11/04/14 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers where instructed to:" Piease examine your inventory immediately to determine if you have any of the products listed." lf so, please remove the affected products; quarantine them in a secure location." Piease complete and return the enclosed Field Recall Response Form, acknowleding your receipt and understanding of this communication " Send all affected products back to PULSION Medical Systems." Upon return of the affected product you will receive a credit note.For questions contact your PULSION Medical lnc Recall Coordinator. |  | 90 | 90 units | US Nationwide Distribution in the states of CT, NJ and NY. The affected product was distributed to one government consignee. | Terminated | October 02, 2015 | 332 | 90 | Not_Computer
Z-1124-2015 | 70087 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | RÂżSCH Curved Reinforced Endotracheal Tube: Rusch Reinforced Endotracheal Tube is designated for oral or nasal intubation and are indicated for airway management. The product may be used to reduce the potential for kinking whenever an unusual positioning of the head or neck is required following intubation. |  |  | December 12, 2014 | 2014 | Teleflex, Inc. | Teleflex Medical has issued an advisory notification for the RUSCH Curved Reinforced Endotracheal Tube because of a discrepancy between the labeled cuff diameter and the actual cuff diameter. The actual product cuff diameter meets required specification. | Labeling mix-ups | Teleflex Medical sent an Urgent Medical Device Notification dated December 12, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Please provide this Advisory Notice to all those who need to be aware of it within your organization. Please consider, clinicians, risk managers, supply chain/distribution centres, etc. in the circulation of this notice. There is no further action required.Clinical recommendation with regard to patient use is to continue the standard clinical practice of basing ET tube selection based on the tube size, carefully checking for the adequacy of ventilation and to change the ET tube size up or down as appropriate when intubating the patient.INSTRUCTION FOR DISTRIBUTORS OF AFFECTED PRODUCTIf you are a distributor, provide this advisory notice to all of your customers who have received product in scope of this action.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 2292361 | 2,292,361 | Nationwide Distribution | Terminated | July 30, 2015 | 230 | 2292361 | Not_Computer
Z-1119-2015 | 70342 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis One; The Artis One is an angiography system developed for diagnostic interventional procedures including, but not limited to pediatric and obese patients. Procedures that can be performed with the Artis One include cardiac angiography, neuro-angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures as well as procedures next to the table for patient extremities. Additionally, angiographic procedures can be performed in the operating room, image guided surgery by x-ray, by image fusion, and by navigation systems. |  |  | January 16, 2015 | 2015 | Siemens Medical Solutions USA, Inc | The possibility exists that the monitor may fail and requires a power circle (shutdown and then power on) to resume operation. The problem is not systematic; but sporadic on single units. | Process control | Siemens notified their customer via Customer Safety Advisory Letter AX049/14/S, dated January 16, 2015, and has initiated a fix with Update Instruction AX048/14/S which will be completed by a Siemens service technician. Customers with questions should contact the Customer Care Center at 1-800-888-7436.For questions regarding this recall call 610-219-6300. |  | 1 | 1 | Nationwide Distribution to MI only | Terminated | April 21, 2015 | 95 | 1 | Not_Computer
Z-1122-2015 | 70438 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Mallory Head 4 Finned Acetabular Shell 50 mm, PN 13-104150, hip prosthesis component (Ringloc) |  |  | January 09, 2015 | 2015 | Biomet, Inc. | Biomet recalled the Mallory Head 4 Finned Acetabular Shell 50mm (Lot 364600) and Ringloc + Acetabular Shell Limited Hole Finned 52 mm (Lot 427840) following a mix up in packaging with the two. | Packaging process control | Biomet notified all domestic and foreign customers via FEDEx overnight on 1/9/2015. Response forms were provided, customers were instructed to notify hospital personnel if product was further distributed, and remove recalled products from circulation. Replacement parts were made available, customers may contact Audrey Daenzer, Field Action Specialist, Regulatory ComplianceBiomet, Inc., audrev.daenzer@biomet.com, or 574-372-1570 Monday- Friday 8 AM to 5 PM. |  | 17 | 17 | Distributed in the states of VA and TX and the country of Japan. | Terminated | October 22, 2015 | 286 | 17 | Not_Computer
Z-1125-2015 | 70363 | No | DXH | Telephone electrocardiograph transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Alive ECG App 2.1.2 ( a medical device application for the Apple iOS operating system, intended to be used with the AliveCor Heart Monitor. The AliveCor Heart Monitor is intended to record, store and transfer single-channel electrocardiogram (ECG) rhythms. The AliveCor Heart Monitor also displays ECG rhythms and detects the presence of atrial fibrillation and normal sinus rhythm (when prescribed or used under the care of a physician). |  |  | January 09, 2015 | 2015 | Alivecor SFO | Alive ECG App version 2.1.2 (intended to be used with the AliveCor Heart Monitor) crashed upon use of the application. | Software change control | AliveCor posted information on their website, Facebook Page and Twitter to alert users of the issue. |  | 5600 | 5600 active users with Alive ECG app for iOS . | Downloaded by Apple users - locations not shared by Apple. | Terminated | February 18, 2015 | 40 | 5600 | Not_Computer
Z-1035-2015 | 70209 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | This correction applies to all GE Healthcare MRI systems with superconducting magnets: Discovery MR450, Discovery MR750, Signa HDx, Signa HDxt, Signa HDi, GE 1.5T Signa HDe, Optima MR360, Brivo MR355, Signa Excite-HD 1.5T &amp; 3T, Signa Excite 1.5T HD Twinspeed, Signa Excite 1.5T HD Echospeed, Signa Excite 1.5T HD Highspeed, 1.5T Signa Infinity TwinSpeed , 1.5T Signa Infinity EchoSpeed Plus, 1.5T Signa Infinity HiSpeed Plus, 1.0T Signa Infinity HiSpeed Plus. 1.0T Signa Infinity SmartSpeed, Signa EXCITE 3.0T, Signa EXCITE 3.0T HD, Signa Excite 1.5T TwinSpeed, Signa Excite 1.5T EchoSpeed, Signa Excite 1.5T HiSpeed, Signa Excite 1.5T SmartSpeed, Signa Excite 1.5T, Signa Excite 3.0T, Signa Contour/I, Signa OpenSpeed, Vectra, MR Max, Optima MR450w GEM, Discovery MR750W GEM, Optima MR450w. |  |  | January 07, 2015 | 2015 | GE Healthcare | At certain sites, the MRU may not be connected to the magnet. In emergency situations, a disconnected MRU could delay removal of a ferrous object from the magnet, potentially resulting in life-threatening injuries. The MRU must be connected to the magnet at all times. | Employee error | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref# 60876 dated January 6, 2015 to affected consignees. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction &amp; Contact Information. Customers were instructed to do the following:As a preventative measure, confirm that MRU is connected to the magnet by performing the following four step test on the MRU.1. Verify the green CHARGER POWER LED is lit.2. Depress and hold the TEST BATTERY switch for 15 seconds. The green BATTERY TEST LED should light and remain lit while the TEST BATTERY switch is depressed.3. Place the TEST HEATER toggle switch in the A position. The green HEATER TEST LED should light. If it does not light, depress TEST HEATER LED switch to verify that the LED is functioning.4. Place the TEST HEATER toggle switch in the B position. Green HEATER TEST LED should light. If it does not light, depress TEST HEATER LED switch to verify that the LED is functioning.If the MRU test does not perform as described in each of the 4 steps above, GEHC strongly recommends that you stop using the system, and immediately call your GEHC representative.Customers with questions may contact their local service representative.For questions regarding this recall call 1-800-437-1171. |  | 12968 | 12,968 (5,708 US, 7,260 OUS). | Worldwide Distribution - US Nationwide in all states in continental USA including DC, PR, GU, and the countries of ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BOLIVIA, BOSNIA &amp; HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR,K ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GEORGIA, GERMANY, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MACEDONIA, MALAYSIA, MALTA, MARTINIQUE, MAURITIUS, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE , UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, UZBEKISTAN, VENEZUELA, VIETNAM, ZAMBIA. | Terminated | September 10, 2015 | 246 | 12968 | Not_Computer
Z-1061-2015 | 70238 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Echostar Spica 1.5T MRI system, for magnetic resonance imaging of the body. |  |  | January 09, 2015 | 2015 | Alltech Medical Systems America Inc | The MRI system safety feature, which lowers the magnetic field in emergency situations, may not work correctly. | Process control | Altech Medical Systems America Inc., notified customers on January 9, 2015, of the problem via telephone. Customers were instructed to immediately cease using the product until the correction can be made. A follow-up letter was sent through the mail, which stated that a service engineer would make a site visit to ramp down the magnet until the issue can be corrected. Customers were asked to complete the attached Acknowledgement and Receipt Form and fax it to (440) 424-2255. For questions contact AMSA at 1-855-200-2672. |  | 2 | 2 units | US Distribution in the state of Maryland. | Terminated | September 15, 2015 | 249 | 2 | Not_Computer
Z-1126-2015 | 70294 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | BodyGuard 323 pump, models |  |  | January 16, 2015 | 2015 | CME America, LLC | CME America is recalling the BodyGuard and BodyGuard 323 Infusion pumps due to the potential for an over delivery. | Software design | CME America sent an "Urgent Medical Device Recall" letter dated January 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Required Action for Users: To date there have been no reports of adverse events or serious injury. However, there is the potential for serious injury or death. CME America recommends that you DO NOT use the BodyGuard 323 pump for PCA infusion. If your devices are currently in use, be sure that the programmed parameters DO NOT involve a patient controlled bolus dose less than 0.1 miln volume (translated into meg or mg).Follow-up Actions by CME America: CME America has identified the root cause of this issue and is recommending installation of a new version of software, USA36B4, for affected BodyGuard 323 infusion pumps. This software is identical to the previously released version, USA36A4.2. However the designation of B4 in the software name will identify updated software in each device. CME America will contact all affected customers within 60 days to initiate the scheduling process for the remediation.Your assistance is appreciated and necessary to prevent use and facilitate remediation of the recalled product. CME America can address any questions regarding the recalled product by calling the CME America recall line at 720-839-2391. |  | 3186 | 3,186 | Worldwide Distribution - Nationwide Distribution and VA/military/govt consignees and the country of Canada. | Terminated | March 31, 2017 | 805 | 3186 | Not_Computer
Z-1127-2015 | 70335 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | MC-500 Multicolor Laser Photocoagulator; |  |  | March 27, 2012 | 2012 | Nidek Inc | The spot size control of the scan delivery unit of the MC-500 Multicolor Laser Photocoagulator may turn by itself. | Component design/selection | Nidek Japan issues a Engineering Change Order and Nidek reps visited affected sites to perform correction. |  | 16 | 16 units in commerce | Nationwide Distribution- including the states of CA, MI, NY, IL, NH, KS, NY, UT, IA, WI, TX, and FL. | Terminated | February 19, 2015 | 1059 | 16 | Not_Computer
Z-1128-2015 | 70374 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Restoris Partial Knee Application (PKA) RIO (TGS 2.0). |  |  | January 22, 2015 | 2015 | Mako Surgical Corporation | When using the MAKOplasty partial knee Arthroplasty application, the burr continues spinning outside of the stereotactic boundary and after the control switches (foot pedal and trigger) cease to be activated | Device Design | On January 22, 2015 the firm sent an IMPORTANT MEDICAL DEVICE CORRECTION asking consignees to complete the attached acknowledgement form and return them by email, fax or mail. The notice also noted the RIO system can continue to be used. The likelihood of occurrence is rare (.035% of cases) and the event is readily detectable and can be resolved. Should you have any questions regarding this Important Medical Device Recall Notice, to contact us at (954) 628-1721. |  | 228 | 228 units | Nationwide Distribution-including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, LA, MA, MD, MI, MO, MS, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV and Hawaii. | Open, Classified |  |  | 228 | Not_Computer
Z-1130-2015 | 70352 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Haemonetics TEG Hemostasis System Level II Control, Catalog number: 8002. |  |  | January 26, 2015 | 2015 | Haemonetics Corporation | Specific lots of Level II Control Kits produce high Maximum Amplitude (MA) results, which are outside of the manufacturer's expected Quality Control range for MA parameter. | Process control | Haemonetics issued on 1/26/15, an Urgent Medical Device Recall to consignees The letter identified the affected product, the reason for the recall, as well as the risk to health. It instructs consignees to inspect their stock and determine if they have any of the affected lots. Customers are to immediately discontinue use and discard as instructed any affected product. The attached customer acknowledgment form should be completed and returned whether or not customers have any affected product. Customers are to contact their local service representative at 1-800-537-2802 if replacement product is needed. If any affected product has been forwarded to other locations, customers are to inform them of the recall. Please call the Technical Support team at 1-800-438-2834, with questions about this action. |  | 4143 | 4,143 kits | Worldwide Distribution -- US, Canada, Czech Republic, Denmark, Great Britain, and Sweden. | Terminated | June 05, 2015 | 130 | 4143 | Not_Computer
Z-1129-2015 | 70207 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Cold Therapy Pad - cold water circulating pad, one unit per package. |  |  | January 07, 2015 | 2015 | Cincinnati Sub-Zero Products Inc | Seals in the water path of the CT-99 cold therapy pad may separate during use, which could lead to a water leak. | Process design | Cincinnati Sub-Zero sent an Urgent Medical Device Recall Field Action Notice dated January 7, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease use and/or further distribution of the affected product. Customers were instructed to obtain a Return Material Authorization Number and return the pad(s) for exchange. Customers with questions were instructed to call 1-800-989-7373 or 513-772-8810. Customers were also asked to commplete and return the enclosed Response Form as soon as possible by fax to (513)772-9119 to acknowledge receipt of the notification and to inform CSZ that they have performed the requested actions.For questions regarding this recall call 513-772-8810, ext 6816 |  | 1390 | 1390 | Nationwide Distribution including FL, IL, TN, GA, VA, MO, NM, WI, MI, NC, IN, OH, MA, CA, WA, PA, KY, MD, HI, and AL. | Terminated | June 19, 2015 | 163 | 1390 | Not_Computer
Z-1137-2015 | 69378 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TDX Electric Wheelchair. Invacare FDX, TDX series and Storm Series (3G) power wheelchairs in 20 to 24 inch width configurations with power recline seating system and a conventional style back. |  |  | January 15, 2015 | 2015 | Invacare Corporation | If the slack in the wires is not routed and secured correctly, flexing of the back pan under the users weight may cause the headrest knob or clamp bracket to pinch, damage or cut the wire. Smoking, sparking, burning or fire may occur if a wire is repetitively pinched at the same place. | Device Design | The firm initiated the recall on 01/16/15 by sending a notification letter to customers through the mail. Communication was sent to providers to contact consignees via mail return receipt to inform them of the recall. Recall notification included a Provider letter, Consumer letter, Provider Response card, Unit Disposition form and tracking sheet and the Wiring guide. The notification instructed the provider to visually inspect all potentially impacted chairs. Not all potentially impacted chairs are expected to have improperly routed and secured wires. |  | 1942 | 1942 | Worldwide Distribution-US (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Canada and Australia. | Terminated | September 23, 2015 | 251 | 2527 | Not_Computer
Z-1135-2105 | 70401 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CoCr Femoral Head - indicated for hip arthroplasty for noncemented use in skeletally mature invididuals undergoing primary surgery for rehabilitating hips damaged as a result of noninflammatory degenerative joint disease |  |  | December 17, 2014 | 2014 | Zimmer Gmbh | Potential compromise of the inner sterile tray. If a compromised tray is detected, a delay in surgery may result. If undetected breach in both the inner and outer trays were to occur, and there is contamination present, a periprosthetic infection is possible, which could lead to revision. | Package design/selection | On 12/17/2014, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital Risk Managers with instructions for returning the affected product. All distributors were notified via electronic mail, and all hospital Risk Managers and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call the customer call center at 1-800-348-2759. |  | 6109 | 6109 | Domestic:AK, AL, AR, AZ, CA, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MS, MA, MD, ME, MI, MN, MO, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, WYInternational: Canada, Japan | Terminated | September 23, 2015 | 280 | 18403 | Not_Computer
Z-1140-2015 | 70510 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Maquet CARDIOSAVE Hybrid Intra-Aortic Balloon Pump (IABP) |  |  | January 30, 2015 | 2015 | Maquet Datascope Corp - Cardiac Assist Division | Power supply malfunction complaints related to suboptimal thermal management. | Under Investigation by firm | Maquet Inc. sent a recall letter/return response form dated 1/30/2015. |  | 1300 | 1,300 units | US Nationwide distribution | Open, Classified |  |  | 1300 | Not_Computer
Z-1072-2015 | 69336 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Plum A+ Infusion System is designed to meet the fluid delivery requirements of today's evolving healthcare environments. It is a cassette based multi-function infusion system. The Plum A+ allows two lines in and one line out. The pump can be used for standard, piggyback, or concurrent delivery. Therapy modes include: Standard Infusions, Multistep Programming, Loading Dose, and Dose Calculation. The Plum A+ is designed to deliver parenteral, enteral, or epidural infusions over a broad range of infusion rates from multiple fluid container types. Each system includes a pumping module and an assortment of disposable IV sets, optional accessories, and an operator's manual. |  |  | July 02, 2014 | 2014 | Hospira Inc. | One lot of alarm assemblies used in Plum A+ and Plum A+3 infusion pumps may fail to sound at all volume levels. | Component change control | Hospira sent an Urgent Medical Device Correction letter dated July 2, 2014, to all affected customers. The letter notified customers that Hospira will be contacting them regarding the completion of an audible alarm test and will replace any alarm assemblies that fail to audible alarm test. Additionally, the letter included instructions for customers to: 1) inform potential users of this notification; 2) complete and return the attached reply form; and, 3) if the products were further distributed, notify those accounts of this recall and ask them to contact Stericycle at 888-912-7350 (Monday - Friday, 8:00 AM - 5:00 PM ET) to receive a reply form. A second URGENT MEDICAL DEVICE CORRECTION letter dated July 16, 2014 was sent to all customers of record because the initial letters may have contained errors in the attached data files listing the potentially affected devices at the customer's' facility. The instructions to the customer were unchanged from the original letter.For questions regarding this recall call 800-441-4100. |  | 3939 | 3,939 infusion pumps | Worldwide Distribution - USA including AK, AL, AR, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico and Brazil, Canada, Colombia, Ecuador, France, Israel, Italy, Mexico, Oman, Peru, Saudi Arabia, Spain, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Open, Classified |  |  | 9109 | Not_Computer
Z-1141-2015 | 70467 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist One Vessel Sealer, IS3000, 6 pack used in conjunction with the da Vinci Surgical System Si IS3000 |  |  | February 11, 2015 | 2015 | Intuitive Surgical, Inc. | Intuitive Surgical has identified the potential for a certain, small portion of Vessel Sealers to exhibitinterference in the back of the instrument grips due to manufacturing variability in some jaws. | Process control | Intuitive Surgical sent an Urgent Medical Device Correction letter dated Feruary 11, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following actions:1. Forward this letter to any members of your medical staff who perform da Vinci procedures in addition to your Risk Manager, OR Director, Purchasing and Biomedical Engineering staff.2. This field correction does not require you to quarantine or return the affected devices. Vessel Sealers can continue to be used safely by adhering to the following guidelines as provided in the Vessel Sealer User Manual:a. Only procedures that incorporate sealing and transection of thin tissue bundles and small diameter vessels are affected by this issue. Prior to the procedure, determine whether this issue may be a factor and choose whether or not to use the device per the following instructions. Alternative methods of sealing and transection are described below.b. During sealing, always inspect tissue for energy effect (i.e. blanching, steam, bubbling) along the entire jaw of the instrument. While audio tones designate seal cycle completion, tissue effect should always be seen prior to proceedingto cut.If no tissue effect is seen, then DO NOT CUT. Re-grasp tissue and re-attempt sealing as described above.If the issue persists, the instrument may be affected and an alternative means of sealing and transecting thin vascular tissue bundles should be used. These other methods include using another da Vinci Vessel Sealer, alternative da Vinci electrosurgical instruments, a da Vinci Harmonic instrument or hand-held laparoscopic sealing device.3. Complete the attached Acknowledgement Form and return it to Intuitive Surgical as instructed.4. Please retain a copy of this letter and the Acknowledgement Form for your files.5. If you choose to not continue to use affected Vessel Sealer instruments, you may retu |  | 3282 | 3282 total units, all part numbers | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, China, Denmark, Ecuador, Finland, France, Germany, Greece, Italy, Netherlands, Norway, Puerto Rico, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | June 01, 2015 | 110 | 3282 | Not_Computer
Z-1146-2015 | 70383 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Set Screw 5.5 mm Rod - Model #7703-1600. |  |  | November 30, 2012 | 2012 | Biomet Spine, LLC | Biomet Spine LLC is recalling the Set Screw 5.5 mm Rod due to the set screw being out of specification | Process control | Biomet Spine LLC (formerly Lanx) sent an Urgent Recall Notification letter to affected customers. The letter identified the affected product, problem, and actions to be taken. For questions contact Lanx Customer Service. |  | 1999 | 1999 | Worldwide Distribution - US Nationwide in the states of: AZ, CA, CO, CT, FL, GA, HI, IA, ID, IN, KY, LA, MA, MD, MI, MO, NC, NE, NY, OH, OR, PA, TN, TX, VA, WI. and the countries of: Italy, and Mexico. | Terminated | May 22, 2015 | 903 | 1999 | Not_Computer
Z-1143-2015 | 70388 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Starburst MRI Semi-Flex Electrode Device, 25cm/Attached Cable/14Ga/6.4F, Product No. H7877001039091, Catalog No. 700-103909; 25cm/14Ga/6.4F, Product No. H7877001026151, Catalog No. 700-102615; and 25cm/14Ga/6.4F, Product No. H7877001026151, Catalog No. 700-102615 |  |  | December 19, 2014 | 2014 | Angiodynamics | Printed mark on the trocar intended to indicate a 5cm distance from the distal tip of the needle is actually located 4cm from the distal tip of the needle. | Other | AngioDynamics sent an Urgent Medical Device Recall - Immediate Action Required letter dated December 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken:Immediately remove the recall product from your inventory (whether in labs, Central Supply, Shipping and Receiving or ANY other location).Segregate this product in a secure location for return to AngioDynamics, Inc.Immediately forward a copy of this recall notification to all sites to which you have distributed affected product.If affected product is located in your institution, please call AngioDynamics, Inc. Customer Service at 1-800-772-6446 between 8:00a.m. and 7:00p.m. (Monday- Friday: Eastern Standard Time) to obtain a replacement or credit for your returned product.Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return) following the directions on this page and the Reply Verification Form. |  | 43 | 43 | Worldwide Distribution - US Distribution to the states of AZ, AR, FL, MI, MN, MO, NY, TX, VA, WA and WI., and to the countries of Hungary and Spain. | Terminated | November 27, 2015 | 343 | 43 | Not_Computer
Z-1208-2015 | 70297 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Comprehensive Reverse Shoulder Glenosphere Mini Baseplate 25mm with Taper Adaptor, |  |  | January 07, 2015 | 2015 | Biomet, Inc. | The taper adaptor is missing from the packaging, which may result in a delay in surgery greater than 30 minutes. | Process control | Biomet issued written notices on 1/7/2015, Requested completion and submission of response/effectiveness forms for Product Part Number: 01 0000589 Product Description: Comprehensive Reverse Shoulder Glenosphere MiniBaseplate 25mm with Taper Adapter, Expiration October 2024. Questions related to this notice should be directed to Audrey Daenzer, Field Action Specialist (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 22 | 11- 6 consumed, 5 with distributors | Domestic: AL, FL, NY, TXForeign: None | Terminated | October 22, 2015 | 288 | 22 | Not_Computer
Z-1145-2015 | 70453 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare qube Compact Monitor, Model 91390. |  |  | January 28, 2015 | 2015 | Spacelabs Healthcare Inc | Spacelabs Healthcare qube Bedside Monitors, Model 91390, are recalled because the firm has received multiple reports of qube monitors failing to boot up or returning to factory default configuration settings following power on or reset. | Device Design | Spacelabs began to send the Urgent Medical Device Correction XPREZZON Bedside Monitor, Model 91393 qube Compact Monitor, Model 91390 letter, dated 28, January 2015, to US consignees.Spacelabs began sending the Urgent Medical Device Correction XPREZZON Bedside Monitor, Model 91393 qube Compact Monitor, Model 91390 letter (translated as necessary) to international customers via email on 2/4/15.Actions to be Implemented by Users:-Do not rely solely on the monitor for critically ill patients.-If customer facility operates any XPREZZON or qube monitors on a standalone basis, rather than via central monitoring, please contact Spacelabs technical support as these monitors will be the first repaired.Spacelabs Healthcare will contact consignees to schedule a convenient time for Spacelabs to replace the PCBAs in all of your facility's affected XPREZZON and qube monitors at no cost.If you have any questions about this corrective action, please contact Spacelabs at the toll free number of 1-800-522-7025 and select 2 for Technical Support, or via direct dial at 1-425-363-5203. |  | 2955 | total 2955 units (2117 units in US and 838 units outside US) | Worldwide Distribution. | Terminated | December 30, 2015 | 336 | 2955 | Not_Computer
Z-1131-2015 | 70340 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HeartWare Controller) |  |  | January 16, 2015 | 2015 | HeartWare Inc | The affected clinical trial Controllers exhibit a higher susceptibility to ESD than newer commercial Controllers. | Device Design | HeartWare sent an URGENT MEDICAL DEVICE RECALL dated January 16, 2015 asking clinicians to quarantine the affected controllers and complete and return the acknowledgement form. HeartWare also asks clinicians to identify patients with affected controllers and review the applicable risks with the patient and as soon as reasonably possible if medically advisable, exchange the affected controller under medical supervision with a new controller. HeartWare also asks that for each patient using an affected controller as their backup to contact the patient and arrange to have the backup controller replaced with a new controller. Lastly, HeartWare asks the firm to complete the patient tracking form and return all quarantined controllers to HeartWare along with the completion form. Should you have any questions or concerns, please contact your local HeartWare representative. Our 24-Hour Clinical Support personnel are also available at 1-888-494-6365. |  | 4845 | 4,845 | Worldwide Distribution - US Nationwide in the states of: AL, AZ, CA, CO, FL, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NY, OH, OR, PA, TX, UT, VA, WA and District of Columbia and countries of: Canada, Poland, Australia, Austria, Italy, DE, New Zealand, South Africa, France, Switzerland, Germany, Brazil, UK, Belgium, Netherlands, Israel, Malaysia, Turkey, Greece, Japan, Singapore, Poland, Lithuania, Sweden, Norway, Finland, and India. | Open, Classified |  |  | 4845 | Not_Computer
Z-1212-2015 | 70545 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | COMPASS SW Version 3.1, Catalog Number CS10-100, medical linear accelerator, radiological |  |  | February 02, 2015 | 2015 | Iba Dosimetry Gmbh | Error in the software. During internal tests of the current development version of the Compass SW it was found that dose reconstruction for DMLC plans when the jaws move or mlc leaves reverse during beam on will not be scaled correctly for all control points. | Software design (manufacturing process) | IBA sent an Field Safety Notice dated January 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken:Please do not use the dose reconstruction function COMPASS 3.1 and 3.1a for DMLC plans from Monaco or Eclipse -11 when the jaws move or MLC leaves reverse during beam on. A correct Version ( COMPASS 3.1b) will be delivered as soon as it is available. The field safety notice including a confirmation of receipt was sent to the affected hospitals; by email and UPS letter on February 2, 2015.Should you have questions or require additional information, please contact the Customer Service Team at South and Middle America, USA and Canada - Phone. +1 901.386.2246Asia Pacific, Australia and New Zealand - Phone. +86 10 8080 9107All other countries - Phone. 49 9128 607 - 38 |  | 124 | 124 units | Worldwide Distribution - US Distribution to the states of AZ and OK., and the countries of Germany, Austria, France, UK, Poland, Finland, The Netherlands, Spain, Turkey, Russia, and South Africa, Algeria, North and South America, Mexico, Brazil, Venezuela, Colombia, Panama, Asia, China, Japan, Philippines, Korea, India, Singapore, Thailand, Taiwan, Hong Kong, Malaysia and Guam. | Terminated | August 18, 2015 | 197 | 124 | Not_Computer
Z-1200-2015 | 70322 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | NAMIC Convenience Kit, 3V WMMII KIT, UPN H965601852521, REF/Catalog No. 60185252, STERILE, Rx ONLY --- For single use only. --- Device Listing Number D136588. |  |  | January 09, 2015 | 2015 | Navilyst Medical, Inc | The labelled expiration dating of the affected product is not fully supported by shelf life verification testing, which may result in a sterility breach. Use of a Fluid Management Convenience Kit, where sterility has been breached, may result in an increased risk of infection. | Packaging | Navilyst Medical sent an Urgent Medical Device Recall Notification letter dated January 9, 2015 to US customers (end users) via Federal Express. On January 20, 2015, Recall Notification Packages were sent via Federal Express to 2 International consignees (distributors). the letter identified the affected product, problem and actions to be taken. Customers are instructed to: " Immediately remove the recall product from your inventory (whether in labs, Central Supply, Shipping and Receiving or ANY other location)." Segregate this product in a secure location for return to Navilyst Medical." Immediately forward a copy of this recall notification to all sites to which you have distributed affected product." Complete and return the Reply Verification Tracking Form. |  | 7 | 7 units | Worldwide Distribution - US Nationwide and the countries of: Canada, and Guam. | Open, Classified |  |  | 16508 | Not_Computer
Z-1215-2015 | 70291 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | The Brainlab Offset Cup Impactor Universal. Part of the Brainlab hip system. An intraoperative image-guided localization system to enable minimally invasive orthopedic surgery. (a stereotaxic instrument) |  |  | January 16, 2015 | 2015 | Brainlab AG | The recommended sterilization and drying parameters are not effective to achieve the minimum required sterilization level. | Reprocessing Controls | FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letters dated December 11, 2014 were sent to all direct accounts (customers). The letters included instructions for customers to utilize the sterilization cycles included in the letter. Customers were further advised to alert users of these devices to the sterilization process changes. Customers with questions can contact Brainlab by telephone at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 114 | 114 offset cup impactors | Worldwide Distribution - US: Nationwide (AZ, CA, ID, IL, MD, MI, MN, OH, OK, OR, TN, WA, WI, WV); Australia, Austria, Bahrain, Belgium, Germany, Hong Kong, India, Italy, Japan, Malaysia, Russian Federation, Saudi Arabia, South Korea, Spain, Switzerland, Thailand, United Arab Emirates, United Kingdom | Open, Classified |  |  | 114 | Not_Computer
Z-1214-2015 | 64784 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals, including radiologists, physicians, technologists, clinicians and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display, manipulate, archive, print and export images when connected with the Centricity PACS infrastructure. |  |  | February 18, 2013 | 2013 | GE Healthcare | Using Merge Exam in single Study Mode may result in Missing Study Record (Cannot Display Exam, Send Exam, etc). Exam merge in Single Study Mode may fail on &amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp | Software design | Consignees were notified of the issue via telephone and site visit by GE Healthcare IT Field Engineers. |  | 128 | 128 systems | Distributed US (nationwide) including the states of AK, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, ME, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI, and WV, and the country of Bermuda. | Open, Classified |  |  | 128 | Not_Computer
Z-1213-2015 | 70456 | toothbrush, manual | EFW | Manual toothbrush. | Dental | Dental | 510(K) Exempt | Manual Toothbrush; SilverCare Junior 2/6 years Antibacterial Toothbrush Made in Italy |  |  | February 03, 2015 | 2015 | Norwex Usa Inc | The toothbrush head can potentially disengage from the handle. | Under Investigation by firm | Norwex sent an Urgent Product Recall letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were instructed to call 1-866-450-7499. For questions regarding this recall call 204-647-0798. |  | 35214 | 35,214 | Nationwide Distribution | Open, Classified |  |  | 35214 | Not_Computer
Z-1044-2015 | 69870 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Multi Diagnost Eleva II |  |  | October 14, 2013 | 2013 | Philips Medical Systems, Inc. | The possibility of the Reference Air Kerma Rate for low dose mode will not be reduced to 50% of the value for the normal mode, but will be approximately 70-80%, leading to a higher patient dose than expected. | Radiation Control for Health and Safety Act | Philips Healthcare Planned Action: 1. Software updates to SW R6.1.3 through Field Change Orders (FCOs) FCO70800140 and FCO70800142.2. A customer notification letter which includes a statement that Philips Healthcare will without charge, remedy the defect or bring the product into compliance.1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Philips Healthcare will implement this CAP by September 1, 2015.For further questions please call (978) 687-1501. |  | 238 | 238 | Worldwide Distribution - USA ( nationwide ) to the states of : AR, AZ, CA, CO, DE, FL, GA, HI, IL, IA, KY, KS, LA, MI, MO, MS, NC, NY, OH, PA, TN, TX, VA, VT, WA, WY and WV., and the countries of Austria, Australia, Colombia, Germany, Switzerland, Turkey, Ireland, Dubia and United Kingdom. | Open, Classified |  |  | 238 | Not_Computer
Z-1224-2015 | 70224 | No | CBN | Gas-scavenging apparatus. | Anesthesiology | Anesthesiology | 510(k) | Servo-I ventilator system CO2 analyzer. Part number 65 23 588. |  |  | January 12, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | A calibration problem was detected in a specific number of the Servo-I ventilator system CO2 analyzer models that may have been caused by the module's internal circuit board function. | Device Design | Maquet sent an Urgent Product Recall Medical Device Action Field Removal Letter dated 1/12/2015 via FedEx. |  | 486 | 486 units | Worldwide distribution - USA nationwide, Austria, Australia, Belgium, Chile, Canada, Colombia, Ecuador, Egypt, Finland, France, Germany, Hungary, Iceland, India, Indonesia, Iraq, Italy, Japan, Libya, Netherlands, Norway, Panama, Peru, Portugal, Romania, Russia, Saudi Arabia, South Africa, Spain, Sweden, Thailand, Tunisia, Turkey, Ukraine, and United Kingdom. | Open, Classified |  |  | 486 | Not_Computer
Z-1219-2015 | 70500 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Segmental System (ZSS) Cemented Stem / ZSS Cemented Stem, Smooth Stem |  |  | February 05, 2015 | 2015 | Zimmer, Inc. | PMMA coating does not meet specifications. Area of the stem may be missing PMMA coating. | Process control | Zimmer sent an URGENT MEDICAL DEVICE RECALL - Lot Specific letter dated February 5, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Your Responsibilities1. Review the notification and ensure affected personnel are aware of the contents.2. Locate and remove from inventory the affected products identified above.3. Return any affected product with missing PMMA coating per the PER process4. Return the Notification Acknowledgment Form (Attachment 1) to corporatequality.postmarket@zimmer.com.5. Please notify Zimmer if the hospital that you have distributed the affected product to has implanted the product. Inaddition, identify the surgeons that have used this product.6. If after reviewing this notification you have further questions or concerns please call the customer call center at1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 1 | 1 unit | US Distribution to the state of TX | Terminated | September 23, 2015 | 230 | 1 | Not_Computer
Z-1216-2015 | 69970 | No | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | RotoRest Delta Advanced Therapy System. |  |  | December 03, 2014 | 2014 | Arjo Hospital Equipment AB | The recalled devices labeling and instructions for use contain unapproved medical claims. | Labeling design | Devices in the rental fleet will be delivered to the customer with revised documents including the updated information. The customers who purchased the devices will be notified of the recall via letter sent by registered mail. The field correction notice will include a recall letter, response form, and updated labeling. |  | 271 | 271 units | Worldwide Distribution. US nationwide, Germany, Austria, France, Italy, India, EU, Mexico, Switzerland, Canada, Saudi Arabia, Kuwait, Middle East, China, Japan, and Qatar. | Terminated | October 30, 2015 | 331 | 1064 | Not_Computer
Z-1223-2015 | 70395 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | Ikaria, INOmax DSIR (Delivery System), Model 10007. Nitric oxide delivery system for use with ventilators. |  |  | January 14, 2015 | 2015 | INO Therapeutics (dba Ikaria) | Potential delivery failure alarm condition. INOmax DSIR with software version 2.0.4 and a certain variant of the monitor display may trigger a delivery failure alarm when the display brightness is set to its lowest level. If this condition occurs, therapy will be interrupted. | Process change control | Consignees were sent on 1/14/2015 an Ikaria "Urgent Field Correction Action" letter dated January 14, 2015. The letter described the problem and the product involved in the recall. The letter provided a number of recommendations to the consginees in regard to use and return of the product. For questions or adverse events they can contact Ikaria Customer Care at 1-877-KNOW-INO (1-877-566-9466) 24 hours a day, 7 days a week. |  | 45 | 45 units | Distribution US nationwide, including AZ, CA, CT, FL, ID, LA, MD, MI, MS, NH, NJ, NV, NY, OK, PA, TX, UT, VA, WV. | Terminated | April 15, 2015 | 91 | 45 | Not_Computer
Z-1226-2015 | 70441 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Vacuum Source Tubing Set for the EndoWrist Stabilizer for da Vinci S (IS2000) and Si (IS3000); |  |  | February 16, 2015 | 2015 | Intuitive Surgical, Inc. | There is the possibility that the Vacuum Source Tubing set for the Endo Wrist Stabilizer may have the outer pouch compromised during shipping, resulting in small pinholes which compromise the sterility of the pouches outer surface. | Under Investigation by firm | Urgent Medical Device Recall letters were sent on February 16, 2015 |  | 706 | 706 units | US and Belgium, Brazil, Canada, China, Finland, France, Greece, India, Italy, Japan, Qatar, Russia, Saudi Arabia,South Korea, Spain, Switzerland, Taiwan, Turkey, United Kingdom | Terminated | July 01, 2015 | 135 | 706 | Not_Computer
Z-1227-2015 | 70128 | dispenser, liquid medication | KYX | Liquid medication dispenser. | General Hospital | General Hospital | 510(K) Exempt | Baxter, Self-Righting Syringe Tip Caps, Self-Righting DISCPAC Syringe Tip Caps. Product codes H93866100, H93869100, H93866025, H93867025, H93867100, H938671025, H938671100, H938673025, H938673100, H938674025, H938674100, H938676025, H938676100, H938677025, H938677100, H938679025, H938679100, H938690025, H93869025, H938693025, H938696025, H93868025, H93868100, H938672025, H938672100, H938675025, H938675100, H938678025, and H938678100. |  |  | December 29, 2014 | 2014 | Baxter Corporation Englewood | Self-Righting Luer Tip Caps were manufactured and shipped with a missing latex symbol on both single and case labels. | Labeling Change Control | Baxter sent an Safety Alert letter dated January 15, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Baxter is requesting that you take the following actions:1.Complete the enclosed customer reply form and return it to Baxter by faxing it to 224-270-5457 or scanning and e-mailing it to fca@baxter.com. Returning the customer reply form promptly will prevent you from receiving repeat notices.2. Please forward a copy of this Safety Alert letter to other facilities or departments within your institutions to ensure that those locations are aware of this action.Action to be taken if you are an end-user and purchased product from a distributor or reseller:1. Please forward a copy of this Safety Alert letter to other facilities or departments within your institutions to ensure that those locations are aware of this action.2. Follow your supplier's reply process. Please do not return the customer reply form to Baxter.Action to be taken if you are a distributor:If you are a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, please notify your customers of this Safety Alert in accordance with your customary procedures. Further information and supportFor general questions regarding this communication, contact The Center for One Baxter at 800-422-9837, Monday through Friday, between the hours of 8:00 AM and 5:00 PM Central Time. |  | 1404230 | 1,404,230 | Nationwide Distribution | Open, Classified |  |  | 1404230 | Not_Computer
Z-1222-2015 | 70397 | No | BSY | Tracheobronchial suction catheter. | Anesthesiology | Anesthesiology | 510(K) Exempt | KimVent Wet Pak, Closed Suction System for Adults, Double Swivel Elbow, 14 Fr. Sterile, Distributed by Kimberly Clark Global Systems, LLC. |  |  | January 16, 2015 | 2015 | Halyard Health | Defect: A hole or crack was found which may cause air to leak. | Packaging process control | Consignees were notified by letter via FedEx on 01/16/2015. |  | 83 | 83 cases (20 tubes per case) | CA, FL, IL, LA, MA, MK, MN, MS, NM, NY, OK and TX and Japan. | Terminated | July 07, 2015 | 172 | 93 | Not_Computer
Z-1147-2015 | 70271 | No | LGW |  |  | Neurology | PMA | The Medtronic, Quadripolar ITREL 3 IPG, Model 7425, is a multiprogrammable device designed for use in the Stimulation of the Central Nervous System, Peripheral Nervous System, or muscles when connected to an extension and/or lead. |  |  | January 09, 2015 | 2015 | Medtronic Neuromodulation | Current labeling indicates that cycling improves device longevity in every programming scenario. However, enabling cycling at certain parameter settings may decrease the device battery longevity. With cycling enabled, longevity estimate displayed on the Model 8840Programmer may be inaccurate. | Other | Consignees were sent a Medtronic {Medical Device Correction" letter dated January 9, 2015. The letter was addressed to Healthcare Professional. The letter described the issue and the product involved in the the recall. The letter also listed the Occurrence of the problem and labeling update. A Physician Reply Form was included which requested consignees to return the form to Medtronics. For questions they can contact Medtronic Neuromodulation Technical Services at 1-8007070944 weekdays 7 am 6 pm Central time. A second "Physician Notification Confirmation Form" was included to be completed by Medtronic Representatives. |  | 396975 | 396,975 (258,805 USA, 138,170 OUS) | Worldwide Distributon- USA( nationwide including DC and PR) and Internationally to Mexico and Canada | Open, Classified |  |  | 396975 | Not_Computer
Z-1232-2015 | 70471 | No | JSK | Supplement for culture media. | Microbiology | Microbiology | 510(K) Exempt | Oxoid Legionella BCYE Growth Supplement SR0110A packaged in packs of 10- 100 ml vials. The responsible firm name on the label is Oxoid Ltd., Basingstoke, Hampshire, England and Oxoid Ltd., Basingstoke, Hants, England. |  |  | February 05, 2015 | 2015 | Remel Inc | Use of the product may not provide adequate recovery and could result in false negative identification of Legionella pneumophila. | Under Investigation by firm | ThermoFisher Scientific sent an Urgent Medical Device Recall letter dated February 5, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter informed the consignees of the recall and the associated risk of false negative identification of Legionella pneumophila. The consignees were instructed to notify any personnel who need to be aware of the potential for false negative id. of Legionella spp. Further, the consignees were instructed to discard any affected product remaining in their inventory and review all results reported using the recalled lot. The recalling firm requested the consignees complete and return the Medical Device Recall Return Response form acknowledging receipt of the notice and disposal of the affected product. If the affected lot was further distributed, the consignees were instructed to notify those entities of the recall and provide a copy of the recall notification letter.If you have any questions, please contact our Technical Services Department at 800.255.6730 (US) or 913.888.0939 ( International) |  | 24 | 24 packs (10 x 100 ml vials) | US Distribution to the states of : MI, ND, SC, TN, and TX. | Open, Classified |  |  | 24 | Not_Computer
Z-1120-2015 | 70011 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Medical System Allura Xper X-Ray Angiographic |  |  | June 06, 2014 | 2014 | Philips Medical Systems, Inc. | In certain circumstances, a software error can lead to a situation where the five minute fluoroscopy audible signal does not sound. | Radiation Control for Health and Safety Act | Philips Medical Planned Action: 1. The affected customers will be sent a customer notification letter that identifies the software corrective actions and indicates that these corrective actions will be provided to the end-user free of charge2. A mandatory Field Change Order with reference FCO72200270 will be released that requires Philips field service engineers to install Software release R7.2.8 which addresses the buzzer issue. The expected date of this FCO will be February 2015. The software release will contain updated timer function code to correct the intermittent fluoroscopy timer malfunction.3. Philips Healthcare Field Service Engineers will schedule an appointment with customers to install the software update.4. The software solutions will be rolled out consecutively, each within a maximum completion time of 6 months. The software releases will be corrected in descending order of the number of affected systems in the field, with approximately 80% of the total installed base corrected by 4 2015. |  | 7439 | 7439 | Worldwide Distribution - USA nationwide including Puerto Rico, Austria, Azerbaijan, Australia, Algeria, Argentina, Albania, Brazil, Belgium, Bangladesh, Bahamas, Bulgaria, Cambodia, Chile, Canada, China, Croatia, Czech Republic, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Reunion, Singapore, Slovenia, South Africa, Spain, Sudan, Sweden, Saudi Arabia, Switzerland, Syria, Taiwan, Thailand, France, Dominican Republic, Colombia, Turkey, Germany, Greece, Hong Kong, Denmark, Ecuador, Egypt, Estonia, Finland, Vietnam, Russian Federation, Iran, Iraq, Ireland, Israel, Italy, India, Indonesia, Japan, Jordan, Kenya, South Korea, Kuwait, Latvia, Lebanon, Lithuania, Macedonia, Malaysia, Mauritius, Mexico, Moldavia, Mongolia, Morocco, Nepal, Ukraine, United Arab Emirates, and United Kingdom | Open, Classified |  |  | 7439 | Not_Computer
Z-1237-2015 | 70434 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco radiation therapy treatment planning system. |  |  | December 19, 2014 | 2014 | Elekta, Inc. | Incorrect CT to ED when using the Monaco Image Statistics Tool. | Other | Elekta sent an Important User Notice on December 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This notice provided a workaround until new software releases. The notice included a acknowledgement form to be returned.For further questions please call (770) 300-9725 |  | 1354 | 1354 | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and to the states of : AL, AK, AZ, CA, CO, CT, FL, ID, IL, IN, KS, KY, LA, MD, MI, MN, MO, MT, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA and WI., and to the countries of Angola, Argentina, Austria, Australia, Bangladesh, Belgium, Bulgaria, Bahrain, Brazil, Canada, Switzerland, Chile, China, Colombia, Cyprus, Czech Republic, Germany, Algeria, Ecuador, Spain, Finland, France, Great Britain, Greece, Hong Kong, Ireland, Israel, India, Iraq, Italy, Jordan, Japan, South Korea, Kazakhstan, Sri Lanka, Lithuania, Libya, Morocco, Myanmar, Malta, Mexico, Malaysia, Namibia, Nicaragua, Netherlands, New Zealand, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Sweden, Singapore, Slovenia, Thailand, Turkey, Taiwan, Venezuela and South Africa. | Open, Classified |  |  | 1354 | Not_Computer
Z-1209-2015 | 70506 | Yes | OAD |  |  | Cardiovascular | PMA | IntellaTip MiFi Open-Irrigated radiofrequency Ablation Catheters: |  |  | February 03, 2015 | 2015 | Boston Scientific Corp | Reports of formation of char adherent to the proximal part of the distal tip electrodes | Device Design | Customer notification letters were sent by overnight mail on February 10, 2015. |  | 14 | 14 units | Distributed in Germany, Great Britain, and the Netherlands. | Terminated | March 21, 2017 | 777 | 14 | Not_Computer
Z-1210-2015 | 69939 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT X-ray imaging device |  |  | September 08, 2014 | 2014 | Xoran Technologies, LLC | Out of tolerance kVp, Unexposed (black) frames, and Out of tolerance Dose. | Radiation Control for Health and Safety Act | Xoran Planned Action :A Xoran Technologies service representative will re-measure and perform calibration to correct all affected systems for kVp and dose issues. For the image quality issue, the service representative will mark the black frame and then attempt to reconstruct an adequate image. CDRH also understands that you are continuing your investigation to determine the root cause and a more permanent corrective action for the 3 identified defects.1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. A customer notification letter which includes a statement that Xoran Technologies, LLC will without charge, remedy the defect or bring the product into compliance.3. Xoran Technologies, LLC will continue to evaluate why the kVp is not always within Xoran specifications, why the unexposed frames continues to occur, and why the dose is not always within Xoran specifications and identify the actions needed to resolve these issues by July 1, 2015. Xoran will communicate the results of this evaluation and any necessary actions to both customers and CDRH.Further questions please call (734) 663-7194. |  | 432 | 432 | Nationwide Distribution | Terminated | February 02, 2016 | 512 | 432 | Not_Computer
Z-1233-2015 | 70436 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | ARROW FiberOptix Intra-Aortic Balloon Catheter Kit, 40 cc 7.5 Fr |  |  | February 05, 2015 | 2015 | Arrow International Inc | Outer package product label incorrectly identifies the product as FiberOptix Flex IAB 7.5Fr rather than the correct FiberOptix IAB 8Fr | Labeling mix-ups | Arrow International notified Users, Distributors, Risk Managers by letter dated February 5, 2015. The letter explains the labeling error and requests immediately discontinue use and quarantine product. Customers were instructed to complete the Recall Acknowledgement Form and return by fax to 1-855-419-8507 whether they have or do not have the affected product. Customers with questions were instructed to contact their local sales representataive or Customers Service at 1-866-246-6990.For questions regarding this recall call 610-378-0131. |  | 165 | 165 units | Worldwide Distribution - USA (nationwide) and Internationally to Belgium | Open, Classified |  |  | 165 | Not_Computer
Z-1228-2015 | 70424 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | CareLink Pro |  |  | November 10, 2014 | 2014 | Medtronic MiniMed Inc. | The firm is informing customers of a software error that might result in an inaccurate display of the Temp Basal time and duration in CareLink Clinical reports with data uploaded from the MiniMed 620G or 640G insulin pump. | Software design | Medtronic MiniMed Inc. sent an Urgent Medical Device Correction letter in December 2014 to all affected customers. The firm states in their letter that patients can continue to use their MiniMed 620G and 640G insulin pump systems, and that there is no risk to patients caused by this issue. The firm instructs that when reviewing the CareLink Pro 4.0 or CareLink Clinicial reports from the MiniMed 620G and 640G insulin pumps not to use the Temp Basal graphic for clinical consideration.For questions regarding this recall call 818-567-4700. |  | 9 | 5 CareLink Pro 4.0 CD ROMs | Worldwide Distribution to Japan only.CareLink Clinical is for clinical trials only (US and International) | Terminated | March 10, 2015 | 120 | 9 | Not_Computer
Z-1229-2015 | 70470 | No | JLS | Progesterone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur Systems Progesterone Kit, (1 -pack) |  |  | January 27, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc | ADVIA Centaur Systems Progesterone Kit Lots Ending in 268 have a high bias on results greater than 30 ng/mL(95.4 nmol/L) | Process control | Siemens issued an Urgent Medical Device Correction (UMDC) was sent to affected Siemens Healthcare Diagnostics customers in the United States on January 27, 2015 to be delivered to customers on January 28, 2015. An Urgent Field Safety Notice (U FSN) was sent out for distribution to all customers outside the United States on January 27, 2015. These notices inform the customer of the potential to observe a high bias on patient sample with results greater than 30 ng/ml when using ADVIA Centaur Systems Progesterone kit lots ending in 268. To mitigate the bias, Siemens instructs customers to dilute samples with values greater than 30 ng/ml when using ADVIA Centaur Systems Progesterone kit lots ending in 268. In addition, Siemens informs customers to continue to follow Urgent Medical Device Correction 10819675, Rev. A and Urgent Field Safety notice 10819674, Rev A (Information Regarding the ADVIA Centaur Systems Calibrator E), dated September 2014.Fax-back forms to be returned to confirm that the customers received theUrgent Medical Device Correction. |  | 9778 | 9778 kits | Worldwide Distribution-US (nationwide), Austria, Bahrain, Belgium, Bulgaria, Brazil, Canada, Chile, Colombia, Croatia, Cyprus, Czech, Republic, Denmark, Egypt, Estonia, France, Georgia, Germany, Great Britain, Greece, Hong Kong, Hungary, India, Iran, Iraq, Ireland, Italy, Japan, Jordan, Kazakhstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Mexico, Morocco, Netherlands, Norway, Oman, Pakistan, Peru, Poland, Portugal, Qatar, Romania, Russian Federation, Serbia, Slovakia, Spain, South America, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and Vatican. | Terminated | March 31, 2016 | 429 | 23477 | Not_Computer
Z-1234-2015 | 70466 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Zip 8i Surgical Skin Closure Device; |  |  | February 05, 2015 | 2015 | Zipline Medical | The manufacturers seal may not be present on the Nylon pouch resulting in lack of assurance for sterility. | Packaging process control | Zipline sent an Urgent Voluntary product Recall letters to customers on February 5, 2015 by e-mail with possible follow up on site by firm representatives. The letter identified the affected product, problem and actions to be taken. Customers are asked to quarantine inventory immediately and examine the affected lots for complete pouch seals. Any product found with an incomplete seal should be destroyed or returned to ZipLine Medical. Please complete the enclosed response form as soon as possible, even if you no longer have stock of the lot numbers subject to this correction. This information is required to reconcile the inventory. Please contact Customer Service at 1-888-326-8999 to coordinate any product returns, replacement stock, or credit notes. |  | 3109 | 3109 - total - all model numbers | Worldwide Distribution - US Nationwide and the countries of Panama, Tehran, Ireland, Sweden, Japan, Hong Kong, Finland, Pakistan, Taiwan, Malaysia. | Terminated | April 10, 2015 | 64 | 3109 | Not_Computer
Z-1240-2015 | 70220 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | The I-Cart is used to store, transport, and power video equipment used in medical procedures. |  |  | December 10, 2014 | 2014 | Richard Wolf Medical Instruments Corp. | The transformers installed on carts used for Richard Wolf medical video systems may overheat causing smoke and black plastic material to leak from the cart chassis. | Device Design | Field Action Notice letters dated December 10, 2014 were sent to both direct accounts (customers). The letters instructed customers to: 1) Instruct the biomed to bypass the integrated power grid on the cart and not to use the electrical system on the affected carts until a qualified Omni Corporation service technician is able to inspect and repair the carts; and 2) Omni Corporation will schedule a qualified Omni field service technician, in conjunction with Richard Wolf and the customer's biomed unit, to visit and replace the transformer on the cart. |  | 83 | 83 carts | Nationwide | Terminated | July 31, 2017 | 964 | 83 | Not_Computer
Z-1241-2015 | 70338 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LFT v40 Femoral Head; Catalog Number 6260-9-032; |  |  | February 03, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Three customer complaints were received for the same batch of V40 LFIT Vitallium femoral heads (part number 6260-9-032, lot 48681201) alleging that the femoral head couldn't be assembled with its corresponding V40 stem trunnion. | Under Investigation by firm | Branches/Agencies were notified of this action by e-mail on December 19, 2014 and were asked to quarantine affected devices. Notification Letters and Product Accountability Forms were sent to Hospital Risk Management and Chief of Orthopaedics via FedEx (with return receipt) by 2/3/2015. |  | 9 | 9 units | Distributed in the states of GA, MA, NC, NY &amp; PA. | Open, Classified |  |  | 9 | Not_Computer
Z-1242-2015 | 70590 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | LASSO NAV Duo Loop eco Catheters, Catalog No. D-1354-01, D-1354-02, D-1354-03, D-1354-04, D-1354-05, D-1354-06. |  |  | February 25, 2015 | 2015 | Biosense Webster, Inc. | Biosense Webster is recalling the LASSO NAV Duo Loop eco Catheters because of reports of difficulty retracting and damage to the spine cover. | Device Design | Biosense Webster sent an Urgent Voluntary Medical Device Recall letter dated February 24, 2015 , to all affected customers. The letter informed customers that Biosense Webster is recalling the LASSO NAV Duo Loop eco Catheters because of reports of difficulty retracting and damage to the spine cover. The letter informed the customers of the problems identified and the actions to be taken. Customers with questions were instructed to contact Biosense Webster sales representatives or call (866) 473-7823, Monday through Friday from 7am to 8pm EST. Customers with questions related to the Voluntary Field Removal Certification Form and its return, are instructed to call (909) 839-8468.For questions regarding this recall call 909-839-8500. |  | 464 | 464 units | Nationwide Distribution | Terminated | October 16, 2015 | 233 | 464 | Not_Computer
Z-1239-2015 | 70502 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical PESSARY KIT-CUBE#2, 1-3/8 in , 35 mm |  |  | February 09, 2015 | 2015 | CooperSurgical, Inc. | Incorrect size printed on the Milex Pessary Kit packaging label | Error in labeling | CooperSurgical sent an Recall Letter dated February 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter asked customers to contact CooperSurgical Product Surveillance Department if they had any questions.At CooperSurgical's expense, arrangements were made to replace any of the affected products that customers had on hand. CooperSurgical provided further instruction on the disposition of the product at the time of replacement.Questions contact CooperSurgical 203.601.5200. |  | 95 | 95 units | Nationwide Distribution | Terminated | January 15, 2016 | 340 | 95 | Not_Computer
Z-1243-2015 | 70464 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DynaLOC Breast Interventional Planning Software Versions 3.1 and 3.2 |  |  | January 21, 2015 | 2015 | iCAD, Inc. | A device malfunction may cause the biopsy needle to fail to reach the intended target. If location confirming radiographs are not taken and properly evaluated, biopsy site mislocation can result in non-target tissue being sampled rather than target tissue. May result in an incorrect diagnosis. | Device Design | iCAD notified Invivo of the recall via letter on January 21, 2015. Invivo sent a Customer Information letter dated January 23, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letters from both Invivo and iCAD (device manufacturer) recommend that DynaCAD Breast Interventional Software not be used with Invivo ds Breast 16 channel lateral grid plate part number 4535-303-31491. Invivo and iCAD are developing DynaCAD Breast Interventional 3.3 that corrects this error and will be available for installation beginning in March 2015.If you need any further information or support concerning this issue, please contact your local Invivo representative:1-877-Invivo1 (1-877-468-4861, option #3)Invivo apologizes for any inconveniences caused by this problem. |  | 153 | 153 devices | Worldwide Distribution - Nationwide Distribution to the states of : AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NY, OR, PA, SC, TN, TX, VA, VT, WA, WI, WV and W., and to the countries of : BELGIUM, GERMANY, GREAT BRITAIN, INDIA, ISRAEL, NETHERLANDS and SWITZERLAND. | Open, Classified |  |  | 153 | Not_Computer
Z-1244-2015 | 70433 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Sim radiation therapy treatment planning system |  |  | December 19, 2014 | 2014 | Elekta, Inc. | Incorrect CT to ED when using the Monaco Image Statistics Tool. | Software design | An Important User Notice (IUN 382-03-FCL-001) was issued to customers on 12/19/2014. The notice included a workaround to be used until a new software release. The notice also included an acknowledgement form to be returned. |  | 907 | 907 | Worldwide distribution. US nationwide (AL, AK, CA, CO, CT, FL, IL, IN, LA, MD, MA, MI, MO, MT, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA, WA, WY, DC), Puerto Rico, Argentina, Austria, Australia, Bulgaria, Bahrain, Brazil, Belarus, Canada, Chile, China, Colombia, Cuba, Cyprus, Czech Republic, Germany, Algeria, Egypt, Spain, France, Great Britain, Greece, Croatia, Indonesia, Israel, India, Iraq, Italy, Jordan, Japan, South Korea, Sri Lanka, Lithuania, Libya, Morocco, Macedonia, Mexico, Malaysia, New Zealand, Philippines, Portugal, Romania, Serbia, Russia, Singapore, Slovenia, Suriname, Thailand, Turkey, Ukraine, Venezuela, South Africa, Kosovo. | Terminated | August 07, 2017 | 962 | 907 | Not_Computer
Z-1245-2015 | 70194 | Yes | FIK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AMIA v1.0 APD System. For Automated Peritoneal Dialysis (APD) therapy for pediatric and adult renal patients. |  |  | April 03, 2014 | 2014 | Baxter Healthcare Corp. | Additional warning in the AMIA APD System labeling for patients vulnerable to unintended Increased Intraperitoneal Volume (IIPV), also known as overfill. The patients include neonates, infants, small children, and any patient with significant heart and/or lung disease. | Error in labeling | Baxter sent IMPORTANT PRODUCT INFORMATION letters dated April 3, 2014 to all peritoneal dialysis healthcare providers of record. The letters advised healthcare providers of the potential to develop Increased Intraperitoneal Volume (IIPV) and instructed the healthcare providers to complete and return the attached Customer Reply Form. Healthcare providers with clinical questions can contact Baxter's Renal Division Clinical Helpline at 888-736-2543, option 2, 8:00 AM - 5:00 PM CT. Healthcare providers with general questions can contact The Center for One Baxter by telephone at 800-422-9837, Monday - Friday, 8:00 AM - 5:00 PM CT. Additionally, an IMPORTANT PRODUCT INFORMATION letters dated April 8, 2014 were sent to all peritoneal dialysis patients of record. The letters advised patients of the potential to develop Increased Intraperitoneal Volume (IIPV) and instructed the patients to complete and return the attached Home Patient Reply Form. Patients with questions about the letter can contact The Center for One Baxter by telephone at 800-422-9837, Monday - Friday, 8:00 AM - 5:00 PM CT. |  | 55 | 55 systems | US: Nationwide in the states of (GA, MA, NC, NY, PA, WA) | Terminated | July 21, 2017 | 1205 | 55 | Not_Computer
Z-1246-2015 | 70507 | No | MBB | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | SMARTSET GHV Gentamicin Bone Cement. A self-curing, radiopaque, polymethyl methacrylate based cement containing antibiotic. It is used for securing a metal or polymeric prosthesis to living bone in arthroplasty procedures. The bone cement has no intrinsic adhesive properties but relies instead on close mechanical interlock between the irregular bone surface and the prosthesis. |  |  | February 11, 2015 | 2015 | DePuy Orthopaedics, Inc. | One lot of SMARTSET GHV Gentamicin Bone Cement is partially agglomerated . This may result in surgical delays while a replacement package of cement is retrieved and mixed. | Under Investigation by firm | DePuy Orthopaedics, Inc. sent an Urgent Information - Recall Notice dated February 6, 2015, to all affected consignees for a voluntary recall of one lot of SMARTSETÂż GHV Gentamicin Bone Cement because the affected cement is partially agglomerated (clumped together). Consignees were instructed to discontinue use of the affected lot of bone cement immediately and destroy according to the consignees hazardous materials protocol or return if Certified Dangerous Goods Shipper, holding an International Air Transport Association (IATA) certification. Customers were instructed to call 574-371-4917 (M-F; 8 am  5 pm EST).For questions regarding this recall call 574-371-4577. |  | 1267 | 1267 | Nationwide Distribution including CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MS, MT, NC, NV, NY, OH, OR, PA, TN, TX, UT, VA, WA, and WV. | Terminated | May 19, 2015 | 97 | 1267 | Not_Computer
Z-1249-2015 | 70477 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | EVOLUTION MP TIBIAL BASE, REF ETPKN8SR, SIZE 8 RIGHT, TYPE PRIMARY, FIT CEMENTED, STYLE KEELED, STERILE R, Rx ONLY, MicroPort Orthopedics, Inc. Knee prosthesis component. |  |  | February 02, 2015 | 2015 | Microport Orthopedics INC. | Some units of the EVOLUTION Tibial Base was affected by a casting tool issue which caused varying retaining wall thickness and keel offset in the finished product. | Device Design | MicroPort Orthopedics sent an Urgent Action Required letter dated January 29, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their stock including all consignment stock to determine if they have any affected product. If they have affected product they should return immediately to the MicroPort Orthopedics Inc. Distribution Center at 11481 Gulf Stream, Arlington, TN 38002. Mark all return shipping boxes with "RECALL" for identification and processing. Customers with questions were instructed to contact MPO's Customer Experience Department at 1-866-872-0211. For questions regarding this recall call 901-867-4324. |  | 17 | 17 units | Worldwide Distribution - USA including AL, AZ, ID, KS, KS, MO, TX, IA, WV and Internationally to Canada and The Netherlands. | Terminated | March 19, 2015 | 45 | 17 | Not_Computer
Z-1252-2015 | 70203 | No | MQM | Bilirubin (total and unbound) in the neonate test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | RAPIDPoint 405 Blood Gas Analyzer; Siemens Material Number: 10310464, 10322347, 10314817, 10317193, 10318999, 10320055, 10322347, 10328278, 10328302, 10336784. |  |  | December 30, 2014 | 2014 | Siemens Healthcare Diagnostics Inc | Neonatal bilirubin (nBili) parameter may have increased variability when the nBili concentration is greater than 12 mg/dl and the tHb concentration exceeds the upper reportable range of greater than 25 g/dl. | Software design | An Urgent Field Safety Notice was sent by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on December 30, 2014 both in the United States and elsewhere for communication with affected customers. In addition, all affected domestic customers were sent a hard copy of the Urgent Field Safety Notice via Federal Express on December 30, 2014. All affected customers outside the US will be provided a copy of an Urgent FieldSafety Notice via hard copy, e-mail, and/or fax as determined by each countrieslocal regulations and procedures. This notice informs consignees of the issue. Consignees are also asked to complete a field effectiveness check form and return it Siemens via fax (312) 275-7795. |  | 3420 | 3420 devices | Worldwide Distribution - US nationwide including Puerto Rico, Canada, Mexico, Albania, Algeria, Argentina, Austria, Bahrain, Belgium, Bosnia Herzegovina, Botswana, Brazil, Bulgaria, Chile, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Greece, Hungary, Iceland, Iraq, Ireland, Israel, Italy, Ivory Coast, Kazakhstan, Kuwait, Latvia, Lebanon, Lesotho, Lithuania, Luxembourg, Macedonia, Netherlands, New Caledonia, Norway, Peru, Poland, Portugal, Qatar, Romania, Russian Fed., Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, U.A.E., United Kingdom, Vatican City, Venezuela, White Russia/Belarus. | Terminated | August 02, 2017 | 946 | 10900 | Not_Computer
Z-1248-2015 | 70514 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Proteus XR/a with wall stand model number 2260354. |  |  | January 30, 2015 | 2015 | GE Healthcare | Potential breakage of the steel cable which supports the bucky device in certain wall stands of Proteus XR/a X-ray imaging systems. A fall of a bucky while the system is in use could result in an injury to a patient or operator. | Device Design | Consignees were sent on 1/30/2015 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref#17108 dated January 30, 2015. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology, and Radiologists/Cardiologists.The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact 800-437-1171. Outside the US contact your local GE Healthcare Service Representative. |  | 192 | 192 (131 US, 61 OUS). | Worldwide Distribution-US (nationwide) including the states of AK, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MD, MI, MO, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI,and WV, and the countries of ARGENTINA, CANADA, CHILE, CZECH REPUBLIC, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, IRELAND, ITALY, JAPAN, NEW ZEALAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SLOVAKIA, SPAIN, and UNITED KINGDOM. | Terminated | September 15, 2015 | 228 | 192 | Not_Computer
Z-1079-2015 | 70241 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Juno DFR X-ray system |  |  | December 14, 2014 | 2014 | Villa Radiology Systems LLC | The system does not provide the appropriate audible signal, permanent activation, and manual override, although the system is in high-level control functionality | Radiation Control for Health and Safety Act | Planned Action by the firm: 1. You will send out a customer notification letter that identified the software corrective actions and that indicates that these corrective actions will be provided to the end-user free of charge.2. You will implement a corrective software patch so that kV and mA manual adjustments from the generator console are inhibited, and consequently no high-level control modality is allowed. In AERC, high-level control is already disabled. A Field Service Engineer is to implement the software patch according to the corrective action plan1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFRThe system shall be modified upgrading the generator power firmware which will inhibit the high-level control bringing the system into compliance.This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. You will implement the corrective software patch to inhibit high-level control according to the corrective action plan by June 22, 2015.For further questions please 203.262-8836 |  | 30 | 30 units distributed in the US | US Distribution to the states of: AZ, AL, CA, CO, CT, DC, FL, IA, IL, IN, MI, MD, MS, MA, MO, NY, OH, PA, UT, TN and TX. | Terminated | March 16, 2017 | 823 | 30 | Not_Computer
Z-1251-2015 | 70432 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DuraDiagnost X-Ray system, Stationary X-Ray System Radiographic |  |  | January 23, 2015 | 2015 | Philips Medical Systems, Inc. | Improper installation of Tube arm, and Wall stand may cause it to fall down, Improper installation of table may cause it to float. | Process control | Philips Healthcare issued notice to foreign Distributors on 1/23/15 to issue and advise customers to stop using the device until a Philips Engineer checks the device. After the engineer verifies that the tube arm, Wall stand and table were properly assembled, the device may be used again.Philips will subsequently install a new Tube Arm and Wall stand fixing structure that is designed to prevent this hazard if the tube arm or Wall stand is incorrectly reinstalled in the future.Philips will subsequently provide the method to adjust the T-nuts of table to make sure it was properly assembled. A Philips Service Engineer will contact affected customers to install the solution. |  | 71 | 71 units | Worldwide Distribution. Argentina, Australia, Burkina-Faso, Chile, China, Colombia, Ecuador, France, Germany, Hungary, Indonesia, Iraq (NL), Kuwait, Libya, Macedonia, Malaysia, Mayotte, Poland, Saudi Arabia, Slovenia, Spain, Switzerland, Turkey, United Kingdom, and Uzbekistan. | Open, Classified |  |  | 71 | Not_Computer
Z-1247-2015 | 68235 | mattress, flotation therapy, non-powered | IKY | Nonpowered flotation therapy mattress. | General Hospital | General Hospital | 510(K) Exempt | Pressure IQ Evolve Mattress; Non-powered mattress replacement system. |  |  | September 01, 2014 | 2014 | Arjo Hospital Equipment AB | The firm received complaints regarding multiple mattresses were showing failures resulting in pressure ulcers. | Device Design | On 9/1/14 the recalling firm sent out a letter to their customers notifying them of an updated IFU for the recalled device. The letter also requested the customers to verify the quantity of the product they currently have on hand. |  | 1034 | 1,034 units | Nationwide Distribution. | Terminated | January 04, 2016 | 490 | 1034 | Not_Computer
Z-1258-2015 | 70287 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Healthcare Caps for capillary 50/60 uL, Package count 200 |  |  | January 22, 2015 | 2015 | Siemens Healthcare Diagnostics Inc | Capillary Caps for RAPIDLyte Multicap and Multicap-S Blood Collection Capillaries may not attach and result in breakage of the glass capillary | Nonconforming Material/Component | Siemens Healthcare issued an Urgent Field Safety Notice via e-mail to affected Siemens Healthcare Diagnostics Regional offices on January 22, 2015 both in the United States and elsewhere for communication with affected customers. Siemens advises customers to use caution whenapplying the Capillary Caps with Siemens Multicap glass and Multicap-S plastic capillaries to minimize the potential for breakage or bending. |  | 6184 | 6184 | Worldwide Distribution. US Nationwide, Australia, Austria, Belgium, Bosnia, Herzegovina, Canada, Colombia, Croatia, Curacao, Czech Republic, Estonia, Finland, France, Georgia, Germany, Great Britain, Hong Kong, Ireland, Italy, Japan, Jordan, Kazakhstan, Kyrgyzstan, Latvia, Lithuania, Mexico, Netherlands, New Zealand, Norway, Poland, Portugal, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates. | Terminated | April 13, 2016 | 447 | 49309 | Not_Computer
Z-1250-2015 | 70473 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | Apollo (cold) Laser Desktop Control Units, Model AP2-DT. |  |  | December 17, 2014 | 2014 | Pivotal Health Solutions | Control units were equipped with an internal mounting kit that does not meet medical safety standards, and are conductive, increasing the risk of electric shock to the user and patient. These units were manufactured prior to Pivotal Health Solutions acquisition of the Apollo product line. | Device Design | The firm, Pivotal Health Solutions, sent a letter dated December 17, 2014 on 12/17/2014 and an amended Pivotal "Urgent Medical Device Recall" letter dated February 9, 2015 to its consignees/customers. The letters described the product, problem and actions to be taken. The consignees/customers were instructed to stop using the unit and contact Pivotal Health Solution's at 1-800-743-7738 to arrange for return and repair asap; immediately examine your device inventory and quarantine any product subject to recall; if you have further distributed the product, identify your customers and notify them at once of this product recall, and complete and return the enclosed RETURN AUTHORIZATION FORM with the units to Pivotal Service Center, 1654 Mardon Drive, Dayton, OH 45432 and the DECLARATION OF CONTAMINATION STATUS form via Fax to: 605-882-8398.If you have any questions, contact Pivotal Health Solutions Service Repair Coordinator at 800-743-7738 or email Robin@PivotalHealthSolusitons.com. |  | 35 | 35 | US Distribution to states of: AZ, AR, CA, GA, IL, IN, IA, KS, MI, MN, MO, NY, OH, OR, PA, UT and WA. | Terminated | March 09, 2016 | 448 | 35 | Not_Computer
Z-1259-2015 | 70603 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Arrow Intra-Aortic Balloon Catheter Kit Fiberoptix (IAB), |  |  | November 08, 2013 | 2013 | Arrow International Inc | Potential for a leak between the catheter and the bifurcate which would allow helium to escape to the atmosphere. | Process control | Arrow International Inc notified accounts by letter dated 11/8/13 describing the problem and requested users to immediately discontinue use and return product. A Recall Acknowledgement form was to be completed and returned. |  | 166 | 166 units | Distribution US nationwide (including AZ, NJ, ME, TX, WA), and Germany. | Terminated | March 16, 2015 | 493 | 166 | Not_Computer
Z-1260-2015 | 70580 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Inserter for Titanium Elastic Nails; The Inserter for Titanium Elastic Nails (TEN) is intended for fixation of diaphyseal fractures where the canal is narrow or flexibility of the implant is paramount. This includes: lower extremity fractures in pediatric and small-statured patients and upper extremity fractures in patients. |  |  | February 03, 2015 | 2015 | Synthes, Inc. | The Inserter has the potential for mechanical failures such as breakage of the main shaft of the Inserter for the TEN, malfunction of the chuck of the Inserter for the TEN in the form of the chuck getting stuck/jammed, breakage of the cross bar of the Inserter for the TEN. | Device Design | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated February 3, 2015, to all affected customers to alert them about the recall and explain the benefits and risks associated with removing this device from the market. It stated that the immediate removal of the product would prevent the performance of emergency surgery, thus DePuy Synthes is not requiring an immediate removal of affected product(s). DePuy Synthes is in the process of developing the replacement &amp; recovery plan for indicated devices. Due to their current but temporary lack of replacement inventory, Synthes proposed to remove and replace the device in increments to prevent any adverse impact on the availability of the device for use on patients in trauma-related surgery events.DePuy Synthes will contact customers as replacements become available.Customers were advised to complete the Verification Section by checking the appropriate box indicating affected product has been located. Also,indicate the number of devices found. If your facility chooses to return the product(s) at this time, please take the following additional steps:Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, include a copy of Verification Section with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132.Send the completed Verification Section to the following fax or email address: Fax: (866) 782-7287 or Scan/email: Synthes5095@stericycle.com Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales Consultant.For questions regarding this recall call 810-719-5000. |  | 1550 | 1550 | Nationwide Distribution | Open, Classified |  |  | 1550 | Not_Computer
Z-1264-2015 | 70562 | stain, papanicolau | HZJ | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Carton: Richard-Allan Scientific CYTO-STAIN; REF 7501R; Quantity: 1x1 Gallon; Size: 1 Gallon/3.8L; LOT 313833; 11/2016; Bottle: Richard-Allan Scientific CYTO-STAIN G; REF 7501G; 1 Gallon/3.8L; Thermo Scientific. |  |  | February 19, 2015 | 2015 | Richard-Allan Scientific Company | The case was correctly labeled as REF 7501R, however the gallon bottle was incorrectly labeled as REF 7501G Cyto-Stain G. The bottle label has the correct lot number of 313833 and EXP date 11/2016. The gallon bottle contains the correct intended contents of 7501R: Cyto-Stain R Stain. | Error in labeling | Urgent Medical Device Recall letters sent to Richard Allan distributors and customers. The recall notification includes the reason for the voluntary recall, risk to health, product information and actions to be taken by the distributor/customer. |  | 132 | 132 | US, Canada and Europe | Terminated | January 12, 2016 | 327 | 132 | Not_Computer
Z-1263-2015 | 70264 | [6 empty] | VITROS Hand-held Barcode Scanner, Catalog Number 6844210, for use with the following VITROS Systems: VITROS 5,1 FS Chemistry System, VITROS 3600 Immunodiagnostics System, VITROS 4600 Chemistry System, VITROS 5600 Integrated System. This product is an accessory for use with VITROS Analyzer Systems. |  |  | December 29, 2014 | 2014 | Ortho-Clinical Diagnostics | In combination with the Batch Programming option (only) for programming samples, the hand-held barcode scanner unexpectedly skips sample cup positions within a sample tray. This issue only occurs when using the Batch Programming option; all other sample programming options function as intended. | Software design | On 12/29/2014, Customer letter (Ref. CL2014-315) was sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and provided with instructions for reconfiguring their Barcode Scanners. |  | 72 | Domestic: 62 units; Foreign: 10 units | Worldwide Distribution-US (nationwide) and the countries of Canada, Australia and England. | Terminated | March 27, 2017 | 819 | 72 | Not_Computer
Z-1262-2015 | 70544 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | - NTL-30 30ml Nutrio Enteral Feeding Syringe (sterile) |  |  | February 13, 2015 | 2015 | Medela Inc. EF Division | The firm is recalling sterile enteral feeding syringes due to determining that the sterility of the affected products may be compromised as a result of a potential breach in packaging. | Packaging | Medela sent a Device Recall letter dated February 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Requested Action of Distributors:1. Stop sale and distribution of product with affected lot numbers.2. Conduct a physical count and record the data on the Product Inventory Form included with the letter.3. Return the Product Inventory Form via fax to 800-995-7867 or e-mail it to syringe@medela.com. The firm should contain a contact name and signature.4. If have the affected product in stock, Medela Customer Service will issue a RMA and schedule a pick-up of the affected product. Contact will be made to ensure product is ready. For further questions please call (888) 399-2455. |  | 1959187 | 1,959,187 | Nationwide Distribution. | Terminated | November 09, 2015 | 269 | 1959187 | Not_Computer
Z-1266-2015 | 70523 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona Trabecular Metal Tibial Plate / Persona TM Tibia-Prosthesis, Knee, Patello/Femorotibial, Semi-Constrained, Uncemented, Porous, Coated, Polymer/Metal/Polymer |  |  | January 28, 2015 | 2015 | Zimmer, Inc. | An increase in complaints of loosening and radiolucent lines. | Under Investigation by firm | Zimmer Inc. is initiating a voluntary recall of Persona Trabecular Metal Tibial following an increase in complaints of radiolucent lines and loosening. All sizes and lots of the affected devices are being removed from distribution. Urgent Medical Device Recall notices were issued to affected distributors, hospitals, and surgeons on 2/16/2015 via mail. Customers are asked to review the notification and ensure affected personnel are aware of the contents. All affected product are to be located and quarantined immediately. The Inventory Return Certification Form should be completed and returned along with the recalled product. Customers may contact the following with any questions: 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 11658 | 11,658 | Worldwide Distribution -- US, including the states of AK, AL, AZ, CA, CO, FL, IL, IN, KS, MA, MI, MN, MO, NC, NJ, NY, OH, OK, PA, TN, TX, UT, VA, WA, and WI; Foreign:Canada, Australia, New Zealand, Korea, Austria, Belgium, Germany, Italy, Luxembourg, Netherlands, South Africa, Sweden, Switzerland, and United Arab Emirates. DOD/VA: HOSPITAL- VA MAINE HLTCR SYTM (402)VA MROC (402) ATTN: PROSTHETICS DEPT TOGUS ME 04330 207-672-9425HOSPITALVA MED CTR (438) 2501 W 22 SIOUX FALLS, SD 57105 605-336-8922HOSPITALVA MED CTR (509)1 FREEDOM WAY AUGUSTA , GA 30904 706-721-6892HOSPITALVA MED CTR (521) 700 S 19TH ST, BIRMINGHAM, AL 35233 378-962-1599HOSPITALVA MED CTR (523) 1400 VFW PARKWAY, WEST ROXBURY, MA 02132 617-862-5977HOSPITALVA MED CTR (528) 3495 BAILEY AVENUE BUFFALO, NY 14215 716-898-6922HOSPITALVA MED CTR (552) 4100 WEST 3RD STREET BLDG # 330 RECEIVING, DAYTON, OH 45428 937-226-5987HOSPITALVA MED CTR (583), 1481 WEST 10TH STREET INDIANAPOLIS, IN46202 317-267-8765HOSPITALVA MED CTR (595) 1700 SOUTH LINCOLN AVENUE, LEBANON, PA17042 717-272-6621HOSPITALVA MED CTR (657) 915 N GRAND BLVD, ST LOUIS, MO63106 314-652-4100HOSPITALVA MED CTR (664) 350 LA JOLLA VILLAGE DR SAN DIEGO, CA92161 858-552-8585HOSPITALVA MED CTR (693) 1111 EAST END BOULEVARD, WILKES BARRE, PA 18711 717-824-3521HOSPITALVA SNHCS (654/118) 975 KIRMAN AVENUE RENO NV89502 702-597-1699 | Open, Classified |  |  | 11658 | Not_Computer
Z-1267-2015 | 70556 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SIEMENS SOMATOM Force with software version VA50A and /or VA50A_FP1 and/or VA50A_SP0; Computed tomography system. |  |  | February 06, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Possibility of image artifacts during data acquisition when using Adaptive Cardio Sequence, Turbo Flash, and Head modes. The update improves system start-up behavior and resuming, ECG handling with visual feedback of correct placement of ECG leads and contact quality, and other safety related issues. | Software design | Siemens sent a Customer Advisory Notice dated February 6, 2015, to all customers via USPS Certified Mail. The notice identified the affected product, problem and actions to be taken. |  | 12 | 12 | US Nationwide Distribution in the states of PA, MA, MN, NY, SC, MD, NY, IL, OH, TX and KY. | Terminated | August 31, 2015 | 206 | 12 | Not_Computer
Z-1265-2015 | 70524 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | MAX-TL 9.0 x 7.0 mm Implant, Catalog No. 15711K; Z-MAX Implant 9.0 x 7 mm Length, Catalog No. Z-MAX9-7. Dental implant. |  |  | January 23, 2015 | 2015 | Southern Implants, Inc | Z-MAX Implant, 9mm diameter, 7mm length labeled package, catalog number Z-MAX9-7, contained a MAX-TL Implant, 9mm diameter, 7mm length. And vice versa. | Mixed-up of materials/components | A customer notification letter dated 1/23/15 was sent to customers to inform them that Southern Implant is recalling the MAX-TL 9.0x7.0mm Implant because it may contain a Z-MAX9-7 Implant, 9mm diameter, 7mm length. Customers are instructed to return the fax back form and to contact customer relations depart at (866) 902-9272. |  | 130 | 130 units | US Nationwide Distribution. | Terminated | May 20, 2015 | 117 | 130 | Not_Computer
Z-1278-2015 | 70270 | applicator, absorbent tipped, sterile | KXG | Absorbent tipped applicator. | General Hospital | General Hospital | 510(K) Exempt | Puritan absorbent foam tipped applicator; Catalog Number 25-1406 1 PF BT. |  |  | December 22, 2014 | 2014 | Puritan Medical Products Co., LLC | Shaft breaks so that adequate patient specimen sample cannot be obtained. | Process control | Consignees were initially notified of the recall via phone on 12/23/2014 and 12/29/2014. They were informed that the product may have been over exposed to gamma radiation, resulting in brittle sticks. They were instructed to not use the product for patient sampling and to return any product remaining in inventory to Puritan. |  | 4500 | 9 cases of 500 applicators each | Distributed in the states of MS, ME, GA, and NC. | Terminated | March 18, 2015 | 86 | 4500 | Not_Computer
Z-1269-2015 | 70512 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MiniCap with Povidone-Iodine Solution, item number 5C4466P, for hemodialysis. |  |  | January 22, 2015 | 2015 | Baxter Healthcare Corp. | Product may have separating or protruding sponges | Under Investigation by firm | Baxter sent an Urgent Product Recall letter dated January 22, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Baxter is requesting that you take the following actions:1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product package or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit. Center for Service can be reached at 888-229-0001, Monday through Friday, between the hours of 7:00 AM and 6:00 PM Central Time. Please have your Baxter eight-digit ship-to account number ready when calling.3. If you received your product directly from Baxter, please complete the enclosed customer reply form, and return it to Baxter by either fax or scanned e-mail.1. If you ordered product from a dealer, wholesaler, or distributor/reseller, please follow your supplier's reply and recall process.2. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product package or shipping carton.3. Contact Baxter Healthcare Center for Service to arrange for return and credit.Center for Service can be reached at 888-229-0001, Monday through Friday, between the hours of 7:00 AM and 6:00 PM Central Time. Please have your ship-to account number ready when calling.4. Please do not return the Baxter customer reply form to Baxter. Reply forms should be returned to your supplier.For clinical questions, contact Baxter's Renal Clinical Helpline at888-736-2543, option 2, Monday through Friday, between the hours of8:00 AM and 5:00 PM Central Time.For general questions regarding this communication, contact The Center for One Baxter at 800-422-9837, Monday through Friday, between the hours of 8:00AM and 5:00 PM Central Time. |  | 4391520 | 4,391,520 units | Worldwide Distribution - Nationwide Distribution and to the countries of : Canada, Austria, Norway, Portugal, Saudi Arabia and Spain. | Open, Classified |  |  | 4391520 | Not_Computer
Z-1275-2015 | 70430 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | BD Posiflush Saline Flush Syringe 10 ml inlcuded in the following device kits: Groshong Midline Bard MYPICC Kit 4F (Single-Lumen) |  |  | August 08, 2014 | 2014 | Bard Access Systems | Some unit packages of BD PosiFlush SF Saline Flush Syringe 10 mL may exhibit open seals, which could impact product sterility. Bard includes these pouched syringes, attached to some configurations of PICC kits, Port Access kits (PAK), PowerGlide kits, Midline kits, and Convenience kits. | Nonconforming Material/Component | The firm, BD Medical, sent an "URGENT VOLUNTARY PRODUCT RECALL" letter dated July 22, 2014 and a follow-up "URGENT: Voluntary Product Recall-BD PosiFlush SF Saline Flush Syringe 10mL with rework instructions for applicable Bard Kits" letter dated August 8, 2014 to its customers and distributors. The letters described the product, problem and actions to be taken. The customers were instructed to immediately review your inventory and quarantine product subject to the recall; complete and return the Recall Response Card form via fax to BD at 1-201-847-6990 or email to Becky_Saggau@bd.com and return all affected products. Note: If you do not have any of the affected lots in your inventory, please complete the Recall Response Card form indicating you have (0) quantity and fax or email.The distributors were instructed to immediately examine your inventory and identify any product subject to this notification. In addition, if you have further distributed this product, please identify your customers and notify them at once of this notification. Complete the attached Recall Reply Form and FAX to Bard Access Systems, Inc. at 1-801-522-5674, or EMAIL to BASD.fieldaction@crbard.com. Even if you do not have any inventory of the affected syringe lots.Please direct any return requests and questions to our coordinator at 800-290-1689. If you have any questions or require assistance with the return of the recalled product and/or availability of replacement product, please contact BD Customer Service at 1-888-237-2762 option 3. |  | 48 | 48 units | US Nationwide Distribution. | Terminated | August 05, 2015 | 362 | 290671 | Not_Computer
Z-1268-2015 | 70348 | No | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | ML 7090 Calibrated LASIK Blade (CLB); CLB CALIBRATED LASIK BLADE 10. |  |  | January 13, 2015 | 2015 | Med-logics Inc | The ML 7090 Calibrated LASIK Blade (CLB) will not fit into the AMO Amadeus Metal Blade Holder. | Process design | The firm notified their consignees by letter on 1/13/15, sent via FedEx. The US customers were instructed to return any remaining product to MED-LOGICS and complete the response form. |  | 360 | 360 units | Distributed in the states of CA, AL, GA, and TN, and the countries of Chile and Germany. | Terminated | October 28, 2015 | 288 | 360 | Not_Computer
Z-1279-2015 | 70673 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Millennium VG, Discovery VH and VariCam Nuclear Medicine systems, all configurations. |  |  | March 05, 2015 | 2015 | GE Healthcare | GE Healthcare has recently become aware of a potential safety issue due to a collimator falling off a VG system during a Quality Control check, while the detectors were rotating. No injuries have been reported to date | Use error | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref# 40865 dated March 5, 2015. The letter was addressed to Managers of Nuclear Medecine, Hospital Administrators / Risk Managers and Managers of Radiology /Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact in the United States: 1-800-437-1171. For other countries, contact your local GE Healthcare Service Representative. |  | 472 | 472 (127 units US, 345 units OUS) | Worldwide Distribution - US Nationwide in the states of: AL, AZ, CA, CO, CT, FL, GA, ID, IL, IN, IA, LA, ME, MD, MA, MI, MN, MS, NE, NY, NC, OK, PA, PR, SD, TN, TX, VA, WA, WV, WI. and the countries of URUGUAY, UNITED KINGDOM,UKRAINE, TURKEY, THAILAND, TAIWAN, SWEDEN, SPAIN, SLOVAKIA, RUSSIA,PORTUGAL, POLAND, PERU, NORWAY, NETHERLANDS, MEXICO, REPUBLIC OF KOREA, JAPAN, ITALY, ISRAEL, INDIA, GREECE, GERMANY, FRANCE, ESTONIA, DENMARK, CZECH REPUBLIC, CYPRUS, COLOMBIA, CHINA, CANDA, BRAZIL, BELGIUM, AUSTRIA, AUSTRALIA, ARGENTINA. | Terminated | August 17, 2015 | 165 | 472 | Not_Computer
Z-1277-2015 | 70143 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medstream 81" (206 cm) 20 drop Universal Administration Sets. |  |  | October 10, 2014 | 2014 | US Infusion Inc dba Trucare Biomedix-USA | Several complaints regarding the IV sets leaking at the Y site. | Unknown/Undetermined by firm | On December 10, 2014 the firm sent an URGENT DRUG RECALL to its consignees. Instructions were to immediately examine inventory and quarantine all product subject to the recall. Customers were asked to complete and return the response form. For returned goods instructions contact Roberta Parness at 866-593-8444 or via email at rparness@tcbiomedix.com. |  | 62700 | 62,700 IV sets | Distributed in the states of TX, IL, MA, OH, CA, and GA. | Terminated | June 23, 2015 | 256 | 62700 | Not_Computer
Z-1261-2015 | 70558 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Autostainer Link 48 (AS480), Autostainer Plus Link Instrument (AS100), Autostainer (S3400), and Autostainer Plus (S3800). Automated slide stainer for in vitro diagnostic use. |  |  | February 12, 2015 | 2015 | Dako North America Inc. | A false negative result affecting the diagnosis may occur. A defect in the syringe assembly stopcock component may, in some circumstances, cause leakage of excess buffer onto slide location 34 on the AS480, or slide locations 35 and 36 on the AS100, S3400, and S3800. Affected dates are from 12/2013-08/2014. | Nonconforming Material/Component | The firm, Dako, sent an "Urgent Medical Device Recall Notification" letter dated 2/12/15 to customer on 02/12/15.The firm's letter advised customers to do the following:The firm stated that customers can continue to use the instrument however should take the following actions:- Discontinue to use the HercepTest on slide location 34- As set forth in Dako's staining protocol, use on-slide positive controls and two drop zones- Watch for any indication of buffer leakage from they syringe assembly. This can be done easily when the tubes are checked for air bubbles in connection with exchange of the buffer bottle- Check for weak staining in regards to HercepTest in slide position 34.Dako's Actions:Dako states that a representative will contact the customers within the next few weeks in order to determine whether the HercepTest has been used for staining slides in position 34.The firm has identified a longer term corrective action plan of installing a drip tray under the stopcock and states that a representative will contact the customer when it is available.Transmission of Notice:The firm requests that customers inform those that need to be aware of the notification and to where the instruments might have been transferred to. The firm states that customers should ensure the organizations maintains awareness of the noticed and recommended steps until the corrective actions have been completed.The firm also asks customers to complete, sign, and return the attached recall form as soon as possible. On 02/23/15 the firm sent out a second notification letter to customers with the following additional recommendations:- Discontinue the use of HercepTest in slide location 34, 35, and 36 depending on the make and model of the instrumentIn addition the firm recommended that customers pay special attention to any indication that leakage has compromised staining.-When staining using the HercepTest protocols, excess leakage from the sy |  | 409 | 409 | Worldwide Distribution: US (nationwide) and countries of: AT, AU, BE, CA, CH, DE, DK, ES, FR, GB, IE, IT, JP, NL, NO, PL, and SE. | Terminated | February 11, 2016 | 364 | 409 | Not_Computer
Z-1282-2015 | 70384 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens LANTIS Oncology Information System Servers; allows the radiation therapist to deliver treatment to patient using the MEVATRON and all available accessories |  |  | January 21, 2015 | 2015 | Siemens Medical Solutions USA, Inc | There is a potential safety risk when using LANTIS server software with operating systems with which it has not been validated or released which can lead to an incorrect treatment to the patient. | Use error | Siemens sent an Customer Advisory Notice, dated January 21, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were informed the risk associated with use of more recent operating systems with LANTIS Server Software. Customers were warned against the use of more recent operating systems to ensure the LANTIS system remains in the validated configuration.In the interest of safety, we ask you to perform the above preventive measures and inform all affected personnel immediately.Please include this Field Safety Notice in your System Owner Manual chapter "Safety Advisory Letters" where it should remain.The relevant National Competent Authority will be informed of this Field Safety Notice.Please forward this safety information to other organizations that could be affected by this measure.If the device has been sold and therefore no longer in your possession, please forward this customer notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this notice until the update has been fully completed.We apologize for any inconvenience and thank you for your cooperation.For further questions please call (610) 219-6300. |  | 87 | 87 | Nationwide Distribution | Terminated | July 30, 2015 | 190 | 87 | Not_Computer
Z-1280-2015 | 70546 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Passport V Monitor, Multiparameter Patient Monitor(with Arrhythmia Detection and Alarms) |  |  | January 23, 2015 | 2015 | Mindray DS USA, Inc. dba Mindray North America | An issue has been identified with Passport V Monitors invasive blood pressure function (IBP) which may provide an incorrect IBP measurement | Process control | Mindray DS USA Inc sent Urgent: Passport V Monitor Field Corrective Action letters dated 1/23/2015 to hospital administrators. The letters informed customers that the IBP module may provide an incorrect IBP measurement. The letter also explained the adverse effects on patients. Customers are asked to either calibrate the IBP following the instructions provided; or, contact a Mindray Service Representative, who can perform the calibration, at 1-800-288-2121. In the case of the second option, it is recommended that the monitor be removed from service until the calibration is performed. If the first option is chosen, customers are to notify a Service Representative upon completion of the action. Customers can direct their questions to Diane Arpino at 201-995-8407. |  | 519 | 519 units | Nationwide Distribution | Open, Classified |  |  | 519 | Not_Computer
Z-1283-2015 | 70341 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Distal Capture Assembly |  |  | February 12, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Potential for weld disassociation of the cross pin, which could lead to loose or disassociated trigger mechanism on the Distal Capture Assembly. | Device Design | On March 20, 2015, Stryker sent Urgent Medical Device Recall Notifications/Urgent Medical Device Recall Notification Acknowledgement Forms dated March 19, 2015 to all affected customers. The letter identified the product, the related issue and the action to be taken by the customers. Customers were asked to return the attached acknowledgement form confirmed they received and reviewed the notification. Customers with questions were instructed to call 201-831-6365. Stryker sent an Urgent Medical Device Correction letter dated February 11, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to return the attached acknowledgement form confirming they have received and reviewed the letter. Customers with questions were instructed to call 201-831-6365. |  | 2030 | 2030 units | Worldwide Distribution - USA (nationwide) Puerto Rico, and Internationally to Canada, Hong Kong, Netherlands, Sweden, Switzerland, Spain, France, Italy, UK, India, Australia, Korea, Panama, Polan, Ireland, Kuala Lumpar, New Zealand and South Africa. | Open, Classified |  |  | 2030 | Not_Computer
Z-1281-2015 | 70455 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Thumb Screw for the Synthes Recon Locking Aiming Arm for Lateral Entry Femoral Nails-EX (part number 03.010.048) and the Thumb Screw for the Aiming Arm for Titanium Cannulated Tibial Nails-EX (part number 03.010.052). |  |  | February 17, 2015 | 2015 | Synthes, Inc. | For certain lots, the incorrect raw material of annealed 17-4PH Stainless Steel was used instead of the correct material of heat-treated 304 Stainless Steel. This may lead to stress-corrosion cracking, which can lead to surgical delay or unintended debris entering the surgical site. | Process control | An Urgent Notice: Medical Device Recall, dated February 17, 2015, was sent to all direct consignees and sales consultants instructing them to take the following actions:Take note of the quantity of affected parts in your possession, and then follow ONE of the options for replacement of the affected product:Option 1: Call DePuy Synthes Repair Department at 1-800-288-6698 to schedule a repair for the affected products.Complete the Verification Section following the instructions provided and send a copy of the completed Verification Section to DePuy Synthes by: Fax: 866-723-2823 or Scan/e-mail: Synthes6734@stericycle.comOption 2:Call DePuy Synthes Repair Department at 1-800-288-6698 to order repair part. The replacement part will be sent directly to your facility.Questions should be directed to 610-719-5450 or a local DePuy Synthes sales consultant. |  | 183 | 183 | Nationwide Distribution | Open, Classified |  |  | 183 | Not_Computer
Z-1289-2015 | 70385 | No | CFN | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | N Latex IgM; In-vitro diagnostic reagent for the quantitative determination of IgM in human cerebrospinal fluid (CSF) and in paired CSF/serum samples by means of particle-enhanced immunonephelometry using the BN II and BN ProSpec Systems. The determination of IgM aids in the evaluation of the patient's immune system. |  |  | January 15, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Complaints have been confirmed on failed proficiency testing due to serum samples recovering too low when using some N Latex IgM reagent kit lots 44018, 44019 and 44042. | Nonconforming Material/Component | The firm, Siemens, sent an "Urgent Medical Device Recall" letter dated January 2015 to its customers on January 13, 2015. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: Discontinue use of the affected lots and discard any remaining material; Identify if IgM quotient was overestimated due to falsely low IgM results with serum sample; Reconsider patient diagnosis taking other patient data into consideration.; When other patient data does not support the result for an intrathecal IgM synthesis or inflammation, re-measurement of samples may be considered using and alternative lot. In addition, Siemens requested customers to review this letter with their Medical Director; to complete and return the attached Effectiveness Check/Product Replacement form via fax to Customer Care Center at (312) 275-7795; retain this letter with your laboratory records, and forward this letter to those who may have received this product. If you have any questions, please contact your Siemens Customer Care Center 800-441-9250 or email: aydee.I.crawford@siemens.com or your local Siemens technical support representative. |  | 26 | 26 | Worldwide Distribution: US distribution in states of: UT , CO, CA , FL , NY , OH , NC and country of : Canada. | Open, Classified |  |  | 26 | Not_Computer
Z-1285-2015 | 70535 | No | IOL | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | GE Healthcare, Treadmill T2100 and Treadmill T2000 with power cord part numbers 408930-001 Rev E, 408930-002 Rev C, or 2028038-008 Rev B. Also Field Replaceable Unit P/Ns 408930-001 Rev E, 408930-002 Rev C. |  |  | February 16, 2015 | 2015 | GE Healthcare | Potential safety issue due to improper assembly associated with the power cords of T2100 and T2000 Treadmills. | Process control | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref#30073 dated February 16, 2015. The letter was addressed to Director of Biomedical Engineering, Chief of Nursing, &amp; Healthcare Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative.UPDATE 3-30-2015:A second revised letter was sent on March 31, 2015. NOTE: this revised notification includes an update to the Affected Product Details section to include all revisions of power cords manufactured within the date range indicated. (The previous notification referenced only revisions B , C , &amp; E ). |  | 13345 | 13,345 (5,495 units US, 7,850 units OUS) | CONTINENTAL USA DISTRIBUTION INCLUDING DC AND PR.OUS: ARGENTINA AUSTRALIA, AUSTRIA, BANGLADESH, BARBADOS, BHUTAN, BOLIVIA, BRAZIL, BRUNEI, CAMBODIA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, DOMINICAN REPUBLIC, ECUADOR, EL SALVADOR, SPAIN, FINLAND, FRANCE, UNITED KIGDOM, GERMANY, GUAM, GUYANA, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, ITLAY, JAMAICA, JAPAN, REPUBLIC OF KOREA, MACAU, MALYAYSIA, MARTINIQUE, MEXICO, MONACO, MYANMAR, NEPAL, NEW ZEALANAD, NICARAGUA, PANAMA, PARAGUAY, PERU, PHILIPPINES, PORTUGAL, RUSSIA. SINGAPORE, SOUTH KOREA, SRILANKA, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, URUGUAY, UZBEKISTAN, VENEZUELA, VIET NAM. | Terminated | June 23, 2017 | 858 | 13345 | Not_Computer
Z-1276-2015 | 70644 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, 10 Fr FlexCath Select Steerable Sheath 990065. |  |  | February 23, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | Clinicians observed debris, appearing to originate from the hemostasis valve on the proximal end of the Steerable Sheath, outside of the patients bodies. | Nonconforming Material/Component | The firm, Medtronic, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 2015 to its consignees/customers. Medtronic representatives, beginning Monday, Feb 23, 2015 hand delivered the letter to US consignees/customers.. The letter was addressed to Physician, Risk Manager and Health Care Professional. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to immediately quarantine all unused product; return all affected product to Medtronic; contact Customer Service at 1-800-848-9300 to initiate product return and to complete and return the Customer Confirmation Certificate to via email to: RS.CFQFCA@Medtronic.com or fax to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. In additions, consignees/customers should share this notification with others in your organization.Your Medtronic field representative may assist the customers with identifying suitable replacement product.If you have any questions, contact Customer Service at 1-800-848-9300. |  | 151 | 151 (4 US, 147 OUS) | Worldwide Distribution: US (nationwide) in states of: NC and IA; and countries of: AUSTRIA, BELGIUM, DENMARK, FRANCE, GERMANY, NETHERLANDS, POLAND, SPAIN, and UNITED KINGDOM. | Terminated | October 07, 2015 | 226 | 151 | Not_Computer
Z-1288-2015 | 70648 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | LP Rotate Foot System prosthetic device for fitting lower extremity amputations. |  |  | March 04, 2015 | 2015 | Ossur H / F | The threads of the outer tube of the shock module of LP Rotate are too short, resulting in insecure assembly. Use of the product may cause the patient to fall with continued use, as the shock module can become loose from the foot blade. | Nonconforming Material/Component | The firm's recall letter states the following instructions:Immediately examine inventory and quarantine product subject to the recall. In addition the firm asks if further distributed to identify the customers and notify them of the recall. The firm states that notification may be enhanced by providing a copy of the letter.The firm states that the recall notice needs to be passed along to all those who need to be aware.The firm also states that if customers are wearing a faulty product that they recommend them scheduling a visit to have the product replaced with a new LP Rotate that is currently available. The firm states that they can also arrange pick-up of defected products.For any questions the firm provides a customer service contact: Jamie Dutter, Senior Customer Service Representative. |  | 2194 | 2,194 (1,464 in U.S. and 730 internationally) | Nationwide within the U.S. and internationally to the following foreign countries; Belgium, United Kingdom, Germany, Czech Republic, France, Israel, Poland, Spain, Australia, Japan, Mexico, Bulgaria, Canada, South Korea, Sweden, Finland, | Terminated | October 28, 2015 | 238 | 2194 | Not_Computer
Z-1287-2015 | 70664 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Signature TKA GDE/MDL Set 04-05 (Femur &amp; Tibia) |  |  | February 04, 2015 | 2015 | Materialise N.V. | Images belonging to a different patient were used for the production of the patient specific surgical guide. The incorrect guide may cause a delay in surgery to accommodate the preparation and use of traditional instrumentation. | Employee error | Materialise N.V. sent an email on February 4, 2015, to the the affected distributor regarding the reason for the recall, affected product, and instructions for removing the affected product. For questions regarding this recall call 321-639-6611. |  | 1 | 1 unit | Worldwide Distribution to the Netherlands only. | Terminated | October 30, 2015 | 268 | 1 | Not_Computer
Z-1284-2015 | 70490 | Yes | OAE |  |  | Cardiovascular | PMA | TactiSys Quartz Pack, PN-004 400. |  |  | November 25, 2014 | 2014 | St Jude Medical | A configuration update needs to be done on TactiSys to appropriately recognize all TactiCath catheters. | Software design | St. Jude Medical sent a Service Bulletin to inform of a TactiSys Configuration File update that needs to be performed. A technical service personnel will visit each consignee and performed the software change on-site. For questions about this Service Bulletin, contact Technical Support(ETechsupport@sjm.com). |  | 89 | 89 | US Nationwide Distribution in the states of: AL, AZ, CA, CO, CT, FL, GA, ID, IL IN, KS, LA, MD, MA, MI, MN,MS, NE, NV,NH, NM, NY, NC, OH, OK OR, PA, TN, TX, UT, VA,WA. | Terminated | May 06, 2015 | 162 | 89 | Not_Computer
Z-1286-2015 | 70402 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CADD" Medication Cassette Reservoir, 50mL, REF 21-7001-24 &amp; 21-7301-24, Sterile EO, Rx only. |  |  | February 02, 2015 | 2015 | Smiths Medical ASD, Inc. | Smiths Medical has become aware of an issue with specific lots of 50mL CADD" Medication Cassette Reservoirs ( Cassette ). Some Cassettes may leak at the sealing area of the pump tube and medication bag. Smiths Medical has received no reports ofserious injury or death related to this issue. | Under Investigation by firm | Smiths Medical sent an "Urgent Medical Device Recall" letter dated February 2, 2015, to all affected customers. The letter was addressed to Risk Safety Manger, Biomedical Professionals, Clinicians who oversee the use of CADD pumps, Distributors, and other users of these devices. The letter described the problem and the product involved in the recall. The letter also addressed the "Advice on Action to be Taken by the User", and "Transmission of this Urgent Recall Notice". Customers were requested to complete and return the Confirmation Form by Fax to 1-800-237-8033 or be email to recall.response@smiths-medical. For issues with these products they can contact Smiths Medical's Global Complaint Department at 1-866-216-8806 or globalcompaints@smiths-medical.com. For questions on the information provided they can contact Customer Service Department at 1-800-258-5361. |  | 48208 | 48,208 (17,040 US, 31,168 OUS). | Worldwide Distribution - US (nationwide) Distribution to the states of : AL, AZ, AR, CA, CO, CT, FL, GA, HI, IL IN, IA, MD, MA, MN, MS, MO, NE, NV, NH ,NJ, NM, NY, NC, OH, OR, PA, SD, TN, TX, UT, VA and WA., and to the countries of : Arab Emirates, Austria, Australia, Belgium, Canada, Switzerland, Colombia, Czech Republic, Germany, Denmark, Spain. Finland, France, United Kingdom (GB) Hungary, Indonesia, Ireland, Italy, Netherlands, Portugal, Sweden, Singapore and South Africa. | Terminated | March 18, 2016 | 410 | 48208 | Not_Computer
Z-1291-2015 | 70551 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | Outer cardboard box labeling: |  |  | February 11, 2015 | 2015 | Sybron Dental Specialties | Cases of Cure Sleeve labeled with part number 4513, which fits 13mm tips, have inner boxes labeled with part number 4500 which fit 7mm and 8mm tips. There is a low risk associated with the use of the product as the sleeve may tear causing a potential for cross contamination. | Packaging | The firm sent out letters to customers on 02/11/15.The firm's letter states that they are willing to replace or credit back customers if they have any of the affected product, and instruct them to contact Metrex Research Customer Care at 1-800-841-1428 in order to receive an RMA number. The firm also requests that customer complete and fax back the enclosed reply form in order to confirm receipt of this advisory notification. The firm also states that if the customer is a distributor that they identify the customers that might have received the affected product, and contact them within 48 hours. |  | 172 | 172 | Worldwide distribution. US Nationwide, Canada, Australia, Netherlands, Switzerland, Japan, | Terminated | May 21, 2015 | 99 | 172 | Not_Computer
Z-1292-2015 | 66737 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | The HomeChoice Pro APD System is a peritoneal dialysis system. |  |  | October 16, 2013 | 2013 | Baxter Healthcare Corp. | The occluder in affected devices 1) may have sharp edges that could cut the pneumatic tubing that connects the occluder blade to the manifold causing an unrecoverable system error, and/or 2) blade may separate from the back plate and cause Slow / No Flow alarm requiring therapy termination. | Process design | An Urgent Device Correction letter dated 10/16/13 was mailed to dialysis centers and medical facilities. A similar Urgent Device Correction letter dated 10/31/13 was mailed to dialysis home patients. The letters requested for customers to examine their devices to determine if the devices are being recalled. If the devices are on the attached list of recalled devices, the customer is to contact Baxter Technical Services at 1-800-553-6898 (7 days a week, 24 hours a day) to arrange for a swap cycler. The letter further requests that customers complete the attached Customer Reply Form and return it to Baxter by fax or e-mail. |  | 75 | 75 units (US) | Worldwide Distribution-US (nationwide) including the states of AK, AR, AZ, CA, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MO, NC, NJ, NV, NY, OH, PA, SC, TN, TX, UT, VA, VT, WA, and WI, and the countries of Australia, Canada, and Mexico. | Terminated | July 21, 2017 | 1374 | 75 | Not_Computer
Z-1293-2015 | 69813 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Crea A and B membranes, packaged 4 per box |  |  | October 29, 2014 | 2014 | Radiometer America Inc | Negative drift on QC during in-use lifetime of CREA membranes can be observed. | Under Investigation by firm | Radiometer sent a Product Notification letter dated November 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to perform the following steps immediately:" Perform one quality control measurement using S7835 AutoCheck 6+, level 1 on already installed Crea membrane units to verify the Crea performance. If the QC result is out of range, new Crea A and Crea B membranes must be installed.Until further notice the following must be carried out every 24 hours:" Perform one quality control measurement using S7835 AutoCheck 6+, level 1 to verify the Crea performance. If the QC result is out of range, new Crea A and Crea B membranes must be installed.Please do not report Crea results unless the membrane unit has been verified as described above.If you already run S7835 AutoCheck 6+ level 1, every 24 hours, as part of your regular QC schedule, you do not need to do anything additional. Please review your QC results to ensure the results are within the QC range.Please complete and return the attached fax form, with your signature as soon as possible or within 7 days.Customers with questions were instructed to contact Radiometer America Technical Support at 1-800-736-0600, option 4.For questions regarding this recall call 440-871-8900, ext 209. |  | 10824 | 10,824 boxes | Worldwide Distribution - USA including NM, NC, TN, OH, NE, CA, MO, GA, MA, TX, WV, PA, NY, VA, MN, IN, ME, MD, IL, MI, KS, WI, FL, WA, CT, OK, CO, and Internationally to Canada, Denmark, Spain, Italy, India, Sweden, Germany, Australia, Hungary, Great Britain, Norway, Uruguay, France, Poland, Switerland, Russian Federation, Netherlands, Czech Republic, Mexico, Brazil, Czechoslovakia, Austria, Singapore, Syria, Phillipines, Egypt, Bolivia, Japan, New Zealand, Turkey, Saudi Arabia, Algeria, Qatar, and Finland. | Terminated | July 02, 2015 | 246 | 10824 | Not_Computer
Z-1295-2015 | 70647 | Yes | MQS | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Transonic Hemodialysis Monitor, Model HD03 -- |  |  | February 17, 2015 | 2015 | Transonic Systems Inc | Customer unable to obtain measurements due to the device seeing no flow or irrelevant flow condition during dialysis procedures. The user saw a zero or a very low flow value displayed on HD03 and/or no measurements produced in other modes (recirculation; access flow; or cardiac output). | Process control | Transonic sent a Medical Device Recall Letters and Response Forms dated February 17, 2015 to the users via e-mail and fax. The recall is being conducted to the user level. The letter identified the affected product, problem and actions to be taken. For questions call 607-257-5300 ext. 326. |  | 31 | Domestic: 21 units, Foreign: 10 units | Worldwide Distribution - USA Nationwide and the countries of: Australia, Canada, and South Korea | Terminated | May 31, 2016 | 469 | 31 | Not_Computer
Z-1294-2015 | 70698 | No | DWQ | External vein stripper. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre Vascular Tumescent Catheter Inversion (TCI) System. The TCI System is intended to provide a system to the removal of an incompetent greater saphenous vein. |  |  | February 24, 2015 | 2015 | LeMaitre Vascular, Inc. | Mislabeled with an incorrect expiration date | Employee error | LeMaitre Vascular issued a Field Safety Notice/Dear Doctor dated 2/24/15 to customers The letter provides information on the recalled product and instructions on how the customer can inspect for the defective devices and return the affected products for replacement/credit. The letter contains a form that is requested to be returned to LeMaitre Vascular, Inc. as a record of notification and reconciliation.Questions, contact 781-221-2266 ext. 183 |  | 54 | 54 units | CA, CO, KS, MN, UT, WA, WI | Terminated | May 21, 2015 | 86 | 54 | Not_Computer
Z-0984-2015 | 69880 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE MV System, Linear Accelerator (LINAC) with SysVC10A software. |  |  | November 19, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There may be an existing dark current phenomenon on ARTISTE LINAC in combination with IMRT or mARC treatments using unflat beams. Software issue. | Software design | Siemens Planned Action: 1. Providing recommendations for user actions in terms of routine dosimetric quality control checks for any IMRT treatment at very high dose rates (Multiple X) as well as for mARC treatments.2. Software update for the Function Controller #3, which controls parts of the beam generation. This software update enables the correct switching between Pulse Forming Network (PFN) and reduced Pulse Forming Network (iPFN) in pause state, to reduce the dark current during the movement as much as possible. This software will be installed by a Siemens service representative.3. Siemens will contact the user to make an appointment to implement the modification system at a date and time when the system will be available. These corrections shall be made free of charge.The deployment of the permanent correction shall be completed by July 1, 2015.For further questions please call (610) 219-6300. |  | 3 | 3 | US Distribution to the states of UT, WI, and NY. | Terminated | April 28, 2015 | 160 | 3 | Computer
Z-1301-2015 | 70286 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | NovaPlus Brand Warm Gel Infant Heel Gel Warmer, Item # V6391 |  |  | January 26, 2015 | 2015 | CooperSurgical, Inc. | Products without expiration date have the remote potential for minor burn | Package design/selection | CooperSurgical issued Recall Notification letter on January 26, 2015, via Fedex confirmed receipt delivery, to all affected customers. Customers were requested to review inventory and return/destroy product that does not have a expiration and lot number. A fax back form to be returned. Customers with questions should call 203.601.5200.CooperSurgical's expense, arrangements were made to replace any of the affected products that customers had on hand. CooperSurgical provided further instruction on the disposition of the product at the time of replacement. |  | 4938266 | 4938266 units | Worldwide Distribution - USA (nationwide) and Internationally to AUSTRALIA, BAHRAIN, BERMUDA, CANADA, FRANCE, LATVIA, LIBYA, MEXICO, NEW ZEALAND, QATAR, SAUDI ARABIA, SINGAPORE, and UNITEDKINGDOM. | Open, Classified |  |  | 52899493 | Not_Computer
Z-1238-2015 | 70458 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Elekta Precise Treatment Table |  |  | January 29, 2015 | 2015 | Elekta, Inc. | It is possible to position the Treatment Table with errors greater than 5mm and no inhibits are displayed, if the positional sensors are not correctly installed during corrective maintenance. Clinical mistreatment may occur if a position error is not detected. | Radiation Control for Health and Safety Act | Planned Action by Elekta:1. Notifications have been sent out. The notification informs users of the possibility treatment table position errors might occur if there is a maintenance error when replacing a potentiometer. The notification recommends establishing a procedure to verify the accuracy of the treatment table position as shown by the treatment room monitor which is performed as part of the daily machine checks. An example procedure is provided.2. Elekta will establish a procedure service engineers will follow after maintenance involving the relevant potentiometers to ensure proper installation. The procedure will be taught using a video which includes a step by step guide to the correct methodology for installation of the potentiometer assembly, a description of the impact for not installing correctly, and a guide to the new verification of correct installation procedure. The video will be distributed through Elektas learning management system (LMS). The LMS maintains a record of service engineers who successfully complete the training.3. Elekta will implement an engineering change to the potentiometer assembly which applies to the manufacture of all new Precise Treatment Tables. This change will eliminate the potential for incorrect installation and will prevent a loose grub screw from causing backlash. This change is expected to consist of a 'D' shaped bore that cannot spin on the shaft of the potentiometer.All actions associated with this correction, including service engineer training, will be completed by June 2015. For further questions please call (770) 300-9725. |  | 3480 | 529 US 2951 Worldwide | Worldwide Distribution - USA (nationwide) Distribution and to the countries of Albania, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Lebanon, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland and Syria. | Open, Classified |  |  | 3480 | Not_Computer
Z-1296-2015 | 70246 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia E is a multi-purpose SPECT system used to detect or image the distribution of radionuclides in the body or organ, using the following techniques: planar imaging, whole body imaging, tomographic imaging for isotopes with energies up to 588 keV. |  |  | April 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | Radial motor drives replaced during a customer service action were incorrectly assembled backwards. | Employee error | Urgent Field Correction- Recall letters dated April 11, 2012 were sent to all potentially affected customers. The letters included instructions for customers to immediately remove the patient from the bed and contact the customer's Service Engineer in the event that one of the following errors is displayed "A motion control error has occurred (E112)" or "A motion control error has occurred (E113)." Customers were advised that a Service Engineer will be inspecting the radial drive gear on the customer's camera system to ensure it is functioning correctly and that the machines can continue to be used as usual until the Service Engineer can perform the inspection of the machine(s).Customers with questions were advised to contact Siemens as 800-888-7436. |  | 419 | 419 systems (worldwide) | Worldwide Distribution-US (nationwide) including the District of Columbia and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, and WV and the countries of AE, AL, AO, AR, AT, AU, BA, BD, BE, BG, BH, BO, BR, BW, BY, CA, CH, CL, CM, CN, CO, CR, CY, CZ, DE, DK, DO, DZ, EC, EG, ES, FA, FI, FR, GB, GE, GR, HK, HR, ID, IE, IL, IN, IQ, IR, IT, JO, JP, KR, KW, KZ, LB, LK, LU, LV, MA, MK, MM, MU, MX, MY, NE, NG, NL, NO, NP, NZ, OM, PE, PH, PK, PL, PS, PT, QA, RO, RS, RU, SA, SE, SG, SI, SK, SV, SY, TH, TN, TR, TW, UA, UG, UY, UZ, VN, YE, and ZA. | Open, Classified |  |  | 5275 | Not_Computer
Z-1306-2015 | 70151 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | Penner Patient Stretcher/Transfer Lift System Models 391000-1 and 392000-1 |  |  | January 30, 2015 | 2015 | Penner Mfg Inc | If the patient sits on either wing, not squarely on the center section of the Stretcher Lift, tipping and injury to the operator or patient could result. The device manual was updated with an additional warning statement. | Use error | Beginning on 1/30/2015, the recalling firm sent letters via certified mail to their customers informing them of the change to the stretcher manual along with a revised manual which incorporated those changes. For questions regarding this recall call 402-694-5003. |  | 158 | 158 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | March 10, 2016 | 405 | 158 | Not_Computer
Z-1307-2015 | 69812 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Lithium Polymer Fiber Optic Module Batteries used in the Intera MR System upgraded to CDAS, Part Number: 9065 and 9067. These rechargeable batteries are used to supply power to the following products: Wireless ECG Module used with the Precess MRI compatible Patient Monitoring System, Wireless SpO2 Module used with the Precess MRI Compatible Patient Monitoring System, Philips Basic Triggering Set, Philips Monitoring Sensor Set, Philips Battery Set |  |  | November 20, 2014 | 2014 | Philips Medical Systems, Inc. | VCG battery ignited in a VCG unit when recharging. | Component design/selection | Consignees were informed of the recall via letter sent on November 20, 2014 via certified mailing. Field Service Engineers (FSE) will replace all affected batteries. |  | 79 | 79 | Distributed in the following states: AL, CA, FL, GA, IL, MA, MN, NC, ND, NE, NJ, NY, OH, OR, PA, TN, TX, UT, WA, WI; OUS: Australia, Austria, Belgium, China, Ecuador, Finland, France, Germany, Italy, Japan, Mexico, Netherlands, slovenia, South Africa, Sweden, Switzerland, United Kingdom | Open, Classified |  |  | 79 | Not_Computer
Z-1211-2015 | 70478 | No | RCE |  |  |  |  | GE Inspection Technologies x¦tome¦x L240 an Cabinet X-ray system |  |  | January 13, 2015 | 2015 | GE Inspection Technologies, LP | It was discovered that the upper left corner-shaped lead shield in the window in one of the doors of the v¦tome¦x L240 installed had lost adhesion and fall out of position, which could lead to gaps in shielding. | Device Design | GE Inspection Technologies, Planned Action: 1. GE will notify customers of the affected cabinet x-ray systems through a Technical Information Letter. GE notified its affected customers via email on January 13, 2015, via letter dated January 14, 2015, and via telephone on January 19, 2015. 2.GE will install an interim insert within 30 days of receipt of this letter. The insert will cover any existing or interim gaps in the original lead shielding on the door.3. GE will replace the doors of affected cabinet x-ray systems no later than 180 day of receipt of this letter.4. These corrections will be made free of charge.The Technical Information Letter used for notification of all dealers and purchasers is to be made within 15 working days of receipt of this letterFor further questions, please call (717) 447-1278 |  | 2 | 2 | US Distribution to the states of OH and PA. | Terminated | August 17, 2015 | 216 | 2 | Not_Computer
Z-1310-2015 | 70686 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Radiation Therapy Treatment Planning System; -- RayStation 3.5, RayStation 4.0, RayStation 4.5 and RayStation 4.7. For RayStation 4.7, the issue applies also to machines with fixed jaws, regardless of MLC/jaw position. |  |  | March 04, 2015 | 2015 | RAYSEARCH LABORATORIES AB | An issue with photon dose calculation for DMLC (Dynamic MLC) plans for machines where the MLC is positioned above the jaws, e.g. some Elekta linacs. The magnitude of the error depends on the beam model output factor corrections and on the individual DMLC plan characteristics. | Under Investigation by firm | The firm, Raysearch Laboratories, sent a "Medical Device Correction/Field Safety Notice" and Reply Form(dated 3/3/2015) on 3/4/2015 via e-mail to the affected customers. The notice describes the product, problem and actions to be taken. The customers were instructed to deprecate any beam models that have been commissioned as DMLC capable for the affected linac types; do not commission beam models as DMLC capable for the affected linac types; educate physics staff and all users; inspect your system and identify all affected units of RayStation, acknowledge this notice by email or complete and return the Reply Form via fax to Americas market, RaySearch Americas Inc., at 888 501 7195, email: freddie.cardel@raysearchlabs.com or Rest of the world, RaySearch Laboratories AB, to email support@raysearchlabs.com.If you have any questions, contact Quality and Regulatory Affairs Manager, at +46 722 366 110 or email david.hedfors@raysearchables.com. |  | 245 | Domestic: 245 units | Nationwide Distribution. | Open, Classified |  |  | 245 | Not_Computer
Z-1290-2015 | 70727 | No | MCX |  |  | Cardiovascular | PMA | Diamondback 360 Coronary Orbital Atherectomy System, Model Number: DBEC-125. |  |  | March 04, 2015 | 2015 | Cardiovascular Systems, Inc. | The OAD saline sheath outer layer may flake off from the catheter due to friction between the outer material layer and corresponding materials it may come in contact with during clinical use. If the material is large enough, it has the potential to cause a flow limiting embolism. | Nonconforming Material/Component | Consignees were sent on 3/4/2015 a CSI "Urgent Medical Device Recall" letter dated March 04, 2015. The letter described the problem and the affected product. Consignees were advised to remove and return the affected product to CSI. The letter also requested consignees to complete and return the Customer Acknowledgement Form by Fax or with the RGA (Returned Goods Authorization). For further information they can contact Customer Service, Cardiovascular Systems, Inc., 1225 Old Hwy 8, New Brighton, MN 55112, 877-274-0901 -Tel, 612-677-3355- Fax. |  | 30 | 30 | Nationwide Distribution-including the states of AZ, FL, GA, LA, MA, MS, MO, NJ, NY, OH, and RI. | Terminated | April 26, 2016 | 419 | 30 | Not_Computer
Z-1304-2015 | 70705 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Xi Surgical System, IS4000; |  |  | March 03, 2015 | 2015 | Intuitive Surgical, Inc. | Correction due to a fault that causes distal SUJ (DSUJ) failures that indicates a mismatch between the positions reported by the redundant position sensors on one of the patient cart arm joints. | Under Investigation by firm | The firm, Intuitive Surgical, sent an "Urgent Medical Device Correction" letter dated March 11, 2015 to its consignees/customers on March 11, 2015 via Fed Ex. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that all affected personnel are fully informed of this notice; forward this notice to your Risk Manager, OR Director, Purchasing Manager, biomedical Engineering staff and members of your medical staff who perform da Vinci Surgery procedures; Identify systems containing affected product at your site; if you encounter this error during a procedure and the fault cannot be recovered, disable the affected arm and contact dVStat; Contact Intuitive Surgical Customer Service; Complete and return the attached Acknowledgment Form to Intuitive Surgical via U.S. Fax +1 (408) 716-3040, or Scan and Email: isi.compliance@intusurg.com, ATTN: REGULATORY COMPLIANCE-Subject line for email: Pot to Encoder;and retain a copy of this notice for your records.An Intuitive Surgical representative will contact you to schedule a remote assessment to determine the health of your system and to schedule an appointment to replace the affected arm.If you need further information or support concerning this Medical Device Correction, please contact your Clinical Sales Representative or Intuitive Surgical Customer Service at the numbers listed below:-North America and South America: 800-876-1310, Option 3 (6 AM to 5 PM PST)-Japan: 0120-56-5635 or 03-5575-1362 (9 AM to 6 PM JST)-South Korea: 02-3271-3200 (9 AM to 6 PM KST)-Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020(8 AM to 6 PM CET) or ics@intusurg.com. |  | 167 | 167 total affected da Vinci Xi systems | Worldwide Distribution: US (nationwide) and countries of: Australia, Belgium, China, France, Germany, India, Israel, Italy, Norway, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | September 08, 2015 | 189 | 167 | Not_Computer
Z-1308-2015 | 70712 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM systems Aera/Skyra/Avanto/Verio with software syngo MR D13A ; indicated for use as magnetic resonance diagnostic devices (MRDD) that produce transverse, sagittal, coronal and oblique cross-sectional images, spectroscopic images and/or spectra, and that display the internal structure and/or function of the head, body, or extremities. Other physical parameters derived from the images and/or spectra may also be produced. Depending on the region of interest, contrast agents may be used. These images and/or spectra and the physical parameters derived from the images and/or spectra, when interpreted by a trained physician, yield information that may assist in diagnosis. The MAGNETOM systems may also be used for imaging during interventional procedures when performed with MR compatible devices such as in-room display and MR-safe biopsy needles. |  |  | March 09, 2015 | 2015 | Siemens Medical Solutions USA, Inc | The gradient output supervision was permanently turned off on the MAGNETOM system, meaning that gradient outputs could exceed IEC60601-2-33 limits and peripheral nerve stimulation could occur. | Nonconforming Material/Component | A Customer Safety Advisory Notice, dated March 6, 2015, was sent to end users to inform them of the potential issue and what measures were being taken to mitigate possible risks. |  | 132 | 132 | Nationwide Distribution. | Open, Classified |  |  | 132 | Not_Computer
Z-1305-2015 | 70687 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, SIGNA 1.5T TWINSPEED, SIGNA INFINITY MRISYSTEM, SIGNA 3.0T INFINITY WITH EXCITE, SIGNA 3.0T WITH EXCITE, SIGNA EXCITE 1.5T, SIGNA EXCITE 3.0T, GE 1.5T AND 3.0T SIGNA HDX MR SYSTEM, Discovery MR750 3.0T, Discovery MR450 1.5T, Discovery MR750w 3.0T, Optima MR450w 1.5T, SIGNA Contour, Brivo MR355, Optima MR360, GE 1.5T SIGNA HDE MR SYSTEM, Signa Openspeed 0.7T MR SYSTEM, Vectra, SIGNA 0.35T OVATION WITH EXCITE, SIGNA 0.2T PROFILE With EXCITE |  |  | March 09, 2015 | 2015 | GE Healthcare | GE Healthcare has become aware of a potential safety issue involving MRI systems due to software versions not being maintained properly at some sites. | Software Manufacturing/Software Deployment | GE Healthcare sent an "Urgent Medical Device Correction" letter dated March 9, 2015, to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact your local service representative or call 800-437-1171 (available 24 hours a day). |  | 9369 | 9,369 (2,937 US, 6,432 OUS). | Worldwide Distribution - USA (nationwide) Distribution including PR, GU and DC., and to the countries of: ALBANIA. ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, COTE DLVOIRE, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND FRANCE, FRENCH POLYNESIA, GEORGIA, GERMANY. GHANA, GREECE, GUADELOUPE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KAZAHHSTAN,KENYA,REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG MACEDONIA, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMA N,PAKISTAN, PALESTINE, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SIERRA LEONE, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TURKEY, UGANDA , UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, UZBEKISTAN, VENEZUELA, VIETNAM and ZAMBIA. | Terminated | September 08, 2015 | 183 | 9369 | Not_Computer
Z-1309-2015 | 70557 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Intraocular Lens (IOL), Part No. EC-3 PAL, Diopter: 17.5, 18.0, 18.5, 19.0, 19.5, 21.5, 23.5, 9.0, 22.5, 23.0, 21.0, 24.0. |  |  | February 13, 2015 | 2015 | Aaren Scientific, Incorporated | Aaren Scientific is recalling intraocular lenses because the instructions for use (IFU) are not fully compliant with US requirements. | Labeling design | A customer notification letter dated 2/13/15 was sent to all customers who received the Intraocular Lens model EC-3 PAL. The letter provides the customers with the problems identified and the actions to be taken. Customers are instructed to confirm receipt of the recall letter and to contact Mark Scheib at (909) 906-5119 or via email at mark.scheib@zeiss.com with any questions or comments. |  | 28 | 28 units | Distributed in the states of OK, KS, FL, and AL. | Terminated | June 29, 2015 | 136 | 28 | Not_Computer
Z-1312-2015 | 70216 | No | FRG | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Tyvek Pouch/Roll with STERRAD Chemical Indicator, P/N 12425 |  |  | December 24, 2014 | 2014 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the Tyvek Rolls with STERRAD Chemical Indicator because they were shipped to customers after its expiration date of May 18, 2013. | Under Investigation by firm | A customer notification letter was sent on 12/24/14 to all customers who purchased the expired Tyvek Rolls with STERRAD Chemical Indicator. The letter informs the customers of the problems identified and the actions to be taken. |  | 2 | 2 units | Saudi Arabia and Portugal | Terminated | September 04, 2015 | 254 | 2 | Not_Computer
Z-1313-2015 | 70247 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Electric Pen Drive (EPD) System, The Synthes Electric Pen Drive is indicated for screw insertion, pin and wire replacement, cutting of bone and metal, drilling, decorticating, shaping, and smoothing of bones and teeth in a wide variety of surgical procedures, including, but not limited to general orthopaedic trauma, foot, hand, maxillofacial, neurosurgical, oral, otolaryngolical, reconstructive, and spine surgery. |  |  | December 19, 2014 | 2014 | Synthes, Inc. | In May 2013, the original recall was initiated due to complaints the Hand Switch for the Electric Pen Drive was not labeled correctly, and the on and lock labels were interchanged. | Process design | Synthes sent an Urgent Medical Device Recall Notice, dated December 19, 2014, to all affected customers to provide updated information on additional lots identified. Customers were instructed to examine their inventory for the affected product, remove from inventory, and return them to Synthes immediately for replacement. Customers were instructed to call Synthes to obtain a Return Authorization (RA) Number, complete the Verification Section at the end of the letter, and return the Verification Form with the product to: Credit/Returns, Synthes 1101 Synthes Avenue Monument, CO 80132. Customers with questions were instructed to call 610-719-5450 or email FieldAction@synthes.com. For questions regarding this recall call 610-719-5000. |  | 249 | 249 | Nationwide Distribution | Open, Classified |  |  | 249 | Not_Computer
Z-1311-2015 | 70608 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC units, Model No. 8015. Infusion pump. |  |  | March 12, 2015 | 2015 | CareFusion 303, Inc. | CareFusion is recalling the Alaris PC unit because of an error code. The error code may occur upon power on during the "Power-On Self Test" due to a keypad issue. | Process control | CareFusion sent an Urgent: Medical Device Recall Notification letter dated March 12, 2015, to all affected customers. The letter informs the customers that CareFusion has received service reports of a system error on the Alaris PC unit 8015. The letter informs customers of the problems identified and the actions to be taken. Customers with recall related questions are instructed to contact CareFusion Support Center at (888) 562-6018, supportcenter@carefusion.com. Customers with adverse event reports are instructed to contact customer advocacy at (888) 812-3266, customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris system are instructed to contact technical support at (888) 812-3229, DL-US-INF-TechSupport@carefusion.com. |  | 56015 | 56,015 units | Worldwide Distribution - USA (nationwide), Canada, European Union, Singapore and Australia. | Open, Classified |  |  | 56015 | Not_Computer
Z-1314-2015 | 70578 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Variable Angle Locking Peg, D:1.8mm x L:14.0mm; Variable Angle Locking Peg, D:1.8mm x L:16.0mm, Variable Angle Locking Peg, D:1.8mm x L:18.0mm; Variable Angle Locking Peg, D:1.8mm x L:20.0mm; Variable Angle Locking Peg, D:1.8mm x L:22.0mm, Variable Angle Locking Peg, D:1.8mm x L:24.0mm; Variable Angle Locking Peg, D:1.8mm x L:26.0mm. |  |  | February 16, 2015 | 2015 | Flower Orthopedics Corporation | Locking pegs were not locking during distal radius procedures. There has been one (1) report of revision surgery scheduled to remove a loosened peg. | Device Design | On February 16, 2015, Flower Orthopedics generated and distributed a field memo/advisory notice (FOCMB-15-02) to all distributors with locking peg inventory on February 6, 2015. The field memo provided technique guidance to end users. All 12 consignees with remaining inventory (ie.not reported as used) were contacted directly via phone and all have agreed to ship the locking pegs back to Flower Orthopedics. Flower Orthopedics determined that a generic statement to all customers was not necessary for the following reasons detailed above and summarized again here: product was distributed only through direct channels, the health hazard evaluation has determined there is a very low risk to health, and all consignees have been directly contacted already and product inventory is being returned. |  | 901 | 901 | Distributed in the states of VA, TX, IL, OK, PA, CA, NY, KY, and FL. | Terminated | August 06, 2015 | 171 | 901 | Not_Computer
Z-1315-2015 | 70090 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Liquid Mouse Monoclonal Antibody CD15 used is laboratories to perform immunohistochemical (IHC) staining. |  |  | December 16, 2014 | 2014 | Leica Microsystems, Inc. | The staining intensity decreases over the shelf life. | Labeling design | Leica sent an Urgent Medical Device Recall Notification letters dated December 16, 2014 to all customers of record. The letters included instructions for customers to: 1) destory any usused and partially used recalled materials or confirm that the materials are no longer in stock; 2) complete and return the attached Field Safety Correction Notice Acknowledgement Form; 3) contact your local Leica representative immediately if a replacement is required; and, 4) pass this notice primarily to the end users where the product has been sold and to all those within your organisation who need to be aware of this manufacturing issue. Customers with questions can contact Leica Microsystems via e-mail at LMGRA@leica-microsystems.com. |  | 309 | 309 units | Worldwide Distribution - US Nationwide in the states of CA, CT, FL, IA, ID, IL, MT, NC, NY, TX, WA and in the countries of: Argentina, Australia, Bosnia and Herzegovina, Brazil, Canada, Chile, Denmark, El Salvador, France, Germany, Hungary, India, Italy, Japan, Netherlands, Nigeria, Portugal, Romania, Russia, Saudi Arabia, Serbia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tunisia, Turkey, United Arab Emirates, United Kingdom. | Open, Classified |  |  | 309 | Not_Computer
Z-1270-2015 | 70443 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare Engstrom Carestation, Engstrom Pro, Cart Only S/5 EC. GE model numbers 1505-8538-000, 1505-9000-000, and 1505-9003-000. Ventilators. |  |  | January 27, 2015 | 2015 | GE Healthcare | Potential failure of the caster mounting hardware, which may result in the ventilator falling/tipping (overbalancing). | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated January 26, 2015, Ref#34063. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171, or your local Service Representative. |  | 3213 | 3,213 (437 US, 2,776 OUS) | Worldwide Distribution. US nationwide, ALBANIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOSNIA &amp; HERZEGOVINA, BRAZIL, CANADA, CAMEROON, CANDA, CHILE, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, CHANA, GREECE, GUATEMALA, HONDURAS, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LEBANON, LIBYAN ARAB JAMAHIRIYA, MACEDONIA, MALAYSIA, MALTA , MEXICO, MOROCCO, NETHERLANDS, OMAN, PANAMA, PAPUA NEW GUINEA, PERU, POLAND, PORTUGAL, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWITZERLAND, SYRIAN ARAB REPUBLIC, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VIETNAM, VENEZUELA, and URUGUAY. | Terminated | February 03, 2016 | 372 | 3213 | Not_Computer
Z-1322-2015 | 70693 | No | ISW | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | The Rheo Knee bionic prosthetic is to be used exclusively for fittings of lower extremity amputations. RHEO KNEE uses Artificial Intelligence to continuously adapt to the users walking style and environment. The RHEO KNEE recognizes and responds immediately to changes in speed, load and terrain, restoring the users ability to walk naturally, comfortably and confidently at any speed. |  |  | February 26, 2015 | 2015 | Ossur H / F | The firm is recalling Rheo Knee bionic prosthetic due to it being discovered during an internal audit of the service line that devices were released for distribution without fully going through the assembly process. | Employee error | The firm, Ossur, notified their customers of the Recall Notice via phone on February 26, 2015. The firm informed the customers about the recalled product, the problem and actions to be taken. The customers were instructed to send the unit in for service so the Acceptance testing can be performed for the unit in accordance to their process. If you have any questions, please contact Regulatory Affairs &amp; Quality Assurance specialist at 949-382-3741 or email: kmontes@ossur.com. |  | 4715 | 4715 units (2 units affected) | Worldwide distribution: US distribution in TN and country of: Austria. | Terminated | November 20, 2015 | 267 | 4715 | Not_Computer
Z-1316-2015 | 70657 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac versions 6.x patient positioning systems are used to position patients during radiosurgery or radiotherapy procedures.. |  |  | February 16, 2015 | 2015 | Brainlab AG | The ExacTrac 6.x Patient Positioning System may incorrectly position the patient when using the ExacTrac Cone Beam CT (CBCT) with a TrueBeam-specific optional subvolume-CBCT. | Software in the Use Environment | BrainLab sent an FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letters dated February 16, 2015, to all affected customers. The letters included instructions for customers to: 1) Do not use any actively re-reconstructed CBCT subvolumes from the Varian TrueBeam system with the Brainlab ExacTrac 6.x CBCT Import &amp; Alignment Software module; 2) Import exclusively original, not modified, CBCT volumes into ExacTrac to use for patient positioning; and, 3) Continue to verify any ExacTrac CBCT-based correction using the ExacTrac X-ray verification and/or an external IGRT procedure as mandatory. Brainlab provided existing potentially affected ExacTrac v.6.x CBCT Import &amp; Alignment Software customers (with a Varian TrueBeam system) with the product notification information. Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact the customer starting August 2015 to schedule the update installation.For further questions please call Customer Hotline + 49 89 99 15 68 44 or 1-800 597-5911 ( For US Customers only) |  | 24 | 24 systems (US); 36 systems (Foreign) | Worldwide Distribution - USA and to the countries of : Australia, Austria, Canada Denmark, France, Germany, Hong Kong, Japan, Russia, Singapore, South Korea and United Kingdom. | Open, Classified |  |  | 24 | Not_Computer
Z-1317-2015 | 70593 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | BD Q-Syte Luer Access Split Septum, 0.16 ml, Sterile, Rx only, Catalog REF #385100, Manufacturer: Becton Dickinson Infusion Therapy Systems Inc. |  |  | February 19, 2015 | 2015 | Becton Dickinson &amp; Company | Becton Dickinson received an elevated number of complaints for blood or infusate/medication leakage due to tearing of the silicone septum. | Process control | Becton Dickinson sent an Urgent Voluntary Product Recall Letter/Recall Response Card/Packing Instructions, dated 2/19/15, to their affected customers via UPS next day mail. The letters identified the affected device and explained the reason for the recall. Customers were to review their inventory for the specific catalog and lot numbers provided and quarantine the affected product. The Recall Response Card form should be completed and faxed back to BD at 1-866-514-2113 or e-mail bd7043@stericycle.com. All affected product should be returned with the completed Recall Response Card following the steps on the enclosed packing instruction. Upon receipt of the returned product, BD will send customers replacement product. If customers have questions or require assistance, they should contact 1-866-800-2917. |  | 315800 | 315,800 units | Distributed to the states of CA, CT, FL, IL, MA, ND, NY, OH and PA. | Terminated | May 25, 2017 | 826 | 315800 | Not_Computer
Z-1320-2015 | 70355 | No | OWD |  |  | Molecular Genetics | PMA | cobasÂż EGFR Mutation Test |  |  | January 13, 2015 | 2015 | Roche Molecular Systems, Inc. | False positive results for Exon 20 insertion mutations are being detected with the cobasÂż EGFR MutationTest, kit batch T08661. | Under Investigation by firm | RMS notified their customers on 1/13/2015. |  | 190 | 190 kits | Foreign | Open, Classified |  |  | 190 | Not_Computer
Z-1318-2015 | 70736 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | DUAFIT(R) PIP 0o CA SIZE 2, REF A60 SP002, STERILE R, Rx only. |  |  | March 09, 2015 | 2015 | In2Bones, SAS | Dimensional non-conformity leading to the impossibility to introduce the implant in the dedicated holder. | Nonconforming Material/Component | The firm notified their sole consignee in the US by email on 03/09/2015 requesting the return of the product, and to notify his consignee(s) to quarantine the unused product. |  | 35 | 35 units | Distributed in the state of TN. | Open, Classified |  |  | 50 | Not_Computer
Z-1321-2015 | 70410 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry Systems Drug Calibrator I; For in vitro diagnostic use in the calibration of Phenobarbital_2 (PHNB_2), Phenytoin_2 (PHNY_2), and Theophylline_2 (THEO_2) methods on the ADVIA Chemistry systems. |  |  | January 19, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Incorrect calibrator lot values were assigned for ADVIA Chemistry Systems Drug Calibrator I, lot 3JD018, for the ADVIA Chemistry Systems Phenytoin 2 (PHNY_2) and Phenobarbital 2 (PHNB_2) assays which was the result of calibrator value miscalculation. | Device Design | Siemens issued an Urgent Medical Device Correction, on/about January 19, 2015, notifying customers about the product, problem, and action to be taken. Customers were instructed to update the calibrator values for Drug Calibrator 1, Lot 3JD018. All customerswere requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within (30) days. Field service personnel have been sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. Customers with questions were instructed to contact their local Siemens technical support representative. For questions regarding this recall call 914-631-8000. |  | 657 | 657 | Worldwide Distribution - US (nationwide) and Internationally to Canada and Mexico. | Terminated | September 21, 2015 | 245 | 657 | Not_Computer
Z-1324-2015 | 70576 | No | OVE | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Endo TCS Driver Straight Hexalobe #6 P/N 5210-1004, Non-Sterile, Manufactured by Titan Spine, LLC. |  |  | February 20, 2015 | 2015 | TITAN SPINE, LLC | The TCS Straight Hexlobular #6 Drivers (5210-1004) tip sheared off during use in procedures that involved patients with dense bones. As a result of this malfunction, Titan Spine, LLC is removing all of the TCS Straight Hex lobular #6 Drivers (5210-1004) Lot# G150101 from the field. | Nonconforming Material/Component | Titan Spine sent an "Urgent Medical Device Recall" letter via Fed Ex dated February 20, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.A replacement TCS Straight Hex lobular# 6 Drivers (5210-1004) will be shipped out to each Titan Spine Area Vice President and Sales Rep. within 4-6 weeks.For further questions please call (262) 242-7801. |  | 26 | 26 | US Distribution to the states of : AZ, CA, CT, FL, GA, ID, IL, KY, MD, MO, NV, NY, OH, OK, OR, TX, UT and WI. | Terminated | April 17, 2015 | 56 | 26 | Not_Computer
Z-1327-2015 | 70599 | No | HGY | Nonpowered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Ameda Dual Short Term HygieniKit Milk Collections System (SKU 17115) - Primary packaging is 12 pouches with Tyvek header (10.75 x 16.375 x 3") within a secondary corrugated shipper (20.75 x 15.50 x 8.5" inner dimensions). |  |  | February 12, 2015 | 2015 | Ameda, Inc. | Devices were not sterilized | Under Investigation by firm | Ameda sent an URGENT: MEDICAL DEVICE RECALL letter dated February 12, 2015 to all affected customers. The letter included instructions for distributors to: 1) quarantine the recalled products; 2) complete and return the enclosed Recall Response Forms within 3 days of receiving the letter; 3) upon receipt of the completed Recall Response Form, an Ameda representative would contact the distributor to make arrangements for the return of the recalled products; 4) Ameda will use the information provided in the returned Recall Response Form to contact the distributor's customers of record to arrange for the return of recalled product that was already shipped to the customers of record; and, 5) replacement product will be shipped to both distributors and their customers of record upon receipt of the completed Recall Response Form. Questions about this recall will be answered by Ameda personnel who can be reached at 847-964-2645. |  | 12 | 12 collection systems | Nationwide Distribution including CA, FL, GA, IA, IL, IN, LA, MA, MD, MO, MS, NC, NJ, NY, OK, PA, RI, SC, TN, TX, UT, VA, and WA. | Terminated | July 21, 2017 | 890 | 1775 | Not_Computer
Z-1323-2015 | 70594 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Coblator II System, 120V. |  |  | February 20, 2015 | 2015 | ArthroCare Medical Corporation | The recalling firm was notified by the contract manufacturer that they had reversed the L2, L3, L4, or L5 components on the P/N09010 RoHS boards of the Coblator II Controller (Model RF8000E). | Packaging process control | ArthroCare sent an Urgent Medical Device Recall letter dated February 2015 to customers. A response form will be included in the letter and a follow up notifcation letter will be sent to non-responding consignees. The letter identified the affected product, problem and actions to be taken. For questions contact your local Returns Department Representative at 800-343-5717, option 3. |  | 15 | 15 units | US Nationwide Distribution in the states of TN, IA, AZ, NE, SD, NY, and OK. | Terminated | February 26, 2016 | 371 | 15 | Not_Computer
Z-1325-2015 | 70605 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen TRABECULAR METAL FEMORAL AUGMENT BLOCK DISTAL ONLY, 20 MM Augment, Size D, Sterile, Rx Only. |  |  | February 04, 2015 | 2015 | Zimmer Trabecular Metal Technology, Inc. | One unit of the 20 mm augments was mislabeled as 5 mm and it was subsequently distributed. | Error in labeling | Zimmer Trabecular Metal Technology sent an email dated 12/10/2014 to their sole customer. The customer was requested to return the affected product. |  | 1 | 1 unit | US Distribution in the state of CA | Terminated | July 16, 2015 | 162 | 1 | Not_Computer
Z-1336-2015 | 70632 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific: Expel" Drainage Catheter with Twist-Loc" Hub APDL Drainage Catheter and APDL Drainage Catheter Kit, Expel" Drainage Catheter APD Drainage Catheter, Expel" Large Capacity Drainage Catheter APD Large Capacity Drainage Catheter. |  |  | February 25, 2015 | 2015 | Boston Scientific Corporation | Boston Scientific is initiating this Medical Device Field Correction of the Expel" APD and APDLDrainage Catheters. To date, Boston Scientific has received one complaint for device fragmentationafter an Expel" APD catheter was implanted in the biliary system. The most serious foreseeable patientrisk for this issue is additional intervention for fragment retrieval using minimally invasive meth | Nonconforming Material/Component | Boston Scientific sent an "Urgent Medical Device Correction" letter dated February 25, 2015, to all affected customers. The letter described the problem and the product involved in the recall. The letter informed customers not to use the Expel APD and APDL Drainage catheters for bile drainage and to complete and return the Customer Acknowledgement Form. The letter also noted that the devices are not being retrieved and they are not required to return them to Boston Scientific. For further questions please call (763) 494-7971. |  | 918 | 918 | Worldwide Distribution - US Distribution to the states of : CA, LA, FL, WI, NY, IL, MD and OH., and to the countries of: Belgium, Czech Republic, Finland, France, Germany, Great Britain, Italy, Spain, Switzerland, Austria, Hungary, Netherlands and Portugal. | Open, Classified |  |  | 918 | Not_Computer
Z-1348-2015 | 70721 | No | FYD | Air-handling apparatus for a surgical operating room. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LASER RESISTANT WAND - STERILE, 7/8" X 8" (2.22cm x 20.32cm), REF 24013, QTY 10 EA, STERILE EO, NO LATEX |  |  | March 10, 2015 | 2015 | Microtek Medical Inc | This recall is being conducted because a cohesive band used to package the devices contains latex and the product is not labeled as required by 21 CFR 801.437; it does not contain Caution: The Packaging of This Product Contains Natural Rubber Latex Which May Cause Allergic Reactions. | Error in labeling | The firm, Ecolab - Microtek Medical, sent an "Urgent: Recall - Mislabeled Product Immediate Action Required" letter dated March 2015 to its US customers on 03/10/2015 via letter and email. The letter describes product, problem and actions to be taken. The customers were instructed to immediately examine your inventory and promptly quarantine the lot(s) listed; complete and return the enclosed Recall Response Form via fax to Quality Assurance Manager, Ecolab-Microtek Medical at (662) 244-3011 or email to: recall@ecolab.com RE: Recall - Mislabeled Product ASAP, and if you have further distributed the products subject to this recall, please notify your customers at once. Upon receipt of the completed response forms, Ecolab-Microtek Medical will issue a return material authorization (RMA) and ask that you return any unopened product listed above to Ecolab-Microtek Medical at the address listed in the attached form.If you have any questions in regard to this recall, please contact Customer Service by phone at 1-800-824-3027 or email; customerservice@microtekmed.com. |  | 6540 | 6540 units | US (nationwide) Distribution. | Open, Classified |  |  | 145758 | Not_Computer
Z-1350-2015 | 70704 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Skintact ECG Electrode FS-50, Catalog no. 58028 , &amp; Stable Base ECG Electrode SBT-60, Catalog no. 50547. |  |  | February 24, 2015 | 2015 | Leonhard Lang Medizintechnik GmbH | The electrodes may fail to transmit signals. | Nonconforming Material/Component | Consignees were sent via e-mail a Leonhard Lang "Important Safety Notice" letter, dated 24 February 2015. The letter provided the Product trade name, Affected lot numbers, Required Action, Target audience, Details of the affected product which included Description of the problem and Measures to be taken by the user. The letter advised consignees to quarantine the affected product, and to complete and return the acknowledgment form to Leonhard Lang. For questions contact either Sophie Hurmann, Sales Department (tel: +43 512 33425 - 45, fax: +43 512 39 22 10, e-mail: sophia.hurmann@leonhardlang.at) , Stefan Rudig, Quality Department ,(tel: +43 512 33425 -15, fax: +43 512 39 22 10 e-mail: stefan.rudig@leonhardlang.at), or Bernhard Ladner,Qualitats Manager (tel: +43 512 33425- 39, fax: +43 512 39 22 10, e-mail: bernhard.ladner@leonhardlang.at) . |  | 364800 | 364,800 | Distributed to FL. | Terminated | July 26, 2017 | 883 | 364800 | Not_Computer
Z-1353-2015 | 70609 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | Vitagel; indicated in surgical procedures (other than in neurosurgical and ophthalmic) as an adjunct to hemostasis when control of bleeding by ligature or conventional procedures is ineffective or impractical. |  |  | February 26, 2015 | 2015 | Orthovita, Inc., dBA Stryker Orthobiologics. | Several shipments of Vitagel product (a surgical hemostat) required to be stored at 2 -8 degrees Celsius (C) were not delivered to customers within the validated delivery time. | Storage | Orthovita, Inc., dBA Stryker Orthobiiologics made phone notifications in February 2015 to all affected customers. Customers were told that the Vitagel they ordered on x date has an internal procedural mandate that says they must deliver in 35 hours and it has now passed that window. As a precaution until we can confirm that Vitagel delivered after shipping for longer is safe for use we are asking that they quarantine the order in question in the event the delivery occurs.For questions regarding this recall call 610-640-1775. |  | 274 | 274 | Nationwide Distribution including NY PA WV MI IL MA TX AR KS FL WI CA MO LA VA IA OH and NC. | Terminated | July 28, 2015 | 152 | 274 | Not_Computer
Z-1354-2015 | 70596 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza. A Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images. It supports the physician in diagnosis and treatment planning. |  |  | February 18, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Possibly incomplete archived studies during pre-fetch. In a server farm setup, when pre-fetch/retrieve operation is performed for partially archived studies, the series that have not yet been archived, will remain unarchived. | Software design | Safety Advisory Notice, dated February 18, 2015, was sent to end users that identified the product, problem, and action to be taken. Users were advised not to manually remove the delete protection in the STS. |  | 4 | 4 | US Distribution: MA, WI, FL | Open, Classified |  |  | 4 | Not_Computer
Z-1351-2015 | 70472 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medcomp Duo-Flow 400XL Catheter, Hemo-Flow Catheter, SLX Catheter, Triple Lumen Infusion Catheter, CT rated PICC, Split Cath III, PC Split Cath III; |  |  | February 05, 2015 | 2015 | Medical Components, Inc dba MedComp | Medcomp has initiated the recall of Duo-Flow 400XL Catheter, Hemo-Flow Catheter, SLX Catheter, Triple Lumen Infusion Catheter, CT rated PICC, Split Cath III, PC Split Cath III because the product shipped was out of specification. One BI (biological indicator) out of 21 on the sterilization load was out of specification. | Process control | On or about 02/03/2015, Medcomp sales representatives notified the affected medical facilities and providing Medcomp's recall letter. This was accomplished by e-mail, telephone, or visit. Distributors for the international accounts were notified by Medcomp via e-mail. A copy of Medcomp's recall letter was attached. For AngioDynamics product- Medcomp notified AngioDynamics by e-mail of the product recall. Every consignee was requested to complete and return Page 2 of the recall letter verifying receipt of the recall and disposition of the product they received. AngioDynamics was requested to provide an account of the disposition for the total number of product they received. |  | 3152 | 3152 | Worldwide Distribution. | Terminated | August 06, 2015 | 182 | 3152 | Not_Computer
Z-1352-2015 | 70579 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes 130 Degree Aiming Arm and 125 Degree Aiming Arm; The 130 Degree Aiming Arm and 125 Degree Aiming Arm are a part of the TFN-ADVANCED (TFNA) System. The TFN-ADVANCED Proximal Femoral Nailing System is intended for treatment of fractures in adults and adolescents (12-21) in which the growth plates have fused. |  |  | February 24, 2015 | 2015 | Synthes, Inc. | The guide sleeve has the potential to jam in the Aiming Arms, which may delay disassembly and removal of instruments. | Component design/selection | DePuy Synthes sent an Urgent Notice: Medical Device Recall letter, dated March 2, 2015, to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to review inventory for the affected parts and following the actions provided in the letter for removal and replacement of these devices. For questions call 610-719-5450. |  | 131 | 131 | US Nationwide Distribution - NY, TX, MD, OR, MS, ND, CA, AZ, MI, IL, PA, TN, AR, SD, SC, FL, VA, OK, and GA | Terminated | September 21, 2015 | 209 | 131 | Not_Computer
Z-1356-2015 | 70763 | No | DTE | External pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Dual Chamber Temporary External Pacemaker, Model 5392. |  |  | March 17, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | Possible performance issue when used with specific AA-sized (LR6) batteries. The negative terminal of a commercially available AA (LR6) battery did not maintain a sufficient connection with the battery drawer electrical contact. This issue could prevent the EPG from powering on or cause the EPG to abruptly lose primary battery power, potentially stopping delivery of pacing therapy. | Device Design | Medtronic sent an "Urgent Medical Device Correction" letter dated March 2015 to affected customers. The letter was addressed to Risk Manager or Health Care Professional. The letter described the Issue and Customer Actions to be taken. The letter requested customers to complete and return the Customer Confirmation Certificate. For questions contact your Medtronic representative, for technical questions call Instrument Technical Service at 800-638-1991. |  | 6040 | 6040 devices | Worldwide Distribution - US Nationwide in the states of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MA, MI, MM, MS, MO, MT, NE NV,N, NM, NY, NC, ND, OH,OK, PA, RI, SC, TN, TX, UT, VT, VA, WA, WV, WI, WY, PR, VI and the countries of:Australia, Austria, Bahamas, Belgium, Bolivia ,Botswana, Brazil ,Brunei, Darussalam, Bulgaria, Burkina Faso, Canada, Colombia, Croatia, Cyprus, Czech Republic, Denmark,, Dominican Republic, Egypt, France, Georgia, Germany, Greece, Guyana, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Jordan, Kazakhstan, Korea, Kuwait, Lebanon, Luxembourg, Macedonia, Malaysia, Mauritius, Morocco, Mozambique, Namibia, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Papua, New Guinea, Peru, Philippines, Poland, Portugal, Qatar, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, St. Pierre and Miquelon,Sweden, Switzerland, Tanzania, Thailand, Trinidad and Tobago, Tunisia, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, Viet Nam, and Yemen.. | Terminated | October 08, 2015 | 205 | 6040 | Not_Computer
Z-1357-2015 | 70692 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | RECLAIM Assembled Implant Inserter Adaptor orthopedic manual surgical instrument for hip implantation. |  |  | February 23, 2015 | 2015 | DePuy Orthopaedics, Inc. | This lot of the RECLAIM Assembled Implant Inserter Adaptor Instrument may fracture during surgery and there is the potential for plastic pieces to be left in the patient. | Under Investigation by firm | The firm issued a voluntary device removal on 02/23/2015 for lot no. SO2018074 of the RECLAIMÂż Assembled Implant Inserter Adaptor. The recall notice informs US Distributors and Medical Professionals about the issue and any actions to take. US Distributors will return affected instruments and notify one Hospital/User Facility with the Hospital Notice and Reconciliation Forms. Reconciliation forms are required for the end-user Hospital/Medical Facility Consignee. Customers may contact Mindy K. TinsleyDirector Wordwide CommunicationsTel: (574) 372-7136Fax: (574) 372-7717E-mail: mtinsley@its.jnj.com |  | 54 | 54 of Lot No: SO2018074 | Worldwide Distribution: US (nationwide) in CA; Ireland and Israel. | Terminated | September 28, 2015 | 217 | 54 | Not_Computer
Z-1355-2015 | 70604 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Various Lots of Expel Drainage Catheter with TwistLoc Hub Biliary Drainage Catheter and Biliary Drainage Catheter Kits |  |  | February 25, 2015 | 2015 | Boston Scientific Corporation | Complaints of device fragmentation after the catheter was implanted in the biliary system. The most serious patient risk for this issue is additional intervention for fragment retrieval using minimally invasive methods and, in some of the reported cases, this was performed to retrieve the device fragments. | Device Design | Boston Scientific sent an Urgent Medical Device Correction - Immediate Action Required letter dated February 5, 2015. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter described the problem and the product involved in the recall. Advised consignees to complete and return the "Recall Removal Reply Verification Tracking Form".Your local Sales Representative can answer any questions you may have regarding this Field Correction notice.For further questions please call (763) 494- 7971. |  | 289 | 289 | Worldwide Distribution - US Distribution to the states of : CA, WI, FL and OH., and to the countries of: Belgium, Czech Republic, Finland, France, Germany, Great Britain, Italy, Spain, Switzerland, Austria, Hungary, Netherlands and Portugal. | Terminated | June 30, 2017 | 856 | 289 | Not_Computer
Z-1359-2015 | 70683 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-ray Systems (Brilliance CT 64-channel w/Essence technology, Brilliance iCT, Brilliance iCT SP, Ingenuity Core, Ingenuity Core128 &amp; Ingenuity CT), Philips Medical Systems, Cleveland, OH |  |  | February 05, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Philips discovered that a software defect exists in marketed product wherein the sign indication of the longitudinal position of some types of scan is inverted. | Software design | Philips sent an Customer Information letter dated February 5, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They informed their customers that they should refrain from using Multi Planar Reconstruction (MPR) derived images for identifying the absolute position of the intervention for feet-first patient orientation. A Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, please contact your local Philips representative.Philips apologizes for any inconveniences caused by this problem. For further questions please call (440) 483-7600. |  | 424 | 424 units | Worldwide Distribution - Nationwide Distribution to the states of : AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV and WY., and to the countries of : Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Canada, Cambodia, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jumeirah, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Training, Tunisia, Turkey, United Arab Emirates,United Kingdom, Uzbekistan and Viet Nam. | Open, Classified |  |  | 424 | Not_Computer
Z-1361-2015 | 70610 | No | BZA | Airway connector. | Anesthesiology | Anesthesiology | 510(K) Exempt | Double Swivel Connector with Suction Port |  |  | February 17, 2015 | 2015 | King Systems Corp. | Product is mislabeled as Double Swivel Connector with Suction Port and is actually Double Swivel Connector without Suction port. If the incorrect component is not identified prior to clinical use, the incorrect connector would be replaced during the case without risk to the patient and could result in a short delay of the procedure. | Mixed-up of materials/components | Kingsystems sent an Urgent Medical Device Recall letter dated February 17, 2015 to affected customers. The letter identified the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call your King Systems customer service representative at 800.642.5464. |  | 900 | 900 units | US Nationwide in the states of IL, IN, KY, MI, PA, and TN | Terminated | October 27, 2015 | 252 | 900 | Not_Computer
Z-1358-2015 | 70666 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Bard 100% Latex-Free Urinary Drainage Bag with Anti-Reflux Chamber and Bard EX-Lok Sampling; Product Code 154006: Lot # NGYJ0603 and NGYJ1671 |  |  | January 30, 2015 | 2015 | C.R. Bard, Inc. | Potential breach of the sterile barrier packaging. | Unknown/Undetermined by firm | Bard sent an Urgent Medical Device Recall letter on January 30, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter also requested that a sub-recall should be conducted if the product was further distributed. The letter also included a response form which is to be returned to Bard.This voluntary recall is only for the lots noted above with the affected product catalog number. Notifications of this recall are being sent to all affected accounts of Bard Medical Division, Inc. If you have further distributed this product, you are required to notify your accounts of the voluntary recall. If you have any questions or need assistance in notifying your accounts about the correction, please call 1-770-784-6471. We appreciate your cooperation and assistance in dealing with this matter and sincerely apologize for any inconvenience that may result from this action. |  | 9620 | 9,620 units | Nationwide Distribution to the states of : AL, AZ, CA, FL, GA, IL, IN, KY, LA, MD, MA, MI, MN, MS, MO, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, VA, WV and WA. | Open, Classified |  |  | 9620 | Not_Computer
Z-1360-2015 | 70676 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Brand Name: PersonaÂż Anterior Referencing Sizer with Locking Boom |  |  | March 10, 2015 | 2015 | Zimmer, Inc. | Zimmer, Inc is voluntarily recalling 22 lots of PersonaÂż Anterior Referencing Sizer with Locking Boom, part number 42-5099-088-10. Devices are being removed from distribution due to nonconforming to print specifications. | Nonconforming Material/Component | Zimmer sent an Urgent Medical device recall letter to all customers on March 10, 2015, via letter and email. The letter identified the product, the problem, and the action to be taken by the customer. Affected units will be inspected. Units that fail inspection will be physically destroyed/scrapped. Units that pass inspection will be re-etched and redistributed. If after reviewing this notification customers have further questions or concerns they were instructed to call the customer call center at 1-877-946-2761. Hours of operation are Monday through Friday, 8:00 a.m. through 5:00 p.m. EST.For questions regarding this recall call 574-372-4487. |  | 688 | 688 | Worldwide Distribution - US including Domestic: Puerto Rico, AZ, NY. OR, ME, NC, OH, TX, CO, IL, TN, PA, HI, CA, KS, WI, VA, MN, NJ, UT, MA, GA, MI, OK, MS, FL, AL, IN, and Internationally to Germany, India, Dominican Republic, Singapore, Korea, Japan, Austrailia, Canada. | Terminated | November 06, 2015 | 241 | 688 | Not_Computer
Z-1364-2015 | 70667 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ziploop Button |  |  | February 26, 2015 | 2015 | Biomet, Inc. | The firm is recalling ToggleLocs and Ziploops manufactured in May and July of 2014 as they were not processed following a standard work process, which resulted in complaints of sterile packaging not being sealed. | Employee error | Biomet sent an Urgent Medical Device Recall Notice letter dated March 4, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediate locate and discontinued use of the product and its return to Biomet. Customers may contact Field Action Specialist, Regulatory Compliance,at (574) 372-1570. |  | 50 | 50 Units | Worldwide Distribution: US (nationwide) Distribution in the states of:TX, GA, NM, NC and OK and the countries of: Japan, and the Netherlands. | Terminated | January 12, 2016 | 320 | 50 | Not_Computer
Z-1362-2015 | 70630 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | GAP Nail 4.8mm (diameter size) of the GAP ENDO-EXO MEDULLARY SYSTEM. Catalog number GAP-N48-16; GAP-N48-18; GAP-N48-26; GAP-N48-28. |  |  | February 20, 2015 | 2015 | Pega Medical Inc. | Failures of the 4.8mm of the GAP Endo-Exo Medullary System have been identified at the level of the distal cortical screw holes. Failures mainly occurred in patients with tibial pseudoarthrosis (CPT) or Osteogenesis imperfecta (OI). Higher failure rates have been detected in tibias with distal fractures that are close to the distal locking holes. | Other | Pega Medical sent the Urgent Field Safety Notice-Recall/Advisory Notice letter, dated 2015/02/20, to US consignees (surgeons) via email.The following actions were advised:1. It is recommended to monitor patients as considered necessary in order to ensure that weight bearing does not take place until full consolidation of the fracture/osteotomy is completed and that weight bearing is applied progressively.2. Report to the distributor or Pega Medical any adverse event observed. In cases where the consolidation has occurred, normal patient follow up can be continued with no serious consequences to the patient. In cases where consolidation has not been achieved after six weeks, replacement of the nail might be necessary.3. All surgeons receiving this notice must acknowledge reception by a written confirmation sent to the distributor or to the contact reference person indicated below.US consignees with questions can call Pega Medical at 450-688-5144 ext. 242, or send fax to 450-688-1977, or send email to egarcia@pegamedical.com. |  | 7 | 7 units distributed in the US; | US Nationwide Distribution in the states of: California, Florida, New York, and Puerto Rico (US). | Terminated | April 30, 2015 | 69 | 7 | Not_Computer
Z-1365-2015 | 70846 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Covidien Trellis 6 and Trellis 8 Peripheral Infusion Systems, Sterile EO, Rx only. Models BVT608010V01, BVT608030V01, BVT612010V01, BVT612030V01, CVT808015, CVT808025, CVT812015, CVT812025, EVT808015, EVT808025, EVT812015, EVT812025, BVT808015, BVT808030, BVT812015, BVT812030, EUT808015, EUT808030, EUT812015, EUT812030. |  |  | March 23, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic is conducting a voluntary recall of all former Covidien Trellis 6 and Trellis 8 Peripheral Infusion Systems due to potential for a sterility breach of the outer packaging or pouch material. A breach of the outer pouch may compromise the outside surface sterility of the inner pouch and does not directly affect the sterility of the device components within the inner pouch. | Package design/selection | Consignees were sent on 3/23/2015 a Medtronic "Urgent Medical Device Recall" letter dated March, 2015. The letter described the problem and the product involved in the recall. Consignees were advised to quarantine and return the product and to complete and return the Customer Confirmation Certificate to Medtronic. For questions contact your Medtronic Sales Representative or Customer Service a 1-800-716-6700. |  | 8269 | 8,269 ( 6,786 US, 1,483 OUS) | Worldwide Distribution: US (nationwide) including all states in continental USA, except DE and including DC and PR; and countries: AUSTRIA, AUSTRALIA, NEW ZEALAND, BELGIUM, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, IRELAND, ITALY, NETHERLANDS, POLAND, PORTUGAL, ROMANIA, RUSSIA, SPAIN, SWEDEN SWITZERLAND, TURKEY, UNITED KINGDOM, CHILE, EL SALVADOR, PANAMA, and CANADA. | Open, Classified |  |  | 8269 | Not_Computer
Z-1376-2015 | 70677 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aespire 7900, Aespire View Anesthesia Machines. Model numbers 1009-9012-000. |  |  | February 27, 2015 | 2015 | GE Healthcare | Specific Aespire 7900 Anesthesia Machines have an increased likelihood of failure of the CPU circuit board. This could result in a blank display or loss of mechanical ventilation. | Process control | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 34064, dated February 27, 2015. The letter was addressed to Chief of Anesthesia, Health Care Administrator / Risk Manager and Director of Biomedical / Clinical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The letter stated that customers can continue to use the affected device in Manual Ventilation mode, and contact their local GE Healthcare Service Representative. GE Healthcare will correct all affected devices at no cost to the customers. A Service Representative will contact customers to arrange for the correction. For questions or concerns contact GE Healthcare Service at 1-800-437-1171, or your local Service Representative. |  | 46 | 46 (10 units US, 36 units OUS) | Worldwide Distribution -- US, including the states of NC and OK; and the countries of CHINA, INDIA, JAPAN, MEXICO, and VENEZUELA. | Terminated | April 16, 2015 | 48 | 46 | Not_Computer
Z-1375-2015 | 70811 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | CD123 (9F5) PE |  |  | March 23, 2015 | 2015 | BD Biosciences, Systems &amp; Reagents | One lot of BD CD123 PE (ASR) has been determined to contain a low amount of CD4 antibody and may result in an unexpected staining pattern. | Under Investigation by firm | Recall letters were sent to affected sites on March 23, 2015 by certified mail. All product in house was placed on ship hold and remaining inventory is in process of being scrapped. |  | 9 | 9 | US Distribution in states of: MO, GA, AZ, NE, TX, VA | Terminated | February 26, 2016 | 340 | 9 | Not_Computer
Z-1371-2015 | 70820 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Fem-Flex II Femoral Arterial Cannula 8, 10, 12 French,Sterile EO, Rx Only, Manufacturer Edwards Lifesciences LLC, Irvine, CA.-Model Numbers: FEMII008A, FEMII008AT, FEMII008V, FEMII010A, FEMII010AT, FEMII010V, FEMII012A, FEMII012AT, and FEMII012V. |  |  | March 23, 2015 | 2015 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling Fem-Flex II Pediatric Femoral Arterial Cannula sizes 8, 10, 12 French because of the potential of tissue damage caused by a protruding wire located at the tip of the cannula. | Process change control | The firm, Edwards, sent an "URGENT - PRODUCT RECALL - ACTION REQUIRED" letter dated March 23, 2015 via FED EX on March 27, 2015 to their customers. The letter described the product, problem, and actions to be taken. The customers were instructed to review entire inventory for the lots listed; complete and return attached acknowledgment form via fax to Edwards Customer Service at 800.422.9329. within three days of receipt of this Field Safety Notice; contact Customer Service at 800.422.3278 to obtain an RGA number and replacement product; return affected product to Edwards Lifesciences, Attn: Cirilo Chaparro, 12050 Lone Peak Drive, Draper, UT 84020, Attention: RECALL, RGA#XXX; and transfer this notice to other organizations if the affected devices have been transferred to any another facilities.If you have any questions that have not been answered by this letter, please call Edwards Customer Services at 800.424.3278 from the hours of 6:00AM - 4:30PM PST; Edwards Customer Service at (800) 268-3993 from 8:00AM  4:30PM Eastern Time or contact your Edwards sales representative concerning the recall. |  | 5125 | 5125 units | Nationwide Distribution. | Terminated | April 17, 2017 | 756 | 5125 | Not_Computer
Z-1374-2015 | 70445 | No | FGY | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Bio-CORE Cannula (Needle) Set |  |  | January 23, 2015 | 2015 | Biologic Therapies, Inc | Devices are misbranded - failure to properly classify and obtain FDA clearance prior distribution | Use error | On January 23, 2015 the firm sent an URGENT: Class II MEDICAL DEVICE RECALL notification asking its consignees to: 1) Quarantine all remaining devices in their inventory and return the product to Biologic Therapies Inc. via UPS Ground. Product returns should be shipped under UPS Account #A197E9 and must reference RMA#15-0115-01. Please ship to the Biologic Therapies Inc. address listed at the bottom of Page 5. To coordinate return shipping, please contact Biologic Therapies, Inc. via the contact information provided.Contact hours: Monday - Friday 8am to 5pm Eastern Standard Time. (352) 304-5149 2) Adverse reactions or quality problems experienced with the use of this product may be reported to the FDAs MedWatch Adverse Event Reporting program either online, by regular mail or by fax. |  |  | 0 (distributed or released) | US Distribution in states of: MA, CA, WA, NJ, TX, IL, and FL. | Terminated | August 06, 2015 | 195 | 2855 | Not_Computer
Z-1369-2015 | 70775 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ACCU-CHEK Connect Diabetes Management App; Instruction Manual |  |  | October 30, 2014 | 2014 | Roche Diabetes Care, Inc. | Roche Diabetes Care has become aware of an issue with the Accu-Chek Connect diabetes management app that could potentially lead to inaccurate bolus advice being provided to the user. A thorough investigation of the situation revealed that this issue may occur if the user changes the screen orientation of the phone from portrait to landscape or vise versa while looking at the Bolus Advisor or Carb | Software design | Urgent Field Safety Notices were sent to the customer and healthcare professional starting on 10/30/2014. The letter provided a description of the reason for the notification; and recommended that all users should update their app immediately. |  | 72772 | The application has been downloaded total of 644 times (Italy-219, South Africa-24, Germany-401) and there are 113 bolus advice activations. | There are no US distributors as the affected device is not marketed in the US. The affected device was distributed in Germany, Italy and South Africa. The application has been downloaded total of 644 times (Italy-219, South Africa-24, Germany-401) and there are 113 bolus advice activations. | Terminated | June 08, 2015 | 221 | 72772 | Not_Computer
Z-1367-2015 | 70571 | No | LXS | Nd:YAG laser for posterior capsulotomy and peripheral iridotomy. | Ophthalmic | Ophthalmic | 510(k) | YC-1800 OPHTHALMIC YAG LASER SYSTEM; |  |  | March 06, 2015 | 2015 | Nidek Inc | Dear Doctor letter and Quick Reference Guide is being sent to users of the YC-1800 to reiterate important safety and cleaning information after complaints of pitting lens that may cause unclear images. | Process control | Dear Doctor letters were sent along with a Quick Reference guide to all affected sites in February. |  | 799 | 799 units | Nationwide Distribution. | Terminated | September 03, 2015 | 181 | 799 | Not_Computer
Z-1368-2015 | 70633 | No | HNL | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | LIS27T - Lacrimal Intubation Set RX Only Sterile |  |  | February 26, 2015 | 2015 | Quest Medical, Inc. | Expiration date incorrectly printed in Manufacturing Date field on both pouch and carton labels | Labeling Change Control | The consignees were notified by letter on 3/2/2015. The recalling firm requested the consignees to return the recalled product for a replacement device. |  | 106 | 106 units | Worldwide Distribution: US (nationwide) including DC and the states of IL, KY, TX, CA, FL, MI, GA, MD, NH, OH, NY, UT, NJ, MO, TN, VA, NC, CO, MI, and PA, and the country of Ireland. | Terminated | February 26, 2016 | 365 | 106 | Not_Computer
Z-1366-2015 | 70665 | No | KZH | Syringe needle introducer. | General Hospital | General Hospital | 510(k) | Autoject E1 Fixed needle device, Model Number AJ1310, UPC Code 384701310010 |  |  | November 24, 2014 | 2014 | Owen Mumford USA, Inc. | The syringe carrier is missing components: a damper and spring. | Process control | Recall letters were issued to customers via FedEx. The letter instructed the customer to check their shelves for the product and specific lot number and to visually inspect (pictures included in letter of a correctly assembled device and one with the missing spring &amp; damper) the device if they have the affected lot. A form was included which was to be filed out and returned. |  | 3100 | 3,100 | US nationwide distribution in the states of AL, AR, CA, FL, GA, IL, MO, NC, OK, TN, and TX. | Terminated | July 31, 2017 | 980 | 3100 | Not_Computer
Z-1377-2015 | 70649 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | CS 8100, CATALOG # (s): 5303045, 5311129, 5303060, 5303110 |  |  | January 08, 2015 | 2015 | Carestream Health, Inc. | Units device head descended unexpectedly | Component design/selection | The firm sent customers "Urgent: Medical Device Recall" letters, dated December 29, 2014. The letter described the problem, as well as, the actions to be taken. A Carestream Health service representative will contact customers to schedule a visit to inspect the affected devices. Questions or concerns should be directed towards the Carestream Customer Care Center in the U.S. at 1-800-328-2910. |  | 3110 | 3,110 units (863 domestically &amp; 2,247 internationally) | Worldwide Distribution -- US, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NV, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, and WV; and the countries of India and Switzerland. | Open, Classified |  |  | 3110 | Not_Computer
Z-1379-2015 | 70690 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | ROTROL P Control for ROTEM delta Thromboelastometry System, Manufactured for: Tem Innovations GmbH. |  |  | October 23, 2014 | 2014 | TEM Systems Inc | Some of vials were found to be partially filled. | Nonconforming Material/Component | Tem sent an Dear Customer letter dated October 23, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. You are entitled to return products delivered to you in the event that they are defective or otherwise not in conformity with your order when you received them. If you observe a low level of diluent in a box of Rotrol P lot number 41819801 at your location, please contact us for a free of charge replacement box from a new lot of Rotrol P.You can reach us at:Email: support@roteminc.comPhone: 919-941-7777, Option 2Toll Free: 1-866-597-1652, Option 2We apologize for any inconvenience you may experience. |  | 7870 | Total 1574 x 5 vials per box | Nationwide Distribution | Terminated | August 13, 2015 | 294 | 7870 | Not_Computer
Z-1380-2015 | 70777 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System IS2000, Arm Drape, 20 Pack. |  |  | March 16, 2015 | 2015 | Intuitive Surgical, Inc. | Cloudy/waxy appearance with potential to transfer waxy substance to the patient, and potential for drape to tear. | Packaging process control | Firm issued field safety notice to all customers on March 16, 2015. Field safety notice includes acknowledgement form to be returned to recalling firm. A second letter was sent on May 12, 2015 to inform users that the date that the drape product unaffected by the issue will not be availalble untisl October/November 2015, not May 2015 as stated in the original communication. |  | 442475 | 442,475 total all drape models | Worldwide Distribution-Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China (including Hong Kong), Colombia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, France, Germany, Greece, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Luxembourg, Malaysia, Mauritius, Mexico, Monaco, Netherlands, New Zealand, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, United States, Uruguay, Venezuela, Vietnam. | Terminated | December 03, 2015 | 262 | 442475 | Not_Computer
Z-1394-2015 | 70748 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | INFX-8000V Interventional Angiography System. CAT-880B catheterization table. |  |  | March 11, 2015 | 2015 | Toshiba American Medical Systems Inc | The table base rotational brake that holds the table top may not always be engaged when any force or side impact is applied to the table top. | Device Design | Toshiba sent a Urgent Medical Device Correction letter dated March 11, 2015, to all customers who purchased the CAT-880B catherization table for the Tosiba Infinix X-Ray Interventional System. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to complete and return the attached form and fax it to (877) 349-3054 or email to raffairs@tams.com. Customers with questions were instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968.For questions regarding this recall call 714-730-5000. |  | 28 | 28 units | Nationwide Distribution | Terminated | September 18, 2015 | 191 | 28 | Not_Computer
Z-1399-2015 | 70591 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Chemistry XPT System, Siemens Material Number 10723034, Software Version 1.0.2, Siemens Material Number 11219493, IVD. |  |  | February 11, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed an issue with ADVIA Chemistry XPT software version 1.0.2. The system may cause samples to remain in an Inprocess state. Test resultson a sample that is held Inprocess will not transmit to the LIS. Manual intervention is necessary to complete the processing of the samples that are held Inprocess. | Software design | Siemens Healthcare Diagnostics sent Urgent Medical Device Correction Letters (dated 2/11/2015) via FedEx to the domestic consignees informing users of the software issue with ADVIA Chemistry XPT Software Version 1.0.2. The letters included work instructions to handle the bug. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within 30 days. For questions about in process samples, the event message 02 981 8728, or need assistance with mitigation solutions, please contact your Siemens Customer Care Center or your local Siemens Technical Support Representative. Urgent Field Safety Notice Letters were sent to the foreign consignees (distribution is determined at the country level). Field service personnel were notified of the Urgent Medical Device Correction Letter. |  | 32 | 16 units: 3 Domestic, 13 Foreign | Distributed in the states of AZ, CA, and WA, and the countries of Germany, Italy, Spain, and UK. | Open, Classified |  |  | 32 | Not_Computer
Z-1396-2015 | 70744 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Instrument Kit for use with OsteoSponge SC grafts. The Instrument Kit is used to insert OsteoSponge SC grafts to fill bone defects in the subchondral region of articulating joints. |  |  | March 09, 2015 | 2015 | Bacterin International, Inc. | Instrument Kits for OsteoSponge SC grafts, OsteoLock Facet Stabilization Dowel, and BacFast HD Facet Dowel have incorrect sterilization parameters in the IFU and on the lid of the kits. Sterilization parameters for the Steam Gravity 121 degree C (250 F) cycle were not validated and the dry time for the Steam Pre-Vacuum 132 degree C (270 F) cycle was omitted. | Process control | The firm, Bacterin, sent a "VOLUNTARY PRODUCT REMOVAL NOTIFICATION" letter and Steps for Voluntary Removal instructions, dated 03/03/2015, to their consignees on 3/9/2015. The letter described the product, problem and actions to be taken.The Consignees were instructed to check with your Inventory and Central Supply Departments to determine if you have the Instrument Kit for use with OsteoLock Facet Stabilization Dowel and BacFast HD Dowel listed; immediately remove all affected product and segregate the product in a secure location for return to Bacterin International, Inc.; complete and return the Tracking Verification Form to Bacterin (even if they do not have affected product to return) to listed address, by email tkarlberg@bacterin.com or fax to 1-406-388-3380; indicate returned products on tracking/verification form; contact the firm via email at tkarlberg@bacterin.com or call 1-406-388-0480 ext. 131 from 8am to 5pm Mountain Time to obtain instructions for returning the Instrument Kits to Bacterin; package the returned goods and send to Bacterin International, Inc. 664 Cruiser Lane, Belgrade, MT 59714.Consignees with further questions about recall, please call Regulator Affairs Manager at 406-388-9480 ext. 128 or send email to cming@bacterin.com |  | 63 | 63 units | US Distribution in states of: Alabama, Arizona, Arkansas, Colorado, Florida, Georgia, Illinois, Indiana, Louisiana, Maryland, Michigan, Minnesota, Mississippi, Missouri, Nebraska, New Jersey, New York, North Carolina, Ohio, Pennsylvania, Puerto Rico, Texas, and Virginia. | Terminated | July 01, 2015 | 114 | 63 | Not_Computer
Z-1398-2015 | 70627 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | QuikLYTE Dilution Check: used to monitor and adjust for the dilution ratio on Dimension clinical chemistry systems with the QuikLYTE module. |  |  | February 24, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | A positive bias beyond 5% resulting in a dilution check failure or an unnecessary dilution check correction if the positive bias is greater than 1% and less than 5%. The dilution check correction would be detected by QC; however, if not detected by QC, the calibration curve would be decreased and cause elevated sodium, potassium and chloride patient results. | Under Investigation by firm | An Urgent Medical Device Recall letter, dated February 2015, was provided to all DimensionÂż system customers who received the affected lot to notify them that Siemens confirmed customer complaints of failing dilution checks when using QuikLYTE Dilution Check lot 4MD707. The letter also instructed them to immediately discontinue use and discard all inventory of Dimension QuikLYTEÂż Dilution Check lot 4MD707. Siemens offered a no charge replacement with a non-impacted lot. |  | 5981 | 5981 | Worldwide Distribution. | Open, Classified |  |  | 5981 | Not_Computer
Z-1397-2015 | 70368 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | CooperSurgical TransWarmer Infant Transport Mattress with WarmGel |  |  | May 01, 2013 | 2013 | CooperSurgical, Inc. | Trans Warmer Infant and Transport Mattress not labeled with exp.date may not meet specification | Package design/selection | CooperSurgical sent a formal recall notification letter dated April 4, 2015 to all affected customers. The letter identified the product problem and the action needed to be taken by the customer.CooperSurgical notified consignees affected by this recall via Fedex with confirmed delivery receipt. The letter identified the affected product and detailed the necessary actions to be taken by customers who carried the affected lots in their inventory. CooperSurigcal will replace any of the affected products thatcustomers had on hand. CooperSurgical provided further instruction on the disposition of the product at the time of replacement.Contact CooperSurgical at 203 601 5200 Ext. 3100 if you have questions.Previously CooperSurgical notified consignees affected by this recall on 5/1/13. |  | 4807566 | 4,807,566 | Worldwide Distribution - US Nationwide - in the countries of:ARGENTINA, AUSTRALIA,BAHRAIN, ,BERMUDA,CANADA,COLOMBIA ,ENGLAND,GERMANY,HUNGARY,INDONESIA, KUWAIT,LATVIA,LIBYA,MALAYSIAMEXICO,NEW ZEALAND,SAUDI ARABIA. TURKEY, UNITED ARAB EMIRATES | Open, Classified |  |  | 4807566 | Not_Computer
Z-1400-2015 | 70719 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Surgi-Tip Transducer Cover Sterile (15.2 x 244 cm) polyethylene telescopically-folded cover with attached (3 cm) NeoGuard tip, REF 610-833. Each cover is packaged in a Tyvek pouch with accessories of gel and fasteners. Individual Tyvek pouches are packaged as a box of 12. Rx only. |  |  | February 26, 2015 | 2015 | Civco Medical Instruments Inc | Leaks were detected in the recalled lot which may compromise the sterility of the product. | Process design | On 3/12/2015, the Urgent Medical Device Recall notice was emailed to all of the affected customers. The notice identified the affected product and stated that the sterile covers from the affected manufacturing lot were not meeting quality levels and may leak. Customers were requested to return all of surgi tip covers from the production lot identified for replacement. To schedule the return and arrange for replacement product to be sent, customers were asked to contact Ms. Amy King, Customer Service Team Leader at 319-248-6520 or info@civco.com. |  | 120 | 120 pouches | Distributed to the states of CA, IL, MA, MI, MO, and WI. | Open, Classified |  |  | 120 | Not_Computer
Z-1402-2015 | 70645 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Vitek 2 AST-P640 REF 418 579 Gram Positive Susceptibility Card, 20 cards per container, IVD. |  |  | February 25, 2015 | 2015 | Biomerieux Inc | The outer carton label and package insert for the recalled product incorrectly lists Streptococcus pneumoniae as an organism for intended use. | Employee error | Biomerieux, Incs sent an Urgent Product Correction Notice on February 25, 2015, to all affected customers. The letter identified the product, the problem, and the actions to be taken by the customer. Customers were instructed to do the following: " Please confirm this letter has been distributed and reviewed by all appropriate personnel within yourorganization." For the referenced test kits (VITEKÂż 2 ASTP635, ASTP640 or ASTP641), review the claimed organisms inthe Package Insert antibiotic table to confirm which organisms can be tested." Refrain from using the referenced test kits to determine the susceptibility of Streptococcus pneumoniae." Reference your Laboratory Standard Operating Procedures to identify card type used for specific organisms." Please store this letter with your bioMÂżrieux instrument documentation.Please complete the attached Acknowledgement Form and return it to bioMÂżrieux, Inc.Customers with questions were instructed to contact their local BioMerieux representative.For questions regarding this recall call 314-731-8526. |  | 2053 | 2053 cartons | Worldwide Distribution to United Kingdom and Turkey. NO U.S. distribution. | Terminated | March 22, 2017 | 756 | 13967 | Not_Computer
Z-1404-2015 | 70723 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Pluer-evac Sahara Chest Drainage System, S-1100-08LF, Dry Suction/Dry Seal Control, Rx Only, Teleflex Medical. |  |  | March 13, 2015 | 2015 | Teleflex Medical | Package contains incorrect sized unit: The adult sized S-1100-08LF product contains an infant sized chest drainage system. | Other | Teleflex Medical sent an Urgent Medical Device Recall Notification dated March 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately discontinue use and quarantine any of the affected listed catalog numbers and to return the product. Complete the enclosed acknowledgement form to Teleflex Medical. The form is to be returned even if they have not products.For further questions contact your local sales representative or Customer Service at 1-866 246-6990. |  | 2148 | 2,148 units | US Distribution to the states of : AK, AL, AR, CA, CO, FL, GA, IL, IN, KY, KS, MA, MD, MI, MN, NY, OH, PA, RI, SC, TX, VA WI and WV. | Terminated | August 14, 2017 | 885 | 2148 | Not_Computer
Z-1408-2015 | 70927 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | VasoNova Vascular Positioning System (VPS ) G4 Console |  |  | March 30, 2015 | 2015 | VasoNova, Inc. | A limited number of VPS G4 Consoles are being recalled due to noncompliance with International Electrotechnical Commission Standard 60101-1-1. | Device Design | Teleflex sent an Urgent Medical Device Recall Notification letters dated on March 30, 2015 to all affected customers. The letters identified the affected product, problem and actions to be taken. The risk to users is identified and users are asked to discontinue use and quarantine affected serial numbers. Teleflex sales reps will help with return of product. The Acknowledgeable form should be completed and returned as per instructions. Questions should be directed to 1.877.236.6869. |  | 17 | 17 | US Nationwide Distribution in the states of AZ, DC, PA, IN, PA. GA, NY, KS, CA, VA, NC. | Terminated | April 11, 2016 | 378 | 17 | Computer
Z-1329-2015 | 70668 | [6 empty] | Ebola GP IgX Blood, Serum, Plasma, Cassette, Catalog No. I-101 |  |  | March 13, 2015 | 2015 | Lusys Laboratories, Inc. | LuSys Laboratories is recalling Ebola Virus One Step Test Kits to stop and prevent further use of these devices. The Ebola Virus One-Step Test Kits have not yet been cleared, approved or authorized by the FDA for diagnostics purposes. | No Marketing Application | LuSys Laboratories sent an Urgent Medical Device Recall letter dated March 13, 2015, to all affected customers informing them that the Ebola Virus One Step Test Kits are not approved by FDA for diagnostics purposes. The letter informs the customers of the problems identified and the actions to be taken. Customers were instructed to discontinue use and return the products to LuSys Laboratories. Customers are instructed to fax back the acknowledgement and Receipt Form to 1-858-866-1688.For questions regarding this recall call 858-546-0902. |  | 1015 | 1,015 units | Worldwide Distribution - US including CA and Internationally to Canada, Sierra Leone, and Denmark. | Terminated | May 31, 2017 | 810 | 2275 | Not_Computer
Z-1413-2015 | 70839 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Bard ConQuest PTA Balloon Dilatation Catheter, 12mm Balloon Size, Product Code: CQ-75124. |  |  | March 23, 2015 | 2015 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular is recalling the Bard ConQuest PTA Balloon Dilatation Catheter because it may be at risk of having deflation related issues. | Process control | Bard sent an Urgent Medical Device Recall Notification letter dated March 23, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter informs the customers that the product code and lot combination listed on the attachment 1 of the recall letter may be at risk of having deflation related issues. Customers are instructed to complete and return the Effectiveness Check Form. Customers with questions are instructed to call Bard Peripheral Vascular's (BPV) at (800) 321-4254 option #2 Ext 2501 (M-F 7am-4pm MST) Customers are instructed to fax the recall and effectiveness check form to BPV at (800) 994-6772. |  | 2620 | 2620 units | Worldwide Distribution - US Distribution including Puerto Rico and the countries of : Austria, Belgium, Germany, France, Greece, Iran, Ireland, Israel, Martinique, Netherlands, Portugal, Russia, Slovenia, South Africa, Spain, United Arab Emirates (UAE), United Kingdom, Canada, Korea, Taiwan, Singapore, Thailand, Chile, Guam, Mexico, Malaysia and Australia. | Terminated | August 28, 2015 | 158 | 2620 | Not_Computer
Z-1412-2015 | 70810 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Natus neoBLUE blanket LED Phototherapy system which includes the following items: |  |  | March 16, 2015 | 2015 | Natus Medical Incorporated | neoBLUE blanket LED Phototherapy System is recalled because Natus has received feedback from some customers about early failure of the neoBLUE blanket system fiberoptic pads. These failures involve discoloration and melting of the fiberoptic bundle at the connector that is inserted in the neoBLUE blanket light box. | Other | Natus sent the PRODUCT RELIABILITY ALERT neoBLUE blanket system performance along with an attached Technical Bulletin, Instructions for Use, and Service Manual to consignees via email during the week of March 16, 2015.Natus emailed the Distribution Partner (DP) letter to all Natus InternationalDistribution Partners, requiring that they notify their customers (at Biomedical Department and Neonatal Intensive Care Unit) with copies of the Field Safety Notice, customer letter, the Technical Bulletin, Instructions for Use, and Service Manual.The technical bulletin asks the customer to visually and tactilely inspect the fiber optic cable connector before use, and to call Natus if they notice a failingpad. This procedure will prevent the complete degradation and therapy will not be compromised.U.S. customers should contact Natus Technical Service at 888-496-2887 with any questions regarding this notification. International customers should contact their local distributors or contact Natus Technical Service at (888) 496-2887 or Technical_Service@natus.com. |  | 2611 | 2611 units (1911 units in the US and 700 units outside the US) | distributed worldwide | Open, Classified |  |  | 2611 | Not_Computer
Z-1410-2015 | 70628 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | OSSEOTITE Tapered Certain Implant Rx Only; 3.25 x 18.5 - 15.0 mm. |  |  | November 07, 2014 | 2014 | Biomet 3i, LLC | Endosseous Dental Implants in contact with a residual machine fluid caused the implants to be discolored. | Manufacturing material removal | Biomet 3i sent an URGENT MEDICAL DEVICE RECALL NOTICE dated November 7, 3014 to all affected customers. The letter identified the product, problem, and actions to be taken by the customer. |  | 620 | 620 implants in total | Worldwide Distribution-US (nationwide) including the states of AL, AZ, CA, CO, CT, DE, FL, GA, IL, KY, LA, MA, MD, MI, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, VA, WA and HI and the countries of Austria, Australia, Canada, France, Germany, Italy, Japan, Netherlands, New Zealand, Portugal, Spain, Turkey, Taiwan, and United Kingdom. | Terminated | May 22, 2015 | 196 | 620 | Not_Computer
Z-1414-2015 | 70821 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Replace Select Tapered TiU NP 3.5x10mm, Article No. 29401. |  |  | March 09, 2015 | 2015 | Nobel Biocare Usa Llc | A material used in the packaging of the dental implants may impart a sticky substance on the top surfaces of the implants and in some cases on the first threads. | Nonconforming Material/Component | Nobel Biocare sent an Important - Field Correction letter dated March 6, 2015, to all affected customers and another recall letter dated dated March 10, 2015, to all affected customers to inform them that Nobel Biocare Quality Management Systems observed an issue with a material used in the packaging of their implants. The firm's internal analysis has confirmed that this material is not according to their standards. The letters provide the problems identified and the actions to be taken. Customers were instructed to return the affected product(s) not yet used to Nobel Biocare to receive a replacement product. Customers were also instructed to complete the attached Acknowledgement form and fax it to 714-282-5073 or email it to aeinfo@nobelbiocare.com. Customers with questions were instructed to contact their customer service department.For questions regarding this recall call 714-282-4800. |  | 131 | 131 units | Worldwide Distribution - US (nationwide) and Internationally to Austria, Australia, Belgium, Bulgaria, Canada, Switzerland, Czech Republic, Germany, Denmark, Estonia, Spain, Finland, France, United Kingdom, Greece, Croatia, Hungary, Ireland, Italy, Lithuania, Luxembourg, Mexico, Netherlands, Norway, Poland, Portugal, Romania, Russian Fed., Saudi Arabia, Sweden, Turkey, and United Arab Emirates. | Terminated | December 30, 2015 | 296 | 7594 | Not_Computer
Z-1442-2015 | 70896 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | ProTrack Microcatheter; |  |  | March 30, 2015 | 2015 | Baylis Medical Corp * | Baylis Medical Company is recalling two lots of ProTrack Microcatheter due to the possibility that the catheter may have circumferential defects along its shaft. | Process control | A recall notification, marked "URGENT" will be sent by Baylis Medical via overnight courier to all customers affected by the recall with a fax-back form. The form is to be faxed back to Baylis Medical confirming receipt of the recall notice. As per instructions on the fax-back form, affected devices are requested to be returned. Baylis Medical will replace the returned affacted catheters with a new catheter. |  | 58 | 58 units | US Nationwide Distribution - US Nationwide in the states of MO, MA, OH, CA, AR, MA, FL, GA, CO, TX, KY, NY, MI, WI, OR | Open, Classified |  |  | 58 | Not_Computer
Z-1441-2015 | 70749 | No | KPY | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | Kenex Radiation Shield installed with GE Healthcare Interventional, Surgery, CT and X-Ray Imaging systems. |  |  | March 09, 2015 | 2015 | GE Healthcare | GE Healthcare has recently become aware of an improper installation of a ceiling suspended pendent supporting a radiation shield and/or surgical lamp. The pendent system was not installed per specified installation requirements. There has been a reported incident of a radiation shield falling from the ceiling causing a minor injury. | Equipment maintenance | GE Healthcare sent an Urgent Medical Device Correction letter dated March 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions in USA contact 1-800-437-1171. For other countries contact your local GE Healthcare Service Representative. |  | 797 | 797 (3 US, 794 OUS). | Worldwide Distribution - US Distribution to the states of : CA and MA., and to the countries of : SPAIN, INDIA, JAPAN, AZERBAIJAN, BULGARIA, BELARUS, SWITZERLAND, CZECH REPUBLIC, GERMANY, EGYPT, FINLAND, FRANCE, UNITED KINGDOM, GUYANA, IRELAND, ISRAEL, ITALY, NIGERIA and NORWAY. | Terminated | December 21, 2015 | 287 | 797 | Not_Computer
Z-1443-2015 | 70822 | No | GDZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Bard Peripherally Inserted Central Catheter (PICC) and Dialysis catheter kits containing Safe-Cut Safety Scalpels. |  |  | June 24, 2014 | 2014 | Bard Access Systems | Potential for inadvertent scalpel stick injury involving the safety scalpel. | Component change control | Bard sent an Urgent Notification letter date June 26, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed on the proper device use and safety mechanism activation of the Safe-Cut Safety Scalpel. For questions call coordinator, Shelly Gilbert, at 1-800-290-1689. |  | 51025 | 51,025 units | Worldwide Distribution - US Nationwide | Terminated | November 05, 2015 | 499 | 51025 | Not_Computer
Z-1444-2015 | 70767 | No | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | Hardy Diagnostics cat no. G327 HardyCHROM SS Agar |  |  | March 04, 2015 | 2015 | Hardy Diagnostics | The firm is recalling HardyCHROM SS, Cat. no. G327 Lot no. 15037 due to a report of possible lack of color development in Shigella organisms in the referenced lot. | Employee error | On 03/12/2015 the firm sent recall notification letters to customers. The firm's letter informs customers of the reason for recall, and requests that they discard any remaining stock of the recalled lot.The firm continues to request that even though the customer has been informed of the necessary actions that they still take a moment to complete, sign, and fax back the enclosed form stating their compliance with the above action to the attention of their Quality Assurance Department at (805) 614-9274. The firm stated that the form is required for their records and need to document the full quantity received, used and/or discarded. For replacements or credit to please contact the Customer Service Department at (800) 266-2222 and select option 1. The firm states that replacements will be sent out at no charge to customers, or a credit will be applied to the invoice.For any additional questions the firm states to contact Technical Services at (800) 266-2222, option 2 or e-mail at techservice@hardydiagnostics.com. |  | 840 | 840 (84 pk/10) | Worldwide Distribution -- U.S., including the states of TX, CA, AZ, WA, PR, and IA; and the countries of Canada and NSW. | Terminated | July 13, 2015 | 131 | 840 | Not_Computer
Z-1446-2015 | 70837 | No | LXT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | LOW PROFILE NON-LOCK SCREW; 2.7mm x 22 mm; REF 131227222; |  |  | March 18, 2015 | 2015 | Biomet, Inc. | A Biomet investigation found that the Low Profile Non-Locking Screw 2.7mmx22mm screw (Part Number: 131227222 Lot: RM105E) was anodized green and should have been anodized gold per the print. | Nonconforming Material/Component | On 3/18/2015, Biomet issued an URGENT MEDICAL DEVICE RECALL NOTICE to affected distributors and one medical facility. Notification included the reason for recall, description of affected product, statement of no expected patient health risk, and actions required of the addressee. Questions related to the recall notification should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 45 | 45 | Nationwide Distribution-including the states of CA, IA, NY, FL, PA, UT, NC, SC, KY, CO, MA, GA, WA, TX, OK, and IL. | Terminated | December 11, 2015 | 268 | 45 | Not_Computer
Z-1445-2015 | 70830 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis Q/ Q zen systems; Artis Q/ Q zen systems are dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed include cardiac angiography, neuro angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | March 19, 2015 | 2015 | Siemens Medical Solutions USA, Inc | potential issue with the Artis Q/ Q zen systems where it is possible that an electrical connection in the equipment cabinet has not been installed correctly. The potentially defective connection in the equipment cabinet will be inspected and, if necessary, corrected. | Employee error | A customer notification, dated March 19, 2015, was distributed to all affected customer via AX018/15/S. Additionally, Siemens has initiated an inspection of Artis Q/ Q zen systems via Update Instruction AX017/15/S. |  | 139 | 139 | Nationwide | Terminated | August 06, 2015 | 140 | 139 | Not_Computer
Z-1447-2015 | 70733 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Monopty Disposable Core Biopsy Instruments and Kits, Monopty 14g x 16cm (22mm), Product Code: 121416; Monopty Biopsy, Product Code: 000441; Monopty 14g x 10cm (22mm), Product Code: 121410; Monopty 14g x 10cm (11mm), Product Code: 211410. |  |  | March 17, 2015 | 2015 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular is recalling the Bard Monopty Disposable Core Biopsy Instruments and Kits because they may be at risk of having activation-related issues. | Nonconforming Material/Component | The firm, Bard Peripheral Vascular, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 3/17/15 to customers. The letter informs customers of the problems identified and the actions to be taken. The customers were instructed to do the following: Do not use or further distribute any affected product; check all inventory locations within your institution; if you have further distributed any of the product, advise them of the recall and have them return the affected product to BPV; remove any idenfied product from your shelves; and complete and return the attached Recall and Effectiveness Check Form to BPV by fax at (800) 994-6772, if you cannot FAX the form, call the BPV telephone number provided on the form and report the required information verbally. BPV will provide replacement product for your returned product.Customers are instructed to call BPV's Recall Coordinator at 1-800-321-4254 Option #2 Ext 2988 (M-F 7am to 4pm MST) or email at HUrandy.hopkins@crbard.comUH. |  | 29342 | 29,342 units | Worldwide Distribution: US (nationwide) and countries of: Denmark, Spain, Finland, France, Ireland, Israel, Italy, Netherlands, Poland, Japan and Pakistan. | Terminated | August 03, 2015 | 139 | 29342 | Not_Computer
Z-1448-2015 | 70131 | No | KTO | Immunoglobulin G (Fc fragment specific) immunological test system. | Immunology | Immunology | 510(K) Exempt | SpermMar Test |  |  | December 22, 2014 | 2014 | Vitrolife Inc | The firm is recalling three batches (FP14A09, FP14A10, FP14A11) of SpermMAR IgA test due to new formulation causing the motility of sperm cells to reduce in certain samples when the new bead solution was added. | Component design/selection | FertiPro sent an Urgent: Medical Device Recall letter dated December 22, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to check if they have stock of the concerned lots and provide a confirmation that they have destroyed the affected products.Distributors were asked to inform their customers and instruct them to do the same.For further questions you may call (858) 824-0888. |  | 45 | 45 kits | US Distribution to the state of : OR., and Internationally to Argentina. | Terminated | May 05, 2015 | 134 | 45 | Not_Computer
Z-1449-2015 | 70847 | No | DXX | Steerable catheter control system. | Cardiovascular | Cardiovascular | 510(k) | Magellan Robotic System; Model No(s). 11132 and 11139. A steerable catheter control system used at clinical sites. |  |  | March 18, 2015 | 2015 | Hansen Medical Inc | Based on investigation of two complaint incidents, all Hansen Medical Magellan Robotic System (Model No. 11132 and 11139) may experience an unexpected event during a retraction operation of catheter. Retraction of catheter may not stop when equipment button is released. | Software design | Hansen Medical sent an Urgent Notice - Field Safety Notice dated March 24, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. We request that you review the information in pages 1 and 2 with Hansen Medical personnel and acknowledge your understanding by signing the Field Safety Notice Training Form on page 3.If you need any further information regarding this issue, please do not hesitate to contact me via telephone at 650-404-2784. |  | 19 | 19 total installed systems | Worldwide Distribution - US Distribution and the countries of : Saudi Arabia, Australia, France, Italy, Germany, and United Kingdom. | Terminated | September 23, 2015 | 189 | 19 | Not_Computer
Z-1451-2015 | 70658 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Master Drug Library Software version 8.0, Product Code 35723V080, to be used with SIGMA Spectrum Infusion System (Pump) version 8.0, Product Code 35700BAX2 |  |  | March 02, 2015 | 2015 | Baxter Healthcare Corp | Loading/Bolus default dose settings in the Master Drug Library and the values shown on the pump during programming may differ. MDL drug dose time in seconds will round to the nearest integer in minutes on the pump dose setup screen (20 sec may show as 1 min on the pump display). The pump will administer drugs as configured. The discrepancy may cause therapy delay or unintended rate of delivery. | Under Investigation by firm | The firm, Baxter, sent an "Urgent Device Correction" letter dated 3/05/2015 with MDL Instructions were sent to all affected users via USPS First Class Mail on 3/05/2015. All four (4) limited launch locations with Master Drug Library (MDL) version 8.0 with SIGMA Spectrum Infusion System version 8.0 were notified via phone call on 3/02/2015. The letter describes the product, problem and actions to be taken. Baxter is requesting customers take the following actions: 1. Follow the steps described in Attachment 1 to confirm that this discrepancy is properly corrected. 2. Complete the enclosed customer reply form and return it to Baxter by faxing it to 224-270-5457 or scanning and e-mailing it to fca@baxter.com. --- Baxter sent a follow-up Communication Letter (dated 3/17/2015) to the initial Urgent Device Correction Letter (dated 3/5/2015) to provide additional clarification and instructions to ensure continued mitigation of this discrepancy as users implement the Spectrum infusion System in their facility. The four launch sites received the follow-up letters via hand delivery; remaining customers received the follow-up letter via hand delivery and/or e-mail.If you have additional questions, please contact your Baxter sales representative, Baxter Healthcare Medina at 800-356-3454 (choose option 1) Monday through Friday, during the hours of 7:00 am to 7:00 pm Eastern Time, or The Center for One Baxter at 800-422-9837, Monday through Friday during the hours of 8:00 am to 5:00 pm Central Time. |  | 58 | 58 MDLs | Worldwide Distribution: US Distribution to states of: CA, CT, FL, IA, IL, IN, LA, MA, MD, MI, MN, MO, NC, NE, NH, NY, SD, TX, UT, WV, including Puerto Rico and country of: Canada. | Open, Classified |  |  | 58 | Not_Computer
Z-1452-2015 | 70758 | No | KGN |  |  | General &amp; Plastic Surgery | 510(k) | IntegraÂż Flowable Wound Matrix |  |  | March 13, 2015 | 2015 | Integra LifeSciences Corp. | A lot of Flowable Wound Matrix that has been distributed to customers in late 2013 / early 2014 was verified as not meeting stability acceptance criteria for the attributes of visual appearance during dispensing and force vs. displacement (force needed for product to be extruded from the syringe). | Unknown/Undetermined by firm | Integra Life Sciences Inc. sent a Recall letter/Recall acknowledgement and return form dated March 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Integra asks customers to review inventory and determine if any of the affected inventory is in possession. The attached Recall Acknowledgement and Return Form should be completed and returned. Customer service will contact the customer upon receipt of the form and provide an RMA number. Questions can be directed to Customer Service at 1-888-301-0203. |  | 287 | 287 packages ( 179 US 108 OUS ) | Worldwide Distribution - US Distribution and to the countries of and to the countries of Greece, France, Italy, Spain, Switzerland, South America, Turkey and United Kingdom. | Terminated | April 13, 2016 | 397 | 287 | Not_Computer
Z-1453-2015 | 70872 | No | OCZ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Talon Grasping Device, 160 cm, US endoscopy. Used for retrieval of foreign bodies, tissue specimens, stones or calculi in endoscopic procedures of the gastrointestinal tract. |  |  | March 10, 2015 | 2015 | US Endoscopy Group Inc | A wire component on the distal grasping assembly of the device became detached.. | Under Investigation by firm | On 03/25/2015 the firm sent recall notification letters to their customers via email. |  | 1695 | 1695 units | Product was shipped to the following states: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI &amp; WV. Product was also sent to the following countries: ARGENTINA, AUSTRALIA, BELGIUM, CANADA, ESTONIA, FRANCE, MEXICO, NEW ZEALAND, SAUDI ARABIA &amp; SWEDEN. | Open, Classified |  |  | 1695 | Not_Computer
Z-1450-2015 | 70829 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | WALLACH LL 100 cryosurgical units: LL100 Cryosurgical - Part Number 900001; LL110 with Pinned Yoke - Part Number 900019-70; LL100 Multi Tip W/TC - Part Number: 900629; LL100 - CO2 - Part Number 900162 |  |  | March 09, 2015 | 2015 | CooperSurgical, Inc. | Internal gas line may malfunction causing the line to burst under pressure. | Under Investigation by firm | CooperSurgical sent an Recall letter dated March 13, 2015, shipped via Fedex with confirmed delivery receipt. The letter identified the product the problem and the action needed to be taken by the customer. No adverse events or injuries have been reported to date in connection to this recall. As stated above, the risks associated with the use of this product involve the potential for a hose to burst while under pressure. The potential hazard associated with the failure mode is negligible; there is no foreseen adverse event due to this failure mode. The event was discovered internally during an investigation. Please complete the attached Acknowledgement and Receipt Form to make repair arrangements. If you have any further questions please feel free to contact me at 203.601.5200 Monday through Friday 8:00am to 5:00pm EST. |  | 255 | 255 units | Worldwide Distribution - Nationwide Distribution to the states of : AK, AL, AR, AZ, CA, CO, DC, FL, GA, HI, IL, IN, KY, LA, MA, MD, MI, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA and WV,. and to the countries of : CANADA, COLOMBIA, EL SALVADOR, ENGLAND, GUATEMALA, HONG KONG, INDONESIA, SINGAPORE and ZIMBABWE. | Terminated | February 01, 2016 | 329 | 255 | Not_Computer
Z-1454-2015 | 70856 | No | OYP | John Cunningham Virus serological reagents. | Microbiology | Microbiology | 510(k) | STRATIFY JCV DxSelect, Model No. EL1950. Assay for detection of antibodies to JC Virus in human serum or plasma. In Vitro diagnostic. |  |  | March 15, 2015 | 2015 | Focus Diagnostics Inc | Focus Diagnostics is recalling the Stratify JCV Dx Select due to the use of a non-conforming batch in the manufacture of JCV DxSelect EL1950 kit lot #27333. This could lead to a potential for false positive JCV results. | Nonconforming Material/Component | Focus Diagnostics sent an Urgent Device Customer Notification letter dated March 15, 2015, was sent to the customer to inform them that Focus Diagnostics is recalling the Stratify JCV DxSelect (EL1950) lot#27333 due to the potential for false positive JCV results. The letter informs the customer of the issues and the actions to be taken. The customer is instructed to complete the enclosed acknowledgement form and email the form to DxTS@focusdx.com or fax back to Focus Diagnostics Technical Services at (562) 240-6526 within 10 business days. If the customer has any questions or require additional information, then they are instructed to contact Focus Diagnostic's Technical Services department at (800) 838-4548, select option 3, between the hours of 7am to 5pm (PST) or send an email to DxTS@focusdx.com. |  | 300 | 300 units | Internationally: Denmark only. | Terminated | July 23, 2015 | 130 | 300 | Not_Computer
Z-1455-2015 | 70739 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | DBX Mix, 20 cc, Product Code: 058200. Demineralized Bone Matrix Mix, bone void filler. |  |  | December 04, 2014 | 2014 | Musculoskeletal Transplant Foundation, Inc. | The seal on the DBX Mix outer PETG/Aclar tray with foil lid packaging for lot identified above may not be fully intact. | Packaging process control | MTF notified their customers on 12/4/2014 and 12/8/2014 via phone and letter. |  | 25 | 25 units | US Nationwide distribution to CA, KY, IL, LA, PA, NC, TN, WV, and Puerto Rico. | Open, Classified |  |  | 25 | Not_Computer
Z-1459-2015 | 70722 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Equinoxe 44mm Short Humeral Head |  |  | February 13, 2015 | 2015 | Exactech, Inc. | Incorrectly packaged. Outer and inner labeling may incorrectly identify the size of the enclosed device. | Labeling mix-ups | Exactech Agents sent a Product Recall Advisory Notice dated February 12, 2015 by mailed to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the Recall Inventory Respponse form and return it to Exactech, per the instructions provided on the notice. |  | 112 | 112 devices | Worldwide Distribution - US Nationwide in the states of AK, AZ, CA, CO, FL, ID, ME, MN, NC, OH, OK, PA, RI, TN, VA and WA. and countries of Brazil and Canada. | Open, Classified |  |  | 112 | Not_Computer
Z-1458-2015 | 70754 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Spacelabs Pediatric Flow Sensor Kit, PN: 376-0561-00. This kit is for Spacelabs Healthcare Blease 700/900 Series Ventilators. |  |  | March 05, 2015 | 2015 | Del Mar Reynolds Medical, Ltd. | Reports of inaccurate low flow readings. Monitored inspiratory tidal volume (VTi) and expiratory tidal volume (Vte) measurements from the pediatric flow sensor are reporting out of specification low compared to the actual delivered volumes being administered to the patient. | Vendor change control | Spacelabs started mailing the Urgent Medical Device Correction- Pediatric Flow Sensor kit, PN: 376-0561-00 used with Spacelabs 700 and 900 series ventilators letter, dated 17 March 2015, via priority service to all US customers.Spacelabs will be emailing the customer letter (translated as necessary) to all international subsidiaries and distributors of record on March 20, 2015.Customers are asked to dispose of all pediatric flow sensors purchased from Spacelabs (even if they shipped outside of the time frame in which affected product shipped). Customers are requested to complete and return the Business Reply Form to Spacelabs Healthcare either by fax to Global Tech Support at 425-363-5758or scan the form and email to techsupport@spacelabs.com Spacelabs will send you at no charge replacement pediatric flow sensors (376-0561-01) for all Spacelabs pediatric flow sensors you purchased. These NEW sensors will be marked with a "-01" molded into the body so you can tell the difference between the old sensors and the new ones.For additional information or technical assistance, please contact:Technical SupportSpacelabs Healthcare, Inc.35301 SE Center StreetSnoqualmie, WA 980651-800-522-7025 and select 2 for Technical Support |  | 1040 | 1040 units total (398 in the US and 642 international) | Worldwide Distribution-US including the states of Louisiana, North Carolina, and Washington and the countries of Argentina, Brazil, China, Colombia, Ecuador, Finland, Great Britain, India, Libya, Mexico, Nicaragua, Oman, Philippines, Saudi Arabia, Slovakia, Sri Lanka, United Arab Emirates, and Vietnam. | Terminated | November 16, 2015 | 256 | 1040 | Not_Computer
Z-1464-2015 | 70773 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | DiaSorin, PRO-Trac II Tacrolimus ELISA kit, REF 32400 and 32400-CN. |  |  | March 16, 2015 | 2015 | Diasorin Inc. | Instability in the PRO-Trac II Tacrolimus ELISA kit which may cause falsely elevated patient results. | Under Investigation by firm | DiaSorin sent a Customer Notification Letter" to affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were asked to complete and return the confirmation form to the manufacturer. For questions contact DiaSorin Inc. Product Support at 1-800-328-1482. |  | 451 | 451 kits | Worldwide Distribution - US in the state of CA and the countries of: CHINA, BELGIUM, ARGENTINA, and INDIA. | Terminated | September 15, 2015 | 183 | 451 | Not_Computer
Z-1457-2015 | 70835 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry Lipase Reagent , Catalog Number B01-4840-01 REF: 01984894. |  |  | March 23, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens internal investigation confirmed that current contamination avoidance settings do not prevent carryover resulting in falsely elevated ADVIA Chemistry Lipase outlier results when running with Triglycerides, Triglycerides_2, Triglycerides_c, Cholesterol concentrated, and DLDL reagents. | Device Design | Siemens issued an Urgent Medical Device Correction on March 23, 2015, notifying direct consignees about the product, problem, and action to be taken. Customers were instructed to follow the directions provided in the customer letter in the event that a Lipase outlier is observed. All customers were requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within (30) days. Field service personnel have been sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. For further questions please call (312) 275-7795. |  | 7814 | 7814 | Worldwide Distribution - US (nationwide) Distribution and to the countries of : Thailand, Pakistan, Vatican, Guadeloupe, Reunion, French Polynesia , Serbia, Martinique, India, China, Singapore, Brazil, South Africa, Mexico, Canada, Israel, Australia, Egypt, Russian Federation, Indonesia, Rep. Korea (S) ,Unit.Arab Emir., Taiwan ,Argentina ,Chile, Malaysia ,Colombia, Peru, Kazakhstan, Paraguay, Vietnam, Algeria, Austria, Belgium, Switzerland, Cyprus, Czech Republic, Germany, Estonia, Spain, Finland, France, United Kingdom, Greece, Hungary, Italy, Lithuania, Luxembourg, Latvia, Netherlands, Norway, Poland, Portugal, Romania, Slovenia, Slovakia and Turkey. | Terminated | December 18, 2015 | 270 | 7814 | Not_Computer
Z-1465-2015 | 70832 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Lineum HF Torque Handle Limit (Lineum) Model Number 14-525035 |  |  | March 12, 2015 | 2015 | Biomet Spine, LLC | Biomet is recalling the Lineum HF Torque Handle Limit (Lineum) and Torque Limiting Handle (Aspen/ Alpine) due to the supplier withdrawing it from the market for component updates or reconfiguration. | Component design/selection | Biomet sent an Urgent Medical Device Recall Notice dated March 12, 2015 via phone and customer letters. The letter identified the affected product, problem and actions to be taken. Customers were asked to follow the instructions on the enclosed "Response Form". For questions call 858-232-2672. |  | 148 | 148 | Worldwide Distribution - US Nationwide | Terminated | September 16, 2015 | 188 | 1022 | Not_Computer
Z-1463-2015 | 69996 | No | HIG | Hysteroscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Fluid Management System P4000; for irrigation and fluid warming in laparoscopic procedures, and distention, fluid warming, and volume/deficit measurements in endoscopic procedures within gynecology, urology, and orthopedic procedures. |  |  | September 01, 2014 | 2014 | Thermedx LLC | To correct software bugs that could affect the ability to accurately measure fluid deficit. | Software design | Thermedx sent an Urgent Product Correction letter dated March 13. 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm informed their customers that the software update will be performed at customer locations by Thermedx trained and authorized personnel. A Thermedx and/or your Sales Representative will contact customers to schedule a time to perform the field service update.Should you have any questions regarding this notification, please call a Thermedx technical service representative. Support is available toll free at 1-888-542-9276 from 8AM to 5PM Monday to Friday Eastern Standard Time or by e-mailing techservice@thermedx.com. Our website may be found at www.thermedx.com. We sincerely apologize for the inconvenience and thank you in advance for your cooperation. |  | 127 | 127 units | US Distribution to the states of : OH, MS, MA, WA, MI, NC, NY, WV, LA, TX and IL. | Terminated | September 14, 2015 | 378 | 127 | Not_Computer
Z-1467-2015 | 70369 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 6" Trifurcated Extension Set, Product Code: AMS-375-1 |  |  | January 15, 2015 | 2015 | Churchill Medical Systems, Inc. | Potential for leaking caused by insufficient bond between needleless device and female luer. | Process design | Vygon sent an Urgent Medical Device Recall letter dated January 15, 2015, to all affected customers via email or courier service. The letter instructs consignees to check stock for products included in the scope of the recall, count and document the affected inventory on the "Recall Acknowledgement and Inventory Return Form" and fax or email the form to Vygon. Vygon customer service will then contact the customer to issue a Return Merchandise Authorization Number so the inventory can be replaced at no cost. Distributors were provided with instructions on how to handle product that was further distributed by them. If after 30-days, product is not returned from each consignee, Vygon will follow-up with the consignee to obtain a status of the product return.Customers with questions were instructed to contact Vygon Customer Service at 1-800-473-5414, or by e-mail at customerservice@vygonus.com. |  | 5200 | 5,200 | Nationwide Distribution including TX, OH, NY, CA, UT, PA, MA, NY, VA, and IL. | Terminated | February 25, 2016 | 406 | 158850 | Not_Computer
Z-1480-2015 | 70710 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Captiva's lumbar intervertebral fusion system Reamers (color code rings on the Reamers). |  |  | February 25, 2015 | 2015 | Captiva Spine, Inc | It is possible for the titanium color-coded ring to detach from the Reamers. | Device Design | On 02/26/2015 the firm sent a Voluntary Correction Notification to its consignees via FedEx. Within that notification the firm asks to please check with your product representative and Central Supply Departments to determine if you have any of the products listed on the notificaiton. If any product is located please call 877-772-5571 from 8:30am (Eastern Time) to receive instructions as to how to remove the color-coded ring OR please follow the enclosed steps for the correction in their entirety. They contain information about returning the removed rings. Enclosed is Form 30.4 Product Tracking and Verification Form. You must complete and return this form even if you do not have any rings to remove and return. Please contact your Captiva Spine customer service representative if you have any questions regarding this correction, any of our products, or would like assistance with the correction. |  | 233 | 233 devices | GA, IN, FL, UT, AZ, CA, and NY | Terminated | June 01, 2015 | 96 | 233 | Not_Computer
Z-1479-2015 | 70488 | No | FTY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SpineFrontier MISquito Percutaneous Pedicle Screw System Surgical Technique; Part Number: DOC80023. The Rod Caliper is a manual surgical instrument that may be used as part of the PedFuse Pedicle Screw System and its MISquito components during percutaneous, minimally invasive surgical (MIS) technique of posterior pedicle screw fixation. |  |  | March 05, 2015 | 2015 | SpineFrontier, Inc. | The PedFuse Pedicle Screw Rod Caliper included with the MISquito Pedicle Screw System may be used incorrectly as the Surgical Technique does not adequately convey the intended use of the design. | Labeling design | SpineFrontier sent an Urgent Advisory Notice letter to customers via email. The letter identified the affected product, problem and actions to be taken. Customers were asked complete and return the Acknowledgement Form to confirm receipt of the recall notification. For questions call 978-232-3990. |  | 5 | 5 devices | US Nationwide Distribution in the states of TX, KS, CO, and MA | Terminated | June 15, 2015 | 102 | 5 | Not_Computer
Z-1481-2015 | 70948 | No | HGS | Intrauterine pressure monitor and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | HSG Trays/HS 5F Catheter Combo; HSG Procedure Tray+ Shapeable 5F HSG catheter. For hysterosonography and hysterosalpingography, gynecological use. |  |  | April 07, 2015 | 2015 | Catheter Research, Inc | HSG Trays and Catheters were mislabeled. Trays with the flexible HS catheter were labeled as containing the shapeable, stiffer HS catheter, and vice versa. | Packaging process control | Catheter Research, Inc. (CRI) has initiated a voluntary medical device removal of HSG Tray/ 5F HS Catheter Combo, and issued letters to customer son 04/07/2015 by e-mail and certified mail. Customers are asked to inspect their inventory and return recalled devices to CRI for replacement. If the devices were further distributed CRI requested customers provide notice of the recall. Customers may contact Babacar Diouf, VP of RA and QS, at (317) 872-0074 ext. 3512. |  | 880 | 880 units | Domestic: MN, CA, OH, IL, FL, MA, WI, VAForeign: NoneDOD/VA: None | Terminated | July 17, 2015 | 101 | 880 | Not_Computer
Z-1512-2015 | 70774 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Xpress System Pedicle Screw Assembly, 5.5 x 55mm. An internal fixation device for spinal surgery consisting of rods, pedicle screw assemblies (screw, cap, yoke) and cross bar connector assemblies. |  |  | March 11, 2015 | 2015 | X Spine Systems Inc | One lot of instruments have etched on the Yoke of the Pedicle Screw a laser marking size of 7.5 x 55mm instead of 5.5 x 55mm as described on the product label. | Nonconforming Material/Component | The firm, X-spine Systems, Inc., sent an "Urgent Medical Device Recall" letter dated March 19, 2015 to their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to notify affected hospital and surgeon of the recall immediately; complete the attached Medical Device Recall Return Response Form and return it with any affected product in inventory to X-Spine Systems immediately via mail: 725 Alexandersville Road, Miamisburg OH 45342-RMA #5288; and No action should be taken with product that has been implanted.Please contact customer service at 800.903.0640 or email orders@x-spine.com with any questions. |  | 14 | 14 units | US Distribution to states of: KS, NV &amp; TX. | Terminated | October 21, 2015 | 224 | 14 | Not_Computer
Z-1501-2015 | 70878 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | Sentec, V-Sign Sensor 2 (REF VS-A/P/N), Reusable / Nonsterile / Waterproof, SenTec AG, CH-4106, Therwil, Switzerland. The SenTec Digital Monitoring System (SDMS) is used for continuous and non-invasive monitoring of transcutaneous carbon dioxide tension (tcPCO2), oxygen saturation (SpO2), |  |  | March 25, 2015 | 2015 | SenTec AG | SenTec is voluntarily replacing V-Sign Sensors 2 due to pre-mature deterioration of a part of the PCO2 measuring unit and thus disabling PCO2-monitoring.This means that the affected sensors reach their end-of-life earlier than normally expected requiring the sensor`s replacement.There is no increased risk for the patient or operator, because the aforementioned deterioration results in the sens | Under Investigation by firm | Sentec sent an Important Information letter via email dated March 25, 2015. The letter described the problem and the product involved in the recall. Customers were instructed to locate the affected sensors at their end-customers and replace them with the new sensors. Replaced sensors are to be returned to SenTec. For questions contact SenTec at + 41 61 726 97 60.Consignee was sent an emai in regard to the recall being expanded. The email was sent on May 08, 2014. The email informed the consignee of the additional lot being affected and provided a list of the serial number of V-Sign Sensors to be replaced. |  | 188 | 129 US, 59 OUS | US Distribution to the state of : MO | Terminated | February 05, 2016 | 317 | 188 | Not_Computer
Z-1486-2015 | 70859 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | CROSSCHECK(R) Plating system, LAPIDUS CX Plate, 1mm Step, Left, REF CCP-LPX1L. |  |  | March 06, 2015 | 2015 | Wright Medical Technology, Inc. | The minor diameter, major diameter, thread pitch, and pitch on the Crosscheck are out of specification, which may result in intraoperative malfunction | Nonconforming Material/Component | Firm has initiated the recall on 03/06/2015 through written communication via Fed Ex P1 delivery to all Wright Medical Technologies distributors (hubs). Follow-up to non-responding consignees will commence 30 days after initiation. Effectiveness checks will be performed through the confirmation of returned product via a validated ERP system, a returned/signed verification form, and in the case of non-responders, a phone call, e-mail or follow up letter, depending on the region and availability of consignee. |  | 102 | 102 units | Nationwide Distribution. | Open, Classified |  |  | 3271 | Not_Computer
Z-1502-2015 | 70787 | No | ECN | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Herbst Universal Nut which is used in a Herbst Appliance |  |  | March 09, 2015 | 2015 | Specialty Appliance Works, Inc. | The Universal Nut, a component of the orthodontic Herbst Appliance was made with 416 SS and not the specified 303 SS. | Under Investigation by firm | Specialty Applicances sent an Urgent Voluntary Productionn Field Correction Notification dated March 9, 2015, to all affected customers who received the Herbst Appliances and for sales supply orders. The letter advised doctors to monitor the appliance during routine follow-up appointments for signs of corrosion, and for sales supply orders to return affected components to the lab for replacement. Customers were asked to complete and return the enclosed Field Correctivve Action Response Form and return via fax to 678-513-7345. To return affected devices call 800-522-4636 to receive instruction as to how to return it to Speciality Applicances. In the event that customers observe signs of corrosion or a fracture, they should contact customer service at 800-522-4636 or their sales representative. For questions customers should call 800-522-4636 For questions regarding this recall call 678-513-4408, ext 226. |  | 14428 | 14,428 Universal Nuts | Worldwide Distribution - US nationwide including Puerto Rico, Brazil, China, Colombia, Germany, Ireland, Mexico, Switzerland, and United Kingdom. | Terminated | January 12, 2017 | 675 | 14428 | Not_Computer
Z-1515-2015 | 70679 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | OptiMedica Catalys Precision Laser System; Model No(s). Catalys-U (US), and Catalys-I (Outside US). |  |  | February 20, 2015 | 2015 | Abbott Medical Optics, Inc. | Device may experience loss of suction during treatment resulting in scoring the cornea during lens fragmentation. | Device Design | Abbott Medical Optics (AMO) sent an Advisory Notice letter dated March 5, 2015 via FedEx overnight delivery to each Catalys Precision Laser System owner. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the attached Acknowledgement Form. For questions contact AMO Technical Support at 1-855-662-0911. |  | 202 | 202 | Worldwide Distribution - US Nationwide and the countries of, Canada, Germany, Switzerland, Australia, Spain, Japan, Great Britain, France, Netherlands, Austria, Singapore, South Korea, India, Thailand, Colombia, Ecuador, Italy, Brazil, Egypt, Tunsia, Belgium, Mexico, Greece, Guatemala, Chile, Turkey, Peru, Hong Kong, Saudi Arabia, Bangladesh, and Dominican Rep. | Open, Classified |  |  | 202 | Not_Computer
Z-1513-2015 | 70851 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CSA Medical truFreeze Console- Cryosurgical Unit Cryogenic Surgical Device |  |  | March 25, 2015 | 2015 | CSA Medical | TruFreeze Console caused a higher rate of liquid nitrogen (cryogen) to be delivered and may cause: stricture,scarring, bradycardia, or pneumothorax | Software design | CSA Medical issued letter dated 3/25/15 advising users of the problem. Users provided with:The mitigations available to the active venting procedures coupled with the extremely unlikely probability of injury, rare risk, thus allowing the physician to continue with active venting procedures.In regards a passive venting procedure may not have sufficient mitigation to allow the physician to identify the potential hazard with sufficient time to preclude a potential for injury. Therefore, passive users will be instructed not to use the system until a software improvement is put into place. Additionally, no catheters or consoles will be shipped to passive venting users of truFreeze.A response form to be signed and returned confirming receipt of the notification.Questions contact: Stephen Mascioli, MD 781-538-4755 smascioli@csamedical.com. |  | 82 | 82 units | Nationwide | Terminated | August 21, 2015 | 149 | 82 | Computer
Z-1456-2015 | 70988 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | OriGen BIOMEDICAL VV13F Reinforced Dual Lumen Catheter STERILE DISPOSABLE: DO NOT RE-USE Manufactured by: OriGen Biomedical. |  |  | March 30, 2015 | 2015 | OriGen Biomedical, Inc. | Potential for clear extension tube to separate from the hub, which could result in blood loss and may require intervention to prevent permanent impairment/damage. | Mixed-up of materials/components | The firm, OriGen, sent an "URGENT; MEDICAL DEVICE RECALL" letter dated 3/30/2015 and response form to all consignees via e-mail. The conisgnees were notified on 3/31/15 and 4/1/15. The recalling firm also posted a recall notification on the ELSO website, which is a public site mainly used by professionals within the ECMO community and press release dated 4/15/2015. The letter described the product, problem and actions to be taken. The consignees were instructed to discontinue use of all OriGen W13F Reinforced Dual Lumen ECMO Catheters from the lot; immediately return any unused product to OriGen Biomedical; and complete and return the attached acknowledgement and receipt form even if you do not currently have product in your inventory. The firm recommends that the consignee use an alternate product in the interim. If you have any questions regarding the return of product that you have on hand or the status of the investigation, please do not hesitate to conact the company at: OriGen Biomedical, Attn: Director, Quality and Regulatory Affairs, 7000 Burleson Rd, Bldg D, Austin, TX 78744 or call +1 512 474 7278 CST Monday-Friday 8:00-5:00. |  | 51 | 51 units | Worldwide Distribution: US (nationwide) and countries of: Spain, UK, Germany, Sweden, and Netherlands. | Terminated | May 03, 2016 | 400 | 51 | Not_Computer
Z-1514-2015 | 70702 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sony Medical Monitor LMD-1951MD 19 inch |  |  | March 05, 2015 | 2015 | Sony Latin America | Monitor has either not turned on or has lost image during clinical use. | Nonconforming Material/Component | SOLA sent an URGENT MEDICAL DEVICE FIELD ACTION dated March 2, 2015, to all affected customers. Within that notification SOLA requests that you send the customer letter included in Attachment 1 of this notification to all customers who have purchased the affected monitors from you. The letter instructs customers who already possess an AC adapter for each of their affected monitors to only use the AC adaptor when the monitors are in use. If you have units in inventory that can be provided with an AC adapter, we will provide under separate cover kits that contain: 1) a sticker warning and 2) an adhesive seal to go over the AC outlet and 3) correction instructions document. If you have units in inventory that cannot be accompanied by an AC adapter, for each of these affected monitors, do not ship the unit until the permanent corrective action described below has been applied. Please complete the MEDICAL DEVICE RECALL RETURN RESPONSE included in Attachment 2 to the contact person identified on the form. Please contact me below if you have any questions. 305-260-4000 Hours of operation: 8am to 8pm (Monday through Friday, excluding holidays). |  | 4764 | 4,764 units | Worldwide Distribution - US Distribution to the states of : FL and NC., and to the countries of : Ecuador, Costa Rica, Brazil, Nicaragua, Venezuela, Peru, Colombia, Dominican Republic, and South Korea. | Open, Classified |  |  | 4764 | Not_Computer
Z-1516-2015 | 70838 | No | MGQ |  |  | General &amp; Plastic Surgery | 510(k) | Iamin Moist Dressing and the Iamin Hydrating Gel: intended to cover a wound or burn on the skin, provide a moist wound environment and protect against abrasion, friction, desiccation and contamination. |  |  | March 26, 2015 | 2015 | PhotoMedex, Inc. | The labeling for the Iamin family of medical devices (e.g. printed and internet text) is not consistent with the intended uses and limitations identified at the time of the clearance to market these devices. | Labeling design | PhotoMedex sent an Urgent product Recall notification, dated March 2015, to direct consignees requesting they discontinue use, complete the response form, and return affected products. |  | 200188 | 200188 | Nationwide and Internationally, including Canada. | Terminated | August 18, 2015 | 145 | 200188 | Not_Computer
Z-1461-2015 | 70853 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TIGERPAW System II, |  |  | March 25, 2015 | 2015 | Laax, Inc. | Increase of field reports involving issues with the TIGERPAW System II resulting in possible tissue tearing on left atrial wall and bleeding during use of the device, which may or may not be caused by incomplete closure of the TigerPaw System II fastener. | Unknown/Undetermined by firm | The firm, LAAx, Inc.on behalf of Maquet Medical Systems, sent an "URGENT MEDICAL DEVICE RECALL (Removal) IMMEDIATE ACTION REQUIRED " letter dated March 30, 2015 to all affected sites by express mail. The letter describers the product, problem and actions to be taken. The customers were instructed to examine your inventory immediately to determine if you have any TIGERPAW System II products; remove the products, quarantine them and place in a secure location; complete and return the MEDICAL DEVICE RECALL (Removal) RESPONSE FORM by e-mailing a scanned copy to:Tigerpaw2015@maquet.com or FAX to: 1-(973)-396-3607. If you have acted as a distributor for the TIGERPAW System product II, immediately contact those accounts, advise them of the recall situationa and have them return all stock to you.Your Marquet Medical System USA Sales Representative will contact you to assist with the return of any product.If you have any additional questions, contact your local Maquet Medical Systems USA Sales Representative or our Customer Service at 1-888-880-2874 (6:00 AM - 5:00 PM PST). |  | 4154 | ~ 4,154 total units in US - both model numbers | Worldwide Distribution: US (nationwide) and countries of:Australia, Bahrain, Belgium, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Ireland, Israel, Italy Netherland, Norway, Poland, Qatar, Slovenia, Spain, Saudi Arabia, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | March 17, 2017 | 723 | 4154 | Not_Computer
Z-1517-2015 | 64225 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | The PCA 3 infusion pump allows clinicians, or patients, to (self) Administer analgesia safely and effectively within clinician programmed limits. |  |  | February 27, 2013 | 2013 | Hospira Inc. | The infusion pump has a clear plastic door that, due to its design, allows for excessive flexing. This in turn permits unauthorized access to the vial of medication. | Device Design | Hospira will contract with Stericycle Inc., in Indianapolis, IN to send letters and a reply form to Hospira customers. Receipt of the letters will be confirmed by the UPS tracking. Customers will be requested to return a reply form to Stericycle indicating the number of pumps on hand and that they have received the letter and notified the users.For questions regarding this recall call 800-441-4100. |  | 8214 | 8,214 units | Worldwide Distribution - US including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico, American Virgin Islands; and Internationally to Canada. | Open, Classified |  |  | 47096 | Not_Computer
Z-1519-2015 | 70863 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Light Sheer Desire Diode Laser System with XC Handpiece Accessory options. |  |  | March 25, 2015 | 2015 | Lumenis Limited | Device software treatment preset parameters for the XC treatment handpiece do not match the Operator Manual, and exceed recommended settings. Operator Manual parameters are lower than indicated for specific hair color and Fitzpatrick skin type resulting in insufficient treatment effect. May result in patient burns and hypopigmentation. | Software design | Lumenis Limited sent an Urgent Safety Advisory Notice dated March 25, 2015, to all affected customers. Customers were instructed that Lumenis is eliminating the potential risk of patient harm through a software upgrade and Operator Manual replacement. The software upgrade must be completed by a Lumenis-certified engineer as soon as possible. In the Interim, customers can continue to use their LightSheer DESIRE system by following the recommendations in the operator manual to always perform a test patch prior to a full treatment. Customers with questions were instructed to contact the Aesthetic Business Unit Product Manager.For questions regarding this recall call 801-656-2690. |  | 117 | 117 units | Worldwide distribution. US nationwide (to OR, NM, NY, IL, and CO), Switzerland, Canada, Italy, Great Britain, United Arab Emirates, France, Belgium, India, Israel, Japan, Portugal, Czech Republic, Kuwait, Germany, Spain, Georgia, Armenia, Slovenia, New Zealand, Saudi Arabia, and Russian Federation. | Terminated | June 15, 2016 | 448 | 117 | Not_Computer
Z-1520-2015 | 70888 | exerciser, non-measuring | ION | Nonmeasuring exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | AmTryke Model 1424 Community Cruiser Hand Cycle. Therapeutic tricycle. |  |  | March 27, 2015 | 2015 | Amtryke LLC | Product is being recalled due to complaints of cracked frames. | Under Investigation by firm | Consignees were notified via letter on 03/27/2015. |  | 53 | 53 units | US nationwide distribution to AL, CA, FL, HI, IL, KY, KS, LA, MI, NY, OK, OR, PA, SC, TX, VA, VT and WV. | Terminated | July 10, 2017 | 836 | 53 | Not_Computer
Z-1521-2015 | 70685 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Software Version 2.8 &amp; Below on VITROS 5,1 FS Chemistry Systems, Catalog Number 6801375, Global Trade Item Number 10758750001132, and VITROS 5,1 FS Chemistry System Refurbished, Catalog Number 6801890, Global Trade Item Number 10758750001644; IVD. |  |  | March 10, 2015 | 2015 | Ortho-Clinical Diagnostics | Software Anomaly; It is possible for the device to process samples with a cartridge other than the intended cartridge, potentially leading to erroneous patient results. If this anomaly occurred previously, an indication would be a series of results that were similar across multiple samples (i.e., results could be believable or outside of the Reportable Range for the intended assay). | Software in the Use Environment | On 3/10/2015, URGENT PRODUCT CORRECTION NOTIFICATION letters (Ref. CL2015-047, dated 3/10/2015) were sent to users notifying them of the issue and advising them of the required actions. Customers are to complete and return the Confirmation of Receipt no later than March 23, 2015. The firm is currently working on a resolution and will issue a follow up notification upon the software's availability. For questions, please contact Customer Technical Services at 1-800-421-3311. |  | 274.1 | VITROS 5.1 System: Domestic - 909; Foreign - 1250; 5,1 Refurbished: Domestic - 66 units, Foreign - 152 | Worldwide Distribution -- USA, Puerto Rico, Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 274.1 | Not_Computer
Z-1522-2015 | 70874 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan -Neg Breakpoint Combo Panel Type 34; |  |  | May 14, 2014 | 2014 | Beckman Coulter, Inc. | Siemens Healthcare Diagnostics has received complaints regarding panel degradation with a portion of the MicroScan Neg Breakpoint Combo 34 (NBP34) panel lot 2014-05-28. | Equipment maintenance | Siemens sent an Urgent Medical Device Recall letter dated May 14, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The risk to health is explained and users are asked to discontinue use and discard any of the recalled lot still in inventory. No other lot numbers of the NBPC 34 panels are affected and may still be used. The Field Correction Effectiveness Check form should have been completed and returned to Siemens.PLEASE FAX THIS COMPLETED FORM TO 302-631-8467. Contact Siemens Customer Care Center or local Siemens Tech support. For further questions please call (916) 374-3031 |  | 681 | 681 boxes | Worldwide Distribution - Nationwide Distribution including Puerto Rico and to the states of: KS, LA, MD, NY, WV,NM, NH, MO, CT, MA, OK, NC. IA, OH. WI, TX, PA, MS and MN and to the countries of : Canada, Kuwait, South Africa and Turkey. | Terminated | April 27, 2015 | 348 | 681 | Not_Computer
Z-1532-2015 | 70893 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Microbiology Systems: Dried Gram Negative MIC/Combo Panels |  |  | August 08, 2014 | 2014 | Beckman Coulter, Inc. | Technical Support Bulletin issued in February 2013 did not mention performance limitation for Proteus mirabilis and Imp when using the CLSIM100-S20 revised interpretive criteria. | Labeling False and Misleading | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated August 7, 2014 to affected customers by trackable method. The letter describe product, problem and actions to be taken. The customers were instructed to confirm receipt of the letter; complete the attached "Field Correction Effectiveness Check" form via fax to Siemens Healthcare Diagnostics at 302-631-8467;TSB 190 has been removed from use and archived from the Siemens Document Library, if you have a hardcopy, discard it appropriately;and follow the listed in instructions and retain and forward this letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center; Quality Systems &amp; Compliance Director at 916-374-3031 or email: juntalan@beckman.com or your local Siemens technical support representative. |  | 84249 | Approx ~(216240 boxes U.S. and 84249 boxes OUS) total | Worldwide Distribution: US (Nationwide) and to countries of: Argentina, Bahamas, Brazil, Canada, China, Ecuador, Egypt, Germany, Lithuania, Mexico, Paraguay, Peru Poland, Portugal, Romania, Russian Federation, Singapore, South Africa, Spain, Taiwan, Thailand, Trinidad and Tobago, United Arab Emirates, Uruguay, and Venezuela. | Terminated | April 27, 2015 | 262 | 84249 | Not_Computer
Z-1526-2015 | 70804 | No | LYJ |  |  | Neurology | PMA | VNS Therapy DemiPulse Generator Model 103 and VNS Therapy DemiPulse Duo Generator Model 104; CYBERONICS, INC. |  |  | March 06, 2015 | 2015 | Cyberonics, Inc | The pulse generators have a lower battery longevity than specified in their design requirement as a result of the devices being inadvertently left in a programmed ON state during manufacture. | Nonconforming Material/Component | On 03/6/2015, the recalling firm sent a recall notification letter to their consignees. The letter included an effectiveness card which is to be signed by the consignee and returned to the recalling firm. |  | 13 | 13 units | AR, TX, GA, CA, MI, IN, FL, AZ, NV | Terminated | October 09, 2015 | 217 | 13 | Not_Computer
Z-1525-2015 | 71021 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | WEBSTER HIS Catheter, 4 Pole Fixed Curve with Auto ID Technology, for mapping of cardiac structure; i.e., stimulation and recording only. |  |  | April 17, 2015 | 2015 | Biosense Webster, Inc. | The Webster HIS catheter, 4Pole Fixed Curve with Auto ID, is not being recognized by the CARTO 3 navigation System. | Software Manufacturing/Software Deployment | A recall letter dated 4/17/15 will be sent to customers by FedEx to inform them that Biosense Webster is recalling Webster HIS catheter, 4Pole Fixed Curve withAuto ID (Catalog No. D1085413). The letter informs the customers of the problems identified and the actions to be taken. Customers with questions related to the recall letter are instructed to contact their Biosense Webster salesrepresentative or call (866) 473-7823, Monday through Friday from 7:00AM to 8:00PM EST. Customers with questions related to the Voluntary Recall Certification Form and the process for returning the form, are instructed contact Elie Ghattas, Field Action Coordinator, at (909) 839-7281. |  | 503 | 503 units | US nationwide distribution. | Terminated | December 10, 2015 | 237 | 503 | Not_Computer
Z-1527-2015 | 70457 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | 9F Plastic Dual Port; The port systems are totally implantable vascular access devices. They consist of an injection port, a self-sealing membrane and a detachable catheter. The port systems provide repeated vascular access and are applicable for delivery of: medications, intravenous fluids, blood products, parenteral nutrition, and for venous blood sampling. |  |  | January 28, 2015 | 2015 | Medical Components, Inc dba MedComp | The dual port with catheters was not covered under FDA clearance. | PMA | Sales reps were notified by email on/about January 28, 2015, and instructed to notify their medical facilities and provide the recall letter. This will be accomplished by either email, telephone, or visit. Every consignee will be requested to complete and return Page 2 of the recall letter verifying receipt of the recall and disposition of the product they received. |  | 325 | 325 | Worldwide distribution. US nationwide, Canada, Brazil, and United Arab Emirates. | Terminated | January 06, 2017 | 709 | 325 | Not_Computer
Z-1523-2015 | 70993 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Anatomical Shoulder Handle for Rasp |  |  | March 18, 2015 | 2015 | Zimmer Gmbh | Complaints of difficulties to attach the Anatomical Shoulder Rasps (high resistance) or, once attached, due to difficulties to remove the handle (seizing up of the two components). This has the potential for delay in surgery. | Nonconforming Material/Component | On 3/18/2015, URGENT MEDICAL DEVICE RECALL Notification - LOT SPECIFIC letters were sent to the affected distributors with instructions for locating, quarantining, and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 30 | 30 units | Distributed to the states of IN, GA, WI, MO, PA, KS, TN, and NY. | Terminated | September 23, 2015 | 189 | 30 | Not_Computer
Z-1536-2015 | 70641 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | The Patient Control Module (PCM) is used in conjunction with a Baxter infusor as a single use device for the control of intermittent bolus doses of medication based on patient demand. |  |  | April 17, 2014 | 2014 | Baxter Healthcare Corp. | Potential for device malfunction resulting in flow when the device should not be flowing | Process control | Baxter Healthcare sent an Urgent Product Recall letter dated April 17, 2014, to the sole US consignee. The letter included instructions for the consignee to: 1) locate and remove all affected product; 2) contact Baxter's Center for Service to arrange for the return and credit; and 3) to complete and return the Customer Reply Form included with the recall notification. Questions about the recall can be answered by calling The Center for One Baxter at 800-422-9837. |  | 7596 | 7,596 patient control modules | Worldwide Distribution - US Distribution to the state of FL., and to the countries of : Belgium, Denmark, Germany, Italy, Japan, Norway, Poland, South Korea, Spain, Sweden and United Arab Emirates. | Terminated | July 21, 2017 | 1191 | 7596 | Not_Computer
Z-1538-2015 | 70960 | No | FAE | Penile rigidity implant. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Genesis Malleable Penile Prosthesis, 9.5 mm, Catalog /Item Number 91-9509SC/519250. |  |  | April 09, 2015 | 2015 | Coloplast Manufacturing US, LLC | Several lots of the 9.5 mm Genesis Malleable Penile Prosthesis were mislabeled. The Retail Box and Patient Labels indicate a size 9 mm, instead of 9.5 mm. The device size is printed directly on the individual prosthesis tray during the packaging process and states the correct size, 9.5 mm. | Error in labeling | Consignees were notified via Field Representatives starting on 4-9-2015 with a Coloplast "Urgent Correction Notification" letter dated April 7, 2015. The letter described the problem and the product involved in the recall. Requested consignees to complete the Return Form if returning inventory. For questions contact Customer Service at 800-258-3476. |  | 11 | 11 (6 USA, 5 OUS) | Worldwide - US Nationwide Distribution in the states of FL, NY, TX and the country of FRANCE | Terminated | September 11, 2015 | 155 | 11 | Not_Computer
Z-1405-2015 | 70924 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Automatic Mobile X-Ray (AMX) Series: Brivo XR285amx |  |  | December 19, 2014 | 2014 | GE Healthcare | An issue was identified with the mAs accuracy check performed on the system that could result in inaccuracies in the 10% to 20% range for exposure times longer than 15 msec. | Radiation Control for Health and Safety Act | GE Healthcare planned action:1. The customer notification letter will be distributed to the facilities that have installed the following AMX series X-ray systems: Models Brivo XR285 AMX, Optima XR200 AMX, and Optima XR220 AMX.2. As stated in the customer notification letter, the correction will be provided free of charge and subsequent communication between the firm and the facility will result in a site visit by a GE Healthcare field engineer to bring the X-ray system into compliance.3. The GEHC field engineer will follow Field Modification Instruction 10899 which provides instructions on how to perform a mAs accuracy check and how to determine if the test results pass predetermined acceptance criteria.If you have any questions or concerns regarding this notification, please call theInformation following phone number: 1-800-437-1171. |  | 6 | 6 units US | US Distribution to the states of: AZ, FL, GA, MO and WI. | Terminated | April 29, 2015 | 131 | 6 | Not_Computer
Z-1535-2015 | 70823 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | CD99 |  |  | March 17, 2015 | 2015 | Biocare Medical, LLC | A drop in staining intensity over time has been observed. These lots may not meet their labeled shelf life. | Unknown/Undetermined by firm | Biocare sent a Medical Device Notification letters dated March 17, 2015 to all affected sites. The letters identify the affected product, problem, actions to be taken and the fact that it affects all identified lots shipped between 02/13/13 and / 01/28/15. The letters instructed customers to immediately discontinue further use of the recalled lots, dispose of them according to state and local regulations and complete the response form and returned by April 7, 2015. Credit will be issued upon receipt of the response form. |  | 887296 | 44: 8 of lot 110912, 15 of lot 041913; 17 of lot 120310 and 4 of lot 110314 | Worldwide Distribution - US nationwide in the states of TX, KS, OK, NJ, LA, GA, NC, SC, VA, and the countries of Spain, Austria, Ecuador, Dominican Republic, Vietnam, Kuwait, Hungary and Israel. | Terminated | June 15, 2015 | 90 | 887296 | Not_Computer
Z-1524-2015 | 70904 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Pos Combo Panel Type 33 (PC33): |  |  | April 08, 2014 | 2014 | Beckman Coulter, Inc. | Siemens investigation confirmed that a portion one lot of panel Pos Combo 33 lot was packaged without desiccants.The substrates on a panel packaged without a desiccant may deteriorate over time due to exposure to additional moisture in the panel pouch. | Process control | Siemens sent an Urgent Medical Device Recall letter dated April 8, 2014, to all affected customers. Letters advised users of the recalled product with catalog number, Siemens Material Number and Lot number as well as the reason for recall. The risk to health was explained as well as the recommendation for discussing the content of the letter with the medical director regarding the need to review previous test results.Customers are asked to discontinue use and discard any remaining inventory of the PC33 panel 2014-09-12. All other lot numbers are not impacted.The attached confirmation receipt is to be completed and returned and the letter should remain with laboratory records. Questions should be directed to Siemens Customer Care Center. For further questions call (916) 374-3031. |  | 664 | 664 panel boxes | Nationwide Distribution including Puerto Rico and to the states of : CT, GA, NC, OK, NY, SC, KY, TN, CA, WA, OR, NJ, OH, PA, FL, MI, TX, WI, MO and NH. | Terminated | April 27, 2015 | 384 | 664 | Not_Computer
Z-1537-2015 | 70836 | No | MYN |  |  | Radiology | PMA | Logicon Caries Detector |  |  | April 06, 2015 | 2015 | GA Industries | GA Industries is recalling the Logicon Caries Detector, version 5.0, for Carestream Dental because it was not submitted for approval to the FDA prior to product release. | Unknown/Undetermined by firm | A Medical Device Correction and Removal Notice dated 3/25/15 will be sent to customers on 4/7/15 to inform them that GA Industries is providing them with a new approved version of Logicon Caries Detector, Version 5.1. The letter informs customers that the original version 5.0 should be uninstalled from their computers and the installation CD be destroyed. The letter informs the customers of the changes and improvements to the product. Customers are instructed to join training sessions conducted by Carestream that demonstrates the new features. Customers are instructed to contact LogiconSupport@Carestream.com for webinar schedule. Customers with questions are instructed to contact GA Industries at logiconsupport@carestream or at (800) 944-6365. |  | 825 | 825 units | US Nationwide Distribution | Terminated | November 23, 2015 | 231 | 825 | Not_Computer
Z-1542-2015 | 71074 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Command Modules, Model 91496-M, upgraded with the Masimo SpO2 Option. |  |  | April 17, 2015 | 2015 | Spacelabs Healthcare Inc | Ultraview SL Command Modules which were upgraded with the Masimo SpO2 Option, Model 91496-M was affected. Customer reported that during the time that the module is resetting (~10 seconds), monitoring of all parameters will be suspended. Upon completion of this sequence, alarms will be reset to their default values. | Software design | Spacelabs Healthcare sent an Urgent - Medical Device Correction letter, dated 17 April 2015, to the US consignees via priority mail. On 4/24/2015, the firm will send the Urgent - Medical Device Correction letter (translated as necessary) via email to international subsidiaries and distributor of record. The Identified the affected product, problem and actions to be taken. Ultraview SL Command Modules which were upgraded with the Masimo SpO2 Option, Model 91496-M are affected. Ultraview SL Command Modules originally shipped with Masimo SpO2 are NOT affected. Spacelabs Healthcare will contact you to schedule a convenient time for Spacelabs to correct all of your facility's affected Command Modules at no cost. For additional information or technical assistance, please contact:Technical SupportSpacelabs Healthcare, Inc.35301 SE Center St.Snoqualmie, WA 980651-800-522-7025 and select 2 for Technical Support |  | 1999 | 1999 units total (1932 in the US and 67 international) | Worldwide Distribution- US Nationwide in the states of AL, DE, GA, ID, IL, KY, LA, MA, ME, MI, MN, MO, NC, NJ, NY, OH, OR, PA, SC, TN, VA, WA and the country of France. | Terminated | November 13, 2015 | 210 | 1999 | Not_Computer
Z-1541-2015 | 70737 | [6 empty] | Portable FAR Infrared Sauna |  |  | December 23, 2014 | 2014 | Therasage, LLC. | Product distributed without 510(k) approval | Labeling False and Misleading | On January 20, 2015 Therasage sent a notification to its consignees asking them to discard any of the labeling materials that state the following: 1) Use of the sauna allows one to "naturally detoxify." 2) Use of the sauna expels toxins from the body 3) Any reference to an exact amount of calories one uses in an hour when enjoying the benefits of this product 4) Although this product is enhanced with a negative ion environment, we cannot say that the product is enhanced with negative ion therapy 5) Any reference to the maximum infrared absorption available cannot be explicitly stated, as it may vary for each individual person. Customers with questions were instructed to call 888-416-4441 or email info@therasage.com.For questions regarding this recall call 561-416-9300. |  | 8 | 8 devices | Nationwide Distribution including SD, IL, TX, OR, NY, WI, GA, and CA. | Terminated | August 06, 2015 | 226 | 8 | Not_Computer
Z-1548-2015 | 70898 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Test Request Distributor (TRD 1.3 and TRD 1.4), a software module used on the Dako Omnis System, an automated slide stainer for in vitro diagnostic use. |  |  | April 01, 2015 | 2015 | Dako North America Inc. | If a user requests slides from the LIS or TPID, then updates a request by changing the test, the TRD software will reject the update. The LIS and printed slide label will display the updated test with the requested change, but the Omnis system will execute the initial test, not reflecting the change. The system will not warn the user that the test request was rejected. | Other | Dako sent a Voluntary Product Recall Notification letter starting the week of April 1, 2015 along with a acknowledgement attachment. The letter identified the affected product, problem and actions to be taken. The firm requests that customers inform those in their organization of the notification including organizations where the potentially affected instruments have been transferred. The firm requests that customers complete, sign, and return the acknowledgement form within one week after receiving the notification. For questions contact your Dako sales representative.On 04/29/15 the firm sent out an updated customer notification letter in order to provide additional clarification and instructions to customers in regards to the recall situation. |  | 33 | 33 (7 in U.S. and 26 Internationally) | Worldwide Distribution - U.S Nationwide in the states of IN and PA and the countries of: Spain, Belgium, France, Netherlands, Denmark, Finland, Sweden, Germany, UK, and Korea. | Terminated | August 05, 2015 | 126 | 33 | Not_Computer
Z-1544-2015 | 70873 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | Integra Long Epidural Steroid Injection Needle (Tuohy Needle), 18 Gauge x 5 inch Removable Wing, Metal Stylet, Calibrations (10BX), Catalogue No. PITRW1850. |  |  | March 30, 2015 | 2015 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | A single lot of 18 Gauge, 3.5 inch Tuohy Needles were incorrectly packaged and labeled as 18 Gauge, 5 inch needles. | Employee error | US Consignees were notified by an Urgent: Voluntary Medical Device Recall letter on 3/30/15 delivered by traceable courier service. The letter identified the affected product and explained the reason for the recall. Customers were asked to review their inventory, identify affected products, and immediately stop use and remove them from service. Customers are to complete the attached form and return via mail or fax. Upon receipt of the form, Customer Service will provide an RMA number and directions for returning the affected products to the firm. Customers should contact Customer Service at 1-800-241-2210 with any questions regarding the instructions. |  | 900 | 90 Tuohy needles (9 cases of 10 each) | Nationwide Distribution. | Terminated | June 24, 2015 | 86 | 900 | Not_Computer
Z-1543-2015 | 71015 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Remote Network Station, Catalog RNS-9703, Model No. RNS-9703-19, RNS-9703-24, Software version 02.40. |  |  | April 14, 2015 | 2015 | Nihon Kohden America Inc | Nihon Kohden America (NKA) is recalling the Remote Network Station (RNS) 9703 because it may fail to sound. | Under Investigation by firm | The firm, Nihon Kohden, sent an Urgent Medical Device Alert" letter dated 4/14 /15 to customers. The letter describes the product, problem and actions to be taken. The customers were instructed to stop using the RNS for primary alarm notification and instead utilize the CNS and or the Bedside monitor for primary alarm notification, until such time NKA can provide corrective measures to your facility; Inform potential users in your organization of this notification, and complete the attached reply form and return it by fax (949) 580-1550 or email it to RNS_Replay@nkusa.com. Customers with questions are instructed to contact NKA Technical Support Center at (800) 325-0283 Ext 6 (available 24 hours a day/7 days/week). |  | 208 | 208 units | US Distribution to states of: AL, CO, DE, FL, GA, HI, KY, IL, IN, MA, MD, MI, MN, MO, NC,, ND, NJ, NY.OH, PA, and WV. | Terminated | October 01, 2015 | 170 | 208 | Not_Computer
Z-1545-2015 | 70906 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Insorb Subcuticular Skin Stapler. INSORB staples are made from an absorbable copolymer which is a synthetic polyester derived from lactic and glycolic acids. It is chemically similar to other surgical glycolide/lactide-based copolymers. Polyglycolic/polylactic acid copolymers degrade in vivo by hydrolysis to glycolic acid and lactic acid which are then absorbed and metabolized by the body. |  |  | March 19, 2015 | 2015 | Incisive Surgical Inc | A previous Turkish distributor over-labeled a portion of the lot with an expiration date of 2016-05. The expiration date at the time of manufacturing was 2014-05. | Labeling False and Misleading | Firm is working with EU Rep on English and Turkish notification letters. |  | 1260 | 210 cartons (containing 6 INSORB staplers per carton) | International Distribution Only - Turkey. | Terminated | July 08, 2015 | 111 | 1260 | Not_Computer
Z-1550-2015 | 70574 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008 Series Hemodialysis Machines: 2008T with the following product code and description: 190395 2008K at Home System - 1st Gen. -Âż Canada; |  |  | February 20, 2015 | 2015 | Fresenius Medical Care Holdings, Inc. | The Acetate value for GranuFlo on the Select Concentrate screen should be 8mEq/l but it is displayed as 4mEq/l. The correct value is shown on the dialysate screen. | Software design | The firm, Fresenius Medical Care, sent an "URGENT - MEDICAL DEVICE CORRECTION" letter dated February 19, 2015 to its Consignees via certified mail on February 26, 2015. The letter describes the product, problem and actions to be taken. The consignees were instructed to sign and return the attached "Reply Form" via fax to: 781-699-9769 Attn: Quality Department or Scan and E-mail to: NOTIFYRA@fmc-na.com to confirm that they received and understand the customer notification. The firm will be providing a software update and stated that in the near future, your local clinic will be contacting you about installing the software update.If you have any questions pertaining to this Customer Notification, please contact your local clinic, 800-227-2572 or the Corrections and Removals Manager at 800-662-1237. |  | 133751 | 133,751 in total (US: 130,337 and OUS: Canada 1,201 and Mexico 2,213) | Worldwide Distribution: US (Nationwide) and countries of: Canada and Mexico. | Open, Classified |  |  | 133751 | Not_Computer
Z-1547-2015 | 70958 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | 4.0mm Variable Angle, Self-Drilling Hexalobe Screw, 14mm |  |  | April 10, 2015 | 2015 | Alphatec Spine, Inc. | During a field audit it was discovered by an employee that a 16mm Trestle Luxe Screw has been both laser marked and labeled as a 14mm Trestle Luxe Screw (4.0mm Variable Angle, Self-Drilling Hexalobe Screw, 14mm). As a result the firm recalling the affected lots codes for this mislabeling. | Employee error | Alphatec Spine sent an Urgent-Medical Device Recall Notification dated April 7, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm first requested that customers notify their consignees of the recall and the risk, and risk mitigation factors.The firm requests that customers take the following field actions:-Upon receipt of the Field Safety Notice to review inventory to determine if affected screw --part number 71240-14, lot 685904 is in their possession.-If the product is in their possession to please abstain from use and contact Alphatec Spine Customer Service (customerservice@alphatecspine.com) immediately for instructions on how to return the affected product. All shipping instructions will be provided, including replace product and/or credit.-Fill out page 3 of the letter to confirm that they have read the notification and have take all necessary actions stated in the notification.-Return a signed copy using one of the methods below:Mail to: Alphatec Spine, Inc. 5818 El Camino Real Carlsbad, CA 92008 Fax to: 1-800 431-7729 For further questions call ( 760) 431-9286. |  | 61 | (244 units in US and 61 units OUS) | Worldwide Distribution - US Distribution to the states of: KS, LA, TN, NJ, KY, CA, FL, OR, OK, AZ, WI, CO, IN, NH, AR, UT, and AL., and to the countries of : Spain, Japan, Belgium, and Saudi Arabia. | Terminated | July 10, 2015 | 91 | 61 | Not_Computer
Z-1546-2015 | 71061 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Female Driver, Part number FDr101, instrument used for the insertion of the Fassier-Duval IM Telescopic System. For orthopedic procedures. |  |  | April 10, 2015 | 2015 | Pega Medical Inc. | The hexagonal tip of the Female Driver instrument for the Fassier-Duval IM Telescopic System is 0.06 mm under the specifications. This may result in a reduction of the total strength of the instrument, possibly leading to deformation, wear, and tip breakage. | Nonconforming Material/Component | The firm sent the Urgent Field Safety Notice-Recall letter, dated April 10, 2015 to consignees via email. The letter listed the affected product as FEMALE DRIVER FASSIER-DUVAL IM TELESCOPIC SYSTEM.Consignees are requested to return devices back to the firm for disposal. Consignees with questions should contact the firm at 450-688-5144 ext. 242 or send email to egarcia@pegamedical.com |  | 9 | 9 units (5 units in the US and 4 units international) | Worldwide distribution. US in Delaware, Florida, Nebraska, and New York; Australia, Brazil, and United Kingdom. | Terminated | June 29, 2015 | 80 | 9 | Not_Computer
Z-1557-2015 | 70753 | No | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | Introduction Pack. A device inserted into an endotracheal tube to aid in the intubation procedure. |  |  | March 13, 2015 | 2015 | Vital Signs Colorado Inc. | Vital Signs Colorado (dba CareFusion) is recalling Vital Signs LightWand Stylet catalog numbers 3910 and 3960 due to a potential patient safety risk involving a breach in the sterile packaging that may not be recognized prior to use. | Packaging | Customers were notified on 3/13/15 via letter to perform an inventory, complete and return the Recall Response Form within 15 days, and to destroy the product on site (do not return to CareFusion). |  | 399 | 399 | Worldwide Distribution. US Nationwide, Australia, Canada, Chile, Colombia, Dominican Republic, Ecuador, Egypt, Germany, Guatemala, Hong Kong, India, Israel, Italy, Japan, South Korea, Lebanon, Mexico, Netherlands, New Zealand, Panama, Peru, Poland, Russia, Slovenia, Turkey, United Kingdom, and Qatar. | Terminated | October 07, 2015 | 208 | 83219 | Not_Computer
Z-1560-2015 | 70834 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | PowerLoc MAX Safety Winged Infusion Set, Product Code 0132215 |  |  | March 13, 2015 | 2015 | Bard Access Systems | Bard Access Systems is conducting a field action due to the potential that an incorrect Maximum Flow rate label associated to the power injection IFU for the PowerLoc MAX Safety Infusion Set (SWIS) was attached to the outside of the unit pouch. | Labeling False and Misleading | The firm, Bard, sent an "URGENT: Notification-Potential Incorrect Infusion IFU Label for PowerLocÂż MAX Safety Winged Infusion Set (SWIS)" letter dated March 25, 2015 to its customers. The letter described the product, problem and actions taken. The firm has corrected the cause of the issue and all future lots will be manufactured with the correct maximum flow rate of 2mL/second. Customers may return for a replacement or keep items.Please direct any return requests and questions to our coordinator at 801-522-5640. |  | 700 | 700 units | US Distribution to states of: GA and MN. | Terminated | June 24, 2015 | 103 | 700 | Not_Computer
Z-1561-2015 | 70938 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO C and AMSCO 400 Steam Sterilizers, STERIS Corporation. |  |  | February 19, 2015 | 2015 | Steris Corporation | STERIS has identified that the control board software in select AMSCO 400 and AMSCO C units will interrupt and cancel a processing cycle should the selected sterilizer shutdown time coincide with a processing cycle. This may result in a procedure delay. | Software design (manufacturing process) | On 3/9/2015, the firm sent Urgent Field Correction Notices, to their customers. The letter identified the affected product and gave a description of the problem. The letter states that the control software is being revised to ensure the utility shutdown feature will not interrupt an ongoing processing cycle. STERIS Field Representatives will contact customers to schedule installation of the new control board on affected units. Customers may direct questions regarding the issue to STERIS Customer Service at 1-800-548-4873; Marie LaFrance, Product Manager, at 440-392-7648, or their local sales representative. |  | 701 | 701 units | Worldwide Distribution -- US, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, I, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY; and, the countries of Afghanistan, Canada, Japan, Kuwait, Mexico, Republic of Korea, &amp; United Arab Emirates. | Terminated | January 20, 2016 | 335 | 701 | Not_Computer
Z-1558-2015 | 70895 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Teleflex Medical Hudson RCI ConchaTherm Neptune Heated Humidifier, Rx Only. |  |  | April 01, 2015 | 2015 | Teleflex Medical | Customer complaints were received of power to the unit ceasing and the display going blank. | Nonconforming Material/Component | Teleflex sent an Urgent - Medical Device Notification letter dated April 1, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They were informed that Teleflex Medical would exchange all of the potentially impacted units in the field with the NeptuneÂż Heated Humidifier with ConchaSmart" Technology. In the meantime, they were instructed to continue to follow hospital and departmental protocol regarding the regular monitoring of the patient and the product during treatment. If the unit display is blank, replace it with another unit and call Teleflex for an immediate replacement. A Recall Acknowledgement Form was attached to be completed, scanned and returned to Teleflex Medical to document receipt of this Notification.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-396-2111. |  | 7228 | 7,228 units | US Nationwide and Internationally to the countries of: Belgium and Japan | Open, Classified |  |  | 7228 | Not_Computer
Z-1553-2015 | 70919 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Small Battery Drive (SBD) Oscillating Saw Attachment Part Number: 532.021 for orthopedic surgery. |  |  | March 24, 2015 | 2015 | The Anspach Effort, Inc. | The potential exists for the Oscillating Saw Attachment to disengage from the Small Battery Drive handpiece. If the attachment disengages from the handpiece intra-operatively, there is a potential for injury to the patient and/or user. | Device Design | DePuySynthes sent a "URGENT NOTICE: MEDICAL DEVICE RECALL- CORRECTION" letter dated March 24, 2015 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The firm requests that customers review the inventory and perform the in-process inspection outlined in the letter attachment, a review and repair will be implemented if necessary during the next visit schedule maintenance interval at the customer site.Customers having questions regarding this notice should contact the firm Complaint Handling Unit Manager at 561-494-3673 or to contact their Depuy Synthes Sales Consultant.For questions regarding this recall call 561-627-1080. |  | 12085 | 12,085 pieces | Worldwide Distribution - US (nationwide including DC, Hawaii, and Puerto Rico) and Internationally to Argentina, Austria, Australia, Belgium, Bahrain, Bolivia, Brazil, Canada, China, Chile, Costa Rica, Czech Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, France, Great Britain, Greece, Guatemala, Hong Kong, Croatia, Hungary, Indonesia, Ireland, Israel, India, Iraq, Iran, Italy, Jordan, Japan, Kenya, Korea, Kuwait, Lebanon, Sri Lanka, Libyan Arab Jamahiriya, Morocco, Macedonia, Mexico, Malaysia, Netherlands, Norway, Nepal, New Zealand, Oman, Ontario, Peru, Philippines, Palestine, Portugal, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, South Africa, El Salvador, Thailand, Turkey, Taiwan, Ukraine, Uruguay, Venezuela, Vietnam and United Arab Emirates. | Open, Classified |  |  | 12085 | Not_Computer
Z-1559-2015 | 70899 | No | MDZ | Medical washer-disinfector. | General Hospital | General Hospital | 510(k) | Caviwave Pro Ultrasonic Cleaning System, |  |  | February 19, 2015 | 2015 | Steris Corporation | Should the Caviwave Pro Ultrasonic Cleaning System experience a temporary interruption in electrical power during a processing cycle, the circulation pump may stop functioning. In addition, it is possible that the ultrasonic generator may stop working. The machine does not provide an alert notifying the user of this device malfunction. | Software design | STERIS sent an Urgent Field Correction Notice letter dated March 9, 2015 customers. The letter identified the description of the product, description of the problem and actions to be taken. For questions contact STERIS Customer Service at 1-800-548-4873 or local sales representative. |  | 491 | 491 units | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH,NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI &amp; WV and the countries of: Australia, Brazil, Canada, France, Mexico, Saudi Arabia, Spain, Taiwan &amp; United Arab Emirates. | Terminated | October 21, 2015 | 244 | 491 | Not_Computer
Z-1574-2015 | 70638 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System (Software Version 3.5.0 and 3.5.1) |  |  | April 17, 2015 | 2015 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | The operator console allows users to deliver therapeutic radiation to patients even though a specific communication error relevant to imaging conditions has occurred. In addition, the treatment record (the delivered radiation record) cannot be saved. | Software design (manufacturing process) | MITSUBISHI HEAVY INDUSTRIES, LTD. has issued a Safety Alert for the MHI-TM2000 Linear Accelerator System dated April 10, 2015, due to a software issue.Customers were advised not to deliver therapeutic X-ray if popup warning message show is displayed during treatment. Instructions were provided if therapeutic radiation is delivered during popup warning. Customers with questions should contact the Customer Service Representative by calling 1-800-597-5911, Monday through Friday or email: us.support@brainlab.com. **CAPA** Correction software for this issue has a tentative availability date of June 2015. Estimated software update completion date would be end of October 2015. |  | 11 | 11 units total | Worldwide Distribution - US including FL, NY, OH, &amp; TX, **Internationally - Belgium, Italy, Korea, France &amp; Germany. | Terminated | April 11, 2017 | 725 | 11 | Not_Computer
Z-1562-2015 | 70841 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 3/16 in. Trocar with Single Drain |  |  | March 26, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the CBCII Wound Drains and Ambulating Bulb for a potential sterility breach in the packaging | Package design/selection | An URGENT MEDICAL DEVICE NOTIFICATION letter was sent to all customers on 03/31/2015. The letter identified the affected product, the reason for the recall, the risk to health, and actions to be taken by the customer or user. All are asked to sign the Business Reply Form stating they have read the Notification and followed all instructions within. Customers were asked to immediately check all inventory and quarantine any affected product, return affected product to Stericycle for destruction using a return label provided by Stericycle. Upon receipt of affected product a credit will be issued to customer account. |  | 2711 | 2,711 boxes (27,110 each) | Worldwide Distribution. US nationwide, Australia, Great Britain, The Netherlands, Spain, Romania, Greece, Switzerland, South Africa, and Japan. | Terminated | December 31, 2015 | 280 | 42806 | Not_Computer
Z-1395-2015 | 70701 | No | IWE | Light beam patient position indicator. | Radiology | Radiology | 510(K) Exempt | GE Discovery MR950 MRI system |  |  | October 24, 2014 | 2014 | GE Healthcare | The alignment lasers are missing the labels required by radiation safety regulations. | Error in labeling | GE Planned action(s) to repair defect or to bring product into compliance:The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 5 units consists of the following:1. New labels will be mailed to all purchasers, with instructions to apply the label in the appropriate location on the unit, or a GE field engineer will be sent to each user site to complete the correction.2. All mailings and service calls will be made at no cost to the purchaser.3. The corrections will be completed by August 1, 2015.If you have any questions, please contact your local Service Representative or call 1 (800) 437-1171. |  | 5 | 5 units (2 in US) | US Distribution to the state of CA., and Internationally to Italy and Japan. | Terminated | May 06, 2015 | 194 | 5 | Not_Computer
Z-1575-2015 | 71009 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen Augment; 5mm Distal Femoral Augment;Part number 00-5990-036-10. |  |  | April 02, 2015 | 2015 | Zimmer, Inc. | The affected lot is being recalled because it was packaged without the locking screw. The locking screw packaged with the augment is used to fix the augment to the femoral component. This has the potential to delay surgery | Mixed-up of materials/components | On 4/2/2015, URGENT MEDICAL DEVICE RECALL LOT SPECIFIC notifications were sent to the affected distributors with instructions for locating, quarantining, and returning the affected product. All distributors were notified via electronic mail, and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 20 | 20 units | Distributed in the states of IN, MN, FL, and IL, and the countries of Argentina and Germany. | Terminated | September 23, 2015 | 174 | 20 | Not_Computer
Z-1370-2015 | 70678 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACSCalibur; |  |  | March 06, 2015 | 2015 | BD Biosciences, Systems &amp; Reagents | The circuit breakers on some BD FACSCalibur instruments may have the incorrect rated amperage (20 amps instead of 2 amps) causing the circuit breaker to fail to open in an over-current situation. | Mixed-up of materials/components | BD Biosciences sent an Important Product Recall letter on March 6, 2015 by certified mail and Field Change Notification (FCN FCB15-01-660312 directing Field Service Engineers to correct the circuit breaker is in process. Customers with questions were instruct to contact BD Customer Support at 1-877-232-8995 (prompt 2) in the United States. For customers outside the US were instructed to contact their local BD Biosciences representative or distributor. For questions regarding this recall call 408-954-6080. |  | 72 | 72 devices | Worldwide Distribution -US including TX, NJ, and OK and Internationally to Hong Kong, Japan, Singopore, and Belguim | Terminated | April 04, 2016 | 395 | 72 | Not_Computer
Z-1576-2015 | 70870 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | CONTIPLEX Continuous Peripheral Nerve Block Tray, CONTIPLEX Stirn |  |  | March 27, 2015 | 2015 | B. Braun Medical, Inc. | There is a potential for a hole to be present in the outer tray of a limited number of pain control trays. This issue may cause the integrity of the sterile barrier to be compromised. A compromised sterile barrier poses a risk of inflammatory response and/or possible contamination of the spinal fluid or epidural space. | Nonconforming Material/Component | BBMI provided a recall notification letter, dated March 27, 2015, to distributors and end users in receipt of suspect product, along with acknowledgement forms detailing the suspect product that was distributed to each account. All customers were provided with instructions for disposition of suspect product in their inventory. Distributors were reminded of their responsibility in executing the recall. Returned product will be sent to the BBMI Distribution Center in Breiningsville, PA. |  | 5000 | 5000 | Nationwide Distribution. | Open, Classified |  |  | 5000 | Not_Computer
Z-1577-2015 | 71118 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Command Modules, Model 91496, with the Masimo SpO2 PCBA, PN: 010-1636-02. |  |  | April 21, 2015 | 2015 | Spacelabs Healthcare Inc | Modules resetting. This failure mode happens when the Spacelabs Command Module is configured with the Masimo SpO2 option (-M), the Spacelabs Respiration option (-R) and the Masimo SpO2 PCBA, PN: 010-1136-02. With these two options and the Masimo SpO2 PCBA present, the module may experience random resets. | Software design | Spacelabs expanded the recall to include additional 29 units and notified one additional new consignee on June 4, 2015.*******************************************************************************************The firm, Spacelabs Healthcare, sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated 27 April 2015, via priority mail to US consignees and email to foreign consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to circulate this notification to all persons concerned by the information it contains; add a copy to the appropriated system user manuals until such time as corrections have been implemented by Spacelabs, and weigh the benefits versus the risks when deciding whether or not to continue to use your modules until they can be upgraded.Only modules shipped or repaired after March 8, 2014 are affected by this recall.Affected command modules will be upgraded with a new version of software by Spacelabs at no cost.Customer with questions about this corrective action, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 1973 | 1973 units total (1345 modules in the US and 604modules international and 3 PCBAs in US and 21 PCBAs international) | Worldwide Distribution:US (nationwide) including states of: AL, AR, AZ, CA, FL, GA, IA, ID, IL, LA, ME, MI, MN, MO, NE, NJ, NY, OH, OR, PA, SC, TN, TX, VA, WA, WY, and Puerto Rico and countries of: Argentina, Australia, Bangladesh, Canada, Chile, France, Great Britain, India, Kuwait, Panama, Poland, Qatar, Saudi Arabia, and Turkey.. | Terminated | November 16, 2015 | 209 | 1973 | Not_Computer
Z-1583-2015 | 71029 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM ARISTOS FX; Multipurpose Radiography System (MPRS) is a dedicated x-ray system with a flat panel detector which allows the acquisition of x-ray exposures without the use of conventional film/screen systems. The MPRS allows radiographic exposures of the whole body including skull, spinal column, chest, abdomen, extremities, and excluding mammography. Radiographic exposures may be taken with the patient in the sitting, standing, or supine positions. The intended use and indications for use of this modified device as described in its labeling have not changed from its predicate device. |  |  | April 14, 2015 | 2015 | Siemens Medical Solutions USA, Inc | If the cable cart and the corrugated hose are unable to move freely, the cable holder could drop down and hit patients or system operators. | Device Design | The firm, Siemens Healthcare, sent a "Safety Advisory Notice" dated April 14, 2015, to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to follow the instructions in the notice; promptly notify and instruct all the staff at your organization making them aware of the problem; and forward this safety information to other organizations that could be affected. If the device has been sold and therefore no longer in your possession, forward the safety notice to the new owner. Additionally, customers were instructed observe this safety notice and comply with the corresponding measures until the update has been fully completed. If you have any questions, contact customer support at 610-219-6300. |  | 43 | 43 | Nationwide Distribution. | Open, Classified |  |  | 43 | Not_Computer
Z-1460-2015 | 70732 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | EOS, Digital radiography system used in general radiographic examinations. |  |  | February 17, 2015 | 2015 | Eos Imaging Inc | When performing calibration, an alert message on the spectral filtration of the X-ray beam may be suppressed. Improper filtration of the X-ray Beam can then occur in exams set up with copper filtration. | Radiation Control for Health and Safety Act | Planned Action of EOS Imaging Inc.:1. Install a new acquisition software version, which will prevent use of the system when improper filter positioning is detected.2. A customer notification letter which includes a statement that EOS Imaging will without charge, remedy the defect or bring the product into compliance.CDRH approves the CAP subject to the followingCDRH approves the CAP subject to the following conditions:1. Provide documentation to show that all filter positioning subassemblies have been replaced, which you state was implemented to prevent improper filter positioning.2. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.3. EOS Imaging will implement this CAP by September 1, 2015.For further questions call (678) 564-5400. |  | 13 | 13 units installed in US | US Distribution to the states of: CA, DE, PA, MN, FL, MO, OH, IN and IL. | Terminated | May 31, 2017 | 834 | 13 | Not_Computer
Z-1581-2015 | 71116 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 volume imaging ultrasound systems at software versions VA16C, VA16D and VA16E. Model 10433816. |  |  | May 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | In some cases, the system is unable to capture a clip or image during a routine scan. | Other | Siemens sent a Customer Advisory Notice letter to affected customers. The letter identified the affected product. problem and how this issue is being resolved. To resolve this issue on all affected systems, a mandatory software update has been developed and is being installed free of charge. |  | 186 | 186 devices | Worldwide Distribution - US Nationwide and the countries of:Angola, Armenia, Australia, Austria, Belgium, Bosnia Herzeg., Brazil, Brunei, Bulgaria, Colombia, Czech Republic, Finland, France, Germany, India, Japan, Kazakhstan, Kenya, Kuwait, Mauritius, Mexico, Norway, Oman, P.R. China, Poland, Rep. of Yemen, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E., United Kingdom. | Terminated | May 06, 2015 | 1090 | 186 | Not_Computer
Z-1578-2015 | 71084 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ 5 Ultrasound System with Pediatric Cardiology option, Model: EPIQ 5G, EPIC 5C, EPIQ 5W. Catalog Number: 795204 / 795205. Part Number: 989605408541 With |  |  | April 10, 2015 | 2015 | Philips Ultrasound, Inc. | When Epiq 5 Ultrasound System, WITH the Pediatric Cardiology option, is set up to Metric and weight and/or height is entered, a unit conversion error may result in the incorrect calculation of the patients Body Surface Area (BSA). | Software design | Philips sent the Urgent-Medical Device Correction-Philips Model EPIQ5 and Ultrasound system letter, dated 2015 APRIL 10, and the Urgent-Medical Device Correction-Philips Model EPIQ7 and Ultrasound system letter, dated 2015 APRIL 10, letters to their consignees.These systems WITH Pediatric Cardiology option are affected.Philips issued two Field Change Orders (FCO) to their Field Service Engineer to install upgraded software: -FCO 79500336 for EPIQ5-FCO 79500337 for EPIQ7Customers with questions can call Philips Customer Service at 1-800-722-9377. |  | 751 | 751 units total (246 units in the US and 505 units outside the US) | Worldwide Distribution - US (nationwide and DC) and Internationally to Norway, India, New Zealand, Bahrain, Sweden, Italy, Australia, France, Netherlands, Switzerland, Germany, Austria, and Turkey. | Terminated | February 29, 2016 | 325 | 751 | Not_Computer
Z-1580-2015 | 71050 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ViewRay System, Radiation Therapy System |  |  | January 15, 2015 | 2015 | Viewray Incorporated | The software was not correctly using the RT (Radiation Therapy) to MR (Magnetic Resonance image) coordinate correction for non HFS (Head First Supine) patient orientations, resulting in slice mismatch error. | Software Design Change | On January 15, 2014, ViewRay personnel discussed the investigation of the issue (slice mismatch error) with the customer. The firm advised the customer that until the issue was corrected, to not treat patients in non HFS orientations, or to use the MR acquired from the ViewRay system as the primary plan series for non HFS scans. ViewRay installed software version 3.4.5.28 on January 20 and 21, 2014, in order to prevent this slice mismatch error from occurring in any patient position. |  | 1 | 1 | US distribution to MO. | Terminated | September 23, 2015 | 251 | 1 | Not_Computer
Z-1585-2015 | 70854 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Aspire HD, Full Field Digital, System, x-Ray, Mammographic; |  |  | February 03, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | User Manual lacks a description for handlingsome errors even though the system can be recovered from such errors | Labeling Change Control | Fuji Film on 2/3/15, issued a User Notification letter and the recovery procedure were hand carried to the user. A copy of Notification Letter is user's acknowledgement of receipt of the manual . |  | 1 | 1 unit | Distributed in HI. | Terminated | May 20, 2015 | 106 | 1 | Not_Computer
Z-1584-2015 | 71122 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 ultrasound systems at software versions VA20A or VA20B. Model number 10433816. |  |  | July 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | In certain situations, loss of post exercise images can occur in Stress Echo. | Other | On 07/11/2012 a field corrective action was initiated via update programs US016/12/P to update all affected SC2000 customers to software version VA20C, which contains the fix to the described issue. Each customer received a site visit to install the software update. |  | 476 | 476 | Worldwide Distribution -- US and Albania, Australia, Austria, Bangladesh, Bulgaria, Canada, Colombia, Denmark, Ecuador, Egypt, France, Germany, Hungary, Iceland, India, Israel, Italy, Japan, Kenya, Kuwait, Malaysia, Mauritius, Mexico, Netherlands, P.R. China, Philippines, Portugal, Rep. of Yemen, Republic Korea, Russian Fed., San Marino, Saudi Arabia, Singapore, Spain, Suriname, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E., United Kingdom, and Venezuela. | Terminated | May 06, 2015 | 1029 | 476 | Not_Computer
Z-1582-2015 | 70951 | No | IYH | X-ray radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 6.x. is software used to place patients at an accurately defined point within the treatment beam of a medical accelerator for stereotactic radiosurgery or radiotherapy procedures. |  |  | March 06, 2015 | 2015 | Brainlab AG | ExacTrac 6.x Patient Positioning System: Potentially incorrect patient positioning when using the ExacTrac Cone Beam CT (CBCT) with a CBCT acquired at a couch angle other than 0.0 degrees. | Software in the Use Environment | FIELD SAFETY / PRODUCT NOTIFICATION letters dated March 6, 2015 were sent to all customers of record. The letters included instructions for customers to: 1) With immediate effect, always make sure the couch is at exactly 0.0 degrees (displayed on the Varian monitor) when acquiring a CBCT set to import into the ExacTrac software; 2) For Varian systems providing an option to prevent a CBCT acquisition for couch angles other than 0.0 degrees, Brainlab recommends selecting that option for use in combination with ExacTrac; and, 3) Continue to verify any ExacTrac CBCT based correction using the ExacTrac X-ray verification and/or an external IGRT procedure as mandatory. The letters further advised that Brainlab will be taking the following actions: 1. Brainlab provides existing potentially affected ExacTrac v.6.x CBCT Import &amp; Alignment Software customers (with a Varian C-Series linac) with this product notification information; and, 2. Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact the customers starting August 2015 to schedule the update installation. Customers with questions about this issue were advised to contact Brainlab at 800-597-5911 or via e-mail by writing at us.support@brainlab.com. |  | 25 | 25 systems (US only); 53 systems (Foreign) | *** US: Nationwide; *** FOREIGN: Austria, China, Denmark, France, Germany, Hong Kong, Japan, Netherlands, Qatar, Russian Federation, South Korea, Switzerland | Open, Classified |  |  | 25 | Not_Computer
Z-1586-2015 | 70952 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | BARD MAX-CORE Disposable Core Biopsy Instrument |  |  | April 09, 2015 | 2015 | Bard Peripheral Vascular Inc | The firm is recalling Bard MAX-CORE Disposable Core Biopsy Instrument due to the risk of having self-activation-related issues. | Process design | Bard sent an Urgent Medical Device Recall Notification letter dated April 9, 2015 to their customers. The firm's letter provided customers with the following instructions:-Do not use or further distribute any affected product.-Check all inventory locations within the institution for BardÂż Max-Core DisposableCore Biopsy Instruments for the recalled product code/lot number combination.-If further distributed any of the product code/lot numbers listed, please immediately contact that location, advise them of the recall and have them return the affected product to BVD address listed.-Remove any identified product from the shelves.-If used the affected product, complete, and return the attached Recall and Effectiveness Check Form indicating no product will be returned.Report any new and/or previously unreported adverse events associated with this recall to the US Food and Drug Administration's ("FDA") MedWatch Program by phone at 1-800-FDA-1088; by fax at 1-800- FDA-0178; by mail to MedWatch. Food and Drug Administration. 5600 Fishers Lane. Rockville. MD 20857-9787: or on line at http://wwvdda.gov/medwatch/report.htm.Once the product affected by this recall has been removed from inventory:-Fill out the Recall and Effectiveness Check Form. Be sure to state the quantities and lot numbers of each recalled product that are in stock. -Call the firm's Recall Coordinator Regan Gealy at 1-770-784-6471 or email at Regan.Gealy@crbard.com. The Recall Coordinator will issue either a Return Authorization (RCL) Number or Consignment Recall Number (CRC) to facilitate the expedient return of the product.-Fax the Recall and Effectiveness Check Form back to at 1-770-784-6469.-BPV will provide replacement product for the returned product.A mailing label is going to be enclosed to return the affected product. Mark the outside package as "RECALLED PRODUCT" and include the RCL or CRC number. All products should be returned to the follo |  | 7995 | 7,995 units | Worldwide Distribution - US Nationwide in the states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WI, WV and the countries of: France, Mexico, Netherlands, Panama, United Kingdom, Switzerland, Spain, Poland, Norway, Iceland, Italy, Greece, Finland, Belgium, Germany, Austria, China, Taiwan, Singapore, Korea, Australia, and Japan. | Terminated | September 25, 2015 | 169 | 7995 | Not_Computer
Z-1591-2015 | 71028 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 113" (287cm) Transfer Set w/MicroClave T-Connector, Check Valve, Anti-Siphon Valve, Smallbore Quadfuse Ext Set w/MicroClave, 4 Clamps (Blue, 2 Red, White), Rotating Luer, Non-DEHP Tubing, Item No. B33303 |  |  | April 20, 2015 | 2015 | ICU Medical, Inc. | ICU Medical is recalling MicroClave T-Connectors because they have potential connection separation issues that could lead to leakage. | Under Investigation by firm | An Urgent Medical Device Recall letter dated 4/17/15 will be sent to customers on 4/20/15 to inform them that ICU Medical Inc. has identified a potential connection separation issue in a small number of isolated Intravenous Administration Transfer Sets. While the separation may be visually detected, if used, this failure mode could lead to leakage. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to return completed recall response form to ICU Medical by Fax at (801) 264-1755 or email to recall@icumed.com. Customers with questions or require assistance relating to the recall are instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address: customerservice@icumed.com. |  | 900 | 900 units | Nationwide in US: SC, NC, MN, CT, KY, MD, FL, VA, TN | Terminated | November 24, 2015 | 218 | 23100 | Not_Computer
Z-1587-2015 | 70933 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Universal Charger |  |  | March 25, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | The Stryker Universal Battery Charger is not transmitting usage data to the Stryker Cloud as designed. The firm is initiating a software correction to address the issue. | Software design | Stryker issued e-mail notices to distributors and reps on March 26, 2015 advising that all battery charges would receive a software update to provide a field correction for the upgrade to units to be able to track device usage data to Stryker as designed. There is no customer action required, the chargers will transmit a message to Stryker once the upgrade is complete. Please contact Kelly Jo Davis at 269-389-2921 with any questions. |  | 223 | 223 chargers | US Nationwide in the states of AL, AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, MD, MI, MO, MS, MT, NC, NH, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, VT, WA, WI and the foreign country of: Canada | Terminated | July 08, 2015 | 105 | 223 | Not_Computer
Z-1596-2015 | 70662 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | ThermoFlo 1 HME, Disposable, Non-sterile, Expiry date 04/17. |  |  | February 19, 2015 | 2015 | Arc Medical Inc | Outer shipping carton was marked with an incorrect expiry date, however inner product carton was labeled correctly. | Process control | Consignees were contacted by telephone on 02/19/2015. |  | 1810 | 1810 units | Distributed in CA and NC. | Completed |  |  | 1810 | Not_Computer
Z-1597-2015 | 70950 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Spinal Elements 5.5 x 45MM Mercury lordosed rod. A component of the Mercury Spinal System, comprised of a variety of screws, hooks, rods, connectors, and staples that are used for attachment to the non-cervical spine. Rods span the distance between screws and hooks and achieve fixation by the mechanical joining of the rods with the screws or hooks. Lordosed rods are pre-curved to better fit the patient anatomy. |  |  | April 09, 2015 | 2015 | Spinal Elements, Inc | Some rods may not meet the appropriate traceability requirements including the lot number. Some rods were found to not include the product identification number (product number and lot number) marked directly on the rod. The lack of marking interrupts the ability to maintain traceability of the device. | Under Investigation by firm | The firm sent customer notification letters on 04/09/15. In this letter the firm requested that all affected lots be returned, and replacement inventory will be made available.The firm requested that customers take the following actions:1. Immediately examine inventory for part number 60100-045.2. Immediately discontinue the use and/or distribution of the affected lots3. Send all affected product back to:Spinal Elements, Inc.ATTN: Julie Lamothe3115 Melrose DriveCarlsbad, CA 920104. If further distributed the product to please identify the customers and notify them as once.5. Complete and return the enclosed "Customer Response Form." Any questions the firm states to contact the Customer Service department at 1.877.SPINAL5 Ext. 213. |  | 588 | 588 | Nationwide Distribution including the states of GA, CA, TX, CO, IN, NC, FL, UT, TN, MO, MI, PA, OH, and OK. | Terminated | June 25, 2015 | 77 | 588 | Not_Computer
Z-1598-2015 | 71190 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Nuclear Gamma Cameras (ADAC Vertex Classic, ADAC VERTEX PLUS/SOLUS/CARDIO, ADAC VERTEX V60) Philips Medical Systems, Cleveland, OH |  |  | March 13, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | During clinical use of a Vertex Plus Gamma Camera, as the detector heads were in the relative 180 degree position, Detector #1 drifted down and made contact with the patients chin during a whole body study acquisition. | Employee error | Philips sent an Customer Information letter dated March 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to :1. Do not lubricate the clutch(es) with oil, grease or any other type of lubricant..2. Follow product Instructions For Use (IFU), CARDIO. SOLUS, ADAC VERTEX, AND ADAC VERTEX PLUS WITH ATLAS USERS MANUAL, 9201-0227E-ENG, REV A- General Precautions that instructs user to monitor the patient during scanning.3. Follow product Instructions For Use (IFU), CARDIC, SOLUS, ADAC VERTEX, AND ADAC VERTEX PLUS WITH ATLAS USER'SMANUAL, 9201-0227E-ENG, REV A- Emergency Patient Removal , Section Safety and Regulatory Agency Compliance, page 2-11, instruction(s) on emergencypatient removal.4. Ensure completion of product Preventive Maintenance (PM) procedure requiring inspection of the radius drive clutches every 3 months for contamination and replacement if contaminated and, testing of the gantry system.If you need any further information or support concerning this issue, please contact your local Philips representative: Technical Support Line: 1-800-722-9377. |  | 918 | 918 units | Worldwide Distribution - US (nationwide) including Puerto Rico and to the states of : AL, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV., and to the countries of : Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Costa Rica, Denmark, Egypt, France, Germany, Ireland, Italy, Japan, Lebanon, Malaysia, Netherlands, Norway, Pakistan, Portugal, Republic of Korea, Russia, South Korea, Spain, Sweden, Syrian Arab Republic, Thailand and United Kingdom. | Open, Classified |  |  | 918 | Not_Computer
Z-1600-2015 | 70975 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | TFN-ADVANCED Proximal Femoral Nailing System (TFNA); for treatment of fractures in orthopedics and trauma. |  |  | April 09, 2015 | 2015 | Synthes, Inc. | TFNA nails from certain lots were assembled with a locking mechanism too close to the top of the nail. Potential patient impact: surgical delay. | Process control | An urgent Field Safety Notification, dated April 9, 2015, was sent to end users and sales consultants to alert them about the issue and request that they take the following actions: 1) review the instructions listed in this notification; 2) forward this notice to anyone in your facility that needs to be informed; and 3) complete the attached Verification Section (page 4 of this notification) and send a copy of the completed Verification Section by: Fax: (888) 446-5160 orScan/email: Synthes3152@stericycle.com. |  | 8175 | 8175 | Distribution US Nationwide and Switzerland. | Open, Classified |  |  | 8175 | Not_Computer
Z-1599-2015 | 71138 | No | PJQ |  |  | Cardiovascular | PMA | VenaSeal, Sapheon Closure System, Model # SP-101, VS-301 and VS-402. For vascular use. |  |  | April 28, 2015 | 2015 | Covidien LLC | Possible sterility breach of the outer packaging (pouch) material which may compromise the outside surface sterility of the sealed inner tray. | Package design/selection | Consignees were contacted by letter via email on April 28, 2015. |  | 1661 | 1661 units | Worldwide distribution. No U.S. accounts; Australia, Canada, France, Germany, Hong Kong, Ireland, Italy, Latvia, Netherlands, Norway, Portugal, Switzerland, UAE and the United Kingdom. | Terminated | March 29, 2016 | 336 | 1661 | Not_Computer
Z-1605-2015 | 71125 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Tapered HD Upgrade Set, REF TSKHDUS, Non-Sterile, Rx Only, |  |  | April 16, 2015 | 2015 | BioHorizons Implant Systems Inc | BioHorizons Tapered HD Upgrade set was packaged with an incorrect drill component manufactured of stainless steel in place of the correct component manufactured from titanium. | Process control | BioHorizons sent an Urgent Medical Device Field Correction letter on 04/16/2015 to affected customers. Non-responding customers will be notified after 30 days. The letter identified the affected product, problem and actions to be taken. For questions contact Customer Care Department at 888-246-8338. |  | 48 | 48 units | US Nationwide in the states of: AL, FL, MS, VA, LA, CA, NC, TX, TN, OH, OK, VT, WA, MD, NY, SD, and SC. | Terminated | August 07, 2015 | 113 | 48 | Not_Computer
Z-1601-2015 | 70868 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Vitek 2 AST-YS06 REF 412 610, Fungal Susceptibility Card, 20 cards per carton, IVD. |  |  | March 27, 2015 | 2015 | Biomerieux Inc | The test kits have Out-of-Range-Low (OORL) Flucytosine QC results for Candida krusei ATCC 6258 | Other | The notification, consisting of the Field Safety Corrective Action (FSCA #2488) packet, was issued electronically to the subsidiaries and distributors on 3/27/15. FSCA #2509 was issued electronically to the subsidiaries and distributor on 4/13/15, which superseded FSCA #2488. |  | 1025 | 1,025/20-card cartons | Distributed only to foreign subsidiaries and distributors, including Canada and Mexico. | Terminated | March 24, 2017 | 728 | 46164 | Computer
Z-1606-2015 | 71211 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B850, B650 and B450 (Bx50). |  |  | May 04, 2015 | 2015 | GE Healthcare | The heart rate could be calculated from pacer pulses without indication that pacemaker detection is OFF in some combinations of Bx50 monitors, a PDM, and CIC/CSCS. Then the monitor does not show the Pacer Off message and there may not be an alarm for asystole. Undetected asystole could result in irreversible changes in the patientÂżs condition and delayed or missed life sustaining patient treatment | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated May 1, 2015, to all affected consignees. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 110 | 110 (83 units US, 27 units OUS) | Worldwide Distribution - US including CO, MI, WI and Internationally to NETHERLANDS, FINLAND, and BELGIUM. | Terminated | September 15, 2015 | 134 | 110 | Computer
Z-1607-2015 | 71135 | No | LMC | Medical image hardcopy device. | Radiology | Radiology | 510(K) Exempt | Nuclear gamma cameras (ADAC VERTEX PLUS/SOLUS/CARDIO, ADAC VERTEX V60 &amp; ADAC CARDIO C60) |  |  | March 11, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm was notified by a customer that the collimator exchange carriage and the collimator storage cabinet were misaligned. | Device Design | Philips sent an Urgent - Medical Device Correction letter dated March 12, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm requested their customers to do the following: " If a misalignment on the top and/or bottom of the guide rail between the cabinet and carriage can be visually detected do not continue with the collimator exchange procedure," If the misalignment is visually detected, be aware that the customer/user cannot remedy this issue and will require service interaction prior to the continued use of the system." This letter should be placed in your User Documentation until otherwise notified.If you need any further information or support concerning this issue, please contact your local Philips Healthcare representative: For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter site ID or follow the prompts). |  | 975 | 975 units | Worldwide Distribution - US (nationwide) to the states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY., and to the countries of : Australia, Austria, Belgium, Brazil, Canada, China, Costa Rica, Egypt, France, Germany, Hong Kong, Ireland, Italy, Japan, Korea, Lebanon, Netherlands, Norway, Pakistan, Portugal, Russia, Singapore, South Korea, Spain,Thailand and United Kingdom. | Open, Classified |  |  | 975 | Not_Computer
Z-1610-2015 | 71057 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer 100 mL Hemovac Device Hemovac Mini Evacuator. 00-2568-000-10. Zimmer Surgical, Inc. |  |  | April 21, 2015 | 2015 | Zimmer Surgical Inc | Devices were placed in a shipper carton without the appropriate number of the Instructions for Use (IFU). Packaging specifications require that each sales unit is accompanied by an IFU. Only one IFU was present in the shipper carton, instead of one IFU per each of the two sales units. | Other | On 4/28/2015 the firm sent Urgent: Medical Device Removal letters to their customers. The letters identified the affected product in addition to stating that the reason for the recall is due to missing Instructions for Use from sales units. The letters discussed the clinical implications or risk to health, and actions required. Customers are to review the notification, identify and quarantine the affected product to prevent further distribution or use, and return the completed response form to CorporateQuality.PostMarket@zimmer.com and the affected product to the address provided. A credit will be issued for unused devices upon receipt. Questions regarding this information should be directed to 330-364-0989. |  | 36 | 36 units | Worldwide Distribution -- US, including the states of AZ, CA, ID, IL, MD, NC, RI, and SC; and, Canada. | Terminated | November 06, 2015 | 199 | 36 | Not_Computer
Z-1608-2015 | 71233 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs qube Compact Monitor, Model 91390. The monitor is functioning as either bedside or central monitors; passively displays data generated by Spacelabs Healthcare parameter modules, Flexports interfaces, and other SDLC based products in the form of waveform and numeric displays, trends and alarms. |  |  | May 08, 2015 | 2015 | Spacelabs Healthcare Inc | The input power can fail even though the power source at the wall has not been interrupted. In addition, there is the potential for the qube Compact Monitor to fail to turn on. This is a retrospective recall related to a recall initiated in 2012 (Z-0403-2013). | Software design | The firm sent the Urgent Medical Device Correction- Spacelabs Healthcare Inc.-qube Compact Monitor, Model 91390, Power Input PCBA Failure letter, dated 8 May 2015, via priority service mail to US consignees.This letter described the affected system, problem, actions to be taken by users and corrective actions conducted by Spacelabs.For additional information or technical assistance, please contact:Technical SupportSpacelabs Healthcare, Inc.35301 SE Center St.Snoqualmie, WA 980651-800-522-7025 and select 2 for Technical Support |  | 5 | 5 units distributed in the US | distributed in North Carolina, Texas, and Wyoming | Terminated | November 05, 2015 | 181 | 5 | Not_Computer
Z-1604-2015 | 71099 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD) Pump Driveline Splice Kit |  |  | April 29, 2015 | 2015 | HeartWare Inc | Failures of the splice repair kit if exposed to excessive force. | Device Design | HeartWare, Inc. sent an Urgent Medical Device Correction letter to all affected consignees on May 11, 2015. The letter identified the affected product, problem, and actions to be taken by the consignees. The firm requested that consignees: 1) review the notice and the Patient Communication contents, 2) forward the notice to individuals within their organization who need to be aware of the notice, 3) identify patients that are currently on support and have undergone a driveline splice repair, 4) distribute the Patient Communication to the affected patients in person, 5) continue to reinforce the messages described in the notice with patients who have experienced a splice repair during their regularly scheduled appointments, 6) complete, sign, and return the "Acknowledge and Completion Form to HeartWare within 30 days of receipt of the letter. Should any questions or concerns arise, please contact there local HeartWare representative. Their 24-hour Clinical Support personnel are also available at 1 (888) 494-6365. |  | 35 | 35 | Worldwide Distribution - US Nationwide in the states of FL, KY, MO, NY, TN,TX, and the countries of: Canada, Germany, Italy, Netherlands, Turkey and United Kingdom. | Open, Classified |  |  | 35 | Not_Computer
Z-1614-2015 | 70784 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Kinemax Stem Extenders, 80 mm, |  |  | February 27, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker received four customer complaints reporting that inside the box labeled part #6476-8-260 (Kinemax Stem Extenders, 80 mm) there was a different part #6476-8-250 (Kinemax Stem Extenders, 40 mm. | Mixed-up of materials/components | Stryker Orthopaedics an Urgent Medical Device Recall Letter/Notification Acknowledgement Form dated April 9, 2015 to their affected customers. The recall letter identified the affected product, problem and actions to be taken. Customers are asked to contact their Stryker Orthopaedics Sales Representative to arrange for return of the product if they have any of the affected items. Customers are asked to complete and return the acknowledgement form within 5 days by emailing to: so_m_product_field_action_response@stryker.com or faxed to: 1-855-251-3635. For question call 201-831-6365. |  | 4 | 4 units | US Nationwide distribution in the states of MA and RI. | Open, Classified |  |  | 4 | Not_Computer
Z-1615-2015 | 70913 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension EXL with LM, Dimension EXL 200, Dimension RxL, Dimension RxL HM, Dimension RxL Max, Dimension RxL Max HM, Dimension Xpand, Dimension Xpand HM, Dimension Xpand Plus, Dimension Xpand Plus HM. |  |  | April 02, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Certain Dimension Integrated Chemistry Systems Instrument Power Configurations may have been installed with a non-standard instrument power plug.The use of a non-standard plug configuration as described may lead to a potential electrical hazard. The likelihood of potential clinical impact or injury due to this issue is extremely low. | Under Investigation by firm | Siemens sent an Urgent Field Safety Notice, dated April 2015, was issued for distribution to affected customers in Japan to notify them of the issue. A Siemens Service Representative will visit customer sites to inspect the Dimension instrument power plug and wall outlet configuration. If the instrument is configured with a non-standard 15 Amp plug/outlet or a 20 Amp twist lock plug, Siemens will correct the issue identified, ensuring that the system is installed according to the instrument requirements. |  | 17801 | 17801 | Distributed in the country of Japan. | Open, Classified |  |  | 17801 | Not_Computer
Z-1612-2015 | 70890 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Fujifilm Synapse Cardiovascular I ProSolv CardioVascular |  |  | February 04, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | Under certain circumstances the patientDemographics in a report exported into the EMR may not match the demographics shown in the corresponding Synapse CV clinical report. | Software in the Use Environment | Fujifilm Medical on 2/4/14 , contacted users and advised to discontinue using HL7 reports pending FMSU determination of root cause. Firm issued follow-up letters on 2/24/14 and 7/24/14 advising users of a mandatory upgrade to version 5.1.1 I base, in-place upgrade, to correct the problems. |  | 16 | 16 units | Nationwide Distribution including PR and the states of IL, IN, FL, GA, MD, MI, MO, MS, NC, NY, OR, and PA, PR and the country of Canada. | Open, Classified |  |  | 16 | Not_Computer
Z-1611-2015 | 70891 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | FujiMedical Synapse Cardiovascular I ProSolv CardioVascular. |  |  | February 04, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | Measurement values in the EMR may not accurately reflect the actual measurement value on the Synapse Cardiovascular clinical report. | Software in the Use Environment | FujiMedical Systems North America issued intial recall letter dated February 4, 2014, advising users of the HL7 report exported to the EMR wherein measurement values presented in the Hl7 reports in your EMR may not accurately reflect the actual measurement value on the Synapse Cardiovascular clinical report.Follow-up letters were issued February 12, 2014, and July 24, 2014. A mandatory upgrade to software version 5.1.1/ base, in-place upgrade required. A customer response form attached to acknowledge receipt and arrange for the on site upgrade. For questions regarding this recall call 203-324-2000. |  | 30 | 30 units | Worldwide Distribution - US including CA, FL, IL, IN, MD, MI, MN, MO, NC, NY, OH, OR, PA, PR, TX, VA, WIForeign: Canada | Open, Classified |  |  | 30 | Not_Computer
Z-1613-2015 | 70995 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | ReTurn belt, wipeable, sizes Small through Extra Large, Plastic pouch, 1 item per pouch. Patient transfer aid. |  |  | March 26, 2015 | 2015 | Handicare USA, Inc. | Complaints were received that the heat-welded seam on the wipeable belt's strap tore when force was applied. | Device Design | Handicare sent an Voluntary Product Recall Notification letter dated April 2, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were requested to quarantine any remaining stock and to contact customers if product was further distributed and to return the affected product for a replacement and/or full refund.Also, enclosed is a response form. Please complete and return the form even if you do not have any product to return. Your Sales representative can assist you in completing this form. This information is essential in order to maintain recall effectiveness information (required by FDA).Thank you for your support. We are committed to re-releasing the wipeable ReTurn belt as soon as we believe that it meets the quality standards that you expect from Handicare. Please contact our Customer Experience team at 1-866-276-5438 if you have any questions regarding the recall or would like assistance with the recall. We regret any inconvenience that this action may cause, but we appreciate your understanding as we take action to ensure patient and customer satisfaction. |  | 64 | 64 | US Distribution to the states of : FL, CA, NJ, NC and MD. | Terminated | August 17, 2015 | 144 | 64 | Not_Computer
Z-1620-2015 | 71027 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | The Integra Licox Pt02 Monitors are compact, portable device for use with Pressure/Temperature, Pressure and Oxygen Partial Pressure Transducer-Tipped Catheters. |  |  | April 14, 2015 | 2015 | Integra LifeSciences Corp. | Integra LifeSciences Corporation is recalling Integra's Camino ICP Monitor (CAM 02) &amp; Licox PtO2 Monitor ( LCX02 &amp; LCX02R) due to complaints that the audible speaker malfunctions. | Device Design | Integra initiated this recall by sending an Urgent Voluntary Medical Device Recall Letter/Recall Acknowledgement Form dated 4/14/15 via traceable courier service to their affected customers. |  | 59 | 59 monitors | Worldwide Distribution. | Open, Classified |  |  | 1814 | Not_Computer
Z-1616-2015 | 71045 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | COULTER DxH Diluent, Catalog No. 628017, Part No. A59956. |  |  | April 13, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter has received an increased number of complaints related to Hemoglobin (HGB) on UniCel DxH 800 and DxH 600 Coulter Cellular Analysis Systems. The use of the DxH Diluent lots listed in the recall notice may eventually result in a compromised HGB chamber. | Nonconforming Material/Component | An Urgent Medical Device Recall letter dated 4/13/15 was sent to customers who purchased the Coulter DxH Diluent. The letter informs customers that Beckman Coulter has received an increased number of complaints related to Hemoglobin (HGB) on UniCel DxH 800 and DxH 600 Coulter Cellular Analysis Systems. The letter informs customers of the problems identified and the actions to be taken.Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Customer Support Center via http://www.beckmancoulter.com/customersupport/support or call (800) 526-7694 in the US and Canada. Customers outside the US and Canada are instructed to contact their local Beckman Coulter Representative. |  | 829888 | 829,888 units total (659,584 units in US) | Worldwide Distribution -- Argentina; Brazil; Brunei Darussalam; Canada; Chile; Colombia; Hong Kong; India; Lebanon; Macao; Mexico; Panama; Philippines; Qatar; Singapore; Taiwan; Thailand; Trinidad and Tobago; United States, including Puerto Rico; and Uruguay. | Terminated | October 01, 2015 | 171 | 829888 | Not_Computer
Z-1622-2015 | 71034 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Homocysteine test system - Homocysteine Reagent. |  |  | March 23, 2015 | 2015 | Roche Diagnostics Operations, Inc. | Customers complained about under-recovery of non-Roche controls and discrepant low patient results with certain Homocysteine reagent lots. This negative bias could, in the worst case, lead to inaccurately low Homocysteine results. An elevated level of Homocysteine is considered an important risk factor in the assessment of peripheral vascular disease. Falsely low values could lead to a delay of di | Nonconforming Material/Component | Roche Diagnostics sent URGENT MEDICAL DEVICE CORRECTION notifications dated 3/23/2015 to affected direct accounts via UPS. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 if you have questions. |  | 2558 | 2558 units | Nationwide Distribution-including PR and the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, MA, MD, ME, MI, MO, MT, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, , SC, TN, TX, VA, WA, and WV. | Terminated | January 19, 2016 | 302 | 2558 | Not_Computer
Z-1623-2015 | 71167 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | CT LightSpeed 16 and Discovery ST Pet/CT DST. Catalogue # 2339985. |  |  | April 22, 2015 | 2015 | GE Healthcare | GE Healthcare has become aware that 3 bolts that secure a power supply to the rotating side of the gantry were over-torqued on certain Lightspeed 16 CT scanner or PET/CT DST Scanner.. In the unlikely event that the component becomes loose on the CT gantry during operation, it could result in serious bodily injury if the component were expelled. There were no incidents or injuries reported as a res | Process design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 25461 dated April 22, 2015. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology/Cardiology, and Radiologists/Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The letter stated that customers can continue to use their device. If the error occurs, then stop using the system and call a service representative to bring the system back into proper operation. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative.UPDATED:Consignees were sent on 5/14/2015 a second GE Healthcare "Urgent Medical Device Correction" letter dated May 14th, 2015. The letter was to incorporate and additional unit being recalled. |  | 4 | 4 USA | Distributed to the states of MN, FL &amp; TX. | Terminated | July 15, 2015 | 84 | 4 | Not_Computer
Z-1621-2015 | 71049 | [6 empty] | MYLA server HP Proliant: DL380-G8 and ML350-G6 version 3.xx computer application software. |  |  | April 14, 2015 | 2015 | bioMerieux, Inc. | The MYLAÂż server could slow down due to the volume ( weight ) of the data to manage and it could also stop processing, impacting the server start-up phase. The problem occurs in high volume settings. The MYLA server slows down and will stop processing completely if any task requires more than 5 minutes to complete. The server then fails to reboot and requires service, provided remotely or on-site | Software design | Consignees were notified electronically on April 14, 2015 and letter via FedEx on 04/20/2015. |  | 730 | 730 units | Worldwide Distribution-Algeria, Argentina, Australia, Austria, Brazil Canada, Chile China, Colombia, Czech Republic, France Germany , Greece, Hungary, India, Italy Japan, Korea, Mexico, Netherlands, Poland, Portugal, Russia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, Angola, Armenia, Kocaeli, Bahrain, Bangladesh, Republic of Belarus, Benin, Bolivia, Bosnia, Brunei, Bulgaria, Burkina Faso, Cambodia, Cameroon, Chad, Democratic Republic of Congo, Costa Rica, Germany Dominican Republic, Dutch Antilles, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Gabon, Georgia, Guam, Guatemala, Guinea, Honduras, Indonesia, Iran, Iraq, Israel, Ivory Coast, Kazakhstan, Kenya, Kuwait, Laos, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Mali, Malta, Morocco, Portugal, Myanmar, Managua, Niger, Nigeria, Oman, Palestine, Panama, Paraguay, Peru, Philippines, Qatar, Romania, Kingdom of Saudi Arabia, Senegal, Slovakia, Sri Lanka, Tanzania, Togo, Tunisia, Ukraine, UAE, Uraguay, Venezuela, Vietnam, Yemen and Zimbabwe. | Open, Classified |  |  | 730 | Not_Computer
Z-1617-2015 | 71016 | No | MHY |  |  | Neurology | PMA | Medtronic DBS Extension Kit for Deep Brain Stimulation, Models 37085 and 37086. Sterile |  |  | April 16, 2015 | 2015 | Medtronic Neuromodulation | Medtronic is providing Important information regarding extension handling during implant procedure after receiving reports of high impedances. This applies to extension Models 37085 and 37086 which can be used with the followingimplantable Neurostimulators: ActivaÂż PC (Model 37601), ActivaÂż RC (Model 37612), and the ActivaÂżSC (Model 37603). | Under Investigation by firm | Medtronic began notifying consignees with an Urgent Medical Device Safety Notification letter beginning April 16, 2015. Consignees will be notified by mail and/or in person, by visit of a Medtronic field representative. The letter described the issue and provided recommendations of following the instructions defined within the Neurostimulator implant manual. A Physician Reply Form was asked to be returned via fax at 800-897-3899 or email to neuro.quality@medtronic.com. Consignees with questions can contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7am - 6pm CST. |  | 3294518983 | 86,851 (US 37,933 / OUS 48,918) | worldwide | Open, Classified |  |  | 3294518983 | Not_Computer
Z-1628-2015 | 70946 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | ABACUS SE (Single-Workstation Edition) and ABACUS ME (Multi-Workstation Edition) |  |  | April 08, 2015 | 2015 | Baxter Corporation Englewood | Baxter Corporation is conducting a field action for the ABACUS SE and ME models due to the possibility that the compounder will load the incorrect formula upon scanning the bag label when two or more different formula files have been created with the same order number. | Software design | Baxter sent a Safety Alert letter dated April 8, 2015 to affected customers via USPS first class mail. The letter identified the affected product, problem and actions to be taken. Customers were asked to 1) contact Baxter Technical Services if you need to restore the ABACUS databse, 2) foword a copy of this Safety Alert letter to other facilities or departments within your institution to ensure that those locations are awae of this action, and 3) complete the customer reply form and return it to Baxter by faxing to 224-270-5457 or scanning and e-mailing it to fca@baxter.com. For questions contact the Center for One Baxter at 800-422-9837. |  | 882 | 882 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 882 | Not_Computer
Z-1627-2015 | 71143 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | PHALINX (TM) HAMMERTOE HANDLE, SIZE: LARGE, STYLE: ANGLED, REF 45304014, CONTENTS 1 EACH, RX ONLY |  |  | April 16, 2015 | 2015 | Wright Medical Technology, Inc. | The term DORSAL was laser etched on the incorrect side on some of the angled handles. | Process control | Wright sent an Urgent Medical Device Recall letter dated April 13, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for the affected product, send it back to Wright Medical as soon as possible. Contact Customers Service at 901-867-4575 for return instructions and replacement inventory. Customers were instructed to complete the attached response immediately, confirming their receipt of this notice and the status of any product(s) on had, and return a copy by email to complaints@wmt.com or fax to 901-867-7401. Customers with questions should call 901-451-6144. |  | 126 | 126 units | Nationwide Distribution | Open, Classified |  |  | 497 | Not_Computer
Z-1555-2015 | 70916 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare DuraDiagnost X- Ray |  |  | February 09, 2015 | 2015 | Philips Medical Systems, Inc. | The system is designed to emit a beep upon termination of an exposure. However, if the system has been powered on for over 12 hours, the system will no longer emit this signal. This is a failure to comply with 21CFR 1020.31(j). | Software design (manufacturing process) | Philips Healthcare Planned action: 1. Philips healthcare will contact customers and initiate a software update to correct the software deficiency causing the problem.2. Philips engineers will install the software.3. Philips engineers will perform testing to ensure the software update was effective.4. Corrections are to be made free of charge5. Corrective actions should be completed by August 15, 2015For further questions call (978) 659-7892 |  | 3 | 3 DuraDiagnost | Worldwide Distribution - Nationwide Distribution and the countries of Australia, Austria, Argentina, Canada, China, Chile, Colombia Czech Republic, Great Britain, Germany, Egypt, Finland, France, India, Ireland, Japan, Malaysia, Macedonia, Mexico, Netherlands, New Zealand, Nepal, Peru, Poland, Russia, Saudi Arabia Singapore, Switzerland, Sweden, Spain, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 57 | Not_Computer
Z-1629-2015 | 71006 | No | CEE | Thyroxine-binding globulin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE/IMMULITE 1000 Systems TBG, catalog LKTB1 |  |  | April 09, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | IMMULITE/IMMULITE 1000 TBG catalog LKTB lot 257 is failing the adjustment slope instrument parameter of &lt;1.80. The resulting slope of the kit is &gt;1.80 resulting in the potential to produce depressed TBG patient values. | Unknown/Undetermined by firm | Siemens sent an Urgent Field Safety Notice dated April 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be Taken by the CustomerDiscontinue use and discard kit lot listed in Table 1.Review your inventory of these products to determine your laboratory's replacement needs.Please review this letter with your Medical Director.Complete and return the Field Correction Effectiveness Check Form attached to this letter within thirty (30) days.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions call (914) 631-8000. |  | 155 | 155 kits | Worldwide Distribution to the countries of : Denmark, Spain, Great Britain, Germany, Bosnia Herzeg, China and Russian Federation. | Terminated | March 22, 2017 | 713 | 155 | Not_Computer
Z-1635-2015 | 70994 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Triathlon femoral component. |  |  | August 15, 2014 | 2014 | Stryker Howmedica Osteonics Corp. | The peel strength of the inner blister may have been below internal validated requirements. The strength of the seal is directly related to long term shelf life of the sterile barrier. | Under Investigation by firm | The firm, Stryker Orthopaedics, notified their Branches/Agencies via email on 8/15/2014 and sent an "Urgent Product Recall" letter dated 8/15/2014 with a Product Recall Acknowledgement Forms to their Branches/Agencies and consignees. The letter describes the product, issue, potential hazards and actions to be taken. The consignees were instructed to complete and return the attached Product Recall Acknowledgment Form within 5 days via fax to 855-251-3635; and return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah NJ 07430.If you have any questions, feel free to contact Sr. Regulatory Compliance Specialist to 201-831-5826. |  | 1147 | 1,147 in total | Worldwide Distribution: US (nationwide) including PR and countries of: Australia, Canada, China, Brazil, France, Hong Kong, India, Italy, Mexico, Netherlands, Singapore, South Africa, Sweden, and Spain. | Open, Classified |  |  | 1147 | Not_Computer
Z-1632-2015 | 71105 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ 7 Ultrasound System versions 1.3.2 or lower, WITHOUT the Pediatric Cardiology option, Model: EPIQ 7G, EPIC 7C, EPIQ 7W. Catalog Number: 795200 / 795201 |  |  | April 14, 2015 | 2015 | Philips Ultrasound, Inc. | When EPIQ 7 Ultrasound System, WITHOUT the Pediatric Cardiology option, is set up to "Metric" and weight and/or height is entered, a unit conversion error may result in the incorrect display of their values and incorrect calculation of the patients Body Surface Area (BSA). | Software design | Philips sent the Urgent-Medical Device Correction-Philips Model EPIQ5 Ultrasound system and the Urgent-Medical Device Correction-Philips Model EPIQ7 Ultrasound system letters, both dated 2015 APR 14, to their US consignees. Philips sent the Urgent-Field Safety Notice-Philips Model EPIQ5 Ultrasound system letter and the Urgent-Field Safety Notice-Philips Model EPIQ7 Ultrasound system letter (both letters dated 2015 APRIL 14) to their foreign consignees. The EPIQ5 or EPIQ7 Ultrasound systems WITHOUT the Pediatric Cardiology option are affected. Philips issued two Field Change Orders (FCO) to their Field Service Engineer to install upgraded software: -FCO 79500340 for EPIQ5 -FCO 79500341 for EPIQ7 Customers with questions can call Philips Customer Service at 1-800-722-9377.Field Service Engineers (FSE) to install software version 1.3.3 per FCO 79500340 instructions for EPIQ5 5 and per FCO 79500341 instructions for EPIQ7. |  | 1692 | 1692 units total (902 units in the US and 790 units outside the US) | Worldwide Distribution | Open, Classified |  |  | 2800 | Not_Computer
Z-1630-2015 | 70828 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Carter-Thomason II Port Closure System; Model Number(s) CTI-1015P, for use in endoscopic and laparoscopic surgery. |  |  | March 20, 2015 | 2015 | CooperSurgical, Inc. | The pad printing around the suture entry holes on the 15mm Suture Guide included in the kit may not have been appropriately manufactured resulting in the padprinting flaking off around the entry holes. | Component design/selection | Consignees were notified of the recall by letter sent via FedEx dated March 20, 2015. The letter instructs consignees to do the following: discontinue use of the affected product, complete the attached acknowledgement and receipt form to return the product and schedule a replacement. |  | 430 | 430 units | US Nationwide distribution. | Terminated | July 29, 2015 | 131 | 430 | Not_Computer
Z-1636-2015 | 71127 | No | OYD |  |  | General &amp; Plastic Surgery | PMA | MelaFind Device, Model MF 100 US, optical diagnostic device for melanoma detection. |  |  | April 02, 2015 | 2015 | MELA Sciences, Inc. | The probability and histogram data within the Melafind's device displayed user interface lacks approval (not included in the PMA supplement). | Software design | The probability and histogram data within the Melafind's device displayed user interface is not included in the PMA supplement |  | 65 | 65 units | AL, CA, CT, CO, FL, HI, IL, MA, MD, MI, NH, NJ, NY, OH, PA, RI, SC, TN, TX, VA and WA | Open, Classified |  |  | 65 | Not_Computer
Z-1642-2015 | 71193 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | DeRoyal(R) Medium Specimen Retrieval Bag, REF 5500-000-003, Rx Only |  |  | April 17, 2015 | 2015 | DeRoyal Industries Inc | The firm received reports of specimen retrieval bags tearing, ripping, and coming apart during use. | Device Design | DeRoyal sent an Urgent: Voluntary Recall letter dated April 17, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to 1. Check your inventory for the product and specific lot number from the attached Recall Response Form.2. If you have affected inventory, please destroy it and dispose of it according to your facilitys procedures.3. Record all information on the included Recall Response and Notice of Destruction forms.4. If you purchase through a distributor, please indicate their name/location on the Notice of Destruction Form so that we may issue credit appropriately. 5. Return both forms to recalls@deroyal.com or 865-362-3716 (fax) even if you no longer have affected inventory.6. Once we have received your Recall Response and Notice of Destruction forms, you will receive credit for destroyed product as indicated on the form.7. If you are a distributor of these products, please forward this recall notice on to your end-users, or provide us with your customer listing and we will gladly contact them for you.For questions regarding this recall, please contact your DeRoyal Sales Representative or DeRoyal Customer Service at 1-800-251-9864. |  | 5359 | 5359 units | US Nationwide Distribution | Terminated | December 29, 2015 | 256 | 5359 | Not_Computer
Z-1643-2015 | 71117 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Pinnacle3 Software Version 10.0, Model 459800200841. |  |  | April 23, 2015 | 2015 | Philips Medical Systems, Inc. | A problem (the dose computed in planning mode is calculated incorrectly ) has been detected in the Philips Pinnacle3 software version 10.0 that, if it were to re-occur, could pose a risk for patients or users. Specifically, the dose engine is being passed the wrong snout position. The snout position is used to determine the penumbra of the beam which includes the calculation of the source si | Software design | Philips Healthcare sent a "Field Safety Notification PROST-T-300-01-06" letter dated 20 April 2015 to their customers. The letter described the Problem Description, Hazard Involved, How to Identify Affected Products, Action to be Taken by Customer / User, Actions Planned by Philips and Further Information and Support (Contact Philips Customer Care Solutions Center at 1-800-722-9377). |  | 4 | 4 | US Nationwide Distribution in the states of MO, FL, OH | Open, Classified |  |  | 4 | Computer
Z-1639-2015 | 71178 | preservative, cytological | LEA | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Mucolytic Agent 1 pint/473 ml; 1 gallon/3.8 l |  |  | April 29, 2015 | 2015 | Richard-Allan Scientific Company | Richard Allan Scientific (RAS) is initiating a recall of Mucolytic agent and Mucolexx following a customer complaint reporting that a lab had found fungal contamination in 2 lots of Mucolexx. | Material/Component Contamination | Richard Allan Scientific (RAS) initiated a voluntary recall of Mucolexx and Mucolytic Agent due to fungal contamination on 5/1/2015 via certified mail to all customers who received the affected. Customers were asked to return a repsonse form, and dispose or return all recalled products. Please contact Sarah Rickert at 269-544-5628 or sarah.rickert@thermofisher |  | 152478 | Product 76166- 116; Product Code 76167-29 | Worldwide Distribution - US Nationwide in the states of: AR; AZ; CA; CO; FL; GA; I A; IL; IN; KY; MA; MN; MO; MS; MT; NC; ND; NJ; NV; NY; OH; OK; PA; PR; SD; TN; TX; VA; WA; WI and the countries of: Bangladesh, Canada, China, Costa Rica, Dubai, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Korea, Netherlands, Panama, Poland, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, and UK. | Terminated | March 09, 2016 | 315 | 10142853 | Not_Computer
Z-1641-2015 | 71215 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ENDOPATH ECHELON" Vascular White Reload for Advanced Placement Tip, |  |  | February 26, 2015 | 2015 | Ethicon Endo-Surgery Inc | During an internal inspection of the production process the firm discovered an issue which may cause the cartridge to deploy an incomplete staple line. | Software in the Use Environment | Ethicon sent an Urgent Medical Device Recall letter dated February 27, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine your inventory immediately to determine if you have recalled product on hand, remove and quarantine the recalled product, communicate the issue to relevant operating room or materials management personnel or anyone else in your facility whoneeds to be informed, if any product included in this recall has been forwarded to another facility, contact that facility to arrange return, and complete the Business Reply Form (BRF) (Attachment B) confirming receipt of thisnotice within three (3) business days. The BRF may be sent to Stericycle by fax at 1-800-807-5967or by email at ees8760@stericycle.com. Please return the BRF even if you do not have affected product.For clinical or product support, please contact your local sales representative orcall our Customer Support Center at 1-877-ETHICON (1-877-384-4266) (7:30 a.m.  6:30 p.m. EST). If you need additional shipping labels or a communications package, contact Stericycle at 1-877-643-8419 (8:00 a.m.  5:00 p.m. EST) and reference Event8760. |  | 6744 | 6,744 devices | Worldwide Distribution - US Nationwide in the states of: AR, AZ, CA, CO, DE, FL, GA, IA, IL, KS, KY, LA, MA, ME, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA &amp; WI and the country of Hong Kong. | Terminated | November 17, 2015 | 264 | 6744 | Not_Computer
Z-1645-2015 | 71101 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Emotion 16, the intended use of computed tomography is to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angels or spiral planes taken at different angles. |  |  | April 20, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Software bugs in VC20B SP0a or SP1 software versions may cause issues that could make it necessary to rescan patients. Syngo Main UI may crash if the Patient Browser is scrolled with arrow keys of the keyboard. Tomo images may freeze on the screen when zoomed in or out under the mode of CAREVision. Sporadic displaying error in WorkStream4D application. DB may lock under heavy and multitasks. | Software design | Siemens sent an Customer Safety Advisory Notice dated April 20, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A remote software update is scheduled for release that will resolve the issue. For further questions, please call (610) 219-6300. |  | 75 | 75 | Nationwide Distribution including Puerto Rico. | Open, Classified |  |  | 75 | Not_Computer
Z-1656-2015 | 71177 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT. |  |  | March 16, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm was informed that while raising the patient couch on the system to perform an exam, the couch unexpectedly descended to the lowest point without being commanded to do so. | Employee error | On 3-18-2015 the firm sent Urgent Field Safety Notices to their customers. |  | 153 | 153 units total | Distributed in China. | Open, Classified |  |  | 153 | Not_Computer
Z-1663-2015 | 71022 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid Waste Bottle Kit, Catalog#10386287. |  |  | March 05, 2015 | 2015 | Siemens Medical Solutions Diagnostics | The bottle assemblies were manufactured with a quality issue where in some cases the smaller diameter opening on both IMMULITE 2000 and IMMULITE 2000XPi Water Bottles and Liquid Waste Bottles is undersized and/or deformed, preventing the cap from closing or fastening securely to the bottle. | Nonconforming Material/Component | Siemens Healthcare Diagnostics sent an Urgent Device Correction Letter/Field Correction Effective Check Form dated 3/5/15 via FedEx to their affected customers. |  | 15 | 15 bottles | Worldwide Distribution. | Open, Classified |  |  | 1417 | Not_Computer
Z-1669-2015 | 71115 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | SIEMENS Luminos Agile Max; a universal imaging system for radiographic and fluoroscopic studies. Using either film cassettes or a digital mobile flat detector, it can perform a range of applications including general R/F, angiography and pediatric examinations. Luminos Agile is applicable to emergency treatment on an outpatient basis, as well as for bedside examinations. |  |  | April 13, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Study and all acquired images deleted when using systems with software version VE10E. Acquisitions have to be repeated. Connection with portable detectors (MAX wi-D, MAX mini) is sporadic in image recovery. Pressing undo button on the RAD subtask card or previous series button on the Image subtask card during image readout. Sporadically, during an automatic or a manual RIS update. | Software design | A safety advisory notice, dated April 13, 2015, was sent to end users with affected devices to alert them about the issue and associated risks,and to provide instructions for avoiding potential risk. A new software version will be released to permanently resolve the issue. |  | 133 | 133 total | Nationwide Distribution. | Terminated | March 20, 2017 | 707 | 133 | Not_Computer
Z-1664-2015 | 71256 | joint, ankle, external brace | ITW | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | Elastic Ankle Support. |  |  | April 13, 2015 | 2015 | United Surgical Associates | The product does not bear the caution label regarding latex content. Potential for Allergic skin reaction | Error in labeling | On 4/13/2015 United Ortho emailed the affected distributor regarding the need for correction due to latex in the product and that updated labeling and instructions would be provided. United-Ortho personnel will travel to Hawaii and accompany the distributor to label the product at the sites to which it was distributed. Records will be maintained at each location. |  | 1278 | 1278 | Domestic: HIInternational; None | Terminated | July 14, 2015 | 92 | 3294 | Not_Computer
Z-1677-2015 | 71059 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | Alere i Influenza A &amp; B |  |  | April 22, 2015 | 2015 | Alere Scarborough, Inc. dba Binax, Inc. | High invalid rate and an increased risk of false negative results with certain lots of the Alere i Influenza A &amp; B kit. | Under Investigation by firm | Alere sent an Urgent Medical Device Recall letter dated April 23, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to stop using the product and discard any unused product. Customers asked to complete and return the Verification form within 10 days to confirm receipt of the recall notification and to request credit.For questions about the information contained in the notification, contact Alere Technical Services, Monday through Friday, 8am to 8pm EST. Telephone: (855) 731-2288 or E-mail: t s.scr@alere.com.For questions regarding the status of your credit, please contact your AI ere i distributor or E-mail: ssc.alerefieldactions@alere.com. |  | 216 | 213 kits plus 3 - evaluation use only kits | Worldwide Distribution - US Nationwide in the states of AL CA CT FL IA ID IL IN KY LA MD MO MS NC NE NJ OH PA SC TN TX and the countries of: Hong Kong and Singapore. | Terminated | January 14, 2016 | 267 | 216 | Not_Computer
Z-1674-2015 | 71067 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | EXPEDIUM Spine System/Synapse System; MODEL: EXPEDIUM Offset Conical Lock Nut 8mm, CATALOG NO. : 1754-91-150, UDI NUMBER (if applicable): GTIN: 10705034032467 |  |  | April 21, 2015 | 2015 | DePuy Spine, Inc. | Device does not mate as intended. | Process design | Consignees were notified of the recall via letter sent by overnight shipment for delivery on April 23, 2015. The letter instructs distributors and customers to review their inventory and return any products with the listed lot number. |  | 183 | 183 devices | US: PA and OUS: SE, BE, IT, PT | Completed |  |  | 183 | Not_Computer
Z-1678-2015 | 71110 | No | BSY | Tracheobronchial suction catheter. | Anesthesiology | Anesthesiology | 510(K) Exempt | KimVent Turbo-Cleaning Closed Suction System for Adults, 10 F T-Piece (Product Code 227105); KimVent Turbo-Cleaning Closed Suction System for Adults, 12 F, DSE (Product Code 2271603); KimVent Turbo-Cleaning Closed Suction System for Adults, 12 F, T-Piece (Product Code 227135). |  |  | April 14, 2015 | 2015 | Halyard Health, Inc | The internal blue flapper valve of certain KimVent Turbo-Cleaning Closed Suction System for Adults may have a potential defect. If present, the defect potentially allows air to leak out of the closed ventilator circuit. | Device Design | Halyard Health Regulatory Affairs sent an Urgent Voluntary Product Recall letter dated April 16, 2015, to all affected consignees. Consignees were instructed to quarantine and discontinue use of the recalled product. Specific return instructions were provided in the Notice. The consignees were also requested to provide a response to to the recall notification. Wholesalers were instructed to follow their internal recall procedures to identify and address affected products. Consignees with questions were instructed to call Halyard Health Regulatory Affairs at 470-448-5700. |  | 220 | 11 cases each containing 20 individual endotracheal tubes | Nationwide Distribution including CA, MA, MI, NC, TX, VA , and WA. | Terminated | July 07, 2015 | 84 | 220 | Not_Computer
Z-1672-2015 | 71085 | No | NJL |  |  | Orthopedic | PMA | LCS COMPLETE RPS Femoral Implants, various sizes and orientation. Knee prosthesis component for orthopedic surgery. |  |  | March 04, 2015 | 2015 | DePuy Orthopaedics, Inc. | The LCS COMPLETE RPS Knee System has higher rates of revision in Australia when the native patella is not resurfaced when compared to other unresurfaced PS implants. If the LCS COMPLETE RPS Knee System is implanted and the native patella is not resurfaced, the patient may experience patello-femoral pain, potentially requiring a secondary procedure. | Unknown/Undetermined by firm | On 3/4/2015, URGENT INFORMATION - DEVICE CORRECTION NOTICE for the LCS COMPLETE RPS Knee System and on 3/19/2015 REVISED URGENT INFORMATION - DEVICE CORRECTION NOTICE for the LCS COMPLETE RPS Knee System (Revised) notifications were sent to the affected distributors and medical professionals. All distributors were notified via electronic mail, and all medical professionals with affected inventory were also notified by distributor or Sales Consultant via in person delivery or mail or electronic mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, please contact Kim Earle, Recall Coordinator, at 574-371-4917 (M-F; 8 a.m. to 5 p.m. EDT). |  | 19059 | 19,059 | Worldwide Distribution-US (nationwide) and the countries of CANADA, AUSTRALIA, AUSTRIA, BELGIUM, UK, FRANCE, GREECE, IRELAND, ISRAEL, ITALY, KOREA, LUXEMBOURG, MALAYSIA, NETHERLANDS, NEW ZEALAND, NORWAY, SOUTH AFRICA, SWITZERLAND, THAILAND. | Terminated | October 30, 2015 | 240 | 47791 | Not_Computer
Z-1676-2015 | 71229 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | D2-40 (Lymphatic Marker); For In Vitro Diagnostic; |  |  | May 06, 2015 | 2015 | Biocare Medical, LLC | Firm discovered that incorrect Instructions for Use were shipped with one lot of D2-40. | Labeling mix-ups | Biocare Medical sent a Medical Device Notification letter dated May 6, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete all actions listed on the enclosed Return Response Form. |  | 19 | 19 units | Worldwide Distribution - US Nationwide in the states of: MI, SC, NC, AZ, CA, MO, WI and Hungary, Spain and Austria. | Terminated | June 04, 2015 | 29 | 19 | Not_Computer
Z-1675-2015 | 70959 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panorama Patient Monitoring Network, Multi- Parameter Patient Monitor (with Arrhythmia Detection and Alarm. Part number 0998-00-0708-01. |  |  | March 12, 2015 | 2015 | Mindray DS USA, Inc. dba Mindray North America | Software anomaly in the Panorama System software version 8.9 that manifests when the Panorama is in use with Mindray Passport V Monitor. Arrhythmia alarms which were previously displayed on both the Passport V and Panorama will not be displayed on the Panorama following a restart of either system (due to communication loss, restart, etc.). | Software design | Mindray sent a Corrective Action letter dated March 12, 2015 to their affected customer via FedEx. The letter identified the affected product, problem and actions to be taken. For questions call 201-995-8407. |  | 4 | 4 units | Worldwide Distribution - US Nationwide in the state of TX. | Open, Classified |  |  | 4 | Not_Computer
Z-1609-2015 | 70920 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CareFusion AVEA ventilator, all models. |  |  | April 03, 2015 | 2015 | Carefusion 211 Inc dba Carefusion | A pressure transducer failure can develop, activating a false Extended High Ppeak or Circuit Occlusion Alarm. The Safety Valve will open to ambient air and the unit will stop ventilating, allowing spontaneous breathing patients to breathe. The malfunction may delay initiation or cease ventilation. Nonbreathing patients will need manual ventilation or to be connected to another ventilator. | Process control | A customer letter dated 4/21/15 will be sent to all customers who received the AVEA Ventilator, all models, to inform them that CareFusion has identified a potential risk associated with AVEA Ventilator caused by a malfunction of a pressure transducer. The malfunctioning pressure transducer is detected and by design ceases ventilation and opens the safety valve to atmosphere allowing spontaneous breathing of patients capable of doing so. The non-breathing patient would need to be manually ventilated.The letter informs the customers of the issues, the potential risk, and the actions to be taken. Customers are instructed to contact CareFusion Recall Support Center at (888) 562-6018, (858) 617-5300, or email supportcenter@carefusion.com for recall related questions.Customers are instructed to contact CareFusion Technical Support/Customer Advocacy at (800) 231-2466, (714) 283-2228, or email support.vent.us@carefusion.com for Product Technical Support Adverse Event Reporting.CareFusion posted a press release on their website on 5/27/15 and it was sent to PRNewswire. The press release provided an update on a global voluntary recall that was initiated on April 21, 2015 to address an issue with certain units of AVEA ventilators. The press release provide the problems and the actions to be taken. The press also list the countries that are affected by the recall. |  | 15905 | 15,905 units total (5,670 units in US) | Worldwide Distribution. US Nationwide including Guam and Puerto Rico, Canada, Albania, American Samoa, Argentina, Australia, Austria, Bahamas, Bahrain, Bangladesh, Barbados, Belgium, Belize, Bermuda, Bolivia, Bosnia Herzegovina, Brazil, Bulgaria, Chile, China (Peoples Rep), Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Dominican Republic, Ecuador, Egypt, El Salvadore, Estonia, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea (Rep), Kuwait, Latvia, Lebanon, Libyan Jamahiriya, Lithuania, Malaysia, Mexico, Moldova, Morocco, Nepal, Netherlands, Nicaragua, Norway, Occupied Palestinian Territory, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Rwanda, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, South Africa, South America, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and Yemen. | Open, Classified |  |  | 15905 | Not_Computer
Z-1671-2015 | 71185 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EliteCore Full Core Biopsy Device. EliteCore 18G, 20cm with and without HiLiter tip, 15 cm with and without HiLiter tip, and 10 cm with and without HiLiter tip. |  |  | April 24, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Potential for the device cannula to overthrow past the intended length. | Device Design | Stryker sent an Urgent Medical Device Recall Notification letter dated May 5, 2015, to all affected customers via Fed Ex overnight. Customers were instructed to qaurantine any product found, return the repsonse form even if there is no product on hand, forward the letter if further distribution of recalled product occurred, return recalled product to Stryker. Customers were instructed to 269-389-2458 with any questions. |  | 1420 | 1,420 Units | US Nationwide Distribution including AL, AR, AZ, CA, CO, FL, GA, IL, IN, KS, KY, LA, MI, MN, MO, MS, NJ, NM, NV, NY, OK, OR, PA, TN, TX, UT, VA, and WI. | Terminated | October 15, 2015 | 174 | 1420 | Not_Computer
Z-1679-2015 | 71289 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RaySearch RayStation 2.5, 3.0, 3.5 and 4.0; Radiation Therapy Treatment |  |  | May 08, 2015 | 2015 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories AB became aware of an issue with the dose calculation for some imported VMAT plans where controls points are defined with wide gantry angle spacing. | Software Design Change | RaySearch sent a Device Correction letter dated May 8, 2015 via e-mailed to their customers. The letter identified the affected product, problem, actions to be taken by users and contact information. |  | 539 | Domestic: 539 units | US Nationwide Distribution | Terminated | April 11, 2017 | 704 | 539 | Not_Computer
Z-1683-2015 | 70672 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Catalys Precision Laser System-Catalys - U - US product; Catalys-1 - International product; software version 3.00.05 |  |  | February 20, 2015 | 2015 | Optimedica Corporation | Software anomaly on OptiMedica CATALYS System version 3.00.05 that may result in incorrect parameters provided for cataract incision surgeon templates and both eyes have the same parameters. | Software design | Abbott sent an Advisory Notice dated March 5, 2015, vis Fed Ex overnight to all affected customers. Out of State accounts will be notified through their distributor after letters are translated into local language.Recall notification letters advise customers of issue with CATALYS System using software version 3.00.05. The letter actually covers two issues, the cataract templates being the same for both eyes, and a loss of suction during treatment that may result in scoring the cornea.In both cases the issue is explained and users are advised to use mitigation listed to prevent injury. AMO is developing an update to software that will correct these issues and users will be advised when it is available for installation.All users of the device should be made aware of the issues and each surgeon is requested to sign the acknowledgement form to verify their receipt and understanding. The acknowledgement for should be returned to Abbott Medical Optics per instructions. Questions in the US may be directed to AMO Technical support at 1-855-662-0911. |  | 171 | 171 distributed Worldwide | Worldwide Distribution - US Nationwide, Germany, Switzerland, Australia, Spain, Japan, Great Britain, France, Netherlands, Austria, Singapore, South Korea, India, Brazil, Italy, Belgium, Egypt, Colombia, Hong Kong, and Canada. | Terminated | December 08, 2015 | 291 | 171 | Computer
Z-1689-2015 | 70681 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Quadrox-iD Adult Diffusion Membrane Oxygenator |  |  | March 05, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | A potential issue was detected that could result in blood leakage from the gas outlet port in all lots of Quadrox-iD Adult Diffusion Membrane Oxygenator with Bioline and Softline Coatings, Custom Tubing Packs and Diffusion Membrane HLS Set Advanced and HIT Set Advanced Models. | Nonconforming Material/Component | Maquet sent a "Field Safety Notice" accompanied by a Response Form to their affected customers on 3/5/2015 via FedEx. The Field Safety Notice identified the affected products/Part numbers and actions to be taken by the user. The customers were asked to complete the Field Safety Notice Response Form to acknowledge that they have received the notice. The completed notice can be faxed to 1-973-396-3607 or sent via email to fieldactions@maquet.com. For any questions, customers can contact their Maquet Sales Representative or Maquet Customer Service at 1-888-627-8383 (press option 2, followed by option 2), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST. |  | 10199 | 10,199 units | Worldwide Distribution - US Nationwide and the countries of: Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Great Britain, Greece, Hungary, Iceland, Ireland, Israel, Italy, Japan, South Korea, Latvia, Lithuania, Luxembourg, Malta, Mexico, Netherlands, Norway, Poland, Portugal, Romania, Saudi Arabia, Serbia Montenegro, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey and Uruguay | Open, Classified |  |  | 10199 | Not_Computer
Z-1691-2015 | 70772 | Yes | OAE |  |  | Cardiovascular | PMA | SMARTABLATE RF Generator System, Catalog No. M490006, M490007, M4900106, and M4900107. |  |  | March 19, 2015 | 2015 | Biosense Webster, Inc. | Biosense Webster is recalling the SMARTABLATE Generator System because Biosense Webster has received two complaints reporting that during an RMT procedure the SMARTABLATE RF Generator inadvertently delivered RF energy. | Device Design | Biosense Webster sent an Urgent Field Safety Notification Letter dated March 18, 2015 via FedEx to all their customers that received the affected SMARTABLATE Generator System. The letter informs the customers of the problems identified and the actions to be taken. Biosense Webster is taking precautionary safety measures of reinforcing the Warnings in the Instructions for Use (IFU) for the SMARTABLATE System Foot Pedal. Customers with questions related to the are instructed to contact their local Biosense Webster sales representative or call (866) 473-7823, Monday through Friday from 7:00AM to 8:00PM EST. Customers with questions related to the voluntary field removal certification form and its return, are instructed to contact Elie Ghattas Field Action Coordinator at (909) 839-7281. |  | 115 | 115 units | Worldwide Distribution - US Nationwide in the state of: CA, IL, KY, NJ, PA and the countries of: Czech Republic, Denmark, Finland, France, Germany, and Sweden. | Terminated | June 17, 2015 | 90 | 115 | Not_Computer
Z-1681-2015 | 71047 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II Implantable Infusion Pumps, models 8637-20 and 8637-40. Sterile. |  |  | April 10, 2015 | 2015 | Medtronic Neuromodulation | Medtronic is conducting a recall of a specific subset of Model 8637-20 and 8637-40 SynchroMed II implantable drug pumps because the audible alarm could cease to function. | Process control | Customers were sent a Medtronic "Urgent Medical Device Removal" letter dated April 2015. The letter described the problem and the product involved in the recall. The letter described the "Required Actions" which included to quarantine and return the product, complete and return the Customer Confirmation Form, and to share the notification with other customers. For questions, they can contact Medtronic Neuromodulation Technical Services at 1-800-707-0933.A second letter was sent to Healthcare Professional. The "Urgent Medical Device Correction" letter dated April 2015 described the Nature of the Device Issue, Scope and Likelihood of the Issue, Recommendations, Patient Management Recommendations and Additional Information.Advise consignees to report any malfunction or adverse event related to a device to Medtronic Neuromodulation Technical Services, at 1-800-707-0933 available weekdays from 7 am - 6pm Central Time, and to FDA'ss MedWatch Program (www.fda.gov/medwatch). |  | 1901 | 1901 (1576 US, 325 OUS) | Worldwide Distribution - US Nationwide including DC AND PR and in the countries of : Argentina, Austria, Belgium, France, Germany, Ireland, Italy, Japan, Luxembourg, Malta, Netherlands, Norway, Panama, Poland, Portugal, Russian Federation, San Marino, Saudi Arabia, South Korea, Spain, Switzerland, and United Kingdom. | Terminated | November 23, 2015 | 227 | 1901 | Not_Computer
Z-1680-2015 | 71044 | No | CCL | Oxygen gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | D-fend, (p/n 876446-HEL), and D-fend+ (p/n 881319-HEL) water trap. Used with Compact Airway Modules E/M-C(Ai)O(V)(X) and Cardiocap/5 monitors. |  |  | April 24, 2015 | 2015 | GE Healthcare | A potential safety issue due to loss of gas monitoring associated with occlusion of specific D-Fend / D-Fend+ water traps used with Compact Airway Modules E/M-C(Ai)O(V)(X) and Cardiocap/5 monitors. This water trap occlusion could cause a loss of respiratory airway gas monitoring. | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref#36111 dated April 24, 2015. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, &amp; Director of Biomedical Engineering, The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. A confirmation form is asked to be returned and GE will replace all affected product. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 2948 | 5,018 (US = 2,948 / OUS = 2,070) | Worldwide Distribution - US Nationwide including Washington DC, Puerto Rico. and the countries of: Algeria, Argentina, Australia, Austria, Bahrain, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic , Egypt, Estonia, Finland, France, French Polynesia, Germany, Greece, Guadeloupe, Guernsey , Guernsey , Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jersey, Jordan, Korea (Republic of), Kuwait, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand , Turkey, Ukraine, United Arab Emirates, United Kingdom, Vietnam, | Open, Classified |  |  | 2948 | Not_Computer
Z-1684-2015 | 70659 | No | LKK |  |  | General Hospital | PMA | CODMAN 3000 Series Refill Set- Huber needles. |  |  | March 05, 2015 | 2015 | Codman &amp; Shurtleff, Inc. | Needle is leaking from hub/needle cannula connection | Process control | Codman Neuro issued recall letter on 3/05/15 requested to immediately check all inventory to determine if you have affected product. Complete the Acknowledgement Form and follow the instructions to report your inventory status and return affected product for credit and replacement (if requested). General inquiries, customer service or field action logistics - please contact Codman Neuro Quality at: CodmanProductComplaints@DPYUS.JNJ.com or 866-491-0974. |  | 5463 | 5463 units | Worldwide Distribution-US (nationwide) and the countries of Canada, France, Germany, and Switzerland. | Open, Classified |  |  | 5463 | Not_Computer
Z-1687-2015 | 67989 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | 100/120V AC, Advanced Perfusion System Platform (APS) The TerumoÂż Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. |  |  | May 01, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Terumo CVS is implementing field correction activities to address the identified causes of the Terumo System 1 experiencing loss of system power. These activities will be implemented in a phased approach beginning late 2015:- Implementing design changes and upgrades to improve the reliability of power switching and battery backup.- Adding a switch protector to the On/Off power switch to preven | Device Design | On 5/1/2015 Terumo issued an URGENT MEDICAL DEVICE CORRECTION letter to users explaining that Terumo has identified causes of the Terumo System 1 experiencing loss of system power. These activities will be implemented in a phased approach beginning late 2015: - Implementing design changes and upgrades to improve the reliability of power switching and battery backup.- Adding a switch protector to the On/Off power switch to prevent inadvertent shut off.- Replacing the On/Off power switch with a new design to prevent switch failure. A Terumo Field Service Representative will contact users to schedule the field correction activities once the activities are ready to be implemented.CUSTOMER INSTRUCTIONS: Review this Medical Device Correction and assure that all users have received notice of this issue.Âż- Complete and return the enclosed Customer Response Form.Âż- Terumo CVS will contact users to schedule the field correction activities.We encourage you to contact us with any questions or concerns:Terumo CVS Customer Service1-800-521-2818Fax 1-734-741-6149Customer Service Hours:Monday  Friday, 8 a.m.  6 p.m. ET |  | 1713 | 1713 units total | Worldwide Distribution-US (nationwide) including DC and the states of AL , AR , AZ , CA , CO , CT , DE , FL , GA , HI , IA , ID , IL , IN , KY , LA , MA , MD , ME , MI , MN , MO , MS , MT , NC , ND , NE , NJ , NM , NV , NY , OH , OK , OR , PA , SC , TN , TX , UT , VA , VT , WA , WI , and WV, and the countries of Mexico , AUSTRALIA , UNITED ARAB EMIRATES (UAE) , Hong Kong , Indonesia , Singapore , Taiwan , Thailand , Philippines , South Korea , CHILE , Vietnam , Malaysia , BELGIUM , Japan , and CANADA. | Open, Classified |  |  | 1713 | Not_Computer
Z-1685-2015 | 71014 | No | HJK | Polymethylmethacrylate (PMMA) diagnostic contact lens. | Ophthalmic | Ophthalmic | 510(k) | Volk 25D Large Clear BIO Lens with case UPC: 8 11608 01010 6, VOLK OPTICAL INC. The Volk Optical BIO Lenses are intended for visualization during diagnosis and laser therapy of the human retina (fundus) using a binocular indirect ophthalmoscope. |  |  | February 27, 2015 | 2015 | Volk Optical Inc | Mislabeling. The incorrect lot number was engraved on the anodized aluminum ring. | Process control | Volk Optical, Inc, sent an Urgent Medical Device Recall letter dated March 30, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The Firm requests that you return the device(s) for replacement, at our cost. The recalled device(s) should be returned to Volk with an assigned Return Merchandise Authorization (RMA) number. Please contact Volk as soon as you receive this notification to setup a RMA for your device(s). Please contact Volk at (440) 942-6161 or toll-free at (800) 345-8655 to speak with a Volk RMA Coordinator. To return a recalled device(s) to Volk, please ask Volks RMA Coordinator for the appropriate Volk shipping account number (e.g., FedEx, UPS, etc.). |  | 6 | 6 units | Worldwide Distribution - US Distribution to the states of : MA and VA., and to the countries of Mexico and Sweden. | Terminated | February 19, 2016 | 357 | 6 | Not_Computer
Z-1682-2015 | 71228 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF Positron Emission Tomography (PET) and Magnetic Resonance Imaging (MRI), 882380, Philips Medical Systems. |  |  | March 11, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | It was noticed that the spectroscopy voxel map did not align with the anatomical images of the phantom. This resulted in incorrect orientation of the Chemical Shift Imaging (CSI) data. | Software in the Use Environment | On 3/11/2015 the firm sent Urgent - Medical Device Correction letters to their customers. The letters identified the affected device, explained the reason for the recall, how to identify affected product, advice on actions to be taken by customer/user, and any actions planned by the firm. Customers were advised to use software provided by Philips to view spectroscopy images; and, not to export and view images using third party software. The firm stated that it will be issuing a software update. If further information or support concerning the issue is needed, customers are asked to contact their local Philips representative or local Philips Healthcare office. |  | 10 | 10 units | Worldwide Distribution -- US, including the states of NY, OH, &amp; TX; and, the countries of Finland, Germany, Japan, Netherlands, South Korea, Spain, &amp; Switzerland. | Open, Classified |  |  | 10 | Not_Computer
Z-1686-2015 | 71296 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRADÂż Stellant CT Injector System with Certegra |  |  | September 20, 2013 | 2013 | Medrad Inc dba Bayer R&I; | The root cause investigation showed that the under-volume hazard can occur when the unit has been configured to prime with contrast, but instead primes with saline due to the configuration data not being sent to the injector head. This may create a non-optimal priming condition that could lead to under-infusion. | Software design | Bayer Healthcare sent a Medical Device Field Corrective Action letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify any affected device. Bayer Service will schedule time to update their software free of charge. Customers with questions were instructed to contact the customer support team at 1-800-633-7237. For questions regarding this recall call 412-767-2400. |  | 133 | 133 | Nationwide Distribution | Terminated | June 16, 2015 | 634 | 133 | Not_Computer
Z-1692-2015 | 70793 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Smiths Medical, BCI Advisor Vital Signs Monitor, Product Reorder No 92M774325A. |  |  | March 21, 2015 | 2015 | Smiths Medical ASD, Inc. | The Auxiliary Serial Input/ Output Modes on the Advisor do not function as expected, resulting in data formatting issues orincorrect data being transmitted in or out of the serial port into a data collection system. | Nonconforming Material/Component | Consignee was sent on 3/20/2015 a Smiths Medical "Urgent Medical Device Recall" letter dated 13 March 2015. The letter was addressed to Supplier Quality. The letter described the issue and the product affected by the recall. Requested consignees to quarantine the affected product and to complete and return the Confirmation Form by Fax to 1-800-237-8033 or by email to recall.response@smiths-medical.com within 10 days of receipt of this letter. Customers shall report any issues with these products to Smiths Medical's Global Complaint Department at 1-866-216-8806 or globalcompaints@smiths-medical.com. For questions they can contact Smiths Medical's Customer Service Department at 1-800-258-5361. UPDATE 4-20-2015 Smiths Medical sent an Urgent Medical Device Recall letter via email to the one consignee on April 10, 2015. The letter identified the additional monitors and requested that those devices be quarantined, return the attached confirmation form, and a customer service representative will contact the consignee to arrange for exchange of the affected Advisor monitors for credit or replacement. |  | 4202030 | 5 UPDATE 4-20-2015 to include 10 additional devices | MN | Terminated | November 20, 2015 | 244 | 4202030 | Not_Computer
Z-1637-2015 | 71191 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Synovis VASCU-GUARD Peripheral Vascular Patch in the following sizes and product codes: 1x6cm, product code 1504026, 0.8x8cm, product code 1504028, 1x10cm, product code 1504030, and 2x9cm, product code 1504032. VASCU-GUARD is packaged between two pieces of foam within a double sterile barrier pouch system. The contents of the unopened, undamaged package are sterile. VASCU-GUARD is intended for use in peripheral vascular reconstruction including the carotid, renal, iliac, femoral, profunda, and tibial blood vessels and arteriovenous access revisions. |  |  | May 02, 2015 | 2015 | Synovis Surgical Innovations, Inc. | Baxter healthcare is recalling specific product codes of Vascu-Guard Peripheral Vascular Patch due to complaints received for difficulty in distinguishing the smooth from rough surface. | Under Investigation by firm | The firm, Baxter, sent an "Urgent Product Recall" letter dated 5/2/2015 via USPS overnight delivery to its customers. The letter identified the affected product, problem and actions to be taken. The customers were instructed to immediately discontinue use and segregate the affected product, locate and remove all affected product from your facility and to return it to Baxter. Customers are to contact Baxter Healthcare Center, 888-229-0001, for Service to arrange for return and credit. In addition, the customers were to complete and return the Customer Reply Form by faxing to 224-270-5457 or emailing to fca@baxter.com.If you have any questions, contact Director, Quality at 651-796-7543 or email: heidi_drafall@baxter.com. |  | 3974 | 3,974 | Nationwide Distribution. | Terminated | February 18, 2016 | 292 | 3974 | Not_Computer
Z-1695-2015 | 71231 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-ray system |  |  | April 13, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | A calibrated torque wrench malfunctioned at the supplier leading to possible incorrect torque of x-ray tube mounting fasteners. | Other | Philips sent an Urgent - Field Safety Notice dated April 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips Healthcare is recommending to stop the use of your system until the X-ray tube mounting fasteners are replaced. In addition, a Philips Field Service Engineer (FSE) will service your system via aMandatory Field Change Order (FCO 72800638). A Philips FSE will contact youto schedule this appointment.If you need any further information or support concerning this issue, please contact your local Philips representative or your local Philips Healthcare office.This notice has been submitted to the appropriate Regulatory Agency.Philips apologizes for any inconveniences this problem has caused your organization. For further questions please call (440) 483-7600. |  | 2 | 2 | Internationally to the countries of China and Norway. | Terminated | October 01, 2015 | 171 | 2 | Not_Computer
Z-1694-2015 | 71243 | No | NDJ | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Cortical Screw for M/DN Intramedullary Fixation. Metaphyseal/Diaphyseal Nail System, Orthopedic surgical device. |  |  | April 29, 2015 | 2015 | Zimmer, Inc. | The entire scope of 4.5mm Cortical Screws listed as compatible with the M/DN system is not compatible with the M/DN system. This has the potential to lead to a situation in which the screw would bind upon entering the nail and have to be removed during the surgery or removal of the screw after fracture healing. | Component design/selection | Zimmer sent an Urgent Medical Device Recall letter dated April, 29, 2015, to all affected customers.On 4/29/2015. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to review the notification and ensure affected personnel are aware of the contents. Locate all affected product and quarantine them immediately. Return the recalled product along with completed Inventory Return Certification Form. Notify Zimmer of any further distribution of the affected product. In addition, identify the surgeons and hospitals that you have identified using the provided spreadsheet. For further questions call the customer call center at 1-877-946-2761. |  | 289616 | 289,616 | Worldwide Distribution - US (nationwide) and Internationally to Canada, Brazil, Mexico, Bahamas, Nicaragua, Singapore, China, Japan, Malaysia, Taiwan, Thailand, Korea, Spain, Belgium, Switzerland, India, Germany, England, Italy, Saudi Arabia, Lebanon, Libya, United Arab Emirates, Algeria, France, United Kingdom, Canary Islands, Iraq, Jordan, Lithuania, Morocco, Macedonia, Malta, Mauritius, Serbia, Russian Federation, and Yemen. | Terminated | December 08, 2015 | 223 | 289616 | Not_Computer
Z-1696-2015 | 71250 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PRI Femoral Impactor Block for knee prosthesis surgery. |  |  | April 30, 2015 | 2015 | Zimmer, Inc. | Potential for uncured adhesive between the metal threaded insert and the impactor block to leak out of the assembled part following the autoclave sterilization process. The Master Bond epoxy serves as a secondary locking mechanism to fix the metal insert to the impactor head. Foreign material has been observed on the impactor block. This has the potential for a biological response. | Process control | On 4/30/2015, URGENT MEDICAL DEVICE RECALL LOT SPECIFIC notifications dated April 30, 2015 were sent to the affected distributors, hospital Risk Managers, and surgeons via courier; all distributors were also notified via electronic mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-800-348-2759 between 8:00 am and 8:00 pm EST. |  | 1689 | 1,689 | Worldwide Distribution-US (nationwide) including PR and the states of AL, AZ, CA, CO, FL, IL, IN, KS, MA, MI, MN, MO, NC, NJ, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI, and the countries of Canada, Korea, Singapore, Australia, Japan, Malaysia, India, Austria, Belgium, Canary Islands, Czech Republic, France, Germany, Iceland, Iran, Israel, Italy, Norway, Russian Federation, Saudi Arabia, South America, Spain, Sweden, Switzerland, United Kingdom, and United Arab Emirates. | Terminated | March 15, 2016 | 320 | 1689 | Not_Computer
Z-1697-2015 | 71279 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | ActCel |  |  | May 18, 2015 | 2015 | Coreva Health Science LLC | Coreva Health Science, LLC. is recalling ActCel Hemostatic Dressing due to customer literature and web sites promoting that it can help control bleeding from open wounds and in body cavities. ActCel is indicated as a prescription device for management of topical wounds and to temporarily control external surface bleeding. | Labeling False and Misleading | On 05/18/15 the firm sent out customer notification letters to those distributors that were identified as marketing ActCel Hemostatic Dressing for use to help control bleeding from open wounds in body cavities.The firm's notification letter provided the following actions to be taken by the customer:- Review their literature and labeling (including internet websites) to determine whether there are any marketing statements for usage in body cavities; -Destroy all hardcopy literature and labeling with the statement for usage in body cavities - Correct all Internet websites eliminating his statement; and- Complete the attached Medical Device Labeling Clarification Return Response form and return it to Coreva Health Science.The firm states to direct any questions to: Ms. Casssie Inglis, Vice President of Marketing.email: cassie@coreva.netphone: 818-685-0951 ext. 119toll free: 800-808-9094 ext. 119Monday through Friday, 9am-4pm PST |  | 180900 | 180,900 | U.S. distribution to the following; TX, NY, FL, VA, WA, and AZ. No foreign distribution.NOTE: The distribution list the firm provided does not include all consignees that purchased the product; rather target the consignees (distributors) who were observed to be using an unapproved marketing use. | Terminated | August 26, 2015 | 100 | 180900 | Not_Computer
Z-1704-2015 | 71206 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | syngo Workflow SLR; The information system syngo Workflow SLR is a digital radiology information system (RIS) with integrated modules for patient administration, examination, reporting, statistics and system administration. The system electronically displays stores, retrieves, transfers, exchanges, and prints. |  |  | May 06, 2015 | 2015 | Siemens Medical Solutions USA, Inc | A potential exists for order transactions from interfaced HIS (Hospital Information System) systems to be lost during a restart of interface processes when using the Order Batching feature, leading to diagnosis delay. | Software design | Siemens sent an Safety Advisory Notice, dated May 6, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were given work around instructions to follow and informed that a software fix would be implemented to permanently resolve the issue.For further questions please call (610) 219-6300. |  | 131 | 131 | Nationwide Distribution | Open, Classified |  |  | 131 | Not_Computer
Z-1701-2015 | 71202 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Tibial Alignment Guide; Orthopedic Stereotaxic Instrument; |  |  | April 19, 2015 | 2015 | Zimmer CAS | Zimmer CAS has determined that the potential exists for the spikes on the iASSIST Tibial Alignment Guide to bend or break during insertion or extraction from the tibia during use. | Device Design | Zimmer distributed notices via certified mail on May 14, 2015. Zimmer is removing affected Tibial Alignment Guide iAssist Knee System product from distribution in a two-phased process. Based upon complaint investigation, Zimmer CAS has determined that the potential exists for the spikes on the iASSIST Tibial Alignment Guide to bend or break during insertion or extraction from the tibia during use. Customerswere asked to review the notifictaion, follow the iASSIST Surgical Knee Technique provided, further distribute the notice for any distributed product, inspect devices in use before and after procedures, report any adverse events, and complete the certificate of acknowledgment and return it to CorporateQuality.PostMarket@zimmer.com.Customers with questions concerning this notice, please contact Customer Service at the following address/phone number.Customer Service / Zimmer CAS75 Queen Street, Suite 3300Montreal, Quebec, Canada H3C 2N6Email : cas-support@zimmercas.com, Telephone : 1-514-395-8883, toll free for North America 1-866-336-7846, Fax : 1-866-978-3801.For questions regarding this recall 574-372-4487. |  | 78 | 78 devices | Nationwide Distribution including AL, AZ, CA, CA, CO, FL, IN, MI, NJ, NY, OH, OR, PA, TX, VA, and WI. | Open, Classified |  |  | 78 | Not_Computer
Z-1700-2015 | 71169 | No | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Swan-Ganz CCO/Sv02/CEDV/VIP Thermodilution Catheters model number 744HF75. Flow-directed pulmonary artery catheters for monitoring cardiac output, mixed venous oxygen saturation, and end diastolic volume when used with the Vigilance Monitor. |  |  | May 05, 2015 | 2015 | Edwards Lifesciences, LLC | The product failed to meet the internal established statistical test requirement for heparin activity. | Process control | On 05/05/2015 Edwards sent their consignees a letter including the following instructions: Once you have verified your inventory, please complete the attached acknowledgement form (evenif you have no inventory) and fax it to Edwards Customer Service at 1-800-422-9329 within three days of receipt of this notification. This is to confirm that you have reviewed this notice and have taken appropriate action. Please call Edwards Customer Service at 1-800-424-3278 to arrange for return of any unused product from the identified lot and obtain information about replacement product. Please call for an RGA prior to returning any unused product and return unused product to the following address:Edwards LifesciencesReturn Goods/RGA # XXXXXX1212 Alton Pkwy.,Irvine, CA 92606-4838 |  | 978 | 978 | Worldwide distribution US nationwide, Canada and Europe (IC, BE, IT, GB, FR, DK, CZ, DE, AT, ES, IE, NO, PL) | Open, Classified |  |  | 978 | Not_Computer
Z-1703-2015 | 71273 | No | GFO | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | Dade Actin Activated Cephaloplastin Reagent; Liquid rabbit brain cephalin with plasma activator for use in the determination of the activated partial thromboplastin time (APTT) and other coagulation procedures requiring an activated partial thromboplastin reagent. |  |  | March 25, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Recovery for assayed controls are out of assigned ranges. A drift &gt;3 seconds in the normal APTT range and/or &gt;15% in the pathological range has been observed. Patient samples with recoveries close to the medical decision points could exhibit a deviation up to 4 seconds for the normal range and up to 33% for the pathological range. | Under Investigation by firm | On March 24, 2015 an Urgent Medical Device Recall letter was mailed to all affected customers to notify them about the issue, associated risks, and actions to be taken. |  | 27501 | 27501 | Worldwide Distribution. US naionwide including Puerto Rico, Barbados and Mexico. | Open, Classified |  |  | 27501 | Not_Computer
Z-1705-2015 | 71314 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via and Syngo.x ; a software solution intended to be used for viewing, manipulation, communication, and storage of medical images. |  |  | May 15, 2015 | 2015 | Siemens Medical Solutions USA, Inc | measurements drawn on the 2nd and subsequent images of the series are not visible on printouts when the series is sent to print. | Software design | Siemens sent a Customer Advisory Notice, dated May 15, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to check the software version of their syngo.via system and consider whether the issue described in the advisory notice impact their clincal coutine and whether to arrange the installation of SY027/15/P (VA30A_HF07). Customers were advised to include the advisory notice in their operator's manual. Customers were also advised to notify anyone at their organization who should be aware of this information.For questions regarding this recall call 610-219-6300. |  | 167 | 167 | Nationwide Distribution | Terminated | March 24, 2017 | 679 | 167 | Not_Computer
Z-1708-2015 | 71096 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista Systems LOCI Thyroid Stimulating Hormone; the TSH method is an in vitro diagnostic test for the quantitative measurement of Thyroid Stimulating Hormone in human serum and plasma on the Dimension Vista System. |  |  | April 20, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Dimension Vista LOCI Thyroid Stimulating Hormone (TSH) Lot 14237AA may produce falsely depressed results ranging from -23% to -96% near the end of the well set due to inadequate reagent volume in a Flex well. | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Recall notice, dated April 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. - Review this letter with your Medical Doctor.- Discontinue use and discard your remaining inventory of the reagent lot listed in Table 1.- Indicate your replacement product needs on the attached Field Correction Effectiveness Check form and fax to (312) 275-7795 within 30 days. Siemens will replace any unused inventory of the affected lot at no charge.We apologize for the inconvenience this situation may cause.. If you have any questions, please contact your Siemens Customer Care Center Technical Solutions at 800-441-9258, or your local Siemens technical support representative. |  | 2256 | 2256 | Worldwide Distribution - Nationwide and to the countries of : Australia, Japan, New Zealand, Canada, Austria, Belgium, Switzerland, Germany, Denmark, Spain, Finland, France, Great Britain, Italy, Netherlands, Portugal, and Slovakia. | Open, Classified |  |  | 2256 | Not_Computer
Z-1709-2015 | 71175 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP, Flush Syringe, 10 mL in 12 mL, Sterile, Rx only, Code No. 2T0806 |  |  | April 27, 2015 | 2015 | MRP, LLC dba AMUSA | Incorrect expiration date on label | Process control | The firm, AMUSA, sent a "Voluntary Recall Notification" dated 4/24/2015 to their consignees via Fed Ex on 04/27/2015. The letter described the product, problem and actions to be taken. The consignees were instructed to return any product remaining in the distribution channel to AMUSA; check with your product representative to determine if you have any of the listed products; if any product is located, call (615) 833-2633 from 8:30 am to 5:00 pm (Eastern Time) to receive instructions as to how to return the product; follow the enclosed Steps for Voluntary Recall in their entirety, and complete and return the enclosed Product Tracking and Verification Form.Please contact your AMUSA customer service representative if you have any questions regarding this correction, any of our products, or would like assistance with the corrections. Customers can also contact the Director, Quality Assurance at 513-325-1940 or email qadirector@amusa.us. |  | 200000 | 200,000 units | US Distribution in states of:OH, PA, AL, IL, TN, and CT. | Terminated | May 02, 2016 | 371 | 200000 | Not_Computer
Z-1710-2015 | 71098 | No | LZS |  |  | Ophthalmic | PMA | EC-5000 Excimer Laser Corneal Surgery System (EC-5000, EC 5000 CXIII, EC-5000 Quest. |  |  | March 24, 2015 | 2015 | Nidek Inc | Multi-stage treatment option for the EC-5000 Operator's Manual was not approved by the FDA. | No Marketing Application | Nidek sent an Device Correction letter dated April 30th, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Nidek will contact customers to schedule an appointment to replace the operators manual and implementing a software lockout device on the device so that the multi stage feature is no longer available.If a Nidek rep does not contact you before April 30, 2015, please call 800-223-9044 to schedule an appointment, All non responding customers.will be contacted three time via Fed Ex letter |  | 114 | 114 units | Us distribution only. | Terminated | April 05, 2016 | 378 | 114 | Computer
Z-1707-2015 | 71176 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products Signal Reagent, Product Code 1072693, Unique Device Identifier No. 10758750008360 VITROS Immunodiagnostics Products Signal Reagent (SR) is supplied to users as a sales unit box containing two Signal Reagent Packs. Each Signal Reagent Pack is labeled with a barcode label. The label should include the name of the reagent, a barcode, lot information and expiry dating. |  |  | May 06, 2015 | 2015 | ORTHO-CLINICAL DIAGNOSTICS | The Immunodiagnostic Products Signal Reagent, Lot 3082 as the reagent packs may contain blank labels. | Error in labeling | On 5/06/2015, Customer Letter and Verification Form (Ref. CL15-095) were sent notifying the users of the issue and advised them of the required actions. The notification was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected product and advised them of the issue and the required actions. On 5/06/2015, Distributor Letter (Ref. DL15-095) was sent via US Postal Service Priority Mail and/or FedEx overnight courier to US Distributor consignees and instructed to notify their consignees of the issue and actions. For questions,questions, please contact Customer Technical Services at 1-800-421-3311 at any time. |  | 898 | Domestic: 898 units | Nationwide Distribution. | Open, Classified |  |  | 898 | Not_Computer
Z-1712-2015 | 71318 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncentra Brachy radiation therapy planning software designed to analyze and plan radiation treatments in three dimensions for the purpose of treating patients with cancer. |  |  | May 21, 2015 | 2015 | Elekta, Inc. | Incorrect dose calculation for Regions of Interest (ROIs) defined on a secondary image series. | Other | Important User Notice, 806-03-BTP-001, was distributed to customers on 5/21/2015, A software patch is in development and is expected to be released end of November, 2015. Elekta service will work with customers to implement the patch on affected devices within 6 months of the patch release. |  | 351 | 351 | AL, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, IN, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NC, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WV, WI, Argentina, Australia, Bangladesh, Belarus, Botswana, Brazil, Canada, Chile, China, Colombia, Ecuador, Egypt, Guatemala, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kosovo, Libya, Malaysia, Mexico, Morocco, Namibia, Nepal, New Zealand, Oman, Pakistan, Peru, Philippines, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Sri Lanka, Suriname, Taiwan, Thailand, Tobago, Ukraine, Venezuela, Vietnam, Zambia | Open, Classified |  |  | 351 | Not_Computer
Z-1711-2015 | 71277 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 3.2 mm Cannulated Drill Bit/ QC 170 mm |  |  | May 13, 2015 | 2015 | Synthes, Inc. | affected parts and lots of the Cannulated Drill Bits have the potential to break during use | Component design/selection | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated May 13, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed if they have any of the identified devices, please take the following steps:" Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number." Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided." Return the Verification Section (page 3 of this letter) with the product to:" Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132." Send a copy of the completed Verification Section to DePuy Synthes by:" Fax: (888) 943-4893 or" Scan/email: Synthes6230@stericycle.comIf you DO NOT have the identified product, please take the following steps:" Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information." Return the documents to DePuy Synthes by:" Fax: (888) 943-4893 or" Scan/email: Synthes6230@stericycle.com" Note: If the Verification Section is answered on behalf of more th An urgent medical device recall notice, dated May 13, 2015, was sent to consignees and sales consultants to in form them of the issue and associated risks. All in receipt of the device were asked to discontinue use and return affected devices along with the response form.Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales |  | 1495 | 1495 US (37 OUS) | Worldwide Distribution - US (nationwide) and Canada. | Open, Classified |  |  | 1495 | Not_Computer
Z-1715-2015 | 71209 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVOTECH Endoscope Cleaner and Reprocessing System 220-230V, P/N 50004-002. |  |  | May 11, 2015 | 2015 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the EVOTECH Endoscope Cleaner and Reprocessor (ECR) because it may not detect an improper endoscope connection on two channels, and subsequently pass the cycle without cancellation if there is a disconnected scope. | Nonconforming Material/Component | A customer notification letter dated 5/7/15 will be sent to all customers on 5/13/15 to inform them that ASP is recalling the EVOTECH Endoscope Cleaner and Reprocessor (ECR) because it may not detect an improper endoscope connection on two channels and subsequently pass the cycle without cancellation if there is a disconnected scope. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the attached Business Reply Card (double-sided) via mail or by fax to Stericycle: 888-879-8144. Customers with questions regarding the business reply card are instructed to contact Stericycle at 888-345-4671 and reference event #7781. |  | 44 | 44 units | Worldwide Distribution-US (nationwide) and the countries of Australia, Singapore, Canada, and Thailand. | Open, Classified |  |  | 869 | Not_Computer
Z-1718-2015 | 71058 | No | ITH | Power traction equipment. | Physical Medicine | Physical Medicine | 510(k) | Extentrac Elite User Manual DCN: 14-001 (Powered decompression table, for traction) |  |  | May 13, 2015 | 2015 | Advanced Back Technologies, Inc. | Item is marketed for use in patients weighing up to 350 lbs, but the device was only approved for use in patients up to 300 lbs | Error in labeling | Advanced Back Technologies, Inc. sent an Urgent Medical Device Brocure &amp; User Manual Recall letter dated May 13, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed the discard the DCN's identified as DCN 13-001, DCN 09-001, DCN 14-001 and replace with the newly issued DCN 15-001, DCN 15-002. The new DCN's would be provided to customers in pdf format via email for issuance. Customers were asked to contact their customers who purchased the affected product and inform them that they should not treat patients with a maximum weight exceeding 300 lbs until the firm is able to make an amendment to their 510(k). Customers with questions should call 631-231-0076. |  | 18 | 18 units | Worldwide Distribution - US including CA, FL, NJ, TX, WA, Internationally: Korea, Russia, Kuwait, Canada, Greece, Lebanon, Turkey, Egypt, Saudi Arabia, China, Armenia, Azerbaijan, Australia. | Open, Classified |  |  | 54 | Not_Computer
Z-1713-2015 | 71267 | No | GJT | Factor deficiency test. | Hematology | Hematology | 510(k) | Factor deficient plasma for the Siemens BCSÂż / BCSÂż XP System. |  |  | May 11, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Discrepancies for Siemens BCS / BCS XP with dilution studies; erroneously high Factor IX results at the lower end of the measuring range. | Under Investigation by firm | An Urgent Medical Device Recall notice, dated May 2015, was sent to affected customers to inform them about the issue and associated risks. Customers were asked to discontinue use, distribute this information to all appropriate personnel, and to complete and fax back to Siemens the Effectiveness Check form. |  | 3575 | 3,575 | Worldwide Distribution -- Austria, Bulgaria, Belgium, Switzerland, Cyprus, Czech Republic, Germany, Denmark, Italy, Latvia, Netherlands, Poland, Portugal, Romania, Sweden, Slovenia, Slovakia, Turkey, Croatia, Bahrain, Belarus, Bosnia-Herzegovina, Guadeloupe, Kuwait, Lebanon, Libya, Macedonia, Montenegro, Qatar, Russia, San Marino, Saudi Arabia, Serbia, South Africa, United Arab Emirates, Argentina, Brazil, Chile, El Salvador, Peru, China, Mexico, Canada and United States. | Open, Classified |  |  | 3575 | Not_Computer
Z-1719-2015 | 71295 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Laksell GammaPlan, a computer based dose planning system specifically designed for use with Leksell Gamma Knife, radiation therapy treatment. |  |  | June 01, 2015 | 2015 | Elekta, Inc. | Memory can become corrupted when creating a fused study via drag and drop in Leksell GammaPlan 10.2. | Software in the Use Environment | The firm, Elekta, sent an "URGENT IMPORTANT FIELD SAFETY NOTICE 100-01-102-015" to all customers on 6/1/2015. The notice described the product, problem and actions to be taken. The customers were instructed to not use drag-and-drop to fuse images as a solution until a permanent fix can be developed and implemented on all affected devices. A solution is in development that will resolve the issue and is expected to be released September 2015. Elekta Service will then work with customers to implement on affected devices - estimated 6 month completion. Customers were also instructed to complete and return the IMPORTANT FIELD SAFETY NOTICE ACKNOWLEDGEMENT form to your local Elekta Office or Representative as soon as possible and within 30 days at the latest. If you have any queries about this Notice, please contact your local Elekta representative or the Global Post market Surveillance Manager at 770-300-9725 or email: Elekta_Global_PMS@elekta.com. |  | 10 | 10 | Worldwide Distribution: US (nationwide) including states: IL, OH, PA, TX, and WI; and countries of: Germany, Japan and United Kingdom. | Terminated | January 12, 2016 | 225 | 10 | Not_Computer
Z-1716-2015 | 71093 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | CRIT-LINE Blood Chamber- an accessory to the Crit line III- Monitors in real time blood parameters |  |  | April 27, 2015 | 2015 | Fresenius Medical Care Holdings, Inc. | Blood chamber connection leaks during the use of the CRIT-LINEÂż Blood Chamber | Under Investigation by firm | Fresenius LLC North America notified accounts by letter sent certified mail on May 22, 2015, expanding the recall to All lots. Previously recall letters were on April 30, 2015, 2nd notice on May 14, 2015. Users are requested to examine inventory, discontinue use and return product. Contact FMCRTG, LLC Customer Service Team 1-800-323-5188 for instructions. A reply form is included as part of the customer communication with instructions to fill out and return via fax or scanned/emailed copy.Confirmation of delivery will be obtained through the use of registered mail. |  | 22640288 | 22,640,288 | Worldwide Distribution - Nationwide Distribution including Puerto Rico., and the countries of : Ireland, Spain, Slovenia, Great Britain, Netherlands, Norway, Mexico, Egypt, Czech Republic and Spain. | Open, Classified |  |  | 22640288 | Not_Computer
Z-1720-2015 | 71112 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Roscoe Mini Neb Compressor User Manual for the NEB-ROS |  |  | April 20, 2015 | 2015 | Compass Health Brands | Roscoe Medical has recently identified the need to update its user manual with additional warning statements regarding proper handling and storage of the nebulizer. These warnings are to better inform end users of proper care when using and storing a Roscoe Medical Nebulizer to reduce user damage (wear &amp; tear) to the power cord. | Use error | Compass Health sent an Urgent Recall Field Correction letters dated April 20, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to: Âż Ensure all affected personnel are fully informed of this notice.Âż Forward this notice to your Regulatory Compliance Manager, Purchasing Manager, Customer Service Manager and Field Technicians.Âż Advise your customers regarding the proper care and storage of the nebulizers.Âż Your customers should discard their NEB-ROS User Manual and immediately download the latest versions HERE.Âż If they are unable to download the most recent version, hard copies can be provided. Please contact Roscoe Medical Customer Support at (800) 376-7263 Monday - Friday 8:00 am EST - 5:00 pm EST.Âż Complete and return the attached Recall Field Correction Response Form to Roscoe Medical within fifteen (15) days of receipt of this field correction notification confirming your acknowledgement. Send completed form to Roscoe Medical via fax number (440) 572-4261 or via email at corporate@roscoemedical.com. |  | 219353 | 219,353 Manuals | Worldwide Distribution US nationwide in the states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV &amp; WY and countries of:: Australia, Canada, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, France, Hong Kong, India, Norway, Spain, Trinidad and Tobago &amp; United Kingdom. | Completed |  |  | 219353 | Not_Computer
Z-1730-2015 | 71142 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Low Volume Extension Set, 72' w/Clamp FLL/MLL APV=1.3ml, STERILE, AND FLUID PATH NONPYROGENIC UNLESS PACKAGE IS DAMAGED OR OPENED. LATEX FREE, SINGLE USE Tyvek pouches, 10 per box, Smiths Medical ASD, Inc. |  |  | May 04, 2015 | 2015 | Smiths Medical Asd Inc | Smiths Medical has become aware of an issue with the Low Volume Extension Set, (72'' Extension Set) where under-sized tubing was used in the production of 1 lot of product. | Process control | Smiths Medical sent Urgent Medical Device Recall Notice letters, dated 4 May 2015, requiring its customers to return all affected unused Extension Sets. Please take the following actions: 1. Customers are to inspect their inventory and quarantine any affected Extension Sets.2. Complete and return the attached Confirmation Form by Fax or by e-mail3. Upon receipt of the completed form, a customer service representative will contact customers to arrange for exchange of their unused affected Extension Sets for credit or replacement.The notice shall be passed on to all personnel who need to be aware, including points of use. If these affected Extension Sets have been further distributed to other persons or facilities, please promptly forward the recipients a copy of this Urgent Medical Device Recall Notice. |  | 6000 | 6,000 units | Distributed to the states of KY, MN, NJ, UT, &amp; WI. | Terminated | January 20, 2016 | 261 | 6000 | Not_Computer
Z-1724-2015 | 71174 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Laboratory Automation System (enGen System), Product Code ENGEN, configured with all of the following system components: Recapper Caps (Product Codes 952238 or 6844286 only), Recapper Module (230V) or Recapper Module (110V) (Product Codes 952036-EG or 952136-EG), Rack Entry and Exit Module and/or Rack Exit Module (Product Codes 952020-EG/6844086 or 952019-EG/6844103), Storage Rack Slim 100 positions or Storage Rack 180 positions (Product Codes 952075 or 952073). |  |  | May 06, 2015 | 2015 | Ortho-Clinical Diagnostics | Recapper Caps may fall off of, or be knocked from sample tubes placed in the storage racks of entry/exit modules. The gripper may grab a tube at the wrong height and drop the tube in the Rack Entry/Exit or Rack Exit module. These situations may result in splashing of bio-hazardous fluid. Sample cross-contamination may occur with an uncapped tube in the rack. | Component change control | On 5/06/2015, a customer letter (Ref. CL2015-092) with recommendation to prevent this issue was sent to all customers in the US with enGen Systems via FedEx or US Priority Mail (for Post Office Box addresses) or by email to subscribers of ORTHO PLUS e-Communications. Foreign affiliate consignees were notified by e-mail on 5/06/2015 informing them of the issue. For questions, please call Customer Technical Services at 1-800-421-3311 at any time. |  | 45 | 45 units worldwide (Domestic: 19 units, Foreign: 26) | Worldwide distribution. US Nationwide, Australia, Bermuda, Brazil, Canada, Chile, China, France, India, Italy, Norway, Portugal, Spain, Sweden, and the United Kingdom. | Open, Classified |  |  | 45 | Not_Computer
Z-1706-2015 | 71054 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | FLOW-i Anesthesia System |  |  | April 02, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | In some Flow-i Anesthesia Systems, patient cassette can get dislodged which could result in gas leakage within the system internal circuit potentially leading to a stop in ventilation. | Under Investigation by firm | The firm, Maquet Inc., sent an "Urgent Medical Device Field Action Field Correction for FLOW-i Anesthesia Systems" letter dated 4/2/2015 via FedEx. The letter describes the product, problem and actions to be taken. The customers were instructed to read the letter and inform staff, who may use the FLOW-i Anesthesia System of the patient cassette locking device issue. Customers are to obtain the serial number of their system; verify the serial number obtained against the serial number listed in the box in the Appendix. A Maquet Service Representative will be contacting you to arrange a convenient time to replace your FLOW-i Anesthesia System patient cassette locking device.If you have any questions regarding this field action, please contact your Maquet Service Representative or Maquet Technical Support at 1-888-627-8383 (Press option 3, followed by option 1 and then option 1 again), Monday through Friday, between the hours of 8:00 a.m. and 5:00 p.m. EST. |  | 1641 | 1641 [80 units (US) 1561 units (OUS)] | Class I Recall - Worldwide distribution -- US, including the states of AL, AZ, DC, FL, MI, MO, NJ, NY, OR, PA, TX, and WV; and, countries of Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Chile, China, Colombia, Croatia, Czech Republic, Ecuador, Finland, France, Germany, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Libya, Mexico, Mongolia, Mozambique, Netherlands, Norway Oman, Panama, Peru, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain Sweden, Switzerland, Syria, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Unites States, and Venezuela. | Terminated | July 06, 2016 | 461 | 1641 | Not_Computer
Z-1698-2015 | 71164 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD). |  |  | April 29, 2015 | 2015 | HeartWare Inc | HeartWare has received complaints relating to damage or bent connection pins within the power supply ports of the VAD controller. | Device Design | HeartWare sent an Urgent Medical Device Correction letter to customers on May 25, 2015. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to:1. Review the enclosed notice and Patient Communication .2. Forward the notice to those individuals within their organization who need to be aware of its contents.3. Identify the patients currently supported by the HVAD System.4. Distribute the "Patient Communication" to their patients directly via FedEx or some other reliable means of communication.5. Continue to reinforce the messages set forth in the notice with their patients during their regularly scheduled appointments.6. Complete, sign and return the Acknowledgement and Completion Form to HeartWare within 30 days of receipt of this letter.For questions contact your local HeartWare representative or the 24-Hour Clinical Support personnel at 1-888-494-6365. |  | 3747 | 3,747 currently implanted worldwide | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA and countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Not_Computer
Z-1728-2015 | 70918 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Software Version 2.8 &amp; Below |  |  | April 06, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly regarding urine samples that require acidified pretreatment. The software has allowed urine samples for assays that require acidified pretreatment to be metered from the same sample container as assays that required no pretreatment when the sample was programmed. | Software design | Ortho Clinical Diagnostics distributed correction notification letters dated April 6, 2015, and response forms to their customers via courier service. All urine patient results that require pretreatment was advised to refer to the notice of correction for instructions. Correction notices should also be posted beside the device user documentation for easy reference. Customers were advised to complete and return the Confirmation of Receipt before April 20, 2015. Customers with questions may contact the Customer Technical Service at 1- 800-421-3311 anytime. |  | 2387 | 2,387 units Total (987 domestically &amp; 1400 internationally) | Worldwide Distribution - US Nationwide, Brazil, Chile, Mexico, Australia, Hong Kong, India, Taiwan, Belgium, Denmark, France, Ireland, Italy, Portugal, Romania, Spain, Sweden and the UK. | Open, Classified |  |  | 2387 | Not_Computer
Z-1699-2015 | 71272 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | M/L Taper with KinectivÂż Technology. |  |  | May 18, 2015 | 2015 | Zimmer, Inc. | Zimmer is initiating a voluntary recall of 64 lots (752 implants total) of M/L Taper with KinectivÂż femoral stems and modular necks due higher than allowed cytotoxicity levels found with the product. Reasonable probability of adverse biological response and subsequent revision | Environmental control | Zimmer, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated May 18, 2015, to all affected consignees. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to review the notification and ensure affected personnel are aware of the contents. Locate all affected product identified and quarantine them immediately. Carry out a physical count of all affected product in their inventory and complete the Inventory Return Certification Form. Email a completed copy to corporatequality.postmarket@zimmer.com. Return the recalled product along with the completed Inventory return Certification Form. Notify Zimmer of any hospitals that they have further distributed the affected product to. In addition, identify the surgeons that have implanted the product. Supply the information for any hospitals that they have identified, as well as the affected surgeons using the provided spreadsheet template. Customers with questions or concerns should call the customer call center at 1-877-946-2761. |  | 752 | 752 | Worldwide Distribution - US Nationwide in the states of AK, AL, AZ, CA, FL, GA, IL, IN, KS, MA, MI, MN, MO, NC, NY, OH, OK, PA, TX, UT, VA, WA, WI and countries of Canada, Australia, Japan, Taiwan, France, Germany, Spain, and Italy. | Terminated | August 02, 2016 | 442 | 752 | Not_Computer
Z-1638-2015 | 71298 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE OEC Brivo 865 plus, Mobile X-ray system laser aimer assemblies (part number 520002DW) |  |  | May 22, 2014 | 2014 | GE OEC Medical Systems, Inc | Laser aimer assemblies were shipped without certification of the overall assembly or submitting a laser product report and therefore do not comply with the US regulations for laser products. | Radiation Control for Health and Safety Act | GE Healthcare Surgery Planned Action(s) to repair the defect or to bring the product into compliance: The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 9 units consisted of the following:1. Replacing the tube side cover with the laser aimer assembly with a tubeside cover without a laser aimer assembly.This correction brought the systems into compliance with applicable US regulations. For further questions please call (801) 536-4615. |  | 9 | 9 | US Distribution to the states of: MD, NC, MT, NJ, UT and NY | Terminated | June 24, 2015 | 398 | 9 | Not_Computer
Z-1725-2015 | 71244 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alkaline Wash, List number 9D31-20, a cleaning solution used by the Abbott Architect analyzers. It is supplied in 500 mL bottles and contains NaOCl, NaOH, and KOH. It is shipped and stored at room temperature and is shipped ready to use. |  |  | May 11, 2015 | 2015 | Abbott Laboratories | The product is leaking from the cap and crystallization was noted as present around the cap. | Under Investigation by firm | Abbott Laboratories sent a Product Recall letter dated May 11, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to wear the appropriate Personal Protective Equipment (PPE) to determine if they have the affected product in their inventory. If customers observe any signs of leakage or loose caps discontinue use and disppose of any remaining inventory of those kits in accorance with their institutions's policies and procedures. If customers do not observe any signs of leakage or loose caps they may continue to use the product following the precautions per the ARCHITECT System Operations Manual and the Safety Data Sheet. Customers were also asked to complete and return the enclosed Customers Reply form. If affected product was forwarded to other laboratories, customers need to inform them of the Product Recalls and provide them a copy of tthe product recall letter. Customers should also maintain a copy of the recall letter. Customers with questions were instructed to contact Customer Service at 1-877-4ABBOTT. For questions regarding this recall call 224-668-1926. |  | 8868 | 8868 units | Worldwide Distribution - US (nationwide), Mexico, Colombia, Brazil, Germany, Hong Kong, Belgium, Czech Republic, Finland, Hungary, Ireland, China, New Zealand, Uruguay, Puerto Rico, St. Lucia, Austria, Croatia, Denmark, France, India, Isreal, Bermuda, Curacao, Cayman Islands, Jamaica, Canada, Bangladesh, Cyprus, Estonia, Greece, Indonesia, Italy, Kazakhstan, Latvia, Macedonia, Norway, Portugal, Russia, Slovakia, Spain, Taiwan, UAE, Kenya, Lithuania, Malaysia, Philippines, Qatar, Saudi Arabia, Slovenia, Sweden, Tunisia, United Kingdom, Kuwait, Luxembourg, Netherlands, Poland, Romania, Singapore, South Africa, Switzerland, Turkey, and Vietnam. | Completed |  |  | 8868 | Not_Computer
Z-1723-2015 | 71344 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Vilex, Inc., FUZE 10.0 mm x 250 mm, Product # FZ100-250T-25. |  |  | April 22, 2015 | 2015 | Vilex In Tennessee Inc | Incidence of breakage is higher than expected. | Device Design | The initial telephone call (4/9/2015) informed the distributor that the product was on hold (not to be used or transferred). The second telephone call (4/29/2015) and letter informed the customer that the product was being removed from market and that the customer was to return the product to the corporate office.The letter was disseminated on 05/05/2015 to distributors via certified mail. The letter contains the instructions for them to follow, which as summarized are: inventory all products on hand, sign off on inventory sheet, send products back to the corporate office with the identified return number on the shipping container, inform the corporate office that the process has been completed and provide the corporate office with the list of the products returned.A second letter was disseminated to sales representatives (distributors), surgeons, and hospitals on 05/28/2015. |  | 4 | 4 units | Distributed in the states of MS, MD, PA, and TX. | Open, Classified |  |  | 24 | Not_Computer
Z-1727-2015 | 71308 | No | LKF | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Ally Uterine Positioning System (Ally UPS) used in the mounting and positioning and holding of uterine manipulators |  |  | May 18, 2015 | 2015 | CooperSurgical, Inc. | Design of device may expose user to injury to fingers or body parts | Device Design | CooperSurgical initiated the recall on May 18, 2015 via Fedex with confirmed delivery receipt. The letter identified the affected product and asked users to discontinue use and complete the Acknowledgement and Receipt Form to schedule a replacement.Questions contact 203.601.5200. |  | 5 | 5 units | Distributed in the states of CA, IL, NC, and NY. | Terminated | August 07, 2015 | 81 | 5 | Not_Computer
Z-1702-2015 | 71156 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | The Alcon Custom Pak. The Alcon Customer Pak is an assemblage of single-use medical devices and accessories designed by medical professionals for specific ophthalmic surgical procedures. Some Paks contain the Devon Light Glove. |  |  | April 28, 2015 | 2015 | Alcon Research, Ltd. | The supplier of the Devon Light Glove, a component of the Alcon Custom PAK, recalled the gloves due to the potential for splits and holes on the glove. Should the user be unaware that the Light Glove is torn/split, surgical site infections are possible. | Nonconforming Material/Component | Alcon sent a Recall letter dated April 28, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter included a return form and instructions for credit. Customers were instructed to review inventory for any affected units, segregat the product to ensure it is not used, and return the attached Response Form via fax or email to Alcon.For questions or need assistance with obtaining replacement products contact Alson Customer service at 800-862-5266. |  | 219680 | 219,680 units | Worldwide Distribution - US Nationwide and the countries of Australia and Canada. | Terminated | March 24, 2016 | 331 | 219680 | Not_Computer
Z-1742-2015 | 70917 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 4600 Chemistry System, Software Version 3.2 &amp; Below. |  |  | April 06, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly allows testing of multiple assays using a single urine specimen, regardless of the pretreatment requirements for each selected assay, potentially leading to biased patient results. | Software design | On April 6, 2015, Ortho Clinical Diagnostics distributed correction notification letters and response forms to their customers via courier service. All urine patient results that require pretreatment was advised to refer to the notice of correction for instructions. Correction notices should also be posted beside the device user documentation for easy reference. Customers were advised to complete and return the Confirmation of Receipt before April 20, 2015. Customers with questions may contact the Customer Technical Service at 1-800-421-3311 anytime..***CAPA***Since the Assay Data Disk (ADD) provides data tables that determine which assays can be tested from the same sample container, an updated ADD is being created to address the anomaly on the VITROS 4600 and 5600. The anomaly will be addressed with an updated ADD is uploaded and the revised ADD mitigates the issue totally. Anticipated date of availability TBD. |  | 255 | 255 units total (131 domestically &amp; 255 internationally) | Worldwide Distribution-US (nationwide), Brazil, Chile, Mexico, Australia, Hong Kong, India, Taiwan, Belgium, Denmark, France, Ireland, Italy, Portugal, Romania, Spain, Sweden and the UK. | Open, Classified |  |  | 255 | Not_Computer
Z-1732-2015 | 71246 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | ClearSight, Model No. EVHRS Heart Reference Sensor. The EV1000 Clinical Platform Nl and ClearSight finger cuffs noninvasively measures blood pressure and associated hemodynamic parameters. |  |  | May 18, 2015 | 2015 | Edwards Lifesciences, LLC | Edward Lifesciences is recalling the ClearSight System because the finger component of the Heart Reference Sensor may be damaged. | Device Design | Edwards Lifesciences sent an URGENT Product Notification letter dated May 12, 2015, to all affected customers. The letter informed customers that Edwards Lifesciences has received reports that a component of the ClearSight System can be damaged if not handled properly. The letter informs customers of the problems identified and the actions to be taken. Customers with questions are instructed to call Edwards Technical Support. For questions regarding this recall call 949-250-2500. |  | 1309 | 1,309 units total (696 units in US) | Worldwide Distribution - US (nationwide) and Internationally to Australia, Estonia, Italy, Romania, Sweden, Austria, Finland, Kuwait, Saudi Arabia, Switzerland, Bahrain, France, Netherlands, Slovenia, Turkey, Belgium, Great Britain, Oman, Slovakia, South Africa, Bulgaria, Greece, Poland, South Africa, United Arab Emirates, Czech Republic, Hungary, Portugal, Spain, Denmark, Ireland, and Qatar. | Terminated | November 06, 2015 | 172 | 1309 | Not_Computer
Z-1734-2015 | 71235 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | AccuChek Spirit insulin infusion pump. |  |  | April 22, 2015 | 2015 | Roche Diabetes Care, Inc. | Roche has become aware of an increase of mechanical errors experienced by customers using the Accu-Chek Spirit Combo insulin infusion pump and Accu-Chek Spirit insulin infusion pump when the insulin pumps display E6 and El0 error messages.If the user does not act upon the E6 and E10 error messages appropriately, insulin delivery will be stopped and if unnoticed may lead to severe hyperglycemia u | Device Design | Roche sent an URGENT MEDICAL DEVICE CORRECTION notification dated April 22, 2015, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer. The recall notification included a description of the reason for the recall, customer responsibilities, training chart, and instructions for responding to the formal recall notification. Please contact the ACCU-CHEK Customer Care Service Center, 24 hours a day, seven days a week at 1-800-688-4578, if you have questions about the information contained in this notification. |  | 35795 | 35,795 | Nationwide Distribution | Open, Classified |  |  | 35795 | Not_Computer
Z-1739-2015 | 71330 | No | KWZ | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | R3(TM) 28 MM I.D., 62 MM O.D. CONSTRAINED LINER, REF 71339162, QTY: (1), STERILE EO. |  |  | May 06, 2015 | 2015 | Smith &amp; Nephew, Inc. | The R3 Constrained outer liner locking groove diameter is 0.030 inches oversized and may not allow the locking ring to engage into the R3 shell. | Process change control | Smith &amp; Nephew sent an Urgent-Product Recall Notification letter dated May 6, 2015. to Sales offices by E-mail and Hospitals via Fed- Ex. The letter identified the affected product, problem and actions to be taken. Customers were asked to quarantine the product and return it. Distributors were asked to notified their customers of the recall and ask them to return the product. |  | 32 | 32 units | Worldwide Distribution - US Nationwide in the states of CA, FL, IL, IN, MI, NE, NY, TN, TX, VA, WI and the countries of Australia, Switzerland, Chile, Canada, and Spain | Open, Classified |  |  | 101 | Not_Computer
Z-1741-2015 | 71255 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Unilateral external fixation ankle clamp |  |  | May 01, 2015 | 2015 | Biomet, Inc. | Ankle clamp was assembled incorrectly; the offset of the pin to the pivot point is reversed. This prevents complete range of motion and if pushed too far may loosen the pin from the bone. A revision surgery may be necessary to adequately reconnect the fixator to the bone. | Process control | Biomet sent an URGENT MEDICAL DEVICE RECALL NOTICE letter dated May 1, 2015 to the affected distributors and implanting surgeons. The recall letter included a description of the reason for the recall, affected product, consignee responsibilities, and distributor instructions for responding to the formal recall notification. Consignees were instructed to return the affected product and continued monitoring of patients. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 19 | 19 | US Nationwide Distribution in the states of CA, TX, FL, NC, MI, GA. | Terminated | December 08, 2015 | 221 | 19 | Not_Computer
Z-1740-2015 | 71232 | No | KAM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | RELIEVA FLEX Sinus Guide Catheter F-70. |  |  | April 28, 2015 | 2015 | Acclarent, Inc. | Acclarent determined that one single RELIEVA FLEX Sinus Guide Catheter was distributed past it's expiry date. | Nonconforming Material/Component | Acclarent sent an Urgent Medical Device Recall letter dated May 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed that if the affected catheter is still in inventory, it should be isolated and returned to Acclarent per instructions.Questions should be addressed to Gabriel Alfageme, Field Action Coordinator at 1-650-687-5398. |  | 986 | 986 | US in the state of Texas | Terminated | June 10, 2015 | 43 | 986 | Not_Computer
Z-1746-2015 | 69954 | [6 empty] | BBL Port A Cul tubes, catalog number 221606. Patient sample collection tubes. BBL Port A Cul tubes, jars and vials contain a reduced transport medium and are intended to maintain the viability of anaerobic, facultative and aerobic microorganisms during transport from the patient to the laboratory. Sterile packages are for collection of specimens in clean areas; e.g., surgical suites. |  |  | November 21, 2014 | 2014 | Becton Dickinson &amp; Co. | The products may exhibit indications of excessive oxidation. Excessive oxidation may affect the viability of anaerobic microorganisms. | Process control | BD Diagnostic Systems sent an Urgent Product Recall letter dated December 2014 to all affected customers. Recall notification was initiated on November 21, 2014, the notification was sent by fax and letters by UPS. Customer notification for the expansion of the recall was initiated on December 18, 2014, the notification was sent by fax and letters by UPS. For questions regarding this recall call 800-638-8663. |  | 145860 | 145,860 tubes | US Nationwide distribution in AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | January 07, 2016 | 412 | 238570 | Not_Computer
Z-1748-2015 | 71198 | No | JSG | Multipurpose culture medium. | Microbiology | Microbiology | 510(K) Exempt | Remel Blood Agar (TSA w/sheep blood) plate REF R01202 packaged 10 plates per pack, 10/10-packs per case (100 plates). |  |  | May 04, 2015 | 2015 | Remel Inc | Product may be contaminated with Listeria monocytogenes. | Material/Component Contamination | The recalling firm sent the recall notification letter (dated May 1, 2015) to all customers on 5/4/2015 via standard mail. |  | 8450 | 845/10/10-plate pack cases | Nationwide Distribution. | Open, Classified |  |  | 8450 | Not_Computer
Z-1750-2015 | 71100 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Astral 100 ( Model No. 27001) &amp; Astral 150 (Model No. 27003) ventilator. |  |  | May 05, 2015 | 2015 | ResMed Ltd. | The Astral device allows clinicians to disable all alarms including those that detect circuit disconnection. Inappropriate disabling of alarms in ventilator dependent patients can lead to hazards of insufficient ventilation and insufficient expiratory pressure which can result in major severity harm. | Use error | A customer notification letter dated 5/5/15 will be sent to customers that the Astral device allows clinicians to disable all alarms including those that detect circuit disconnection. The letter provides the customers with the hazards involved and the actions to be taken. Customers are instructed to ensure alarms are appropriately set up for ventilator dependent patients. Customers with questions are instructed to contact Larissa D'Andrea at (858) 836-6837 or email larissa.d'andrea@resmed.com. |  | 1519 | 1519 units | Worldwide distribution. US nationwide, Uruguay and Canada. | Terminated | April 07, 2016 | 338 | 1519 | Not_Computer
Z-1745-2015 | 71393 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | AGFA DX-D 100 is a mobile digital radiography x-ray system |  |  | May 29, 2015 | 2015 | AGFA Healthcare Corp. | When using the DX-D 100 unit, sporadic unintended movements caused by an electrostatic discharge of the unit to the ground may occur. | Other | AGFA Healthcare sent an "URGENT FIELD SAFETY NOTICE" letter dated May 29, 2015, to all affected customers. The letter was either emailed or sent via FedEx to the US customers. When using the DX-D 100 unit, sporadic unintended movements caused by an electrostatic discharge of the unit to the ground may occur.Please distribute this information within your facility to all those who need to be aware of the notice. Please complete the feedback form as soon as possible and return it to us.We thank you for your careful attention to this issue and your continued support. If you have any questions about this matter, please contact your local Agfa HealthCare organization: 1-877-777-2432 and please reference PR1411200006 and PR1501190003. |  | 233 | 233 units | US Distribution to the states of : AL, CA, DC, FL, GA, IL, IA, KY, LA, MD, MA, MI, MN, MS, MO, NE, NJ, NY, NC, OH, OK, OR, SC, TN, TX, VA, WV, WI and WY. | Open, Classified |  |  | 233 | Not_Computer
Z-1751-2015 | 71241 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD) For use as a bridge-to-cardiac transplantation in patient who are at risk of death from refractory end-stage left ventricular heart failure. |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Discolored and Cracked Driveline Outer Sheath. | Device Design |  |  | 3747 | 3,747 currently in use | AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, and VA, Foreign: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxemburg,Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK, | Open, Classified |  |  | 3747 | Not_Computer
Z-1744-2015 | 71365 | No | KIH | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | The Optigun Ratchet is a hand-held cement gun for use with Optivac. The Optigun Ratchet is used to deliver bone cement from the Optivac. |  |  | May 14, 2015 | 2015 | BIOMET FRANCE S.A.R.L. | The pin which maintains the knob button, may disconnect and become lost. This may result in a delay of surgery to obtain a new Optigun and detect/retrieve the pin, and may necessitate revision surgery if the pin falls into the surgical site and is not detected. | Nonconforming Material/Component | BIOMET sent an URGENT ADVISORY NOTICE dated May 14, 2015, to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to check the product to identify to determine if the pin is defective, if defective they should return it to Biomet Warsaw, following the directions on the Response Form and confirm receipt of the notice by sending back the response form. Questions related to this notice should be directed to (574) 372- 1570 Monday through Friday 8 a.m. to 5 p.m. |  | 3 | 3 | Nationwide distribution to MD only | Terminated | October 27, 2015 | 166 | 3 | Not_Computer
Z-1749-2015 | 71297 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | ORTHOFIX Firebird Spinal Fixation System 6.5mm x 55mm Multi-Axial Screw, Standard Implant Grade Titanium Alloy Rx Only Provided Non-Sterile Do Not Reuse |  |  | May 18, 2015 | 2015 | Orthofix, Inc | The Firebird Spinal Fixation System 6.5 mm x 55 mm Bone Screw, Self-Tapping is currently mislabeled at the distal tip of the bone screw showing a length measurement of 50 mm instead of 55 mm. | Unknown/Undetermined by firm | Orthofix sent an Urgent Medical Device Recall Notification letter dated May 15, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Affected customers were instructed to inspect their inventory, and identify and remove the recalled product to send back to Orthofix. The distributors are also instructed to provide recall notification to hospitals, surgeons, or other distributors that they have further distributed affected product to.Contact your local Orthofix representative or Orthofix customer service representative at 888-298-5700 for return instructions and shipment information. |  | 45 | 45 units | US Distribution to the states of : OK, MD, CA, FL, MI, SC, CO, NY and PA., and Internationally to Asturias Spain. | Terminated | October 30, 2015 | 165 | 45 | Not_Computer
Z-1759-2015 | 71134 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16 Slice (Uninterruptible Power Supply (UPS) accessory devices) Computed Tomography X-ray systems and Diagnostic Imaging systems. |  |  | March 06, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm was notified that the failure of the UPS accessory devices used in conjunction with Computed Tomography X-ray systems and Diagnostic Imaging systems. may result in battery acid leakage, overheating and/or emission of fumes. | Nonconforming Material/Component | On 3/9/2015 the firm sent Urgent Field Safety Notices Medical Device Correction letters to their customers. |  | 125 | 125 units | Worldwide Distribution-US (nationwide) including DC &amp; Puerto Rico and the following countries: Albania, Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Bulgaria, Cambodia, Cameroon, Cayman Islands, China, Colombia, Democratic Congo, Croatia, Cuba, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Ethiopia, Finland, France, Gabon, Georgia, Germany, Ghana, Greece, Guatemala, Hong Kong, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Kuwait, Kyrgyzstan, LAO People's Dem. Rep., Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malawi, Malaysia, Malta, Mauritius, Monaco, Mongolia, Morocco, Myanmar, Namibia, Nepal, Netherlands, New Zealand, Nicaragua, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Philippines, Poland, Portugal, RÂżunion, Romania, Russia, Rwanda, Saudi Arabia, Senegal, Serbia &amp; Montenegro, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syria, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Vietnam, Yemen, and Zambia. | Terminated | March 13, 2017 | 738 | 4590 | Not_Computer
Z-1757-2015 | 71171 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Cobalt-Chrome Femoral Components, 32MM MOD HEAD COCR -6MM NECK; 32MM MOD HEAD COCR -3 MM NECK; 32MM MOD HEAD COCR STD NECK ; 32MM MOD HEAD COCR +3MM NECK; 32MM COCR MOD HD +6MM NO SKIRT |  |  | April 16, 2015 | 2015 | Biomet, Inc. | Biomet hip joint, knee joint and shoulder joint replacement components device labeling contains incorrect GTIN (Global Trade Identification Number) codes. | Labeling mix-ups | Biomet sent an Urgent Medical Device Recall Notice letter dated April 29, 2015 to affected customers via Fed-Ex. The letter identifed the affected product, problem and actions to be taken. Customers were asked to complete the attached response form and return it along with recalled product to Biomet. For questions callt (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 1315 | 1315 | Worldwide Distribution - US Nationwide in the states of AR, CA, IL, IN, PA, FL, KS, LA, WI, VA, MA, MD, NJ, UT, NC, TX, AL, MI, KY, OH, SD, MO, GA, NV, WA and the countries of: Europe, Canada, Japan, Chile, Korea, Mexico, New Zealand, Trinidad/Tobago, Thailand, Argentina, Australia, and China. | Open, Classified |  |  | 2632 | Not_Computer
Z-1758-2015 | 71269 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences IntraClude Intra-Aortic Occlusion Device, model number ICF100, 10.5 Fr (3.5 mm) x 39.4" (100 cm) ThruPort System. Packaged sterile and non-pyrogenic in a sealed, peel-type pouch |  |  | May 05, 2015 | 2015 | Edwards Lifesciences, LLC | Potential device damage on the strain relief near the hub area, which may result in blood leakage. | Device Design | Urgent Medical Device Safety Notice letters were sent via Fed-Ex by May 05, 2015. As part of the customer instructions in the safety notice, customers were requested to review the notice to understand the potential hazard, complete and return the acknowledgement form, and forward copies of the Safety Notice to any organization where the potentially affected devices have been transferred. The device will only be distributed to customers who have returned the acknowledgement form. Questions should be directed to Edwards Customer Service at 800-424-3278. |  | 3360 | 3,360 units. | Nationwide Distribution -- US, including the states of Alabama, Arizona, California, Colorado, Connecticut, Florida, Georgia, Idaho, Illinois, Kentucky, Louisiana, Massachusetts, Maine, Michigan, Missouri, Mississippi, North Carolina, Nebraska, New Jersey, New York, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, Wisconsin, and West Virginia. | Open, Classified |  |  | 3360 | Not_Computer
Z-1726-2015 | 71242 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD). |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Retraction of Pins within the driveline connector. | Device Design | Urgent Medical Device Correction letters, dated APR2015, were sent to users. The letters identified the affected device, as well as the reason for the recall and risks involved. Users were asked to review and familiarize themselves with the notice and "Patient Communication." The notice should be forwarded to anyone within customers' organization that need to be aware. Customers are to identify their patients who are currently supported by the HVAD System and distribute the "Patient Communication" to them. A HeartWare representative can assist with that process is needed. Customers should continue to reinforce the messages in the notice with their patients. The "Acknowledgement and Completion Form" should be completed, signed, and returned within 30 days of receipt of the notice. Questions or concerns should be directed to a local HeartWare representative. Clinical Support is also available at 1-888-494-6365. |  | 3747 | 3,747 currently in use | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, and VA; and, the countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Not_Computer
Z-1753-2015 | 71238 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Gyrus ACMI Diego Elite Tubeset, intended for cutting, coagulation, drilling, debriding, and removal of bone, and soft and hard tissue in general ENT, Sinus/Rhinology, Nasopharyngeal/Laryngology, and Head and Neck procedures. Catalog No. TS101DC, Lot numbers JC924165 and below; Rx Only, STERILE EO |  |  | May 08, 2015 | 2015 | Gyrus Acmi, Incorporated | Potential lack of sterility assurance. | Under Investigation by firm | The firm notified their direct accounts of the recall by Fed Ex Second Day on 05/08/2015. The letter requesting disposal of the product was extended to hospitals or medical practices that purchased the product. Non-responding consignees will be contacted again if necessary. |  | 22250 | 22,250 boxes | Nationwide and Canada, Australia, Germany, New Zealand, Japan, India, Great Britain | Open, Classified |  |  | 34270 | Not_Computer
Z-1797-2015 | 71161 | No | MGQ |  |  | General &amp; Plastic Surgery | 510(k) | Amerigel Hydrogel Gauze Dressing 2 x 2, A5002. |  |  | February 25, 2015 | 2015 | Amerx Health Care Corp. | Nonapproved labeling claims (Sterile) | Labeling False and Misleading | On March 9, 2015 the firm contacted its customers via telephone. Within that conversation the firm asked its customers to examine their inventory to determine if they have any of the affected lot numbers on hand and notify the firm for an immediate product substitution, credit or refund. Customers were asked to discontinue dispensing any product on hand, and notify any customers that the product may have been further distributed to. The firm also asked that the response form be completed and returned. If no contact was made via telephone, the firm sent a certified letter asking the firm(s) to do the same actions as the phone request. If you have any questions please call 1-800-448-9599 or e-mail marketcorrection@amerxhc.com. |  | 6491 | 6,491 units | Worldwide Distribution -- US, including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WV, WY and Hawaii; and, the countries of Canada and Bahrain. | Open, Classified |  |  | 6491 | Not_Computer
Z-1794-2015 | 71278 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic Passive Planar Blunt Probe Part Number 960-556, also included in Part Number Kits: 9732540, 9732622, 9733174, 9733935. As part of kits, the product is either packaged with other instruments or into a holding tray. The Passive Planar Blunt Probe is an optically tracked, reusable instrument used in navigated surgery and is included with Medtronic cranial instrument kits. It is designed for use in patient registration and palpating brain tissue during surgical procedures. |  |  | May 05, 2015 | 2015 | Medtronic Navigation, Inc. | Passive Planar Blunt Probe (Lot# 150223) was manufactured with a sharp tip rather than a blunt tip. | Mixed-up of materials/components | Medtronic sent an Medical Device Safety Alert letter dated May 5, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action to be taken by the user:"Examine your inventory and if any of the affected products are found, please quarantine them for return to Medtronic."A Medtronic representative will follow-up to arrange for removal and replacement.If you have any questions regarding this action, please contact Medtronic Navigation Technical Services at 1-800-595-9709. |  | 18 | 18 | US Distribution to the states of : CA, MO, LA and PA., and Internationally to Ireland. | Terminated | July 02, 2015 | 58 | 18 | Not_Computer
Z-1796-2015 | 71405 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | ELI 380 Electrocardiograph. |  |  | May 29, 2015 | 2015 | Mortara Instrument, Inc | Mortara Instrument, Inc. has recently becomeaware of a potential safety hazard involving our ELI380 electrocardiograph. When used in a particularworkflow, acquired ECG waveforms for one patientmay become associated with the patientdemographics for a different patient when therecord is transmitted to a records managementsystem. | Software design | Mortara Instrument, Inc. sent an "Urgent Medical Device Correction" letter dated May 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised that users in their facilities be made aware of the safety notification and that required actions listed on the letter be taken immediately. A Field service Engineer will be dispatched to their facility to update the affected unit(s) on-site. For questions they can contact Technical support at 1-888-667-8272 or Techsupport@mortara.com or contact your local service Representative. |  | 176 | 176 | Worldwide Distribution - US Distribution to the states of : AZ, CA, FL, GA, IL, MI, NE, SC, UT, WA and WI., and to the countries of : CHINA, PORTUGAL, CANADA, OMAN, MALAYSIA, NEW ZEALAND, INDIA, POLAND, LEBANON, SWITZERLAND, FINLAND, ROMANIA,UNITED ARAB EMIRATES, UNITED KINGDOM, AUSTRALIA, THE NETHERLANDS, ITALY, FRANCE, GERMANY, THAILAND,SAUDI, ARABIA, NORWAY and IRELAND. | Terminated | August 05, 2015 | 68 | 176 | Computer
Z-1795-2015 | 71264 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | I-PortalÂż NOTC and VNG; a device used to measure, record, or visually display the involuntary movements (nystagmus) of the eyeball, indicated for use as a diagnostic tool to assist trained clinicians in their analysis of vestibular disorders, which requires the separation of central vs. peripheral nervous system deficits. |  |  | April 30, 2015 | 2015 | Neuro Kinetics, Inc. | complaints of system malfunction and unintended, sudden movement at start up. No injuries reported. | Software design | Neuro Kinetics sent a Field Correction Advisory letter dated April 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with the affected product were instructed to contact Customer Service if a malfunction is encountered and not to use the device until it has been assessed and corrected by Nuerokinetics. Customers with questions should call 412-963-6649. |  | 140 | 140 | Worldwide distribution - US (nationwide) and Internationally to Canada and Mexico. | Terminated | April 17, 2017 | 718 | 140 | Not_Computer
Z-1731-2015 | 71239 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD). |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Internal Controller Alarm Battery failures. | Device Design | Urgent Medical Device Correction letters were sent to customers. The letter identified the affected product, as well as the reason for the recall and risks involved. HeartWare requests that customers review and familiarize themselves with the letter and "Patient Communication," and ensure that all necessary individuals are aware. Patients currently being supported by the device should be identified and given the "Patient Communication." A HeartWare representative can provide assistance with that process if needed. The "Acknowledgement and Completion Form" should be completed, signed, and returned within 30 days of the receipt of the letter. Customers can contact their local HeartWare representative with questions or concerns. Clinical Support personnel are also available at 1-888-494-6365. |  | 3747 | 3,747 curently implanted | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, and VA; and, the countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Not_Computer
Z-1808-2015 | 70974 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | T/S2 0.9% Sodium Chloride Injection, USP, AMUSA Sterile Field Flush Syringe, 10 mL in 12 mL Single Use Syringe, Rx only, Part no. 2S0706 |  |  | February 09, 2015 | 2015 | MRP, LLC dba AMUSA | Incorrect labeling of the expiration date | Process control | The firm, AMUSA, Inc., initiated the recall by telephone to the sole consignee on 02/09/2015. The consignee were informed of product, the problem and actions to be taken. The consignees verified that they were in possession of the product and that all the product has been destroyed.If you have any questions, call 513-235-1940 or email: qadirector@amusa.us. |  | 1080 | 1,080 units (3 cases) | US distribution to CA only. | Terminated | June 22, 2015 | 133 | 1080 | Not_Computer
Z-1801-2015 | 71247 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | T.A.L.O.N (Tactically Advanced Lifesaving Intraosseous Needle) 15G x 38.5 mm 10 NEEDLE SET STERILE EO Single Use Rx only Vidacare Corporation |  |  | May 12, 2015 | 2015 | Vidacare Corporation | Potential breach in sterility | Device Design | Teleflex sent an Urgent Medical Recall Notification dated May 12, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter directed the customers to quarantine the affected product and return it to the manufacturer. If you have any further questions please call 1-866-246-6990. |  | 2104 | 2,104 boxes | US Distribution to the states of : NC, MD, MS, TX and MO., and Internationally to Canada and Denmark. | Terminated | April 11, 2016 | 335 | 2104 | Not_Computer
Z-1800-2015 | 71276 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.7mm Cannulated Locking Screws and 3.7mm Cannulated Conical Screws; . the Cannulated Screws are indicated for fixation of fractures, osteotomies, and non-unions of the clavicle, scapula, olecranon, humerus, radius, ulna, pelvis, tibia, and fibula, particularly in osteopenic bone. |  |  | May 13, 2015 | 2015 | Synthes, Inc. | Certain lots of the 3.7mm Cannulated Locking Screws and 3.7mm Cannulated Conical Screws were found to have been labeled incorrectly. | Employee error | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated May 13, 2015, to all affected customers to advise them about the issue and associated risks. These accounts were asked to review inventory, immediately remove the affected lots that are still in the original packaging from stock, and to complete and return the response and return authorization forms. Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes consultant. |  | 54 | 54 | Nationwide Distribution | Open, Classified |  |  | 54 | Not_Computer
Z-1798-2015 | 71253 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard M Unicondylar Tibial Trial |  |  | April 22, 2015 | 2015 | Biomet, Inc. | The Left and Right orientation of the trial was incorrectly etched on one side. When the instrument is rotated about the axis of the handle 180 degrees to view the other side, the etch on the handle now shows R on the same side as the keel marked LM and the L on the same side as the keel marked RM. This is incorrect. Please note that the keel is not etched on this side by design. | Process control | On 4/22/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications, dated April 22, 2015 were sent to the affected distributors with instructions for returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers are required to immediately locate and remove the affected product from circulation, follow the instructions on the "Response Form," and e-mail a copy of the completed form to the designated e-mail address prior to returning the product. Additionally, if the affected product has been further distributed, the necessary personnel must notified via the "Dear Risk/Recall Manager" notice. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 27 | 27 | Worldwide Distribution -- US, including the state of IL, NJ, Canada, Netherlands, and Chile. | Terminated | March 22, 2016 | 335 | 27 | Not_Computer
Z-1806-2015 | 71293 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BVI Custom EyesÂż Ophthalmic Procedure Kits with the component -Devon Light Glove |  |  | May 01, 2015 | 2015 | Beaver-Visitec International Inc. | Devon Light Glove in procedure kits may have holes/slits compromising sterility of the product | Nonconforming Material/Component | Beaver-Visitec International Sales Ltd. (BVI) sent an Urgent Field Safety Notice dated May 2, 2015, to all affected customers. Customers were instructed to quarantine the affected kits and return to BVI for subsequent replacement of product. A Verification Form was provided to confirm receipt of the recall notice. For questions regarding this recall call 791-906-7950. |  | 268 | 268 kits | Nationwide Distribution including KY, GA, MI, and NJ. | Open, Classified |  |  | 2072 | Not_Computer
Z-1799-2015 | 71364 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Pull Reduction Device For 4.3mm Percutaneous Drill Guide; The pull reduction device with quick coupling is placed through the guide sleeve and plate holes to pull or push bone fragments relative to the plate. This instrument can be used for: Minor Varus-valgus adjustments (approximately 2Âż - 4Âż), Translational adjustments, Stabilization of plate bone orientation during insertion of the first screws, alignment of segmental fragments and pre-drilling dense or thick cortical bone before placing a 5.0 mm locking screw. |  |  | May 28, 2015 | 2015 | Synthes, Inc. | Certain affected parts and lots of the Pull Reduction Devices may have been manufactured to an incorrect hardness specification which could result in intraoperative breakage. | Component design/selection | DePuy Synthes sent an Urgent Notice Urgent notice dated May 28, 2015, to all affected customers to advise them of the defect and patient risks associated with its use. Customers were instructed if they have any of the identified devices, please take the following steps:Âż Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number.Âż Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affectedproduct has been located. Also, please indicate the number of devices found and note the ReturnAuthorization Number. Please include your name, title, address, telephone number and signature in thespaces provided.Âż Return the Verification Section (page 3 of this letter) with the product to:Âż Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132.Âż Send a copy of the completed Verification Section by:Âż Fax: (888) 943-5206 orÂż Scan/email: Synthes2912@stericycle.comIf you DO NOT have the identified product, please take the following steps:Âż Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicatingthat no affected product has been located. Please include your name, title, address, telephone number andsignature in the spaces provided. This return documentation acknowledges your receipt of medical deviceremoval information.Âż Return the documents by:Âż Fax: (888) 943-5206 orÂż Scan/email: Synthes2912@stericycle.comÂż Note: If the Verification Section is answered on behalf of more than one facility and/or individual, pleaseclearly indicate the name and address of the facility and/or individual on page 3 of the notification.Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 3766 | 3747 US, 19 Foreign | Worldwide Distribution-US (nationwide) and internationlly to Canada | Open, Classified |  |  | 3766 | Not_Computer
Z-1807-2015 | 71432 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips Ultrasound, Model Q-Station, with software version 3 or higher, Catalog number: 795088; Part Number: 989605382391 |  |  | May 15, 2015 | 2015 | Philips Ultrasound, Inc. | Philips Healthcare has discovered a problem in the Philips Ultrasound Q-Station version 3.0 or higher that couldresult in measurements from a study (Structured Report) to be appended to subsequent studies for other patients. | Software design | PhilipsHealthcare sent an URGENT - Medical Device Correction Philips Ultrasound Q-Station, Q-STATION STRUCTURED REPORTING AND MEASUREMENTS letter, FCO 79500335, dated 2015 May 15, to all affected consignees.Philips is sending a CD with Q-Station software version 3.3.2 to consignees for a free of charge. Consignees must uninstall previous versions of Q-Station and install version 3.3.2 on your system carefully following the installation instructions that accompany the CD. It is important that consignees register their Q Station 3.3.2 installation as described in those instructions, so that Philips can confirm that they have completed this correction.Consignees were instructed to complete the enclosed Customer Reply Form to confirm that they have (1) read and understood this important Medical Device Correction and (2) have either installed Q-Station version 3.3.2 or declare that this FCO 79500335 does not apply to their institution. Consignees were asked to return the enclosed Customer Reply form to Philips Healthcare within 5 days of receipt of the replacement software. Consignees with questions can call local Philips representative or Philips Customer Service at 1-800-722-9377. |  | 125 | 125 units (30 in US and 95 international) | Worldwide Distribution - US (nationwide) and Puerto Rico; internationally to Australia, Canada, Colombia, France, Germany, Great Britain, India, Italy, Mexico, Netherlands, Norway, Pakistan, Poland, Russia, Singapore, Spain, Switzerland, Thailand, and Zimbabwe. | Open, Classified |  |  | 125 | Not_Computer
Z-1811-2015 | 71081 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015, Infusion pump. software version 9.17 |  |  | May 13, 2015 | 2015 | CareFusion 303, Inc. | An issue with the cancel functionality used during atypical infusion programming to cancel user inputted values. An infusion may start that is greater than or less than the hospital established limits for the specific medication. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated May 13, 2015 to customers. The letter identified the affected product, problem and the recommended actions to be taken. Customers are informed that CareFusion has released a software upgrade to address this issue. CareFusion will contact all affected customers within 60 days to initiate the scheduling process for the software upgrade installation. Customers with recall related questions are instructed to contact CareFusion Support Center at (888) 562-6018, 7am-4pm PT, supportcenter@carefusion.com. Customers with adverse event report questions are instructed to contact customer advocacy at (888) 812-3266, 24 hrs/day, 7 days a week, customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris System are instructed to contact Technical Support at (888) 812-3229, 6am-5pm PT, Mon-Fri, dl-us-inf-techsupport@carefusion.com. |  | 63 | 63 units | US Nationwide Distribution | Open, Classified |  |  | 63 | Not_Computer
Z-1693-2015 | 71306 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Neusoft NeuViz 64 Multi-slice CT Scanner System |  |  | March 08, 2014 | 2014 | Neusoft Medical System Co | The A-plane compensator is made of Teflon which will become fragile due to X-ray exposure passing through it. The thinnest part of the compensator is subject to radiation degradation and as a result has been found to fail prematurely. Failure is indicated by pin or ring artifact in the image | Radiation Control for Health and Safety Act | Neusoft planned action:1. Notification of all dealers and purchasers is to be made within 15working days of receipt of this letter in the manner specified in 21 CFR1003.21 and 1003.22. This office and the Food and Drug Administration(FDA) district office coordinator noted below are to be included in thenotification.2. Neusoft Medical Systems Co., Ltd will implement actions to monitor theexposures of the tube and when the exposure number is close to300,000 seconds, the A-plane compensator will be replaced. The Foodand Drug Administration (FDA) district office coordinator should benotified when the last A-plane compensator is replaced.3. The Customer Information Letter has been sent to customers as acorrective action. The maximum exposure number of 5 US systems is191830 seconds as of Apr.20, 2015, so the firm will monitor the datacontinuously and replace the A-plane compensators before the definedthreshold.4. Neusoft will replace all compensators that have exceeded the 300,000seconds free of charge and will replace all compensators that fail duringwarranty and the compensators failed less than 300,000 secondsthough out of warranty free of charge.You may proceed or continue with implementation of the CAP (subject to the conditions noted above) For further questions please call (281) 453-1205 |  | 5 | 5 installed in US | Worldwide Distribution - US Distribution including Puerto Rico and to the states of : IL, MA, TX and SC., and to the countries of : China, Malaysia, Lebanon, Ecuador, Egypt, United Arab Emirates,Kazashstan, Tunisia, Angola, Brazil, Uzbekistan, Lebanon, Cote D' ivoire, South Africa, Montenegro and Tajikistan. | Open, Classified |  |  | 5 | Not_Computer
Z-1810-2015 | 71268 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Navigation System II-Cart, 7700-100-000 with Articulated Arm Camera, 7700-103-001 and Articulated Arm Camera, 7700-103-001U |  |  | May 04, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Stryker initiated a device recall of the Navigation System II-Cart, Camera Articulated Arm because there is a potential that the weld seam between the main stud and the two flaps on the Articulated Arm was not welded in the correct location, which may result in a potential failure of the arm joint. | Process control | Stryker sent an Urgent Medical Device Recall Notification letter dated May 18, 2013 to Domestic customers notice via FedEx overnight and Sales Representatives via email. The International Subsidiaries will also be notified via email. Customers are instructed to immediately review the Recall Notification adn complete the following: Immediately check all stock areas and/or operating room storage for affected equipment found. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many affected items are currently in inventory. Please complete and return the BRF even if they dont have any affected product on hand. If customers have further distributed this product, they are asked to forward this letter and the attached BRF to all affected locations. Please indicate each location and serial number(s) on the BRF.Fax (866-521-2762) or email (kellyjo.davis@stryker.com) the completed BRF to Stryker Instruments Regulatory Department, Attn: Kelly Jo Davis. A Stryker Representative will contact the facility to set up a time to perform the additional inspection of the weld seam on the Navigation System II-Cart, Articulated Arm Camera. |  | 72 | 72 | Worldwide Distribution - US Nationwide in the states of AR, CO, FL, GA, IN, KS, KY, MA, MI, NC, NE, NJ, NY, OH, OK, PA, TX, WA, WY and the countries of Canada, Europe, and Asia. | Terminated | May 16, 2016 | 378 | 72 | Not_Computer
Z-1812-2015 | 71285 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Thoracic Grasper, 5 MM, for use with da Vinci S (IS2000) and SI (IS3000) Surgical Systems. |  |  | May 21, 2015 | 2015 | Intuitive Surgical, Inc. | The coating of the 5 mm Thoracic Grasper main tube may experience damage with the possibility of fragments of coating material dislodging and potentially falling into patient. | Nonconforming Material/Component | Urgent Medical Device Recall letters were sent to all consignees on May 21, 2015.The letter describes the affected product and issue involved. Additionally, the risk to health, as well as the actions to be taken were also discussed in the letter. Customers are to ensure that all affected personnel were informed of the notice. All affected product are to be identified and returned by contacting Customer Service for a RMA. The Acknowledgment Form is to be completed and returned using the instructions provided. If further information or support concerning this recall is needed, customers are to contact their local Clinical Sales Representative or Intuitive Surgical Customer Service. |  | 2436 | 2,436 units | Worldwide Distribution -- US, Belgium, Brazil, Canada, Chile, Colombia,, Denmark, France, Germany, India, Netherlands, Panama, Singapore, South Korea, Switzerland, Taiwan, Turkey, and United Kingdom. | Open, Classified |  |  | 2436 | Not_Computer
Z-1816-2015 | 70164 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BD Viper LT System, catalog number 442839 when used in conjunction with the BD ProbeTec Chlamydia trachomatis (CT) Qx Assay (cat. #442959) and the BD ProbeTec Neisseria gonorrhoeae (GC) Qx Assay (cat. #442842). In vitro diagnostic. |  |  | November 07, 2014 | 2014 | Becton Dickinson &amp; Co. | Potential for false positive Chlamydia trachomatis (CT) results and Neisseria gonorrhoeae (GC) results when processing the CT Qx and GC Qx Assays on the BD Viper LT Instrument. | Device Design | BD sent an Urgent Field Corrective Action letter dated November 2014, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer. Customers were advised: a. Include the date recall was initiated, if it is already underway or the date your firm plans to start the recalling process. b. How do you plan to notify all the consignees affected by this recall? (press release, letter, telefax, telephone, e-mail, visit, etc.) The customer was initially contacted via a phone call and subsequently provided with a written communication. c. How do you plan to monitor the number of consignees non-responding to the recall communication? The single US customer was contacted by phone and provided the written communication.d. How do you plan to do effectiveness checks of this recall? The affected instrument will be corrected.e. Date your firm ceased further distribution of the product(s). Product was not placed on hold. This issue does not affect new customers. f. How do you plan to store the recall product? BD has not requested that instruments or reagents be returned.g. How do you plan to dispose of the recall products? BD is not requesting that customers return the instruments or reagents.If you require further assistance or clarification, please contact the BD Technical Services Department at 1-800 638-8663. |  | 7 | 7 Viper LT Systems | Worldwide distribution to US (state of LA), Hong Kong, Austria, China, France, and the UK. | Terminated | May 09, 2016 | 549 | 7 | Not_Computer
Z-1814-2015 | 71329 | No | HDW | Contraceptive diaphragm and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | MILEX OMNIFLEX DIAPHRAGM SIZE 80 |  |  | May 20, 2015 | 2015 | CooperSurgical, Inc. | A diaphragm size 80 labeled box was incorrectly packaged with a size 85. | Labeling mix-ups | CooperSurgical notified consignees by letter on 5/18/15 via Fedex with confirmed delivery receipt.. Accounts requested to discontinue use and complete the attached Acknowledgement and Receipt Form for replacement. CooperSurgical provided further instruction on the disposition of the product at the time of replacementIf you have any further questions please feel free to contact 203.601.5200.. |  | 200 | 200 units | AK, CA, DC, FL, IL, ME, MT, NE, NJ, NM, NY, OK, OR, P A, VA, W A, and WIForeign: CANADA, ENGLAND, and POLAND | Terminated | January 20, 2016 | 245 | 200 | Not_Computer
Z-1817-2015 | 71398 | No | DRQ | Transducer signal amplifier and conditioner. | Cardiovascular | Cardiovascular | 510(k) | CLEARSIGN II Amplifier for LABSYSTEM PRO EP Recording System, 120 channels, using firmware version 2.08. The channel numbers in the amplifier are as follows: |  |  | June 02, 2015 | 2015 | Boston Scientific Corporation | Boston Scientific has received complaints indicating that, when using the blood pressure (BP) channels on the CLEARSIGN II Amplifier, the surface Electrocardiogram (ECG) channels become over-written to a variable degree, with the result that it appears shifted from baseline on the system's output screen. This may, in turn, manifest as an uninterpretable ECG signal in the affected channel. | Software design | Boston Scientific sent an Important Field Safety Notification letter to all affected customers beginning June 3, 2015, by 24 hour overnight delivery. The letter identified the product the problem and the action needed to be taken by the customer. INSTRUCTIONS:1. Please read carefully the Field Safety Notice letter and immediately post this information in a visible location near the product to ensure it is easily accessible to all users of the device.2. Please complete the attached Verification Form even if you do not have any product to return.3. When completed, please return the Verification Form to your local Boston Scientific Office to the attention of ÂżCustomer_Service_Fax_NumberÂż on or before DATE 2015.4. If you have products to return, and once a loaner unit is available, a Boston Scientific representative will coordinate with you to package the product being returned using a special shipping container.Your Competent Authority is being notified of this Field Safety Notice.We regret any inconvenience that this action may cause and we appreciate your understanding as we take action to ensure patient safety and customer satisfaction. If you have any questions or would like assistance with this Field Safety Notice, please contact your local Sales Representative. For further questions please call (978)-805-3200 |  | 43 | 43 units | No US distribution. Worldwide Distribution to the countries of : Austria, Belgium, France, Germany, Great Britain, Greece, Italy, Netherlands, Poland, Romania, South Africa, Spain, Switzerland and Turkey. | Open, Classified |  |  | 43 | Not_Computer
Z-1815-2015 | 71343 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Camino Intracranial Pressure Monitoring Catheter with Licox IMC Bolt Fitting, Catalog No. 110-4L |  |  | May 27, 2015 | 2015 | Integra LifeSciences Corp. | Integra LifeSciences is recalling the Camino Intracranial Pressure Monitoring Catheters because they may have been incorrectly assembled. Specifically, the o-ring in the bolt adapter may be in the wrong location and as a result, there is a potential CSF/blood could leak from the bolt adapter. | Under Investigation by firm | Integra sent an Urgent Medical Device Recall letter dated May 27, 2015, to all customers who received the Camino Intracranial Pressure Monitoring Catheters because they may have been incorrectly assembled. Specifically, the o-ring in the bolt adapter may be in the wrong location and as a result, there is a potential CSF/blood could leak from the bolt adapter. The letter informs the customers of the problems identified and the actions to be taken. Customers were instructed to complete the customer response form and return it by email or fax to: FCA3@integralife.com or FAX to 1-609-275-9445. Customers with any questions regarding the instructions are instructed to contact Customer Service at 1-888-601-0203. For questions regarding this recall call 609-936-6822. |  | 1218 | 1218 units | Worldwide Distribution - US (Nationwide) and Internationally to ARGENTINA, BRAZIL, BULGARIA, CANADA, COLUMBIA, COSTA RICA, FRANCE, GERMANY, ITALY, MEXICO, GERMANY, ITALY, MEXICO, PANAMA, SOUTH AFRICA, SPAIN, SWITZERLAND, UNITED ARAB EMIRATES, AND UNITED KINGDOM | Terminated | December 09, 2015 | 196 | 1218 | Not_Computer
Z-1813-2015 | 71304 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | CADDÂż- Solis Ambulatory Infusion Pump, Model 2111 &amp; 2112, Version 1.0, Version 2.0 and Version 3.0 &amp; Model 2120, Version 1.0. |  |  | May 26, 2015 | 2015 | Smiths Medical ASD, Inc. | Smiths Medical has become aware of an issue with an intermittent occurrence of binding of the locking assembly on some CADDÂż-Solis pumps. Binding of the Cassette/ Keypad Lock can occur after latching the CADDÂż-Solis Medication Cassette Reservoir or Administration Set disposable to the pump. When binding occurs, it can prevent the key from fully rotating the Cassette/ Keypad Lock to the locked (e | Nonconforming Material/Component | Consignees were sent on 5/26/2015 a Smiths Medical "Urgent Medical Device Recall Notice" letter dated 26 May 2015. The letter described the problem and the device involved in the recall. The letter provided "Advise on Action to be taken by the User" and " Transmission of this Urgent Medical Device Recall Notice". Consignees were requested to complete and return the Confirmation Form. For questions they can contact Smiths Medical's Customer Service Department at 1-800-258-5361, press 2, then press 1. |  | 4426 | 4426 (2427 US, 1999 OUS) | Worldwide Distribution-US (nationwide) including the states of AL, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, LA, MA, MI, MN, MS, MO, NE, NM, NY, NC, ND, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, and WI, and the countries of AT (Austria), AU (Australia), BE (Belgium), CA (Canada), CH (Switzerland), DE (Germany), DK (Denmark), FR (France), GB (Great Britain), HK (Hong Kong), LB (Lebanon), NL (Netherlands), NZ (New Zealand), QA (Qatar), SA (Saudi Arabia), and SE (Sweden). | Open, Classified |  |  | 4426 | Not_Computer
Z-1820-2015 | 70675 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Steriseal-Cannula Ophthalmic; |  |  | March 02, 2015 | 2015 | Hill-Rom, Inc. | The metal cannula attached to the handpiece can rotate and become dislodged from the plastic handpiece. | Use error | Customers were notified via telephone and e-mail by 03/02/2015. The firm is requiring customers to complete a submitted response form and certificate of destruction for recalled devices by 03/31/2015. If recalled devices were further distributed the firm is asking customers to conduct a sub recall and forward notice of the recall. Public contact is Aspen Medical Europe Ltd, QA Department, Lynda McConaghy-QARA Manager, Lynda.McConaghy@aspenmedicaleurope.com , or +44 (0)1527 587716. |  | 66691 | 66691 units | Worldwide Distribution-US (nationwide) including the states of MI &amp; PA, and the countries of Belgium, England, Italy, Czech Republic, Turkey, Greece, Ireland, Scotland, Spain, and Lebanon. | Open, Classified |  |  | 66691 | Not_Computer
Z-1818-2015 | 71132 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Viking M, L, and XL Mobile Lifts. |  |  | March 20, 2015 | 2015 | Hill-Rom, Inc. | Complaints (including one reported death) allegedly of the lift arm drifting down suddenly. If the lift arm assembly is manually lifted, the actuator can become damaged and get stuck in the highest position. If a patient is lifted into the sling while the actuator is stuck, there is a potential risk of a free fall of the patient, resulting in minor or potentially catastrophic injuries. | Labeling design | On 3/20/2015, URGENT Field Safety Notice notifications, dated 3/19/2015 were sent to US customers by e-mail and or courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers are forward a copy of the letter to any other appropriate facility personnel. Additionally, the attached form should be completed and returned to Hill-Rom. Upon receipt, the firm will contact customers to schedule an inspection of the potentially affected devices. For questions, call 800-445-3720 (Mon - Thurs, 8 am - 6:30 pm; Friday, 8 am - 6 pm). |  | 24798 | 24, 798 | Worldwide Distribution -- AE, AO, AR, AT, AU, AZ, BD, BE, BG, BH, BR, CA, CH, CL, CN, CO, CS, CZ, DE, DK, EG, ES, FI, FR, GR, HK, HR, HU, ID, IE, IL, IN, IQ, IR, IS, IT, JO, JP, KR, KW, LB, LI, LT, LU, LV, MA, MT, MX, MY, NL, NO, NZ, OM, PH, PL, PT, QA, RO, SA, SE, SG, SI, SK, SY, TH, TN, TR, TW, UK, US, VE, and ZA. | Terminated | August 04, 2017 | 868 | 24798 | Not_Computer
Z-1821-2015 | 71332 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Sure T, contact, contact detach, neria, neria detach, neria multi and thalaset |  |  | May 21, 2015 | 2015 | Unomedical As | Unomedical has found that in rare cases the steel needle can break during use, interrupting the delivery of medication including insulin. | Nonconforming Material/Component | Each consignee was sent an "Urgent Field Safety Notice Regarding Steel Cannula Infusion Sets" on 5/21/2015 via Express Mail. This notification identified the affected product as well as the reason for the recall. It also contained instructions on how to use the product. Customers were to contact their health care provider if they suspected that a needle had broken off and remained under the skin. Diabetic patients were to monitor their blood sugar and non diabetic patients were to monitor symptoms which would indicate drug delivery had failed. Questions should be directed to the persons identified in the letter. |  | 22976870 | 22,976,870 units | Worldwide Distribution -- US, including the states of CA, IL, MN, and PA; and, the countries of Australia, Bahrain, Belgium, Canada, China, Cypress, Denmark, Estonia, France, Germany, Greece, Hong Kong, Hungary, Israel, Italy, Kuwait, Lebanon, Malaysia, Maldives, Oman, Schweiz, South Africa, Spain, Sweden, Switzerland, Netherlands, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 22976870 | Not_Computer
Z-1819-2015 | 71262 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ORTHOFIX Centurion POCT System Drill Guide PN 69-1010 and PN 69-1012. Drill Guides are used to interface with various drill tips during the drilling procedure of bone preparation. |  |  | May 11, 2015 | 2015 | Orthofix, Inc | There is a possibility that the thumb button remains in the depressed (unlocked) position allowing the threaded sleeve (drill stop) to unintentionally rotate and change the depth setting. | Device Design | Orthofix sent an Urgent Medical Device Product Advisory Notification letter dated May 11, 2015, to the affected distributors and hospitals. The consignees were directed to inspect thier inventory and perform checks before surgical use of the drill guides. Customers were instructed to provide a copy of the Product Advisory Notification letter if product was further distributed. Customers with questions were instructed to contact their local Orthofix representative and to for returns they should contact Orthofix customer service representative at 888-298-5700. |  | 36 | 36 units | Nationwide Distribution including CA, VA, FL, ID, CO, OR, PA, OK, SC, and TX. | Terminated | February 17, 2016 | 282 | 36 | Not_Computer
Z-1827-2015 | 71160 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NaturaLyte Liquid Bicarbonate Concentrate (Dialysate Concentrate for Hemodialysis (liquid), 6.4 Liter Bottle |  |  | May 20, 2015 | 2015 | Fresenius Medical Care Holdings, Inc. | Bacterial contamination. | Process control | Fresenius Medical Care sent an Urgent Medical Device Recall letter dated May 15, 2015, to all affected customers. Users were requested to Immediately examine stock to determine whether they have any NaturalyteÂż Liquid Bicarbonate Concentrate of the lots." If any product of these lots is found, discontinue use immediately." Place all units in a secure, segregated area." If affected product was on the machine prior to patient treatment, perform a[Heat Disinfect] program." Your dialysis schedule should not be interrupted. If interruption of yourdialysis schedule is expected, please discuss your options with your healthcare provider." Contact FMCNA Customer Service Team at 1-800-323-5188 forinstructions on how to return the recalled product." Promptly fill out and return the attached reply formAdditional medical concerns or questions, please contact Medical Informationand Communication: 855-616-2309 or Website: www.fresenius-medinfo.com.For questions regarding this recall call 800-662-1237. |  | 1856619 | 1,856,619 Bottles | Nationwide Distribution | Open, Classified |  |  | 1856619 | Not_Computer
Z-1826-2015 | 71519 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | Merit Medical Prelude Pro Sheath Introducter, 6F (2.0mm) 11cm, REF PRO-6F-11-038, Sterile EO |  |  | June 04, 2015 | 2015 | Merit Medical Systems, Inc. | Some of these introducer sheaths were packaged with a 0.035 inch dilator, rather than the labeled 0.038 inch. This affects one lot, H767887, of catalog number PR0-6F-11-038. | Packaging process control | Merit Medical sent an Urgent Product Recall letter on June 4, 2015, to all affected consignees. Consignees were instructed to immediately quarantine any affected devices and discontinue use, ensure all personnel to whom devices were distributed are made aware of this field action, and instructions to contact their Merit representative. An email notification was sent to Merit sales representatives with accounts that had received affected product. The affected sales representatives were provided a Customer Response Form applicable to their account and a representative copy of communications sent directly to their affected accounts. Merit sales representatives were instructed to visit the account and assist with the return of the affected units back to the Merit Field Assurance Department. Customers with questions were instructed to contact their Merit Sales Representative or call (801) 208-4365. |  | 1445 | 1,445 | Worldwide Distribution - US including TN, NE, MA, SC, WI, UT and Internationally to Belgium, Switzerland, United Kingdom, and Germany. | Terminated | October 26, 2015 | 144 | 1445 | Not_Computer
Z-1823-2015 | 71357 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Fortex" Pedicle Screw System Continuous Radius Rod 5.5 - 50mm, X-Spine Systems, Inc., 452 Alexandersville Road Miamisburg, OH 45342 Phone: 937-847-8400 |  |  | April 07, 2015 | 2015 | X Spine Systems Inc | The Continuous Radius Rod should measure 50mm in length but actually measures 45mm in length. | Other | X-Spine sent an Urgent Medical Device Recall Letter dated April 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken:1. The affected hospital and surgeon should be notified of this recall immediately.2. Complete the attached Medical Device Recall Return Response Form and return it with any affectedproduct in inventory to X-Spine Systems immediately.a. Reference RMA # 5410b. Use enclosed shipping label3. No action should be taken with product that has been implanted.Please feel free to contact me at 800.903.0640 ext. 137. |  | 75 | 75 units | Worldwide Distribution - US Distribution to the states of : FL, ID, IL, KS, LA, MI, MO, NJ, OH, PA and TX., and to the countries of Austria, Germany and Italy. | Terminated | October 01, 2015 | 177 | 75 | Not_Computer
Z-1825-2015 | 71449 | No | IOL | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | GE Healthcare, Treadmill T2100. |  |  | March 12, 2015 | 2015 | GE Healthcare | GE Healthcare has recently become aware of a potential safety issue due to uncontrolled motion of the walking belt on your T2100 Treadmill. | Device Design | General Healthcare sent an Urgent Medical Device Correction letter (GEHC Ref# 30074) dated March 12, 2015, to all affected consignees. The letter was addressed to Director of Biomedical Engineering, Chief of Nursing and Healthcare Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 11423 | 11,423 (5,540 units US, 5,883 units OUS | Nationwide Distribution including Guam, DC, and Puerto Rico.OUS: VIET NAM, VENEZUELA, UZBEKISTAN, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TRINIDAD &amp; TOBAGO, THAILAND, TAIWAN EL SALVADOR, SRILANKA, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, RUSSIA, PORTUGAL PHILIPPINES, PERU, PARAGUAY, PANAMA, PAKISTAN, NEW ZEALAND, NICARAGUA, NEPAL, MYANMAR, MOROCCO, MEXICO, MARTINIQUE, MALAYSIA, LEBANON, KUWAIT, REPUBLIC OF KOREA, KAZAKHSTAN, JAPAN, JAMAICA, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, HON GONG, GERMANY, FRANCE, FINLAND, ETHIOPIA, EGYPT, ECUADOR, DOMINICAN REPUBLIC, COSTA RICA, COLOMBIA, CHINA, CHILE , CAYMAN ISLAND , CANADA, CAMBODIA, BRUNEI DARUSSALAM, BRAZIL, BOLIVIA, BHUTAN, BANGLADESH, BAHRAIN, AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA. | Open, Classified |  |  | 11423 | Not_Computer
Z-1822-2015 | 71257 | No | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ULTRA-FAST FIX Knot Pusher Suture Cutters |  |  | April 30, 2015 | 2015 | Smith &amp; Nephew, Inc. | Sterility of device maybe compromised | Package design/selection | Smith &amp; Nephew sent an Urgent - Product Recall 1st Notification letter dated April 27, 2015, via Federal Express on April 30, 2015. Users are requested to inspect inventory and locate any unused devices from the above listed product and batch numbers, and quarantine them immediately. Complete the the Inventory Return Certification and obtain a RGA and obtain replacement product.Contact Smith &amp; Nephews Field Action Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to obtain a return authorization (RA) number |  | 202279 | 202,279 units | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico and to the countries of : Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, Chile, China, Colombia, Costa Rica, Dubia UAE, Germany, Denmark, Spain, Finland, France, United Kingdom, Greece, Guam, Hong Kong, Israel, India, Italy, Japan, Korea, Mexico, Malaysia, Netherlands, Norway, Portugal, Qatar, Sweden, Singapore,Thailand, Turkey, Taiwan, Venezuela and South Africa. | Open, Classified |  |  | 202279 | Not_Computer
Z-1824-2015 | 71313 | No | HIB | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | McKesson Medi-Pak Vaginal Speculum, 10/bag |  |  | April 13, 2015 | 2015 | CooperSurgical, Inc. | Small sized specula were incorrectly packaged in printed bags for the medium sized speculum | Packaging change control | CooperSurgical sent an Urget Medical Device Recall letter dated April 13, 2015, to all affected customers. The letter identified the product, the problem, and the actions to be taken by the customer. Customers were instructed to discontinue use, discard, or return the affected to product to Cooper Surgical. Customers were asked to complete the attached Acknowledgment and Receipt Form for replacement. Customers with questions were instructed to call 203-601-5200. |  | 9400 | 9400 units | Nationwide Distribution including IL, FL, GA, MA, MO, TX, and WA. | Terminated | February 05, 2016 | 298 | 9400 | Not_Computer
Z-1828-2015 | 71196 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Custom Cardiovascular Procedure Kits; Over-Pressure Safety Valves. |  |  | May 06, 2015 | 2015 | Terumo Cardiovascular Systems Corp | The Overpressure Safety Valve is not cleared with an indication for use on the arterial side of the extracoporeal bypass circuit. It is intended for use only on the venous side. There is a remote possibility of minimal blood loss, fluid weeping, or dripping on the user. | Component design/selection | Consignees were notified of the recall by phone beginning on 5/7/2015 and a Urgent Medical Device Recall letter, dated 5/7/2015. The letter identified the affected device, reason for recall, potential hazard, and correction. Customers were asked to review the notice and ensure that all users are aware of the issue. In addition, they are to confirm the quantity of procedure kits for return; and, complete and return the Customer Response Form. Customers can call Terumo CVS Customer Service at 1-800-521-2818 with questions. |  | 163 | 163 devices | Distributed to the states of MO and FL. | Terminated | July 27, 2015 | 82 | 163 | Not_Computer
Z-1829-2015 | 71450 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Surface Applicator Set with Leipzig-style Cone. |  |  | May 18, 2015 | 2015 | Varian Medical Systems Inc | Varian has discovered that there is a discrepancy in the absolute dose rate given with the Leipzig-style Surface Applicator's Instructions for Use (IFU): Dose CharacterizationGM11010080 2012-09-06. The actual dose rate of the applicator is approximately 14% higher than the rate published within the IFU. Use of the dose rate as provided by the IFU without modification or independent confirmat | Process design | All the consignees affected by this recall will be notified by:The Urgent Medical Device Correction Letter/ Field Safety Notification will bedistributed to all affected users, with a description of the problem.The Letter will also be distributed to the Varian Sales, Marketing, and Serviceorganizations, informing them of the issue.Additionally, the letter will be posted to the customer support site:http:/ /www.MyVarian.com. |  | 46 | 46 | Worldwide Distribution. | Terminated | January 26, 2016 | 253 | 46 | Not_Computer
Z-1830-2015 | 71287 | No | KGY | Organ bag. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LapSac Surgical Tissue Pouch. |  |  | May 14, 2015 | 2015 | Cook Inc. | Includes Information not stated in the intended use. | Labeling Change Control | Cook Medical sent an "Recall Communication Updated Instructions for Use" letter dated May 15, 2015, to all affected customers . The letter identified the product the problem and the action needed to be taken by the customer. Cook has requested that customers do the following: The intended use for the device is being modified from "Used to isolate tissue" during, or prior to, surgical removal and/or morcellation to "Used to isolate tissue" during, or prior to, surgical removal .A contraindication being added is as follows:" Not for use with power morcellators.The warnings are being modified to read as follows:" The LapSac Surgical Tissue Pouch is not intended for use with sharp surgical instruments." Care should be taken to avoid damaging, puncturing or tearing the LapSac Surgical Tissue Pouch." Use caution when removing potentially malignant or harmful tissues, to isolate contamination from surrounding tissue.The instructions for use are being modified to read as follows:1. Following aseptic technique, introduce the LapSac Surgical Tissue Pouch into the abdominal cavity.2. Place tissue to be removed from the body into the pouch. 3. Using the drawstring, carefully close the LapSac Surgical Tissue Pouch. Please note that if excess force is used, the drawstring may break or tear the LapSac.4. Remove the LapSac Surgical Tissue Pouch from the body, taking care not to spill contents of the bag.Please report any Adverse Event to Cook Medical Customer Relations (800) 457-4500 or 1-812-339-2235. Monday through Friday between 7:30 a.m. and 5:00 p.m. Eastern Daylight Time or email at CustomerrelationsNA@cookmedical.com; and to FDA. The customer notice must be passed on to appropriate personnel within each organization or to any organization where the potentially affected devices have been transferred.If customers have further questions regarding labelling changes for this device, please contact Cook Sales Representative or Cook M |  | 123184 | 123,184 devices were distributed in commerce | Worldwide Distribution - US (nationwide) including PR and to the states of: AZ, AK, AL, AR, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MS, MO, ME, MT, NC, ND, NE, NH, NJ, NY, NM, OH,OK,OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY., and to the countries of : China, Korea, Denmark, Mexico, New Zealand, India and Colombia. | Open, Classified |  |  | 123184 | Not_Computer
Z-1831-2015 | 71385 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SERFAS 90 degree Energy Probe, Part Number 279-350-101; |  |  | June 03, 2015 | 2015 | Stryker Endoscopy | Stryker Endoscopy is recalling all non expired SERFAS 90 degree Energy Probes due to reports of fragments of the probe breaking off into the patient. | Device Design | Stryker sent an Urgent Medical Device Recall letter dated June 5, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:1. Inform individuals within their organization who need to be aware of this device removal.2. Review inventory of lots of part number 279-350-101 and determine if you have the affected product (all non-expired devices) in stock. Response is required.3. If no product is found, complete acknowledgement form located on the Stryker Endoscopy recall website endorecall.stryker.com by logging in using the account number and zip code on this letter.4. If you do have product, segregate the product and call Stryker customer service at 1-800-624-4422 Option 3 to arrange for product return and issuance of credit.For questions regarding this recall call 408-754-2664. |  | 22063 | 22,063 devices | Worldwide Distribution - US (nationwide) and Internationally to US Argentina, Australia, Bolivia, Brazil, Chile, China, Colombia, Hong Kong, India, Italy, Japan, Republic of Korea, Malaysia, Netherlands, Poland, Romania, Singapore, South Africa, Spain, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | February 10, 2016 | 252 | 22063 | Not_Computer
Z-1809-2015 | 71254 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Lifesaver Single Patient Use Manual Resuscitator |  |  | May 14, 2015 | 2015 | Teleflex Medical | The intake port may be blocked which can cause the bag to fail to fill. | Nonconforming Material/Component | Teleflex Medical sent an Urgent Medical Device Recall Notification letter dated May 14, 2015, to all consignees asking them to immediately discontinue distribution and quarantine the recalled product. The letter also requested a sub-recall of the product. The recall letter also included a response form which is to be returned to Teleflex by emailing it to recall@teleflex.com or fax it to 1-855-419-8507, Attn: Customer Service. Customers with questions were instructed to contact their local sales representative or Customer Service at 1-866-246-6990.For questions regarding this recall call 610-378-0131. |  | 9333 | 9,333 units | Worldwide Distribution - US Nationwide in the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, IL, IA, KS, LA, MA, MI, MN, MS, MO, NE, NH, NY, NC, OH, OK, PA, RI, SD, TN, TX, VA, WA, including Puerto Rico and the countries of Australia, Bahamas, Canada, Guatemala, and Mexico. | Open, Classified |  |  | 9333 | Not_Computer
Z-0351-03 | 25150 | [6 empty] | Medtronic Cardioblate Radio Frequency Ablation Pen, model number 60811C47 |  |  | November 11, 2002 | 2002 | Medtronic Cardiac Surgery Technologies | The radio frequency ablation pens contained unreacted epoxy resin which was observed to leak out of the back-end of the device. | Other | A letter dated November 11, 2002 was sent to the affected consignees who were requested to quarantine any unused pens for retrieval by Medtronic representatives. |  | 148 | 148 pens | All of the consignees were in foreign countries in Europe, Asia and Latin America. | Terminated | July 02, 2004 | 599 | 148 | Not_Computer
Z-0494-05 | 31217 | No | RAD |  |  |  |  | X-Ray High Voltage Generator, Model KXO-80G used on Image Intensified Fluoroscopoc X-Ray System DUA-450D |  |  | September 07, 2004 | 2004 | Toshiba American Med Sys Inc | Non-compliance with Radiation Standards. Incorrect display of the radiographic time, due to software anomaly. | Other | Firm sent recall letters to customers and will be following up with field corrections which will be made on-site by the firm's own employees. |  | 494 | 494 | Nationwide. | Terminated | February 10, 2012 | 2712 | 494 | Not_Computer
Z-0352-03 | 25226 | [6 empty] | Surgical Mesh |  |  | December 11, 2002 | 2002 | Synovis Surgical Inovation Div. of | A gap in the exterior seal may lead to sterility concerns with the interior pouch. | Other | The sites and/or distributors that received suspect devices were contacted by telephone on 12/11/02 and given instructions on how to return/replace the devices. In addition, a letter dated 12/11/02 was sent to each consignee. |  | 28 | 28 devices | Product shipped nationwide, no foreign consignees | Terminated | May 28, 2003 | 168 | 28 | Not_Computer
Z-0792-03 | 26046 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemstarÂż 7 Therapy I.V. Infusion Pump, List#: 13000-04; 2.9 Software Version |  |  | April 01, 2003 | 2003 | Abbott Laboratories | Products contain software deficiencies which could allow over infusion thereby administering excess medication. | Other | On 4/1/03, the firm issued letters via Federal Express Overnight Delivery to all its direct consignees, informing them of the affected devices and providing instructions on the recall. |  | 1310 | 1310 units | Product received nationwide distribution to approx. 120 direct consignees. The firm identified 3 Govt. accounts: (1) Nthrn VA Comm. Hospital, 601 S. Carlin Spg. Rd., Arlington VA 22204, (2) Veterans Admin Med Ctr., 4100 W. 3rd St., Dayton OH 45428, (3) GW University Hospital, 901 23rd St. NW # 5306N, Washington DC 20037. The recall was appropriately extended to the user level; i.e., the hospital/medical centers, physicians and nurses who received the recalled product.There is no known Canadian distribution. | Terminated | October 09, 2003 | 191 | 1310 | Not_Computer
Z-0412-03 | 4397 | [6 empty] | Cryo-Tip T-1910 Exo-Endocervical, Part Number 90501 (as labeled) |  |  | June 19, 2002 | 2002 | ACMI Corporation | Derm Bevel Cryo-Tips, Part Number 90512, were mislabeled as Exo- Endocervical Cryo-Tips, Part Number 90501. | Other | The recalling firm telephoned the 2 consignees on June 19, 2002, informed them of the recall and asked them to return the product they had. |  | 7 | 7 Cryo-tips | Seven units had been distributed to five consignees, but three consignees had returned four units on 5/22, 5/31 and 6/13/02 prior to the recall because they were mislabeled and not the devices needed by the consignees. This left 3 units at 2 consignees in the States of California and Texas. | Terminated | January 01, 1990 | -4552 | 7 | Not_Computer
Z-0382-03 | 3265 | [6 empty] | Medtronic Model 550 Bio-Console Extracorporeal Blood Pumping |  |  | September 25, 2001 | 2001 | MEDTRONIC PERFUSION SYSTEMS | Motor controller printed circuit boards with a capacitor installed in reverse can result in a shut down of the product pump. | Other | September 25th, 2001 letter about the problem followed by field correction of the units. |  | 126 | 126 units | The devices were distributed to 30 domestic consignees (mostly hospitals) located nationwide in the United States. The product was also distributed to 4 Medtronic foreign subsidiaries and from there to foreign consignees located worldwide. | Terminated | January 01, 1990 | -4285 | 126 | Not_Computer
Z-0353-03 | 4411 | [6 empty] | AXIS/IRIX Gamma Camera System, controlled with Odyssey computers, |  |  | October 08, 2001 | 2001 | PHILIPS MEDICAL SYSTEMS | A broken gantry wire may prevent the contact sensors from stopping motion of the unit. | Other | The firm sent a "Manditory Service Letter" (ML#437), dated 10/8/2001, to their Field Service Engineers, instructing them to immediately contact their accounts and schedule the units for repair. A software upgrade, version 8.9B, will be installed on all affected units in the field. The 8.9B retrofit kit (part #53000011) will be sent out after a service order has been generated. |  | 653 | 653 devices installed | The units were shipped to consignees located nationwide and worldwide. The following U.S. VA Medical Centers are involved: BUFFALO, NY; HINES, IL; CHICAGO, IL; MONTGOMERY, AL; TUCSON, AZ; SAN DIEGO, CA; SAN FRANCISCO, CA; LONG BEACH, CA; WILMINGTON, DE; DECATUR, GA; IOWA CITY, IA; N. CHICAGO, IL; BILOXI, MS; CLEVELAND, OH; CINCINNAYI, OH; DAYTON, OH; BIG SPRING, TX; HUNTINGTON, WV; ALBUQUERQUE, NM; COLUMBIA, SC; DENVER, CO; DES MOINES, IA; PHOENIX, AZ; PRESCOTT, AZ; FRESNO, CA; LEXINGTON, KY; SAN JUAN, PR; OKLAHOMA CITY, OK; and ENRICO, CA. | Terminated | April 15, 2004 | 920 | 653 | Not_Computer
Z-0407-03 | 4607 | [6 empty] | EXPO 32 Software versions 1.0, 1.1B, and 1.2 |  |  | May 20, 2002 | 2002 | BECKMAN COULTER, INC. | Software contains errors that might affect Cytometry instruments | Other | Recall letters were sent May 20, 2002 which indicate potential problems that would be noted depending on the equipment being used and communicates instructions to prevent the problems until the next version of the software is issued. |  | 33 | 33 | Only domestically to the states of CA, FL, HI, KS, KY, MA, MD, MI, NC, NJ, OH, OR, SC, TX, UT, WA. | Terminated | November 22, 2004 | 917 | 33 | Not_Computer
Z-0471-03 | 4795 | [6 empty] | Kendall Argyle 3.7 Fr Polyurethane Umbilical Vessel Catheter Kendall Argyle 3.7 Fr Polyurethane Umbilical Vessel Catheter |  |  | October 02, 2002 | 2002 | TYCO HEALTHCARE GROUP LP/KENDALL | Lumen of Catheter may not accomodate the designated blood gas sensor | Other | Tyco/Kendall notified Direct Accounts (Hospitals and Distributors) on 10/2/02 by letter advising accounts to examine inventory, quarantine recalled and return. Distributors are requested to subrecall.On 1/22/03, the firm extended the recall to ALL Lots which included the addition of one lot number: 426038. |  | 734 | 734 cases (10/case) | Nationwide. AR, GA, IL, IN, MO, MN, NC, NJ, SC, TX, WI, VA.International: UK | Terminated | May 02, 2012 | 3500 | 734 | Not_Computer
Z-0445-03 | 4403 | [6 empty] | IMx Glycated Hemoglobin Reagent Sets, list 1A86-88; |  |  | July 05, 2002 | 2002 | ABBOTT LABORATORIES, INC. | Low Control reads high, out of insert range | Other | Recall letters dated 7/2/02 were sent to all customers receiving IMx Glycated Hemoglobin Reagent Sets on 7/5/02. Customers were requested to cease use of the affected lots and destroy the remaining product. |  | 38877 | 38,877 kits | The sets were distributed nationwide to 1101 hospital/laboratory accounts, including 22 direct sale government accounts. The sets were also shipped through Abbott subsidiaries internationally to Canada, Bermuda, Korea, India, Hong Kong, Honduras, Taiwan, Bahamas, Japan, Argentina, Costa Rica, Nicaragua, Mexico, Panama, Brazil, Jamaica, Guatemala, Colombia, Chile, Uruguay, Dominican Republic, El Salvador, Great Britain, Pakistan, Guyana, Arabia, Haiti, Trinidad, Grenada, St. Lucia, Antigua, Curacao, Aruba, Tortola, Anguilla, Cayman, Virgin Islands and Australia. | Terminated | January 01, 1990 | -4568 | 38877 | Not_Computer
Z-0395-03 | 4719 | [6 empty] | Digene's Rapid Capture System (RCS) robotic microplate processor, |  |  | September 11, 2002 | 2002 | DIGENE CORP. | IVD instrument may measure incorrect sample volume for clinical testing | Other | Recalling firm notified consignees by letter dated 9/11/2002. Firm advises in the letter that users should not use RCS to process cervical brush specimens until firm has adjusted the instrument. |  | 5 | 5 units | Product distributed to end users in MD, IN and OR and to foreign consignee in Japan.US Gov't consignee:Oregon Public Health Lab1717 Southwest 10th St.Portland, OR 97201Foreign consignee:Mitsubishi8-5-1 Chuo, Ami-machiInashiki-gun, Ibaraki300-0332 Japan | Terminated | January 01, 1990 | -4636 | 5 | Not_Computer
Z-0334-03 | 4713 | [6 empty] | HomeChoice and HomeChoice PRO Automated Peritoneal Dialysis Systems; Manufactured for Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A., Made in U.S.A.; the following models are affected: |  |  | September 19, 2002 | 2002 | BAXTER HEALTHCARE CORP. | possible overfill | Other | Safety Alert letters dated 9/19/02 were mailed to all Home Care Coordinators on 9/19/02 and to all Home Dialysis Patients on 9/20/02, informing them of the possible overfill situation and advising them that if a manual drain is required during fill, to make sure that the manual drain is completed without interruption in order to prevent an overfill situation. Addendum pages were provided for the patient Help Cards, Patient At-Home Guide and Trainer's Guide to include the warning not to stop or bypass a manual drain during fill. |  | 12667 | 12,667 units | Nationwide to 892 hospitals, 88 VA/government accounts, 1 dealer, 9 physicians, 5 home care suppliers, 2,041 dialysis centers and 14,707 home patients. The device was also distributed internationally through Baxter subsidiaries in Canada, Japan, Argentina, Australia, Belgium, Brazil, Chile, China, Colombia, Czech Republic, Guatemala, Honduras, Hong Kong, India, Israel, Italy, Jordan, Lebanon, Malaysia, Malta, Mexico, New Zealand, Oman, Panama, Peru, Philippines, Qatar, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Taiwan, Thailand, Turkey, United Arab Emirates and the United Kingdom. | Terminated | July 16, 2015 | 4683 | 12667 | Not_Computer
Z-0406-03 | 4712 | [6 empty] | Myocardial Temperature Probe, YSI 400 Series, 15mm, packaged in |  |  | September 17, 2002 | 2002 | DEROYAL SURGICAL | Surgical device is mislabeled with incorrect size. | Other | Firm notified consignees 9/17/2002 by certified letter. The letter instructed to return product to firm's distribution center for replacement. Firm included a response form for consignees to fax numbers of product to be replaced. The letter additionally instructed distributors to further notify their customers or provide customer contact information. |  | 3000 | 3000 units | Product was distributed to health care facilities and distributors nationwide | Terminated | January 01, 1990 | -4642 | 3000 | Not_Computer
Z-0383-03 | 4419 | [6 empty] | BD DTX Plus Pressure Monitoring Kits and Extension Sets. Packaged |  |  | July 08, 2002 | 2002 | BECTON DICKINSON INFUSION THERAPY SYSTEMS, INC. | Bond between the tubing and female or male luer fitting may fail. | Other | On 7/8/2002, all consignees were notified by telephone and FAX. |  | 8572 | 7122 pressure monitoring kits, 1450 extension sets | Nationwide.Domestic distribution to AR,CA,CO,LA,MO,NJ,NY,OH,OR,PA, RI,TN,VA, WA,WV. Includes VA facility in OH. | Terminated | January 01, 1990 | -4571 | 8572 | Not_Computer
Z-0329-03 | 4792 | [6 empty] | Trade Name: ABC Fixation Pin. System Name: ABC Cervical Plating |  |  | September 06, 2002 | 2002 | AESCULAP, INC. | pin manufactured with sharp tip and not hardened, pin may be susceptible to bending | Other | The recalling firm issued recall leters dated 9/11/02 to the direct accounts via certified mail. Recall notifications dated 9/11/02 were also sent the their sales representatives. |  | 216 | 216 kits | The kits were shipped to hospitals in CA, GA, IN, MN, MS, NJ, PA, SC, TN, TX, UT, and WA. The kits were also shipped to 16 sales representatives. There were no foreign and/or government/military accounts. | Terminated | January 01, 1990 | -4631 | 216 | Not_Computer
Z-0331-03 | 4853 | [6 empty] | Stryker Wedge turning frame stretcher, model 965. 965 |  |  | October 23, 2002 | 2002 | STRYKER CORP., | Weld failure-anterior and posterior frame may dislodge from stretcher litter | Other | The firm sent each customer an undated urgent device correction letter on October 23, 2002 asking them to remove the stretchers from service until corrected. Stryker service technicians will visit each account with upgrade kits around November 2002. |  | 29 | 29 | Arkansas, District of Columbia, Georgia, Illinois, Louisiana, Maine, Missouri, Ohio, Utah, West Virginia, Hong Kong, Japan, Latin America, Sweden and Switzerland.U.S. GOVERNMENT ACCOUNTSWalter Reed Army Medical Center, Bldg 2 Room 2E07A, Washington, DCFort Gordon, Fort Gordon, GAHill AFB, Hill AFB, UT | Terminated | January 01, 1990 | -4678 | 29 | Not_Computer
Z-0333-03 | 4839 | [6 empty] | Cryolife Valve Allograft. |  |  | October 04, 2002 | 2002 | CRYOLIFE, INC. | Donor tissue had possible sepsis secondary to pyogenic brain abscess. |  | The implanting physician was notified of the circumstances via letter on/about October 4, 2002 |  | 1 | One individual heart valve allograft. | One consignee located in INJohn Fehrenbacher, M.D., Methodist Professional Center, 1801 North Senate Blvd., #755, Indianapois, IN 46202-5300. | Terminated | January 01, 1990 | -4659 | 1 | Not_Computer
Z-0328-03 | 4828 | [6 empty] | Epidural Catheterization Kit, model number AK-05503L. The kit is |  |  | October 15, 2002 | 2002 | ARROW INTL., INC. | mislabeled - package may contain incorrect gauge catheter | Other | The recalling firm issued recall letters dated 10/14/02 informing their customers of the problem and the need to return it. The product label states that a 19 gauge needle is part of the kit when a 20 gauge needle is actually in the kits for this recalled lot. |  | 720 | 720 kits | The kits were sold to two distributors in IL, eight hospitals in CA, IN, KY, and OH; and one government account in NE. There are no foreign accounts.The government/military account is VAMC, Omaha, NE. | Terminated | January 01, 1990 | -4670 | 720 | Not_Computer
Z-0365-03 | 4825 | [6 empty] | LC+, LCV+, and LCA Angiographic X-ray Systems, Models 2168678, 2168678-2, 2219151, 2219151-2, 2219151-3, 2124824, 2168678-3, 2219150, 2219150-2, 2219150-3, 2107015, 2115906, 2168677, 2219152, |  |  | October 28, 2002 | 2002 | GENERAL ELECTRIC MEDICAL SYSTEMS | The motor gearshaft for the image intensifier may break allowing the image intensifier to slide down to its end stop. | Other | Field modification instructions were issued on 10/28/02 to GE field engineers who then corrected the devices at hospitals. |  | 1361 | 1361 systems | The affected products had been shipped to 610 hospitals located nationwide in the U.S. and to 675 foreign hospitals located worldwide. | Terminated | July 02, 2004 | 613 | 1361 | Not_Computer
Z-0397-03 | 4820 | [6 empty] | Medex Vented Flow Clip Nitroglycerin Set, Pv22ml, sterile, packaged |  |  | October 16, 2002 | 2002 | MEDEX, INC. | Product contains latex, however the label incorrectly states 'Latex Free'. | Other | On 10/17/2002, the firm notified their direct accounts and sub-accounts by telephone and facsimile. The customers were instructed to remove the product from service and return to Medex. |  | 304 | 304 cases (6,080 units) | The product was shipped to distributors and hospitals located in IL, AL, FL, UT, NY, and CA. | Terminated | December 30, 2003 | 440 | 304 | Not_Computer
Z-0411-03 | 4819 | [6 empty] | Side Fire Laser Fibers, Models DBLF-60SF and HBLF-60SF DBLF-60SF and HBLF-60SF |  |  | October 15, 2002 | 2002 | Laser Peripherals, LLC | The laser fibers were manufactured with an expired epoxy sealant which could result in degradation and misdirection of the laser beam. | Other | The recalling firm telephoned all of the consignees on October 15, 2002 and then sent recall letters, dated October 15, 2002, to the consignees requesting that the consignees return unused product. |  | 159 | 159 laser fibers | The products were shipped to 17 consignees (mostly hospitals and some distributors) located nationwide in the United States. There are no U.S. Government accounts. The products were also shipped to 2 distributors in Canada and Japan. The foreign consignees are:InstruMed Surgical2255 Dunwin Drive, Unit # 1Mississauga, Ontario L5L 1A3Canada905-820-0902M &amp; M Co. LtdMutoh Asakusabashi Bldg.1-16-6 Yanagigashi, Taito-kuTokyo, 111-0052 Japan81-3-3865-6575 | Terminated | January 01, 1990 | -4670 | 159 | Not_Computer
Z-0326-03 | 4814 | [6 empty] | Qwikwash Tubing (Air, Waste, Water), list 6208-23; an accessory for |  |  | October 10, 2002 | 2002 | ABBOTT LABORATORIES, INC. | visual discoloration of water tubing assembly | Other | Recalled by telephone calls on 10/10/02 and follow-up letters sent 10/11/02. Accounts were requested to examine their Abbott Commander PPC and Qwikwash systems, as well as their stocks of the tubing, for tubing that does not have the TYGON logo imprinted on the tubing. They were instructed to destroy any found and report the number destroyed back to Abbott. |  | 1343 | 1343 | The tubing was shipped intermingled with tubing from another supplier nationwide to 248 laboratory accounts, including 20 government direct sale accounts. The tubing was also shipped to Abbott subsidiaries in Canada, Germany and Singapore. | Terminated | January 01, 1990 | -4665 | 1343 | Not_Computer
Z-0454-03 | 4812 | [6 empty] | A) Roche/Hitachi MODULAR ANALYTICS E 170 Module |  |  | October 18, 2002 | 2002 | ROCHE DIAGNOSTICS CORP., | Fluid leak may cause smoke/damage analyzer | Other | The firm sent an urgent product correction notice dated 10/18/02 to all customers alerting them of the problem and advising them to be alert for the smell or sight of smoke. The firm is replacing the fittings as part of the correction. |  | 116 | 116 | United States and CanadaU.S. Government accountsBrooke Army Medical Center, 3851 Roger Brooke Dr., Rm 430-1, Ft. Sam Houston, TXPortland WV Med Ctr, 3710 SW US Veterans Hosp. Rd., Portland, ORVA Med Ctr, 700 S 19th St., Birmingham, ALVA Med Ctr, 1601 SW Archer Rd, Gainesville, FLVA Med Ctr, 10 Calle Casia, San Juan, PR Canadian consignee: Roche Diagnostics Canada, Laval, QC | Terminated | January 01, 1990 | -4673 | 116 | Not_Computer
Z-0396-03 | 4803 | [6 empty] | Holdex tube holder Item#450261 |  |  | September 20, 2002 | 2002 | Greiner Vacuette North America, Inc. | Holdex tube holder malfunction | Other | Letters were sent to consignees on/about 09/20/2002. |  | 103921 | Lot# A020208-9800 pieces, Lot#A050113-23800 pieces | Nationwide.AR, CA, CT, FL, GA, KS, LA, MA, MD, ,MI, MO, NH, NJ, NM, NY, OK, OH,OR, PA, SC, TX, VA.Veterans Administration, Medical Center, Calm VA-1 689, West Haven, CT. State of Georgia, Central State Hosptial, Lawrence Building, Milledgeville, GA. Sate of California, California State University, Dr. Larry Mai/Anthropology, Long Beach, CA.It appears no foreign accounts are involved. | Terminated | January 01, 1990 | -4645 | 103921 | Not_Computer
Z-0327-03 | 4680 | [6 empty] | Alumina V40 Head |  |  | June 26, 2002 | 2002 | HOWMEDICA OSTEONICS CORP. | V40 Alumina Heads may not assemble properly onto V40 hip stems | Device Design | On June 26, 2002, Howmedica sent out recall letters to their customers instructing them to remove and quarantine affected product from their inventory. |  | 190 | 190 (+676 28mm heads) | Howmedica distributed this product throughout the nation to various hospitals. There were 4 international custmers in United Kingdom, Australia, France, and Italy. | Terminated | June 29, 2010 | 2925 | 190 | Not_Computer
Z-0332-3 | 4498 | [6 empty] | Little Soothers Cold pack. Each boxed set (unit) contains 3 patches. |  |  | July 25, 2002 | 2002 | Avon Products, Inc. | Microbial contamination. | Other | Voice message on 7/25/02, letter on 7/29/02 to Avon Sales Representatives advising of the problem and requesting return of product. |  | 1609 | 1609 units, each containing three Little Soothers. | Sales representatives ("Avon Ladies") nationwide, for further distribution/sales to their customers (consumers). | Terminated | August 24, 2012 | 3683 | 1609 | Not_Computer
Z-0811-03 | 4552 | [6 empty] | Phillips ADAC brand C-PET and Allegro Imaging Systems DICOM 3.0, C-PET rev. 7.3.5. |  |  | July 24, 2002 | 2002 | PHILLIPS ADAC | Resulting images may be labeled incorrectly on PETView and receiving system. | Other | On 7/24/02, the firm issued letters via Federal Express to all its direct consignees, informing them of the affected devices and providing instructions on the recall. |  | 104 | 104 | Distribution of the product is worldwide and nationwide. Distribution of CPET was made between 6/7/02 to present. Distribution of Allegro was made between 4/15/02 to present.Product was distributed to the following foreign countries: China, Canada, Australia, Egypt, Spain, Belgium, Italy, Germany, Austria, France, Japan, and Korea. Product was also distributed to: National Naval Medical Center1602 Skipwith Road Richmond, VA301-295-2357and Wilford Hall Air Force2200 Bergguist DriveLackLand AFBSan Antonio, TX210-292-2165 | Terminated | January 01, 1990 | -4587 | 104 | Not_Computer
Z-0011-05 | 30059 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Model 512 |  |  | September 14, 2004 | 2004 | Medtronic MiniMed | Several complaints of pumps exhibiting A47 alarms were received. Investigation revealed the alarms occurred when the pump attempted to display text that exceeded the maximum display length of the pump screen. The investigation revealed this only occurs when the pump is programmed to display information in Spanish. | Other |  |  |  |  | Nationwide &amp; Spain | Terminated | March 03, 2005 | 170 |  | Not_Computer
Z-0409-03 | 4882 | [6 empty] | Ohmeda Medical Giraffe(tm) Omnibeds; and Ohmeda Medical Giraffe(tm) |  |  | October 21, 2002 | 2002 | OHMEDA MEDICAL | Electrical defects in neonatal incubators may lead to unintended movement | Other | Firm notified medical facilities' biomedical engineering department directors by certified letter 10/21/2002 that devices are usable with following precautions: (1) disable function of the elevating footswitch using the control panel and (2) place the 2 enclosed warning labels on incubators. Firm advised they will retrofit the incubators after evaluating possible solutions. Reply form enclosed.On 10/28/2002, firm notified consignees of a planned field correction to the OmniBed and Incubator devices by retrofitting isolation transformers to the units. The letter also advises that the firm has a retrofit kit available for correcting rechargeable battery back-up. The letter further advises that the units can continue to be used until corrective action is taken. |  | 1364 | 1364 units | Distribution to direct account medical facilities nationwide and international distributors.Foreign DistributorsCustomer NameAddress Line 1CityZipCountry DATEX-OHMEDA PTY LTD 149 ARTHUR STREET HOMEBUSH NSW 2140 AUSTRALIA SANITAS G M B H HOLUNDERSTRASSE 598 WALS B. SALZBURG A 5071 AUSTRIA MEDA N.V. OEYVAERSBOSCH 12 AARTSELAAR 2630 BELGIUM DYNAMED HEALTH CARE SYSTEMS 235 SHIELDS COURT MARKHAM ONTARIO L3R 8V2 CANADA CHINA HUA YANG TECHNOLOGY AND TRADE CORP (GROUP) BEIJING 100031 CHINA ESTIO SHANGHAI CO LTD RM 2202, ITALY BLDG SHANGHAI 200023 CHINA THE ORIENTAL SCIENTIFIC INSTR.I/E GRP,YINDU BLDG#67 BEIJING 100036 CHINA KIVEX LTD USSEROD KONGEVEJ 157 HORSHOLM 2970 DENMARK INSTRUMENTARIUM MEDICAL OU LIIVALAIA STR 22 TALLINN 10118 ESTONIA INSTRUMED INSTRUMENTARIUM OY ESPOO FIN-02630 FINLAND INSTRUMED PO BOX 300 HELSINKI FIN-00031 FINLAND INSTRUMENTARIUM CORP. INSTRUMED HELSINKI FIN-00510 FINLAND EUROCARE SRL RUE DES COMMERES COIGNIERES F-78310 FRANCE DATEX-OHMEDA GMBH DR ALRED-HERRHAUSEN DUISBURG 47228 GERMANY INTERMEDICA MANOS TZANIDAKIS ATHENS GREECE HONG KONG OXYGEN CHUN YAT ST, O INDUSTRIAL ES HONG KONG HONG KONG A KARLSSON HF BRAUTARHOLT 28 REYJAVIK 121 ICELAND PT MURTI INDAH SENTOSA JL SULTAN ISKANDAR JAKARTA 12240 INDONESIA OXYGENCARE LTD CORRIG ROAD DUBLIN 18 IRELAND OXYGENCARE LTD CORRIG ROAD DUBLIN 18 IRELAND DATEX-OHMEDA S.P.A. VIA CASSANESE, 100 MILANO ITALY AIR WATER INC 4019 OAZA YAMATO NAGANO 390-1701 JAPAN CHOONGWAE PHARMA CORPORATION 698, SHINDAEBANG-DONG SEOUL 156-757 KOREA, REPAL-ESSA MEDICAL &amp; SCIENTIFIC EQUIPMENT CO SAFAT 13036 KUWAIT INTERTRAME PUEBLA #14 COL. PE~NON V.E. MEXICO DF 15520 MEXICO SERVICIOS DE INGENIERIA EN MEDICINA SA DE CV,CUMBRES DE MEXICO D.F. 03630 MEXICO DATEX-OHMEDA B.V. KANTEMARSWEG 18 HOEVELAKEN 3871 AP NETHERLANDMEDISINK UTSTYR PO BOX 42 STOUNER OSLO 9 N-0913 NORWAY MUSCAT PHARMACY PO BOX 438 MUSCAT OMAN PRACTIKA PHILIPPINES, INC. 3RD FLR,DISPOPHIL BLDG, PASIG CITY 1605 PHILIPPINEPROMED SA UL KRAJEWSKIEGO 18 WARSAW PL 01-520 POLAND IBERDATA EQUIPAMENTOS SA RUA DR.RICARDO JORGE NO.4 AMADORA 2700-301 PORTUGAL CENTRE PERINATAL MEDICINE BUILDING 4 MOSCOW 109004 RUSSIAN FEABDULREHMAN ALGOSAIBI GTB ALGOSAIBI BLDG,AIRPORT RD RIYADH SAR-11411 Please see paper file for complete record. | Terminated | October 11, 2005 | 1086 | 1364 | Not_Computer
Z-0330-03 | 4896 | [6 empty] | Detection Reagent 1 of Digene's Hybrid Capture(TM) System Human |  |  | October 08, 2002 | 2002 | DIGENE CORP. | IVD test kit reagent may give false negative results | Other | Consignees were notified beginning 10/8/02 by phone and fax to discontinue use of the lot and reply to Digene representative with number of replacement reagents needed to retest patient specimens. Firm requests users to fax back the acknowledgement section of recall notification. Firm listed Digene's representative in Europe at 44 207 348 3500. |  | 217 | 217 kits | Kits were not distributed in the US. Kits were distributed to end users and distributors internationally.Genomed Ltd., United KingdomJK Trading, Czech RepublicInstitut fur Infektionskrankheiten, SwitzerlandAlpha Medical, SlovakiaHospital Universitario de Salamanca, SpainASL No. 18 Alba/Bra - Regione Piemonte, ItalyS.L. Durvis, SpainDr. Freischem &amp; Kollegen, GermanyDr. Med. Friedhelm Strube, GermanyLabor Dr. Peter Nae, GermanyUniversitatsklinikum Dusseldorf, GermanyDr. Med Eberhard Schneckenburg, GermanyDr. Med. Walter Franken, GermanyMedzinisches Labor Dr. Falke &amp; Partner, GermanyPathologisches Inst. Drs. Uhl &amp; Rosenkranz, GermanyArtzpraxis f. med. Mikrobiologie, GermanyZytologisches Labor Stocker, GermanyDr. Med. Ingo Scharmann &amp; Partner, GermanyDr. Med. Horparcy Gemeinschaftspraxis, GermanyDr. Med. Dietlinde Broer-Ahlers, GermanyGemeinschaftspraxis Leipzig, GermanyLandeskrankenhaus Feldkirch, AustriaApparategemeinschaft Allg auGmbH, GermanyDr. Menelaos Zaloumis, GermanyLabor 28, GermanyUniklinik Marburg Apotheke, GermanyHerm Dr. med. Reiss, GermanyFrauenklinik Hilden, GermanyDr. Med. Uwe Mundhenke, GermanyGem. Praxis. Drs. Loer, Winterhoff, GermanyDr. Thomas Weyerstahl, GermanyLabor Drs. Haubold Hofmann Jung Petihmann, GermanyFrau Dr. Med. Krompotic, GermanyDr. Med. Dietlinde Broer-Ahlers, GermanySteiermarkische Krankenanstaltenges, AustriaInstitut f. Medizinische Diagnostik, GermanyDr. T. Hermann / Dr. L. Holzer, GermanyDr. Josef Felixberger, GermanyZytologisches Labor Stocker, GermanyKlinikum der J. Gutenberg Uni Mainz, GermanyDr. Peter J. Stadler, Germany | Terminated | December 16, 2003 | 434 | 217 | Not_Computer
Z-0320-03 | 4890 | [6 empty] | Medtronic RF Enhancr SC Single-Curve Steerable Ablation Catheters, |  |  | October 29, 2002 | 2002 | MEDTRONIC, INC. | In some of the catheters, the catheter manipulator wire detached from the catheter insulating tip anchor during operation of the catheter deflection mechanism and after detachment penetrated the catheter wall and protruded from the shaft. | Other | A recall letter dated October 29, 2002 was delivered to the affected consignees in the United States by Medtronic representatives who retrieved the product. |  | 500 | 500 catheters | The catheters were distributed to 34 domestic consignees (hospitals) located nationwide in the United States. The product was also distributed to 2 Medtronic foreign subsidiaries and from there to foreign consignees located worldwide. | Terminated | January 01, 1990 | -4684 | 500 | Not_Computer
Z-0394-03 | 4898 | [6 empty] | 9 Fr. Percutaneous Sheath Introducer Kits, catheter introducer. |  |  | November 11, 2002 | 2002 | Arrow International | possibilty of leakage | Other | The recalling firm issued recall letters dated 11/11/02 to all their customers. The customers were instructed to examine their stock and cease using the affected lots. The letter also informed the customers that replacements would be shipped free of charge. |  | 25130 | 25130 kits | The product was shipped nationwide to 224 hospitals, 72 distributors and three sales represenatives. The product was also shipped to an international subsidiary account and an international distributor account. The government/military accounts are: VAMC, Houston, TX; VAMC, Portland, OR; VAMC, Birmingham, AL; VAMC, San Diego, CA; VA West Side Hospital, Chicago, IL; NIH, Bethesda, MD; VAMC, Decatur, GA; and VAMC, Minneapolis, MN.The international accounts are: Cecil E Mayo Ltd, 34 Morley Avenue, Rosenberg, NSW 2018, Sydney, Australia and Arrow Medical Products Ltd, 2300 Bristol Circle, Unit 1, Oakville, Ontario L6H 5S3, Canada. | Terminated | December 01, 2003 | 385 | 25130 | Not_Computer
Z-0389-04 | 28581 | [6 empty] | Lodox Statscan digital radiographic system. |  |  | May 10, 2004 | 2004 | Lodox Na Llc | The labels do not have proper identification and certification and the firm failed to provide adequate test and quality control procedures to establish certification by the manufacturer. | Other | The device in Illinois was re-exported. The device in Maryland was tested on 5/10/04 by a trained Lodox representative and proper labeling was reportedly attached. |  | 2 | 2 | Illinois and Maryland. | Terminated | October 06, 2004 | 149 | 2 | Not_Computer
Z-1350-02 | 4490 | [6 empty] | Disetronic D-TRON Insulin Infusion Pump, Models 8100001 D-TRON |  |  | June 17, 2002 | 2002 | Disetronic Medical Systems, Inc. | Some pumps started programming a bolus that was not initiated by the pump user, and the delivery of an unintended bolus was made unless interrupted by the device user. | Other | A notification letter dated June 17, 2002 was sent to patients who were using the pumps to direct them to check the buttons on their pumps to make sure they spring back when pushed and to make sure the acoustic alarms are turned on. All pump users then received instructions and mailing materials to return their pumps for inspection and replacement, if necessary. During February 2003 the firm changed their recall action from return for testing and possible replacement to recall and replace all of the pumps. They started sending a "Product Recall" letter about this to domestic patients on February 11, 2003. |  | 3357 | 3357 insulin infusion pumps in the U.S. | The pumps, imported into the United States, were distributed nationwide. The are an estimated 3,000 patients using the pumps in the U.S. | Terminated | June 01, 2004 | 715 | 3357 | Not_Computer
Z-1064-04 | 29393 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Osmetech OPTI-CCA Blood Gas Analyzer, Single Use, BP7564 OPTIÂżE-Glu Cassettes.The cassettes are labeled, Roche OPTICCA Cassettes E-Glu. There are 25 cassettes in each box. |  |  | June 22, 2004 | 2004 | Osmetech Inc | In some cases patient samples with high glucose concentrations, above the reportable range of the OPTI-CCA Analyzer, may be erroneously reported as 'Low'. | Other | Consignees were notified by letter on 6/22/04. A separate Recall Notice was mailed to distributors outside the United States with instructions on issuing the recall notice to their affected customers in the appropriate language. |  | 1087 | 1087 boxes | Nationwide and Argentina, Austria, Bolivia, France, Greece, Germany, Korea, India, Italy, Jordan, Netherlands, Sweden, Switzerland. | Terminated | January 12, 2005 | 204 | 1087 | Not_Computer
Z-0930-04 | 28972 | No | JCW | Penile inflatable implant. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | AMS Ambicor Penile Prosthesis, Penoscrotal Approach, Length: 14 cm, Diameter: 11 mm, Product Number 72401450 |  |  | April 29, 2004 | 2004 | American Medical Systems | The product is mislabeled as to length. | Other | Recall letters, dated April 29, 2004, state that it is important for the consignees to contact the firm's customer service for an exchange of the recalled products for new product. The letters recommend that the consignees send an attached letter to physicians who implanted the products. |  | 5 | 5 units | The affected products were distributed to consignees in the States of California and Texas, in Puerto Rico, and in the countries of Costa Rica and Taiwan. | Terminated | October 31, 2005 | 550 | 5 | Not_Computer
Z-0527-03 | 4679 | [6 empty] | Spinal 26 Gauge Quincke with Bupivacaine, Epinephrine, and |  |  | September 13, 2002 | 2002 | Abbott Laboratories | Mis-stenciling on tray. Tray stenciled Saddle Block instead of Spinal | Other | Letters to be sent out 09/12/2002. |  | 2980 | 2980 units | AR, CA, FL, MA, ME, NH, NM, PA, PR. There are no government nor foreign consignees. | Terminated | May 03, 2004 | 598 | 2980 | Not_Computer
Z-0874-05 | 31976 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nichols Advantage TSH - Third Generation |  |  | April 28, 2005 | 2005 | Nichols Institute Diagnostics | Does not meet performance claims in the labeling. | Other | Firm sent recall letter and response form April 28, 2005. Firm will follow-up to nonrespondants. |  | 1072 | 1072 | Nationwide, Germany, Greece and Spain | Terminated | March 16, 2012 | 2514 | 1072 | Not_Computer
Z-0847-05 | 31864 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | Fresh Look * RadianceÂż * Moonlight, (phemfilcon) Contact Lenses * -2.75 median, Median Dia:14.5 * Rx only * Ciba Vision, Duluth, GA. The product has been distributed in twin packs containing six pairs of lenses. |  |  | April 21, 2005 | 2005 | Ciba Vision Corporation | The actual power of the lenses may be different from what is indicated on the package labeling. | Other | Consignees were notified via traceable mail on April 21, 2005. |  | 2250 | 2250 units | Nationwide, Europe | Terminated | October 21, 2005 | 183 | 2250 | Not_Computer
Z-1440-05 | 32931 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Aqueduct Hematology Workflow Manager, software |  |  | August 10, 2005 | 2005 | Beckman Coulter Inc | Potentail for misidentification of historical patient results within the Aqueduct Hematology Workflow Manager software. In instances where the workflow manager software receives the same patient identifier for more than one patient the Aqueduct Workflow Manager software merges all results into one Patient File. The may cause test results to be associated with the wrong patient. | Other | A Product corrective action letter to be mailed the week of August 8, 2005, to all customers directing that each patient is identified with a truly unique patient identifier. |  | 36 | 36 | AK, AZ, CA, DE, FL, IN, KS, MA, MI, MN, MO, NC, NY, OH, PA, TX, VA, WA | Terminated | April 03, 2012 | 2428 | 36 | Not_Computer
Z-1065-04 | 29737 | [6 empty] | Sollux Systeme commercial tanning beds; Soleil Systems, Inc. |  |  | August 02, 2004 | 2004 | Soleil Systems, Inc. | The sunlamp products failed to incorporate a timer system and a control for termination of radiation emission in accordance with the Federal Sunlamp Product Performance Standard. | Other | Soleil Systems sent recall notification letters dated 6/28/04 to their dealers on 8/2/04, informing them of the noncompliances of failing to incorporate a permanently mounted timer and on/off switch in the sunbeds. Soleil will bring each sunbed into compliance and requested the dealers to review their sales records for any Sollux beds received between 4/28/04 and 11/01/03, and supply Soleil with a list of names and addresses of end users who received those units. Soleil sent out end user notifications dated 8/20/04 beginning on 8/23/04 and initiated the retrofit of the sunbeds on 10/25/04. |  | 70 | 70 beds | The sunbeds were distributed through dealers in Missouri, Texas, Virginia, Kentucky and Illinois | Terminated | March 28, 2006 | 603 | 70 | Not_Computer
Z-1168-2007 | 38273 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Urological Black Silicone Filiform Double Pigtail Ureteral Stent Set, 7.0Fr/24cm, sterile, Order Number (REF) 133724. |  |  | June 29, 2007 | 2007 | Vance Products Incorporated | The radiographic band may detach from the stent and remain in the patient after removal of the stent. | Other | Consignees were notified of the recall via letter dated 6//29/07, which instructed them to quarantine stock on hand and to fax number on hand back to Cook for return authorization. |  | 75 | 75 | Worldwide, including USA, Canada, and Japan. | Terminated | June 11, 2008 | 348 | 75 | Not_Computer
Z-1264-2007 | 44735 | No | LZK | Chin prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Spectrum Designs Medical-Nasal Dorsal Columella Implant, Catalog Number: S400-314S and Lot #31053; Serial Numbers: BAG612 ,BAG6I3, BAG6I4, BAG652, BAG654 and BAG655 |  |  | July 24, 2007 | 2007 | Spectrum Designs Inc | Mispackaging/Mislabeling: a portion of packaged Nasal LOT 31053 contained chin implants instead of the nasal implant. | Other | The firm initiated the recall on 07/24/2007 by contacting the customer (distributor) who filed the original complaint of mislabeling by email. The firm followed-up by sending recall letters (dated July 24, 2007) via FedEx or certified mail with signature confirmation to the one other direct account who received the recalled product on 08/03/2007. Customers were instructed to check their inventory for mislabeled products and call Spectrum at (800)239-6399 to report the number of this product being returned. They were given an address for product return. The firm indicated they will examine returned products, and scrap after examination. |  | 130 | 130 | Nationwide: California and China | Terminated | July 29, 2011 | 1466 | 230 | Not_Computer
Z-0289-2008 | 44752 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | A1CNOW+ with sampler 10-Test Kits (Product No: 3022), Metrika, Inc., Sunnyvale, CA 94085. (This test provides quantitative measurement of the percent of glycated hemoglobin (%A1C) levels in capillary (fingerstick) or venous whole blood samples. The test is used by professional healthcare providers to monitor long-term glycemic control in people with diabetes. The product is sold as a kit with a disposable monitor and reagents for ten A1C tests and is suitable for use at the point-of-care. |  |  | August 31, 2007 | 2007 | Metrika Inc | Incorrect Test Results: Due to manufacturing errors regarding temperature compensation factors, the device will give erroneously high results if the test is performed below 70 degrees Fahrenheit. | Process control | A September 14, 2007, Urgent Device Recall letter was issued to all Distribution Center Managers to return recalled product for replacement. Customer notification was made via a Dear Healthcare Professional letter dated Sept. 14, 2007, also titled Urgent Medical Device Recall. |  | 4286 | 4,286 | Nationwide including states of AK, AL, CA, CO, CT, DC, FL, GA, HI, IL, IN, KS, KY, MA, MD, MI, MN, MS, NC, NJ, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI and Puerto Rico. | Terminated | March 25, 2008 | 207 | 4286 | Not_Computer
Z-0288-2008 | 45333 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott Clinical Chemistry Aeroset/Architect (c Systems)Bilirubin Calibrator, List Number (LN) 1E66-03 &amp; List Number (LN) 1E66-04, Abbott Laboratories, Abbott Park, IL 60064 |  |  | August 10, 2007 | 2007 | Abbott Laboratories Inc. | Incorrect Calibrator values: changes in a raw material used in the Bilirubin Calibrator value assignment process contributed to variability in the preparation of the standard. | Material/Component Contamination | An August 10, 2007, Product Correction letter was provided to all ARCHITECT cSystem and AEROSET System customers. The letter advises that changes in a raw material used in the Bilirubin Calibrator value assignment process contributed to variability in the preparation of the NIST standard 916a. Revised Bilirubin values for the calibrator lot numbers in this recall are attached to the recall notice. Six necessary actions were listed in the recall notice as: 1) Identify the lot numbers of Clinical Chemistry Bilirubin Calibrator(s) currently in use. 2) Discard the value sheet and replace with the appropriate values provided with the recall letter. 3) Enter the revised Bilirubin values. 4) Calibrate the assays. 5) Evaluate your quality control. 6) Evaluate your reference range and any shifts in patient results. |  | 7033 | 7,033 units total worldwide | Worldwide: USA, Canada, Mexico, Guatemala, Puerto Rico, Germany, Hong Kong, Singapore, Australia, New Zealand, Jamaica, Venezuela, Columbia, Ecuador, Chile, Uruguay, Argentina, Brazil, Thailand, Trinidad &amp; Tobago, Bahamas, Panama, Cayman Islands, and El Salvador | Terminated | April 19, 2012 | 1714 | 7033 | Not_Computer
Z-1653-2008 | 46686 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL700 Series analyzer - Human Blood Analyzer (Laboratory Device). |  |  | January 18, 2008 | 2008 | Radiometer America Inc | Reported calibration errors are not displayed on the analyzers parameter bar screen or on the patient results when the calibration number is between 32768 - 65535 and 98303 - 131070. This may cause the operator to perform measurements and obtain questionable results. | Software design | On February 28, 2008, a Field Action Notes Letter was issued advising them about a Mandatory Software Update. Customers were notified of the problem by fax and letter (Fed Ex). The Recall Letter instructed the user to disable the parameter bar feature and only rely on the traffic light system indicator which properly indicates system errors. The letter includes instructions to temporarily fix the problem and notice of a update of the software module that updates the parameter bar. If you have any questions contact Vincent Sigmund at (440)-871-8900 ext. 209. |  | 2481 | 2481 Units | Worldwide Distribution - USA and Canada - Products were sold in all fifty (50) US States. including twenty-nine 29 government accounts and 138 Canadian accounts. | Terminated | February 24, 2012 | 1498 | 2481 | Not_Computer
Z-1363-2008 | 47340 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | BIOMET 3i LOCATOR ABUTMENT, Catalog/Item Number: LOA002, 2.0mm, BioMet 3i, Palm Beach, FL. |  |  | February 07, 2008 | 2008 | Biomet 3i, Inc. | Mispackaging: The Locator Abutment labeled as LOA002 may contain Catalog/Item # LOA003 product and visa versa | Other | Each USA customer was contacted by telephone and then by letter, Urgent Medical Device Recall, dated 2/15/2008, sent by fax. International customers were first contacted by E-mail then by fax letter also. Each customer was instructed to check their respective inventory for the affected part and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory for product return. They were also instructed to return the affected product to Biomet 3i. |  | 100 | 100 | Worldwide. | Terminated | July 10, 2008 | 154 | 197 | Not_Computer
Z-0627-2009 | 50139 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Excelsior SL-10 Microcatheter, placed in sterile pouch. Manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | October 07, 2008 | 2008 | Boston Scientific Corporation | Packaging issue may result in failure to maintain sterile barrier. | Package design/selection | Consignees were notified by letter dated October 7, 2008. The letter identified the issue and affected product. Customers were instructed to: 1.) Discontinue use of and segregate recalled product, 2.) Complete and return the enclosed Reply Verification Tracking Form, and 3.) Package/ship the recalled product and send to the Boston Scientific Corporation (BSC).Return the completed Reply Verification Tracking Form via fax (800-836-4804) to the BSC Customer Service Call Center.A follow-up letter dated October 8, 2008 was also issued to customers. The letter notified customers of the correct date range of expiration for product within the scope of the recall which is September 2008 through October 2010. There were no other changes identified in the follow-up letter. |  | 21423 | 21,423 units | Worldwide Distribution. | Terminated | March 02, 2010 | 511 | 21423 | Not_Computer
Z-0906-2009 | 50737 | No | EMG | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Obwegeser Ramus Clamp w/ wide shield, 20 cm, 8", Right. ORAL-MAXILLO-CRANIO-FACIAL SURGERY, Original OBWEGESER Instruments. Walter Lorenz Surgical, Inc., Subsidiary of Biomet Microfixation, Inc. NOTE: None distributed domestically. |  |  | May 21, 2008 | 2008 | Biomet Microfixation, Inc. | The Biomet Microfixation Obwegeser Ramus Clamp, Wide Right 01-7996 were incorrectly etched as the left part number 01-7997. | Labeling mix-ups | Customers were contacted by e-mail notification with read receipt. Customers were notified in the e-mail to immediately discontinue use of the recalled product and return the affected items to Biomet Microfixation, Inc. The e-mail also included an Inventory Reconciliation form for customers to complete and return via fax (904-741-1801) to Biomet Microfixation, Inc. For questions related to this notice, please call 1-800-874-7711 or 904-741-4400 extension 448, Monday through Friday, 8 a.m.-4 p.m. Eastern Standard Time (EST). |  | 4 | 4 | International Distribution including countries of Ireland and Japan. | Terminated | September 11, 2009 | 478 | 4 | Not_Computer
Z-1620-2011 | 51208 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BD FACS Sample Prep Assistant (SPA and SPA II) Instrument, |  |  | January 03, 2009 | 2009 | BD Biosciences, Systems &amp; Reagents | To prevent a needle stick while replacing the probe, the firm has initiated a Field Correction, and issued an improved Users Guide procedure (User Bulletin, Part Number 645050). | Nonconforming Material/Component | The firm, BD Biosciences, sent an information letter dated December 2008 with an enclosed New Probe Replacement Service/User Bulletin, to the customers. The letter described the product, problem and action to be taken. The customers were instructed to review the attached User Bulletin (Part Number 645050); ensure that their laboratory follows the updated procedure for removing and replacing probe; place the User Bulletin in their BD FACS SPA or SPA II User's Guide; always follow the instructions in User Bulletin to properly remove and replace probe, and complete and return the Effectivity Form via fax to 408-577-1409. Should you have any comments or questions, please contact BD Biosciences at 1-877-232-8995 in the US. Outside the US, contact your regional BD office or field service engineer for assistance. |  | 70 | 70 | Worldwide distribution: USA and countries including: Hong Kong, Australia, Belgium, Canada, El Salvador, Hong Kong, Japan, and Mexico. | Terminated | March 10, 2011 | 796 | 70 | Not_Computer
Z-1555-2009 | 52007 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Stratus CS Acute Care Troponin I CalPak. |  |  | April 28, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc. | The firm received and confirmed that some lots of Cardiac Troponin I (cTnl) CalPaks (CCTNI-CR) may exhibit calibration issues that cause lower than typical signal during calibration with the current population of cTnl TestPaks (CCTNI). This issue may result in failed calibrations due to "Slope Error" messages or falsely elevated patient sample and quality control (QC) results on calibrations that | Labeling False and Misleading | Siemens Healthcare Diagnostics, Inc. issued an Urgent Field Safety Notice dated April 29, 2008 to all affected customers instructing them to discontinue use of 3 CalPak lots and reduce the shelf-life for the remaining lots.For further information, contact Siemens Healthcare Diagnostics, Inc. at 1-302-631-7672. |  | 4415 | 4415 cartons | Worldwide Distribution - including United States, United Arab Emirates, Austria, Bermuda, Belize, Canada, Germany, Egypt, India, Japan, Saudi Arabia and Bengal. | Terminated | July 14, 2009 | 442 | 4415 | Not_Computer
Z-1625-2009 | 52189 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific EndoVive 20F (6.7 mm) Safety PEG Kit, Push Method for use with sterile Xylocaine ampule, latex free, sterile, Boston Scientific Corp., Natick, MA; Cat. no./REF 6647, inner pouch UPN M00566470 and outer carton UPN M00566471. |  |  | May 05, 2009 | 2009 | Boston Scientific Corp | The user may be unable to pass the device over the guidewire for placement in the patient, resulting in an inability to use the device. | Process control | Boston Scientific Corporation issued an "Urgent Medical Device Recall - Immediate Action Required" notice dated May 5, 2009 as follow-up to their "Recall of the Safety and Standard Push PEG Kits" dated December 30, 2008. The notice outlined the issue and Consignees were instructed to locate and return all affected product to the firm. In addition, the firm requests that users complete and return a Reply Verification Tracking Form via fax to 1-508-683-5578.For further questions, contact the Boston Scientific Corporation Field Action Team by phone at 1-508-683-4678. |  | 1490 | 1490 | Worldwide Distribution -- United States, India, Japan and Slovenia. | Terminated | October 13, 2010 | 526 | 10259 | Not_Computer
Z-1716-2009 | 52217 | No | JDL | Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | CONSERVE, TOTAL A-CLASS HEAD W/BFH TECH., REF 38AM-3635, O.D. Size: 36mm, Neck Length: +3.5mm, Taper: SLT, Sterile R, 1 each, Rx only, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002. |  |  | May 14, 2009 | 2009 | Wright Medical Technology Inc | Incorrect packaging, swapped sizes | Packaging process control | Wright Medical Technology, Inc. notified hospitals, surgeons and their distributors of the recall situation via an "Urgent: Medical Device Voluntary Recall" notice dated May 14, 2009 explaining the problem, the hazard involved and requesting confirmation that the product was removed from their inventory. The letter included a "Fax Back" response page to ensure the receipt of the notice.For further questions, contact Wright Medical Technologies, Inc. at 1-800-874-5630. |  | 15 | 15 units | Worldwide Distribution -- Including states of NC, TN, IA, IN, MD, WI, WV, OK, WA, PA, TN, TX, FL and CA and country of Canada. | Terminated | December 09, 2009 | 209 | 29 | Not_Computer
Z-0541-2011 | 52039 | No | KLC | Dental chair and accessories. | Dental | Dental | 510(k) | A-dec Decade Plus 1221 Dental Chair |  |  | April 01, 2009 | 2009 | A-Dec Inc | The design of the Decade Plus 1221 chair backrest creates a possible pinch-point between the chair backrest and the underside of the armrests. | Device Design | A-dec sent a letter dated March 23, 2009 to consignees/customers (doctors and network dealers). The consignees were informed that the design of the Decade Plus 1221 chair backrest creates a possible pinch-point between the chair backrest and the underside of the armrests. The potential for this occurs when returning the chair to the upright position. The consignees were advised to reprogram the Position 0/Entry-Exit preset to stop the back before it reaches a full upright position. This increases the gap between the armrest and the back slightly. To reprogram a chair, the firm included excerpts from a chair Instructions for Use (IfU) guide that detail the steps to reprogram a chair. Consignees can contact the firm's Customer Service at 503-538-9471, 1-800-547-1883, or via fax at 503-538-0276.The firm will replace the chair back assembly for all affected dental chairs. |  | 2723 | 2723 units | Worldwide distribution: US and Canada. | Terminated | December 16, 2010 | 624 | 2723 | Not_Computer
Z-1717-2009 | 52428 | No | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | M-19 (Iridium 192): Source Production and Equipment Company, St. Rose, LA. |  |  | March 03, 2008 | 2008 | Source Production &amp; Equipment Inc | Radioactive source broke during use. | Other | Source Production and Equipment, Inc. notified the Distributor by phone and letter March 2008 and requested the return of all affected units. |  | 3 | 3 units | Nationwide Distribution -- GA, NV and NY. | Terminated | August 19, 2009 | 534 | 3 | Not_Computer
Z-1846-2015 | 71359 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/ zeego systems; |  |  | May 22, 2015 | 2015 | Siemens Medical Solutions USA, Inc | There exists a possible position sensor fault in the swivel base axis not being detected by the system software. When moving, the system could potentially exceed its usual speed, resulting in a collision. No injuries reported. | Software design | Siemens sent an Important Customer Safety Notice letter dated May 22, 2015, to affected customers. The letter identified the affected product, problem and actions to be taken. The letter also informed customers that Siemens would contact them to schedule the installation of a software update to resolve the problem. |  | 475 | 475 | US Nationwide Distribution | Open, Classified |  |  | 475 | Not_Computer
Z-1835-2015 | 71394 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Optical Access System; 5mm x 100mm and 12mm x 100mm; |  |  | June 01, 2015 | 2015 | Stryker Sustainability Solutions | Stryker Sustainability Solutions is recalling Trocars because they have received reports indicating that some may fracture at the distal tip resulting in fragments at the surgical site. | Under Investigation by firm | On 06/01/15 a customer notification letter was sent out by the firm.The notification letter instructs customers to stop using the affected lots of the recalled product, and should consider other trocar models reprocesses by Stryker as an alternative. The letter notes that attached is a Effectiveness Check Form and complete list of affected lots shipped to the facility.A Stryker Sustainability Solutions Sales Representative will be sent to check the inventory of the affected lots. Once the inventory is checked the affected lots will be removed. The firm will complete a Recall Effectiveness Check Form . The form is to be completed even if no affected product is found.The firm requests that the form be returned to the sales representative or fax to 480-809-4975. If the form indicates that affected devices remain in inventory a prepaid shipping label will be issued for the return of the product. Customers will receive a credit for all affected devices returned.The firm states that adverse reactions or quality problems experiences with the use of the product may be reported to the firm's hotline 888-888-3433 x5572, or at www.stryker.com/productexperience as well as through FDA's MedWatch Adverse Event Reporting program. |  | 147351 | 147,351 | Distributed Nationwide including Puerto Rico. | Open, Classified |  |  | 147351 | Not_Computer
Z-1839-2015 | 71373 | No | GAL | Absorbable surgical gut suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon Fast Absorbing Plain Gut - MultiPass Needles, Sterile, Absorbable Surgical Suture, 5-0 (1.5 metric), 18" (45 cm), Conventional Cutting, Lot #JD6163, Exp: July 2019. |  |  | May 08, 2015 | 2015 | Ethicon, Inc. | Ethicon Fast Absorbing Surgical Gut (Plain) Suture is being recalled because the labeling on the box and foil are incomplete. | Labeling mix-ups | Ethicon initiated this voluntary recall by sending an Urgent: Medical Device Recall (Removal) Letter/Business Reply Form (BRF) dated May 8, 2015 to their affected customers via UPS. Stericycle will be handling returns. The letters identified the affected device and the reason for the recall. Customers were asked to examine their inventory immediately to determine if they had affected devices on hand. If so, they are to remove the affected product and communicate the issue to any necessary personnel. Customers with questions were advised to contact their local sales representative or call the Customer Support Center at 1-877-384-4266.For questions regarding this recall call 908-218-0707. |  | 1922 | 1,922 units | Worldwide Distribution -- US (nationwide) and Canada. | Terminated | June 06, 2017 | 760 | 1922 | Not_Computer
Z-1841-2015 | 71291 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | VersaTREK Instrument Series 240 REF 6240, Manufactured for Remel Inc. |  |  | May 19, 2015 | 2015 | Remel Inc | Use of the recalled product may result in false positive reports. | Release of Material/Component prior to receiving test results | Remel, Inc. a part of Thermo Scientific sent a Product Correction Notification dated May 19, 2015, to all affected customers. The notification states the problem with the device and that the recalling firm's Technical Services department will be contacting the customer to schedule a field service of the device to correct the issue. Customers were instructed to notify any personnel who need to be notified of the potential for a false positive resulting from rare inconsistent barometric pressure readings. Return the attached Medical Device Recall Return Response form acknowledging receipt of the notice of the affected product. If product was further distributed customers were instructed to contact those entities, advise them of the situation and provide them with a copy of the letter. Customers with questions were instructed to contact the Technical Services Department at 800-642-7029. For questions regarding this recall call 770-409-0713. |  | 143 | 143 devices | Worldwide distribution: US (nationwide) to states of AZ, CO, FL, GA, IA, ID, LA, MI, MO, PR, TN, TX, VA, and WA and the countries of Canada, People's Republic of China, Egypt, Finland, France, Indonesia, Israel, Jordan, Kenya, Republic of Korea, Latvia, Romania, Russia, Spain, Thailand, United Kingdom, Uruguay, and Vietnam. | Terminated | March 24, 2017 | 675 | 164 | Not_Computer
Z-1845-2015 | 71537 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Prestige Plus 185cm J-Tip: |  |  | June 04, 2015 | 2015 | Volcano Corporation | During a relabeling procedure at the manufacturing facility, the products were labeled with a shelf life in accordance with the date of the relabeling and not in accordance with the original shelf life date at the time of manufacture | Labeling mix-ups | Volcano sent an Urgent Medical Device Recalls lettersdated June 4, 2015 via Federal Express to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine product inventory, quarantine the affected product, complete the attached Customer Return form, and contact Volcano Customer Service at 800-228-4728, option 4 to arrange for the return of any product in inventory. |  | 162 | 162 total (both devices) | US Nationwide Distribution and one account in Canada. | Terminated | October 27, 2015 | 145 | 162 | Not_Computer
Z-1832-2015 | 71311 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | EVAC 70 XTRA with Integrated Cable REF EICA5872-01 |  |  | May 15, 2015 | 2015 | ArthroCare Medical Corporation | Potential component failure causing device inoperability | Component design/selection | Smith &amp; Nephew sent an Urgent Product Recall notification letter dated May 15, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory, quarantine immediately and return to the address listed on the inventory Return Certification Form. |  | 4111 | 4,111 units | US Nationwide Distribution | Open, Classified |  |  | 6466 | Not_Computer
Z-1837-2015 | 71294 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom Surgical Kits containing Medtronic Covidien Devon(TM) Light Gloves: |  |  | April 27, 2015 | 2015 | DeRoyal Industries Inc | Surgical kits contain Medtronic Covidien Devon(TM) Light Gloves which were recalled for splits or holes | Nonconforming Material/Component | The firm, DeRoyal, sent an "URGENT! RECALL NOTICE" letter dated 4/27/2015 to their customers (i.e., end-user level, its direct accounts and hospitals) via overnight on 04/27/2015. The letters described the product, problem and actions to be taken. The customers were instructed to do the following: a. send Recall Instructions and a listing of their accounts who were shipped affected product (also included were labels, Training forms, Record of Destruction Forms)b. notify Distributors and Hospitals with a list of products shipped to their account with instructions to place any remaining inventory in quarantine.c. document the lots/quantities labeled at their accounts and complete and return this information on the "Affected Product Listing" to DeRoyal via fax to 865-362-3716 or email to recalls@deroyal.com even if your distributor no longer has any inventory.d. conduct training at the hospitals and this training will be documented on the Recall Training form. e. leave Record of Destruction Forms at the hospital for completion/return to DeRoyal as products are discarded/destroyed.If you have questions or need assistance with the recall, please contact me at 865-362-6465 or tmarsee @deroyal.com. |  | 235785 | 235,785 units | Worldwide Distribution: US (nationwide) and country of Guatemala. | Terminated | December 29, 2015 | 246 | 235785 | Not_Computer
Z-1843-2015 | 71363 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | IMMULITE 2000/IMMULITE 2000 XPi Italian Cypress 3g Allergy, Catalog # T23L4 (40 test), Siemens Material Number (SMN) 10386079, IVD. |  |  | May 06, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | IMMULITE 2000/2000XPi Italian Cypress (T23L4), Lots 206 and 208, fail to meet specifications to support the shelf-life stability of 2 years (24 months) expiration date as indicated on the products label. | Under Investigation by firm | Siemens Healthcare Diagnostics issued an Urgent Medical Device Recall Letter (IMC-15-11A.US, dated May 6, 2015) in the US via FedEx, and an Urgent Field Safety Notice (IMC-15-11A.OUS, May 2015) outside the US on May 6, 2015. Customers were instructed to immediately discontinue use of and discard IMMULITE 2000/IMMULITE 2000 XPi Italian Cypress allergen, lot 206. The shelf-life of IMMULITE 2000/IMMULITE 2000 XPi Italian Cypress allergen, lot 208, has been shortened to June 30, 2015. Discontinue use and discard lot 208 after June 30, 2015. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Recall letter within 30 days in the US and OUS and fax to Customer Care Center at 312-275-7795. |  | 1566 | Domestic: 17 kits, Foreign: 1549 kits | Worldwide Distribution-Distributed in the states of Florida, New Mexico and New York and the countries of Australia, Belgium, Chile, Costa Rica, Croatia, Czech Republic, France, Germany, Greece, Guatemala, Israel, Italy, Japan, Kazakhstan, Latvia, Netherlands, Portugal, Romania, Russian Federation, Slovakia, and Spain. | Open, Classified |  |  | 1566 | Not_Computer
Z-1840-2015 | 71406 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The spare component Anti-HBs Cutoff Calibrator, 10 mLU/mL, Catalog number 26154. |  |  | May 13, 2015 | 2015 | Bio-Rad Laboratories, Inc. | The spare component Anti-HBs Cutoff Calibrator, Catalog number 26154, is recalled because the expiration date of 2016-04-30 printed on label in error. | Incorrect or no expiration date | Bio-Rad Laboratories Inc, sent an Urgent Product Correction letter dated May 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The recalled calibrator is a spare component which is sold separately from the MONOLISA Anti-HBs EIA Kit and the kit is not affected by this recall.Customers were informed that testing was performed and the Anti-HBs 10 mLU/mL Calibrator was acceptable to use through the expiration date of the kit dated 2015-11-30.Customers with questions can call the firm at 1-800-224-6723, option #2, then #3. |  | 20 | 20 vials | US Distribution to the states of : AZ, CA, ID, MO, and NH. | Terminated | November 05, 2015 | 176 | 20 | Not_Computer
Z-2060-2010 | 55224 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Terumo HX2 Temperature Management System 120V, 50/60 Hz catalog number 809810, Terumo Cardiovascular Systems Corp., Ann Arbor MI. |  |  | June 11, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The HX2 Temperature Management System mixing valve may fail resulting in an unusable water channel and then the machine would go into standby mode. | Device Design | The firm, TERUMO Cardiovascular Systems, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 11, 2010, to the consignees. The letter describes the product, reason for correction, what to do in the event of failure, potential hazard, correction, affected population and action to be taken by the customers. The customers are instructed to review the correction notice, assure that all users are aware of this notice, and complete, sign and fax the attached Customer Response Form to Terumo CVS Customer Service at 800-292-6551. The Terumo CVS Customer Service Dept. will contact you upon receipt of the Customer Response Form to schedule the replacement. The firm plans to inspect and correct all units in the field within 5 months.We encourage you to contact Terumo CVS Customer Service at 1-800-521-2818, Monday-Friday, 8 AM-6 PM EST, with any questions or concerns. |  | 100 | 100 | Nationwide Distribution: AR, CA, CO, CT, FL, HI, IA, IL, KS, LA, MD, MI, MO, MS, NY, OH, PA, TX, UT, WA and WI. | Terminated | April 13, 2011 | 306 | 100 | Not_Computer
Z-2067-2010 | 55868 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. |  |  | May 26, 2010 | 2010 | Ge Healthcare It | There is a potential safety issue associated with the use of GE Centricity Laboratory software where results from interfaced external reference laboratories are not displayed or stored correctly. This happens when the test is set up in Centricity Laboratory as a decimal (numeric) type result and the test result sent from the reference lab includes a comma in the result field. | Software design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated May 26, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the patient safety issue involving results from interfaced external reference laboratories that are not displayed or stored correctly if an unexpected alpha character is encountered in a result message that is expected to have a numeric value. This happens when the test is set up in Centricity Laboratory as a decimal (numeric) type result and the result sent from the reference lab includes a comma in the result field. The letters provided the users with specific safety instructions to follow until the software is updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to the their GE Service Representative or the GE Helpdesk at 888-778-3375. |  | 37 | 37 units | Worldwide distribution: California, Georgia, Michigan, Minnesota, New York, Ohio, Oregon, Tennessee and Texas, and internationally to Australia, Canada, China, England, India, Malaysia, New Zealand, Qatar, Scotland and Singapore. | Terminated | July 15, 2011 | 415 | 37 | Not_Computer
Z-2093-2010 | 55578 | No | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | The "RS-FBG Full Back Conductive Garment" with electrode attachments for use with interferential and muscle stimulation with a Transcutaneous Electrical Nerve Stimulation device (TENS) and/or muscle stimulator. |  |  | May 17, 2010 | 2010 | International Rehabilitative Sciences, Inc. | Device distributed with labeling and IFU outside of their 510(k) clearance. | Other | The firm, RS Medical, mailed "URGENT NOTICE OF VOLUNTARY CORRECTIVE ACTION" letters dated May 14, 2010, to customers on May 17, 2010 and letters to physicians/consignees on May 19, 2010. The letters included an updated RS-FBG Operation Manual and stated what the update covered; i.e. that the device is indicated for use on the "middle and upper back". The customers were ask to please read your new manual and replace your original RS-FBG Full Back Conductive Garment Operation Manual with this updated version and discard your old RS-FBG Garment Operation Manual. RS Medical stated that your RS-FBG Garment is not affected by this Corrective Action, only the Operation Manual. RS Medical informed the customers that you may continue to use the RS-FBG Full Back Garment with your RS Medical Stimulator as prescribed by your healthcare provider.If you have any questions or concerns, please call our Toll Free Customer Service number at 1-866-417-1293, email us at RSFBG@rsmedical.com or visit our special web page at www.rsmedical.com/rsfbginfo. |  | 190073 | 108134 units, 81,939 subject to recall | Nationwide distribution. | Terminated | August 15, 2012 | 821 | 190073 | Not_Computer
Z-2094-2010 | 56033 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Rolko Trapeze handles, part 148671, for use with Care Assist bed P1170 and Hill Rom 1000 bed P1160, Hill Rom Batesville, IN. |  |  | June 11, 2010 | 2010 | Hill-Rom, Inc. | The firm received a complaint regarding the plastic housing used to support the trapeze handle. The firm investigated and found that the manufacturer of the trapeze handle didn't correctly process the plastic component. Therefore the plastic had the potential to break prematurely. | Process control | The firm, Hill-Rom, sent an "Urgent Field Safety Notice" letter dated June 11, 2010, to its customers. The letter describes the product, problem and action to be taken by the customers. Hill-Rom provided the quantity of handles needed to replace handles customers received during the affected period. The customers were instructed to immediately remove and replace all Trapeze Handles that were delivered during the affected period; to replace the Trapeze Handle Assemblies (with a manufacturing date from January 2007 to January 2009) used on their Patient Helper with the enclosed unit(s); complete and return the Customer Response Form as soon as possible, or by July 26, 2010; advise the firm immediately if they have relocated the beds and/or the accessories to another location, and forward the letter to any other facility personnel they deem appropriate. If you have any questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. |  | 319 | 319 | Worldwide distribution: USA and Canada. | Terminated | February 18, 2011 | 252 | 319 | Not_Computer
Z-2104-2010 | 56161 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Merit Laureate Hydrophylic Guide Wire, Angled, Made In Ireland, CAT No: LWSTDA35260EX, 260cm, Sterile EO, Merit Medical Systems, Inc., South Jordan, UT 84095. |  |  | June 28, 2010 | 2010 | Merit Medical Systems, Inc. | Guidewires have decreased lubricity at the proximal end, inhibiting proper operation. | Process control | The firm, Merit Medical System, Inc., notified the Sales representatives by e-mail on June 28, 2010, and instructed them to contact all of their accounts and to advise them to discontinue use and return any unused products to Merit. A Product Recall Notice was to be completed by each facility under the direction of the rep and all activities were to be completed by July 9, 2010. Questions should be directed to Penny Goldman, Senior Product Manager at 801-208-4198 or Jackie Preece, Manager, Customer Service at 801-208-4365. |  | 199 | 199 units | Nationwide distribution: CA, CO, FL, GA, IL, IN, KS, MA, MD, MI, MN, NC, NH, NJ, NY, PA, and SD. | Terminated | October 19, 2010 | 113 | 964 | Not_Computer
Z-0492-2011 | 56405 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit Model 8000 only when used with Alaris Pump module or Alaris Syringe module |  |  | August 02, 2010 | 2010 | CareFusion Corporation | Than recall was initiated because CareFusion has identified identified a potential risk associated with the Alaris PC unit formerly known as Medley PC unit (PC Unit) model 8000 only when used with Alaris Pump module (Pump module) or Alaris Syringe module' (Syringe module), This recallnotification details the potential risk and recommended steps for users to take if they encounter a specific PC | Device Design | CareFusion will send by registered return receipt mail the customer notification letter to all customers usingthe affected Alaris PC Unit beginning August 24, 2010. The customer notification letter will be addressed tothe Directors of Nursing, Risk Management, and Biomedical Engineering. Customers will be required toconfirm receipt of the notification by returning the Recall Response Card to CareFusion by postage-paid, selfaddressedmail, fax or email. The recall notice informed the customers of the reason for the notice, the unitsaffected, the potential risk, a step-by-step description of the issue and the required actions for users.The customers were also told that CareFusion is working to release a software update to address this display error and will contact your facility as soon as the software update is available. CareFusion does not require that you return your devices. Clinicians should weigh the risk/benefits to the patients beforecontinuing the use of the device.Customers were also told that any adverse reactions experienced with the use of this product, and/or quality problems shouldalso be reported to the FDA's MedWatch Program. The recall notice provided contact information for CareFusion's Recall Center, Customer Advocacy and Technical Support units. |  | 94146 | 94,146 units | Nationwide in the United States, Canada, Mexico, and the European Union | Terminated | July 25, 2013 | 1088 | 94146 | Not_Computer
Z-2583-2010 | 56343 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | SynFix-LR Spinal Implant System (Implant Holder for SynFix-LR and surgical technique brochure) |  |  | July 23, 2010 | 2010 | Synthes USA (HQ), Inc. | A potential for the implant holder tip to break off and lodge in the SynFix-LR implant, which pose a potential risk to the patient. | Error in labeling | Synthes Spine Hospital sent an ' URGENT: MEDICAL DEVICE LABELING CORRECTION" letter dated August 20, 2010 to all customers. The letter described the labeling issue and actions to be taken by the customers. The customers were instructed to review the supplementary labeling provided; contact their Synthes Spine Sales Consultant to arrange for provision of the Holding Forceps (p/n 388.401) prior to each SynFix-LR procedure; and complete and return the enclosed Verification Section form to Synthes via fax at 610-251-9005 or email: FieldAction@Synthes.com.If you have any questions, please call 1-800-620-7025 x 3500 or contact your Synthes Spine Sales Consultant. |  |  |  | Worldwide Distribution: USA and the countries of Canada, Japan, Switzerland, Australia, Mexico, Malaysia, New Zealand, and Singapore. | Terminated | December 18, 2014 | 1609 |  | Not_Computer
Z-0350-2011 | 56753 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI |  |  | September 01, 2010 | 2010 | Elekta, Inc. | It is possible to do couch corrections twice if both Couch Move Assistant in MOSAIQ and Remote Automatic Table in XVI are configured when in SYNERGISTIQ mode. | Other | Elekta sent an "Important Notice A340" dated September 1, 2010 to all customers. The notice described the product, problem and actions to be taken by the customer. The customers were instructed to make a decision if XVI or MOSAIQ does the corrective table movement; configure and ensure that the system only uses RATM or CMA for the corrective table movement, when in SYNERGISTIQ mode, and complete and return the IMPORTANT NOTICE Confirmation of Receipt form via fax at +44(0)1293 654401. If you have any questions contact your local service representative at +44(0)1293 654200. |  | 23 | 23 units | Nationwide distribution: AZ, CA, DC, IL, KY, MA, MI, MS, NV, NY, OH, OR, PA, VT, and WV. | Terminated | December 14, 2010 | 104 | 23 | Not_Computer
Z-1650-2011 | 57404 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | BIO-RAD BioPlex 2200 Syphilis IgG Control Set. Lot # 46636 |  |  | November 23, 2010 | 2010 | Bio-Rad Laboratories | An upward shift in the control values can result in a control value that falls outside of the current manufacturing ranges. | Other | The firm, Bio-Rad, sent a "Voluntary Field Correction: QC Ranges for Syphilis IgG Control Set Lot#46636" letter to Domestic customers via FedEx on November 24, 2010 and to the International customers via e-mail on November 23, 2010. The letter described the product, problem and actions to be taken by customers. The Customers were advised to temporarily establish their own means and acceptable ranges for Control Lot 46636 for the r15 and r47 markers. The firm will provide adjusted recommended means and acceptable ranges. Additionally, the customers were to complete and return the Customer Response Form for USA and/or the Subsidiary Response Form (Outside US) via fax to (510) 741-4846, Attn: Bio-Rad BPD-Product Support.If you have any questions about this notification, contact your local Bio-Rad Technical Support. |  | 202 | 202 | Worldwide distribution: USA and countries including: Canada, Czech Republic, France, Germany, and Italy. | Terminated | August 04, 2011 | 254 | 202 | Not_Computer
Z-1519-2011 | 57107 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Pacific Hemostasis Âż Thromboplastin-DS, catalog #100354, packaged in 4 mL lypholyzed vials, 10 bottles/carton. Vials are labeled, in part ***Fisher Diagnostics a division of Fisher Scientific Company, LLC Middletown, VA 22645-1905 USA*** |  |  | October 06, 2010 | 2010 | Fisher Diagnostics, A Company of Fisher Scientific LLC | An in vitro diagnostic reagent used in coagulation testing may cause inaccurate clotting times in quality control and patient plasma samples. | Nonconforming Material/Component | The firm, ThermoFisher Scientific, sent two "Urgent Medical Device Removal" letters one via Federal Express on October 20, 2010 and certified mail for domestic customers on October 21, 2010. The letters described the product, problem and action to be taken by the customers. The customers were instructed to indicate the quantities of all affected product currently in stock at their facility; complete and return the attached Response Form/table to Fisher Diagnostics via fax to (540) 869-8129 or by mail to the attention of Technical Support; dispose of any remaining inventory of the affected lot numbers; provide a copy of this letter to their customers, and file this letter for future reference.If you have any questions regarding this notification, please contact our Technical Support staff at 800-528-0494, Option 2, or email techsupport.diagnostics.mtn@themofisher.com. |  | 1190 | 1190 kits | Worldwide distribution: USA including states of: Alaska, Florida, Kentucky, Ohio, New Jersey, New York, and Texas; and countries including: Argentina, Dubai, Italy, Mexico, Philippines, Romania, and Turkey. | Terminated | November 02, 2011 | 392 | 1190 | Not_Computer
Z-1741-2011 | 57705 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CR Tibial Insert, Size 5, 7mm, REF 163-2507, Sterile R, Ortho Development, Draper, Utah 84020. |  |  | January 07, 2011 | 2011 | Ortho Development Corporation | Components of a replacement knee system could be mislabeled. | Error in labeling | The firm, Ortho Development, sent a "Field Notification Memo" date January 7, 2011 to its customers. The memo described the product, problem and action to be taken. The customers were instructed to promptly return any remaining parts of the listed product; locate all the parts in their inventory; identify any of the listed parts that may have been used in surgery; verify the quantity they have on hand and the quantity used in surgery and reconcile them with the quantity sent; if discrepancy exists, notify customer service resolution; prepare items for return and contact customer service to obtain RMA number. Once reconciled, arrangements would be made for sending new inventory and the return of unused parts.If you have any questions regarding this issue, please notify Regulatory Affairs Manager at (801) 619-3450. |  | 14 | 14 units (US), 2 (Intnl) | Worldwide distribution: USA including states of: AZ, CA, FL, MI, TX, and UT; and country of: Japan. | Terminated | December 27, 2011 | 354 | 72 | Not_Computer
Z-1728-2011 | 57814 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Hc2 System Software Suite 4.0 version 3.0, available as component in Qiagen user guide, catalog #5050-00162, and as a component within a Qiagen PC, available as catalog # 5050-11140A. |  |  | January 12, 2011 | 2011 | QIAGEN Gaithersburg, Inc. | The software component of a diagnostic medical device may be defective and could cause incorrect patient values. | Software design | The firm, Qiagen, sent a "Notice of Product Field Action" letter dated January 12, 2011 to all customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue the process of running multiple assay protocols on a single microplate to eliminate the potential for inaccurate patient result reporting, and to complete and return the acknowledgment section via fax to QIAGEN 661-702-3854 or scan the document and email to techservice-na@qiagen.com. The firm informed users that they will provide further detailed instructions for a workaround and software fix. A QIAGEN representative will contact the customers who indicated that they run multiple assays on a plate and are currently using the HC2 System Software Version 3.0 (Suite 4.0), to assist in the identification of those patients that may have been affected by the software defect.If you have any questions or concerns, please contact Qiagen Technical Services at 800-344-3631, Option 2. |  | 163 | 163 | The recalled product was sold to medical laboratories nationwide and in Canada. | Terminated | November 14, 2011 | 306 | 163 | Not_Computer
Z-1895-2011 | 57948 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Coulter PrepPlus 2 (part # 378600) is a microprocessor-controlled pipetting and diluting system. |  |  | January 18, 2011 | 2011 | Beckman Coulter Inc. | The recall was initiated after erroneous, though potentially credible, results due to intermittent reagent level sense failures could be generated. During these failures, the instrument fails to produce a warning that a reagent vial is low or out of reagent. The low or absent levels of reagent used during these failures may lead to insufficient antibody being dispensed into a daughter tube there | Other | The firm, Beckman Coulter, sent an "URGENT: PRODUCT CORRECTION" letter dated January 18, 2011 to all their customers. The letter described the product, problem and actions to be taken. The customers were instructed to: 1) Follow PrepPlus 2 Instructions for Use (Part # 4277264-BA) for all reagent and specimen requirements prior to initiating a run on the PrepPlus 2 including:A) Ensuring all dead volume requirements for reagents as indicated in 7.11 and 7.12 of the Sample Processing sections and 12.2 of the Troubleshooting Guide. Appropriate volume of reagent in all vials per the planned number of tests to be run must contain at least the volume listed on the Worklist (plus an additional 10% of the full volume of the container to account for dead volume)".B) Ensuring specimen processing dead volume requirements are met as indicated in section B of Specimen Tube Size to ensure the minimum amount of specimen needed.C) Monitoring for any foaming/bubble accumulation in the reagent vial.2) A review of all clinically relevant results generated from samples prepared on the system be conducted. Abnormally dim or unexpected negative sample staining patterns should be reviewed to ensure correlation with clinical symptomatology and other diagnostic data. If the clinical information Is in accordance with the clinical data released, the validity of the results could be considered accurate and no retrospective action may be needed. Ifinconsistent data is observed, please adhere to good laboratory practices within your institution and consider data re-analysis and re-testing.The customers were also instructed to share this information with their laboratory staff and retain it as part of their laboratory Quality System documentation; if they have forwarded any of the affected product(s) listed above to another laboratory, provide a copy of the letter to them, and complete and return the enclosed PRODUCT CORRECTION RESPONSE FORM via fax to: (786)639-7500/7501/750 |  | 369 | 369 units (140 in the U.S.) | Worldwide distribution: USA and countries including: Australia, Barbados, Belarus, Belgium, Canada, Croatia, France, Germany, Ireland, Italy, Japan, Malawi, Mozambique, Namibia, Poland, Puerto Rico, Russian Federation, Slovenia, South Africa, Spain, Swaziland, Sweden, Switzerland, Ukraine, United Kingdom, Zambia. | Terminated | September 02, 2014 | 1323 | 369 | Not_Computer
Z-1885-2011 | 57743 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | MDI(R), S.A.L.T. Airway Device, Supraglottic Airway Laryngopharyngeal Tube, Intended for use with 6.5mm -9.00mm ET tubes, Including SALT(TM) airway, ET tube clamp, Latex-free securing strap, Tongue blade, Lubricating jelly, NON STERILE, REF 79-580, Microtek Medical, Inc., Columbus, MS 39702, www.microtekmed.com, Manufactured in the Dominican Republic of U.S. Materials |  |  | January 05, 2011 | 2011 | Microtek Medical Inc | Product contains Triad Lubricating Jelly which was recalled. | Other | The firm, Microtek Medical, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated January 17, 2011 to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine their inventory and place any S.A.L.T. Airway Device (79-580) in quarantine; if they have distributed any of the product, they should identify their customers and notify them at once of this recall; complete and return the enclosed URGENT: MEDICAL DEVICE RECALL-RETURN RESPONSE form via fax to: (662)244-3011 or email: tom.bonner@ecolab.com, and return any unused product inventory to Microtek Medical. The firm will provide the customers with a replacement product as soon as product becomes available.If you have any questions in regards to this recall, please feel free to contact the V.P.-Regulator Affairs at 1-888-844-0988 or email: Tom.bonner@ecolab.com. |  | 2230 | 223 cases of 10 units | Nationwide distribution: AZ, CA, CT, FL, IN, MS, NC, NH, PA, TX and WI. | Terminated | June 20, 2011 | 166 | 2230 | Not_Computer
Z-2046-2011 | 58240 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac High Energy Medical Linear Accelerator, a Trilogy Radiotherapy Delivery System; Model #s: H14, H18, H27, H29, HCX; Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | February 24, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | The wedge angle labeling on the wedge tray may not correctly match the wedge body. However, the wedge body is labeled with the correct wedge angle. | Labeling mix-ups | The firm, Varian Medical, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 24, 2011, to all affected users. The letter described the product, problem and actions to be taken. The customers were instructed to IMMEDIATELY inspect all wedges to verify that the wedge angle label on the wedge body matches the wedge angle on the wedge tray; if any wedge is identified with mismatched labeling follow the written instructions provided in the letter; always measure wedge data at installation of new or replacement wedges; verify all treatment accessories of "matched" linear accelerators, and please advise the appropriate personnel working in your Radiotherapy Department of the content of this letter.Varian has enhanced their manufacturing quality control process to further ensure correct labeling of wedges. Also, Varian will replace any wedge that is identified as mislabeled, at the request of the customer. If you require further clarification, please feel free to contact your local Varian Customer Support Service Manager at: 1-888-VARIAN5 (888-827-4265) for USA and Canada; +14 41 749 8844 Europe; North American: support-americas@varian.com; Australia/New Zealand:support-anz@varian.com; Europe: support-emea@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.apps.helpdesk@varian.com; Latin America:soporte.al@varian.com. |  | 5541 | 5541 units | Worldwide distribution. | Terminated | November 05, 2012 | 620 | 5541 | Not_Computer
Z-2048-2011 | 58331 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Percutaneous Sheath Introducer Kit |  |  | March 25, 2011 | 2011 | Arrow International Inc | Labeling discrepancy - the outer corrugate label as well as a card inside the package state the product contains dry natural rubber - this is not duplicated on the unit packaging. | Other | The firm, Arrow International, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated March 25, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately quarantine any products with the listed lot numbers; complete and return the Immediate Attention Request form via fax to 800-343-2935 indicating how much product they have in stock and whether they need labels. In the event that an alternative approach is needed, the customers are to contact Arrow Customer Service for more information. Once the labels have been received and applied the product can be released from quarantine.An Arrow's Sales Representative will pay a personal visit to their customers and obtain the quantity and the amount of labels they need. If you have any other questions, feel free to contact your local sales representative or Customer Service at 866-396-2111. |  | 840 | 840 | Nationwide distribution: USA including states of: PA and NJ. | Terminated | January 29, 2014 | 1041 | 840 | Not_Computer
Z-2754-2011 | 59111 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic 6F Zuma Guide Catheter with AL 1.5 Curve with side holes |  |  | June 08, 2011 | 2011 | Medtronic Vascular | Mislabeled; Product labeled as 6F Zuma Guide with AL 1.5 curve with side holes may contain a 5F Zuma Catheter with SAL 1.0 curve | Labeling Change Control | The firm, Medtronic, sent an "Urgent Medical Device Recall Notice(for US)" letter dated June 7, 2011 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to segregate the product so that it will not be used. A Medtronic Sales Representative will visit the customers to account for all units they received and initiate a credit as appropriate.If you have any questions, please call (978) 739-0371 or (978) 739-3259. |  | 5 | 5 units | Nationwide distribution: USA including states of: CA, FL, and MA. | Terminated | April 17, 2012 | 314 | 5 | Not_Computer
Z-0368-2012 | 60419 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer ZMR Femoral Spout &amp; Cone Bodies and Trabecular Metal Hip Prosthesis Femoral Stems: |  |  | November 16, 2011 | 2011 | Zimmer Inc. | Lack of assurance of proper metal fatigue strength due to metal grain structure anomaly. In 2006 a recall was initiated as a result of the grain anomaly; however, these 4 lots had been split off during normal production and had not been included in the 2006 recall. | Nonconforming Material/Component | On 11/16/11, Zimmer sent letters titled "URGENT MEDICAL DEVICE RECALL" to distributors via electronic mail and to risk managers and surgeons via FedEx indicating that the product must be immediately removed from use and quarantined." Stop using the affected lots and place affected product in quarantine." Review available product on the shelf" Identify those part/lot combinations listed in the recall communication" Immediately remove any implants found that match part/lot combinations listed in the recall communication and quarantine" Return affected lots and signature of acknowledgement to Zimmer |  | 1 | 1 unit (Distributed to Demark) | Worldwide distribution: USA (NY), AUSTRALIA, SWITZERLAND, and DENMARK | Terminated | August 28, 2013 | 651 | 4 | Not_Computer
Z-0870-2012 | 60786 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Model 73X Uni-Vent |  |  | November 04, 2011 | 2011 | Impact Instrumentation, Inc. | A piece of internal tubing in the 73X portable ventilator is of incorrect specifications for the device. | Component design/selection | Impact Instrumentation Inc. sent an URGENT Device Product Recall letter dated November 3, 2011 to all affected consignees. The letter identified the product, problem, instructions for product returns, and actions to be taken. The letter states that Impact will arrange for expedited return, at no charge to the customers of the affected units. The notice instructs customers to complete and return the Recall Tracking Form. If you need additional information, please call ImpactÂż Customer Service at (800) 969-0750 |  | 57 | 57 units | Worldwide Distribution - USA (nationwide) | Terminated | May 23, 2013 | 566 | 57 | Not_Computer
Z-0916-2012 | 60762 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | Drystar Mammo. |  |  | March 02, 2010 | 2010 | AGFA Corp. | Labeling problem | Error in labeling | The firm, AGFA Corporation spoke to the customer directly via telephone concerning the mislabeled affected products. The firm also contacted the customer via e-mail on April 13, 2011 explaining the outcome of the customer complaint. Replacement product was sent to the customer. |  | 2 | 2 boxes | Product distributed in Florida. | Terminated | February 14, 2012 | 714 | 2 | Not_Computer
Z-0014-2012 | 59949 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Dynamic Multileaf Collimator (DMLC) |  |  | May 25, 2011 | 2011 | Elekta, Inc. | It is possible to load different patients in the record and verify (R&V;) system and in the multilead control software (MCS). | Device Design | Elekta, Inc. sent an "IMPORTANT NOTICE" letter dated May 25, 2011 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The letter recommends several precautions customers should use until Elekta finds a solution. Customers will be contacted when a solution is found. |  | 27 | 27 | Nationwide Distribution-USA including the state of CT, FL, GA, IA, MN, MO, NJ, NY, ND, OH, OK, PA, RI, TX, VA , and WA. | Terminated | January 14, 2014 | 965 | 27 | Not_Computer
Z-3031-2011 | 59335 | No | JKD | C-peptides of proinsulin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT C-Peptide Reagent Kit; a chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of C-Peptide in human serum, plasma and urine on the ARCHITECT i System; Produced by Biokit S.A., 08186 Barcelona, Spain for Abbott Diagnostics Division, Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA and Abbott, 65205 Wiesbaden, Germany; list 3L53-25 |  |  | July 11, 2011 | 2011 | Abbott Laboratories | C-Peptide lots 03610K000 and 01711A000 have the potential to generate falsely elevated results with certain patient samples and with non-Abbott Controls. | Nonconforming Material/Component | Abbott Laboratories sent a "PRODUCT RECALL" letter dated July 11, 2011 to all affected customers. The letter describes the product, problem, and actions to be taken by the customers.The letter instructs customers to discontinue use of the lots and destroy any remaining inventory of the lots according to laboratory procedures. A Customer Reply form was provided to customers to complete and return via fax at 1-800-777-0051.Customers can contact Abbott Customer Service at 1-877-4ABBOTT for questions regarding this notice. |  | 3049 | 3,049 kits | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Algeria, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, India, Indonesia, Ireland, Italy, Japan, Jordan, Latvia, Libya, Lithuania, Luxembourg, Malaysia, Mauritania, Mexico, Netherlands, Oman, Pakistan, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Korea, Spain, Switzerland, Taiwan, Trinidad &amp; Tobago, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | February 10, 2012 | 214 | 3049 | Not_Computer
Z-0336-2011 | 56939 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | ENDOPATHÂż XCEL" with OPTIVIEW" Technology trocar |  |  | September 27, 2010 | 2010 | Ethicon Endo-Surgery Inc | This recall has been initiated because Ethicon Endo-Surgery is aware of an increase in reports of leaking in the ENDOPATHÂż XCEL" with OPTIVIEW" Technology trocar, which potentially can lead to loss of pneumoperitoneum during a surgical procedure. Pneumoperitoneum leaks may lead to conversion from laparoscopic to open surgery. | Device Design | Ethicon Endo-Surgery sent an URGENT: DEVICE RECALL EVENT 2244 letter, dated October 8, 2010. The letter identified the product, the problem, and the action to be taken by the customer. The customers were instructed to stop using the product, examine all inventory to determine if they had the affected product on hand, and remove the affected products from inventory. Ethicon Endo-Surgery requested that each customer complete the Business Reply Form and fax it back to Stericycle at (877) 272-6141 with in 3 business days and also return any affected product to Stericycle.For question regarding this recall call (877) 598-5706 and reference 2244. |  | 671 | 671 lots totalling 2,479,134 instruments | Worldwide distribution: USA, ARGENTINA, AUSTRALIA,BELGIUM,BRAZIL,CANADA,CHILE,CHINA,COLOMBIA,COSTA RICA, DOMINICAN REPUBLIC,ECUADOR, EL SALVADOR, GREECE, GUATEMALA, HONG KONG, ISRAEL, ITALY, JAPAN,JORDAN, KOREA, REPUBLIC OF, LEBANON, NETHERLANDS ANTILLES, NICARAGUA, PANAMASAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICASPAIN, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATESURUGUAY, and VENEZUELA. | Terminated | January 25, 2012 | 485 | 671 | Not_Computer
Z-2095-2010 | 55465 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Evolis-Femoral Sizer Posterior Reference |  |  | September 11, 2009 | 2009 | Medacta Usa Inc | Medacta USA Inc is initiating a recall on the Femoral Sizer, an orthopedic manual surgical instrument, because the "L" and "R" markings are mislabeled which would cause incorrect usage. | Process design | The firm, Medacta USA Inc, sent an "Urgent Safety Notification" letter dated August 5, 2009, to the customers. The letter describes the product, problem and the action to be taken by the customers. The customers were instructed to check their inventory for devices listed, if found, to immediately sequester them and DO NOT use them in surgery. The customers were ask to please pack the affected femoral sizer(s) posterior reference for shipment to Medacta USA. Medacta informed the customers that Medacta USA personnel will contact them by telephone to provide them with an RMA number and the reworked femoral sizer(s) posterior reference will be returned as soon as they are available.If you have any questions, please feel free to contact the Medacta USA office at 805-437-7085 or contact Ms. Heather Neely directly at 805-886-1069. |  | 5 | 5 | Worldwide distribution: USA including states of Texas and Idaho and country of France. | Terminated | February 29, 2012 | 901 | 5 | Not_Computer
Z-3163-2011 | 59427 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoMobile, Hi-Art System, H-0000-0003, TomoTherapy 1240 Deming Way, Madison, WI 53717 |  |  | July 07, 2011 | 2011 | Accuray Incorporated | TomoTherapy Inc. is sending this Field Safety Notice to make you aware of an anomaly which may affect the performance of the TomoMobile Hi-Art System.TomoTherapy has discovered that when attempting to open the TomoMobile shielding doors, while the door hinges are in the locked position, the hinges may fail allowing the door to disengage from the shielding. | Device Design | The firm, TomoTherapy, sent an "Urgent Medical Device Correction" letter dated July 7, 2011 to its consignees/customers. The letter described the Issue, Product Affected, Recommended Action (advised consignees to not attempt to open the shielding doors when the hinges are in the locked position. The Hinges should always be in the unlocked position when opening the doors) and Resolution. A TomoTherapy field service engineer or an authorized service representative will contact the customers when a resolution is available.If you have any questions, please contact the TomoTherapy Technical Solutions Center by email at support@tomotherapy.com or phone at: North America: 1-866-368-4807; Belgium: 0800 38783; France: 0805 631 565; Germany: 0800 000 7541; Italy: 800 986 399; Netherlands: 0800 020 1364; Spain: 800 300049; Switzerland: 0800 001927; United Kingdom: 0808 238 6035; China/CNC (Northern): 10 800 712 1701; China (Southern): 10 800 120 1701; Hong Kong: 800 967912; Japan: 0044 22 132374; Singapore: 800 1204 683; South Korea: 0079 81 4800 7204, all other locations: +1 608 824 2900 or +32 2 400 44 44. |  | 4 | 4 | Worldwide distribution: USA including states of: AZ, CA and OK; and country of: United Arab Emirates | Terminated | June 29, 2012 | 358 | 4 | Not_Computer
Z-0036-2014 | 66335 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran Mini CT Scanner |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | A description of the defect in the product or the manner in which the product fails to comply with an applicable Federal standard; The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 ÂżÂż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evalu | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. |  | 320 | 320 | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1507 | 320 | Not_Computer
Z-0419-2015 | 69584 | No | KNH | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Gyrus ACMI Disposable Falope-Ring Band Applicator Kits. REF 006889-901 and REF 005280-901 (with 8 mm Disposable Trocan). Sterile EO. Single Use Only. The disposable Falope-Ring Band Applicator is a sterile, two-band, dual-incision instrument. The two-band applicator enables the physician to occlude both Fallopian tubes without removing the applicator from the abdominal cavity. The Falope-Ring Band and Applicator Kits are indicated for use for female sterilization (permanent contraception). |  |  | October 21, 2014 | 2014 | Gyrus Acmi, Incorporated | All packages of Falope Ring Band Applicator kits are being recalled due to compromises in sterile packaging that could be associated with bacterial contamination that might lead to patient infection. | Process control | Olympus sent an Urgent Medical Device Recall letter dated October 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Olympus requires you to take the following action: I. Immediately cease any further use of any affected product you have, remove it from your inventory and quarantine it until it is shipped back to us. 2. Call your Olympus customer service representative at 1-888-524-7266 to obtain a Returned Goods Authorization so that you may return the product with no charge to you. Olympus will issue a credit t or replacement to your facility y for any returned product. 3. Please note on the enclosed questionnaire that you have received this information. 4. Fax the completed questionnaire to 484-896-7128 regardless of whether you have any affected inventory at your facility. ln addition, if you may have further distributed this product, please identify your customers, notify them at once of this product recall, and appropriately document your notification process. Your notification to your customers may be enhanced by including a copy of this recall notification letter. |  | 74792 | 74,792 units | Worldwide distribution. US Nationwide, Canada, Vietnam, and Europe | Completed |  |  | 74792 | Not_Computer
Z-1729-2011 | 57631 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension(R) Creatinine Flex(R) Reagent Cartridges |  |  | December 27, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Firm has confirmed the potential for falsely elevated creatinine results with specified CREA Flex(R) reagent cartridge lots when processing tests from a well that has been punctured. This issue may affect creatinine results for quality control (QC) and patient samples. | Material/Component Contamination | The firm, SIEMENS, sent an "Urgent Field Safety Notice" dated December 2010 to all customers. The letter described the product ( Dimension(R) CREA Flex(R) Reagent Cartridges DF33A lots BB1286, CA1260 and FB1251), problem and actions to be taken. The customers were instructed to please discontinue use and discard any remaining inventory of the product; complete and return the fax form to the Customer Service Support for replacement of the product; if they must continue use of the product until they receive a replacement, open a new well set every 8 hours to minimize any impact to CREA test results, and complete and return the FIELD CORRECTION EFFECTIVENESS CHECK form via fax to the Technical Solution Center at 302-631-8467.If you have any questions, contact the Technical Solution Center at 800-441-9250. |  | 20177 | 20,177 | Worldwide distribution: United States and the following foreign countries: Austria, Belgium, Bosnia and Herzegovina, Canary Islands, Croatia, Czech Republic, Egypt, Estonia, France, Germany, Great Britain, Greece, Hungary, Italy, Latvia, Lithuania, Kuwait, Macedonia, Morocco, Norway, Poland, Portugal, Qatar, Romania, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland, Turkey and the United Arab Emirates. | Terminated | December 21, 2013 | 1090 | 20177 | Not_Computer
Z-0041-2014 | 66344 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT CT Scanner ID ENT00001 |  |  | December 14, 2009 | 2009 | Xoran Technologies, Inc. | Failure of the gantry to rotate. The complaint alleged that the scanner would emit radiation but would not rotate. | Radiation Control for Health and Safety Act | The service report in response to this complaint stated that the Head select relay board, chip and fan were replaced. The machine (ENTOOOO I) was subsequently tested. All tests were successful. |  | 320 | 320 | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1732 | 320 | Not_Computer
Z-2080-2012 | 62432 | No | IZH | Mammographic x-ray system. | Radiology | Radiology | 510(k) | Bella Blanket Protective Coverlet, for Mammography, 18cm x 25cm or 24cm x 30cm, Product Code: 311. The purpose of the product is to remove the cold by placing this sheet between the patients breast and the imaging receptor plate. |  |  | July 02, 2012 | 2012 | Beekley Corporation | Imaging artifact can occur on mammogram result | Nonconforming Material/Component | Beekley Medical issued "Urgent: Medical Device Recall" notifications dated June 29, 2012 that were sent via certified mail USPS on July 2, 2012. The product issue was was described and customers were provided the recommended actions. Customer Service number is 1-800-233-5539. |  | 86000 | 86,000 units | International Distribution including Nationwide(USA), and the countries of Japan and Canada. | Terminated | July 09, 2013 | 372 | 86000 | Not_Computer
Z-1948-2013 | 65472 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Colorimetric (CM)/ Rate (RT) incubator wear pads. Product Codes: 1) 6801375, 2) 6801890, 3) 6802445, 4) 6802413 |  |  | June 06, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling CM/RT wear pads which are a component in the CM/RT incubator due to becoming worn before their replacement. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated June 6, 2013 via FedEx to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that an OCD trained service personnel will contact them and schedule a replacement of the CM/RT wear pads. OCD requested customers to complete and return the enclosed Confirmation of Receipt form no later than June 17, 2013. Customers are asked to place a copy of the notification by each VITROS System in their facility. Foreign affiliates were notified by email on June 6, 2013 of the issue and instructed to notify their consignees of the issue and actions. If you have any questions or need additional information, please call Customer Technical Services at 1-800-421-3311. |  |  |  | Worldwide Distribution - USA Nationwide in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, including Puerto Rico and the countries of Mexico, Singapore, Venezuela, England, Germany, Italy &amp; Spain. | Open, Classified |  |  |  | Not_Computer
Z-0039-2014 | 66338 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran Minicat00158 |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evaluation of the hazards reasonably related to defect or the failure to comply with the Federal standard; An evaluation of the health ha | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  |  |  | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1507 |  | Not_Computer
Z-0040-2014 | 66340 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT00202 |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xorans specifications. | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  | 320 | 320 | Worldwide Distribution - Nationwide | Terminated | September 11, 2014 | 1507 | 320 | Not_Computer
Z-0756-2013 | 64010 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Buffered Sodium Citrate (9NC) Plus Blood Collection Tubes; Sterile; IVD; REF 363083; 13 X 75 mm, 2.7 mL |  |  | November 12, 2012 | 2012 | Becton Dickinson &amp; Company | BD is conducting a recall of the BD Vacutainer 2.7mL Plus Citrate Blood Collection Tube due to extended Partial Thromboplastin Time (aPTT) test results reported in two customer complaints. | Material/Component Contamination | BD sent "ATTENTION: PRODUCT RECALL" letters dated November 12, 2012 via UPS to distributors and customers instructing them to examine their inventory immediately and discontinue the shipment of the affected product. All customers were instructed to return all products from the lot affected, by following the instructions on the enclosed packing slip,and BD will send them replacement products at no charge. Contact the firm at 10201-847-4267 for assistance. |  | 2944700 | 2,944,700 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, India, Japan, Nepal, Malaysia, Singapore, Brazil, Costa Rica, and Virgin Islands. | Terminated | September 06, 2013 | 298 | 2944700 | Not_Computer
Z-0445-2013 | 63410 | No | LLL | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | The Is-anti-Sm test kits****Catalog No: 720-240*** |  |  | August 06, 2012 | 2012 | Diamedix Corporation | Diamedix Corporation recalled their Is-anti-Sm Test Kit due to the Positive Control recovering high out of specification. | Nonconforming Material/Component | Diamedix sent a Customer Notification letter to all affected customers via Certified Mail. The letter identifed the affected product. problem and actions to be taken. The letter instructed customers to contact Diamedix Technical Service Department at 1-800-327-4565 or 305-324-2300 to arrange for the return and replacement of any remaining affected product. For questions call 305-324-2314. |  | 65 | 65 Test Kits | Worldwide Distribution - USA Nationwide including the states: CA, ID, NJ, NY, OH, PA, PR, RI, TN, and TX, FL, and the countries of Italy, Chile and Ecuador | Terminated | August 08, 2014 | 732 | 65 | Not_Computer
Z-2149-2014 | 68665 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SIEMENS ADVIA Centaur CP Immunoassay System BNP kit; SMN 10309044 - 100 test |  |  | June 23, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Complaints of failed calibration. | Other | Siemens sent an Urgent Medical Device Correction (UMDC) letter to all affected Siemens Healthcare Diagnostics customers in the United States on June 23, 2014 via Federal Express. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on June 23, 2014 via hard copy, email, and/or fax per each country's local regulations. These notices inform the customer of the potential for low calibrator deviation failures when using ADVIA CentaurÂż Systems BNP kit lots ending in 172 on the ADVIA CentaurÂż CP due to an error on the master curve card. Siemens informs customers if a successful calibration is achieved, the results obtained are valid. If customers are unable to obtain a successful calibration on the ADVIA CentaurÂż CP, they should switch to an alternative lot of BNP assay. The notices should be reviewed with the facility's laboratory or medical director.Fax-back forms will be used to confirm that the customers received the UMDC. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Medical Device Correction and completed the required actions. Customers with questions were instructed to contact their Siemens Customer Care Center or their local Siemens technical support representative.For questions regarding this recall call 877-229-3711. |  | 1191 | 1,191 kits | Worldwide Distribution - USA (nationwide including Puerto Rico) and Internationally to Canada, Albania, Argentina, Australia, Bosnia Herzeg., Brazil, Chile, Colombia, Egypt, Georgia, Hong Kong, India, Indonesia, Iran, Iraq, Israel, Japan, Kazakhstan, Lebanon, Malaysia, New Caledonia, Oman, P.R. China, Pakistan, Philippines, Qatar, Rep. of Yemen, Republic Korea, Russian Fed., Saudi Arabia, Serbia, Singapore, South Africa, Taiwan, Thailand, U.A.E., Ukraine, Venezuela, Vietnam, Austria, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece,Hungary,Iceland, Ireland, Italy, Latvia, Lithuania, Netherlands, Norway, Poland, Portugal, Romania,Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 31, 2016 | 647 | 2152 | Not_Computer
Z-1793-2012 | 61332 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Columbus Tibia Plateau (component of the Columbus REVISION Knee System) |  |  | February 21, 2012 | 2012 | Aesculap Implant Systems | During an internal quality inspection, the tibia plateau was found to be out of specification. There is a potential for misalignment between the tibia plateau and attachment of the extension stem. | Nonconforming Material/Component | Aesculap Implant Systems (AIS) initiated their recall of this product by first telephoning consignees and then on February 21, 2012 by sending an IMPORTANT RECALL NOTIFICATION letter dated February 21, 2012 to all consignees. The letter identified the affected product, problem and actions to be taken. The notification told consignees not to use the affected product and to return any stock to AIS (USA) promptly. For questions contact Scott Stephan at (610) 984-9239 or Christian Gabriel at (610) 984-9300. |  | 8 | 8 | Nationwide Distribution including PA and TX. | Terminated | January 15, 2014 | 694 | 8 | Not_Computer
Z-0694-2012 | 60851 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | MERZ AESTHETICS, 27G X 1.0 " (25MM) Cannula, Sterile R, QTY 20, Model 9015M1X20, Distributed by Merz Aesthetics, In., Manufactured by TSK Laboratory. |  |  | December 16, 2011 | 2011 | Merz Aesthetics, Inc. | Merz Aesthetics is initiating a medical device recall on TSK Closed Tip Cannula (27GA x 25mm) due to a labeling error on the secondary (outer) packaging. The outer package label states that the carton contains Cannulas that are 27GA x 25 mm in size instead of 27G x 38 mm. Length of the cannula is a matter of user preference and is unlikely to impact patient safety. | Packaging process control | Merz Aesthetics Inc, initially contacted customers by telephone on December 16, 2011. They were informed that the wrong size cannula were sent with their last order and would need to be returned and a replacement order would be sent. . Merz Aesthetics sent a "Urgent Medical Device Recall" letter dated January 6, 2012. The letter identified the product the problem and the action needed to be taken by the customer. In addition, if you have further distributed this product, please identify your customers and notify them of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.You may contact Merz Aesthetics Customer Service at (866) 862-1211 or (262) 835- 3333 ,for instructions regarding the return and immediate replacement of the product, or if you have any further questions or concerns.. |  | 5 | 5 boxes (20 units per box) | Nationwide Distribution including the states of CA, NY, OK, and TN. | Terminated | July 27, 2012 | 224 | 5 | Not_Computer
Z-0037-2014 | 66337 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT CT Scanner |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | A description of the defect in the product or the manner in which the product fails to comply wvith an applicable Federal standard; The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evalu | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  | 320 | 320 | Worldwide Distribution - Nationwide | Terminated | September 11, 2014 | 1507 | 320 | Not_Computer
Z-0137-2014 | 66353 | No | JCA | Bleeding time device. | Hematology | Hematology | 510(K) Exempt | Surgicutt fully automated incision making device for bleeding time determination. ITC 8 Olsen Avenue, Edison, NJ 08820. 732.548.5700. www.itcmed.com. A subsidiary of Thoratec Corporation. |  |  | August 22, 2013 | 2013 | ITC-Nexus Dx | ITC has determined that the blade in their Surgicutt with Blotting Paper (Product Code: SUB50I; Lot: AK301M) devices may be loose. | Nonconforming Material/Component | ITC sent a Medical Device Recall letter dated September 20, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove/quarantine affected product in their inventory and to return the affected product to ITC. Customers with questions were instructed to call 1-800-631-5945.Questions regarding this recall call 732-548-7000, ext 4476. |  | 300 | 6 boxes (Each box contains 50 Surgicutt Devices plus 55 sheets of Blotting Paper). | Nationwide Distribution including Arizona andTexas.. | Terminated | June 09, 2014 | 291 | 300 | Not_Computer
Z-1917-2015 | 71240 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD) |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Abnormal Power Source Switching. | Software design | HeartWare sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the enclosed Patient Communication and familiarize the contents, forward the notice to those individuals withing their organization who need to be aware of its contents, identify their patients currently supported by the HVAD System, distribute the Patient Communication to their patients via FedEX or some other reliable means of communication. Customers with questions were instructed to contact their HeartWare representative should they need any assistance with the process. Customers were asked to complete, sign, and return the Acknowledgement and Completion Form to HeartWare. Customers with quesitons were instructed to call 1-888-494-6365. |  | 3747 | 3,747 currently in use. | US including AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxemburg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Not_Computer
Z-1912-2015 | 71305 | No | MAL | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Gelseal Straight Diameter 8mm Length 15cm. |  |  | May 15, 2015 | 2015 | Vascutek, Ltd. | Product was labeled as a 30cm graft, when in fact the graft in the box was a 15cm graft. | Error in labeling | On 5/15/2015, Vascutek issued a URGENT FIELD SAFETY NOTICE to their consignee informing them of the corrective action of the affected product. The information included details of the affected devices, including potential risks and actions to be taken by medical staff and distributors. |  | 12 | 12 | Distributed in the state of Michigan (USA), and the countries of United Kingdom, Australia, Poland, and Germany. | Terminated | October 20, 2015 | 158 | 12 | Not_Computer
Z-1914-2015 | 71517 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic DLP Femoral Arterial Cannula, models 57414, 57417, and 57421, and DLP Femoral Cannula Replacement Kit, Models 96017 and 96021. Sterilized using Ethylene Oxide. |  |  | June 12, 2015 | 2015 | Medtronic Perfusion Systems | Medtronic is recalling select lots of DLP Femoral Arterial Cannula product due to recent reports in which the user had difficulty or was unable to connect the cannula to the perfusion circuit tubing. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter, dated June 2015 to customers beginning June 12, 2015. The letter identified affected product, described the issue, requested that unused inventory be quarantined and returned to Medtronic. Customers should contact Customer Service at 800-854-3570 and reference RADAR #241958 to initiate a return and credit of product. Additionally a Customer Confirmation Certificate should be completed and faxed to 651-367-0612. |  | 1304 | 1304 devices | Worldwide Distribution - US Distribution including the states of : AK, AZ, CA, CO, IN, KS, LA, MN, MT, NM, NY, OH, OR, PA, TX, UT, and WI., and to the countries of : Belgium, Brazil, Canada, Colombia, Denmark, France, Germany, Greece, Hungary, India, Ireland, Italy, Korea, Morocco, Netherlands, Poland, Qatar, Russian Federation, Saudi Arabia, Slovakia, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | March 09, 2016 | 271 | 1304 | Not_Computer
Z-1849-2015 | 71451 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | The VariSource iX series afterloader systems are computer controlled remote electro/mechanical systems used for medical purposes, for placing a cable incorporating an irradiated iridium seed internally or close by a malignant tumor or tumor bed in a practice known as brachytherapy. |  |  | May 18, 2015 | 2015 | Varian Medical Systems Inc | After an application freeze and restart of the VariSource iX series by power cycling, the Partial Fraction generated by thesystem will not be correct. The application freeze issue affects only the systems equipped with magnetic Hard Disk Drives (HDD). | Software Design Change | All the consignees affected by this recall will be notified by:The Urgent Medical Device Correction Letter/ Field Safety Notification will bedistributed to all affected users, with a description of the problem.The Letter will also be distributed to the Varian Sales, Marketing, and Serviceorganizations, informing them of the issue.Additionally, the letter will be posted to the customer support site:http:/ /www.MyVarian.com. |  | 55 | 55 | Worldwide Distribution. | Open, Classified |  |  | 55 | Not_Computer
Z-1915-2015 | 71361 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | DePuy Synthes 6.5 mm Titanium Recon Screws w/T25 Stardrive (70mm, 75mm, 80mm, 85mm, 90mm, 95mm, 105mm); part of the Expert Lateral Femur Nail System and indicated to stabilize Subtrochanteric fractures, Ipsilateral neck/shaft fractures, Femoral shaft fractures, Impending pathologic fractures, and malunions/Non-unions. |  |  | May 27, 2015 | 2015 | Synthes, Inc. | certain lots of the Titanium Recon Screws (6.5mm ) were found to contain the incorrect insert GP0760-A (Synthes Titanium Buttress Locking Plate System), instead of the correct insert GP0706-E (Synthes Titanium Intramedullary Nails and Components). | Process control | DePuy Synthes sent an Urgent Medical Device Safety Notification, dated May 27, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions Required:Please take the following actions:Âż Review the information listed in this notification.Âż Forward this notice to anyone in your facility that needs to be informed.Âż Complete the attached Verification Section (page 3 of this notification). Please include your name, title, address, telephone number and signature in the spaces provided.Send a copy of the completed Verification Section by:Âż Fax: (888) 912-2182 orÂż Scan/email: Synthes4844@stericycle.comÂż If the Verification Form is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual in page three (3) of this notification.Âż Maintain a copy of this notice.If you have any questions, please call (610)719-5450.Thank you for your attention to this issue. |  | 276 | 276 | Nationwide Distribution. | Terminated | December 18, 2015 | 205 | 276 | Not_Computer
Z-1847-2015 | 71342 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic Navigated Cannulated Taps |  |  | May 26, 2015 | 2015 | Medtronic Navigation, Inc. | Medtronic Navigation is initiating a field correction due to the potential for injury which could occur as a result of cleaning the Medtronic Navigated Cannulated Taps. | Error in labeling | Consignees were notified via Urgent Medical Device Correction Notification on 5/26/15. The letter identified the affected product and described the issue. This field correction does not require the return of the subject products. Revised instructions for use are being provided to the identified consignees. |  | 15798 | 15,798 | Worldwide Distribution. | Terminated | October 23, 2015 | 150 | 15798 | Not_Computer
Z-1848-2015 | 71560 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Slipped Capital Femoral Epiphysis (SCFE) Drivers of the Free-Gliding SCFE Screw System with catalogue number SCF-MLD265 and SCF-MLD273. |  |  | June 15, 2015 | 2015 | Pega Medical Inc. | SCFE Drivers of the Free-Gliding SCFE Screw System are recalled because the handle of the instruments may detach from the shaft if high torques are applied in counter-clockwise rotation. | Reprocessing Controls | Pega Medical sent an Urgent Field Safety-Recall letter via email on June 15, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Advise on action to be taken by the customer: Âż Identify the device with the same catalogue number and lot numberÂż Return the device to Pega Medical Inc. to the address shown below using the FedEx account.For further questions please call the firm at 450-688-5144 x. 242. |  | 4 | 4 units | US Distribution to the states of : FL and TN. | Terminated | December 29, 2015 | 197 | 4 | Not_Computer
Z-1916-2015 | 71424 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Bullet-Tip PEEK VBR/IBF System, 32mm(l) 13mm(H) Implant 32-13-32 UDI (01)008464680341 08(17)191218(10)200827; T-Plus PEEK VBR/IBF System, 10mm(W) x 27mm(l) x 13mm(H) Implant, 6degree lordotic 30-T-13-6 UDI (01) 00846468032920(17) 200116(10) 202726; and T-Plus PEEK VBR/IBF System, 10mm(W) x 36 mm(l) x 13mm(H) Implant, 6degree lordotic 30-T-1036-13-6 UDI (01)00846468032760(17)200115(10)202582. These products are part of the lnterbody Fusion Vertebral |  |  | May 21, 2015 | 2015 | RTI Surgical, Inc. (dba Pioneer Surgical Technology, Inc.) | Three lots of the Bullet -Tip VBR 13mm x 32 mm implants are mislabeled with a light green colored label, used for 16 mm trial instead of a lime green label, used for a 13 mm trial. | Error in labeling | RTI Surgical sent an" Urgent Medical Device Voluntary Recall " Notification dated May 21, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Immediate Actions: We are aware that one of this product was distributed to you. This one device has already been returned to RTI Surgical so no further action is needed. This notice is for your records.RTI takes this error seriously and has initiated a corrective and preventative action investigation to prevent recurrence. We apologize for any inconvenience this may have caused. For further questions please call (906) 226-4489. |  | 18 | 18 | US Distribution to the states of : FL and NY | Terminated | October 29, 2015 | 161 | 18 | Not_Computer
Z-1921-2015 | 71182 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cholesterol Gen.2. |  |  | May 07, 2015 | 2015 | Roche Diagnostics Operations, Inc. | Possible drug interference when using the Trinder reaction method on multiple Cobas Analyzers. Results of Trinder tests may be falsely decreased when the tests are run with urine samples (Creatinineplus and Uric Acid applications only) or blood samples collected before sufficient time has lapsed after the administration of the drugs Acetaminophen and N-acetylcysteine. | Under Investigation by firm | Roche Diagnostics initiated a voluntary recall and the Urgent Medical Device Correction notices were sent via UPS Ground on 05/07/2015 to customers. Customers were notified of the interference and the method sheets and package inserts will be updated with the new interference claims. The firm plans to monitor any non-responding accounts and address those customers with follow-up attempts to distribute recall notices. The firm plans to conduct level B effectiveness checks via fax and phone. Until the new method sheets are available, the Urgent Medical Device Correction serves as the updated labeling. Roche is not requesting return of the affected product. Customers may contact Todd SieskyCorporate Communications Director, Phone: 317.521.3911, Fax: 317.521.4696, E-mail: todd.siesky@roche.com. For technical questions, contact the Customer Support Center at 1-800-428-2336 |  | 6255 | 6,255 analytical units distributed | US (nationwide) including PR. | Terminated | March 24, 2017 | 687 | 6255 | Not_Computer
Z-1919-2015 | 71292 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | LifeHealth, IRMA TRUPOINT model 443914, typically used in point-of-care settings. For the in vitro measurement of various critical care analytes in human whole blood. These analytes includes: pO2, pCO2, pH, iCa, Na, K and Cl. |  |  | April 29, 2015 | 2015 | LifeHealth, LLC | The barometric pressure measurement output of this particular IRMA TRUPoint SN 443914 device will drift over time. | Nonconforming Material/Component | Consignees were sent via e-mail a LifeHealth LLC " Urgent" letter dated April, 29th, 2015. The letter described the Purpose of this letter, Reason for the Voluntary Recall, Risk to Health, and Actions to be taken by the Customer. Advised consignees to immediately stop using the recalled product and return it to LifeHealth. Requested consignees to complete and return the Acknowledgement and Receipt Form. |  | 7 | 7 | Worldwide distribution. US in the state of MI; INDONESIA, UNITED ARAB EMIRATES, LITHUANIA, POLAND, and PERU. | Terminated | October 20, 2015 | 174 | 7 | Not_Computer
Z-1932-2015 | 71302 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ICT Serum Calibrator REF 146-03. |  |  | July 22, 2014 | 2014 | Abbott Laboratories | ICT Serum Calibrator may generate lower than expected Potassium Quality Control (QC) and patient results on the Abbott Architect cSystem. | Process control | Customers who received shipments of the affected lots were sent a Product Correction letter on July 22, 2014. |  | 11215 | 11,215 units | Worldwide Distribution-US (nationwide) including Puerto Rico, and the countries of Afghanistan, Algeria, Andorra, Angola, Argentina, Australia, Austria, Bahamas, Bangladesh, Belarus, Belgium, Bermuda, Bosnia &amp; Herzegovina, Brazil, Brunei, Bulgaria, Burkina Faso, Canada, Cayman Islands, Chile, China, Colombia, Croatia, Curacao, North Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malawi, Malaysia, Mali, Mauritania, Mauritius, Mexico, Montenegro, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Qatar, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Tajikistan, Tanzania, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, Turks &amp; Calicos, UAE, Uganda, United Kingdom, Uruguay, Venezuela, Vietnam, and Zimbabwe. | Terminated | July 21, 2015 | 364 | 11215 | Not_Computer
Z-1918-2015 | 71346 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | StoneLight 30 Laser System (NS3000). |  |  | May 13, 2015 | 2015 | New Star Lasers, Inc. | During internal testing a situation was discovered where when changing the settings, a new setting could not be accepted and potentially result in a discrepancy in the energy delivered (higher or lower) from the value displayed on the Graphical User Interface display. | Software design (manufacturing process) | Medical Device Correction letters were sent to all affected consignees by Federal Express on May 13, 2015. The letter identified the affected product as well as the issue and risk involved. The software will be upgraded to address the issue; and, the firm will contact customers to coordinate installation of the upgraded software. Customers may continue to use their system safely by taking the precautions provided in the letter. Customers are to complete the attached response form. Questions should be directed to Nina Davis at 916-677-1909. |  | 52 | 52 units | Worldwide Distribution -- US, including the states of CA, IL, AZ, FL, OH, MA, MI, and AL; and the country of the Netherlands. | Terminated | October 28, 2016 | 534 | 52 | Not_Computer
Z-1934-2015 | 71484 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Emotion 6; produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angels or spiral planes taken at different angles. |  |  | June 04, 2015 | 2015 | Siemens Medical Solutions USA, Inc | A potential safety issue involving two (2) broken screws and one (1) bowed, which is used to fix the Balance Weight of the DMS in the gantry, was found during a service call at the customer site. | Device Design | A safety advisory notice, dated June 4, 2015, was sent to direct accounts explaining the problem and potential risk to patients and operators. Siemens engineers will begin replacing the bolts in May 2015. |  | 40 | 40 total | Nationwide Distribution. | Terminated | March 24, 2017 | 659 | 40 | Not_Computer
Z-1935-2015 | 71023 | chest drainage kit | PAD | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | Enforcement Discretion | Ocean Water Seal Chest Drain, Single Collection W/AC, |  |  | May 07, 2015 | 2015 | Atrium Medical Corporation | Chest drains have a missing o-ring | Process control | Atrium Medical issued recall letter dated 5/7/15 to inform customers of the steps to take should they encounter affected product. If the O-ring is not present on the patient tubing:a. Do not use the chest drain.b. Replace the entire chest drain including the patient tubing. Open upanother chest drain package and repeat steps 1 to 3. c. Return any units with missing O-rings to Atrium.If the O-ring is present then reconnect in-line connector and proceed with set upas normal. Complete and return Field Action Reply Form to the following e-mail address: connector@atriummed.com or you may fax it to 1-603-386-6590. If you have any questions, please contact your local Atrium/Maquet chest drainage representative or Atrium Customer Service at 1-800-528-7486(603-880-1433), Monday through Friday between 8:00 am and 5:00 pm (EST). |  | 390 | 390 cases | NationwideCanada, Germany | Terminated | April 26, 2016 | 355 | 2080 | Not_Computer
Z-1940-2015 | 71383 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Accurus 2500 Standalone Vitreous Probe REF 8065741018; Accurus 23 GA Standalone Vitreous Probe REF 8065750821; Accurus 25+ 2500 CPM Standalone Vitreous Probe REF 8065751122; Anterior Accurus Probe with Infusion Needle REF 8065803650; Single Use Sterile Probes Rx only. |  |  | June 02, 2015 | 2015 | Alcon Research, Ltd. | Insufficient seal on the outside packaging, potentially affecting the sterility of the product. | Packaging | Recall notification letters were sent to consignees on June 2, 2015. |  | 35475 | 35,475 units | Worldwide Distribution-US (nationwide) and the countries of Argentina, Australia, Austria, Brazil, Canada, Chile, China, Colombia, Czech Republic, France, Germany, Hong Kong, Hungary, India, International APL, Italy, Japan, Kenya, Korea, Kosovo, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Poland, Russia, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | April 20, 2017 | 688 | 35475 | Not_Computer
Z-1942-2015 | 71204 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Data Management System; a clinical data management system that assists medical laboratory professionals with preanalytic and post-analytic functions in conjunction with multiple instruments, the laboratory information system (LIS) and Siemens StreamLAB Automation Solutions. |  |  | April 29, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | System Limitations and Software Issues related to the following features: Custom reference ranges, User-initiated sample and result actions, Auto-verification and quality control rules, and Sample reports. | Software design | Siemens sent an Urgent Field Safety Notice (UMDC # ISW-15-01.A.US) on May 5, 2015 to all affected customers. The letter instructed customers to follow specific work around instructions, review this letter with their Medical Director, complete and return the Field Correction Effectiveness Check attached to this letter within 30 days, retain the letter with laboratory records, and forward it to those who may have received this product.For questions regarding this recall call 800-441-9250. |  | 1355 | 1355 | Worldwide Distribution - US (nationwide) and Internationally to Austria, Belgium, Canada, Switzerland, Germany, Denmark, Spain, France, Hong Kong, Italy, Japan, Korea, Malaysia, Netherlands, Norway, New Zealand, Philippines, Portugal, Saudi-Arabia, Slovenia, and Taiwan. | Terminated | May 11, 2017 | 743 | 1355 | Not_Computer
Z-1850-2015 | 70985 | Yes | NYU | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Palindrome SI Chronic Catheter Kit 14.5 Fr x 19cm |  |  | April 23, 2015 | 2015 | Covidien LLC | Catheters may not meet antimicrobial specification | Nonconforming Material/Component | Medtronic (Covidien) issued recalll letter dated April 24, 2015, via FedEx overnight delivery to all affected customers. Customers were requested to quarantine and discontinue use of the affected devices. Customers were instructed to complete A Recalled Product Return Form for product returns. Customer Service will respond with a Return Goods Authorization (RGA) number and shipping documentsCustomers with questions were instructed to contact their Medtronic representative or Medtronic Customer Service, Monday through Friday, 8am  6:30pm ET, at (800)-882-5878.For questions regarding this recall call 203-492-5000. |  | 96994 | 72,363 US - 24,631 Internationally | Worldwide Distribution - US (nationwide) and Internationally to Argentina, Austria, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Hong, Kong, India, Iran, Ireland, Israel, Italy, Libya, Luxembourg, Malaysia, Mexico, Morocco, Netherlands, Norway, Oman, Panama, Philippines, Poland, Portugal, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, Thailand, UAE, and United Kingdom. | Terminated | July 26, 2017 | 825 | 96994 | Not_Computer
Z-1941-2015 | 71137 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Extractor Pro RX Retrieval Balloon Catheter 9-12 mm Above Retrieval Balloon Catheter; Universal Product Number (UPN) M00547000 |  |  | April 27, 2015 | 2015 | Boston Scientific Corporation | One lot of the Extractor" Pro RX Retrieval Balloon Catheter (9-12mm balloon) contains an incorrect syringe size (12-15mm), which is larger than the size identified on the label. | Employee error | Boston Scientific sent an Urgent Medical Device Recall Removal letter dated April 27, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify any affected product in their inventory, quarantine the affected product immediately and return to Boston Scientific in accordance with the enclosed recall instructions. Customers will receive replacements for all recalled product that is returned to Boston Scientific. Customers were instructed to forward the recall notification to their customers. Customers with questions were instructed to contact their local sales representative.For questons regarding this recall call 763-494-1133. |  | 39 | 39 units | Worldwide Distribution including the following countries: France, Great Britain, Spain, Ireland.. | Terminated | January 27, 2016 | 275 | 39 | Not_Computer
Z-1939-2015 | 71455 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Sechrist Low Profile Wheeled Stretcher, Model No. 21762-110-33H, 21762-110-36H, 21762-110-41H, 21762-135-33HJ, 21762-010-36H, 21762-010-41H, |  |  | June 01, 2015 | 2015 | Sechrist Industries Inc | Sechrist Industries is recalling the Sechrist Low Profile Wheeled Stretcher because it may experience an unexpected incline angle during the height adjustment of the Low Profile Wheeled Stretcher when lifting the patient upward or when lowering the Wheeled Stretcher to a height at which the patient can be loaded/unloaded. | Device Design | Seachrist sent an Urgent: Medical Device Correction letter dated June 1. 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Our records indicate the Sechrist Low Profile Wheeled Stretcher identified in the attached serial number list (Attachment A) are affected by the field correction and are located at your facility. Within the following weeks, a Sechrist trained Field Service Technician will schedule a visit to your facility to install a Hydraulic Wheeled Stretcher Alignment kit which will greatly minimize the degree of the incline angle that can be experienced after it is installed.Customers are informed that within the following weeks, a Sechrist trained Field Service Technician will schedule a visit to their facility to install a Hydraulic Wheeled Stretcher Alignment kit which will greatly minimize the degree of the incline angle that can be experienced after it is installed. For further questions, please call: (714) 579-8318, or (714) 579-8311.We apologize for any disruption that this correction may cause. Sechrist is dedicated to ensuring the safety and quality of our products. If you have any questions about what to do, please contact us. |  | 81 | 81 US Distribution | US Distribution and Internationally to Japan and Jordan. | Terminated | October 20, 2015 | 141 | 81 | Not_Computer
Z-1952-2015 | 70742 | No | EJI | Cartridge syringe. | Dental | Dental | 510(K) Exempt | Milestone Wand STA Handpiece for the delivery of local anesthesia to dental tissue. Model Number STA-5050-305. The Wand STA handpiece is a single use syringe used in conjunction with the CompuDent STA computer controlled dental injection system. |  |  | February 11, 2015 | 2015 | Milestone Scientific, Inc. | Wand Handpiece was mislabeled as STA handpiece. The Wand handpiece cannot be installed in the STA instrument. | Nonconforming Material/Component | Milestone Scientific notified their customers via e-mail and phone on 2/11/2015. |  | 165850 | 165,850 units | Worldwide distribution. US in FL, IL, NY WA and Puerto Rico, Israel, Croatia, Italy, Australia, Germany, Colombia, UK, Sweden, and UAE. | Open, Classified |  |  | 165850 | Not_Computer
Z-1949-2015 | 71448 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | INSTRUMENT ARM DRAPE, IS4000, 20 PACK; used on the da Vinci Xi Surgical System. |  |  | June 16, 2015 | 2015 | Intuitive Surgical, Inc. | Field Safety Notification to advise of an issue that may occur involving the combination of the da Vinci Xi Drapes and Xi instruments. A pinched drape may cause a message on the Vision Cart touchscreen - Missing reload error. | Device Design | Intuitive Surgical sent an Urgent Medical Device Notification letter on June 16 , 2015 to all affected consignees. Consignees were instructed to take the following actions:1. Ensure that all affected personnel are fully informed of this notice. Forward thisnotice to your Risk Manager, OR Director, Purchasing Manager, BiomedicalEngineering staff and members of your medical staff who perform da Vinci Surgeryprocedures.2. Complete and return the attached Acknowledgment Form to Intuitive Surgicalusing the instructions provided.3. Please retain a copy of this notice with your system.4. Refer to the da Vinci Xi System User Manual for more information on draping thesystem.Consignees with questions or needing spport should contact their Clinical Sales Representative or contact Intuitive Surgical Sustomer Service at the numbers listed below:Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:Âż North and South America: (800) 876-1310, Option 3 (6 AM to 5 PM PST) or mail:customersupport-servicesupport@intusurg.comÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 AM to6 PM CET) or ics@intusurg.comÂż South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)Âż Japan: 0120-56-5635 or 003-5575-1362 (9 AM to 6 PM JST). |  | 310 | 310 | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Belgium, Czech Republic, France, Germany, Greece, India, Israel, Italy, Japan, Netherlands, Norway, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | October 05, 2015 | 111 | 310 | Not_Computer
Z-1947-2015 | 71475 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Etest Cefotaxime CT 256 WW and Etest Cefotaxime CT 256 US. in vitro diagnostic. |  |  | June 12, 2015 | 2015 | BioMerieux SA | QC results were out of range resulting in false resistant strains. | Process control | BioMerieux sent an Urgent Product Removal Notices dated June 12, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Required actions:ÂżÂż Please distribute this information to all appropriate personnel in your laboratory, retain a copy in your files, and forward this information to all parties that may use this product, including others to whom you may have transferred our product.ÂżÂż Discontinue using and discard any remaining inventory of EtestÂżÂż CEFOTAXIME CT 256 lots reported in table 1.ÂżÂż Discuss any concerns you may have regarding previously reported results with your Laboratory Medical Director to determine the appropriate course of action.ÂżÂż Complete and return the Acknowledgement Form in Attachment A by Fax to confirm receipt of this notice and to request product replacement.bioMÂżÂżrieux is committed to providing our customers with the highest quality product possible. We sincerely apologize for any inconvenience that this may have caused you. If you require additional assistance or have any questions, please contact your local bioMÂżÂżrieux Clinical Customer Service organization at (800) 682-2666. |  | 674 | 674 | Worldwide Distribution - USA (nationwide) to the states of AL, AR, CA, CO, FL, GA, IL, IN, KY, MD, MI, MN, MO, NE, NY, NC, OH, OK, OR, PA, TN, TX, UT, VA, WA, WV and WI., and to the countries of: Austria, Australia, Canada, Switzerland, Germany, Spain, France, United Kingdom, Greece, Hong Kong, India, Italy, Japan, Netherlands, Poland, Sweden, Thailand, Turkey, Saudi Arabia, Philippines, Slovenia, Kuwait, Brazil, Chile, Czech Republic, Algeria, South Africa, Kosovo, Costa Rica, Belarus, Guam, Honduras and Vietnam. | Open, Classified |  |  | 674 | Not_Computer
Z-1943-2015 | 71336 | No | OCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Velocity" Biopsy Valve, 25 units per box, |  |  | February 17, 2015 | 2015 | US Endoscopy Group Inc | The company has received reports of structural non conformance at the connection between the valve body and the irrigation tube, resulting in separation of the components during use. | Under Investigation by firm | US Endoscopy sent an Urgent Recall Notification Letter dated to their customers. The letter identified the affected product, problem and actions to be taken. US Endoscopy requests that their customers:1. Remove all unused affected Velocity Biopsy Valves from your inventory.2. Complete the attached Customer Response Card and return to the attention of Kia Pennington, via fax 1-440-639-4495 or email kpenning@usendoscopy.com.3. Mark the shipping container with code: RMA xxxxxxxx QIF, and return the product to US Endoscopy via Fed Ex Account #293799547.For questions contact Kia Pennington in Customer Service at 1-800-769-8226. |  | 2173 | 2,173 units | Worldwide Distribution - US Nationwide in the states of: CA, CO, IA, MA, OH, TX &amp; WV and the countries of: Brazil, Costa Rica, Croatia, Czech Republic, Egypt, France, Germany, India, Indonesia, Italy, Latvia, Lithuania, Philippines, Poland, Romania, Russian Federation, Slovenia &amp; South Africa. | Terminated | September 21, 2015 | 216 | 2173 | Not_Computer
Z-1945-2015 | 71526 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Gentle Threads Interference Screw 9x35mm Full Thread. |  |  | June 03, 2015 | 2015 | Biomet, Inc. | Part Number: 905620 Lot: 824610 Gentle Threads Interference Screw 10x25mm Round Head is mixed with the Part Number: 905627 Lot: 753410 Gentle Threads Interference Screw 9x35mm Full Thread. If the incorrect screw is used various failures may occur which may ultimately lead to revision surgery or a delay in surgery. | Mixed-up of materials/components | On 6/3/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors with instructions for locating and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 41 | 41 units | Distribution-Domestic: None. International: Australia, Mexico, Chile, Costa Rica, Netherlands, and Brazil. | Terminated | December 22, 2015 | 202 | 188 | Not_Computer
Z-1948-2015 | 71531 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Natural Knee II Revision Femoral Stem, 125mm, item 6215-12-185. packaged in configuration 803/791. Knee Prosthesis component. |  |  | October 02, 2014 | 2014 | Zimmer, Inc. | Failed to meet established requirements for sterile barrier integrity due to a breach of the inner cavity. If the outer tray is compromised in addition to the inner tray a periprosthetic infection may occur. The risk associated with a periprosthetic infection is a revision or potential multi-stage revision to treat the infection. | Package design/selection | On 10/2/2014, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital Risk Managers. All distributors were notified via electronic mail, and all hospital Risk Managers and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, and consignee responsibilities. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-800-348-2759. Hours of operation are Monday through Friday, 8 a.m. through 8 p.m. ET. |  | 352 | 352 | US Nationwide Distribution, including DC and the states of AL, AZ, CA, CO, FL, GA, IL, IN, KS, LA, MD, MI, MN, MO, MT, NC, ND, NJ, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA, WI, and WY. | Terminated | July 01, 2015 | 272 | 352 | Not_Computer
Z-1946-2015 | 70992 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Axis-Shield Liquid Stable (LS) 2-Part HOMOCYSTEINE REAGENT; 100 test format, 1000 test format, and 200 test format; and Private Label: (Distributed by BECKMAN COULTER, for professional use only, on the BECKMAN COULTER AU platforms (AU400, AU480, AU680 and AU5800)) |  |  | March 20, 2015 | 2015 | Axis-Shield Diagnostics, Ltd. | Deterioration in performance of the Liquid Stable (LS) 2-Part HOMOCYSTEINE REAGENT resulting in under recovery. | Under Investigation by firm | Axis Shield sent an Urgent Field Safety Notice letter dated March 20, 2015 to their customers via email and written notice sent with confirmed delivery receipt. The letter identified the affected product, problem and actions to be taken. Recommended Action to Customer is as follows: Confirm receipt of FSN to Axis-Shield. Do not use the affected Reagent lots referenced in the safety notice. In the event of holding one of the affected lots, immediately contact Axis-Shield Product Support Department to arrange for replacement material and return any defective product to Axis-Shield. Complete reconciliation form. Complete acknowledgement of the safety notice within 10 working days. Ensure safety notice has been distributed and reviewed by all appropriate personnel within your organization. If further distribution of affected lots to another organization has been made, inform them of this Product Recall and provide a copy of this safety notice to them. Retain a copy of the safety notice for their records. For any issues or concerns contact Axis-Shield via e-mail; MDL-AXD-Product-Support@alere.com. |  | 127 | 52 kits + 75 kits Beckman Coulter Brand OUS | Worldwide Distribution - US Nationwide in the states of: KY, WA, CA, and PA and the country of Germany. | Terminated | October 27, 2015 | 221 | 127 | Not_Computer
Z-1961-2015 | 71073 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The Rapidlab 1260 system is a blood gas analyzer used for laboratory testing of blood gases, electrolytes, metabolites and CO-oximetry in arterial, venous and capillary whole |  |  | May 06, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc | D50 and D51 Diagnostic error codes are not functional. | Software design | Siemens sent an Urgent Field Safety Notice April 2015, to all affected customers. Customers were notified of the recall beginning on May 6, 2015 via Federal Express. Foreign customers were notified starting on May 5, 2015 via method determined by each countries local regulations and procedures. The notification informs customers of the issue and asks customers to do the following: Ensure that the Glucose and Lactate sensor contacts are aligned and completely seated with the contacts on the analyzer, Analyze a minimum of 2 levels of quality control material after installing the sensors, as indicated in the RAPIDLab 1200 Operator's Guide, Review this letter with your Medical Director. Complete and return the Field Correction Effectiveness Check Form attached to this letter within 7 days. Retain this letter with your laboratory records, and forward this letter to those who may have received this product. For further questions please call (800) 242-3233. |  | 122 | 122 devices | Worldwide Distribution - Nationwide Distribution including Puerto Rico and to the countries of : Argentina, Australia, Austria, Belgium, Bosnia, Herzog,Brazil, Brunei, Canada, Chile, Colombia, Croatia, Czech Republic, Egypt, Finland, France, Germany,Georgia, Greece, Hong Kong, Iceland, India, Indonesia, Ireland, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, P.R. China, Peru, Poland, Portugal, Qatar, Republic Korea, Romania, Russian Fed, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E., United Kingdom, Venezuela and Vietnam. | Terminated | August 02, 2017 | 819 | 2835 | Not_Computer
Z-1953-2015 | 71129 | No | MMS | Patient lubricant. | General Hospital | General Hospital | 510(k) | Dr. Sheffield's LUBRI GEL 3 oz (85 g) |  |  | April 27, 2015 | 2015 | Sheffield Pharmaceuticals, LLC | Recovery of high count of gram positive rods including single colonies of B. cepacia. | Process control | A letter was sent to inform customers of the voluntary recall of Sheffield Pharmaceuticals Lubri-Gel Personal Lubricant Jelly. The letter described the issue and asked customers to quarantine all related inventory immediately and provide their inventory counts on the specific lots. Questions or concerns related to the recall or letter can be directed to James Turner at James.Turner@sheffield-pharmaceuticals.com. |  | 37636 | 37,636 pieces | Worldwide Distribution -- CA, NJ, FL, PR, NJ, MI, NY, NC, FL, IL, GA, MN, KS, and NC; and, the countries of Jamaica and Georgia. | Open, Classified |  |  | 37636 | Not_Computer
Z-1957-2015 | 71362 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CentraLink" Data Management System; software is a network solution provider and multi-system data manager for the instruments and lab automation systems (LAS) within the lab. The CentraLink software consolidates data |  |  | May 19, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has determined that the sample query function that includes Instrument or Instrument Group as search criteria may not return all samples from the Dimension Vista systems if onboard aliquot support rules are configured in the CentraLink system. | Under Investigation by firm | Siemens sent an Urgent Medical Device Correction letter dated May 19, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be Taken by the CustomerWhen using any of the CentraLink sample queries, do not select Instrument or Instrument Group criteria for Dimension Vista 500 or Dimension Vista 1500. This would also include any preconfigured queries, commands or tool bar buttons that may use these criteria.Your local Siemens service provider will be contacting you to verify and/or update the configuration of your CentraLink system.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions please call (800) 441-9250. |  | 1941 | 1941 | Worldwide Distribution - Nationwide Distribution and to the countries of : Belgium, Canada, Czech Republic, Denmark, France, Germany, Italy, Netherlands, Saudi Arabia, Spain, Switzerland and United Kingdom. | Open, Classified |  |  | 1941 | Not_Computer
Z-1956-2015 | 71521 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | JustRight 5mm Reload, JR-REL25-2.0 Reload is packaged in cases of 12., placed in a tray and sealed with a Tyvek lid. |  |  | June 11, 2015 | 2015 | Justright Surgical, LLC | JustRight Surgical is recalling various lots of JustRight 5mm Stapler product because they may deliver incomplete staple formation and incomplete tissue closure. | Process design | JustRight Surgical sent an Urgent Medical Device Recall letter dated June 12, 2015 to aaffected customers and distributor via an email. The letter identified the affected product, problem and actions to be taken. Custers were instructed Review, complete, and return the enclosed Acknowledgement Form directly to ÂżRepÂż, your Sales Representative or email to orders@justrightsurgical.com or fax to (720) 287-7135.The following assistance is available:" For questions related to the issue or product replacement, please contact your local sales representative, ÂżRepÂż, or contact JustRight Surgical Customer Service at (866) 683-1743, Monday through Friday from 7:30 AM to 5:00 PM MST." For questions related to product pick-up by JustRight Surgical, please also contact your local sales representative or Customer Service (866) 683-1743. Return instructions are provided on the Acknowledgement Form." For questions related to the Acknowledgement Form and its return, please contact the JustRight Surgical Director of Quality Systems at (720) 287-7146. |  | 1452 | 1452 | Distribution - US Nationwide and Ireland. | Terminated | August 26, 2015 | 76 | 2274 | Not_Computer
Z-1965-2015 | 71434 | No | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Hamilton Medical Infant Flow Sensor, single use, (3.1m), Part number 260179. The infant flow sensor is an accessory for the Hamilton Medical Ventilators. |  |  | June 01, 2015 | 2015 | Hamilton Medical, Inc. | An issue has been discovered with the Hamilton Medical Infant Flow Sensor, single use. When a ventilation mode with adaptive volume control (APVcmv/CMV+) is used, the ventilator may adjust the ventilation on the inaccurate flow measurement and decrease the amount of ventilation delivered. | Device Design | Hamilton Medical Inc sent a Medical Device Safety Notices letter dated May 8, 2015 by traceable method to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact Technical Support Manager at 800-426-6331, Ext 215. |  | 154 | 154 total, all part numbers | US Nationwide Distribution. | Terminated | August 12, 2015 | 72 | 154 | Not_Computer
Z-1972-2015 | 71507 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | SYNGO IMAGING (VERSION V30 and V31); Syngo Imaging is a Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images, including digital mammography images. |  |  | June 11, 2015 | 2015 | Siemens Medical Solutions USA, Inc | During certain clinical workflows safety issues may occur involving monitor assignment order and incorrect display and measurements for images with non square pixel size. | Software design | Siemens sent a Customer Advisory Notice, dated June 11, 2015, to their customers. The letter identified the affected product, problem and actions to be taken. Customers were also asked to check the software version and consider whether the issue impacts clinical routine and if a software update is necessary. |  | 5 | 5 | US in the states of NC and NE | Terminated | March 24, 2017 | 652 | 5 | Not_Computer
Z-1959-2015 | 71469 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound Systems between software versions VA16A and VA30A and with the Stress Echo Option. |  |  | May 29, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | When performing a 2D Stress Echo study, some following keystrokes result in a potential loss of data, where some selected clips may not be saved as part of the study. | Software design | Siemens sent an Customer Safety Advisory Notification on May 29,2015, to all affected customers. The letter identified the product the problems and the action needed to be taken by the customer.Customer Safety Advisory notifications inform users of the potential issue with the ACUSON SC2000 systems with software versions between VA16! and VA30A and the Stress Echo feature.Letters describe the steps that result in the potential loss of data and suggest that the following steps be taken to remedy the situation:1. Select Stress Echo Selection2. If applicable relabel the clips for all Stages and Views3. Do not return to Stress Echo Selection to select additional relabeled clips.The issue will be resolved in an upcoming software release for customers who did not receive the original software update. Until the necessary repairs have been completed, all personnel in the Organization who need to be aware of the issue should be notified. For further questions please call (650) 694-5398. |  | 573 | 573 devices | Worldwide Distribution - US Distribution and to the countries of : Albania, Angola, Armenia, Australia, Austria, Bangladesh, Belarus, Bolivia, Bosnia Herzog, Brazil, Brunei, Bulgaria, Canada, Colombia, Czech Republic, Ecuador, Egypt, Finland, France, Germany, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Lithuania, Malaysia, Maldives, Mauritius, Mexico, Netherlands, Nigeria, Oman, P.R. China, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed, San Marino, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Suriname, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E., United Kingdom and Venezuela | Open, Classified |  |  | 573 | Not_Computer
Z-1958-2015 | 71270 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Cios Alpha; The Cios Alpha is a mobile X-ray system designed to provide X-ray imaging of the anatomical structures of patient during clinical applications. Clinical applications may include but are not limited to: interventional fluoroscopic, gastro-intestinal, endoscopic, urologic, pain management, orthopedic, neurologic, vascular, cardiac, critical care and emergency room procedures. The patient population may include pediatric patients |  |  | May 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc | under certain circumstances the Cios Alpha system may freeze during a procedure. | Software design | Siemens sent an Safety Advisory Notice dated May 5 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They were alerting them about the issue and risks involved. Customers were provided work-around instructions until a software fix is released. For further questions please call (610) 219-6300. |  | 37 | 37 | Nationwide Distribution | Terminated | January 13, 2017 | 619 | 37 | Not_Computer
Z-1967-2015 | 71352 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Infinite F500, in vitro diagnostic. |  |  | May 29, 2015 | 2015 | Tecan US, Inc. | A firmware bug causes incorrect values to be displayed for specific measured luminescence values. | Device Design | Tecan sent an Urgent Field Corrective Action letter dated May 29, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached Customer Response Form. For questions contact your local Tecan Helpdesk. |  | 868 | 868 in total | US Nationwide Distribution in the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, NE NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, including DC | Terminated | August 18, 2017 | 812 | 868 | Not_Computer
Z-1954-2015 | 71505 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Treatment Planning and Delivery System Software version 3.6. ViewRay. |  |  | May 07, 2015 | 2015 | Viewray Incorporated | The software was failing to determine new patient locations if imaging is not enabled during treatment. | Software design | On May 7, 2015 the firm sent Customer Advisory Notification Letters to their customers. |  | 3 | 3 | Distributed in the states of CA, MO &amp; WI. | Terminated | February 04, 2016 | 273 | 3 | Not_Computer
Z-1960-2015 | 71497 | No | ETN | Surgical nerve stimulator/locator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Spine Wave Access Kit, Catalog Number 10-1530 |  |  | May 04, 2015 | 2015 | Spine Wave, Inc. | Mislabeled outer carton:Monopolar Probe is listed as 200 x 1.5mm when it should be listed as 280 x 1.5mm | Packaging change control | Spine Wave sent an Urgent: Medical Device Recall letter dated June 11, 2015 to affected customers. The letter identified the affect product, problem and actions to be taken. Customers were instructed to acknowledge receiptof notification by returning the enclosed Recall Notice Acknowledgement Form in the stamped envelope provided.For questions, contact Spine Wave directly at 203-944-9494and ask for the Recall Coordinator or by email at rsmith@spinewave.com. |  | 121 | 121 units | US Nationwide Distribution in the state of CA, FL, GA, LA, MA, MD, MI, MO, NC, NH, OH, OR, PA, TX, and VA. | Open, Classified |  |  | 121 | Not_Computer
Z-1966-2015 | 71572 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision 500D and Advantx Legacy/Legacy-D Radiographic and Fluoroscopic Systems. |  |  | June 12, 2015 | 2015 | GE Healthcare | The SFD/IDD safety mechanism may not engage properly at the lower range of SFD/IDD travel. This mechanism is installed to hold and prevent the SFD/IDD from falling due to a counterweight cable failure. The lower range of travel is 1.1 in (28mm) from the lowest point of normal SFD/IDD vertical compression. A fall of the SFD/IDD could result in an injury to a patient or operator. | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref#10902 dated June 12, 2015 to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology and Radiologists/Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 4053 | 4,053 (3,804 US; 249 OUS) | Worldwide Distribution - US Nationwide including PR, GU and DC and the countries of: Afghanistan, Argentina, Australia, Austria, Belgium, Bermuda, Canada, Cayman Islands, Chile, Colombia, Czech Republic, Dominican Republic, Ecuador, Egypt, France, Germany, Ghana, Guatemala, Honduras, Indonesia, Ireland, Israel, Italy, Jamaica, Jordan, Korea (Republic of), Kuwait, Libyan Arab Jamahiriya, Malaysia, Malta, Mexico, New Zealand, Panama, Romania, Russia, Saudi Arabia, Spain, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Yemen. | Terminated | December 21, 2015 | 192 | 4053 | Not_Computer
Z-1973-2015 | 71508 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima IGS 320. |  |  | May 29, 2015 | 2015 | GE Healthcare | GE Healthcare has recently become aware of a potential safety issue that may result in the loss of C-arm gantry motion capability on fluoroscopic Interventional imaging systems (Innova, Optima and Discovery). | Nonconforming Material/Component | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter on 5/29/2015- GEHC Ref#12231- dated May 29, 2015 to their Consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and contact Information. The customers were instructed that you may continue to use the system and avoid rapid direction changes with the gantry motions and prevent potential collisions with other equipment. If the failure occurs during interventional examination, please ensure that you have established procedures for handling patients in case of the loss of gantry motion as stated in your product labeling.GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 3 | 3 | Worldwide Distribution: US (nationwide) including states of: CA, DE, FL, GA, IL, IN, KS, MS, NV, OH, OK, OR, PA, TN, TX, VA, WA and WI; and countries of: ARGENTINA, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, FRANCE, GERMANY, HONG KONG, HUNGARY, INDIA, INDONESIA, ITALY, JAPAN, RUSSIA, UNITED KINGDOM, and VIETNAM. | Terminated | October 28, 2015 | 152 | 83 | Not_Computer
Z-1986-2015 | 71443 | No | ISD | Measuring exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Primus (THE BTE WS30), model PrimusRS (PRRS) and Primus+ (PR30) a measuring exercise equipment device designed to provide active and passive torque resistance to patient motion during rehabilitation therapy. The Primus that includes the optional Chop / Lift Bar. |  |  | April 09, 2015 | 2015 | BTE Technologies, Inc. | On the Primus that includes the optional Chop / Lift Bar, there is a remote possibility that the Chop / Lift Bar could get accidently detached from the snap hook in specific use scenarios when the bar is rotated and twisted around the snap hook. The snap hook could open, resulting in possible injury to the user due to a fall. | Device Design | BTE sent an Medical Device Correction letter dated April 10, 2015, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer.In order to correct this issue, please immediately install the provided retrofit kit according to BTE instructions 751W2519 included in the kit. From this point on, the ring on the Chop / Lift Bar shall not be used to connect the Chop / Lift Bar to the snap hook on the 191 Tool or any other device with a snap hook. Only the Eye part of the Jaw and Eye swivel will be used to connect to the snap hook.Please pass this Medical Device Recall notice to all those who need to be aware of it within the customer organization. Once the correction is completed, notify BTE using the provided Field Correction Confirmation card.We apologize for any inconvenience. Please do not hesitate to call BTE Customer Support at 410-850-0333 if you have any questions or concerns. |  | 510 | 510 Devices | Worldwide Distribution - Nationwide Distribution and to the countries of : Austria, Belgium, Canada, France, Germany, Greece, Hong Kong, Italy, India, Japan, South Korea, Mexico, P.R. China, Poland, Republic of South Africa, Romania, Saudi Arabia, Singapore, Spain, Switzerland, Thailand, and United Kingdom. | Open, Classified |  |  | 510 | Not_Computer
Z-1987-2015 | 71208 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Philips GoSafe Communicator 7000L |  |  | April 22, 2015 | 2015 | Lifeline Systems, Incorporated | Home communicator Model 7000L may not sense the phone line under certain unusual landline phone conditions. | Software design | Philips Lifeline notified all consignees of the recall via phone starting on April 22, 2015. The call informs consignees of the issue and actions to be taken. For questions, consignees can contact customer service at 1-800-451-0525 ext. 1608. Consignees were also sent a letter to all consignees instructing them that Lifeline would send out a qualified Installiation Representative for an equipment exchange. |  | 3189 | 3189 devices total, 38 are subject to recall | Nationwide Distribution including AK AZ CA CT FL IA IL KY MA MD MT NC NE NH NY OH OR SC VA WA and WI. | Completed |  |  | 3189 | Not_Computer
Z-1988-2015 | 71428 | No | MRD |  |  | General &amp; Plastic Surgery | 510(k) | Natrelle Re-sterilizable Breast Implant Sizer, Size: 410 cc, Style No. MSZFX410. |  |  | June 08, 2015 | 2015 | Allergan Inc | Allergan is recalling the NATRELLE 410cc Re-Sterilizable Breast Implant Sizer because the expired product was shipped to the user level. | Error in labeling | A recall letter dated 6/8/15 was sent to all customers who received the NATRELLE 410 cc Re-Sterilizable Breast Implant Sizer. The letter provides the problem identified and the action to be taken. Customers are instructed to contact GENCO Pharmaceutical Services if the recall actions on the recall letter are unclear. For product returns, customers are instructed to contact GENCO at (877) 319-8966, 7am to 5pm CST. Customers with credit/reimbursements are instructed to contact Allergan at (800) 811-4148, 7am to 5pm PST. Customers with medical inquiries are instructed to contact (800) 433-8871, option 2, 8am-5pm, PST. Customers with adverse events/product complaints are instructed to call (800) 624-4261, option 3, 8am-5pm CST.On June 12, 2015, Actavis sent follow up letters to customers to include one additional serial number. |  | 9 | 9 units | CA, IN, MO, NJ, VA | Terminated | December 16, 2015 | 191 | 9 | Not_Computer
Z-1990-2015 | 71351 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The PowerLED / X-TEN / VOLISTA Surgical Light Systems are surgical light systems intended to be used to provide visible illumination of the surgical area or the patient. Ceiling surgical light systems typically include one or more arms that are able to rotate 360 degrees around a central axis. |  |  | May 29, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Maquet Suspension Arms used with PowerLED/X-TEN/Volista Surgical Light Systems may grind at the junction of the suspension arms which can cause paint chips or particles to fall into the surgical field, when the suspension arms are mounted above the surgical table. | Device Design | Maquet Medical Systems sent an Urgent Medical Device Field Correction letter dated May 29, 2015, to all affected customers via Fedex. The letter identified the product the problem and the action needed to be taken by the customer.The customer was informed that the correction will include inspection of PowerLED / X-TEN / VOLISTA Surgical Light Systems mounted above the surgical table and replacement of the suspension arms if found to be defective. As a remedial measure, pending replacement of the defective suspensions, Maquet personnel will install a protective plastic strip at each junction where the arms are grinding. The correction will be performed by an authorized Maquet Service Technician. The customer will be contacted by a representative of the Maquet Service team to schedule an on-site inspection and possible service of the PowerLED / X-TEN / VOLISTA Surgical Light Systems. If the customer has any questions or require additional information, they can contact their local Maquet representative, or Maquet Customer Service at 1-888-627-8383 (option 3), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST. |  | 646 | 646 units | Worldwide Distribution - Nationwide Distribution and to the countries of : Argentina, Austria, Australia, Belgium, Bulgaria, Bahrain, Bolivia, Brazil, Belarus, Canada, Switzerland, Chile, China, Colombia, Croatia, Hungary, Indonesia, Ireland, Iraq, Jordan, Japan, Czech Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, France, Greenland, Greece, Hong Kong, United Arab Emirate, United Kingdom, Kenya, Kuwait, Sri Lanka, Luxembourg, Libya, Mongolia, Martinique, Mexico, Malaysia, Namibia, Nigeria, Netherlands, Norway, Oman, Panama, Peru, Philippines, Pakistan, Poland, Saint Pierre and Miquelon, Portugal, Reunion, Romania, Russia, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, Senegal, Thailand, Tunisia, Turkey, Trinidad and Tobago, Taiwan, Uruguay, Venezuela, Serbia, Yemen, Mayotte and South Africa. | Open, Classified |  |  | 646 | Not_Computer
Z-1989-2015 | 71457 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Nakoma-SL Anterior Cervical Plating System's Variable Single Barrel Drill Guide |  |  | June 09, 2015 | 2015 | Alliance Partners LLC | Inadequate interference between the screw and the plate when screw implanted into the plate at angulation greater than 7 degrees can result in the screw advancing through the plate during implantation. | Device Design | Alliance Spine sent an Urgent Medical Device Field Action Notification letter dated June 15, 2015 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the response form and fax to 210-314-2524. |  | 21 | 21 units | US Nationwide Distribution in the states LA, TX, and PR. | Terminated | October 09, 2015 | 122 | 21 | Not_Computer
Z-1996-2015 | 71130 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Endovenous Kit, Endo Pack and Endovenous Laser Kit - contains Devon Light Glove |  |  | April 20, 2015 | 2015 | Stradis Healthcare, LLC. | Convenience kits being recalled due to the recall by the manufacturer (Medtronic) of a component (Devon Light Glove) contained in the kits. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter dated April 16, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine and discontinue use of the affected device. Complete the Recalled Product Return Form and fax to 800-895-6140 or email to feedback.customerservice@Covidien.com. Customers with questions regarding return of the product should call 1-727-642--4124. |  | 15 | 15 | Distribution US nationwide to AL, AK, CA, DE, FL, GA, HI, ID, IL, IN, KS, KY, ME, MD, MA, MI, MS, MO, NM, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA; Canada and Dubai | Open, Classified |  |  | 204 | Not_Computer
Z-2043-2015 | 71404 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano s5/s5x/CORE Mobile Intravascular Ultrasound Imaging and Pressure System; Product/Part numbers: s5 - 804200-002 and 804200-014; |  |  | September 23, 2013 | 2013 | Volcano Corp | A limitation in interoperability between the Philips Xper Flex Cardia Physiomonitoring System and Volcano's sS/Si, Core, and ComboMap systems. This limitation may impact the calculation of Fraction Flow Reserve (FFR) within the Volcano System when the system is connected to the Xper Flex Cardia system via its aortic pressure auxiliary output. | Nonconforming Material/Component | Volcano and Philips sent an Advisory Notice letter dated September 23, 2013 to all impacted customers. The letter identified the affected product, problem and actions to be taken. For question contact your Volcano and Philips representative. |  | 227 | 227 total units affected by the issue - all models | Worldwide Distribution - US Nationwide and the countries of Belgium, United Kingdom, Denmark, Netherlands, Sweden, Iceland, France, Saudi Arabia, Australia, Philippines and Thailand. | Terminated | July 09, 2015 | 654 | 227 | Not_Computer
Z-2047-2015 | 71420 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | PowerPICC SOLO Catheter with Sherlock 3CG TPS Stylet 4F Maximal Barrier Tray with Microintroducer |  |  | May 27, 2015 | 2015 | Bard Access Systems | Bard Access Systems is recalling PowerPICC Solo Catheter Kits due to incorrect expiration dates on the primary labeling. | Labeling Change Control | BARD sent an Urgent Notification letter dated May 27, 2015, to all affected consignees via FedEx. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed that they may choose to return the affected product to Bard Access Systems or use with the understanding that the correct expiration date is November 2015 (2015-11). Customers with questions were instructed to call 1-800-290-1689. For questions regarding this recall call 801-522-5974. |  |  |  | Nationwide Distribution | Terminated | February 10, 2016 | 259 |  | Not_Computer
Z-2052-2015 | 71456 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Block Versions Used by Dental Offices often referred to as chairside milling: |  |  | June 15, 2015 | 2015 | 3M Company - Health Care Business | 3M is initiating a Field Correction to remove the crown indication from all Lava Ultimate products. These products will continue to be indicated for onlays, inlays and veneers. | Device Design | 3M ESPE Dental sent an URGENT: MEDICAL DEVICE CORRECTION letter dated June12, 2015 to affected customers. The letter described the affected product, problem and action to be taken. The letter instructed customers to complete and return the Acknowledgement Return/ Refund Form. For questions call 3M ESPE Dental Customer Care center at 1-800-634-2249. |  | 1800000 | 1.8 million (US 1,017,834) | Worldwide Distribution - US: Nationwide and the countries of: Australia, India, Japan, Korea, New Zealand, Singapore, Thailand, Vietnam, Canada, Bulgaria, Czech Republic, Estonia, Hungary, Israel, Latvia, Lithuania, Poland, Romania, Slovakia, Switzerland, Turkey, China Hong Kong, Taiwan, Brazil, Colombia, Costa Rica, Guatemala, Mexico, Egypt, Morocco, South Africa, Austria, Belgium, Denmark, Finland, France, Germany, Greece, Italy, Netherlands Norway, Portugal, Spain, Sweden, UK and Ireland. | Terminated | April 26, 2016 | 316 | 1800000 | Not_Computer
Z-2051-2015 | 71440 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TITANIUM END CAP WITH T40 STARDRIVE RECESS, 0MM EXTENSION FOR TITANIUM TIBIAL NAILS-EX; Synthes Tibial Nail System EX is intended to stabilize fractures of the proximal and distal tibia and the tibial shaft; open and closed tibial shaft fractures; certain pre-and post-isthmic fractures; and tibial malunions and non-unions. |  |  | June 03, 2015 | 2015 | Synthes, Inc. | A specific part and lot number was packaged and shipped prior to the completion of a required internal inspection. | Employee error | The firm, Depuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" notification letter, dated June 3, 2015, to direct accounts/customers. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory and immediately remove the affected lots from your stock; If you DO have any of the identified devices, please take the following steps:ÂżÂż Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA)Complete and Return the Verification Section with the product to:ÂżÂż Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132.ÂżÂż Send a copy of the completed Verification Section even if you do not have any product to DePuy Synthes by:ÂżÂż Fax: (610) 430-7083 or ÂżÂż Scan/email: Fieldaction@synthes.com Number. If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 3 | 3 | US Distribution to states of: MO and SC. | Terminated | September 15, 2015 | 104 | 3 | Not_Computer
Z-2046-2015 | 71511 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | A.L.P.S. Foot Locking Calcaneus Plate,Small-Right. |  |  | June 04, 2015 | 2015 | Biomet, Inc. | The thread location on one of the thread holes is offset from the axis of the predrill hole. The threads are too deep on one side and too shallow on the other. The locking screw may back out of the plate if it does not achieve proper purchase. This may require a revision surgery to replace or remove the screw and/or plate. The threads, not uniformly loaded, may shear off, causing metal slivers. | Process change control | On June 4, 2015 an URGENT MEDICAL DEVICE REMOVAL letter was sent to all consignees. This action requires the immediate location and discontinued use of the product and its return to Biomet.The following actions are REQUIRED: Immediately locate and remove the identified device(s) listed below from circulation; Carefully follow the instructions on the enclosed "Response Form"; Email a copy of the Response Form to audrey.daenzer@biomet.com prior to return of product; Use priority carrier for your shipment; and If you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear Risk/Recall Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. However, you are charged with the location and return of these products. Please confirm receipt of this notice by sending back the response form within three (3) business days. Thank you in advance for your assistance and prompt attention. On behalf of Biomet, I apologize for any inconvenience this may cause. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 75 | 75 units | Distributed in the states of TX, MD, PA, NJ, NC, MA, AZ, GA, FL, and AR. and the countries of Costa Rica CP and The Netherlands. | Terminated | December 22, 2015 | 201 | 75 | Not_Computer
Z-1950-2015 | 71197 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Shiley Neonatal, Pediatric and Long Pediatric Tracheostomy Tube, Cuffless. Model Numbers: 2.5NEF,3.0NEF,3.5NEF,4.0NEF,4.5NEF,2.5PEF,3.0PEF,3.5PEF,4.0PEF,4.5PEF,5.0PEF,5.5PEF,5.0PELF,5.5PELF,6.0PELF,6.5PELF,2.5NCF,3.0NCF,3.5NCF,4.0NCF,4.5NCF,2.5PCF,3.0PCF,3.5PCF,4.0PCF,4.5PCF,5.0PCF,5.5PCF,5.0PLCF,5.5PLCF,6.0PLCF,6.5PLCF,3.0NEF-P,3.5NEF-P,4.0NEF-P,4.5NEF-P,3.0PEF-P,3.5PEF-P,4.0PEF-P,4.5PEF-P,5.0PEF-P,5.5PEF-P, |  |  | May 08, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Reports where patients who recently switched from the current Shiley Neonatal and Pediatric products to the affected products experienced discomfort and in some cases, breathing difficulties, corresponding to a negative effect on oxygen levels. | Device Design | Consignees were notified 5/8/15 via letter. Consignees were asked to return products. |  | 69461 | 69,461 units total | Worldwide Distribution. US nationwide, Canada, Australia, Belgium, Chile, Germany, Israel, Italy, Japan, Saudi Arabia, Singapore, South Africa, Turkey, and Uruguay. | Terminated | April 12, 2017 | 705 | 69461 | Not_Computer
Z-2056-2015 | 71147 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | IntegraÂż Long Epidural Steroid Injection Needle (Tuohy Needle), 20 gauge x 5 inch. Plastic Hub, Metal Stylet, calibrations (10BX). IntegraÂż Long Epidural Steroid Injection Needle (Catalogue #PITRW2050). Peel Pouch. The IntegraÂż Long Epidural Steroid Injection Needle is packaged in a labeled and sterilized sealed tyvek / plastic pouch. The sterilized pouches are subsequently packaged in a case of 10 that is labeled, sealed and distributed to customers. |  |  | May 01, 2015 | 2015 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences Corp. d.b.a. Integra Pain Management Announces Voluntary Recall of Tuohy Needles (20 Gauge, 5 inch) due to incorrect packaging (labeled as 18 Gauge, 3. 5 inch). | Labeling design | Integra sent an Urgent Voluntary Medical Device Recall letter dated May 1, 2015, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Integra is asking customers to do the following:1. Review your inventory and determine if you have Tuohy Needle, 20 Gauge X 5 inches (Cat. No. PITRW2050) with an affected Lot Number that is listed on page 2. If so, stop using them immediately and remove them from service.2. Complete the attached form. If you do not have affected product, check the box: I do not have any product(s) with a Lot Number from the list below.3. If you do have affected product, check the box:  I do have product(s) with a Lot Number from the list below.4. Complete the other information on the form and return it by email or fax as indicated on the form. Keep a copy of the form for your records.When your form is received, Customer Service will contact you and provide an RMA number and directions to return and replace the affected products. Should you have any questions regarding these instructions, please contact Customer Service at 1-800-241-2210. For questions regarding this recall call 609-936-6822. |  | 123 | 123 Tuohy needles | Nationwide Distribution | Terminated | December 04, 2015 | 217 | 123 | Not_Computer
Z-2055-2015 | 71048 | lens, fundus, hruby, diagnostic | HJI | Diagnostic Hruby fundus lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | Lombart 90D Clear with case |  |  | February 27, 2015 | 2015 | Volk Optical Inc | Mislabeling. The incorrect lot number was engraved on the anodized aluminum ring. The correct lot number is BE03356; however, BD03356 was engraved on the ring in error. | Other | Volk Optical sent a Medical Device Recall letter on March 30, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and quarantine the affected product. If product was further distributed customers need to identify their customers and notify them at once. Customers were instructed to return the affected device for replacement. Customers with questions were instructed to contact Volk at 440-942-6161 or toll free at 800-345-8655. For questions regarding this recall call 440-510-0745. |  | 78 | 78 units | Distribution to VA only | Terminated | February 04, 2016 | 342 | 78 | Not_Computer
Z-2053-2015 | 71474 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Songer Spinal Cable System, Single Cable, Crimp, Bar, Stainless Steel, Sterile. |  |  | May 29, 2015 | 2015 | RTI Surgical, Inc. (dba Pioneer Surgical Technology, Inc.) | The box end label is incorrectly labeled as Titanium instead of Stainless Steel on one lot of implantable Songer Spinal Cable System, Single Cable, Crimp, Bar, Stainless Steel, Sterile, distributed by DePuy Spine Inc. | Labeling Change Control | On 5/29/15, RTI Surgical sent DePuy/Synthes Spine Inc. with an URGENT: Medical Device Voluntary Recall Notification identifying the recalled product, reason for recall, potential risks, and immediate actions required. |  | 25 | 25 | Distributed in MA. | Terminated | March 17, 2016 | 293 | 25 | Not_Computer
Z-2054-2015 | 71315 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics Dimension Vista Systems, used with the Dimension Vista Intelligent Lab System software version 3.6.1. In vitro diagnostic analyzer. |  |  | May 21, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Two software issues may occur in Vista software version 3.6.1. Issue #1: Samples may stop processing without notification. Issue #2: Timing issue in Vista 1500 may cause wrong reagent or no reagent delivery. This may lead to unflagged, unexpected low results and results flagged with assay errors. | Software design (manufacturing process) | Siemens sent an Urgent Medical Device Correction letter dated May 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be Taken by the Customer:For the corresponding issues listed above, please perform the following:Issue 1: If you experience this situation, sample processing can be resumed using the step(s) below:Step 1: From the Operation Icon select PAUSE, then immediately return to the Operation Icon and select Cancel PAUSE. Sample aliquotting should immediately resume.Step 2: If sample processing does not resume after Step 1, wait until the instrument displays System Ready, then restart the system software per the Vista Operator Guide,Section 2, Controlled Shutdown-Restart procedure.Issue 2:To prevent the Reagent Servers from temporarily losing synchronization, Siemens recommends setting the Auto Dispose option to No for methods assigned to Reagent Server 2 (see instructions below). These empty or expired reagent cartridges will now remain on the instrument and require the operator to unload them.Please consult your Vista Operators Guide or method Instructions for Use when results with flags and/or errors occur.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center-Technical Solutions at (800) 441-9250 or your local Siemens technical support representative. |  | 2315 | 2315 | Worldwide Distribution - Nationwide Distribution including Puerto Rico., and the countries of : Austria, Australia, Belgium, Canada, Denmark, France, Finland, Germany, Japan, Italy, New Zealand, Netherlands, Czech Republic, Portugal, Republic Korea, Norway, Spain, Switzerland, Slovakia, Saudi Arabia, Slovenia, and United Kingdom, | Open, Classified |  |  | 2315 | Not_Computer
Z-2057-2015 | 71234 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Hamilton-G5 Ventilators with software versions V2.40/2.41 |  |  | March 24, 2015 | 2015 | Hamilton Medical, Inc. | Customer reports that the ventilator display can freeze. Ventilation continues but the information is no longer displayed on the screen and the user can no longer operate the device. | Device Design | *** Please note *** This is not a new recall. An error occurred, and this recall was originally classified as a class I on July 10, 2015. No further action is required.Hamilton Medical AG notified their customers with a letter dated February 19, 2015 and had a fix (SW 2.42) that was available as of February 20th 2015. Hamilton Medical proceeded to send out notification letters the 24th of March 2015 once they compiled the complete and validated list of customers that had the affected software. Customers with questions were instructed to call 800-426-6331, ext.215. |  | 248 | 248 | Nationwide Distribution | Terminated | December 23, 2015 | 274 | 248 | Not_Computer
Z-2059-2015 | 71528 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Altiva / Exactech Spine, Large Set Screw, Catalog Number 28202, Lot Number AS05107. |  |  | March 06, 2009 | 2009 | Exactech, Inc. | Incorrect thread form on the screw. | Device Design | Consignees were notified with a Product Recall Notice on 03/06/2009 that all screws must be checked to ensure that any discrepant products to be removed from the field. |  | 749 | 749 devices | Distributed in the states of FL, GA, IL, NY, TX, and WA. | Terminated | July 14, 2015 | 2321 | 749 | Not_Computer
Z-2058-2015 | 71423 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | A.L.P.S., 3.5mm x 60 mm Cortical Locking Screw. Bone fixation screw |  |  | June 04, 2015 | 2015 | Biomet, Inc. | Sterile packaging was not sealed. The implantation of an unsterile screw may lead to infection. | Packaging | On 6/4/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors with instructions for removing the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday 8 a.m.to 5 p.m. |  | 50 | 50 units | Domestic: None.International:Netherlands and Japan | Terminated | March 22, 2016 | 292 | 50 | Not_Computer
Z-0811-2016 | 71008 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope Video Laryngoscope Reusable Blades (GVL 3, 4, 5 and AVL 2, 3, 4, 5). |  |  | April 13, 2015 | 2015 | Verathon, Inc. | Verathon has revised the Operations and Maintenance Manual (OMM) to specifically state compatibility limitations of the reusable video laryngoscope Blade validated through simulateduse cycles, such that exceeding this number of cycles in the specific chemical agent may affect the potential life of the product. | Nonconforming Material/Component | Verathon sent an Urgent Medical Device Safety Alert letter dated April 13, 2015 and Safety Alert Reply Form to their consignees on April 17, 2015Users are informed that Verathon implemented a Safety Alert regarding updates to the following Operations and Maintenance Manuals (OMM):" GVL Operations and Maintenance Manual: 0900-1204-08-60" AVL Operations and Maintenance Manual: 0900-4200-02-60For a copy of the most recent GVL or AVL Operations and Maintenance Manual, please visit http://verathon.com/support.If you have any questions regarding the Operations and Maintenance Manuals revisions noted above, please contact your Verathon representative or Verathon Customer Care at 800.331.2313 or 425.867.1348. You may also email customerservice@verathon.com.PLEASE NOTE: NO ADDITIONAL ACTIONS NEED TO BE TAKEN. |  | 9088 | total 9088 units ( 5786 units in the US and 3302 units International) | Worldwide Distribution. | Completed |  |  | 9088 | Not_Computer
Z-2061-2015 | 71606 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Agility. |  |  | June 26, 2015 | 2015 | AGFA Healthcare Corp. | Firm received a complaint that the hospital had sent two accession numbers over the same morning and on the Agility side the patient name was incorrect. | Other | PRB0046372 was created to address the corrective actions related to this issue. On 6/26/2015, an email communication, as an effectiveness check was sent to the two customers. Acknowledgment, via email, that the communication was received was requested from the consignees. |  | 2 | 2 | Distributed in the state of MO. | Terminated | January 27, 2016 | 215 | 2 | Not_Computer
Z-2062-2015 | 71271 | No | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | 0210-312-000-Woundcare Replacement kit with soft tip; 0210-318-000; Woundcare Replacement with fan tip; 0210-318-100- Woundcare Replacement Kit; 0210-318-200-Woundcare Replacement with retract coax tip. Accessory to Stryker InterPulse Irrigation System, used for wound debridement, soft tissue debridement, and cleansing of medical, clinical, or surgical sites. |  |  | May 08, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Potential sterility breach in the packaging. | Package design/selection | Stryker sent Customer Notification Letters on 05/20/2015 through Certified mail via USPS.Customers are asked to do the following: Immediately review this Recall Notification. Immediately check all stock areas and/or operating room storage and quarantine any affected product found. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if you do not have any affected product on hand. If you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF. Fax (866-521-2762) or email (julie.forsyth@stryker.com) the completed Business Reply Form to Stryker Instruments Regulatory Department.Upon receipt of the Business Reply Form, if you have recalled product on hand, Stryker will email a pre-paid shipper. This shipper can be used to return the recalled WoundCare Replacement kits to Stryker. |  | 5328 | 5,328 each (444 boxes) | Nationwide Distribution-including the states of FL, IA, IL, IN, LA, MD, MN, MO, NC, NM, NY, PA, TN, TX, VA, and WI. | Terminated | October 19, 2015 | 164 | 5328 | Not_Computer
Z-2060-2015 | 71530 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | ZNN Antegrade Femoral Nail (ZNN AF). Orthopedic internal fixation device. |  |  | June 17, 2015 | 2015 | Zimmer, Inc. | A single distributed ZNN Greater Trochanter Femoral Nail may have been dented. Potential for fatigue of implant prior to sufficient fracture healing, which may result in need for implant removal and revision of fracture fixation to assure proper healing and avoid fracture malunion or nonunion. | Process control | On 6/17/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital Risk Managers with instructions for returning the affected product. All distributors were notified via electronic mail, and all hospital Risk Managers and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761. Hours of operation are Monday through Friday, 8 a.m. through 5 p.m. EST. |  | 8 | 8 | Worldwide distribution. US states of NC and AK; Taiwan, Japan, and Germany. | Terminated | March 15, 2016 | 272 | 8 | Not_Computer
Z-2065-2015 | 71529 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Oxford Femoral Drill Guide Lateral IM Adapter |  |  | June 03, 2015 | 2015 | Biomet, Inc. | Oxford Femoral Drill Guide Lateral IM Adapter may have an undersized diameter that can cause an interference with PN: 32-422822, Oxford Microplasty IM Link. | Component design/selection | On 6/3/2015 an "URGENT MEDICAL DEVICE RECALL NOTICE" letter was sent to the affected distributors with instructions for locating and returning the affected product. The recall letter included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 18 | 18 | Worldwide Distribution - US Nationwide in the states of NC, IN, PA, OH and the country of the Netherlands. | Terminated | March 15, 2016 | 286 | 18 | Not_Computer
Z-2066-2015 | 71493 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | G7 Variant Elution Buffer HiS no. 1 (M) component of the Tosoh Automated Glycohemoglobin Analyzer HLC 723G7. |  |  | February 10, 2015 | 2015 | Tosoh Bioscience, Inc. | The G7 Variant Elution Buffer HiS no. 1 (M) is incorrectly labeled with an expiration date of October 2015. The correct expiration date is October 2014. | Error in labeling | Tosoh Bioscience sent an Urgent Recall Notification letters dated February 10, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discard any remaining pouches of the product, complete and return the attached response form. For questions call Tosoh Technical Support at 1-800-248-6764. |  | 45 | 9 boxes/5 pouches | Us Nationwide Distribution in the state ID including PR. | Terminated | October 01, 2015 | 233 | 45 | Not_Computer
Z-2064-2015 | 71407 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Apex ARC Hip Stem, HA Coated, Size 5; Product Code: HC-21005-HA |  |  | June 03, 2015 | 2015 | Omnilife Science Inc. | Product may breach the inner and outer sterile pouches during shipping or handling. | Packaging | All product is consigned to US Agents until implanted or returned to OMNI. All US Agents holding the potentially affected inventory were notified that the product is being recalled via letter on June 3, 2015. The letter instructs them to immediately discontinue use of the product and contact OMNI Customer Service at 800-448-6664 to request a return materials authorization, and to return the Acknowledgement and Receipt form by FAX to 508-819-3390 or email to cflores@omnils.com. |  | 146 | 146 devices | US Nationwide Distribution in the states of: UT, CA, FL, TX, MA, GA, OK, VA, CT, OK, and WY. | Terminated | September 08, 2015 | 97 | 146 | Not_Computer
Z-2067-2015 | 71527 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dual Delivery 4DD-1 of GYC-1000 Green Laser Photocoagulator system |  |  | March 25, 2015 | 2015 | Nidek Inc | Accessories to the GYC-1000 laser were missing Laser Aperture labels. | Error in labeling | Nidek sent Recall Notifications to doctors on March 31, 2015, by Federal Express. Letters were specific to the accessory involved.NIDEK detected that the laser aperture label has not been affixed on the delivery unit of GYC-1000, Green Laser Photocoagulator System and the device does not meet FDA regulation 21 CFR Subchapter J.In order to comply with FDA regulation, Nidek requests that the customer affix the enclosed laser aperture label to the Dual Delivery 4DD-1 of GYC-1000, according to the instructions attached.Please see the enclosed Laser Aperture label, instruction for affixing to the device and the response card.Nidek requests that you complete and return the response card as soon as possible. For additional information, please call Nidek at 1-800-223-9044 or visit http://usa.nidek.com |  | 22 | 22 units | US Distribution to the states of : CA, IL, MN, MO, OR, IL, UT RI, MA, WA, NE, NY, NJ, TX and WI. | Terminated | July 14, 2015 | 111 | 69 | Not_Computer
Z-2072-2015 | 71549 | No | HYA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | A.L.P.S. - Medium Bone Plate Forceps. Orthopedic manual surgical instrument. |  |  | June 17, 2015 | 2015 | Biomet, Inc. | Medium Bone Plate Forceps was assembled with the incorrect grip arm tip. Potential that the surgeon will be unable to provisionally clamp the plate to the patient's bone. | Vendor change control | On 6/17/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors with instructions for locating and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 223 | 223 | Domestic: TX NJ, NC, KY, MA, SD, MO, AL, IN, GA, ARInternational: Netherlands, Japan, Costa Rica | Terminated | March 22, 2016 | 279 | 223 | Not_Computer
Z-2070-2015 | 71501 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Exactech Ambassador Cervical Plate System |  |  | May 21, 2015 | 2015 | Exactech, Inc. | Shipped to a single consignee prior to completion of final inspection. | Release of Material/Component prior to receiving test results | Exactech contacted their consignee by telephone on May 21, 2015, of this recall for their two products and instructed him to recover the devices from the hospital. Devices were recovered from the hospital May 22, 2015, and returned to the firm on May 26. All devices have been recovered. |  | 4 | 4 devices. | Nationwide Distribution to Florida only. | Terminated | July 16, 2015 | 56 | 6 | Not_Computer
Z-2063-2015 | 71573 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | July 01, 2015 | 2015 | Elekta, Inc. | A problem exists in MOSAIQ resulting in the incorrect field size being sent to the treatment machine for stereotactic plans using cones. | Software design | Elekta sent an Urgent Important Field Safety Notice 371-01-MSQ-003 on July 1, 2015, to all affected customers. The notice informed users of the specific product and version numbers affected and that the problem only exists if all the following criteria is met - (1) Running Mosaiq 2.41; (2) with at least one linac equipped with an Agility 160 MLC head (Versa HD, Infinity, Axesse (EOS) or Synergy); and (3) who are treating with Stereotactic Cones. The notice also advises users of the workaround to avoid the problem. The customers were also instructed to return the Important Field Safety Notice Acknowledgment form. For questions regarding this recall call 770-670-2422. |  | 74 | 74 Total (45 US) | Worldwide Distribution - US including AZ, CA, GA, ID, IL, MO, NJ, NY, NC, OR, SC, TN, WV, and Internationally to Austria, Brazil, Bulgaria, Canada, Colombia, Germany, Israel, Italy, Japan, Mexico, Russia, Taiwan, and Venezuela. | Open, Classified |  |  | 74 | Not_Computer
Z-2082-2015 | 70783 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Autopen Classic Automatic insulin delivery pen, 3.0ml x 2 unit 2-42 units, AN3800, Lot 7JH, NDC 08470-3800-01, |  |  | December 11, 2014 | 2014 | Owen Mumford USA, Inc. | The dials in the Autopens are not working - the dial does not stay dialed, but springs right back to 0. As a result, the patient could not dial up a dose of medicine to deliver via an injection. | Process control | Owen Mumford, Inc. sent an Urgent Medical Device Recall Notification letter dated December 19, 2014, to all affected customers via Fed Ex. The letter instructed customers to check their shelves for the product and specific lot number. The customer was instructed to contact Customer Service to obtain an RMA and a pre-paid, pre-addressed label for return of the devices. A form was included for the documentation of the inspection of their inventory, whether any are in stock and will be returned, and if they will provide a customer list of how they sold any devices to or that they will notify their customers of the issue with the product. They were instructed to fill it out and fax back to 770-977-2866. Customers with questions were instructed to contact Customer Service Department at 1-800-421-6936. For questions regarding this recall call 770-977-2226, ext 13. |  | 225 | 225 units | Nationwide Distribution including FL, IL, NV, NJ, OH, and TN. | Terminated | July 11, 2017 | 943 | 225 | Not_Computer
Z-2083-2015 | 70684 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | The FLOW-i Anesthesia System Axillary O2 and Suction Module, Device Part number: 66 79 847 |  |  | March 02, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Maquet has identified that in some FLOW-i Anesthesia System Auxiliary O2 and Suction Modules, the suctioning function cannot be activated or controlled by the suction unit on/off switch leading to a decrease or loss of suctioning capability of the modules. | Under Investigation by firm | Maquet sent an Urgent Medical Device Field Action Removal and Replacement Form/Device Removal Response Form dated March 2, 2015 to their affected customers via FedEx. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed read through the letter and inform all members of their staff who may use the affected device. Fill out the attached Device Removal Response Form with requested information, sign and date and fax/email the Response Form using the fax/email information provided. Customers with questions were instructed to contact their Maquet Service Representative or Maquet Technical Support at 1-888-627-8383 (press option 3, followed by option 1 and then optionn 1 again. |  | 104 | 104 units | Worldwide Distribution - US (nationwide) and Internationally to Australia, Belgium, Brazil, China, Colombia, Croatia, Czech Republic, Ecuador, Finland, France, Germany, Hong Kong, India, Israel, Japan, Mexico, Netherlands, Peru, Poland, Portugal, Romania, Russia, Serbia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand and United Arab Emirates. | Open, Classified |  |  | 104 | Not_Computer
Z-2075-2015 | 71559 | No | HXY | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Allofix Insertion Kit, 2.0 mm, Single Drill, |  |  | May 27, 2015 | 2015 | Musculoskeletal Transplant Foundation, Inc. | Musculoskeletal Transplant Foundation is recalling a certain lot of Allofix Insertion Kits because they are labeled as having a 2.0mm diameter drill however it may contain a 2.4 mm diameter drill. | Process control | Musculoskeletal Transplant Foundation notified their affected customers of this recall by sending them a Voluntary Recall Notification Letter dated May 27, 2015 or via telephone. The letter identified the affected product, problem and actions to be taken. For questions contact Customer Service at 1-800-433-6576. |  | 117 | 117 units | US Nationwide Distribution in the states of AZ, CA, CO, FL, IL, MI, MN, NC, NJ, NV, PA, TN, TX, UT, WA and WV. | Open, Classified |  |  | 117 | Not_Computer
Z-2076-2015 | 71593 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Sorin Stockert Heater-Cooler 3T, 230 V |  |  | June 15, 2015 | 2015 | Sorin Group USA, Inc. | Potential colonization of organisms, including Mycobacteria, in Sorin Heater Cooler Devices, if proper disinfection and maintenance is not performed per Instructions for Use. | Error in labeling | Sorin Group issued a Field Safety Notice dated June 15, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and identify any affected devices. For each unit customers were instructed to determine if the device has been maintained according to the Instructions for Use. If yes, customers should strictly adhere to the new Instructions for Use. Customers were also provide with a Response form to confirm they received, read and understood the Field Notice. Customers were instructed to return the completed form to assist in monitoring the effectiveness of the communication. For technical support customers should call 1-800-221-7943, ext 6355. For questions regarding this recall call 303-467-6306. |  | 2837 | 2837 | Worldwide Distribution - US (nationwide) and Internationally to AL, AT, BE, BG, CH, CY, CZ, DE, DK, EE, ES, FI, FR, GB, GR, HR, HU, IE, IT, LT, LU, LV, MC, MK, MT, NL, NO, PL, PT, RO, RS, SE, SK, AE, AR, AU, AZ, BD, BH, BR, BY, CA, CL, CN, CO, CR, DZ, EC, EG, ET, GE, GY, HK, ID, IL, IN, IQ, IR, JO, JP, KR, KW, KZ, LB, LK, LY, MA, MN, MU, MX, MY, NG, NP, NZ, OM, PA, PE, PH, PK, PR, PS, QA, RE, RU, SA, SG, SV, SY, TH, TN, TR, TT, TW, UA, VN, ZA. | Open, Classified |  |  | 5055 | Not_Computer
Z-2073-2015 | 71515 | adaptor, hygiene | ILS | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Etac Supporter toilet armrest |  |  | June 15, 2015 | 2015 | Etac Supply Center Ab | The Supporter toilet armrest has two Fast Nuts fixing the product to the toilet porcelain. The firm has reports of the two Fast Nuts coming loose and causing the Supporter to become unsteady. | Nonconforming Material/Component | Etac sent an Urgent Notice of Field Correction beginning June 15, 2015, to all affected consignees. The letter directed the consignees to do an in field correction with replacement Fast Nuts. Defined replacement instructions were included with the replacement nuts. Customers with questions were instructed to contact their Etac distributor or the Etac coordinator, Snug Seal Inc. at 1-800-336-7683. For questions regarding this recall call 1-800-367-6844. |  | 25 | 25 units | Nationwide Distribution including GA, CT and PA. | Terminated | July 14, 2017 | 760 | 25 | Not_Computer
Z-2074-2015 | 71552 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Abutment Retrieval Instrument Zirconia CC RP/WP, Article No. 37882 and Abutment Retrieval Kit, Article No. 37508 (Article No. 37882 was assembled in this kit). Used to remove a zirconia abutment if the abutment is stuck in the implant connection due to a tight seat. Endosseous dental implant instrument. |  |  | June 22, 2015 | 2015 | Nobel Biocare Usa Llc | One dimension of the affected instrument is incorrect. Therefore this instrument cannot be used to remove the metal adapter as intended. | Under Investigation by firm | A customer notification letter dated 6/22/15 will be sent to all customers who purchased the Abutment Retrieval Instrument Zirconia CC RP/WP because an internal investigation showed that one dimension of the affected instrument is incorrect. Therefore this instrument cannot be used to remove the metal adapter as intended. The letter provides the problems identified and the actions to be taken. Customers are instructed to complete the acknowledgement form and return it to their local Nobel Biocare Subsidiary with the affected instrument. Customers with the affected instrument that is a part of a kit, Nobel Biocare Representative will contact them directly. Customers with any remaining questions, are instructed to contact their Nobel Biocare Representative. |  | 149 | 149 units total | Worldwide Distribution-US (nationwide) and the countries of Australia, Austria, Belgium, Canada, Finland, France, Germany, Greece, Hong Kong, Hungary, Italy, Luxembourg, Netherlands, New Zealand, Norway, Portugal, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | September 18, 2015 | 88 | 149 | Not_Computer
Z-2094-2015 | 69760 | Yes | NOU | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Breas Vivo 50 Home Care Ventilator |  |  | November 07, 2014 | 2014 | Breas Medical AB | Unintended treatment termination could result from a keypad malfunction in some situations. The device erroneously interprets this as a Stop Treatment Instruction. An alarm will not sound, or be registered. Accessories and monitoring equipment connected to the Vivo 50 will stop functioning as the device enters stand-by mode. | Device Design | Breas Medical sent Urgent Release: Field Safety Notice letters dated November 7, 2014 and Field Safety Notice dated November 21, 2014to their customers. The letters identified the affected product, problem and the actions to be taken. |  | 846 | 846 US | Worldwide - United States Nationwide Distribution and the countries: Argentina, Belgium, Chile, Cyprus, Denmark, Ecuador, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Malaysia, Netherlands, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, Thailand, Turkey, UAE and United Kingdom. | Terminated | April 07, 2016 | 517 | 846 | Not_Computer
Z-2084-2015 | 69804 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | NDI Disposable Reflective Marker Spheres for Brainlab IGS (Image Guided Surgical) system. |  |  | June 10, 2015 | 2015 | Northern Digital Inc. | Inadequate curing of adhesive, may cause the two halves of the spheres to separate. | Process control | NDI sent an Urgent Customer Notification letter dated June 10, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the Acknowledgement Form attached to the notification letter and return using the enclosed pre-paid FedEx label. For questions, contact NDI at 1 (877) 634-6340 ext. 544 (if calling from North America) or +1 (800) 634 634 00 ext. 544 (if calling from outside North America) . |  | 6120 | 524, 700 spheres (41773 = 440,820 spheres; 41774 = 77,760 spheres and 41813 = 6,120 spheres) | US Nationwide Distribution | Open, Classified |  |  | 6120 | Not_Computer
Z-2090-2015 | 71502 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Flow Cytometers; used on all BD FACSCount, BD FACSCalibur, BD FACSCanto, BD FACSCanto II, BD LSR II, BD LSRFortessa,BD FACSAria, BD FACSAria II, BD FACSAria Ill, BD lnflux, BD FACSArray,BD Pathway 855, BD Pathway 435 and all SORP instruments. |  |  | June 15, 2015 | 2015 | BD Biosciences, Systems &amp; Reagents | Instruments installed between June 2007 and September 2011 do not contain labeling pursuant to the Food and Drug Administration's (FDA's) Guidance on Laser Products- Conformance with IEC 60825-1 and IEC 60601-2-22 (Laser Notice No. 50). | Labeling Change Control | BD Biosciences sent an Important Product Correction letter dated June 2015 to all affected customers. Customers were informed that a Field Service Engineer will schedule a time to perform the necessary correction. Customers with questions were instructed to contact BD Customer Support at 877-232-8995 (prompt 2). For questions regarding this recall call 408-954-6307. |  | 2188 | 2188 units | Nationwide Distribution | Open, Classified |  |  | 2188 | Not_Computer
Z-2092-2015 | 71366 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Reamer / Irrigator / Aspirator (RIA): RIA Drive Shafts, RIA Tube Assemblies, sterile, RIA Reamer Head, sterile (12.0mm-19.0mm); intended to clear the medullary canal of bone marrow and debris and to effectively size the medullary canal for the acceptance of an intramedullary implant or prosthesis |  |  | May 28, 2015 | 2015 | Synthes, Inc. | It was reported that the RIA Drive Shaft, Tube Assembly, and Reamer Head have the potential to break when incorrectly assembled or used improperly. Precautionary statements are being added to the Reamer / Irrigator / Aspirator (RIA) Surgical Technique Guide . | Labeling False and Misleading | DePuy Synthes sent an Urgent Medical Device Safety Notification dated May 28, 2015, to direct accounts and sales consultants to advise them of the patient risk associated with misassembly or misuse of the device. Customers and sales force were instructed to remove the Surgical Technique Guide J4352 (all versions) dated 11/2011 or earlier and replace it with the updated Technique Guide (DSUS/TRM/0515/0620, Date 06/2015),and to return the completed verification response form.If you have any further questions, please call (610) 719-5450. Thank you for your attention to this matter. |  | 132564 | 132564 | Nationwide Distribution and Internationally to Canada and Switzerland. | Open, Classified |  |  | 132564 | Not_Computer
Z-2091-2015 | 71492 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Helical Blade Inserter for Trochanteric Fixation Nail-ADVANCED (TFNA); intended to secure the head element to the head element inserter to prevent disengagement during head element insertion. Manual orthopedic surgical instrument. |  |  | June 10, 2015 | 2015 | Synthes (USA) Products LLC | The height of one of the three guiding pins of the Helical Blade Inserter for the TFNA is oversized. The Helical Blade Inserter is therefore unable to pass through the Blade/Screw Guide Sleeve (part 03.037.017). | Process control | A recall notification letter, dated June 10, 2015, was sent to sales consultants and Synthes distribution sites requesting return of affected devices. |  | 32 | 32 | Nationwide Distribution. | Open, Classified |  |  | 32 | Not_Computer
Z-2085-2015 | 71555 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Patient Handling System (Motion Control Software), |  |  | April 01, 2015 | 2015 | Viewray Incorporated | ViewRay received a report that the couch moved unexpectedly into the bore after performing a RTCS reboot. | Software Design Change | ViewRay sent a Customer Advisory Notification Letters dated April 1, 2015 to thier customers. The letter identified the affected product, problem and actions to be taken. For questions or concerns regarding this issue, contact ViewRaylM Customer Support at support@viewray.com or call the ViewRaylM Customer Support center at 855-286-8875. |  | 3 | 3 | US Nationwide Distribution in the states of CA, MO &amp; WI | Terminated | February 04, 2016 | 309 | 3 | Not_Computer
Z-2095-2015 | 71585 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | Alere i flu, Alere i, Alere Influenza A &amp; B; Part Number (PN) 425-024 |  |  | June 26, 2015 | 2015 | Alere Scarborough, Inc. dba Binax, Inc. | Risk of false negative results due to microbial growth in the Sample Receiver of one lot of the Alere i Influenza A &amp; B kit | Process change control | Alere sent a Urgent Medical Device Recall letter dated June 26, 2015, to all affected consignees. Customers were instructed to cease use and discard any unused pouched Sample Receivers and Transfer Cartridges contained in Alere" i Influenza A &amp; B PN 425-024, Lot 0073853. Pouched Sample Receivers and Transfer Cartridges are identified in the kit as part number 425-431 lot number 074005 or 074108. Replacement Sample Receivers and Transfer Cartridges will be provided to the customer. Customers with questions were instructed to call Alere Technical Services at 855-731-2288 or email ts.scr@alere.com. |  | 257 | 257 Kits | Nationwide Distribution including AR CA CO CT FL GA IA IL IN KS KY LA MA MD ME MO MS NC NE NJ NY OH PA SC TX VA WA WY. | Terminated | September 15, 2015 | 81 | 257 | Not_Computer
Z-2093-2015 | 71350 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | NDI Disposable Reflective Marker Spheres for Brainlab IGS (Image Guided Surgery) system. |  |  | May 20, 2015 | 2015 | Northern Digital Inc. | Inadequate curing of adhesive, which may cause the two halves of the spheres to separate. | Process control | NDI sent an Urgent Medical Device Recall letter dated May 20, 2015 to their customers via Certified Mail. The letter identified the affected product, problem and actions to be taken. For question from Canada or the United States contact NDI at 1-877-634-6340 ext 544 and for outside North America call 1-800 634 634 00 ext 544. Customer Service can also be reached at spheres@ndigital.com. |  | 32400 | 32,400 spheres | Worldwide - US Nationwide Distribution and the countries of Australia and Canada. | Open, Classified |  |  | 32400 | Not_Computer
Z-2087-2015 | 71628 | No | PGY | Picture archiving and communications system. | Radiology | Radiology | 510(k) | RadiForce GX540, 5MP Monochrome LCD Monitor, Part Code No. 0FTD1826A5, 0FTD1826B5, 0FTD1826C5, UPC 690592034752. |  |  | July 07, 2015 | 2015 | Eizo Corporation | EIZO is recalling the RadiForce LCD Monitors could display abnormally while connecting to particular graphics boards with DisplayPort. | Device Design | EIZO sent An Urgent Field Safety Notice dated July 7, 2015, to all customers who purchased the RadiForce RX440, RX650, RX850 and GX540. The notice informs customers that the RadiForce monitors could display abornormally in 6-bit (64 gradations) instead of normal 8-bit (256 gradations) or 10-bit (1024 gradations) while connecting to particular graphics boards with DisplayPort. The notice provides the customers with the actions to be taken and contact information (1-800-800-5202, M-F, 8am-4pm, PST, us.feedback@eizo.com, 5710 Warland Drive, Cypress, CA 90630). |  | 19 | 19 units | Nationwide Distribution including OH, FL, IL, TX, OH, NY, AZ, and NJ. | Terminated | November 30, 2016 | 512 | 24 | Not_Computer
Z-2086-2015 | 71599 | No | FZO | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 16mm Chisel Blade; used in the LCP Anterior Ankle Arthodesis Plate System and the Hip Preservation Surgery Set. Orthopedic manual surgical instrument. |  |  | June 24, 2015 | 2015 | Synthes (USA) Products LLC | It was discovered that a 10mm Chisel Blade was etched as a 16mm Chisel Blade. | Employee error | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated June 24, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Recipients were asked to check inventory and return affected devices along with the completed response form.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 100 | 100 | US Distribution to the states of : NJ, VA, OH, IN, MI, FL, WI, CA, TN, MD, NC, WA, UT, ID, MO, NV and KS. | Open, Classified |  |  | 100 | Not_Computer
Z-2097-2015 | 71446 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Magnus Hybrid OR Table Column (1180.01B2) used in conjunction with Siemens Artis Zeego Angiography System. Maquet GmbH &amp; Co. |  |  | June 10, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | There is a potential issue that could result to collision between the C-arm of the angiography system and the MAGNUS OR table system. After switching on the angiography-system, in rare cases, MAGNUS column provides wrong position data for the longitudinal movement leading to a collision. | Device Design | Maquet sent a field correction letter dated June 8, 2015, via Fed Ex. Customers are advised of the issue with the device and provided instructions. Questions or additional information can be obtained from the local Maquet representative or Maquet Technical Support at 1-888-627-8383 (option 3 followed by option 1) Monday through Friday between the hours of 8:00 am and 6:00 pm EST. |  | 12 | 12 units | Worldwide Distribution - US Nationwide Distribution and to the countries of : France, Australia, Belgium, Brazil, Switzerland, Chile, China, Germany, Spain, United Kingdom, Hong Kong, Ireland, Israel, India, Italy, Japan, Kuwait, Kazakhstan, Mexico, Netherlands, Norway, Russia, Sweden, Turkey, Taiwan and South Africa. | Terminated | April 18, 2016 | 313 | 12 | Not_Computer
Z-2098-2015 | 71582 | No | HDX | Vaginal stent. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Milex Vaginal-Hymenal Silicone Dilators Set of 4 |  |  | June 19, 2015 | 2015 | CooperSurgical, Inc. | Incorrect expiration date on outer carton kit label | Incorrect or no expiration date | CooperSurgical sent an Urgent Medicall Device Recall dated June 19, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of the affected products and complete the attached Acknowledgement and Receipt Form for a replacement. For questions call 203.601.5200. |  | 40 | 40 kits | US Nationwide Distribution in the states of: CA, FL, GA, HI, IL, KY, MD, MN, NC, NY, OH, TX, VT, and W A. | Terminated | February 01, 2016 | 227 | 40 | Not_Computer
Z-2099-2015 | 71427 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 0702-001-350Q, Diverter Valve Assembly |  |  | May 22, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is initiating a recall of the Neptune 2 Waste Management System due to complaints related to installation failures of the Neptune 2 diverter valve assembly during docking. | Nonconforming Material/Component | Stryker sent an Urgent Medical Device Notification distributed customer notices on 06/09/2015 via certified mail. Customers were asked to review the notice, complete and return the response card, and forward the notice in the event the recalled device was further distributed. Customers were asked to fax the completed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com, and to please continue to rotate all Neptune2 Rovers. If Rovers are removed from circulation for more than 40 days, it is recommended that you contact Stryker Repair if any issue is experienced putting the Rover back into use. |  | 9245 | 9,245 | Worldwide Distribution in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN. KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NH, NY, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and the countries of: Canada, Europe, Asia Pacific, EEMEA. | Terminated | September 12, 2016 | 479 | 9245 | Not_Computer
Z-2096-2015 | 69762 | No | GCB | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Safety PICC Insertion Tray with Chloraprep OneStep - Model #5026CL-9 - 12 pr case. Packaged for Physician use only. |  |  | November 05, 2014 | 2014 | Custom Healthcare Systems, Inc. | An INCORRECT item number is packed within the OUTSIDE "PIGGYBACKED" PACKAGE. The interior sealed STERILE KIT is correct and NO adverse issues affect that part of the item. The NON-STERILE OUTSIDE "PIGGYBACKED" PACKAGE mistakenly contains ITEM #8881579121 - Syringe prefilled, Saline Flush, 09% Sodium Chloride and is a NONSTERILE PRODUCT that CANNOT BE DROPPED INTO THE STERILE FIELD. | Mixed-up of materials/components | Custom Healthcare System sent an Urgent Product Recall letter dated November 5, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.PLEASE CHECK YOUR INVENTORY AND QUARANTINE ANY AFFECTED PRODUCT MATCHING THE DESCRIPTIONS ABOVE. Call 804 (421-5959) with regards to returning the affected products. Return Authorizations will be issued and product pickup and return will be scheduled. |  | 293594 | 3 Cs-Lot#92414 , 25 Cs-Lot#100214, 24 Cs-Lot#100914 | US Distribution to CT. | Terminated | January 12, 2016 | 433 | 293594 | Not_Computer
Z-1838-2015 | 71056 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PROBEAT WITH DISCRETE SPOT SCANNING SYSTEM |  |  | April 21, 2015 | 2015 | Hitachi America, Ltd., Power Systems Division | At a PROBEAT III site installed in Japan, the operator of the system found that the radiation was delivered inconsistent with the treatment plan and stopped the irradiation. The irradiation control system may be potentially affected when rebooting a part of the control system. | Device Design | The recalling firm notified the one consignee about the incident that occurred in Japan on 4/7/15, then gave them an update on 4/13/15 and 4/21/15. The written notification had details for the consignee on how to prevent the device from malfunctioning. The recalling firm also had held conversations with the consignee to discuss the issue and revised work instructions via telephone conference and on site Hitachi personnel. |  | 3 | 3 | Distribution in US (Texas) and Japan | Terminated | March 24, 2016 | 338 | 3 | Not_Computer
Z-2105-2015 | 71468 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | EVac 70 XTRA with Integrated Cable and PROcise EZ View with Integrated Cable |  |  | June 02, 2015 | 2015 | ArthroCare Corporation | Wands inspected with the adaptor fixture were released to the field not meeting the dielectric (Hipot) specifications. | Nonconforming Material/Component | Smith&Nephew; sent an Urgent Notification letter dated June 10, 2015 via mail/courier service with delivery tracking/confirmation to affected customers. A response form is included in the notification letter. A follow up letter will be sent to non-responding consignees. Customers were instructed to complete the acknowledgement form and return via fax or email. |  | 733 | 733 devices | Worldwide Distribution - US Nationwide and the countries of Australia, Belgium, Germany, Spain, United Kingdom, Ireland, India, Italy and Singapore. | Open, Classified |  |  | 733 | Not_Computer
Z-2103-2015 | 71498 | No | LCJ |  |  | General &amp; Plastic Surgery | 510(k) | Natrelle 133 Series Tissue Expanders |  |  | June 18, 2015 | 2015 | Allergan Inc | Allergan is recalling the NATRELLE 133 and CUI Series Tissue Expanders because the product shipped was beyond the 47 month shelf life. | Incorrect or no expiration date | A recall letter will be sent on 6/18/15 to all customers who received the NATRELLE 133 and CUI series tissue expanders. The letter provides the problem identified and the action to be taken. Customers are instructed to contact GENCO Pharmaceutical Services if the recall actions on the recall letter are unclear. For product returns, customers are instructed to contact GENCO at (877) 319-8966, 7am to 5pm CST. Customers with credit/reimbursements are instructed to contact Allergan at (800) 811-4148, 7am to 5pm PST. Customers with medical inquiries are instructed to contact (800) 433-8871, option 2, 8am-5pm, PST. Customers with adverse events/product complaints are instructed to call (800) 624-4261, option 3, 8am-5pm CST. |  | 167 | 167 units total (114 units in US) | US in the state of Canada | Terminated | December 16, 2015 | 181 | 167 | Not_Computer
Z-2102-2015 | 71429 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Stryker SelectCore Variable Throw Biopsy Device, 14G, 10CM |  |  | June 02, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is recalling the Stryker SelectCore Variable Throw Biopsy Device due to the potential for for the stylet to eject from the SelectCore device during actuation. | Employee error | Stryker distributed Customer Notification Letters via Fed Ex overnight to be delivered on 06/10/2015. Customers are asked to Immediately review this Recall Notification. 2. Immediately check all stock areas and/or operating room storage and quarantine any affected product found.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the BRF even if you do not have any affected product on hand.Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in this Notification.4. If you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF.5. Fax (866-521-2762) or email (julie.forsyth@stryker.com) the completed Business Reply Forms; and return the recalled SelectCore devices(s) to Stryker. |  | 55 | 55 units | US Nationwide Distribution in the states of: TX, NJ and MS. | Terminated | September 03, 2015 | 93 | 55 | Not_Computer
Z-2100-2015 | 71430 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ring Transjugular Intrahepatic Access Set |  |  | May 28, 2015 | 2015 | Cook Inc. | Cook Inc. is initiating a voluntary recall of the Ring Transjugular Intrahepatic Access Set and the Flexor RTPS Introduce Sheath due to complaints of the dilators being too long. | Component change control | Written recall communication letters were shipped via FedEx on June 05, 2015 to all US Consignees. Customers were asked to complete and return response forms and recalled products. Recalled devices will be held for final disposition and destruction. Customers wree asked to forward notices if recalled devices were further distributed. Effectivenes checks will be conducted by mail, phone and visits. Customers may contact Cook Medical Relations (800) 457-4500 or 1 (812)- 339-2235 Monday through Friday from 7:30 AM to 5:00 PM or e-mail at CustomerrelationsNA@cook medical.com. |  | 661 | 661 | Worldwide Distribution - US Nationwide in the state of: AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NV, NY, OH, OK, PA, SD, TN, TX, UT, VA, WA, WI, WV and the countries of Canada and EU. | Terminated | May 30, 2017 | 733 | 661 | Not_Computer
Z-2119-2015 | 71113 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hysto Pack, Non-sterile |  |  | April 16, 2015 | 2015 | Covidien LLC | Devon Light Gloves contain splits or holes compromising the sterility | Process control | Medtronic/Covidien initiated recall on April 16, 2015 via FedEx and certified mailCustomers were instructed to discontinue use of the product and return all inventory and kits to Medtronic (Covidien) -Attention: Devon" Light Glove 110 Kendall Park Lane. Atlanta, GA 30336.If you have any questions contact your Medtronic representative or Medtronic Customer Service, Monday through Friday, 8am to 6:30pm ET,at (800) 882-5878. |  | 23560 | 23560 | Worldwide Distribution-US (nationwide) including the Virgin Islands and the countries of Canada, Australia, New Zealand, Costa Rica, Austria, Belgium, Bulgaria, Czech Republic, Denmark, Finland , France, Germany, Netherlands, Sweden, Ireland, ISRAEL,Spain, Italy, Jordan, Kenya, Kuwait, United Kingdom, Lebanon, Liechtenstein, Luxembourg, Monaco, South Africa, Norway, Poland, Portugal, Qatar, Russia, Saudi Arabia, Slovenia, Switzerland, Turkey, and UAE. | Terminated | June 28, 2017 | 804 | 27407865 | Not_Computer
Z-2152-2015 | 70961 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Polaris 4.75 TI Spinal System, 4.0 x 20 MM, Closed MAS, Closed Head Multiaxial Screw, Non-Sterile, Rx Only, REF 14-583020, Biomet Spine and Bone Healing Technologies. Spinal fixation component. |  |  | March 12, 2015 | 2015 | Biomet Spine, LLC | Multiaxial screws may not meet internal requirements related to fatigue strength. Compression testing result and surface treatment are out of specification. | Nonconforming Material/Component | Biomet Spine issued an Urgent Medical Device Recall Notice/Response Form dated 3/12/2015 to their affected customers. |  | 4765 | 4,765 units | Distribution US Nationwide (AL, CA, CO, FL, GA, LA, MA, MD, MI, MO, NY, OH, PA, TX, WA, WI) and The Netherlands. | Terminated | December 21, 2015 | 284 | 4765 | Not_Computer
Z-2069-2015 | 70163 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Various OxyTOTE Portable Oxygen System regulators, Valve Integrated Pressure Regulator, one unit per package. |  |  | January 31, 2015 | 2015 | Western / Scott Fetzer Company | The firm has become aware of the possibility of ignition and rupture (explosion) of a compressed gas oxygen cylinder, which can cause serious injury and/or death. | Component design/selection | Western/Scott Fetzer sent an Urgent Medical Device Recall Notice notification dated January 31, 2015 to their customers via U.S. mail. The notification identified the affected product, problem and actions to be taken. The OxyTOTE/oxyQuik/AirTOTE Distributors and Oxygen Service Providers were instructed to:1) Notify and provide the letter to all accounts/customers to which any OxyTOTE/oxyQuik/AirTOTE products have been distributed. Alternately, please provide a list of customers who received affected product and Western/Scott Fetzer will notify them on your behalf. 2) Referring to the instructions for OxyTOTE/oxyQuik/AirTOTE Users and User Facilities, follow up with each account/customer to ensure that they locate and identify all affected product and establish the number of units in their possession subject to the recall. 3) Contact Western/Scott Fetzer to schedule remediation for any affected OxyTOTE/oxyQuik/AirTOTE product included under this notice. We will begin scheduling product for remediation on February 17, 2015. 4) Please complete the Recall Acknowledgement and Receipt Form and return it to Western/Scott Fetzer as soon as complete responses are available.The OxyTOTE/oxyQuik/AirTOTE Users and User Facilities (Hospitals, Health Care Facilities, Health Care Service providers, etc.) were instructed to: 1) Locate and identify any OxyTOTE/oxyQuik/AirTOTE product in your possession. 2) Review the instructions to determine if the product IS or IS NOT marked and if remediation is needed. 3) Contact your OxyTOTE/oxyQuik/AirTOTE supplier or Medical Gas Provider directly to schedule the remediation of product in your possession.For questions or concerns regarding this notification, please visit the website for additional details or contact Western/Scott Fetzer:Website: www.westernenterprises.com/recallBy Email: recall@WesternEnterprises.comBy Phone: (800) 783-7890, extension 2516By Fax: (440) 835-8283By Mail: Recall Coordinator, Wester |  | 162678 | 162,678 units | Worldwide Distribution -- US, including the states of Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Delaware, Florida, Georgia, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Maine, Michigan, Minnesota, Missouri, Mississippi, Montana, North Carolina, North Dakota, Nebraska, New Hampshire, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, Wisconsin, West Virginia, and Wyoming; and, the countries of CANADA, CHILE, GERMANY, ISRAEL, LEBANON, MACAU, MEXICO, PUERTO RICO, and SINGAPORE. | Open, Classified |  |  | 162678 | Not_Computer
Z-2155-2015 | 71509 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Cochlear Nucleus Sterile Silicone Template |  |  | June 05, 2015 | 2015 | Cochlear Americas Inc. | Cochlear Americas is recalling Nucleus Sterile Silicone Template CI24RE/CI422 part number Z421736 because expired product was distributed. | Process control | Cochlear sent an Urgent Medical Device Recall letter dated June 5, 2015 via Fed-Ex shipment. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return affected products for replacement. |  | 26 | 26 | US Nationwide Distribution to CO, OH, OR, MI, and CA. | Terminated | August 05, 2015 | 61 | 26 | Not_Computer
Z-2123-2015 | 71574 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Precision MPi fluoroscopic imaging systems. |  |  | June 12, 2015 | 2015 | GE Healthcare | A reported incident of a monitor that fell from the suspension. A fall of a monitor could result in an injury to a patient or operator. Recall affects the single and dual LCD monitor suspensions used on Precision 500D Classical R &amp; F system, Precision RXi 23A/32A Analog System, Precision RXi Digital system and Precision MPi fluoroscopic imaging systems. | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter dated June 12, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Perform a visual check before and after each use for the presence of the four (4) mounting screws as shown in the pictures below. Check the secureness of these screws by hand. If these screws are present and secure, you may continue use of the system.If any of these screws are missing or can easily be turned by hand, follow the precautions below before you continue use with the monitor(s).1. Position the monitor suspension in the most frequently used position and limit further movement as much as possible.2. Clearly inform (e.g. through signage and verbal instructions) the users and other auxiliary personnel who may come in contact with the system to not move the monitor suspension once in position.3. Contact your GE Healthcare representative to inform them of any loose or missing screws.For questions or concerns regarding this notification, please contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 788 | 788 (701 US; 87 OUS) | Worldwide Distribution - US Nationwide Distribution including GU and DC., and to the countries of : Australia, Bulgaria, Canada, Chile, Dominican Republic, Germany, Ireland, Italy, Kuwait, Lebanon, Portugal, Romania, Russia, Saudi Arabia, Slovakia, Spain, Taiwan,Turkey, United Kingdom, Venezuela and Yemen. | Terminated | August 03, 2017 | 783 | 788 | Not_Computer
Z-2153-2015 | 71411 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Bausch &amp; Lomb SofPort Advanced Optics Aspheric Lens |  |  | May 26, 2015 | 2015 | Bausch &amp; Lomb Surgical, Inc. | Haptics breaking during lens loading and insertion. | Device Design | Bausch + Lomb sent an Urgent Medical Device Recall letter dated May 29, 2015 to their custometers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to: segregate the identified IOLs, return the IOLs through your B&L; representative or by using the prepaid shipping label included in this package, order replacement product through customer service (800) 338-2020 or your B&L; representative and complete &amp; return the acknowledgement form included in this package. Customers were advised that if replacement product is needed for a scheduled surgery before your full return can be processed, please contact your B&L; representative to expedite shipment of the IOL. For questions contact Bausch + Lomb at (800) 338-2020. |  | 72773 | 72,773 units. 52,787 units (US) - 19,986 units (Foreign) | Worldwide Distribution - US Nationwide in the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MD, MN, MO, MS, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, Hawaii and Puerto Rico and countries of: Canada, Brazil, Argentina, Austria, Belgium, Switzerland, Germany, Spain, France, Great Britain, Georgia, Greece, Kyrgyzstan, Norway, Russia, Ukraine, United Arab Emirates, China, Australia, India Malaysia, Korea, and Taiwan. | Open, Classified |  |  | 72773 | Not_Computer
Z-2137-2015 | 71421 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ATTUNE FB PS ARTICULATION SURFACE SIZES 3-8 |  |  | June 12, 2015 | 2015 | DePuy Orthopaedics, Inc. | Potential for the Balseal, a small wire spring coil located on the post feature of the ATTUNE Knee Tibial Articulation Surface Instruments used in trialing to become damaged and disassociate (come off). If the Balseal is separated from the post, it has the potential to enter the surgical site and be left in the patient if the surgeon is unaware of the disassociation. | Device Design | URGENT INFORMATION: MEDICAL DEVICE CORRECTION was issued for all lots of the affected devices. The device correction is to inform US Distributors and Medical Professionals about the issue, how to prevent the issue, and any actions to take. US Distributors will provide Hospitals/User Facilities with the Medical Professional Notice and Acknowledgment Cards. US Distributors will also work with the sales consultants to ensure all consultants complete sales training regarding this issue. |  | 34676 | 34676 units | Worldwide Distribution - US including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and Internationally to CANADA , AUSTRALIA, AUSTRIA, BELGIUM, ENGLAND/ UK, FRANCE, GERMANY, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JAPAN, MALAYSIA, NETHERLANDS, NEW ZEALAND, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SOUTH KOREA, SPAIN &amp; IBERIA, SWEDEN, SWITZERLAND, THAILAND, UAE, CHILE, COLOMBIA, MEXICO, PERU, and PUERTO RICO. | Terminated | November 12, 2015 | 153 | 280591 | Not_Computer
Z-2125-2015 | 71380 | No | JJJ | Beta or gamma counter for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | WIZARD2 2-detector, 550 samples |  |  | June 03, 2015 | 2015 | Perkinelmer | Error in WIZARD2 Barcode ID Label #023 Contentmay produce erroneous results | Nonconforming Material/Component | Perkin Elmer issued the recall letter on June 3, 2015. Action to be taken: Until the PerkinElmer correction has been completed, we request you immediately destroy all barcode ID labels #023 in your possession. The WIZARD Gamma Counter barcode label binder contains a total of 20 barcode labels identified as #023 (i.e. 2 labels on each of 10 pages). Please ensure you remove thebarcode labels ID #023 from all barcode label binders in your possession. Please be sure to inspect all sample racks to remove and destroy any barcode ID labels #023 already affixed to racks. A WIZARD2 Gamma Counter barcode label binder with corrected barcode labels ID #023 will be provided upon request and return of the completed Correction Response Form to PerkinElmer. Correction Response Form and return it by scan and e-mail to QMresponse.SG@perkinelmer.com |  | 53 | 53 units | Worldwide Distribution-US (nationwide) and the countries of ARGENTINA AUSTRALIA AUSTRIA BELGIUM BRAZIL CANADA CHILE CHINA COSTA RICA CROATIA CZECH REPUBLIC DENMARK ETHIOPIA FINLAND FRANCE GERMANY GREECE HONG KONG HUNGARY INDIA IRELAND ISRAEL ITALY JAPAN KOREA KUWAIT LUXEMBOURG MEXICO MYANMAR NETHERLANDS NORWAY PHILIPPINES POLAND REUNION SAUDIA ARABIA SINGAPORE SOUTH AFRICA SPAIN SWEDEN SWITZERLAND TAIWAN THAILAND TURKEY UK. | Terminated | August 04, 2017 | 793 | 422 | Not_Computer
Z-2149-2015 | 71569 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | THOMAS by Gardner Denver, MEDI-PUMP(R) Aspirator, Model 1240 |  |  | June 04, 2015 | 2015 | Gardner Denver Thomas Inc | Vacuum pump may not perform to specification. | Nonconforming Material/Component | Gardner Denver Thomas sent an Important Medical Device Recall letter dated May 7, 2015 to affected distributors. The letter identified the affected products, problem and actions to be taken. For questions call your Customer Service Representative or Sales Representative or the Recall Coordinator at 920-451-6423. |  | 160 | 160 units | Internationally Distributed in the countries of :Mexico, Peru, El Salvador, Chile, Colombia, Vietnam, and Malaysia. | Open, Classified |  |  | 1038 | Not_Computer
Z-2147-2015 | 71597 | No | MPO | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(k) | Infant Heel Warmer and Infant Gel Warmer, sold under the brand names CooperSurgical and Fisher brand as follows : |  |  | July 06, 2015 | 2015 | CooperSurgical, Inc. | Product marketed without a 510(k) | No Marketing Application | CooperSurgical sent an " Urgent WarmGel Infant Heel Warmer" letter dated June 30, 2015, FEDEx, confirmed delivery receipt on July 6, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.CooperSurgical is advising all WarmGel users to review your inventory of this product, remove it from your inventory and discontinue its use. informing them of the reason to cease using the product and either discard or return product to CooperSurgical. Accounts are provided with a response form to complete. Any questions contact CooperSurgical at 203 601 5200 Ext. 310 |  | 6344700 | 6,344,700 | Worldwide Distribution - Nationwide Distribution and to the countries of : AUSTRALIA, BERMUDA, ECUADOR , FRANCE, NEW ZEALAND, SAUDI ARABIA and UNITED KINGDOM. | Open, Classified |  |  | 6344700 | Not_Computer
Z-2133-2015 | 71654 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano s5/s5i/CORE Intravascular Ultrasound Systems software versions 3.3 and 3.4 (with and without iFR Scout Technology) and v.3.4 software kits; |  |  | June 22, 2015 | 2015 | Volcano Corporation | A software issue was discovered where an inaccurate FFR/iFR value could be calculated under certain circumstances. | Software design | Urgent Medical Device Correction letters were sent on June 22, 2015 to all affected sites via FED Ex. |  | 4007 | A total of 4007 devices, all models | Worldwide Distribution-US (nationwide) and regions in Europe and the countries of Japan, Canada, South Africa, Taiwan, Thailand, Singapore, Philippines, Australia, New Zealand, El Salvador, Malaysia, Mexico, Lebanon, Kuwait, South Korea, Kenya, Jordan, Iran, India, Israel, Hong Kong, Egypt, Algeria, Dominican Republic, Costa Rica, Chile, Brazil, Bolivia, and Argentina | Terminated | April 14, 2017 | 662 | 4007 | Not_Computer
Z-2148-2015 | 71635 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | The EKOS EkoSonic Control Unit is intended exclusively for use with the EkoSonic Endovascular Device and the EkoSonic MACH4 Endovascular Device. |  |  | June 24, 2015 | 2015 | EKOS Corporation | The Connector Interface Cable (CIC) was not recognized by the EkoSonic PT-3B control unit. | Vendor change control | Ekos sent an "Urgent Voluntary Medical Device and Replacement" letter dated June 24, 2015, to US customers. The letter identified the product the problem and the action needed to be taken by the customer. The replacement CIC and recall letter were sent by FEDERAL EXPRESS directly to the person responsible for maintaining the EkoSonic System. Tracking numbers and delivery notifications will be retained for each shipment. Letters to Risk Management and Purchasing were sent by REGISTER MAIL. The return cards will be retained for each letter.Customers with questions can contact EKOS Customer Service at 1-888-400-3567. |  | 46 | 46 units | Nationwide Distribution | Terminated | August 25, 2015 | 62 | 46 | Not_Computer
Z-2170-2015 | 71534 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Fiber Cleaver; |  |  | May 28, 2015 | 2015 | American Medical Systems Innovation Center - Silicon Valley | validation data related to cleaning instructions and sterilization methods and the methods defined in instructions for use (IFU) for the Sureflex Reusable Fibers and the Reusable Stripper and Cleaver Accessories do not meet the new guidance on qualifying reusable products | Reprocessing Controls | Urgent Medical Device Field notification letters were sent on June 9, 2015. Translation for countries out of the US will be conducted. Communications are being coordinated by the Prostate Health QA Director.Urgent Medical Device Field notification Letters advised customers of the issue and list all affected products with name and model numbers. Any SureFlex Reuseable Fibers which have not been used but are in inventory and with the 3 year shelf life should be returned as per instructions enclosed.Any of the SureFlex devices in inventory that have been used should be discarded or returned.All Reuseable Stripper and Cleaver Accessories that have not been used should be returned. Used devices should be destroyed or returned. Questions to 1-408-456-3320. The Acknowledgement form should be completed and returned. |  | 16710 | 16,710 total all affected devices | Worldwide Distribution - US Nationwide in the countries of: Argentina France Morocco, South Africa,Australia, Germany, Netherlands, Spain,Austria, Greece, New Zealand, Sri Lanka,Belgium, Guatemala, Norway, Switzerland,Bolivia, India, Panama, Thailand,Brazil, Ireland, Peru, Turkey,Canada, Israel, Poland, United Kingdom,Chile, Italy, Portugal, United States,Colombia, Korea, Qatar, Uruguay,Costa Rica, Lebanon, Romania,Finland, Malaysia and Saudi Arabia. | Terminated | January 19, 2016 | 236 | 16710 | Not_Computer
Z-2189-2015 | 71533 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | T-Lok Bone Marrow Biopsy Needle 8G x 4 Ext Can, Part Number DBMNJ0804TL; T-Lok Bone Marrow Biopsy Needle 11G x 4 Ext Can, Part Number DBMNJ1104TL; Bone Marrow Biopsy Needle 13G x 4, Part Number DBMNJ1304 |  |  | June 18, 2015 | 2015 | Argon Medical Devices, Inc | Potential defect in the forming of the primary package of the blister that may compromise the sterility of the product inside the tray. | Packaging process control | The recalling firm's sales force was notified of the recall on 6/17/15. The recalling firm will begin to send out Urgent Medical Device Recall letters by registered mail to affected customers on 6/18/15. Customers are instructed to inspect and quarantine and return any affected product and complete the Recal Response Form. |  | 10290 | 10290 units | US | Open, Classified |  |  | 16680 | Not_Computer
Z-2190-2015 | 71525 | No | DWO | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Mermaid Medical M-Biopsy Coaxial Introducer Needle, Coaxial Needle for M-Biopsy 18Gx10cm-PRO-MAG, RX Only. Single packed device in a tyvek pouches. 10 single packed devices in a |  |  | June 15, 2015 | 2015 | Mermaid Medical A/S | Mermaid Medical A/S is recalling some of their M-Biopsy Coaxial Introducer Needles because of the potential for small plastic pieces to be present in the luer of the cannula. | Process control | The firm, Mermaid Medical, notified both customers directly via email with written letters dated June 12 and June 15, 2015 followed up by sales and customer service for product return to MMI. The customers were instructed to "Please coordinate the return of the mentioned products with corresponding lot numbers with our Mermaid Medical's Customer Care the recalled product: either by phone at 1{720}420-9999 or by email: customercare@mermaidmedical.com".If you have any questions please do not hesitate to contact Vice President QA &amp; Regulatory Affairs toll free at (855) 427-6799 or (720) 420-9999 or email: customercare@memaidmedical.com. |  | 320 | 320 | US Distribution to FL including PR. | Terminated | December 07, 2015 | 175 | 320 | Not_Computer
Z-2197-2015 | 71699 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF Ready CT/PET System. |  |  | July 01, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Philips has identified four (4) software defects in the Tumor LOC software application that may potentially result in irradiation of healthy tissue or non-irradiation of diseased tissue. | Software design | On 7/6/2015 the firm sent Field Safety Notices to their customers. |  | 1 | 1 unit | Worldwide Distribution - US Nationwide, Australia, Belgium, Brazil, Canada, China, Denmark, France, Germany, Israel, Italy, Japan, Lebanon, Netherlands, Portugal, Saudi Arabia, South Africa, Switzerland, Thailand, Venezuela &amp; Vietnam. | Open, Classified |  |  | 86 | Not_Computer
Z-2200-2015 | 71689 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Pinnacle3 Software Version 9.0, 9.2, 9.4 and 9.6, Model Numbers 453560446041, 459800091001, 459800220161, 459800232931, 459800235871, 459800338451. |  |  | July 31, 2014 | 2014 | Philips Electronics North America Corporation | Philips, Pinnacle Radiation Treatment Planning System version 8.0h, 8.0k, 8.0m, 8.0n, 9 0, 9 2 9 4, 9 6 is being recalled because the dose may be inconsistent with the density of a density-overridden ROl. | Software design | Consignees were sent on 7/31/2014 a Philips "Urgent Field Safety Notice" dated 2014-July 20. The letter described the problem, product involved in the recall and the action that should be taken by the customers. For questions contact Customer Care Solutions Center at 1-800-722-9377.UPDATE: A second Philips letter dated 2 July 2015 was sent in August 6, 2015. This letter updated the recall to incorporate the additional products being recalled. |  | 1642 | 1642 | Worldwide Distribution-US: All states in the US including DC and PR. OUS: Australia, Austria, Bangladesh, Belgium, Brazil, China, Denmark, France, Germany, Hong Kong, India, Indonesia, Ireland, Italy, Japan, Malaysia, Mexico, Netherlands,New Zealand, Oman, Philippines, Singapore, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, United Kingdom, Viet Nam.UPDATED: Canada, Bulgaria, Egypt, Hungary, Jordan, Luxembourg, Poland, Reunion, Romania, Russia, Sauth Arabia, Turkey. | Open, Classified |  |  | 1642 | Not_Computer
Z-2202-2015 | 71473 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT c8000 System List Number 01G06-98 |  |  | June 11, 2015 | 2015 | Abbott Laboratories | The ARCHITECT c8000 instrument contains tubing that does not meet specifications and has the potential to leak. | Mixed-up of materials/components | Abbott notified thier affected consignees via phone or visit on 5/29/15. A follow up Product Recall notification letter dated June 4, 2015 was sent to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to immediately discontinue use of the product, Abbott representative will be make all necessary arrangements for replacing your impacted instruments(s), follow your laboratory protocol regarding the need to review previously reported patient results and retain this notification for laboratory records. For questions contact Customer Service at 1-877-4ABBOTT. |  | 5 | 5 units | US Nationwide Distribution in the states of CA, AL, NY and WI. | Terminated | February 09, 2016 | 243 | 5 | Not_Computer
Z-2201-2015 | 71650 | No | GXH | Cryogenic surgical device. | Neurology | Neurology | 510(k) | iovera 155 Smart Tip; |  |  | July 01, 2015 | 2015 | Myoscience Inc | The expiration date on the outer box label and the pouch label for the iovera Smart Tip product is incorrect. Instead of indicating the correct expiration or 2015-12. the labels were incorrectly labeled 2016-05. | Incorrect or no expiration date | Myoscience sent an Urgent Medical Device Correction letters dated July 1, 2015 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are requested to quarantine any affected product at their site. Myoscience will make arrangements to either correct the incorrect expiation date label or remove the product and replace it with correctly labeled product immediately.Question or concerns may be directed to Customers Service at 510-933-1500. |  | 36 | 12 boxes - 3 boxes of 10 count and 9 boxes of 5 count. | US Nationwide Distribution in the states of TX, NY, MD, CO and CA. | Terminated | September 08, 2015 | 69 | 36 | Not_Computer
Z-2205-2015 | 71227 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Head Holder shipped with Ingenuity CT Computed Tomography X-ray Systems |  |  | March 17, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | A head holder released for use with other CT systems but not yet qualified for use on the Ingenuity CT, which was shipped with this system. | Mixed-up of materials/components | On 6/4/2015, the firm mailed Customer Information Letters. |  | 66 | 66 units | Worldwide Distribution-Product was shipped to the following states: AR, AR, AZ, CA, CO, CT, FL, GA, IA, IN, KS, LA, MA, MD, MI, MO, MS, NC, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, VA, VT, WA, WV &amp; WY, and to the following countries: Argentina, Australia, Bangladesh, Belgium, Brazil, Bulgaria, China, Colombia, Denmark, Ecuador, France, Germany, India, Indonesia, Israel, Japan, Latvia, Libya, Myanmar, Norway, Oman, Poland, Republic of Korea, Russian Federation, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, United Kingdom &amp; Viet Nam. | Open, Classified |  |  | 200 | Not_Computer
Z-2191-2015 | 71263 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | American Catheter / Cholangiogram Catheter 9100 Series. |  |  | March 17, 2015 | 2015 | American Catheter Corp | Catheters were not sealed before sterilization. and was not detected through subsequent acceptance activities, product release and distribution. | Process control | Consignees were contacted by telephone of the affected lot number and sent a Recall Script dated 3/17/2015 giving instructions how to respond to this recall. |  | 890 | 890 devices. | Distributed in the states of AL, AK, CA, CO, FL, GA, HI, IL, KS, LA, MI, MO, MT, NJ, OH, OK, SC, TN, TX, WA, and WI. | Completed |  |  | 890 | Not_Computer
Z-2192-2015 | 71577 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | HardyCHROM VRE Agar-RSR |  |  | June 24, 2015 | 2015 | Hardy Diagnostics | Hardy Diagnostics is recalling HardyCHROM VRE Agar due to lack of 510(k) clearance. | No Marketing Application | The firm initially called customers to inform them of the recall. The firm's phone script stated that they were discontinuing the product and as part of the process were recalling anything that still might be within expiration date. The phone script stated that the firm was going to be sending something in writing.On 06/24/15 the firm sent out a notification letters to customers to inform them that they were discontinuing the product, and were requested anything remaining within expiration to be discarded. The firm requests that customers complete, sign, and fax back the enclosed form stating their compliance with the above action. For any questions the firm states to contact Technical Services Department at 800-266-222, option 2 or via email at TechService@HardyDiagnostics.com. |  | 1500 | 1,500 | Distributed to one U.S. customer in SC. | Terminated | August 28, 2015 | 65 | 1500 | Not_Computer
Z-2208-2015 | 71615 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Silhouette Lift brochure |  |  | July 02, 2015 | 2015 | Silhouette Lift, Inc. | Silhouette Lift is recalling the Silhouette Suture brochure because it does not include a complete list of risks and side effects. Therefore, it can potentially mislead the patient regarding the safety of the product. | Labeling False and Misleading | Silhouette lift sent an "Correction of Information for Silhouette Suture" letter dated July 2, 2015, in response to FDA untitled letter. The letter identified the product the problem and the action needed to be taken by the customer.Action Requested:All brochures for Silhouette Suture should be destroyed and no longer be used for discussions with patients.Customers with questions are instructed to call (951) 734-7701. |  | 2260 | 2,260 units | US Distribution to the states of : CA, TX, IA, MA, RI, GA, WI, PA, VA, MO, FL, NE, MI, NY and CO., and Internationally to Canada. | Terminated | February 05, 2016 | 218 | 2260 | Not_Computer
Z-2207-2015 | 71337 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Tni-Ultra Assay; SMN 10317708 - 100 tests and SMN 10317709 -500 tests |  |  | May 19, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc | System-to-system bias between the Tni-Ultra assay onthe ADVIA Centaur CP and ADVIA Centaur/ADVIA Centaur XP/ADVIA Centaur XPT systems. | Under Investigation by firm | Siemens sent an Urgent Medical Device Correction letter dated via Federal Express to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Field Correction Effectiveness Check attached to the letter.For questions contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 21159 | SMN10317708 = 20544 kits &amp; SMN 10317709 = 21159 kits | Worldwide distribution - US Nationwide in the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, MN, MS, MT, NC, NJ, NM, NY, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WY, WV, and Puerto Rico; and the countries of Algeria, Austria, Bahrain, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Hungary, Iceland, Indonesia, Ireland, Italy, Jordan, Latvia, Lebanon, Lithuania, Luxembourg, Macedonia, Malta, Morocco, Netherlands, Norway, Pakistan, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Slovakia, Slovenia, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, AR, AU, BD, BR, CA, CL, CN , CO, CR, DO, EC, HK, ID, IE, IL, IN, JP, KR, LK, MX, MY, NI, NP, NZ, PE, PH, PY, SA, SG, TH, TW, UY, VN, and ZA. | Open, Classified |  |  | 21159 | Not_Computer
Z-2254-2015 | 71124 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-ray Systems Brilliance CT 64-channel with Essence technology. |  |  | February 05, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | 5 issues. Fast Get Ready incorrectly enabled during service procedures. Contrast annotation missing for some images when manual contrast was administered. Dot artifact present, intermittently, after startup. Heart Rate measured value incorrectly remains constant for some heart rate change conditions. DoseRight incorrectly enabled when the Exam Card specified disabling DoseRight. Tube heat predict | Software change control | On 2/5/2015 the firm sent Customer Information Letters to their customers. |  | 4 | 4 Units | Worldwide Distribution-US (nationwide) including the states of AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV &amp; WY, and to the following countries: Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 424 | Not_Computer
Z-2260-2015 | 70897 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MiniMed NGP 640G 1.8ml (mmol/L), Model No. MMT-1511; NGP 640G 1.8ml (mg/dL), Model No. 1512; NGP 640G PLGM 3ml (mmol/L), Model No. MMT-1711; NGP 640G PLGM 3ml (mg/dL), Model No. MMT-1712. |  |  | June 19, 2015 | 2015 | Medtronic MiniMed Inc. | Medtronic MiniMed is recalling the MiniMed 620G and 640G insulin pumps because there are certain scenarios where the set Bolus screen will not timeout, which could cause confusion by showing a bolus amount that is no longer appropriate. | Under Investigation by firm | A customern notification letter sent to customers on 6/19/15 who received the MiniMed 640G insulin pumps. The purpose of the letter is to inform customers that when using Bolus Wizard with the MiniMed 640G insulin pumps, that there are certain scenarios where the set Bolus screen will not timeout, which could cause confusion by showing a bolus amount that is no longer appropriate. Customers are informed that Medtronic MiniMed has updated the user guide which instructs to respond promptly to set Bolus screens, and remind not to activate delivery on a bolus based on a blood glucose value that is more than 12 minutes old. The letter informs the customers of the summary of changes in the MiniMed 640G insulin pump user guides. |  | 1936 | 1936 units | Australia, Czech Republic, Denmark, Finland, France, Ireland, Netherlands, Norway, Sweden, United Kingdom. | Terminated | October 02, 2015 | 105 | 1936 | Not_Computer
Z-2261-2015 | 71715 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ottobock Pheon 3R62 Knee Joint |  |  | July 16, 2015 | 2015 | Otto Bock Healthcare GmbH | Otto Bock Healthcare GmbH is voluntarily recalling all 3R62 Pheon polycentric prosthetic knee joints distributed since February 2014 because the locking function may fail under certain conditions which may result in a fall. | Component design/selection | The firm, Ottobock, sent an "URGENT MEDICAL DEVICE RECALL 3R62 Pheon" letter dated 16 July 2015 to its customers. The letter described the product, problem, and actions to be taken. The customers instructed to read this notice in its entirety; examine your stock for 3R62 Pheon devices; quarantine 3R62 Pheon devices and prevent further distribution; contact Ottobock customer service at (800) 328-4058-option 1 or (512) 806-2000 -Option 1, to obtain a return authorization for those devices, and complete and return the MEDICAL DEVICE RECALL RETURN RESPONSE Acknowledgement and Receipt Form via email to: USRecalls@ottobock.com "Pheon Recall"; fax (801) 974-6718-"Pheon Recall" or mail to: Otto Bock HealthCare LP, ATTN: PHEON RECALL, 3820 W. Great Lakes Drive, Salt Lake City, UT 84120-B7 is re-designing the locking function to prevent this failure from occurring. The re-designed devices will be available for replacement sometime within the next few months. If you have any questions contact Director, Regulatory &amp; Quality Compliance at (801) 974-6618 or email: Lex.Pearce@ottobock.com. |  | 518 | 518 (439 US; 79 OUS) | Worldwide Distribution: US (nationwide) and countries of: AU, BE, CN, CZ, IT, NL, and SK. | Terminated | June 21, 2017 | 706 | 518 | Not_Computer
Z-2209-2015 | 71459 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | InjectionPlasty 1.0 Bone Void Filler; an injectable and biocompatible calcium phosphate bone void filler. Sterilized using irradiation. The single-use InjectionPlasty 1.0 Kit contains the necessary components for mixing and delivery of the bone void filler. The InjectionPlasty 1.0 sterile kit contains: Calcium Phosphate Powder, Dilute Sodium Silicate Liquid, a Mixing System, Targeting K-Wires and a Delivery Needle (Cannula). |  |  | May 29, 2015 | 2015 | Skeletal Kinetics, Llc | Reports of venous uptake, DVT, pulmonary embolism in patients who were implanted with Injection Plasty Bone Void Filler. | No Marketing Application | Skeletal Kinetics sent letters to distributors by overnight mail with instructions to forward a Dear Doctor notification letter on May 14, 2015. Second letters dated May 29, 2015, advised that the firm is removing the product from the marker. Distributor letters asked that the Dear Doctor letter be forward within 24 hours of receipt of the letter. All Injection Plasty kits in their possession and those with customers should be retrieved and returned by Fed Ex. Dear Doctor letters advised that their distribution agent has been requested to return all unused kits. Questions may be directed to 408-366-5002 of duran@skeletalkinetics.com. |  | 67 | 67 kits | US Distribution including Puerto Rico. | Terminated | July 24, 2015 | 56 | 67 | Not_Computer
Z-2253-2015 | 71524 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | OxySure Model 615 disposable replacement cartridges, for portable oxygen generator. |  |  | June 12, 2015 | 2015 | Oxysure Systems, Inc. | Flow rate of the recalled replacement cartridges does not meet the draft FDA standard of a minimum 6 LPM for a minimum of 15 minutes. | Under Investigation by firm | The recalling firm will notify their affected customers via email beginning the week of 6/15/15. The recall notification will have a response for to be completed by the consignee. The recall notification calls for the customers to dispose of the affected product according to the MSDS. |  | 1005 | 1005 units | Nationwide throughout the US and Chile, Trinidad, China | Open, Classified |  |  | 1005 | Not_Computer
Z-2210-2015 | 71564 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Cortical Pin 2.0 mm,Sterile, FREEZE DRIED: Store at ambient temperature. Do Not freeze. |  |  | May 27, 2015 | 2015 | Musculoskeletal Transplant Foundation, Inc. | Musculoskeletal Transplant Foundation (MTF) is recalling one lot of MTF's Allograft Cortical Pins because the lot was labeled as a 2.0 mm diameter pin and it may contain a 2.4 mm diameter pin. | Labeling mix-ups | Musculoskeletal Transplant Foundation initiated this recall by either telephoning their customers or by sending a recall letter dated 5/27/15 to the customer. |  | 22 | 22 units | Distributed in the states of AZ, CO, MI, MN, NC, OH and TX. | Terminated | December 17, 2015 | 204 | 22 | Not_Computer
Z-2218-2015 | 71608 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Medical Device Exchange - Versaport Bladeless Trocar w/ Fixation |  |  | February 11, 2015 | 2015 | EXP Pharmaceutical Services Corp | EXP did not register as a medical device establishment, list the devices being recalled, orestablish its own quality system and instead relied on the fact that its vendors were registered and had their own quality systems. | Other | EXP first notified their customers to quarantine EXP devices February 11-17, 2015. Customers were contacted through telephone, e-mail, and written communications. Customers were instructed to recover and quarantine all EXP products from their inventories. EXP believes that they have contacted all customers to whom products were shipped between April 26, 2014, and December 9, 2014.A second letter is scheduled to be sent in July 2015 (date unknown at this time). This letter will advise users of the recall of due to lack of registration with the FDA. |  | 1 | 1 | Distributed in DC and the states of CO, CA, FL, TX, and AR. | Terminated | December 30, 2015 | 322 | 174 | Not_Computer
Z-2266-2015 | 71541 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Invivo Expression MRI Patient Monitoring System. |  |  | June 17, 2015 | 2015 | Philips Electronics North America Corporation | The device's labeling inaccurately indicates Endotracheal as a body temperature application site. The FlexTEMP reusable temperature sensor is not intended to be used for endotracheal placement. The application site should be listed as Esophageal. | Labeling False and Misleading | An URGENT Medical Device Correction was sent to all consignees by letter June 2015 regarding the Invivo Expression MRI Patient Monitoring System FlexTEMP Application Site explaining the reason for the affected labeling. |  | 697 | 697 devices. | Worldwide Distribution-US (nationwide) including the states of AL, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OR, OK, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY, and Alaska, Hawaii and the District of Columbia, and the countries of Argentina, Australia, Austria, Canada, Czech Republic, Denmark, Egypt, France, Germany, Hong Kong, Iceland, Israel, Italy, Japan, Netherlands, Norway, Oman, Poland, Portugal, Russian Federation, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 697 | Not_Computer
Z-2267-2015 | 71583 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | OxiMax N-65 Handheld Pulse Oximeter. N65, N65-1, N65P, and N65P-1. For continuous or spot check monitoring of functional arterial oxygen saturation (SpO2) and pulse rate. |  |  | June 29, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Potential missing segments on the display that can result in misinterpretation of data. | Device Design | Metronic sent an Urgent Medical Device Correction letter dated June 29, 2015 to affected customers. The letter identified the adfected product, problem and actions to be taked. Customers where instructed to return the attached verification form. For questions contact your Medtronic representative. |  | 211426 | 211,426 units | Worldwide Distribution - US Nationwide including Canada. | Open, Classified |  |  | 317257 | Not_Computer
Z-2269-2015 | 71376 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Prism Medical P-440 Portable Ceiling Lift; PN 303070 and 303070R (refurbished); Patient lift with weight specification of 440 lbs (200 kg). |  |  | June 24, 2015 | 2015 | ErgoSafe Products, LLC (DBA) Prism Medical | There is a potential the sling loops may not stay attached to the carry bar hooks. | Other | Prism Medical sent a Field Safety Notice to affected customers via certified first class US mail. The notice identified the affected product, problem and actions to be taken. For questions contact your local Prism Medical dealer or Prism Medical Customer Service Representative 1-800-891-6502 |  | 7105 | 7105 lifts | Worldwide Distribution - US Nationwide and the countries of:Canada, Great Britain, Australia, South Korea, and Japan. | Terminated | April 03, 2017 | 649 | 7514 | Not_Computer
Z-2262-2015 | 71663 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | FOUNDATION KNEE SYSTEM NON STERILE Femoral A/P Cut Block - Type II REF 800-01-139 |  |  | June 24, 2015 | 2015 | Encore Medical, Lp | The markings on the distal face of the instruments are mis-oriented by 180Âż. This includes Anterior and Posterior, which the surgeon uses to correctly place the cut block on the resectioned face of the femur. The cut block is not symmetric, so correct orientation of the instrument in the A/P direction is critical to making proper anterior, posterior, and chamfer cuts to fit the femoral implant. | Component design/selection | DJO Surgical sent an " Urgent Field Safety Notice" dated June 30, 2015, to the affected customer. The recalling firm requested the customer return the affected product for a replacement. For further questions, please call (512) 834-6302. |  | 10 | 10 units | US Distribution to TX and GA. | Open, Classified |  |  | 10 | Not_Computer
Z-2265-2015 | 71636 | No | EIH | Porcelain powder for clinical use. | Dental | Dental | 510(k) | IPS Empress Universal Glaze Spray 270 mL, Product Code/Model/Catalog Number 609432AN, Rx only. Porcelain glaze spray used in the fabrication of dental restorations |  |  | July 09, 2015 | 2015 | Ivoclar Vivadent, Inc. | Complaints were received of coarse particles in the glazed surface of IPS e.max Ceram restorations. A basic component of the glass contains a higher number of large grain sizes than usual. | Nonconforming Material/Component | Ivoclar Vivadent sent an URGENT MEDICAL DEVICE RECALL Letter (dated 7/8/2015) and Response Form to affected customers via Certified Mail Return Receipt Request. The letter identified the affected product, problem and actions to be taken. Customers were instructed to notify dentists of the recall and request return of product. A notification letter was sent to dentists on 7/24/2015 requesting they destroy the product on hand (if any) and report back to Ivoclar Vivodent. For questions in the United States contact US Headquarters Ivoclar Vivadent Customer Services at 1-800-533-6825. |  | 27 | Domestic: 27 units | US Distribution in the states of: PA, IN, IA and SC. | Open, Classified |  |  | 27 | Not_Computer
Z-2271-2015 | 71592 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Fabius MRI Anesthesia Machine; Fabius MRI is an inhalation anesthesia machine for use in MRI environments in operating, induction and recovery rooms. |  |  | June 22, 2015 | 2015 | Draeger Medical, Inc. | the Fabius MRI or parts of the system were attracted by the magnetic field of the MRI. If the Fabius MRI is positioned too close to the MRI, its parts can be loosened as a result of the magnetic force of the MRI. | Labeling False and Misleading | Draeger Medical sent an Urgent Medical Device Recall Letter, dated June 2015, and Instructions for Use Supplement (one supplement for each Fabius MRI Anesthesia Machine at the facility) to end users. The letter identified the affected product, problem and actions to be taken. For questions regarding the recall letter contact Michael Kelhart at 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of the DrÂżger Fabius MRI anesthesia machine contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 373 | 373 | US Nationwide Distribution | Terminated | May 05, 2017 | 683 | 373 | Not_Computer
Z-2263-2015 | 71570 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Paragon Laboratory Management |  |  | June 15, 2014 | 2014 | McKesson Technologies, Inc. | For Paragon Laboratory Management 12.1 and 12.1.1 releases, if used with Microsoft SQL Server 2012, the Final Cumulative Report may incorrectly display the data. | Software design | McKesson issued a Clinical Alert to notify customers of the issue on June 15, 2015, and requested that customers disable Cumulative Reporting. McKesson is identifying customers currently in a live environment with Microsoft SQL Server 2012 and the affected configuration. McKesson will provide those customers with corrective software, and provide assistance in generating correct reports and updating those to the patient folder. For further questions, please call (404) 338-3556. |  | 128 | 128 | Nationwide Distribution including Puerto Rico and to the states of : CA, CO., CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OK, OA, PA, SC, SD, TN, TX, VT, VA, WV and WY. | Open, Classified |  |  | 128 | Not_Computer
Z-2272-2015 | 71595 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Synthes TI Matrix Pre-Bent Maxillary Plate / 2mm ADVMNT/0.8mm THK/L and Synthes TI Matrix Pre-Bent Maxillary Plate / 2mm ADVMNT/0.8mm THK/R. The Synthes Matrix ORTHOGNATHIC Fixation System is intended for use in selective trauma of the midface and craniofacial skeleton; craniofacial surgery; reconstructive procedures; and selective orthognathic surgery of the maxilla, mandible and chin in adults and children. |  |  | June 24, 2015 | 2015 | Synthes (USA) Products LLC | The affected part and lot numbers of the TI Matrix Pre-Bent Maxillary Plates may potentially have the incorrect laser etch denoting the orientation. | Process control | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated June 24, 2015, to all affected customers. Customers were asked to check their inventory and return affected devices with the completed response form. Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales Consultant. |  | 31 | 31 | Nationwide Distribution including CA, UT, CO, NJ, WA, VT, KY, WY, CT, NY, MS, and MD. | Open, Classified |  |  | 31 | Not_Computer
Z-2280-2015 | 71212 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | ANSPACH - Fluted Drum, Side Cut only 10MM x 14MM , REF L-10DRC-G1 |  |  | April 13, 2015 | 2015 | The Anspach Effort, Inc. | G1 Dissection Tools are not compatible with all attachments listed on the labeling; G1 cutting tools are labeled with a Specified Shaft Exposure which may not be consistent when used with each attachment on the product label. Shaft exposure may be greater. | Device Design | Anspach sent an Urgent Notice: Medical Device Recall (Correction) letter of Anspach G1 Dissection Tools dated May 6, 2015, to all affected customers. Customers were not required to return their inventory of GI Dissection tools. Customers were instructed to complete, sign, and return the attached reply form confirming receipt of the letter by fax or email. Customers were also instructed to forward the letter as appropriate immediately. Customers with questions should contact the Complaint Handling Unit Manager at 561-494-3673 or contact their Anspach Sales Representative.For questions regarding this recall call 561-627-1080. |  | 65 | 65 tools. | Worldwide Distribution - US including AZ, CA, CO, FL, GA, IL, KY, LA, MN, MO, NC, NJ, NY, OH, OR, PA, TX, and Internationally to Canada and China. | Open, Classified |  |  | 5320 | Not_Computer
Z-2279-2015 | 71602 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Hardy Diagnostics |  |  | June 26, 2015 | 2015 | Hardy Diagnostics | Hardy Diagnostics is recalling MacConkey Agar with Ciprofloxacin, Brain Heart Infusion Broth with Ciprofloxacin 1 ug, MDR Acinetobacter, BEA with Azide and Vancomycin, BEA Agar with Vancomycin, Gentamicin, and Amphotericin B, BEA Broth with Vancomycin, and VRE Broth due to lack of 510(k) clearance. | No Marketing Application | The firm initially called customers to inform them of the recall. The firm's phone script stated that they were discontinung the product and as part of the process were recalling anything that still might be within expiration date. The instructions provided via phone were to discard any remaining plates that are in stock in order for credit. The phone script stated that the firm was going to be sending something in writing. On 06/25-29/15 the firm sent out notification letters to customers to inform them that they were discontuing the product, and were requested anything remaining within expiration to be discarded. The firm requests that customers complete, sign, and fax back the endlosed form stating their compliance with the above action. The notification states if the laboratory was unable to be reached via telephone and would like a credit to contact the Customer Service Department at 800-266-2222, option 1. For any questions the firm states to contact Technical Services Department at 800-266-222, option 2 or via email at TechService@HardyDiagnostics.com.Follow-up calls will be made to customers who do not respond. |  | 900 | 900 each (90 pk/10) | U.S. Nationwide Distribution and the country of: Canada | Terminated | September 11, 2015 | 77 | 194840 | Not_Computer
Z-2314-2015 | 71687 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S 1000, ACUSON S 2000, or ACUSON S 3000 ultrasound systems with software version C3, C3, C3, or C1. |  |  | June 25, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Potential measurement error on ACUSON S Family ultrasound system when using the Doppler manual trace measurement tool in full screen format that has a low probability of misdiagnosis. | Device Design | Customer Safety Advisory notifications were sent by certified mail to US customers on June 25, 2015. Siemens Service organization coordinate the implementation of the field correction with the distributors for all countries.Letters advised users of the potential miscalculation on ACUSON S1000, ACUSON S2000, or ACUSON S3000 ultrasound systems with software version VC30A, VC30B, VC30C, or VC31A. The manual trace tool may incorrectly calculate the Time Average Velocity and subsequently incorrectly calculate the Pulsatility index. This could lead to potential misdiagnosis.To mitigate the problem users should not perform measurement utilizing the manual Doppler trace tool in a full screen display format.This issue will be resolved in a future software release. If you have any questions, please contact your Siemens Service for more information.Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization who need to be aware of this issue.Siemens is continuing to ship in forward production and will be including the customer safety advisory notice with the shipment. Siemens will update the software for those customers as well once it is released. For further questions please call (650) 694-5398. |  | 1789 | 1789 units | Worldwide Distribution - Nationwide Distribution including Puerto Rico and to the countries of : Algeria, Angola, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia Herzeg., Brazil, Bulgaria, Canada, Chile, Colombia, Croatia, Cyprus, Czech Republic, Dem. Rep. Congo, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Malaysia, Mexico, Morocco, Nepal, Netherlands, New Zealand, Norway, P.R. China, Peru, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad,Tobago, Turkey, Turkmenistan, U.A.E., Ukraine, United Kingdom, Uzbekistan, Venezuela and Vietnam. | Open, Classified |  |  | 1789 | Not_Computer
Z-2315-2015 | 71605 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Aequalis Humeral Nail Targeting Jig, Product Code: 9020060. Orthopedic manual surgical instrument. |  |  | June 25, 2015 | 2015 | Tornier, Inc | Recall for the Aequalis IM Nail instrumentation set (Tray Number 9020000) due to several reports that the mast of the targeting jig (Part Number 9020060) is separating from the jig boom. Separation will impact alignment and can affect the ability to fixate Aequalis IM Nail screws. | Nonconforming Material/Component | Consignees were sent on 6/25/2015 a Tornier "Urgent Medical Device Recall" letter dated June 24, 2015. The letter described the problem and the product affected by the recall. The letter provided information on jig inspection and referred consignee to the attached "Perfect Circles' surgical technique that should be used when separation is observed. Requested consignees to complete and return the questionnaire to FieldAction@tornier.com. For questions they can contact Customer Service Department at 1(888)494-7950. |  | 99 | 99 | US: AL, AK, AZ, CA, CO, CT, FL, IL, IN, ME, MD, MI, MO, NC, OH, PA, TX.OUS: CANADA AND FRANCE. | Open, Classified |  |  | 99 | Not_Computer
Z-2322-2015 | 71550 | No | DJZ | Breath-alcohol test system. | Toxicology | Toxicology | 510(K) Exempt | Mission |  |  | June 17, 2015 | 2015 | Acon Laboratories, Inc. | Acon Laboratories, Inc. is recalling Mission Breath Alcohol Detector due to a labeling amendment. | Labeling False and Misleading | On 06/17/15 the firm sent out customer notification letters to their distributors. The letter identified the affected product, problem and actions to be taken. In the letter the firm requests that Mission Breath Alcohol Detectors be quarantined. The firm asks that distributor complete the "fax back" form to allow for the return of unexpired product for a replacement. Once the form is received the firm is going to issue a return authorization number as well as provide a return label.The contact noted in the letter is Qiyi Xie MD, MPH, Sr. Officer Regulatory and Clinical Affairs at 858-857-8011. |  | 1283 | 1,283 kits | U.S. Nationwide Distribution in the states of: NY, CA, FL, VA, LA, NJ, MN, PA, MI, OH, TX, VA, UT, AL, and GA | Terminated | December 08, 2015 | 174 | 1283 | Not_Computer
Z-2321-2015 | 71688 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 and ACUSON S3000 ultrasound systems with the Virtual Touch IQ option; |  |  | June 25, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Potential measurement error on ACUSON S Family ultrasound system. When repositioning the Virtual Touch IQ region of interest from the original (default) location, the lateral position of the measurement tool may not align with the lateral position of the shear velocity data. | Software Design Change | On June 25, 2015 Customer Safety Advisory Notification letters were sent to all affected customers by Certified mail. The letter identified the affected product, problem and actions to be taken. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. For questions contact your Siemens Service for information. |  | 629 | 629 units | Worldwide Distribution - US Nationwide and the countries of Algeria, Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Croatia, Czech Republic, Denmark, Dominican Rep., France, Germany, Greece, Hong Kong, India, Indonesia, Israel, Italy, Japan, Kenya, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, P.R. China, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., Ukraine, United Kingdom. | Open, Classified |  |  | 629 | Not_Computer
Z-2316-2015 | 71300 | No | HXD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SYMMETRY SURGICAL DISPOSABLE TOWEL CLAMP, REF 19-4137, STERILE EO, Rx only. |  |  | June 16, 2015 | 2015 | Symmetry Surgical, Inc. | Lack of sterility assurance. | Package design/selection | The firm initiated their recall by via US certified mail, with return receipt requested, beginning on 06/18/2015. The Urgent Medical Device Recall letter, dated June 16, 2015, identified the affected product, as well as the reason for the recall. The letter discussed the potential harm associated with the issue, and requested that the consignees examine their inventory and discontinue use immediately. Customers should promptly call (800) 251-3000 for a return material authorization (RMA) number to support the return of the impacted lots. Symmetry Surgical Customer Service should be contacted with questions or any other needs. |  | 2708 | 2,708 units | Worldwide Distribution -- US, including the states of CA, VA, TX, FL, OH, NY, MS, KY, UT, GA, PA, MO, IN, NV, AL, MI, MA, and LA; and, the countries of Germany and Mexico. | Open, Classified |  |  | 2708 | Not_Computer
Z-1991-2015 | 71649 | No | RBS |  |  |  |  | Olympus Scientific Solutions Americas Analytical X-ray System. Delta, Delta Element, Delta Plus, DPO-4000-XX-U-EN-EN-PM-SP12, DPO-6000-XX-U-EN-EN-AP-SP10. |  |  | June 09, 2015 | 2015 | Olympus Scientific Solutions Americas | Higher than published scatter around the probe head when testing some samples. At the point the customer would interact directly with the instrument (Trigger position), most readings were lower. However, measurements around the probe head were higher particularly on samples know to be highly scattering (plastics, soils) with minimal increases on high density samples (steel, aluminum). | Component change control | Olympus Scientific Solutions planned action:1. Notify product managers as to potential issues 2. Distribute new table and notice to instrument owners 30 JUL 153. Post new table on website for access 15 JUN 15 |  | 1047 | 1047 | Nationwide Distribution. | Open, Classified |  |  | 1047 | Not_Computer
Z-2326-2015 | 71709 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S4C Occiput Torque Wrench F/Set Screw; used to tighten the S4 Cervical Spinal System clamping screws with a defined torque 2.8 Nm for tightening the clamping screw in the screw head/hook head/occipital plate/transverse connector and other connectors. |  |  | July 07, 2015 | 2015 | Aesculap, Inc. | A FW103R torque wrench that was improperly reassembled after reprocessing caused a screw to break during implantation due to too much torque. The firm did not include instructions for reassembly after reprocessing. | Labeling design | Aesculap sent Field Safety Notification Letters, dated July 7, 2015, to end users, sales consultants, and distributors to advise them about the issue. Enhancements to the instructions for use were made available on the firm's website. |  | 10 | 10 | Distributed US (nationwide) and to Canada. | Terminated | April 17, 2017 | 650 | 10 | Not_Computer
Z-2325-2015 | 71756 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic Passive Biopsy Needle Kit, Part Number 9731754 - Passive Biopsy Fee Per Use (FPU) Kit. |  |  | July 17, 2015 | 2015 | Medtronic Navigation, Inc. | The adhesive between the sphere assembly and the biopsy needle shaft may be compromised, causing the sphere assembly to move on the needle shaft. This may result in inaccurate navigation and improper placement of the needles biopsy window in the brain. | Process change control | Medtronic sent an Urgent Medical Device Recall letter dated July 17, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for any of the affected lots and quarantine them if found. They will be asked to complete a confirmation form indicating they have completed this action and are to email or fax the form to Medtronic. Customers will be asked to contact Medtronic to obtain a return material authorization number and arrange for no-charge replacements of impacted products. Customers with questions were instructed to contact Medtronic Technical Services at 800-595-9709. For questions regarding this recall call 720-890-3302. |  | 1415 | 1415 total | Worldwide Distribution - US (nationwide) and Canada. | Terminated | April 07, 2016 | 265 | 1415 | Not_Computer
Z-2327-2015 | 71740 | No | OVE | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Titan Spine Endoskeleton Locking Bone Screw Assembly, which is part of the Endoskeleton TCS Interbody Fusion Device, spinal implant. Part Numbers / Product Description: 5302-3514 / 03.5 x 14mm Locking Screws; |  |  | July 20, 2015 | 2015 | TITAN SPINE, LLC | Titan Spine is recalling specific Endoskeleton TCS Integrated Locking Bone Screw Assembly due to the locking bone screw collars not 'locking'. | Nonconforming Material/Component | Physicians and Distributors were sent on 7/20/2015 separate Titan "Urgent Medical Device Recall" letters dated July 20, 2015. The letters described the problem and the product involved in the recall. Advised consignees to complete and return the enclosed response form. For questions they contact Jane Rodd Senior AQ/RA Manager at 262-242-7801. Distributors were advised to examine their inventory and quarantine the product. If they have further distributed the product they should identify their customers and advise Titan Spine of their contact information. |  | 137 | 137 units | Distributed in the states of CA, TX, OK, MA, NY, IN, OR, and GA. | Terminated | October 20, 2015 | 92 | 137 | Not_Computer
Z-2328-2015 | 71584 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor |  |  | June 05, 2015 | 2015 | Philips Electronics North America Corporation | Multiple software and hardware issues with device that can affect its function. | Device Design | On June 29, 2015, Philips sent an Urgent Medical Device Correction notification/Field Safety Notice (FSN) to inform customers of the issue. The FSN identifies details of the units affected, gives instructions on actions to be taken by the customer, and identifies what action Philips plans to take to remedy the issue. An upgrade consisting of both software and hardware will be provided free of charge for all units affected by these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrades. Philips is asking customers to follow the "Action to be Taken by Customer/User" section of the Urgent Medical Device Correction notification/Field Safety Notice.For questions contact your local Philips representative. |  | 13168 | 13,168 devices | Worldwide Distribution - US Nationwide including PUERTO RICO, Canada, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BULGARIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICA, DOMINICAN REPUBLIC, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GABON, GERMANY, GREECE, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN (ISLAMIC REPUBLIC OF), IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, Jordan, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LIBERIA, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACAU, MALAYSIA, MALTA, MAURITIUS, MEXICO, MOROCCO, MYANMAR (Burma), NAMIBIA, NEPAL, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION, RWANDA, SAUDI ARABIA, SENEGAL, SERBIA, SEYCHELLES, SINGAPORE, SLOVAKIA (Slovak Republic), SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, United Republic of TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VIETNAM and ZIMBABWE. | Open, Classified |  |  | 13168 | Not_Computer
Z-2331-2015 | 71712 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph mCT S(40)-4R, Model Number 10248671 |  |  | July 14, 2015 | 2015 | Siemens Medical Solutions USA Inc. | Emission computed tomography systems and computed tomography xray systems with optional ceiling mounts may be missing a locking ring. | Packaging change control | The firm initiated their recall on 07/14/2015 by telephone to the sole US consignee. The update is scheduled for 07/15/2015. |  | 1 | 1 unit | Distributed in the state of Texas and the countries of Germany, France, Russia, Denmark, and Switzerland. | Terminated | March 03, 2016 | 233 | 12 | Not_Computer
Z-2340-2015 | 71711 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens Artis zee systems; |  |  | June 29, 2015 | 2015 | Siemens Medical Solutions USA, Inc | The possibility exists for a large amount of liquid to seep through the top cover of the table and penetrate the components inside the table. If this occurs, table movements may no longer be possible. No injuries reported. | Device Design | A customer safety notice, dated June 29, 2015, was sent to direct accounts to alert them about the issue and that Siemens' service organization will be contacting them to perform the corrective action. |  |  |  | US Nationwide | Open, Classified |  |  |  | Not_Computer
Z-2330-2015 | 71609 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | S2 Fluid Pack on the Cobas b 221 system- Catalog Number 03260925184 |  |  | June 24, 2015 | 2015 | Roche Diagnostics Operations, Inc. | The affected S2 fluid packs may generate a transponder error when the fluid pack is loaded on to the cobas b 221 system. The error prevents the fluid from being recognized by the system and the instrument cannot start measurements until the affected S2 Fluid Pack is replaced. | Process control | Roche Diagnostics sent an Urgent Medical Device Correction dated July 2, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are asked to do the following: 1. Keep in the inventory a backup S2 Fluid Packs from an unaffected lot, or at least four additional S2 Fluid Packs from the affected lots when product from unaffected lots is not available.2. If the S2 Fluid Pack generates a Transponder error message, contact the Roche Support Network Customer Support Center at 1-800-526-2272 (provide the lot and serial number printed on the side of the affected S2 Fluid Pack).3. After contacting Roche Support Network Customer Support Center, dispose of the affected S2 Fluid Pack per local guidelines. Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-526-2272, if you have questions about the information contained in this Urgent Medical Device Correction. |  | 407564839 | 19 affected devices of lot 21450781 were distributed in the US. | US Distribution: CT, TX, LA, ND, MI, NJ, OH, OK, PR, SC, and KY. | Terminated | January 12, 2016 | 202 | 407564839 | Not_Computer
Z-2345-2015 | 71659 | No | HTZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Battery Power Line II Oscillator Model 530.710, orthopedic surgical power tool. |  |  | July 01, 2015 | 2015 | The Anspach Effort, Inc. | Saw head separated from the drive unit or began to come loose. | Nonconforming Material/Component | DePuy Synthes sent an " Urgent Notice Medical Device Recall (Removal) letter dated July 1, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.ACTION REQUIRED:Our records show that your facility has the affected product(s) subject to this removal. The Anspach Effort, on behalf of DePuy Synthes asks that you review your inventory and immediately remove any affected BPLII Battery Oscillator devices from stock. In addition to this letter, your DePuy Synthes Sales Consultant has been provided with new product(s) to immediately replace the affected product(s).If you have any questions regarding this removal, please contact the Complaint Handling Unit Manager at 561-494-3673 or contact your DePuy Synthes Sales Consultant. |  | 360 | 360 handpieces (Currently there are 9 affected units remaining in the US and 17 units OUS all others were reworked) | Worldwide Distribution - US Distribution to the states of : AR, LA, OH, and WA., and the countries of : China, France, Italy, Russia, South Africa, Switzerland, and United Kingdom. | Open, Classified |  |  | 360 | Not_Computer
Z-2346-2015 | 71510 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Atrium 36Fr Right Angled Firm PVC Thoracic Catheter, (Pkg of 10) |  |  | June 19, 2015 | 2015 | Atrium Medical Corporation | Inner pouches incorrectly labeled as part 8136, 36Fr Right Angled Thermosensitive PVC catheter, but are part 15236-36Fr Right Angled Firm PVC Catheter | Employee error | Atrium Medical notified accounts by letter dated 6/18/15, explaining the mislabeled inner units. If any product is discovered with incorrect labeling, the product may be used or product may be returned for credit. contact your local Atrium/Maquet chest drainage representative or Atrium Customer Service at 1-800-528-7486 (603-880-1433), Monday through Friday between 8:00 am and 5:00 pm (EST). |  | 62 | 62 cases (10/cs)=620 units | AZ, CA, LA. IN, NC, TN, TX, VA | Terminated | March 21, 2016 | 276 | 62 | Not_Computer
Z-2318-2015 | 71692 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Beacon Tip Royal Flush Plus High-Flow Catheter. Angiographic catheter. |  |  | July 02, 2015 | 2015 | Cook Inc. | Reports of catheter tip splits and or separation. Loss of device function, medical intervention to retrieve separated segment, or complications resulting from separated tip occluding blood flow to end organs may result. | Component design/selection | Cook Medical sent an Urgent Medical Device Recall letter dated July 2, 2015, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification and instructions for returning the affected product. All customers were notified via courier. Customers with questions were instructed to contact Cook Medical Customer Relations at (800) 457-4500 or (812) 339-2235.Firm issued press release 8/3/2015. |  | 268169 | 268169 | Worldwide Distribution - US (nationwide including Puerto Rico and US Virgin Islands), Algeria, Antigua, Armenia, Austria, Australia, Bahamas, Belgium, Bangladesh, Bulgaria, Bahrain, Brazil, Brunei Darussalam, Canada, Canary Islands, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, El Salvador, Finland, France, French Polynesia, Georgia, Germany, Ghana, Great Britain, Greece Guadeloupe, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jersey, Jordan, Kuwait, Kazakhstan, Republic of Korea, Kuwait, Lebanon, Lithuania, Luxemburg, Latvia, Libya, Macau, Malaysia, Malta, Martinique, Mexico, Mongolia, Morocco, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Nicaragua, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Sri Lanka, Sweden, Switzerland, Spain, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Open, Classified |  |  | 967886 | Not_Computer
Z-2264-2015 | 71719 | Yes | NPT |  |  | Cardiovascular | PMA | EnVeo R Loading System, Model No. LS-EnVeoR-23, LS-EnVeoR-2629, LS-EnVeoR-23-C, LS-EnVeoR-2629-C. |  |  | July 15, 2015 | 2015 | Medtronic Cardiovascular Surgery-the Heart Valve Division | Medtronic is recalling the EnVeo R Loading Systems because of the presence of particulate. | Process control | An Urgent Medical Device Recall Letter dated July 2015 will be be sent to customers who purchased the EnVeo R Loading System because Medtronic has received 8 reports related to this issue out of 7347 potentially affected units. The letter informs customers of the problems identified and the actions to be taken. Customers are instructed to complete the customer confirmation certificate and to email RS.CFQFCA@Medtronic.com or fax to Medtronic at (651) 367-0612 to the attention of Customer Focused Quality. |  | 6912 | 6,912 units total (540 units in US) | Worldwide Distribution - US Nationwide, Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Greece, Hong Kong, Ireland, Israel, Italy, Kuwait, Latvia, Lithuania, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Turkey, United Arab Emirates and United Kingdom. | Terminated | May 20, 2016 | 310 | 6912 | Not_Computer
Z-2344-2015 | 71697 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S Family ultrasound systems utilizing the 18L6 HD and/or 6C1 HD biopsy guidelines |  |  | June 29, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | When using the CIVCO biopsy attachments in combination with the ACUSON S Family on-screen guidelines, the biopsy needle may traverse outside the on-screenguidelines. | Under Investigation by firm | Siemens sent a Customer Safety Advisory Notification letters to affected customers via certified mail with return receipt. The letter identified the affected product, problem, potential risk to patients and actions to be taken. Seimens instructed that this notice be shared with all personnel within your organization who need to be aware of this issue. For questions contact your Siemens service representative. |  | 4140 | 4140 total | Worldwide Distribution - US Nationwide and the countries of Algeria, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Brazil, Brunei, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Czech Republic, Dem. Rep. Congo, Denmark, Dominican Rep., Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Lithuania, Luxembourg, Macedonia, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, P.R. China, Peru, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, U.A.E., Ukraine, United Kingdom, Venezuela and Vietnam. | Open, Classified |  |  | 4140 | Not_Computer
Z-2339-2015 | 71788 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Signa Infinity MR System; Signa MR/i MR System; SmartSpeed; EchoSpeed Plus; HighSpeed Plus, Signa Infinity MR System with EXCITE Technology, Excite II, Signa 3T Infinity with Excite, Signa Excite 1.5T and Signa Excite 3T, Signa Excite HD, Signa HD 1.5T and Signa HD 3.0T, Signa HDx, Signa HDxt, Signa HDi. Whole body magnetic resonance diagnostic imaging scanners. |  |  | July 10, 2015 | 2015 | GE Healthcare | A gradient cable in the ACGD cabinet may be missing a spacer that prevents connecting the gradient cable terminals in the wrong polarity on X gradient amplifier. This could cause images to be flipped left-right with incorrect orientation annotation. | Nonconforming Material/Component | GE Healthcare sent an Urgent Medical Device Correction letter dated July 10, 2015, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Consignees with questions were instructed to contact GE Healthcare Service at 1-800-437-1171 or their local Service Representative. |  | 82 | 82 (39 USA; 43 OUS) | Worldwide Distribution - US (nationwide) and Internationally to UNITED ARAB EMIRATES, ARGENTINA, BRAZIL, INDIA, EGYPT, CHINA,FINLAND, FRANCE, UNITED KINGDOM, HUNGARY, JAPAN, INDONESIA, KOREA, MEXICO, NEW ZEALAND, CANADA, AUSTRALIA, ITALY, RUSSIAN FEDERATION, TAIWAN, and SOUTH AFRICA. | Terminated | November 19, 2015 | 132 | 82 | Not_Computer
Z-2348-2015 | 71575 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 18inch Dual Port/Single Bladder Disposable Tourniquet Cuff, Zimmer Surgical, Inc. Tourniquets are intended to be used by qualified medical professionals to exert enough pressure on the arterial blood flow in a limb to produce a bloodless operating field. Single bladder tourniquet cuffs are generally used for operations lasting less than 90 minutes. |  |  | June 24, 2015 | 2015 | Zimmer Surgical Inc | The inside of the folded Instructions For Use (IFU) pamphlet was missing the usage instructions and warnings. | Labeling False and Misleading | Zimmer sent an Urgent Medical Device Correction letter dated June 30, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification, identify the affected product and replace the IFU with the enclosed corrected IFU.Customers were also instructed to complete the attached Response Form and return it via email to CorporateQuality.PostMarket@zimmer.com. Customers with questions were instructed to call 330-364-0989. |  | 61 | 61 units | Worldwide Distribution - US including AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV &amp; WY and Internationally to Canada. | Open, Classified |  |  | 391 | Not_Computer
Z-2347-2015 | 71814 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo, 9.10, 9.20.0, 9.20.1,9.20.2, 9.30, 9.40.0, 9.40.1, 9.40.2 with Massimo PHASEIN End Tidal CO2 (EtCO2) module. Merge Hemo (formerly named HeartSuite Hemodynamics) monitors, measures, and records physiologic data from a human patient undergoing a cardiac catheterization procedure. |  |  | January 21, 2015 | 2015 | Merge Healthcare, Inc. | A system freeze-up of the Merge Hemo system, that included the PHASEIN End Tidal CO2 unit, resulted in the user needing to move the patient to another cath lab. | Software design | Merge Healthcare sent an Important Product Information letter dated January 21, 2015, to all affected consignees. The letter described the Issue, Affected Product, Instructions, Product Updates, and Contact Information. For questions contact Merge Customer Service at (877) 741-5369 or support@merge.comA second UPDATE letter was sent April 16, 2015. The letter included the Product Updates.For questions regarding this recall call 866-261-5761. |  | 451 | 451 | Nationwide Distribution including IL, SC, WI, AL, CT, CA, LA, NM, TX, OK, GA, IA, OH, FL, CO, TN, NY, NC, MA, AR, MO, HI, ND, AK, UT, WV, WA, AZ, MI, PA, ME. | Terminated | December 09, 2015 | 322 | 451 | Not_Computer
Z-2352-2015 | 71694 | No | GBQ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ReSolve Locking Drainage Catheters |  |  | June 29, 2015 | 2015 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is voluntarily conducting a recall due to a potential discrepancy between the unit labeling of the Merit ReSolve Drainage Catheters and the actual device in thepackage. | Employee error | Merit Medical Systems sent an Urgent Product Recall Notice dated July 10, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions required of you:1. Please immediately determine if any of the devices are within your facility, quarantine them, and discontinue use. Due to the nature of this label discrepancy, you will need to physically verify whether product in inventory is within the scope of this recall. Your Merit Sales Representative can assist you with this determination.2. Ensure that all personnel to whom the devices were distributed are made aware of this field action.3. Please work with your Merit Sales Representative to arrange product return.4. If you have any questions concerning this communication, please don't hesitate to contact your Merit Sales Representative or call (801)208-4365. |  | 950 | 950 | Nationwide Distribution including the Virgin Islands. | Terminated | October 26, 2015 | 119 | 950 | Not_Computer
Z-2323-2015 | 71630 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | ReSTOR intraocular lens and ReSTOR Toric intraocular lens. |  |  | April 15, 2015 | 2015 | Alcon Research, Ltd. | There are reports of post - operative inflammation and/or toxic anterior segment syndrome (TASS) after cataract surgery. | Process design | Alcon Research sent recall notification letters dated April 16, 2015, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to stop using the affected product and place on hold until their sales representatives visit their site for withdrawal. Follow up communication activities included personal contact via phone and site visits. For questions regarding this recall call 817-293-0450. |  | 45391 | 45,391 | Distribution Japan, No US distribution. | Terminated | June 30, 2017 | 807 | 45391 | Not_Computer
Z-2351-2015 | 71614 | No | GTY | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | PathoDx Strep B Grouping Latex, REF 62031, 3 ml packaged in plastic dropper bottle with a pink cap. |  |  | June 30, 2015 | 2015 | Remel Inc | A reagent present may produce weak or slow reactions; continued use may result in a failure to identify target species. | Under Investigation by firm | On 4/28/2015, the recalling firm's European subsidiary notified the recalling firm of the recall. On 6/30/2015, the recalling firm notified their customers of the recall via letter, sent USPS standard mail. |  | 4420 | 4420 bottles | Nationwide Distribution, including Puerto Rico and the country of Canada. | Open, Classified |  |  | 6279 | Not_Computer
Z-2356-2015 | 71718 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Trumpf Ceiling Mounted Surgical Lighting Systems,TruVidia SD, User Manual #1400191. |  |  | July 27, 2015 | 2015 | Trumpf Medical Systems, Inc. | Complaints received of the front joint of the spring arm AC 2000 could break at the interface to the monitor. | Process control | Consignees were notified by letter on/about 7/27/2015. |  | 9287 | 9,287 units total | Worldwide Distribution-US (nationwide) including Puerto Rico, and the countries of Angola, Antilles, Argentina, Austria, Australia, Azerbaijan, Bosnia and Herzegovina, Belgium, Bulgaria, Bahrain, Brasilia, Belarus, Canada, Switzerland, Chile, China, Columbia, Cyprus, Czech Republic, Germany, Denmark, Algerian, Ecuador, Estonia, Egypt, Spain, Finland, France, Gabon, United Kingdom, Georgian Republic, Greece, Hong Kong, Croatia, Hungary, Indonesia, Ireland, Israel, India, Iraq, Iran, Iceland, Italy, Jordan, Japan, Kenya, South Korea, Kuwait, Kazakhstan, Lebanon, Liechtenstein, Lithuania, Luxembourg, Latvia, Libya, Morocco, Macedonia, Myanmar/Burma, Malta, Mauritius, Mexico, Malaysia, Netherlands, Norway, New Zealand, Peru, Philippines, Pakistan, Poland, Palestine, Portugal, Qatar, Romania, Russia, Saudi Arabia, Sudan, Sweden, Singapore, Slovenia, Slovakia, San Marino, Syria, Thailand, Turkmenistan, Tunisia, Turkey, Taiwan, Ukraine, United Arabian Emirates, Uruguay, Venezuela, British Virgin Islands, Vietnam, Kosovo, Serbia and South Africa.* | Terminated | May 03, 2017 | 646 | 9287 | Not_Computer
Z-2358-2015 | 71613 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | InvacareÂż ProntoÂż Air Personal Transporter (power wheelchair), Invacare Corporation, One Invacare Way, Elyria, Ohio, 44035-4190, (800) 333 -6900. |  |  | February 02, 2015 | 2015 | Invacare Corporation | The batteries used in the InvacareÂż ProntoÂż Air Personal Transporter models power wheelchairs have the possibility to overheat during charging which could potentially result in leaking battery acid and/or swelling battery, loss of use, chemical burns and/or property damage. | Device Design | Invacare sent an " Urgent Medical Device Recall " letter on April 6, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.You should immediately contact your provider at the number below so a replacement batteries can be installed.If you have any questions concerning these instructions, please call Invacares recall support line at (877) 413-6008, Monday  Friday, 8 a.m.  5 p.m. EST, for assistance. Visit www.invacare.com/recall for more information about this recall. |  | 1057 | 1,057 finished devices | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico, and to the states of : AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA and WI., and to the countries of : AUSTRALIA, CANDA, FRANCE, NEW ZEALAND, SOUTH AFRICA, TRINIDAD AND TOBAGO | Terminated | September 14, 2015 | 224 | 1057 | Not_Computer
Z-2360-2015 | 71677 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon PS X3 Tibial Insert; No. 3 Triathlon TS Plus Tibial Insert X3 Poly 19 mm |  |  | May 29, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics received a report that 5537-G-319 No 3. Triathlon TS Plus Tibial Insert X3 Poly 19 mm i insert was in a package labeled as a 5532-G-511 Triathlon PS Stibialism Insert. | Under Investigation by firm | Stryker Orthopaedics sent an Urgent Medical Device Recall Notification letter dated July 9, 2015 and Product Accountability Forms via return receipt. The letter identified the affected product, problem, potential hazard, risk mitigation and actions to be taken. Contact Stryker Orthopaedics Sales Representative to arrange for return of the affected product. For questions call 201-831-5826. |  | 26 | 26 units | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 26 | Not_Computer
Z-2366-2015 | 71431 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Oasis Coil Extension Cable used with Hitachi Oasis MRI System |  |  | April 25, 2014 | 2014 | Hitachi Medical Systems America Inc | The firm received a complaint stating the Coil Extension Cable accessory would overheat and be a burn hazard if touched. | Nonconforming Material/Component | Hitachi sent an Urgent: Device Removal Notice letter date April 24, 2014 to their customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers to discontinue using the Coil Extension Cable and that a Field Service Engineer will schedule a visit to remove the cable. For questions call 800-800-3106 ext 3720. |  | 173 | 173 units | US Nationwide in the states: AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NJ, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI &amp; WY and the country of Mexico. | Open, Classified |  |  | 173 | Not_Computer
Z-2361-2015 | 71665 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | i-Stat PT/INR Cartridges |  |  | June 15, 2015 | 2015 | Abbott Point Of Care Inc. | Certain lots of PT/INR cartridges may generate a higher than typical number of PT/INR star outs potentially delaying the generation of patient results for PT/INR. | Under Investigation by firm | Abbott Point of Care sent a recall letter dated June 15, 2015 via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the enclosed Business Reply Card (BRC). A Proof of Destruction From (FG-4533 in included with letter and is required to be returned to APOC as part of the credit process. Customers with questions were instructed to contact APOC Technical Support at 800-366-8020, Option 1 or their APOC representative. |  | 423168 | 423,168 cartridges | Worldwide Distribution - US (nationwide) and Internationally to China, Lebanon, Costa Rica, Singapore, Brazil, Jordan, Russia, Greece, Austria, India, Nigeria, United Kingdom, Korea, Saudi Arabia, Morocco, Slovenia, Portugal, Tunisia, Kazakhstn, Qatar, U.A.E. Israel, Taiwan, Mexico, Argentina, Poland, Italy, Hong Kong, Kuwait, Switzerland, Turkey, Ireland, Chile, Germany, and Australia. | Terminated | February 19, 2016 | 249 | 423168 | Not_Computer
Z-2364-2015 | 71674 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Perfectum Tuberculin Glass Syringe, Non sterile |  |  | July 09, 2015 | 2015 | Cadence Inc. | Graduation marks of the syringe are not centered with the numeric position marks on the syringe barrel. | Device Design | Cadence issued the recall letter, specific to the customer dated 7/9/15 via registered mail on July 17, 2015. Accounts are requested to Recalled products must not be used.2. Locate and quarantine all affected products.3. Contact Cadence Science, Inc. using the Response Form provided.4. If this information is received by a dealer, wholesaler, or distributor/reseller that distributed any of these products to other facilities, then notify their customers of this action so that those customers can locate and remove all affected products. Any questions contact: Phone: 401-942-1031 |  | 766 | 766 units | NationwideForeign Canada, Denmark. Italy | Open, Classified |  |  | 2431 | Not_Computer
Z-2363-2015 | 71445 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ReUnion TSA Shoulder System, Standard SR Humeral Head Trial, 48X18, Catalog number 5901-S-4818. Orthopedic manual surgical instrument. |  |  | May 12, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | A package labeled as Part #5901-S-4818, Lot MAC7C14, standard humeral head trial, contained part # 5901-E-4818 Lot MAC7C14, offset humeral head trial. | Process control | Stryker Orthopaedics notified their branches/agencies via email on 5/12/2015 and sent notification letters and product accountability forms to hospital risk management/branches on 7/15/2015 and 7/16/2015 |  | 21 | 21 units | US Nationwide | Open, Classified |  |  | 21 | Not_Computer
Z-2359-2015 | 71729 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 8.5mm Reamer Head for Flexible Intramedullary Nailing System; An orthopedic manual surgical instrument is a nonpowered hand-held device intended for medical purposes to manipulate tissue or for use with other devices in orthopedic surgery. This generic type of device includes the cerclage applier, awl, bender, drill brace, broach, burr, corkscrew, countersink, pin crimper, wire cutter, prosthesis driver, extractor, file, fork, needle holder, impactor, bending or contouring instrument, compression instrument, passer, socket positioner, probe, femoral neck punch, socket pusher, reamer, rongeur, scissors, screwdriver, bone skid, staple driver, bone screw starter, surgical stripper, tamp, bone tap, trephine, wire twister, and wrench. |  |  | July 01, 2015 | 2015 | Synthes (USA) Products LLC | The possibility exists for intraoperative reamer head breakages which could also allow for un-retrieved fragments of non-implant grade material: additional lot numbers identified. No new adverse events reported. | Other | An urgent notice of medical device recall, dated July 1, 2015, was sent to end users and sales consultants to alert them about the device defect and potential risk to patients. Recipients were asked to complete the verification form and return it along with any affected devices their posession. |  | 4329 | 4329 | Worldwide Distribution. | Open, Classified |  |  | 4329 | Not_Computer
Z-2373-2015 | 71724 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Naturalyte Liquid Bicarbonate Concentrate 6.4 liter bottle |  |  | July 16, 2015 | 2015 | Fresenius Medical Care Holdings, Inc. | Expansion of recall to lots manufactured in 2015 due to the potential for bacterial contamination. | Process control | Fresenius Medical NA issued an Urgent Expanded Medical Device Recall letter on 7/16/15 to clinics and patients. The letter identified the affected product, explained the reason for the expansion, and instructed users to immediately examine their stock to determine whether there was any Naturalyte Liquid Bicarbonate Concentrate beginning with 15AMLB, 15BMLB, 15CMLB.Users are to discontinue use immediately if any product of these lots is found, .All units are to be placed in a secure, segregated area.If affected product was on the machine prior to patient treatment, a [Heat Disinfect] program should be performed.The attached Reply Form should be completed and returned.Contact your FMCNA Customer Service Team at 1-800-323-5188 for instructions on how to return the recalled product. |  | 2264028 | 2,264,028 | Nationwide Distribution | Open, Classified |  |  | 2264028 | Not_Computer
Z-2320-2015 | 69849 | No | RER |  |  |  |  | GE Optional Laser Centering Device X-ray generator. |  |  | November 10, 2014 | 2014 | GE Inspection Technologies, LP | GE has identified a potential failure to comply with 21 C.F.R. Âż 1040.10 and 21 C.F.R. Âż 1010.2 for the alignment laser centering device. | Radiation Control for Health and Safety Act | GE's Planned Action for repair: 1. Replacement compliant laser centering devices will be shipped to purchasers, with prepaid shipping envelopes for purchasers to return the noncompliant units to GE.2.All mailings will be made at no cost to the purchaser; and3.The corrections will be completed by November 1, 2015Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and1003.22. This office and the Food and Drug Administration (FDA) district office coordinator are to be included in the notification. For further questions please call (717) 447-1278. |  | 25 | 25 | US Distribution to the states of : NY, CA, WA, PA, MD, AL, FL and LA. | Open, Classified |  |  | 25 | Not_Computer
Z-2376-2015 | 71658 | No | RCE |  |  |  |  | PSW 1203.U/1, PSW 1301.U, PSW 1307.U, PSW 1403.U, and PSW 1501.U Pacemaker Programmer software. Software approved for use with BIOTRONIK device programmers (trade names: Renamic and ICS 3000). |  |  | June 30, 2015 | 2015 | BIOTRONIK, Inc. | Ventricular packing: LV software programming versions for BIOTRONIK CRT-P and CRT-D devices are recalled because they are not approved for use in the US. | Software Manufacturing/Software Deployment | BIOTRONIK sent the Important Product Information letter and the Acknowledgement of Important Product Information Form, dated June 20, 2015, to affected physicians with the following recommendations: -Review programming settings for your patients with CRT-P or CRT-D devices.-If you have a patient with a device programmed to Ventricular pacing: LV, arrange to have the patients device reprogrammed to an appropriate approved mode (i.e. "Ventricular pacing: RV" or "Ventricular pacing: BiV").-Physicians are requested to sign the Acknowledgement of Important Product Information Form and immediately fax to 800-913-6993 or email to UScompliance@biotronik.com.BIOTRONIK'S representative will update your programmer with approved 1502.U programmer software. Following this update, the Ventricular pacing: LV feature will no longer be available. However, a device previously programmed with"Ventricular pacing: LV will remain programmed in this mode unless and until the pacing mode setting is changed. If you have any questions concerning this Important Product Information, please contact the BIOTRONIK Technical Services Department:Telephone: 1-(800) 284-6689 (24-hour service) |  | 768 | 768 units (software) | Nationwide Distribution. | Terminated | December 24, 2015 | 177 | 768 | Not_Computer
Z-2375-2015 | 71642 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | Wall Chart for STERRAD 100NX System, P/N 10104. |  |  | July 17, 2015 | 2015 | Advanced Sterilization Products | Advanced Sterilization PRoducts (ASP) is recalling the Wall Chart for STERRAD 100NX System because it incorrectly states "APTIMAX Instrument Trays can be enclosed in either polypropylene wrap or in Tyvek Pouches with STERRAD Chemical Indicators." | Under Investigation by firm | A Urgent recall letter will be sent to customers to inform them that Advanced Sterilization Products (ASP) is recalling the Wall Chart of the STERRAD 100NX System because it incorrectly states "APTIMAX Instrument Trays can be enclosed in either polypropylene wrap or in Tyvek Pouches with STERRAD Chemical Indicators. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete a business reply card for their acknowledgement of the customer letter and documentation of the removal and replacement of the wall chart. Customers are instructed to complete and return the attached Business Reply Card via mail or fax to Stericycle: 877-497-3125. Customers with questions are instructed to contact SteriCycle at 866-737-5389 and reference event #6677. |  | 1465 | 1465 units | Nationwide Distribution. | Open, Classified |  |  | 1465 | Not_Computer
Z-2381-2015 | 71707 | adaptor, hygiene | ILS | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Etac Relax wall mounted shower seat |  |  | July 12, 2015 | 2015 | Etac Supply Center Ab | The recalling firm has received reports of a malfunction of the Relax wall mounted shower seat where the wall bracket has cracked or broken. | Process design | Etac sent an Urgent Field Safety Notice dated July 2015 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that there was no reason for an immediate stop of using Relax, however, as an intermediate precautionary action awaiting replacement, consumers were requested to be observant and inspect the Relax to see if there are visible signs of a crack. Customers with questions were advised to contact their Etac distributor or the Etac coordinator at 1-800-336-7684 or fax 1-704-882-0751. |  | 13 | 13 units | Nationwide Distribution including FL, AZ, and CA. | Open, Classified |  |  | 13 | Not_Computer
Z-2382-2015 | 71666 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Celsite Implantable Access Port System, model ST301. The Celsite Implantable Access Port Systems (Celsite port systems) are implantable port and catheter systems that allow safe, repeated access to the patient's bloodstream. The port chamber and catheter design can be used for the administration of medication and fluids. The Celsite system consists of an access port with a silicone septum, which is connected to a catheter using a connection ring. The triangular shaped access port has a low profile nose, finger stops on the side of the housing, and a round base. Celsite access ports have suture holes or suture zones to secure placement during implantation. |  |  | June 16, 2015 | 2015 | B. Braun Interventional Systems | The manufacturer, B. Braun medical France, received endotoxin test results that are out of specification for the peelable sheath (A1537). | Unknown/Undetermined by firm | An Urgent Medical Device Recall letter, dated 6/16/2015 was sent to the 2 consignees via express mail. The letter explained the issue and requested the hospital review their inventory for the affected model and lot. If any quantities of the lot remained in inventory, the product was to be returned to BIS. A BIS sales representative will personally visit each account and complete an inventory sheet. Customers with questions can contact Paul O-Connell, President at 1-847-274-0097 |  | 11 | 11 units | CA and NY only. | Terminated | November 20, 2015 | 157 | 11 | Not_Computer
Z-2380-2015 | 71849 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioJet AVX Thrombectomy Set: UPN 105039-001, Catalog No.105039, Sterilized Using Ethylene Oxide. For breaking apart and removing thrombus from vascular system, used with the AngioJet Ultra System in breaking apart and removing thrombus from A-V access conduits. |  |  | July 31, 2015 | 2015 | Boston Scientific Corporation | AngioJet products were distributed with an incorrect use-by (expiry) date. The correct date is Nov-2015, but the units are labeled with a use-by date of Dec-2015. | Process control | Boston Scientific sent an Urgent Medical Device Recall Removal letter dated July 31, 2015, to all affected consignees. The letter described the problem and the affected product. Consignees were advised to segregate the affected product and return it to Boston Scientific. They were also requested consignees to complete and return the Account Reply Verification Tracking Form. Contact your Sales Representative for any questions.For questions regarding this recall call 763-494-1556. |  | 9 | total of 9 units | Worldwide Distribution - US including AZ, MO, OH and Internationally to GERMANY and SPAIN. | Terminated | February 05, 2016 | 189 | 9 | Not_Computer
Z-2385-2015 | 71722 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Magic3 Hydrophilic-coated Female Intermittent Urinary Catheters (6"), 12 Fr, (30/box), Product Code 51612. |  |  | July 07, 2015 | 2015 | C.R. Bard, Inc. | Potential breach of the sterile barrier packaging. | Unknown/Undetermined by firm | A recall notice was sent to consignees on 7/7/2015 asking them to discontinue use of the product. The letter to the distributors requested a sub-recall. Both letters included a response form to be returned to Bard. |  | 91140 | 91,140 each | Nationwide Distribution-including AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NV, NH, NJ, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, WI AND DC. | Terminated | July 21, 2017 | 745 | 694907 | Not_Computer
Z-2390-2015 | 71714 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | RadPro 40kW Digital X-Ray Panel and Mobile/Portable Diagnostic X-Ray. |  |  | July 13, 2015 | 2015 | Sedecal S.A. | Electrostatic energy may be stored in Systems that can cause a short-cicuit which, in turn, can generate an Electomagnetic Pulse which might disturb the electronics of the Digital Motion Control Board producing uncontrolled movements and such uncontrolled movements could cause the system to bump into the user or ppatient. | Device Design | FIELD CORRECTION NOTIFICATION letters dated July 13, 2015 were sent to all dealers. The letters instructed the dealers: to schedule the replacement of the current arrestor by a new design of resistive arrestor with a resistance of 1 M ohm between frame and floor (ground). Although the metallic arrestor is only implemented in the equipments manufactured from April 2013 till March 2015, the upgrade kit has to be implemented in all the installed base of Mobiles with the same configuration in order to ensure that even the arrestors installed as spare part are replaced. Sedecal will provide the necessary upgrade kits to all the customers or distributors. *** U.S. Distributor, Virtual Imaging (a Canon U.S.A. Company) will use letter entitled, FIELD CORRECTION NOTIFICATION, RADPRO 40kW - -DIGITAL MOBILE X-RAY SYSTEM and dated 7/13/2015 to communicate to end users. The letter states "Once Virtual Imaging receives the resistivearrestor, we will contact you to schedule the replacement on the Affected System that you service."Please contact Virtual Imaging, Inc.'s Technical Support for any questions at 561.893.8500 or technicalsupport@vifla.com. **** |  | 361 | 361 systems (U.S.); 5 systems (Foreign) | Distributed US (nationwide) and in Canada. | Open, Classified |  |  | 361 | Not_Computer
Z-2362-2015 | 71655 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Syringe Pump, Model No. 8110. Delivers fluids. |  |  | July 21, 2015 | 2015 | CareFusion 303, Inc. | Channel Error code is displayed on the PC unit with an audio and visual alarm, and on the syringe module. After the error is cleared on the PCU, the syringe pump is unresponsive to key presses until the next power cycle, or the module is detached and reattached. | Process control | An Urgent Medical Device Recall Notification letter was sent on 7/20/15 to customers who purchased the Alaris Syringe Module (Model 8110) to inform them that CareFusion has identified an issue with the Alaris Syringe module manufactured between March 2014 through September 2014. The letter informs customers of the problems identified and the actions to be taken. CareFusion will replace the syringe drive train assembly on the affected units and adjust, if required. CareFusion will contact all affected customers within 60 days to initiate the scheduling process for the remediation. Customers are instructed to contact CareFusion Support Center (888-562-6018, 7am-4pm PT, M-F, supportcenter@carefusion.com) for recall related questions. Customers are instructed to contact Customer Advocacy (888-812-3266, 24hrs/day, 7days/wk, customerfeedback@carefusion.com) for adverse event reports. Customer are instructed to contact Technical Support (888-812-3229, 6am-5pm PT, M-F, dl-us-inf-techsupport@carefusion.com) for technical questions. Customers are instructed to complete and return the enclosed customer response card to expedite the corrective action process.A urgent recall letter for the expansion of the affected units was sent on 10/15/15 to inform customers of the expanded recall. The letter informs customers that the firm has identified an issue with the Alaris Syringe module model 8110 after receiving reports of customers experiencing an channel error on the Alaris Syringe module Model 8110. The letter informs customers that the recall is an expansion of affected units outlined in the CareFusion recall previously issued on 7/20/15. The letter provides the customers with the affected units, issue, potential risk, and actions to be taken.The corrective action for the expanded serial numbers will not change. The affected Alaris Syringe modules will be remediated with a hardware inspection and an adjustment, if required. CareFusion will contact all affected c |  | 68460 | 68,460 units | Worldwide Distribution-US (nationwide) and the countries of Australia, New Zealand, South Africa, Canada, Middle East, Malaysia, and Taiwan. | Open, Classified |  |  | 68460 | Not_Computer
Z-2206-2015 | 71696 | No | RCG |  |  |  |  | American Science &amp; Engineering Inc,Gemini Cabinet X-ray system. |  |  | May 22, 2015 | 2015 | American Science &amp; Engineering, Inc. | During factory testing, we discovered that Gemini systems could, in rare circumstances, experience a fault that leads to a state in which X-rays are being produced but the X-ray indicator lights (located on the operator's console and on both ends of the tunnel) are not illuminated. | Vendor change control | American Science Engineering, Inc (AS&E;) is developing a fix for this problem that includes updating a software configuration on the X-ray source. You will be notified by field service as soon as this update is available.Please inform all operators of the system about this communication, so they are aware of the need to shut down and re-start the system should this fault occur.We appreciate your understanding. if you have any questions, please feel free to contact AS&E; via e-mail at service@as-e.com. If you are under a service contract, you can also contact your customer service representative. For further questions please call (978) 262-8700. |  | 125 | 125 systems | Worldwide Distribution - US Distribution to the states of : NY, FL, TN, PA, NJ, DE, ME, TX, CT, MS, NC, MD, MO, VA, DC, NV and GA., and to the countries of : Saudi Arabia, Indonesia, Canada, Netherlands, Hong Kong, Lebanon and Kazakhstan. | Open, Classified |  |  | 125 | Not_Computer
Z-2343-2015 | 71702 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | PS500 of an IACS Workstation Critical Care (Evita Infinity V500); intended for the ventilation of adult, pediatric and neonatal patients. |  |  | July 13, 2015 | 2015 | Draeger Medical, Inc. | The battery capacity of the optional PS500 of an IACS Workstation Critical Care (Evita Infinity V500) did not last as long as expected. | Nonconforming Material/Component | Draeger Medical, Inc. sent an Urgent Medical Device Recall Letter, dated July 2015, to end users to inform them of the issue and provide interim instructions to avoid the associated risk to patients. All PS500 batteries will be exchanged free of charge. |  | 2081 | 2081 | Nationwide Distribution. | Terminated | March 01, 2017 | 597 | 2081 | Computer
Z-2329-2015 | 71675 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 980 Ventilator System, Universal (with neonatal functionality enabled) and Neonatal Models. |  |  | July 16, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Reports in which tidal volumes reaching patients were lower than set tidal volumes in neonatal Volume Control Plus (VC+) Mode with active humidification. This situation may potentially lead to respiratory compromise if not recognized. | Software design | Consignees will be notified via letter on July 17, 2015. No products will be returned, the reconfiguration of affected ventilators will occur at customer sites. |  | 657 | 657 units | Worldwide Distribution-US (nationwide) and the countries of Canada, United Arab Emirates, Argentina, Australia, Bahrain, Brazil, Chile, Colombia, Costa Rica, Egypt, Ireland, Israel, Jordan, Japan, Kuwait, Lebanon, Morocco, Mexico, Panama, Peru, Qatar, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Turkey, and Uruguay. | Open, Classified |  |  | 657 | Not_Computer
Z-2402-2015 | 71496 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Winco Inverness 24-Hour Treatment Recliner - with Nylon Casters |  |  | April 02, 2015 | 2015 | Winco Mfg., LLC | Shipment of foam for chairs was received classified as Fire Retardant but was in fact standard foam and would not meet Fire Retardant standards. | Nonconforming Material/Component | Winco sent a field correction letter dated April 2, 2015, to all affected customers The letter identified the product, the problem, and the action to be taken by the customer. Winco instructed customers that they will replace the recliners at no charge to the customer. Winco will provide detailed instructions for the replacement of the seats by their facility staff. Winco will extend the remaining warranty on the entire chair for a period of 6 months as compensation. If product was further distributed customers were instructed to notify those customers of the field correction. Customers were asked to complete and return the enclosed field correction response form as soon as possible. Customers with questions should contact Customer Care Monday-Friday 8:30 AM-5:00 PM EST at 800-237-3377 or email QA-RA@wincomfg.com.For questions regarding this recall call 352-854-2929. |  | 29 | 29 chairs | Nationwide Distribution including CA, FL, KS, KY, LA, MA, MD, MN, MS, NJ, NY, OK, TN, TX, and WA. | Open, Classified |  |  | 133 | Not_Computer
Z-2389-2015 | 71678 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Triathlon X3 Tibial Bearing Insert - CR |  |  | May 29, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker has received a report which identified that a Triathlon Tibial Bearing Insert CS X3 Size #2 9mm (5531-G-209) was reported to be in a pack that was labelled as a CR insert (5530-G-209). | Labeling mix-ups | Stryker Orthopaedics sent a recall notification via email on 5/29/2015 to their local branches. Notification letters and Product Accountability Forms were sent to Stryker branches/agencies with return receipts on 7/2/2015. |  | 10 | 10 units | OH, SC, RI, WV | Open, Classified |  |  | 10 | Not_Computer
Z-2405-2015 | 70741 | Yes | MPB | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RenLane Renal Denervation Catheter; Cat No.: D135601. |  |  | April 15, 2014 | 2014 | Cordis Corporation | Potential Damage may occur to the helical tip of the RENLANE Renal Denervation Catheter, if the hemostasis valve of the arterial access device is not fully open before insertion or withdrawal of the catheter. | Labeling Change Control | Cordis sent an Urgent Field Safety Notice dated April 15, 2014 providing additional labeling to their consignees in Germany only, with a follow-up letter on August 4, 2014. Cordis explained When deploying the catheter via a compatible arterial access device, if a Tuohy-Borst type adjustable hemostasis valve is used, ensure the valve is fully open before insertion or withdrawal of the catheter. This is especially important, since the helical tip at the distal end of the RENLANE Renal Denervation Catheter has a larger diameter compared to the proximal catheter shaft. Although the event was not a Removal, all lots involved have expired. The product is no longer commercialized, and was never commercialized beyond the 10 customers in Germany. |  | 98 | 98 units | Distributed in Germany only. | Terminated | August 26, 2015 | 498 | 98 | Not_Computer
Z-2407-2015 | 71727 | Yes | BTS | Bronchial tube. | Anesthesiology | Anesthesiology | 510(k) | Bronchial Double Lumen Tube Set (Left), Sterile |  |  | November 21, 2014 | 2014 | Teleflex, Inc. | Customer complaints reporting that the cobb connector detached from the main connector prior to use. | Other | Recall letters were sent on 11/21/2014 to consignees requesting them to discontinue use and quarantine the products. The letter includes a reply form to be returned. The letter to the distributors requested a sub-recall, and included a recall letter to be sent to their accounts. |  | 143750 | 6,991 ea. - US &amp; 136,759 ea. - INTL | AL, AK, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MD, MA, MI, MS, MO, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV. WI, DC, Puerto Rico, Canada, Bolivia, Ecuador, Hong Kong, Malaysia, Philippines, Taiwan, Trinidad &amp; Tobago | Terminated | August 15, 2017 | 998 | 143750 | Not_Computer
Z-2404-2015 | 71762 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/ zeego systems; dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. |  |  | June 30, 2015 | 2015 | Siemens Medical Solutions USA, Inc | in case a system error occurs and the system enters the "Bypass Fluoro" mode while the X-ray locking function is active, the only way to exit the X-ray locking function, would be to either resolve the root cause of the system being in "Bypass Fluoro" or to restart the system. | Software design | A customer advisory notice, dated June 30, 2015, was sent to consignees informing them of the corrective action that would be performed to prevent the possible loss of the image generation function. Siemens plans to contact end users to install a software patch. For questions regarding this recall call 610-219-6300. |  | 586 | 586 | Nationwide Distribution | Open, Classified |  |  | 586 | Not_Computer
Z-2406-2015 | 71673 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | Cell Search Circulating Tumor Kit. Intended for the enumeration of circulating tumor cells (CTC) of epithelial origin in whole blood. |  |  | May 15, 2015 | 2015 | Veridex, LLC | Complaints of the presence of artifacts that appears as small bead like structures in the image gallery when processing both control cells and patient samples as well as the failure of high control cells. | Under Investigation by firm | Janssen Diagnostics, LLC sent a urgent product correction notification/customer confirmation of receipt on 5/15/2015 via Fed Ex. |  | 1560 | 1560 kits | Nationwide Distribution. | Terminated | March 30, 2016 | 320 | 1560 | Not_Computer
Z-2422-2015 | 71911 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | IBC Suction Wand Catalog Number 1990S |  |  | August 07, 2015 | 2015 | International Biophysics Corp. | It was discovered that the heat seal placed onto sterile pouches may not meet strength acceptance criteria. | Packaging process control | The recalling firm began notifying their affacted customers on 7/2/2015 |  | 100 | 100 units | Worldwide - United States Nationwide and the countries of: Germany, Norway, Turkey, Chile, Pakistan, and Republic of South Africa | Terminated | March 01, 2016 | 207 | 11201 | Not_Computer
Z-2416-2015 | 71710 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | ADVIA Chemistry Acetaminophen reagents; for in vitro diagnostic use in the quantitative determination of acetaminophen in human serum and plasma (lithium heparin) on ADVIA Chemistry systems. |  |  | June 30, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | A change in the concentration of N-acetylcysteine (NAC) that may cause interference for the ADVIAÂż Chemistry Systems Acetaminophen assay for certain lots. | Unknown/Undetermined by firm | Siemens sent an Urgent Medical Device Correction letter dated June 30, 2015, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer. Siemens Healthcare Diagnostics is conducting a voluntary recall for the ADVIA Chemistry Acetaminophen reagent, for all in date and future kit lots. Customers were informed of the revised interference levels of N-acetyl cysteine (NAC) of 200 mg/dL, for the ADVIA Chemistry Acetaminophen assay.Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within 30 days.Field service personnel will be sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. For further questions call 1 (800) 441-9250. |  | 1032 | 1032 | Worldwide Distribution - US (nationwide) and to the countries of : Canada, Netherlands, Great Britain, Norway, Sweden, Spain, Italy, Germany, France, Finland, Belgium, Portugal, Czech Republic, Israel, Poland, Turkey and Bahrain. | Open, Classified |  |  | 1032 | Not_Computer
Z-2421-2015 | 71783 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 2.0/2.4 Cannulated Screw Countersink Part # 210-24-002 NON-STERILE and 3.0/4.0 Cannulated Screw Countersink Part # 210-40-002 NON-STERILE |  |  | July 20, 2015 | 2015 | Trilliant Surgical Ltd. | The cutting efficiency of the Countersinks was out of specification. | Nonconforming Material/Component | Trilliant sent an Important Medical Device Advisory Notice letter dated February 15, 2013 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers of the option to exchange affected instruments at no charge. To initiate the exchange customers were instructed to contact your local Trilliant Surgical Sales Representative, contact the Customer Service Department directly at 1-800-495-2919, or via email at orders@trilliantsurgical.com. |  | 102 | 102 units | US Nationwide Distribution in the states of: NY, CT, TN, TX, OH, IN, LA, NC, MO, MN, GA, FL, SC, WA, WI, IA, PA, CA, ND, OK, NJ, MI and MD. | Open, Classified |  |  | 102 | Not_Computer
Z-2418-2015 | 71869 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Sciatic Nerve Retractor and Sciatic Nerve Retractor Long; Synthes Sciatic Nerve Retractor and Sciatic Nerve Retractor Long are system specific instruments to the 3.5 mm Low Profile Pelvic System Retractor Set. |  |  | July 29, 2015 | 2015 | Synthes (USA) Products LLC | Potential for micropores to form, allowing fluid to enter the hollow handle. | Process control | Urgent recall letters, dated July 29,2015, were sent to end users and sales consultants to inform them of the issue, potential risks, and to provide instructions for returning affected devices and the response form. |  | 1113 | 1113 | Nationwide and internationally, including Canada | Open, Classified |  |  | 1113 | Not_Computer
Z-2417-2015 | 71739 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | Radiometer Fixation Kit N20 for the TCM CombiM monitoring system, REF 905-873. The Fixation Kit N20 is used to apply tc Sensor 84 and tc Sensor 54 on the skin of a patient. TC sensor 84 and 54 are used with the TCM CombiM monitoring system and the Philips TcG10 module. The TCM CombiM monitoring system is intended for continuous transcutaneous monitoring of carbon dioxide (tcpCO2) and oxygen (tcpO2) partial pressures. |  |  | July 16, 2015 | 2015 | Radiometer America Inc | When the fixation ring is taken off the paper liner there is a potential for fractions of the paper liner's surface coating to stick to the surface of the ring. | Under Investigation by firm | Radiometer sent a Product Notification letter dated July 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.In the letter the firm requests that inventory be checked and removed for any the identified lot numbers for Fixation kit N20. The firm requests that the products be discarded. The letter states to complete p.2 and return it to their Radiometer representative. If you have any questions contact Radiometer America Technical Support at 1-800-736-0600 opt.4. |  | 4244 | 4244 kits | Worldwide Distribution - US Distribution and to the countries of : Canada, Spain, France, Great Britain, Philippines, Netherlands, Japan, China, Australia, Denmark, South Africa, Germany, Qatar, Turkey, Hong Kong, Costa Rica, Austria, Finland, Indonesia, Norway, Malaysia, Belgium, Czech Republic, Hungry, United Arab Emirates, Poland, Colombia, Belgium and Malta. | Terminated | January 29, 2016 | 197 | 4244 | Not_Computer
Z-2420-2015 | 71897 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic EOPA Arterial Cannula, Model # 77620. Lot: 2015060544 intended to deliver oxygenated blood to the heart during cardiopulmonary bypass procedures for up to 6 hours in duration. |  |  | July 24, 2015 | 2015 | Medtronic Perfusion Systems | EOPA Arterial Cannula Devices in this lot were shipped without the guidewire. | Packaging process control | On July 24, 2015 Medtronic sales representatives began verbally notifying the customers of the issue, requesting quarantine and return of un-used, affected productConsignees were sent a Medtronic "Urgent Medical device Recall" letter dated August 2015. The letter was addressed to Risk Manger and Health Care Professional. The letter described the problem and the product involved in the recall. Advised consignees to quarantine and remove the product from their inventory and return it to Medtronic. Requested consignees to complete and return the Customer Confirmation Certificate via email to RS.CFQFCA@Medtronic com or fax to Medtronic 651-367-0612 to the attention of Customer Focused Quality. For questions contact your Medtronic Sales Representative or Lifeline Technical Services at 1-877-526-7890. |  | 221 | 221 (111 US, 110 OUS) | Worldwide distribution. US nationwide including: FL, GA, IL, MI, NY, NC, PA, WV; Germany, Hong Kong, and Italy. | Terminated | April 04, 2016 | 255 | 221 | Not_Computer
Z-2425-2015 | 71713 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Tibial Nail Wrap Holds 11 Nails and Synthes Intramedullary Nail Wrap Holds 10 Nails; the Nail Wrap is intended to be used as a protective sheath specifically for Universal Femoral Nails, Universal Tibial Nails, Unreamed Tibial Nail, and the Stainless Steel Tibial Nail during steam sterilization by the hospital. |  |  | July 01, 2015 | 2015 | Synthes (USA) Products LLC | The wrap is for single use for nail sterilization but does not have sufficient testing for the device to be considered a multi-use item. | Under Investigation by firm | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated July 8, 2015, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer.ACTIONS REQUIREDWe have on record that your facility has received the product(s) subject to this recall. DePuy Synthes asks that you review your inventory and immediately remove the affected lots from your stock.Please take the following actions:If you DO have any of the identified devices, please take the following steps:Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number.Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note theReturn Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided.Return the Verification Section (page 3 of this letter) with the product to Credit/Returns, DePuy Synthes,Send a copy of the completed Verification Section to DePuy Synthes by: Fax: 888-670-4162 or Scan/email: Synthes5485@stericycle.comNote: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information.Return the documents to DePuy Synthes by: Fax: 888-670-4162 orScan/email: Synthes5485@stericycle.comNote: If the Verification Section is answered on |  | 1487 | 1487 | Worldwide Distribution - US Distribution including Puerto Rico and to the countries of : Canada, Switzerland and New Zealand. | Open, Classified |  |  | 1487 | Not_Computer
Z-2433-2015 | 71720 | No | LWK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Medline Electrical Clipper with Charging base |  |  | July 08, 2015 | 2015 | Medline Industries Inc | Charging base of surgical clippers overheats and smokes. | Process control | Medline Industries sent a " Urgent Remedial Action Destruction Form " letter via first class dated July 9, 2015, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customer.***URGENT RECALL IMMEDIATE ACTION REQUIRED SURGICAL CLIPPER CHARGER BASE DYND70800 ***Medline Industries, Inc. is recalling specific lots of the Medline Surgical Clippers Charger Base (base only). This recall is being conducted due to a potential defective component on the circuit board causing the charger base to overheat which could result in a charger base malfunction. These lots were distributed October 2013 through March 2015. After further investigation we have determined the root cause to be a defective circuit board. This defect has since been corrected. REQUIRED ACTION: 1. Immediately check your stock for item number DYND70800 and the affected lots listed on the destruction form, dispose of the affected charger base only (do not dispose of the clipper). 2. Please return the completed enclosed destruction form listing the quantity of affected charger base(s) you destroyed. Even if you do not have any affected product please complete and return the form, as indicated on the destruction form. When we receive your completed destruction form we will send you replacement charger base(s). 3. If you are a distributor, promptly notify any of your customers that may have received affected product. If you have any questions, please contact 866-359-1704. As the distributor of this product, we sincerely regret the inconvenience. We, like you, place the health and safety of your patients first and foremost. This recall is being made with the knowledge of the U.S. Food and Drug Administration. |  | 7062 | 7062 units (U.S.); 540 units Foreign | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico and to the countries of : United Arab Emirates, Barbados, Bahamas, Chile, Costa Rica, Italy, Lebanon, and Singapore. | Open, Classified |  |  | 7062 | Not_Computer
Z-2431-2015 | 71841 | No | LSL | Neisseria spp. direct serological test reagents. | Microbiology | Microbiology | 510(k) | Xpert CT /NG Vaginal/Endocervical/ Specimen Collection Kit: |  |  | July 27, 2015 | 2015 | Cepheid | Firm determined that some lots of CT/NG Swab 50 and CT/NG Urine collection kits contain a sub component with an expiry date that is earlier than the kit expiration date. | Incorrect or no expiration date | Notification letters dated July 28 and 29th to all affected customers via Federal Express.. |  | 776 | 729 kits US, 47 kits ROW. | Worldwide distribution. US nationwide, including North Mariana Islands, Aruba, Canada, Chile, Colombia, Costa Rica, Cyprus, Guam, Guatemala, Italy, Malawi, Micronesia Federated States, Panama, San Marino, South Africa, Spain, Switzerland, and Zimbabwe. | Terminated | September 28, 2015 | 63 | 2607 | Not_Computer
Z-2434-2015 | 71633 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer SST Plus Blood Collection Tubes, 5.0 mL, 13 x 100 mm, Sterile. |  |  | June 23, 2015 | 2015 | Becton Dickinson &amp; Company | Some of the tubes were manufactured with stoppers that did not meet current manufacturing specifications. Stoppers were manufactured with a reduced cycle cure time. | Process control | Becton Dickinson notified their affected customers of this recall by sending them an Urgent Product Recall Letter dated June 23, 2015 via UPS 2nd Day. |  | 238 | 238, 000 (Lot #5133959) and 243,000 (Lot #5133957) | Distributed in FL, IL, MA, MI, MO, PA and TN. | Terminated | April 06, 2016 | 288 | 238 | Not_Computer
Z-2371-2015 | 71543 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Medley LVP Frame Membrane. |  |  | June 03, 2015 | 2015 | Elite Biomedical Solutions LLC | Administration of inappropriate quantities of fluid can result, with the potential to cause injury or death. | Other | Elite Biomedical Solutions, sent a "Product Advisory Notices" letter dated May 21, 2015 to their customers. On June 3, 2015 the firm sent an Urgent Medical Device Recall letter to their customers. The letter identified the affected product , problem and actions to be taken. And on June 12, 2015 a press release was issued via ECRI (Emergency Care Research Institute) to all hospitals in the US. The notification letter requested customers to: 1) inspect and quarantine affected products, 2) identify your customers and notify them at once of this product recall, and 3) complete and return the enclosed response forms. For any questions, call Elite Biomedical Solutions, LLC, Quality Manager, at 1-855-291-6701. |  | 609 | 609 pieces | Nationwide Distribution in the states of AZ, CA, GA, IL, IN, KS, KY, MA, MD, MN, MO, MS, NC, NY, OH, PA, SC, TN, TX, VA &amp; WI. | Terminated | March 07, 2017 | 643 | 955 | Not_Computer
Z-2370-2015 | 71360 | Yes | CBI | Tracheal/bronchial differential ventilation tube. | Anesthesiology | Anesthesiology | 510(k) | HUDSON RCI, Sheridan, SHER-I-SWIV and SHER-I-SWIV/FO Double Swivel, Rx Only, Sterile. |  |  | June 01, 2015 | 2015 | Teleflex Medical | The double swivel connector may crack or separate on the endobronchial tube. | Unknown/Undetermined by firm | Teleflex Medical sent an Urgent Medical Device Recall Notification letter dated June 1, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were directed to contact any customers who had received the product included within the scope of this recall. Customers were directed to immediately discontinue use and quarantine any products with the referenced catalog numbers and lot numbers listed in the notification. Additionally, customers were asked to return any affected product together with a completed Recall Acknowledgement Form, for consolidation and return to Teleflex Medical. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 188195 | 188,195 total units | Worldwide Distribution - US (nationwide) and the countries of: The Bahamas, Belgium, Argentina, Canada, China, India, Japan, South Korea, Mexico, New Zealand and Vietnam. | Open, Classified |  |  | 188195 | Not_Computer
Z-2437-2015 | 71867 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens Intevo Series system, a radiological imaging system that is a combination of a single photon emission computed tomography (SPECT) camera system for nuclear medicine images, and a computed tomography (CT) camera system for x-ray images. |  |  | July 31, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Fuses with the improper amperage rating were installed in the Symbia S, Symbia T Series and Symbia lntevo systems. These fuses, in conjunction with a specific failure mode of the Power Controller assembly, could lead to the overheating of the assembly. | Component change control | Siemens has sent an undated Customer Advisory Notice letter to all direct accounts (customers). The letters include instructions for customers to: 1) Continue to use the system(s) while waiting for the field correction to be scheduled and performed unless the user observes a burning smell or sees evidenve of smoke. the user should engage the emergency stop, power down and discontinue use of the system, and contact the local service representative. 2) Ensure that this advisory is placed in the system's instructions for use. Siemens reminds the user to observe the patient during studies. 3) If the product was sold to another party, forward the letter to the party. A local service engineer will contact the customer to schedule the servicing of the affected units by September 2015 which will consist of replacing the affected fuse with the correct amperage rating fuse. If you have any questions regarding the important notice, please contact your local service representative by calling 1-800-888-7436. Europe, Middle East and Africa: +49 9131 940 4000. Asia and Australia: +86 (21) 3811 2121. |  | 185 | 185 | Worldwide Distribution - USA including AL, AK, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico and Internationally to Algeria, Argentina , Australia , Austria , Bangladesh , Belgium , Brazil , Bulgaria , Canada , Chile , China, Colombia , Costa Rica , Croatia , Czech Republic , Denmark , Ecuador , Egypt , Estonia, Finland , France , Germany , Hong Kong , Hungary , Iceland , India , Indonesia , Iran , Ireland , Israel , Italy , Japan , Jordan , Kazakhstan , Korea, Kuwait , Latvia , Libya , Lithuania , Luxembourg , Malaysia , Mexico , Morocco , Myanmar , Netherlands , New Zealand , Norway , Oman , Pakistan , Peru , Philippines , Poland , Portugal , Qatar , Romania , Russia, Saudi Arabia , Singapore , Slovakia , Slovenia , South Africa , Spain , Sweden , Switzerland , Taiwan , Tanzania, Thailand , Tunisia , Turkey , United Arab Emirates , Ukraine , United Kingdom , Venezuela , Vietnam | Open, Classified |  |  | 3068 | Not_Computer
Z-2427-2015 | 71676 | No | KXE | Wrist joint ulnar (hemi-wrist) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | uHead, Ulnar Implant, Standard Stem, Size 1, Sterile, Rx Only. Intended for replacement of the distal radioulnar joint following ulnar head resection arthroplasty. |  |  | June 24, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker is recalling rHead, uHead, Sigmoid Notch, Remotion, Radio Capitellum because packaging integrity (sterile barrier) of the packaging type KIT I may potentially be compromised by transportation. | Packaging change control | Branches/Agencies were notified by Stryker of this action by e-mails on June 24, 2015 and an Urgent Medical Device Removal Letter/Acknowledgement Response Form dated June 25, 2015 was sent to the attention of the Risk Manager. |  | 16992 | 16992 total | Worldwide Distribution. | Open, Classified |  |  | 16992 | Not_Computer
Z-2441-2015 | 71780 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | iChemVELOCITY Automated Urine Chemistry Catalog No. 800-3049, 800-3050, 800-3079, 800-3080, 800-3530, 800-7167, 800-7190, 800-7713,800-7714, 800-7719, 800-7720. An in vitro diagnostic device used to automate the urine chemistry analysis profile using iChemVELOCITY Urine Chemistry Strips. |  |  | July 15, 2015 | 2015 | Iris Diagnostics | Iris International is recalling the iChemVELOCITY systems because they do not maintain the on-board strip stability claim of 5 days at 18-28 degrees celsius at 20-80% relative humidity (RH) as stated in the iChemVELOCITY product labeling. | Nonconforming Material/Component | An Urgent Medical Device Recall letter dated 7/15/15 was sent to all customers who purchased the iChemVELOCITY Automated Urine Chemistry System using iChemVELOCITY Urine Chemistry Strips REF 800-7212 and 800-7212-001. As a result of internal investigations following customer complaints, Iris International has determined that the iChemVELOCITY systems do not maintain the on-board strip stability claim of 5 days at 18-28ÂżC at 20-80% relative humidity (RH) as stated in the iChemVELOCITY product labeling. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions or concerns regarding the recall notice, are instructed to contact Beckman Coulter Customer Support at http://www.beckmancoulter.com or call (800) 526-7694 in US and Canada. |  | 2448 | 2,448 units | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Bahrain, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Croatia, Czech Republic, Ecuador, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Jordan, Republic of Korea, Kuwait, Lebanon, Luxembourg, Macao, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Open, Classified |  |  | 2448 | Not_Computer
Z-2448-2015 | 71941 | No | HRM | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LX3 Floor Stand for Endure Luxite Envision Optical System; Endure LuxOR Q-Vue. |  |  | August 11, 2015 | 2015 | Alcon Research, Ltd. | Alcon is initiating a voluntary recall of all LX3 Microscope Floor Stands after receiving reports concerning loss of functionality due to issues with the cabling between the power, host and illumination modules, potentially causing the system to display a system message, and/or resulting in illumination failure. | Device Design | Alcon sent an Urgent Medical Device Correction letter dated August 11, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact Alcon at 800-862-5266. |  | 688 | 688 units | Worldwide Distribution - US Nationwide and the countries Argentina, Australia, Austria, Belarus, Belgium, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Dominican Republic, France, Germany, Greece, Guatemala, Hungary, India, Italy, Japan, Jordan, Korea, Lebanon, Malaysia, Mexico, Mongolia, New Zealand/Fiji, Oman, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine and United Kingdom. | Open, Classified |  |  | 688 | Not_Computer
Z-2439-2015 | 71395 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PTW UNIDOS webline: 1) REF T10023, 2) REF T10022 and 3) REF T10023 Firmware version 2.05. Dosimeter for radiation therapy, diagnostic radiology and health physics. |  |  | May 27, 2015 | 2015 | PTW-Freiburg | In some cases, especially with high dose-per-pulse-rates typical for particle (specifically proton) therapy, measuring errors can result. Individual measurement intervals can be lost that are not visible to the user when there is a high dose-per-pulse rate resulting in a too-low dose being displayed on the monitor. This could result in a radiation over-exposure. | Software design | PTW issues Urgent Safety Notice due to an error with firmware version 2.05 of the UNIDOS webline in some cases, especially with the high dose-per-pulse-rates typical for particle therapy, measuring errors can result. |  | 242 | 242 firmwares | Distributed in CA, MI, MA, NC, NJ, NY, PA, TN, TX, PA &amp; WI. | Terminated | March 31, 2017 | 674 | 242 | Not_Computer
Z-2445-2015 | 71461 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | APEX INTERFACE Hemi 54X and Flared 56X, 3 Hole Acetabular Shells; Model: H5-12354, H5-11356. |  |  | June 04, 2015 | 2015 | Omnilife Science Inc. | The device may have improper screw hole placement due to inaccurate location of the index line. | Process design | Omni sent an Urgent Medical Device Field Correction letter dated June 4, 2015 via email to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review inventory for affected product and complete and return the attached Acknowledgement and Receipt Form by FAX to 508-819-3390. |  | 112 | 112 | US Nationwide Distribution in the states of IL, VA, CO, UT, NY, FL, MA, OK, CA, TN and WY. | Terminated | September 28, 2015 | 116 | 112 | Not_Computer
Z-2442-2015 | 71816 | No | MDS |  |  | Clinical Chemistry | PMA | Guardian Real-Time Monitor, Model No. CSS7100. Indicated for continuous or periodic monitoring of glucose levels in the fluid under the skin, in adults (18 years and older) with diabetes mellitus, for the purpose of improving diabetes management. |  |  | July 22, 2015 | 2015 | Medtronic MiniMed Inc. | Medtronic MiniMed is recalling the Guardian Real-Time Continuous Glucose Monitoring System because it has a language translation error that impacts the Finnish language. Specifically, the Monitor has a Finnish translation error in Predictive Alerts setting: the on-screen Finnish translation for LOW/HIGH is reversed as HIGH/LOW. | Software design | A Urgent Medical Device recall letter was sent to customers on 7/22/15 to inform them that the Guardian Real-Time Continuous Glucose Monitoring System System has a language translation error that impacts the Finnish language. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns are instructed to contact Nordic Support Line 0-800-164-064. |  | 254 | 254 units | Distributed in Finland. | Terminated | September 18, 2015 | 58 | 254 | Not_Computer
Z-2438-2015 | 71950 | No | ISW | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | Otto Bock Healthcare Products GmbH Genium X3 |  |  | August 12, 2015 | 2015 | Otto Bock Healthcare Product | Otto Bock Healthcare Products GmbH is recalling the new Genium X3 knee joint prosthesis due to issues with the resistance-control of the hydraulics, which may fail under conditions. | Process control | The consignee was notified via e-mail and telephone on 12 August 2015. The affected product was identified, as well as the reason for the recall. The firm will replace the affected device with a new one. A return authorization and label will be provided for the recalled product to be returned as soon as possible. |  | 1 | 1 | Distributed to the state OH. | Terminated | October 28, 2015 | 77 | 1 | Not_Computer
Z-2440-2015 | 71898 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 6.0.x |  |  | August 05, 2015 | 2015 | Brainlab AG | ExacTrac 6.0 Patient Positioning System:Display of potentially incorrect Digitally Reconstructed Radiograph (DRR) for x-ray correction andverification. | Software in the Use Environment | Brainlab intends to distribute the Field Safety Notice / Product Notification letter CAPA-20150713-001446, dated July 20, 2015, to the customers. The letter provides instructions for the user on how to address this issue, and to avoid a patient and treatment target position potentially different than intended due to this issue.The Field Safety Notice / Product Notification letter states that as generally required, always verify that the correct DRRs for the current patient are displayed in ExacTrac and an accurate fusion to the X-Ray images is possible for patient positioning. In case incorrect DRRs are displayed, exit and re-start the ExacTrac software. To minimize the probability of this issue occurring for the affected ExacTrac versions: -Reduce the amount of contours exported to ExacTrac as far as possible.-Avoid exporting large contours, such as couch top models, to ExacTrac.-After every (re-)start of the ExacTrac application on any workstation, always first load a phantom plan not containing contours. The Field Safety Notice / Product Notification letter states that as a general reminder the user should always follow the instructions and warnings as described in the user guide. In the context of this notification please specifically consider the safety notes relevant for X-ray correction and verification:-The image overlay functions must be used to verify fusion accuracy in both image views, especially if the images contain a series of similar structures such as vertebrae.-Do not perform patient treatment unless accurate image fusion is possible.-Large correction shifts (e.g. greater than 10 mm) may indicate incorrect automatic fusion. In such cases, verify the correction shift a second time.Brainlab Corrective Action:1. Existing potentially affected ExacTrac 6.0.x customers receive the product notification information.2. Brainlab will provide a software revision (ExacTrac v. 6.0.6) with this issue solved to affe |  | 361 | 361 systems | Worldwide Distribution - USA including the states of AL, AZ, AK, CA CO, CT, FL, GA, ID, IL IN, KY, LA, MD, MA, MI, MN, NE, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, and WI; and, the countries of Australia, Austria, Belgium, Brazil, Canada, China, Denmark, Finland, France, Germany, Hong Kong, India, Israel, Italy, Japan, Netherlands, Panama, Qatar, Russia, Slovenia, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | January 05, 2017 | 519 | 361 | Not_Computer
Z-2443-2015 | 71784 | No | MPO | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(k) | Covidien Argyle Infant Heel Warmers (no tab) |  |  | July 31, 2015 | 2015 | Covidien LLC | Infant Heel warmers may break during activation resulting in spillage of the contents on patients or healthcare worker causing potential skin/eye irritations | Package design/selection | Medtronic/Covidien issued an Urgent Medical Device recall letter dated July 31, 2015 via mailed Certified. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to quarantine and discontinue use of the affected products.. For questions or concerns,contact Medtronic representative or Customer Service at (800)-882-5878. |  | 68800 | 68,800 units | Worldwide Distribution - US Nationwide and the countries: Canada, Mexico, Panama, and Japan. | Open, Classified |  |  | 233549 | Not_Computer
Z-2447-2015 | 71786 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | cobas b 123 Fluid Pack COOX REF 05170036001 400 |  |  | July 14, 2015 | 2015 | Roche Diagnostics Operations, Inc. | Low PO2 results. QC failures of the PO2 parameter, affecting primarily Levels 1 and 2, caused by a calibration issue with the PO2 parameter. This issue may not be detected since QC results can be below mean values, but still within 2 standard deviations (SD) limits. Potential for erroneously low PO2 results in patient samples, especially in blood samples with PO2 values below 50 mmHg | Nonconforming Material/Component | Roche sent an Urgent Medical Device Correction letter dated July 14, 2015 and the Fax Back form via UPS ground to the customers list. The letter identified the affected product, problem and actions to be taken. Consignees are instructed to discontinue use of affected Fluid Packs and discard them per local guidlines; and complete and return the enclosed faxback form and fax it to 1-888-943-5171 to receive replacement product for Fluid Packs from the affected lot. For questions contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-526-2272. |  | 82 | 82 Fluid Packs | US Nationwide Distribution in the states of LA, SC, CA, TX, NE, VA and WV including: Puerto Rico. | Terminated | January 19, 2016 | 189 | 82 | Not_Computer
Z-2449-2015 | 71848 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | CARESTREAM DRX-EVOLUTION X-Ray System |  |  | August 07, 2015 | 2015 | Carestream Health Inc | Software defect that impacts image alignment when using the automatic stitching option for long length images. The defect could cause stitching inconsistencies which may go undetected on a radiograph. | Software design | Carestream sent an Urgent Medical Device Correction letters dated July 29, 2015 and Customer Notification Acknowledgement Forms to customers via Certified Mail. The letter identified the affected product, problem and actions to be taken. Carestream will update the software on your device to correct this issue. For questions or concerns, contact the Carestream Customer Care Center in the U.S. at 1-800-328-2910, available 7 days per week on a 24 hour basis. Outside of the U.S., please call your local Service support number. |  | 236 | Domestic: 76 units, Foreign: 160 units | Worldwide Distribution - US Nationwide and in the countries: Australia, Canada, Chile, China, India, Korea, Malaysia, New Zealand, Singapore, Switzerland, Taiwan and Thailand. | Open, Classified |  |  | 236 | Not_Computer
Z-2450-2015 | 71831 | No | JZG | Beta-2-microglobulin immunological test system. | Immunology | Immunology | 510(K) Exempt | Dimension Vista BÂż2-Microglobulin Flex Reagent Cartridge (B2MIC); |  |  | July 23, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | customers are receiving a high number of Abnormal Assay Errors or are not able to calibrate B2MIC when using Flex reagent cartridge B2MIC various lots. In failure mode, there is the potential for falsely depressed beta2-Microglobulin results on the Dimension Vista system due to assay reagent contamination. | Environmental control | Siemens sent an Urgent Medical Device Correction letter dated July 2015 to all affected customers. The letter identified the affected product, problem, risk to health, and actions to be taken by customers. For questions contact your Customer Care Center - Technical Solutions at 1-800-441-9250 or your local Siemens technical support representative. |  | 1655 | Total 1655 units | US Nationwide Distribution in the states of TN, AZ, CT, Fl, CA, NY and WI | Terminated | March 24, 2017 | 610 | 1655 | Not_Computer
Z-2452-2015 | 71812 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 3.0/4.0 Cannulated Drill Bit; a component of the Tiger Cannulated Screw System |  |  | July 20, 2015 | 2015 | Trilliant Surgical Ltd. | The cannulation of the 3.0/4.0 Cannulated Drill Bit inner diamter is too small to allow for the guide wire to pass completely through. | Nonconforming Material/Component | Trilliant Surgical sent an Important Medical Device Advisory Notice to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Customers were instructed they have the option to exchange the affected drill bits at no charge. To initiate the exchange, customers were instructed to contact their local Trilliant Surgical Sales Representative or contact the Customer Service Department directly at 1-800-495-2919 or via email at orders@trilliantsurgical.com. |  | 121 | 121 units | Nationwide Distribution including NY, TX, OH, IN, FL, NH, MN, SC, OK, NJ, CA, NC, MO, GA, WI, IA, TN, WA, and Puerto Rico. | Open, Classified |  |  | 121 | Not_Computer
Z-2453-2015 | 71826 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | BD CD56 PE Cy 7: NCAM16.2 ASR |  |  | July 17, 2015 | 2015 | Becton, Dickinson and Company, BD Biosciences | Three lots of CD56 PE Cy 7, Catalog number 335809, contain a low amount of HLA-DR PE antibody which could result in an unexpected staining pattern. | Material/Component Contamination | BD Bioscience sent an Urgent Product Recall Letter dated July 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer Actions : 1. If you have CD56 PE-Cy7 integrated into a Laboratory Developed Test (LDT) or panel, evaluate the impact of a false positive in the PE channel the interpretation of on the results.2. Discard all remaining stock of CD56 PE-Cy7, Catalog Number 335809, lot number 4300708, 4304864, and/or 5013643, and request replacement product. BD will replace product at no cost.3. As a requirement of our Regulatory Tracking Process, we request that you complete the enclosed Tracking/Verification form, even if you no longer have the affected reagents listed in your inventory. Please return the Tracking/ Verification form to BD Biosciences, per the instructions (mail, fax, or email).We apologize for any inconvenience that this may have caused and are working toward a long-term solution to this issue. If you require further assistance, please contact BD Customer Support at 877.232.8995 (prompt 3) in the United States. For customers outside the US, contact your local BD Biosciences representative or distributor. |  | 362 | 362 units | US Distribution and Internationally to Taiwan. | Terminated | March 11, 2016 | 238 | 362 | Not_Computer
Z-2456-2015 | 71815 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Concave Reamer Head, 16 mm; Concave Reamer Head, 18 mm; Concave Reamer Head, 20 mm; Concave Reamer Head 22 mm; Convex Reamer Head, 16 mm; Convex Reamer Head, 18 mm; Convex Reamer Head 20 mm; Convex Reamer Head 22 mm, used with Gridlock Plating system, for use in trauma and reconstructive procedures in small bones. |  |  | July 20, 2015 | 2015 | Trilliant Surgical Ltd. | The cannulation of the Concave and Convex Reamer Heads may be too small to allow the Guide Pin to freely pass through. | Component design/selection | The recalling firm sent out written notification of the recall on 7/3/14. The letter instructed the consignees to exchange affected instruments. |  | 638 | 638 units | Nationwide throughout the US | Open, Classified |  |  | 638 | Not_Computer
Z-2451-2015 | 71876 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Trauma Ex-Fix System (Small, Medium, Distraction Osteogenesis (DO) and Large) ; (S) intended to stabilize and provide treatment for fractures of the small bones, such as the hand, wrist, forearm, foot, and ankle; (M) for the construction of an external fixation frame for the treatment of pediatric and adult fractures; )L) for use to provide treatment for long bone and pelvic fractures that require external fixation; and (DO) for fracture fixation (open and closed),pseudoarthrosis or non-unions of long bones, limb lengthening by epiphyseal or metaphysical distraction, correction of bony or soft tissue deformities and correction of segmental bony or soft tissue defects in adult and pediatric patients. |  |  | July 29, 2015 | 2015 | Synthes (USA) Products LLC | Labeling changes were made related to MR conditions that align with standard definitions for MR Safe, MR Conditional, and MR Unsafe. This recall identifies additional Trauma External Fixation System part numbers that should be considered MR Conditional. One serious injury reported. | Labeling design | Urgent recall letters, dated July 31, 2015, were sent to end users and sales consultants to notifiy them about the issue, the associated risks, and to provide instructions for returning the affected devices and response form. |  | 5306080 | 5306080 | Worldwide Distribution-including US (nationwide) and Canada. | Open, Classified |  |  | 5306080 | Not_Computer
Z-2455-2015 | 71470 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Offset Femoral Alignment Guide 4mm x 6Âż; Product Code: KS 67024 |  |  | June 09, 2015 | 2015 | Omnilife Science Inc. | The markings on both versions of the OFFSET FEMORAL ALIGNMENT GUIDES are incorrect. Specifically, the markings on the alignment guides are reversed from the intended position that should provide proper alignment for the offset revision stems when used with the femur. | Device Design | Omni sent an Urgent Medical Device -Field Safety Notice letters to their customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to Return the attached Acknowledgement and Receipt Form by FAX to 508-819-3390. |  | 15 | 15 | Worldwide Distribution - US Nationwide in the states of OK, CO, OH, MA and the countries of Spain and Australia. | Terminated | September 28, 2015 | 111 | 15 | Not_Computer
Z-2461-2015 | 71556 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Wheel caster set screw for Ferno iNX Wheeled Stretcher. Ferno iNX Wheeled Stretcher for pre-hospital patient transport. |  |  | May 01, 2015 | 2015 | Ferno-Washington Inc | The firm discovered an improperly installed set screw in one caster arm assembly on a unit in production. The screw assists in preventing the bearing cup on the caster assembly from backing out of the caster arm during use. | Process control | Ferno's first action of May 1, 2015 was to telephone all customers to alert them to the potential issue and explain our planned inspection. Ferno's Technical Support Manager emailed a copy of the Service Notice and iNX Equipment Evaluation Form to the customers as well. |  | 26 | 26 units | Distributed in the states of AZ, CO, IA, MD, OH, PA &amp; WA, and the countries of Australia, Canada &amp; United Kingdom. | Terminated | November 06, 2015 | 189 | 26 | Not_Computer
Z-2459-2015 | 71708 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | CIDEX Activated Dialdehyde Solution, 4 bottles/case, P/N: 2266 |  |  | July 27, 2015 | 2015 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the CIDEX Activated Dialdehyde Solution because it may lack the adequate data to support the entire 24-month shelf life. | Device Design | A Urgent Product Notification Letter dated July 27, 2015, will be sent to customers and distributor who purchased the CIDEX Activated Dialdehyde Solution because it may lack the adequate data to support the entire 24 month shelf life. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the requested information on the enclosed postage paid Business Reply Card and return it to Stericycle. Customers with questions or need additional packing labels are instructed to contact Stericycle at (877) 550-0701, M-F, 8am-5pm ET. Customers are instructed to report complaints or suspected problems with CIDEX Activated Dialdehyde Solution by contacting (888) 783-7723. |  | 23426 | 23,426 cases | US (nationwide) Distribution | Open, Classified |  |  | 23426 | Not_Computer
Z-2462-2015 | 71854 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | PICC Kits with Peel-Away Sheath Component; catheter, intravascular, therapeutic, ultrasound |  |  | July 30, 2015 | 2015 | Arrow International Inc | the catheter peel-away component hub tabs may prematurely detach. One injury reported. | Unknown/Undetermined by firm | Teleflex sent an Urgent Medical Device Recall Notificatin letters dated July 30, 2015 to their affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to: 1. immediately discontinue use and quarantine any products with the catalogand lot number provided in the notice. 2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507Attn: Customer Service or email it to recalls@teleflex.com. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product toArrow International. 3. Customers were also asked to complete the enclosed Recall Acknowledgment Form if there is no affected product, this will allow for documented receipt of this letter. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 52109 | 52109 | US Nationwide Distribution and internationally to Canada | Open, Classified |  |  | 52109 | Not_Computer
Z-2460-2015 | 71682 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Carepak Sheer Adhesive Bandages 50 Sterile Bandages Assorted Sizes |  |  | June 10, 2015 | 2015 | ASO, LLC | Contain 0.1% Benzalkonium Chloride in the 5/8 bandages inside the assortment which is not listed as an active ingredient in the product label. | Process control | Customers were contacted by phone and e-mail on 06/10/2015; First class letter was sent on 06/23/2015. |  | 655 | 655 cases. | US Distribution to Rhode Island only | Open, Classified |  |  | 655 | Not_Computer
Z-2457-2015 | 71414 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation 4.7, version 4.7.1, used with Structure Definition or Automatic Breast Planning. Radiation Therapy Treatment Planning System, for treatment planning and analysis of radiation therapy. |  |  | May 27, 2015 | 2015 | RAYSEARCH LABORATORIES AB | Issue with the algorithm for ROI contraction in RayStation 4.7 when non-uniform contraction distances are used. The ROI contraction tool uses six distances as input: right/left, inferior/superior and posterior/anterior. These contraction distances are, in error, pairwise reversed, i.e. right is interpreted as left, inferior is interpreted as superior, and posterior is interpreted as anterior. | Software design | On May 27, 2015 RaySearch Laboratories distributed Medical Device Correction notices to their customers via courier service. Within the Medical Device Correction notice, RaySearch Laboratories included a workaround to the problem for their customers. Customers should not use structure templates containing derived ROI definitions with non-uniform contractions. They should avoid using non-uniform contractions. In the case non-uniform contractions must be used, enter switched distances manually and review the result carefully before using it for treatment planning. Lastly, do not use humerus shielding in Automatic Breast Planning. Customers are also expected to complete and return the lower portion of the correction notice via fax to 888-501-7195. **CAPA** Correct the software problem in patch version 4.7.2 of RayStation 4.7, scheduled for market release June 2015. |  | 126 | 126 programs | Distributed in CO, IL, MI, NC, OH,TN, and WA. | Open, Classified |  |  | 126 | Computer
Z-2419-2015 | 71777 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Getinge 86 Series Washer/Disinfector with G1 control system software, Model 8666 and Model 8668. For use in healthcare facilities and laboratories where reusable items such as instruments, utensils, anesthesia sets, and glassware are handled for decontamination purposes. |  |  | June 26, 2015 | 2015 | Getinge Disinfection Ab | Getinge Disinfection AB has initiated a Field Correction due to the potential for unauthenticated remote access to the product over the network. | Software Manufacturing/Software Deployment | Getinge Disinfection AB has initiated a Field Correction due to unauthenticated remote access to the product over the network. |  | 20 | 20 devices | Distributed in CA, CT, GA, IA, ID, MD, MI, MN, NY, PA, TX, and WA. | Open, Classified |  |  | 20 | Not_Computer
Z-2458-2015 | 71889 | No | OFF | Venous blood pressure manometer. | Cardiovascular | Cardiovascular | Enforcement Discretion | Central Line &gt;34 Weeks Kit . Cardiology department. |  |  | August 03, 2015 | 2015 | Medline Industries Inc | The print plate artwork for this lot was found to state: Central Line &lt;34 Weeks Kit displaying an incorrect sign in front of the 34. The description for this device should have stated Central Line &gt;34 Weeks Kit. The defect may lead to the product to be used on patient population that is &lt;34 weeks rather than intended population of &gt;34 weeks. | Error in labeling | Medline Industries contacted the account directly via phone on August 4, 2015, they confirmed no affected product is in their facility. The lot numbers are easily identified on the outside of each kit. Second and third follow up communications will not be sent as we have confirmed the facility does not have any affected product.For questions regarding this recall call 847-643-4129. |  | 12823 | 320 units for lot 14DB9277 &amp; 320 units for lot 14PB2878 | US including PA | Terminated | January 05, 2017 | 521 | 12823 | Not_Computer
Z-2464-2015 | 71813 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Tiger Cannulated Tray Lid Part # 210-00-001-1 NON-STERILE; Combined Tiger/Headless Screw Part # 210-99-101 NON-STERILE; These are components of the Tiger Cannulated Screw System and the Tiger Headless Cannulated Screw System |  |  | July 20, 2015 | 2015 | Trilliant Surgical Ltd. | Sterilization tray lids for the Tiger and Tiger Headless Cannulated Screw Systems were reported to occasionally be difficult to remove from the sterilization tray. | Component design/selection | Trilliant sent an Important Medical Device Adivsory Notice dated November 6, 2013, to all affected customers. The letter instructed customers that they have the option of exchanging affected instruments for trays that meet specification at no charge. To initiate the exchange, customers were instructed to contact their local Trilliant Surgical sales representative or the Customer Service Department directly at 1-800-495-2919, or via email at orders@trilliantsurgical.com. |  | 71 | 71 units | Nationwide Distribution | Open, Classified |  |  | 71 | Not_Computer
Z-2465-2015 | 71913 | No | HSN | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Salto Talaris Tibial Guide, Part # MJU333 ,Tray/Kit Code YKAL11, Non sterile. |  |  | August 06, 2015 | 2015 | Tornier, Inc | Tornier has initiated a voluntary recall for the Salto Talaris Tibial Guide (part number MJU333) due to acomplaint that occurred during physician training, that the guide cannot be assembled to the Salto Talaris Talar Pin Setting guide. This issue will cause the impossibility the time of putting the Talar pin in the guide. | Nonconforming Material/Component | Distributors were sent on 8/6/2015 a Tornier "Urgent - Medical Device Recall" letter, dated August 5, 2015. The letter described the problem and the product involved in the recall. Consignees were advised to review their inventory and the inventory of all reps within their area for the recalled device. The letter requested that consignees complete and return the questionnaire to FieldAction@tornier.com. A Tornier representative will organize the collection and replacement of the devices. For questions they can contact Customer Service Department at 1 (888) 494-7950. |  | 5 | 5 | Distributed to CA, ME, NC, PA, and TX. | Terminated | April 08, 2016 | 246 | 5 | Not_Computer
Z-2463-2015 | 71804 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | HardyCHROM MRSA |  |  | July 09, 2015 | 2015 | Hardy Diagnostics | Hardy Diagnostics is recalling HardyCHROm MRSA due to a report of the potential for breakthrough of the methicillin-sensitive Staphylococcus aureus (MSSA) QC strain, ATCC 29213 on the identified lot. | Unknown/Undetermined by firm | The firm initially notified customers via phone to inform them of a recall. The phone scripts asks customers if they have had any issued with the QC testing. The firm states that our of a precaution they are asking customers to discard the remaining plates of the identified lot, and that credit or replacement plates will be offered for any remaining.The firm followed-up phone calls with a written notification letter. The letter states to complete, sign, and fax or email back the enclosed form stating compliance with the above action. The letters states that is the customer could not be reached via phone and would like replacements or credit to contact the Customer Service Department at 800-266-2222, option 1.Any questions call the Technical Services Department at 800-266-2222, option 2 or email at techservice@hardy diagnostics.com. |  | 2820 | 2,820 (282 pk/10) | Distributed in the states of VA, PA, NY, NJ, IL, OH, and MD. | Terminated | November 10, 2015 | 124 | 2820 | Not_Computer
Z-2467-2015 | 71829 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Quickie IRIS Wheelchair. model EIZ4-2 in combination with option code 183M02 - MONO Backrest system with Dynamic Backrest option. |  |  | July 27, 2015 | 2015 | Sunrise Medical (US) LLC | Quickie IRIS and Zippie IRIS tilt-in-space wheelchairs with MONO Backrest System with Dynamic Backrest may break over time resulting in a fall or injury to occupant. | Device Design | Sunrise Medical sent an Urgent Medical Device Field Correction letter dated July 27, 2015, to all affected dealers. The letter identified the product, the problem, and the action to be taken by the dealer. Dealers were instructed to immediately contact their customers to make arrangements to have the correction made with a replacement kit which will be supplied by Sunrise Medical to each dealer. Each dealer is requested to send back to Sunrise Medical the Acknowledgment and Response Form(s) by fax, email or regular mail once the work is completed. Customers with questions were instructed to contact Sunrise Medical Regulatory Affairs at (888) 208-4901. |  | 64 | 64 total wheelchairs, both models. | Worldwide Distribution - US (nationwide) and Internationally to Australia, Germany and Canada. | Terminated | December 10, 2015 | 136 | 64 | Not_Computer
Z-2374-2015 | 71551 | No | PDR |  |  | Anesthesiology | PMA | SEDASYS Computer-Assisted Personalized Sedation System, SEDPRU01. |  |  | May 13, 2015 | 2015 | Ethicon Endo-Surgery Inc | Ethicon has found that disinfecting methods not specified in the Instructions For Use (IFU) have led to the degradation of the plastic material of the SEDASYS System Control Unit. | Device Design | Sedasys, a Division of Ethicon Endo-Surgery, Inc. sent an Urgent Customer Notification letter dated May 27, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:1. Please distribute this information to all staff within your facility who use the SEDASYS System2. Complete the Business Reply Form (BRF) (Attachment A) confirming receipt of this notice within three (3) business days. The BRF may be given to your sales or service representative or sent to Sedasys by email at SedasysTechSupport@its.jnj.com.3. Keep this notice visibly posted in your facility for awareness.4. If you have any questions or concerns regarding the condition of your SEDASYS System including the Control Units, please contact your Sales or Service Representative.If you need clinical or product support, please contact your local sales representative or call our Customer Support Center at 1-800-SEDASYS (1-800-733-2797). If you need additional communications packages, contact 1-800-SEDASYS and reference Event 052715SED01.For questions regarding this recall call 513-337-7439. |  | 14 | 14 units | Nationwide Distribution including OH, RI, TX &amp; WA. | Terminated | May 02, 2016 | 355 | 14 | Not_Computer
Z-2471-2015 | 64928 | Yes | KDJ | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MiniCap Extended Life PD Transfer Set Extra Short (4") with Twist Clamp. Product Code: 5C4483. |  |  | February 25, 2013 | 2013 | Baxter Healthcare Corp. | Labeling Revision: Additional instructions for the use of peritoneal dialysis transfer sets including the sanitization the connection, healthcare provider monitoring of thyroid function, caution against reuse or reprocessing of the devices, warning against use if tip protectors are not in place, and advisement that the product does not contain natural rubber latex. | Labeling design | Baxter sent an IMPORTANT PRODUCT INFORMATION letters dated February 25, 2013 to affected customers via USPS. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached reply form and return it to Baxter by faxing it to 224-270-5457 of scanning the form and e-mailing it to fca@baxter.com. For general questions about this communication can be answered by contacting The Center for One Baxter at 800-422-9837 (Monday - Friday, 8:00 AM - 5:00 PM CST). Clinical questions can be answered by calling Baxter's Renal Division Clinic Help line at 888-736-2543 x 2 (Monday - Friday, 8:00 AM - 5:00 PM CST). |  | 260244 | 260,244 units | Worldwide Distribution - US: Nationwide + American Samoa + District of Columbia + Guam + Northern Mariana Islands + Puerto Rico; *** FOREIGN: Antigua, Argentina, Australia, Bahamas, Barbados, Bermuda, Brazil, Canada, Chile, Colombia, Cuba, Curacao, Dominican Republic, Ecuador, El Salvador, France (New Caledonia), Guatemala, Guyana, Honduras, Hong Kong, India, Indonesia, Jamaica, Japan, Macau, Malaysia, Mexico, New Zealand, Nicaragua, Panama, Peru, Singapore, Suriname, Taiwan, Thailand, Trinidad and Tobago, United Kingdom (Cayman Islands), Uruguay, Venezuela | Terminated | March 08, 2016 | 1107 | 1543011 | Not_Computer
Z-2468-2015 | 71805 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Stryker |  |  | July 10, 2015 | 2015 | Stryker Sustainability Solutions | Stryker Sustainability Solutions has received reports indicating that devices from one lot (lot# 3574704) were labeled as 100mm models (CB5LT) while the devices inside of the packaging were 75mm models (CB5ST). | Labeling mix-ups | On 07/10/15 the firm sent out notification letters to all customers. The letter states for the customer to discontinue use of the affected lot, and to use other trocar lots reporcessed by the firm. The firm requests that the Recall Effectiveness Check Form be completed. The firm is going to be sending a sales representative to the facilities to check inventory for the affected trocars. A Recall Effectiveness Check will also be completed. If the form indicates that affected devices remain in inventory a prepaid shipping label will be issued for the return of the product. Customers will receive credit for all affected devices returned. The form must be completed even if no affected devices are found in inventory. Stryker Sustainability Solutions Complaint Hotline: 1-888-888-3433 x5555 |  | 90 | 90 devices | U.S. Nationwide Distribution to the following states of:: NV, NM, TX, and MI. | Terminated | December 04, 2015 | 147 | 90 | Not_Computer
Z-2476-2015 | 71735 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP, 10mL in 12mL., Single use, Rx only, Sterile Solution. |  |  | July 02, 2015 | 2015 | MRP, LLC dba AMUSA | Procedures for the acceptance and control of in-process product have not been adequately established. | Packaging process control | The firm notified their consignees of the recall by Fed Ex on 07/02/2015. The firm requested the return of the product. Distributors were instructed to notify their customers. |  | 551 | 551 cases | Distributed in the states of CA. PA, VA, TX, MA, MD, WV, UT, FL, and TN. | Terminated | April 03, 2017 | 641 | 551 | Not_Computer
Z-2475-2015 | 71834 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Synthes Application Instrument for Sternal ZIPFIX; indications for use include primary or secondary closure/repair of the sternum following sternotomy or fracture of the sternum to stabilize the sternum and promote fusion |  |  | July 22, 2015 | 2015 | Synthes (USA) Products LLC | The end cap may loosen and detach making the instrument non-functional. No injuries reported. | Device Design | Urgent recall notification letters, dated July 22, 2015 were sent to end users and consultants to inform them of the issue, the associated risks, and provided instructions to return affected devices along with the response form. |  | 451 | 451 | Nationwide and internationally to Canada. | Open, Classified |  |  | 451 | Not_Computer
Z-2470-2015 | 71964 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity Universal Viewer |  |  | August 11, 2015 | 2015 | GE Healthcare | Images from the Centricity PACS-IW with Universal Viewer andCentricity Universal Viewer may be missing when a system parameter MapRoute is set to a value greater than 1. | Software design (manufacturing process) | The firm sent out an Urgent Medical Device Correction letter dated 8/11/2015 to all affected customers. Letter states that No User activity required. GE Healthcare has completed a remote inspection of your system and found that this system parameter is currently set to the default setting of 1. The firm will correct all affected systems by a providing software update.The letter also states that if the customer has any questions they should call GE Healthcare Service at 1-800-437-1171 or their local service representative. |  | 26 | 26 systems | Worldwide Distribution - U.S. nationwide in the states of: AK, CA, CO, DE, FL, HI, ID, NY, PA, SD, TN and TX; France, Germany, Greece, India, Israel, Malaysia, Mexico, Northern Ireland (UK), Panama, Poland, Russia, South Africa, Saudi Arabia, Switzerland, Turkey, UAE and Uruguay. | Terminated | January 05, 2017 | 513 | 108 | Not_Computer
Z-2477-2015 | 71691 | No | HKO | Ophthalmic refractometer. | Ophthalmic | Ophthalmic | 510(K) Exempt | OPD-Scan III Refractive Power/Corneal Analyzer Opthalmic |  |  | June 15, 2015 | 2015 | Nidek Inc | Software bug was found where there was no difference in Total and Corneal high-order aberrations, but differences were found in Internal high-order aberrations. | Software design | Dear Doctor letters mailing started on July 15, 2015 - first 100. They will be mailed out as the correction is received from Nidek, Japan.Dear Doctor letter will include software upgrade kit including CD or USB, instruction for upgrade and pre-paid response card to be returned to Nidek Inc. |  | 3836 | 3836 | Nationwide Distribution. | Terminated | April 25, 2016 | 315 | 3836 | Not_Computer
Z-2479-2015 | 71840 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | RAPIDPoint 500 Blood Gas Analyzer, |  |  | July 28, 2015 | 2015 | Siemens Healthcare Diagnostics Inc | When both ports (Serial and Ethernet) are configured to transmit data, the data stream from one port could potentially affect the data stream from the other port causing the data management system to include either: duplicate data, missing data, or data from a different patient. | Software design | Siemens Healthcare issued Urgent Field Safety Notice sent by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on July 28, 2015, both in the United States and outside of the United States, for communication with affected customers. This notice informs customers of the issue with the RAPIDPoint 500 dual port LIS transmission feature.Users are informed to: If the system is currently set up with dual LIS transmission enabled, one of the ports (either Serial or Ethernet) needs to be disabled. Instruction provided with the notification. Customer to complete Field Effectiveness Check Form to confirm receipt of the notice.The resolution of this issue will be implemented in the next software release for the RAPIDPoint 500 system.If you have any questions, contact Siemens Customer Care Center or your local Siemens technical support representative. Please fax this completed Customer Care form to (312) 275-7795. |  | 5241 | 5241 units | Worldwide Distribution - US (nationwide) Distribution and to the countries of : Albania, Algeria, Armenia, Australia, Austria, Bahrain, Bangladesh, Belgium, Bosnia Herzeg, Brazil, Bulgaria,Canada, Chile, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany,Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Italy, Japan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Mexico, Netherlands. New Zealand, Norway, P.R. China, Peru, Poland, Portugal, Qatar, Republic Korea, Romania, Russian Fed, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tunisia, Turkey, U.A.E., United Kingdom and Vietnam. | Open, Classified |  |  | 5241 | Not_Computer
Z-2478-2015 | 71970 | No | MTD | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | FIRMap 60mm Catheter; |  |  | August 12, 2015 | 2015 | Abbott Electrophysiology | The expiration date was incorrectly printed on the pouch/box labels of 19 products that were shipped to customer sites. The manufacturing date was accidentally used (copied) for the expiration date. | Incorrect or no expiration date | The firm, Abbott Electrophysiology, sent an "URGENT PRODUCT RECALL" letter dated August 11, 2015. A customer recall notification letter was distributed to firm representatives who will visit each consignee and inspect inventory for affected product. Letters were provided to reps on August 12, 2015. The letter describes the product, problem and actions to be taken. The customers were informed that the product they have in stock will be return to the firm and will be provided a replacement product free of charge.If you have any questions, contact Quality Director at 650-681-1744. |  | 86 | 86 units | US Distribution in states of: KY, IN, MO, NY, OH, TX, LA, IL, PA, NJ, WA, VA, FL, GA, CT, CA, AZ, CO, and TN. | Terminated | September 23, 2015 | 42 | 86 | Not_Computer
Z-2481-2015 | 71835 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | COULTER ISOTON 4 Diluent, Catalog No. 8547148. The diluent provides the ability to analyze portions of the diluted blood sample for different blood cell types, such as red blood cells and platelets. |  |  | July 22, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter is recalling the LH 750 and LH780 Hematology Systems because the use of the diluent lots may eventually result in a compromised Hgb cuvette. | Nonconforming Material/Component | An Urgent Medical Device Recall letter dated 7/22/15 was sent to the affectedcustomers to inform them of the issue and actions to take. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with any questions are instructed to contact Beckman Coulter's Customer Support Center, http://www.beckmancoulter.com/customersupport/support, (800) 526-7694 in US and Canada. Customers outside of the US and Canada, are instructed to contact their local Beckman Coulter Representative. |  | 710 | 710 units | Distributed US (nationwide) including Puerto Rico and U.S. Virgin Islands and the countries of Bermuda, Canada, Guatemala, Hong Kong, Mexico, Panama, Taiwan, and Trinidad-Tobago. | Terminated | December 18, 2015 | 149 | 409752 | Not_Computer
Z-2482-2015 | 71611 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Contour Next Control Solution for Contour Next Blood Glucose Test System. |  |  | June 22, 2015 | 2015 | Bayer Healthcare, LLC | If the control solution bottle is not thoroughly shaken prior to use, the control test result may be higher than the control range and may not be recognized and marked as a control test by the blood glucose meter. Also, the control solution, used as a quality control test in blood glucose meter systems, may read outside the specified range too high or too low. | Device Design | On 6/22/2015, Bayer Healthcare began sending URGENT - MEDICAL DEVICE FIELD CORRECTIVE ACTION notifications to affected Health Care Providers, Pharmacies, and Users. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers are instructed to always check the expiration, not to use and discard of expired control solutions. Also, only CONTOUR NEXT control solution should be used with the CONTOUR NEXT test strips. The control solution bottle should be shaken about 15 times before each use. The letter instructed customers to refer to their meter's user guide for further testing instructions. If test results outside the control range are received, another test should be performed following the instructions provided in the letter. If the problem persists, customers should contact Bayer Diabetes Support. Bayer will replace the control solution or provide new control ranges for the test strips, depending on the situation. The firm has implemented improvements to provide more consistent control solution product. |  | 5000000 | 5 million bottles of control solution | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Croatia, Denmark, Finland, France, Germany, Greece, Ireland, Italy, Luxembourg, Netherlands, Norway, Portugal, Slovenia, Spain, Sweden, Switzerland, UK, Albania, Azerbaijan, Bulgaria, Czech Republic, Estonia, Hungary, Korea, Latvia, Lithuania, Philippines, Romania, Serbia, Slovakia, and Thailand. | Open, Classified |  |  | 5000000 | Not_Computer
Z-2483-2015 | 71850 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vision RingLoc Shell 50 mm Size 23. Hip joint replacement prosthesis. |  |  | July 28, 2015 | 2015 | Biomet, Inc. | The device was assembled without screw hole plugs. | Under Investigation by firm | Zimmer Biomet issued notices on July 28, 2015, via FedEx overnight mail and e-mail to customers. Customers were asked to return response forms and recalled devices to Zimmer Biomet. Customers with questions may call Seth Pixton, Field Action Specialist, at 574-372-1570, Monday through Friday from 8:00 am to 5:00 pm, or e-mail at: seth.pixton@zimmerbiomet.com. |  | 5 | 5 | Distributed in Canada and Japan. | Open, Classified |  |  | 5 | Not_Computer
Z-2487-2015 | 71988 | No | JBF | Erythrocytic glucose-6-phosphate dehydrogenase assay. | Hematology | Hematology | 510(k) | Pointe Scientific G6PD Controls |  |  | May 19, 2015 | 2015 | Pointe Scientific, Inc. | A reduction in the reconstituted stability has been identified.Clinicians may notice a drop in recovery of the control after reconstitution, which could result in an out of specification result prior to the current reconstitution claim of 7 days. For this reason we suggest discontinuing use of the control set. | Nonconforming Material/Component | The firm, POINTE SCIENTIFIC, issued a "TECHNICAL BULLETIN" on 5/19/2015, notifications were sent to the affected distributors via electronic mail. The recall notification included a description of the reason for the recall, affected product, and actions to be taken. The customers were instructed to: discontinue use of the control set and to incorporate this bulletin into your Quality Control files.The firm noted that this product has been placed on indefinite back-order and a replacement product is not currently available. For questions contact Pointe Scientific's Technical Service Department at 1-800-757-5313. |  | 16632 | 2772 total vials (374 kits of 6 vials, 528 invidually sold vials) | Worldwide distribution: US (nationwide) including states of: PA, CA, KY, FL, NJ, DE, MI, MN,and IL; and countries of: United Arab Emirates, Austria, Italy, Israel, Greece, and Australia. | Open, Classified |  |  | 16632 | Not_Computer
Z-2488-2015 | 71862 | No | MDI |  |  | Orthopedic | 510(k) | VEPTR Vertical Expandable Prosthetic Titanium Rib, inferior cradles also known as caudal rib supports. The VEPTR/VEPTR II devices mechanically stabilize and/or correct thoracic deformities to allow the chest and lungs to grow. |  |  | July 29, 2015 | 2015 | Synthes (USA) Products LLC | The last hole on certain VEPTR Inferior Cradles may be partially formed. Fully lengthening the construct to the partially formed last hole may result in mechanical failure of the construct due to rib sleeve breakage, and may require surgury to exchange components. Effects may also include pain, loss of deformity correction, and/or soft tissue/visceral injury. | Process control | Urgent recall letters, dated July 29, 2015, were sent to end users alerting them to the problem, the associated risk to patients, and provided instruction to return affected devices withthe response form. |  | 3262 | 3262 | Worldwide distribution including US nationwide, Canada and Germany. | Open, Classified |  |  | 3262 | Not_Computer
Z-2490-2015 | 71967 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 7 Hole VL Gridlock Fibula Plate, Part # 300-60-001 NON-STERILE SINGLE USE ONLY; 9 Hole VL Gridlock Fibula Plate, Part # 300-60-002 NON-STERILE SINGLE USE ONLY; 11 Hole VL Gridlock Fibula Plate, Part # 300-60-003 NON-STERILE SINGLE USE ONLY. For use in trauma and reconstructive procedures in small bones. |  |  | August 06, 2015 | 2015 | Trilliant Surgical Ltd. | The affected parts subject to the recall are out of specification, resulting in in the loss of the optimal locking screw functionality of the plates. The thread depth of the screw holes is insufficient, potentially limiting the ability to engage and lock with associated mating locking screws. | Nonconforming Material/Component | Trilliant Surgical sent an Urgent Medical Device Recall letter dated August 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter instructed and confirmed consignees to discontinue use of the affected lots, quarantine, and return product (if not yet done so) in accordance with an included acknowledgement sheet. For further questions, please call 1(800) 495-2919 , 8:00am - 5:00pm CST |  | 28 | 28 units | US Distribution to the states of : FL, OH, NJ, MN, WI, NM and CA. | Open, Classified |  |  | 28 | Not_Computer
Z-2489-2015 | 71855 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow International, Inc. PICC Kits and VPS Access Kits with Touhy Borst Component Catheter, Intravascular, Therapeutic, Ultrasound, Long Term Greater than 30 Days |  |  | July 30, 2015 | 2015 | Arrow International Inc | saline leakage from unsecured connections between the Touhy Borst adaptor and the hub of the catheter. No injuries reported. | Under Investigation by firm | Arrow International sent an Urgent Medical Device Recall Notification letter dated July 30, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:If you have affected stock, immediately discontinue use and quarantine any products with the catalogand lot number listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507Attn: Customer Service or email it to recalls@teleflex.com. This will allow us to document the amountof product you have on hand for return. A customer service representative will contact you with aReturn Goods Authorization (RGA) Number and will provide instructions for the return of product toArrow International.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to1-855-419-8507, Attn: Customer Service or email it to recalls@teleflex.com. This will allow us to document your receipt of this letter.Customers with questions were instructed to call their local sales representative or Customer Service at 1-866-246-6990.For questions regarding this recall call 610-378-0131. |  | 92902 | 92902 | Worldwide Distribution - US nationwide and Internationally to Canada. | Open, Classified |  |  | 92902 | Not_Computer
Z-2492-2015 | 71966 | No | BSY | Tracheobronchial suction catheter. | Anesthesiology | Anesthesiology | 510(K) Exempt | KimVent* Closed Suction System for Adults, 12 F, T-Piece (Product Code 2155); KimVent* Closed Suction System for Adults, 14 F, T-Piece (Product Code 2205, 8169, 8248, 22015, 22025, 220125, 220135) |  |  | July 29, 2015 | 2015 | Halyard Health, Inc | The thumb valve of some KIMVENT Closed Suction Systems remained partially open after suctioning was completed, even though the valve appeared to be closed or in the "LOCKED" position. | Process control | Halyard sent a Product Advisory Notice on July 29, 2015, via FEDEX with a second letter dated August 12, 2015, being issued via FEDEX. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to report any occurrence of a valve leak to Halyard Customer Service at 1-844-425-9273, option 1, option 3. |  | 19250 | 19,250 cases | Worldwide Distribution - US (nationwide and Puerto Rico and Internationally to Canada | Terminated | January 12, 2016 | 167 | 19250 | Not_Computer
Z-2491-2015 | 71882 | No | OUO | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/MR, Diagnostic Imaging System , Philips Medical Systems |  |  | July 23, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare has become aware of a problem in which the Ingenuity TF PET/MR PET Reconstruction Server (PRS) database may lock up after an MR acquisition is completed and before a PET acquisition is begun, even though no actual database lockup issue has been reported in the field-installed Ingenuity TF PET/MR systems. | Software design | On July 23, 2015 the firm sent Field Safety Notices to their customers. |  | 11 | 11 | Product was shipped to the following states: NY, OH &amp; TX. Product was also shipped to the following countries: Finland, Germany, Hong Kong, Japan, Netherlands, South Korea, Spain &amp; Switzerland. | Open, Classified |  |  | 11 | Not_Computer
Z-2486-2015 | 71803 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Bright Embrace, a Single Bili Light for newborn phototherapy |  |  | August 31, 2013 | 2013 | Physician Engineered Products, Inc. | Upper limit for ambient temperature for use of the device was printed incorrectly in the instructions | Device Design | Physician Engineered Products, Inc. sent an Bright Embrace Recall notification dated September 10, 2013, to notify customers of the recall of the Bright Embrace .PEP is issuing a Return Authorization to you. Upon receipt PEP will reimburse to you:1) The cost of the device(s) 2) Shipping costs-both from and to PEPFor further questions please call (207) 935-1256. |  | 97 | 97 units | Worldwide Distribution - US Distribution to the states of : CA, CO, FL, KY, MA, MD, NM, PA, TN, TX and VA., and to the countries of : Egypt, Honduras, Saudi Arabia and Singapore. | Terminated | September 02, 2015 | 732 | 97 | Not_Computer
Z-2529-2015 | 71893 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Mahurkar Elite Acute Dual Lumen Catheter Tray, High Flow Pre-Curved with IC* Safety Components, 13.5 Fr x 20 cm |  |  | August 06, 2015 | 2015 | Covidien LLC | Sterility may be compromised | Packaging process control | Medtronic/Covidien sent an Urgent Medical Device Recall letter August 7, 2015 to affected customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter instructs customers to quarantine, discontinue use and return the affected product as instructed in the letter. For questions contact your Medtronic representative or Customer Service at (800)-882-5878 |  | 4 | 4 kits | US Nationwide Distribution | Terminated | June 29, 2017 | 693 | 420002 | Not_Computer
Z-2538-2015 | 71653 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Panoramic X-ray Model PC-1000 |  |  | June 04, 2015 | 2015 | Panoramic Rental Corp. | Jackscrew nut part failure potentially leading to device malfunction and patient/operator injury. | Equipment maintenance | The firm, Panoramic Corporation, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated 7/17/2015 via certified mail to customers on 7/20/2015, with response forms for customers to complete and return. The firm previously issued "IMPORTANT SERVICE BULLETIN" notices in 2003 and 2008 via FED-EX. The letter describes the product, problem and actions to be taken. The customers are instructed to select to use several options including replacing the failing part with one containing the safety feature, or trading in their machine for a newer one; discontinue using the devices immediately regardless the option you select, and complete and return the attached response card via email to recall@pancorp.com or fax to 260-489-5683. If the machine has been sold or donated to another dental outfit, the customer should notify Panoramic so that they can update their records. When the customer is reached, the firm will confirm whether or not the machine is still in service.Customers with questions may contact 800-654-2027 Monday through Thursday 8:00 AM to 6:00 PM and Friday 8:00 AM to 5:00 PM EST to discuss these options and determine which is best for their practice. |  | 7885 | 7885 | Worldwide distribution: US (nationwide) all 50 states, including Puerto Rico; and countries of Antigua, Armenia, Aruba, Austria, Canada, China, Colombia, Ecuador, Egypt, England, France, Germany, Greece, Haiti, Honduras, Indonesia, Japan, Kenya, Lithuania, Mexico, Panama, Saudi Arabia, Singapore, Taiwan, Turkey, UAE, Uganda, Venezuela, Vietnam, and Virgin Islands. | Terminated | May 30, 2017 | 726 | 7885 | Not_Computer
Z-2541-2015 | 71895 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | CrossCHECK Locking Screw in the following sizes: |  |  | July 21, 2015 | 2015 | Wright Medical Technology, Inc. | These screws have a potential of passing through the plate in the case when torque is applied in excessive amounts, which may result in intraoperative malfunction. | Employee error | Wright sent an Urgent Medical Device Recall letter dated July 20, 2015, to all affected customers. Customers were asked to complete the attached response form confirming receipt of the notice and the status of any products on hand, and return a copy to Wright by email (complaints@wmt.com or fax (901)867-7401. Also, if customers have any product in their inventory the product should be sent back to Arlington, TN (ALM). For questions regarding this recall call 901-290-5662. |  | 107 | 107 units | Nationwide Distribution | Open, Classified |  |  | 11588 | Not_Computer
Z-2524-2015 | 71872 | No | FYD | Air-handling apparatus for a surgical operating room. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Neptune 2 Waste Management System Docking Station (120V and 230V), New and Refurbished |  |  | July 20, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | The Fluid Coupling Sleeve, a component within the Neptune Waste Management System Docker, may be out of specification. This can cut or dislodge the PTFE seal and cause a leak between the Rover and Docker couplers during docking or leaking from the Rover coupler after docking. The potential hazard is waste fluid and/or fresh water dispensed from the Neptune 2 Rover at the incorrect time. | Nonconforming Material/Component | Stryker sent an Urgent Medical Device Notification letter dated August 4, 2015 via FedEx to all domestic customers, and all international cutsomers and sales reperesentatives were notified via e-mail. The letter identified the affected product, problem and actions to be taken. Customers are asked to complete the response form and return it to Stryker whether they have affected devices or not. Fax the completed, signed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762, or scan and email a copy to kellyjo.davis@stryker.com. For questions regarding this recall, contact Stryker Instruments at 269-389-2921. |  | 124188 | 524, 237 of these are non-conforming | Worldwide Distribution - US Nationwide in the states of: AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY and countries of: Austria, Canada, Hong Kong, Netherlands, Finland and Spain. | Terminated | November 07, 2016 | 476 | 124188 | Not_Computer
Z-2383-2015 | 71838 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DigitalDiagnost Release 4.0.3, Release 4.1/4.1.1. excluding systems with SP1 |  |  | March 23, 2015 | 2015 | Philips Electronics North America Corporation | The system is designed to emit a beep upon termination of an exposure. However, if the system has been powered on for over 12 hours, the system will no longer emit this signal due to a defect in the Microsoft Windows 7 operating system. | Radiation Control for Health and Safety Act | Philips Healthcare sent an Urgent Field Safety Notice letter dated June 25, 2015 to customers. The letter identified the affected product, problem and actions to be taken. Philips plans to install a software upgrade in affected systems, which will eliminate this issue. Philips Healthcare Field Service Engineers will schedule an appointment with customers to install the software update. For questions contact your local philips representative. |  | 62 | 62 | US Nationwide Distribution to the states of : Arizona, Minnesota, State of Washington, Wyoming and Ohio. Internationally to Canada. | Open, Classified |  |  | 62 | Not_Computer
Z-2546-2015 | 71963 | Yes | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | MediTrace Cadence Adult Zoll, Preconnect |  |  | August 19, 2015 | 2015 | Covidien LLC | Multi-Function Defibrillation Electrodes, Pre connect-potential for damage to the wire insulation may result in a potentially increased risk for reduced or no patient therapy and patient and/or clinician burns | Process control | Medtonic/Covidien issued recall letter on August 19, 2015. Customers were notified via Federal Express and the letter instructs customers to discontinue use of the product and return all inventory to Medtronic (Covidien), 2 Ludlow Park Drive, Chicopee, MA 01022, Attention: Field Action Returns.Contact your Medtronic representative or Customer Service at (800)-882-5878. |  | 114000 | 114,000 sets | NationwideForeign:CanadaAustraliaDenmarkGermanyNew Zealand | Terminated | May 24, 2017 | 644 | 177660 | Not_Computer
Z-2543-2015 | 71785 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | cobas Liat Analyzer, for in vitro diagnostic use. |  |  | July 24, 2015 | 2015 | Roche Molecular Systems, Inc. | Due to a supplier issue, the incorrect green and amber LED components were used in the manufacture of the LED printed circuit board assemblies (PCBA) in certain cobas Liat Analyzers. | Nonconforming Material/Component | Roche sent an Urgent Medical Device Recall letter dated July 24, 2015, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.Actions Required Identify the affected systems in your inventory by referencing the serial numbers in the table on page 1. You will be contacted to return affected systems for immediate replacement.Complete the attached fax form and fax it to 1-317-521-4815.File this UMDR for future reference.Please contact the Roche Support Network Customer Support Center cobasÂż Liat hotline at 1-800-800-5973, if you have questions about the information contained in this letter. |  | 5 | 5 pieces | US Distribution to the states of :TX and NC. | Terminated | December 23, 2015 | 152 | 5 | Not_Computer
Z-2540-2015 | 71972 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SMR Shoulder/Spalla Trial Heads ECC .4 mm Adaptor, REF 9013.30.021; SMR Shoulder/Spalla Trial Heads ECC .8 mm Adaptor, REF 9013.30.031; for use in shoulder implant surgery. |  |  | August 14, 2015 | 2015 | Limacorporate S.p.A | The dimensions of the mating features of the recalled products are out of specification causing the device to be difficult to assemble. | Nonconforming Material/Component | Lima Corporate sent an Urgent Attention Recall letter on August 11, 2015, to US distributors and Sales Agents. The affected customers were instructed to examine their inventory, return the affected lots to the recalling firm, and complete a response form. |  | 93 | 93 units | US Distribution to the states of: OH, TX, NC, NY, CA, FL, KS, AL, MI, MT, CO, MO, NV, WA, MN and LA., and Internationally to Australia. | Terminated | February 26, 2016 | 196 | 203 | Not_Computer
Z-2544-2015 | 71619 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Catalog Number/Product Code 6801375, Unique Device Identifier No. 10758750001132, and VITROS 5,1 FS Chemistry System Refurbished, Catalog Number/Product Code 6801890, Unique Device Identifier No. 10758750001644, IVD --- Ortho Clinical Diagnostics Inc. |  |  | June 30, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly may occur during the installation of an Assay Data Diskette (ADD) using Software Versions 2.2.1 through 2.8. If the user is utilizing User- Configured (Modified) parameters, ALL default settings are restored in the VITROS 5,1 FS Chemistry System from the ADD instead of retaining the user modified parameters. | Software design | The firm, Ortho Clinical Diagnostics, sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter dated 6/30/15 and response forms to customers via FedEx overnight courier, e-mail (ORTHO PLUS e-Communications) and/or via US Postal Service overnight letter on 6/30/2015. Foreign affiliates were informed by email on 6/30/2015 of the issue and instructed to notify their consignees of the issue and required actions. The letter described the product, problem, and actions to be taken. The customers were instructed to follow the required actions and complete and return the Confirmation of Receipt-Response Required form, no later than 7/8/15, by fax at 1.888.557.3759 or 1.585.453.4110, or scan PDF and email to: ConfirmationAdmin@its.jnj.com. Customers with any additional questions, can contact Customer Technical Services at 1-800-421-3311 at any time. |  | 210 | Catalog 6801375: Domestic 919; Foreign 1251; Catalog 6801890: Domestic 63, Foreign 147 | Worldwide Distribution: US (nationwide) including Puerto Rico, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 210 | Not_Computer
Z-2719-2015 | 71978 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza; picture archiving and communications system |  |  | July 06, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Potential issue leading to data loss and patient data mix-up | Software design | Siemens sent a Customer Information letter, dated July 6, 2015, to customers The letter identified the affected product, problem and action to be taken. The letter informed cusotmers about the Hotfix software update that was being released. |  | 37 | 37 | US Nationwide Distribution - in the states of CA, OR, FL, MO, NY, WA, MD, MA, NY, WI, KS, MN and PA. | Terminated | January 06, 2017 | 550 | 37 | Not_Computer
Z-2718-2015 | 71881 | No | PDU | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Volcano Pioneer Plus Re-Entry Catheter; |  |  | July 22, 2015 | 2015 | Volcano Corporation | Due to a manufacturing defect, there is the possibility that a small wire could extend through the catheter shaft of the Pioneer Plus intravascular Ultrasound Guided Re-Entry Catheters. | Under Investigation by firm | The Volcano Corporation sent an Urgent Medical Device Recall letter dated July 22, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately quarantine and return all affected product in their inventory. Customers were also asked to complete the attached form and fax back to 866-723-2482. Customers with questions were instructed to call 866-861-8923.For questions regarding this recall call 978-439-3586. |  | 5042 | 5042 devices | Worldwide Distribution - US (nationwide) Canada, Europe, Saudi Arabia, Iran, Egypt, and Qatar. | Terminated | April 28, 2016 | 281 | 5042 | Not_Computer
Z-0417-2015 | 69545 | No | JZH | Properdin factor B immunological test system. | Immunology | Immunology | 510(K) Exempt | MicroVue Bb Plus EIA is packaged as a kit which contains all reagent components to perform the test. In vitro diagnostic. |  |  | October 16, 2014 | 2014 | Diagnostic Hybrids Inc | Samples are quantitating incorrectly, with approximately 50% higher concentration values than previously reported and historically documented. | Device Design | Quidel sent an Urgent Medical Device Recall letter dated October 16, 2014, to all consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately review their inventory and determine if they have of the affected product. If so, they should remove it from their inventory and destroy it immediately by discarding it into their normal biomedical waste stream. Complete the enclosed Certificate of Destruction. Upon receipt of the Certificate of Destruction, Quidel, will then schedule their shipment of replacement product. If consignees further distributed the affected product, they should identify the customers and provide them with the Urgent Medical Device Recall. Consignees were instructed to monitor and reconcile the quarantine of product with the customers as this information may need to be provided to the authorities. For questions regarding this recall call 740-589-3300.For technical support consignees should contact technicalsupport@quidel.com or call 800-874-1517. |  | 171 | 171 | Worldwide Distribution - USA, including AL, CO, CT, FL, KY, MN, NC, NV, OH, PA, and WA; and, the countries of Canada, China, India, and Switzerland. | Terminated | September 03, 2015 | 322 | 171 | Not_Computer
Z-2548-2015 | 71721 | No | OEQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | Enforcement Discretion | ANGIODRAPE PK 4/CS |  |  | July 10, 2015 | 2015 | Customed, Inc | Recall initiated due to inadequate validation/qualification of sterilization processes, and package integrity for products released for distribution as sterile. | Packaging | Customed issued an "URGENT VOLUNTARY MEDICAL DEVICE RECALL" notification letter dated July 10, 2015 to all affected customers. The letter identified the product, problem, and actions to be taken by the customer. |  | 192 | 192 packs | Distributed Only in Puerto Rico. | Terminated | April 12, 2017 | 642 | 358405 | Not_Computer
Z-2732-2015 | 72024 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Magic3 Intermittent Catheter with Sure Grip Hydrophilic Coated Silicone Coude tip, Product Code 50614. The catheters are packaged in film-to-film pouches and packaged 30 per box. |  |  | August 24, 2015 | 2015 | C.R. Bard, Inc. | Potential breach of the sterile barrier packaging. | Unknown/Undetermined by firm | The firm, Bard, sent an "URGENT: MEDICAL PRODUCT RECALL" letter dated August 20, 2015 to customers via FedEx on 8/24/2015. The letter describes the product, problem and actions to be taken. The customers were instructed to not use or further distribute any affected product; check inventory and remove any identified affected product from your shelves; complete and return the Recall and Effectiveness Check Form by September 19, 2015, even if you have used or no longer have possession of the affected product, via fax to 1-770-784-6469 or email a scanned copy to Regan.Gealy@crbard.com ( if you cannot fax or email the form, call 1-770-784-6471 and report the required information verbally); and if you have further distributed any units of the affected lot, identify your customers and notify them at once of this product recall, and return product to Bard Medical Division, 8195 Industrial Blvd., Covington, GA 30014. BMD will provide replacement product or credit for your return product.If you have any questions, please call Vice President, Quality Assurance at 770-784-6119. |  | 8970 | 8,970 each | US Distribution to states of:CA, FL, IL, MA, MN, NH, NJ, NY, NC, OH, OK, PA, TX, and WA. | Open, Classified |  |  | 8970 | Not_Computer
Z-2734-2015 | 72026 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Xi Surgical System P4 software;with Xi Stapler; model number IS4000 : |  |  | August 17, 2015 | 2015 | Intuitive Surgical, Inc. | Complaints regarding the Exposed Knife Blade recoverable fault for the Xi Stapler with Xi System P4 version software. | Software design | Intuitive sent an Urgent Medical Device Correction letters dated August 21, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. For questions contact Clinical Sales or ISI Customer service at 800-876-1310 Option 3. |  | 198 | 198 affected devices | Worldwide Distribution - US Nationwide and in the countries of Italy and Norway. | Terminated | February 08, 2016 | 175 | 198 | Not_Computer
Z-2730-2015 | 71832 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | Neuro Kinetics, Inc. (NKI) I-PortalÂż NOTC system with the OVAR research test; a rotary chair system used in vestibular and neuro-otologic diagnostic testing. This device provides no diagnoses nor does it provide diagnostic recommendations.The OVAR test has not yet been cleared by the FDA, and until then can only be used as an investigational device. |  |  | July 21, 2015 | 2015 | Neuro Kinetics, Inc. | The devices that included the OVAR research test were not cleared for marketing and mTBI research package was sold with the incorrect investigational label. | No Marketing Application | Neuro Kinetics sent notification letters to end users on/about July 24, 2015, to alert them to the reason for recall and inform them of the corrective action planned by Neurokinetics.Action to be taken by you the Customer/User:If you have already initiated OVAR research (an IRB approved study with non-significant risk designation is ongoing), please reply to this email and include a copy of the IRB approval letter or the IRB approval information (e.g., investigators name, IRB name, IRB location, and study reference number). In return, NKI will mail you a field correction package with the necessary and appropriate Investigational Device labeling.If you have not initiated OVAR research (an IRB approved study is not ongoing), please discontinue use of the OVAR test and reply to this email to confirm that you are not using the OVAR test as part of an IRB approved study. In return, NKI will mail you a field correction package to disable the OVAR test functionality, which can be restored at the time you have an IRB approved study. Disabling the OVAR research test will have no effects on the functionality of the clinical tests included with your NOTC system. There will be no perceivable changes to the system outside of not being able to run the OVAR test.For questions regarding this recall call 412-963-6649. |  | 45 | 45 | Worldwide Distribution - US (nationwide) | Open, Classified |  |  | 45 | Not_Computer
Z-2722-2015 | 71907 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | Transonic Flow Probe, 6 mm, Catalog Number HQC6FMV -- |  |  | July 27, 2015 | 2015 | Transonic Systems Inc | Software defect; programming errors in the blood flow probes EPROM. Transonic received several complaints from users indicating that when the flow probe is connected to a flow meter the meter reads Used Probe or Irrelevant Use Number. | Software Manufacturing/Software Deployment | Transonic sent an Urgent Medical Device Recall letters and Response Forms dated July 27, 2015 to customers via Certified Mail, email and/or fax. The letter identified the affected product, problem and actions to be taken. Customers are instructed to contact Transonic Systems' Customer Service at 1-800-353-3569 or 1-607-257-5300 to arrange for replacement (repair/exchange) immediately. |  | 1 | Domestic: 1 unit | Worldwide Distribution - US Nationwide and the countries of: Canada, Japan, Russia and The Netherlands. | Terminated | July 18, 2016 | 357 | 143 | Computer
Z-2733-2015 | 71878 | No | MSW | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Access Thyroglobulin Assay; 100 determinations, 2 packs, 50 tests/pack. For the quantitative determination of thyroglobulin levels in human serum and plasma using the Access Immunoassay Systems. |  |  | July 28, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter, Inc. is recalling Access Thyroglobulin reagent packs because some may exhibit greater than 10% imprecision at concentrations &lt;10 ng/mL. | Under Investigation by firm | On 07/28/2015 the firm sent out customer notification letters via first class mail and e-mail. The firm's letter contained the following instructions:1. Discard all Access Thyroglobulin reagent pack lots listed in the letter.2. Evaluate Access Thyroglobulin results &lt;10 ng/mL when using the reagent pack lot numbers listed to determine if controls and/or patients were higher than expected or did not match the clinical status of the patient.3. At the discretion of the Laboratory Director, notify clinicians that it is possible the Access Thyroglobulin sample results reported by the laboratory were affected by this issue.The firm's notification letter requests that the information be shared with laboratory staff. If any affected products have been forwarded provide a copy of the letter to them. The firm requests that customers complete and return the Response Form within 10 days.Any questions contact the Customer Care Center:1. From website at http://www.beckmancoulter.com2. By phone: call 1-800-854-3633 in the United States and Canada.3. Outside the United States and Canada, contact local Beckman Coulter representative |  | 3035 | 3,035 (U.S.) and 7,123 (International) | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Algeria, Australia, Austria, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Czech Republic, Egypt, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Korea, Republic of, Lebanon, Lithuania, Mexico, Morocco, New Zealand, Oman, Philippines, Portugal, Qatar, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Vietnam and Yemen. | Terminated | April 13, 2016 | 260 | 3035 | Not_Computer
Z-2484-2015 | 71651 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | OmniPodÂż Insulin Management System. Only the OmniPodÂż device is affected. Catalog Number(s): POD-ZXP420 (US). Intended for subcutaneous delivery of insulin at set and variable rates for the management of diabetes mellitus in persons requiring insulin. |  |  | July 13, 2015 | 2015 | Insulet Corporation | OmniPodsÂż (Pods) have a higher rate of failure causing:Cannula fails to deploy/retract causing insulin to be pumped or the audible alarm is displayed on the PDM and pod will not deliver insulin | Nonconforming Material/Component | The firm, Insulet Corporation, sent an "URGENT: Voluntary Device Recall" letter dated 7/13/2015 to each consignee/customer via Email notification (where possible) on 7/13/15 as well as Federal Express. Customers not responding to the email or Federal Express will receive follow up phone calls. The letter described the product, problem and actions to e taken. The customers were instructed to do the following: If you have Pods from any of these lots, set them aside, we will replace them at no charge; arrange for return and replacement by choosing one of the following methods:1. Complete the online form at https://www.myomnipod.com/podinfo (fastest method).2. Return the enclosed reply card by mail or fax to 1-855-407-3729.3. Contact Insulet Customer Care by calling 1-855-407-3729 (any time day or night).If you do not have any unused Pods form these lots, reply using one of the methods above.If you have any question, contact Acting VP RA/QA at 978-600-7000.EXPANDED RECALL:Insulet Corp. issued on 11/2/15 the URGENT: Field Safety Notification via Email notification and Federal Express. Letter describes problem of that certain lots of the pod's needle failed to deploy or there is a delay in the deployment of the needle. Customers not responding to the email or Federal Express will receive additional mailing and/or follow up phone calls. Call Customer Care at 1-855-407-3729 if you have any questions regarding this Field Safety Notification |  | 51059.9 | 24829 boxes Expanded: 26,230.9 boxes | Worldwide Distribution: US (Nationwide) and country of: Switzerland. | Terminated | January 26, 2017 | 563 | 51059.9 | Not_Computer
Z-2739-2015 | 71843 | No | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA CentaurÂż Folate Ascorbic Acid/Ascorbic Acid Diluent; SMN 10316877. |  |  | July 27, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc | Low red blood cell control recoveries and low results for red blood cell patient samples when using the ADVIA Centaur Systems Folate assay. | Process design | An Urgent Medical Device Recall (UMDR) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on July 27, 2015 to be delivered to customers on July 28, 2015. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on July 27, 2015. These notices inform the customer to discontinue use of ADVIA Centaur Folate Ascorbic Acid/Ascorbic Acid Diluent kit lots 54942, 56230 and 70673and dispose of the remaining inventory per local regulations. The ADVIA Centaur Systems Folate red blood cell sample and control negative bias is minimal in the lower range of the assay and is more pronounced with increasing folate concentrations. The negative bias could lead to a misclassification of a patient as folate deficient. Siemens does not recommend a look back of previously tested patients. Fax-back forms will be used to confirm that the customers received the Urgent Medical Device Recall. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Medical Device Recall and completed the required actions. |  | 466 | 466 | Worldwide Distribution-US (nationwide) including the states of CA CT DE FL GA IA IL IN KS KY MA MD MI MO NJ NM NV NY OH PA RI SC TN TX VA WA WI and OUS: Belgium Estonia Finland France Germany Great Britain Ireland Italy Pakistan Spain Sweden Switzerland Turkey. | Terminated | February 26, 2016 | 214 | 466 | Not_Computer
Z-2737-2015 | 71934 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray DPM 7 Monitor, Multi Parameter Patient Monitor (with Arrhythmia Detection and Alarms) |  |  | July 27, 2015 | 2015 | Mindray DS USA, Inc. dba Mindray North America | The DPM 7 Monitors may display a black screen. | Under Investigation by firm | Mindray sent an Urgent Medical Device Correction / DPM 7 Monitor letter, dated July 27, 2015, to their affected customers via certified mail and return receipt. The letter identified the problem as well as the affected product, adverse effects on patients, and corrective action. The firm will repair the potentially affected monitor at no cost to the customer. Customers are to contact their Mindray Service Representative at 1-800-288-2121 to arrange for the repair. |  | 13 | 13 units | Distributed to the states of CT, IA, KY, MS, PA, UT and WA. | Open, Classified |  |  | 13 | Not_Computer
Z-2738-2015 | 71936 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 2.4, 3.0, and 4.0 mm Gridlock Screw Driver Bit, Model Number 310-30-003, Lot Number TSL002605, NON-STERILE. For use with the Gridlock Plating System for trauma and reconstructive procedures of small bones. |  |  | July 31, 2015 | 2015 | Trilliant Surgical Ltd. | Gridlock screw driver bits for the Gridlock Plating System were identified to be out of specification and have the potential for shallow engagement between the driver and the screw, resulting in an increased likelihood of stripping screw heads intraoperatively. Specifically, the minor diameter of the drive feature is above the specification, potentially reducing the engagement between the driver | Component design/selection | The firm sent out a written notification of the recall on 7/31/2015. |  | 108 | 108 units | NY, OH, CA, SC, NJ, AZ, VA, IN, FL, MN, OK, AR, MO, WI, GA, SD, IA, IL, TX | Open, Classified |  |  | 108 | Not_Computer
Z-2735-2015 | 72025 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MEDIVATORS HEMOCOR HPH700, HPH-1400TS, HPH-1000TS. |  |  | August 18, 2015 | 2015 | Medtronic Perfusion Systems | Hemocor High Performance Hemoconcentrators demonstrate low ultrafiltration performance that is below product specification. These potentially affected units have been distributed as stand-alone devices and in specific lots of Medtronic Perfusion Tubing packs. | Nonconforming Material/Component | Consignees were sent a Medtronic "Urgent Medical Device Recall" letter dated August 2015. The letter was addressed to "Risk Manager". The letter described the problem and the product involved in the recall. Informed consignees that no action is needed for product that has been used, but patients should continue to be managed in accordance with their standard patient management protocol. For unused product they recommend to immediately quarantine and return the affected product to Medtronic. Contact Customer Service at 800-854-3570 and reference RADAR #253277 to initiated a return and credit of product. Requested consignees to complete the Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. If the product has been forwarded to another facility, consignees are advised to notify that facility. For questions contact Medtronic Lifeline Technical Services at 877-526-7890 or your Medtronic representative. |  | 1594 | 1594 total (1306 US, 288 OUS) | Worldwide Distribution-US (nationwide) including the states of AL, AR, CA, CO, FL, HI, ID, IN, KS, LA, MA, MS, MO, NY, NJ, OH, PA, TN, TX, UT, and WI, and the countries of Bolivia, Denmark, Dominican Republic, Italy, Latvia, Lebanon, Mexico, Panama, Singapore, Venezuela, and Zambia. | Open, Classified |  |  | 1594 | Not_Computer
Z-2741-2015 | 71725 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | DimensionÂż CREA (Creatinine) and Dimension VistaÂż CREA (Creatinine); The CREA method used on the DimensionÂż clinical chemistry system is an in vitro diagnostic test intended for the quantitative determination of creatinine in human serum, plasma and urine. |  |  | July 15, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed that the Dimension CREA and Dimension Vista CREA may produce falsely depressed results up to - 12.7% and -22% respectively, for a creatinine concentration of 1.5 mg/dL when hemolysis exceeds 200 mg/dL and 300 mg/dL respectively. | Device Design | Siemens sent an Urgent Medical Device Correction letter, dated July 2015, was provided to all DimensionÂż CREA and Dimension VistaÂż CREA customers who received the products to notify them of the issue, the potential risk to health, and actions to be taken by customers.If you have any questions please contact your Siemens Customer Care Center Technical Solutions at (800) 441-9250, or your local Siemens Technical Support Representative. |  | 32165 | Total number Dimension Vista - 32,165, Dimension - 238,634 | Worldwide Distribution - US (nationwide) and to the countries of : Afghanistan, Algeria, Argentina, Armenia, Australia, Austria, Bahamas, Bahrain, Bangladesh, Belgium, Bolivia, Bosnia Herzegovinia, Bulgaria, Brazil, Canada, Canary Islands, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, El Salvador, France, Georgia, Germany, Ghana, Great Britain, Greece, Guatemala, Hungary, Iraq, Ireland, Honduras, Hong Kong, Iceland, Indonesia, India, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Lithuania, Libya, Macedonia, Malaysia, Mali, Mexico, Myanmar, Morocco, Netherlands, Nepal, Nicaragua, New Zealand, Norway, Oman, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Tunisia, Turkmenistan, United Arab Emirates, Uruguay, Venezuela, Vietnam and Yemen. | Open, Classified |  |  | 32165 | Not_Computer
Z-2740-2015 | 71149 | [6 empty] | DermaPen Pro Model, Micro Needling System, a motorized AC powered surgical instrument for dermabrasive use. Each motorized handpiece is packaged within a small box. Each disposable needle cartridge is packaged individually and then placed into boxes of 30 quantities per box. |  |  | February 05, 2015 | 2015 | Derma Pen, LLC | Derma Pen initiated a field communication for Dermapen after receiving an FDA warning letter for distributing a device without FDA clearance or approval. | No Marketing Application | Derma Pen sent a notification letter to its customers on February 5, 2015. The purpose of the notification was to inform customers that the FDA issued a warning letter and the steps Derma Pen was taking to address the FDA' s concerns.For further questions please call (800) 981-5402. |  | 1466 | 1,466 units | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico and The Virgin Islands and to the countries of : Canada, Brazil, Colombia, Ecuador, Costa Rica, Australia, Hong Kong, Netherlands, Mexico, Romania, Sweden, Peru, Greece, United Kingdom and United Arab Emirates. | Open, Classified |  |  | 1466 | Not_Computer
Z-2744-2015 | 71940 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Biomet Parsippany, NJ. 4 0 MM X 14 MM Variable Screw C-Tek Maxan Anterior Cervical Plate System. Lot 375440. |  |  | August 05, 2015 | 2015 | Biomet Spine, LLC | Some of P/N 14-521614B Lot 375440 was mislabeled as P/N 14-521614 Lot 375440. | Labeling mix-ups | Zimmer Biomet sent an Urgent Medical Device Recall Notice dated August 5, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by customer. Customers were instructed to immediately locate and discontinue use of the affected product. This action requires the immediate location and discontinued use of the item and its immediate return to Zimmer Biomet Spine. Specifically, customers were REQUIRED to take the following steps:" Immediately locate and remove from circulation the items consigned/loaned to youraccount as identified on page two of this notice." Carefully follow the instructions on the enclosed "Response Form"." If you have further distributed these items to medical facilities, you MUST notify them of this action. This letter MUST be given to the person responsible for receiving recall notices. However, you are charged with the location and return of these items.Health care professionals and consumers may report serious adverse events or product quality problems with the use of this product to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail, fax or phone:Online:Fax:Phone:Regular Mail:www. fda .gov /MedWatch/report.htm(800) FDA-0178(800} FDA-1088Use postage-paid FDA form 3500 available at www.fda.gov/Medwatch/getforms.htmand mail to Medwatch, 5600 Fishers Lane, Rockville, MD 20852-9787.Questions related to this notice should be directed to 303-501-8400 Monday through Friday, 8 a.m.to 5 p.m. MST.Zimmer Blomet |  | 7 | 7 parts | Nationwide Distribution including Distributed nationwide to OH, LA, TX, and GA. | Terminated | February 10, 2016 | 189 | 7 | Not_Computer
Z-2742-2015 | 71937 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI 500 Blood Parameter Monitoring System. |  |  | August 07, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Inaccuracies in SvO2, temperature, pH, pCO2, pO2, Hematocrit, and Potassium readings following a software upgrade to version 1.69. | Software design | A voluntary Urgent Medical Device Correction notice that clearly explains initial in-vivo calibration requirements, device operating ranges, and temperature measurements was sent on 08/17/2015, via express mail to consignees of CDI System 500 v1.69. Following the initial notice, Terumo will be updating the Operators Manual and will send the new manual to each consignee when available. One manual per unit at each facility will be provided. Customers with questions and return response forms may contact: Terumo Recall Email: tcvs.recall@terumomedical.comTerumo Recall Fax: 734-741-6149Terumo CVS Customer Service: 1.800.521.2818, Monday - Friday, 8 a.m.  6 p.m. ET. |  | 4638 | 4638 | Worldwide Distribution-US (nationwide) including DC and PR and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, and the countries of Australia, Belgium, Canada, Chile, Colombia, Hong Kong, Indonesia, Japan, Malaysia, Mexico, Philippines, Singapore, South Korea, Taiwan, Thailand, United Arab Emirates. | Open, Classified |  |  | 4638 | Not_Computer
Z-2748-2015 | 71880 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Elcam Medical Minimal Residual Volume Luer activated Swabbable stopcock (MRVLS) for pressure monitoring and intravenous fluid administration; Manufactured by: Elcam Medical ACAL, Israel. Sterile, Quantity 20. |  |  | July 19, 2015 | 2015 | Elcam Medical, Inc. | Punctured blister packages were detected during the packaging process at the manufacturing site. Potential for device contamination. | Packaging | Elcam Medical sent "Urgent: Stopcock Recall" notices and "Recall Return Response Forms" dated 7/19/2015 to their customers. The recall notification informed the customers of the reason for the recall and the actions they must take. Customers were instructed to cease distribution and quarantine any affected products; return to Elcam Medical the enclosed Recall Return Response Form; and notify their customers as per their recall procedures. |  | 7721580 | 7,721,580 units | Texas, New Mexico and Pennsylvania | Open, Classified |  |  | 7969680 | Not_Computer
Z-2753-2015 | 72036 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Needle Driver; used with the da Vinci Si Surgical System (IS3000); |  |  | August 08, 2015 | 2015 | Intuitive Surgical, Inc. | In certain circumstances involving select da Vinci Si Single-Site instruments, the jaws of grip actuated instruments may become fixed in a closed position. | Device Design | Intuitive Surgical sent an Urgent Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:1. Ensure that all affected personnel are fully informed of this notice. Forward this notice to your Risk Manager, OR Director, Purchasing Manager, Biomedical Engineering staff and members of your medical staff who perform da Vinci Surgery procedures.2. Replace any instrument that shows signs of damage and contact Intuitive SurgicalCustomer Service.3. Complete and return the attached Acknowledgment Form to Intuitive Surgical usingthe instructions provided.4. Please retain a copy of this notice as well as the attached instructions with your Single-Site User Manual.If you need further information or support concerning this issue, please contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:Âż North and South America: (800) 876-1310, Option 3 (6 AM to 5 PM PST) or mail:customersupport-servicesupport@intusurg.comÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 AM to 6PM CET) or ics@intusurg.comÂż South Korea: 02-3271-3200 (9 AM to 6 PM KSTJ)Âż Japan: 0120-56-5635 or 003-5575-1362 (9 AM to 6 PM JST)For questions regarding this recall call 408-523-2100. |  | 89 | 89 total units | Worldwide Distribution - US (nationwide) and Internationally to Australia, Belgium, Brazil, Canada, Chile, China, Cyprus, Denmark, France, Germany, Greece, Israel, Italy, Monaco, Netherlands, Panama, Romania, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Terminated | December 03, 2015 | 117 | 89 | Not_Computer
Z-2747-2015 | 71985 | No | NXT |  |  | Orthopedic | PMA | Smith &amp; Nephew BIRMINGHAM HIP (TM) RESURFACING DYSPLASIA CUP W/IMPACTOR, REF Numbers: 74120246 (46 MM), 74122248 (48 MM), 74120250 (50 MM), 74122252 (52 MM), 74120254 (54 MM) |  |  | June 03, 2015 | 2015 | Smith &amp; Nephew, Inc. | The firm observed that the BIRMINGHAM HIP(TM) Resurfacing (BHR) System had revision rates which were higher than established benchmarks. | Device Design | Smith &amp; Nephew sent an notification letter dated June 3, 2015 to customers via E-mail and Fed- Ex. The letter identified the affected product, problem and actions to be taken. Customers were asked to quarantine the product and return it. Distributors were asked to notified their customers of the recall and ask them to return the product. |  | 624 | 624 units | Worldwide Distribution - US Nationwide and countries of: Australia, Austria, Bahrain (UAE), Belgium, Canada, China, Denmark, Finland, Germany, Hong Kong, Italy, Japan, Malaysia, Mexico, Netherlands, New Zealand, Norway, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand and United Arab Emirates (UAE). | Open, Classified |  |  | 16778 | Not_Computer
Z-2759-2015 | 71877 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Biomet Spine Lineum OCC Starter Flexible and Lineum Flexible Inner Shaft, Model Numbers 14-525050 and 14-525057, reusable devices distributed with bubble wrap inside a plastic pouch. |  |  | July 22, 2015 | 2015 | Biomet Spine, LLC | Biomet has initiated a recall on Biomet Spine Lineum OCC Starter Flexible and its inner shaft instrumentsafter identifying that these instruments were distributed prior to completion of all design control activities. | Device Design | The firm, Zimmer Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated 7/23/2015 to their Consignees/Customers via certified letter. The letter described the product, problem and actions to be taken. The customers were instructed to immediately locate and discontinue use of device; remove from circulation; carefully follow instructions on enclosed "Response Form" and complete and return via fax to 303.501.8444; if you have further distributed these devices to medical facilities, you MUST notify them of this action; and return device and form to Zimmer Biomet Spine, 310 Interlocken Pkwy, Ste 120, Broomfield, CO 80021. Questions related to this notice should be directed to the Quality Engineer at 303-501-8457 Monday through Friday, 8 a.m. to 5 p.m. MST. |  | 49 | 49 | Worldwide Distribution: US (nationwide) in states of: NY, MO, WI, IN, CA, NC, and OH, and to countries of: Japan and Netherlands. | Terminated | December 10, 2015 | 141 | 49 | Not_Computer
Z-2761-2015 | 72122 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Hemocor High Performance Hemoconcentrator, HPH700, sterilized with ethylene oxide |  |  | August 03, 2015 | 2015 | Medivators, Inc. | Hemoconcentrators may exhibit low ultrafiltration performance that is below product specification. | Process control | Medivators sent an "Urgent Medical Device" recall letter, dated August 4, 2015, via email. The letter described the issue, identified the affected lot, requested the product be quarantined, and to contact Medivators for return instructions and credit or no-charge replacement product. A response form was requested to be returned to Medivators. Customers with questions can contact Medivators Product at 763-553-3348. |  | 840 | 840 units | US Distribution to the states of CA and MI. | Terminated | November 04, 2015 | 93 | 840 | Not_Computer
Z-2762-2015 | 72019 | No | NAB | Nonresorbable gauze/sponge for external use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | XGAUZE Trauma Dressing, catalog number FIN-0002-01. |  |  | August 12, 2015 | 2015 | Revmedx Inc | The XGAUZE Trauma Dressing was recalled because the packages may acquire small tears or pinholes during shipping due to use of shipping containers that lacked adequate protective materials. If present, these small tears or pinholes may disrupt package integrity and render the device unsterile. | Package design/selection | RevMedx, Inc. sent the Urgent Medical Device Correction and Removal XGAUZE Trauma Dressing letter to US Customers on August 12, 2015.Customers are instructed to inspect packages for signs of damages or loss of vacuum. The affected packages will appear puffy and the dressing will be loose inside. The non-affected packages will have a good vacuum intact seal and the packaging material should contour snugly against the dressing.Customers are advised to return any packages appear to have damaged or loss of vacuum. Any undamaged packages may be used as indicated.Customers with question can call the firm at (503) 218-2172 and fax the Return Response Form via fax to (503) 218-2274 or send via email to info@revmedx.com. |  | 5004 | 5004 units | US Distribution to the states of : Arizona, Idaho, North Carolina, Tennessee and Virginia. | Terminated | September 17, 2015 | 36 | 5004 | Not_Computer
Z-2760-2015 | 72000 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | UCLA Abutment Hexed Castable Cylinder 3.4mm |  |  | July 29, 2015 | 2015 | Biomet 3i, LLC | MUCC2, non-hexed, castable abutment cylinders, has been labeled and distributed as a MUCC1C, hexed, castable abutment cylinders for this one lot only. | Process control | Biomet sent a Medical Device Recall Notice dated July 27, 2015, to all their clinicians informing them about the recall and to complete the attached recall return response form and any unused product to be replaced at no charge. Contact information 1-800-342-5454 |  | 61 | 61 pieces | Worldwide Distribution including Italy and Uruguay. | Open, Classified |  |  | 61 | Not_Computer
Z-2764-2015 | 71953 | No | LBR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific autoHDL Reagent H7545 |  |  | March 13, 2015 | 2015 | Pointe Scientific, Inc. | Crystals may be observed in the R1 reagent. Crystals may be a result of the reagent being frozen during transport. The crystals do not impact product performance. | Nonconforming Material/Component | Pointe Scientific sent a Field Correction Notice on March 13, 2015, to all affected customers via fax, email, and mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Any questions concerning the corrective action may be directed to Technical Service Department at 1-800-757-5313. |  | 40280 | H7545-R1, 8.75 L; 12-H7545-162, 4 kits; H7545-1000, 3 kits; HDL600, 6 kits | Worldwide Distribution - US including MI, FL, CA, TX, CT, PR and Internationally to Canada. | Terminated | October 29, 2015 | 230 | 40280 | Not_Computer
Z-2763-2015 | 71935 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | CIVCO Shoulder Retractors, Reusable non-sterile hand grip, REF 302070, packaged 1 set per box, RX. |  |  | August 19, 2015 | 2015 | Med Tec Inc | The wrist straps used for positioning the patient are too long | Vendor change control | Med Tec Inc sent an Urgent Field Correction Notice dated August 18, 2015, via e-mail on August 19, 2015, and made phone calls on August 19, 2015, to all affected customers requesting return of the product. Customers were instructed to forward a copy of the notification letter to all customers who have received the product. Customers with questions were instructed to reach their representative by calling CIVCO Medical Solutions, adiation Oncology, at 800-842-8688 or 319-248-8688 between 8:00AM and 5:00PM CST. For questions regarding this recall call319-248-6502. |  | 18 | 18 sets | Worldwide Distribution - US including NE, SC, and KS. Foreign distribution was made to Australia, China, Venezuela, Spain, Turkey, and Hong Kong | Open, Classified |  |  | 18 | Not_Computer
Z-2769-2015 | 71956 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Voyant Open Fusion Device |  |  | July 31, 2015 | 2015 | Applied Medical Resources Corp | Potential to improperly characterize tissue within the jaws, which may result in nonoptimal power delivery. | Device Design | The firm sent out customer notification letters to International customers the beginning of July 31st, 2015 and domestic customers the week of August 3rd, 2015. The firm's letter requested customers complete the following actions:- check inventory for recalled product- complete the attached Recall Notification Confirmation Form to acknowledge the recall and indicate if facility is returning or has already used any of the products- Return the confirmation form(s) to Applied Medical by emailing torecall60669166@appliedmedical.com or faxing to (949) 713-8936.- Return affected product and a copy of the confirmation form(s) to Applied MedicalFor product return questions contact Karen Mitchell, Sales Operations Supervisor, RGA Dept. at (949) 713-8622 or by e-mail at kmitchell@appliedmedical.com.For regulatory questions contact Lauren Contursi, at (949) 713-8767 or bye-mail at lcontursi@appliedmedical.com. |  | 101 | 101 units (U.S) and 38 units (International) | Worldwide Distribution - U.S Nationwide in the states of:FL, GA, NV. and countries of: France, Netherlands, Spain, and United Kingdom. | Terminated | November 20, 2015 | 112 | 101 | Not_Computer
Z-2768-2015 | 71993 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MedTest DX Control Reconstitution Fluid |  |  | May 15, 2015 | 2015 | MedtestDx, Inc. | Vial to vial variation in the fill volume that could affect control recovery issues and cause a control result to fall outside the suggested recovery ranges. An under-fill of the vials was found ranging as low as 4.63 ml. When used to reconstitute HNC/HEC, control values were shown to be out of range running on the high end. | Process design | On 5/15/2015, MedTest DX Field Correction Notice notifications were sent to the affected customers via electronic mail and US mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Any questions concerning the corrective action may be directed to MedTest Technical Service Department at 1-800-757-5313. |  | 79 | 79 kits | US Nationwide Distribution. | Terminated | May 23, 2016 | 374 | 79 | Not_Computer
Z-2767-2015 | 72015 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | NICO Myriad Console |  |  | July 29, 2015 | 2015 | Nico Corp. | Two cables within the console were inadvertently twisted during manufacturing. The individual wires within the cables may break and/or contact one another where the cables terminate into two connectors. When these wires break or contact one another (i.e., short), various types of malfunction may occur with the console. One type of malfunction may result in unexpected aspiration increase. | Process control | The firm, Nico Corporation, sent an "URGENT: NICO Console Recall" notification dated 7/31/15 via fax to its customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The customers were instructed to return the console and complete and return the Recall Response Form via fax to: 317-683-0305 or email to: Jay.Dittman@niconeuro.com. The firm will provide customers with a Myriad Console loaner kit.If you have any questions, contact Director of Quality and Regulatory, at 317-660-7118 ext 104 or email to: jay.dittman@nicneuro.com. |  | 6 | 6 units | US Distribution to states of: MD, IN, CA, and NC. | Terminated | September 22, 2015 | 55 | 6 | Not_Computer
Z-2775-2015 | 71994 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Angiographic Injector. Medrad Mark 7 Arterion Injection System, which utilizes Software Version SW 005.006_SH; to be used specifically for the purposes of injecting contrast medium and common flushing solutions into humans for angiographic studies. |  |  | August 04, 2015 | 2015 | Bayer Healthcare | Software Version SW 005.006_SH, resulted in the removal of purge enforcement from traditional New-Case, Power Up and Syringe Change use cases while the injector head is in the upright position. | Software design | Direct accounts were notified by UPS Notification letter dated August 4, 2015 and advised when using the Arterion Injection System to ensure that purges are completed to avoid the injection of air. |  | 71 | 71 | Worldwide Distribution. US Nationwide, Germany, France, UK, Netherlands, Sweden, and Canada. | Open, Classified |  |  | 71 | Not_Computer
Z-2772-2015 | 72054 | No | KIH | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Aliquot Delivery Kit (Syringe and Plunger); for use in open and percutaneous procedures to facilitate placement of bone cement in weakened or diminished bone |  |  | August 20, 2015 | 2015 | Orthovita, Inc., dBA Stryker Orthobiologics. | Two isolated instances of holes in the outer header bag component of the double barrier sterile packaging configuration | Under Investigation by firm | Stryker Sent an Urgent Product Recall letter dated August 20, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken by the Customer/User:Our records indicate that you have received the above referenced instruments. It is Stryker Orthobiologics responsibility as the manufacturer to ensure that customers who may have received these affected instruments also receive this important communication. Please assist us in meeting our regulatory obligation by faxing back the attached Product Recall Acknowledgment Form as soon as possible to 215-253-5020.We regret any inconvenience this action may cause you and if you have any questions, please contact Stryker Orthobiologics at 610-640-1775 x 5299. |  | 9805 | 9805 | Worldwide Distribution - Nationwide Distribution and to the countries of : Australia, Belgium, Denmark, France, Germany, Italy, Malta, Netherlands, South Africa, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Open, Classified |  |  | 9805 | Not_Computer
Z-2771-2015 | 71657 | No | FWB | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sony LMD-1951MD, Sony LMD1951MD/BS, Sony LMD-1951MC (Foreign) 19 inch medical grade liquid crystal display (LCD) Medical Monitors with light-emitting diode (LED) backlight technology. |  |  | January 13, 2015 | 2015 | Sony Medical Systems | Tthe medical monitor has either not turned on or has lost image during clinical use. | Device Design | Sony Corporation sent an URGENT: MEDICAL DEVICE FIELD ACTION field letter dated January 13, 2015. Customers are asked to affix the warning sticker to the rear of the monitor and place the adhesive seal over the AC inlet. If the firm does not possess an AC adaptor (AC-110MD) for each of the affected monitors, stop using the monitor and take the monitor out of service until the permanent corrective action described below is implemented.SEL will be replacing the power board of all affected units with a redesigned power board that does not exhibit the appropriate failure mode. Questions can be directed to 866-789-SONY (7669) between 8:00 am and 8:00 pm Monday through Friday. |  | 4952 | 4952 units (1016 US; 3936 OUS) | Worldwide Distribution - US Distribution to the states of : AZ, CA, NC, PA, IA, OH and FL., and to the countries of : China, Japan, Taiwan, South Korea and Indonesia. | Open, Classified |  |  | 4952 | Not_Computer
Z-2770-2015 | 71851 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Lumenis VersaCut+ Tissue Morcellator GA-0007500 |  |  | July 30, 2015 | 2015 | Lumenis | There is a probability the hand piece may operate inconsistently and unpredictably during the morcellation procedure resulting in damage to non-target tissue. | Device Design | Lumenis sent an Urgent Medical Device Recall letter on August 7, 2015, to all consignees. Consignees were instructed to destroy the operator manual, and return the affected device to Lumenis along with the completed Recall Verification form. Consignees with questions were instructed to call 801-656-2663. |  | 162 | 54 devices with 108 hand pieces | Worldwide Distribution - US including CA, PA, MA, FL, and NY, and Internationally to Italy, Germany, Australia, China, Great Britian, United Kingdom, Taiwan, France, Peru, Mexico, Turkey, Spain, and Russia. | Open, Classified |  |  | 162 | Not_Computer
Z-2776-2015 | 71971 | No | DXO | Catheter tip pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Edwards Pressure Monitoring Kit with TruWave disposable pressure transducer. A sterile, single-use kit that monitors pressures when attached to pressure monitoring catheters. For use on patients requiring intravascular, intracranial, or intrauterine pressure monitoring. |  |  | July 31, 2015 | 2015 | Edward Lifesciences, Llc | Flow restriction issues in a trifurcated IV set included in the pressure transducer kits. A restriction of the flow of IV fluid would result in either an inability to prime the system before use or difficulty in flushing the line after a blood draw. | Device Design | Edwards sent an Urgent Product Recall Letter on August 10, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The firm's letter requests that customers return any of the units that they currently have in inventory. The letter requests that customers verify the inventory that has been pre-populated on the acknowledgment form, and complete the remainder of the form. The firm states to fax back the acknowledgment form within three days of receipt of notification to Edwards Customer Service at 1-800-422-9329. The firm requests to call Edwards Customer Service at 1-800-424-3278 to arrange for return of any unused product from the implicated lots as well as to obtain information about product replacment. The notification states to call for an RGA prior to returning any unused product and return unused product to the firm at Edwards Lifesciences Return Goods. Any questions the firm states to contact Edwards Customer Service at 1-800-424-3278. |  | 13077 | 13,077 (U.S.) and 12,306 (International) | Worldwide Distribution - US nationwide, Australia, Austria, Canada, China, Colombia, Croatia, Denmark, France, Germany, Great Britain, Italy, Netherlands, New Zealand, Sweden, and Slovenia. | Open, Classified |  |  | 13077 | Not_Computer
Z-2774-2015 | 71887 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | CIDEX OPA Solution, P/N 20390 for use as a high level disinfectant for reprocessing heat sensitive medical devices. |  |  | August 10, 2015 | 2015 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is reminding all CIDEX OPA Solution customers to follow the warnings and rinsing procedure in their current labelling around the criticality of proper rinsing of medical device following disinfection with CIDEX OPA according to the CIDEX OPA Solution IFU (Instruction For Use). | Under Investigation by firm | A Urgent Medical Device Field Safety Notice will be sent to customers on 8/10/15 to inform them to follow the warnings and rinsing procedure in their current labelling around the criticality of proper rinsing of medical device following disinfection with CIDEX OPA according to the CIDEX OPA Solution IFU (Instruction For Use). The letter provides the customers with the description of the issues and the recommended actions to be taken. Customers are instructed to complete and return the attached Business Reply Card (double-sided) via mail or by fax to Stericycle. For questions regarding the business reply card, customers are instructed to contact Stericycle. Customers with any complaints or suspected problems with CIDEX OPA Solution related to the field safety letter, are instructed to contact 1-888-783-7723. |  | 318632 | 318,632 cases | US | Open, Classified |  |  | 318632 | Not_Computer
Z-2773-2015 | 71974 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pointe Scientific Creatinine Reagent |  |  | December 12, 2014 | 2014 | Pointe Scientific, Inc. | Creatinine kit C7539-150 contains Creatinine reagent R1 that is labeled as Creatinine reagent R2 (30ml) and reagent R2 vial is labeled as R1 (120ml). With manual procedure, no result will be reported as the control will be out of specifications. | Process control | On 12/12/2014, Field Correction Notice notifications were sent to the affected distributors via electronic mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Any questions concerning the corrective action may be directed to Pointe Scientifics Technical Service Department at 1-800-757-5313. |  | 93 | 93 kits | Worldwide distribution. US nationwide, Singapore, Kenya, Bolivia, Mexico, Greece, Guyana, and Philippines. | Terminated | October 29, 2015 | 321 | 93 | Not_Computer
Z-2779-2015 | 71910 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | The Leica ASP200S/ASP300S Tissue Processor. Pathology and Hematology Departments. |  |  | August 05, 2015 | 2015 | Leica Microsystems, Inc. | The Leica ASP200S/ASP300S Tissue Processor have an incorrect internal wiring to the remote alarm plug, which can cause the remote alarm to not function correctly in the case of an instrument failure. | Nonconforming Material/Component | Leica Biosystems sent an Urgent Medical Device Recall Notification letter dated August 4, 2015, to all affected customers to provide instructions on what the customer should do with the recalled product. These include 1) As an immediate countermeasure, the firm recommends that the customers do not use the remote alarm function on the Leica ASP200S/ASP300S in its current state. 2) The customers local Leica representative will contact them to arrange for a Leica service engineer to visit their facility to correct the internal wiring on their tissue processor and to check the remote alarm function. 3) The Medical Device Recall Notification needs to be passed on to all those who need to be aware within the custromers organisation or to any organisation where the potentially affected devices have been transferred. Please maintain awareness of this Medical Device Recall Notification and resulting action to ensure effectiveness of the corrective action. 4) The letter asks the customer use the enclosed Medical Device Recall Notification Response Form to confirm that they have received and understand the Medical Device Recall Notification. it states "Please confirm receipt of this notice by signing and dating the attached Acknowledgement Form and faxing it back to us at 1-847-236-3747, or scan and email it to LMGRA@leica-microsystems.com" |  | 901 | 901 | Worldwide Distribution - US including U.S. including AZ, CA, CT, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MN, MO, MT, NC, NJ, NM, NY, OH, OR, PA, TX, VA, WA, WI, WV and District of Columbia. Internationall to AE, AT, AU , BD , BE, BF, BG,BR, CA, CL, CN, CO, CR, DE, DK, EC, EG, ES, ET, FI, FR, GB, HU, ID, IN, IT, JM, JP, KR, LK, MA, MT, MY, NL, NZ, PA, PE, PH, PL, PT, RO, RU, SA, SE, SG, TH, TR, TW, VN and ZA. | Terminated | January 05, 2017 | 519 | 901 | Not_Computer
Z-2781-2015 | 71909 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | NebuTech Nebulizer, Model number: Series 8660, 8960, 8961, 8966, 8984. Nebulizer used to generate aerosols delivered directly to the patient for breathing. Designed to be used in either a hospital or homecare environment. |  |  | August 04, 2015 | 2015 | Salter Labs | The Inspiratory valve cap on the nebulizer reservoir tower may be loose and may detach before or during treatment. | Component change control | Salter Labs sent a Medical Device Recall letter dated July 24 2015. The letter identified the affected product, problem and actions to be taken. The letter states that all affected product can be returned to Salter Labs for replacement at no charge with product specifically inspected for a tighter fit on the inspiratory valve cap. Customers were instructed to complete and return the Return Response Form as soon as possible. For any questions contact Salter Labs Customer Support at 800-421-0024 or the local Salter Labs representative. |  | 8975 | 8,975 | Worldwide Distribution - US Nationwide and in the countries of Greece, UK, Belgium, Australia, and Switerzland. | Terminated | May 06, 2016 | 276 | 8975 | Not_Computer
Z-2777-2015 | 72136 | No | LPB |  |  | Cardiovascular | PMA | Medtronic CryoConsole, models 106A3, 106E2, and 106A2-K |  |  | September 01, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | An issue with a USB memory component contained within a subset of CryoConsoles can result in extended procedure time. | Component design/selection | Medtronic began hand delivering an Urgent Medical Device Correction letter on September 1, 2015, to all affected customers. The letter described the issue. Customers were advised to continue using the device, per the operator's manual, and if a system notice appears, they are to carefully read and follow the instructions in the screen of the console. Medtronic personnel will track contact with the customer via Customer Contact Record. Customers with qeustions were instructed to contact their Medtronic AF Solutions representative.For questions regarding this recall call 1-877-464-2796 Technical Services. |  | 103 | 103 units | Worldwide Distribution - US (nationwide) and Internationally to Algeria, Australia, France, Germany, Greece, Iceland, Ireland, Israel, Italy, Lebanon, Malaysia, Martinique, Netherlands, Poland, Romania, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | April 08, 2016 | 220 | 103 | Not_Computer
Z-2782-2015 | 72158 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Fluorescence Imaging Procedure Kits used with the da Vinci Si and Xi Surgical Systems; |  |  | September 09, 2015 | 2015 | Intuitive Surgical, Inc. | B. Braun 2 gang 4-way stopcocks in the FluorescenceImaging Procedure Kit may experience damage to the sterile packaging during shipping. The damage canresult in tears to the packaging and has the potential to cause a breach in sterility of the supplied stopcocks. | Package design/selection | Intuitive sent an Urgent Medical Device Correction letters dated September 9, 2015 to all customers via traceable method. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to :1. Ensure that all affected personnel are fully informed of this notice. Forward this notice to your Risk Manager, OR Director, Purchasing Manager, BiomedicalEngineering staff, anesthesiologist and members of your medical staff who perform da Vinci Surgery procedures.2. Prior to use, open and inspect current supply of Fluorescence Imaging Procedure Kits. If kits include B. Braun stopcocks, discard stopcocks and replace with other stopcocks3. Continue to perform procedures using other stopcocks4. Continue to follow directions provided in Florescence Imaging Procedure Kit IFU (p/n 550976), If a sterile item in a Fluorescence Imaging Procedure Kit becomes compromised in any way, do not use the compromised item; replace it with a new item.5. Complete and return the attached Acknowledgment Form to Intuitive Surgicalusing the instructions provided.6. Please retain a copy of this notice near your Fluorescence Imaging Procedure Kit inventory.Questions or concerns should be directed to ClinicalSales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:" (800) 876-1310, Option 3 (6 AM to 5 PM PST) or mail: customersupportservicesupport@intusurg.com |  | 1200 | 1200 kits | US Nationwide Distribution | Terminated | November 02, 2015 | 54 | 1200 | Not_Computer
Z-2783-2015 | 72094 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound System with software versions VB10B and lower; |  |  | August 19, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | The ACUSON SC2000 ultrasound system considers uppercase/lowercase differences in the same patient name as unique patient instances when registered on the same ultrasound system. If these differences are not corrected at the time of registration, the system does not capture images or clips. | Software design | Siemens sent a Customer Safety Advisory Notification letter on August 19, 2015 by certified mail. The letter identified the affected product, problem and actions to be taken. The letter states that this issue will be resolved in a future software release. If you have any questions,please contact your customer care center for information regarding timelines and status. Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization who need to be aware of this issue. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. |  | 2039 | 2039 devices | Worldwide distribution. US Nationwide and the countries of Albania, Angola, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia Herzeg., Brazil, Brunei, Bulgaria, Canada, Colombia, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Georgia, Germany, Hungary,, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Lithuania, Malaysia, Maldives, Malta, Mauritius, Mexico, Netherlands, New Zealand, Nigeria, Norway, Oman, P.R. China, Pakistan, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed., San Marino, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Suriname, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Turkmenistan, U.A.E., Ukraine, United Kingdom, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 2039 | Not_Computer
Z-2731-2015 | 71905 | No | LOZ |  |  | Cardiovascular | PMA | Temporary Total Artificial Heart (TAH-t) with Freedom Driver System |  |  | August 06, 2015 | 2015 | SynCardia Systems Inc. | SynCardia Systems, Inc. is recalling Freedom Drivers due to a specific component of the drive mechanism may fail and cause the drive mechanism to stop pumping; resulting in the loss of life-sustaining function. | Component design/selection | SynCardia sent an Urgent Medical Device Recall letter dated August 6, 2015 to affected customers. The letter identified the purpose, reason for the voluntary recall, risk to health, how to recognize that the device may fail, actioms to be taken by the customer/user, product and distribution information and type of action by the company. Customer were instructed to completed and return Recall Acknowledgement and Receipt Form to SynCardia. A SynCardia Clinical Support Specialist or other SynCardia representative will assist customers with replacement product and returning affected product.Contact Information For QuestionsMary Pat SloanSenior Vice PresidentGlobal Certification and LogisticsSynCardia Systems, Inc1992 E Silverlake RdTucson, AZ 85713Direct 520-547-7470Toll-free 866-480-1122 Ext 1308Monday-Friday 7:30 AM to 5:00 PM Pacific/Arizona timeHotline 24/7 866-771-9437www.syncardia.com |  | 56 | 56 drivers (21 currently in use on patients) | Worldwide Distribution -- US Nationwide in the states of AR, NC, CA, OH, PA, TX, MI, and VA; and the countries of Germany, Turkey and France. | Terminated | November 04, 2015 | 90 | 56 | Not_Computer
Z-2778-2015 | 71695 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujifilm Medical Endoscope for the duodenum and upper G.I. tract. ED-530XT, ED-250XT5, ED-250XL5, ED-450XT5, and ED-450XL5 Endoscope Operation Manuals. |  |  | June 30, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | The suction button on the ED-530XT endoscope may become lodged (stuck) in the active position following the use of x-ray contrast agents during ERCP procedures. | Device Design | FujiFilm sent a Voluntary Field Correction/Verification Form dated June 30, 2015 to their affected customers. |  | 256 | 256 units | Distributed in the US and Canada. | Terminated | May 25, 2017 | 695 | 256 | Not_Computer
Z-2785-2015 | 71477 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Radiation Therapy Treatment Planning System software |  |  | May 29, 2015 | 2015 | RAYSEARCH LABORATORIES AB | An error caused by a floating point precision problem tends to cause a problem with the display of isodose lines for dose and dose difference and color tablesfor dose, dose difference and PET images | Software design | The firm, RaySearch Laboratories, distributed a "Field Safety Notice, Medical Device Correction #7106" notice dated May 29, 2015,and reply forms to their customers via email. The notice describes the product, problem and actions to be taken. The Customers were provided a workaround as a temporary fix. Customers were also instructed to educate all users, inspect your product and identify all installed units, confirm you have read and understood the notice by contacting Quality and Regulatory Affairs Manager; and complete and return the reply form via email to freddie.cardel@raysearchlabs.com or by fax to 888-501-7195. The solution will be resolved in the next version of RayStation, scheduled for release in June 2015. Customers were advised to share the notice to all those who need to be aware within the organizations. Customers can contact Quality and Regulatory Affairs Manager at +46 8 510 530 12 or via email at david.hedfors@raysearchlabs.com. |  | 20 | 20 devices | Worldwide Distribution: US (nationwide) to state of WA; and countries of: France, Sweden, Austria, Spain, Germany, Italy, UK, New Zealand. | Terminated | March 13, 2017 | 654 | 20 | Not_Computer
Z-2780-2015 | 72046 | No | DPT | Extravascular blood flow probe. | Cardiovascular | Cardiovascular | 510(k) | ThruPort Knot Pusher. This device is sold individually (Model KP1) and as part of the ThruPort Valve Placement Pack (Model VPP). packaged sterile and non-pyrogenic in a sealed, peel-type pouch. |  |  | August 26, 2015 | 2015 | Edwards Lifesciences, LLC | The configuration of the slot at the tip of the knot pusher may inhibit intended ease of use. Fraying or splitting of the suture could occur, which could compromise the integrity of the suture knot. | Process control | Edwards sent an Urgent-Product Recall - Action Required letter dated August 26, 2015, to all affected customers. Letters were sent via Fed-Ex. Please check your inventory for the affected Lot Numbers listed above and return the enclosed confirmation form. If your inventory includes affected product, please contact Edwards Lifesciences to obtain an RGA number and return the affected product.We apologize for the inconvenience caused by this action and appreciate your attention to this matter. If you have questions that have not been answered by this letter, please call Edwards Customer Service at (800) 424-3278 from 6:00AM  4:30PM Pacific Time. |  | 759 | 759 units | Worldwide Distribution - US (nationwide) Distribution and to the countries of : Canada, Japan and Europe. | Open, Classified |  |  | 759 | Not_Computer
Z-2796-2015 | 71933 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 3.5mm Non Locking Cortical Screw, in the PediLoc Tibial Plate System |  |  | July 28, 2015 | 2015 | OrthoPediatrics Corp | The 00-0903-2510, 3.5mm x 10mm Non Locking Cortical Screw, Lot # 150192-J package may contain a different screw. The screw in the package is a 00-0903-2526, 3.5mm x 26mm, and is etched correctly. | Packaging process control | OrthoPediatrics has initiated a Voluntary RECALL of Part # 00-0903-2510, 3.5mm Non Locking Cortical Screw, Lot # 150192-J, initially by contacting customers via telephone, and then by e-mailing notification letters on 7/28/2015. Customers returned the 5 mislabled screws, which were repacked with correct labeling. All recall activities were complete as of 8/14/2015. |  | 5 | 5 | Domestic: TN, IAForeign: EnglandVA/DOD: None | Terminated | March 17, 2016 | 233 | 5 | Not_Computer
Z-2794-2015 | 71976 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Dynamic Locking Screw System (DLS) 3.7mm &amp; Dynamic Locking Screw (DLS) 5.0mm; intended for use in long bone fractures, the fixation of osteopenic bone, the fixation of osteotomies, and for the fixation of non-unions and malunions. |  |  | August 11, 2015 | 2015 | Synthes (USA) Products LLC | DLS pin breakage during planned implant removal, after uneventful and successful healing of the fracture. This notice to additional consignees identified after recall Z-0390-2014, related to complaints of breakage located at the bottom of the screws and under the locking screw head. | Labeling False and Misleading | Synthes sent an Urgent Notices of Medical Device Recall letter, dated August 11, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to return affected devices and response forms.For further questions please call (610) 719-6500. |  | 40038 | 40038 | Nationwide Distribution | Open, Classified |  |  | 40038 | Not_Computer
Z-2793-2015 | 71957 | No | LHC | Powered heating unit. | Physical Medicine | General Hospital | 510(K) Exempt | Microtek Medical, Inc. ORS Fluid Warming and Slush Drapes, Model #'s ORS-100, ORS-110, ORS-130, ORS-300, ORS-301, ORS-320, ORS-321, ORS-325, ORS-330, ORS-331 |  |  | August 18, 2015 | 2015 | Ecolab Inc | The sterile packaging may contain small (less than 250 microns wide) channels in the seal of the pouch. If channels are present, the package integrity is compromised and there is a remote chance that a non-sterile drape may be used during a surgical procedure. | Package design/selection | Consignees were sent on 8/18/2015 an Ecolab "Urgent Medical Device Recall" letter dated August 18, 2015. The letter was addressed to Recall Coordinator or Risk Manager. The letter described the problem and the product involved in the recall. Advised consignees to quarantine all affected lots, and notify their customers of the recall which is to be carried out to users level. Requested consignees to complete and return the enclosed Recall Response Form. For questions they can contact Customer Service at 1-800-824-3027 or e-mail customerservice@microtekmed.com. |  | 5172816 | 5,172,816 units | Worldwide Distribution -- US, including DC and PR; and, the countries of SAUDI ARABIA, LEBANON, GERMANY, HONG KONG, SINGAPORE, CANADA, and JAPAN. | Terminated | August 09, 2017 | 722 | 5172816 | Not_Computer
Z-2791-2015 | 72010 | [6 empty] | Terumo Custom Cardiovascular Procedure Kit X-Coated |  |  | August 24, 2015 | 2015 | Terumo Cardiovascular Systems Corp | Custom Cardiovascular Procedure Kits contain recalled component Medivators HPH700 High Performance Hemoconcentrators | Nonconforming Material/Component | Terumo CVS issued an Urgent Medical Device Recall Notice dated August 24, 2015, to consignees via express mail. The letter advised users to immediately quarantine affected Cardiovascular Procedure Kits and stop using the Hemocor HPH700 Hemoconcentrator within the kit.Replacement sterile, stand-alone Hemocor HPH700 Hemoconcentrators will accompany the letter so that users may replace the affected emoconcentrators at the point of use.The option of replacing the hemoconcentrator at the point of use is offered to facilitate imminently scheduled cardiopulmonary bypass procedures. Users are requested to return the unused hemoconcentrators to Terumo CVS. Alternatively, users may return the entire Cardiovascular Procedure Kit to Terumo CVS for replacement. Customers with questions were instructed to call 763-553-3348. |  | 19 | 19 kits | Nationwide Distribution including CA, CT , FL , WI. | Terminated | February 05, 2016 | 165 | 124 | Not_Computer
Z-2797-2015 | 72079 | No | EGS | Dental handpiece and accessories. | Dental | Dental | 510(k) | Contra-Angles CA1:1L EVO 15 |  |  | August 27, 2015 | 2015 | Bien Air Dental SA | Possible loosening of the dental tool, which may lead to potential swallowing of the tool. | Other | Bien Air Dental sent an Urgent Medical Device Safety Alert on August 27, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease use of the mentioned device.The letter requests that customers inspect their inventory for the product those and affected within the scope, and immediately quarantine them. The firm states to fill out the attached REPLY FORM and returned it fax to +41 32 344 64 91.If customers have further distributed any of the subject devices to other organizations, please forward this note to them immediately, to stop using the affected devices and to place them in quarantine. All the devices at customer or end-user locations shall be also returned to Bien Air.Please return the affected products to Bien-Air Dental (the exact address for return will be communicated to you via Bien-Air Sales Manager). A Bien-Air Sales Manager will also be in contact with customers.Upon receipt of the completed REPLY FORM, a Bien-Air Dental Representative will contact customers to arrange for the return and/or replacement of any non-conforming units.Should customers require any further information or have any queries on the matter please contact Bien-Air Dental SAs Regulatory Affairs at: +41 32 344 6464 |  | 112 | 112 (U.S.) | Worldwide Distribution - US distribution to one distributor in FL. | Terminated | March 31, 2016 | 217 | 112 | Not_Computer
Z-2800-2015 | 72038 | No | GDX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PenBlade Safety Scalpel, for the incision of human skin and tissue, and to cut or trim sutures and catheters up to 6Fr. Models PB-M-10, PB-M-11P, PB-M-15 |  |  | August 18, 2015 | 2015 | Zien Medical Technologies, Inc | Cracks were discovered in the PET tray of a small population of sterile product. These cracks create a risk of loss of sterile barrier for the product. Sampling and root cause analysis indicated that the risk could extend throughout the lots manufactured from September 2014-July 2015. This led to the decision to recall the product. | Packaging process control | All consignees have been notified via email with an attached copy of the recall letter. Telephone calls with be made to all customers and distributors that do not respond directly to email notification to ensure compliance to the recall. |  | 28284 | 28,284 | Worldwide distribution to U.S. nationwide, Germany and England. | Terminated | December 09, 2015 | 113 | 28284 | Not_Computer
Z-2802-2015 | 71842 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | PET Discovery LS |  |  | July 17, 2015 | 2015 | GE Healthcare | Potential for expulsion of the Power Supply on Lightspeed CT scanner or Discovery ST PET CT or Discovery LS PET CT scanner. | Process control | GE Healthcare sent an "Urgent Medical Device Correction" letter to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. GE Healthcare will correct all affected products at no cost to you. A GE Healthcare representative will contact you to arrange for the correction. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. Ref# 25463 |  | 6 | 6 | Worldwide Distribution - US (nationwide) and to the countries of : ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, HUNGRY, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, MEXICO, NORWAY, POLAND, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN and UNITED KINGDOM. | Terminated | March 08, 2016 | 235 | 440 | Not_Computer
Z-2798-2015 | 72008 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Regard Custom Surgical Pack, Item Number: 800628; Sterile, CV0805 - CABG B. Used in cardiovascular surgical procedures. |  |  | August 19, 2015 | 2015 | Resource Optimization &amp; Innovation Llc | Surgical kits contain a surgical scrub product which may be chemically contaminated and lack sterility assurance. | Material/Component Contamination | On 8/19/2015, the recalling firm sent its two consignees an email with the recall notice attached. The notice notified their customers that they were affected by the recalling firm's supplier's recall. The notice instructed the customers to dispose of the recalled component and acknowledge receipt of the recall notice. |  | 24 | 24 kits | Distributed to one customer in Lake Charles, LA and one customer in Spartanburg, SC. | Terminated | April 05, 2017 | 595 | 91 | Not_Computer
Z-2807-2015 | 70858 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVIS EXERA II Duodenovideoscope OLYMPUS TJF TypeQ180V. |  |  | March 26, 2015 | 2015 | Olympus America Inc. | New and validated reprocessing procedures have been issued for the Olympus TJF-Q180V duodenoscope ("TJF-Q180V"), consisting of revised manual cleaning and high level disinfection procedures. | Labeling design | The firm, Olympus, sent an "Urgent Safety Notification" dated March 26, 2015, to direct consignees/customers. The notification described the product, problem and actions to be taken. The notification contained new reprocessing procedures. The customers were instructed to immediately implement new precleaning and high level disinfection procedure; ensure that all personnel involved in the reprocessing are completely knowledgeable and thoroughly trained on the new reprocessing instructions; complete and return the enclosed questionnaire via fax to (484) 896-7128 and implement the new cleaning instructions when the new brushes arrive on/about May 8, 2015.If you have question, contact our Technical Assistance Center (TAC) at 1-800-848-9024, option 1, Monday-Friday between 7 AM EST - 8 PM EST. For additional information on this matter, contact V.P., Regulatory/Clinical Affairs &amp; Quality Assurance at (484) 896-5688 or by email at laura.storms@olympus.com. |  | 6109 | 6109 | Worldwide Distribution: US (nationwide) and countries of: Canada and Mexico. | Open, Classified |  |  | 6109 | Not_Computer
Z-2804-2015 | 71928 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | The Leica ASP6025 Tissue Processor. Pathology and Hematology Departments. |  |  | July 27, 2015 | 2015 | Leica Microsystems, Inc. | The ASP6025 Tissue Processor instrument is incorrectly getting into a state of bottle empty during retort filling, after passing the fill lever sensor 2 and prior to reaching levelsensor 3, thus not completing a 5 litre fill. | Nonconforming Material/Component | Leica sent an Urgent Medical Device Recall Notification letter dated August 11, 2015, to all affected customers via USPS 1st class mail ..Thel letter advised users to 1) do not use the instrument in 3 basket (5L) mode. Instead it advises the customer to change settings to use only the 2 basket (3.8L) mode only. 2)Prompt acknowledgement is requested by the customer 3) when a permanent solution is found a Leiica representative will contact the customer. 4) The Medical Device Recall Notification needs to be passed on to all those who need to beaware within the customers organisation or to any organisation where the potentially affected devices have been transferred.The customer is asked to sign the Medical Device Recall Notification Response Form to confirm that they have received and understand the Medical Device Recall Notification.The customer is requested to confirm receipt of the notice by signing and dating the Acknowledgement Form and faxing it back to Leica at 1-847-236-3747, or scan and email it to LMGRA@leicamicrosystems. com |  | 573 | U.S. Quantity Distributed: 163 systems; Foreign Quantity Diistributed: 410 systems. | Worldwide Distribution - US including AR, AZ, CA, CT, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MN, MO, MT, NC, NE, NJ, NM, NV, NY, OH,PA, SC, TN, TX, VA, WA, WI, WV and District of Columbia. Foreign distribution to United Arab Emirates, Australia, Belgium, Brazil, Canada, Switzerland, Chile, China, Colombia, Czech Republic , Germany, Spain, France, England, Georgia, Indonesia, Ireland, Israel, India, Italy, Jordan, Japan, Korea, Kuwait, Lithuania, Myanmar, Mexico, Malaysia, Netherland, Norway, Peru, Philippines, Poland, Qatar, Reunion, Romania, Serbia, Saudi Arabia, Singapore, Slovenia, Thailand, Turkey, Taiwan, and South Africa. | Terminated | January 05, 2017 | 528 | 573 | Not_Computer
Z-2806-2015 | 71706 | No | FET | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujifilm Processor EPX 2500 High Definition Endoscopy; (Sales Brochure). |  |  | June 30, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | The EPX-2500 Operation Manual and sales brochure incorrectly identify that the ED-530XT endoscope is meant to be used with the EPX-2500 Processor. The Operation Manual and sales brochure state that when the ED-530XT endoscope is used with the 2500 processor, a "super image" is displayed when it is actually a "standard image" that is displayed. | Device Design | Fujifilm Medical Systems USA, Inc. initiated this recall by sending a Voluntary Field Correction Letter and Tracking/Verification Form dated June 30, 2013 to their affected customers. |  |  | multiple | Nationwide US and Canada | Terminated | May 25, 2017 | 695 |  | Not_Computer
Z-2809-2015 | 71819 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | MicroScan LabPro Information Manager System, Version 1.0 to Version 4.11, Catalog No. 10714149, 10714150. Intended to manage both microbial identification (ID) and antimicrobial agent susceptibility testing (AST). |  |  | July 17, 2015 | 2015 | Beckman Coulter Inc. | The software incorrectly allows the operator to manually edit the carbohydrate substrates when manually reading dried overnight gram negative panels with an ID Hold status. | Unknown/Undetermined by firm | An Urgent Medical Device Recall Letter dated 7/17/2015 was sent to all customers who purchased the MicroScan LabPro Information Manager System, Version 1.0 to Version 4.11. Beckman Coulter is recalling the MicroScan LabPro Information Manager System because the software incorrectly allows the operator to manually edit the carbohydrate substrates when manually reading dried overnight gram negative panels with an ID Hold status. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Customer Support Representative: http:www.beckmancoulter.com, (800) 677-7226 in US and Canada. Customers outside of US and Canada, are instructed to contact their local Beckman Coulter representative. |  | 2456 | 2,456 units total (1,034 units total) | Worldwide distribution: United States (including Puerto Rico), Bahamas, Belgium, Canada, Chile, China, Colombia, France, French Guiana, Germany, Greece, Guadeloupe, India, Israel, Italy, Japan, Kazakhstan, Lithuania, Mexico, Netherlands, Panama, Paraguay, Poland, Portugal, Romania, Russia, Singapore, South Africa, Spain, Turkey, , US Virgin Islands, and Northern Mariana Islands, Uruguay. | Open, Classified |  |  | 2456 | Not_Computer
Z-2811-2015 | 72139 | No | GXN | Preformed nonalterable cranioplasty plate. | Neurology | Neurology | 510(k) | Synthes Cranial Flap Tube Clamp and Crimping Device |  |  | August 31, 2015 | 2015 | Synthes (USA) Products LLC | The Synthes Cranial Flap Tube Clamp and Crimping Device for Cranial Tube Clamp had been labeled MR Safe although they do not have the testing protocols currently required to designate them as MR Safe. | Labeling False and Misleading | DePuy Synthes sent an URGENT:FIELD SAFETY NOTIFICATION / MEDICAL DEVICE LABELING CORRECTION letter, dated August 31, 2015 to affected customers. The letter identified affected product, problem and actions to be taken. Customers were instructed to:"Update and review the package insert (GP1345-D)." Discard outdated package insert noted in the table above." Forward this Field Safety Notification to anyone in your facility that needs to be informed, especially personnel that conduct MR testing." If the package insert has been forwarded to another facility, contact that facility." Complete the attached Verification Section (page 3 of this notification). Please include your name, title, address, telephone number and signature in the spaces provided. Send a copy of the completed Verification Section by, Fax: 877-884-7821 or Scan/email: Synthes3911@stericycle.com" If the Verification Form is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual in page three (3) of this notification." Maintain a copy of this notice.For questions call 610-719-5450. |  | 105882 | 105882 | US Nationwide and internationally, including Canada. | Open, Classified |  |  | 105882 | Not_Computer
Z-2808-2015 | 71757 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Marketing brochures for the PowerPICC and PowerPICC SOLO catheters |  |  | July 20, 2015 | 2015 | Bard Access Systems | Bard Access Systems is recalling marketing brochures for the PowerPICCÂż and PowerPICCÂż SOLOÂż catheters due to incorrect indication for use contained in the brochure. | Labeling False and Misleading | Customer letter notification with reply form will be sent via Federal Express with delivery confirmation for US consignees on 7/20/15 informing them of the incorrect indication for use contained in the brochure. |  | 1811 | 1811 | Distributed to domestic and foreign conference attendees. | Open, Classified |  |  | 1811 | Not_Computer
Z-2810-2015 | 71754 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident 10 degrees X3 Insert 36 mm ID |  |  | June 25, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | it was discovered that item 623-10-36I Lot 3Y0KHP, although meeting all device specifications, was determined to be unacceptable due to a documentation issue. However the parts were inadvertently shipped to the field. | Nonconforming Material/Component | Stryker Orthopaedics notified their branches via email on 6/25/2015. Notification letters and Product Accountability Forms were mailed to Stryker Branches/Agencies on 817/2015 and hospital risk management on 8/18/2015. |  | 16 | 16 units | Distributed in WI. | Open, Classified |  |  | 16 | Not_Computer
Z-2817-2015 | 72071 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CD4 (SK3), FITC; |  |  | August 27, 2015 | 2015 | Becton, Dickinson and Company, BD Biosciences | CD4 FITC label contains an error in the Spanish small text - it states CD8 FITC. | Under Investigation by firm | Urgent Product Recall notification letters were sent to each customer on August 27, 2015 by certified mail. |  | 128 | 128 units | Worldwide Distribution. US nationwide (TX, CA, NC, OH, SD, PA, CT, UT, MD, TX, GA), Argentina, Taiwan, Australia, Japan, China, New Zealand, Chile and Colombia. | Open, Classified |  |  | 128 | Not_Computer
Z-2815-2015 | 72044 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens RAPIDPoint 500 Measurement Cartridge Lactate (250) tests |  |  | August 24, 2015 | 2015 | Siemens Healthcare Diagnostics Inc | RAPIDPoint 500 measurement cartridge serial numbers 2517102517 through 2519000012, negative sodium bias on both QC and patient results | Process control | Siemens Healthcare sent an Urgent Field Safety Notice letter dated August 2015 to all affected customers. The notice identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Field Correction Effectiveness Check Form attached. For questions contact your Siemens Customer Care Center or ther local Siemens technical support representative. |  | 982 | 982 cartridges | Worldwide Distribution - US Nationwide and the countries of Canada, Albania, Austria, Australia, Bosnia Herzeg., Bangladesh, Belgium, Brazil, Switzerland, Chile, China, Colombia, Czech Republic, Germany, Denmark, Denmark, Estonia, Spain, Finland, France, Great Britain, Georgia, Greece, Croatia, Hungary, Ireland, India, Italy, Lithuania, Luxembourg, Latvia, Macedonia, Malaysia, Netherlands, Norway, New Zealand, Poland, Portugal, Paraguay, Romania, Sweden, Singapore, Slovenia, Slovakia, Turkey, Taiwan, South Africa, Unit.Arab Emirates | Terminated | May 23, 2016 | 273 | 4323 | Not_Computer
Z-2812-2015 | 71915 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE, ONCOR Avantgarde, ONCOR Impression Plus, ONCOR Expression, PRIMUS, PRIMUS Plus, Syngo RT Therapist; therapeutic treatment of cancer |  |  | August 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc | A software fix has been released to prevent automatic movement resulting in a collision safety risk for patients. | Software design (manufacturing process) | Siemens sent Notification letters to affected customers. The letter identified the affected product, problem and actions to be taken. The letter advise customers that the software fix was rolled out to address multiple safety issues. |  | 26 | 26 | Worldwide Distribution and US Nationwide | Open, Classified |  |  | 26 | Not_Computer
Z-2819-2015 | 72035 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer, Reprocessed Zimmer Tourniquet Cuffs,34" Single Hose Single Bladder Brown (PLC), 34", Model (Reprocessor) Number: ZIM60-7075-106, Reprocessed by Sterilmed. |  |  | August 24, 2015 | 2015 | STERILMED, INC. | Sterilmed, Inc. is recalling its reprocessed Zimmer Tourniquet Cuff, Part #ZIM60-7075-106, lot number 1846885 due to mislabeling. The affected reprocessed product is labeled as single hose devices but the devices inside the packaging are dual hose devices. | Labeling False and Misleading | Sterilmed sent an "Urgent Medical Device Recall" letter dated August 24 2015. The letter was addressed to Materials Manger, Risk Manger and Operating Room Supervisor. The letter identified the affected product, problem and actions to be taken. Customers are requested to return all affected product, sign and return the enclosed Customer Response Questionnaire and the Customer Response Form per the instructions on the Forms. For questions contact Sterilmed's Customer Care Team at 888-541-0078. |  | 43 | 43 | US Nationwide Distribution in the states of CA, KY, MI, OH, UT and WI. | Terminated | March 11, 2016 | 200 | 43 | Not_Computer
Z-2818-2015 | 72051 | No | ODG | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVIS EXERA II Gastrovideoscope OLYMPUS GF type UCT180; designed for endoscopic real-time ultrasound imaging, ultrasound guided needle aspiration and other endoscopic procedures within the upper gastrointestinal tract and surrounding organs. |  |  | August 21, 2015 | 2015 | Olympus Corporation of the Americas | Insufficient information for specific cleaning accessories used for the cleaning and reprocessing of the GF-UCT180 ultrasound endoscope | Device Design | Olympus sent an Urgent Device Safety Information letter dated August 2015, to all affected customers. Corrected instructions were provided with the letter and users were asked to begin implementing them immediately. Customers were asked to indicate on the enclosed questionnaire and fax the completed from to 484-896-7128. Customers with questions were instructed to call 484-896-5688. |  | 1014 | 1014 | Nationwide Distribution | Terminated | May 08, 2017 | 626 | 1014 | Not_Computer
Z-2784-2015 | 71995 | Yes | NVZ |  |  | Cardiovascular | PMA | Merlin@home RF Remote Monitoring Transmitter Model EX1150. |  |  | August 18, 2015 | 2015 | St Jude Medical Cardiac Rhythm Management Division | Some devices exhibit backup VVI operation due to a backup reset. | Software change control | The notification material is in the form of a Physician Letter was delivered to physicians by FedEx overnight service. The mailing went out on August 17, 2015 for receipt on August 18, 2015.The firm's notification letter identifies clinical implications.Clinical ImplicationsIn the event that an Ellipse, Fortify Assura, Unify Assura, or Quadra Assura ICD enters the back-up safety mode, the nominal operational settings will be VVI pacing mode, 67 ppm, 5.0v/0.6ms with bipolar pacing output and defibrillation settings of aVF detection rate of 146 bpm and 36J high voltage therapy. In the event an Assurity or Allure pacemaker enters backup mode, it will have output settings of VVI pacing mode, 67 ppm, 5.0v/0.6ms with unipolar pacing.This anomaly can only occur when the patient is being actively monitored by a Merlin@home bedside transmitter. If a device enters backup mode, the ICD will deliver a patient vibratory alert and the pacemaker will deliver a patient audible alert. If backup operation is encountered, St. Jude Medical Technical Services (1-800-722-3774) can assist with non-invasively restoring the device to normal operation.The firm's notification also idenfities recommendations and mitigations.Recommendations and MitigationsOnce available, a Merlin@home transmitter software update will be performed automatically over its telephone, broadband or cellular connection without requiring any action from you or your patients. No changes to the patients remote or in-clinic follow up schedules are required. In the event a patients device reverts to back-up mode, the firm recommends bringing the patient back in the clinic to clear the condition and return the device to full functionality.If you have any questions or concerns contact the local St. Jude Medical representative or St. Jude Medical's Technical Services Department. |  | 259214 | 259,214 transmitters (U.S. distribution) 23,604 (foreign distribution) | Worldwide Distribution-US (nationwide) and foreign distribution to the following: Japan, France, Austria, Canada, Italy, Spain, Netherlands, South Africa, Great Britain, Australia, Sweden, Belgium, Finland, Germany, Portugal, Singapore, Norway, Poland, Ireland, Czech Republic. Denmark, Hong Kong, Saudi Arabia, Hungary, Switzerland. | Open, Classified |  |  | 259214 | Not_Computer
Z-2743-2015 | 71925 | No | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PROFEMUR PLUS CoCr MODULAR NECK, REF: PHAC1254, SIZE LONG, ORIENTATION 8 degree VAR/VAL. 12/14 SLT Taper. |  |  | August 04, 2015 | 2015 | MicroPort Orthopedics, Inc. | Unexpected rate of postoperative fractures resulting in the need for revision surgery. | Under Investigation by firm | Micro Port Orthopedics sent a Voluntary Device Product Recall letter dated August 7, 2015, via Fed Ex to all affected customers. The letter identified the product, the problem, and action to be taken by the customer. Customers were instructed to locate and return effected product immediately. Customers with questions were instructed to call 1-866-872-0211. Customers were also asked to return the enclosed form immediately even if they do not have any affected product to return. |  | 10489 | 10,489 units | Worldwide Distribution - US (nationwide) and Internationally to AT, BE, CA, CN, CZ, DE, ES, FI, FR, GB, IT, KR, LT, LV, NL, NO, PE, PL, PT, RU, and ST. | Open, Classified |  |  | 10489 | Not_Computer
Z-2820-2015 | 72132 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Guide Wires for orthopedic procedures. |  |  | August 27, 2015 | 2015 | Zimmer, Inc. | Product labeling of these sterile guide wires (external carton label and patient record label) states a 10-year expiration date. Package testing supports a 5-year expiration date for these same guide wires. | Package design/selection | Zimmer sent an "URGENT MEDICAL DEVICE RECALL" notifications dated August 27, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Your Responsibilities:1.Review the notification and ensure affected personnel are aware of the contents.2.Locate part number and lot number combinations listed in attachment 1 and quarantine the devices immediately.3.Carry out a physical count of all affected product in your territory, including hospitals in your territory with directsales as well as consignment, and complete the Inventory Return Certification of Acknowledgement (Attachment 2). Email a completed copy of Attachment 2 to corporatequality.postmarket@zimmerbiomet.com.4.Return the recalled product along with the completed Attachment 2. Clearly mark the outside carton of each product return shipment made as Recall. 5.Please notify Zimmer Biomet of any hospitals and surgeons that have consumed the product in the past 12 months that are not included in either the distribution history or inventory system data information. Supply the information forany hospitals and surgeons that you have additionally identified by using the provided spreadsheet Additional Account template. The template, and the list of known consignees from the distribution history and inventory systems data, will be emailed to you for completion and return to: corporatequality.postmarket@zimmerbiomet.com.6.If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm ESTFor further questions about this recall, please call ( 1-800-447-5633). |  | 112568 | 112,568 units | Worldwide Distribution - US (nationwide) including Puerto Rico and Virgin Islands., and to the countries of Canada, Brazil, Algeria, Mexico, Austria, Taiwan, Belgium, Hong Kong, Croatia, Singapore, Denmark, Bahamas, Egypt, Panama, Finland, Dominican Republic, France, Guatemala, Germany, Australia, Ireland, Japan,Italy, Haiti, Jordan, Malaysia, Kenya, India, Lebanon, Luxembourg, Malta, Mauritius, Morocco, Netherlands, Norway, Poland, Portugal, Russia, Saudi Arabia, Serbia, South Africa, Spain, Sweden, Switzerland, United Kingdom and United Arab Emirates. | Terminated | March 15, 2016 | 201 | 112568 | Not_Computer
Z-2825-2015 | 72086 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow Kits with Medtronic Covidien 0.9% Sodium Chloride Flush Syringes |  |  | August 27, 2015 | 2015 | Arrow International Inc | Manufacturer of 0.9% Sodium Chloride Flush Syringes recalled due to compromised sterility of the outer barrel, Arrow purchased recalled items that were packaged with certain products. | Process control | An urgent recall notification letter, dated August 27, 2015, was sent to direct consignees to inform them of the issue and provided instructions for removing the syringes from the kit and returning product along with the response form. |  | 73714 | 73714 | Nationwide and internationally. | Open, Classified |  |  | 73714 | Not_Computer
Z-2823-2015 | 72183 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | P/N: 9735023, DRIVER 9735023 SOLERA 5.5/6.0 MAS; P/N: 9735024, DRIVER 9735024 SOLERA 5.5/6.0 MAS CAN; |  |  | September 14, 2015 | 2015 | Medtronic Navigation, Inc. | Specific use conditions can increase the possibility of a screwdriver tip deforming or breaking during surgery. | Labeling design | Medtronic sent an Urgent Medical Device Correction Notification letter dated September 14, 2015, to all affected consignees via FedEx. Revised Instructions for Use were provided to add additional warnings related to the careful inspection of the instruments and the importance of knowledge of operating procedures, patient selection, and product information. Consignees with questions were instructed to to call the local Medtronic Navigation Technical Services Representative or call 1-800-995-9709. |  | 2346 | 2,346 | Worldwide Distribution - US (nationwide) and Internationally to Canada; United Arab Emirates; Australia; Belgium; Switzerland; China; Germany; Denmark; Spain; Finland; France; Great Britain; Greece; India; Israel; Italy; Japan; Netherlands; New Zealand; Poland; Portugal; Sweden; Singapore; Thailand; Slovakia; Turkey; and South Africa. | Terminated | December 03, 2015 | 80 | 2346 | Not_Computer
Z-2822-2015 | 72106 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ASP-1021LipoTower with Aspiration Pump and Tumescent Pump. |  |  | June 12, 2015 | 2015 | MicroAire Surgical Instruments, LLC | There is potential risk of electrical shock if lead wires inside the device become detached and the ground connection fails. | Device Design | The firm has issued FDA recall letters to all consignees. The letter have been sent out by UPS to expedite the delivery. |  | 1189 | 168, ASP-1021 LipoTower with Aspiration Pump and Tumescent Pump | Nationwide Distribution. | Open, Classified |  |  | 1192 | Not_Computer
Z-0004-2016 | 71858 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Philips GoSafe Mobile Help Button; 7000MHB. |  |  | July 23, 2015 | 2015 | Lifeline Systems, Incorporated | Irregular battery disconnection resulted in issues with Mobile Health Button (MHB) Charging. | Process control | Customers were also notified of the recall via an Urgent: Medical Device Recall sent, by US Mail, between 8/19/2015 and 8/20/2015. The letter described the reason for the recall and how to recognize the device failure. The customer is asked to do the following: 1. Review and confirm the attached list of affected units for your program. Contact the affected subscribers listed and notify them that a service visit will be scheduled so that their units can be replaced within the next few weeks. 2. If any subscribers report any issues with their MHB, contact Philips for an expedited unit replacement and instruct them to discontinue use. 3. Philips Customer Care will follow-up with customers via telephone to ensure that they have received and understood the notification. During this follow-up call, we will also confirm arrangements for unit replacements. |  | 8616 | 8,616 devices | Nationwide Distribution -- AK, AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KS, KY, MA, MD, MI, MN, MT, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TX, UT, VA, VT, WA, and WI. | Open, Classified |  |  | 8616 | Not_Computer
Z-0001-2016 | 71987 | No | FDI | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Histolock Resection Device, US Endoscopy. |  |  | July 14, 2015 | 2015 | US Endoscopy Group Inc | The company has determined the sterility cannot be assured for the affected lot. | Employee error | The firm, US Endoscopy, sent an "URGENT: HISTOLOCK" RESECTION DEVICE RECALL" letter and response form to their customers on July 14, 2015. The letter described the product, problem and action to be taken. The customers were instructed to complete the response form and return to your US Endoscopy Product Specialist and retain a copy of the letter for your records.If you have any questions, contact Sr. Manager, Reg. Affairs &amp; Quality Compliance at 440-358-6251 or email at: ccochara@usendoscopy.com. |  | 41 | 41 units | US Distribution to states of: AL, AR, CA, IL, IN, MS, OH, OK &amp; TN. | Open, Classified |  |  | 41 | Not_Computer
Z-0003-2016 | 72167 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | SOUNDSTAR eco 8F and 10F Diagnostic Ultrasound Catheters. For intra-cardiac and intra-luminal visualization of cardiac and great vessel anatomy and physiology as well as visualization of other devices in the heart. |  |  | September 10, 2015 | 2015 | Biosense Webster, Inc. | Image disappeared from the cardiac ultrasound system when the CARTO 3 EP Navigation System needed restarting while the patient was experiencing pericardial effusion. Affects the CARTOSOUND Module of the CARTO 3 EP Navigation System when used with the SOUNDSTAR eco 8F and 10F Diagnostic Ultrasound Catheters. New precautions added. | Device Design | A Customer Notification Letter and Acknowledgement Form was distributed to customers on 09/10/15. One letter and form was sent via express mail, addressed to the Risk Management Office and one additional letter and form was sent by the firm's account representative to the Electrophysiology (EP/Cardiology Lab directly at the customer site.The firm's notification letter stated that they want to emphasize the following statement "The intra-cardiac ultrasound image will disappear if the CARTO 3 EP Navigation System power is disrupted and this may present a safety issue if the EP is using the ultrasound to monitor the patien during EP procedure. The ultrasound image will not reappear until the CARTO 3 System is restored."The firm will be updating the product labeling to further reinforce the precautionary statement.The firm requests that customers take the following actions:1. Read the Field Safety Notification carefully.2. Pass on this notification to anyone in the facility that needs to be informed of this issue, including appropriate clinical personnel involved in the use of SOUNDSTARÂż eco Catheters.3. Review, complete, sign and return the attached Acknowledgement Form in accordance with the instructions on the form. 4. Maintain a copy of this letter with the product.5. Maintain awareness of the Field Safety NotificationFor any questions the firm states to contact the BWI sales representative. |  | 1035 | 1035 (U.S.) | Worldwide Distribution. US nationwide, Italy, Canada, Japan, China, Singapore, Colombia, England, Russia, Spain, France, Germany, Finland, Ireland, Australia, Sweden, Hungary, India, Brazil, Slovenia, Netherlands, Mexico, and South Africa. | Open, Classified |  |  | 1035 | Not_Computer
Z-2826-2015 | 71983 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Ferno iNX Wheeled Stretcher, Integrated Patient Transport &amp; Loading System |  |  | August 03, 2015 | 2015 | Ferno-Washington Inc | Hole elongation in the telescoping load frame channel. A hole elongation could affect the ability of the retention pin to stay seated which is necessary to preclude movement of the telescoping load frame. | Device Design | Ferno began calling their customers on 8/3/2015. The firm informed their customers of the following: To provide optimal Customer Experience and minimal disruption to your service, the iNX field action involves the following steps:" Exchange of your current iNX with a new iNX;" Participation in a brief training session; and" Inspection of your Ferno Model 175 antler and rail fastening system as requested by the Dept. of Transportation.As you await delivery of your new iNX, it is important you ensure proper function of the fastening system and iNX telescoping load frame. |  | 72 | 72 units | Worldwide Distribution - US Nationwide in the states of: CO, IA, IL, IN, KY, MD, MN, MO, NC, NJ, NY, OH, OK, OR, TX, VA, WA &amp; WY and the countries of CANADA, AUSTRALIA, UNITED KINGDOM &amp; NEW ZEALAND. | Terminated | January 20, 2016 | 170 | 72 | Not_Computer
Z-2795-2015 | 71996 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Disposable Reflective Marker Spheres |  |  | August 18, 2015 | 2015 | Brainlab AG | Disposable Reflective Marker Spheres (DRMS) for Brainlab Image Guided Surgery (IGS) Systems: DRMS spheres may separate at the mid-point where the two halves of the sphere are sealed together. | Process control | The firm, Brainlab AG, sent out a" FIELD SAFETY NOTICE/ PRODUCT NOTIFICATION" dated 7/23/2015 to its customers. The firm sent this letter to mitigate the potential effects of the affected Disposable Reflective Marker Spheres (DRMS) when used with Brainlab Image Guided Surgery (IGS) systems. The firm advised the medical users that they should take the following steps prior to each procedure using Brainlab DRMS spheres for Brainlab IGS systems: -Prior to opening the package containing the spheres, visually inspect the spheres in the blister pack for any signs of separation between the two halves of the sphere. If there is any indication that the spheres may have separated, use a different blister pack for the procedure and repeat this step with the new pack. Discard any and all affected spheres and inform your Brainlab support representative about the occurrence of the issue. If there is no visual sign of sphere separation, proceed to the next step. -Prior to the start of the procedure, screw each DRMS sphere onto the tool and tighten. Upon tightening, inspect again to determine whether any separation between the two halves is evident. If separation is evident, remove and discard the affected sphere. Replace your surgical gloves, then replace the separated sphere with a new sphere, inspecting it as above. If no separation is evident, proceed with the clinical procedure. -Be sure to discard and replace your surgical gloves any time that a sphere with evident separation is found upon affixing the sphere to the tool, and every time after a separated sphere was touched.-Caution: Do not affix the sphere to the tool over the surgical area to minimize the risk of a sphere, or any portion thereof, falling onto the patient or into the surgical wound. Also avoid affixing the sphere over other sterile areas (such as surgical trays) to minimize potential contamination in case of sphere separation. -Reminder: In general, do not use defective or deformed marker |  | 369659 | For 28 US consignees, 9,900 Disposable reflective marker spheres ***For 451 non-US consignees, 359,280 Disposable reflective marker spheres | Worldwide Distribution: US (nationwide) including states of: TX, CA, TN, NY, LA, IL, MA, PA, OH, AZ, ID, CT, MI, and KY; and countries of: Algeria, Argentina, Australia, Bahrain, Belarus, Bolivia, Brazil, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, Germany, Hong Kong, Hungary, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Malaysia, Malta, Mexico, Moldova, Morocco, Myanmar, Netherlands, Philippines, Poland, Portugal, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, , Venezuela, and Vietnam. | Open, Classified |  |  | 369659 | Not_Computer
Z-2766-2015 | 72045 | No | NGV |  |  | Cardiovascular | PMA | Boston Scientific, WATCHMAN Access System Access Sheath with Dilator, 14F Single curve( M635TU10060, TU1006) , 14F Double curve (M635TU20060, TU2006) , &amp; 14F Anterior curve (M635TU40060, TU4006). For cardiac use. |  |  | August 25, 2015 | 2015 | Boston Scientific Corporation | Cross-threading of the hemostasis valve may occur if it is tightened with the dilator in place, possibly preventing subsequent sealing of the valve when desired. This Field Correction reinforces the Directions for Use (DFU) and gives further guidance on correct use of the hemostasis valve to avoid cross-threading and to securely seal the valve, minimizing potential for undesired blood leakage. | Use error | Consignees were sent a Boston Scientific "Urgent Field Correction" letter dated August 25, 2015. The letter was addressed to Materials manager / Field Action Contact. The letter described the problem and the product affected. The Field Correction reinforces existing Directions for USE (DFU) and provides further guidance regarding the correct us of the hemostatis valve in order to avoid cross-threading and to securely seal the valve, minimizing the poential for undesirable blood leakage. The consignees were informed of the action required which stated that no product is required to be returned and advised to consignees that they must complete the Customer Acknowledgement Form and return it to Boston Scientific (MapleGroveFieldAction Center@bsci.com or Fax to Field action Center 1-866-213-1806). The letter also provided a section of the Updated Directions for Use which were highlighted in blue. If the consignee is a distributor the Field Correction is to the hospital level. For questions, customers may contact their Local Sales Representative. |  | 29169 | 29,169 units total | Worldwide Distribution-US (nationwide) and the countries of Argentina, Australia, Austria, Belgium, Canada, Chile, China, Colombia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Great Britain, Greece, Hong Kong, Hungary, Indonesia, Iran, Ireland, Israel, Italy, Kazakhstan, Kuwait, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Oman, Panama, Philippines, Poland, Portugal, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, White Russia. | Open, Classified |  |  | 29169 | Not_Computer
Z-0022-2016 | 72245 | No | DTB | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | PMA | Medtronic Medical Device Identification Card for SureScan pacemaker patients. The front of this card identifies the implanted pacing system components currently registered to the patient. The back of the card lists patient's following physician as well as specific messaging regarding the MR conditional status of the device components, where applicable. This card is provided to the patient by Medtronic after device registration. |  |  | September 15, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | Some Medical Device Identification Cards provided to SureScan pacemaker patients indicate they have a complete Magnetic Resonance (MR) Conditional system, when in fact, not all of their implanted leads have been FDA approved as MR Conditional. | Process control | On September 15, 2015, Medtronic began mailing an Important Medical Device Identification Card Notification letter, including a corrected Medical Device Identification Card, to impacted patients. The letter described the issue, provided a new corrected Identification Card, asked that the old card be destroyed immediately, and to begin using the new card. Patients can contact Medtronic Patient Services at 1-800-551-5544, option 1, followed by option 3, Monday - Friday, 8am - 5pm Central time. |  | 777 | 777 cards | US nationwide distribution. | Terminated | April 14, 2016 | 212 | 777 | Not_Computer
Z-0024-2016 | 72238 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Periarticular Insertion Handle for 4.5mm LCP Proximal Tibia Plates and Periarticular Aiming Arm for 4.5mm LCP Proximal Tibia Plates |  |  | September 14, 2015 | 2015 | Synthes (USA) Products LLC | Affected lots of Periarticular Aiming Arms and Insertion Handles (Part Numbers 03.120.001, 03.120.002, 03.120.004 and 03.120.005) for the 4.5 mm LCP Proximal Tibia Plates are not able to be secured to the subsequent versions of these devices as intended. | Device Design | An urgent notice of medical device recall, dated September 14, 2015, was sent to end users and sales consultants to alert them to the issue and potential risks. Return of the completed response form and the affected devide was also requested. |  | 172 | 172 | Nationally | Open, Classified |  |  | 172 | Not_Computer
Z-0023-2016 | 72212 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS IW by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources. Images and data can be stored, communicated, processed and displayed within the system or across computer networks at distributed locations. |  |  | September 16, 2015 | 2015 | GE Healthcare | Images may be missing when a system parameter MapRoute is set to a value greater than 1. | Software Design Change | The firm, GE Healthcare, an "Urgent Medical Device Correction" by mail or hand delivered to all sites. The letter instructs the customer to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. Until GE Healthcare conducts remote inspection of the systems "MapRoute" parameter value. If the value is found to be greater than the recommended setting of 1, GE will reset this system parameter to a value of 1. The letter informs the user facilities not to make any changes to this parameter value.Additionally, until a product modification is available to correct this issue, the following actions are recommended. The firm is recommending user facilities to: 1. Utilize the image count within the QC process to alert the user of a discrepancy in the number of transmitted images from the modality to the number of images available in the PACS IW viewer.a. If a discrepancy is identified, attempt to retransmit the exam to PACS.b. If retransmission is unsuccessful, contact your GE Healthcare Service representative for assistance and instructions as to the impacted exam. Urgent cases impacted by this issue should be interpreted at the modality.2. If DICOM storage commit is configured and in use, no commit notification is sent to the modality for images that were impacted by this issue.GE Healthcare will correct all affected systems by providing a software update at no cost to the user facilities. If the user facilities have questions, they are encouraged to communicate questions or concerns regarding this notification to GE Healthcare Service at 1-800-437-1171 or their local Service Representative. |  | 921 | 921 systems are impacted. | Worldwide Distribution: US (nationwide) including states of: AL AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, ON, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and Puerto Rico; and countries of: Argentina , Austria, Bahrain, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile , China, Colombia, Costa Rica, Ecuador, Egypt, France, Germany, Great Britain, Hong Kong, Hungary, India, Israel, Italy, Japan,Korea, Kuwait, Lebanon, Libya, Lithuania, Malaysia, Mexico, Netherlands, New Caledonia, Nigeria, PanamÂż, Peru, Philippines, Poland, Qatar, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, Venezuela, and Vietnam. | Terminated | January 05, 2017 | 477 | 921 | Not_Computer
Z-0020-2016 | 72091 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Definition Flash; intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angles or spiral planes taken at different angles. |  |  | August 14, 2015 | 2015 | Siemens Medical Solutions USA, Inc | software bug issues for SW-Version VA48A_SP0. The following safety issues were resolved: 1) Correction to improve visual warning and error indication son the gantry display. 2) Correction to improve acquisition data in order to optimize image quality. 3) Correction to improve robustness and general system behavior in some exception handling procedures. 4) Correction to improve auto post processin | Software design | A customer advisory notice, dated 8/14/15, was sent to direct accounts advising them of the issue to observe the safety notice and comply with the corresponding measure until the update has been fully completed. |  | 113 | 113 total | Nationwide Distribution. | Open, Classified |  |  | 113 | Not_Computer
Z-0025-2016 | 72068 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Universal Compact Head Ring Adapter Plate (UCHRAP), a component of the Universal Compact Head Ring Adapter, UCHRA. The UCHRA is a component of the IntegraÂż CRWÂż Precision Arc Stereotactic System. |  |  | August 21, 2015 | 2015 | Integra LifeSciences Corp. | Integra identified that the UCHRAP component (Arc Adapter Plate) would not assemble properly to the UCHRAR component (ARC Adapter Ring) of the Compact Head Ring Adapter due to an error in the manufacturing drawing. | Process design | Integra sent an Urgent-Voluntary Medical Device Recall letter dated August 19, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm instructed their customers to:1. Review your inventory and determine if you have an UCHRAP Adapter Plate with an affected serial number listed on the next page. If you do not, it is not affected and you can continue to use it.2. If you do have product with a serial number listed on the next page, stop using it immediately and remove it from service. 3. Complete the attached form. If you do not have product with a serial number listed, check the box, I do not have any product with a serial number on the list below.4. If you do have product with a serial number listed, check the box, I do have product(s) with a serial number below. Circle the lot numbers and record the total quantity of all lot numbers you have.5. Complete the other information and return it by email or fax as indicated on the form or by providing it to your Integra Neuro Specialist.If an Integra Neuro Specialist has not already done so, Customer Service will contact you and provide an RMA number and directions for returning and replacing the product. Should you have any questions regarding these instructions, please call : (513) 533-7932. For any other questions, please call (609) 936-6822. |  | 6 | 6 units | Worldwide Distribution - US Distribution to the states of : NJ, TX and WY., and to the countries of : Algeria, Germany and India. | Open, Classified |  |  | 6 | Not_Computer
Z-0105-2016 | 72107 | No | HEE | Obstetric anesthesia set. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical LEEP REDIKIT-Used as a surgical preparatory kit for Loop Electrosurgical Excision Procedure |  |  | September 01, 2015 | 2015 | CooperSurgical, Inc. | Kit Label incorrect :Labeled with Lidocaine HCI drug cartridges Lidocaine HCl (2%) but contains vials Lidocaine HCl (1%) | Labeling Change Control | CooperSurgical issued an "Recall Notification" dated September 1, 2015, to all affected customers via FedEX with confirmed delivery receipt. The letter identified the product, the problem, and the action needed to be taken by the customer.Customers asked to discontinue use of affected products and complete the attached Acknowledgement and Receipt Form for a free replacement.Contact Product Surveillance at 203.601.5200 ext. 3300 |  | 305 | 305 units | US Distribution to the states of : AL, AZ, CA, GA, IA, IL, KS, LA, MA, MD, MN, MO, MS, NY, OH, PA, TN, TX, VA, and VT. | Terminated | January 29, 2016 | 150 | 305 | Not_Computer
Z-0107-2016 | 72137 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Force; computed tomography x-ray system. |  |  | August 18, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Software and firmware bugs | Software design | Customers will be notified via the system interface and e-mail or U.S. postal mailing via a Customer Advisory Notice. If necessary the advisory letter may be hand delivered. The letter informed customers that the safety related issues identified were resolved by SW-Version VA50A_SP2. The update will be remotely pushed to affected systems. The letter also advised customers to notify anyone in their organization that should be aware of the information. In the event that the device had been sold, the notice should be forwarded to the new owner, and the firm should be informed of the new owner. |  | 21 | 21 | Nationwide Distribution | Terminated | May 08, 2017 | 629 | 21 | Not_Computer
Z-0027-2016 | 72121 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Universal Locking System, Titanium 3.5mm. |  |  | August 11, 2015 | 2015 | Zimmer, Inc. | It was found through review of open Complaints and stock investigations that a single lot of Zimmer Universal Locking System 3.5 mm Stainless Steel Base and Trays has incorrect silk screening as applied by the supplier. The T-plate catalog numbers noted in the tray are listed as the 00-4946-xxx-xx series (titanium plates); however, the correct numbers should be 00-4936-xxx-xx (stainless steel plat | Process control | On 8/11/2015, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 a.m. and 5:00 p.m. EST. |  | 18 | 18 units | Distributed in NC, AL, KS, TX, UT, CA, and SINGAPORE. | Terminated | March 15, 2016 | 217 | 18 | Not_Computer
Z-0106-2016 | 71979 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Syngo RT Oncologist, an optional accessory to the linear accelerator system and permits localization, contouring, segmentation, image review, and review and approval of treatment plan parameters. In addition, it includes tools and administrative functions to aid in the diagnosis, staging, and prescription of radiation therapy. |  |  | August 11, 2015 | 2015 | Siemens Medical Solutions USA, Inc | This update is intended to provide a software update for the Syngo RT Oncologist, which is currently running SW version 4.2 or 4.3. The safety-related issues, which were described in the Customer Safety Notice distributed as UFSN-RTT/RTO 4.3 Adaptive Targeting [Auto-registration]. The update contains several important safety and performance fixes. | Software design | The firm, Siemens, sent Recall letters, dated August 2015, that describe the product, the problem, and correction, will be hand-delivered to affected sites at the time the fix is implemented by a Siemens technician.If you have any questions, contact Regulatory Technical Specialist at 610-448-6471 or email: marlynne.galloway@siemens.com |  | 14 | 14 | US Distribution to states of: MA, MO, OH, PA, UT, and WI. | Open, Classified |  |  | 14 | Not_Computer
Z-0026-2016 | 71899 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Ortho OS Solutions UltOS 3.5mm locking screw 12mm to 50mm Ortho Solutions LTD, Unit 5 West Station Business Park, Spital Road, Maldon, Essex CM96FF Rx Only a) Sterile b) Non-Sterile |  |  | July 14, 2015 | 2015 | Ortho Solutions Inc | Ortho Solutions Inc is initiating a recall on behalf of UltOS 3.5mm locking screw. The locking thread feature of the screw was not manufactured to specification. As a consequence there is a risk that a screw may not lock onto the plate at the time of primary insertion into the plate and the bone. In the event of this occurring the non functioning screw would need to be replaced with an alternativ | Nonconforming Material/Component | The firm, Ortho Solutions Group, sent an "Urgent Medical Device Recall" notifications and "Effectiveness or Recall Declaration Forms" dated 14 July 2015. The notification informed the consignee of the issue, clinical risks, how to identify the affected product, and the actions the consignees were requested to take. Requested actions included reviewing product on-hand, ceasing distribution of affected product and arrangement for product to be returned to Ortho Solutions, Inc. within 2 working days. Consignees were also requested to fill out the Recall Declaration Form and return it to Ortho Solutions, Inc. via fax to: 201-335-0759 by 16 July 2015.If you should have any queries relating to this matter or require any additional information, contact General Manager at 201-906-6175. |  | 875 | 875 units | US Distribution to states of: CA, CO, GA, IL, LA, NC, NY, TX and WA. | Open, Classified |  |  | 875 | Not_Computer
Z-0005-2016 | 71763 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Celesteion PCA-9000A/2 PET/CT System |  |  | May 08, 2015 | 2015 | Toshiba American Medical Systems Inc | It was found that if specific operations are performed in multi-phase helical scanning, the acquired raw data may not be saved due to a software problem. | Device Design | Planned Action by Toshiba Medical America Systems to address the defect:a. A notification will be sent to the customer with affected system in compliance with 21 CFR1 003.21 upon approval by the Secretary.b. To correct this issue revised software will be installed on the system to prevent this occurrence from happening. The consignee will be contacted by their Toshiba service representative for an appointment when the revised software becomes available.c. The customer will be advised to do the following until the corrective measures have been taken, i.e., revised software installed on their systems. If you have any questions please contact your local Toshiba Representative at (800) 521-1968. |  | 1 | 1 | US Distribution to the state of Nevada. | Terminated | December 03, 2015 | 209 | 1 | Not_Computer
Z-0028-2016 | 71904 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | regard Item Number: 800091, Sterile, LD0692 - C Section - surgical kit containing cover light handle soft, RX. The responsible firm on the label is ROi, LLC, St. Louis, MO. |  |  | April 22, 2015 | 2015 | Resource Optimization &amp; Innovation Llc | The kits contain an equipment glove component which may contain splits or holes compromising the sterility. | Process control | ROi initiated a recall of the custom kits via e-mail to their customers on April 22, 2015. The e-mail included their recall letter dated April 17, 2015. and the supplier's recall letter dated April 16, 2015. The customer was requested to discard the recalled component from their procedure tray at the time of set up and replace it with sterile stock. Customers were instructed to contact any ROi custom pack solutions manager with any concerns. The recalling firm issued a second e-mail notification on 10/19/2015 to the recall coordinator of the distributor. The notice informed them they had been receiving uncorrected product and reminded them to discard the defective component at the time the kit is set up for surgery. Any future distribution will be labeled with an alert to remove the defective component until the kits contain acceptable light handle covers.For questions regarding this recall call 314-364-6561. |  | 26 | 26 kits | Distribution was made to LA, MS, and TX. | Open, Classified |  |  | 38056 | Not_Computer
Z-0110-2016 | 72090 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Free T3 Reagent Kit, P/N A13422 |  |  | August 31, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter, Inc. is recalling Access Free T3 Reagent Kit because they were filled incorrectly and contain insufficient quantity of reagents in one or more of the pack wells. | Under Investigation by firm | Beckman Coulter sent an " Urgent Medical Device Recall" letter dated August 31, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter informs the customers that some Access Free T3 reagent packs (p/n A13422) of lot 431433 were filled incorrectly and contain insufficient quantity of reagents in one or more of the pack wells. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with any questions regarding the notice are instructed to contact Customer Support Center at http://ww.beckmancoulter.com; call (800) 854-3633 in US and Canada. Customers outside of US and Canada are instructed to contact their local Beckman Coulter representative. |  | 16648 | 16,648 units total (1874 units in US) | Worldwide Distribution - US (nationwide) and to the countries of : Australia, Brazil, Canada, China, Colombia, Costa Rica, Hong Kong, Israel, Japan, Kuwait, Macao, Malaysia, Mexico, New Zealand, Saudi Arabia, Singapore, Taiwan, Thailand and Viet Nam. | Open, Classified |  |  | 16648 | Not_Computer
Z-0112-2016 | 72221 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco Radiation Treatment Planning System. |  |  | September 15, 2015 | 2015 | Elekta, Inc. | Dose and MU are incorrect when CT images are viewed from the head, and, when using multiple prescriptions with forced densities. | Device Design | The firm issued an URGENT IMPORTANT FIELD SAFETY NOTICE on 09/15/2015. The letter identified the affected device, as well as the reason for the recall, and clinical impact. Customers were informed that both issues would be resolved in Monaco Patch Release 5.10.01. The attached acknowledgement form should be completed, signed, and returned to a local Elekta Office or representative within 30 days. Questions regarding the notice should be directed to a local Elekta office. |  | 154 | 154 units | Worldwide Distribution -- US, including the states of CA, IN, MI, MO, TX and WI; and, the countries of Australia, Canada, Germany, Greece, India, Netherlands, New Zealand, Turkey, and the UK. | Terminated | April 04, 2017 | 567 | 154 | Not_Computer
Z-0109-2016 | 72176 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MICroSTREP plus 2 Panel, Part No. B1027-202 |  |  | September 10, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter is recalling MICroSTREP plus 1 and MICroSTREP plus 2 panels because a manufacturing error may lead to antimicrobic degradation and the potential for elevated or false resistant results. | Equipment maintenance | Beckman Coulter Inc. sent an Urgent Medical Device Recall letter dated Setember 10, 2015 to all customers who purchased the MICroSTREP plus 1 and MICroSTREP plus 2. The letter identified the affected products, problem and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions regarding the notice, are instructed to contact their Customer Support Representative:website: http://www.beckmancoulter.com or by phone: 1-800-677-7226 in the United States and Canada. |  | 438 | 438 units total (433 units in US) | Worldwide Distribution - US Nationwide and the countries of Bahamas, Belgium, Canada, Ecuador, Germany, Italy, Republic of Korea, Mexico, Peru, Poland, Romania, Saudi Arabia, South Africa, Thailand, United Sates (including Puerto Rico), and Uruguay. | Terminated | January 22, 2016 | 134 | 438 | Not_Computer
Z-0111-2016 | 72113 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Specialist 2 Intramedullary (SP2 IM) Rod 400mm Instrument (PN 96-6120) |  |  | September 08, 2015 | 2015 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is initiating a voluntary recall of select lots of the Specialist 2 Intramedullary Rod 400mm Instrument (PN 96-6120) due to complaints of the rod fracturing and leaving a portion of the rod in patients. | Device Design | DePuy Synthes sent an Urgent Information Recall Notice to all affected customers. Instruments are not lot traceable in the system, so the company sent notices by e-mail on September 8, 2015, and US Distributors will hand deliver notices to all Medical Facilities that may have purchased the affected instrument. The US Distributor or US Sales Consultant will deliver the Medical Professional Notice and Reconciliation Form to the affected Medical Facilities. If assistance is needed, Sales Consultants will assist customers with inspection of all instruments, return of affected lots, receiving credit and returning the completed Medical Facility Reconciliation Forms to DePuy Orthopaedics, Inc. The affected instruments will be returned to the quarantine warehouse located at 700 Orthopaedic Drive, Warsaw, IN 46582. For questions regarding this recall call 574-371-4917. For questions regarding this recall call 574-267-8143. |  | 7488 | 7488 | Worldwide Distribution - US including AL; AZ; CA; CO; CT; DC; FL; GA; HI; IA; IL; IN; KY; LA; MA; MD; ME; MI; MN; MO; MS; MT; NC; NE; NH; NJ; NV; NY; OH; OR; PA; RI; SC; TN; TX; UT; VA; VT; WA; WI; and WV and Internationally to ARGENTINA; AUSTRALIA; AUSTRIA; BELGIUM; BRAZIL; CANADA; CHILE; COLOMBIA; CZECH REPUBLIC; EGYPT; FRANCE; GERMANY; GREECE; INDIA; ISRAEL; ITALY; KOREA; Mexico; NEW ZEALAND; NORWAY; P R CHINA; PANAMA; PERU; POLAND; RUSSIA; SINGAPORE; SLOVENIA; SOUTH AFRICA; SWITZERLAND; UAE; UK; URUGUAY; VENEZUELA; and VIETNAM. | Open, Classified |  |  | 7488 | Not_Computer
Z-0114-2016 | 72170 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Disposable Paper Filter for Sterilization Containers |  |  | September 17, 2015 | 2015 | Stryker Craniomaxillofacial Division | Stryker has become aware that the inhomogeneity of the filter paper might potentially compromise the ability of the filter to maintain a sterile barrier during post-sterilization shelf-life. | Device Design | An Urgent Medical Device Recall letter has been sent to the customers. The affected product was identified, as well as the reason for the recall and risk to health. The letter states that customers should inform users of the Medical Device Removal and pass the notice to all appropriate individuals within the organization; scrap all affected products available at the location or return it to Stryker Craniomaxillofacial, Attention: Recall Coordinator, at the address provided. If there are questions regarding the impacted product, don't hesitate to Contact Stryker Customer service and refer to PFA #2015-032. Complete and sign the enclosed Business Reply Form and fax or copy to: (877) 648-7114 or e-mail a copy to CMF-custserv@stryker.com. |  | 915 | 915 packs | Worldwide distribution -- US, Canada, France, Germany, Italy, Japan, Norway, Sweden, Switzerland, and the Netherlands. | Terminated | February 26, 2016 | 162 | 915 | Not_Computer
Z-0113-2016 | 72242 | No | FEQ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MAJ-1606 Instrument Channel Adaptor; a single use adapter which enables a connection between the endoscope's instrument channel port and the O FP-2 pump's tubing. |  |  | September 11, 2015 | 2015 | Olympus Corporation of the Americas | small puncture marks in a sterile package for an accessory to an Olympus device | Process control | A letter of urgent device safety information, dated September 18, 2015, was sent to direct customers to inform them of the issue and instructions to cease use immediately. Customers were also asked to return the completed response form along with affected devices. |  | 31060 | 31060 | US (nationwide) and Canada. | Terminated | March 13, 2017 | 549 | 31060 | Not_Computer
Z-0017-2016 | 71923 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba Aquilion CT System TSX-302A |  |  | May 08, 2015 | 2015 | Toshiba American Medical Systems Inc | If two specific operations are performed in multi-phase helical scanning, the acquired raw data may not be saved. | Radiation Control for Health and Safety Act | Toshiba America Medical Planned Action to bring defect into Compliance:a. A notification will be sent to all consignees with affected systems in compliance with 21 CFR1 003.21 upon approval by the Secretary.b. To correct this issue revised software will be installed on each system to prevent this occurrence from happening. Each consignee will be contacted by their Toshiba service representative for an appointment when the revised software becomes available.c. Each consignee will be advised to do the following until the corrective measures have been taken, i.e., revised software installed on their systems.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Toshiba America Medical Systems, Inc. will implement this CAP byFebruary 1, 2016.If you have any questions please contact your local Toshiba Representative at (800) 521-1968. |  | 459 | 459 | US Distribution including Puerto Rico | Open, Classified |  |  | 459 | Not_Computer
Z-0116-2016 | 72247 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Distal Access Catheters - 044 - 115 cm: |  |  | August 19, 2015 | 2015 | Concentric Medical Inc | One Distal Access Catheter intended for distribution in Korea, was found at a US location. | Employee error | Issue was discovered by a sales representative during a visit for a case in Kentucky, where multiple Distal Access Catheters (DAC's) were being used. It was noticed that the reference number on one DAC carton did not match, even through the product labeling appeared to be the same.The mismatched unit was removed and returned to Concentric (Stryker).A formal recall letter was not issued as the product was returned after discovery.For further questions please call (510) 413-2263. |  | 1 | 1 | US Distribution to one location in KY. | Terminated | October 09, 2015 | 51 | 1 | Not_Computer
Z-0115-2016 | 72166 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MRIdianÂż ViewRayÂż Radiation Therapy System, ViewRayÂż |  |  | August 27, 2015 | 2015 | Viewray Incorporated | ViewRay discovered that in the event that an encoder breaks or fails on the Patient Handling System (PHS, or couch), when attempting to restart the system, the couch could move unexpectedly. | Device Design | ViewRay sent an Customer Advisory Notice dated August 27, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.It is recommended that if you have a patient on the couch and you receive the fault ''The patient table has experienced an unintended movement. This may be due to the patient or operator leaning on the table.Clear the fault and retry. If the problem persists call Service" on the TPDS,first attempt to clear the fault.If it does not clear and requires you to restart the couch software (same switch as restarting the RTCS) you should manually remove the patient from the couch before restarting the system.A software patch to address this issue is planned for early 2016. In the meantime please practice our recommendation of removing the patient from the couch prior to restarting the RTCS. If you have any questions or concerns regarding this issue, please do not hesitate to contact ViewRay Customer Support at support@viewray.com or call the ViewRayÂżÂż Customer Support center at 855-286-8875. |  | 4 | 4 units | US Distribution to the states of : CA, MO and WI., and Internationally to Korea. | Terminated | June 09, 2017 | 652 | 4 | Not_Computer
Z-0117-2016 | 71991 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | NicoletOne Software v5.94, Catalog/Part Number: 482-649600. |  |  | September 11, 2015 | 2015 | Natus Neurology Inc | When using the NicoletOne v5.94 software, after exiting the impedance check function and returning to the EEG screen, the impedance check signal remains active in waveform, obscuring the EEG signals. | Software change control | Natus Neurology sent an Ugent Medical Device Corrrection Notification letter dated September 9, 2015 to all affected consignees. The letter described the problem and the product involved in the recall. Consignees were informed consignees that a Field Service Technician will contact them to install the new software. Once the software is installed they will ask the Recipient or Field service Technician to complete and return the Verification Form.For questions regarding this recall call 608-829-8600. |  | 101 | 101 (88 US, 13 OUS) | Worldwide distribution. US nationwide, INDIA, ROMANIA, DENMARK, JAPAN, AUSTRALIA, RUSSIA, BANGLADESH, SRI LANKA, BRUNEI, and UNITED ARAB EMIRATES. | Terminated | March 08, 2016 | 179 | 101 | Not_Computer
Z-0120-2016 | 72165 | No | OUT |  |  | Neurology | PMA | -Pipeline Embolization Device |  |  | September 04, 2015 | 2015 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | The firm is recalling Pipeline and Pipeline Flex Embolization Devices from U.S. since the devices were shipped with an EU version of the Instructions for Use. | Process control | The firm, Medtronic, sent a "Notification of Recall" letter dated 09/04/15 to its customers. The letter described the product, problem and actions to be taken. The firm's letter instructed that customers return the Pipeline Embolization Devices to Medtronic for replacement with the US version of product. The firm asks customers to take the following steps:1) Stop using this product immediately.2) Segregate this product from other Inventory.3) Fill out the reply form and fax to the completed form directly to Customer Service at (800) 716-6700 or email to: andrew.schramm@medtronic.com. Customer Service will issue an RGA number for the return of the affected product back to Medtronic Neurovascular.4) Upon receipt of an RGA number, write the RGA number on the FEDEXslip and send the product back to: Medtronic Neurovascular, Attention: Product Recall, 9775 Toledo Way, Irvine, CA 92618. Prepaid and preaddressed FEDEX slips are enclosed for convenience.5. Once they receive the returned units, product replacement units will be shipped directly to customers. The local Neurovascular representative can also be available to assist in expediting the exchange of product.For any questions contact directly at 949-680-1340. |  | 23 | 23 devices | US Distributions to states of: TN, KY, VA, WI, MO, IL, UT, OR, and MD. No foreign consignees. | Terminated | May 18, 2017 | 622 | 23 | Not_Computer
Z-0119-2016 | 72123 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural Nail CM Long. Intramedullary Fixation Rod. |  |  | September 03, 2015 | 2015 | Zimmer Gmbh | In about 5% of tested samples, the outer tray was found to be compromised. Risks: A damaged outer tray detected during surgery may cause a delay in surgery to obtain another implant. If a compromised packaging is not detected prior or during a surgery by the user, it may lead to periprosthetic infection. This may result in revision or multi-stage revision to treat the infection. | Package design/selection | On 9/3/2015, URGENT MEDICAL DEVICE RECALL notifications were sent via courier to the affected distributors and Hospital Risk Managers with instructions for locating and inspecting the affected devices and contacting Zimmer if damage is observed. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have questions concerning this notification, please call 1-877-946-2761. |  | 57603 | 57,603 units | US Nationwide distribution. | Terminated | March 15, 2016 | 194 | 57603 | Not_Computer
Z-0126-2016 | 72243 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Screw Inserter; for use with Intramedullary Rod, orthopedic device. Used to secure the head element to the head element inserter to prevent disengagement during head element insertion. |  |  | September 21, 2015 | 2015 | Synthes (USA) Products LLC | Certain lots of the Screw Inserter may be etched with incorrect graphics. The orientation of the etched image of the oblique angle on the Screw Inserter may not correspond to the tip of the angle of the instrument. | Device Design | DePuy Synthes sent an urgent notice of medical device recall letter dated September 21, 2015, to all affected customers to notify them about the issue, the potential risks, and to provide instructions for return of the completed response form and affected devices. Customers with questions were instructed to call 610-719-5450. |  | 793 | 793 | US Nationwide Distribution | Open, Classified |  |  | 793 | Not_Computer
Z-0127-2016 | 72244 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | 4FR Single Lumen Pressure Injectable PICC; intravascular therapeutic catheter, for long-term use greater than 30 days. |  |  | September 18, 2015 | 2015 | Arrow International Inc | Labeling inaccuracy in the Instructions for Use (IFU) and Patient Card. The IFU and Patient Card with these products state the PICC can withstand up to ten (10) pressure injections, however, the PICC has only been tested to withstand five (5) pressure injections. | Under Investigation by firm | An urgent medical device notification, dated September 22, 2015, was sent to direct accounts to alert them about the issue, the potential risks, and provide instructions for returning the completed response form and affected devices. |  | 31590 | 31590 | Worldwide distribution. US Nationwide, Australia, Canada, and New Zealand. | Open, Classified |  |  | 31590 | Not_Computer
Z-0125-2016 | 72133 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Inserter for Titanium Elastic Nails (TEN), orthopedic device |  |  | September 03, 2015 | 2015 | Synthes (USA) Products LLC | potential for mechanical failures such as breakage. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated September 3, 2015, to all affected customers. The letter identified the product the problem and the action need to be taken by the customer. ACTIONS REQUIRED:We have record that your facility has received the product(s) subject to this recall. It is important to note that the removal of the product would prevent the performance of emergency surgery, thus DePuy Synthes is not requiring an immediate removal of affected product(s). DePuy Synthes has developed a replacement &amp; recovery plan for indicated devices and replacement parts are becoming available. If you have questions on replacement timing and availability, please call DePuy Synthes at 1-800-479-6329.If you have questions regarding the return and recovery of the part, we ask that you contact your local DePuy Synthes Sales Consultant. |  | 1550 | 1550 | US (nationwide) distribution and one Internationally to Australia. | Open, Classified |  |  | 1550 | Not_Computer
Z-0122-2016 | 72120 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Cholestech LDX Multianalyte Controls, PN 88769 and 88773. |  |  | August 27, 2015 | 2015 | Alere San Diego, Inc. | Alere San Diego is recalling the Alere Cholestech LDX Multianalyte Control because the assigned control ranges for the Total Cholesterol (TC) and Triglyceride (TRG) analytes are incorrect. | Error in labeling | Customer notification letters dated 8/25/15 and 9/3/15 were sent to customers to inform them that Alere San Diego is recalling the Alere Cholestech LDX Multianalyte Control because the assigned control ranges for the Total Cholesterol (TC) and Triglyceride (TRG) analytes are incorrect. Customers are informed that the Alere Cholestech LDX Total Cholesterol and Triglyceride assays continue to perform as expected. However, if customers' facility has used the control lot to test Total Cholesterol or Triglycerides as part of their Quality Control program, Alere San Diego recommends that they verify previous results or perform control testing using the replacement Expected Value Card or replacement control materials. The letters also provide customers with actions to be taken. Customers with questions are instructed to contact: Alere San Diego, Inc. 9975 Summers Ridge Road San Diego, CA 92121 U.S.A. Phone: 877-308-8289, FAX: 866-333-9839, E-mail: Verifications.ts@alere.com. |  | 190 | 190 kits | Nationwide Distribution. | Terminated | December 22, 2015 | 117 | 190 | Not_Computer
Z-0123-2016 | 72018 | No | OQY | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | The Ebb Complete Tamponade System, model #CTS-1000 and Code #OQY. |  |  | January 05, 2015 | 2015 | Clinical Innovations, LLC | Clinical Innovations, LLC is voluntarily recalling the Ebb Complete Tamponade System (CTS-1000) Lot Number 1214-F-403.) to address the potential for a balloon leak which may affect the ability to deliver therapy, when needed. | Other | Consignees were notified via email on 08/21/15. All consignees were asked to acknowledge the recall and return devices. |  | 827 | 827 units | Nationwide to MS, MD, TN, NC, IN, FL, AL, and LA, and to France. | Terminated | April 21, 2016 | 472 | 827 | Not_Computer
Z-0121-2016 | 72168 | No | IMJ | Shortwave diathermy. | Physical Medicine | Physical Medicine | 510(k) | Monode, Shortwave Accessory |  |  | August 21, 2015 | 2015 | DJO, LLC | Overheating and burning of the wire insulation inside the applicator. | Device Design | The firm sent out two customer notification letters; one to distributors and one to direct users. Letters were sent September 1, 2015 via USPS certified return receipt.Distributors were asked to obtain a list of their customers with contact information that have purchase or received the monode accessory or Intelect Shortwave device that included the monode as an accessory item. Once the list has been assembled the firm requests that it be forwarded to DJO Global Chattanooga Product Support by phone 877-733-4888, or by e-mail at monode.support@djoglobal.com. The firm's notification states that customers will be sent an upgraded monode accessory along with a pre-paid shipper for the return of the current used unit.Any questions the letter states to contact Product Support per the contact information above or William Fisher, Global Regulatory Compliance at 760-734-3126.Letters sent to direct users received the same instruction of returning the product for a replacement. Any questions the letter states to call 877-733-4888 or email at monode.support@djoglobal.com. |  | 2887 | 2887 | Worldwide Distribution - US Nationwide in the states of WY, OH, WA, NC, FL, NE, GA, KS CA, MO, IN, MA, IL, OK, AZ, TX, AR, UT, NJ, WI, PA, SC, WV, NY, MI, MN, VA, AK, TN, OR, NC, NH, KY, CO, NV, ID and the countries of El Salvador, China, Ecuador, South Korea, Chile, Israel, Peru, Nicaragua, Bolivia, Indonesia | Terminated | October 24, 2016 | 430 | 2887 | Not_Computer
Z-0124-2016 | 72330 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.7MM/3.5MM VA-LCP Medial Distal Humerus Plates 10H/RT/186mm- Sterile and Synthes 2.7MM/3.5MM VA-LCP Extended Medial Distal Humerus Plates 10H/RT/189mm-Sterile |  |  | September 28, 2015 | 2015 | Synthes (USA) Products LLC | The 2.7MM/3.5MM VA-LCP Humerus Plates were found to be labeled incorrectly. Part number 02.117.410S was incorrectly etched with part number 02.117.610S | Labeling mix-ups | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated 9/28/2015 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to"Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number and replacement product(s)."Complete the Verification Section (page 3 of this letter), (even if you do not have the product) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided."Return the Verification Section (page 3 of this letter) with the product to:"Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132."Send a copy of the completed Verification Section by:"Fax: 844-294-7175 or"Scan/email: Synthes4672@stericycle.com"Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450. |  | 36 | 36 | US Distribution to states of:OH, IL, PA, WV, FL, AZ, CA, AL, ND, KY, MN, NJ, MI, GA, IN, TN, VA, NC, and LA. | Open, Classified |  |  | 36 | Not_Computer
Z-0133-2016 | 72078 | No | LZA | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | Ambitex NMD400 Nitrile Exam Gloves size medium. General Hospital and Personal use. |  |  | August 25, 2015 | 2015 | Cardinal Health | Ambitex Nitrile NMD400 Exam Gloves size medium on exam hold were inadvertently placed into commerce. The product was under Exam Hold because samples pulled from a shipment manufactured by the firm failed the water leak test performed by FDA as described at 21 CFR 800.20. | Nonconforming Material/Component | The firm sent an Urgent: Product Recall letter dated 8/25/2015. The firm is requesting the user discontinue using and return any remaining stock of the AMBITEX Nitrile Exam Gloves because the gloves were inadvertently shipped prior to U.S. FDA releasing them for sale.Additionally, the firm requests that the customer quarantine the affected examination gloves; and return a copy of the enclosed Acknowledgment Form confirming their receipt of the Urgent recall via fax to 216-651-9760. Customer Service should be contacted at 800-GLOVES-0 to arrange for the return and credit of any on hand product that customer may have. If the customer has further distributed the affected gloves they are advised to notify their customers of the recall. |  | 34 | 34 cases. | Nationwide Distribution. | Open, Classified |  |  | 34 | Not_Computer
Z-0130-2016 | 72223 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System. |  |  | August 31, 2015 | 2015 | Elekta, Inc. | Incorrect Treatment Delivery Using Third Party Fixed Wedges. | Software design | Consignees were notified via URGENT FIELD SAFETY ALERT, on 08/31/2015. |  | 1440 | 1440 units | Worldwide Distribution-Domestic: AK, AL, AZ, CA, CO, CT, DC, FL, IN, IL, KY, MA, MD, MI, MN, MO, NJ, NC, ND, NY, OH, OK, PA, PR, TN, TX, VA, WA and WI; International: Albania, Argentina, Australia, Austria, Bahamas, Bahrain, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Botswana, Brazil, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Ecuador, Egypt, Estonia, France, Gabon, Germany, Greece, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Libya, Macedonia, Malaysia, Mali, Malta, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Suriname, Switzerland, Taiwan, Thailand, Turkey, The United Kingdom, Ukraine, Venezuela and Vietnam. | Open, Classified |  |  | 1440 | Not_Computer
Z-0131-2016 | 72208 | No | PHX | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | AltiVate Humeral Socket Shell Trial, Part Number 804-06-052 |  |  | September 15, 2015 | 2015 | Encore Medical, Lp | Difficulty assembling the shell trial to the broach and the retaining ring disassembled from the screw. | Device Design | The firm, DJO Global, sent an "URGENT FIELD SAFETY NOTICE" letter dated 9/16/15 to affected consignees/customers. The letter described the product, problem, and actions to be taken. The consignees/customers were instructed quarantine the device, pass the notice on to all who need to be aware within your organization, contact customer service to place a replacement order and receive an RMA number; and to complete a response form and return the recalled device for a replacement.If you have any questions, please call (512) 834-6302 or by email at desiree.hubby@djoglobal.com. |  | 20 | 20 units | US Distribution to states of: PA, FL, ID, and CA. | Terminated | January 10, 2017 | 483 | 20 | Not_Computer
Z-0132-2016 | 71260 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ONE-LINK NON-DEHP Y-TYPE MICROBORE CATHETER EXTENSION SET NEEDLE-FREE IV CONNECTOR WITH NEUTRAL FLUID DISPLACEMENT |  |  | May 21, 2015 | 2015 | Baxter Healthcare Corp | Potential for a leak at the tubing to luer bond. | Packaging | Baxter sent their consignees a recall letter dated 05/21/2015 with the following instructions:Actions to betaken if product was purchased directly fromBaxter :1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.Baxter Healthcare Center for Service can be reached at 888-229-0001 betweenthe hours of 7:00 am and 6:00 pm Central Time, Monday through Friday. Pleasehave your Baxter 8-digit ship-to account number, product code, lot number(s) and quantity of product to be returned ready when calling.3. Complete the Âżenclosed Baxter customer reply form and return it to Baxter byfaxing it to 224-270-5457 or scanning and e-mailing it to fca@baxter.com .Returning the Baxter customer reply form promptly will prevent you from receiving repeat notices.4. If you distribute this product to other facilities or departm ents within yourinstitution, please forward a copy of this communication to them.5. If you are a dealer, wholesaler, distributor/reseller, or Original EquipmentManufacturer (OEM) that distributed any affected product to other facilities,please conduct a consumer-level recall of the affected product that you distributed to customers.Action to be taken if product was purchased from a distributor:1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.The Baxter Healthcare Center for Service can be reached at 888-229-0001between the hours of 7:00 am and 6:00pm central time, Monday through Friday.3. Pl ease note that the Baxter customer reply form is not applicable. If a reply form is provided by your distributor or wholesaler, please return it to the supplier according to the |  | 23600 | 23,600 | Domestic Only | Terminated | May 10, 2017 | 720 | 23600 | Not_Computer
Z-0134-2016 | 72134 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | 5.0mm Unit Rod 270mm, 5.0mm Unit Rod 290mm, 5.0mm Unit Rod 310mm, |  |  | September 04, 2015 | 2015 | Synthes (USA) Products LLC | This product was produced using a finishing process not identified as part of the manufacturing specification. The process used with the lots subject to this Recall was a bead blast process. (Bead Blasting vs. Shot Peened). | Process control | An Urgent Notice: Medical Device Recall, dated September 4, 2015, was sent to end users to alert them about the issue and possible risk to patients. Customers were requested to follow the actions to be taken for if they have affected product or not; complete the response form, and return affected product. Customers can call 610-719-5450 or a local Synthes Sales Consultant with any questions. |  | 492 | 492 | Nationwide Distribution | Open, Classified |  |  | 492 | Not_Computer
Z-0135-2016 | 68228 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | All Philips Ingenia, Intera, Achieva and Multiva MR systems using R5.1i and R5.1 .2 version of software. Indicated for use as a diagnostic device. It can produce cross-sectional images, spectroscopic images and/or spectra in any orientation of the internal structure of the head, body, or extremities. |  |  | May 05, 2014 | 2014 | Philips Medical Systems, Inc. | In spine clinical workflows, cross reference lines may be used to determine the position of slices. In cases, where MobiView fused Images are used to show the cross reference lines, the cross reference lines may be positioned incorrectly. | Software design | Consignees were notified of the recall by letter sent May 2014. They were instructed to do the following: Do not perform planning or review planning of transversal images on fused sagittal images which were generated through Mobi View postprocessing Instead, review the planning of transversal scans only on the unfused stations. A Field Change Order with reference FC0781 00431 is being released that PHILIPS requires Philips field service engineers to disable the MobiView post processing software. When possible this will be done remotely. Once it becomes available, a software update will be installed which will address the reference line positioning issue with MobiView Should you need to communicate with Philips with regard to this program, please reference FC0781 00431. If you need any further information or support concerning this issue, please contact your local Philips representative or the Technical Support Line: 1-800-722-9377. |  | 317 | 317 | Worldwide Distribution-US: AL AR AZ CA CO FL GA IL IN KS KY MA MI NY OH OR PA SC TX VA VT WA OUS: Argentina Australia Austria Bangladesh Belgium Brazil Canada Chile China Colombia Denmark Ecuador Estonia Ethiopia France Germany India Indonesia Iraq Ireland Italy Japan Jordan Korea, South Malaysia Mexico Netherlands New Zealand Norway Oman Palestine Poland Russian Federation South Africa Spain Sweden Switzerland Taiwan Turkey United Arab Emirates United Kingdom. | Open, Classified |  |  | 317 | Not_Computer
Z-0137-2016 | 72366 | No | GWO | Preformed alterable cranioplasty plate. | Neurology | Neurology | 510(k) | Thinflap System Case |  |  | August 06, 2015 | 2015 | Biomet Microfixation, LLC | The lids on implant caddies within P/N 17-5000 and 19-5000 could not be opened by the user without application of undesirable/unexpected excessive force. | Component change control | Biomet Microfixation made phone calls on August 6, 2015 and an initial notice letter sent via FedEx on August 6, 2015, to all distributors. During the phone conversation the distributors confirmed that the inventory invoiced to them is currently in their possession and were asked to remove the lids from the trays and return them back to Biomet Microfixation. In conjunction to the Initial Notice, replacement lids were sent to their business address. FollowÂżup communication would be distributed until all units are returned to Biomet Microfixation.Distributors were instructed to do the following: 1. Immediately locate and remove from circulation the items invoiced to your account as identified on this notice.2. Remove the lids from the quarantined units and replace with the lids (15-5001-18 or 19-5001-18) received with this Initial Notice. The new lids are being distributed to address the issue with the trays.3. Return the lids removed from the quarantined units to Zimmer Biomet CMF and Thoracic team with the labeling information provided with this notice.RETURN THIS LID: lot number(101337 or 101~65) located on the back of the lidDO NOT RETURN NEW LID: Part numbert 15-5001-18 or 1~5001-18 respectively4. Use the FedEx label in this packet to return the lids; this will be Zimmer Biomet's method of tracking your return. If distributors have further distributed this product, they must notify the recipient of this action by providing the recipient a copy of this notice. This letter must be given to the personnel responsible for receiving field safety notices. Please understand it is your responsibility to locate and quarantine these products. Customers with questions were instructed to call at 1-800-874-7711 or 904-741-4400, extension 9574, Monday through Friday, 8am-4pm ET. |  | 5 | 5 units | Nationwide Distribution including CA, FL, OK, NY, NC, and IL. | Open, Classified |  |  | 42 | Not_Computer
Z-0152-2016 | 72112 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | PORT-A-CATH Plastic Hub Needles, 0.9mm (20G) x 19mm (3/4 in.) 90-degree. REF 21-2287-24 |  |  | September 14, 2015 | 2015 | Smiths Medical ASD, Inc. | The label on the lid stock of the individual needle packages for a portion of Lot number 46262 contains the incorrect part number, part number description and product image. The part number should be 21-2287-24 Port-A-Cath Plastic Hub Bent Needle. | Labeling mix-ups | The firm, Smiths Medical, sent an "URGENT MEDICAL DEVICE RECALL" letter dated 9/14/2015 to customers. The letter described the product, problem, and actions to be taken. The customers were instructed to Inspect your inventory for and quarantine the affected product, complete and return attached Confirmation form via fax to: 1-651-747-4959 or email to: recall.response@smiths-medical.com within 10 days; upon receipt of completed form, customer representative will contact you and arrange for exchange of your unused affected product; send a copy of letter and confirmation form to your customers identified, replacement product will be arranged for your customers; provide email confirmation of your customer notification to Smiths Medical at recall.response@smiths-medical.com, and notify downstream customers who may have received the product. If you should have any questions regarding this information, please contact Smiths Medical's Customer Service Department at 1-800-258-5361. |  | 130 | 130 Cartons (1560 needles) | Worldwide Distribution: US (nationwide) to CA only and countries of: Australia, Belgium, Germany, Hungary, Italy, Sweden, and Switzerland. | Open, Classified |  |  | 130 | Not_Computer
Z-0150-2016 | 72237 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Perifix Continous Epidural Anesthesia Pain Control Trays; Injection of anesthetics to provide regional anesthesia for pain management. |  |  | August 27, 2015 | 2015 | B. Braun Medical, Inc. | Potential discoloration of the DuraPrep solution from unintended exposure to the sterilant during sterilization performed by the supplier. | Under Investigation by firm | BBMI provided written notification, dated August 27, 2015, to all customers in receipt of suspect product, along with acknowledgement forms detailing the suspect product that was distributed to each account. These customers include both BBMI direct end customers and BBMI direct distributors. Customers were notified by US Postal Service Certified Mail with registered return receipt mail or FedEx Priority. All customers were provided with instructions for disposition of suspect product in their inventory. Distributors shall be reminded of their responsibility in executing the recall. |  | 433 | 433 | Nationally | Open, Classified |  |  | 433 | Not_Computer
Z-0149-2016 | 72017 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Natus neoBLUE 3 Phototherapy system or Neonatal Phototherapy, Catalogue Part Number: 010066; Top Assembly Part Number: 001103 |  |  | August 05, 2015 | 2015 | Natus Medical Incorporated | The neoBLUE Instruction For Use (IFU) and Service Manuals for the neoBLUE 3 were revised to show the cleared intensity statement settings so the output intensity aligns precisely with the information listed in the cleared 510(k) submission. | No Marketing Application | Natus Medical starting sending the Product Labeling Change NeoBlue LED Phototherapy System letter, dated August 2015, to all affected customers. The firm also attached the electronic copy of the revised User Manual with the letter.Customers are informed that the user and service manuals related to intensity for the neoBLUE 3 system are revised. Also, while the neoBLUE system does allow the light output to be adjusted with the potentiometers to accommodate different treatment heights, the firm removed the phrase "adjustable to &gt;50 ÂżW/cmÂż/nm" from the specifications.If you would like a CD disk containing the updated User Manual and the Service Manual at no charge, please call Natus Technical Service at 888-496-2887. |  | 4560 | 4560 units | Nationwide Distribution. | Terminated | March 17, 2016 | 225 | 4560 | Not_Computer
Z-0143-2016 | 72108 | No | NPJ | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ConForMIS iDuo Knee Bicompartmental Knee Replacement System , iDUO G2, Right Medial |  |  | August 31, 2015 | 2015 | ConforMIS, Inc. | May contain small amounts of ethylene glycol residue | Process design | ConForMIS issued a Press on August 31, 2015 to the general public The sales force was notified through email. The sales personnel for affected accounts that had upcoming surgeries scheduled were asked to confirm that the affected surgeries were cancelled and that the devices were not used in surgery (Formal notification letters will be distributed to the affected accounts by Tuesday, September 8, 2015. The user required to confirm their understanding of the recall and to return any unused product to ConforMIS. Distribution of the letters conducted in two phases as described below:Phase 1-Surgeons with product for which the surgery has not yet been completedPhase 2-Surgeons for whom product was confirmed to have beenconsumed/implanted. An updated Surgeon letter dated September 8, 2015 issued includes additional information for devices that have been used in surgery.Contact your local ConforMIS representative or Amita Shah, at 17813459164. |  | 12 | 12 units | Distributed Nationwide and in Germany, Switzerland, and UK. | Terminated | March 02, 2016 | 184 | 939 | Not_Computer
Z-0154-2016 | 72273 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TELETOM EQUIPMENT MANAGEMENT SYSTEM 4 and 6 series |  |  | September 23, 2015 | 2015 | Berchtold Corp. | During the use of the TELETOM Series 4 &amp; 6 there is potential for the controls and brakes not to work properly due to fluid intrusion within the control switches and/or the control board assembly. | Device Design | Stryker sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A Stryker Representative would visit each account for product inspection. During the inspection the Stryker Representative will remove the control switches and replace them with new control switches. Customers were instructed if any malfunction prior to inspection were noticed customers should call local Technical Support immediately at 800-243-5135. |  | 2129 | 2129 units | Worldwide Distribution - US (nationwide), Canada, and Germany | Open, Classified |  |  | 2129 | Not_Computer
Z-0153-2016 | 72011 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | OneTray Sealed Sterilization Containers |  |  | February 03, 2014 | 2014 | Innovative Sterlization Technologies LLC | The wrong sterilization time was on the label. The label lists Steam Gravity 10 - 34 minutes at 270 degrees F instead of Steam Gravity 17 - 34 minutes at 270 degrees F. | Process control | The firm sent an Urgent Medical Device Recall letter dated September 14, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to fill out the attached response SASE mailer and return to or fax to the address/number identified below:Innovative Sterilization Technologies, LLC, 7625 Paragon Road, Suite A Dayton, OH 45459, 937-286-3571 (C), 937-630-4346 (F), 937-619-0138 x 202 (Office), dbillman@ISTSterilization.com |  | 4621 | M2104 - 25 units; M2408 - 84 units | US Nationwide Distribution in the states of IL, IN, LA, OH, PA, TX &amp; WV. | Terminated | February 04, 2016 | 731 | 4621 | Not_Computer
Z-0156-2016 | 72205 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Smith &amp; Nephew MOSAICPLASTY DP- Disposable Harvesting System 4.5 mm, Sterile |  |  | September 10, 2015 | 2015 | Smith &amp; Nephew, Inc., Endoscopy Div. | Sterility of device is compromised due to breach in the packaging | Package design/selection | Smith &amp; Nephew issued Recall Notification Letter dated September 10, 2015 via overnight Federal Express. Customers were requested to inspect inventory and locate any unused devices from the above listed product and batch numbers, and quarantine them immediately. Complete the Inventory Return Certification Form on the following page, indicating the lot numbers and quantities that need to be returned and include your phone and fax number in the spaces provided. If you do not have product to return, please place an x in the column No product to return. Send the completed form to the Smith &amp; Nephew's Field Action Department via email at FieldActions@smithnephew. com or fax +1-901-566-7975 to obtain a return authorization (RA) number.Return any affected product to the address listed on the Inventory Return Certification Form. Please indicate the RA number on your return shipment. |  | 1287 | 1287 units | Worldwide Distribution - US Nationwide, Canada, Austria, Australia, Belgium, China, Spain, Finland, France, Great Britain, Greece, Germany, Hungary, India, Italy, Japan, Korea, Lithuania, Latvia, Netherlands, Poland, Portugal, Sweden, Thailand, Turkey, Ukraine and Venezuela | Open, Classified |  |  | 4483 | Not_Computer
Z-0159-2016 | 72305 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | CC cartridges for IRMA TRUPOINT Blood Gas Analyzer, PN 039903. |  |  | September 16, 2015 | 2015 | LifeHealth, LLC | May report values outside the published performance specifications for pH. After a certain point in shelf life, the pH sensor may give results that are falsely elevated. The discrepancy is greater as shelf life progresses and can be seen on the aqueous controls before the blood. | Nonconforming Material/Component | Consignees were sent a LifeHealth "Urgent Medical Device Recall" letter dated September 16, 2015. The letter described the problem and the product involved in the recall. The letter described the Risk to Health, Action to be taken and Other Information. Consignees were advised to discontinue use and destroy all stock on hand. Requested consignees to complete and return the Verification Form to LifeHealth. For additional information contact Nancy Ring at 1-855-762-8978 or 651-638-1000, Monday through Friday, 8:30 AM to 4:30 PM, Central Time or see our website at www.lifehealthmed.com. |  | 116475 | 4659 boxes of 25. (116475 cartridges) | US: Nationwide (AL, AR, CA, CO, FL, IL, IN, IA, KS, KY, MD, MA,MI, MN, MS, MO, NE, NJ, NY,NC, OH, OK, PA, TN, TX, VA)OUS: ALBANIA, BELGIUM, BOLIVIA, CANADA, CHILE, CHINA, DENMARK. DOMINICAN REPUBLIC, ECUADOR, EGYPT, ENGLAND, ESTONIA, PHILIPPINES, FRANCE, GERMANY, GREECE, ICELAND, INDONESIA, IRELAND, ITALY, JAMAICA, JAPAN, MOROCCO, NETHERLANDS, NORWAY, PERU, PORTUGAL, SOUTH AFRICA, SOUTH KOREA, SPAIN, TAIWAN, TUNISIA, TURKEY, UNITED ARAB EMIRATES, VIETNAM. | Terminated | March 21, 2016 | 187 | 159850 | Not_Computer
Z-0161-2016 | 72365 | Yes | LLB | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Nipro Safe Touch TULIP Safety Fistula Needle, FT+152530TP, for use as a blood access device for blood purification and for other treatments requiring an extracorporeal circuit of larger volumes of blood. |  |  | August 13, 2015 | 2015 | Nipro Medical Corporation | Some 1 1/4" needles were packaged and labeled in the 1" blister package. | Mixed-up of materials/components | Customers were issued a RMA 60001203 on August 18, 2015. |  | 50000 | 50,000 pieces | Indiana | Open, Classified |  |  | 50000 | Not_Computer
Z-0118-2016 | 72138 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | CIOS ALPHA; image intensified fluoroscopic x-ray system |  |  | August 17, 2015 | 2015 | Siemens Medical Solutions USA, Inc | patient procedure interruption due to a potential system failure | Device Design | Customer Safety Advisory Notice XP038/15/S, dated August 17, 2015, was sent to direct accounts to alert them to the issue and provide work-around instructions to avoid potential risks. |  | 44 | 44 | Nationwide Distribution. | Open, Classified |  |  | 44 | Not_Computer
Z-0162-2016 | 72425 | No | OSF | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Signature Patient-Specific Surgical Guides; patient specific instruments used as a guide during surgical procedure for total and partial knee arthroplasty. Part of the Signature Personalized Patient Care System. |  |  | September 23, 2015 | 2015 | Materialise N.V. | Potential inaccuracies in the alignment of the guides. Surgery was inaccurately planned because the patient image data was incorrectly processed. | Use error | On September 16, 2015 Materialise NV identified four cases where not all available image sets were used to design the case. On September 18, 2015, prior to the surgery dates Materialise informed (phone call) the SignatureServices Manager of Zimmer Biomet that inaccuracies in the alignment of the knee might be observed with the use of the guides due to this issue. Biomet was asked to segregate the packages for these four cases and to return them to the Materialise US office. On September 23, 2015 this wasfollowed up by e-mail notification letter with response forms.Biomet was requested to contact Materialise customer service to return these four cases and arrange a replacement set of patient specific surgical guides Monday through Friday, 8:00 AM to 4:30 PM, Eastern Time. E-mail : ortho-quality@materialise. beCustomer ServiceMaterialise USA LLC44650 Helm CourtPlymouth, Michigan 48170734-259-7033 |  | 4 | 4 | Distributed in MI, ID, TX, and in France. | Terminated | February 12, 2016 | 142 | 4 | Not_Computer
Z-0165-2016 | 72216 | No | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Exprt Precision System: Revision Knee, Model Number 160-010-726/738 |  |  | September 16, 2015 | 2015 | Encore Medical, Lp | The labeling is missing the size/diameter information. | Labeling design | DJO Surgical sent an Urgent Field Safety Notice dated September 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter requested the consignees complete and return a response card and return affected devices for replacement. For further questions call (512) 832-6302. |  | 209 | 209 units | US Distribution to the states of : SC, WI, RI, NY, MO, AR, OK, GA, FL, AL, CA, AZ and CO. | Terminated | January 04, 2016 | 110 | 209 | Not_Computer
Z-0164-2016 | 72311 | No | JIF | Ammonia test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Sentinel MULTIGENT Ammonia Ultra Reagent, In-Vitro Diagnostic for ammonia measurement; List Number: LN 6K89-30. Intended for the in vitro quantitative determination of Ammonia (NH3) in human plasma. |  |  | September 22, 2015 | 2015 | Sentinel CH SpA | Results for ammonia are below the linear range of the assay for samples that have been collected in sodium or lithium heparin tubes. | Process control | Sentinel Urgent Field Safety Notice issued on 9/22/15 by Abbott to customers, including their US customers.Actions to be Taken1. Immediately discontinue use of lot 50083Y600 and destroy any remaining inventory.2. Order a replacement Ammonia Ultra reagent lot.3. Follow your individual laboratory protocol regarding reviewing previously reported patient results.4. Please retain this communication for future reference. If you have forwarded any MULTIGENT Ammonia Ultra reagent lot 50083Y600 to another laboratory, please provide a copy of this letter to them.ContactsIf you or any of the health care providers you serve have any questions regarding this information, U.S. Customers should contact Customer Service at 1-877-4 ABBOTT. |  | 765 | 765 kits | Nationwide Distribution. | Terminated | March 30, 2016 | 190 | 765 | Not_Computer
Z-0163-2016 | 72154 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i2000SR System |  |  | September 09, 2015 | 2015 | Abbott Laboratories, Inc | Incorrect tubing. | Nonconforming Material/Component | Abbott Laboratories executed a Technical Service Bulletin on August 14, 2015.. The firm sent out a Product Correction notice on September 9, 2015 to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed that an Abbott representative will make all necessary arrangements for replacing tubing on the impacted instruments. Customers were asked to retain a copy of the notice for their records. Customers with questions were instructed to contact Customer Service at 1-877-4ABBOTT. Customers outside the US should contact their local area Customer Service.For questions regarding this recall call 877-422-2688. |  | 65 | 65 units | Worldwide Distribution - US (nationwide) and Internationally to France, Netherlands, Greece, Italy, Belgium, India, Austria, Germany, and Japan | Open, Classified |  |  | 65 | Not_Computer
Z-0166-2016 | 72013 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Perkin Elmer Specimen Gate, Screening Center. Data management of neonatal screening test results and demographics by qualified laboratory personnel in newborn screening programs. |  |  | August 24, 2015 | 2015 | Perkin Elmer Life Sciences Inc | It would be possible for an abnormal analytical test result value from an affected newborn to be reported with a default result interpretation of normal. The analytical test result value would be correct but the interpretation of the test result value would be incorrect. | Software design | Perkin Elmer issued the Recall Letter on 8/24/15 delivered to the customer by email .Software Services will be asked to make sure the customers receive the Recall Letter and implement the recommended corrective measures: Ensure all laboratory routine sample testing procedures are followed to avoid procedural anomalies. In the event a procedure anomaly occurs, review the final analytical result and the corresponding interpretation to confirm both are accurate before releasing the report. Once a PerkinElmer reporting error query is available, continue to run the query until the final corrective action is deployed to your site. The final corrective action will involve an installation update to the software.For further information, please contact your local PerkinElmer representative or specimen.gate.support@perkinelmer.com |  | 6 | 6 units | US distribution to FL, GA, and NV; and Canada | Open, Classified |  |  | 6 | Not_Computer
Z-0168-2016 | 71207 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ONCOR Avant Garde, ONCOR Impression Plus, ONCOR Expression, ONCOR Impression, PRIMUS HI, MEVATRON M2/PRIMUS Mid-Energy, ARTISTE; intended use is to deliver x-ray radiation for therapeutic treatment of cancer. |  |  | May 05, 2015 | 2015 | Siemens Medical Solutions USA, Inc | potential safety risk of collision of the LINAC gantry with the patient or the treatment table in case of automatically sequenced treatment delivery techniques using SIMTEC" Auto Field Sequence Option. | Software design | Siemens sent Safety Advisory Notices, dated May 5th and 22nd, 2015 to all affected customers to inform them about an update of the Control Console software to correct defects affecting the Control Console with their digital Linear Accelerator. Customers were instructed to ensure everne at their organization is aware of the information in the Advisory Notice. If product was further distributed, please forward Safety Notice to all those affected.For questions regarding this recall call 610-219-6300. |  | 158 | 158 | Nationwide Distribution | Terminated | May 05, 2017 | 731 | 158 | Computer
Z-0183-2016 | 72104 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Chemistry XPT Chemistry Systems, Siemens Material Number (SMN) 10723034. |  |  | August 28, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | ISE module may produce discrepant, unflagged results for Sodium, Potassium or Chloride when the respective electrode is left unplugged after replacement, washing or maintenance activities. The calibration and 2-level QC will fail if the electrode is left unplugged. The IFU states to calibrate the ISEs after replacement and to run two levels of controls before running patient samples. | Device Design | Urgent Medical Device Correction letters (dated 8/27/2015) and Field Correction Effectiveness Check Forms were sent to customers via FedEx starting on 8/28/2015, informing them of the potential for discrepant ISE results. Customers must follow the instructions provided in the letter. The Actions to Be Taken section of the letter notes that customers must ensure that the electrodes are fully plugged in after replacement, washing or maintenance activities, followed by performing an ISE Calibration and running 2 levels of Quality Control material. Urgent Field Safety Notices were sent to customers outside the USA. The Field Correction Effectiveness Check form should have been completed and returned. A copy of the letter should be kept with laboratory records; and, the letter should be forwarded to anyone who may have received the affected product. Customers should contact their Siemens Customer Care Center or their local Siemens technical support representative with questions. |  | 74 | Domestic: 15 units; Foreign: 59 units | Worldwide Distribution -- US, Australia, Bahrain, Denmark, Egypt, Egypt, Finland, France, Germany, Italy, Netherlands, New Zealand, Norway, Republic Korea, Singapore, Slovakia, Spain, and the United Kingdom. | Open, Classified |  |  | 74 | Not_Computer
Z-0193-2016 | 72274 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | IntegraÂż Bone Marrow Biopsy Trays Catalogue No. 3404253 |  |  | September 22, 2015 | 2015 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra received an Urgent Drug Recall Notice from Hospira Inc. for their 1% Lidocaine HCL Injection, USP, 10 mg/mL, Lot 44-359-DK. Hospira Inc. is recalling vials of Lot 44-359-DK, 1% Lidocaine HCL Injection, USP, 10 mg/mLm distributed by Hospira from February 2015 to March 2015 due to a confirmed complaint of visible, partially embedded particulate within a single-dose glass teartop vial. | Process control | Integra sent an Urgent Medical Device / Drug Recall Notice dated September 22, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed.to forward copies of the Integra and Hospira recall notification letters to their customers. Confirmation of recall notification delivery and instructed to stop use of devices and return to Integra.If you have questions of a customer service nature or beyond the information in the Hospira Inc. URGENT DRUG PRODUCT RECALL NOTICE, please contact either your local Integra Pain Management Sales Representative or customer service at 1-800-241-2210. |  | 470 | A total of 470 packages (47 cases) | US Distributed to the state of : MD. | Terminated | January 29, 2016 | 129 | 470 | Not_Computer
Z-0181-2016 | 72424 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco - Radiation Treatment Planning used to make treatment plans for patients with prescriptions for external beam radiation therapy. |  |  | October 16, 2015 | 2015 | Elekta, Inc. | Unintended update of Dose and MU and Incorrect Assignment of Bolus. | Other | Elekta sent Important Field Safety Notice 382-01-MON-003 to all customers on 10/16/2015. The notice informs users of the specific product and version numbers affected by both issues. The affected devices for both issues are those running Monaco 5.10.01. Customers are instructed to complete and return acknowledgement form. Solution/fix has been released in Monaco patch 5.10.02 and all affected users will be made aware of its availability by the end of October 2015. |  | 154 | 154 | CA, IN, MI,MO, TX, WI, Australia, Canada, Germany,m Greece, India, Netherlands, New Zealand, Turkey and United Kingdom | Terminated | April 04, 2017 | 536 | 154 | Not_Computer
Z-0170-2016 | 72374 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARIA Radiation Oncology, versions 10, 11 , 13.0 and 13.5 with Clinical Assessment License. |  |  | September 28, 2015 | 2015 | Varian Medical Systems, Inc. | An anomaly was identified with the ARIA for Radiation Oncology software with a Clinical Assessment license. Admin Instructions may not print on drug order prescriptions or be transmitted with e-Rx. | Software design (manufacturing process) | Varian sent an Urgent Medical Device Corrections letter dated September 29, 2015 via trackable method to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that Varian is developing a technical correction for this issue. Customers will be contacted by a Varian Field Service. A copy of the notification is to be kept with the most current product labeling and all appropriate personnel in the radiation oncology department should be aware of this issue. For questions contact your local Varian Medical Systems Customer Support District orRegional Manager |  | 1971 | 1971 units | Worldwide Distribution - US Nationwide and the countries of: Canada, Switzerland, South Korea, Germany, Spain, Ireland, Spain, France, Belgium, United Arab Emirates, United Kingdom, Norway, Italy, Trinidad and Tobago, Netherlands, Philippines, Portugal, Denmark, Sweden, Belgium, Saudi Arabia, Thailand, Australia, India, Taiwan, South Africa, Israel, Australia, Kuwait, Finland, Tunisia, Algeria, Argentina, Bangladesh, Brazil, Chile, China, Colombia, Costa Rica, Czech Republic, Guam, Guatemala, Hong Kong, Indonesia, Japan, Kazakhstan, Kuwait, Lebanon, Luxembourg, Madagascar, Moldova, Morocco, New Zealand, Panama, Poland, Qatar, Romania, Russia, Singapore, South Africa, Tunisia and Turkey. | Terminated | May 12, 2017 | 592 | 1971 | Not_Computer
Z-0192-2016 | 72360 | No | EZX | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Synthes Threaded Plate Holder- Long for 3.5mm Locking Hold; Part Numbers 324.031. |  |  | October 07, 2015 | 2015 | Synthes (USA) Products LLC | The affected part number and lot listed above is incorrectly etched with the wrong part number and may not fit into a Small Fragment Plate. | Mixed-up of materials/components | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated 10/5/2015 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:" Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number and replacement product(s)." Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided." Return the Verification Section (page 3 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Even if you do not have any product." Send a copy of the completed Verification Section by: Fax: 844-721-3048 orScan/email: Synthes4704@stericycle.com." Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes Sales Consultant. |  | 25 | 25 | Worldwide distribution: US (nationwide) and country of: Canada. | Open, Classified |  |  | 25 | Not_Computer
Z-0189-2016 | 72052 | No | LHC | Powered heating unit. | Physical Medicine | General Hospital | 510(K) Exempt | Regard, Item Number 830500003, Sterile, CV0578C - CV PACK - Custom surgical kit for surgery. |  |  | August 25, 2015 | 2015 | Resource Optimization &amp; Innovation Llc | Custom surgical kits contain a component (ORS fluid warming and slush drapes) that was recalled due to compromised package integrity potentially affecting sterility assurance. | Packaging | On August 25, 2015, the recalling firm sent an email to their direct customers informing them of the recall. The email instructed their customers to examine their inventory and determine whether they had any of the affected devices. The recalling firm attached the recall notice they received from their supplier.On September 4, 2015, the recalling firm sent a second email clarifying that they were instructing their customer to discard the supplier component . The notice also asked the customer to respond with how many affected devices they had in their inventory at the time they received the notice. |  | 17 | 17 kits | Distributed in Texas and Louisiana. | Open, Classified |  |  | 506 | Not_Computer
Z-0167-2016 | 72345 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 2.7 mm Locking Screw SLF-TPNG with T8 Stardrive Recess 12mm. The Synthes modular Mini Fragment LCP System is intended for fixation of fractures, osteotomies, nonunions, replantations, and fusions of small bones and small bone fragments. |  |  | October 06, 2015 | 2015 | Synthes (USA) Products LLC | Specific lots of the SS 2.7mm locking screw have been mislabeled as SS 2.4mm locking screw. | Packaging | Synthes sent an URGENT NOTICE: MEDICAL DEVICE NOTIFICATION UPDATED INFORMATION: CORRECTED LOT NUMBER SS 2.7mm Locking Screw, self-tapping letter dated January 13, 2014, to the affected customer.Please take the following actions:If you DO have any of the identified devices, please take the following steps:Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number.Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided.Return the Verification Form (page 3 of this letter) with the product to:Credit/Returns, Synthes.If you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information.Return the documents to Synthes by: Fax: 1-866-792-5446 orScan/email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. For further questions please call 610-719-6500. |  | 98 | 98 | US Distribution to the state of : TN | Open, Classified |  |  | 98 | Not_Computer
Z-0182-2016 | 72372 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 ultrasound systems |  |  | October 07, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Siemens will provide all customers a new version of system software (VA35E or VB10C depending on the hardware configuration). Hardware fixes will be implemented to address situations where the metal component used to engage the wheel lock becomes disconnected, and where the system locks up due to misalignment between the bezel and MPI board, in all units that have not already been corrected. | Device Design | On August 19, 2015, Siemens Ultrasound began sending a customer notification to users, advising them to correct the uppercase/lowercase differences in the patient name at the time of registration until the issue is resolved with a future software release. Siemens believed that the other issues in this correction will begin approximately November 30, 2015 and December 31, 2015. Field service reps will visit sites to install software updates or hardware corrections. |  | 2099 | 2,099 units | Worldwide Distribution -- US and Albania, Angola, Argentina, Armenia, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia Herzeg., Brazil, Brunei, Bulgaria, Canada, Colombia, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, El Salvador, Finland, France, Georgia, Germany, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Lebanon, Lithuania, Malaysia, Maldives, Malta, Mexico, Netherlands, New Zealand, Nigeria, Norway, P.R. China, Pakistan, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Turkey, Turkmenistan, U.A.E., Ukraine, United Kingdom, Uzbekistan, Venezuela, and Vietnam. | Open, Classified |  |  | 2099 | Not_Computer
Z-0178-2016 | 72282 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Ingenuity Flex Computed Tomography X-ray systems. Intended to produce cross-sectional images of the body. |  |  | September 14, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm discovered Ring/Dot artifact due to X-ray measurement error. | Software design | On 9-23-2015 the firm sent URGENT Field Safety Notices to their customers. The letter identified the products, problem, and actions to be taken by the customers. |  | 67 | 67 units | Worldwide Distribution-US (nationwide) including DC and the states of AL, AR, AS, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NV, NY, OH, OK, OR, SC, TN, TX, UT, VA, VT, WI &amp; WV, and the countries of Albania, Algeria, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Ecuador, Egypt, Ethiopia, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Italy, Japan, Jordan, Kazakhstan, Latvia, Lebanon, Libya, Lithuania, Malaysia, Mexico, Mongolia, Myanmar, Netherlands, Nigeria, Norway, Oman, Panama, Philippines, Poland, Portugal, Romania, Russia, Rwanda, Samoa, Saudi Arabia, Singapore, South Korea, Spain, State of Palestine, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan &amp; Viet Nam. | Open, Classified |  |  | 643 | Not_Computer
Z-0172-2016 | 72211 | No | DXH | Telephone electrocardiograph transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare CardioCall ECG Event Recorder, Models VS20, VS20x4, ST80, ST80x4, PCBAs PN # 670-1670-00 |  |  | September 03, 2015 | 2015 | Spacelabs Healthcare Inc | When the battery is inserted into the CardioCall ECG Event Recorder, the device beeps twice and is followed by a continuous ringing sound that only stops when battery is removed. The device has a self-test failure and will not operate normally. | Vendor change control | Spacelabs Healthcare sent an Urgent-Medical Device Correction letter, dated September 17, 2015 to US consignees (Hospital Administrator/Biomedical Manager/Office Manager). The firm plans to email a customer letter (translated as necessary) to international subsidiaries and distributors of record on 9/25/15. Customers were informed Spacelabs Healthcare will send a replacement CardioCall ECG Event Recorder(s) at no cost and a prepaid method to ease the return of the affected CardioCall device(s) back to Spacelabs.For additional information or technical assistance contact:Global Technical SupportSpacelabs Healthcare, Ltd.1 Harforde Court, John Tate RoadHertford SG13 7NWUnited Kingdom+ 44 (0) 1992 507700For USA Technical Support call 1-800-522-7025 and select 2 for Technical Support |  | 275 | 275 devices (20 in US and 255 international) and 113 PCBAs (international only) | Worldwide Distribution - US Distribution in the states of Nebraska, South Carolina, and Washington and in the following countries of AUSTRALIA, BELGIUM, BOLIVIA, CANADA, CHILE, DENMARK, GERMANY, GUATEMALA, HONG KONG, ITALY, NETHERLANDS, NEW ZEALAND, POLAND, PORTUGAL, QATAR, SINGAPORE, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 275 | Not_Computer
Z-0169-2016 | 72228 | No | HIB | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Welch Allyn KleenSpec Corded Illumination System, Vaginal Speculum Lighting System, Model Number 78x Series Illuminator, REF 901070, Catalog Numbers 78810, 78812, 78814 and 78816, Rx ONLY |  |  | October 02, 2015 | 2015 | Welch Allyn Inc | The firm became aware that the KleenSpecÂż 788 Corded Illuminator has the potential to be incorrectly connected to a commonly available AC power cord instead of the illuminator low voltage transformer power cord which is supplied with the illuminator. This may result in a shattered lamp bulb and a potential electric shock hazard. | Device Design | On 10/02/2015, Welch Allyn commenced mailing of the Urgent Medical Device Correction letters (dated 10/01/2015) via UPS Certified Next Day Air Saver Mail. The letters inform customers of the issue. To help customers avoid any such occurrences, Welch Allyn is providing supplementary User Precautions which re-emphasize safe use of this product and providing additional specificity to the warnings already on the labeling and packaging. |  | 194464 | Domestic: 179,401 units, Foreign: 15,063 units | Nationwide, Puerto Rico and foreign distribution to the following countries: Australia, Bahamas, Bahrain, Belgium, Bermuda, Cambodia, Canada, Czech Republic, China, Cyprus, Denmark, Egypt, France, Germany, Great Britain, Greece, Guam, Guatemala, Iceland, India, Israel, Japan, Kenya, Kuwait, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Poland, Russia, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Sri Lanka, Thailand, Turkey, United Arab Emirates. | Open, Classified |  |  | 194464 | Not_Computer
Z-0195-2016 | 72363 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom Pak 9319-27. Single-use medical devices and accessories designed by medical professionals for a specific ophthalmic surgical procedure. |  |  | September 21, 2015 | 2015 | Alcon Research, Ltd. | The recalled lot of custom paks are labeled as latex free and contain latex gloves rather than the specified non-latex gloves. | Mixed-up of materials/components | The recalling firm notified the consignee on 9/21/15 via phone call, and with a follow up visit to the consignee site. The recalled units were recovered. |  | 5 | 5 units | Distributed in VT. | Terminated | March 24, 2016 | 185 | 5 | Not_Computer
Z-0194-2016 | 72362 | Yes | KDJ | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter 15L Cycler Drainage Bag Sterile, nonpyrogenic fluid path |  |  | October 06, 2015 | 2015 | Baxter Healthcare Corp. | Leakage from the inlet port of the 15L Cycler Drainage Bags. | Under Investigation by firm | Baxter sent an Urgent Product Recall letter dated October 8, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The recalling firm will ask customers to locate and separate all affected products, contact the recalling firm to arrange for a return, and complete the customer reply form. Further questions, please call (800) 422-9837. |  | 19440 | 19,440 units | US Distribution and Internationally to Canada. | Open, Classified |  |  | 19440 | Not_Computer
Z-0180-2016 | 72229 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Reverse Shoulder Prosthesis Stem Impaction Fixture REF 804-03-053 |  |  | September 14, 2015 | 2015 | Encore Medical, Lp | During the assembly, the impaction forces caused the polymer, black acetal copolymer from the Impaction Fixture to wear off on the lateral surface of the Humeral Stem titanium plasma spray coating. | Device Design | DJO Global sent an Urgent Field Safety Notice dated September 15, 2015, to all affected consignees. The notice identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed become familiar with the new techniques and notify all of their applicable users of the RSP Monoblock and Turon stems of these revised methods. When each user has been notified of these new techniques, they should complete the form that was provided with the notification indicating that they have been informed of these changes and understand the new methods. Consignees were asked to complete and return forms. Consignees with questions should call 512-834-6255. |  | 634 | 634 units | Nationwide Distribution | Open, Classified |  |  | 911 | Not_Computer
Z-0196-2016 | 72448 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MR Surgical Suite II Table, Model Number(s): M8074SS, M8053SS, M0074SS. |  |  | September 23, 2015 | 2015 | GE Medical Systems, LLC | GE became aware of an issue with the table transfer release mechanism of the MR Surgical Suite Option. The issue may prevent patient transfer between the MR and surgical tables. | Process control | GE sent an "Urgent Medical Device Correction" letter dated September 23, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.You may continue to use the table. Monitor for any issues in releasing the table. If you see this issue, contact a GE Field Engineer.For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 23 | 23 (11 US; 12 OUS) Update 11-4-2015: 35 (14 US; 21 OUS) | Worldwide Distribution - US Distribution including the states of : CA, CO, FL, IL, MI, OH, TX and WI., and to the countries of : AUSTRIA, BRAZIL, FRANCE, GERMANY, JAPAN, POLAND and TURKEY. UPDATE 11-4-2015: Additional states: AZ, MN. Additional countries: Canada, China, Czech Republic, South Africa, and United Kingdom | Terminated | March 28, 2016 | 187 | 23 | Not_Computer
Z-0171-2016 | 72321 | No | LIZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD CD3 (SK7) FITC; IVD; Catalog # 349201 and 340542 |  |  | September 30, 2015 | 2015 | Becton, Dickinson and Company, BD Biosciences | Two lots of CD3 (SK7) FITC label contains an error in the Spanish Small text. | Under Investigation by firm | BD sent an "URGENT PRODUCT RECALL" letter dated September 30, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer Actions: 1. Discard all remaining stock of CD3 (SK7) FITC, catalog number 349201, lot number 5148598, and/or catalog number 340542, lot number 5113983, and request replacement product. BD will replace product at no cost.2.As a requirement of our Regulatory Tracking Process, we request that you complete the enclosed Tracking Verification form, even if you no longer have the affected reagents listed in your inventory. Please return the Tracking Verification form to BD Biosciences, per the instructions (mail, fax, or email).We apologize for any inconvenience that this may have caused and are working toward a long-term solution to this issue. If you require further assistance, please contact BD Customer Support at 877.232.8995 (prompt 3, 3) in the United States. For customers outside the US, contact your local BD Biosciences representative or distributor. |  | 120 | 120 units | US (nationwide) Distribution to the states of : TX, GA, IN, TN, CA, KY, WA, IL, MO, NY, CT, MN, NV, KS, WI, NH, FL, PA, OH, OR, SC. OK, NC and LA., and Internationally to Mexico. | Open, Classified |  |  | 120 | Not_Computer
Z-0197-2016 | 69082 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GEMINI PET/CT System, a diagnostic imaging system for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. |  |  | May 01, 2008 | 2008 | Philips Medical Systems (Cleveland) Inc | The reason for this inspection is to determine if the Spindle Block Servo Controllers on the CT subsystem of potentially affected GEMINI PET/CT systems are susceptible to premature failure. | Unknown/Undetermined by firm | Philips field service personnel will conduct customer visits in order to inspect systems. For further questions please call (440) 483-7600. |  | 264 | 264 units | Worldwide Distribution - US Distribution to the states of : AL, AZ, CA, CO, DE, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MO, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, TN, TX, UT, VA, WA, WI, WV, and WY., and to the countries of : Canada, Argentina, Australia, Austria, Belgium, Brazil, China, Denmark, Finland, France, Germany, Hong Kong, Israel, Italy, Japan, Jordan, Luxembourg, Mexico, Netherlands, Poland, Portugal, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey and UK. | Terminated | November 06, 2015 | 2745 | 264 | Not_Computer
Z-0200-2016 | 72115 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Achieve Cables, model 990066. The sterile, single use only Electrical Cable (Model 990066) provides the conduction elements from the proximal end of the Achieve Catheter handle to standard shielded ECG pins that connect into standard electrophysiology recording and pacing equipment. Medtronic Ablation Frontiers Achieve Mapping Catheter is indicated for multiple electrode electrophysiological mapping of the cardiac structures of the heart. |  |  | September 29, 2015 | 2015 | Medtronic Inc. | 64 units of Achieve Electrical Cables were shipped with a potential sterility breach. | Under Investigation by firm | Medtronic sales reps began notifying customers on 9/29/2015. Customers were told of the issue and were requested to quarantine and return un-used product. A follow-up Urgent Medical Device Recall letter was hand delivered beginning 10/6/2015. The letter again described the issue, identified affected product and asked that un-used product be returned to Medtronic. A response form was asked to be returned to Medtronic. Customer with questions can contact Medtronic AF Solutions at 888-843-8301 |  | 64 | 64 units | Distributed in DC and the states of AR, CA, CO, DE, GA, IA, KY, ME, MI, MN, NY, OH, PA, TN, TX, and VA. | Terminated | March 09, 2016 | 162 | 64 | Not_Computer
Z-0202-2016 | 72289 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products Calibrator Kit 9 Calibrator, 2 mL, Product Code 8568040, Unique Device Identifier No. 10758750005031, IVD -- Ortho Clinical Diagnostics, Inc. --- VITROS Chemistry Products Calibrator Kit 9 is used to calibrate VITROS 250/350/950/5,1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System for the quantitative measurement of ACET, CRBM, DGXN, and PHYT. |  |  | September 28, 2015 | 2015 | Ortho-Clinical Diagnostics | The firm is conducting a product correction due to an increase in U90-382 or 6LU condition codes (wash errors) when using VITROS Calibrator Kit 9, Lot 0954. | Under Investigation by firm | Customer letter (Ref. CL2015-191) was sent on 9/28/2015 via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) and /or ORTHO PLUS e-Communications to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them to discontinue use of the affected calibrator lot, discard the remaining inventory of the affected lot and recalibrate VITROS DGXN and PHYT Slides using Cal Kit 9 with an alternative lot of calibrator kit 9 upon availability. Distributor letter (Ref. DL2015-191) was sent on 9/28/2015 via US Postal Service Priority Mail and/or FedEx overnight courier to US Distributor consignees and instructed to notify their consignees of the issue and the required actions. The foreign affiliates were notified via e-mail on 9/28/2015. |  | 3991 | Domestic: 2443 kits; Foreign: 1548 kits | Nationwide, Puerto Rico and foreign distribution to the following countries: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain and Venezuela | Open, Classified |  |  | 3991 | Not_Computer
Z-0201-2016 | 72357 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Novation Element Press-Fit, Cementless, HA Coated, Collared, Standard Offset, Size 9, Femoral Hip Implant Stems, Intended for press-fit fixation |  |  | September 23, 2015 | 2015 | Exactech, Inc. | Eight (8) pieces did not receive the specified Hydroxyapatite (HA) coating. | Under Investigation by firm | The firm, Exactech, issued an "IMPORTANT NOTIFICATION-PRODUCT RECALL ADVISORY NOTICE" dated September 23, 2015, by phone and by mail to their consignees. The notice described the product, problem and actions taken. The consignees were instructed to 1) Immediately cease distribution or use of these products. 2) Extend this information to your accounts that may have this product in their possession. 3) Identify and quarantine any of the subject Novation Element Pres Fit Femoral stem in your inventory. 4) Complete and Return the attached Recall Inventory Response Form to Exactech,per the instructions provided on the notice itself, within 5 business days, via fax to: 352-337-3915 or email kaya@exac.com. Return products to: Exactech Distribution Center, 2411 NW 66th Court, Gainesville, Florida 32653, Attention: Exactech inventory representative.Any questions contact your Exactech inventory representative at kaya@exac.com or 1-800-392-2832. |  | 8 | 8 | US Distribution to states of: AZ, CO, FL, and MD. | Completed |  |  | 8 | Not_Computer
Z-0198-2016 | 72398 | No | MJC | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Antibacterial Hydro Personal Catheter, Male 12 Fr, Product Code 63512, sold in boxes of 30 units |  |  | October 02, 2015 | 2015 | C.R. Bard, Inc. | Incorrect unit labeling. The case label and shelf box are correct, Product Code 63512 and Catheter Size 12 French. The unit label (catheter package) may state Product Code 63516 and Catheter Size 16 French. | Error in labeling | Bard sent a Customer Notification Letter dated September 29 to customers via FedEx with proof of delivery notice. The letter identified the product, the problem, and the action to be taken by the customer. The consignees were instructed to carry out the notification to the user level. Customers with questions were instructed to contact BMD at 1-770-784-6471. |  | 4170 | 4,170 each | Nationwide Distribution including CA, FL, IL, MO, NJ, OR, PA, TX &amp; WA. | Open, Classified |  |  | 4170 | Not_Computer
Z-0199-2016 | 72326 | No | IKK | Isokinetic testing and evaluation system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Robotic hand accessory to InMotion Arm, rehabilitation robot. |  |  | May 05, 2012 | 2012 | Interactive Motion Technologies, Inc. | Unsupervised patient could potentially mis-use the device and remove or rotate a protective cover and expose a potential pinch hazard. | Component design/selection | Interactive Motion Technology ( IMT) sent an Important Patient Safety Information letter dated March 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Summary of Actions to be taken For Hand Robot Cover B, firm recommends continued vigilance for both Adult and Pediatric Patients while in contact with this group of devices. For Hand Robot Covers A or C, at clinician discretion based on expected patient population, firm recommends continued use when appropriate of the IMT supplied Hypertonia Spring Assembly, or if not appropriate, insertion of the IMT Zero Force Spring Assembly. IMT will make this kit available for each Hand Robot upon request, or insert the Assembly when the Hand Robot is shipped to IMT using the special shipping container. If you have any questions, please contact your local IMT Representative or IMT Technical Services at support@interactive-motion.com. |  | 8 | 8 | Worldwide Distribution - US Distribution to the states of : Texas and Tennessee., and to the countries of : China, Brazil, Hong Kong, Austria and South Korea. | Terminated | July 13, 2017 | 1895 | 8 | Not_Computer
Z-0204-2016 | 69712 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillators |  |  | November 19, 2014 | 2014 | Philips Medical Systems, Inc. | 1. Device will perform the weekly automated tests hourly, which could cause the therapy capacitors to degrade sooner than intended and 2. While connected to AC or DC power and with no battery installed or the battery installed has a charge level of less than 10%, the Ready for Use (RFU) indicator will not provide the expected low battery indication | Process control | Philips Healthcare sent an " Urgent Field Safety Notice " dated November 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A hardware upgrade will be provided free of charge for all units affected by this issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:A Philips authorized service provider will service the device and ensure it is operating within specification. Prior to receiving service, you can continue to use your MRx provided that the following steps are followed. These steps will prevent further accelerated degradation of the therapy capacitor, and will eliminate the above described behavior for no/low battery conditions.To correct the problem, set the configuration of the MRx to factory defaults in configuration mode.Please contact your local Philips Service Representative for further questions, or call (978) 687-1501. |  | 1553 | 1553 units | Worldwide Distribution - USA (nationwide), Australia, Austria, Bahrain, Belgium, Brunei Darussalam, Canada, China, Finland, France, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Republic of Korea, Libyan Arab Jamahiriya, Lithuania, Mexico, Namibia, Netherlands, Palestine, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 1553 | Computer
Z-0203-2016 | 72312 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | MaxAn 4.0mm and 4.5mm fixed screws, used in the MaxAn Anterior Cervical Plate System. The MaxAn Anterior Cervical Plate System is a decompression-based technique for cervical spine stabilization. |  |  | September 29, 2015 | 2015 | Biomet Spine, LLC | The packaging of 4.5 mm diameter MaxAn screws identifies/labels the screws as 4.0 mm diameter screws. | Labeling mix-ups | Zimmer Biomet Spine sent an Urgent Medical Device Recall Notice dated September 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If the device has already been implanted, written confirmation of this will be required and the necessary parties will be contacted. They will store returned product in quarantine in an isolated location until they can undergo inspection to verify if they are conforming or not and then will be destroyed or retained. Necessary parts will be destroyed so that it is visually obvious they are no longer functional. For further questions call (303) 443-7500 ext.244. |  | 17 | 17 units | US Distribution to the states of : NY, MO, and FL | Open, Classified |  |  | 17 | Not_Computer
Z-0205-2016 | 71518 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Arthroscopic Energy 50Âż Probe with Suction, Catalog number: AES-50S |  |  | June 17, 2015 | 2015 | ConMed Corporation | Probe programming error. The incorrect program can generate more heat at the surgical site, resulting in an elevated probability of tissue damage. | Software design (manufacturing process) | On June 17, 2015 Con Med Corporation distributed Urgent Device Recall Notification Letters and Recall Reply forms to their customers via courier service. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory for any of the affected lot numbers referenced on the recall notice. If any of the recalled product are in stock, customers should discontinue the use or distribution of these items. If product was further distributed, customers should notify those individuals of this recall action. Customers with questions can call Gregory Connell at 727-399-5276 or email AES50S@conmed.com. Customers should also complete and return the Recall Reply Form by mail to the address listed on the recall notice. |  | 276 | 276 devices | Worldwide - US distribution in the states of FL, OH, and the country of Spain. | Terminated | April 11, 2017 | 664 | 276 | Not_Computer
Z-0208-2016 | 72296 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | e.cam Gamma Camera, an emission computed tomography system. It is intended to detect or image the distribution of radionuclides in the body or organ, using the following techniques: planar imaging, whole body imaging, tomographic imaging for isotopes with energies up to 588keV. |  |  | June 26, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | Potential failure mode related to the radial brake, where the detector was able to drift, led to a re-design of the brake redundancy circuit. The redesigned Brake Redundancy Circuit will force even a defective brake to engage in the unpowered state by applying a reverse charge to the brake and ensuring the brake has demagnetized. | Component design/selection | The firm, Siemens, sent out an undated "CUSTOMER SAFETY ADVISORY NOTICE" on 04 April 2015 to its consignees informing them of the e.com system safety update. The notice described the product, problem and actions to be taken. The notice stated the following:In order for an injury to occur, the following series of events must be present;1. The detector configuration angle must be 180 degrees.2. The detector with the failed brake circuit must be above the patient.3. The detector with the faulty brake circuit must be under power (in motion).4. System power must cycle within a very short time period, preventing the brake circuit from charging.The customers were instructed to continue to use their e.com system and to ensure the safety advisory is placed in the systems instructions for use and to observe patients while system is in use. If they experience difficulty or failure while performing the daily touch pad test, they are advised to discontinue use and contact their local representative. If unintended movement is detected they are advised to remove the patient and discontinue using the system and contact their local service representative. If you have any questions regarding this important safety notice, please contact your local Service representative at the contact numbers provided below." America: 1-800-888-7436" Europe, Middle East, and Africa: +49 9131 940 4000" Asia and Australia: +86 (21) 3811 2121Local Service representative at the contact numbers provided below."America: 1-800-888-7436"Europe, Middle East, and Africa: +49 9131 940 4000"Asia and Australia: +86 (21) 3811 2121 |  | 32296 | 32296 e.cam Gamma Camera systems | Worldwide Distribution: US (nationwide) and countries of: United Arab Emirates, Albania, Argentina, Austria, Australia, Bosnia and Herzegovina, Bangladesh, Belgium, Bolivia, Brazil, Belarus, Canada, Switzerland, Chile, Cameroon, China, Colombia, Costa Rica, Czech Republic, Germany, Denmark, Dominica, Ecuador, Egypt, Spain, Finland, France, United Kingdom, Georgia, Greece, Croatia, Canary Islands, Ireland, Israel, India, Iraq, Iran, Islamic-Republic of, Italy, Jordan, Japan, Korea-Republic of, Kuwait, Lebanon, Sri Lanka, Luxembourg, Latvia, Morocco, Myanmar, Mauritius, Mexico, Malaysia, Netherlands, Norway, Nepal, New Zealand, Oman, Peru, Philippines, Pakistan, Poland, Portugal, Qatar, Serbia, Russian Federation, Saudi Arabia, Sweden, Slovakia, El Salvador, Syrian Arab Republic (Syria), Thailand, Tunisia, Turkey, Taiwan, Republic of China, Ukraine, Uganda, Uzbekistan, and South Africa. | Open, Classified |  |  | 32296 | Not_Computer
Z-0213-2016 | 72290 | No | EBC | Pit and fissure sealant and conditioner. | Dental | Dental | 510(k) | Helioseal F Sealant Refill 1 x 1.25 g, REF 558519AN, and Helioseal F Assortment (kit) 5 x 1.25 g, REF 558518AN. Dental sealant. -- ivoclar vivadent clinical --- Helioseal F Refill is provided in a syringe packed sealed foil bag. Helioseal F is a light-curing, white-shaded fissure sealant featuring fluoride release. Helioseal fissure sealants are suitable for the sealing of pits, fissures and foramina caeca. |  |  | October 07, 2015 | 2015 | Ivoclar A. G. | A defect occurred in the production process of Helioseal F Sealant; the particles in the faulty component seem to be larger than the size defined in the formula. This results in difficulties during extrusion of the material. | Under Investigation by firm | Ivoclar Vivadent sent an Urgent Medical Device Recall Letter dated October 7, 2015, with Response Forms via Certified Mail, Return Receipt Requested. Dealers were also contacted via telephone/e-mail and informed to expect the letter. Dealers are requested to check their stock and return any affected material for credit, plus notify their customers of the recall. For questions, please contact Ivoclar Vivadent Customer Service at 1-800-533-6825. |  | 337 | Domestic: 337 units | Nationwide Distribution including NJ, PA, SC, FL, IN, IA, WI, IL, MO, TX, CA and WA. | Open, Classified |  |  | 337 | Not_Computer
Z-0210-2016 | 72217 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO 400 and AMSCO C Small Steam Sterilizers |  |  | September 03, 2015 | 2015 | Steris Corporation | In AMSCO Small Steam Sterilizers equipped with the CS-iQ feature, the sterilizer software inadvertently prevents the capability to operate the foot pedal, which allows the Operator to open and close the sterilizer door. | Software design | On 9/21/2015 the firm sent Recall Correction Notices to their customers. |  | 32 | 32 units | US nationwide distribution. | Terminated | January 22, 2016 | 141 | 32 | Not_Computer
Z-0151-2016 | 72319 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON-G5 Ventilator with software versions between V2.00 and V2.31. |  |  | April 22, 2014 | 2014 | Hamilton Medical, Inc. | Ventilation and alarms of a HAMILTON-G5 ventilator can be suppressed unintentionally after the activation of a suctioning maneuver by the operator; this situation can occur regardless of the selected patient group (neonatal, pediatric and adult). | Device Design | Hamilton Medical issued a letter dated April 22, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were notified that their Hamilton Medical Account Manager would contact them to coordinate the deactivation of the Ventilation Suppression feature. Customers were asked to review the Medical Device Field Safety Corrective Action.Customers with questions were instructed to call 800-426-6331, Ext 215. |  | 1128 | 1128 | Nationwide Distribution | Terminated | August 24, 2017 | 1220 | 1128 | Not_Computer
Z-0207-2016 | 72463 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROW PICC Kit featuring Chlorag+ard Technology |  |  | October 16, 2015 | 2015 | Arrow International Inc | Possibility that the catheter peel-awaycomponent hub tabs may prematurely detach when the practitioner begins to peel apart the sheath body from thecatheter. Although the issue does not pose a significant risk to the patient, it may cause a delay of treatmentwhen placing the catheter. | Unknown/Undetermined by firm | Arrow International sent an Urgent Medical Device Recall Notification letter dated October 16, 2015, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignees. Consignees were instructed to immediately discontinue use and quarantine any effected products. Consignees were instructed to return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507 Attn: Customer Service or email it to recalls@teleflex.com. A customer service representative will contact consignees with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Arrow International. For consignees with no affected product, should also complete the enclosed Recall Acknowledgement Form.For questions customers should call 1-866-246-6990. |  | 12157 | 12,157 | World Wide Distribution - US including AR, AL, TX, MA, CA, IA, MO, NY, AL, VA, RI, IN, TN, FL, KY, NH, SC, MD, WA, OK, IL, ME, OR, OH, PA, MI, KS, LA, TN, WV, NE, NJ, NV, GA, SD, CT, MS, CO, and Internationally to Canada and Belgium. | Open, Classified |  |  | 12157 | Not_Computer
Z-0209-2016 | 72488 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | iOS Galileo Application Version 1.1.1 or lower that programs the Triton Smart Ankle; 1C66 Triton Smart Ankle, an external prosthetic ankle/foot. |  |  | October 23, 2015 | 2015 | Otto Bock Healthcare GmbH | A software issue can make the foot move into dorsiflexion, unnoticed by the user, who could fall. When sitting, the prosthetic foot can go into the Relief Function to move the footplate towards the ground. When operating machines with foot pedals (e.g. a car) the foot could get stuck under the foot pedal and block it or, with the foot on the pedal, it could cause unintended vehicle acceleration. | Software design | OttoBock sent an Urgent Medical Device Correction letter dated October 23, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The US Agent will notify US customers via letter on October 23, 2015, of the correction on behalf of the Manufacturer. Customers will be requested to contact all users and advise them of the hazards when operating machines or vehicles with foot pedals. They may forward a copy of the insert if needed. Customers will also be requested to contact all users with TF-fittings, where the settings were done via iOS Galileo app., inform the user about this issue, and assure that settings in the device are as expected.The customer will be requested to complete a reply form and return to the US Agent. For further questions call (800) 328-4058. |  | 321 | 273 in US: 48 Foreign | Worldwide Distribution - US (nationwide) and to the countries of : Austria, Belgium, Germany, Israel, Luxembourg and Sweden. | Terminated | March 21, 2017 | 515 | 321 | Computer
Z-0206-2016 | 72161 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | BAUSCH + LOMB Biotrue ONEday lenses (nesofilcon A), Product Order Number Biotrue SVS - 7201, STERILE, Rx Only; The lenses are available in 5 pack, 30 pack and 90 pack sizes; Biotrue ONEday Foil - Part Number 8098001, Biotrue ONEday 5 pack Carton International - Part Number 8101403, and Biotrue ONEday 90 pack Carton International - Part Number 8098203. |  |  | September 03, 2015 | 2015 | Bausch &amp; Lomb Inc Irb | The firm determined that 25 lots had been manufactured with a minor surface defect occurring in the optical surface on each lens. Complaints of Halo or Poor Visual Acuity. | Process control | On 9/3/2015, the distributors received the Medical Device Recall Letter (dated 9/2/2015) via e-mail. Medical Device Recall Letters (dated 9/2/2015) and Customer Acknowledgement Cards were hand delivered to the direct customers via Bausch + Lomb Sales Representatives starting on 9/14/2015 or sent via FedEx if consignee is not a B+L account. The letter informs the accounts of the recall and instructs them to direct any consumers who would like to report a product complaint, arrange for a product exchange, or obtain more information on the product recall to contact the Bausch + Lombs Consumer Affairs Team at 1-800-553-5340 or via e-mail at BLCustomerCare@bausch.com. Customers may also contact Bausch + Lomb Customer Service at 1-800-828-9030 to arrange return of affected product. |  | 183920 | Domestic: 99,260 lenses; Foreign: 84,660 lenses | Distributed US (nationwide), Canada, Japan and Thailand. | Open, Classified |  |  | 183920 | Not_Computer
Z-0219-2016 | 70002 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx monitor/defibrillator |  |  | December 23, 2014 | 2014 | Philips Medical Systems, Inc. | MRx monitor/defibrillator could reboot at an indeterminate time, potentially causing therapy to be interrupted or delayed | Software design | Philips Healthcare issued on 12/23/14 The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.A software upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the Action to be Taken by Customer/User section of the UrgentMedical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to receiving the software upgrade. If the MRx reboots during clinical use, continue to treat the patient according to your organizations protocol, reinitiating therapy if necessary. |  | 75693 | 75,693 units | Worldwide Distribution - US Nationwide and the countries of AFGHANISTAN, ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOWINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CAMBODIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, GHANA, GREECE, GUADELOUPE, HONG KONG, ICELAND, INDIA, INDONESIA, IRAN (ISLAMIC REPUBLIC OF), IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, Jordan, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LAOS, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MACAU, MACEDONIA, MALAYSIA, MALTA, MEXICO, MONGOLIA, MOROCCO, MOZAMBIQUE, MYANMAR (Burma), NAMIBIA, NEPAL, NETHERLANDS, NETHERLANDS ANTILLES, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, UNITED REPUBLIC OF TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, TURKMENISTAN, UGANDA, UKRAINE, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 75693 | Not_Computer
Z-0218-2016 | 72236 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Digital Lightbox, BrainLAB system, image processing, radiological |  |  | September 21, 2015 | 2015 | Brainlab AG | Potentially incorrectly displayed objects when actively deselecting a fused reference dataset. | Software design | Brainlab sent A FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated 8/27/2015 to medical users regarding Content Manager, Patient Browser and the DICOM Viewer. The Product Notification letter decribes the corrective action information and to advise the medical users of the action Brainlab is taking to address the issue. If you require further clarification, the firm advises to contact the local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers)E-mail: support@brainlab.com (for US customers: us.support@brainlab.com)Fax: Brainlab AG: + 49 89 99 15 68 33Address: Brainlab AG (headquarters), Kapellenstrasse 12, 85622 Feldkirchen, Germany. |  | 897 | 897 | Worldwide Distribution - US Nationwide and the countries of: Algeria, Argentina, Australia, Austria, Belgium, Brazil, Cambodia, Canada, Chile, China, Colombia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Lithuania, Malaysia, Malta, Mexico, Moldova, Morocco, Myanmar, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Qatar, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom and Vietnam. | Open, Classified |  |  | 897 | Not_Computer
Z-0217-2016 | 72322 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Envoy 500 system includes the following product reference numbers and names: |  |  | September 22, 2015 | 2015 | ELITech Clinical Systems SAS | ELITech clinical Systems reagents based on Trinder reaction used on the Envoy 500 instrument system are recalled due to potential drug interference. No complaints were received. | Under Investigation by firm | The firm sent recall letters or Technical Bulletins to US consignees on 9/22/15 via registered mail. Technical Bulletin #1022 (for EON 100 users)Technical Bulletin #2008 (for EON 300 users)Technical Bulletin #5064 (for ENVOY users)Technical Bulletin #3009 (for Selectra users)ELITechGroup is not requesting return of the affected product. The method sheets and package inserts will be updated with the new interference claims. When available, the revised labeling will be provided to the customer for customer's use until the new IFU is included in the product packaging. Consignees are requested to return the Confirmation of Receipt form back to the firm by FAX to 401-642-9001 or send email to customercare@elitechgroup.com or send via US Post: Vital Diagnostics, 27 Wellington Rd., Lincoln, Rl.If you have any questions, please contact Vital Diagnostics Technical Support at1-800-345-2822 or 855-ELITECH |  | 13455 | 13455 systems | US Nationwide Distribution | Terminated | March 04, 2016 | 164 | 15813 | Not_Computer
Z-0227-2016 | 72390 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical Display, Model No. MDSC-2224 (24 inch), Model No. MDSC-2226 (26 inch) and Model No. MDSC-2242 (42 inch) |  |  | May 28, 2015 | 2015 | Barco Inc. | User settings were no longer on the monitor's memory after these systems were powered off and back on. | Process design | Barco sent an "Urgent Medical Device Correction Notice" dated May 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This notice contained "Interim Corrective Action" and information about a "Final Corrective Action". The notice included a reply notice to be sent back to Barco and also requested a sub-recall. For further questions, please call (678) 475-8156. |  | 462 | 462 devices in US | US Distribution to the states of : AZ, CA, FL, GA, IL, MA, NE, NJ, OH, OR and TX., and Internationally to Canada. | Open, Classified |  |  | 462 | Not_Computer
Z-0222-2016 | 72495 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker UniVise Spinous Process Fixation Plate System Inserter, Stryker Spine, Non-Sterile. Reference Number 48590100. Manual Surgical Instrument. |  |  | September 08, 2015 | 2015 | Stryker Spine | The Inserter Inner Shaft would not fit through the Inserter. The inner diameter of the cannulated shaft of the inserter was undersized. | Process control | Stryker Spine issued an "Urgent Product Removal Notifications /Customer Response Forms dated September 11, 2015. to all affected customers via Fed Ex priority overnight. The notification instructed customers on how to identify affected product; the related issue with the affected product; potential hazards; risk mitigations; and actions needed to be taken. Customers were asked to examine their inventory and related hospital location inventories; and reconcile any UniVise Inserters with Catalog numbers and Lot numbers (referenced in the notification). Customers were asked to fax a copy of the Customer Response Form to Regulatory Compliance (at 855-632-9049) or email it to Spine-RegulatoryActions@Stryker.com within 5 days of receipt of the notification.For further questions please call (201) 760-8346. |  | 41 | 41 units | Nationwide Distribution. | Open, Classified |  |  | 41 | Not_Computer
Z-0229-2016 | 72053 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Regard Custom Surgical Pack, Item Number 880239, Sterile IV Pack. Custom procedure trays contain surgical instruments and accessories intended for use during a surgical procedure. |  |  | August 25, 2015 | 2015 | Resource Optimization &amp; Innovation Llc | The kits contain pressure monitoring kits or sets which are being recalled as a result of flow restriction issues in the trifurcated IV set. | Component design/selection | ROi sent a Notice of Correction/Removal letter dated August 25, 2015, to all affected customers. The letter notified customers that ROi recently received notice of a voluntary recall from Edwards Lifesciences regarding their Transducer Kit. ROi also sent an email to their customer on September 25, 2015, informing them of the recall. Attached to the email was the component manufacturer's recall notice and a recall notice from the recalling firm. The recalling firm's recall notice instructed their customer to respond to the notice indicating whether they had any of the recalled product remaining and to destroy any remaining product in their inventory.For questions regarding this recall call 417-820-2793. |  | 255 | 255 packs | Product was distributed to one customer in Oklahoma. | Terminated | April 06, 2017 | 590 | 255 | Not_Computer
Z-0221-2016 | 72384 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Extension Set, Model No. 20029E. Used to administer fluid and medications through a needle or catheter inserted into the patient's artery or vein. |  |  | October 09, 2015 | 2015 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Extension Set due to disconnection and leakage issues. | Process control | An urgent medical device recall notification was sent to customers on 10/9/15 who purchased the SmartSite Extension Set. The letter informs the customers that CareFusion has identified potential risks with model code 20029E, lot numbers 1503541, 15015769 and 15045950 due to disconnection and leakage issues. The letter also provides the customers with the potential risk involved and the actions to be taken. Customers are instructed to complete and return to CareFusion the recall response card. Customers with recall related questions are instructed to contact CareFusion Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact Customer Advocacy at (888) 812-3266 or email customerfeedback@carefusion.com. Customers with technical questions regarding technical questions are instructed to contact Technical Support at (888) 812-3229. |  | 21000 | 21,000 units | Distributed US (nationwide) and in Canada. | Open, Classified |  |  | 21000 | Not_Computer
Z-0223-2016 | 72459 | No | MIH |  |  | Cardiovascular | PMA | FLAIR Endovascular Stent Graft, Product Code: FAS08070, FAF08070, Implant Size: 8mmx70mm |  |  | October 19, 2015 | 2015 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) is recalling the Bard Flair Endovascular Stent Graft because of the harm associated with failure to deploy or deployment related issues that could occur during use. | Device Design | Bard sent an Urgent Medical Device Recall Notification letter dated October 19, 2015, to customers who received the Bard Flair Endovascular Stent Grant. The letter informs customers that the harm associated with failure to deploy or deployment related issues that could occur during use may pose an incremental risk to the already inherit risk associated with stent graft procedures. The letter also provides customers with the actions to be taken. Customers were instructed to call Bard Peripheral Vascular's (BPV) Recall Coordinator at 1-800-321-4254 Option #2 Ext 2501 (M-F 6am to 3pm MST). Customers were instructed to fax the Recall and Effectiveness Check Form to BPV at 1-800-994-6772. |  | 292 | 292 units | Nationwide in US | Terminated | April 12, 2016 | 176 | 292 | Computer
Z-0224-2016 | 72224 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Robotic Arm Interactive Orthopedic (RIO) system component PN#205143 (locking pin) and PN# 112220 (connector). Indicated for use in surgical knee resurfacing and total hip Arthroplasty joint replacement procedures. |  |  | July 16, 2015 | 2015 | Mako Surgical Corporation | The locking pin on the RIO Base Array is damaged during assembly onto the RIO Base Array Connector. | Device Design | As of July 16, 2015, Mako communicated by mail the nature of the recall to their Product Specialists and the OR Administrators describing the specifics regarding the issue for correction. |  | 1322 | 1,322 units | Worldwide Distribution-US (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, MI, MS, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, WY and HI, and the countries of Italy, Japan, Hong Kong, Scotland, Germany, Turkey, Greece, Australia, Singapore, Taiwan, and South Korea. | Open, Classified |  |  | 1322 | Not_Computer
Z-0220-2016 | 72097 | [6 empty] | EnVision FLEX/HRP visualization reagent found in the following kits: |  |  | September 04, 2015 | 2015 | Dako North America Inc. | There is a defect in one lot of a buffer, non-specific background staining that has, in some instances, interfered with interpretation of the staining results. | Other | Dako sent an customer notification letter on September 4, 2015, to all affected customers via mail. The letter identified the product the problem and the action needed to be taken by the customer.The customers/distributors are instructed in the Recall Notification Letters to 1) remove any unused kits bearing the affected lot numbers listed in their letter from diagnostic use 2) destroy any unused kits bearing the affect lot numbers listed in their letter, and 3) determine any possible impact of this change in performance on patient results from the affected lots.The letter states actions should be taken within 10 days and that customers/distributors should confirm they have received the information by returning the attached, fully completed Device Recall Form and return it to Dako QA Vigilance by email at dako.dkvigilance@dako.com.The firm lists on their letter for customers to contact their Dako representative in there are any questions. For further questions, call (805) 566-6655. |  | 300 | 300 units | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico and to the states of : TX, SD, CA, FL, MA, MD, WV, NC, OH, DE, MI, NJ, CO, HI, AR, NY, LA, GA, NV, RI, IL, OK, CT, MO, VA, and DC., and to the countries of : Canada, Brazil and Costa Rica. | Terminated | April 25, 2016 | 234 | 300 | Not_Computer
Z-0228-2016 | 72096 | No | GAX | Nonpneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MATResponder Tourniquet; Catalog number MATR: UPC 78353 00008; |  |  | August 31, 2015 | 2015 | Pyng Medical Corporation | Once tested, it is required that the tourniquet be returned to its Ready position then placed in its packaging. Eachtourniquet was tested for functionality, but in some cases, the tourniquet was not returned to its Ready position prior to packaging. | Process control | PYNG Medical sent Recall Notification letters dated August 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed:Please immediately cease use of the MATResponder and MATCombat tourniquet lots specified above. Inform all of your teams to cease use of or distribution of these lots.Please immediately place all MATResponder Tourniquets and MATCombat Tourniquets that you have in inventory into a quarantined location. For further questions please call (604) 303-7964.. |  | 6728 | 6728 | Worldwide Distribution - US Distribution and to the countries of : Colombia, France, Canada, Italy, Germany, Finland, United Kingdom, Australia, The Netherlands, Japan, Hungary, Indonesia, Israel, Mexico, Poland and Spain. | Open, Classified |  |  | 6728 | Not_Computer
Z-0231-2016 | 72220 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Catalog Number/Product Code 6801375, Unique Device Identifier No. 10758750001132, and VITROS 5,1 FS Chemistry System Refurbished, Catalog Number/Product Code 6801890, Unique Device Identifier No. 10758750001644. Intended for use in the in vitro quantitative measurement of a variety of analytes of clinical interest. |  |  | September 10, 2015 | 2015 | Ortho-Clinical Diagnostics | Software Anomaly during ADD Installation on VITROS 5,1 FS Chemistry Systems using Software Versions 2.2.1 through 2.8. The user modified parameters would not be saved as expected, but remained changed to the default values obtained from the ADD disk. | Software design | The firm, Ortho Clinical Diagnostics, sent a "FOLLOW UP to URGENT PRODUCT CORRECTION NOTIFICATION" letter dated 9/10/2015 and response forms to customers via FedEx overnight courier, e-mail (ORTHO PLUS e-Communications) and/or via US Postal Service overnight letter on 9/10/2015. Foreign affiliates were informed by email on 9/10/2015 of the issue and instructed to notify their consignees of the issue and required actions.The letter described the product, problem, and actions to be taken. The customers were instructed to follow the required actions and complete and return the Confirmation of Receipt-Response Required form, no laterthan 9/18/15, by fax at 1.888.557.3759 or 1.585.453.4110, or scan PDF and email to: ConfirmationAdmin@its.jnj.com. Customers with any additional questions, can contact Customer Technical Services at 1-800-421-3311 at any time. |  | 210 | Catalog # 6801375: Domestic 919; Foreign 1251; Catalog 6801890: Domestic 63, Foreign 147 | Worldwide Distribution: US (nationwide) including Puerto Rico, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 210 | Not_Computer
Z-0225-2016 | 72472 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS-IW with Universal Viewer is a device that displays medical images (including mammograms) and data from various imaging sources. Images and data can be viewed, communicated, processed and displayed within the system or across computer networks at distributed locations. Typical users of this system are trained professionals, including but not limited to radiologists, physicians, nurses, medical technicians, and assistants. |  |  | September 28, 2015 | 2015 | GE Healthcare | Inaccurate distance measurements with magnified projection X-ray images. | Software design | An URGENT MEDICAL DEVICE CORRECTION letter was sent to all medical users on 9/28/2015 regarding Centricity PACS-IW with Universal Viewer and Centricity Universal Viewer- Inaccurate distance measurements with magnified projection X-ray images.GE Healthcare explains that for all projection X-rays, GE Healthcare recommends that the user perform the necessary measurements on the standard contact images, and not rely on measurements made on the magnified views unless the measurements have been independently verified for each X-Ray machine installed or manually calibrated using a radio-opaque fiducial of a known size. Attention should be given to prior cases acquired before the application of the safety instructions given in this communication as they may be affected by the safety issue. GE Healthcare will correct all affected systems by providing a software update at no cost to the customer. GE Healthcare advises customers if they have any questions or concerns regarding the notification, to please contact GE Healthcare Service at1-800-437-1171 or their local Service Representative. |  | 5 | Centricity PACS-IW with Universal Viewer Versions 5.0 SP2 and higher 1052 units | Worldwide Distribution-US (nationwide) including the states of AZ, CA, CO, DE, FL, GA, IA, IL, IN, LA, MA, MI, MN, MO, NJ, NY, OR, PA, SC, SD, TN, TX, VA, WA, WI and WV, and the countries of Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Ecuador, France, Germany, Ghana, Hong Kong, India, Israel, Italy, Japan, Kenya, Korea, Malta, Malaysia, Namibia, Netherlands, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Turkey, Taiwan, United Arab Emirates, United Kingdom and Venezuela. | Open, Classified |  |  | 11 | Not_Computer
Z-0241-2016 | 72098 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | DeRoyal Sterile Custom Sterile Custom Surgical Kits containing 3M DuraPrep (TM) Surgical Solution, Catalog No. 8630, Lot Nos: 2016-09EE or 2016-09EF, or 2016-09EG, or Catalog No. 863, Lot No. 2016-08DC: |  |  | August 31, 2015 | 2015 | DeRoyal Industries Inc | Custom Sterile Surgical kits contain 3M DuraPrep (TM) Surgical Solution, Catalog Number 8630, Lot Number 2016-09EE which was recalled due to unintended exposure to the sterilant during package sterilization. | Other | DeRoyal initiated their recall on 08/31/2015 to the end-user level by sending overnight letters to its direct accounts and hospitals. Distributors were asked to put the product in quarantine, place the provided labels on the products and complete and return the spreadsheet of how many they labeled. They were asked to return the spreadsheet even if they did not have inventory. Hospitals were provided with labels to affix to the kits with direction to discard the recalled component and notice of destruction forms to be completed and returned. |  |  |  | SC, FL, OH, NJ, OK, MN, AR | Terminated | April 12, 2016 | 225 |  | Not_Computer
Z-0237-2016 | 72434 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | TEKIA Single Piece Hydrophilic Acrylic Intraocular Lenses, Model 811 |  |  | September 29, 2015 | 2015 | Tekia, Inc. | Tekia is recalling the Single Piece Hydrophillic Acrylic Intraocular Lenses (IOLs) because they may be labeled with the incorrect diopter power. | Error in labeling | Tekia sent an Field Safety Notice dated September 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Customers were advised that Tekia is recalling a limited a number of Single Piece Hydrophilic Acrylic IOLs. If the customer has any questions, the customer is instructed to contact a TEKIA Customer Service or QA Representative at (949) 699-1300. Customers with product complaints or adverse events regarding the use of Single Piece Hydrophilic Acrylic IOLs are instructed to inform TEKIA by phone, email, or fax. |  | 50 | 50 units | Internationally to Vietnam. | Terminated | December 03, 2015 | 65 | 50 | Not_Computer
Z-0239-2016 | 72408 | No | DHX | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Access CEA Assay, Part No. 33200. For the quantitative determination of Carcinoembryonic Antigen (CEA) levels in human serum. |  |  | October 13, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access CEA reagent packs because they were filled incorrectly. | Under Investigation by firm | An Urgent Medical Device Recall letter dated 10/13/15 was sent to all customers who purchased the Access CEA Reagent Kit to inform them that Beckman Coulter has determined that the Access CEA reagent packs (P/N 33200) were filled incorrectly. The packs contain insufficient quantity of reagents in one of the pack wells. The letter informs the customers of the impact and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact Beckman Coulter's Customer Technical Support at http://www.beckmancoulter.com, (800) 854-3633 in the US and Canada. Customers outside of the US and Canada are instructed to contact their local Beckman Coulter representative. |  | 11406 | 11,406 units total (1,660 units) | Worldwide Distribution-US (nationwide) and the countries of Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Croatia, Czech Republic, Egypt, France, Germany, Ghana, Gibraltar, Greece, Guam, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Republic of Korea, Lebanon, Mexico, Morocco, Myanmar, Netherlands, Pakistan, Poland, Portugal, Russian Federation, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Vietnam. | Terminated | April 13, 2016 | 183 | 11406 | Not_Computer
Z-0232-2016 | 72225 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific SQ-RX Model 1010 subcutaneous pulse generator, (S-ICD). |  |  | September 17, 2015 | 2015 | Boston Scientific Corporation | Devices shipped to US in dual-channel rather than US approved single-channel RF Telemetry communication mode. While this does not impact therapy delivery of device; this can impact likelihood of successfully establishing telemetry during first scan of device when using a single-channel programmer. Telemetry can be established with additional scans; once established no further difficulty is expecte | Process design | One consignee was contacted via a Boston Scientific "Medical Device Retrieval" letter dated September 17, 2015. The letter was addressed to Hospital Administrator. The letter described the problem and the product involved in the recall. Informed consignee that their local Boston Scientific sales representative will be retrieving the product form their inventory. For questions they can contact their local sales representative or Boston Scientific Technical service at 1-800-227-3422.A second Boston Scientific "Important Medical Device Information" letter dated October 2015 was hand delivered to physicians starting on 10/15/2015. The letter described the problem and provided information on Clinical Considerations, Affected Population, Recommendations and Further Information. For questions they can contact their Boston scientific representative or the Americas Technical service at 1-800-227-3422. |  | 3 | 3 | Distributed in the US to MA and PR. | Terminated | March 08, 2016 | 173 | 3 | Not_Computer
Z-0128-2016 | 71958 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Villa Sistemi Rotograph EVO 3D X-ray |  |  | August 05, 2015 | 2015 | Villa Radiology Systems LLC | It has been discovered that due to lack of information accompanying documents the product fails to comply with 21CfR 1020.33(c),1020.33(d),1020.33(g),and 1020.330). | Other | Villa Radiology Systems Planned Action:1. The current users manuals have been updated and addendums to the user manuals have been developed for currently installed Rotograph Evo3D and Imax Touch 3D. The revisions and addendums contain all information required under 21 CFR 1020.33(c), 1020.33(d), 1020.33(g), and1020.33(j).2. New Rotograph Evo 3D and Imax Touch 3D systems will ship with the updated user manual.3. Addendums will be provided to all end usersCDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. |  | 8 | Rotograph EVO 3D - 8 installed | US Distribution | Terminated | March 22, 2017 | 595 | 8 | Not_Computer
Z-0238-2016 | 72395 | research use only/microbiology | OTT |  |  | Microbiology | Enforcement Discretion | Focus Diagnostics Anaplasma Phagocytophilum IFA IgM Test Kit, Model No. IF1450M in vitro diagnostic. |  |  | October 02, 2015 | 2015 | Focus Diagnostics Inc | Focus Diagnostics is recalling the Anaplasma phagocytophilum IFA IgM kit because of the potential increase in invalid results. | Device Design | A recall letter dated 10/2/15 was sent to distributors who distributed the Anaplasma phagocytophilum IFA IgM (IF1450M) kit lots 26431 and 27011. The letter informs the customers that Focus Diagnostics received customer complaints of IF1450M with low reactivity, which may result in higher invalid rate. The letter also provides the customers with Focus Diagnostics' recommendations and actions to be taken by the distributors. Customers with questions or require additional information are instructed to contact Focus Diagnostics' Technical Services department at (562) 240-6550 from 7am to 5pm (PST) or send an email to DxTS@focusdx.com. |  | 38 | 38 kits | Non US (worldwide): Czech Republic, Denmark, Germany, Italy, Spain and Colombia | Terminated | January 29, 2016 | 119 | 38 | Not_Computer
Z-0234-2016 | 72386 | No | OUB | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | miraDry System consisting of the MD4000-MC Console, the MD4000-HP miraDry Handpiece, and the MD4000-BT miraDry bioTip and accessories. |  |  | October 09, 2015 | 2015 | Miramar Labs, Inc. | Update to labeling for the miraDry Console to remind users to only use approved lubricants in a thin layer to prevent potential burns and blisters. | Labeling design | Miramar Labs sent Medical Device Correction letters to all users on October 9, 2015 to inform them of the labeling clarification by UAPS First Class Certified mail or courier such as Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that the firm added a new warning to their labeling. The firm included an updated copy of the MD4000-MC User Manual with the new warning included and asked that customers destroy their older copy. Customers with questions were instructed to call Customer SService at 1-855-647-2379 or email info@miraDry.com |  | 277 | 277 units | Nationwide Distribution | Terminated | December 23, 2015 | 75 | 277 | Not_Computer
Z-0236-2016 | 72407 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Small Notch Titanium Reconstructive Plate set (SNTRP). Ti PLATE/ROD/PLATE 231mm 6HOLES/5.0mm ROD/17HOLES. Orthopedic implant. |  |  | October 08, 2015 | 2015 | Synthes (USA) Products LLC | The Synthes SNTRP set was released for sale as a Trauma product in 1996 with specific trauma indications and contraindicated for certain spine applications. In 1999, a module was released containing a small offering of the SNTRP plates. Synthes determined that the product in the SNTRP module is being used in spine surgeries, which are a contraindicated use. | No Marketing Application | Synthes sent an Urgent Notice Medical Device Letter dated October 8, 2015. to all affected customer. The letter identified the product the problem and the action needed to be taken by the customer.Please take the following actions:If you DO have any of the identified devices, please take the following steps:"Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number, and signature in the spaces provided."Return the Verification Form (page 3 of this letter)If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indicating that no affected product has been located. Please include your name, title, telephone number, and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information."Return the documents to Synthes by:"Fax: (866) 240-5364 or"Scan/email: Synthes4642@stericycle.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. Thank you for your attention to this issue. |  | 20494 | 20,494 | US Distribution to the states of :MO, NJ, MN, MS, AZ, CA, LA, OK, VA, OH, AR, FL, PA, VT, IN, AL, WA and NY. | Open, Classified |  |  | 20494 | Not_Computer
Z-0245-2016 | 72037 | No | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | RÂżsch Flexi-Slip Endotracheal Tube Stylet with Soft Distal Tip, Sterile, Rx only, Teleflex Medical. |  |  | August 24, 2015 | 2015 | Teleflex Medical | Due to reports of the plastic coating of the stylet splitting and/or breaking off of the stylet. | Unknown/Undetermined by firm | Consignees were notified by letter on 08/24/2015. They were instructed to immediately discontinue use and quarantine any products with the listed catalog numbers. They were to return the product, complete the enclosed Recall Acknowledgement Form and fax or e-mail to recalls@teleflex.com, even if they do not have any of the affected stock. Distributors were instructed to communicate the recall to any of their customers who have received the affected product and to have the customers return any affected product to them with the completed Recall Acknowledgement Form. |  | 233641 | 233,641 units | Worldwide Distribution -- U.S. Nationwide; International Distribution: Argentina, Austria, Belgium, Belgrade, Cameroon, Canada, Chile, Costa Rica, Croatia, Czech Republic, Ecuador, Denmark, Estonia, Finland, France, French Polynesia, Great Britain, Greece, Guyana, Hong Kong, Ireland, Italy, Jordan, Kuwait, Latvia, Lebanon, Martinique, Netherlands, Panama, Peru, Poland, Portugal, Reunion, Romania, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Venezuela, and the United Arab Emirates. | Open, Classified |  |  | 233641 | Not_Computer
Z-0242-2016 | 72331 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Triton Infusion Pump (model 300000) and Triton fp Infusion Pump (model 400000) (proprietary name: AFF (Mark I) Volumetric Infusion Pump). Used to pump fluids into a patient in a controlled manner. |  |  | October 01, 2015 | 2015 | WalkMed Infusion, LLC | The device might fail to detect air in line while infusing the medication Venofer. This may occur when a drop of Venofer gets stuck in the tubing adjacent to the bubble detector and the Total Volume Limit is programmed higher than the actual bag volume. | Device Design | A Field Notification (letter) was provided to users on 10/01/15 to remind them of the following:use air elimination filters, as appropriate, not to over program into the device the volume to be infused, remove air from the IV bag prior to starting the infusion, issues associated with the use of Venofer or similar drugs. No products are being returned. |  | 4695 | 4695 units | Distributed in the states of AL, CA, CO, DE, FL, GA, IL, MA, MD, MI, MN, MS, NC, NJ, NM, NV, NY, OH, PA, TN, TX, and WA. | Open, Classified |  |  | 4695 | Not_Computer
Z-0244-2016 | 72518 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Cook Medical CXI Support Catheter |  |  | October 21, 2015 | 2015 | Cook Inc. | Cook Medical has received six product complaints associated with reports of the 2.6FR curved catheter being mixed with the 4.0FR straight catheter and vice versa. Investigation revealed the products were mixedduring processing. | Labeling mix-ups | Cook Medical sent an Urgent: Medical Device Recall letter dated October 20, 2015, to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, customer responsibilities, and instructions for responding to the formal recall notification. Should you have any questions or concerns, please do not hesitate to contact Cook Medical Customer Relations at 1-800-457-4500 or 1-812-339-2235. |  | 38 | 38 | Worldwide Distribution - US Distribution to the states of : CA, FL, ID, IL, MA, NY, OK, TX, UT and VA., and to the countries of Canada, Austria, Germany, Spain, France, Hungary, Sweden and UK. | Open, Classified |  |  | 38 | Not_Computer
Z-0235-2016 | 72456 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SSS Reprocessed Zimmer, Synthes, Stryker and Linvatec Drill Bits and Blades |  |  | October 14, 2015 | 2015 | Stryker Sustainability Solutions | It was reported that packaging was not properly sealed. | Employee error | Stryker sent an Urgent Medical Device Recall letter dated October 14, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm's letter states that the use of the affected lots of the recalled product should be discontinued. Other bit and/or blade lots reprocessed by Stryker Sustainability Solutions should be considered as alternatives. A Recall Effectiveness Check Form including the quantity of affected devices shipped is enclosed.The letter states to complete the Recall Effectiveness Check Form and indicate if any devices from the affected lots remain in inventory. The form needs to be completed even if no affected product is found.Return the completed and signed form to the local Stryker Sustainability Sales Representative or email to to SSSPFA@stryker.comIf the form indicates that affected devices remain in inventory, a prepaid shippinglabel will be issued for the return of the product. Customers will receive a credit forall affected devices returned. For further questions, please call (888) 888-3433. |  | 1202 | 1202 | US ( nationwide) Distribution to the states of : MN, OH, CA, IN, WI, FL, GA, MO, IA, MT, LA, KS, TX, TN, WY, MD, NJ, WA, PA, SC, NY, MI, AT, HI, MN, AR, AZ, IL, IA, KY, MS, UT, NC, VA, CO, ME and ID., Internationally to Canada. | Open, Classified |  |  | 1202 | Not_Computer
Z-0243-2016 | 72502 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Amphirion Deep OTW PTA Balloon Catheter: |  |  | October 23, 2015 | 2015 | Medtronic Vascular, Inc. | Medtronic has discovered that select lots of the Amphirion Deep PTA Balloon Dilation Catheter have a potential breach in the pouch seal enclosing the sterilized product. | Equipment maintenance | Beginning October 23, 2015, Medtronic will send the customer letter andconfirmation form to all identified US customers via 2-day UPS mail delivery. Internationally, communications will be sent per locally approved methods. |  | 5895 | 1652 units in US; 4243 units OUS. | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Australia Austria Belgium Brazil China Colombia Costa Rica Czech Republic Denmark Dominican Republic Egypt France Georgia Germany Hong Kong Hungary India Iran Ireland Israel Italy Japan Jordan Kazakhstan Korea Kuwait Latvia Lebanon Lithuania Luxembourg Mexico Netherlands Norway Oman Pakistan Panama Peru Philippines Poland Portugal Russian Federation Saudi Arabia Singapore Slovakia Slovenia South Africa Spain Sri Lanka Sweden Switzerland Syria Taiwan Thailand Trinidad and Tobago Turkey United Arab Emirates United Kingdom and Uruguay.. | Terminated | July 31, 2017 | 647 | 5895 | Not_Computer
Z-0233-2016 | 72462 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Myelo-Nate CSF Sampling Kit, Neonatal/Pediatric cerebralspinal fluid sampling sterile tray, including a lumbar puncture needle, four collection vials, an iodine swabstick package, a fenestrated drape, a polylined drape, three gauze pads, and one spot bandage. |  |  | October 16, 2015 | 2015 | Utah Medical Products, Inc | Utah Medical is conducting a recall of Myelo-Nate CSF Sampling Kit due to an internal post-sterilization inspection that found a sterile finished tray where the seal between the Tyvek lid and the PETG tray on oneside of the tray had been compromised. | Package design/selection | Consignees notified on 10/15/15 via letter and email, instructed to return product to distributor. |  | 21403 | 21,403 trays | Worldwide Distribution-US (nationwide) including Puerto Rico and the states of CA, CO, MD, WI, FL, GA, IL, IA, ID, IN, KS, KY, LA, MI, MN, MO, MS, NC, NE, NJ, NV, OH, OK, OR, PA, SC, TN, TX, VA, WA, and WI, and the countries of Canada, United Kingdom, Saudi Arabia, Jordan, United Arab Emirates, Kuwait, and Costa Rica. | Open, Classified |  |  | 21403 | Not_Computer
Z-0250-2016 | 72500 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | NovoCut Suture Manager; Ceterix catalog number CTX-001. |  |  | October 13, 2015 | 2015 | Ceterix Orthopedics, Inc. | The NovoCut Suture Manager device may cut the suture during knot tensioning and this may cause the suture knot to unravel. | Device Design | Ceterix Orthopaedics sent an Urgent Field Safety notification letter dated October 13, 2015, to all affected sites. The product was identified, reason for recall explained and letter contains list of product lot numbers affected. Customers were instructed to identify and quarantine the affected product, contact Ceterix Orthopaedics to arrange a no cost return, and complete and return the attached acknowledgement form. Customers with questions were instructed to call 1-888-305-0615 or email: docs@ceterix.com. For questions regarding this recall call 650-396-3478. |  | 1364 | 1364 units | Nationwide Distribution | Terminated | February 12, 2016 | 122 | 1364 | Not_Computer
Z-0251-2016 | 72376 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Shimadzu C-arm Table, Catalog No. IVS-110, Medical Imaging Table. |  |  | September 30, 2015 | 2015 | Shimadzu Medical Systems | The C-arm may interfere with the table top under certain circumstances. | Software design | An Urgent Voluntary Medical Device Recall notice was sent on 9/30/2015 to all the customers who purchased the Shimadzu C-arm Table (Catalog No. IVS-110). The notification informs the customers of the description of the issue, the actions to be taken by Shimadzu to correct the issue, and the actions to be taken by the customers. Customers with questions or concerns about the device correction, are instructed to contact their local Shimadzu representative or contact ShimadzuNational Technical Support at (800) 228-1429. |  | 26 | 26 units total (4 units in US) | Worldwide Distribution. US (Texas, Wisconsin, Ohio, and Connecticut), Iran, India, Saudi Arabia, Serbia, Thailand, Germany, Croatia, Turkey, Lebanon, Korea, and China. | Open, Classified |  |  | 26 | Not_Computer
Z-0249-2016 | 72381 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics, MIS MODULAR DISTAL CAPTURE TRIATHLON MIS INSTRUMENTS, REF 6541-5-723, Lot AdminECR19205, Non-Sterile, The Triathlon MIS Distal Capture is used when a surgeon elects to use a capture for the distal femoral resection in a Triathlon primary MIS TKA. Once the Triathlon MIS distal Resection Guide (cat. no. 6541-5-721/722, for left or right preparation, respectively) has been secured to the femoral, the MIS distal capture can optionally be attached to the distal resection guide to complete the cut, per Trithion MIS surgical protocol TRI ATH-SP-5. Surgical instrument for knee prosthesis implantation. |  |  | September 29, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker is recalling various lots of the Trithion MIS Modular Distal Capture because complaints were received that specific lots of this product may disassociate during use. | Nonconforming Material/Component | Stryker notified their Branches/Agencies of this recall by email on September 29, 2015 and notifiction letters and product accountability forms were sent via UPS (with return receipt) on September 29, 2015 to branches and on September 30, 2015 to hospitals. Stericycle wil be handling the returns. Customers with questions were instructed to call 201-831-5826.For questions regarding this recall call 201-831-5272. |  | 3444 | 3,444 units | Nationwide Distribution including AZ, DE, ID, IL, KS, LA, MI, NH, NM, NJ, NY, OH, PA,TX, VA, VT, WA and WI. | Open, Classified |  |  | 3444 | Not_Computer
Z-0247-2016 | 71998 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Maquet Getinge Group, Adjustable Positioning Elements (1007.28A0) Manufacturer: Maquet GmbH, Germany. Patient positioners. |  |  | August 17, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | Two serious thromboembolic events were reported involving the Adjustable Positioning Elements (Part Number 1007.28A0) in conjunction with Universal Frame (Part Number 1007.24F0) for lumbar spine surgery. The hip elements may have a negative effect on patients- venous backflow under certain conditions and may have been a contributing factor to thromboembolic events. | Under Investigation by firm | MAQUET sent an "Urgent Product Recall Medical Device Removal" notice and "Medical Device Removal Response Form" dated August 17, 2015 via UPS to affected customers. The notification identified the issue and the affected product. The notification instructed customers to read the notification and disseminate the information provided to all the users of the device; examine their inventory to determine if they have the Adjustable Positioning Elements; if so, remove the affected products, quarantine them and keep them in a secure location for shipment back to Maquet; and fill out and sign the accompany response form and send via email or fax to Maquet. If the customers have any questions regarding this field action, they were instructed to please contact their local MAQUET Representative or call MAQUET Customer Service at 1-888-627-8383 (Press option 2, then followed by option 3) Monday through Friday, between the hours of 8:00 a.m. and 5:00 p.m. EDT. |  |  | United States (3 sets); Rest of World (47 sets) | Worldwide Distribution - US including AL, CA, MO and Internationally to Australia, Brazil, Canada, China, Croatia, Denmark, France, Germany, India, Italy, Japan, Mexico, Netherlands, Russia, Singapore, South Africa and United Arab Emirates. | Terminated | October 19, 2016 | 429 |  | Not_Computer
Z-0252-2016 | 72160 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Smooth Muscle Action (SMA) Immunohistochemistry reagent. Catalog numbers: CME305A-0.1ml; CME305B - 0.5 ml, PME305AA - 6.0 ml prediluted and ready to use. They are the same product with the same intended use. |  |  | September 08, 2015 | 2015 | Biocare Medical, LLC | Certain catalog numbers and lots of Smooth Muscle Actin may stain some cell types that are typically expected to be negative for SMA. | Device Design | Biocare sent a Medical Device Recall Notification letter dated September 8, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue and dispose of all affected product. Customers were asked to complete an return the enclosed Return Response Form. Customer's accounts would be credited for all affected product.For questions regarding this recall call 925-603-8033. |  | 151 | 151 units | Worldwide Distribution - US including CO, FL, OH, SC, AZ, AL, FL, CT, NJ, CA, DC, KY, NC, TX, ID, MA, MD, OR, LA, PA, KS, NY, UT, GA and Internationally to Spain, Sweden, Switzerland, Hungary, Italy, Germany, Turkey, Canada, India, Bolivia, Mexico, Indonesia and Australia. | Terminated | January 11, 2016 | 125 | 151 | Not_Computer
Z-0253-2016 | 72204 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | MILEX PESSARY KIT; Model Numbers: MXPES00; MXKPES00; MXKPREG03; MXKPGSK07 |  |  | September 14, 2015 | 2015 | CooperSurgical, Inc. | The products have been identified to have incorrect size prints on the MILEX PESSARY KIT packaging. While the Pessary pouch is marked with correct size, an incorrect label onthe packaging does not meet the required release specification. | Labeling mix-ups | Cooper Surgical sent an Urgent Medical Device Recall letter dated September 14, 2015 via FedEx with confirmed delivery receipt. The letter identified the affected product, problem and actions to be taken. Customers were asked to discontinue use of the affected products and complete the Acknowledgement and Receipt form for a free replacement. For questions contact Product Surveillance at 203.601.5200 ext. 3300. |  | 144 | 144 units | Worldwide Distribution - US Nationwide in the states of AR, CA, CO, GA, IL, KY, LA, MA, MD, MI, MT, NC, NY, OH, OK, PA, RI, SD, VA, WA, WI,WV and in the countries of FRANCE and VIETNAM. | Terminated | January 08, 2016 | 116 | 144 | Not_Computer
Z-0255-2016 | 72301 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, SIGNA 3.0T. |  |  | August 31, 2015 | 2015 | GE Healthcare | When performing head or neck scans, the currently displayed SAR values could be lower than the actual SAR in the head as predicted by SAR modeling. The predicted HEAD SAR value from the modeling suggests the actual SAR in the head could exceed IEC60601-2-33 limit of 3.2 W/kg for some scans. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated August 31, 2015. The letter was addressed to Hospital Administrators/Risk Managers, Radiology Department Managers, &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 19 | 19 | Worldwide Distribution - US: Nationwide including DC, PR and GU and the countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BERMUDA, BOSNIA HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, GREECE, HONG Kong, HUNGARY, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, KAZAKHSTAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, LUXEMBOURG, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PERU, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VENEZUELA. | Open, Classified |  |  | 1509 | Computer
Z-0262-2016 | 69714 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | November 19, 2014 | 2014 | Philips Medical Systems, Inc. | MRx Defib can be susceptible to one or both issues: 1. The C02 Inlet Port associated with end-tidal carbon dioxide (EtCO2) monitoring on MRx Monitor/Defibrillators can be pushed into the MRx housing, making it inaccessible. 2. The handle can separate from the MRx housing due to breakage of mounts on the rear case. | Device Design | Philips Heatlhcare issued on November 19, 2014, an Urgent Medical Device Correction notification/Field Safety Notice. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.A software upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to receiving the hardware upgrade provided you follow the steps given.For questions regarding this recall call 978-687-1501Customers with questions were instructed to contact their local Philips representative. |  | 27918 | 27,918 units | Worldwide Distribution - US (nationwide) and Internationally to ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CROATIA (local name: Hrvatska), CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NETHERLAND ANTILLES, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UGANDA, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, AND YEMEN. | Open, Classified |  |  | 27918 | Computer
Z-0263-2016 | 72171 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Samsung GM60A-40S |  |  | September 14, 2015 | 2015 | NeuroLogica Corporation | When the system is driving over a non-conductive floor and then over a conductive, grounded part, proper control of the drive behavior may not be possible. The system can be stopped by releasing the drive handle or by pressing the emergency switch. Proper behavior is re-enabled when the system is rebooted. | Device Design | An Urgent - Field Safety Notice identifying the affected products, problem involved, root cause, and potential hazards, was sent to customers. In the event the issue occurs, users can release the drive handle or press the emergency switch and the unit will stop. Then, reboot the system. If unintended movements occur, please contact local representative. To correct the issue, service engineers will visit to apply the corrective action. For further information or support concerning this problem, contact local representative. |  | 7 | 7 devices | Distributed to the states of NC, OH, MA, and MO. | Terminated | January 19, 2017 | 493 | 7 | Not_Computer
Z-0261-2016 | 72297 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Tapers LFIT V40 Vitallium femoral heads. hip prosthesis component. |  |  | September 08, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Complaints that the femoral head could not be assembled with its corresponding V40 stem trunnion at the time of surgery. | Process design | Stryker sent an Urgent Medical Device Recall Notifiications dated September 9 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are asked to complete the attached form and return within 5 days by using email strykerortho6632@stericycle.com or FAX 1-888-965-5803.For further questions please call ( 201) 831-5000. |  | 201 | 201 | Nationwide Distribution | Terminated | July 25, 2017 | 686 | 201 | Not_Computer
Z-0264-2016 | 72475 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE VC150 Vital Signs Monitor |  |  | October 22, 2015 | 2015 | Innokas Medical Oy | Monitor may shut down unintentionally without restarting. | Software design | The firm, Innokas Medical, sent a "Urgent Field Safety Notice(s)" dated October 22, 2015 and October 23, 2015 by e-mail to their consignees. The notice informed the distributor of the device, about the deficiency, and correction of the device. The notice included the instructions regarding what to do with the product and information that the devices will be updated. The consignees were instructed to complete and return the Medical Device FSCA/Recall Return Response Acknowledgement and Receipt Form via fax to +358 8 562 3151, mail to: Innokas Medical Oy, Vihikari 10, FI-90440-Kempele, Finland or email to: feedback@innokasmedical.fi.If you have any questions, contact the QA/RA Director at +358 44 562 3131. |  | 424 | 424 units (38 in US) | Worldwide Distribution: US (nationwide) and Europe. | Open, Classified |  |  | 424 | Not_Computer
Z-0270-2016 | 72413 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Radius 5.5 to 3.5 Rod to Red Connector, REF 48664530. Intended to provide additional support during fusion using autograft or allograft in skeletally mature patients. |  |  | September 25, 2015 | 2015 | Stryker Spine | Stryker Spine is recalling Radius Rod to Rod Connector 5.5 to 3.5 mm (Catalog #486614530) because it does not have an associated 510(k) submisstion and subsequent clearance, which is required as this is a class 2 device. | No Marketing Application | Stryker Spine initiated this recall by sending an "Urgent Product Removal Letter and Product Accountability Form dated September 25, 2015 to the Branch Manager, Agency Manager, Quality Contact via FedEx priority overnight. |  | 42 | 42 units | Nationwide Distribution including the states of AK, AL, AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, MT, NC, NM, NY, OH, OK, OR, PA, TN, TX, UT, WA and WI. | Open, Classified |  |  | 42 | Not_Computer
Z-0268-2016 | 72550 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Perfusion Tubing Packs, Models TL7M64R3, TL7P17R1 &amp; HY2W85R9, Sterilized using Ethylene Oxide. |  |  | October 28, 2015 | 2015 | Medtronic Perfusion Systems | This recall is being initiated due to a potential breach of the Tubing Pack Header Bags manufactured in these select lots. The seal in question is between the Tyvek and the Poly side of the header bag. | Packaging process control | Medtronic sent an Urgent Medical Device Recall letter dated October 2015 to all affected consignees. The letter described the problem and the product involved in the recall. Consignees were advised to quarantine and return all affected product and to complete and return the Customer Confirmation certificate. For questions contact your Medtronic representative or contact Medtronic's national Answering service at 800-633-8766. |  | 496 | 496 | Nationwide Distribution including US: AL, IL, IN, and WI. | Open, Classified |  |  | 496 | Not_Computer
Z-0265-2016 | 72406 | No | LNS | Extracorporeal shock wave lithotripter. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MODULARIS VARIOSTAR; Lithotripter device designed to treat urolithiasis. |  |  | September 24, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Display freeze of MODULARIS hand control results in information not being updated on the display. Current treatment data is not shown to the user. | Software design | Siemens sent a Field Safety Notice (XP044/15/S) dated September 24, 2015 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 15 | 15 | Distributed in PR and the states of MO, NC, GA, MS, LA, and KY. | Open, Classified |  |  | 15 | Not_Computer
Z-0267-2016 | 72368 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Biomet 3i Narrow Right Angle Square Driver Tip. Driver tips are attached to a surgical drilling unit to carry components such as implant mounts, healing abutments and screws for placement and removal. For dental and restorative procedures. |  |  | August 10, 2015 | 2015 | Biomet 3i, LLC | RASQ3N, a square, abutment screw driver tip has been labeled and distributed as RASH3N, a hexed, abutment screw driver tip. Additionally, the RASQ3N received the RASH3N's laser marking "RASH3N" during manufacturing. | Labeling mix-ups | Beginning on or about 08/10/2015 customers will be notified in writing via FEDEX and instructed to to check their inventory for the identified product and to return any unused quantiles along with the recall response. Biomet 3i will promptly issue replacement product. |  | 103 | 103 pieces | AK, AZ, CA, CO, CT, FL, GA, IA, ID, IL, KY, LA, MA, MD, MI, MS, NC, NE, NJ, NY, OH, OR, PA, SD, TN, TX, UT, VA, WA, and District of ColumbiaArgentina, Australia, Belgium, Canada, Colombia, Greece, Hong Kong, Ireland, Italy, South Korea, and Kuwait. | Open, Classified |  |  | 199 | Not_Computer
Z-0269-2016 | 72396 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Ventilator Tubing |  |  | September 23, 2015 | 2015 | Instrumentation Industries Inc | As a distributor of tubing manufactured by Smooth Bore Plastics, we are initiating a product recall for specific lots of KC 036, KC 060 and KC 072 reusable ventilator tubes due to a possible manufacturing defect: A few tubes may have the potential for delamination to occur between the layers of the tube wall which, should it occur, would result in leakage | Nonconforming Material/Component | Instrumentation Industries notified consignees by e-mail and telephone of the recall. The firm also sent a Notice of Product Recall letter dated September 23, 2015, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were asked to determine if they have any of the affected product and contact the firm to make arrangements for return to Instrumentation Industries, Inc. Replacement product will be sent as soon as possible after all quality inspections have been completed. If product was further distributed consignees were asked to provide recall information to their customers. Customers with questions should call 412-854-1133, ext 433. |  | 567 | 567 | Worldwide Distribution - US including PA, FL, GA, WA, OH, CA, KA and Internationally to Canada. | Open, Classified |  |  | 567 | Not_Computer
Z-0272-2016 | 72449 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Easy Dial Oxygen Regulator, 168708D and 168715D. |  |  | October 08, 2015 | 2015 | Precision Medical, Inc. | These lots of Oxygen Dial Regulators could have the settings label off by one position. When the indicator arrow points to OFF position, the regulator is actually on the half liter setting (0.50), for the 15 liter. For the 8 liter, instead of being in the OFF position, the regulator will be in the quarter liter setting (0.25). | Error in labeling | The firm sent "Attention Field Corrective Action" letters, dated Oct 8, 2015, to customers. The letter identified the affected device and explained the reason for the recall. The letter provided instructions on how to check the regulator; and asked customers to call the number provided for a replacement if a regulator is labeled incorrectly. |  | 315 | 315 | Distributed in the US to the states of PA, MA, MO, MT, ID, MI, MN, ND, KY, NY, NC, CA, AL, AZ, LA, OH, TX, GA, IN, and OK. | Open, Classified |  |  | 315 | Not_Computer
Z-0273-2016 | 72477 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ECG Out Cable,for use with a Philips ALS monitor/ defibrillator and bedside monitor for Synchronized cardioversion and Transcutaneous pacing. M1783A/M5526A 12-pin Sync cables, used with HeartStart MRx and HeartStart XL and HeartStart XL+ |  |  | November 19, 2014 | 2014 | Philips Electronics North America Corporation | When a Philips monitor/defibrillator is receiving an ECG signal from an auxiliary bedside monitor via a sync cable, the following can occur if the monitor/defibrillator experiences interference from electrical fast transients (EFTs) while connected to AC power: On the HeartStart MRx and HeartStart XL, EFT noise can be misinterpreted as an R-wave. On the HeartStart XL+, EFT noise can disable ECG mo | Device Design | Philips sent an Urgent - Field Safety Notice dated November, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips plans to take to remedy the issue. A replacement cable will be provided free of charge for all units affected by this issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the: Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: Prior to receiving a replacement sync cable, you can continue to use affected sync cables, provided that they are only used to connect the ECG Out jack on the monitor/defibrillator to the ECG (input) port on a bedside monitor. The issues described in this letter do not impact this application of the sync cable.Affected sync cables should not be used to connect the ECG Out jack of a bedside monitor to the ECG (input) port on the monitor/defibrillator.For further questions please call (800) 722-9377. |  | 18604 | 18,604 | Worldwide Distribution - US (nationwide) and to the countries of : Argentina, Australia, Austria, Bahrain, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Ethiopia, Finland, France, Gabon, Germany, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kenya, Republic of Korea, Latvia, Lebanon, Macau, Malaysia, Mexico, Netherlands, New Zealand, Nigeria, Norway, Oman, Philippines, Poland, Portugal, Qatar, Reunion, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela and Yemen. | Open, Classified |  |  | 18604 | Not_Computer
Z-0212-2016 | 72409 | No | MCX |  |  | Cardiovascular | PMA | RotaWire Elite and wireClip Torquer, Guidewire and Guidewire Manipulation Device, Extra Support. REF H802233301 (5-pack outer package), REF H802233300 (single unit). |  |  | October 09, 2015 | 2015 | Boston Scientific Corporation | Boston Scientific is recalling its recently released RotaWire Elite Guidewire and wireClip Torquer Guidewire Manipulation device that is used in conjunction with the Rotablator Rotational Atherectomy System because of complaints of wire fracture. | Component change control | The firm, Boston Scientific, sent an "Urgent Medical Device Recall Removal - Immediate Action Required" letter dated 10/9/2015. Letters were sent via overnight delivery (FedEx or other equivalent service). The letter stated the issue, identified affected devices, stated that further distribution or use of the devices should cease immediately, and affected product should be returned to Boston Scientific in accordance with the enclosed recall instructions. Complete and return the Recall Instructions via email: MapleGroveFieldActionCenter@bsci.com or Fax to: Field Action Center 1-866-213-1806. Healthcare professionals and consumer may report serious adverse events (side effects ) or product quality problems with the use of this product to Boston Scientific by calling 1-800-811-3211.If you have any questions, please contact Boston Scientific Quality Systems at 763-494-1133, email: MapleGroveFieldActionCenter@bsci.com or your local sales representative can answer any questions that you may have regarding this recall removal. |  | 955 | 955 single units (191 5-packs) | Worldwide Distribution: US Distribution in the states of AZ, CA, DE, FL, GA, IL, LA, MD, MI, MN, NE, NY, PA, SC, TX, WA, and WI, and the countries of Belgium, Germany, Great Britain, Italy, Japan, and South Africa. | Open, Classified |  |  | 955 | Not_Computer
Z-0274-2016 | 72423 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom 100 Low Bed (GPAC) AC Powered adjustable hospital bed |  |  | October 07, 2015 | 2015 | Hill-Rom, Inc. | Hill-Rom is voluntarily recalling the Hill-Rom 100 Low Bed due to complaints of patient injury when exiting the beds set at the lowest position setting. | Device Design | Hill-Rom initiated Phase 1 of the recall on October 2, 2015, which included notices of the recall being issued to customers via certified mail, which included safety information on how to properly use the bed, and instructions to forward the information on if devices have been transferred. Phase 2 for the recall will include the distribution of caps for the edges of the bed to be placed on by the customers. Customers may contact Hill-Rom Technical Support at 800-445-3720. |  | 12251 | 12251 units | Worldwide Distribution - USA (nationwide) to the states of : AL, AR, AZ. CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI and WV., and to the countries of : Canada and France. | Open, Classified |  |  | 12251 | Not_Computer
Z-0275-2016 | 71121 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GMK Hinge knee prosthesis; Model Number: 02.09.TA410; Orthopedics and Orthopedic Surgery |  |  | April 17, 2015 | 2015 | Medacta Usa | Incorrect parts: Packages of the GMK Hinge Tibial Augmentation size 4-10 mm have 5 mm screws instead of 10 mm screws. | Employee error | On 4/16/15 all distributors with recalled product were immediately notified of the recall by phone. On 4/20/15, all distributors were further notified of the recall via email. The URGENT MEDICAL DEVICE RECALL letter dated 4/17/2015 the firm described the initiation of a recall for regarding the GMK HINGE Tibial Augmentation Size 4-10mm Reference 02.09TA410 / Lot 145129.The firm provided the following recall actions to be performed:1.Check your inventory for the device listed above. If you find it, IMMEDIATELY place it in quarantine and DO NOT use it in surgery.2.If you have further distributed this device to other customers please notify me and initiate their return to Medacta USA.3.We will provide you with an update on a replacement within the next few days.The firm also provided a contact for customers requiring additional information.Mike Loiterman Director of Regulatory, Quality and Compliance Medacta USA Phone: 312548-9971Fax: 312-546-6881MLoiterman@medacta.us.com |  | 4 | 4 units (US) | Distributed in the states of IL, OH and UT, and the countries of in Austria, Spain, Australia, France and Switzerland. | Terminated | January 05, 2017 | 629 | 4 | Not_Computer
Z-0276-2016 | 72469 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Synthes Application Instrument for Sternal ZIPFIX. Indications for use include primary or secondary closure/repair of the sternum following sternotomy or fracture of the sternum to stabilize the sternum and promote fusion. |  |  | October 19, 2015 | 2015 | Synthes (USA) Products LLC | Certain lots of the Application Instrument for Sternal ZIPFIX that were previously recalled on August 19, 2015 and subsequently repaired and returned to customers between August 26, 2015 and September 23, 2015. The screws on the repaired instruments have the potential to break or loosen. DePuy Synthes is requesting the return of the repaired device in exchange for a new device. | Device Design | Communication to the accounts having record of receiving this product was sent on October 19, 2015. |  | 267 | 267 | Distributed in the states of KY, UT, CA, WI, OH, CO, IL, TX, OK, WA, MI, LA, MA, PA and NE. | Open, Classified |  |  | 267 | Not_Computer
Z-0278-2016 | 72285 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | HARMONIC ACE+ 7 Laparoscopic Shears with Advanced Hemostasis (23cm, 36cm, 45cm length), Ethicon Endo-Surgery. Indicated for soft tissue incisions when bleeding control and minimal thermal injury are desired. |  |  | September 23, 2015 | 2015 | Ethicon Endo-Surgery Inc | Complaints of holes in Tyvek which can compromise the sterile barrier provided by the packaging. | Device Design | On October 5, 2015, the firm sent Urgent Medical Device Recall Removal Letters to their customers. |  | 441490 | 441,490 pieces | Distributed US (nationwide) including the District of Columbia &amp; Puerto Rico and in the countries of ARUBA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, ECUADOR, HONG KONG, INDIA, JAPAN, KOREA, MALAYSIA, MEXICO, PANAMA, PERU, PHILIPPINES, SINGAPORE, SINGAPORE, THAILAND, TAIWAN &amp; VIETNAM. | Open, Classified |  |  | 441490 | Not_Computer
Z-0279-2016 | 72552 | No | PAH | Surgical mesh. | General &amp; Plastic Surgery | Obstetrics/Gynecology | 510(k) | MiniArc Pro Single incision Sling System, Model number / Product Reference 9000261, Lot 936405xxx. Designed to treat female stress urinary incontinence. |  |  | October 26, 2015 | 2015 | Astora | The lot number on the outer carton label did not match the lot number printed on the tray label or identification labels inside the box. The identification labels inside the box are placed on medical records and other documents associated with the procedure. Use of incorrectly identified labels may prevent direct traceability of the implanted unit in the future. | Labeling mix-ups | Consignees were hand delivered by sales representatives a Astora Women's Health "Urgent Recall Notice". The letter was addressed to Physician, Health Care Professional, Nurse &amp; Risk Manager. The letter described the problem and the product involved in the recall. Consignees were advised to check their inventory and return all affected product, notify those that might be affected and to complete and return the Acknowledgement Form. For questions contact ASTORA Women's Health Customer Service at 1-844-879-0787. |  | 52 | 52 | Distributed in the states of AZ, CA, DE, IL, KS, NE, NJ, NY, OH, PA, TX | Terminated | March 09, 2016 | 135 | 52 | Not_Computer
Z-0277-2016 | 72418 | No | GBW | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Taut Intraducers 10/BX7.5 FR 3.5, Product Code PI-93. To provide an access port into the peritoneal cavity. |  |  | October 14, 2015 | 2015 | Teleflex Medical | The defective component can cause gas leakage through the device. This can result in loss of pneumoperitoneum, potentially necessitating removal of the defective device and insertion of another device. | Nonconforming Material/Component | Recall letters were sent to customers and distributors. The letter requested that the accounts discontinue use and quarantine for return any product. The distributors were requested to conduct a sub-recall. The letters included a Recall Acknowledgement Form which was to be returned to Teleflex. |  | 2050 | 2050 ea. | Nationwide Distribution including the states of AK, AZ, AR, CA, CT, FL, GA, ID, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MO, MT, NE, NJ, NY, NC, OH, OK, PA, TN, TX, VA, VT, and WI. | Open, Classified |  |  | 2050 | Not_Computer
Z-0287-2016 | 72378 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Needle Guide Starter Kit, Reusable non-sterile bracket with Infiniti Plus (12, 14, 16, 18GA) needle guides, Rx ONLY. REFs 674-030, 674-030-T, 674-031, 674-031-T, 674-035, 674-036, 674-044, 674-052, 674-059, 674-061. Large Gauge Infiniti Plus Needle Guides, used in needle-guided (or catheter) procedures with diagnostic ultrasound transducers. |  |  | October 21, 2015 | 2015 | Civco Medical Instruments Co. Inc. | Sterility of the product cannot be assured. | Process design | Civco Medical Solutions sent an Urgent Medical Device Recall letter dated October 19, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all affected needle guides. If the needle guide is contained within a kit, only the needle guide from the kit needs to be returned. The communication contains a response form to be completed and returned by the customer. The customer filling out the response form is asked to acknowledge receipt of the recall notice and indicate the number and type of affected product they are returning. Customers withh questions were instructed to call 800-445-6741 or 319-248-6757 or email order@civco.com.For questions regarding this recall call 319-248-6502. |  | 508 | 508 boxes | Worldwide Distribution - US (nationwide) and Internationally to AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHILE, CHINA (MAINLAND), CHINA (TAIWAN), CZECH REPUBLIC, DENMARK, ECUADOR, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, ISRAEL, ITALY, JAPAN, KUWAIT, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PUERTO RICO, QATAR, SOUTH AFRICA, ROMANIA, SAUDI ARABIA, SINGAPORE, SOUTH KOREA, SPAIN, SWITZERLAND, TRINIDAD AND TOBAGO, TURKEY, UGANDA, UKRAINE, and UNITED KINGDOM. | Open, Classified |  |  | 2624 | Not_Computer
Z-0281-2016 | 72506 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker PCA 10 Degree Hooded Acetabular Insert ID 22mm, |  |  | October 15, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Six (6) PCA 10 Degree Hooded Acetabular Insert-Duration 22 mm (ID) have been distributed in the US, although the 22 mm size was inadvertently omitted from 510(k) #K963612. | Under Investigation by firm | Stryker sent an Urgent Medical Device Recall Notification letter dated October 15, 2015, to all affected customers. The notification instructed consignees on how to identify affected product; the related issue with the affected product; potential hazards; risk mitigations; and actions needed to be taken. Customers were instructed to quarantine the affected product. Consignees were asked to complete and return the attached Product Recall Acknowledgement Form within 5 days and either email (strykerortho7984@stericycle.com or Fax (1-866-672-0627 ) the response form back. Customers with questions were instructed to call 201.831.5272. |  | 1 | 1 unit | Distribution to Virginia only. | Terminated | June 08, 2017 | 602 | 1 | Not_Computer
Z-0283-2016 | 72411 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX, PACS, Picture Archive and Communications System, IMPAX CV 7.8.x and IMPAX CV 12.x. Corrects certain demographic information (Patient Name, Patient Medical Record Number and/or Accession Number in previously signed reports that have since changed. |  |  | September 01, 2015 | 2015 | AGFA Healthcare Corp. | Customers have experienced IMPAX CV software, specifically, Demographics Manager (DM), not producing correct pdf reports after the reports, at the time of physician signature, were correct. | Software design | An Urgent Field Safety Notice was sent to accounts via email on 9/1/2015. The letter described the safety alert. Customers were informed of the problem and that Service would be contacting them to arrange a date to deploy the corrective action and install the Service Update at their facility. An acknowledgement was included which was to be faxed back or emailed back to Agfa. |  | 19 | 19 | Distributed in the states of CA, KY, NH, NY, NC, OH, OR, PA, SC, SD, TN, TX, WA and the country of Canada. | Open, Classified |  |  | 19 | Not_Computer
Z-0291-2016 | 71818 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter HomeChoice and HomeChoice Pro APD systems, designed to provide Automated Peritoneal Dialysis (APD) therapy for pediatric and adult renal patients. |  |  | June 29, 2015 | 2015 | Baxter Healthcare Corp. | Loud operating sounds, which was unacceptable to the end users when the device was powered on. The HomeChoice devices continue to perform as intended, as the issue is only associated with the noise coming from the device and not the function of it. May delay treatment if changing to a different system. | Device Design | Baxter sent an Important Product Information communication to affected customers (patients and clinicians) via U.S.P.S., first-class mail on June 29, 2015. The letter states that Baxter Healthcare Corporation is updating the HomeChoice Patient At-Home Guide to include a discussion of the expected operating sounds associated with the HomeChoice and HomeChoice PRO devices. The new labeling will address the following:The HomeChoice cycler utilizes an air (pneumatic) pump; some sounds related to the operation of this pump are expected as part of normal operation of the HomeChoice cycler. There will be sounds like humming, swishing, clicking, and venting (air being released) that are normal for the operation of the cycler. In certain portions of the therapy, the sound level is expected to increase which is also part of normal operation. If there is a significant change in the sound level or a new, previously unheard sound when using the HomeChoice cycler, please contact your doctor and/or nurse or Baxter Technical Services. The firm advises customers to keep the letter with the Patient At-Home Guide.The firm also requests that customers complete the enclosed home patient reply form (enclosure) and return it to Baxter by faxing the reply form to 224-270-5457, scanning and e-mailing it to fca@baxter.com or by mail in the enclosed, preaddressed, stamped envelope. Returning the home patient reply form promptly will prevent you from receiving repeat notices. If the customer has any questions about their Peritoneal Dialysis (PD) therapy, they are advised to please contact their doctor and/or nurse.Additionally the letter states that any adverse reactions or quality problems experienced with the use of this product may be reported using one of the following options:" Calling Baxter Product Surveillance at 800-437-5176, Monday through Friday, between 8:00 am and 5:30 pm Central Time Baxter will advise customers of the additional product information regarding |  | 29355 | 29355 | Nationwide Distribution, District of Columbia, and Puerto Rico. | Open, Classified |  |  | 29355 | Not_Computer
Z-0282-2016 | 71248 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | BrainLAB Image Guided Surgery (IGS) System, Stereotaxic Instrument navigation software, Cranial ENT version 2.1, and Spine and Trauma 3D versions 2.0 and 2.1 An intraoperative image-guided localization system to enable minimally invasive surgery. |  |  | May 08, 2015 | 2015 | Brainlab AG | Instances of data sets not being accurately registered to the patient anatomy were observed. | Nonconforming Material/Component | A FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated April 30, 2015 was sent to all direct accounts (customers). The letters included instructions for customers to: To reduce the risk of an inappropriate automatic registration, Brainlab, in cooperation with NeuroLogica, has developed an improved integration- and verification-workflow to be followed, and additionally wants to re-iterate important use instructions, as described below and specified in detail in the Appendix of this document. When using Brainlab's navigation software to perform automatic registration of CT image data sets in combination with NeuroLogica BodyTom, please ensure that the following conditions are always fulfilled: 1) Ensure suitability of OR Ensure that the operating room where the BodyTom is used in combination with the Brainlab navigation has a continuously even floor. The flatness must not exceed 1mm (0.04 in) for 1m (39.4 in). Please take into account that the flatness may change over time due to wear of the floor coating material and/or negative impacts under the load of equipment used in the room. Ensure that the OR still fulfills the initial specifications before use. 2) Ensure a valid calibration of the BodyTom Automatic registration with the BodyTom may only be used if a valid calibration and verification of the BodyToms mechanical position is given. Both BodyTom and Brainlab calibration must be performed on the same position and movement path that is used for the navigated clinical procedure. The use of Brainlab navigation system is restricted to the one room and movement path where the calibration has been performed. If multiple calibrated rooms are listed on the BodyTom, please ensure that the correct room is selected on the BodyTom workstation when preparing the scan. As part of the Brainlab Corrective Action (see below) you will be contacted by a Brainlab employee to assist in the selection and marking of a suitable movement path and to perform th |  | 12 | 12 systems total (10 in U.S., 2 in Australia) | Distributed in the US and Australia. | Open, Classified |  |  | 12 | Not_Computer
Z-0280-2016 | 72385 | No | LYJ |  |  | Neurology | PMA | Cyberonics Vagus Nerve Stimulator (VNS) Therapy AspireSR Generator, Model 106 |  |  | October 07, 2015 | 2015 | Cyberonics, Inc | Certain Model 106 Pulse Generators demonstrate delays in sensing during use of the 'Verify Heartbeat Detection' feature and exhibit potential for decreased battery longevity. | Manufacturing material removal | Cyberonics sent written recall notifications beginning October 14, 2015, to all affected customers. The letter identified the product the problem and the action to be taken by the customer. The notice also directs customers to contact Cyberonics to schedule additional on-site troubleshooting. The on-site troubleshooting proposed by Cyberonics will utilize a feature included within the Model 250 Version 9.0 Programming Software as an alternate method of verifying heart rate and determining that the appropriate heart beat sensitivity setting is programmed. The Version 9.0 software displays an instantaneous heart rate when the user selects the 'Refresh' feature. This heart rate is measured directly by the generator and displayed on the programmer.For further questions, please call (866) 882-8804. |  | 27 | 27 units | US Distribution to the states of : OR, NY, GA, PA, CA, NC, KS, FL, KY, MI, TX, IL, OH, IN, UT, CO and WV. | Open, Classified |  |  | 27 | Not_Computer
Z-0284-2016 | 71680 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Prism Medical C-300 Consumer Ceiling Lift; PN 330050, 330050R, 330053, 330053R, and 330057; weight specification 300 lbs. /136 kg. |  |  | September 08, 2015 | 2015 | ErgoSafe Products, LLC (DBA) Prism Medical | Drive shaft and over speed arm may simultaneously fail potentially resulting in a patient falling during use | Device Design | The firm, Prism Medical, sent a "Field Safety Notice (FSN 0008)" letter dated 9/8/2015 to all consignees by email via their sales portal. The letter describes the product, problem and actions to be taken. The consignees were instructed to call Prism Medical directly at 1-314-219-8655 and indicate that you are participating in the 300 series lift recall; provide your existing 300 series lift model number and serial number; follow the instructions provided; complete and return the Acknowledgement Form via scanned &amp; signed to bschoenman@prismmedicalltd.com or fax to: 314-692-7858, Attention: 300 series recall.For further information please contact: Prism Medical Quality department at: 1-314-219-8655 or North American Director for Quality &amp; Regulatory Affairs at 1-314-219-8614. |  | 561 | 561 lifts | Worldwide Distribution: US (nationwide) and countries of: Australia, Canada, South Korea, Taiwan, Mexico, Japan, and Great Britain. | Terminated | May 04, 2017 | 604 | 2282 | Not_Computer
Z-0298-2016 | 72471 | No | JXM | Benzodiazepine test system. | Toxicology | Toxicology | 510(k) | 1 Step 12 Panel Cup, Item No. NBCA-12M-W |  |  | October 20, 2015 | 2015 | Ameditech Inc | Ameditech is recalling drug abuse tests because they have shown reduced reactivity. | Nonconforming Material/Component | Ameditech Inc. sent an Urgent Medical Device Recall Letter to al affected customers on October 20, 2015, to inform them that Ameditech is recalling affected lots of Ameditech's Drugs of Abuse Tests because they have shown reduced reactivity for up to three of the nineteen Benzodiazepine (BZO) compounds. The letter informs customers of the actions to be taken. Customers were instructed to complete and return the customer verification form to (858) 677-0243 or scan and email to amt.fieldaction@alere.com. For questions regarding this recall call 858-535-1968. |  | 787 | 787 kits | Worldwide Distribution - US (nationwide) and Internationally to Australia. | Open, Classified |  |  | 149975 | Not_Computer
Z-0294-2016 | 72484 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 132 Gallon GranuFlo Dissolution Unit - P/N G047-80101. |  |  | October 23, 2015 | 2015 | Fresenius Medical Care Renal Therapies Group, LLC | Dissolution units and replacement motors units do not provide full thermal protection for both the Main and Starter coils resulting in the generation of smoke and/or potentially lead to fire. | Device Design | Fresenius issued an Urgent - Medical Device Correction notification, dated October 21, 2015, on 10/27/15. The letter identified the affected device and the reason for the recall. It also contained instructions (on how to perform the required wiring modification) along with a Reply Form. Customers are to return the form after completing the modification. Inquiries regarding the correction should be directed to Fresenius Technical Services at 800-227-2572, referencing the Field Action number FA-2015-01-W. |  | 918 | 918 units | Nationwide Distribution. | Open, Classified |  |  | 2258 | Not_Computer
Z-0293-2016 | 72451 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Neptune Plus, HEMOSTASIS PAD WITH ANTIMICROBIAL BARRIER, For hemostasis following needle/sheath removal, Part number 8870-03 |  |  | October 14, 2015 | 2015 | TZ Medical Inc. | The IFU of the Neptune Plus Pads, does not match that submitted in the original 510k. Need to remove: for haemostasis following needle/sheath removal | Labeling False and Misleading | TZ Medical Inc. began sending the Urgent: Medical Device Correction 8870-03 Neptune Plus letter, dated 10-7-2015, to consignees in October 2015. The firm is notifying all users of Neptune Plus wound dressing to a correction to the Instructions for Use (IFU). Consignees should discard the IFU in your current inventory and replace with new IFU. Consignees with questions should call TZ Medical Inc. at (503) 639-0282. |  | 32170 | 32,170 units | US nationwide distribution. | Open, Classified |  |  | 32170 | Not_Computer
Z-0297-2016 | 72427 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Bulk Loader Module for cobas p 512/612 |  |  | September 22, 2015 | 2015 | Roche Diagnostics Operations, Inc. | Possible for sample tubes in the loader to be opened and contaminate the system as well as adjacent samples during processing. | Device Design | Roche Diagnostics Operations, Inc. sent an Urgent Medical Device Correction letter to affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were asked to retain the Urgent Medical Device Correction letter until a Roche Field Service Representative has installed the BLM modification kit. Consignees were also asked to complete the attached form and fax to 1-877-906-8982, ext 2015. Consignees with questions were instructed to call 1-800-428-2336.For questions regarding this recall call 317-521-3911. |  | 1 | 1 | Distributed to TX only. | Terminated | February 03, 2016 | 134 | 1 | Not_Computer
Z-0292-2016 | 72575 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | ECG Electrodes: |  |  | November 02, 2015 | 2015 | Leonhard Lang Medizintechnik GmbH | There is a risk ECG electrodes of specific lots fail to transmit signals. There is a risk that the affected electrodes will transmit a poor signal or no signal at all before reaching the indicated expiration date. | Nonconforming Material/Component | Leonhard Lang emailed an "Important Safety Notice" letter dated November 2, 2015, to all affected customers on November 2, 2015.. The letter identified the product the problem and the action needed to be taken by the customer. The letter provided information on: 1. Description of the defect; 2. Actions and time frame of the recall; 3. Replacement of the electrodes &amp; contact information. For questions call : +43 512 33425-39. |  | 16410800 | 16,410,800 (10,226,600 US:, 184,200 OUS) | Worldwide Distribution - US Distribution to the state of : FL., and to the countries of : Australia, Austria, Croatia, Czech Republic, Estonia, Finland, France, Germany, Great Britain, Greece, Ireland, Italy, Japan, Malta, Norway,Poland, Portugal, Romania, Serbia, Spain, Switzerland, Thailand, Turkey and United Arab Emirates. | Open, Classified |  |  | 16410800 | Not_Computer
Z-0230-2016 | 72491 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Inspiration Ventilator System, i Series and LS 12.1 LCD. The Inspiration 7i / 5i Ventilator System is to be used by healthcare professionals in hospitals or healthcare facilities and intra-hospital transport. |  |  | October 13, 2015 | 2015 | Event Medical LTD | A component on the systems power board can fail in a manner which can cause the ventilator to shut down without alarming. | Component design/selection | eVent Medical sent an Urgent Field Safety Notice dated October 12, 2015, to all affected customers on October 13, 2015, to inform them that eVent Medical is recalling Inspiration Ventilators i Series and LS 12.1 LCD because a component on the systems power board can fail in a manner which can cause the ventilator to shut down without alarming. The letter provides the customers with the problems identified and the actions to be taken. Customers with questions are instructed to call (949) 900-1917, x232. |  | 6441 | 6441 units total (251 units in US) | Worldwide Distribution - US (Nationwide) and Internationally to:AlgeriaArgentinaAustraliaAustriaBahrainBangladeshBarbadosBelarusBelgiumBoliviaCanadaChileChinaColombiaCzech RepEcuadorEgyptEnglandEstoniaFinlandGermanyGhanaGibraltarGreeceGuatemalaHollandHong KongHungaryIndiaIndonesiaIranIraqIrelandIsraelItalyJapanJordanKenyaKoreaKuwaitLaosLatin AmericaLebanonLibyaLithuaniaMalaysiaMexicoMoroccoNepalNetherlandsNew ZealandNorwayPakistanPalestinePanamaPeruPhilippinesPolandPortugalRussiaSaudi ArabiaSerbiaSingaporeSouth AfricaSpainSri LankaSudanSwitzerlandSyriaTaiwanThailandTurkeyUAEUkraineUSAVenezuelaVietnamYemenZimbabwe | Terminated | June 22, 2017 | 618 | 6441 | Not_Computer
Z-0318-2016 | 72450 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | SilverStat, sterile 2" x 2 " calcium alginate hemostasis pad with sliver nitrate, Part number 8870-07. |  |  | October 14, 2015 | 2015 | TZ Medical Inc. | During a review of the IFU of the SilverStat Pads, it was found that the IFU did not match that submitted in the original 510k | Other | TZ Medical Inc. sent an Urgent: Medical Device Correction 887007 SilverStat letter, dated October 13, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discard the Instructions for Use (IFU) in their current inventory and replace with new (IFU). Consignees with questions should call TZ Medical Inc. at (503) 639-0282. |  | 10180 | 10,180 units | US in the states of AL, AR, CA, DC, FL, GA, IA, MD, NC, NE, NH, PA, RI, SC, TN, TX, UT, and WV. | Open, Classified |  |  | 10180 | Not_Computer
Z-0322-2016 | 72537 | No | NTR | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Thermo Scientific p21 WF1 Ab-3 (DCS-60.2) |  |  | October 27, 2015 | 2015 | Lab Vision Corporation | It was discovered that the labeling of MS-230 antibody incorrectly describes the concentration - label reads 0.2mg/ml but it is actually 0.4 mg/ml. | Labeling design | The firm, ThermoFisher Scientific, sent an "URGENT MEDICAL DEVICE RECALL-Lab Vision Corporation" letter dated 10/29/2015 to its customers on 10/30/2015. The letter describes the product, problem and actions to be taken. The customers were instructed to Upon identification of the affected lot, notify Lab Vision of the quantity previously used and the quantity still in inventory; review any associated test results to ensure an accurate diagnosis was able to be made; If routine laboratory use of this product follows the Lab Vision Data Sheet, the antibody may safely be used as is; hold any affected material in quarantine until the new labels are received-affected product may also be returned to Lab Vision, and complete and return the attached Recall Return Response Acknowledgement &amp; Receipt Form via email to: sarah.rickert@thermofisher.com or Fax to: 269-372-2674, Attn: Quality Assurance &amp; Regulatory Affairs Manager. A new label with the correct concentration can be sent upon request, for relabeling of the product at the customers facility.If you have any questions, please contact Quality Assurance &amp; Regulatory Affairs Manager at 269-544-5628 or sarah.rickert@thermofisher.com. |  | 5 | 5 vials | Worldwide Distribution: US distribution to states of: CA and WA, and countries of: Serbia, Germany and Netherlands. | Open, Classified |  |  | 5 | Not_Computer
Z-0324-2016 | 72529 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | SCFE Driver (SCF-MLD273 &amp; SCF-MLD265) instrument used for the insertion of the Free Gliding SCFE Screw System. |  |  | October 20, 2015 | 2015 | Pega Medical Inc. | SCFE Driver's threaded shaft broke close to the knob when a manual force was applied during the fixation of the screw. | Unknown/Undetermined by firm | The firm sent the Urgent Field Safety Notice-Recalls to US consignees via email. Consignees are advised to return affected devices back to the firm. Customer with questions can contact Enrique Garcia at 450-688-5144 ext. 242 or send email to egarcia@pegamedical.com |  | 4 | 4 units distributed in the US | Distributed in the states of FL and TN. | Open, Classified |  |  | 4 | Not_Computer
Z-0320-2016 | 72344 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator; M3535A, M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6. |  |  | October 14, 2015 | 2015 | Philips Electronics North America Corporation | The following MRx software issue has been identified: MRx model M3535A with software version F.03.06 and earlier, and model M3536A with version T.00.05 and earlier may stop the automated Ready-For-Use (RFU) test in an abnormal state when the device is turned off. | Software design | On 10/14/2015, Philips sent an Urgent Field Safety Notice to each known affected customer. The Urgent Medical Device Correction notification/ Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Philips will provide a software upgrade for M3535A (software version F.03.07 or higher) and M3536A (software version T.00.06 or higher) to correct this issue. A software upgrade will be provided free of charge for all affected products. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrades.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: All MRx users should be informed that if the pattern appears, they can continue to use the device. Once the MRx is turned on, the pattern will disappear and the device will function normally in all monitoring and therapy modes. |  | 81161 | 81,161 | Worldwide Distribution - US (nationwide) including Puerto Rico, and the countries of AFGHANISTAN, ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOVINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CAMBODIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, GHANA, GREECE, GUADELOUPE, HONG KONG, Hungary, ICELAND, India, INDONESIA, ISLAMIC REPUBLIC OF IRAN, Iraq, IRELAND, ISRAEL, ITALY, JAPAN, Jordan, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LAOS, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MACAU, MACEDONIA, MALAYSIA, MALTA, Mauritius, Mexico, MONGOLIA, MOROCCO, MOZAMBIQUE, Myanmar (Burma), NAMIBIA, NEPAL, NETHERLANDS, NETHERLANDS ANTILLES, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, State of Palestine, PANAMA, PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH, AFRICA, SPAIN, SRI, LANKA, SWEDEN, SWITZERLAND, TAIWAN, UNITED REPUBLIC OF TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, TURKMENISTAN, UGANDA, UKRAINE, UNITED, ARAB, EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, and YEMEN | Open, Classified |  |  | 81161 | Not_Computer
Z-0321-2016 | 72541 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Tool Legend 16cm 2.4 mm wire pass for Midas Rex Legend Dissecting Tool. Used by surgeons to drill holes. |  |  | October 21, 2015 | 2015 | Medtronic Sofamor Danek Usa, Inc - Dallas Distribution | The wire pass hole is not present on the tools. | Process control | A recall letter was mailed to consignees on October 21, 2015. |  | 9 | 9 pieces | Distributed in the states of California, New York, and Texas and the country of Belgium. | Open, Classified |  |  | 9 | Not_Computer
Z-0319-2016 | 72576 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging VB36D_HF02. Radiological image processing system. |  |  | October 28, 2015 | 2015 | Siemens Medical Solutions USA, Inc | To provide supplementary information regarding the release of the syngo Imaging Software Version VB36D_HF02. The software provides improvements for all syngo Imaging installations running on the software Version VB36D. | Software change control | Siemens sent an Customer Information Letter dated October 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If the device has been sold and therefore no longer in your possession, please forward this notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions, please call: (610) 219-6300. |  | 3 | 3 | US Distribution to the states of : NC, NE and OH. | Open, Classified |  |  | 3 | Not_Computer
Z-0323-2016 | 72553 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 2.4MM locking screw SLF-TPNG with Stardrive Recess 18 mm. Intended for fixation of fractures, osteotomies, nonunions, replantations, and fusions of small bones and small bone fragments, particularly in osteopenic bone. |  |  | November 02, 2015 | 2015 | Synthes (USA) Products LLC | It was discovered that the above part number and lot of 2.4mm Locking Screws may contain the incorrect package insert (GP2613, VA-LCP Distal Radius System). The correct package insert for the above part is GP2615 (Modular Mini Fragment LCP System). | Process control | Synthes sent an "URGENT FIELD SAFETY NOTIFICATION" letter dated November 2, 2015. Customers were asked to take the following actions:-Review the correct package insert (GP2615, attached).-If found, discard incorrect package insert associated with the locking screw noted in the table above.-Forward this Field Safety Notification to anyone in your facility that needs to be informed.-If the package insert has been forwarded to another facility, contact that facility.-Complete the attached Verification Section (page 3 of this notification). Please include your name, title, address, telephone number and signature in the spaces provided. ---Send a copy of the completed Verification Section by:-Fax: 877-907-7514 or Scan/email: Synthes7766@stericycle.com-If the Verification Form is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual in page three (3) of this notification.-Maintain a copy of this Notification. |  | 229 | 229 | Distributed in the states of IA, OH, CO, FL, TX, MA, NY, OR, SC, VA, WV, TN, MO, PA, CT, GA, WA, CA, AL, , IN, MI, NJ, VT, ID, KS, MN, and MD, and the country of Canada. | Open, Classified |  |  | 229 | Not_Computer
Z-0325-2016 | 69127 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Ingenia, Intera, Achieva and Achieve dStream MR systems on Software versions R5.1.7 and R5.1.8 with the Mobi View software option. Diagnostic imaging system. |  |  | August 21, 2014 | 2014 | Philips Medical Systems, Inc. | When a fused series of a sagittal, coronal or radial multi station scan is generated in MobiView, the resultant image order is reversed. This occurs with software release R5.1.7/R5.1.8. | Software Design Change | Philips Healthcare sent an Customer Information Letter dated July 2014, on August 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips will issue a Field Change Order to install updated software in a Service Pack that makes the slice numbering consistent for fused and unfused images.If you need any further information or support concerning this issue, please contact your local Philips representative: 1-800-722-9377 |  | 307 | 307 units | Worldwide Distribution - USA (nationwide) and to the countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, Denmark Estonia, Ethiopia. France, Germany, Hong Kong, India, Indonesia, Iraq, Ireland, Italy,Japan, Jordan, Korea, Republic of, Libya, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, Oman, Palestinian territory, Peru, Poland, Russian Federation, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 307 | Not_Computer
Z-0326-2016 | 72543 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage System One S Cycler (High Permeability Hemodialysis System) |  |  | October 29, 2015 | 2015 | NxStage Medical, Inc. | Ultrafiltration (UF) Volume software error inaccurate fluid removal | Software Design Change | NxStage issued issued a Field Safety Notice (FSN) on 10/29/15 to consignees who have NX1000-5 or a NX1000-5-A cycler. The FSN included Instructions for steps to be taken so that the cycler and treatments are not affected by the software error.NxStage will be sending a representative to the facility to attach detailed instructions on each cycler as a user reference. Return the attached reply form to acknowledge receipt of this letter.Questions or comments, please contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243). |  | 323 | 323 units | Worldwide Distribution. US Nationwide, Netherlands, and Sweden. | Open, Classified |  |  | 339 | Not_Computer
Z-0328-2016 | 72197 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CONSULT(TM) diagnostics hCG CONTROLS Kit, Part/Reorder Number 5011, IVD |  |  | September 14, 2015 | 2015 | Biochemical Diagnostics Inc | CONSULT diagnostics hCG Controls have discrepant storage temperatures listed on the kit labels when compared with the individual vial labels and the package insert. | Labeling False and Misleading | Biochemical Diagnostics sent a Notice of Correction Letter (dated August 26, 2015) to customers via USPS 1st class mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return the verification form, per the notification. For questions contact Alere Technical Service at 866-216-0094 or by email at verifications.at verifications.ts@alere.com. |  | 4473 | Domestic: 4,473 kits | US Nationwide Distribution | Terminated | January 26, 2016 | 134 | 4473 | Not_Computer
Z-0332-2016 | 72416 | Yes | NYU | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Palindrome Precision Chronic Catheter Kit Symmetrical Tip, Tal VenaTrac" Stylet 14.5 Fr/Ch (4.8mm) X 33 cm |  |  | October 14, 2015 | 2015 | Covidien LLC | Incorrectly packaged with a 90 degree bend at the distal tip | Packaging process control | The firm, Medtronic (Covidien), sent an "URGENT MEDICAL DEVICE RECALL" letter dated 10/16/2015 to their customers. The letter described the product, problem, and actions to be taken. Customers were instructed to quarantine and discontinue use of the product, return affected product and complete and return the RECALLED PRODUCT RETURN FORM via Fax to Medtronic to (800)-895-6140 or email to: feedback.customerservice@Covidien.com/Distributor to (203) 492-7719 or email to: PalCathFCA@Covidien.com.Questions or concerns contact your Medtronic representative or Customer Service at (800) 882-5878. |  | 150 | 150 units | Worldwide distribution: US (nationwide) and countries of: Canada, Belgium, Chile, Panama, Denmark, France, Germany , Netherlands,Spain, Switzerland, and United Kingdom. | Terminated | June 02, 2017 | 597 | 965 | Not_Computer
Z-0334-2016 | 72548 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | BD CD64 APC-R700; |  |  | October 30, 2015 | 2015 | Becton, Dickinson and Company, BD Biosciences | Three lots of CD64 (MD22) are contaminated with CD4 antibody. | Process control | Urgent Product Recall letters were sent out on October 30, 2015 via certified mail to notify customers of the issue. |  | 23 | 23 vials | Distributed in the states of FL, WA, PA, IL, GA, RI, TX, ND, and IA. | Open, Classified |  |  | 23 | Not_Computer
Z-0341-2016 | 72288 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Lyophilized Mouse Monoclonal Antibody CD10 (CD10). Immunohistochemistry. in vitro diagnostic. 0.1ml NCL-CD10-270 and 1mL NCL-CD10-270 are intended for the qualitative identification by light microscopy of CD10 molecules in paraffin sections. |  |  | September 03, 2015 | 2015 | Leica Microsystems, Inc. | These Lot numbers are not stable up to the expiry date on the product labeling. | Nonconforming Material/Component | Leica sent a MEDICAL DEVICE RECALL NOTIFICATION letter dated September 3, 2015 via mail to all of the consignees. The letter identified the affected, product, problem and actions to be taken. Customers are instructed to discontinue use of the reagent lots listed in the recall notification. It is requested that customers appropriately destroy any unused or partially used affected lots of the reagent, and indicate on the attached Medical Device Recall Notification Acknowledgement Form that this action has been undertaken. Customers were asked to confirm receipt of the notice as soon as possible by signing and dating the attached Medical Device Recall Notification Acknowledgement Form and fax it back to1-847-236-3747, to receive youralternative replacement for applicable orders. For questions contact your local Leica representative. |  | 66 | U.S: 1 unit of CD10-270-CE-S &amp; 66 units of CD10-270-CE. Foreign: 83 units of CD10-270-CE-S &amp; 219 units of CD10-270-CE | Worldwide Distribution - US Nationwide, Australia, Austria, Brazil, Canada, China, Czech Republic, Finland, France, Germany, Hungary, Indonesia, Ireland, Israel, Italy, Japan, Netherlands, Norway, Romania, Singapore, Slovenia, South Korea, Spain, Sweden, Switzerland and UK. | Open, Classified |  |  | 66 | Not_Computer
Z-0271-2016 | 72377 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Expansion of previous recall of AcrySof IQ ReSTOR and AcrySof IQ ReSTOR Toric IOL; now to include AcrySof IQ Toric models SN6AT6, SN6AT7, SN6AT8, and SN6AT9. |  |  | September 29, 2015 | 2015 | Alcon Research, Ltd. | Continued increase in reports of post-operative inflammation in patients who received AcrySof IQ Toric IOL. | Process design | The firm, Alcon, sent an "FDA Notification-Update to Recall Z-2323-2015" letter, dated October 1, 2015, to its Consignee in Japan who have received AcrySof IQ Toric IOL models SN6AT6 to SN6AT9 in order to initiate the recovery of product. The letter described the product, problem and actions to be taken. The consignee was instructed to stop using these immediately and put on hold until sales representatives visit their site for withdrawal. Recall does not affect AcrySof IQ Monofocal IOLs or AcrySof IQ Toric IOL models SN6AT3, SN6AT4 and SN6AT5. |  | 43651 | 43,651 units | International Only Distribution -- Japan. | Terminated | June 30, 2017 | 640 | 43651 | Not_Computer
Z-0340-2016 | 72392 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | ADVIA Chemistry Hemoglobin A1c_3, A1c_3M Reagent kits. Used on the ADVIA 1200, 1650, 1800, 2400, and XPT Chemistry Systems. For in vitro diagnostic use in the quantitative determination of Hemoglobin A1c, a diabetes marker, in whole blood on the ADVIA Chemistry systems. |  |  | September 24, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Reagent in these lots may demonstrate an increased occurrence of high %HbA1c bias, a positive bias averaging 0.6% HbA1c units, ranging from -0.1% to 1.1% HbA1c units. The maximum bias was observed at higher %HbA1c concentrations. QC samples may exhibit a similar bias. | Under Investigation by firm | Siemens has issued an Urgent Medical Device Recall in the US and an Urgent Field Safety Notice Outside the US on September 24, 2015, instructing customers to discontinue use, discard reagent and request no charge replacement kits for these products. |  | 3031 | 3031 | Distributed in the states of AZ, NY, CA, IN, LA, MN, OH, MI, NM, and GA, and the countries of Mexico and Canada. | Terminated | January 13, 2017 | 477 | 3031 | Not_Computer
Z-0338-2016 | 72626 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | rainbow Reusable Sensors. RAINBOW DCI-DC3, SpO2/SpCO/SpMet, Part No. 2201; RAINBOW DCI-dc8, SpO2/SpCO/SpMet, Part No. 2407; RAINBOW DCI, SpO2/SpCO/SpMet, Part No. 2696; RAINBOW DCIP, PED, SpO2/SpCO/SpMet, Part No. 2697. rainbow reusable sensors are indicated for spot check or continuous non-invasive monitoring of arterial saturation (SpO2), arterial carboxyhemoglobin saturation (SpCO), arterial methemoglobin (SpMet), and pulse rate. |  |  | October 28, 2015 | 2015 | Masimo Corporation | These sensors were manufactured with incompatible configurations. This could result in sensors that will provide either no readings or inaccurate readings for the SpCO and SpMet parameters. | Finished device change control | Masimo sent an Urgent Product Recall letter on October 28, 2015, to all affected customers to inform them that Masimo has identified specific lots of sensors that were manufactured with incompatiable configurations. The letter provides the customers with the problems identified and the actions to be taken. Customers are instructed to complete the tracking/verification form and email to Masimo. Customers with questions were instructed to contact Masimo Technical Services at http://Service.Masimo.com. For questions regarding this recall call 949-297-7000. |  | 3476 | 3476 units total (3322 units in US) | Worldwide Distribution - US (nationwide) and Internationally to Australia, Costa Rica, France, Germany, Kuwait, Martinique, Netherlands, Norway, South Korea, and Turkey. | Terminated | May 18, 2016 | 203 | 3476 | Not_Computer
Z-0339-2016 | 72399 | No | JGJ | Magnesium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista System (MG) ; Flex reagent cartridge Catalog number K3057, SMN 10445158. For the quantitative measurement of magnesium in human serum, plasma and urine on the Dimension Vista System. |  |  | September 30, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Erroneous low results on a small number of reagent wells. There is the potential for under-recovery on certain wells for quality control (QC) and patient sample results. The under-recovery ranged from -0.3 mg/dL [0.12 mmol/L] to -1.6 mg/dL [0.66 mmol/L]; equivalent to a maximum negative bias of -88%. | Nonconforming Material/Component | An Urgent Medical Device Recall dated September 2015 was provided to all Dimension Vista MG customers who received the affected lot to notify them of the issue, the potential risk to health, and actions to be taken by customers.1. An Urgent Medical Device Recall was provided to all customers who received the affected lot to notify them of the following:Siemens has received customer complaints for Dimension VistaÂż MG lot 15063BA, regarding sporadic depressed serum/plasma and QC results.Discontinue use and discard their remaining inventory of the affected lot.Review the letter with their Medical Director. Siemens will replace any unused inventory of the affected lot. |  | 3601 | 3601 | Distributed US (nationwide) including DC and the states of IL, CA, GA, LA, TX, TN, FL, SC, IA, TN, DE, NJ, WA, MS, KY, NY, WI, OH, OH, MO, UT, KS, VA, AR, WV, MA, PA, AL, CO, IN, MT, NM, ME, AZ, NM, MO, and NC, and the country of Canada. | Open, Classified |  |  | 3601 | Not_Computer
Z-0335-2016 | 72554 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage System One S Cycler - Model no. NX1000-3. |  |  | October 29, 2015 | 2015 | NxStage Medical, Inc. | Ultrafiltration Volume (UFV) may not decrease during treatment-software error. The UF pump may continue to run and remove fluid even after the target UF volume has been removed | Software design | NxStage issued a Field Safety Notice (FSN), dated October 29, 2015, to clinic and home patients with a NX1000-3 or NX1000-4 cycler. For patients performing short daily hemodialysis, the FSN included instructions for steps to be taken so that the cycler and treatments are not affected by the software error. Patients performing nocturnal hemodialysis were instructed to contact their Center for a replacement NX1000-1 Cycler. A software update will be released to correct the error. Customers should complete and return the reply form. Contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243) with questions or comments. |  | 2134 | 2,134 units | Worldwide Distribution. | Open, Classified |  |  | 2274 | Not_Computer
Z-0342-2016 | 72618 | Yes | NKE |  |  | Cardiovascular | PMA | Medtronic InSync III, model 8042, 8042B and 8042U. Implantable pulse generator pacemaker. |  |  | November 09, 2015 | 2015 | Medtronic Inc. | Medtronic has confirmed an issue with InSync III CRT-pacemaker devices related to long-term battery performance. The root cause is unexpected high battery impedance. | Device Design | Medtronic sent an Urgent Medical Device Correction letter dated November 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Medtronic records indicate you are following one or more patients with this device. Please see the enclosed Physician/Patient Detail Report.We regret any difficulties this may cause you and your patients. We will continue to monitor device performance and provide regular updates in our product performance report available at wwwp.medtronic.com/productperformance/. Medtronic Patient Services is available to assist patients at 800-551-5544.Medtronic will offer a supplemental device warranty. Contact your sales representative for terms and conditions. If you have any questions, please contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 96787 | 96,787 | Worldwide Distribution - USA (nationwide) including Puerto Rico and to Algeria, Argentina, Armenia, Australia, Austria, Bangladesh, Belarus, Belgium, Bermuda, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, French Polynesia, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary , India, Indonesia, Iran, Ireland, Israel, Italy, Japan Jordan, Kazakhstan, Kenya, Kuwait, Korea, Latvia, Luxembourg, Lebanon, Lithuania, Macedonia, Malaysia, Malta, Martinique, Mauritius , Mexico, Morocco, Netherlands, New Zealand, Norway,Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Trinidad, Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Vietnam. | Open, Classified |  |  | 96787 | Not_Computer
Z-0343-2016 | 72590 | No | DZC | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Copper NiTi Preformed Archwires; Copper NiTi 40 BAF .016X.022 LO Pk10, Part No. 211-0952. Intended to provide tooth movement. |  |  | November 04, 2015 | 2015 | Ormco/Sybronendo | Ormco Corporation is recalling Copper NiTi Preformed Archwires because the product transformation temperature as listed on the labeling is incorrect. | Employee error | A recall letter dated 11/5/15 was sent to all their customers who purchased the Copper NiTi Preformed Archwires because it has been brought to Ormco's attention that the product transformation temperature as listed on the labeling is incorrect. The wire is labeled as having 40 degrees Celsius transformation temperature, but the actual wire has a transformation temperature of 29 degrees Celsius. The letter provides the customers with the problems identified and the actions to be taken. Customers are instructed to complete and fax back the enclosed acknowledgement form. |  | 161 | 161 units | Distributed US (nationwide) and the countries of Canada, Japan, Mexico, Russia, and Peru. | Terminated | January 19, 2016 | 76 | 161 | Not_Computer
Z-0345-2016 | 72454 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | HYDRO-TEMP(R) Neonatal Temperature Skin Sensor, REF HNICU-32, 50 units per box, NON STERILE, Rx only. |  |  | October 14, 2015 | 2015 | DeRoyal Industries Inc | DeRoyal manufactured four lots of HNICU-32, Hydro-Temp(R) Neonatal Temperature Skin Sensor incorrectly. The pins contained in the probe were soldered incorrectly. When probe is connected to the proper port, this defect results in the inability to read or display temperatures on the infant warming system. | Process control | DeRoyal issued their recall on 10/14/2015 to the end-user level by sending overnight letters to its direct accounts and hospitals. The consignees were instructed to destroy and discard the units. In the case of further distribution, the firm requested that the consignee either notify its customers of the recall or to submit a customer list for DeRoyal to contact them. |  | 800 | 16 cases of 50 units (800 eaches) | Distributed in the states of VA, AL, FL, and CA. | Terminated | July 01, 2017 | 626 | 800 | Not_Computer
Z-0344-2016 | 72466 | No | FOX | Infusion stand. | General Hospital | General Hospital | 510(K) Exempt | Two-stage and three-stage I.V. poles are shipped as an optional accessory to a bed or stretcher. The I.V. poles are stationary, height adjustable stands intended to hold infusion liquids, infusion accessories, and/or other medical devices. They may be used on the following parent devices: Model 2030 (EPIC II), model 3005 S3 beds and stretcher. |  |  | October 13, 2015 | 2015 | Stryker Medical Division of Stryker Corporation | It was identified by a customer complaint that some I.V. poles wereretracting/lowering from an extended position to the low height unexpectedly. Aninvestigation was initiated which identified that the I.V. Latch Housing wasnonconforming. potential hazards:" Suspended I.V. pole retracts unexpectedly" Suspended I.V. pole retracts unexpectedly with pump/bags attached | Process control | On 10/13/2015, URGENT MEDICAL DEVICE NOTIFICATION letters were sent to the affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If you have any urgent questions or concerns, please, contact Stryker Customer Service at (1-800-327-0770). Our normal business hours are Monday-Friday 8 a.m-5 p.m. (ET). |  | 19876 | 19,876 | Worldwide Distribution-US (nationwide) including DC and Guam, and the countries of New Zealand Canada Chile Hong Kong Colombia Argentina Brazil France Germany Italy Netherlands Spain Switzerland United Kingdom India Japan South Korea Mexico South Africa Poland Singapore. | Open, Classified |  |  | 19876 | Not_Computer
Z-0354-2016 | 70370 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Xper vascular systems R7.6 |  |  | February 03, 2015 | 2015 | Philips Medical Systems, Inc. | Monitor Ceiling Suspension system may fall | Process control | Philips sent an Urgent Field Safety Notice letter dated January 8, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. The actuator of all affected systems will be replaced free of charge. All affected customers will be contacted in order to schedule the implementation of this action. For questions contact your local Philips representatice: Technical Support Line: 1-800-722-9377. |  | 22 | 22 units | Worldwide Distribution - US Nationwide and the countries of: Canada,Australia, Austria, Brazil, Chile, China, Denmark, France, Germany, India, Italy, Japan, Mexico, Netherlands, South Africa, Spain, Switzerland and United Kingdom. | Open, Classified |  |  | 176 | Not_Computer
Z-0359-2016 | 72586 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Altrus Thermal Tissue Fusion Energy Sources, Catalog Numbers: 1) 60-9500-120 &amp; 2) 60-9500-INT. Thermal Ligation and Cutting Energy Source. |  |  | November 05, 2015 | 2015 | ConMed Corporation | Miscalibration of the Energy Source monitoring function, which may incorrectly sense the handpiece temperature as higher than the actual handpiece's temperature. This causes some devices to generate a "Check Seal" message with an audible alarm and to stop the heating cycle. | Process control | On November 4, 2015, ConMed Corporation distributed Urgent Device Recall Notification letters dated November 5, 2015 to their customers via overnight courier service. Customers should stop the use of these devices immediately. The recall notice recipient should contact all organizations within their facility and any other facilities that may have received the affected products. It is imperative that all end users of these devices receive this notice and respond immediately. Customers with questions can call Patricia Cotter, ConMed Recall Coordinator 315-624-3237 (8AM - 5PM EST, Monday through Friday), or fax to 315-624-3225 or email altrus@conmed.com. Customers who have any of the devices subject to recall (listed on the recall notice-attachment I) should immediately contact their sales representative. The sales representative will deliver a replacement unit to your facility and will pick up and return the recalled devices to ConMed Corporation. If you are unable to reach your sales representative, please contact Patricia Cotter at the numbers above. If you do not have any devices to return, customers should complete the Business Reply Form (attachment II), indicating you have no devivces and fax it to 315-624-3225, Attn: Patricia Cotter. |  | 9 | 9 devices | Distributed in the states of CA, and TX and in the country of Mexico. | Open, Classified |  |  | 9 | Not_Computer
Z-0355-2016 | 72596 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | U-SDN 2 Part Ultra Modified Seldinger Needles, NS/37-0387. |  |  | October 27, 2015 | 2015 | Procedure Products, Inc. | The firm had received a complaint alleging that a medical professional had difficulty inserting a guide wire into a needle | Nonconforming Material/Component | The firm sent the Urgent Medical Device Recall letter to their consignees via email on 10/27/15.Customers with questions can call Doug Rowley at 360-693-1832 or send email to doug@procedureproducts.com |  | 1100 | 1100 pieces | Distributed in the states of MT and NY. | Terminated | April 01, 2016 | 157 | 1100 | Not_Computer
Z-0357-2016 | 72564 | No | HJO | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | IOL Master 5.5: Software versions 7.5 and 7.7; |  |  | October 27, 2015 | 2015 | Carl Zeiss Meditec AG | IOL Master software versions 7.5 and 7.7 calculation printouts and exported reports can contain the wrong IOL power data. | Software design | Urgent Medical Device Correction letters will be mailed on Friday, November 13, 2015 as the IOL SW update material is now available. Letter informs customers of the issue and provide instructions for identifying the software version and for installing the software update with illustrations. All customers are requested to complete and return the confirmation form to Stericycle as soon as the upgrade is complete or it has been determined that your unit is not running the affected version of the software. |  | 49 | 49 | Nationwide Distribution. | Open, Classified |  |  | 3862 | Not_Computer
Z-0356-2016 | 72539 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | ProTrack Microcatheter. Model/Catalog Numbers CIC38-145, CIC35-145. Sterile, single use device intended to be used as a standard infusion wire or as an exchange guide for exchanging guidewires. Coaxial Injectable catheter guide wire. Cardiovascular. |  |  | October 28, 2015 | 2015 | Baylis Medical Corp * | Microcatheter may have circumferential defects (cracks) along its shaft. | Process control | Recall notification letter was sent to affected consignees on 10/28/2015 via courier. The letter contained a fax-back form to be completed by the consignee and sent to the quality department of the recalling firm. |  | 433 | 433 units | Nationwide Distribution. | Open, Classified |  |  | 433 | Not_Computer
Z-0362-2016 | 72498 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Duracon Flat Tibial Wedge Cemented use only Sterile. Tibial Wedge is a component of Total Stabilizer Knee which consists of femoral component, tibial insert, modular offset adaptor, femoral/tibial stem extenders and tibial wedges. Tibial wedges are intended to be used with the Universal tibial baseplate to augment bone loss on the surface of the tibia. |  |  | September 21, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has received two customer complaints regarding Duracon Tibial Wedge implants puncturing the packaging's Tyvek lid(s). | Under Investigation by firm | On September 18, 2015, Stryker notified their Branches/Agencies of their action via email. The Branches/Agencies were instructed to quarantine affected devices. Urgent Medical Device Recall Notification Letters/Urgent Medical Device Recall Notification Acknowledgement Forms dated September 21, 2015 were sent to Branches/Agencies via UPS on 9/21/2015. Urgent Medical Device Recall Notification Letters/Urgent Medical Device Recall Notification Acknowledgement Forms dated September 21, 2015 were sent to Hospital Risk Managers via UPS on 9/23/2015. The notification instructed consignees on how to identify affected product; the related issue with the affected product; potential hazards; risk mitigations; and actions needed to be taken. Consignees were asked to contact their Stryker Orthopaedics Sales Representative to arrange for return of any recalled product. Customers were asked to complete and return the attached Product Recall Acknowledgement Form within 5 days and either email (strykerortho7984@stericycle.com or Fax (1-866-672-0627) the response form back. A point of contact was provided in case the customer had any questions 201.831.5826. |  | 2577 | 2,577 units | Distributed US (nationwide) and the countries of Australia, Canada, Germany, Hong Kong, Italy, Japan, Malaysia, New Zealand, Netherlands, Poland, Romania, Spain, Sweden and Switzerland. | Terminated | June 08, 2017 | 626 | 2577 | Not_Computer
Z-0363-2016 | 72574 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | DeRoyal sterile custom surgical kits: |  |  | September 10, 2015 | 2015 | DeRoyal Industries Inc | These custom sterile surgical kits contain the component, ECOLAB ORS Fluid Warming and Slush Drape, which was recalled due to the possibility that the sterile packaging may contain small channels in the seal of the pouch. | Other | DeRoyal issued their recall on 09/10/2015 to the end-user level by sending overnight letters to its direct accounts and hospitals. Distributors were asked to put the product in quarantine, place the provided labels on the products and complete. Hospitals were provided with labels to affix to the kits with the direction to discard the recalled component and notice of destruction forms to be completed and returned. |  | 836 | 836 units | Distributed in MO, MA, NC, CT, Kuwait. | Open, Classified |  |  | 836 | Not_Computer
Z-0361-2016 | 72524 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Model 5001102 MPS Delivery Set w/ arrest agent and additive cassettes, heat exchanger &amp;10 ft. delivery tubing; Model 5001102-AS MPS Delivery Set with 6 ft. delivery tubing; Model 7001102 MPS Low Volume Delivery Set for use with cardiopulmonary bypass equipment. |  |  | October 28, 2015 | 2015 | Quest Medical, Inc. | The products have been found to intermittently exhibit a seal failure during use. | Equipment maintenance | The recalling firm issued a press release on 10/28/2015. The press release was posted on FDA website on 10/29/2015. The recalling firm emailed affected consignees on 10/28/15 and sent out recall letters via USPS beginning 10/30/15. |  | 18761 | 18,761 units | Distributed US (nationwide) and the countries of Canada, Japan, Saudi Arabia, Kuwait, South Africa, Switzerland, France, and Italy. | Open, Classified |  |  | 18761 | Not_Computer
Z-0364-2016 | 69090 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Patient Support Material Number: 9890 010 87431 used with: BuckyDiagnost ; Model: 704060/706037 |  |  | August 19, 2014 | 2014 | Philips Medical Systems, Inc. | The hook does not securely hold the footplate in vertical position. Therefore the footplate can fall or swing down and may hurt the operator or patient. | Device Design | The Field Safety Notice FCO70600086 will be mailed on August 19, 2014. Philips Medical Systems will send out certified letters and will have the receipts returned for tracking purposes as well as have a Field Service Engineer visit each site to provide the fix when the Field Change Order is released. |  | 1517 | 1517 systems, not all units have the patient support accessory | Distributed in the states of MN, KY, MA, VA, TX, NC, HI, FL, AZ, CT, and WY and the countries of Argentina, Australia, Austria, Canada, Denmark, Germany, Netherlands, Russia, Saudi Arabia, Singapore, Spain, and Switzerland. | Open, Classified |  |  | 1517 | Not_Computer
Z-0368-2016 | 72591 | No | KHJ | Ear, nose, and throat synthetic polymer material. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | PROLARYN PLUS Injectable Implant 1.0 cc Plus Calcium Hydroxlyapatite (CaHA) Injectable Implant, Model No. 8044M0K5. An injectable implant used to treat vocal fold insufficiency. |  |  | October 26, 2015 | 2015 | Merz North America, Inc. | The expiration date encoded in the UDI barcode on the label on the product carton and the peelable patient identification label is incorrect. | Other | Letters were issued 11/4/2015 to consignees. The letter informs the consignees that they can continue to safely use the affected products by referring to the USE BEFORE date. They letter does not require for any product to be returned. Also, there is no sub-recall required as the product was not distributed to distributors or wholesalers. |  | 9439 | 9,439 units | Distributed US (nationwide) including the Virgin Islands and South Africa. | Terminated | May 08, 2017 | 560 | 17154 | Not_Computer
Z-0366-2016 | 72585 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | Radiesse (+) Lidocaine Injectable Implant 1.5cc, P/N: 8063M0K1. Indicated for subdermal implantation for the correction of moderate to severe facial wrinkles and folds. |  |  | October 26, 2015 | 2015 | Merz North America, Inc. | The expiration date encoded in the UDI barcode on the label on the product carton and the peelable patient identification label is incorrect. | Other | Letters dated 11/4/2015 were sent to customers with tracking and delivery confirmation. The letter informs the customer that they may continue to safely use the affected product by referring to the USE BEFORE date. |  | 1616 | 1,616 units | Nationwide Distribution including the states of AZ, AR, CA, CO, CT, FL, GA, IL, IN, KS, LA, MD, MA, MI, MN, MO, NV, NH, NJ, NY, NC, OH, OK, PA, SC, TN, TX, UT, VA, WA, and WI. | Terminated | May 08, 2017 | 560 | 1616 | Not_Computer
Z-0374-2016 | 70000 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Information Center iX (release A.00, A.01, and A.02) are impacted by this issue: |  |  | December 19, 2014 | 2014 | Philips Medical Systems, Inc. | Sp02 and/or Non Invasive Blood Pressure (NBP) alarms may become disabled without visual notification | Software Design Change | The firm, Philips Healthcare, issued a "URGENT-Field Safety Notice" on 12/19/14 to inform customers of the issue, identify details of the units affected, instructs on actions to be taken by the customer and identify what action Philips plans to take to remedy the issue. Philips will provide a software correction free of charge.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice-To prevent this issue from occurring, customers/users should:1) Disable the Change Bed Label field in Patient Management. This is accomplished by changing the feature from Optional to Hidden2) All patient transfers must be performed using the Transfer key in the Manage Patient application.Remember that the most reliable method of patient monitoring combines close personal surveillance with correct operation of monitoring equipment.If you need any further information or support concerning this issue, please contact your local Philips Healthcare Customer Service representative or the Customer Care Solutions Center at (800) 722-9377 with questions or concerns about this Device Correction. |  | 10013 | 10,013 units | Worldwide Distribution: US (nationwide) and countries: AUSTRALIA, AUSTRIA, BELGIUM, BULGARIA, CANADA, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, HUNGARY, ICELAND, INDIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, LUXEMBOURG, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 10013 | Not_Computer
Z-0389-2016 | 72534 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre NovaSil Silicone Single Lumen Embolectomy Catheter 6F, 80 cm, Model 1601-68. Indicated for the removal of arterial emboli and thrombi. |  |  | October 21, 2015 | 2015 | LeMaitre Vascular, Inc. | Manufacturing issue that could cause a pinhole in the formed end of the packaging, which may lead to compromised sterility. | Packaging process control | Consignees were notified of the recall on October 21, 2015 via letter. Consignees were asked to do the following: Identify all catheters with the affected LOT and model numbers in inventory; quarantine all unused catheters from the affected LOTs and record this action in the attached form; send the form via mail, email or fax to Customer Service who will then issue an RMA-number (Return Merchandise Authorization number) for the return shipment and replacement of the products; pass this notice on to all those who need to be aware within your organization and to any organization to which the potentially affected devices have been transferred. For questions concerning this safety notice, please contact Salva Kozin, Quality Assurance Engineer, at 781-221-2266 ext. 183. |  | 4315 | 4315 total | Worldwide Distribution. US nationwide including PUERTO RICO, and Australia, AUSTRIA, BELGIUM, COLOMBIA, COSTA RICA, Czech Republic, DENMARK, EGYPT, FRANCE, GERMANY, ICELAND, Italy, JAPAN, JORDAN, Latvia, LUXEMBOURG, MOROCCO, NETHERLANDS, Norway, Oman, PERU, PORTUGAL, Poland, Slovenia, Spain, Switzerland, Sweden, THAILAND, United Kingdom, United Arab Emirates, and VIETNAM. | Open, Classified |  |  | 4315 | Not_Computer
Z-0372-2016 | 72655 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew PERI-LOC VOLAR DISTAL RADIUS PLATE STANDARD 3 HOLE SHORT, LEFT - 62 MM, STERILE R, Rx only, REF Numbers: 71863110 Bone fixation plate |  |  | October 20, 2015 | 2015 | Smith &amp; Nephew, Inc. | The devices were packaged and laser marked anatomically incorrect. The Left plates were laser marked and packaged as Right plates and the Right plates were laser marked and packaged as Left plates. | Employee error | Distributors, notified by Fed- Ex on 10/20/2015, were asked to quarantine the product and return it. |  | 9 | 9 units | International onlyGermany, Japan, Russia | Open, Classified |  |  | 9 | Not_Computer
Z-0371-2016 | 72198 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Valeant Pharmaceuticals North America Model No. Spag-2 |  |  | September 14, 2015 | 2015 | Valeant Pharmaceuticals North America LLC | SPAG-2 6000 Lot J3522070 was released to commerce without being tested for particle size per quality control procedure SPAG-2515. | Under Investigation by firm | Valeant Pharmaceuticals International issued recall letter/return response form dated 9/11/2015 via 2nd day air service Customers were requested to check their inventory for the SPAG-2 device and sgregate it for return shipment. |  | 61 | 61 units | AL, AZ, CA, CO, IL, OH, MO, NC, NY, PA., SC, UT and TX | Terminated | June 27, 2017 | 652 | 61 | Not_Computer
Z-0376-2016 | 72604 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | AMS 4674 Custom Pack. For use in a general clinical procedure. |  |  | November 02, 2015 | 2015 | Windstone Medical Packaging, Inc. | AMS 4674 Custom Pack is recalled because this kit containing U-SDN 2 Part Ultra Modified Seldinger Needle which is recalled by supplier. | Process control | Aligned Medical Solutions (AMS) sent the Procedure Products -Urgent Medical Device Recall Letter, dated November 02, 2015 to consignees.This letter listed the items affected by recall as: AMS 4674 Custom Pack and AMS 4675 Angio Pack. Both direct account and sub account are informed that the part # NS/37 -0387 01 U-SDN 2 Part Ultra Modified Seldinger Needle (introducer needle) is recalled because some was unable to accept guide wires up to .038" as intended. The inner diameter was found to be out of specification in some cases. All other components in these two AMS kits are not affected by this recall.Customers with questions about this recall can call Vicki Davis at 321-527-7714 or send email to vdavis@alignedmedicalsolutions.com. |  | 40 | 40 kits | Distributed in AL and VA. | Open, Classified |  |  | 40 | Not_Computer
Z-0375-2016 | 72555 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Trochanteric Fixation Nail- Advanced System-11MM/130 DEG TI CANN TFNA 420MM/Right Sterile (The TFN-ADVANCED Proximal Femoral Nailing System) The TFN-ADVANCED Proximal Femoral Nailing System is intended for treatment of fractures. Intramedullary fixation rod, orthopedic. |  |  | November 02, 2015 | 2015 | Synthes (USA) Products LLC | TFNA nails may have the locking mechanism too close to the top of the nail. This may 1. Prevent the cannulated connecting screw from fully tightening the insertion handle to the nail resulting in a loose or toggling nail in the handle. Or 2. Cause the bottom of the cannulated connecting screw to tighten against the top of the locking mechanism preventing advancement of the locking mechanism. | Process control | The firm, DePuy Synthes, sent an "URGENT NOTICE:MEDICAL DEVICE FIELD SAFETY NOTIFICATION" letter dated 11/2/2015 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: Review the instructions listed in this notification; forward this notice to anyone in your customer's facility that needs to be informed; complete the attached Verification Section and send a copy of the completed Verification Section by: Fax: (866) 629-6186 or Scan/email: Synthes3826@stericycle.com, if the Verification Form is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual, and maintain a copy of this notice.If you have any questions, please call (610)719-5450. |  | 5 | 5 | US Distribution to states of: CA, FL, MI and WI. | Open, Classified |  |  | 5 | Not_Computer
Z-0391-2016 | 72598 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | RBT, Internal Implant, Laser-Lok, 5.0 x 15mm, 5.7 Platform, REF LPBR5015, Rx only, Sterile R, BIOHORIZONS(R), |  |  | October 30, 2015 | 2015 | BioHorizons Implant Systems Inc | An incorrect label reading 10.5mm and not the specified 15mm was placed on the implant inner vial packaging. | Employee error | BioHorizons sent an Urgent Medical Device Field Correcton letter dated Friday, October 30, 2015, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to complete the enclosed Customer Reply Form and fax to 205-484-2195 or scan to PDF and e-mail to lpbr5015@biohorizons.com. The sole international consignee was notified Monday, 02 November 2015. Customers with questions were instructed to contact the Customer Care Department at 888-246-8338. For questionsn regarding this recall call 205-986-1205. |  | 32 | 32 units | Worldwide Distribution - US including CA, CO, NY, OH, WA, AZ MI, MD, GA, NC, and Internationally to Iran. | Terminated | February 04, 2016 | 97 | 32 | Not_Computer
Z-0395-2016 | 72436 | No | HCC | Biofeedback device. | Neurology | Neurology | 510(K) Exempt | The "Pacifier Activated Lullaby" (PAL) System composed of three major components: 1) the PAL Player Module (aka Player Unit), 2) the PAL Pacifier Sensor, and 3) power supply. The system is designed to encourage and reinforce effective non-nutritive sucking of premature infants. |  |  | October 01, 2015 | 2015 | Powers Medical Devices, LLC | PAL unit has undergone significant technological changes since the previous 510(k) and labeling is not consistent with the cleared indications of the 510(k). | No Marketing Application | Powers Medical Devices, LLC intends to notify customers via certified U.S. Mail and supply the updated Instruction Manual via U.S. Mail or other tracked service. Instructions to customers include: 1) destroy the previous Instruction Manual, and replace it with the new one provided, 2) attach the new product label over the existing labeling without covering up the Model # and serial # of the device, 3) certify the destruction and replacement of the Instruction Manual nd replacement of the product label to the company using the form and self-addressed stamped envelopw provided. |  | 128 | Total of 128 units | Nationwide Distribution. | Open, Classified |  |  | 128 | Not_Computer
Z-0396-2016 | 72549 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | DGDE, DirectGuide Drill Extender. Instrument for use with dental implants. A manufactured device that adds a 6mm extension for a 25mm and 28mm tissue height capability suitable for individual patient needs. The thru hole in the DGDE allows existing surgical drills to be inserted and locked in place. |  |  | October 28, 2015 | 2015 | Implant Direct Sybron Manufacturing, LLC | Implant Direct Sybron Manufacturing, LLC is recalling 26 lots of DGDE DirectGuide Drill Extender because of a possible dropping of the drill. | Unknown/Undetermined by firm | On 10/28/15 customer notification letters were sent.Within the letter the firm states that if customers have any of the affected product listed to contact the firm's Customer Care at 1-888-649-6425. A return label is provided with notification.The firm states that if customers are an authorized Implant Direct Sybron Manufacturing distributor they request that they identify their customers that have been shipped the affected product, and contact them within 48 hours of receiving notification in order to obtain the product.The letter states if any customers or patients have experienced any issues as a result of the affected products to report the incident to the FDA through he MEDWATCH reporting system. |  | 478 | 478 | Worldwide Distribution- US (nationwide) including the states of CA, FL, TX, MO, MI, MA, NV, OR, PA, AZ, IL, VA, OH, GA, and SC, and the country of Canada, Poland, Serbia, Italy, Germany, Cairo, Jordan, El Salvador, Kuwait, Costa Rica, Switzerland, Colombia, United Kingdom, Austria, Netherlands, Chile, and Egypt. | Terminated | April 08, 2016 | 163 | 478 | Not_Computer
Z-0397-2016 | 72711 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | IBS - Osteosynthesis screws for bone fixation Model S25 ST014. |  |  | November 18, 2015 | 2015 | In2Bones, SAS | One batch of osteosynthesis screws were manufactured with the incorrect Torx imprint on the screw head. | Nonconforming Material/Component | IN2Bones - USA sent an Field Safety Corrective Action Medical Device Batch Recall letter dated November 17, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Recommended actions :We therefore you to follow the instructions here below:1. Identify all I.B.screws with the batches subject of this recall that might still be in your inventory and quarantine them2. Inform and distribute this Recall Notification to all relevant persons within your organization3. To distributors only : Identify all I.B.screws with the batches subject of this recall that were delivered to your customers, and if relevant instruct them to also follow these instructions (identification and quarantine)4. Fill in and return the fax back form enclosed. With this form, you will certify that you have received this Recall Notification and intend to, comply with the recommendations listed.This fax back form will enable ln2Bones to conduct effectiveness checks.For any question,plea1se contact our Quality and Regulatory Affairs team at: +33 4 72 29 26 26/ +33 7 6188 44 30 or by email:qualite@in2bones.com. |  | 20 | 20 units | US Distribution to the state of :TN | Completed |  |  | 20 | Not_Computer
Z-0398-2016 | 72667 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Thin-Flex Venous Cannula, Model Number:TF292902A |  |  | November 04, 2015 | 2015 | Edwards Lifesciences, LLC | Edwards Lifesciences initiated a field action for the Thin-Flex Venous Cannula (Model TF292902A) because the connector size is incorrectly described in the Edwards product guide as 3/8 inch. The actual and correct connector size is 1/2 inch. | Under Investigation by firm | Edwards Lifesciences sent an Urgent Medical Device Safety Notice dated November 17, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer Instructions1. Review this field safety notice to understand the potential hazard.2. Meet and review with appropriate clinical staff at your hospital.3. Complete and return the acknowledgement form attached to this letter via fax to (800) 422-9329 within 5 business days of receiving this notice.4. Distribute this notice within your organization or to any organization where the potentially affected devices have been transferred.5. Product return is not required. Actual product is acceptable for use and is not being recalled as part of this field action.Your assistance is appreciated and necessary to ensure that this notice is reviewed and understood.We appreciate your attention, and apologize for the inconvenience caused by this matter. If you have questions that have not been answered by this letter, please call Edwards Customer Service Monday through Friday at (800) 424-3278 from 8:00AM - 4:00PM Pacific Time. |  | 8915 | 8,915 units | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : United Arab Emirates , Belgium, Czech Republic, Spain, France, United Kingdom, Italy, Kuwait, Netherlands, Romania, Saudi Arabia, Sweden, Singapore, Thailand and South Africa. | Open, Classified |  |  | 8915 | Not_Computer
Z-0400-2016 | 69121 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT (computed tomography) scanners, Big Bore configuration, Model number 728243. software version 2.0 The products are intended to provide cross sectional images of the human body and visualization of the internal organs of the body. |  |  | April 25, 2006 | 2006 | Philips Medical Systems (Cleveland) Inc | Brilliance Big Bore- An anomaly was identified with software version 2.0 tumor localization application when printing work sheets. this causes information regarding relative locations of marked regions of interest on the patient work sheets to be incorrect and may cause mistreatment. | Software design | Philips Medical Systems sent a product safety notice to customers and corrective software for the issue was released to the field on July 25, 2006. Service Reps will contact each customer to install the new version of software. For questions regarding this recall call 440-483-7600. |  | 54 | 54 units | Worldwide distribution, including US states of CA and NV. | Terminated | December 15, 2015 | 3521 | 54 | Not_Computer
Z-0399-2016 | 70931 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | The eCareManager system. Software intended for use in data collection, storage and clinical information management with independent bedside devices, and ancillary systems that are connected either directly or through networks. For use in a hospital environment. eCareManager allows clinicians to generate a PDF display of an order (new, change, discontinue), which is printed at the bedside and/or hospital pharmacy. This printed PDF should be reviewed by the clinician and verified by the hospital pharmacist and entered into the hospital pharmacy system. |  |  | March 27, 2015 | 2015 | Visicu, Inc. | A software defect may cause incorrect medication order change. If the user decides to edit the order prior to electronic signature AND selects the return to Previous Screen function, the previously selected medication is replaced with whatever medication is in the number one spot on the complete medication list. At the stage where the order is signed, the medication displayed may be incorrect. | Software design | Philips Visicu sent an "Field Safety Notification" to all affected customers still using an affected version of eCareManager. (4 Customers) 03/27/2015 The letter identified the problem and circumstances in which the defect can occur; recommend actions the customer should take to prevent risks for patients; and list the actions planned by Philips to correct the problem."Philips Hospital to Home has taken corrective action to address the software defect in all supported versions of eCareManager software. Philips Hospital to Home is working with affected customers to complete the necessary upgrade.Please contact your local Philips Representative for further questions. |  | 43 | 43 | US distribution only. | Open, Classified |  |  | 43 | Not_Computer
Z-0409-2016 | 69055 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance 64 Computed Tomography X-Ray System, Model number 728231, running version 2.6.0. Intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | March 03, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | Multiple software defects which could impact the performance or reliability of the system. These impacts include, but are not limited to: error messages which prevent operation, incorrect scan region, couch movement failure, inaccurate CTDI reporting, image artifacts, lost images. There is a risk of misdiagnosis. | Software design | On March 3, 2011, the firm sent Field Service Notices to their customers. |  | 1765 | 1,765 units | Worldwide Distribution. US nationwide, including Puerto Rico, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Equatorial Guinea, Finland, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia &amp; Montenegro, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Vietnam &amp; Yemen. | Terminated | December 15, 2015 | 1748 | 1765 | Not_Computer
Z-0404-2016 | 72659 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer, Cuff Tourniquet Dual Hose Single Bladder Maroon Reprocessed by Sterilmed, Sterile EO, Rx only. |  |  | November 13, 2015 | 2015 | STERILMED, INC. | Sterilmed, Inc. is voluntarily recalling all unexpired Reprocessed Tourniquet Cuffs. Sterilmed cannot assure the sterility of all units in the recalled lots. The introduction of non-sterile product into a sterile operative field raises the risk of potential infection. Sterilmed has not received any reports of Adverse Events for this issue. | Process control | Consignees were sent a Sterilmed "Urgent Medical Device Recall: letter dated November 13, 2015. The letter was addressed to Materials Manager, Risk Manager and Operating Room Director. The letter described the product affected, the description of the problem and the Actions Required. Requested consignees to quarantine and return affected product and to complete and return the Business Reply Form. For questions contact Sterilmed's Customer Care Team at 888-541-0078. |  | 174960 | 174,960 | US: Nationwide | Open, Classified |  |  | 174960 | Not_Computer
Z-0403-2016 | 72713 | No | HET | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Olympus ENDOEYE HD II Video Telescopes. Models WA50040A ENDOEYE HD II, 10 mm, 0 degree, WA50042A ENDOEYE HD II 10 mm, 30 degree, W7100603 Repair cable 10 mm 0/30 degree. The ENDOEYE video Telescopes are used with other supporting equipment for endoscopy and endoscopic surgery within the thoracic and abdominal cavities including the female reproductive organs. |  |  | November 17, 2015 | 2015 | Olympus Corporation of the Americas | A damaged temperature sensor at the ENDOEYE tip caused the distal end to become abnormally hot. Excessive heating of the ENDOEYE distal end could result in patient or user injury. | Unknown/Undetermined by firm | Olympus Corporation of Americas sent a letter dated November 25, 2015, to all affected customers to inspect the serial #'s to identify if the affected lots were in inventory and discontinue of the affected lots. Then return all aftected lots for repair and service. Customers were instructed to contact the Customer Care Center at 800-537-5739, Option 2 for instructions on returninng the ENDOEYE for repair and service. The form should be faxed to 484-896-7128. Customers with questions were instructed to call 484-896-5688. |  | 182 | 182 | US Nationwide Distribution including these states: NE, IL, FL, TX, OH, PA, IN, CA, NY, AL, ND, GA, AK, MI, WA, MS, NJ, IA, AZ, and RI. | Open, Classified |  |  | 182 | Not_Computer
Z-0402-2016 | 72364 | No | KGC | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | One Step Button, Low Profile Initial Placement PEG Kit, gastro-enterostomy tube. |  |  | September 11, 2015 | 2015 | Boston Scientific Corp | Boston Scientific Corp. initiated a voluntary recall of specific lots of the EndoVive One Step Button, including the Decompression Tube and Low Profile Replacement Button devices because they may contain decompression tube adapters with a protrusion that could impede its ability to be introduced or removed from the PEG Button. | Under Investigation by firm | Boston Scientific Corp. sent an "Urgent Medical Device Recall Removal - Immediate Action Required" letter dated September 10, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If you identify any product from the affected lot within your inventory, please segregate the affected product immediately and return it to Boston Scientific in accordance with the enclosed recall instructions. You will receive replacements for all recalled product that is returned to Boston Scientific.If you are a distributor, or hospital please note that this recall is to the customer level. Please notify any customer who has received affected product.We are notifying affected worldwide regulatory authorities of this Recall Removal as required.Please read carefully through the enclosed Recall Removal Instructions. Your local Sales Representative can answer any questions that you may have regarding this Recall Removal.We regret any inconvenience that this action may cause, and we appreciate your understanding as we take action to ensure patient safety and customer satisfaction. We are committed to continuing to offer products that meet the highest quality standards that you expect from Boston Scientific.Health care professionals and consumers may report serious adverse events (side effects) or product quality problems with the use of this product to Boston Scientific by calling 1-866-868-4004. |  | 22 | 22 units | US Distribution to the states of : IN, NV, CA, PA, CT, TN, NY, WI, and WA, and Internationally to Ireland. | Open, Classified |  |  | 22 | Not_Computer
Z-0401-2016 | 72617 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Revolution CT Scanners. |  |  | October 23, 2015 | 2015 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential issue with routine head scans on your Revolution CT scanner. A potential hazardous situation can occur during a routine head scan with possible artifacts that may emulate pathology between the brain tissue and bone in the head images. No injuries have been reported to date related to this issue. | Software design | The firm, GE Healthcare, sent an"Urgent Medical Device Correction" letter GEHC Ref#25465 dated October 23, 2015 to its Consignees. The letter described the Safety Issues, Safety Instructions, Product Correction, Affected Product Details and Contact Information. The consignees were instructed to follow the instructions in existing CT User Manual and the safety instructions listed in the letter. GE will be delivering an addendum to the User Manual that reiterates these mitigations. If you have any questions or concerns regarding this notification, please contact GEHealthcare Service at 1-800-437-1171 or your local Service Representative. |  | 125 | 125 (US = 37; OUS = 88) | Worldwide Distribution: US (nationwide) including states of: AL, CA, CT, FL, GA, HI, IL, NY, NC, OH, OR, PA, TX, UT, VA, WA, WI; and countries of: AUSTRALIA, BELGIUM BRAZIL, CANADA, CHINA,DENMARK, FRANCE, GERMANY, HONG KONG, INDONESIA, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, NORWAY, PERU, POLAND, SAUDI ARABIA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Open, Classified |  |  | 125 | Not_Computer
Z-0411-2016 | 72714 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Persona EM Distal Rod, orthopedic manual surgical instrument for use with Persona Knee Prosthesis, item 42-5399-002-00, part of the Persona Tibial Resection System. |  |  | November 24, 2015 | 2015 | Zimmer, Inc. | Complaints that the ratchet teeth of the tube are not latching onto the rod. | Under Investigation by firm | Zimmer Biomet initiated a voluntary recall of the PersonaÂż EM Proximal Tube and PersonaÂż EM Distal Rod due to complaints that the ratchet teeth of the tube are not latching onto the rod via certified mail on 11/24/2015. The firm requests customers quarantine recalled devices, return the customer response forms, and assist in returning the devices to the Zimmer Biomet sales reps. Customers with questions may contact 1-800-348-2759 between 8:00 am and 8:00 pm EST. |  | 156 | 156 | Distributed in the states of AZ, CA, ID, IL, IN, KS, MA, MI, MN, MO, NC, NJ, NY, OH, PA, TN, TX, VA, and WA, and the country of Germany. | Terminated | March 24, 2017 | 486 | 156 | Not_Computer
Z-0405-2016 | 71926 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Brilliance 64 Computed Tomography X-ray system |  |  | April 14, 2015 | 2015 | Philips Electronics North America Corporation | During scans with specific protocol steps, the software unexpectedly sets the Axial ( 2 axis ) scan length to a different length than the user set in the plan. | Radiation Control for Health and Safety Act | Philips Healthcare Planned Action: 1. Philips Healthcare will contact customers and initiate a software update to correct the defect.2. Philips Healthcare engineers will install the software and perform testing to ensure the software update was effective.3. The customer notification letter which includes a statement that Philips Healthcare will without charge, remedy the defect or bring the product into compliance.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Philips Healthcare will implement this CAP by April 1, 2016.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. |  | 7 | 7 | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Canada, China, Colombia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Gibraltar, Great Britain, Greece, Hong Kong, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Latvia, Lebanon, Libya, Lithuania, Malaysia, Martinique, Mayanmar, Mexico, Mozambique, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Republic of Korea, Reunion, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, State of Palestine, Sweden, Switzerland, Taiwan, Thailand, Tanzania, Tunisia, Turkey, United Arab Emirates, Uzbekstan and Vietnam. | Open, Classified |  |  | 322 | Not_Computer
Z-0414-2016 | 72740 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima IGS 320. |  |  | November 13, 2015 | 2015 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential safety issue related to a non-recoverable loss of displayed imaging (loss of monitor video) involving Optima CL323i &amp; Optima IGS 320 systems. | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref#12237 dated November 13, 2015 to customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 45 | 45 | Worldwide Distribution - US Nationwide in the states of: AR, AZ, DC, FL, GA, IL, IN, LA, MI, NC, NE, NJ, NV, NY, OH, SC, TN, TX, WA. and the countries of Algeria, Bangladesh, Bosnia and Herzegovina, Brazil, China, Dominican Republic, Ethiopia, Georgia, Germany, India, Indonesia, Italy, LebanonMartinique , Morocco, Poland, Republic of Moldova, Russia, SpainThailand, Trinidad, Turkey, and Uzbekistan. | Open, Classified |  |  | 153 | Not_Computer
Z-0360-2016 | 72649 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Perseus A500 Anesthesia Workstation, MK06000. |  |  | November 12, 2015 | 2015 | Draeger Medical, Inc. | There is a potential problem with the device main power switch of the Perseus A500 Anesthesia Workstation. The device power switch may spontaneously fail during use, causing one of the following to occur: 1. The workstation has a fresh gas delivery failure and a ventilator failure with respective alarms. 2. The workstation shuts down completely with an audible alarm. | Nonconforming Material/Component | Draeger Medical, Inc. sent Urgent Medical Device Recall Letters, dated November 2015, to consignees on November 10, 2015. The letter identified the affected device and the reason for the recall. The device power switch of the affected workstations must be replace. Customers should refer to the list of affected serial numbers provided. A Drager Service Representative will contact customers to schedule the replacement free of charge, as soon as new switches are available. Until the replacement takes place, the workstation may only be operated under continuous supervision. Increased attention is necessary to notice the potential failure case early. All affected users should be informed. Customers can contact Michael Kelhart at 1-800-543-5047, (press 1 at the prompt, then 2, then 32349), with questions pertaining to the letter. Contact Drager Service Technical Support at 1-800-543-5047 for questions regarding the operation or servicing of the Drager Perseus A500. |  | 34 | 34 | US Distribution -- To the states of WI, PA, OH, ME, IL, and NY. | Terminated | April 05, 2017 | 510 | 34 | Not_Computer
Z-0416-2016 | 72718 | No | NHC | Ventricular catheter. | Neurology | Neurology | 510(k) | VentriClearÂż II Ventricular Drainage Catheter Set. Allows external access and drainage of cerebrospinal fluid (CSF) from the ventricles of the brain. |  |  | November 19, 2015 | 2015 | Cook Inc. | Cook Medical initiated a voluntary recall of specific lot numbers of the VentriClear II Ventricular Drainage Catheter Set due to the device being sterilized twice, which may result in the minocycline and rifampin catheter coating to be less than what is indicated on the labeling. | Under Investigation by firm | Cook Medical issued letter via certified mail on 11/19/2015, for their voluntary recall of the VentriClearÂż II Ventricular Drainage Catheter Set due to being sterilized twice. The Distributor was instructed to forward the recall letter on to its customers, and quarantine any unused recalled devices, and return them to Cook Medical. Customers with questions may contact: Cook Medical Customer Relations 1-800-457-4500 or 1-812-339-2235. Mondaythrough Friday between 7:30a.m. and 5:00p.m. Eastern Daylight Time or email atCustomerRelationsNA@cookmedical.com. |  | 35 | 35 | Distributed in Tennessee, Florida, Michigan, and Pennsylvania and in Saudi Arabia. | Open, Classified |  |  | 35 | Not_Computer
Z-0417-2016 | 72775 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Cournand Modified 2-PC 19G, 5.4cm or (Guidewire Introducer Needle) |  |  | November 23, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | The Guidewire Introducer Needle (0910-192-000) was shipped after the documented expiration date or with minimal time until expiration.There is a potential risk of a soft tissue infection to the patient if device is used after expiration date. | Process control | Stryker sent an Urgent Medical Device Recall Notification letter dated November 24, 2015, to all affected customers via courier with instructions for identifying and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions regarding this recall, please contact Stryker Instruments:Monday-Friday 8am-5pm (EST)269-389-2921 |  | 120 | 120 units | Nationwide Distribution including PA and SC. | Terminated | September 06, 2016 | 288 | 120 | Not_Computer
Z-0415-2016 | 72682 | No | PIF | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Catalog No. INF0020-A  EnteraLite Infinity Enteral Pump Delivery Set with ENFit Connector System and Transitional Stepped Connector (spiked set); Catalog No. INF0500-A - 500ml EnteraLite Infinity Enteral Pump Delivery Set with ENFit Connector System and Transitional Stepped Connector; Catalog No INF1200-A - 1200ml EnteraLite Infinity Enteral Pump Delivery Set with ENFit Connector System and Transitional Stepped Connector; Catalog No GR1200-A  1200ml Enteral Feeding Delivery Set with ENFit Connector System and Transitional Stepped Connector (gravity set). |  |  | November 20, 2015 | 2015 | MOOG Medical Devices Group | Potential for leaking between the purple ENFit connector and the white transitional stepped connector due to cracking or breaking of the purple connector. | Unknown/Undetermined by firm | MOOG Medical sent an URGENT Field Safety Notice, dated November 20, 2015, to its customer. The letter identified the affected device, provided a description of the problem, and discussed the risk to health. The customer was asked to contact all known distributors and inform them that MOOG Medical will cease production of all its enteral administration sets using ENFit connectors and transition back to the previous revision of the product codes that do not include the ENFit connector (INF0020, INF0500, INF1200, and GR1200). MOOG Medical will produce the previous revision until a solution to the problem is found. All necessary users should be made aware of the notice. Christopher Dodge, Manager, Regulatory Affairs, should be contacted with questions or for additional info at 801-264-1001, ext. 112 or cdodge@moog.com. |  | 12014130 | 12,014,130 sets | Distributed to the state of MN. | Open, Classified |  |  | 12014130 | Not_Computer
Z-0419-2016 | 72731 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCI Adult Universal Ventilator Circuit, Duel Heated Limb and Hudson RCI Adult Heated Wire Circuit Kit with Column |  |  | April 17, 2014 | 2014 | Teleflex Medical | Product is being recalled due to complaints that the inspiratory and expiratory tubes are transposed. | Process control | Teleflex sent an "Urgent Medical Device Recall Notification" dated April 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Our records indicate that you have received products that are subject to this voluntary recall. We are now notifying our customers to take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the product and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866- 804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 2600 | 2,600 ea | US Distribution to the states of : AL, AZ, CA, GA, IL, KY, LA, MA, MI, MN, NJ, NY, NC, OH, OK, OR, PA, TN, TX and WA. | Terminated | December 16, 2015 | 608 | 2600 | Not_Computer
Z-0393-2016 | 72535 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | OmniPodÂż, Insulin Management System (US) |  |  | November 02, 2015 | 2015 | Insulet Corporation | Pod's needle mechanism fails to deploy or there is a delay in the deployment of the needle mechanism. | Process change control | Insulet issued on 11/2/15 an URGENT: Field Safety Notification via Email notification and Federal Express. Letter describes the problem that certain lots of the pod's needle failed to deploy or there is a delay in the deployment of the needle. Customers not responding to the email or Federal Express will receive additional mailing and/or follow up phone calls.Call Customer Care at 1-855-407-3729 if you have any questions regarding this Field Safety Notification. |  |  | 26,230.9 boxes | Worldwide Distribution - US Nationwide and countries of Switzerland, Germany, and Israel. | Terminated | March 07, 2017 | 491 |  | Not_Computer
Z-0418-2016 | 72258 | No | FDT | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujifilm Medical Systems Duodenoscope, Model Numbers ED-200XU, ED-200XT, ED-250XT, ED-310XU, ED-420XL, ED-410XT, ED-410XU, ED-450XL, and ED-450XT. Medical endoscope for the duodenum and upper G.I. tract. Open elevator wire channel duodenoscopes. |  |  | August 31, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | A precautionary measure because the Duodenoscopes have been redesigned to incorporate a closed elevator wire channel that facilitates reprocessing. | Device Design | FujiFilm mailed a Medical Device Recall Letter and Tracking/Verification Form dated August 31, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to check their inventory for affected product, if found, affected product should be removed from circulation and returned to Fugifilm. Customers were also asked to complete and return the enclosed tracking/verification form to Fujifilm Medical Systems USA, Inc., Endoscopy Division (FMSU-ESD), 10 High Point Drive, Wayne, NJ 07470 or via fax to 973-633-8818. Customers with questions were instructed to contact their Fujifilm sales representative. For questions regarding this recall call 973-686-2482 |  | 79 | 79 units (US) and 985 units (Internationally) | US including CA, CO, FL, IA, IL, KS, LA, MD, MN, MO, NC, TN and WA and Internationally to Asian Pacific, EU, Latin America, and Japan. | Open, Classified |  |  | 79 | Not_Computer
Z-0421-2016 | 72642 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope AVL Video Laryngoscope (AVL), Part Number AVL 3- 0574-0115; AVL 4-0574-0116, AVL 5- 0574-0117. |  |  | November 30, 2015 | 2015 | Verathon, Inc. | Certain models and serial numbers of GlideScope GVL and AVL laryngoscope blades may separate, due to inadequate bonding of the silicone adhesive affixing the halves of the blade. The potential for blade separation may not be readily visible during routine inspection before or after intubation. | Process control | The Urgent Medical Device Recall letter, dated November 30, 2015, had an error on the dates of manufacture for the AVL devices. Once Verathon discovered the incorrect information, the notification letter was revised and it was sent on December 7, 2015 to any customers that received the previous letter between 11/30/15 and 12/2/15. ****************************************************************************************Verathon sent an Urgent Medical Device Recall letter dated November 30, 2015, to all affected consignees. Verathon will, at no cost to you, test your GlideScope GVL and/or AVL blades that fall within the serial number ranges identified to determine whether the blades exhibit the defect that may result in separation of the blade. For all blades that we determine to exhibit this defect, Verathon will replace the blade, free of charge, with a blade of the same model and size. If the blade passes the testing without any evidence of the defect, we will return it to you in its current condition. Please note that this testing will not return blades to their original, factory condition. Verathon will provide you with a loaner blade for each blade in your possession that is subject to this recall, so you will not experience any interruption in use of your GlideScope system.To comply with this recall notice for the affected models and serial numbers of GlideScope GVL and A VL blades, please take the following actions:" Fill out the attached Recall Response Form and return it to Verathon by fax, (425) 883-2896, or email, CSNotifications@verathon.com . Please return the form even if you do not have any blades subject to the recall." Contact Customer Care to arrange for delivery of a loaner blade and return of your recalled blade. Customer Care will issue you an "RMA" number and arrange for shipment of a loaner blade. Do not send us your current blade(s) until you receive the loaner blade(s).Should you have any questions about th |  | 448 | 448 units (189 units in US and 259 units outside US) | Worldwide Distribution - US (nationwide including Puerto Rico) and Internationally to Bahrain, Belgium, Bermuda, Brunei Darussalam, Bulgaria, Canada, China , Czech Republic, Denmark, France, Georgia, Germany, Greece, India , Indonesia, Iran , Israel, Italy , Japan, Kuwait, Lebanon, Malaysia, Morocco, Palestinian Territories, Panama, Peru, Philippines, Republic of Serbia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Taiwan, Thailand, Turkey, and U.A.E. | Open, Classified |  |  | 448 | Not_Computer
Z-0412-2016 | 72686 | No | FLL | Clinical electronic thermometer. | General Hospital | General Hospital | 510(k) | Digital Temple Thermometer labeled under: Bestmed, Good Neighbor (Amerisource Bergen), Kroger, Medline, Meijer, Premier Value, Safeway, Life Brand, Target, Top Care, Best Choice, Western Family. Device is packaged in plastic blister with cardboard insert, 2 AAA batteries installed, a Quick Start Guide, and an Instruction Manual booklet. |  |  | November 12, 2015 | 2015 | Bestmed, LLC | An incorrect calibration of the thermometer that causes the device to display inaccurate and frequently lower temperatures than the actual body temperature of the user. | Process control | Consignees were notified verbally about 11/12/15. Written notification is planed and will include instructions on the identification of the Affected Device, potential health risk, method for return, and manner of restitution for the cost of the Affected Device. |  | 305548 | 305,548 | US Nationwide Distribution and the country of Canada. | Open, Classified |  |  | 305548 | Not_Computer
Z-0426-2016 | 72643 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MYLA MASTER DVD 3.2 CLI DL380 |  |  | November 03, 2015 | 2015 | bioMerieux, Inc. | MYLAÂż software connected to a Laboratory Information System (LIS) and a VITEKÂż 2 system has the potential to link a test result to an incorrect patient record with the same specimen ID. | Software design | bioMerieux sent an Urgent Product Correction Notice letter dated November10, 2015 via FedEx to affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your local bioMÂżrieux Clinical Customer Service organization at (800) 682-2666. |  | 626 | 626 systems in total | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CT, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and countries of: Foreign accounts: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, France, Germany, Greece, Hungary, India, Italy, Japan, Korea, Netherlands, Poland, Portugal, Russia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Turkey, United Kingdom, Angola, Puerto Rico, Republic of Belarus, Republic of Belarus, Benin, Bosnia, Brunei, Bulgaria, Burkina Faso, Cambodia, Cameroon, Chad, Democratic Republic of Congo, Costa Rica, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Gabon, Guam, Guatemala, Guatemala, Guinea, Honduras, Indonesia, Iraq, Israel, Ivory Coast, Kazakhstan, Kenya, Kenya, Kuwait, Laos, Latvia, Lebanon, Libya, Lithuania, Malaysia, Malaysia, Mali, Portugal, Myanmar, Managua, Niger, Nigeria, Palestine, Philippines, Philippines, Qatar, Kingdom of Saudi Arabia, Senegal, Serbia, Slovakia, Togo, Tunisia, Ukraine, UAE, Vietnam, Vietnam, Ivory Coast and Zimbabwe. | Open, Classified |  |  | 626 | Not_Computer
Z-0448-2016 | 72646 | No | LRZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Tray Set containing multiple instruments and may contain the flexible bone awl, flexible drill and/or flexible screw driver in varying amounts depending on order. Tray set is not sealed. |  |  | November 10, 2015 | 2015 | Aesculap, Inc. | The flexible shaft of the instruments may break during orthopedic spinal surgery. There is a risk of a delay in surgery or that a fragment from the shaft could be left in a patient. | Component design/selection | The firm, Aesculap Implant Systems, sent a "URGENT MEDICAL DEVICE RECALL" letter dated Nov. 23, 2015 to all consignees via certified letter with return receipt. Distributors and sales reps will be notified via e-mail with return receipt. The letter described the product, problem and actions to be taken. The consignees were instructed to review the product recall notification in its entirety and ensure all users your organization and other concerned persons are informed; determine your current inventory; do not destroy any affected product; immediately remove and return the affected instruments-these instruments may be located within Set ST0485; use the U joint screwdriver, ME014R in place of the SJ706R of which is already part of set ST0485; If you require Customer Service assistance please call: Customer Service: 866-229-3002 or Email: SetWorld_Orders.BBMUS_Service@aesculap.com., and complete and return the Product Removal Acknowledgement Form to Aesculap Quality Assurance Department via mail to: Aesculap Inc., 615 Lambert Pointe Drive, Hazelwood, MO 63042; fax to: 610-791-6882 or email to: val.strawn@aesculap.com within two (2) weeks of receipt, even if the total inventory in your possession is zero (0). A replacement for the SJ607R will be sent to you when one becomes available (ME190R). If you have any questions, contact via email: val.strawn@aesculap.com/Tel: 610-984-9414 or email: Gordon.young@aesculap.com/Tel: 610-984-9043. |  |  | unknown | Worldwide Distribution: US (nationwide) and country of: Canada. | Open, Classified |  |  |  | Not_Computer
Z-0434-2016 | 72773 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 600 Coulter Cellular Analysis System, Software Version 1.1.1.0, Part No. B23858 |  |  | December 01, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter is recalling the DxH 800 Slidemaker Stainer Software version 3.0.2.0 and DxH 600 software version 1.1.1.0 because the software for the DxH systems allows the creation of multiple orders with the same Specimen identification (ID) but different Patient identification when manually editing pendingorders at the System Manager. | Software design | Beckman Coulter sent an Urgent Medical Device Recall letter dated December 1, 2015, to all affected customers to inform them that Beckman Coulter is recalling the the DxH 800 Slidemaker Stainer Software version 3.0.2.0 and DxH 600 software version 1.1.1.0 because of the possibility of sample misidentificaton. Customers were informed that during internal testing, Beckman Coulter determined that the software for the DxH systems noted above allows the creation of multiple orders with the same Specimen identification (ID) but different Patient identification when manually editing pending orders at the System Manager. The issue does not occur for edits of released results or for test orders requested through host transmission. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions were instructed to contact Customer Support Center at www.beckmancoulter.com/customersupport/support or via phone (800) 526-7694 in US and Canada. Customers outside of US and Canada are instructed to contact their local Beckman Coulter Representative.For questions regarding this recall call 714-961-5321. |  | 666 | 666 units total (327 units in US) | Worldwide Distribution - US (nationwide) and Internationally to Australia, Austria, Belarus, Belgium, Bosnia and Herzegovina, Brunei Darussalam, Canada, China, Colombia, Croatia, Czech Republic, Ecuador, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Republic of Korea, Kuwait, Lebanon, Lithuania, Macao, Malaysia, Mayotte, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Norway, Oman, Panama Philippines, Poland, Portugal, Puerto Rico , Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Republic, United Kingdom, Uruguay, and Vietnam. | Open, Classified |  |  | 666 | Not_Computer
Z-0447-2016 | 72725 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System. |  |  | November 24, 2015 | 2015 | Elekta, Inc. | First Databank (FDB) is a third party product incorporated into MOSAIQ. FDB routinely modifies drugs in their database to reflect market and formulary changes. In FDB Framework v2.0, those drugs are not being included in drug-drug, drug-allergy and duplicate drug therapy interaction checking. | Software design | Elekta sent an Urgent Important Field Safety Notice 371-01-MSQ-004 to all affected customers on 11/24/2015. The notice informs users to the issue, the work around and the solution. The customers are instructed to complete and return the Important Field Safety Notice Acknowledgement form to their local Elekta Office or Representative, as soon as possible but no longer than 30 days from receipt. |  | 211 | 211 | Nationwide Distribution -- AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WY, &amp; Puerto Rico. | Open, Classified |  |  | 211 | Not_Computer
Z-0437-2016 | 72352 | No | OYC |  |  | Clinical Chemistry | PMA | MiniMed 620G, 3.0mL mg/dL, Model No. MMT-1710 |  |  | September 25, 2015 | 2015 | Medtronic MiniMed Inc. | The pump drive motors may experience a malfunction which would result in a pump error message alarm notifying you that insulin is no longer being delivered. Possible cause is either lack of greasing or deficient greasing in a ball bearing assembly of the motor. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall Notification dated September 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letters stated that the pump drive motors may experience a malfunction which would result in a pump error message alarm notifying the patient that insulin is no longer being delivered. The letters also inform the customers of what actions are to be taken. Customers with questions are instructed to contact Medtronic MiniMed's customer support team. For further questions please call (818) 576-4700. |  | 53 | 53 units | Internationally Distribution only to the countries of : Australia, Austria, Belgium, Chile, Czech Republic, Denmark, Finland, France, Germany, Greece, Iceland, India, Ireland, Italy, Japan, Kenya, Luxembourg, Namibia, Netherlands, Norway, Slovenia, South Africa, Spain, Sweden, Switzerland, United Kingdom. | Terminated | February 05, 2016 | 133 | 1105 | Not_Computer
Z-0446-2016 | 72690 | No | DXO | Catheter tip pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Intravascular (IV) Administration Sets, Model No. 54IV12, 54IV22, 54IV32, MK2450, 54IV42, MK132, MK2156, MK2270, MK2292, MK2300, MK2374, MK2394, MK2508, MK2227 |  |  | November 13, 2015 | 2015 | Edwards Lifesciences, LLC | Edward Lifesciences is recalling Intravascular (IV)Administration Sets due to the requirement of an additional FDA submission to support commercialization (clearance) of the devices. | Under Investigation by firm | An urgent product notification letter dated 11/13/15 was to customers to inform them that during an internal documentation review, the requirement of an additional FDA submission to support commercialization (clearance) of the devices was discovered. Customers are informed that the Edwards Lifesciences Intravascular (IV) Administration sets are temporarily unavailable for sale. Customers are informed that there is no increased risk to the patients' health or safety by using products in their current inventory and there is no impact to the IV sets included in pressure monitoring kits (with TruWave pressure transducer or VAMP system). Customers are informed that if they are in need of a substitute IV set, then to call customer service at (800) 424-3278. Customers instructed to respond to the customer response form by faxing it to (800) 422-9329. |  | 147350 | 147,350 units | US | Open, Classified |  |  | 147350 | Not_Computer
Z-0435-2016 | 72276 | applicator, ent drug | LRD | Ear, nose, and throat drug administration device. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | AbsorbENT, Sterile Compressed Absorbent Sponge, Nasal Pack, 8.9cm(L) x 3.0cm (W) x 1.5cm (H), RH-7413-10, Sterilized by Gamma Radiation, Rx only. |  |  | September 03, 2015 | 2015 | Summit Medical, Inc. | Summit Medical is recalling Doyle P/F Nasal Pack w/Airway, string 8 cm x 3 cm x 1.5 cm, Catalog # RH-7413-10, Lot # 160889 &amp; 161354 because the product does not have the split in the packing. | Process control | Consignees were sent a Summit Medical "Medical Device Recall" letter, dated September 2, 2015. The letter described the problem and the product involved in the recall. It also advised consignees to immediately examine their inventory, quarantine the product, and to complete the attached response form, as well as follow the instructions provided regarding the affected product. For questions, contact Nicole Dove at 651-789-3921 or ndove@summitmedicalusa.com. |  | 9 | 9 boxes (10 /box) | Distributed to CA. | Open, Classified |  |  | 9 | Not_Computer
Z-0444-2016 | 72616 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Unolok Blunt Fill Needle 18G x 1.5" Red, Sterile, Single use, Non-injectable needle, |  |  | November 06, 2015 | 2015 | Myco Medical Supplies Inc | Particulates: Complaints of rubber coring from the closure of injection vials. | Under Investigation by firm | MYCO Medical sent an Advisory Notice dated November 6, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory, if any affected product found MYCO request that the customer return the product to MYCO Medical following the instructions provided on the Materials for Return Form. Customers with questions were instructed to call Customers Service at 919-460-2535/ customerservice&amp;mycomedical.com; or cal the Quality Department at 919-460-2535/ QA@mycomedical.com. |  | 2401419 | 2,401,419 units (100ea/bx, 12bx/cs, 1200ea/cs) | Nationwide Distribution including AZ, CA, CO, FL, GA, IL, IN, KY, LA, MA, MI, NJ, NY, NV, OK, PA, TX and UT. | Terminated | May 03, 2017 | 544 | 2401419 | Not_Computer
Z-0445-2016 | 72583 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Alaris Pump SmartSite Infusion Set, Model No. 2420-0500. |  |  | November 05, 2015 | 2015 | CareFusion 303, Inc. | CareFusion is recalling the Alaris Pump Module SmartSite Infusion Set because the disconnection and leakage may occur during infusion. | Under Investigation by firm | An urgent recall letter dated 11/6/15 will be sent to all customers who purchased the Alaris Pump Module SmartSite Infusion Set because CareFusion has received reports of breakage of the spin collar on the male luer. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return to CareFusion the enclosed recall response card. Customers with questions related to recall are instructed to contact CareFusion Support Center at (888) 562-6018, 7am-4pm PST. Customers with adverse event reports are instructed to contact customer advocacy at (888) 812-3266, customer feedback@carefusion, 24 hrs/day, 7days/wk. |  | 1445180 | 1,445,180 units | Worldwide Distribution-US (nationwide) and the countries of Canada, Australia, and New Zealand. | Open, Classified |  |  | 1445180 | Not_Computer
Z-0443-2016 | 72525 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Gentamicin Antimicrobial Susceptibility Testing Disc (CN10), REF# CT0024B packaged in packs of 5 cartridges - 50 disc/cartridge. |  |  | November 06, 2015 | 2015 | Remel Inc | Antibiotic concentration may be insufficient which may result in incorrect test results. | Nonconforming Material/Component | ThermoFisher Scientific sent an Urgent Medical Device Recall letter dated November 6, 2015, to all affected customers via first class mail. The recall notice identified the problem with the device and asked customers to inspect their stocks for the affected lots. The notice instructed customers to destroy any remaining inventory from the affected lots, contact the recalling firm for replacement, and complete the accompanying Acknowledgement form. Customers with questions were instructed to contact the Technical Services Departmentn at 800-255-6730 (US) or 913-888-0939 (International). For questions regarding this recall call 913-895-4077. |  | 290 | 58/5/50 discs/catridges/packs | Nationwide Distribution including IL, GA, FL, NC, CA, AL, MA, IN, MO, NY, TX, VA, KS, and IA. Product was also distributed to Puerto Rico. | Open, Classified |  |  | 290 | Not_Computer
Z-0449-2016 | 72729 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Label reading in part as: "NexGen System, Complete Knee Solution; Reference Number 5980-37-01; Tibial Component, Precoat, Stemmed Size 3; |  |  | December 02, 2015 | 2015 | Zimmer Manufacturing B.V. | Complaints were received from Thailand reporting that there was no white paper (Tyvek) seal on the inner package of P/N:00-5980-037-01 unit from lot #62460264 | Packaging | An Urgent Medical Device Recall Letter, dated 12/2/15, was sent to the firm's consignees. All distributors will be notified via electronic mail. Hospital risk managers, surgeons, and distributors with product will be notified via courier. |  | 179 | 179 | US Distribution and the country of Thailand. | Open, Classified |  |  | 179 | Not_Computer
Z-0450-2016 | 72741 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | A.L.P.S. 4.0mm Cancellous Locking Screws |  |  | November 20, 2015 | 2015 | Biomet, Inc. | These screws have a single lead (instead of a triple lead) thread. The screw may sit proud in the plate, If the screw does not lock into the plate; the screw may back out. If the screw sits proud, soft tissue irritation may occur. If the screw does not lock into the plate and backs out, a revision surgery may be necessary. | Process change control | On 11/20/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors, hospital risk managers, and implanting surgeons via courier. Distributors were instructed to locate and return the affected product, hospitals to locate and discontinue use, and implanting surgeons to monitor patients. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 182965 | 182, 965 | AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, PRCHILE, COSTA RICA, NETHERLANDS, CANADA | Open, Classified |  |  | 182965 | Not_Computer
Z-0451-2016 | 72683 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobile X-ray system MobileDaRt Evolution/FDR Go Software |  |  | November 10, 2015 | 2015 | Shimadzu Medical Systems | The firm is recalling the Shimadzu Mobile X-ray system because an image may not transfer to image server properly. | Software design | Shimadzu sent an Urgent Voluntary Medical Device Recall Notice dated November 15, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter was sent to customers to inform them that Shimadzu is recalling the Mobile X-ray system MobileDaRt Evolution because of reports that an image was not transferred to image server properly. Customers with questions are instructed to contact their local Shimadzu representative or contact Shimadzu National Technical Support at (800) 228-1429. |  | 297 | 297 units | US Distribution and Internationally to Canada | Open, Classified |  |  | 297 | Not_Computer
Z-0392-2016 | 72702 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Chariot ST-CCV Guiding Sheath, Sterilized using ethylene oxide: |  |  | November 19, 2015 | 2015 | Boston Scientific Corporation | Boston Scientific is voluntarily recalling its Chariot" Guiding Sheath. To date, Boston Scientific has received fourteen complaints for shaft separation, four for distal shaft separation.The most severe outcome of this failure is embolism of device fragments. | Nonconforming Material/Component | Consignees were sent a Boston Scientific "Urgent Medical Device Recall Removal" letter dated November 19, 2015. The letter was addressed to Materials Manager /Field Action Contact. The letter described the problem and the product involved in the recall. Advised consignees to segregate and return affected product to Boston Scientific. Requested consignees to complete and return the "Recall Removal Reply Verification Tracking Form" to MapleGroveFieldActionCenter@bsci.com or fax to 1-866-213-1806. |  | 7985 | 7985 | Worldwide Distribution - US (nationwide) including DC and Puerto Rico, and in the states of AK, AL, AR. AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, WV, and WY and the countries of Austria, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Hong Kong, Hungary, Ireland, Italy, Luxembourg, Macedonia, Norway, Poland, Portugal, Slovakia, Slovenia, Spain, Sweden, Switzerland, and Ukraine. | Terminated | February 23, 2017 | 462 | 7985 | Not_Computer
Z-0455-2016 | 72632 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Alaris EtCO2 Module, Model 8300. A capnograph that continuously monitors end tidal carbon dioxide (EtCO2), fractional inspired carbon dioxide (FiCO2), and respiratory rate (RR). |  |  | November 12, 2015 | 2015 | CareFusion 303, Inc. | CareFusion is recalling the Alaris EtCO2 module model 8300 because it may have been tested during service using a higher than specified concentration of CO2. The effect to the EtCO2 module tested with a higher concentration of CO2 could lead to an incorrect measurement of CO2. | Other | CareFusion sent an Urgent: Medical Device Recall Notification Important letter dated November 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The Alaris EtCO2 Module ( model 8300) is being recalled because it may have been tested during service using a higher than specified concentration of CO2Customers with questions regarding clinical inquiries, product complaints or clinical troubleshooting are instructed to contact Customer Advocacy at (888) 812-3266 or email at customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris are instructed to contact Technical Support at (888) 812-3229 or by email at DL-US-INF-TechSupport@carefusion.com. |  | 157 | 157 units | US Distribution. | Open, Classified |  |  | 157 | Not_Computer
Z-0452-2016 | 72688 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer UNIVERSAL Power System Loaner and Modular Electric/Battery Double Trigger Handpiece. Rx only Made in Switzerland |  |  | November 10, 2015 | 2015 | Zimmer, Inc. | The firm identified a malfunction which made it possible for a handpiece to start by itself when the power source is connected. | Device Design | Zimmer sent an Urgent Medical Device Correction letter dated November 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Review this notification.2. You will receive copies of notification sent directly to hospital risk managers and surgeons in your territory. As necessary, review and facilitate understanding of this notification by those entities.3. Further identify hospital risk managers and/or surgeons in your territory who should also receive notification of this update, and supply the information to the entities you have identified.4. If after reviewing this notification you have further questions or concerns, please call 330-364-0989 between the hours of 8 a.m. and 5 p.m. EST, Monday through Friday.Zimmer instructed their direct customers that they will upgrade the device with the new electronic during their annual preventative maintenance activities. Until the new upgrade is installed, the firm instructed their customers to follow the Instructions for Use. |  | 87 | 87 | US Distribution to the states of : AR, CA, CO, CT, FL, GA, HI, IL, KS, KY, LA, ME, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, PA, RI, SC, TX, UT, VA, WA, WI, WV and WY. | Terminated | February 19, 2016 | 101 | 176 | Not_Computer
Z-0456-2016 | 72700 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer K2 EDTA (K2E) 5.4mg Plus Blood Collection Tubes. |  |  | November 17, 2015 | 2015 | Becton Dickinson &amp; Company | Some tubes were manufactured with the stopper not fully inserted into the tube. This may result in the stopper popping-off in a small number of instances which could lead to blood exposure and in some cases to blood splatter or spraying. | Under Investigation by firm | BD sent an "Urgent Product Recall notifications/Recall Response Forms dated November 17, 2015 via email and/or UPS Next Day Mail on November 17, 2015 to their customers. The notification identifies the affected product and problem and actions to be taken. Customers are requested to take the following actions (1) Immediately review their inventory and quarantine the product subject to the recall and immediately discontinue the shipment of the affected product (2) complete the enclosed Recall Response Firm and fax it bat to BD at 1.866.873.0312 or emai to bd8554@stericycle.com (3) Return all affected products with the completed Recall Response Form following the instruction on the enclosed packing instruction and (5) if distributed, identify your customers and notify them immediately of this product recall. If the customers had any questions, they were instructed to contact BD at 1.888.628.0732 between 8 AM and 5 PM Monday through Friday. |  | 4835000 | 4,835,000 units | Worldwide Distribution - US Nationwide and the countries of: Canada, Indonesia, Malaysia and South Korea. | Open, Classified |  |  | 4835000 | Not_Computer
Z-0454-2016 | 72654 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Aegis Torque Handle CATALOG NO. : 2871-10-800. A spinal surgical instrument used on the AEGIS Anterior Lumbar Plate System. |  |  | November 20, 2015 | 2015 | DePuy Spine, Inc. | Manufacturing error within the Torque Handle may prohibit some handles from being able to fully assemble with the tightener shaft. The design of the inspection gage does not account for the additional shift length | Device Design | DePuy sent an Urgent Voluntary Product Recall Notification dated November 17, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They were asked to do the following: 1. Review inventory located within your facility and quarantine any products listed in this recall notice. Please contact your DePuy Synthes Spine* Sales Consultant to return any inventory located within your facility. 2. Review, complete, sign and return the attached business reply form to in accordance with the directions on the business reply form within 5 business days of receipt of this notification.3. Forward this notice to anyone in your facility that needs to be informed.4. If any product listed below has been forwarded to another facility, contact that facility and provide them with this letter. 5. Keep a copy of this notice.For further questions please call (508) 828-3774 |  | 16 | 16 devices | US Distribution to the states of : NC, MI, CO, MD, FL, ID, TX, CA and NY. | Terminated | July 03, 2017 | 591 | 16 | Not_Computer
Z-0457-2016 | 72619 | No | DRL | Defibrillator tester. | Cardiovascular | Cardiovascular | 510(k) | DA-2006P - Defibrillator / Pacer Analyzer; |  |  | October 28, 2015 | 2015 | BC Group International Inc | Device is not functioning as intended: Two wires running to the ECG distribution circuit board are reversed resulting in inverted waveforms. | Employee error | The recalling firm issued return authorization numbers to the two customers who submitted complaints regarding the malfunctioning device. The customers sent back the units and the recalling firm replaced their devices with new units. On 10/29/2015, the recalling firm contacted their third customer by phone and informed them that the three devices they had in their possession were miswired which would cause the inverted ECG waveform. The recalling firm issued a return authorization number to their customer for the three devices. The recalling firm told their customer they would repair the devices and return them. |  | 5 | 5 units | US Nationwide Distribution in the states of South Carolina, Florida, and Canada. | Open, Classified |  |  | 5 | Not_Computer
Z-0467-2016 | 72861 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM systems, magnetic resonance diagnostic devices (MRDD). |  |  | December 07, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Possible installation error of the quench lines of superconducting magnets, used to vent helium gas in the rare event of a quench. Any restriction, obstruction, or improper installation of the quench line may cause the helium gas to be blown directly into the magnet room or other areas. This could lead to displacement of oxygen, cause cold burns, and/or pose a risk to users, patient or others. | Other | Siemens sent an Customer Advisory letter dated December 7, 2015, certified mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter notifies customers of the potential injuries, which may result of improperly installed quench lines. The letter indicates it is the system's owner who is responsible for installing the quench line per the requirements. If Siemens installed the quench lines, the consignee is directed to contact Siemens, who will inspect the quench lines during next maintenance visit. Siemens will monitor service reports for affected sites. For further questions please call ( 610) 219-6300. |  | 3752 | in US 3752 units | Nationwide including : DC, PR, Virgin Islands and Bermuda | Terminated | June 21, 2017 | 562 | 3752 | Not_Computer
Z-0468-2016 | 72752 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | Factor VIII Chromogenic Assay, REF/Catalog No. B4238-40, Siemens Material Number (SMN) 10445729, IVD. --- CLASSIFICATION NAME: Test, Qualitative and Quantitative Factor Deficiency |  |  | November 19, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Factor VIII Chromogenic Assay Lot 44637 is showing an increased occurrence of irregular curves when used for measurement on Sysmex CA-1500, CA-7000, CS-2000i, CS-2100i, CS-5100 Systems. This could lead, under certain circumstances, to an erroneously increased result with a Factor VIII Chromogenic Assay. | Employee error | Siemens sent an Urgent Medical Device Correction Letter dated November 19, 2015, to all affected customers . The letter identified the product the problem and the action needed to be taken by the customer.Actions to be Taken by Customers: If you are using the affected lot of Factor VIII Chromogenic Assay on a Sysmex CA-1500 or CA-7000 System, please discontinue use. - Due to a different detection principle, the use on BCS and BCS XP is not affected and you can continue using the product on BCS / BCS XP. - Complete and return the Field Correction Effectiveness Check attached to this letter within 30 days. - Review your inventory of this product to determine your laboratorys replacement needs and to provide information to Siemens for reporting to the authorities. --- Corrective Action: The Factor VIII Chromogenic Assay Lot Number 44637 kits were placed on product hold on August 19, 2015, at the Siemens Distribution Centers. For further questions, please call ( 914) 524 - 2495. |  | 675 | Domestic: 52 kits, Foreign: 623 kits | Worldwide Distribution - US Distribution to the state of : Massachusetts and to the countries of : Austria, Czech Republic, France, Germany, Hungary, Italy, Poland, Serbia, Spain and Sweden. | Terminated | August 15, 2017 | 635 | 675 | Not_Computer
Z-0461-2016 | 72738 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Reagent Management System (RMS) |  |  | November 10, 2015 | 2015 | Siemens Healthcare Diagnostics Inc. | Siemens hasidentified RMS refrigeration compressors that are missing the safety cover on the electricaltermination block. The exposed electrical hazard on the electrical termination block is locateddirectly behind the RMS Waste Container. | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Correction (UMDC # DI-16-01.A.US) and Urgent Field Safety Notice (UFSN # DI-16-01.A.OUS) letters dated November 2015 to all affected customers. The letters identified the affected product, problem and actions to be taken. For questions contact your local Siemens technical support representative. |  | 151 | 151 | US Nationwide Distribution: FL, NY, SC, TX, LA, AZ, WI, IA, IL, OH, AL, AK, NH, OK, MD, CA, NM, MS, KY, NJ, MN, GA, CO, MA, SD, TN, MO, PA, MT, SC, NE, KS, VA, AL, NV, WY, OR, PR, MO, OK, GU, HI, AR, ME, VT, CT, NC and the countries of Canada and Mexico. | Open, Classified |  |  | 264 | Not_Computer
Z-0466-2016 | 72811 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Oscillating Flat - TS ST Hub; |  |  | December 08, 2015 | 2015 | Synvasive Technology Inc | Two lots of oscillating saw blades incorrectly identify the blade cutting width as 25mm instead of the correct specification of 19mm. | Labeling False and Misleading | Zimmer sent an Urgent Medical Device Removal letter on December 10, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed: I. Review this notification.2. You will receive copies of notification sent directly to hospital risk managers and surgeons in your territory. As necessary, review and facilitate understanding of this notification by those entities.3. Further identify hospital risk managers and/or surgeons in your territory who should also receive notification of this update, and supply the information to the entities you have identified.4. If after reviewing this notification you have further questions or concerns, please call 330-364-0989 between the hours of 8 a.m. and 5 p.m. EST, Monday through Friday.Letters to Rick Managers are similar and the actions requested by them are:1. Reviewing this notification and forwarding to all people within your organization that may utilize this product.2. Identifying and quarantining the affected product to prevent further distribution or use.3. Completing the attached Response Form and returning it via e-mail toCorporateQua1ity.PostMarket@zimmerbiomet.com. Returning affected product to:Zimmer SurgicalAttn: QA!RA Dept. - RecallPlease include a copy of the Response Form with the shipment.4. Upon receipt of affected product, Zimmer will credit your account for unused Zimmer Oscillating Flat Blade - TS ST Hub, 75mm x 19mm, xl.OO/l.27mm, lots 43054 and 45104 only. Please return a copy of the completed response form along with your returned product to ensure proper credit.Important: Please distribute this notification to all personnel within your organization who need to be aware. If you have further distributed affected product(s), please provide the customer 's information on the Business Response Form to Zimmer.Letters also included a FAQ sheet regarding the labeling error. |  | 14 | 14 | US Distribution to one location each in CT, MA and CA. | Open, Classified |  |  | 14 | Not_Computer
Z-0462-2016 | 72721 | Yes | OAD |  |  | Cardiovascular | PMA | IntellaTip MiFi XP Asymmetric (N4) Curve, |  |  | November 20, 2015 | 2015 | Boston Scientific Corp | During routine monitoring of complaints, BSC noted that the IntellaTip MiFi XP N4 complaint rate for distal tip bends/kinks was above the anticipated rate in BSC's product risk documentation. | Device Design | Boston Scientific sent an Urgent Medical Device Removal letters to all affected customers dated November 20, 2015 via Fed Ex or other equivalent service. The letter identified the affected products, problem and actions to be taken. Customers were instructed to contact their local Sales Representative to answer any questions. |  | 1556 | 1556 - all models | Worldwide Distribution - US Nationwide in the state of: PA, FL, IN, MI, TN, AL, IL, KY, MI, TX, OH, MO, MA, MN, ID, AL, TN, VA, CA and internationally to Austria, Australia, Belgium, Switzerland, Germany, Great Britain, Greece, Hong Kong, Israel, Iran, Italy, Japan, Kuwait, Latvia, Malaysia, Netherlands, Saudi Arabia, Singapore, Thailand, and South Africa. | Open, Classified |  |  | 1556 | Not_Computer
Z-0469-2016 | 72748 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Hook Blade, Angled. Part number 7058A. Sterile disposable blade for use with Endoscopic Tarsal Tunnel Decompression (ETTD) System. |  |  | November 24, 2015 | 2015 | Instratek, Incorporated | The sterile blade will not fit into the handle provided with the ETTD system. | Employee error | Instratek contacted customers by phone call either by direct contact or voice mail message on November 23, 2015. that a recall was being initiated. Formal letters were sent via UPS to the Customers annoucing the recall on November 24, 2015.Immediately examine your inventory and quarantine product subject to recall. In addition, if you may have further distributed this product,please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.You should contact lnstratek's Customer service for a return authorization number and further instructions for return of unused product. |  | 39 | 39 units | US distribution to the states: TX, UT, NY and CO. | Open, Classified |  |  | 39 | Not_Computer
Z-0442-2016 | 72755 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | EtestÂż PIP/TAZO/CON-4 PTC 256 - Foam of 100 tests US. |  |  | November 19, 2015 | 2015 | Biomerieux | Due to the potential for reporting erroneous results for Piperacillin/Tazobactam; specifically false susceptible results. | Process change control | Consignees (Subsidiaries and Distributors) of bioMÂżrieux, impacted by the identified issue, were notified of the FSCA and corrective action (FSCA notification) electronically on November 24th, 2015. Additionally, the bioMÂżrieux United States subsidiary notified US consignees (distributors/end-users) of the FSCA notification via FedEx on December 2nd, 2015. |  | 2422 | 2422 units | Worldwide Distribution- US Nationwide including DC and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GS, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, WV and WY, and the countries of Argentina, Australia, Austria, Brazil, Canada, Chile, Colombia, Czech Republic, Estonia, France, Germany, Hungary, India, Israel, Italy, Japan, Kazakhstan, Kenya, Kingdom of Saudi Arabia, Kuwait, Lebanon, Lithuania, Malaysia, Macedonia, Oman, Netherlands, Pakistan, Paraguay, Philippines, Poland, Portugal, Qatar, Romania, Russia, Serbia, Singapore, Slavonia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, Uruguay, the United Kingdom, Vietnam and Venezuela. | Terminated | June 19, 2017 | 578 | 11591 | Not_Computer
Z-0479-2016 | 72703 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VAPRÂż TRIPOLAR 90" Degree Suction Electrodes (Model/Catalog No: 225028) |  |  | November 23, 2015 | 2015 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Mitek identified that on VAPR Tripolar 90 Degree Suction Electrode the ablation and coagulation buttons are colored incorrectly. The wiring for each button function was confirmed as correct and the non-conformance impacted button color only | Process control | DePuy sent an Urgent Field Safety Notice dated November 23, 2015, mailed FEDEx to all affected Customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were asked to take the following steps: 1. Review inventory located within your facility and quarantine any products listed in this product recall notification. Please follow the instructions on the Business Reply Form to return any affected inventory located in your facility. 2. Review, complete, sign and return the attached business reply form in accordance with the directions on the business reply form within 5 business days of receipt of this notification. 3. Forward this notice to anyone in your facility that needs to be informed. 4. If any potentially affected product has been forwarded to another facility, contact the firm. 5. Keep a copy of this notice. For further questions please call (508) 880-8100. |  | 127 | 127 | Worldwide Distribution - US Distribution to the states of : AL, CA, CT, FL, KY, MN, MS, ND, NJ, NM and TX., and to the countries of: France, Netherlands and Norway. | Terminated | April 12, 2017 | 506 | 127 | Not_Computer
Z-0475-2016 | 72695 | No | LZJ |  |  | Orthopedic | 510(k) | Total Toe Instrument Kit, Labeled in part: "Instrument Set, 15mm Toe, Dorsal Flange Combi" Part Number: 9000-3002 |  |  | June 09, 2014 | 2014 | Arthrosurface, Inc. | An oversight in the implementation of the design change record resulted in outdated components being shipped to customers. The component did not pass a half-cycle sterilization. | Process change control | The distributors were verbally notified, and replacement Rev F Hex Drivers sent to the locations of the kits. The kits had been moved in preparation for upcoming surgeries, but the Drivers were replaced prior to any additional surgery taking place. No affected product remains in the field. For further questions, please call (508) 520-3003. |  | 3 | 3 units | US Distribution to the states of : PA, WV and CA. | Terminated | February 05, 2016 | 606 | 3 | Computer
Z-0476-2016 | 72726 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Dressol-X, Radiopaque D-502, NDC 11004-502-40, Dental Use Only, Active Ingredients: A.S.A. (Acetylsalicylic Acid) + Eugenol |  |  | November 18, 2015 | 2015 | Rainbow Specialty &amp; Health Products | Rainbow Specialty &amp; Health Products initiated a field action for Dressol-X and Dressit-X because the labeling does not currently list the percentage of drug components. | Labeling design | Consignees have been made aware of the recall by FAX, e-mail and via telephone calls on 11/18/2015. Requested that distributors contact their clients, by telephone initially and then by email with written correspondence requesting them to recall the product from their clients. |  | 2775 | 2775 | Nationwide. No foreign, govt/VA/military consignees. | Open, Classified |  |  | 2925 | Not_Computer
Z-0470-2016 | 72786 | No | JSC | Microbiological assay culture medium. | Microbiology | Microbiology | 510(K) Exempt | FB Broth, 10ml |  |  | November 18, 2015 | 2015 | Hardy Diagnostics | The firm is recalling FB Broth Cat. no. K31, Lot no. 15231 because of potential Burkholderia fungorum contamination. | Under Investigation by firm | Hardy Diagnostic notified Customers via phone on November 18, 2015, followed up with written notifications being sent via first class, email, or facsimile.Customers were instructed to destroy the recalled product according to local regulations for nonhazardous waste if uninoculated or as medical waste if inoculated. and to discard any remaining tubes of the recalled lot. Any remaining product on hand the firm stated would be replaced. The firm stated to customers that they would follow-up the phone conversation with written communication and requested a full name, and email or fax number.We sincerely apologize for the inconvenience. Thank you in advance for completing the attached form. Please feel free to contact our Technical Services Department if you have not received your replacements or if you have additional questions about this notification. Technical Services can be reached at (800) 266-2222, option 2 or via email at TechService@HardyDiagnostics.com. |  | 500 | 500 25pk/20) | US. Distribution to the states of : CA, WI, MD and UT.. | Terminated | March 28, 2016 | 131 | 500 | Not_Computer
Z-0436-2016 | 72754 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Optional PS500 Power Supply Unit for the Evita V500 Ventilator and Babylog VN500 Ventilator. |  |  | December 01, 2015 | 2015 | Draeger Medical, Inc. | The firm became aware of cases in which the battery run times of the optional PS 500 power supply unit with the Infinity Workstation Critical Care (Evita Infinity V 500) were unexpectedly short due to the design of the charging algorithm in the current software. Devices used for patient transport will be a priority. | Software design | Draeger Medical, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated November 2015 to all affected consignees. The letter identified the product, problem, and actions to be taken by the customers. The Battery Charging Software will be updated and, if necessary, the batteries replaced in all affected devices Free of Charge. |  | 2422 | 2422 | US Nationwide-including DC and PR, and the states of AK; AR; AZ; CA; CO; CT; DE; FL; GA; HI; IA; ID; IL; IN; KS; LA; MA; MD; ME; MI; MN; MO; MS; MT; NC; ND; NE; NH; NJ; NM; NV; NY; OH; OK; OR; PA; SC; SD; TN; TX; UT; VA; VT; WA; WI; and WV. | Open, Classified |  |  | 2422 | Not_Computer
Z-0471-2016 | 72829 | No | HWQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Black Mamba Suture Passer |  |  | December 02, 2015 | 2015 | Biomet, Inc. | There may be a separationand an insufficient weld between the spring and washer. The weld between the spring and washer may fracture,causing the Mamba Suture Passer instruments jaw to not actuate properly.6 complaints have been received in which 3 have been reported with a delay up to 10 minutes. | Nonconforming Material/Component | On 12/2/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications were sent to the affected distributors and hospital risk managers via courier with instructions to locate and return the affected products. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 50 | 50 | FL, NJ, WI, UT, TX, KY, KS, OH, GA, LA, AL, NY, NC, AR, MO, VA, CA, MN, IN, HI | Open, Classified |  |  | 63 | Not_Computer
Z-0480-2016 | 72747 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Synaptive Medical BrightMatter Navigator System (Guide System). Part number SYN-0026. The Guide system is comprised of a cart labelled SYN-0024 and a cart labelled SYN-0025, Operator Cart. A planning and intraoperative guidance system to enable open and percutaneous computer assisted surgery. The system is indicated for medical conditions requiring neurosurgical cranial procedures. It can be used for intra-operative guidance where a reference to a rigid anatomical structure can be identified. |  |  | December 04, 2015 | 2015 | Synaptive Medical, Inc. | Out of tolerance for radio frequency emissions. At the 150-1000MHz frequency, the testing indicated the BrightMatter Navigation system was up to 20dB uV/meter higher than the applicable IEC 60601-1-2:2007 (Ed3.0) standard specification. | Component design/selection | Consignees were sent via courier on 12/4/2015 a Synaptive "Urgent Medical Device Recall" letter dated 30 November 2015. The letter described the problem and the device involved in the recall. Informed consignees that a technical support team member from Synaptive Medical will conduct the field correction. They also requested consignees to complete and return the Recall Response Form. For questions contact Thas Yuwaraj, PhD, .PEng Monday through Friday, 9 AM to 5 PM, Eastern Time at 1-844-462-7246. |  | 8 | 8 | US: FL, WI, IL, LA, WA. | Terminated | March 04, 2016 | 91 | 8 | Not_Computer
Z-0473-2016 | 71883 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Vancomycin Screen Agar |  |  | June 29, 2015 | 2015 | Becton Dickinson &amp; Co. | The affected expired lots of product may have exhibited breakthrough growth of vancomycin susceptible Enterococcus facaelis ATCC 29212 near the end of shelf life. The in-date product may also exhibit breakthrough growth of vancomycin susceptible Enterococcus facaelis ATCC 29212 near the end of shelf life. | Under Investigation by firm | Becton Dickinson (BD) sent US customers a Urgent Product Recall letter dated June 2015, via UPS.Customers were requested to return response forms indicating acknowledgement of the recall and quantity of remaining product that was at their facility. The acknowledgement form included with the follow up letter that is faxed back to BD will be reconciled with the customer list. BD will contact 10% of non-respondents to verify receipt of the customer letter.Any product returned to BD will be quarantined and discarded.If further assistance is needed, please contact BD Technical Services at 1-800-638-8663. |  | 60840 | 60,840 | Worldwide Distribution - US (nationwide) Distribution to the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, WI and WV., and to the countries of : Canada, Colombia, India, Mexico and Europe. | Terminated | August 12, 2016 | 410 | 61610 | Not_Computer
Z-0552-2016 | 72848 | Yes | BYE | Positive end expiratory pressure breathing attachment. | Anesthesiology | Anesthesiology | 510(k) | Airways Development LLC WaterPAP Positive Airway Pressure Device Reference #WP-7700. |  |  | September 24, 2015 | 2015 | Airways Development LLC | Airways Development LLC has received a complaint about a canister leaking from WaterPAP Lot 8605A. During a review with their supplier it could not be determined if this was an isolated incident or if the whole lot is impacted. | Nonconforming Material/Component | Airways Development Inc. sent Product Alert notices dated 9/24/2015 to their customers via email. The alert informed customers about the related product issue and how to identify affected product. Airways Development is replacing any affected product left in their customers stock. Customers were also instructed to notify their customers if the product was further distributed. A point of contact was provided if the customers had any related questions. |  | 1464 | 1,464 units | Nationwide | Open, Classified |  |  | 1464 | Not_Computer
Z-0481-2016 | 72742 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis systems with large display from Siemens Medical Solutions USA, Inc. Interventional Fluoroscopic X-ray for angiographic procedures. |  |  | October 30, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Due to improper soldering in a limited number of large display bypass module components of a specific production lot, a loss of video signal could potentially occur. This has not been observed in the field and only sporadic cases have been observed in factory screening. | Software design | Siemens sent an Important Customer Safety Notice dated October 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The issue will be remedied with a hardware update AX025/15/S. The affected bypass module hardware will be exchanged with unaffected production lots to eliminate the problem. For further questions call: (610) 219-6300. |  | 17 | 17 | US Distribution to the states of : IL, WI, MA, FL, AL, TN, ND, OH,TX, CA, NJ and MI. | Terminated | March 28, 2017 | 515 | 17 | Not_Computer
Z-0543-2016 | 72681 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Cardinal Health Tiny Toes; 11470-010T Infant heel warmer |  |  | November 20, 2015 | 2015 | Philips Electronics North America Corporation | Heel warmers may cause first and second degree burns. | Labeling design | Cardial Health was notified of the recall via letter on 11/20/2015. Cardinal Health notified all of their customers as of 12/1/2015. Philips instructed Cardinal Health do take the following steps: 1.) Check for and destroy any inventory of Cardinal Health Tiny Toes. 2.) Complete the Distributor Reply Form included in this letter to document this action. Fax to Philips Healthcare at +1 404-855-4900. Please do not return any product to Philips. 3.) Determine which of your customers may have affected product. 4.) Ask your customers to follow the instructions in the section entitled: ACTIONS FOR CARDINAL HEALTH CUSTOMERS/USERS TO TAKE. 5.) As you receive signed Customer Reply Forms from your customers, please forward a running monthly summary to us. Please email the summary to Jennifer Maglio at jennifer.maglio@philips.com. To maintain complete documentation, please retain your customers completed reply forms. Philips may request access to particular reply forms to respond to an audit inquiry.Actions for Cardinal Health Customers to Take: 1.) Discontinue use of Cardinal Health Tiny Toes. 2.) Check for any inventory of Cardinal Health Tiny Toes. 3.) Destroy all products on hand (eg. by activating or puncturing each unit). 4.) Complete and sign the Customer Reply Form provided in this letter or Cardinal Healths version thereof, to notify Cardinal Health of how many Cardinal Health Tiny Toes were destroyed. Submit your signed reply form only after you have completed all steps |  | 7327100 | 7,327,100 individual units | Moberly, MO | Terminated | August 07, 2017 | 626 | 7327100 | Not_Computer
Z-0545-2016 | 72684 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-ray system |  |  | October 29, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Perfusion scan feature may not be available on machines running software versions 4.1.2, 4.1.3, and 4.1.4. Customers previously using the basic axial perfusion software (versions preceding 4.1.2) without buying the helical perfusion license key could not access the basic axial perfusion feature after upgrading to software version 4.1.2. | Software design | On 10/30/2015 the firm sent Customer Information letters to their Consignees. |  | 51 | 51 Units | Worldwide distribution. US nationwide, Australia, Austria, Bangladesh, Belgium, Brazil, Cambodia, Canada, China, Egypt, France, Germany, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Korea, Lebanon, Libya, Lithuania, Malaysia, Martinique, Myanmar, Mexico, Mozambique, Netherlands, New Zealand, Norway, Palestine, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uzbekistan &amp; Vietnam. | Open, Classified |  |  | 436 | Not_Computer
Z-0482-2016 | 72687 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MicroCutter Xchange 30 Blue Cartridge, FG-025320; |  |  | November 06, 2015 | 2015 | Cardica, Inc. | Cardica, Inc. has received reports that use of the Blue Cartridges with the MicroCutterXCHANGE 30 Stapler was associated with incomplete firing of staples, potentially resulting in anincomplete transection or anastomosis of tissue during the surgical procedure. | Process design | An Urgent: Medical Device Removal letter, dated November 17, 2015, was sent via Fed Ex. The letter identified the affected device, described the reason for the recall, and provided the instructions to be followed. Customers were asked to check their inventory, and remove and quarantine any affected product. The notice should be forwarded to relevant personnel, and the affected product should be returned along with the completed Reply Form. A credit will be issued for any returned product. A sales representative can be contacted with questions. |  | 7384 | 568 boxes or 6816 cartridges | Worldwide Distribution -- USA, distributed to the states of Alabama, California, Colorado, Florida, Georgia, Iowa, Massachusetts, Maryland, Minnesota, New York, Ohio, Pennsylvania, Rhode Island, Texas, Washington, and Wisconsin ; and, to the countries of Austria, Belgium, France, Germany, Ireland, Italy, Luxembourg, Netherlands and United Kingdom, Switzerland, and Qatar. | Terminated | February 12, 2016 | 98 | 7384 | Not_Computer
Z-0512-2016 | 72651 | No | JXM | Benzodiazepine test system. | Toxicology | Toxicology | 510(k) | ImmuTest Cup 7-panel w/ Adulteration, Item No. IMCA-7P |  |  | November 16, 2015 | 2015 | Ameditech Inc | Ameditechs Drugs of Abuse Tests listed above have shown reduced reactivity for up to two (2) of the nineteen (19) Benzodiazepine (BZO) compounds for which specificity information is provided in the product insert. The particular BZO compounds for which there may be reduced reactivity are Clonazepam and Chlordiazepoxide. | Nonconforming Material/Component | A Medical Device Labeling Correction letter was sent on 11/16/15 to distributors and end users to inform them that Ameditechs Drugs of Abuse Tests have shown reduced reactivity for up to two (2) of the nineteen (19) BZO compounds for which specificity information is provided in the product insert. The letter informs the customers that the particular BZO compounds for which there may be reduced reactivity are Clonazepam and Chlordiazepoxide. Customers with questions are instructed to contact Ameditech at (844) 616-4453 or email at amt.fieldaction@alere.com. |  | 60 | 60 kits | Nationwide Distribution. | Open, Classified |  |  | 161537 | Not_Computer
Z-0551-2016 | 72828 | No | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | Terumo Surflo Winged Infusion Set with Filter and Needle Protection (Surshield), Product Code: SV-S25FL35. |  |  | November 23, 2015 | 2015 | Terumo Medical Corp | Potential improper bonding between the tube and luer for certain lots of Surflo Winged Infusion Set with Filter and Needle Protection (Surshield). | Process control | Terumo N.V. Europe notified their US Importer (Termumo Medical Corporation, Somerset, NJ) of the voluntary recall on 11/23/2015 via phone and email. Terumo Medical Corporation in turn notified their sole consignee on 11/24/2015 by email that included the manufacturers Recall Notification. The notification informed the customer of the reason for the recall; how to identify affected product; to immediately identify and quarantine any available stock; and complete and return the attached Reply Form. |  | 80500 | 80,500 units | California | Terminated | August 01, 2017 | 617 | 80500 | Not_Computer
Z-0550-2016 | 72838 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging XS is a Picture Archiving and Communication System (PACS) |  |  | November 17, 2015 | 2015 | Siemens Medical Solutions USA, Inc | For the Syngo Workflow SLR System with Software Ver: VB10C: Printouts may be printed in incorrect anatomical size when using syngo Imaging XS filming application in conjunction with a printer not released for anatomical print usage. This may happen due to an invalid combination of printer and the syngo Imaging XS printer configuration file. | Software in the Use Environment | 11/17/2015, CSAN sent to customers with a list of suitable printers for the anatomical printing. |  | 42 | 42 | AL; CA; CO; CT; FL; GA; IA; IL; IN; KS; MD; MI; MO; NC; NE; NJ; NV; NY; OH; PA; TN; TX; WI; | Open, Classified |  |  | 42 | Not_Computer
Z-0554-2016 | 72661 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Hemo with software versions 10.0 LA. |  |  | October 26, 2015 | 2015 | Merge Healthcare, Inc. | The firm, Merge Healthcare, sent out Merge HEMO V10.0 &amp; Merge CARDIO V10.1 software to 3 customers before the product was completely validated in house. | Process control | Consignees were either contacted via phone 9/2015 or in person 7/31/2015. The customer communications involved the need to upgrade the consignees software. Merge Healthcare did not script the conversations because they opted to hold one-on-one management level discussions to carefully explain the importance of moving to the Generally Available (GA) release. |  | 2 | 2 | Distributed in the states of IL, NC, and VT. | Open, Classified |  |  | 3 | Not_Computer
Z-0553-2016 | 72892 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, 3.0T 6 Channel Phased Array Flex Coil, M0050SS. |  |  | December 07, 2015 | 2015 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential safety issue with the 3.0T 6 Channel Flex Coil used with the 3T MR750W Surgical Suite Scanners.Coil overheating can occur when the device is used in Mode 2 setup. This could lead to a serious patient thermal injury. There have been no injuries reported as a result of this issue. | Nonconforming Material/Component | Consignees were sent on 12/7/2015 a GE Healthcare "Urgent Medical Device Correction" GEHC Ref#60888 letter dated December 7, 2015. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 90 | 90 (54 US; 36 OUS) | Worldwide Distribution- US Distribution including the states of AZ, CA, CO, IL IN, MI, MN, NY, OH, OK, SD, TX, WI, and the countries of CANADA, CHINA, CZECH REPUBLIC, ITALY, JAPAN, NETHERLANDS, NEW ZEALAND, PORTUGAL, ROMANIA, SOUTH AFRICA, SWITZERLAND, TURKEY. | Open, Classified |  |  | 90 | Not_Computer
Z-0556-2016 | 72760 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Granada Biphasic Broth |  |  | November 23, 2015 | 2015 | Biomerieux Inc | FDA has determined that the product requires 510(k) clearance to continue marketing/distribution in the United States. | No Marketing Application | The firm, Biomerieux, sent an "Urgent Product Removal Notice" dated November 23, 2015 to its consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to confirm that the letter has been distributed and reviewed by all appropriate personnel, discontinue and discard any remaining inventory, store letter with your bioMerieux documentation and complete the attached Acknowledgement Form via fax to: (919) 620-3306 attention of: bioMerieux NA ComOps QA/RA.If you have any questions or concerns, please contact your local bioMÂżrieux Clinical Customer Service organization at (800) 682-2666. |  | 24 | 24 cartons | US Distribution to states of: CA, GA, IN, MA, MN, MO, NY, OH and PA. | Terminated | February 24, 2016 | 93 | 24 | Not_Computer
Z-0557-2016 | 72801 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ Oncology Information System |  |  | December 09, 2015 | 2015 | Elekta, Inc. | Incorrect drug dosage due to "Age Limit" and patient weight data item issue. | Software design | The firm, Elekta, sent an "URGENT IMPORTANT FIELD SAFETY NOTICE 371-01-MSQ-005" to all customers on 12/9/2015. The notice informs users of the specific product and version numbers affected by both issues, also includes a workaround for both issues. A copy of the Important Field Safety Notice is to be kept with the most current labeling and all personnel working with the product should be made aware of the content of the letter. The customers are instructed to return the acknowledgement form to their local Elekta as soon as possible-Fax: 408-830-8003.If you have any queries about this Notice, please contact your local Elekta office, call 408-830-8000 or toll free: 855-693-5358 or www.elekta.com. |  | 399 | 399 | Worldwide Distribution: US (nationwide) including states of: AK, AZ, AR, CA, CO, CT, DE DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, ,MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI, and WY; and countries of: Antiqua &amp; Barbuda, Australia, Austria, Bahama, Bahrain, Canada, Czech Republic, Germany, Hungary, Saudi Arabia, Sweden, Turkey, &amp; United Kingdom. | Open, Classified |  |  | 399 | Not_Computer
Z-0558-2016 | 72779 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Pos Combo Panel Type 21, Part No. B1017-201 |  |  | December 01, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter has confirmed falsely negative results for Voges-Proskauer (VP) (pale-pink/brown/colorless) with Quality Control (QC) American Type Culture Collection (ATCC) organism Staphylococcus aureus ATCC 29213 in a portion of affected lots of MicroScan Pos Combo and MicroScan Pos Breakpoint Combo panels. The expected result is positive. | Device Design | The firm, Beckman Coulter, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 1, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to discontinue use and discard any remaining inventory of lots listed; return the attached form for product replacement; retain your inventory of the other lot numbers of the panel types, as they are not impacted by this issue; discuss content of the letter with your laboratory director/staff, retain the notification as part of your laboratory QS documentation; if you have forwarded any of the affected products to another laboratory provide them a copy of the letter, and complete and return the enclosed Response Form within 10 days via mail to: Beckman Coulter, Inc., 1584 Enterprise Blvd., West Sacramento, CA 95691, Attn: Quality Systems &amp; Compliance or Fax: 916-374-2119 or email: MicrobiologyCustomersHCUS@beckman.com.Customers with questions are instructed to contact their Customer Support Representative at http://www.beckmancoulter.com or by phone at (800) 677-7226 in US and Canada. |  | 284 | 284 units total (96 units in US) | Worldwide Distribution - USA (nationwide) Distribution and to the countries of : Austria, Belgium, Brazil, Canada, Germany, Italy, Japan, Kazakhstan, Kuwait, Mexico, Myanmar, Netherlands, Norway, Panama, Poland, Portugal, Romania, Russian Federation, Spain, Taiwan, Thailand and United Kingdom. | Open, Classified |  |  | 284 | Not_Computer
Z-0567-2016 | 72868 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer PLUS K2EDTA Tubes (Plastic). 13 x 75 mm x |  |  | December 03, 2015 | 2015 | Becton Dickinson &amp; Company | Some of the Blood Collection Tubes in these lots were manufactured with insufficient EDTA additive. An insufficient EDTA additive may lead to erroneous platelet results, specimen recollection, or a delay of treatment. | Process change control | The firm, BD, issued an "URGENT PRODUCT RECALL" notifications/Recall Response Forms to their customers dated 12/3/2015 via email and/or UPS 2nd Day Mail. The notification identified the reason for the recall; how to identify affected product; and actions to be taken. The customers were instructed customers to immediately review their inventory; complete the enclosed Recall Response Form and fax (1-866-873-0312 it to BD or email it to bd4964@stericycle.com; and return all affected products with the completed Recall Response Form following instructions on the enclosed packing instruction, even if you do not have any affected lot in your inventory. If customers have any questions or require assistance with the return of the recalled product and/or availability of replacement product, they were instructed to contact 1-888-628-0732 between 8AM and 5 PM ET Monday through Friday. |  | 29815600 | 29,815,600 units in total | Worldwide Distribution: US (Nationwide) and the countries of: European Union, Peoples Republic of China, Japan and Republic of Korea. | Open, Classified |  |  | 29815600 | Not_Computer
Z-0566-2016 | 72712 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Valve XS Atrium Retractor |  |  | November 17, 2015 | 2015 | Aesculap, Inc. | AAG received a customer complaint that corrosion had been found by a hospitals central sterile service department on the ball joint of an FC424R blade. Investigation revealed that the corrosion was due to incorrect material used in the manufacturing of the disk spring of the ball joint by the supplier | Mixed-up of materials/components | The firm, Aesculap, sent an "URGENT MEDICAL DEVICE CORRECTION OR REMOVAL" letter dated 11/17/2015 to its customer. The letter described the product, problem and actions to be taken. The customer was instructed to review the product recall notification in its entirety; do not destroy any affected product; if you have no inventory remaining, notate that on the Product Removal Acknowledgement Form; contact our customer service department for an RGR at 866-229-3002 or email: SetWorld_Orders.BBMUS_Service@aesculap.com and immediately return affected product; complete and return the Product Removal Acknowledgement Form to Aesculap Quality Assurance department by fax to 610-791-6882 or email to val.strawn@Aesculap.com within two(2) weeks of receipt, even if the total inventory in your possession is zero (0)Sales Representative will make an appointment to come and exchange the effected product with an interium solution (FC360-Atrium Lift Retractor Set) that will allow any scheduled or unexpected cases, which would require the use of this product. The newly manufactured product will ship automatically when it's available.Then a visit will be scheduled to retrieve the loaner FC360R sets.If you have any questions, contact Quality Specialist at 610-984-9414 or email to val.strawn@aesculap.com or Product Manager at 610-984-9150 or email to mark.eppedio@aesculap.com. |  | 41 | 41 | US Distribution to state of: NY. | Open, Classified |  |  | 41 | Not_Computer
Z-0563-2016 | 72542 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ceiling Mounted Surgical Light System TruLight 3000/5000 and Ceiling Mounted Surgical Light System iLED3/iLED5 |  |  | November 02, 2015 | 2015 | Trumpf Medical Systems, Inc. | The firm has received 15 complaints over ten years related to the connection of the spring arm to the top horizontal arm of the lighting system that while positioning the light head enabled the spring arm and light head to fall. | Other | TRUMPF sent an " Urgent Field Safety Information" letter dated December 8, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken by the customer:1. Perform an inspection of the lighting systems (see Appendix 1 for instructions). If any of the symptoms listed below are observed, immediately remove the light from operation and contact your Trumpf Medical Service Representative or designee.Signs that indicate a possible fall of the spring arm:The movement of the light and spring arm is difficult for the user,The lights have stopped being operational, and/orA gap is noticed between the spring arm and the central axis (see the attached Appendix 1 for an example).2. Follow the User Manual: No unauthorized servicing of the spring arm is permitted. Snap rings are one use only and replacements must be purchased from your authorized Trumpf Medical partner. If you have any questions concerning this request or the procedure as outlined, please contact TRUMPF Medical Systems, Inc. at 888-474-9359 |  | 20713 | 20,713 | USA (nationwide) Distribution to the states of : AL, AK, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KY, ME, MA, MN, MO, NE, NV, NJ, NM, NY, NC, ND, OH, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI and WY., and Internationally to Canada. | Open, Classified |  |  | 20713 | Not_Computer
Z-0564-2016 | 72658 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Radiographic Transparency for Certain and External Connection Tapered Implants. Images that assist with measurements for dental implants. |  |  | November 16, 2015 | 2015 | Biomet 3i, LLC | Images contained within radiographic transparency ART 1036 are undersized. | Software design (manufacturing process) | BIOMET3i sent an" Urgent Medical Device Recall letter " dated November, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Responsibilities:1. Please review this notice and check your inventory for the affected item by comparing any on-hand ART 1036 Transparencies to a standard 8.5" x 11" sheet of paper. The non-conforming ART 1036 is approximately 7.8" x 10.2".2. Immediately quarantine and remove all affected items from service.3. Complete the attached Business Reply Form and either: a. fax it to 561-514-6316 orb. email it to postmarket@biomet.com orc. return it along with the affected product using the included shipping label4. Return affected product to BIOMET 3i.For assistance or other questions that you may have relative to this notice, please contact BIOMET 3i at 1- 800-342-5454 or 1-561-776-6700. Team members are available to assist you 8:00am to 6:00pm (Eastern), Monday through Friday. |  | 686 | 686 | Worldwide Distribution - US Distribution including Puerto Rico., and to the states of : CA, CO, FL, IL, MA, MI, MO, NY, OH, SC, TX, VA, WI, WV and Hawaii., and to the countries of : Australia, Germany, France, Greece, Northern Mariana Islands, Taiwan and Uruguay. | Open, Classified |  |  | 686 | Not_Computer
Z-0565-2016 | 72810 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | RenovoCath RC 120; |  |  | November 25, 2015 | 2015 | RenovoRx, Inc. | Incorrect version of Instructions for Use were shipped to two US customers. | Labeling Change Control | RenovoRx notified customers by phone and by mail of this issue. This notification requests that the customers remove the affected product from their inventory and remove and replace the incorrect Instructions for Use with the correct Instructions for Use provided by RenovoRx. The customers are requested to return the incorrect Instruction for Use to RenovoRx for disposition and are requested to complete a Recall Confirmation Form enclosed to document the completion of the requested Correction. For further questions, please call (650) 284-4433. |  | 11 | 11 units | US Distribution to the states of : CA and FL. | Terminated | February 29, 2016 | 96 | 11 | Not_Computer