## Sheet1
Recall Number | Recall Event ID | Device Type | Product Code | Regulation Description | Medical Specialty | Review Panel | Submission Type | Trade Name/Product | Main Name | Recall Class | Date Posted | Year Posted | Recalling Manufacturer | Reason for Recall | FDA Determined Cause | Action | Consumer Instructions | Clean Quantity | Quantity in Commerce | Distribution | Recall Status | Termination Date | Time to Terminate | Merged Quantities | Fault Class | Failure Mode | Action Class | Action Category
Z-0321-2007 | 36854 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | RYM-5001 InVision-PlusÂż Neutral Displacement I.V. Connector, packaged 100 units per box, 5 boxes per case (500 units/case); labeled as Sterile-Nonpyrogenic; assembled in Mexico for Rymed Technologies, Inc., Austin, TX. |  |  | May 25, 2006 | 2006 | Rymed Technologies, Inc. | Weld failures on IV connector (area of female luer and spike body) | Other | Consignees were notified by e-mail and/or phone on 05/25/06 regarding the recall. Distributors were asked to return product. |  | 28479 | 28,479 units. | Nationwide Recall | Terminated | March 26, 2007 | 305 | 79328 | Not_Computer
Z-0206-2007 | 36690 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Ultra Blood Glucose test strips, LifeScan, a Johnson &amp; Johnson Co., 50 test strips, 2 vials of 25 test strips |  |  | October 19, 2006 | 2006 | Discount Diabetic Supply | Counterfeit-blood glucose test strips (manufacturer unknown) | Other | The recalling firm notified its consignees of the problem and the recall by telephone on 10/19/2006. The firm followed this with a written notice on 10/30/2006. |  | 240 | 240 units | Nationwide distribution --- States of CA, TX, TN, GA, AL, and OK | Terminated | December 22, 2009 | 1160 | 240 | Not_Computer
Z-0339-2007 | 36693 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Bivona Adult Tracheostomy Tube, I.D. 8.5 mm, O.D 11.0 mm, Length 88 mm, TTS 11.8 mm, Mfrd by Smiths Medical Critical Care; Product code 670185. |  |  | August 31, 2006 | 2006 | Smiths Medical ASD, Inc. | Mislabeled- The outer diameter (OD) of the shaft is labeled as 11.0 m on the inner, outer and tray label in error. The OD is actually 11.8 mm. | Other | Distributors and hospital consignees were notified via recall letter dated 8/25/06 to return all unused product to the firm. Distributors were asked to notify their customers of the recall. |  | 197 | 197 | Worldwide Distribution ---- including USA states of Arkansas, California, Georgia, Illinois, Kansas, Minnesota, Missouri, Ohio, Pennsylvania, Texas and countries of Canada and Netherlands. | Terminated | May 04, 2007 | 246 | 197 | Not_Computer
Z-0452-2007 | 36774 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | LifeScan brand OneTouch Sure Step- Blood Glucose Monitoring System Test Strips, 50 strips per package, |  |  | September 29, 2006 | 2006 | Lifescan Inc | Inaccurage test results: testing may lead to no test results, error message or inaccurately low test results. | Other | The firm has issued notification on 9/29/06 to consumers and hospital end users known to the firm and direct accounts that received the recalled product. The firm also posted the notice on the website, and customer service scripting was developed to notify callers. |  | 844920786 | 26,398 units (both recall numbers Z-0452/3-2007 | Nationwide to Al, AZ, CA, CN, DE, FL, GA, IL, IN, KT, KY, LA, MA, MD, MI, MN, MS, MO, NE, NJ, NM, and TX. | Terminated | November 27, 2007 | 424 | 844920786 | Not_Computer
Z-0450-2007 | 36728 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Bayer Advia 70 Hematology Analyzer, manufactured by Drew Scientific, Inc., Dallas, TX. |  |  | September 20, 2006 | 2006 | Drew Scientific Group PLC | System has given inaccurate readings when running samples that trigger an extra rinse due to very high WBC, RBC or Platlet counts. | Other | Firm sent 'Technical Bulletin #39' to all customers on 09/20/06. Users advised of potential for inaccurate readings and provided with instructions on how to eliminate problem. |  | 545 | 545 units. | Worldwide distribution ---- including USA states of FL, MA, MN, ND, NM, NY, Puerto Rico and TX and countries of Austria, Bangladesh, Chile, China, Czech Republic, Egypt, France, Hong Kong, Italy, India, Iran, Korea, Mexico, Morocco, Phillipines, Portugal, Russia, Serbia and Montenegro, Spain, Syria, Thailand, Turkey, United Arab Emerates (Dubai), United Kingdom and Vietnam, West Indies and Yemen. | Terminated | March 10, 2008 | 537 | 844 | Not_Computer
Z-0469-2007 | 36898 | No | LCX | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Confirm Clearly Smart Pregnancy Test (Confirm Clearly Starter Kits, Product Code 30136 and Confirm Clearly Refills Kits, Product Code 30138) |  |  | November 17, 2006 | 2006 | Mizuho USA, Inc. | False positives-the firm has received a higher than expected rate of complaints regarding false-positive results. | Other | On November 14, 2006, Mizuho USA and SSL decided to cease further distribution of the product, and to contact SSL''s customers and ask them to return any unsold units. Mizuho USA informed SSL Americas by telephone.On November 21, 2006, SSL Americas sent Recall Notification Letters to each of its customers by US First Class, Certified, Return-receipt mail. The letter included a 'Fax Back' Notification to be use by the firm for recall effectiveness checks. |  | 474000 | 474,000 Kits | Worldwide-USA and Canada | Terminated | May 23, 2008 | 553 | 474000 | Not_Computer
Z-0266-2007 | 36863 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Basic/Profile and One Touch Ultra Blood Glucose Test Strips |  |  | October 14, 2006 | 2006 | Core Care Technologies, Inc | Counterfeit product (manufacturer unknown) | Other | The firm identified the customers who had been sent OneTouch test strips since August 2006 and contacted each by phone to identify the lot number of the test strips in the customer's possession. A letter was sent on 11/21/2006 by first class mail which identified the lot numbers identified in the letter from the law firm and providing instructions to follow if the affected lots were found. |  |  | unknown | Nationwide. | Terminated | June 06, 2007 | 235 |  | Not_Computer
Z-0338-2007 | 36934 | No | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences-Swan-Ganz,Thermodilution Venous Infusion Port (VIP) Catheter with AMC Thromboshield (An Antimicrobial*Heparin Coating), REF:831HF75 |  |  | November 27, 2006 | 2006 | Edwards Lifesciences Technology SARL | Misbranding-One lot of Swan-Ganz Thermodilution VIP Catheters, Model 831HF75, lot 246HC197 was labeled with different model information (831F75) on the edge label . The edge label states the model is 831F75 which incorrectly indicates that product does NOT contain an antimicrobial heparin coating. (The tray lid label correctly states the model is 831HF75, containing the heparin coating) | Other | Edwards initiated the recall in the US by sending a Customer Letter by Federal Express on November 28, 2006, along with a Q&A; to their customers Service and Direct Customers to facilitate the uderstanding of the recall and the request for verification and return of all unused product. |  | 1110 | 1,110 units | Worldwide Distribution ----USA and countries of Columbia and Venezuela. | Terminated | January 23, 2012 | 1883 | 1110 | Not_Computer
Z-0444-2007 | 36306 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett KnightStar 330 Bi-Level Ventilator, Catalogue Numbers: Y-KS330-NA, Y-KS330-SD, DSY-KS330-NA, DSY-K5330-SD, DLY-KS330-NA &amp; DLY-KS330-SD |  |  | August 31, 2006 | 2006 | Puritan Bennett Corporation | Directions for Use-Puritan Bennett has determined that a service interval should be defined for the device. The firm has also determined that an exhalation filter should be employed when the device is in use. | Other | Domestic customers of record will be advised of this matter by letter sent via Certified US Mail, Return Receipt. The letter will provide recipients with simple instructions on how they may return their units for service, and provides scheduled dates for service for each of their devices, based upon Nellcor Puritan Bennett shipment records. |  | 7754 | 7,754 (in US) | Worldwide distribution ---- including USA and countries of Argentina, Australia, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Egypt, France, Germany, Greece, Hungary, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Mexico, Netherlands, Norway, Poland, Portugal, Peru, Puerto Rico, Russia, Saudi Arabia, Scotland, Serbia &amp; Montenegro, Singapore, Slovak Republic, South Africa, Spain, Switzerland, Syria, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | March 05, 2014 | 2743 | 7754 | Not_Computer
Z-0448-2007 | 36706 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Bayer Advia 70 Hematology Analyzer, manufactured by Drew Scientific, Inc., Dallas, TX. |  |  | August 24, 2006 | 2006 | Drew Scientific Group PLC | Flags and alerts on the user interface will not automatically transfering to the LIS. | Other | Firm sent 'Product Notification #22' to all customers on 08/24/06. Users advised of potential risk to patient health and provided with instructions on how to eliminate risk by disabling the auto transfer feature and manually transfering data until a software patch is made available |  | 545 | 545 units. | Worldwide distribution --- including USA states of FL, MA, MN, ND, NM, NY, Puerto Rico and TX and countries of Austria, Bangladesh, Chile, China, Czech Republic, Egypt, France, Hong Kong, Italy, India, Iran, Korea, Mexico, Morocco, Phillipines, Portugal, Russia, Serbia and Montenegro, Spain, Syria, Thailand, Turkey, United Arab Emerates (Dubai), United Kingdom and Vietnam, West Indies and Yemen.- | Terminated | March 10, 2008 | 564 | 844 | Software | Device Operation | Software update | Software Update
Z-0466-2007 | 36868 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | BOLD/INI-CLIP 2.2mm Drill Bit with AO Attachment. |  |  | November 10, 2006 | 2006 | Integra LifeSciences Corp. | Incorrect drill length-Cannulated drill bits were manufactured to incorrect specifications (shorter length) resulting in inadequate preparation of bone prior to insertion of the UNI-CLIP Staple. | Other | An Urgent Recall Notice was sent on 11/10/2006 to both hospitals/distributors and sales representatives by first class mail. |  | 107 | 107 in US. | Nationwide to distributors, hospitals and Integra Sales Reps. | Terminated | April 16, 2008 | 523 | 107 | Not_Computer
Z-0293-2007 | 36392 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 4 inch diameter Roller Pump; Model 801040. |  |  | October 06, 2005 | 2005 | Terumo Cardiovascular Systems Corp | An incorrect serial number was placed on one roller pump. | Other | The six inch roller pump labels were corrected during in-house servicing on 10/6/05 and 10/24/05. A corrected nameplate for the 4 inch roller pump with the incorrect serial number was printed and shipped to the foreign distributor on 10/11/06, and applied in the field by the international distributor on 10/18/06. |  | 1 | 1 | Worldwide-Michigan and Singapore. | Terminated | February 01, 2007 | 483 | 3 | Not_Computer
Z-0471-2007 | 36958 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Expo32 ADC Software Version 1.1C, Part Number 6418337-New Users Software Kit, and Part Number 175454 Upgrade Software Kit used with the Coulter EPICS XL Flow Cytometers |  |  | November 01, 2006 | 2006 | Beckman Coulter Inc | The firm has confirmed that the ratio parameter assignments (numerator and denominator) of a protocol may inadvertently change after the Parameter Selection Dialog box is accessed. This issue is limited to only protocols utilizing the Ratio parameter. | Other | A Product corrective Action (PCA) letter was mailed on 11/1/06 to Expo 32 software accounts. They are informed that the ratio parameter assignments (numerator and denominator) of a protocol may inadvertently change after the Parameter Selection Dialogue box is accessed. The issue is limited to only protocols utilizing the Ratio parameter. An immediate user preventive action to remedy this issue each time acquisition of samples that require a protocol utilizing the ratio parameter is desired is given within the letter: 1) Perform acquisition of samples in batches which use protocols containing the same ratio parameter assignments. 2) Before running samples, open the protocol intended for use and open the Parameter Selection dialog box to ensure the correct ratio parameter assignments for numerator and denominator are selected. Make any necessary changes and save the protocol. 3) Run samples. (Letter was sent by US mail). A response form is included. |  | 28 | 28 units | Nationwide including the states of MA, AL, CO, MI, IN, MD, MO NY, PA, TX, AR, CA, NJ, FL, and TN | Terminated | February 27, 2012 | 1944 | 28 | Computer
Z-0454-2007 | 36821 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Phenobarbital- Reagent-Clinical Chemistry--List number: 1E08-20-Abbott Diagnostics Division |  |  | September 26, 2006 | 2006 | Abbott Laboratories Inc. | Erratic elevated results and or the inability to calibrate due to imprecision. | Other | Abbott sent out a product recall letter on 9/26, 2006, to all consignees that received lot numbers 39074HW00, 41008HW00 and 42063HW00. The letter states that the firm has identified a Phenobarbital performance issue that is demonstrated by erratic elevated results and/or the inability to calibrate due to imprecision. Until the cause of this issue is identified, they have discontinued shipment of Phenobarbital reagent. It will be necessary to identity an alternate means for Phenobarbital testing during this time. AXSYM and TDxITDxFLx Phenobarbital are available as alternate testing options (consignee may contact the local sales person or Customer Service). Consignees also advised if they forwarded any of the Clinical Chemistry Phenobarbital lots listed above to another laboratory to provide a copy of the letter to their customer. If there are questions, they are advised to call Customer Service (US consignee) or local sales person (foreign consignee). Letter advises if user currently using or have inventory of lot number 39074HW00, expiration 9/30/06, to run two levels of quality control (QC) every 6 hours and be alert to QC trends. For those using or have inventory of lot numbers 41008HW00 &amp; 42063HW00, they are advised to discontinue use of these lots and destroy them following laboratory procedures. Abbott will provide reimbursement for the destroyed kits. A customer reply form is provided. |  | 448 | Worldwide: Lot #41008HW00 - 162 kits; Lot #42063HW00 - 136 kits; Lot #39074HW00 - 150 Kits | Worldwide distribution ---- USA and countries of of Canada, Chile, Germany, Australia | Terminated | June 14, 2011 | 1722 | 448 | Not_Computer
Z-0299-2007 | 36624 | No | LGW |  |  | Neurology | PMA | Advanced Bionics Precision Linear Leads, part of the Precision Spinal Cord Stimulation System. Model Numbers: SC-2138-30, SC-2138-50, SC-2138-50T and SC-2138-70. |  |  | September 21, 2006 | 2006 | Advanced Bionics Corporation | A small number of unimplanted Precision Linear Leads may have been assembled with incorrect electrode material. At high stimulation levels, the metal may corrode and dissolved metals may enter the patient. | Other | In United States, the clinician recall notification letters were distributed by Fed-Ex or U.S. Postal Service certified mail on 09/21/2006. For the other regions, the in-country representatives started providing notification to the clinicians on 09/21/2006.The recall notification letters will instruct physicians to immediately return the unimplanted units. |  | 18 | 18 | Worldwide. | Terminated | January 25, 2012 | 1952 | 18 | Not_Computer
Z-0445-2007 | 36626 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | SPS Blue Line Tube Clamp Adapters, Product No. BL120M |  |  | October 03, 2006 | 2006 | Southern Prosthetic Supply Co. | Device Fracturing: recent design and process changes leave the device susceptible to fracturing, which could cause the prosthesis to fail | Other | On October 6, 2006, SPS faxed a recall notification to the 187 customers who had purchased one or more BL120M Tube Clamp Adapters from SPS between August 29 and September 20, 2006. On October 9, SPS contacted those customers that had purchased five or more Adapters via telephone to confirm they had received the faxed recall notification. On October 10, SPS mailed another copy of the recall notification, via certified mail, return receipt requested to all of its customers that had purchased one or more BL120M Tube Clamp Adapters. The customers were told to send the BL120M Adapter back to SPS following replacement. |  | 943 | 943 units | Nationwide | Terminated | March 07, 2008 | 521 | 943 | Not_Computer
Z-0447-2007 | 36649 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | REMstar Heated Humidifer Main PCA Repair Kit, Part number 1006265 |  |  | October 25, 2006 | 2006 | Respironics, Inc. | AC inlet connector failure that can result in thermal damage to the humidifier enclosure and/or ignition of materials external to the device | Other | The recalling firm issued a recall letter dated 10/25/06 to their domestic and Canadian dealers informing them of the problem and the need to contact the end users. These dealers were given two options: 1) contact the end user by forwarding a letter the recalling firm provided or 2) contacting the third party company with their distribution list. The recalling firm issued a recall letter dated 10/25/06 regarding the service repair kit and the need to contact end users that had the device repaired using this kit. The recalling firm issued a recall letter date 10/25/06 to the international customers informing them of the problem, contact the end users, and the need to destroy the recalled part. |  | 235 | 235 units | Worldwide Distribution ---- including USA and countries of Denmark, Austria, France, Canada, Netherlands, Japan, Hong Kong, Austrailia, Great Britain, Spain, Belgium, Germany, Poland, Ireland, Portugal, Saudia Arabia, Korea, Mexico, Finland, Taiwan, Philippines, Iceland, Brazil, Kuwait, New Zealand, Chile, India, Egypt, Norway, Turkey, Jordan, Thailand, Hungary, Russia, Peru, South Africia, Israel, Czech, Pakistan, Lebanon, Venezuela, Sweden, United Arab Emirates, Qatar, Slovakia, Bulgaria, Bermuda, and Colombia. | Terminated | December 15, 2009 | 1147 | 172235 | Not_Computer
Z-0467-2007 | 36871 | No | MBB | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | Simplex P with Tobramycin. Antibiotic PMMA Bone Cement. |  |  | November 07, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | The expiration date on the external label is incorrect in that it does not identify the earliest expiration date of the subcomponents. The External label provides for an incorrect longer expiration date than the actual subcomponents will support. | Other | A recall notification was sent by Federal Express on 11/7/06. The firm will follow-up with telephone calls and follow-up letters as necessary. |  | 813 | 813 units | Nationwide | Terminated | February 28, 2007 | 113 | 813 | Not_Computer
Z-0468-2007 | 36883 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Re-Flex VSP and Low Profile Re-Flex VSP external prosthetic foot modules with integrated shock absorber, Model numbers RFXS0002, RFXS0003, RFLS0006 and RFL50007 |  |  | November 03, 2006 | 2006 | Ossur | Inner Tube Breakage: distributed low quality units can lead to breakage of the inner tube, resulting in an abrupt rotation of the shock module, compromising function and stability of the prosthetic foot | Other | By letter on Nov 9, 2006. An initial letter was sent to all involved customers stating the invoice number, the purchase order number and the original shipping date. The problem is explained and the detailed exchange process is outlined. Contact information is included (ref. enclosure 2.1). At the same time as the mailing, all customers are being contacted proactively by a customer service representative. |  | 138 | 138 worldwide (86 in US) | Nationwide in US plus Worldwide distribution in Canada, Austria, Switzerland, Czech Republic, Germany, Spain, France, United Kingdom, Greece, Israel, Italy, Netherlands and Norway. | Terminated | May 28, 2008 | 572 | 138 | Not_Computer
Z-0478-2007 | 36187 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Minstrel Patient Lift with SR weighing scale; a non-AC-powered patient lift; Arjo, Inc., 50 N. Gary Ave., Roselle, IL 60172; model number HMB002-US |  |  | August 31, 2006 | 2006 | Arjo, Inc. | The hanger bar assembly may detach from the lift during use due to fatigue failure of the bolt. | Other | Arjo sent an Urgent Device Field Correction Recall letter dated 8/31/06 to all affected end users to the attention of the Administrator/Risk Manager via FedEx on the same date. The letters informed the accounts of the potential for the hanger bar to detach from the scale due to fatigue failure of the pivot bolt. The accounts were advised that an Arjo Field Service Technician would contact them within five business days to schedule a visit to the account to perform an upgrade of their Minstrel lifts. Any questions were directed to Arjo Quality Department at 800-323-1245, ext. 6140. |  | 76 | 76 lifts | Nationwide | Terminated | October 16, 2008 | 777 | 76 | Not_Computer
Z-0486-2007 | 36304 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | GammaMed 12i radionuclide applicator systems, distributed by Varian Medical Systems, Palo Alto, CA |  |  | September 11, 2006 | 2006 | Varian Medical Systems | Medical device for patient treatment does not meet electrical safety standards. | Other | The recalling firm notified end users by e-mail and letter on 09/11/06 and advised that some devices did not comply with electrical safety standards. The firm stated that they plan to make modifications to improve electrical safety performance but users may continue to use the device to treat patients. The notification included a response form for users to fax back verification of receipt. |  | 31 | 31 total units | Worldwide distribution, including NC, GA, IN, IL and AZ in USA, Germany, South Korea, Switzerland, Belgium, and Brazil. | Terminated | May 18, 2007 | 249 | 32 | Not_Computer
Z-0479-2007 | 36636 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BHM Kwiktrak Fixed Gate component of the Kwiktrak rail system used with the Maxi Sky 600 Ceiling Patient Lift; Manufactured by BMH Medical Inc., 2001 Tanguay Street , Magog, Quebec, Canada J1X 5Y5. |  |  | October 23, 2006 | 2006 | Arjo, Inc. | The Kwiktrak Gate stopper may fail and allow the ceiling lift to pass through to the end of the track and fall to the floor. | Other | Arjo Inc. USA sent recall letters to the direct accounts on 10/30/06, informing them that if the ceiling lift is moved too fast or too forcefully into a closed gate, the gate stopper can be damaged, causing its function to fail. One consequence can be that the gate may remain open, allowing the ceiling lift to pass through to the end of the track. The letters requested the customer to visually inspect their Kwiktrak Gate installation used with their Maxisky-600 Ceiling Lift to determine if there are any signs of impact or burr marks on the closed gate stopper. Illustrations of a damaged and normal gate stopper were included. The customers were requested to stop use of the Kwiktrak track system if any signs of impact or burr marks are found, and contact the Arjo Quality Department at 1-800-323-1245, ext. 6140 for special instructions. If no signs of damage were found, the customers were reminded to make sure that they manually drive the ceiling lift in a gentler manner to prevent damage to the gate stopper, making sure that all personnel using the ceiling lifts are properly trained. The accounts were informed that Arjo personnel will contact them in the next few weeks to install an improved system that will eliminate the risks of failure. |  | 3 | 3 gate systems | Michigan, Ohio and Illinois | Terminated | January 09, 2010 | 1174 | 3 | Not_Computer
Z-0481-2007 | 36740 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Viasys V.I.P. Bird Sterling pediatric ventilator, Catalog Number 15654 |  |  | October 20, 2006 | 2006 | Viasys Respiratory Care, Inc. | The reason for this correction is due to the potential of a repeat ''failure to cycle'' event. | Other | On November 2, 2006, a product notification letter was mailed to all customers in the United States via overnight carrier (Federal Express). Consignees are informed in the notification letter that service calls are in process of being scheduled to update domestic consignees. On November 4, 2006, a product notification letter was mailed to all international customers via registed mail. Consignees are informed in the notification letter that replacement software, replacement instructions, and installation verification forms are included in their recall package. |  | 539 | 539 | Worldwide including USA, Canada, Estonia, Belgium, Saudi Arabia, Brazil, Thailand, United Arab Emirates, Syria, Portugal, Jordan, Italy, Latvia, Hong Kong China, Ecuador, Jamaica, Dominican Republic, New Zealand, Australia, Czech Republic, Bolivia, Austria, Greece, United Kingdom (England), Turkey, Chile, Japan, India, Mexico, Uruguay, Algeria, Vietnam, Costa Rica, Ireland, Switzerland, Bahamas, Singapore, South Africa, Philippines, Spain, France, Peru, Lebanon, Taipei, Malaysia, Netherlands, South Korea, Israel &amp; Slovak Republic | Terminated | September 03, 2008 | 684 | 2017 | Computer
Z-0465-2007 | 36849 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Depuy Preservation Uni-Compartmental Knee Fixed Tibial Tray, cemented, Right Medial/left Lateral, Size 5, sterile; Part number 1498-23-005. |  |  | November 15, 2006 | 2006 | Depuy Orthopaedics, Inc. | Chamfer Step Defect-The surgeon may have difficulty in assembling /seating the polyethylene insert into the tray due to a Chamfer step defect. | Other | The firm's sales force was notified of the recall via email sent 11/15/06. The sales force was instructed to deliver a recall letter to personnel in those facilities that had purchased the products, and the recall letters instruct the facilities to quarantine the product and respond to Depuy concerning product returns. The sales force was provided with a Dear Doctor letter for issuance to implanting surgeons at the sales forces' discretion. The recall was extended to Sizes 1-5 on Dec 5, 2006 with a second recall letter. |  | 558 | 558 for Recalls Z-0456-0465-2007 | Class II Recall - Worldwide Distribution ---USA and countries of Argentina and Colombia. | Terminated | April 30, 2008 | 532 | 558 | Not_Computer
Z-0482-2007 | 36802 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BHM Reacher for Portable Ceiling Lifts; a portable ceiling lift accessory used to connect the lift to the BHM Track or KwikTrak Trolley; Manufactured by BMH Medical Inc., 2001 Tanguay Street , Magog, Quebec, Canada J1X 5Y5. The reacher comes in two models: part #700.08310 - 24'' handle length, and part #700.08320 - 36'' handle length |  |  | November 13, 2006 | 2006 | Arjo, Inc. | The carabineer may not be correctly attached to the reacher and the portable patient lift may drop. | Other | BHM sent recall letters to Arjo''s direct accounts on 11/13/06 via first class mail, informing them that fastening the carabineer to the exterior of the closed ring portion of the reacher is contrary to the manufacturer''s instructions and is an unsafe practice. The letters informed the users that Health Canada had received notification of an incident involving a reacher marketed by a competitor, and had conducted an investigation into all similar reachers available on the Canadian market. Enclosed with the letter was a copy of the BHM Reacher User''s Guide as well as a warning sticker that visually describes the safe method for connecting the carabineer to the reacher. The accounts were requested to place the warning stickers on their reachers, review the enclosed User''s Guide in detail, and ensure that all personnel using the reacher receive refresher training on its proper use and have access to the supplied User''s Guide. Once these measures have been completed, the account was requested to complete the enclosed response form and return it to BHM by fax, e-mail or mail. |  | 886 | 886 reachers | Nationwide | Terminated | January 19, 2010 | 1163 | 886 | Not_Computer
Z-0484-2007 | 36928 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Triple Channel Volumetric Infusion Pumps; Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015 U.S.A., Made in Singapore; product codes 2M8153, 2M8153R, 2M8163, 2M8163R |  |  | December 06, 2006 | 2006 | Baxter Healthcare Corp. | Baxter has identified the potential for a corruption of the memory chip in the pump to occur due to linear accelerator radiation exposure. If corruption occurs, it could result in an interruption of therapy with an audible and visual alarm notification to the user. | Other | Baxter sent the 12/6/06 Safety Alert letter to all Colleague Infusion Pump customers, to the attention of the Director of Radiation Oncology with copies to the Medical Physicist and Vice President of Nursing, via first class mail to advise them that the Colleague pump is not recommended for use in Linear Accelerator radiation therapy suites. The letters informed the users that there is a potential for a corruption of the memory chip to occur due to linear accelerator radiation exposure. If corruption occurs, it could result in an interruption of therapy with an audible and visual alarm notification to the user. The letter listed the error codes that may occur if the pump has been exposed to this type of radiation: 18:200, 18:416, 36:416, 19:113, 41:113, 314:432 and 302:435:XXXX:0006. If the pump displays any of the codes, the user was requested to submit the pump to their service facility. Failure code 199 may also occur and requires reconfiguration of the pump settings by the user's biomedical services organization. Any questions were directed to Baxter Global Technical Services at 1-800-843-7867, prompt 2. |  | 80212 | 80,212 pumps | Worldwide including USA, Puerto Rico, Aruba, Australia, Austria, Bahrain, Belgium, Bermuda, Brazil, Canada, Cayman, Chile, China, Colombia, Curacao, Denmark, Dominican Republic, El Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Honduras, Ireland, Israel, Italy, Jamaica, Korea, Kuwait, Lebanon, Macedonia, Malaysia, Malta, Mexico, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | June 30, 2011 | 1667 | 347529 | Not_Computer
Z-0488-2007 | 36889 | No | CEP | Luteinizing hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT LH Reagent, each kit contains bottles of Anti-Beta LH coated Microparticles and bottles of Anti-Alpha LH acridinium-labeled Conjugate; list numbers 6C25-22 (4 x 100 tests), 6C25-27(1 x 100 tests), and 6C25-30 (4 x 500 tests); Abbott Laboratories, Abbott Park, IL 60064-3500 |  |  | November 22, 2006 | 2006 | Abbott Laboratories MPG | The lots of ARCHITECT LH Reagent List 6C25-22, have the potential to generate patient specimen results that are elevated. Correlation studies have demonstrated an upward shift in slope for patient results of approximately 13% on average when compared to other currently available reagent lots and an average bias from patient specimens of 13% to 17% has been observed. | Other | Abbott initiated a field correction recall by letter dated 11/22/06 to all customers receiving the affected reagent. The accounts were informed of the elevated patient results with the affected lots of reagents. The accounts were requested to assess the inventory for the affected lots of reagents, discontinue use of the lots and destroy the remaining inventory of the lots, recording the number of kits on the customer reply form and faxing it to Abbott by 12/6/06, indicating that the letter was received and understood. Credit will be issued for the number of kits destroyed, as indicated on the reply sheet. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. |  | 2960 | 2,960 kits | Worldwide distribution including U.S.A., Puerto Rico, Canada, Mexico, Venezuela, Colombia, Chile, Uruguay, Argentina, Brazil, Dominican Republic, Germany, Japan, Singapore, Hong Kong, Taiwan, Thailand, South Korea, Australia, New Zealand and Panama. | Terminated | December 09, 2009 | 1113 | 2960 | Not_Computer
Z-0489-2007 | 37010 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | LCP Volar Distal Radius Plate Extra Articular 4H HD-RT-Long, LCP Volar Distal Radius Plate 2.4 mm, Catalog number 242.465 |  |  | December 05, 2006 | 2006 | Synthes | Hole is malformed, leading to reduced construct strength. | Other | All Sales Consultants were notified of the recall via teleconference on 12/5/06. The telephone script, part number and affected lot were communicated to the consultants during the teleconference. They were instructed to retrieve the affected product from the field. |  | 11 | 11 units | The products were shipped to medical facilities in NJ, TN, LA, and KS. The products were also shipped to sales reps in PA, SC, and CA. The product was also shipped to Canada. | Terminated | February 13, 2007 | 70 | 11 | Not_Computer
Z-0485-2007 | 37003 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Fisher &amp; Paykel IW930 CosyCot Infant Radiant Warmer (with integrated bed) fitted with 9001W005J Elevator Module |  |  | October 09, 2006 | 2006 | Fisher &amp; Paykel Healthcare Ltd | This firm has received Product Complaint Reports concerning bending of the pivot plate. The elevator base of the CosyCot Infant Warmer incorporates a pivot plate which forms part of the mechanism allowing the CosyCot to raise and lower in height. A bent pivot plate generally results in the Warmer being unable to elevate properly. Other component failures may result. | Other | Fisher &amp; Paykel Healthcare commenced sending out Product Correction Notifications letter via U.S. Postal Service certified mail October 9, 2006. This Product Correction Notice explained the nature of the problem and include a reminder of the maximum recommended weight of accessories. Two FAX Response Forms were provided, Part A for initial receipt of the Product Correction Notice and Part B for confirmation of the completed upgrade. |  | 282 | 273 units in US &amp; 9 in Canada | Nationwide and Canada | Terminated | April 26, 2012 | 2026 | 282 | Not_Computer
Z-0470-2007 | 36956 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | International Technidyne Corporation (ITC) ProTime Microcoagulation System Instrument-Catalogue Number-Protimeint &amp; Protimepro, Protimeintrf &amp; Protimeprorf, Protimeintl &amp; Protimeprol. |  |  | December 06, 2006 | 2006 | International Technidyne Corp. | Incorrect Test Results-ProTime ITC''s Microcoagulation System: Professional Instrument may report incorrect test results when more than 50 tests are in the memory and the technician uses the AUTO Send or PRINT RESULTS feature. | Other | Medical Device Correction letters were sent on 12/6/2006 by first class mail along with a replacement Operator's manual.On 3/19/07 the firm sent and additional letter to customers in order to ensure that self testers, as well as health care professionals and distributors were notified. |  | 1541 | 1,541 units | Nationwide to health care professionals and ITC distributors | Terminated | April 26, 2007 | 141 | 1541 | Software | Output/Calculation | Update the manual | Other
Z-0475-2007 | 36973 | No | KWL | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Howmedica Hip Fracture Stem System |  |  | November 28, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Lack of assurance of Sterility (package integrity): Anomalies in seal width and strength that may affect the integrity of the sterile barrier. | Other | Recall notification letters and product accountability forms were sent on 11/29 via Fed Ex with return receipt. |  | 126 | 126 units | Nationwide | Terminated | May 08, 2007 | 161 | 1657 | Not_Computer
Z-0355-2007 | 36731 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | TraXis Fixed Rasp in sizes: 7mm, 9mm, 11mm, 13mm and 15mm; for use with TraXis vertebral body replacement system |  |  | November 07, 2006 | 2006 | Abbott Spine | Tip of devices may dislodge while in use during surgery. | Other | All consignees that have the affected TraXis Fixed Rasps and TraXis Fixed Trials will be contacted via a documented phone script 11/07/2006. Affected inventory will be exchanged for re-worked instruments. |  | 66 | 66 units. | Worldwide, including USA, Spain, Mexico, United Kingdom, Germany, and Cyprus | Terminated | October 11, 2007 | 338 | 142 | Not_Computer
Z-0356-2007 | 36887 | No | CFN | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | Roche IGGT COBAS INTEGRA Immunoglobulin G (Turbidimetric); Catalog no. 20766631322. in vitro diagnostic. |  |  | December 15, 2006 | 2006 | Roche Diagnostics Corp. | Falsely elevated IGGT results may be reported because the R2 reagent of the Tina-quant Gen.2 albumin assay has a carryover effect on the Immunoglobulin G (turbidimetric) assay on the COBAS INTEGRA 700 and 800 analyzers and the R1 reagent of the Tina-quant Hemoglobin A1c assay has a carryover effect on the Immunoglobulin G (turbidimetric) assay on the COBAS INTEGRA 800 analyzer. | Other | Device correction letters dated 12/15/06 instructing users to perform an extra wash cycle between applications when using both the ALBT2 and IGGT cassettes on the COBAS INTEGRA 700 and 800 analyzers. The recall was extended to via device correction letters dated 2/28/07 instructing users to perform an extra wash cycle between applications when using both the A1C2 and IGGT cassettes on the COBAS INTEGRA 800 analyzer. |  | 18905 | 18,905 | Nationwide. | Terminated | April 23, 2008 | 495 | 18905 | Not_Computer
Z-0352-2007 | 36713 | No | LEG | Neuroleptic drugs radioreceptor assay test system. | Toxicology | Toxicology | 510(k) | ADVIA 1650 Chemistry System, Automated Clinical Chemistry Analyzer - human serum, plasma and urine tests. The following are the measured parameters that are within the scope of this ''field correction'': Affected assay: Valproic Acid. Affecting assays: Digoxin; Phenytoin; Gentamicin; Theophylline. |  |  | September 08, 2006 | 2006 | Bayer Healthcare, LLC (Diagnostics Division) | Bayer received a customer complaint where the Valproic Acid Quality control material recovered lower than target when it was run in a panel of other Therapeutic Drug (TDM) assays. The under recovery observed by the customer was approximately -25%. Investigation into this issue was initiated promptly following receipt of this complaint. | Software design | On 9/08/06, a Support Bulletin was sent to all affected Bayer HealthCare LLC Branches in order to inform them of the issue and to provide them with instructions for the 'field correction'. On 9/08/06, a Customer Bulletin was sent to all affected Bayer HealthCare LLC Branches both in the US and worldwide for communication with the affected end user customers. The bulletin advises of the problem and the corrective action. A Completion Confirmation (Fax-back) Form was included to ensure that all customers worldwide have been notified and have complied with the mandatory action. The Customer Bulletin is to be sent to all affected end user customers. Bayer Technical Service Specialists will follow-up with all of the ADVIA Chemistry System customers to make sure they have understood the bulletin and to ask that they return the FAX-back Completion Confirmation Form to Bayer HealthCare. If Bayer HealthCare does not receive the form in a timely manner, specialists in the Bayer Technical Care Center will contact the customer through whatever means possible in an effort to obtain 100% effectiveness on this communication. |  | 1184 | 1184 instruments total, all products | Worldwide, including USA, Puerto Rico, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Ireland, Israel, Italy, Korea, Malaysia, Mexico, The Netherlands, New Zealand, Norway, Poland, Portugal, Russia, Serbia, Singapore, Slovakia, South Africa, Spain, Saudi Arabia, Sweden, Switzerland, Syria, Taiwan, Turkey, and United Kingdom. | Terminated | May 23, 2012 | 2084 | 1184 | Not_Computer
Z-0455-2007 | 36831 | No | NAE | Hydrogel wound dressing and burn dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | HiZel, Trayless-sterile hydrogel wound dressing. Sizes 2.4 x 2.4'', 2.4 x 4.7'', &amp; 4.7 x 4.7''. Product is a transparent, flexible sheet. Packaged in a peelable pouch. |  |  | November 09, 2006 | 2006 | Embeem Inc. | Foreign material- this device, a sterile wound care product, was found to be contaminated with clumps of brown/black foreign material | Other | Recall letters were sent out on 11/9/2006 to the retail level |  | 1258 | 1,258 pieces | There are 47 customers which includes two large distributors/accounts and the rest are internet individual purchasers. These internet purchases amounts are 6 to 18 pieces at a time. There are no government accounts and no international accounts generating from EMBEEM. | Terminated | July 21, 2008 | 620 | 1258 | Not_Computer
Z-0358-2007 | 36984 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | I-Flow ON-Q PainBuster (100 ml, 2 mI/hr) Pump, PM012, Lot #662449 and 662742 |  |  | November 29, 2006 | 2006 | I-Flow Corporation | The pump flow rate labeling may not match the package labeling. The label on filter may not match the label on the top of the pump or the package labeling. | Other | A Recall Notification Letter was faxed to affected customers on November 29, 2006. Affected customers were notified of the identity of product subject to recall, the reason for recall, and instructed to quarantine immediately any remaining inventory.A Customer Response (FAX back) form was included with the Recall Notification Letter (dated November 29, 2006). Affected customers were requested to return the FAX back form to I-Flow within 48 hours of receipt. The fax back form indicates the quantity of remaining product from each lot the customer may have. |  | 3475 | 3475 | Nationwide | Terminated | December 11, 2009 | 1108 | 3475 | Not_Computer
Z-0357-2007 | 36655 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Fibrin Analysis Catheter Testing System includes the component BD Vacutainer Luer-Lok Access Device; Catalogue #364902 as a finished sterile device, Non Pyrogenic in unopened undamaged package, Disposable/ Single use/ Do Not Reuse; for use with multi lumen; (4 Venous Catheter sizes) 20 cm. 12 fr. Central; 16 cm. 7-7.5-8 fr. Central; 20 cm 7.7.5-8 fr Central; 16cm 12 fr. Central. Manufactured in the USA for: CRS Medical Diagnostics, Inc., Pewaukee, WI. |  |  | October 18, 2006 | 2006 | Continental Medical Labs,Inc | Fibrin Analysis Catheter Testying System contains the ''Luer Lok Access Device'' that fails to lock securely to certain catheter devices because the male luer taper surface is not within the specification of the taper. | Process control | Consignee was notified by certified mail on 10/19/06. |  | 840 | 840 kits | Product was distributed to one consignee located in Wisconsin | Terminated | January 18, 2007 | 92 | 840 | Not_Computer
Z-0343-2007 | 36785 | No | MJK | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Welch Allyn PIC 50 Portable Intensive Care System, Multi-Parameter Monitor/Defibrillator; a portable 12V internal battery powered defibrillator; Manufactured by MRL Inc., a WelchAllyn Company, 1000 Asbury Drive, Buffalo Grove, Illinois 60089-Part Numbers-971081,971082,971083,971084. |  |  | November 15, 2006 | 2006 | MRL, Inc., A Welch Allyn Company | Shock Delay:The monitor-defibrillator may display an 'ECG Comm' error message on the display during use, which may prevent or cause an unacceptable delay the delivery of a defibrillation shock. This delay could result in failure to resuscitate the patient. | Other | MRL sent Urgent Medical Device Recall letters dated 11/8/06 via certified mail on 11/15/06 to the direct accounts informing them of an intermittent connection between the J5 cable and its mating connector on a circuit board inside of the device that could result in a malfunction that could prevent or unacceptably delay the delivery of a defibrillating shock, which possibly could result in failure to resuscitate the patient. The accounts were provided a list of the affected units'' parts/serial numbers and were requested to inspect their PIC 50 units for the affected serial numbers and return the identified units for servicing. Until their PIC 50 is serviced, the accounts were instructed to check the unit before each shift by using the self-test, which is automatically initiated by powering up the device. If the 'ECG Comm' error occurs during use, cycling the power off-on may clear the error. Even if power cycling clears the device, the accounts were instructed to contact Welch Allyn - MRL to obtain priority servicing of the affected unit. The accounts were also instructed to respond to the notification within 30 days by fax or mail, or by calling Welch Allyn - MRL Technical Support at 1-800-462-0777, to schedule their PIC 50 for preventative service. Welch Allyn will provide the accounts with a loaner PIC 50 after receiving the form with instructions to send back the PIC 50 in the same shipping box accompanying the loaner unit. Any questions were directed to Welch Allyn-MRL Technical Support at 1-800-462-0777. |  | 510 | 510 units | Worldwide-USA, Australia, Belgium, Canada, Egypt, Finland, France, Germany, Ireland, Italy, Jordan, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sultanate of Oman, United Arab Emirates and the United Kingdom. | Terminated | October 31, 2007 | 350 | 510 | Hardware | Treatment/Delivery/Therapy | Repair | Repair
Z-0349-2007 | 36843 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | DRLock Volar Plate Screw System, Model #DVR-900. The product is a kit contained in an aluminum tray case that contains the following: IMPLANTS: |  |  | August 03, 2006 | 2006 | OrthoHelix Surgical Designs Inc | The 48 distal locking-screws (ranging from 10-24mm) contained in the kit are not within established specifications, specifically the minor diameter of the screws are undersized. | Other | The recalling firm contacted all consignees by phone on or about 8/3/06. |  | 180 | 180 | The product was distributed to 5 consignees located in the U.S.: Archway Medical, Columbia, SC; Gulf Coast Surgical, Stafford, TX; Vaugham Medical, Wheat Ridge, CO; Belemy Medical, Cleveland, OH; and Crystal Clinic, Akron, OH. | Terminated | September 10, 2012 | 2230 | 180 | Not_Computer
Z-0364-2007 | 36991 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | i-STAT PT/INR Cartridge prothrombin time test in vitro diagnostic. |  |  | December 12, 2006 | 2006 | Abbott Point of Care Inc. | Patients treated with the antibiotic CUBICIN can cause a concentration dependent false prolongation of prothrombin time (PT) and elevation of INR (International Normalized Ratio) when using the i-STAT PT/INR cartridges. | Other | Customers were notified by letter sent by regular mail on 12/13/2006. |  | 337080 | 14,045 boxes of 24 | Nationwide to hospitals and health care facilities | Terminated | August 15, 2007 | 246 | 337080 | Not_Computer
Z-0362-2007 | 37071 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Suros ATEC Breast Biopsy System handpiece (9 gauge cannula 12 centimeters in length with a 20 millimeter aperture) for use with the ATEC Breast Biopsy and Excision System, disposable, sterile; Part no. 0912-20. |  |  | December 15, 2006 | 2006 | Suros Surgical Systems Inc. | The handpiece package may have an incorrect part number on the inner package. Error could result in over-penetration of the biopsy needle. | Other | Consignees were notified by phone beginning on 12/15/06 and by follow-up letter dated 12/19/06. |  | 1249 | 1249 handpieces | Nationwide. | Terminated | May 04, 2007 | 140 | 1249 | Not_Computer
Z-0361-2007 | 37011 | No | DKJ | Salicylate test system. | Toxicology | Toxicology | 510(k) | Salicylate-SL, in vitro diagnostic, Cat. No.511-20, Lot No. 29892 |  |  | December 05, 2006 | 2006 | Diagnostic Chemicals, Ltd. | Upon subsequent testing on different instruments (as required by other customers), DCL detected a non-linear response on certain analyzers. A decision was made to recall the product due to the fact that DCL customers sell the product to different end users, DCL cannot be certain which instruments the end users employ in their laboratories | Other | For product shipped by DCL to one customer notification letter dated 12/5/06 sent out to customer to return/discard product. Replacement will be sent to customer. For product shipped to DCL (USA) warehouse notification sent to discard kits remaining in stock. For product remaining at DCL Canada in-house notification sent out to discard in-house kits. |  | 55 | 55 kits shipped to one customer | PA | Terminated | January 27, 2016 | 3340 | 55 | Not_Computer
Z-0340-2007 | 36333 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept Custom Tubing Pack. Do Not Reuse. The design and components of this product were specified by the user from an array of components and designs offered by the manufacturer. Any change in the cardiopulmonary bypass procedure that compromises the function of this Custom Tubing Pack is the responsibility of the user. The tubing is made of polyvinyl chloride (PVC). Portions of the tapered tubing are used in the vent, suction and arterial roller heads of the heart lung machine. Manufacturer: Medtronic, Inc., Minneapolis, MN 55432-5604 USA. Tube supplier: Natvar (Tekni-Plex Co.), 19555 E. Arenth Ave., City of Industry, CA 91748. |  |  | September 13, 2006 | 2006 | Medtronic Perfusion Systems | An out of specification condition exists with the tapered tubing found in some Custom Tubing Packs. The wall thickness is undersized and does not meet specifications. | Other | An Important Customer Notification, dated 09/06/06, was sent to affected customers. The letter identifies the affected model and lot numbers. Medtronic requests that affected product be returned. A suplemental pack can be built and shipped to the customer upon request. Medtronic will replace all returned product. A Recall Certificate is asked to be completed by customer and faxed to Medtronic. |  | 146 | 146 | Distributed in 6 states: CA, FL, IL, IN, LA and TN | Terminated | February 03, 2007 | 143 | 146 | Not_Computer
Z-0363-2007 | 36256 | No | KFX | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | Electronic Otto Bock Compact leg prosthesis System, 3C95/3C85, to be used exclusively for the exoprosthetic fitting of amputations of the lower limb. Microprocessor-controlled prosthetic knee joint used by lower extremity amputees. |  |  | August 10, 2006 | 2006 | Bock,Otto,Orthopedic Ind,Inc | A correction of the Otto Bock Compact prosthetic knee joint, material numbers 3C93 and 3C85, is being conducted. A limited number of Compacts have a defective solder connection between the contact wires and battery terminals. If exposed to strong vibrations these Compact knee joints have an increased probability of malfunction and a resultant risk to the Compact wearer. | Other | An Urgent Correction Notice, dated August 11, 2006, was sent to customers who purchased the Compact Prosthetic knee joint. The letter identifies affected units, asks for all in stock devices to be returned to Otto Bock, and for patients fitted with an affected device, a loaner knee will be provided as a temporary replacement while their device is serviced. |  | 176 | 176 | Nationwide. | Terminated | November 26, 2008 | 839 | 176 | Not_Computer
Z-0342-2007 | 36672 | No | LYJ |  |  | Neurology | PMA | Handhelds-Cyberonics VNS Therapy System, Model 250-''Handhelds'' preprogrammed with software v6.0, v6.1, v6.1.7, v7.0 and v7.1.3; manufactured by Cyberonics, Inc., Houston, TX. |  |  | October 17, 2006 | 2006 | Cyberonics, Inc | During programming, pulse generator may be inadvertently set to 8.0 mA output, regardless of the mA range selected by the clinician | Other | Firm sent safety alert to physicians on 11/14/06. Safety Alert informs physicians of the problem and provides instructions on how to avoid an event and correct an event should it occur. |  | 7404 | 7,404 Handheld PCs pre-programmed with software. | Worldwide Distribution ---- USA and countries of Argentina, Australia, Austria, Belgium, Canada, Cyprus, Czech Rep, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, Ireland, Israel, Italy, Japan, Jordan, Malaysia, Mexico, Netherlands, Norway, Poland, Portugal, S. Korea, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Kingdom, and Venezuela. | Terminated | October 02, 2009 | 1081 | 10942 | Other | Output/Calculation | Instructions to avoid | Safety Notice/Insructions
Z-0372-2007 | 37089 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Ventilator System SERVO-i V3.1 PC1777 |  |  | May 23, 2006 | 2006 | MAQUET Inc. | Reported failures of the SERVO-i ventilators user interface system causing technical errors and increased fault frequencies. Switches may be activated unintentionally, or a function not activate when the switch is selected. Direct access knobs may not function or change the parameter in the wrong direction compared to how they are turned. | Other | Maquet sent an ''Extended Warranty'' letter on May 23, 2006. On March 13, 2007, Maquet issued a Recall Notification letter to all customers stating that they will replace all PC1777 boards in the serial number range 2400-3300, which have not previously been replaced. A safety alert letter was included outlining the failure indication and effect for the user of the failures, as well as, how to identify whether the Servo-i unit has shown any failure indications. |  | 1970 | 1, 970 units | Worldwide | Terminated | April 11, 2008 | 689 | 1970 | Other | Device Operation | Replace board | Remove or Replace
Z-0350-2007 | 37023 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Jostra Heater Cooler Unit HCU30, system version 1.02-Low voltage machine |  |  | November 10, 2006 | 2006 | MAQUET Inc. | The HCU 30 system potentially will not be have the power capacity to run the compressor when the heaters are on during operation. This may lead to a situation where the user may not be able to cool the patient and/or cardioplegia sufficiently. | Other | Letters were mailed to customers on 11/10/2006 by Federal Express. |  | 13 | 13 units | Hospitals in CA, NC, FL, WY, NY and WV.S | Terminated | February 05, 2007 | 87 | 13 | Not_Computer
Z-0378-2007 | 37051 | No | KAC | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Bivona Laryngectomy Tube, Cuffless, Silicone, I.D. 9.5 mm, O.D 20.0 mm, Length 85 mm, Mfrd by Smiths Medical Critical Care; REF BOSL1L. |  |  | December 15, 2006 | 2006 | Smiths Medical ASD, Inc. | The label incorrectly states the outer diameter of the tube is 20.0 mm when it is actually 12.0 mm. | Other | Consignees were notified via recall letter dated 12/15/06. Distributors were requested to provide the firm with final user name and address information so that they could notify them of the recall. The firm is advising users that they can either return the product for replacement or continue to use it now that they are aware of the labeling error. |  | 117 | 117 | Worldwide, including USA, Canada, Netherlands and United Kingdom. | Terminated | December 04, 2007 | 354 | 117 | Not_Computer
Z-0376-2007 | 37097 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Bio-Rad D-10 Dual Program Calibrator Diluent, a component of the Bio-Rad D10 Dual Program diagnostic test. This is a specific part of the kit, and only the calibrator diluent is defective. It comes in two kits. |  |  | November 17, 2006 | 2006 | Bio-Rad Laboratories Inc | The Diluent in the D-10 Dual Program Calibrator reorder packs is defective, and can cause early calibration failure or elevated control and sample recovery results. | Other | The firm has issued phone calls and faxes to its US consignees starting 11/17/2006 and International notices began on 11/20/2006 explaining the reason for the recall and requesting the product be disposed. |  | 943 | 900 kits, 43 sets distributed. | Worldwide, including USA, India, Canada, France, Sweden, the UK, Hong Kong, Singapore, Germany, Brazil, Italy, China, Costa Rica and Panama. | Terminated | November 28, 2007 | 376 | 943 | Hardware | Device Operation | Remove | Remove or Replace
Z-0373-2007 | 37121 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation; for diagnostic image analysis. GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | January 12, 2007 | 2007 | GE Healthcare Integrated IT Solutions | The Centricity PACS RA1000 Workstation Software may manifest certain mammography image display problems only if the images were acquired through GE Senographe versions 2000D, DS or Essential. | Other | GE Healthcare IITS sent recall letters dated 1/12/07 to all customers who have the combination of GE Centricity PACS RA1000 Workstation software versions 2.1.X.X, or 3.0.X.X combined with the GE Senographe 2000D, the Senographe DS, or Senographe Essential mammography scanners, informing them that the entire breast image or a portion of the breast image may not be displayed in the monitor regions and may not be evident to the end-user. To mitigate this problem, the users were provided a Quick Reference Guide detailing how to use the Pan tool to move the image left, right, up and down in order to display any portion of the image that may not be displayed on the screen. The customers were requested to make sure that all necessary physicians and personnel at their location are made aware of the issues and understand the risk mitigation provided. A software patch to permanently resolve the issue will be installed at no charge when it becomes available. Any questions were directed to the GE Customer Care Center at 800-437-1171. |  | 198 | 198 units | Worldwide including USA, Australia, Belgium, Canada, China, Denmark, France, Germany, Hong Kong, Israel, Italy, Japan, Korea, Kuwait, Malaysia, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, Taiwan, Turkey and the United Kingdom. | Terminated | February 03, 2010 | 1118 | 198 | Software | Display/Image | Software update | Software Update
Z-0377-2007 | 37115 | No | FED | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Flexor DL Dual Lumen Ureteral Access Sheath with AQ Coating, sterile, REF FUS-095013-DL. |  |  | January 05, 2007 | 2007 | Cook Urological, Inc. | Lack of assurance of sterility, as the packages may not have been sealed. | Other | Consignees were notified via recall letter dated 1/5/07 and requested to return the product. |  | 4 | 4 | Massachusetts, Maine, New Hampshire and Puerto Rico. | Terminated | May 01, 2008 | 482 | 4 | Not_Computer
Z-0368-2007 | 37012 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura FD 10 FC, X-Ray System, angiographic - a monoplane, digital Allura FD (flat detector) with a 10' Image Intensifier (II). |  |  | December 14, 2006 | 2006 | Philips Medical Systems North America Co. Phillips | Movement of the AD5 table or Lateral C-arm (LARC). The table and/or LARC will not move upon request, will move slowly when movement is requested or will not stop moving when requested. | Other | On 12/14/05 and 12/20/06 the firm issued a letter dated 12/12/06 to their customers. The letter explains the problem and states that the firm will be performing mandatory upgrades to correct the situation. Consignees will be contacted to arrange for service. |  | 140 | 140 devices total | Units were distributed to hospitals and medical centers nationwide. | Terminated | October 14, 2011 | 1765 | 140 | Not_Computer
Z-0365-2007 | 37013 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sterile Neolon 2G Latex-Free Powder-Free Neoprene Surgical Gloves, Does Not Contain Natural Rubber Latex; 1 pair of gloves per package, 100 packages per case; Manufactured for Medline Industries, Inc., Mundelein, IL 60060 U.S.A., Made in Malaysia |  |  | December 18, 2006 | 2006 | Terang Nusa Sdn Bhd | Some of the packages labeled as latex-free actually contain latex gloves. | Other | Terang Nusa telephoned Medline Industries on 12/13/06 requesting them to retrieve the affected lot of size six gloves from their customers. Medline Industries notified their customers by letter dated 12/20/06, informing them of the possible latex gloves being labeled as the Neolon 2G Neoprene Surgical Gloves, and requesting them to inspect their inventories for the suspect size and lot of gloves. The accounts were instructed to complete the enclosed response questionnaire, indicating the amount of affected gloves found, and fax it back to Medline at 847-949-2643 to obtain a return goods authorization. Any questions were directed to Kathy Dunne at 847-359-1704. |  | 43 | 43 cases | Worldwide including USA, Canada, and Panama | Terminated | August 12, 2009 | 968 | 43 | Not_Computer
Z-0359-2007 | 36990 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bard Composix Kugel Hernia Patch- |  |  | January 10, 2007 | 2007 | Davol, Inc., Sub. C. R. Bard, Inc. | PET recoil ring breakage-Memory recoil ring may break which could potentially lead to bowel perforation and or chronic enteric fistula | Other | Davol Inc. notified Hospital Administrators/Risk Health Managers,distributors, and surgeons of the expanded recall by letter beginning 1/10/07. Product under recall is to be returned to Davol for replacement. Davol is posting recall information to their website. |  | 32581 | 32,581 units | Nationwide | Open, Classified |  |  | 49184 | Not_Computer
Z-0374-2007 | 36922 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Davol Salute Fixation System, 38cm Reusable Shaft (Laparoscopic) with Handle. To attach mesh to the abdominal wall. |  |  | December 27, 2006 | 2006 | Davol, Inc., Sub. C. R. Bard, Inc. | Inadequate directions for use in the Salute System. Labeling -Revised to provide updated instructions for use and proper servicing | Other | Davol notified users by a first letter on 12/27/06 providng a introductory letter explaining the Enhanced Servicing Program and changes made to the IFU (re: servicing, tip inspection, end of wire, and test fire protocol). A second letter will issue to accounts to identifying the serial number to be serviced and returned to Davol. |  | 4787 | 4787 (includes both types) | Worldwide, including USA, Austria, Italy, Israel, Finland, Switzerland, Portugal, Japan, France, Spain, Greece, Belgium, Germany, Sweden, Holland, Denmark, UK, and South Africa. | Open, Classified |  |  | 4787 | Not_Computer
Z-0381-2007 | 36963 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Integrated Centrifugal System Control Unit; Catalog number 801046. |  |  | November 08, 2006 | 2006 | Terumo Cardiovascular Systems Corp | Following installation of software version 1.30 in October 2006, the firm confirmed a complaint that the new software caused a discrepancy in the pump speed calculation and caused the pump to stop during clinical use and a ''connect motor'' message to be displayed. | Other | Customers were notified by phone, fax, email or by a Terumo representative of these problems between 11/08/06 and 11/13/06 and instructed to not use the unit until new software had been installed. All version 1.30 software was replaced and version 1.40 installed prior to 12/4/06. |  | 35 | 35 | Nationwide. | Terminated | June 11, 2007 | 215 | 35 | Software | Output/Calculation | Software update | Software Update
Z-0389-2007 | 37006 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 Protege GPS Self-Expanding Nitinol Stent Biliary System. SERB65-09-60-120. |  |  | December 06, 2006 | 2006 | Ev3, Inc | Two individual ev3 Protege GPS Biliary Stent System devices were incorrectly labeled as SERB65-09-60-120 and the actual device is SERB65-12-60-80. Length is shorter than expected. | Other | Letter 12/06/2006. An ev3 representative visited the affected hospital to determine the status of the recalled device. If the device is unused, the representative is to retrieve the device for return to ev3. The representative is to complete the Device Recall Field Action Form, along with a hospital representative signature, and fax the completed form to ev3. |  | 2 | 2 | IL, PA | Terminated | February 03, 2007 | 59 | 2 | Not_Computer
Z-0390-2007 | 37019 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | ACS Rotating Hemostatic Valve, 0.096 inch, for use with a dilatation catheter - Reference Number 23242 |  |  | December 05, 2006 | 2006 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Specific lots of the .096' and .115' RotatingHemostatic Valves (RHV) and the Duostat .096' Rotating Hemostatic Valve (RHV) potentially have an incomplete seal in the packaging pouch | Other | On December 5, 2006, the firm began notifying all affected customers, affiliates and distributors of the recall via letter. The letter advises that the firm has voluntarily recalled specific lots of the .096' and .115' Rotating Hemostatic Valves (RHV) and the Duostat .096' Rotating Hemostatic Valve (RHV) due to a potentially incomplete seal in the pouch. The letter advises Abbott Vascular will be replacing all returned product and that appropriate corrective actions will be implemented. |  | 24365766 | 6090751 - 791 devices, 6090752 - 656 devices, 6091151 - 381 devices, 6091152 - 132 devices | Worldwide, including USA, Canada, Germany, Italy, Greece, New Zealand, Netherlands, France, Japan, Spain, and Bulgaria. | Terminated | July 24, 2012 | 2058 | 85277830 | Not_Computer
Z-0393-2007 | 37021 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | G2 Filter System - Jugular/Subclavian Delivery Kit, for use in the prevention of recurrent pulmonary embolism. Catalog Number RF-320J |  |  | December 20, 2006 | 2006 | Bard Peripheral Vascular Inc | This action is being taken because BPV has recently received three complaints for the device that involve reports of introducer sheath tip damage. The damage may result in loss of tip integrity. | Other | Recall letters were sent to all consignees via Federal Express overnight delivery with signature receipt required on December 20, 2006. Each letter included a Recall and Effectiveness Check Form for each consignee to fill out and fax to Bard Peripheral Vascular, Inc (BPV). Customers will be asked to immediately remove product from inventory and isolated from use and contact BARD for return authorization. All returned product will be destroyed. |  | 5761 | 5761 units | Worldwide including USA, Australia, and Canada | Terminated | October 23, 2008 | 673 | 11770 | Not_Computer
Z-0395-2007 | 37126 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific brand iLabÂż Ultrasound Imaging System, for intravascular use; |  |  | December 01, 2006 | 2006 | Boston Scientific Target | The product fails to comply with applicable electronic product performance standard, in that the electromagnetic energy is in excess of the labeled conformance standards for radiated emissions. | Other | A customer notification was sent December 1, 2006 indicating the need for the correction. |  | 24 | 24 units | Nationwide and Portugal. | Terminated | December 05, 2007 | 369 | 24 | Hardware | Treatment/Delivery/Therapy | Correction | Repair
Z-0397-2007 | 37143 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PortraitÂż PSR3 System, an electro-surgical device used in dermatological applications |  |  | October 02, 2006 | 2006 | Rhytec Inc. | Inadequate Directions for Use: Software Upgrade to set maximum energy level at 2.5 Hz | Other | Rhytec notified users by letter dated 10/2/06 of the software upgrade. A service representative will perform the update. |  | 234 | 234 units | Canada, Australia, Brazil, Korea | Terminated | February 02, 2016 | 3410 | 234 | Software | Output/Calculation | Software update | Software Update
Z-0398-2007 | 36458 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Product is Mucotrol Concentrated Oral Gel Wafer contained in HDPE Bottle. |  |  | September 25, 2006 | 2006 | Belcher Pharmaceuticals Inc. | Processing and cleaning process validation were not completed before production. Product had some GMP failures related to the quality system. | Other | Consignees were notified of recall via letter dated 09/25/2006. |  | 1097 | 1097 bottles | Nationwide. | Terminated | July 22, 2008 | 666 | 1097 | Not_Computer
Z-0421-2007 | 36967 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Desktop ProÂż R6.0 &amp; R6.1, Linear Medical Accelerator. |  |  | November 23, 2006 | 2006 | Elekta, Inc. | Unexpected Diaphragm movement when manual field size modifications are not saved. | Software design | Consignees were notified by letter on/about 11/23/2006 |  | 37 | 37 units | Worldwide distribution ---- including USA states of AZ, AL, CA, CO, FL, GA, IA, MI, NV, OH, OK, PA, WA, WV and Country of Canada . | Terminated | May 14, 2008 | 538 | 37 | Not_Computer
Z-0416-2007 | 36996 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i2000SR Processing Module, for in vitro diagnostics, List Number 3M74-01; Temperature Controller Board Part #78560-109; distributed by Abbott Laboratories, Irving, TX. |  |  | December 12, 2006 | 2006 | Abbott Laboratories, Inc | The re-designed ARCHITECT Temperature Controller Board has been identified as being susceptible to electromagnetic interference (EMI) in the laboratory. This can cause a board reset condition, stopping temperature control function of the board. Situation does not stop analyzer operation and does not generate an error condition alerting operator when condition occurs. | Other | Firm contacted consignees by telephone and fax with instructions to verify proper functioning of the Temperature Controller Board starting December 13, 2006. |  | 648 | 648 units. | Worldwide, including USA, Argentina, Australia, Brazil, Canada, Chile, Columbia, Dominican Republic, Germany, Hong Kong, Jamaica, Japan, Korea, Mexico, New Zealand, Peru, Puerto Rico, Singapore, Taiwan, Thailand, Trinidad, Tobago, Uruguay, and Venezuela. | Terminated | September 24, 2007 | 286 | 705 | Hardware | Device Operation | Instructions to verify | Safety Notice/Insructions
Z-0388-2007 | 37001 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Mach 1 Peripheral Guide Catheter, 6F, 55cm, LIMA Tip Shape, UPN # M001196620, Catalog # 19-662. Made in Mexico: Edificio G1 Local B, H, Parque Industrial La Mesa, Tijuana, Baja California 22650 MEX. |  |  | December 13, 2006 | 2006 | Boston Scientific | Excess strands of resin may exist in the lumen near the hub of certain batches/lots of the Mach 1 Guide Catheters. If this excess resin is present in the catheter, there is the potential of an embolization resulting from the excess resin detaching from the device during the procedure. | Other | An Urgent Medical Device Recall letter, dated 12/13/06, and a Clarification of Recall Notification letter, dated 12/14/06, were sent to all affected hospitals. The letter describes the issue, affected product, asks for customers to discontinue further distribution or use of any remaining product and asks for product to be returned. Replacement product will be available for any/all returned product. |  | 18 | 18 | Nationwide. | Terminated | June 28, 2007 | 197 | 51 | Not_Computer
Z-0417-2007 | 37065 | No | DDC | Thyroid autoantibody immunological test system. | Immunology | Immunology | 510(k) | AccessÂż Immunoassay Systems Thyroglobulin Antibody Reagent Kit, Part Number 33890 |  |  | October 26, 2006 | 2006 | Beckman Coulter Inc | Low end imprecision affecting both S0 calibrators and patient samples which can lead to inaccurate low level patient results when using the Access Thyroglobulin Antibody assay. | Other | A Product Corrective Action (PCA) letter was mailed on Oct 27, 2006 to all Thyroglobulin Antibody customers. They are informed that there has been an increase in customer feedback regarding low end imprecision affecting both S0 calibrators and patient samples, which can lead to inaccurate low level patient results when using the Access Thyroglobulin Antibody assay. The letter informs the customer that the firm recommends that they use an alternate method for testing TgAb. (Letter was sent by US mail). Customers are asked to complete and return the enclosed response form within 10 days. |  | 474 | 474 units | Nationwide and Canada | Terminated | April 30, 2012 | 2013 | 474 | Not_Computer
Z-0396-2007 | 37139 | No | MQU | Eye sphere implant. | Ophthalmic | Ophthalmic | 510(k) | Perry-Kolberg Titanium Motility/Support System (Threaded Sleeve and Flat Peg), Model Number 100045 |  |  | December 20, 2006 | 2006 | Integrated Orbital Implants Inc | The threads on the threaded sleeve were manufactured with left-handed threads, requiring a modification to the surgical technique for proper insertion. Failure to properly identify the affected components and apply the modified insertion technique will result in an inability to insert the sleeve into the implant. | Other | All domestic customers were notified by phone call on 12//20/2006 which was followed up with a fax and the letter sent registered mail. Internationally, customers were notified by fax, email and mail.Customers were instructed to check their inventory of the product by lot number and visual inspection. They were instructed to isolate the affected inventory. They were provided with alternate instructions on how to use the product as it is, if needed. They were provided with a reply card to complete and fax back to the company. They were instructed that when replacement product becomes available, they will be contacted to arrange for the exchange of the affected product. |  | 234 | 234 | Worldwide, including USA, United Arab Emirates, Canada, Mexico, Costa Rica, Taiwan &amp; Romania. | Terminated | July 12, 2007 | 204 | 234 | Not_Computer
Z-0399-2007 | 37024 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon X3 CR Tibial Bearing Insert |  |  | December 04, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | A product mix-up in which the labeling indicates a Triathlon X# CR Tibial Bearing Insert Size #2 - 11MM, however, the device inside the package is actually a Triathlon X3 CR Tibial Bearing insert Size # 7 - 11MM and vice versa. | Other | Notification letters and Product Accountability forms were sent on 12/8/2006 by Federal Express with return receipt. |  | 23 | 23 units | Nationwide and Canada. | Terminated | February 21, 2007 | 79 | 47 | Not_Computer
Z-0410-2007 | 37057 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD BeaverÂż Mini-Blade, Mini/Bulk 69, Catalog Number: 374823 |  |  | December 26, 2006 | 2006 | Becton Dickinson and Company | Product sterility may be compromised due to incomplete package seal | Other | On 12/26/2006 BD Medical, Ophthalmic Systems mailed a recall notification letter to US Distributors and US OEM''s. Customers are instructed to return any affected products to BD for a replacement. On 1/10/07, the firm expanded the recall to additional lots. Those lots are noted with an asterisk. |  | 43600 | 43,600 units | Worldwide, including USA, Asia, Australia, Canada, Mexico, Taiwan, Singapore, Korea, and Japan. | Terminated | March 07, 2016 | 3359 | 664050 | Not_Computer
Z-0411-2007 | 37066 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Rocker Reducer, part of the SiLo Spinal System, REF 8880014, Rx only, Material: Stainless Steel, qty: 1 ea, non-sterile, Medtronic Sofamor Danek, 1800 Pyramid Place Memphis, TN 38132, USA, Manufactured at: Bartlett, Tennessee, USA |  |  | December 18, 2006 | 2006 | Medtronic Sofamor Danek Instrument Manufacturing | Due to a manufacturing error, the proximal end may deform causing the protrusions that engage the associated implant to break off of the instrument during use. | Other | The firm sent out the recall letter via Federal Express to direct consignees on 12/18/2006. The letter explained the problem, requested discontinuation of use and stated that sales representatives would be coordinating the retrieval and replacement of the device. |  | 58 | 58 units | Worldwide, including USA, the Netherlands, Canada, Australia, and the Bahamas. | Terminated | July 16, 2007 | 210 | 58 | Not_Computer
Z-0418-2007 | 37083 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 with CXP software, Versions 2.0 &amp; 2.1: Part #722645, Software Kit; CXP version 2.1 New User; Part #723113 Software Kit; CXP version 2.1 Upgrade; Part #629637 Software Kit; CXP version 2.0 New User; Part #629636 Software Kit; CXP version 2.0 Upgrade. The Cytomics FC 500 is a system for the qualitative and quantitative measurement of biological and physical properties of cells and other particles. |  |  | August 29, 2006 | 2006 | Beckman Coulter Inc | Two workflow scenarios associated with renaming regions may produce incorrect results. | Other | A Product Corrective Action (PCA) letter was sent on the week of August 29, 2006 to all accounts that have an FC500 or have ordered CXP or MXP software to inform them that a) incorrect statistics are generated when a polygonal region with vertical re-entrant vertices is renamed using the Region Properties dialog box. b)when a region is copied from one plot to another, and then the new region is renamed, the statistics are not updated with the new region name. (Letter was sent by US mail). Letter includes immediate user preventive action to remedy both issues. A response form is included for consignees to return to recalling firm. |  | 397 | 397 nationwide | Nationwide &amp; Canada | Terminated | April 19, 2012 | 2060 | 623 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0414-2007 | 37162 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AutoSuture Bladed Trocar, equipped with safety shield. Model number 179094. Available with smooth or threaded sleeve in sizes 5-15mm inner diameter and 70-100mm length. |  |  | October 19, 2006 | 2006 | Ascent Healthcare Solutions, Inc. | The tip of the trocars may separate before and while in use. In some instances the trocars have been found with exposed blades upon opening the packaging. | Process design | The recall was initiated on 10/19/2006 by direct distribution of the recall notice to direct accounts by the firm's sales force. The firm also posted the notification on their web site. All customers were instructed to examine their inventories for the product and to remove and quarantine all affected product. Instructions on returning the product to the firm were also included along with a response form. |  | 1566 | 1566 | Nationwide | Terminated | October 31, 2008 | 743 | 1566 | Not_Computer
Z-0412-2007 | 36154 | No | DSQ |  |  | Cardiovascular | PMA | Thoratec Dual Drive Console a drive component of the Thoratec Ventricular Assist Device, Model Numbers, 10025-2600-005, 10025-2601-007, 10025-2602-006, Thoratec Corporation, 6035 Stoneridge Drive, Pleasanton, CA 94588 |  |  | September 24, 2004 | 2004 | Thoratec Corp | The module power switch (at the back of console) may break while attempting to turn the Dual Drive Console OFF or ON. | Other | The firm issued on 1/25/07 an URGENT: MEDICAL DEVICE CORRECTION-notice informing its consignees about the 'Power Switch Replacement Instructions' and arrangements for their service department replacement of the switches. |  | 331 | 331 | Worldwide | Terminated | April 18, 2008 | 1302 | 331 | Not_Computer
Z-0379-2007 | 37099 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Ventilator System Servo-i V3.1 |  |  | May 23, 2006 | 2006 | MAQUET Inc. | The Connection of certain nebulizers to the built in driver has lead to reduced or no nebulization effect on the Servo-i Ventilator | Other | Maquet Critical Care issued an ''Extended Warranty'' letter on 5/23/2006. On March 13, 2007, Maquet modified its recall strategy and issued a Recall Notification letter accompanied by a Safety Alert letter to all customers with Servo-i ventilators, serial numbers in the 14000-22772 range, with Servo nebulizer installed. The Safety Alert letter provides customers with information on how to identify if the unit is affected, and how to identify if the unit has shown any failures, as well as, any precautions. Firm initiated recall is on-going |  | 2436 | 2,436 units | Worldwide | Terminated | March 23, 2009 | 1035 | 2436 | Not_Computer
Z-0432-2007 | 36490 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 100V -120V, 15 A (circuit breaker), 50/60 Hz (20A power source required) base; Model 801763. |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | A 'battery needs service' message may be displayed, or the remaining battery time may incorrectly be reported out as zero, even when the battery is fully functional. | Other | Consignees were sent a letter dated 11/10/06 informing them of the nature of the issues and that the firm would contact them to schedule a service call to install the software upgrade package. The software upgrade installation began at customers on 10/16/06, prior to issuance of the letter to consignees. |  | 489 | 489 | Worldwide distribution ---- including USA and countries Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 696 | Software | Alarm/Message | Software update | Software Update
Z-0436-2007 | 36491 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; Central Control Monitor; Catalog number 802100. |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | Pump may stop without an audible tone as a result of quickly clearing the internal pump alarm condition or overpressure alert/alarm condition. | Other | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06. |  | 683 | 683 | Worldwide distribution ---- including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 1379 | Software | Alarm/Message | Software update | Software Update
Z-0438-2007 | 36494 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 4 inch diameter Roller Pump; Model 801040. |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | The roller pump will pause in response to a pressure alarm condition, as designed, but then may not resume operation as expected after the alarm condition has been cleared. | Other | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06. |  | 1399 | 1,399 | Worldwide distribution ---- including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 3350 | Computer
Z-0493-2007 | 36496 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; CAPIOX Control Module; Catalog number 811113 (Not distributed within the United States). |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | The pump may lose communication with the central control monitor, causing the monitor to display a service message "System Computer Needs Service" and to become non-operational. | Other | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06. |  | 50 | 50 | Worldwide distribution --- including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 3743 | Computer
Z-0494-2007 | 36498 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; Air Bubble Detection Module; Catalog number 802110. |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | The air bubble detection system may falsely continue to sound the air alarm after the user has disabled it. | Other | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06. |  | 806 | 806 | Worldwide distribution --- including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 806 | Software | Alarm/Message | Software update | Software Update
Z-0437-2007 | 36492 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 6 inch diameter Roller Pump; Model 801041. |  |  | October 16, 2006 | 2006 | Terumo Cardiovascular Systems Corp | The large roller pump may experience false "overspeed' error message when starting rotation, which will prevent the pump from operating; until the condition is cleared by power cycling the roller pump. | Other | Consignees were notified of the problem via letter dated 11/10/06 and informed that they would be contacted to schedule a service call to install software upgrades. Installation of software upgrades began at customer locations on 10/16/06. |  | 1951 | 1,951 units | Worldwide distribution ---- including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Egypt, Germany, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Republic of China, Singapore, South Africa, South Korea, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 995 | 1951 | Software | Alarm/Message | Software update | Software Update
Z-0477-2007 | 36932 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Ultra Blood Glucose Test Strips-Lot # 2691191, LifeScan |  |  | October 09, 2006 | 2006 | Matrix Distributors, Inc. | Counterfeit Test Strips (manufacturer unknown) | Other | Matrix Distributors contacted by telephone those pharmacies to which they had shipped the affected lots. These call were conducted from 10/9 to 10/13/2006. A follow-up letter was issued to the pharmacies extending the recall to the user level. |  | 800 | 800 units | Product was distributed to 96 pharmacies throughout the United States. | Terminated | January 07, 2008 | 455 | 800 | Not_Computer
Z-0380-2007 | 36241 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | NicoletOne version 5.20 software (also known as NicoletOne Version 5.2 Software) used with the NicoletOne LTM, Sleep, nEEG, vEEG and ICU monitoring Systems. |  |  | August 25, 2006 | 2006 | Nicolet Biomedical Div of Viasys Healthcare | A software anomaly exists in the 5.20.1038 NicoletOne LTM, Sleep, nEEG, vEEG and ICU Monitor systems using the M, and/or C series amplifiers. If this anomaly occurs, the system will display the data for channel one in all channels on NicoletOne LTM, Sleep, nEEG, vEEG or ICU Monitor systems. | Other | An Advisory Notice, dated August 25, 2006, was sent to all affected customers. The notice described the issues, states to discontinue use of 5.20.1038 software until updated 5.21.1039c software is installed. |  | 75 | 75 | Worldwide, including USA, Belgium, Chile, Denmark, Italy, Japan, Poland, and Spain. | Terminated | December 10, 2011 | 1933 | 75 | Software | Output/Calculation | Software update | Software Update
Z-0500-2007 | 36925 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Atrium Express Dry Seal Chest Drain with Sterile Fluid Path Package |  |  | November 08, 2006 | 2006 | Atrium Medical Corporation | Chest drainage tubing incorrectly assembled and may disconnect from patient during use | Other | Atrium Medical notiifed accounts by fax to hospitals and phone/email to Distributors on 11/08/06. Product is requested to be returned. |  | 49 | 49 cases ( 294 units) | Nationwide, including CA, and PA. | Terminated | November 27, 2007 | 384 | 49 | Not_Computer
Z-0504-2007 | 36855 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Stille Sonesta Video Fluoroscopy Table 6210; manufactured by Stille AB, Solna, Sweden. |  |  | November 09, 2006 | 2006 | Stille-Sonesta, Inc | Bolts may back out of their joints causing the back section of the table to tilt. | Other | Notifications letters sent to consignees on 11/20/06. Users are instructed to check the bolts located under the table. The Notice includes instructions with pictures for checking the bolts and tightening them if necessary. Stille AB has developed a kit for replacing the bolt that is involved and will send a service engineer out to the site to do the replacement. |  | 66 | 66 units. | Nationwide. | Terminated | June 20, 2007 | 223 | 66 | Not_Computer
Z-0426-2007 | 37085 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Sara Active Sling - Standard; item KK53210.0 - model KA1210 (with safety); manufactured by Medibo Medical Products NV, Hanmont, Achel, Belgium, distributed by Arjo Inc., Roselle, IL 60172 |  |  | January 09, 2007 | 2007 | Arjo, Inc. | There is the potential for the slings to come apart, potentially allowing injury to the patient. | Other | Arjo sent Urgent Device Field Correction Recall Notice letters dated 1/9/07 via FedEx to the end user accounts who received the affected slings, advising them of the potential for the slings to come apart. They were instructed to remove the affected slings from use and replace them with the slings enclosed with the letter, complete the enclosed customer response form, and return the form and the original recalled slings, using the pre-paid call tag provided. If they have any questions or require additional, they were instructed to contact the Arjo Quality Department at 800-323-1245, ext. 6140. |  | 13 | 13 slings | Nationwide | Terminated | September 02, 2008 | 602 | 95 | Not_Computer
Z-0507-2007 | 37028 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The LightSpeed VCT System is intended for head and whole body X-ray Computed Tomography applications. |  |  | December 08, 2006 | 2006 | GE Healthcare | When the manual cradle release feature on the table is used to move a patient out of the gantry, a discrepancy of up to 6mm between the numerical display on the gantry and the actual position of the table could be introduced. | Other | A Product Safety Advisory Notice, dated December 8, 2006, was issued to all customers informing them of the safety issue and production correction. Customers are advised to operate the CT cradle as intended with the system''s electronic controls. If a critical need arises to release the cradle during a radiation therapy simulation scan, they are advised to independently verify patient position prior to marking the patient for therapy to avoid the potential inaccuracy. A software update will be installed to correct this problem by a GE representative. |  | 1256 | 1,256 units | Worldwide, including USA, Puerto Rico, Albania, Argentina, Austria, Australia, Belarus, Belgium, Bulgaria, Brazil, Canada, Chile, China, Columbia, Denmark, Finland, France, Germany, Great Britain, Greece, Guam, Hong Kong, Honduras, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia , Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Reunion Island, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain , Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Venezuela, Vietnam, and Yugoslavia. | Terminated | November 30, 2007 | 357 | 1278 | Software | Display/Image | Software update | Software Update
Z-0501-2007 | 37002 | No | NRP | Sirolimus test system. | Toxicology | Toxicology | 510(k) | IMx Sirolimus Reagent Pack; list 5C91-21; 100 test pack containing 1 bottle (9.7 mL) Anti-Sirolimus Antibody Coated Microparticles, 1 bottle (9.7 mL) Sirolimus Alkaline Phosphatase Conjugate, and 1 bottle (10 mL) 4-Methylumbelliferyl Phosphate, 1.2 mM; Produced by Axis-Shield Diagnostics, Ltd., Dundee, UK for Abbott Diagnostics Division, Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA |  |  | August 11, 2006 | 2006 | Abbott Laboratories MPG | The sirolimus values may shift higher after storage at 2-8ÂżC or after one freeze/thaw cycle of specimens. | Other | Abbott initiated a field correction recall by letter dated 8/11/06 to all customers receiving the affected reagent. The accounts were informed of the shift in sample values after storage at 2-8*C or with freeze thaw of specimens with the affected lots of reagents. The accounts were requested to review the revised sample handling information included in the attachment to the letter with their laboratory personnel and retain a copy of the letter and attachment for their records. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. |  | 2800 | 2,800 kits | Worldwide, including USA, Puerto Rico, Poland, Croatia, Latvia, Estonia, Serbia &amp; Montenegro, Japan, Turkey, Singapore, China, India, Taiwan, Thailand, South Africa, South Korea, Australia, New Zealand, Norway, Italy, Portugal, Spain, Greece, Panama, Hungary, Slovakia, Russia, Romania, Cyprus, Germany, Lithuania, Saudi Arabia, Jordan, Israel, Lebanon, Kuwait, Oman, Qatar, UAE, Egypt, Mexico, Guatemala, Venezuela, Colombia, Peru, Chile, Argentina, Brazil, El Salvador, Canada, Ireland, United Kingdom, Sweden, Netherlands, Belgium, France, Austria, Czech Republic, and Switzerland | Terminated | January 18, 2008 | 525 | 2800 | Not_Computer
Z-0509-2007 | 37180 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Candela Corporation Vbeam Aesthetica Laser System, Model/Catalog # 9914-08-0320 Vbeam Aesthetica w/ gray enclosure and gray display |  |  | January 15, 2007 | 2007 | Candela Laser Corporation | Internal component may overheat causing the possibility of risk of fire or smoke | Other | Candela Corporation notified US consignees by letter (w/return receipt) identifying the potential problem on 1/17/2007. Candela services will contact the customers and schedule upgrades. International packages were sent to each distributor via Fed-x/DHL tracking on 1/17/2007. Distributors were asked to forward individual serial number packages to their customers. Upgrade material will be forwarded to distributors upon which they will schedule upgrades with their customers. |  | 474 | 474 laser systems total (all models) | Worldwide including USA, Europe, Far East, Canada, Japan, Australia, Middle East, Mexico | Terminated | February 03, 2014 | 2576 | 474 | Not_Computer
Z-0427-2007 | 37113 | No | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | OÂżOptixÂż (lotrafilcon B) Soft Contact Lenses, Rx only, Ciba Vision. The product is distributed in 3 packs (3 lenses per package) and 6 packs (6 lenses per package). |  |  | January 11, 2007 | 2007 | Ciba Vision Corporation | Reduced Ion Permeability | Other | Consignees were notified by traceable mail on 01/12/2007. Consignees were advised to locate any of the affected lots and return them to CIBA Vision in accordance with the attached instructions. Returned lenses will be destroyed. |  | 11400000 | 11.4 million lenses | Worldwide, Including USA, Canada, Australia, Benelux, UK and Germany | Terminated | March 21, 2008 | 435 | 11400000 | Not_Computer
Z-0515-2007 | 37191 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC Uroview 2800 fluoroscopic x-ray system with 3 Phase Power Distribution Box option (Wonder Box) , Model Numbers: 882082-01, 882082-02, 886765-01, GE Healthcare, Surgery, Salt Lake City, UT. |  |  | January 15, 2007 | 2007 | GE OEC Medical Systems, Inc | Uroview 2800 system with 3 Phase Power Distribution Box (Wonder Box) option can present a shock hazard to the user if the access door to the power box is not completely closed when the unit is powered on. . | Other | Consignees were notified by letter on 01/15/2007. Instructions were given to position and not move the dual monitors until inspection and correction by GE personnel. For the affected units with a 3 phase power distribution box, users were cautioned to ensure that the access door to the power box is completely closed when the unit is powered on to prevent a shock hazard. |  | 277 | 277 units | Worldwide, including USA, Canada, China, France, Germany, Hong Kong, Italy, Japan, Singapore, South Africa, Switzerland, and Venezuela. | Terminated | December 09, 2011 | 1789 | 702 | Not_Computer
Z-0428-2007 | 37182 | No | GBW | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Taut Intraducer Peritoneal Catheter, 12 Fr x 10 Fr x 20.3 cm; a sterile, Rx peritoneal catheter for single use, individually packaged in a Tyvek/clear pouch; Taut Inc., 2561 Kaneville Ct., Geneva, IL 60134; the individually packaged, sterilirzed catheter was sold in two configurations: |  |  | January 19, 2007 | 2007 | Taut, Inc. | Taut, Inc. has become aware that some package seals are not intact, compromising the sterility of the catheters. | Other | Taut sent Urgent Medical Device Recall letters dated 1/19/07, via UPS over night, on the same date, informing them that the two lots of catheters may have package seals that are not intact. Customers were instructed to check their stocks for the affected lots, remove any found and return all affected product immediately. The PI-128 is also a component of the System One Comprehensive Lap CBDE Kist, catalog number 50000, which is a convenience kit. The customer is being instructed to remove the PI-128, Intraducer Peritoneal Catheter contained in the System Once Comprehensive Lap CBDE Kit in the eight kit lots listed, and return the PI-128 catheters. All other components of the kit may be used. As the PI-128, Intraducer Peritoneal Catheter, can also be ordered as a replacement component, the customer is also being instructed to check all opened System One Comprehensive lap CBDE kits, regardless of lot number, to ensure that it does not contain the impacted PI-128 lots. The customers were requested to complete and return the enclosed Acknowledgment &amp; Stock Status Form, which includes an option to obtain replacement product for the PI-128, immediately faxing it back to Linda Todd at 919-433-4986. Any questions were directed to Linda Todd at 1-800-334-9751, ext. 4951. |  | 5769 | 5,769 catheters | Worldwide, including USA, Guam, Australia, Iceland and Italy | Terminated | August 12, 2009 | 936 | 5769 | Not_Computer
Z-0502-2007 | 37192 | No | NCX | Thermal regulating system. | Cardiovascular | Neurology | 510(k) | Alsius Cool Line Catheter, Model CL-2295A |  |  | January 01, 2007 | 2007 | Alsius Corporation | Testing by Alsius has identified that 6-10% of catheters from the three affected lots may exhibit a significant reduction in their heat exchange capacity. | Other | Customer letters were sent starting January 1, 2006 via FEDEX and DHL. Verification of receipt of the notification letter to level A will be undertaken by using tracked delivery requiring signature upon receipt . The Customer letters will notify consignees of the nature of the problem and provide instructions to return the product. |  | 77 | 77 | Worldwide including USA, Austria, Canada, France, Greece, Japan, Norway, Spain &amp; Switzerland | Terminated | June 22, 2007 | 172 | 202 | Not_Computer
Z-0429-2007 | 37009 | No | MFK | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | AcrySofÂż ReSTORÂż Intraocular Lens |  |  | December 14, 2006 | 2006 | Alcon Laboratories, Inc | Intraocular lenses exposed to extreme temperatures while in storage. | Other | The recall of the fifteen re-shipped lenses was conducted by telephone calls to the nine customers that had received the lenses. These calls were made on December 14, 2006. |  | 4 | 4 units | Nationwide. | Terminated | April 09, 2008 | 482 | 15 | Not_Computer
Z-0517-2007 | 37263 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | The recalled product is the Lateral biopsy plate, part #2414346. The biopsy plate is labeled in part: Lateral Biopsy Plate, part 2414346, For Use with the Liberty 9000 8-Channel Breast Coil, USA Instruments, Inc., Aurora, OH, sterile. The breast biopsy plate is used with GE coils during Magnetic Resonance breast biopsy procedures. The lateral biopsy plate is packaged and sold as a part of catalogue E8800BG, M3087JG, G3087JH, M3335LA. |  |  | January 31, 2007 | 2007 | USA Instruments Incorporated | The packaging used for the Breast Biopsy plate may exhibit small holes or tears | Other | The recalling firm notified consignees via certified letter, dated 1/31/07. |  | 2200 | Approximately 2200 units for all products being recalled | Worldwide, including USA, Canada, Mexico, Argentina, Australia, Austria, Bhutan, Brazil, Chile, China, Denmark, Finland, Germany, Great Britain, Greece, Guatemala, Honduras, Hong Kong, Indonesia, Ireland, Italy, Japan, Korea, Malaysia, Netherlands, New Zealand, Panama, Puerto Rico, Singapore, South Africa, Spain, Switzerland, Taiwan, Thailand, and Turkey. | Terminated | April 13, 2012 | 1899 | 2200 | Not_Computer
Z-0505-2007 | 36866 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | LATERAL BIOPSY PLATE - For use with the Liberty 9000, 8-Channel Breast Coil. The Lateral Biopsy plate is sent to customers in the following configurations: 1) The plate is purchased in sets of 5 as part number E8800BG 8 Channel Breast Biopsy Lateral Grids (pack of 5 grids) or 2) the plate is purchased in pairs with the following part numbers: M3087JG, which contains Liberty 9000 8-Channel Breast Coil, 2 plates and a manual; G3087JH, which contains 2 plates and a manual; M3335LA, which contains a 3T HD 8 channel Vibrant Breast Array, 2 plates and a manual. |  |  | November 21, 2006 | 2006 | USA Instruments Incorporated | The packaging used for the Breast Biopsy Plate may exhibit small holes or tears. | Other | Customers were notifed by certified letter, dated 11/28/06. |  | 250 | 250 units | Nationwide. | Terminated | April 13, 2012 | 1970 | 250 | Not_Computer
Z-0498-2007 | 37215 | [6 empty] | Anaerotest for microbiology |  |  | August 17, 2006 | 2006 | EMD Chemicals Inc | Lot # OC553145 of Anaerotest failed to remain in specification. QC indicator shows an anaerobic environment even if oxygen is present. | Other | A Recall Letter was issued by EMD Chemicals on 8/17/2006 to discontinue product usage and for product return. |  | 8 | 8 rolls | Product was only sent to one US customer: Becton Dickinson, Cayey. Pureto Rico. | Terminated | June 06, 2008 | 659 | 40008 | Not_Computer
Z-0422-2007 | 36722 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | iE33 Ultrasound System (System, imaging, pulsed doppler, ultrasonic and system, imaging, pulsed echo, ultrasonic) |  |  | November 02, 2006 | 2006 | Philips Ultrasound, Inc. | Acoustic output intensity may exceed maximum specified limits and the patient contact temperature may exceed 43 degree C (when the system is used with the S12-4 transducer in CW-mode) | Other | On 11/2/06 the firm sent a letter dated October 31, 2006 to their customers. The letter explains the problem and states that new software is being released and will be available sometime in November. Customer service representatives will visit consignees to upgrade their units. The firm later determined they sent letters out to consignees that did not have affected units. Those consignees will be notified the week of 11/20/06 to disregard the previous letter. |  | 1325 | 619 domestic units, 706 international units | Worldwide | Terminated | July 14, 2009 | 985 | 1325 | Software | Output/Calculation | Software update | Software Update
Z-0423-2007 | 36974 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | OsSatura TCP (pure tricalcium phosphate) 3mm granules packaged in glass jars, Part Numbers: 05-6010-050 5cc, 05-6010-100 10cc, 05-6010-150 15cc &amp; 05-6010-300 30cc |  |  | August 31, 2006 | 2006 | Teknimed SA | The firm has determined an incorrect Instructions for Use (IFU) was packaged within the product box. | Other | Teknimed notified IsoTis OrthoBiologics via electronic recall letter dated August 31, 2006. IsoTis was instructed to notify their consignees of this field correction and they were asked to be responsible for identifying and making any corrections necessary at the medical user level. Teknimed supplied the correct IFU to IsoTis so that they could perform a field correction of the affected product. October 1, 2006 IsoTis OrthoBiologics sent their consignees a recall notification that included the IsoTis notification letter, a copy of the Teknimed recall letter, IsoTis Field Correction Instruction Sheet and replacement TCP IFU's for the affected lots. |  | 1681 | 1681 units | Nationwide | Terminated | April 23, 2012 | 2062 | 1681 | Not_Computer
Z-0525-2007 | 37120 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Skytron Infinity Series (also called Astro series) minor surgery light with dual lightheads, 3 bulbs in each of 2 - 19' diameter lightheads, Manufactured by Dai-ichi Shomei Co., Ltd., Made in Japan; Models IN1919 and IN1919EL and IN1919EL/PT. |  |  | January 08, 2007 | 2007 | Skytron, Div. The KMW Group, Inc | If the mounting plug is improperly adjusted, the unit could separate, come apart, and fall from the mount. | Other | The firm initially notified consignees of this problem on 11/23/1999 and again on 11/01/2002 and most lights have already been corrected. Remaining consignees with lights in need of correction were notified of the problem via letter dated 1/8/07 and instructed to contact the distributor to have their lights inspected and modified. |  | 5000 | 5,000 for all models. | Nationwide and Canada. | Terminated | November 18, 2009 | 1045 | 5000 | Not_Computer
Z-0520-2007 | 37213 | No | KNQ | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Quantum TTC Biliary Balloon Dilation Catheter, Wilson Cook Medical, GI Endoscopy, 4900 Bethania Station Road, Winston-Salem, NC 27105 |  |  | January 24, 2007 | 2007 | Cook Endoscopy | The graphic label is incorrect, it lists 6mm/ 18FR. The product label is correct, it lists 10mm/ 30FR. | Other | Consignees were notified by UPS certified mail and email on/about 01/31/2007. They were instructed to return any affected product to Cook Endoscopy. A response form was attached to be completed and returned to Cook Endoscopy. |  | 12 | 12 units | Worldwide, including NY and TX in USA, Australia, Canada and India. | Terminated | July 26, 2012 | 2010 | 12 | Not_Computer
Z-0548-2007 | 37231 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia S Series SPECT (Single-Photon Emission Computed Tomography) System with a Pinhole Collimator; Siemens Medical Solutions USA, Inc., Molecular Imaging, 2501 N. Barrington Road, Hoffman Estates, IL 60195-5203 USA; part number 08717741 |  |  | January 30, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Symbia systems, running on e.soft version 5.5 or earlier, could allow a fully-extended patient pallet to come into contact with the pinhole collimator, resulting in system damage. | Other | Siemens sent Customer Safety Advisory Notice/Urgent Field Correction Recall letters dated 1/30/07 to all users of Symbia systems with a pinhole collimator and e.soft software versions 5.5 or earlier. The accounts were informed of the potential for a fully-extended patient pallet to come into contact with the pinhole collimator, resulting in system damage. The customers were advised that a Siemens Customer Service Representative would be sent to their site to upgrade their system software to version syngeo MI Applications 2006A within 30 days. The upgrade will be done free of charge. |  | 16 | 16 units | Worldwide distribution --- including USA states of Alabama, California, Georgia, Indiana, Kentucky, Michigan, Ohio, Pennsylvania, Texas, Virginia and countries of Australia, Bolivia, Canada, China, France, Germany, Hong Kong and Spain. | Terminated | November 30, 2007 | 304 | 27 | Software | Device Operation | Software update | Software Update
Z-0554-2007 | 37349 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Cmax Surgical Table, Model No. 2182625, Manufactured by STERIS Corporation, 2720 Gunter Park East, Montgomery, AL 36109 USA, TEL: 334 277 6660, FAX: 334 271 5450 |  |  | February 06, 2007 | 2007 | Steris Corporation | A variation in suppliers welding processes, at the inner section of the hydraulic lift column, of the surgical table, may result in the table top becoming unstable, which could result in injury to patients or staff. | Other | A certified letter was sent to all affected consignees on February 6, 2007 notifying them of the issue and informing them of the recall and the replacement of the affected tables with new tables. |  | 46 | 46 tables | USA and Canada | Terminated | September 11, 2009 | 948 | 46 | Not_Computer
Z-0536-2007 | 36864 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microalbumin Multi-Calibrator Set, in vitro diagnostic. |  |  | September 06, 2006 | 2006 | Diagnostic Chemicals, Ltd. | Microalbumin Multi Calibrator mislabeled label on Level 5 and Level 6, cap label correct | Other | Diagnostic Chemicals Ltd, Canada notified customers by telephone and fax on 9/06/06 and advised to use the cap label to determine concentration of the calibrators. Accounts were asked to confirm and acknowledge receipt of the letter. |  | 38 | 38 kits | Nationwide. | Terminated | May 14, 2007 | 250 | 38 | Not_Computer
Z-0522-2007 | 37005 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Polyethylene tibial posterior stabilized components. Optetrak All Poly Tibial PS Component. Knee prosthesis. Catalog number : 204-11-13 |  |  | December 07, 2006 | 2006 | Exactech, Inc. | The outside profile on this manufactured lot of devices is oversized and does not meet specifications. | Other | Letter dated 12/08/06. The firm representative retrieved all four devices from the field. |  | 4 | 4 units | Distributed to one hospital in New York. | Terminated | July 11, 2012 | 2043 | 4 | Not_Computer
Z-0521-2007 | 37067 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Tom Tessier Osseous Microtome: Plug for Bone Mill Systems (Parts 01-15400 and 01-15401), Stryker Leiblinger Inc., Kalamazoo, MI; Part 01-15407. |  |  | December 22, 2006 | 2006 | Stryker Leibinger USA | Metal shavings come off of the plug as a result of friction between the plug and the cutting cylinder, and may be mixed with the milled bone chips. | Other | Consignees were notified via recall letter dated 12/22/06 which instructed them to quarantine the product and to contact the firm for further instructions. |  | 26 | 26 | Nationwide. | Terminated | June 04, 2008 | 530 | 26 | Not_Computer
Z-0542-2007 | 37014 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LifeShield Latex-Free Primary I.V. Pump Set with Distal Microbore Patient Line and Gravity Flow Prevention Valve, Convertible Pin, 110 Inch with High-Pressure Filter, Orange Polyethylene-Lined/Light Resistant Tubing and Option-Lok for use with Omni-Flow Medication Management System; Hospira, Inc., Lake Forest, IL 60045 USA, Made in Costa Rica; list 12163-01 |  |  | December 20, 2006 | 2006 | Hospira Inc. | The tubing can separate from the set at the filter inlet post. | Other | Hospira sent product recall letters dated 12/20/06 to the accounts via Federal Express overnight letter on the same date. The letters informed the accounts that the tubing can separate from the set at the filter inlet post due to an inadequate solvent bond. The accounts were requested to examine their inventory for the affected lot, quarantine any found and call Hospira Customer Care at 1-877-946-7747 to arrange for return of the product for replacement. |  | 7656 | 7,656 sets | Nationwide. | Terminated | March 24, 2008 | 460 | 7656 | Not_Computer
Z-0523-2007 | 37077 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Command Module with Masimo Sp02 Option, a multiparameter patient monitor. |  |  | January 03, 2007 | 2007 | Spacelabs Healthcare, Incorporated | Potential for modules to reset, alarms returning to default settings, invasive pressure losing its labels and zero calibration and Sp02 alarm limits being frozen. This can result in delay in treatment. | Other | On 1/3/07 the firm issued a letter to their customers advising of the issues and provided guidance to the consignee to enable the Check Setup feature that will provide an audible prompt whenever the monitor resets. The letter advises that consignees will be contacted to schedule an update, once new software has been validated. |  | 2694 | 1347 devices; 1067 domestic and 280 international | Nationwide and worldwide. | Terminated | February 15, 2011 | 1504 | 2694 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0546-2007 | 37252 | No | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | LifeSync ECG System - Disposable leadwear, Ls-222, Ls-223 |  |  | January 10, 2007 | 2007 | GMP Companies/Lifesync Corporation | When the V-lead is separated from the rest of the lead set, the laminate may tear, exposing the dielectric layer causing a break in the silver ink. If the product is used during a defibrillation after such a de-lamination, it could burn the skin and may compromise the efficacy of the defibrillation. | Other | The recall was initiated on January 15, 2007. The firm''s representatives, hand carried, Device Recall Letter to Account Representative and Risk Manager. Distributors who received the Device Recall Letter, purged their stock of the affected devices and notified their customers.Recall Communication was hand carried to each account with 100% compliance by LifeSync Clinical/Sales Personnel. EFFECTIVENESS CHECKS of this recall will be achieved through Return Material Authorization via telephone with each affected account. |  | 272 | 272 cases | Nationwide | Terminated | July 27, 2009 | 929 | 272 | Not_Computer
Z-0543-2007 | 37104 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Terumo L Series 1863 type arterial perfusion cannulae, straight open tip, 1/4'' connector, arterial cannula, 10 FR, 17.5 cm (7'') long; Catalog no. L7350. |  |  | December 07, 2006 | 2006 | Terumo Cardiovascular Systems Corp | The cannula connector is an incorrect size: It is actually 1/4" by 3/8" instead of the correct 1/4" by 1/4". | Other | Consignee was notified via telephone on 12/7/06 and returned the product. |  | 50 | 50 | Pennsylvania. | Terminated | December 05, 2007 | 363 | 60 | Not_Computer
Z-0537-2007 | 36988 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Biosense Webster PREFACE Guiding Sheath, Multipurpose Short, for the introduction of intravascular electrophysiology catheters into any cardiac chamber. Catalog Number 301803MS |  |  | November 17, 2006 | 2006 | Biosense Webster, Inc. | Biosense Webster, Inc. has received a number of complaints regarding the radiopaque tip separating or partially detaching from the PREFACE Sheath during use in the left atrium. | Other | Initially a Medical Device Correction Advisory letter was mailed on November 20, 2006 to all domestic and international consignees recommending to customers that the Preface Sheath was not to be used with the NaviStar ThermoCool Catheter or any 8 French or larger catheter.On December 9, 2006, a Recall Customer letter and Certification form were sent out to all customers and Johnson &amp; Johnson affiliates who had purchased the product in the past three years. For customers in the US, two letters and forms were sent out -one addressed to the EP/Cardiology Lab and the second to the Risk Management Office.For the US, customers will return product directly to Biosense Webster afler receiving a return authorization number. Returned product will be segregated and labeled as quarantined product. Returned product will be used as necessary for the root cause investigation. Any product not used at the conclusion of the investigation will be destroyed.For international customers, product will be returned to the Biosense Webster warehouse in Roden, Netherlands or to the Biosense Webster US warehouse (depending on order placement) if export issues don''t present a problem. If product cannot be easily returned to Biosense Webster, the customers will be instructed to destroy the product and provide notification that the destruction was performed. |  | 4565 | 4565 | Worldwide distribution, including USA, Australia, Brazil, Canada, China, Colombia, Hong Kong, Indonesia, Israel, India, Japan, Korea, the Netherlands, New Zealand, Pakistan, Singapore, Thailand, and Taiwan. | Terminated | June 25, 2009 | 951 | 32317 | Not_Computer
Z-0553-2007 | 37289 | No | FLD | Sorbent hemoperfusion system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Clark Biocompatible Hemoperfusion Cartridge, sorbent: Heparinized polymer over acitivated carbon, 250mL, Clark Research and Development, Inc. Folsom, LA 70437 |  |  | January 19, 2007 | 2007 | Clark Research and Development Inc | Lack of sterility assurance based on lack of sterility validation for labeled sterilization directions | Other | The firm notified its consignees of the problem and the recall via letter (priority mail with delivery confirmation)dated 01/19/2007 beginning on 01/19/2007. This letter requested that the consignee cease the use of this product and contact the recalling firm for return instructions. Follow-up with non-reponders will be made by facsimile. |  | 276 | 276 | Nationwide | Terminated | July 17, 2007 | 179 | 292 | Not_Computer
Z-0345-2007 | 36711 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Krystal Klear Rinse (China only) |  |  | September 15, 2006 | 2006 | Advanced Medical Optics, Inc. | Due to a potential sterility issue with the product container. Three lots sold in Japan were found to have bacterial contamination, which compromised sterility.Because of this production-line issue at the firm's plant in China, AMO is recalling lots that were manufactured on the same production lines during the same production period. | Other | Beginning on September 15, 2006, a product notification letter was sent via facsimile or communicated via telephone with each of the customer accounts. Beginning on September 19, 2006, an AMO Japan sales representative visited each of the customer accounts obtaining COMPLETE Amino Moist Multi-Purpose Solution subject to recall. Letters dated November 21, 2006 were sent from AMO, Santa Ana, California beginning Novemer 27, 2006. |  | 3125885 | 182,678 units for all US products, 2,943,207 units worldwide | Worldwide including USA, Australia, China, Hong Kong, India, Indonesia, Japan, Korea, Malaysia, New Zealand, Philippines, Singapore, Sri Lanka, Taiwan and Thailand. | Terminated | February 17, 2010 | 1251 | 3125885 | Not_Computer
Z-0535-2007 | 36605 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | FastCath Transseptal Catheter Introducers with Hemostasis Valve, DAIG division of St. Jude Medical, Minnetonka, MN, consists of a radiopaque sheath and dilator; Product # 406850. |  |  | October 05, 2006 | 2006 | St. Jude Medical / Daig Division | Limited number of FastCath Transseptal Catheter Introducers (REF #406850) are susceptible to the irrigation extension tube on the side port becoming disconnected when a pull force is applied. | Other | The affected customers were notified on 10/5/06 by letter. |  | 40 | 40 Units | Nationwide. | Terminated | May 25, 2007 | 232 | 40 | Not_Computer
Z-0560-2007 | 37153 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Alcon Custom-PakÂż, part #10975-02, containing BD BeaverÂż Mini-Blade (part #BD6400 or part #BD6900); individual blades repackaged into custom ophthalmic surgical packs and resterilized by Alcon Laboratories, Inc., Houston, TX. |  |  | December 22, 2006 | 2006 | Alcon Laboratories, Inc | Individually packaged surgical blades, labeled as sterile, found with incomplete seals; blades purchased for use in custom surgical packs. | Other | Customers notified of recall via phone on 12/22/06, with follow-up fax and letter. A revised notification was made via phone and fax to each customer on 12/26/06. Customers instructed not to use blades and to return blades for a replacement. Recall to the user/hospital level. |  | 4 | 4 custom surgical packs. | Custom surgical packs sold to hospitals in AR, MN, NY, TX and WI. | Terminated | June 06, 2007 | 166 | 412 | Not_Computer
Z-0545-2007 | 37160 | No | CDD | Vitamin B[bdi1][bdi2] test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Sample Management System software for in vitro diagnostic testing |  |  | January 15, 2007 | 2007 | Siemens Medical Solutions Diagnostics | Under limited circumstances, an incorrect patient result could be printed on the optional chartable patient report. | Other | A recall communication was sent on 1/15/2007 to all customers by fax. |  | 87 | 87 units | Worldwide, including USA, Australia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Cyprus, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Greece, Ireland, Israel, Italy, Japan, Mexico, Netherlands, Norway, Poland, Portugal, Russia, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, UAE and United Kingdom. | Terminated | January 18, 2008 | 368 | 87 | Software | Output/Calculation | Notification | Safety Notice/Insructions
Z-0565-2007 | 37234 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | CELL-DYN Diluent/Sheath Reagent, for use with CELL-DYN Ruby, Sapphire and 4000 systems. List #01H73-01. Packaged in 20 Liter cubitainers. |  |  | December 22, 2006 | 2006 | Abbott Laboratories | Results for tests of RBC, MCV, RDW, MPV and PLT may be falsely elevated or falsely decreased when the Cell Dyn Diluent/Sheath reagent is used with Cell-Dyn 4000, Ruby and Sapphire systems. | Other | Recall initated 12/22/2006. Product letters were sent to consignees, as well as a customer reply form. The letters were sent with customer reply forms. Responses will be tracked via the response forms included with letters. |  | 652 | 652 (20 L. each) cubitainers | Nationwide, including California and Nevada. An Abbott Laboratories facility in Hong Kong. | Terminated | November 28, 2007 | 341 | 652 | Not_Computer
Z-0495-2007 | 36980 | No | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb Lightning High Speed Vitrectomy Cutter, Catalog #CX4804, 1/20-gauge high speed vitrectomy cutter per package. The firm name on the label is Bausch &amp; Lomb incorporated, Rochester, NY. |  |  | December 14, 2006 | 2006 | Bausch &amp; Lomb Inc | The cutter tip could break during a procedure | Other | Recall letters were issued via e-mail to surgeons on 12/14/06 and via overnight mail on 12/14/06 to surgeons and operating room supervisors requesting immediate segregation and return of the product. |  | 8289 | Cat. CX4804 - 3,485 Cutter Tips | Worldwide distribution ---- including USA and countries of Canada, China, Japan, Australia, Hong Kong, India, Malaysia, New Zealand, Singapore, Korea, Taiwan, Thailand, Argentina, Brazil, Mexico, and The Netherlands. | Terminated | April 07, 2009 | 845 | 45288 | Not_Computer
Z-0547-2007 | 37307 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Basic/Profile and One Touch Ultra-Blood Glucose Test strips- 50-count |  |  | February 02, 2007 | 2007 | Safety &amp; Supplies Co. Inc. | Counterfeit product (manufacturer unknown) | Other | A recall notification was sent on February 1, 2007 via first class mail to all customers. |  |  | Unknown | Nationwide to pharmacies. | Terminated | June 08, 2007 | 126 |  | Not_Computer
Z-0582-2007 | 37248 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD Integra 1ml Insulin Syringe with Retracting Precision Guide Needle |  |  | January 27, 2007 | 2007 | Becton Dickinson &amp; Company | The needle becomes detached from the hub/syringe | Other | Recall notification letters were sent via UPS second day air on January 27, 2007. Distributors were asked to provide end user contact information to recalling firm. |  | 247824 | 247,824 units | Nationwide | Terminated | July 24, 2007 | 178 | 247824 | Not_Computer
Z-0566-2007 | 37287 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter LH700 Series Hematology Analyzers, Part Number 6605632 with 2B Software (2B3, 2B4 and 2B5) |  |  | September 12, 2006 | 2006 | Beckman Coulter Inc | There is a potential for erroneous yet credible results for Body Fluids on the LH700 series hematology analyzers when a cassette label fails to read and the subsequent sample is cycled in Body Fluid mode. | Other | A Product Corrective Action (PCA) letter was mailed on Sep 12, 2006 to all LH 750 accounts to inform them that there is an issue with COULTER LH 700 Series analyzer software that could affect the Body Fluid analysis and that this this anomaly is present only in Software revisions of 2B. The letter instructs the customer on how to determine their software version.Immediate user preventive action was provided in the letter which is to reset the analyzer when a cassette label read error occurs. The letter advises that this information be shared with laboratory staff and be retained as part of their Quality System documentation. A customer response form was also included. |  | 1987 | 1893 units in the US; 94 units to Canada | Nationwide and Canada | Terminated | April 19, 2012 | 2046 | 1987 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0574-2007 | 37170 | No | DSQ |  |  | Cardiovascular | PMA | AB5000 Console Circulatory Support System |  |  | January 08, 2007 | 2007 | Abiomed, Inc. | Unit may alarm ''Low Flow'' and ''Low Pressure'' due to manufacturing material in the pressure source. | Other | Abiomed notified customers by letter dated 1/08/07. Service reps will inspect units and replace the defective components. |  | 55 | 55 units | Worldwide, including USA, Japan, Germany, Australia, and Mexico. | Terminated | February 07, 2008 | 395 | 55 | Not_Computer
Z-0567-2007 | 37250 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Wright Medical Technology, GladiatorÂż Bipolar Hip Prosthesis, REF: GLBP-2241, Size 41mm O.D., use with 22 mm Head |  |  | January 03, 2007 | 2007 | Wright Medical Technology Inc | Product specification failure which could result in improper locking of the cup onto the femoral head. | Other | The firm initiated the recall by telephoning its distributors on January 3, 2007 to request the return of the products. They followed with letters (mailed 1/15/07) to hospitals, implanting surgeons, and distributors requesting return of the product and informing them of the revision surgery that was a result of the problem. |  | 10 | 10 units | Nationwide and Japan | Terminated | January 15, 2009 | 743 | 90 | Not_Computer
Z-0576-2007 | 37298 | No | GZX | Microsurgical instrument. | Neurology | Neurology | 510(K) Exempt | AccuGuide Injection Needles (Surface electrodes) |  |  | January 22, 2007 | 2007 | Medtronic Xomed, Inc. | A labeling discrepancy was internally identified for the surface electrode component of REF 8263210 AccuGuide EMG Injection Needle, 30G X 25MM. The outside box label correctly indicates 'Content: 5 sterile Injection Needles and 10 non-sterile Surface electrodes'; however, the pouches for the 10 surface electrodes inside the box are incorrectly labeled as 'sterile'. | Other | Each consignee contacted via telephone 01/22/2007, and then sent a follow-up fax letter 02/02/2007, which explained the issue, and the options for dispositioning any product that they still had in stock.Since firm has determined that the risk of keeping and using these devices is negligible to no risk, as they are intended to be and typically used in a non-sterile manner, the firm is offering the consignees the option to keep their devices or have them replaced. |  | 48 | 48 boxes | Nationwide | Terminated | July 07, 2012 | 1993 | 48 | Not_Computer
Z-0413-2007 | 37070 | No | JJI | Enzyme analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i System Assay CD-ROM, in vitro diagnostic analyzer, U.S. Version 21.0; for use on the ARCHITECT i System; Abbott Laboratories, Abbott Park, IL 60064; list number 06E58-21 |  |  | January 02, 2007 | 2007 | Abbott Laboratories MPG | The ARCHITECT AUSAB and ARCHITECT Anti-HCV assay parameter default interpretation screens when using ARCHITECT Assay CD-ROM US Version 06E58-21 does not align with the result interpretation options in the ARCHITECT AUSAB Reagent Package Insert (PI) 34-4162/R1 and ARCHITECT Anti-HCV PI 34-4152/R1. | Other | On 1/2/07, Abbott sent Product Correction, Immediate Action Required letters dated 12/29/06 to all ARCHITECT AUSAB (list number 1L82-25) and ARCHITECT Anti-HCV customers (list number 1L79-25) who have the affected assay CD-ROM. The end users were informed that for the ARCHITECT AUSAB and ARCHITECT Anti-HCV assays, the default interpretations in the Assay Parameters window, installed from ARCHITECT I System Assay CD-ROM US Version 06E58-21, are not align with the result interpretation options in the package inserts. The accounts were advised to reconfigure the names of the interpretations in Assay Parameters window, following the steps in the enclosed attachments A and B. The accounts were requested to keep a copy of the letter on file until the revised Assay CD-ROM is received. The accounts were requested to complete the customer reply form, indicating that the letter was received, understood and disseminated to the laboratory manager/supervisor responsible for Architect AUSAB and/or Anti-HCV testing, and fax it to Abbott by 1/12/07. |  | 51 | 51 affected analyzers | Nationwide | Terminated | December 09, 2009 | 1072 | 51 | Computer
Z-0583-2007 | 37079 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic 2T10R1 Intersept Custom Tubing Pack. Pack 2T10R1 consists of two PVC lines; one of them is 132 inch in length and is capped. The second line is 102 inch in length and has a one way vacuum relief valve assembled in one end. Tubing is used in the vent roller heads of the heart lung machine. This tubing is connected to the cannula in the patient in order to divert blood to the circuit when the patient is on bypass. |  |  | January 02, 2007 | 2007 | Medtronic Perfusion Systems | One lot of Custom Perfusion Sets, Catalog 2T10R1 was incorrectly assembled. The one way vacuum relief valve on the 102 inch line is attached backwards, thus preventing flow. | Other | There was only one customer. The customer was sent a letter on December 29, 2006, requesting that the impacted product be disposed of or returned to Medtronic for disposal. |  | 17 | 17 | Illinois | Terminated | February 24, 2008 | 418 | 17 | Not_Computer
Z-0577-2007 | 37314 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | EKOS EndoWave Infusion System and EKOS Lysus Infusion Systems - drug delivery catheters. EKOS Corporation Bothell, WA |  |  | February 09, 2007 | 2007 | Ekos Corp | Ekos Corporation received three reports that distal radiopaque marker bands on the EndoWave Drug Delivery Catheter became detached from the catheter during use and remained in the patient. | Other | On 2/9/07 the firm faxed and sent via certified return receipt an URGENT: Device Recall letter to their consignees. The letters requested to immediately examine stock and segregate all recalled products. In addition, the firm''s representatives will contact, via phone or visit, each of the consignees. Product is to be returned to the firm. |  | 446 | 446 distributed | Nationwide. | Terminated | May 07, 2008 | 453 | 446 | Not_Computer
Z-0596-2007 | 37369 | No | KDC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | JariTrak Table Clamp, 1 inch diameter, vertical post, Catalog Number 206-160. The JariTrak Table Clamp with vertical post (206-160) is part of the Table Mounted Hardware that is used with the JariTrak&#153; Retractor System. The JariTrak&#153; self-retaining retractor system consists of the Table Clamp with Vertical Post, Flexible or Horizontal Bar with Universal Joint, Oval or Segment Rings, Tilting Blade Clamps and various blades. The product is supplied non-sterile and may be packaged individually or as part of the JariTrak&#153; retractor set. |  |  | November 14, 2006 | 2006 | J. Jamner Surgical Instruments, Inc. | The JariTrak Table Clamp may not securely fasten to the bed rail of the OR table during set up of the retractor system. In the event that the clamp is not securely fastened to the bed rail, the system may move during surgery. The firm received several complaints from hospitals. | Other | On 11/14/06, an "URGENT-PRODUCT NOTIFICATION" letter (dated 11/14/06) and accompanying Fax Back Form were sent via Certified Mail Return Receipt Requested, to each customer that purchased or was supplied with the 206-160 Table Clamp. The letter instructs each surgical instrument coordinator (user) to review the inventory of 206-160 JariTrak" Table Clamps at the facility. The letter also provides instructions for the proper installation of the clamp to the bed rail. The customer is also asked to complete the Fax Back Form and fax it to back to Customer Service. J. Jamner Surgical Instruments, Inc. (J. Jamner) will replace any Table Clamps at the request of the customer. J. Jamner will account for all Fax Back replies and product returns. A second notification letter will be sent to any non-responding customers. Table Clamps returned from customers and from J. Jamner's inventory will be returned to the manufacturer, Koros USA Inc. or be destroyed by J. Jamner. |  | 30 | Domestic: 30 units | Nationwide | Terminated | January 10, 2008 | 422 | 30 | Not_Computer
Z-0580-2007 | 37394 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Defibtech Sentry Semi-Automatic External Defibrilator (AED) sold as 'Lifeline AED' brand -Model DDU-100A (version 2.002 and earlier) |  |  | February 17, 2007 | 2007 | Defibtech, LLC | Device Malfunction-The self-test software may allow a self-test to clear a previously detected low battery condition. If this occurs, the operator may be unaware of the low battery and the device may be "unable" to deliver a defibrillation shock, which could result in failure to resuscitate a patient. | Software design | Defibtech LLC initiated notification to customers/distributors on 2/17/07 by email. Hard copy letters followed on 2/22/07 with acknowledgement card. Users provided with instructions for 'Immediate Recommendations' until the unit''s software is upgraded. Distributors, their customers, and direct accounts instructed to locate affected units and upgrade these units in the field. Distributors, theircustomers, and direct accounts shall return to Defibtech documentation confirming unitsoftware upgrade was completed..Press was issued by the firm on March 6, 2007 |  | 22197 | 22,197 units | Worldwide distribution ---- USA and countries of Canada, Uruguay, Jordon, Turkey United Kingdom, Saudi Arabia, Sweden France, South Africa, Cyprus, Spain, Germany India, Japan, Israel, Australia, Norway, Philppines Singapore, Poland, Germany, Italy, Switzerland Netherlands, Indonesia, United Arab Emirates Malaysia, Austria, Ireland, Kingdom of Bahrain, and Korea. | Open, Classified |  |  | 36219 | Software | Alarm/Message | Software update | Software Update
Z-0578-2007 | 37389 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis MP, Angiographic x-ray system, model number 5904466 |  |  | January 31, 2007 | 2007 | Siemens Medical Solutions USA, Inc | C-arm gantry could rapidly descend without command during use. | Other | In response to this issue, a Siemens Service Representative will visit the affected consignees beginning 02/01/2007, to check the spindle and adjust the spindle, if necessary. |  | 62 | 62 units | The products were shipped to medical facilities nationwide. | Terminated | May 05, 2009 | 825 | 70 | Not_Computer
Z-0597-2007 | 37393 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | ThromCat Thrombectomy Catheter System. 7F 2.5 -7.0 mm 150 cm. Catalogue number 63000-01. It is a single-use, disposable device that is used to perform percutaneous maceration and removal of thrombus and restoration of blood flow. |  |  | February 07, 2007 | 2007 | Kensey Nash Corp | Face seal may wear excessively and cause particulate matter. | Other | On 2/7/07, a voice mail message was sent to all KNC Sales Representatives instructing them to contact affected accounts immediately. On 2/9/07, recall notices were faxed to all affected accounts that had affected product. The recall notices instructed hospitals to quarantine the affected lots and to contact their KNC Sales Representative to arrange for retrieval of the product. |  | 165 | 165 units | Nationwide. | Terminated | June 12, 2007 | 125 | 165 | Not_Computer
Z-0575-2007 | 37270 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | StarClose Vascular Closure System (Clip Applier &amp; Exchange System), |  |  | January 31, 2007 | 2007 | Abbott Vascular, Inc. | Premature release of the vessel locator wings, which stabilize the device prior to clip deployment, will result in no hemostasis. | Other | On 1/31/07, the firm initiated the recall and its notification was via letters explaining the reason for the recall and requesting the product be returned to its sales rep. |  | 97952 | 97,952 units | Worldwide distribution, including USA, Puerto Rico, and 410 other countries. There were no shipments to Canada and Mexico. | Terminated | February 26, 2008 | 391 | 97952 | Not_Computer
Z-0604-2007 | 37068 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Wallstent Endoscopic Biliary Endoprosthesis- 8mm x 40mm, UPN H965430100 |  |  | December 20, 2006 | 2006 | Boston Scientific Corporation | Guidewire restriction during use (excessive adhesive in the guidewire lumen) | Other | On December 20, 2006, Boston Scientific notified all US consignees by letter sent via Fed Ex Priority mail. Customers are instructed to remove affected product from inventory and return it to Boston Scientific. The notification also requested the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product in inventory. Boston Scientific International Regulatory Affairs was notified and is responsible the monitoring foreign recall. |  | 8511 | 8,511 for Recalls Z-0604-0609-2007 | Worldwide distribution --- USA and countries of Austria, Belgium, France, Czech Republic, Germany, great Britain, Greece,Ireland, Italy, Luxemburg, Netherlands, Norway, Poland, Portugal, Spain, Switzerland, and Sweden. | Terminated | December 12, 2007 | 357 | 8511 | Not_Computer
Z-0588-2007 | 37346 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | FRAZIER Surgical Suction Instrument, 8 Fr. with Control Vent and Obturator, REF 0033080. MANUFACTURED FOR CONMED CORPORATION 310 Broad Street, Utica, NY 13501. ASSEMBLED IN MEXICO. |  |  | November 10, 2006 | 2006 | Conmed Corporation | Firm was made aware of instances where the sterile barrier of the instruments had been compromised. | Packaging process control | On 11/10/06, ConMed Corp. sent "URGENT: DEVICE RECALL" letters (dated 11/10/06) to the US and foreign consignees by UPS overnight. The letters instructed the consignees to inspect their stocks of the instruments and that they return all devices. The letters also instructed the distributors to contact their customers that received the instruments and conduct sub-recalls. Additionally, when requested, ConMed notified the distributor's customers (conducted sub-recalls) by UPS overnight mail. This recall action is to the end user level. Accompanying each letter was an effectiveness check form to be completed by the consignee, and faxed back to ConMed Corp. The instruments are to be returned to ConMed Corp. in Utica, NY. For questions, the consignee can contact Ms. Nancy Crisino by phone at (315) 624-3078 or by fax at (315) 624-3089. |  | 711245 | 711,245 units | Worldwide | Terminated | July 31, 2012 | 2090 | 7401479 | Not_Computer
Z-0594-2007 | 37356 | No | GHH | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Clearview Simplify D-dimer, in vitro diagnostic for qualitative detection of D-dimer in human whole blood and plasma. |  |  | February 15, 2007 | 2007 | Inverness Medical Professional Diagnostics | Lot PX018A is exhibiting a decrease in sensitivity affecting the qualitative result around the cut off. | Other | Recall notification letters were sent to all customers on 2/14/07 by first class mail. |  | 382 | 382 kits | Worldwide, including USA, Puerto Rico, and Canada. | Terminated | May 08, 2007 | 82 | 382 | Not_Computer
Z-0595-2007 | 37363 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | Vysis LSI ATM/p53: D13S319 /CEP 12/13q34 DNA Probe Set; fluorescence in situ hybridization (FISH) analyte specific reagents, in vitro diagnostic; Abbott Molecular Inc., Des Plaines, IL 60018 USA; list 05J83-001; |  |  | December 22, 2006 | 2006 | Abbott Molecular | The LSI D13S319 SO/13q34 SA/CEP 12 SG probe of the two mixture DNA-Probe was contaminated with LSI ATM SG/p53 SO probe. This contamination would result in 10 FISH signals in each normal cell rather than the expected 6 FISH signals in each normal cell. | Other | Abbott Molecular telephoned and/or sent recall letters to the Abbott FISH accounts on 12/22/06, advising them that an unexpected signal pattern with the Vysis LSI ATM/p53: D13S319 /CEP 12/13q34 Probe was observed, confirmed and identified as a result of contamination in a subset of vials of D13S319 /CEP 12/13q34 from lot number 79500. This contamination results in 10 FISH signals in each normal cell (4 SO, 4 SG, 2 SA) rather than the expected 6 FISH signals in each normal cell (2 SO, 2 SG, 2 SA). The customers were advised that lot 80178 was also being included in the recall, and were requested to review the information carefully and assess the risk to their patient population; discontinue and destroy any unused product immediately; complete and return the enclosed reply form; review the information with laboratory personnel; and communicate the information to health care providers served in cases where the affected product were used and reported. Abbott Molecular provided replacement product for that product destroyed, upon receipt of the reply form. Any questions were directed to Abbott Molecular Technical Service at 800-553-7042, option 2. |  | 109 | 109 probe sets | Worldwide including USA, Austria, Canada, Czech Republic, Denmark, Finland, France, Germany, Italy, Netherlands, Norway, Spain, and the United Kingdom. | Terminated | April 17, 2008 | 482 | 109 | Not_Computer
Z-0614-2007 | 37404 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gyrus ACMI Tripolar Ultra Cutting Forceps: 10 mm X 32 cm-Reference Number: 006689-901 |  |  | February 19, 2007 | 2007 | Gyrus ACMI Corporation | Product sterility may be compromised due to failure of the tray seal packaging | Packaging process control | Gyrus ACMi issued recall letters on 2/19/07. Each account is asked to remove recalled product and return product . |  | 184 | 184 boxes (5 units/box) | Worldwide distribution ----USA and countries of Canada, The Netherlands, Japan, Portugal, Columbia, South Africa, Australia, and Spain. | Terminated | February 24, 2008 | 370 | 722 | Not_Computer
Z-0616-2007 | 37479 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The iSight PACS (picture archiving and communications system) a software package used with general purpose computer hardware to acquire, store, distribute, process, and display images/associated data throughout a clinical environment. |  |  | February 14, 2007 | 2007 | Stentor Inc A Phillips Medical Systems Co. | A defect may cause patient image orientation markers to be incorrectly labeled on Multi-Planar Reformation (MPR) images. Left and Right Anatomical orientation markers may be reversed. | Other | The firm has issued a notification letter to its consignees on 02/14/2007. This notification letter contains preventive measures. |  | 54 | 54 units | Worldwide Distribution ---- USA and countries of Australia, Netherlands, and Canada. | Terminated | December 03, 2007 | 292 | 54 | Software | Display/Image | Notification | Safety Notice/Insructions
Z-0610-2007 | 37147 | No | NGY | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | 19 Gauge Trocar Cannula Replacement Plug, Catalog Number 617.32; Manufactured by Alcon Grieshaber, Ltd., Schaffhauxen, Switzerland CH8203. |  |  | January 18, 2007 | 2007 | Alcon Research, Ltd | Product is misbranded; 19 gauge Schlera Plugs distributed in containers labeled as 19 gauge Trochar Canula Plugs. | Other | Recalled by letter on January 22, 2007. |  | 15 | 15 units. | Nationwide- Hospitals in MD and MN. | Terminated | September 24, 2010 | 1345 | 15 | Not_Computer
Z-0617-2007 | 36962 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Manual Drive Unit for Sarns (Terumo) Centrifugal Perfusion System; Catalog No. 164268. |  |  | December 12, 2006 | 2006 | Terumo Cardiovascular Systems Corp | Internal friction may cause the centrifugal manual drive to be inoperable when a disposable pump is attached. | Other | Consignees will be notified via letter dated 12/12/06, which provides instructions for identifying defective manual drives and to contact the firm if their unit is defective. |  | 2696 | 2696 | Woldwide, including USA, Canada, and Mexico | Terminated | June 19, 2008 | 555 | 2696 | Not_Computer
Z-0621-2007 | 37423 | No | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | Remel MacConkey Agar w/Sorbitol, growth medium, Catalog R01556, packaged 10/box. The firm name on the label is Remel, Lenexa, KS. |  |  | February 21, 2007 | 2007 | Remel, Inc. | False negative results for sorbitol fermenting bacteria such as E. Coli 0157. | Other | Recall letters dated 2/19/07 were issued on 2/21/07 via regular mail explaining the reason for recall and requesting all remaining inventory of the affected lot be discarded. A Product Inventory Checklist was enclosed to indicate the amount of inventory the customer discarded. This was to be returned via FAX or mail in the return envelope provided. |  | 694 | 694/10-pack boxes | Nationwide. | Terminated | March 06, 2008 | 379 | 694 | Not_Computer
Z-0620-2007 | 37362 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Vented Paclitaxel Set; a sterile fluid pathway with polyethylene lined tubing, non-DEHP pump segment, 0.22 micron downstream high pressure filter, injection site, male Luer lock adapter; for use with rigid nonvented solution containers on Colleague and Flo-Gard 6201 and 6301 series pumps; Baxter Healthcare, Deerfield, IL 60015; product code 2C7553 |  |  | February 16, 2007 | 2007 | Baxter Healthcare Corp. | Baxter has received reports of disconnections/leaks during use of vented Paclitaxel administration sets. | Other | Baxter sent the 2/16/07 Urgent Product Recall letter to affected customers, to the attention of the Director of Nursing with copies to the Materials Manager, via first class mail to advise them of the potential disconnections/leaks during use of the Vented Paclitaxel Administration Sets. The accounts were instructed to immediately discontinue use and segregate the listed affected lot numbers, and contact Baxter Center for Service at 1-888-229-0001 to arrange for the return and credit for the affected product. Any questions were directed to the Center for One Baxter at 1-800-422-9837. |  | 200309 | 200,309 units | Worldwide, including Switzerland, Australia, Canada, Hong Kong, India, Korea, Mexico, Malaysia, New Zealand, Panama, Phillipines, Singapore, Thailand and Taiwan. | Terminated | March 17, 2008 | 395 | 215886 | Not_Computer
Z-0627-2007 | 37422 | No | KMK | Intravascular catheter securement device. | General Hospital | General Hospital | 510(K) Exempt | StayFix Fixation Device for Percutaneous catheters, REF 680ME, (small) 5F - 14F, Merit Medical, Unomedical, Ltd., Great Britain. |  |  | February 15, 2007 | 2007 | Merit Medical Systems, Inc | Sterility of some units may be compromised due to damaged packaging. | Other | Consignees were notified by letter on 02/23/2007. They were instructed to stop use of and return all affected product to Merit. |  | 195001 | 195,001 units | Worldwide, including USA, France, Germany, and United Kingdom. | Terminated | November 09, 2011 | 1728 | 250808 | Not_Computer
Z-0631-2007 | 37442 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Vaxcel Implantable Vascular Access System - Titanium Mini-Port with PASV Valve and 8F Silicone Catheter, REF/Catalog No. 45-220, UPN/Material No. M001452200, |  |  | September 27, 2006 | 2006 | Boston Scientific Corp | Sterile barrier may have been compromised. Firm received one complaint which demonstrated that the inner and outer seals of the port tray packaging had been compromised. | Other | On 9/27/06, Boston Scientific Corp. sent Urgent Medical Device Recall - Immediate Action Required letters (dated 9/27/06) via Federal Express to the 84 end users (medical facilities). The letters informed the users of the recall of the products due to a compromised sterile barrier. The letters instructed the users to immediately remove the affected lots from inventory, segregate them, contact Boston Scientific at 1-800-862-1284 to obtain a Return Authorization (RGA) number, return the product in an appropriate shipping carton, place the label on the outside of the shipping box, write the RGA # on the outside of the shipping carton, and return the recalled product to Boston Scientific Corp. Distribution Center, Attn: QA Returns, 500 Commander Shea Blvd., Quincy, MA 02171. Accompanying each letter was a Reply Verification Tracking Form to be completed by the user, and mailed or faxed back to Boston Scientific. |  | 42 | Domestic - 42 | Nationwide | Terminated | November 26, 2007 | 425 | 236 | Not_Computer
Z-0637-2007 | 37102 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | COOK Triple Lumen Central Venous Catheter Tray with Cook SpectrumÂż GlideÂż Antimicrobial Catheter with EZ--PassÂż Hydrophilic Coating: Reorder Number C-UTLMY-701J-RSC-ABRM-HC-FST-A. |  |  | January 02, 2007 | 2007 | Cook, Inc. | These catheters do not include the proximal sidehole as required by the specification. | Other | Consignees were notified by phone on 1/2/07 and instructed to return the product. |  | 100 | 100 | Nationwide, including the states of Florida, Louisiana, Maine, Massachusetts, Minnesota, Missouri, New Jersey and North Carolina. | Terminated | December 05, 2007 | 337 | 100 | Not_Computer
Z-0612-2007 | 37272 | No | MWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Guidant Acrobat Vacuum Stabilizer, Model number-OM-9100 |  |  | December 01, 2006 | 2006 | Guidant Cardiac Surgery | Vacuum Tubing Failure- vacuum tubing fails before the labeled two year shelf life. This may result in partial or complete vacuum loss during surgery. | Other | The firm has issued by FedEx letters dated 12/1/2006, to all consignees who received the product in the US, and has also informed their foreign distributors in the same manner on December 4, 2006. The firm''s field representatives were notified via teleconference on December 1, 2006. A hard copy of the notice was delivered to field reps on December 4 as well. |  | 57691 | 57,691 units distributed total | Product was distributed to a total of 529 consignees, 482 domestic and 47 international. Product was sent Nationwide, Europe and Asia. | Terminated | December 04, 2007 | 368 | 57691 | Not_Computer
Z-0639-2007 | 37450 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | CELL-DYN 4000 Diluent/Sheath Reagent, |  |  | May 12, 2004 | 2004 | Abbott Laboratories | Incorrect Test Results-test results may show a higher than expected platelet background count when used on the CELL-DYN 4000 System and report patient results that are unacceptable (out-of-range). | Other | On May 12, 2004, the firm initiated the recall and its notification was via letters explaining the reason for the recall and requesting the product be disposed. |  | 6893 | 6,893 units | Class II- Worldwide Distribution-USA including the states of AL, CA, CO, CT, DE, FL, GA, IL, IN, KY, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NY, OR, PA, SC, TN, TX, VA, WA, WI, and WV and countries of Argentina, Brazil, Hong Kong, Japan, Malaysia, Singapore, Taiwan, and Thailand. | Terminated | May 01, 2007 | 1084 | 6893 | Not_Computer
Z-0636-2007 | 36685 | [6 empty] | Brilliance CT System, Big Bore Configuration, Model #728243. (version 2.2.1) |  |  | October 02, 2006 | 2006 | Philips Medical Systems (Cleveland) Inc | Two conditions have been identified with the Philips Medical System Brilliance Bore ,version 2.2.1 Tumor Localization, that may lead the user to misinterpret the displayed data. | Other | Customer advisory notification letters dated 10/4/06 and 10/25/06 were sent to consignees providing short term work around instructions. Software modifications installed by the recalling firm will follow as the long term correction. |  | 17 | 17 units | Class II Recall-Nationwide Distribution-USA including states of KY, MD, TX, RI, NJ, MA, LA, VA, PA, SC, CA, OH, and FL. | Terminated | March 26, 2012 | 2002 | 17 | Software | Output/Calculation | Software update | Software Update
Z-0640-2007 | 37444 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | CELL-DYN 4000 Diluent/Sheath Reagent, List Number 01H73-01, Distributed and Manufactured by: |  |  | February 24, 2004 | 2004 | Abbott Laboratories | Incorrect Test Results-product may show a higher than expected platelet background count when used on the CELL-DYN 4000 System and report patient results that are unacceptable (out-of-range). | Other | Recall Letter-On 2/24/04, the firm initiated the recall and its notification explaining the reason for the recall and requesting the product be destroyed |  | 9800 | 9,800 units | Worldwide Distribution-USA including states of AL, AR, AZ, CA, CO, CT, DC, FL, GA, lA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MS, NC, ND, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WI, and WV. and countries of Australia, Argentina, Brazil, Hong Kong, Japan, Korea, Malaysia, Singapore. | Terminated | May 01, 2007 | 1162 | 9800 | Not_Computer
Z-0632-2007 | 36872 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Flo-Gard 6201 Volumetric Infusion Pump, catalog #2M8063, 2M8063R (refurbished); Baxter Healthcare Corporation, Deerfield, IL 60015 |  |  | November 27, 2006 | 2006 | Baxter Healthcare Corp. | There is the potential for non-detection of, and no alarm for, upstream occlusion for Flo-Gard Volumetric Infusion Pumps, if an occlusion occurs above a flexible chamber such as a drip chamber, Buretrol or blood filter. | Other | Baxter sent the 11/27/06 Urgent Device Correction letter to all Flo-Gard Volumetric Infusion Pump customers, to the attention of the Vice President of Nursing and the Biomedical Engineer, via first class mail to alert them to the potential for non-detection of, and no alarm for, upstream occlusion under specific conditions for Flo-Gard Volumetric Infusion Pumps if the occlusion occurs above a flexible chamber (such as a drip chamber, Buretrol or blood filter), and to provide the addendum to the operator manual to address these specific conditions. The letters included the mitigations and risks associated with each of the conditions that could potentially cause this situation. This information was included in the letter, as well as the Operator''s Manual Addendum, which Baxter requested to attach a copy to all Flo-Gard Operator''s Manuals in their facility. Baxter instructed the customers to share this information with all staff members, services and facilities that use Flo-Gard pumps. Any questions were directed to Baxter''s Global Technical Services at 1-800-843-7867, prompt 2. Although Baxter previously notified customers on 3/5/99 that they would no longer support or service the Flo-Gard 6200 and 6300 devices as of 12/1/99, these product codes were included in this letter due to a similar potential for this issue. Should the customer still have these 6200 or 6300 devices in their possession, Baxter recommended that they be taken out of service due to obsolescence of the appropriate components. The accounts were requested complete the enclosed reply form acknowledging receipt of the letter and dissemination of the information to their staff. |  | 177468 | 177,468 pumps | Worldwide including USA, Puerto Rico, Argentina, Australia, Bahrain, Belgium, Bermuda, Brazil, Canada, Chile, China, Colombia, Costa Rica, Denmark, Ecuador, Egypt, El Salvador, Guatemala, Hong Kong, Honduras, India, Indonesia, Ireland, Italy, Korea, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Peru, Philippines, Poland, Russia, Singapore, Slovakia, Slovenia, South Africa, Sweden, Taiwan, Thailand, Turkey, United Kingdom and Venezuela. | Terminated | January 19, 2010 | 1149 | 219566 | Not_Computer
Z-0497-2007 | 37150 | No | LFD |  |  | Dental | 510(k) | Salivart Oral Moisturizer, packaged in 2.5 fluid ounce (73.9 mL) spray cans, Part Number: 0386-0009-75; Manufactured by Gebauer Company, Cleveland, Ohio |  |  | January 19, 2007 | 2007 | Gebauer Company | Microbial Contamination-certain lots of product failed USP &lt;61&gt; Microbial Limits Testing for total aerobic count during 6 month stability testing. | Other | The recalling firm sent recall letters dated 1/23/2007 to their Distributor Customers and Non-distributor customers who received Salivart product from Lot Code: 06AA003. both letters include an attached recall response form which the customers are instructed to destroy the product, complete a form and fax back it to the recalling firm. The response forms provide verification that the form has been received by the customer(s). With the initiation of the RECALL EXPANSION for more lot codes to be recalled on 2/14/2007, the firm issued new recall letters dated, 2/13/2007, 2/15/2007 and 2/22/2007. These four letters provide the same basic information and also have attached recall response cards, however, they list the additional 6 lot codes subject to recall. The firm issued a Press Release concerning this recall through Cleveland (Business Wire) on 3/1/2007. |  | 34545 | 34,545 cans | The product was distributed throughout the U.S. to wholesalers, retail accounts (i.e. pharmacies, dental offices, and medical facilities) and full size samples of product were given away at trade shows to medical, and dental offices and some end users consumers). A total of 240 units of product derived from Lot Code: 06AA001, were distributed to a single Canadian distributor under the assigned Lot Code: 06BB002 | Terminated | December 17, 2007 | 332 | 34545 | Not_Computer
Z-0638-2007 | 37318 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | VIASYS VELA Ventilator, Catalog Number-16186; Viasys Respiratory Care, Inc. |  |  | January 31, 2007 | 2007 | Viasys Respiratory Care, Inc.dba Bird Products | The graphical user interface may become slow in its response and in some cases non-responsive. Under some circumstances this situation may cause the ventilator to delay the annunciation of alarms. | Other | On January 31, 2007, a product notification letter was mailed to all customers in the United States via overnight carrier (Federal Express). Consignees are informed in the notification letter that service calls are in process of being scheduled to update domestic consignees. On January 31, 2007, a product notification kit (including a product notification letter, download software, download instruction sheets and installation verification fax-back sheets) were mailed to all international customers via register mail. Consignees are informed in the notification letter that replacement software, replacement instructions, and installation verification forms are included in their recall package. Domestic Consignees will have product corrected by VIASYS service personnel. Self address postage paid envelopes will be provided with notifications to aid in effectiveness checks |  | 878 | 878 | Worldwide Distribution ---USA and countries of Bangladesh, Bosnia, Brazil, China, Czech Republic, Ecuador, Estonia, Finland, France, Guatemala, Hong Kong, India, Israel, Italy, Japan, Latvia, Lithuania, Malaysia, Morocco, Netherlands, Portugal, Russian Federation, Taiwan, Thailand, Turkey &amp; Vietnam. | Terminated | September 18, 2008 | 596 | 878 | Software | Device Operation | Software update | Software Update
Z-0642-2007 | 36979 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Medi-Tech(R) Katzen (TM) Infusion Wire, for the infusion of therapeutic agents (i.e. heparin, saline, thrombolytic agents, etc.) in the peripheral vasculature. The wire allows for the delivery of agents in either a 'pulse-spray' or 'slow weep' technique. Order No. REF: 46-193 (M001461930). |  |  | December 12, 2006 | 2006 | Boston Scientific | Boston Scientific is voluntarily recalling one lot/batch of Katzen Infusion Wires because they have identified that the label on the carton may indicate a different length device than what is actually in the carton. The affected batch was manufactured using the Katzen Core Assembly of 146 cm, instead of a 177cm assembly. | Other | An Urgent Medical Device Recall letter, dated December 12, 2006 was sent to affected hospitals Risk Manager beginning 12/12/2006. The letter describes the issue and product affected, requests that further distribution or use of any remaining product cease immediately, and provides information for returning product to Boston Scientific Replacement product will be issued for all recalled product returned from affected facilities. A Reply Verification Tracking Form is requested to be completed and returned to Boston Scientific by Wednesday, December 27, 2006. |  | 10 | 10 units | US only to: FL, GA, IN, NY. | Terminated | July 20, 2007 | 220 | 10 | Not_Computer
Z-0626-2007 | 37106 | No | IRT | Powered heating pad. | Physical Medicine | Physical Medicine | 510(K) Exempt | HoMedics Thera P Model King size Moist/Dry Heating Pad with Auto Shut-off; Model No. HP-500 (UPC 031262003652). |  |  | February 09, 2007 | 2007 | HoMedics, Inc. | Heating pads may have a loose electrical connection, which could result in excessive heat being generated that poses a risk of patient burns, fire, and property damage. | Other | Press release issued on 2/9/07 instructing consumers to return the product to store of purchase or Homedics. Retail stores notified via recall letter issued on 2/9/07. Retail stores are being instructed to accept returns and to post placards to notify customers of the recall. |  | 292108 | 292,108 of all models | Nationwide and Canada. | Terminated | January 04, 2010 | 1060 | 292108 | Not_Computer
Z-0644-2007 | 37141 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/VIEW C.30a Software update (P/N 514515-1) diagnostic data management system, bioMerieux, Inc., P. O. Box 15969, Durham, NC 27704-0969 USA |  |  | December 19, 2006 | 2006 | bioMerieux, Inc. | Bottle data is not sent to BacT/ALERT 3D instrument. | Other | A customer notification advising customers of the issues and the precautions were sent on/about December 20, 2006. |  | 1233 | 1233 disks | Worldwide, including USA, Canada, UK, Italy, France, Spain, Germany, Poland, Chile, Japan and the Netherlands. | Terminated | December 08, 2009 | 1085 | 1233 | Not_Computer
Z-0643-2007 | 37133 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERTÂż FA Culture Bottles, 30 ml, for in vitro diagnostic use, BioMerieux, Inc., Durham, NC 27704 |  |  | January 11, 2007 | 2007 | bioMerieux, Inc. | Bacillus sp. contamination was detected in inoculated and uninoculated bottles of BacT/ALERT FA Culture bottles. | Other | Consignees were notified by letter on/about 01/12/2007. |  | 2136 | 2136 cases (100 bottles per case) | Worldwide, including USA, Austria, Denmark, Finland, France, Greece, Guyana, Italy, Libya, Netherlands, Norway, New Caledonia, Poland, Portugal, Singapore, Slovenia, Spain, Tanzania, and Sweden. | Terminated | December 08, 2009 | 1062 | 2136 | Not_Computer
Z-0556-2007 | 37322 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9800 Fluoroscopic X-ray System, Model Number A349855, GE Healthcare, Surgery. |  |  | February 09, 2007 | 2007 | GE OEC Medical Systems, Inc | During routine service, a cover may have been installed without the required proper lead shielding. | Other | Consignees were notified by telephone on 02/09/2007 and told to discontinue use of the units until they could be inspected by GE personnel. GE personnel visited all 8 sites on 02/12/2007. |  | 11 | 11 units of all models | Nationwide to AL, AZ, IA, OH, TX, VA, and WV. | Terminated | June 26, 2007 | 137 | 11 | Not_Computer
Z-0646-2007 | 37473 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Thyrotest TSH POC Rapid membrane test. In vitro diagnostic. |  |  | January 24, 2007 | 2007 | ThyroTec Inc | Positive control showed little or no positive line. | Other | On 1/24/07 the recalling firm notified their only customer of the problem. On 2/5/07, the distributor issued recall letters to their customers informing of the problem and the need to destroy the kits. The kits from lot number 02106-06 were returned to the recalling firm. |  | 484 | 484 units | Nationwide. | Terminated | June 04, 2007 | 131 | 484 | Not_Computer
Z-0645-2007 | 37390 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur CP System, Catalog/Part Number 086-A001. Automated Immunoassay Analyzer. --- Classification of the Device: Under 21 CFR, Part 862.2160, the ADVIA Centaur CP is a discrete photometric chemistry analyzer for clinical use and is classified as a Class I (general control) device. --- The ADVIA Centaur CP is part of Bayer's Immunochemistry analyzer family, which also includes the ADVIA Centaur and the ACS:180. Based on FDA guidance for deciding when regulatory filings are warranted for device modifications, as well as, the FDA Instrument Family and Replacement Reagent Policy Guidance, Bayer determined that the changes associated with the ADVIA Centaur CP did not affect the safety and effectiveness of the predicate device (ACS:180) and therefore no device modification 510(k) was warranted. The ADVIA Centaur CP instrument uses the same reagents as the ADVIA Centaur. The ACS:180 was cleared by FDA under 510(k) # K902336, and the ADVIA Centaur was cleared by FDA under 510(k) # K971418. The ADVIA Centaur CP system was released for sale in November 2005 and distribution of these systems began at that time. All ADVIA Centaur CP Systems are impacted. There is no expiration date for these systems. |  |  | November 09, 2006 | 2006 | Siemens Healthcare Diagnostics | Firm received complaints for signal 2 &amp; 4 errors, and shifts in Relative Light Units (RLU). Investigations showed that the On-Board Stability (OBS) of the acid &amp; base reagents may be compromised after 4 days after installation onto the Centaur CP system due to evaporation. | Other | A Support Bulletin was sent by e-mail to all affected Bayer HealthCare LLC Branches on 11/09/06 in order to inform them of the issue and to provide them with instructions for this field correction. Two Customer Bulletins were prepared and sent by e-mail to Bayer HealthCare LLC Branches on 11/09/06 both in the US and outside the US for communication with the affected customers. The first bulletin advised customers of the problem and the corrective action. A Confirmation (Fax-back) Form was included to ensure that all customers worldwide have been notified and have complied with the mandatory action. The second Customer Bulletin describes an optional alternative mitigation, which requires a Bayer representative to go to the customer site to implement. |  | 612 | 306 units: 101- Domestic, 205 - Foreign | Worldwide | Terminated | July 30, 2012 | 2090 | 612 | Not_Computer
Z-0641-2007 | 37477 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart FR2+ Defibrillator |  |  | February 28, 2007 | 2007 | Philips Medical Systems | Failure of tactile switches on one of two printed circuit board assemblies (PCBA) used in the FR2+ device. Potential delay or prevention of shock treatment due to PCBA board contamination, causing intermittent switch operation. | Nonconforming Material/Component | On 02/28/2007 the firm sent letters via return receipt to all domestic and foreign customers. The letter advises the customer of the issue and if the user encounters the problem, to press the button again with additional force. In addition, the firm sent with the letter a new Quick Reference Card. Customers are instructed to replace the old Quick Reference Card with the new one which instructs them to press 'On' and 'Shock' buttons firmly. |  | 9099 | 9099 units | Distribution is worldwide. Devices are used by first responders (EMT, EMS, Fire and Police personnel, etc.), and by trained personnel on commercial airlines and business offices. | Terminated | November 13, 2013 | 2450 | 9099 | Hardware | Treatment/Delivery/Therapy | Update manual | Other
Z-0655-2007 | 37366 | No | DTW | Cardiopulmonary bypass level sensing monitor and/or control. | Cardiovascular | Cardiovascular | 510(K) Exempt | Terumo Perfusion System 8000 Base, 4 pump, 115v; Model 8000; Catalog No. 16400. |  |  | January 12, 2007 | 2007 | Terumo Cardiovascular Systems Corp | The hand crank may crack in the bearing/shaft area due to a manufacturing defect. (Note: Some of the hand cranks subject to this recall were shipped as replacements for hand cranks recalled in the Summer of 2006.) | Other | Consignees were notified by phone between 1/12/07 and 2/23/07 and informed replacement hand cranks would be shipped. |  | 7 | 7 | Worldwide, including USA, Belgium, Hong Kong, Japan, Taiwan and United Arab Emirates. | Terminated | June 19, 2008 | 524 | 83 | Not_Computer
Z-0654-2007 | 37376 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Colorado 2 Spinal System, Anterior Plate, REF8619500, for export only, manufactured at: Warsaw, Indiana USA, Rx only |  |  | February 02, 2007 | 2007 | Medtronic Sofamor Danek USA Inc | Bone screws may pull though the plate. | Other | Medtronic Global Quality initiated the recall via letter on 02/02/2007. Sub-recalls will be managed by Medtronic EMEA quality (Europe, Middle East, and Africa) in the affected geographies. Distributors and surgical facilities will be contacted as well as regulatory bodies in each country. |  | 293 | 293 units | France, Germany, Italy, UK, and Tunisia. No US distribution. | Terminated | August 11, 2007 | 190 | 1700 | Not_Computer
Z-0650-2007 | 37069 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen Complete Knee Solution Minimally Invasive Solutions MIS Technology Femoral Finishing Guide, instrumentation used in total knee arthroplasty procedures. Catalog No. 82015272521. |  |  | October 10, 2006 | 2006 | Zimmer Inc. | The device was cut to an incorrect angle, which may result in an incorrect bone cut. | Other | Product was removed from customer location during visit on 10/10/2006. |  | 1 | 1 | One medical facility in Michigan. | Terminated | December 04, 2007 | 420 | 1 | Not_Computer
Z-0660-2007 | 37436 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 8mm EndoWrist Bipolar Maryland Instrument, Part Number: 400172-04, |  |  | March 24, 2005 | 2005 | Intuitive Surgical, Inc. | The products were incorrectly programmed as training instruments, which allows for 30 uses (surgical procedures) instead of 10 uses. | Other | The firm initiated its recall in March 24, 2005, when the firm sent a letter via Federal Express mail to all 18 consignees, informing them of the issue, how to identify affected product, and instructions to return the affected instruments for a full credit. The form included a reply form that consignees were requested to fill out and return. |  | 8 | 8 | Product was distributed to 12 consignees in the US, who received 19 instruments, in FL, IN, LA, MN, NC, NH, NJ, NY, TN and TX. In addition, 5 consignees outside the US received product as well, in Italy, Sweden, France, Singapore and Germany. | Terminated | November 28, 2007 | 979 | 8 | Other | Device Operation | Return | Remove or Replace
Z-0652-2007 | 37347 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo Dynamics 5.0 Workplace. Medical image report and archive system. |  |  | January 31, 2007 | 2007 | Siemens Medical Solutions, USA, Inc | The cardiac calculation feature may incorrectly calculate a derived V max value (reported out lower than it should be) and certain values imported into worksheets and reports may be incorrect due to a software defect. | Other | Distributors were issued a recall letter dated 1/29/07 and asked to notify their customers by issuing their subaccounts a similar letter. When available, new software will be provided. |  | 71 | 71 | Nationwide. | Terminated | April 30, 2008 | 455 | 71 | Software | Output/Calculation | Software update | Software Update
Z-0667-2007 | 36947 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Disposable components of the superDimension Bronchus System: Locatable Guide Kit (AKI00017-01) and Extended Working Channel (AKI00041-01). |  |  | November 29, 2006 | 2006 | SuperDimension, Inc. | The action is being taken in response to certain incidents that have been reported to the company, one of which concerned an adverse event in which a patient incurred a pneumothorax during a procedure in which a superDimension Bronchus System was used. The biopsy results for the patient were taken from a different area than the confirmed navigation site that was indicated by the system. | Other | An Urgent Medical Device Field Action letter, dated 11/30/2006, was sent to medical facilities who have the superDimension Bronchus System installed. The letter infoms the facility to immediately stop using the Bronchus System, and move all inventory of the disposable system components to a controlled quarantine area. A superDimension representative will contact the consignees to schedule a time to retrieve the quarantined inventory, at which time a verification form will be completed. |  | 2000 | Up to 2,000 each Locatable Guide Kit and Extended Working Channel | Worldwide, including USA, Austria, Belgium, Canada, France, Germany, Hungary, Israel, Italy, Netherlands, South Africa, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | November 20, 2008 | 722 | 2074 | Not_Computer
Z-0675-2007 | 37365 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Trabecular Metal Reverse Shoulder System Instrumentation Conical Reamer, instrument for shoulder arthroplasty, REF no. 00-4309-021-00. |  |  | February 01, 2007 | 2007 | Zimmer Inc. | The reamer may lock up on the threads of the reamer body rather than spin freely as intended by design. | Other | Zimmer sales staff was notified by letter dated 2/1/07 to pull any of the product from stock and to provide the name and address of their hospital customer accounts for further notification by Zimmer. |  | 4 | 4 | California, Ohio and Texas. | Terminated | May 04, 2007 | 92 | 4 | Not_Computer
Z-0662-2007 | 37232 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | RF Denervation Probe, 5cm, temperature sensing electrodes designed for use in radiofrequency lesion procedures for the treatment of pain. |  |  | January 19, 2007 | 2007 | Smith &amp; Nephew, Inc. Endoscopy Division | Product is non-sterile but labeled incorrectly as sterile. | Other | Smith &amp; Nephew notified accounts by letter sent via Federal Express January 19, 2007 |  | 56 | 56 units | Worldwide, including USA, Denmark, France, UK, Germany, Korea, Switzerland, Italy, Spain, and South Africa. | Terminated | January 18, 2008 | 364 | 497 | Not_Computer
Z-0668-2007 | 37384 | No | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Matrix2 Detachable Coils, vascular embolization device. |  |  | February 07, 2007 | 2007 | Boston Scientific Target | The product may be missing the temperature indicator label on the inner pouch. | Other | The firm initiated the recall on 02/07/2007, and initially notified consignees via letter. |  | 57 | 57 units | Worldwide including USA, Argentina, Australia, Brazil, Czech Republic, Germany, Spain, Finland, France, Great Britain, Greece, Ireland, Netherlands, and New Zealand. | Terminated | November 27, 2007 | 293 | 57 | Not_Computer
Z-0661-2007 | 37452 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | MicroMaxx LAP/12-5 MHz Transducer, Laparoscopic Ultrasound Transducer |  |  | February 05, 2007 | 2007 | Sonosite, Inc. | Laparoscopic transducer exceeds the limits specified for radio-frequency emissions by up to 10db. | Other | On 2/5/2007 firm sent by FED-Ex notification letters to firm representatives. Representatives began hand delivery of recall letter to consignees and retrieval of LAP transducer for return delivery to firm. The letter advised customers that the LAP transducer exceeded the limits specified for radio-frequency emissions. The firm also advised that exceeding the limit may cause the LAP transducer to interfere with surrounding equipment that may be susceptible to radiated emissions and cause surrounding equipment to not perform as expected. |  | 6 | 6 Devices | Distribution to USA (Tennessee &amp; Ohio), France, Korea, and Germany. | Terminated | June 01, 2007 | 116 | 6 | Not_Computer
Z-0670-2007 | 37431 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive brand Da Vinci S Surgical System, an Endoscopic Instrument Control System, Model SS2000 Patient Side Cart, Model PS2000, Part Number: 380267-01, IS2000, Top Level, 4-Arm; Distributed and Manufactured by: Intuitive Surgical, Inc. 950 Kifer Road Sunnyvale, CA |  |  | April 12, 2006 | 2006 | Intuitive Surgical, Inc. | Da Vinci S Surgical System's spinal pin could limit mechanical motion of the arm and render the system unavailable for surgical use, and may result in system failure and the need to convert to an alternative surgical technique. | Other | The firm did not release notification to its consignees. The reworks were arranged as part of routine service visits initiated on 4/12/06. |  | 24 | 24 units | Nationwide. | Terminated | December 05, 2007 | 602 | 24 | Not_Computer
Z-0669-2007 | 37405 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | BDÂż Sensi-DiscÂż Ticarcillin TIC-75, in vitro diagnostic, catalog # 231619, package of 10 cartridges, labeled in part ***BD Diagnostic Systems, 7 Loveton Circle, Sparks, MD 21152*** |  |  | February 19, 2007 | 2007 | Becton Dickinson &amp; Co. | Antibiotic susceptibility testing reagent was not manufactured according to specifications and may cause ineffective treatment regimen in patients with infections. The recalling firm received customer reports of false susceptibility results. | Other | The recalling firm ceased distribution of the product on 01/15/07 and subsequently notified US distributors by faxed letter and international consignees by e-mail on 02/19/07. The notification advised distributors to discontinue and return remaining inventory for replacement. The consignees were instructed to return the accompanying response form indicating inventory counts for credit. The notice also requested distributors to provide within three days the end-user lists of those receiving the product for subsequent direct notification by the recalling firm. The notification to consignees was initiated on 02/22/07 by overnight courier. The consignee notification advised that the product will give false susceptible results when testing with the identified QC organism using the Kirby Bauer method. Consignees were instructed to immediately discontinue use and return prouct identified by lot 6167435, EXP 06/30/08 and reply by returning the enclosed response form. The firm recommended to end users, primarily laboratory and testing facilities, that patient isolates yielding susceptible results with the identified lot of Ticarcillin Sensi-disc which had not been tested with the QC organism, be retested and the physicians notified. |  | 5280 | 528/10 packs | Worldwide, including USA, Australia, Canada, Hong Kong, Philippines, Belgium, Japan, Mexico, Singapore, Chile, India, and Venezuela. | Terminated | July 09, 2007 | 140 | 5280 | Not_Computer
Z-0672-2007 | 37533 | No | KYG | Ocular surgery irrigation device. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD Visitec Irrigating Cytosome, for use in cataract surgery. |  |  | March 02, 2007 | 2007 | Becton Dickinson and Company | Mislabeled: Carton label incorrectly labeled as Ref 581618 (.40 X 16mm) , may contain product Ref 581617 (.50 x 16mm) | Other | BD Opthalmic notifiied accounts by telephone on 3/2/07 and followed with letter. Accounts are requested to return product. |  | 1500 | 1500 units | Nationwide. | Terminated | March 05, 2008 | 369 | 1500 | Not_Computer
Z-0665-2007 | 37568 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Cook ATB Advance PTA Dilatation Catheter, French size 5, rated burst pressure 15 atm, inflated diameter 7 mm, balloon length 4 cm, recommended introducer French size 6, sterile, REF (Global product no.) G26887, REF (order no.) ATB5-35-40-7-4.0. |  |  | February 28, 2007 | 2007 | Cook, Inc. | Mislabeled as to size: The outer label reads 7-4.0 balloon, when the balloon is actually 5-2.0. | Other | Consignees were notified via telephone on 2/28/07 and instructed to return the product. |  | 20 | 20 units | Illinois, North Carolina and South Carolina. | Terminated | May 29, 2007 | 90 | 20 | Not_Computer
Z-0671-2007 | 37490 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK CR Plus defibrillator. The device is a battery operated, portable automated external defibrillator (AED). The devices are designed for non-professional users and are primarily used in public access defibrillator environments such as airports, schools, or public libraries. |  |  | March 02, 2007 | 2007 | Medtronic Emergency Response Systems, Inc. | LIFEPAK CR Plus defibrillators with software version 1.16 or lower may miscalculate the HLC battery capacity and result in the CHARGE-PAK battery charger icon/indicator to display and falsely indicate battery needs replacement. | Software design | Firm sent letters on 03/02/2007 dated 03/2007 via certified return receipt to consignees with affected LIFEPAK CR Plus devices. The firm advised consignees to upgrade their devices' software, and indicated that each consignee will be provided a user-installable software upgrade for their device(s) within 45 days. |  | 13835 | 13,835 units | Distribution is worldwide. The device is designed for non-professional users and are primarily used in public access defibrillator environments such as airports, schools, or public libraries. | Terminated | April 26, 2012 | 1882 | 13835 | Software | Alarm/Message | Software update | Software Update
Z-0673-2007 | 37600 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Prostate Biopsy Needle, Model: NAC-1825BB, Soxe 18 GA x 25 CM, Single use, Ethylene sterilized, Non-pyrogenic, Remington Medical, Inc., Alpharetta, GA 30005. |  |  | January 19, 2007 | 2007 | Remington Medical Inc. | Pouches were not sealed prior to sterilization. | Other | Consignees were contacted by telephone on 01/19/2007. Each customer was informed the sterility of the product may be compromised as a result of unsealed pouches and requested to visually check their inventory. |  | 192 | 192 units | Nationwide. | Terminated | May 11, 2007 | 112 | 192 | Not_Computer
Z-0704-2007 | 31058 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | InstaTrak 3500 Plus system with cranial application (aid to the surgeon for locating anatomical structures in the human body) |  |  | January 26, 2005 | 2005 | GE OEC Medical Systems,Inc | Tracking inaccuracy of 5-10 mm may result when the shunt guide accessory is used with version 5.1 software. | Other | Consignees were visited by GE personnel to upgrade the software, beginning in April, 2004, to version 5.2 |  | 91 | 91 systems (version 5.1 software) | Worldwide distribution --- USA including states of AK, AZ, AL, CA, CT, IA, IL, MD, ME, MO, MT, MS, NC, NJ, NY, OH, PA, RI, SD, UT, TX, VA, WA and WI and countries of France, United Kingdom, Japan, Italy, India and Canada | Terminated | July 08, 2011 | 2354 | 91 | Software | Output/Calculation | Software update | Software Update
Z-0684-2007 | 37168 | No | FFK | Electrohydraulic lithotriptor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Swiss Lithocast Ultra Ultrasound Handpiece (EMS- PN:EL-236) used in combination with Swiss Lithoclast Ultra System . |  |  | March 16, 2005 | 2005 | E.M.S. - Electro Medical Systems SA | Handpiece may become dislodged causing cable damage and possible electric shock. | Other | EMS Lyon, France (manufacturer) notified the single US Distributor by letter to contact their accounts and return device for replacement. Boston Scientific, Natick, Ma conducted the sub-recall by letter dated 4/22/05 to all USA accounts. |  | 160 | 160 units | Nationwide | Terminated | February 02, 2016 | 3975 | 160 | Not_Computer
Z-0685-2007 | 37247 | [6 empty] | Plicator EPS Plication System |  |  | January 26, 2007 | 2007 | NDO Surgical, Inc. | Device arm may fail to open after being deployed and require surgical intervention to remove | Process design | NDO Surgical notified accounts by letter on 1/26/07 advising to quarantine units. Replacement units are being sent and recalled devices are requested to be returned. |  | 9 | 9 units | Worldwide distribution, USA including states of AZ, GA, NH, NY, TN and country of Germany. | Terminated | December 12, 2007 | 320 | 9 | Not_Computer
Z-0686-2007 | 37277 | [6 empty] | Aspyra, Inc. CyberMED Version 2.1 Pharmacy Software |  |  | January 31, 2007 | 2007 | Aspyra, Inc. | Medications prescribed for one patient were printed on the Medication Administration Record (MAR) of another patient. | Other | Current customers were mailed a Recall Letter January 31,2007 explaining the issue and informing the customers of an upcoming CyberMED Version 2.1 maintenance software release to correct the potential issue. A fax-back Acknowledgement of Receipt form was sent with the Recall Letter. |  | 10 | 10 | Nationwide to: CA, GA, IN, MI, MN, MS, NC, ND, &amp; OK | Terminated | July 22, 2011 | 1633 | 10 | Software | Output/Calculation | Software update | Software Update
Z-0688-2007 | 37291 | No | MNB |  |  | Obstetrics/Gynecology | PMA | HTA Console System, 230 Volt, English, Product Number: 560010 (HTA system is a software-controlled hysteroscopic thermal endometrial ablation system) |  |  | January 24, 2007 | 2007 | Boston Scientific Corporation | Defective integrated circuit board could result in loss of the system pump and patient injury (hot fluid 90 degree C into uterus) if in use during the recirculation/heating phase of treatment | Other | Boston Scientific notiifed accounts by letter dated January 24, 2007 via federal Express Mail. Accounts were requested to examine inventory, remove and return product. |  | 14 | 14 units | Worldwide-USA, UK, Denmark, Italy, and France | Terminated | December 12, 2007 | 322 | 83 | Not_Computer
Z-0695-2007 | 37306 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gambro Prismaflex Continuous Renal Replacement System, Catalog Number 602314700, Gambro Lundia AB, Sweden. |  |  | January 04, 2007 | 2007 | Gambro Renal Products, Inc. | Under certain conditions, * an excessive amount of anticoagulant may be infused into the patient. (* interruption of a self-test by an alarm and obstruction of the access lines) | Device Design | Customers were notified by letter on 12/19/2006. They were told that Gambro would schedule the required upgrades to their equipment and not to use the current syringe for delivering anticoagulant. |  | 153 | 153 units (Domestic) | Nationwide distribution -- including states of AL, CA, FL, GA, IN, MD, ME, MI, MS, NC, NH, NY, OH, PA, SC, TX, VA, and WI. | Terminated | January 03, 2012 | 1825 | 153 | Not_Computer
Z-0700-2007 | 37462 | No | LGW |  |  | Neurology | PMA | The Precision Implantable Pulse Generator, Model Number SC-1110 |  |  | January 30, 2007 | 2007 | Advanced Bionics Corp | The Precision Implantable Pulse Generator with firmware Revision 3.01 incorrectly reports an error condition for a valid output from its self internal check. | Other | March 2007 by letter to physicians who had implanted the units notifying them of the recall and future upgrade by wireless RF telemetry. |  | 57 | 5 unimplanted devices; 52 implanted devices | Nationwide including Puerto Rico | Terminated | January 25, 2012 | 1821 | 57 | Software | Output/Calculation | Upgrade wireless RF telemetry | Repair
Z-0698-2007 | 37403 | No | EQH | Ear, nose, and throat fiberoptic light source and carrier. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Boston Scientific SpyGlass Light Cable |  |  | February 07, 2007 | 2007 | Boston Scientific Corporation | Light Degradation in transmission-A film may deposit on the front face and limit transmission of light | Other | Boston Scientific notified accounts by letter dated 2/1/07 requesting return of product. |  | 13 | 13 units | Nationwide Distribution --- including states of CA, CO, FL, OH, MA, MN, MO NY, PA, SC,TX, and WA . | Terminated | November 15, 2007 | 281 | 152 | Not_Computer
Z-0599-2007 | 36399 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 220V -240V, 7A (circuit breaker), 50/60 Hz (10A power source required) base; Model 801764. (Not distributed within the United States). |  |  | November 01, 2005 | 2005 | Terumo Cardiovascular Systems Corp | The power manager subsystem may not initialize properly during start-up, which will cause the battery status monitoring and battery switching operations to not be fully functional. | Other | Firm repaired the affected units beginning in November 2005, upon receipt of complaints of failure. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that they will be visited and their units upgraded as replacement components become available. |  | 485 | 485 of both model numbers | Worldwide distribution ---- Nationwide and countries of Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Malaysia, Mexico, Pakistan, Philippines, Republic of Georgia, Russia and Saudi Arabia. | Terminated | June 23, 2009 | 1330 | 485 | Not_Computer
Z-0600-2007 | 37017 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 4 inch diameter (small) Roller Pump; Model 801040. |  |  | September 01, 2004 | 2004 | Terumo Cardiovascular Systems Corp | The internal ribbon cable may loosen from the roller pump display board, resulting in loss of local control and/or local display. | Other | The firm corrected units in the field or brought them back for evaluation and repair as complaints were received between 9/04 and 4/06. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that they will be visited and their units upgraded as replacement components become available. |  | 1484 | 1,484 of both models | Worldwide: USA, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Korea, Mexico, Malaysia, Singapore, South Africa, Taiwan, Thailand and United Arab Emirates. | Terminated | July 07, 2009 | 1770 | 1484 | I/O | Device Operation | Replace component | Remove or Replace
Z-0603-2007 | 37049 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 220V -240V, 7A (circuit breaker), 50/60 Hz (10A power source required) base; Model 801764. (Not distributed within the United States). |  |  | August 01, 2005 | 2005 | Terumo Cardiovascular Systems Corp | The power supply may fail to function due to a defective capacitor. (Note; this device incorporates a redundant separate power supply not affected by this capacitor) | Other | The power supplies were replaced as complaints were received between August 2005 and February 2006. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that their units will be corrected when replacement components become available. |  | 534 | 534 of both model numbers | Worldwide Distribution: USA, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Jordan, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore and South Africa. | Terminated | August 04, 2009 | 1464 | 534 | Not_Computer
Z-0706-2007 | 37453 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | CELL-DYN 3200 Diluent/Sheath Reagent, |  |  | September 29, 2004 | 2004 | Abbott Laboratories | Elevated platelet background count results-some products may show a higher than expected platelet background count when used on the CELL-DYN 3200 System and report patient results that are unacceptable-out-of-range | Other | On September 29, 2004, the firm initiated the recall and its notification was via letters explaining the reason for the recall and requesting the product be disposed. |  | 27411 | 27,411 units | Class II Recall-Worldwide Distribution- USA including states of IL, TX, GA, and PA. and countries of Argentina, Australia, Bahamas, Barbados, Brazil, British Virgin Islands, Colombia, Dominican Republic, Germany, Hong Kong, Japan, Korea, New Zealand, Puerto Rico, Singapore, Taiwan, Thailand, Trinidad &amp; Tobago, Uruguay, and Venezuela. | Terminated | May 01, 2007 | 944 | 27411 | Not_Computer
Z-0701-2007 | 37596 | No | KDC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SCANLAN International Mobin-Uddin Vein Holder, Sterile, Disposable, single-use, catalog #1001-761, Lot AZ115806. |  |  | February 19, 2007 | 2007 | Scanlan International Inc | Incorrectly assembled by manufacturer | Other | The consignees were notified of the product recall by letter dated February 19, 2007. Consignees were asked to return product to the manufacturer and replacement product is being offered. |  | 35 | 35 units domestically ( 70 units to OUS) | Worldwide distribution --- USA including states of IN, MO and country of Japan | Terminated | June 27, 2007 | 128 | 35 | Not_Computer
Z-0707-2007 | 37632 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | 1,25-Dihydroxyvitamin D RIA Kit -For the quantitative determination of 1,25 Dihydroxyvitamin D in serum or EDTA plasma-DiaSorin, Stillwater, Minnesota 55082-0285, USA |  |  | June 05, 2006 | 2006 | Diasorin Inc. | Potential for the DiaSorin Kit 1,25-Dihydroxyvitamin D RIA Kit Control 1 (lot 548520) &amp; 2 (lot 548521) to recover out of the defined range, LOW. (If one or both of the kit controls recover outside the defined range, the run is considered invalid) | Other | Product recalled by letter dated 6/01/2006 |  | 715 | 715 | Worldwide Distribution-USA including states of UT, CO, IL, NY, NV, MI, NC, MN, VA, GA, CA, KS, VT, CT and IN, and Countries of Canada, Sweden, Germany, Belgium, Italy, France, Norway, Australia, South Korea, Taiwan, Israel, New Foundland, India, China, Brazil, Austria, Saudi Arabia, and Spain | Terminated | June 25, 2007 | 385 | 715 | Not_Computer
Z-0703-2007 | 37615 | No | FAE | Penile rigidity implant. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AMS DURA-II Universal Tips, 2cm, Product Number 72009361, Sterile, American Medical Systems, Inc. Minnetonka, MN |  |  | February 19, 2007 | 2007 | American Medical Systems | Mis-Labeling:The Dura-II 2 cm Universal Tips and the Dura-II 4 cm Distal Tips were incorrectly labeled with the opposite labels. | Other | The consignees were notified of the recall by letter starting Feb 7, 2007. Consignees were asked to return product to American Medical Systems. |  | 5 | 5 devices | Nationwide | Terminated | June 27, 2007 | 128 | 5 | Not_Computer
Z-0708-2007 | 37454 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | CELL-DYN Sapphire and CELL-DYN 4000 Diluent/Sheath Systems- List No. 01H73-01, Distributed and Manufactured by: Abbott Diagnostics Division, 5440 Patrick Henry Drive, Santa Clara, CA |  |  | April 17, 2006 | 2006 | Abbott Laboratories | High platelet background count when using the CELL-DYN 4000 System and/or CELL-DYN Sapphire System; with patient results unacceptable -out-of-range. | Other | On 4/17/06, the firm initiated the recall and its notification letters explained the reason for the recall and reported Abbott Field Service personnel will install and replace filters at all recall sites |  | 56075 | 56,075 units | Worldwide Distribution-USA including states of AL, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, and WV and Countriers of Argentina, Australia, Brazil, Hong Kong, Panama, Peru, Puerto Rico, Singapore, South Korea, Taiwan, and Venezuela. | Terminated | November 28, 2007 | 590 | 56075 | Not_Computer
Z-0713-2007 | 37018 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | ThermaSure Medical Device Drying Cabinet, series 1000 |  |  | December 21, 2006 | 2006 | Cenorin | Potential for user to receive an electrical shock. | Other | Letter dated 12/20/2006 notifying customers of potential shock hazard under certain fault conditions and that a corrective action upgrade is being developed. |  | 8 | 8 units | Worldwide. | Terminated | November 09, 2007 | 323 | 159 | Not_Computer
Z-0714-2007 | 37358 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | ProWick Shoulder Postoperative Dressing and Cold Therapy System, Model number AR-1625 |  |  | February 01, 2007 | 2007 | Integra Biotechnical LLC | During routine new product testing, it was discovered that the product may not be sterile, although sterilized in accordance with validated parameters. | Other | An email and/or fax recall notification was sent starting on 02/1/2007 by Arthrex, Inc. (the Primary Distributor) . A product recall notice in the form of a letter was mailed via US Postal Certified Mail, return receipt requested by Arthrex, Inc. to all of the Distributor''s customers starting on 02/1/2007, including those who have not ordered or received the Recalled Product. Customers were asked to 'segregate and protect the product from shipment, and return all unused products to Arthrex, Inc.' |  | 1769 | 1769 | Nationwide | Terminated | August 03, 2012 | 2010 | 1769 | Not_Computer
Z-0710-2007 | 37597 | No | KMF | Liquid bandage. | General Hospital | General &amp; Plastic Surgery | 510(k) | Sea-Clens Wound Cleanser, 6 fl. oz., Manufactured by Coloplast Corp., Marietta, GA, #11701-159-36, #1063. |  |  | February 14, 2007 | 2007 | Coloplast Corp Skin Care Div | Coloplast Corp. is recalling Sea-Clens brands Sea-Clens wound cleaner which were found to have some particulates (foreign materials). | Other | The initial consignees of the product were contacted and notified of the recall February 14, 2007. The consignees were asked to contact their customers immediately and return all product to the manufacturer. |  | 7848 | 654 boxes, 12 bottles per box for a total of 7,848 bottles | Nationwide. | Terminated | June 25, 2007 | 131 | 7848 | Not_Computer
Z-0715-2007 | 37507 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Contact lens solution. The product is distributed under the following trade names: Bausch &amp; Lomb ReNuÂż Multi-Plus Multi-Purpose Solution, Equate Multi-Purpose Solution and Target Brand Multi-Purpose Solution. The product is distributed in up to three sizes: 4 fl.oz., 12 fl. oz. and 16 fl. oz |  |  | March 05, 2007 | 2007 | Bausch &amp; Lomb Inc | A higher than expected amount of trace iron in some bottles of finished product that over time affects the product's stability and color. | Other | The firm issued a press release on 03/05/2006. Consignees were notified on 03/06/2007 via a customized Recall Notification Package via overnight courier with tracking and delivery confirmation capability. Bausch &amp; Lomb Sales Representatives will visit the home offices of the major wholesale and distributors concurrent with the mailing of the notification packages. |  | 1500000 | 1.5 million units | Worldwide, including USA, Argentina, Brazil, Canada, Chile, Costa Rica, Korea, Mexico, Taiwan and Venezuela. | Terminated | January 09, 2008 | 310 | 1500000 | Not_Computer
Z-0716-2007 | 37642 | [6 empty] | QRS Card 12 Lead Resting &amp; Stress ECG Machine with Blue Tooth (wireless) interface. Cardiology Suite (CS 4.0) Software is used with the machine. |  |  | March 05, 2007 | 2007 | Pulse Biomedical Inc | No 510(k) marketing clearance for the product | Other | On 3/5/07, the recalling firm telephoned their customers to inform them of the problem and the need to return to the product. As a follow up to the telephone calls a letter dated 3/5/07 was sent via certified mail. |  | 8 | 8 units | The product was shipped to four direct customers in CA, NY, TN and US military (Cuba). The product was also shipped to Chile, Greece, and the Netherlands | Terminated | August 29, 2007 | 177 | 8 | Not_Computer
Z-0674-2007 | 37188 | No | BTI | Portable air compressor. | Anesthesiology | Anesthesiology | 510(K) Exempt | Medical Air Compressor |  |  | January 26, 2007 | 2007 | Draeger Medical, Inc. | Compressors not providing sufficient supply pressure to connected medical ventilator | Other | The recalling firm issued a Medical Device Recall Letter dated 1/22/07 to their customers via Certified Mail. An authorized service representative will inspect each potentially affected device. If a device is found to have an affected dryer assembly, it will be repaired free of charge. |  | 67 | 67 units | Nationwide. | Terminated | June 15, 2007 | 140 | 67 | Not_Computer
Z-0718-2007 | 37486 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Deltec branded CADD-Prizm PCS II ambulatory infusion pump and the CADD-Prizm VIP system ambulatory infusion pumps; (Both) Model 6101; Order #21-8861-01 (PCS II) and 21-8821-01 (VIP). Smiths Medical MD, Inc., St. Paul, MN 55112 USA; RX ONLY. |  |  | January 19, 2007 | 2007 | Smiths Medical MD, Inc. | The year of manufacture had been incorrectly entered as 1980 into the pumps software during the manufacture process. As a result the pump may go into an immediate 'Clock battery is Low/Service Immediately' or 'Clock Battery needs service soon' alarm message indicating that the pump requires servicing. | Other | All consignees who have received the affected CADD-Prizm PCS II ambulatory infusion pumps and the CADD-Prizm VIP System ambulatory infusion pumps were contacted via telephone by February 1, 2007. The problem will be explained to the consignee and they will be asked to return the pumps to Smiths Medical MD in order for the problem to be corrected. |  | 27 | 27 | Nationwide. | Terminated | May 25, 2007 | 126 | 27 | Software | Device Operation | Remove | Remove or Replace
Z-0719-2007 | 37598 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 30ml Syringe Luer-Lok tip |  |  | February 15, 2007 | 2007 | Becton Dickinson &amp; Company | The unit label on a limited number of syringes is incorrect, citing a 20 ml size. All other levels of labeling are correct, identifying the product as 30 ml size. | Other | Distributors were sent recall notification letters on February 1, 2007Recall notification letters were sent to all direct accounts and potential end users via UPS second day air on February 15, 2007. |  | 332400 | 332,400 units | Nationwide. | Terminated | July 19, 2007 | 154 | 332400 | Not_Computer
Z-0722-2007 | 37334 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Clinical Chemistry Phosphorus in vitro diagnostic, List Number: 7D71-30 and 7D71-20 |  |  | January 30, 2007 | 2007 | Abbott Laboratories Inc. | Patient results are falsely decreased by up to 15% at these levels: 1) Serum Plasma phosphorus greater than 8.0 mg/dL (2.60 mmol/L) 2) Urine phosphorus greater than 80.0 mg/dL (25.80 mmol/L) | Other | A Product Correction letter dated January 30, 2007 provided to all Clinical Chemistry Phosphorus customers that have received lot numbers 42020HW00 and 44037HW00 stating that patient results are falsely decreased by up to 15% at these levels: Serum/Plasma phosphorus greater than 8.0 mgldL (2.60 mmol/L) and Urine phosphorus greater than 80.0 mg/dl. (25.80 mmol/L). Modified linearity ranges are provided in the letter that are to be implemented when using the two lots. Instructions are also provided for changing the linearity range on the Architect cSystem as well as the Aeroset System. Customers are instructed that If they have forwarded any of Information the Clinical Chemistry Phosphorus lots listed above to another laboratory to provide a copy of this letter to them. A customer response form is provided. |  | 87722 | 42020HW00 - 1,351 units worldwide; 44037HW00 - 314 units worldwide | Worldwide, including USA, Canada, Trinidad &amp; Tobago, Bahamas, Panama, Cayman Islands, Jamaica | Terminated | August 10, 2011 | 1653 | 87722 | Not_Computer
Z-0721-2007 | 37638 | No | MEQ |  |  | Gastroenterology/Urology | PMA | Urologix Microwave Delivery System 2.5 - 3.5 For use with prostatic urethra lengths of 2.5-3.5 cm only. Model TA1321D. |  |  | March 16, 2007 | 2007 | Urologix, Inc. | The catheter in the kit matches the description on the kit and pouch label; however, an incorrect label was placed onto the catheter during manufacturing. | Other | A letter dated 3/16/07 was FedEx overnight to consignees. The letter described the product, problem and gave instructions for returning and replacing the affected product. |  | 21 | 21 | Nationwide and the Netherlands. | Terminated | May 25, 2007 | 70 | 117 | Not_Computer
Z-0705-2007 | 37167 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL800 FLEX Series analyzer with FLEXQ module-equipped with software version 5.2 to 5.27 |  |  | January 16, 2007 | 2007 | Radiometer America Inc | In some situations pre-registered patient data is not reset when leaving the pre-registration mode in the software. This occurs when a Safe Pico blood sampler associated with a specific patient is pre-registered in the analyzer by scanning a barcode. | Other | Radiometer America issued a recall letter to their U.S. customers dated 1/16/2007. The letter informs the customers of this software anomaly and the conditions underwhich analyzer results can be stored with incorrect Patient Demographics. The letter goes on to inform the customers that a Radiometer Service Representative will be visiting each customer location shortly in order to upgrade the affected software. In the meantime, the letter recommends that the customer(s) not pre-register their samples. The letter is accompanied by an attached FAX FORM which the customer is asked to sign and complete in order to verify receipt of the recall letter and to fax back to the recalling firm. |  | 1345 | 64 in U.S. &amp; 1,281 to foreign accounts | Worldwide Distribution-USA including states of NY, KS, SD, OH, TX, KY, NC, IL, MO, AL, UT, CO, CA, NM, and MI and Country of Denmark | Terminated | May 08, 2012 | 1939 | 1345 | Software | Output/Calculation | Software update | Software Update
Z-0728-2007 | 37446 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Estradiol Reagent, each kit contains bottles of coated Microparticles and bottles of Conjugate; 400 Test Kit-list number 2K25-20 (4 x 100 tests) and 100 Test Kit-list number 2K25-25 (1 x 100 tests); Abbott Laboratories, Abbott Park, IL 60064-3500 |  |  | February 19, 2007 | 2007 | Abbott Laboratories | A small number of reagent kits were incorrectly configured. The affected kits contain either two bottles of conjugate or two bottles of assay diluent instead of one bottle of each. | Other | Abbott initiated a field correction recall by letter dated 2/19/07 to all customers receiving the affected reagent. The accounts were informed of the incorrectly configured reagent kits. The accounts were requested to assess the inventory for the affected lots of reagents, inspect each kit and destroy any kits incorrectly configured, recording the number of kits inspected and the number of kits discarded on the customer reply form and faxing it to Abbott by 3/2/07, indicating that the letter was received and understood. Replacement kits will be issued for the number of kits destroyed, as indicated on the reply sheet. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. |  | 1944 | 1,944 kits | Worldwide distribution including U.S.A., Puerto Rico, Canada, Mexico, Venezuela, Chile, Uruguay, Argentina, Brazil, Germany, Japan, Singapore, Hong Kong, Taiwan, Thailand, South Korea, Australia, New Zealand and Panama. | Terminated | January 09, 2008 | 324 | 1944 | Not_Computer
Z-0723-2007 | 37383 | [6 empty] | Hamilton brand Disposable Precision Tips, |  |  | January 30, 2007 | 2007 | Hamilton Co | Some of the packaged tips have increased force during aspiration, which can lead to pipetting errors such as LLD error messages, and also has risk of causing a missed pipetting. | Other | The firm initiated its recall on January 30, 2007. The initial notification was sent via certified mail and e-mail. Subrecalls will be conducted by the distributor. |  | 37 | 37 cases | Nationwide (via a single distributor in New Jersey) | Terminated | November 28, 2007 | 302 | 37 | Not_Computer
Z-0724-2007 | 37582 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Smith &amp; Nephew Suture Passer 12 inch |  |  | March 01, 2007 | 2007 | Smith &amp; Nephew, Inc. Endoscopy Division | Product is labeled as reusable but there is a potential that the cannulated handle of the product cannot becleaned adequately following use. | Other | Smith &amp; Nephew notified accounts by letter on 3/1/07, users are advised to destroy product and return unused product in their inventory. A reply form is required to be completed. |  | 12292 | 12, 292 units for all recalls Z-0724-0726-07 | World wide- USA and Canada, Belgium. Germany, Denmark, Spain, Finland, France, Great Britain, Greece, Italy, Netherlands, Portugal, Austria, Ireland, Sweden, Argentina. Chile, Columbia, Brazil, China. Japan, Korea, Thailand, Hong Kong, Malaysia. Taiwan, Singapore, India, , Dubai, Mexico, Australia, Egypt. Switzerland, New Zealand and Norway | Terminated | July 14, 2008 | 501 | 12292 | Not_Computer
Z-0727-2007 | 37641 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Sunrise Medical-Hoyer Advance Patient Lifts (Portable and Folding Patient Lift) Sunrise Medical, Inc. |  |  | February 21, 2007 | 2007 | Joerns Healthcare Inc. | Unapproved Design Change (by base supplier) consisting of a reduction in the number of mast-to-base welds (from four welds to two welds). | Other | Recall initiated by letter dated 2/21/2007. Sunrise Medical will schedule onsite appointments to complete the base replacement. |  | 581 | 581 units | Worldwide Distribution --- USA including states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and country of Canada. | Terminated | October 07, 2010 | 1324 | 581 | Not_Computer
Z-0733-2007 | 37465 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Bio-Rad-TOX/See Drug Screen Test (25 tests/box), Catalog number 194-5230; Product is distributed by Bio-Rad Laboratories, Inc., Hercules, CA |  |  | February 12, 2007 | 2007 | Bio-Rad Laboratories Inc | The product may produce a faint line which could be interpreted as a false negative test result . | Other | The firm intends to contact its subsidiary distributors and sales staff via Fax (Feb 12, 2007) and phone call. In addition, the recall notification will be posted on the firm's intranet site. Customers are instructed to destroy recalled product following standard laboratory practices. |  | 177 | 177 Kits | Worldwide distribution --- USA and countries of Australia and Italy. | Terminated | February 04, 2008 | 357 | 177 | Not_Computer
Z-0735-2007 | 37585 | No | NTI | Drug metabolizing enzyme genotyping system. | Toxicology | Toxicology | 510(k) | AmpliChip CYP450 Test, CE-IVD, M/N: 4381866190, Roche Molecular Systems, Inc. |  |  | March 09, 2007 | 2007 | Roche Molecular Systems, Inc. | The current package insert for the AmpliChip CYP 450 test, lists an incorrect part number for DNase I. The DNase is listed in the "Other Materials Required" section as 'DNase I rec., RNase-free, P/N 04716728001 (Roche Applied Science)" . The Part Number listed for the DNase I listed is incorrect and of lower specific activity than the correct/validated DNase. | Other | A product advisory notice was issued on March 9, 2007 to all affiliates, indentifying the problem and listing corrective actions followed by a customer letter (Urgent Medical Device Correction) issued on 3/21/2007, requesting several actions be performed including discontinued use of incorrect DNAse I. |  | 89 | 89 | World wide Distribution --- USA including states of MI, IN, NC, CA KY, IN, VA, and FL.and countries of Netherlands, Germany, Sweden, Denmark, Italy, Spain, Belgium, Switzerland, and Hungary | Terminated | August 23, 2007 | 167 | 218 | Not_Computer
Z-0737-2007 | 37602 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | CoaguChek brand PT Test Strips; U.S. Catalog Number 3116247 (48 strip pack-professional use). |  |  | February 28, 2007 | 2007 | Roche Diagnostics Corp. | The product was erroneously distributed to home users in contrast to its use for healthcare professionals only. | Other | All consignees were notified by telephone on/or about 2/28/07 and some received a follow-up letter dated 3/2/07 or 3/5/07. |  | 10 | 10 | Nationwide-California, Georgia, Louisiana, Maryland, Oklahoma and Texas. | Terminated | June 13, 2007 | 105 | 10 | Not_Computer
Z-0730-2007 | 37628 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Flexxus&#153; Endoscopic Biliary Stent, 7.5 Fr |  |  | March 14, 2007 | 2007 | Conmed Endoscopic Technologies, Inc. | Blue safety clip incorrectly placed on the stent may prevent deployment | Other | ConMed Endoscopic Technologies notified consignes by letter dated 3/14/07 requesting them to return recalled devices. |  | 1 | 1 UNIT | Worldwide distribution, USA including states of DE, OH, MI, DC and countries of Belgium Japan. | Terminated | May 15, 2007 | 62 | 12 | Not_Computer
Z-0736-2007 | 37599 | No | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | Roche COBAS Integra DIG, Digoxin Reagent for use on the COBAS INTEGRA models 400, 400 plus, 700 and 800 analyzers, as well as, Cobas c 501 analyzer; Catalog Number 20737836322. |  |  | March 16, 2007 | 2007 | Roche Diagnostics Corp. | The lower detection limit (LDL) may be a higher value than stated in the labeling. Note: To resolve this issue, the Digoxin LDL claim has been revised to 0.3 ng/mL for all reagent lots. | Other | A recall letter dated 3/16/07 was sent to each consignee advising them that the LDL claim in the label has been revised to 0.3 ng/ml for all reagent lots. |  | 4735 | 4,735 | Nationwide. | Terminated | June 12, 2008 | 454 | 4735 | Not_Computer
Z-0694-2007 | 37687 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | BioHorizons 0.050 Hex Driver, Long, REF 300-351, BioHorizons Implant Systems, Inc., Birmingham, AL 35243, non-sterile, Rx only |  |  | March 12, 2007 | 2007 | BioHorizons Implant Systems Inc | Premature Wear; The material from which the hex driver was made was improperly hardened during manufacturing and could lead to premature wear. | Other | Letters and replacement drivers were sent to the consignees (both domestic and abroad) on 3/12/2007 explaining the problem and requesting that the consignee discard the instrument and use the replacement. They were delivered via United Parcel Service (UPS) 3rd day package delivery. |  | 436 | 436 units | Worldwide Distribution -- USA and countries of Canada, Norway, Australia, Spain, Brazil, Pakistan, Israel, Latvia, Japan, Columbia, Thailand, Turkey, Italy, South Korea, India, Taiwan, Ukraine, Russia, United Arab Emirates, Saudi Arabia, Hong Kong, Romania, UK, Kuwait, and Lithuania | Terminated | September 04, 2007 | 176 | 2732 | Not_Computer
Z-0744-2007 | 37254 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Mega Diagnostics Alcohol Standard (Cat.# A7504-STD) packaged in 5ml vials. |  |  | September 20, 2006 | 2006 | Mega Diagnostics | Potential for microorganism growth in the alcohol standard. | Other | The firm sent a recall notification to their customer in Indonesia and advised them to dispose of the product |  | 6 | 6 units | Country of Indonesia | Terminated | June 04, 2012 | 2084 | 18 | Not_Computer
Z-0746-2007 | 37427 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Vaxcel Implantable Vascular Access System - Plastic Standard Port With 9F Polyurethane Catheter, Catalog/REF No. 45-360, UPN/Material No. M001453600, Single Use Only, STERILE, EO, Boston Scientific MEDI-TECH (617) 972-4000/(800) 225--3238, 480 Pleasant Street, Watertown, MA 02472 --- The port is provided as a kit, with various single use, disposable items necessary for achieving vascular access and implanting the port catheter. Each kit is provided in a plastic (PETG) tray with a heat-sealed tyvek lid. This inner tray is then placed in a second outer plastic (PETG) tray, along with the Directions for Use, and also sealed with a tyvek lid which is then labeled. The port kit is placed in a cardboard carton, with additional labels and sealed. |  |  | June 13, 2006 | 2006 | Boston Scientific Corp | Mispackaging: some kits of peelable sheaths/dilators may contain a 10F sheath instead of a 9F sheath--(Recalling Firm was notified by their supplier B. Braun Medical of the error) | Other | Boston Scientific sent Urgent Medical Device Recall Notification- Immediate Action Required letters- (dated 6/13/06) to all 11 end users via Federal Express. The letters informed the end users that some of the peelable sheaths/dilators in the Vaxcel Implantable Vascular Access kits may incorrectly contain 10F sheaths instead of 9F sheaths which had been supplied by B. Braun Medical Inc. Improper sizing of the sheath for the catheter may result in difficulty /inability to insert the catheters. The letters instructed the end users to immediately remove the affected lots from inventory, segregate the recalled product, contact Boston Scientific at 1-800-862-1284 to obtain a Return Authorization (RGA) number, pack the product in an appropriate shipping carton, affix the shipping label to the outside of the shipping carton, write the RGA number on the outside of the shipping carton, and to return the recalled product to Boston Scientific''s Distribution Center Attn: QA Returns, 500 Commander Shea Blvd., Quincy, MA 02171. Accompanying each letter was a Reply Verification Tracking Form to be completed by the end user and mailed or faxed to Boston Scientific even if the user did not have the recalled lots on hand. |  | 5 | 5 | NY, CA, MD, WA, TN. | Terminated | April 23, 2007 | 314 | 27 | Not_Computer
Z-0748-2007 | 37606 | No | JNO | Trypsin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | DiaSorin TRYPSIK Kits, Catalog Number P2573-(assay used as a procedure for the quantitative determination of trypsin-like immunoreactiviy (TLI) in human serum or plasma samples) |  |  | February 15, 2007 | 2007 | Diasorin Inc. | Assay Failure: Assay failures may occur due to the kit control running outside of the established range (invalidating the assay) | Other | The consignees &amp; distributors were notified by letter on 2/15/07 . The Distributors are instructed to forward the notification to the customers who have received the affected products. A response form will be requested from the customers. |  | 100 | 100 kits | Worldwide Distribution --- USA including states of UT, CO, IN, NC, CA, FL and countries of Canada, Spain, Germany, Belgium &amp; Italy. | Terminated | July 15, 2007 | 150 | 100 | Not_Computer
Z-0747-2007 | 37528 | No | LOM |  |  | Microbiology | PMA | VITROSÂż Immunodiagnostics Products HBsAg Reagent Pack, Catalog # 6801322: 1 Reagent Pack box per sales unit (100 tests per box), and Catalog # 6802450: 5 Reagent Pack boxes per sales unit (100 tests per box), IVD. Ortho-Clinical Diagnostics Amersham UK. |  |  | February 01, 2007 | 2007 | Ortho-Clinical Diagnostics | Complaints of an increase in ''Reactive'' results with patient samples collected in sodium citrate or EDTA plasma collection tubes when using these lots of VITROS HBsAg Reagent Pack compared to samples collected in other tube types. | Other | On 2/01/07, all US end-user consignees who received the VITROS HBsAg Reagent Packs, were informed of this issue by Federal Express overnight letter, and informed that they may continue to use their current lot of VITROS HBsAg Reagent Pack providing that they followed the instructions in the letter. INSTRUCTIONS: Please do the following if you choose to continue using sodium citrate and/or EDTA plasma collection tubes: - Do not freeze samples collected in sodium citrate or EDTA plasma collection tubes prior to processing the assay. - Process samples collected into sodium citrate plasma tubes within 12 hrs. of collection. -Continue to perform supplemental testing to confirm all HBsAg results that repeat as 'Borderline' or 'Reactive', as directed in the Package Insert. - Complete and return the attached Confirmation of Receipt Form upon receipt of this notification. - Forward the information in this notification if you have distributed the product outside your facility |  | 48514125 | Cat. #6801322 (1 Reagent Pack box per sales unit): 11,761 units; Cat. #6802450 (5 Reagent Pack boxes per sales unit): 4125 units | Worldwide Distribution --- USA including Puerto Rico and countries of Tortola (British Virgin Islands) and Anguilla. | Terminated | August 03, 2012 | 2010 | 48514125 | Not_Computer
Z-0749-2007 | 37698 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobilett XP, mobile X-ray system, Model number 1818454 |  |  | March 21, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Mounting bolts for tank fork assembly and collimator flange may become loose. | Other | Siemens Service Personnel are visiting each consignee, beginning March 21, 2007, to replace the required mounting bolts and tighten them according to required torque specifications. |  | 186 | 186 units | The products were shipped to medical facilities nationwide. | Terminated | August 07, 2007 | 139 | 227 | Not_Computer
Z-0754-2007 | 37742 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | RITAÂż StarburstÂż MRI SemiFlex (3-5 cm Diameter/25cm Length), Electro-surgical Device, Manufactured by: RITA Medical Systems, Inc., One Horizon Way, Manchester, GA 31816 |  |  | April 03, 2007 | 2007 | Rita Medical Systems, Inc. | The product may have a cracked tray which can compromise the sterility of the product. | Other | Notification by letter of the AngioDynamics, Inc. sales force and the domestic and international distributors began on April 2, 2007. They were asked to return all affected devices to the Manchester, GA division of AngioDynamics, Inc. A reply form was attached to be completed and returned via fax to AngioDynamics. |  | 2021 | 2021 devices | Worldwide, including USA, Austria, Canada, Croatia, Denmark, Egypt, Estonia, France, Finland, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Latvia, Netherlands, Norway, Poland, Republic of Ireland, Romania, Russia, Saudi Arabia, Singapore, Solomon Islands, South Africa, Spain, Switzerland, UK, and United Arab Emirates. | Terminated | January 20, 2010 | 1023 | 5103 | Not_Computer
Z-0761-2007 | 37518 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Bio-logic-Ceegraph/Sleepscan Netlink Traveler, Bio-logic Systems Corp., a nebus company, Mundelein, IL 60060; Model 580-T2ASM2. (a Digital EEG/Sleep Recorder Electroencephalograph that includes a built-in pulse oximeter, body position sensor, snore monitor, chest, abdominal and air flow transducers and electrode array for EEG, EMG and EOG) |  |  | February 26, 2007 | 2007 | Bio-Logic Systems Corp | Battery and Battery Compartment; incorrect placement in the Netlink Traveler unit which caused a short in the wiring of the battery pack and overheating. | Other | On 2/26/07, Bio-logic Systems shipped replacement batteries, marked with part number 520-LIBAT4-3 and a label reading 'THIS SIDE OUT', and a new battery compartment cover with a label inside the cover to clarify the correct installation of the battery to each of the affected customers. The accounts were requested to remove and replace the old batteries, part 520-LIBTA4-2, and compartment covers, and return the old batteries and covers to Bio-logic Systems in the enclosed pre-paid shipping container, along with the completed accounting form. |  | 130 | 130 battery packs (112 domestic USA, 18 foreign) | Worldwide, USA , Canada, China, Hong Kong, India and Israel. | Terminated | August 26, 2008 | 547 | 130 | Battery | Physical Safety Hazards | Replace battery | Remove or Replace
Z-0731-2007 | 37189 | No | NFB | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | DeVilbiss PD1000A Pulse Dose Oxygen Conserving Cylinder |  |  | January 24, 2007 | 2007 | Sunrise Medical | Cylinder could lose oxygen at high pressure after the filling process | Other | The recalling firm began to notify customer via telephone on 1/24/07. An ''Urgent Medical Device Recall Letter'' was sent on 1/29/07. Customers were advised what model and serial numbers were affected and how to contact the firm and/or dealer to have the unit replaced. The recalling firm identified additional serial numbers that were involved in the recall therefore another Urgent Medical Device Recall Letter dated 6/12/07 was sent to the customers that received the addiitonal units. |  | 4169 | 4,169 units (Z-0731/0732-2007) | Worldwide-original shipments to distributors/dealers USA and Canada. | Terminated | August 27, 2009 | 946 | 4169 | Not_Computer
Z-0765-2007 | 37281 | No | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb Millennium Microsurgical System TSV 25 Ga. High Speed Vitrectomy Cutter, Catalog No. CX5825, packaged individually in a Tyvek pouch, 6/box, RX. The firm names on the label are Bausch &amp; Lomb Incorporated, Rochester, NY, and Bausch &amp; Lomb Incorporated, Kingston-upon-Thames, UK. |  |  | February 09, 2007 | 2007 | Bausch &amp; Lomb Inc | The cutter probe tip breaks during use. | Other | On 2/12/07, express mail letters were addressed to operating room supervisors to notify them of the recall. Additional E-mail notifications were issued on 2/14/07 to doctors and letters were addressed to them on 2/20/07 to make them aware of the recall, although they do not directly receive the product. The communications instructed them to not use the affected product and to return it, along with the business reply card. |  | 187506750 | 32,190 units of Catalog No: CX 5825 | Class II Recall - Worldwide Distribution --- USA including states of MO, NC, FL, CA, and TX and countries of Canada, China, Japan, Australia, Hong Kong, India, Singapore, Korea, Taiwan, Thailand, Argentina, Brazil, Mexico, and The Netherlands | Terminated | April 07, 2009 | 788 | 187520435 | Not_Computer
Z-0767-2007 | 37759 | No | GES | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Miltex Stainless Steel Disposable Scalpel #22, Part Number 4-422. Each box contains 10 individually packaged scalpels. |  |  | April 09, 2007 | 2007 | Miltex, Inc. | Sterility may be compromised based on incomplete package seals (manufacturer notified Miltex) | Other | The recalling firm issued a recall letter dated 4/9/07 informing the customers of the problem and the need to return the product. The letter was sent via UPS with a return receipt required to verify delivery. Distributors were notified to sub-recall. |  | 400 | 400 boxes | Nationwide to distributors in CA, FL, IA, ID, NM, NJ, OH, PA, RI, and VA. | Terminated | August 07, 2007 | 120 | 400 | Not_Computer
Z-0756-2007 | 37495 | No | MLN | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | PADPRO ADULT Radiotranslucent Pad w/ZollÂż connector. Defibrillation/Pacing/Cardioversion/Monitoring Electrodes, REF/Catalog # 2516Z. RX ONLY. Made in USA for: CONMED CORPORATION Utica, New York, 13502. |  |  | March 08, 2007 | 2007 | Conmed Corporation | Firm was notified of three incidents over the past 15 months in which customers opened the pouch &amp; found an incorrect connector. This error makes the product unusable. The Zoll and Medtronic connectors are not interchangeable and are unique to their specific device. | Other | On 3/08/07, "URGENT DEVICE RECALL" letters (dated 3/05/07) will be sent via UPS to the US distributors and direct user accounts. The letters inform them of the problem and of the recall. The letters instruct them to immediately discontinue using the affected lots, inspect their stocks, and to return them as soon as possible. The letters instruct the consignees to notify their customers of the recall of the devices to the user level if they have been distributed, and request the return of the products. Additionally, when requested, ConMed will notify distributor's customers (sub-accounts) by the same means. Returned devices will be automatically replaced by ConMed Corporation. A separate letter is being sent to the foreign accounts today instructing them to return the devices back to ConMed Corporation for final disposition. The letters also instruct the consignees to complete the accompanying reply form and fax it to Nancy Crisino at (315) 624-3089. ConMed has contracted Stericycle in Indianapolis, IN, to handle the recall and returns from the domestic accounts. |  | 9712 | 9712 pairs | Worldwide, including USA, Canada, Romania, Chile, and Belgium | Terminated | July 16, 2012 | 1957 | 10632 | Not_Computer
Z-0760-2007 | 37607 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 840 Ventilator System, Catalogue Numbers: |  |  | March 12, 2007 | 2007 | Nellcor Puritan Bennett | In rare instances, the safety valve may remain open after an occlusion is resolved. If the safety valve does not reset, the ventilator will detect a circuit disconnect condition. When this happens, the ventilator will annunciate a high-urgency audible alarm and the ventilator will not resume normal operation. | Other | The correction will extend to the end user. All users will be contacted by mail or through personal contact. In the U.S., all customers were mailed the 'Urgent Device Correction' letter along with a serial number listing of those units which records show were consigned to them beginning on March 12, 2007. The customer communication informs customers of the issue and provides instruction for continued safe use of the device pending the correction. |  | 3756 | 3756 | Worldwide, including USA, ARGENTINA, AUSTRALIA,BELGIUM, BRAZIL, CANADA, CHILE, COLOMBIA, EGYPT,FRANCE, GERMANY, GREECE, GUATEMALA, IRAN,IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN,KOREA, LEBANON, MEXICO, MOROCCO, NICARAGUA, PANAMA, POLAND, PUERTO RICO, ROMANIA, RUSSIA,SINGAPORE, SOUTH AFRICA, SPAIN, TURKEY, UK,&amp; YUGOSLAVIA | Terminated | December 13, 2012 | 2103 | 3756 | Hardware | Device Operation | Correction | Repair
Z-0759-2007 | 37604 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Medipro OP4 Prosthetic Knee Joint (external for prosthesis, not implant). |  |  | August 11, 2006 | 2006 | Medi Manufacturing Inc | Joints contain a component manufactured with an alloy that may not comply with manufacturer's specifications. | Other | Each consignee was contacted by phone on 8/11/2006. A follow up letter was sent to consignees on 8/11/2006 providing details on who to contact and how to return the affected devices for replacement. A self addressed, stamped response form was included to be completed and returned to Medi. |  | 8 | 8 units | Nationwide. | Terminated | April 12, 2011 | 1705 | 8 | Not_Computer
Z-0755-2007 | 37691 | No | LLM | Epstein-Barr virus serological reagents. | Microbiology | Microbiology | 510(k) | DiaSorin ETI-EBNA-G kits, Catalog No. P001607A. In vitro diagnostic. |  |  | March 14, 2007 | 2007 | Diasorin Inc. | A lot to lot inconsistency was noted on kits of ETI EBNA G. | Other | All consignees will be notified by phone followed by a notification letter sent by fax March 14, 2007. The distributors are instructed to forward the notification to the customers who have received the affected product. |  | 203 | 203 | Nationwide and Canada. | Terminated | July 15, 2007 | 123 | 203 | Not_Computer
Z-0769-2007 | 37760 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Magnetom Espree System with OR Table |  |  | March 30, 2007 | 2007 | Siemens Medical Solutions USA, Inc | If table top is not seated properly it can lock up/stick. | Other | The recalling firm issued a Safety Advisory Letter to their customers on March 30, 2007, informing them of the problem and steps to take to avoid the problem. |  | 3 | 3 units | The product was shipped to medical facilities in NY and TX. | Terminated | August 07, 2007 | 130 | 3 | Not_Computer
Z-0768-2007 | 37590 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medfusion 3500 Syringe Infusion Pump with PharmGuard Medication safety Software versions 4.1.3 and 4.1.4. Medex, 4350 Green Parkway, Suite 200. Duluth, GA 30096 |  |  | February 03, 2007 | 2007 | Smiths Medical MD, Inc. | Software anomaly that causes the pump to experience a sudden and unexpected failure. This anomaly occurs during the pump's startup sequence when the pump's Preventative Maintenance Alarm goes off, preventing the pump from completing its power up routine. Lock-up occurs and the pump will not function. | Other | Consignees were initially notified of the recall via a Customer Information Bulletin which was sent via electronic mail 2/3/07 and confimed with a telephone conversation. A Product Recall Notification and a Product Recall Bulletin will be sent to all consignees via the mail on 2/16/07. |  | 23519 | 23,519 | Nationwide, Canada, and Italy. | Terminated | February 10, 2011 | 1468 | 23519 | Software | Device Operation | Notification | Safety Notice/Insructions
Z-0758-2007 | 37591 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Conditional Fiber Optic Laryngoscope Kit, which contains laryngoscope blades, handles and 2 battery packs for use in Magnetic Resonance environments; Distributed by Teleflex Medical, Research Triangle Park, NC 27709; catalog number 005852300 |  |  | March 01, 2007 | 2007 | Teleflex Medical | The MRI Battery is out of specification for its magnetic content. | Nonconforming Material/Component | Teleflex telephoned the account on 3/1/07, and faxed them a follow-up recall letter on the same date. The account was informed that the batteries did not meet specification, and was requested to check the batteries in the laryngoscope kit shipped to them for the affected battery lot numbers, and return the affected batteries to Teleflex for replacement. Any questions were directed to Linda Todd at 1-800-334-9751, ext. 4951. |  | 1 | 1 kit | Georgia | Terminated | May 03, 2007 | 63 | 1 | Not_Computer
Z-0773-2007 | 36826 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/ALERT Classic System Software, Blood Culturing System |  |  | August 14, 2006 | 2006 | bioMerieux, Inc. | Plastic bottle samples assigned less sensitive default glass bottle algorithm instead of the plastic bottle algorithm on BacT/ALERT Microbial Detection Systems. | Other | Consignees were notified via A Field Corrective Action Notification on/about 10/18/2006. A second notification was sent on October 31, 2006. |  | 10790 | 10790 units, total for all models | Worldwide, including USA, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Dutch Antilles, Egypt, Ecuador, Estonia, Faeroe Islands, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kenya, Korea, Latvia, Lithuania, Luxembourg, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania , Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Suriname, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Terminated | December 08, 2009 | 1212 | 10790 | Software | Output/Calculation | Corrective Action Notification | Repair
Z-0778-2007 | 37719 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GEM Premier 4000 Laboratory Analyzer |  |  | March 27, 2007 | 2007 | Instrumentation Laboratory Co. | Instrument in Micro Sampling Mode failed to aspirate and no error code was generated and an incorrect patient report could result | Other | Instrumentation Labs on 3/2707 notified Service Reps to perform on site mandatory upgrade. The foreign counterparts were advised by email on 3/21/07 |  | 210 | 210 | Worldwide, including USA, Australia, Austria, Greece, Spain, Netherlands, Great Britain, France, Germany, Italy, and Sweden | Terminated | May 23, 2012 | 1884 | 210 | Not_Computer
Z-0779-2007 | 37762 | No | GIR | Russell viper venom reagent. | Hematology | Hematology | 510(k) | American Diagnostica Inc. ACTICLOT dPT |  |  | February 28, 2007 | 2007 | American Diagnostica, Inc. | Stability of product not assured. An increase of clotting time for both normal and abnormal plasmas | Other | American Diagnostica notified accounts by letter on 2/28/07 and 3/02/07 and international distributors on 2/28/07. Users are requested to discontinue use and destroy. Kits will be replaced. |  | 249 | 249 kits | Worldwide, including USA, Canada, Germany, UK, Italy, Israel, Greece, and the Netherlands. | Terminated |  |  | 249 | Not_Computer
Z-0780-2007 | 37510 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Doro Products Adapter, Product number 5979-300. The adaptor system is attached to OSI''s Spinal Frame, which is an operating room tabletop that is attached to a table base. The Frame is used to support patients during spinal surgery. |  |  | February 14, 2007 | 2007 | Orthopedic Systems Inc | Material used to manufacture accessory bracket does not meet product specification. The screw is normally composed of ''Torlon'', which is yellow. The part was made of "Tercite", which is blue. | Other | A notification letter was sent to all 8 consignees on 02/16/2007. The firm will send service representatives to review any devices on site and replace with the proper part if needed. |  | 8 | 8 units distributed. | The products were distributed to 8 consignees in the United States, to ID, IL, NY, and CA. | Terminated | December 06, 2007 | 295 | 8 | Not_Computer
Z-0770-2007 | 37702 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 83 plus Mini-flex with accessory crate, Custom Ultrasonics, Inc. |  |  | December 12, 2005 | 2005 | Custom Ultrasonics, Inc. | The device was marketed without a 510k | PMA | On 12/12/05, the recalling firm telephoned their sales representatives to inform them of the problem. Letter were hand delivered to the hospitals by the sales representatives. |  | 4 | 4 units | Worldwide-The device was shipped to medical facilities in CA, GA, VA. and the UK. | Terminated | February 20, 2013 | 2627 | 4 | Not_Computer
Z-0771-2007 | 37813 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Nemio XG, ultrasound system, SSA-590A with version 1.5 software, Toshiba American Medical Systems, Inc., Tusin, CA. |  |  | December 19, 2006 | 2006 | Toshiba American Med Sys Inc | Transducer Surface Temperature : It was found during testing that when using the system in the 2DCDI + PW mode, with the Scale Graph function active, that the transducer surface temperature could exceed the 27 degrees Celsius limit as specified in the IEC safety standard. | Other | Notification letters were sent to customers on December 19, 2206. Customers were instructed to avoid using Scale Graph function with live images, and instructed to call up the Scale Graph after images are frozen. Customers were informed that a Toshiba Customer Service Engineer will contact them to schedule a software upgrade when software becomes available. If customers have any questions they were instructed to contact their Toshiba Customer Service Engineer or call (800) 421-1968 and ask for the Regulatory Affairs Department. Customers were also instructed to distribute this information to all users of the product and also with the Bio-medical Department at their facility. |  | 37 | 37 | Nationwide Distribution including the states of CA, FL, GA, ID, MA, MN, NC, NY, PA, RI, TN, TX, WA, and WI. | Terminated | February 15, 2012 | 1884 | 37 | Software | Physical Safety Hazards | Software update | Software Update
Z-0772-2007 | 37747 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | LumexÂż Walkabout Lite Four-Wheel Rollator, Manufactured for: Graham-Field Health Products, Atlanta, Georgia 30360. |  |  | April 02, 2007 | 2007 | Graham Field Atlanta | The caster bolt stem/screw may loosen over time causing the wheel to come off. | Other | Graham-Field notified all of its Durable Medical Equipment (DME) dealers involved in the distribution of these products by certified US Mail on April 3, 2007. They were made aware of the need to adjust the product to lessen the risk of injury and were provided a re-work kit. The distribution of DME dealers includes the announcement of the field correction to the DME dealer, the letter to send to the end-users, rework supplies, rework instructions and proof of compliance cards. Graham-Field instructed the DME dealers to adjust any impacted products on site, as well as, any delivered by end-users, customers or other impacted third parties to the DME dealer for correction. DME dealers will be compensated for performing the corrective action. The installation of the rework kit involves the installation of lock washers and Locitite and is described in more detail in the attached rework instructions. Graham-Field does not have direct access to end user names, however it plans to reach end-users by financially supporting a certified mailing from DME dealers to their individual customers notifying them of this field correction. A follow-up 60 day certified mailing will be performed for non-responders. |  | 13300 | 13,300 units | U.S. Nationwide, Costa Rica, Ecuador, El Salvador, Guatemala, and Santo Domingo | Terminated | January 12, 2010 | 1016 | 13300 | Not_Computer
Z-0781-2007 | 37819 | No | FWF | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SOPRO Camera Control Units, NTSC, SV420P; REF (Catalog) #72200505 v1.1; Made in France; Manufactured for Smith &amp; Nephew, Inc., Endoscopy, Andover, MA 01810. |  |  | January 19, 2007 | 2007 | Smith &amp; Nephew, Inc Endoscopy Division | Surgical camera control units were not designed and manufactured to pass the IEC 60101 testing for electrical fast transients (EFT) | Other | Recalling firm sent Recall notification letters to customers (Jan 2007) and distribution centers requesting device return. |  | 77 | 77 units. | Foreign Distribution Only-No U. S. A. distribution International distribution to: Argentina, Austria, Australia, Belgium, Brazil, China, Columbia, Denmark, Dubai, France, Germany, Greece, Italy, India, Japan, Korea, Malaysia, Netherlands, Norway, Portugal, South Africa, Spain, Sweden, Turkey and United Kingdom. | Terminated | September 02, 2008 | 592 | 102 | Not_Computer
Z-0783-2007 | 37855 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | B.Braun HyperFormer Pump Dispensing Set, Catalog Number V9900-02. The product is shipped in cases containing 10 sets. |  |  | April 11, 2007 | 2007 | B. Braun Medical, Inc. | The integrity of the sterile product barrier may be compromised | Other | An Urgent Product Removal Notice was issued to the distributors on 4/11/07. The notice informs the distributors of the problem, the need to conduct a sub recall and to return the product. |  | 180 | 180 units | Nationwide distribution --- including states of GA, MA, OH, and TX | Terminated | August 07, 2007 | 118 | 180 | Not_Computer
Z-0786-2007 | 37437 | No | JCW | Penile inflatable implant. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | AMS 700 CX MS Pump IZ Preconnected penile Prosthesis, product length 18 cm, product number 72404232. Penoscrotal Approach - with InhibiZone, 1 MS Pump, 2 Cylinders. |  |  | February 09, 2007 | 2007 | American Medical Systems | The carton labeling does not match the product packaged inside. | Other | AMS has contacted every account (via e-mail 2/9/2007) that either purchased or implanted the affected product. Unused pieces have been recovered and scrapped. |  | 5 | 5 | Worldwide, including USA, Canada and England. | Terminated | February 24, 2008 | 380 | 26 | Not_Computer
Z-0791-2007 | 37780 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i2000SR Processing Module, for in vitro diagnostics, List Number 3M74-01; manufactured by Abbot Laboratories, Irving, TX. |  |  | March 28, 2007 | 2007 | Abbott Laboratories, Inc | If the ARCHITECT i2000/i2000sr System becomes contaminated with microbes generating folate-like by-products, the ARCHITECT Folate assay may experience calibration failures and/or shifts in results. | Other | Notification Letters were sent to all customers who received the ARCHITECT Folate Reagent on 03/28/07. Customers were provided with instructions on how to conduct a Wash Buffer Unload procedure and were requested to perform this action daily. Customers were also asked to contact their local Abbott Customer Support Rep to schedule an Internal Decontamination procedure on the ARCHITECT System. Firm recalling to the user (laboratory) level. |  |  |  | Folate Reagent distributed to the following countries: Australia, Austria, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Dominican Republic, Finland, France, Greece, Hungary, Ireland, Israel, Italy, Jordan, Lebanon, Luxembourg, Netherlands, New Zealand, Norway, Oman, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Serbia and Montenegro, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Trinidad and Tobago, Turkey, UAE, United Kingdom. No US distribution. | Terminated | September 25, 2007 | 181 |  | Not_Computer
Z-0793-2007 | 37854 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | CIBASoftÂż Visitint Soft Contact Lenses, SKU #749591001385, Power -6.50, Diameter 13.8, Base Curve, 8.9, Expiration 2011-10. |  |  | April 30, 2007 | 2007 | Ciba Vision Corporation | Base curve of the lenses do not meet specification. | Other | Consignees were contacted via traceable mail on 4/30/2007. The initial notification was sent to Eye Care Professionals. The firm subsequently identified one distributor and one retailer. The recall letter was revised and sent to these consignees on May 2, 2007. Specific return instructions were attached along with a Recall Tracking Form that was to be completed and returned to CIBA Vision. They were instructed to return any of the affected product on hand to CIBA Vision for replacement. |  | 22 | 22 lenses | Nationwide. | Terminated | March 13, 2008 | 318 | 22 | Not_Computer
Z-0784-2007 | 37105 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare Engstrom Carestation Critical Care Ventilator with the High Performance Display Unit with SW version 3.x and 4.11; identification numbers 1505-9000-000 and M1057699 (upgrade). Datex-Ohmeda, Inc., a General Electric Company, doing business as GE Healthcare, P.O. Box 7750, Madison, WI 53707-7550, USA, www.gehealthcare.com. |  |  | February 02, 2007 | 2007 | Datex - Ohmeda, Inc | There are two unique conditions with the Engstrom Carestation that require corrective action. These include an anomaly in the control of the safety relief valve and a situation where the display unit may reset. | Other | Device Correction Letter to customers was mailed on 2/2/07 giving information about the correction and a short term recommendation until version 4.15 could be installed by a GE representative who would be contacting the customer. |  | 446 | 446 | GA, RI, WI, TX. Foreign - AUSTRALIA, BRUNEI DARUSSALAM, BRAZIL, COLOMBIA, COSTA RICA, SPAIN, FINLAND, FRANCE, UNITED KINGDOM, INDIA, ITALY, KOREA, MEXICO, MALAYSIA, PHILIPPINES, SAUDI ARABIA, SWEDEN, SINGAPORE, GERMANY. | Terminated | October 16, 2008 | 622 | 446 | Software | Display/Image | Software update | Software Update
Z-0777-2007 | 37710 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Safety-Lok Blood Collection Set with Preattached Holder. Color green. 21 Gauge. Translucent, integrated protective shield that provides one-handed activation immediately after use, 0.75 in needle length, 12 in tubing, Luer adapter and holder, wings, sterile. |  |  | March 20, 2007 | 2007 | Becton Dickinson &amp; Company | The BD Vacutainer Safety-Lok Blood Collection Set with Preattached Holder has a problem where the non-patient needle (in the holder) may become separated from the hub and remain the stopper of an evacuated blood collection tube. | Other | Recall letters were sent to customers on 3/11/2007 via a third party, West Inc., Appleton, WI who will coordinate the mailing to all consignees. |  | 147200 | 147,200 units | Nationwide and Canada. | Terminated | July 25, 2007 | 127 | 147200 | Not_Computer
Z-0795-2007 | 37805 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Gynecare PROLIFT Total Pelvic Repair System |  |  | April 20, 2007 | 2007 | Ethicon, Inc. | The pre-printed outer box incorrectly identifies the product as GYNECARE PROLIFT Total Pelvic Floor System (Product Code PFRT01). The contents of the box actually contain GYNECARE PROLIFT Anterior Pelvic Floor System (Product Code PFRA01) | Other | Recall notifications were sent by 2nd day air UPS delivery on April 20, 2007 to all hospitals Director of Surgery, OR Managers and Materials Managers. |  | 160 | 160 units to US | Nationwide | Terminated | March 04, 2008 | 319 | 160 | Not_Computer
Z-0794-2007 | 37837 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Medi-Tech Stainless Steel Greenfield Vena Cava Filter, order #50-400, lot #7826572 |  |  | October 06, 2006 | 2006 | Boston Scientific | One lot of product, Stainless Steel Greenfield Vena Cava Filter, may have the incorrect expiration date of 2008-07 and the correct date is 2008-06. | Other | Consignees were notified by letter (10/06/2006) which was accompanied by a Reply Verification Letter. All consignees were asked to cease use of product upon receipt of the letter and to fax the Reply Verification Letter to recalling firm and return product to firm. |  | 24 | 24 units | Nationwide. | Terminated | July 03, 2012 | 2097 | 24 | Not_Computer
Z-0799-2007 | 37064 | [6 empty] | Maxi Move Patient Lift with Scale; Manufactured by Arjo Hospital Equipment AB, Eslov, Sweden; Model KMBB4OSU2FUS |  |  | January 05, 2007 | 2007 | Arjo, Inc. | The hanger bar assembly may detach from the Maxi Move lifters with scale, due to an incorrect torque setting on the two bolts holding the scale in place. | Other | Arjo sent Urgent Device Field Correction Recall Notice letters dated 1/5/07 to the end user accounts who received the affected lifts, advising them of the potential for the hanger bar assembly to detach. They were informed that an Arjo Field Service Technician would call them within five business days to schedule the time for him to visit and perform the inspection and correction to the lifts involved. If they did not hear from the technician within that time, they were instructed to contact the Arjo Quality Department at 800-323-1245, ext. 6140. The accounts were also requested to complete and return the enclosed customer response form, acknowledging the receipt and understanding of the letter. |  | 121 | 121 lifts | Nationwide | Terminated | September 02, 2008 | 606 | 284 | Not_Computer
Z-0800-2007 | 37635 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | 3M ESPE Rely X Veneer Cement Refill Syringes, A3 Opaque/Yellow Opaque Shade, manufactured by 3M. ESPE is a radiopaque , color stable light-cured resin cement indicated for bonding veneers fabricated of porcelainor composite. |  |  | February 12, 2007 | 2007 | 3M Espe Dental Products | Cement in 3M ESPE RelyX Veneer Cement Refill A3 Shade syringes, may not cure to the degree expected. | Other | All consignees including distributors were notified of the recall via telephone or mail on 2/27/2007. Communication with distributors included the product being recalled, the problem , recommended stop usage or distribution and directions for return shippment. Further, the consignees were also sent a letter which described the product and problem and what to do if the product was still availabe or had been used. |  | 99 | 99 | Worldwide distribution ---- USA including states of IN, PA, NY, CA, NV, AZ and country of Malaysia. | Terminated | February 24, 2008 | 377 | 99 | Not_Computer
Z-0801-2007 | 37893 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Biohorizons 3.0mm Overdenture Implant System, REF 30120S2, (3 mm x12 mm-2mm collar), Biohorizons Implant Systems, Inc., Birmingham, AL 35243 |  |  | April 18, 2007 | 2007 | BioHorizons Implant Systems Inc | Dental implant system is missing ball attachment system. | Other | Letters explaining the problem and MBAS systems were sent via United Parcel Service (UPS) 3rd day package delivery to the consignees (both domestic and abroad) on 04/18/2007. |  | 34 | 34 units | Worldwide distribution ---- USA and country of Italy | Terminated | August 11, 2007 | 115 | 34 | Not_Computer
Z-0803-2007 | 37007 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE LightSpeed VCT Select / LightSpeed Pro32, Model 5126093, CT Scanner system. |  |  | December 21, 2006 | 2006 | GE Healthcare | Oil may leak onto the port window of the CT tube and may result in image artifacts that could lead to misdiagnosis. | Other | A Field Modification Instruction was issued to the field to inspect all systems and replace the tube window screws and the window filter, as well as clean any traces of oil in the tube window area. A 12/21/2006, Urgent Product Correction Notice/Customer letter has also been distributed informing the customers of the safety issue and the product correction. |  | 158 | 158 | Worldwide Distribution: USA in the following states and locations: AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI, Guam and Puerto Rico. Foreign countries include ; Albania, Algeria, Argentina, Austria, Australia, Belgium, Bulgaria, Brazil, Belarus, Canada, Chile, China, Columbia, Czechoslovakia, Denmark, Finland, France, Germany, Great Britain, Greece, Honduras, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Martinique, Mexico, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Reunion Island, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Venezuela, and Vietnam | Terminated | July 03, 2012 | 2021 | 1817 | Not_Computer
Z-0811-2007 | 37865 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Certain MicroMiniplant Straight Healing Abutment. Dental implant. Model numbers: ISMHA32, ISMHA33, ISMHA34 and ISMHA36. |  |  | April 25, 2007 | 2007 | Biomet 3i | The Straight Healing Abutment did not fully seat to the restorative platform of the implant, leaving a gap between the abutment and the implant. | Other | Each consignee was notified on 4/25/2007 by telephone using a prepared script. A follow letter was sent by fax. Each consignee was instructed to check their respective inventory for the affected product and to return a response card and any affected product to the firm. |  | 362 | 362 units | Distributed worldwide, including USA, Spain and the Netherlands. | Terminated | November 18, 2008 | 573 | 362 | Not_Computer
Z-0810-2007 | 37761 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche PTH (parathyroid hormone test system) for use in conjunction with Roche cobas e 411, cobas e 601, Elecsys 1010, Elecsys 2010 and Modular E 170 analyzers; catalog no. 11972103122. |  |  | April 19, 2007 | 2007 | Roche Diagnostics Corp. | Negative bias caused by hemolysis interference at the low end of the reference range could cause erroneous diagnosis of hypoparathyroidism. | Other | The firm issued a letter dated 04/19/07 to all users informing them of the change in the hemoglobin interference claim. |  | 14503 | 14,503 | Nationwide. | Terminated | June 11, 2008 | 419 | 14503 | Not_Computer
Z-0812-2007 | 37868 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | STA Compact and STA Compact CT |  |  | January 18, 2007 | 2007 | Diagnostica Stago, Inc. | Result of complaints of field failures of the silicon bridge rectifier, resulting in smoke billowing out of the unit and burnt circuit boards, wires and connectors. | Other | A recall notification has not been sent. The firm did issue a technical bulletin to their representatives.--Update: based on FDA input the firm issued a Recall Letter to all users dated May 25, 2007 |  | 693 | 693 units | Product was distributed to hospitals and laboratories nationwide. | Terminated | December 10, 2007 | 326 | 693 | Not_Computer
Z-0816-2007 | 37838 | [6 empty] | Cranex Tome panoramic x-ray unit with Spiral Tomography for cross-sectional imaging. manufactured by SOREDEX, PO Box 250, 00031 Helsinki, Finland. |  |  | April 28, 2006 | 2006 | Instrumentarium Dental, Inc. | There may be a material defect in the up/down movement axle of the unit. This may cause the upper part of the unit to suddenly come down. | Other | A recall letter, dated 04/28/2006, was sent to affected customers . The letter describes the issue, and requests that no patient operation is permitted until the unit is checked by authorized service technician. If the service technician noted wear of the screw, the assembly was replaced. |  | 9 | 9 | Worldwide, including USA, Canada, Bermuda, and Colombia. | Terminated | February 24, 2008 | 667 | 9 | Not_Computer
Z-0815-2007 | 37653 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare-Americas Signa 3.0T EXCITE Whole Body Magnetic Resonance System |  |  | June 01, 2006 | 2006 | GE Medical Systems LLC | To provide the 3.0T MRI systems users with the 8-channel brain coil the proper coil weighting software. The proper coil weighting factors for the eight-channel brain coil were not included in the software. | Other | The corrective action is already completed and all customers have been notified via a personal visit to each site, to install a software patch beginning June, 2006. |  | 30 | 30 units | Worldwide, including USA, China, Finland, Germany, Israel, Italy, Korea, Spain, and Taiwan. | Terminated | February 24, 2008 | 633 | 30 | Software |  | Software update | Software Update
Z-0814-2007 | 37136 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BacT/VIEWÂż Software, all versions. in vitro diagnostic database software. |  |  | December 19, 2006 | 2006 | bioMerieux, Inc. | BacT/VIEW will be unable to run scheduled macro applications unless the user touches the BacT/VIEW touch screen or clicks a function after resynchronizing the clocks. | Other | A customer notification was sent on/about December 19, 2006. Customers will be sent a new BacT/VIEW software version C.30b QTR 1 2007. |  | 1233 | 1233 units | Worldwide, including USA, Canada, UK, Italy, France, Spain, Germany, Poland, Chile, Japan and the Netherlands. | Terminated | December 08, 2009 | 1085 | 1233 | Software | Device Operation | Software update | Software Update
Z-0817-2007 | 37852 | ph buffer | JCC | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Medtronic Buffer Solution pH 1.07 500ml |  |  | April 06, 2007 | 2007 | Medtronic Neuromodulation | Medtronic pH 1.07 buffer solution lot #2610596 was mislableled. | Other | Customers were sent a recall notice, a customer detail sheet for this product lot, and a reply form to be used when returning the suspect bottles. Distributor was provided with a formatted templates to be used in recalling the supsect bottles from their customers and instructions on how to manage the returns and ship the bottles back to Medtronic. |  | 238 | 238 | Nationwide. | Terminated | March 02, 2008 | 331 | 238 | Not_Computer
Z-0821-2007 | 37848 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Definium AMX 700 Digital mobile x-ray system, GE Healthcare 3000 North Grandview Waukesha, WI |  |  | December 21, 2006 | 2006 | GE Medical Systems LLC | When using the Tech Switch on exposures &gt;750ms, &gt;100 kVp and &gt; 100 mAs the exposure terminates prematurely. | Other | Visit to each consignee beginning December 21, 2006. |  | 6 | 6 units | Product was distributed in the following states: MS, NE, NJ, and, TX | Terminated | February 24, 2008 | 430 | 6 | Not_Computer
Z-0709-2007 | 37385 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Precision Link Software (Version 2.1), (Version 2.2), (Version 2.3), (Version 2.4), (Version 2.5), and (Version 2.6); Serial Number(s): ALL; |  |  | December 18, 2006 | 2006 | Abbott Diabetes Care, Inc. | If the time and date are not set correctly on both the user's meter and computer, there is the potential for incorrect times and dates to be assigned to the results in the Precision Link Software. | Other | Firm initiated its recall on 12/21/2006. The firm provided customer letters with additional information (eg. flyer or insert) to customers that they had addresses for via fedex. In addition, the product notification was posted on the firm's internet website. Further, a flyer is to be included as a supplement to the user manual. |  | 67601 | 67,601 units | Worldwide Distribution: USA as well as abroad to Argentina, Aruba, Australia, Austria, Bahamas, Belgium, Canada, Colombia, Croatia, Curacao, Czech/Slovak Republics, Dominican Republic, Ecuador, El Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Ireland, Italy, Japan, Korea, Latvia, Lithuania/Estonia (Baltics), Mexico, Netherlands, New Zealand, Norway, Poland, Puerto Rico, Portugal, Saudi Arabia, Serbia, Bulgaria/Macedonia, Singapore, Slovenia, South Africa, Spain, St. Maarten, Sweden, Switzerland, Turkey, UK and Venezuela. | Terminated | December 20, 2010 | 1463 | 67601 | Other | Output/Calculation | User manual update | Other
Z-0826-2007 | 37843 | No | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BinaxNOW A&B; Kit 10 Test Kit (Japan) |  |  | March 28, 2007 | 2007 | Binax, Inc. dba IMPD | Tip of the foam swab could detach from the shaft causing injury to the patient. | Other | Inverness Medical issued recall notification by letter on 3/28/07. Customers are instructed to discard the NP swabs provided in the kits. The firm will replace the NP swabs. In the meantime customers are to use the tests with nasal wash/aspirate samples (influenza) or nasal wash samples (RSV). Response letters are to be signed by customers and returned to Inverness Medical. |  | 3814 | 3814 kits | Nationwide and Japan. | Terminated | September 08, 2008 | 530 | 13934 | Not_Computer
Z-0813-2007 | 37871 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 5ml Syringe Luer-Lok Tip |  |  | April 09, 2007 | 2007 | Becton Dickinson &amp; Company | Small cracks in plastic trays of one lot of BD 5mL Syringe Luer Lok Tip, Bulk Sterile Convenience Paks that may adversely affect tray integrity and steriity. | Other | Recall communications were sent on April 2, 2007 and May 3, 2007 by UPS second day air to all direct consignees and potential customers. |  | 174300 | 174,300 units | Nationwide. | Terminated | August 09, 2007 | 122 | 174300 | Not_Computer
Z-0807-2007 | 37756 | No | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Product is an EO Sterile, Polyester braided and coated nonabsorbable surgical suture with double needle, size 2/0 (3.0 Metric), NSH 26 mm TAPER Needle. Product is wrapped in a five-fold card with inserted flap ends (overwrap). A label is placed on the overwrap and the packet is then inserted into a tyvek/mylar pouch. One end of the overwrap card is printed with 'Open'. This end is positioned in the pouch so it is closest to the peel-open end. No suture is exposed or visible through the overwrap which must be torn to have access to the suture and needle. Product is boxed (12 to a box) and shrink wrapped. |  |  | April 03, 2007 | 2007 | C P Medical | Potential for sterility of product to be compromised. Suture package seals incomplete or missing. | Other | On 04/03/2007 the firm notified its Polyester product (Polybond) customer by telephone and followed up on 04/09/2007 with a letter via certified mail requesting the distributor to discontinue sales of the affected product lot, to notify their customers of the issue, and return stock to CP Medical. The firm provided the distributor with a letter dated 04/09/2007 to be provided to the distributor's customers requested those customers to discontinue use of the affected product lots and return stock to their distributor who in turn are to return product to CP Medical.On 04/20/2007 the firm notified its Polypropylene and Polydioxanone product customers by letter dated 04/17/2007 via certified mail requesting distributors to discontinue sales of the affected product lot, to notify their customers of the issue, and return stock to CP Medical. The firm provided the distributor with a letter dated 04/18/2007 to be provided to the distributor's customers requested those customers to discontinue use of the affected product lots and return stock to their distributor who in turn are to return product to CP Medical. |  | 14641 | 121 boxes (12/box) Lot I0625120, and 121 boxes (12/box) Lot I0627120-01 | Firm distributes products to human and veterinary medical device distribution centers worldwide. Polybond Polyester sutures, product CP523A, were distributed to one distribution center in South Africa.Polypropelene sutures, product 8665, were distributed to veterinary distribution centers in OH, OR, and NE for nationwide distribution.Polydioxanone sutures, product M398A, were distributed to human and veterinary distribution centers in AZ, IA, OH, LA, MA, MI, MO, NE, and TX for nationwide distribution. | Terminated | August 01, 2008 | 486 | 29569 | Not_Computer
Z-0838-2007 | 37900 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD AutoShield Pen Needle (Re order Number : 329300) |  |  | April 23, 2007 | 2007 | Becton Dickinson &amp; Company | Needle Sticks: Complaints were received regarding Needle Stick Injuries while administering insulin to patients. (Labeling will be revised to provide clear direction on product usage) | Other | April 23, 2007, Recall letters were sent via UPS 2nd day air and/or hand delivered to all direct consignees and potential customers of the product. |  | 496200 | 496,200 units | Nationwide :Distribution centers in NJ, MO, TX, OH, NM, IL, ND, FL, KS, VA, LA, and GA. | Terminated | April 16, 2008 | 359 | 496200 | Not_Computer
Z-0827-2007 | 37903 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Endura No-React Dural Substitute 2 cm x 10 cm |  |  | May 04, 2007 | 2007 | Integra LifeSciences Corp. | Integra LifeSciences Corporation d/b/a Integra NeuroSciences, the distributor of EnDura No-React Dural Substitute, manufactured by Shelhigh, Inc., Union, NJ, is recalling this product based on an FDA Public Health Notification and FDA requested recall letter to Shelhigh citing sterility and other manufacturing concerns regarding this product manufactured by Shelhigh. Integra is a distributor of t | Other | On May 4, 2007, an Urgent Recall Notice and accompanying recall acknowledgement form was sent to all customers that had purchased or been supplied EnDura No-React Dural Substitute product since sale of the product was initiated in 2003. Recall Coordinators at each facility are asked to review inventory for EnDura products and to call Integra Customer Service for an authorization number for return of product. |  | 12736 | 12,736 for all recalls Z-0827-0832-2007 | Nationwide Distribution to hospitals and medical centers | Terminated | September 24, 2008 | 509 | 12736 | Not_Computer
Z-0820-2007 | 37873 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROSÂż Immunodiagnostic Products Troponin I Reagent Pack, REF/Catalog No. 194 9882. A Troponin Immunoassay Method, IVD. Ortho-Clinical Diagnostics, Amersham UK |  |  | April 27, 2007 | 2007 | Ortho-Clinical Diagnostics | Quality control and patient results can be lower than expected when the true Troponin I value is less than or equal to 0.2 ng/mL which may result in the device failing to detect myocardial injury. | Nonconforming Material/Component | On 4/27/07, Ortho-Clinical Diagnostics in Rochester, NY, notified the end users of the recall of the product by phone, fax, and FedEx overnight letter. Clinical labs in possession of these lots have been contacted by OCD, and were instructed to (a) discontinue use of this product and (b) notify healthcare providers who ordered the test in recent weeks. Replacement product has been provided to clinical labs with the affected product lots. Clinical laboratories with questions may contact the company at 1-800-421-3311. On 4/27/07, OCD foreign affiliates were informed of this action by e-mail and instructed to inform their customers who received these reagent packs of this issue. Ortho-Clinical Diagnostics in Raritan, NJ, issued a nationwide press release on 5/04/07. |  | 5066 | Lot 3151: 1,750 units, Lot 3170: 3,316 units | Worldwide-including USA, Canada, and Bermuda | Terminated | September 07, 2012 | 1960 | 5066 | Not_Computer
Z-0834-2007 | 37701 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | FX Spare Shieldwire Inflators (3.0 - 5.0 mm), Catalog number 60042-02. The product is sold 5 per box. |  |  | March 28, 2007 | 2007 | Kensey Nash Corp | Insufficient wire sealing | Other | On 3/26/07, the recalling firm left voice mail message with all sales representatives instructing the sales representatives to contact all TriActiv FX accounts immediately. Recall notices were faxed on 3/28/07 to all accounts which received the product. Recall notices instructed the hospitals to quarantine all FX and to contact their sales representative to arrange for retrieval of the product if they had not already been contacted. |  | 109 | 109 units | Worldwide, including USA and Germany. | Terminated | August 30, 2007 | 155 | 547 | Not_Computer
Z-0818-2007 | 37856 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Ultrasite Horizon Pump IV Set. The product is shipped in cases containing 24 sets. |  |  | April 18, 2007 | 2007 | B. Braun Medical, Inc. | Secondary infusion backs up into primary container. | Other | The recalling firm issued Product Removal Notices on 4/20/07 via certified priority mail. The notice informed the customer of the problem and the need to return the product. |  | 2500000 | 2.5 million for all products | Nationwide and Canada. | Terminated | April 22, 2008 | 370 | 2500000 | Not_Computer
Z-0847-2007 | 37902 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Custom Kit, Left Heart Kit, REF/CAT No.: K09T-09576, Rx only, Merit Medical Systems Inc., South Jordan, Utah. |  |  | April 17, 2007 | 2007 | Merit Medical Systems, Inc. | Male luer in fluid administration sets may crack due to inadequate molding process | Other | Consignees were notified by letter on 04/17/2007 and instructed to quarantine and return any unused product. |  | 17 | 17 units | CA, IL, NJ, NY. No military, government or foreign distribution. | Terminated | May 05, 2008 | 384 | 41251 | Not_Computer
Z-0869-2007 | 37723 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova 4100/4100 IQ. Digital Fluoroscopic Imaging System (Vascular Angiography System) |  |  | March 26, 2007 | 2007 | General Electric Med Systems LLC | Measurement errors during the computation of vessel diameters while using the Stenosis Analysis software (also known as Quantitative Coronary Analysis function-QCA) using the auto-calibration function on Innova Digital (located in control room) or in-room InnovaCentral/Touch-Screen | Other | A letter dated Jan. 18, 2007, was sent to consignees describing the problem, product ,safety instructions and contact information. Software upgrades were promised for correction. |  | 70 | 70 | Worldwide Distribution: USA states including-AZ,CA,CO,DE,FL,GA,IL,IN, MD,MN,MO,NJ,NV,NY,OH,OK,OR,PA,SC,SD,TX, UT, VA,WA,WI, and countries of Canada, Belgium, China, Czech Republic, Finland, France, Germany, Hungary, India, Isreal, Italy, Japan, Korea, New Zealand, Norway, Russia, Spain, and United Kingdom. | Terminated | October 21, 2008 | 575 | 288 | Software | Output/Calculation | Software update | Software Update
Z-0854-2007 | 37605 | No | LZS |  |  | Ophthalmic | PMA | LADAR6000 Excimer Laser System |  |  | February 21, 2007 | 2007 | Alcon Refractive Horizons, Inc. | Observed 'central islands' in patients following custom myopia laser procedures using the LADAR6000 Excimer Laser. (On May 11, 2007, Alcon notified consignees that the device software programs for CustomCornea Myopia and CustomCornea Myopia with Astigmatism of the device would be deactivated) | Other | Consignees were notified of the recall via two letters. On Feb. 21, 2007, consignees were requested to immediately stop performing all custom cornea myopia procedures and to return an acknowledgement card indicating they have received the alert and understand the required actions. On May 11, 2007, Alcon notified consignees that the device software programs for CustomCornea Myopia and CustomCornea Myopia with Astigmatism of the device would be deactivated. After meeting with CDRH's Office of Compliance on August 27, 2007, the recalling firm committed to total product removal of the device. Notification will begin in Sept 2007 and will be made by personal visits to each consignee. |  | 102 | 102 | Worldwide Distribution-including USA, Czech Republic, Spain, Australia, Italy, Hong Kong, Vietnam, Singapore, South Korea, and Canada | Terminated | March 21, 2017 | 3681 | 102 | Other | Physical Safety Hazards | Remove | Remove or Replace
Z-0889-2007 | 37886 | No | LJE |  |  | Gastroenterology/Urology | 510(k) | Boston Scientific PCNL (Percutaneous Tract) Kit with Nephromax Balloon; |  |  | April 11, 2007 | 2007 | Boston Scientific Corporation | Mislabeled: Outer box unit has correct part number but product is incorrectly identiified as a Percutaneous Tract Kit with Nephromax Balloon. The box actually contains a Percutaneous Tract Kit with Amplatz Dilator and Sheath Kit | Other | Boston Scientific notified customers on 4/11/07 by letter via Federal Express mail. Users are requested to return product in inventory. |  | 15 | 15 kits | Worldwide Distribution; USA including states of OH, MS, NM, MD, and TX and Foreign country of Trinidad | Terminated | December 12, 2007 | 245 | 15 | Not_Computer
Z-0890-2007 | 37997 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Sienet MagicStore. Radiological Digital Image Communications System |  |  | May 02, 2007 | 2007 | Siemens Medical Solutions USA, Inc | System could crash and result in data loss. | Other | The recalling firm issued a Safety Advisory letter to their customers (May 2, 2007) to inform them of the problem. A company representative will be contacting the customers to arrange for replacement of the RAID component. |  | 19 | 19 units | The products were shipped to medical facilities in CA, CO, FL, MS, NH, NJ, NY, and WA. | Terminated | September 17, 2008 | 504 | 19 | Other | Device Operation | Replace component | Remove or Replace
Z-0873-2007 | 37793 | [6 empty] | Light Source 300 Watt Xenon Lamp used in the following models: |  |  | April 05, 2007 | 2007 | Vision Systems Group, A Div of Viking Systems | Incorrect power entry module assembled into unit may cause burning to insulation, compromised electrical emissions and premature light source failure | Nonconforming Material/Component | Viking notified 2 distributors and 1 end user by telephone on 3/15/07. Viking representatives performed on site visit to perform inspection and replace units if needed. |  | 4 | 4 | Nationwide to CA and NC | Terminated | June 07, 2007 | 63 | 4 | Not_Computer
Z-0896-2007 | 37806 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Surge Medical Solutions 15 Fr. (5.0mm) Retrograde Cardioplegia Cannula, self-inflating balloon 18 mm, handle stylet, sterile, REF CODE: RSH-M014S. |  |  | April 17, 2007 | 2007 | Surge Medical Solutions, LLC | The device may leak between the two lumens or at the stop cock, which could result in inaccurate coronary sinus pressure readings. | Other | Consignees were notified and requested to return the recalled products via recall letter dated 4/17/07. |  | 2211 | 2,211 for all Recalls Z-0892-0904-2007 | Worldwide, USA, Argentina, Germany, Ireland, Italy, Lebanon, Saudi Arabia, Turkey, and the United Kingdom. | Terminated | March 10, 2009 | 693 | 2211 | Not_Computer
Z-0906-2007 | 37857 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | c.cam Gamma Camera, Model 9VIR1200; Siemens Medical Solutions USA, Inc., Molecular Imaging, 2501 N. Barrington Road, Hoffman Estates, IL 60195-5203; Made in Denmark; |  |  | April 30, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Chair Breakage: Under extreme load, the chair back of the c.cam emission computed tomography system may bend and eventually break. This could be caused by the accumulated stress from heavy patients seating or stretching themselves in the chair or simply from long time use. | Other | Important Safety Notice letters dated 4/30/07 on Siemens letterhead were sent to the end users, informing them of the potential risk that the c.cam chair may bend or break under heavy load. They were told that a Siemens service representative will contact them within the next few days to arrange the repair of the chair at no cost to the end user. To ensure the chair does not become damaged, SIemens recommended that the c.cam not be used for patient scanning until the update has been performed on the chair. Any questions regarding the Important Safety Notice were directed to their service representative or Siemens' office at: 1-800-888-7436 (USA). |  | 413 | 413 cameras | Worldwide-USA, Canada, France, Germany, Israel, The Netherlands and Portugal. | Terminated | January 09, 2009 | 620 | 413 | Not_Computer
Z-0905-2007 | 37811 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | FramelessArray software, Version 1.0. ; Medical Device firmware incorporated with the RadioCamera Extracranial System, distributed by Varian Medical System, Palo Alto, CA. |  |  | March 14, 2007 | 2007 | Varian Medical Systems Oncology Systems | Localization error; with planning data transferred to the Optical Guidance Platform via DICOM RT; The software incorrectly computes the center of the CT volume, resulting in a potential axial error ranging from 0.3 mm to 1.5 mm, affecting both Fractionated and Stereotactic Radiosurgery (SRS) Treatments | Software design | The recalling firm notified consignees by certified letter on 03/14/07. The notification was flagged as "Urgent Medical Device Correction" and advised end users that a localization error may occur when using the FramelessArray and affect its accuracy for Frameless treatments. Users were instructed to verify software in the devices and if running FramelessArray version 1.0, to discontinue use of the software when Stereotactic Radiosurgery plans are transfered via DICOM RT. The notification offered that firmware would be upgraded during scheduled service visit. |  | 178 | 178 devices | Worldwide : The software was distributed to direct accounts in the US and internationally distributed to accounts in Canada, Australia and Spain. | Terminated | July 03, 2008 | 477 | 178 | Software | Treatment/Delivery/Therapy | Firmware update | Software Update
Z-0891-2007 | 37309 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CIC Pro Clinical Information Center, Central Station, Software Version 4.1. , GE Medical Systems Information Technologies, Inc., 8200 W. Tower Ave. Milwaukee, Wisconsin USA. (System hardware consists of a processing unit and a monitor) |  |  | January 24, 2007 | 2007 | General Electric Medical Systems Information Technology | Alarm and Reboot problems: Two separate issues could occur with the CIC Pro system when used with Telemetry. Each issue occurs in software version v4.1 only; A) the system may experience a loss of audible alarms when operating system memory resources reach 170MB and/or B) the system may intentionally or unexpectedly reboot . | Other | GE Healthcare sent a certified letter titled "Urgent Medical Device Correction" to customers on 3/13/07 that explained the short and long term solutions and that the company would be providing a software update that will rectify the issue specific to CIC Pro v4.1 software. |  | 5190 | 5,190 Units | Nationwide | Terminated | December 10, 2011 | 1781 | 5190 | Software | Alarm/Message | Software update | Software Update
Z-0872-2007 | 37792 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Well Leg Holder With Boot, part of the Supine Universal Hip Distractor System; Ref (Part) #72200632; Manufactured for Smith &amp; Nephew, Inc., Endoscopy Divsion, Andover, MA 01810. |  |  | November 01, 2006 | 2006 | Smith &amp; Nephew, Inc Endoscopy Division | The device was not manufactured to correct specification causing interference when used in conjunction with Patient Transfer Board. | Other | Firm notified consignees by letter and phone call on 11/1/2006. Consignees asked to return out-of-spec Well Leg Holder for a replacement. |  | 12 | 12 | Nationwide Distribution: USA including states of WI, CA, MO, IL, TX, MA, WA, NM, FL, NH, and SC and the UK | Terminated | June 16, 2008 | 593 | 12 | Hardware | Device Operation | Replace | Remove or Replace
Z-0841-2007 | 37876 | No | MIE |  |  | Cardiovascular | Contact ODE | CryoValve, Pulmonary Valve &amp; Conduit |  |  | April 17, 2007 | 2007 | CryoLife, Inc. | Donor tissue was released prior to information received regarding a pericardial effusion. | Other | Both physicians were contacted by telephone on 4/17 &amp; 18/2007. Both tissues were returned to CryoLife |  | 1 | 1 tissue | AZ, CO | Terminated | December 31, 2007 | 258 | 1 | Not_Computer
Z-0856-2007 | 37538 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Precision Xtra Blood Glucose monitoring system. Individual test meter and test strip kit. |  |  | March 05, 2007 | 2007 | Abbott Diabetes Care, Inc. | The front button covers may detach from the meter. | Other | The firm intends to send letters to all customers that have registered their units under the firm's warranty database and the firm's first consignees. In addition, this notification will be placed on the company website, with e-mail and phone contacts as needed. Letter dated March 5, 2007. |  | 300 | approximately 300 lots | Product was distributed worldwide. | Terminated | April 04, 2008 | 396 | 300 | Not_Computer
Z-0861-2007 | 37665 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Infant Nasal CPAP Cannula, Size 4; an Rx infant nasal prong continuous positive airway pressure (CPAP) cannula one inspiratory and one expiratory elbow connector; 10 units per case; Teleflex Medical, Research Triangle Park, NC 27709 U.S.A., Made in Mexico; The cannulas are sold as components in the following configuration: catalog No. 1694: Infant CPAP Set intended for use with mechanical ventilators, which includes one cannula with one inspiratory and one expiratory elbow connector, two 10mm to 7.5mm adapters, one knit cap and two 6 inch sections of Velcro securing tape; |  |  | March 23, 2007 | 2007 | Teleflex Medical | The luer-port in the expiratory elbow may be occluded, which interferes with a ventilator's ability to detect an accurate pressure reading. | Other | Teleflex Medical sent Urgent Medical Device Recall letters dated 3/23/07 via UPS 2nd Day Air to the direct accounts, informing them that the luer-port in the expiratory elbow may be occluded, which interferes with a ventilator's ability to detect an accurate pressure reading. The accounts were requested to cease use of the affected units, remove them from use/stock, quarantine and destroy them. If the accounts further distributed the product, they were requested to conduct a sub-recall of the affected units from their customers. The accounts were instructed to complete and return by fax the enclosed Acknowledgement &amp; Stock Status Form, indicating the number of units destroyed and if they are requesting replacement product. Any questions were directed to Linda Todd at 1-800-334-9751, ext. 4951. |  | 13530 | 13,530 for all products | Worldwide, including USA, Puerto Rico, Australia, Bolivia, Brazil, Canada, Chile, China, Colombia, Dominican Republic, El Salvador, Germany, Guatemala, Hong Kong, India, Israel, Kuwait, Malaysia, Mexico, New Zealand, Paraguay, Peru, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, and Venezuela | Terminated | November 16, 2009 | 969 | 13530 | Not_Computer
Z-0862-2007 | 37782 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Smith + Nephew, Birmingham Hip Resurfacing, Acetabular Cup W/ Impactor, 46 mm cup Ref.: 74120146 |  |  | March 13, 2007 | 2007 | Smith &amp; Nephew Inc | The carton label reads BHR Acetabular Cup 60mm contained a 58mm cup. | Other | The consignees (hospital CEOs) were notified of the problem and the recall by letter (Fed Ex) on 03/13/2007. |  | 1 | 1 unit | Europe. No domestic distribution. | Terminated | October 28, 2011 | 1690 | 5 | Not_Computer
Z-0839-2007 | 37821 | [6 empty] | Silver Speed Hydrophilic Guidewire, for placement of vascular catheters. Model # 103-0601-200. |  |  | April 05, 2007 | 2007 | EV3 Neurovascular | Two customer complaints were received from physicians in Japan indicating the hydrophilic coating was observed flaking from the guidewire while immersed in hydrating media following use during aneurysm embolization procedures. | Other | A recall plan was developed for 100% recovery of guidewires within affected lot numbers. The plan includes notification to U.S., European, and Rest-of-World customers and distributors. A recall notification letter was prepared 04/05/2007 which summarizes the coating problem and potential risks, identifies the affected lot numbers and quantities shipped to the customer, and provides instructions for return of unused guidewires. |  | 2120 | 2120 | Worldwide, including USA, Italy, United Kingdom, Switzerland, Germany, Belgium, South Africa, Canada, Slovenia, China, Austria, France, Spain, Jordan, Mexico, Portugal, Colombia, Greece, Brazil, Venezula, Ireland, Australia, Malaysia, Saudi Arabia, South Korea, Lithuania, Netherlands, Hungary, and United Arab Emirates | Terminated | August 23, 2012 | 1967 | 2120 | Not_Computer
Z-0840-2007 | 37849 | No | FFS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Smith &amp; Nephew, 420XL Xenon Light Source, 180 Watt; for endoscopic use. Catalog/Part #72200510; manufactured for Smith &amp; Nephew Endoscopy, Andover, MA 01810. |  |  | April 09, 2007 | 2007 | Smith &amp; Nephew, Inc Endoscopy Division | Several units were manufactured using a power supply that contains a 1 amp fuse where a 5 amp fuse is required. When the unit's power supply draws more that 1A, the fuse will fail causing the light to cease operating. The problem can occur anytime: from the initial turning on (powering up), to anytime during use. | Nonconforming Material/Component | Notification sent 04/09/07 via FedEx to distribution centers, sales reps and user facilities. Users asked to send light source to Smith &amp; Nephew GMBH Service Center in Germany for repair (contact Ute Dieling at +49 (0) 7462-208-313). |  | 67 | 67 units. | Product sent to distributors in Argentina, Austria, Chile, Denmark, Dubai, France, Greece, Hong Kong, India, Italy, Japan, Korea, Malaysia, Norway, South Africa, Spain, Sweden and United Kingdom. No US distribution. | Terminated | September 02, 2008 | 512 | 67 | Battery | Device Operation | Repair | Repair
Z-0868-2007 | 37826 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ProfemurÂż R, Revision Hip system, Proximal Body (part), STD 4, REF: PPW38354, Cementless, Titanium Plasma Spray Coated, Sterile, Rx only. Hip implant. |  |  | March 27, 2007 | 2007 | Wright Medical Technology Inc | The titanium plasma coating was found to have missing fragments. | Other | The firm sent out recall notices to implanting surgeons by letter dated 03/27/2007 (via FedEx) to notify the surgeons of the issue. The hospitals and distributors were also notified of the issue and given instructions for return of the product. |  | 4 | 4 units | Nationwide, Puerto Rico, &amp; Japan | Terminated | June 30, 2008 | 461 | 22 | Not_Computer
Z-0913-2007 | 37775 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova 2100 IQ Digital Fluoroscopic Imaging System (Versatile cardiac and vascular imaging system.). |  |  | February 05, 2007 | 2007 | General Electric Med Systems LLC | Unintended Gantry Movements : GE Healthcare has recently become aware of accidental, operated-initiated activation of gantry movements associated with the new, metallic Smart Box controller, that may impact patient or staff safely. | Other | Consignees were contacted via a letter "Urgent: Safety Notice" dated 2/5/07. The letter described the safety issue, affected product, safety instructions and product correction contact information. -This is the short-term fix, GE is working on a correction that will be installed at each site as the long term correction |  | 201 | 201 | Class II Recall - Worldwide Distribution --- USA including states of WV, WI, WA, VA, UT, TX, TN, SD, SC, PA, OR, OK,OH, NY, NV, NJ, NE, NC, MT, MO, MN, MI, ME, MD, MA, LA, KY, KS, IN, IL, ID, HI, GU, GA, FL, DE, CT, CO, CA, AZ, AL, AK and countries of UNITED KINGDOM, TURKEY, SYRIA, SWITZERLAND, SWEDEN, SPAIN, SAUDI ARABIA, RUSSIA, POLAND, NORWAY, NEW ZEALAND, THE NETHERLANDS, MOROCCO, MALAYSIA, KOREA, JORDAN, JAPAN, ITALY, ISREAL, IRELAND, INDIA, HUNGARY, GREECE, GERMANY, FRANCE, FINLAND , DENMARK, CZECH REPUBLIC, CHINA, CANADA, BELGIUM, AUSTRALIA, and ALGERIA. | Terminated | December 17, 2011 | 1776 | 446 | Not_Computer
Z-0916-2007 | 37788 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | AcuMatch M-Series Straight Fluted Stem Segment and AcuMatch M-Series Curved Fluted Stem Segment. Catalog numbers 150-07-19 and 150-07-44. |  |  | July 21, 2006 | 2006 | Exactech, Inc. | The packaging was mislabeled: Exterior and interior packaging on the device was labeled with catalog number 150-07-19 15mm x 200mm, however, the actual part inside the packaging was a 21mm x 200mm (catalog number 150-07-44) device. | Other | The consignees were contacted by fax and email (dated July 21, 2006) to inform them of the recall requesting return of product |  | 6 | 6 | World wide-Domestically distributed to one account in Alabama and internationally to Columbia and Australia. | Terminated | May 06, 2012 | 2116 | 6 | Not_Computer
Z-0917-2007 | 37795 | No | KLB | Tobramycin test system. | Toxicology | Toxicology | 510(k) | Roche Online TDM Tobramycin, for use on Roche/Hitachi analyzers 911,912, 917 and Modular P; Catalog No. 04642511190. (Note: Not distributed within the United States) |  |  | April 05, 2007 | 2007 | Roche Diagnostics Corp. | Reported results may have a positive bias of up to 113% because inadequate cleaning instructions may result in carry-over of the following reagents: Online TDM carbamazepine, Online TDM gentamycin, Online TDM phenytoin, Online DAT cocaine II, Online DAT methadone II, Online DAT opiates II and Online DAT THC II. | Other | Foreign affiliates were notified via email letter dated 4/5/07 requesting that to notify their customers of the need to implement a special wash cycle to prevent reagent carry over. |  | 474 | 474 total (Z-0917-0918-2007) | International Distribution ONLY-Australia, Canada and Germany. | Terminated | April 30, 2008 | 391 | 474 | Not_Computer
Z-0920-2007 | 37796 | No | LEG | Neuroleptic drugs radioreceptor assay test system. | Toxicology | Toxicology | 510(k) | Roche VALP2, Online TDM Valproic acid , for use with the Roche/Hitachi Cobas c 501 analyzer; Catalog No.04491041190. |  |  | April 30, 2007 | 2007 | Roche Diagnostics Corp. | Probe Wash: Reported results may have a positive bias because inadequate analyzer probe wash instructions may result in carry-over from the following assays: Creatinine Kinase L and Creatine Kinase-MBL. | Other | Consignees were issued a recall letter dated 4/30/07 instructing them to perform a R2 reagent probe wash when using CKL and CKMBL with the vancomycin or valproic acid assays. |  | 992 | 992 | Nationwide. | Terminated | December 17, 2007 | 231 | 1654 | Not_Computer
Z-0926-2007 | 37984 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | NeatNick Sweeping Action Heel Lancet -Full-Term- |  |  | May 11, 2007 | 2007 | Hawaii Medical LLC | Lancet may misfire and an inadequate blood sample may result; additionally, the blade may not retract after use and result in a needle stick to the user/patient | Other | Hawaii Medical notified the initial distributor and end users by telephone and follow-up Recall Notification Letter on May 15, 2007. |  | 286376 | 286,376 units | Nationwide | Terminated | July 22, 2008 | 438 | 344434 | Not_Computer
Z-0925-2007 | 37864 | No | MJK | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | LIFEPAK 20 defibrillator/monitor. The device is an AC powered defibrillator device with battery backup. The defibrillators are primarily used by healthcare providers in device user facilities such as hospitals and clinics. |  |  | April 27, 2007 | 2007 | Medtronic Emergency Response Systems, Inc. | White Screen Lock up: LifePak 20 defibrillators with software version 48 or version 52 may experience an intermittent timing issue during the power on sequence. This may lock up the device resulting in the monitor screen going white, and may cause a potential delay or prevention of shock treatment. | Other | On 04/27/2007 firm sent letters via return receipt to all domestic and foreign customers. The letter recommends for customers to turn the defibrillator power off and back on to restore normal operation in the event of a white screen condition. In addition, the letter advises that a representative from the firm will contact customer to make arrangements to provide a software upgrade to correct the issue. |  | 859 | 859 units | Distribution is Worldwide. Devices are used by healthcare providers in hospital or healthcare clinic settings. | Terminated | March 23, 2012 | 1792 | 859 | Software | Device Operation | Software update | Software Update
Z-0922-2007 | 36616 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | EP Healing abutment- Product Code-ITHA52-a pre-manufactured prosthetic component directly connected to an endosseous dental implant. |  |  | October 16, 2006 | 2006 | Biomet 3i | Package Integrity; sterility compromised-Incomplete seals on the nylon bags in which the product is packed. | Other | Each consignee will be contacted by email or telephone followed by a faxed letter. Each consignee was instructed to check their respective inventory for the recalled part and corresponding lot number. They were directed to return the product to the recalling firm. |  |  |  | Worldwide. | Terminated | December 30, 2008 | 806 |  | Not_Computer
Z-0933-2007 | 37818 | No | MCX |  |  | Cardiovascular | PMA | St. Jude Medical-CPS Direct SL- slittable CRT lead delivery tool. -Model Numbers 410110, 410111, 410112, 410113, 410114, 410115, 410116, 410120,410121. 410122, 410123, 410124, 410125 and 410126. |  |  | March 20, 2007 | 2007 | St Jude Medical CRMD | Device Separation- product may separate at the mapping collar transition joint located at the proximal (light blue) marker band end of the tool during the implant procedure. | Other | By Letter dated March 20, 2007, customers were requested to returned affected product immediately to St. Jude Medical. St. Judes Sales Representative will be assisting customers in retrieving the subject inventory from their shelves.The firm instructed customers that if they should require further information regarding this issue, to contact the firm's Technical Service department at (800) 722-3774 or customers can contact local St. Jude Medical representative. |  | 14000 | 14,000 | Worldwide-including states of AZ, CA DC, FL GA, IN, KS, LA, MA, MD, ME, MI, MN, MS, NC, NH, NY, OH, PA, RI, TN, VA, and WI and countries of Austria, Belgium, Denmark, Finland, Netherlands, Sweden, France, Germany, Italy, UK, Spain, Portugal, Poland, Switzerland, Australia, New Zealand, Thailand, Hong Kong, Korea, Malaysia and Singapore. | Terminated | February 29, 2012 | 1807 | 14000 | Not_Computer
Z-0855-2007 | 37570 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ARCHITECT STAT Troponin-I Reagent; list numbers 2K41-20 (400 test kit), 2K41-25 (100 test kit) and 2K41-30 (2000 test kit); each kit contains bottles of microparticles, conjugate and assay diluent; Abbott Laboratories, Abbott Park, IL 60064 |  |  | March 08, 2007 | 2007 | Abbott Laboratories MPG | False Negative Results : The analytical sensitivity claim of less than or equal to 0.01 ng/mL (ug/L) might not be met for all lots of Architect STAT Troponin-I reagents, thus false negative results may occur with negative results less than 0.1 ng/mL | Other | Abbott sent a Product Correction letter on 3/8/07 to all Troponin-I customers who received reagent lots in the last 12 months. The letters informed the accounts that the analytical sensitivity claim of |  | 9438 | 9,438 kits | Worldwide: USA including Puerto Rico, and internationally to Canada, Brazil, Germany, Japan, Taiwan, Hong Kong, Thailand, South Korea, Australia, New Zealand, Mexico, Colombia, Chile, Argentina and Trindad &amp; Tobago. | Terminated | July 16, 2008 | 496 | 9438 | Not_Computer
Z-0739-2007 | 37767 | No | IZW | Diagnostic x-ray beam-limiting device. | Radiology | Radiology | 510(K) Exempt | Philips Collimators, Model numbers: 9896 010 00614, 9896 010 00615, 9804 607 00301, and 9804 602 83501. Product is a Collimator/ Beam Limiting Device and are replacement components for the Bucky Diagnost TH X-ray system. |  |  | March 15, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | The Collimator field-defining light beam does not meet the light field contrast ratio requirement of 4:1. System is non-compliant with performance standard (1020.31 (d)(2)(iii) 21 CFR | Other | Firm issued letter on 03/20/2007 to consignees advising them that firm representatives will contact consignee to replace collimators at no charge. |  | 6 | 6 | Nationwide; Units were distributed to hospitals, radiologists, and health care clinics in IA, LA, ND, PA, and TX, and Washington DC. | Terminated | November 15, 2010 | 1341 | 6 | Not_Computer
Z-0960-2007 | 37671 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Synovis Peri-Strips Dry with Veritas Collagen Matrix Circular Staple Line Reinforcement,( PSD 21-E-V), 21mm, Ethicon Circular stapler. Each package contains the following components: Cartridge and anvil assemblies, Cartridge cone, PSD Gel, &amp; C-clip. |  |  | March 16, 2007 | 2007 | Synovis Surgical Inovation Div. of | Package Integrity/Sterility Compromised: The foil pouch in the Peri-Strips Dry with Veritas Collagen Matrix Circular Staple Line Reinforcement has a potential for minute defects which may compromise the sterility of the outside of the tray in which the device is packaged. | Other | Recall notification letters were sent to hospitals via Federal Express on 03/16/2007. This letter describes the product being recalled and problem. It requests consignees to complete and fax back the product return form. EXPANDED Another recall letter dated 04/23/2007 was sent to inform customers of 17 aditional lot numbers being recalled. A third recall letter dated 05/05/2007 was set to hospitals to inform them of one additional Lot/Serial number that was affected. This letter also requests a product return form be completed and faxed to Synovis Surgical Innovations. |  | 1722 | 1722 | Nationwide including states of AZ, CA, CT, FL, IA, IL, IN, KY, MA, MI, MN, MO, NJ, NY, OH, OK, PA, TN, TX, VA, and WI | Terminated | February 24, 2008 | 345 | 3902 | Not_Computer
Z-0934-2007 | 37847 | No | MHY |  |  | Neurology | PMA | Medtronic DBS Lead Kit for Deep Brain Stimulation, 3389S-40, Medtronic, Inc. Minneapolis, MN 55432-5604 |  |  | April 16, 2007 | 2007 | Medtronic Neuromodulation | Mislabeling: The inner package labeling incorrectly specified the model number of the lead as 3387-40. The outer package labeling is correct and the package actually contains the correct lead (3389S-40). | Other | Medtronic field personnel visited affected customers and retrieved product from this lot, unless the product was needed for an imminent scheduled surgery. Telephone scripts and leave behind letters to each consignee were initiated in April 2007. |  | 31 | 31 kits | Nationwide including states of CA, CO, IL, IN, MD, NC, OH, PA, and WV. | Terminated | February 24, 2008 | 314 | 31 | Not_Computer
Z-0936-2007 | 37869 | Yes | MVK |  |  | Cardiovascular | PMA | Electrode Belt, component of the LifeVest Wearable Defibrillator, Product number 10A0889-A01. Specifically the 3 Therapy Electrode Pads attached to the Electrode Belt. |  |  | April 27, 2007 | 2007 | Zoll Lifecor Corporation | The electrodes do not properly release conductive gel | Other | Patients currently using the product should continue to do so while the product removal is underway. Product replacements will begin with the belts currently in the possession of the sales reps. By replacing these first any new patients will be receiving new or refurbished belts with new Therapy Electrodes that have not been exposed to methanol. As patients complete their use of the LifeVest and return their device to the firm the belt's refurbishment process will include replacement of all 3 Therapy Electrodes. The firm expects this process to be completed in approximately 3 months based on the average use of time and current production capabilities. The process may take an additional month for devices distributed overseas. Physicians with active patients were notified via letter sent certified mail on 5/3/07. Long term use patients (patients using the device longer than 3 months) will be notified of the product removal via telephone and follow up letter. Each long term patient will be provided with a replacement belt. They will be instructed to return their existing belt in the same shipping box. Due to the low residual risk no notification is planned for short term use patients. Prescribing physicians have the option to notify their patients if they choose to do so. |  | 944 | 944 belts | Worldwide: The product was shipped to patients in AZ, CA, CT, FL, GA, IL, IN, KY, LA, MA, MI, MN, MO, MT, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WI, and WV. The product was also shipped to Germany, United Kingdom, Italy, Spain, France, and Denmark. | Terminated | March 06, 2008 | 314 | 944 | Not_Computer
Z-0921-2007 | 37774 | No | FRW | Patient scale. | General Hospital | General Hospital | 510(K) Exempt | IMS Patient Scale used on Sunrise Medical's Hoyer Presence and Stature Patient Lifts (Cradle), Models: HOY-4pt-WSC and HOY-StatureWSC |  |  | January 30, 2007 | 2007 | Integrated Measurement Systems, Inc. | Patient Scale may become detached, which can cause a patient to fall and result in injury. | Other | The scale manufacturer, IMS, in conjunction with the lift manufacturer, Joerns Healthcare (previously Sunrise Medical), will initiate a field fix strategy consisting of shipping replacement 4-point cradles without scales, for use until a replacement can be designed, tested and validated. Joerns will utilize their Customer Service Help Line (1-800-826-0270) on behalf of IMS, to inform customers of the voluntary field correction, address any questions regarding the IMS scale field fix strategy and to provide additional information on return of affected IMS Patient Scales. Customers were notified by telephone beginning on January 30, 2007. |  | 149 | 149 units | Worldwide | Terminated | March 02, 2008 | 397 | 149 | Not_Computer
Z-0937-2007 | 37905 | No | KNQ | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | ConMed Eliminator Esophageal PET Balloon Dilator; Outer Diameter (OD) size 8mm (24 Fr), length 8cm; product code/end item #000341; Single Use, Non-Sterile; distributed by ConMed Endoscopic Technologies, Billerica, MA 01821. |  |  | April 16, 2007 | 2007 | Conmed Endoscopic Technologies, Inc. | Mislabeling issues including; incorrect balloon diameters, lot number discrepancies, and incorrect type of dilator. | Other | Notification Letters were sent to domestic consignees on 04/16/07 and to international consignees on 04/17/07. Consignees instructed to discontinue product use, sub recall to accounts sold product, and to contact Stericycle, Inc. at 1-800-668-4391 for instructions on returning product. |  | 68 | 68 units. | Worldwide-Devices distributed to medical facilities and distributors Nationwide (USA) and to distributors in the following countries: Australia, Belgium, Canada, Croatia, Denmark, Germany, India, Italy, Japan, Mexico, Poland, Spain and Sweden. | Terminated | July 30, 2010 | 1201 | 12185 | Not_Computer
Z-0950-2007 | 37918 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i2000 Processing Module, for in vitro diagnostics; List Number 8C89-01; v2.10, v2.11, v2.12, v2.20, v2.20DB, v2.30, v2.60 and v3.00; distributed by Abbott Laboratories, Irving, TX. |  |  | April 30, 2007 | 2007 | Abbott Laboratories, Inc | One or more of four software issues : 1) calculation error for Architect cSystem assays using the Spline math model. 2) The system configuration option, Run controls onboard reagents by Kit, does not function correctly for assays requiring a standard sample dilution. 3) Infrequent data may be missing from automatically or manually printed reports generated on the Architect system. 4) Clinical Chem | Other | Notification letters sent to consignees on 04/30/07. Consignees provided with work around instructions for the software glitches and asked to respond via fax using Customer Reply form acknowledging receipt, understanding and completion of correction letter. This is the short term fix pending software updates, validation,and full root cause determination. |  | 170 | 170 units. | Worldwide distribution, including: USA, Argentina, Australia, Bermuda, Brazil, Canada, Chile, China, Columbia, Costa Rica, Dominican Republic, El Salvador, Germany, Guetamala, Honduras, Hong Kong, Jamaica, Japan, Korea, Mexico, New Zealand, Panama, Peru, Puerto Rico, Singapore, South Korea, Taiwan, Thailand, Trinidad/Tobago, Uruguay, and Venezuela. | Terminated | March 19, 2008 | 324 | 3344 | Computer
Z-0959-2007 | 37994 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Biomet 3 i, Disposable Twist Drills, Model Number DT2715. |  |  | May 14, 2007 | 2007 | Biomet 3i | The depth indicator laser line markings on the drills are too faint for clinicians to read during use. | Other | Each domestic consignee was contacted on 5/14/07 by phone using a prepared script. A follow up letter was then faxed to each consignee. International consignees were contacted by email followed by a faxed letter. Each consignee was instructed to examine their inventory for the recalled product. They were requested to complete and fax back a response form to Biomet and to return any affected product. |  | 147 | 147 | World Wide: USA states including: NY, MA, MD, NJ, IL, DE, RI, CA, MI, VA, KS and fL. Internationally to Brazil, Germany, France, Great Briton Japan, Greece and Italy. | Terminated | September 11, 2009 | 851 | 147 | Not_Computer
Z-0935-2007 | 37863 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | HUMIDIFLOW Respiratory Gas Humidifier; Manufactured by: Porous Media Corporation 1350 Hammond Road St Paul, MN 55110, Part Number HMK-RM |  |  | April 26, 2007 | 2007 | Porous Media Corporation | The Intake filter of the Humidiflow kits may became detached from the module and fall onto the compressor cooling fan during patient treatment. (If the filter were to come in contact with the fan blades this could stop the fan from operating and cause the concentrator to overheat and interrupt the flow of oxygen to the patient) | Other | Customers were notified via a letter dated April 26, 2007. The letter describes the problem and product. Included with the letter, is a correction kit and installation instructions. The letter also urges the consignee to take an action to fill the attached form to confirm that the new intake filter restaint has been installed. |  | 178 | 178 | WORLDWIDE DISTRIBUTION: USA including states of GA, MO, FL, UT, CO, TX, MT, MN, PA, OH, AL, SC, VA, KY, IA, MA, TN, IL, NM, UT, NJ, IN, and WI, and countries of CANADA, MEXICO, and JAPAN. | Terminated | December 10, 2011 | 1689 | 178 | Not_Computer
Z-0836-2007 | 36484 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK CR Plus defibrillator |  |  | October 05, 2006 | 2006 | Medtronic Emergency Response Systems, Inc. | Premature battery depletion due to electrical component failure | Nonconforming Material/Component | Per request from Japan, the firm initiated action on 10/5/06 replacing the devices in Japan with new devices. On 3/2/07 the firm issued a Safety Alert letter dated February 2007 to consignees (with the exception of Japan), advising that the defibrillators have an identified, lead-free component that is susceptible to failure within the first 250 days after manufacture. Failure of this component might cause a short that will drain the battery and eventually render the device inoperative. The letter further states that the firm strongly recommends that the consignee immediately and on a regular basis, inspect their defibrillator(s) per the LIFEPAK CR Plus/EXPRESS Operating Instructions - Section 5 Caring for the Defibrillator. The letter also provides a website to allow consignees to enter their device serial number(s) to determine if their LIFEPAK CR Plus/LIFEPAK EXPRESS are included in the action. This recall strategy is subject to modification pending the performance of devices remaining on the market. |  | 13741 | Total of CR Plus 13,741 units worldwide; amount does not include units distributed to Japan. | Distribution is worldwide. The device is designed for non-professional users and are primarily used in public access defibrillator environments such as airports, schools, and public libraries. | Terminated | August 08, 2013 | 2499 | 14747 | Not_Computer
Z-0955-2007 | 37919 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Angiojet XMI Rapid Exchange, catheters, 135 cm, 4 F, Single Use Only, Sterilized with Ethylene Oxide, POSSIS MEDICAL, INC. 9055 Evergreen Boulevard NW Minneapolis, MN 55433-8003 USA-- (XMI - Rapid Exchange Rheolytic Thrombectomy Catheter) is one component of the AngioJet Rheolytic Thrombectomy System (AngioJet System). |  |  | May 01, 2007 | 2007 | Possis Medical, Inc | Mislabeling: XMI (OTW) catheters were mislabeled as XMI-RX units and AVX catheters were mislabeled as Spiroflex catheters. | Other | Consignees were notified of the problem on May 1, 2007 via a Urgent Medical Device Recall Letter. The consignees were instructed to quarantine the product until a Possis Representative contacts them. The representative will examine the units and determine which ones to be returned to Possis. They were also asked to fill and return the attached form. |  | 45 | 45 | Nationwide including states of AK, AR, AZ, CA, CT, FL,IL, IN, KY, MD, ME, MI, MN, MS, NC, ND, NJ, NY, OH, OR, PA, SC, TN, TX, WA, and WI. | Terminated | July 15, 2007 | 75 | 114 | Not_Computer
Z-0957-2007 | 37953 | [6 empty] | BD Microtainer Tube with Glycolytic Inhibitor, Catalog Number 365992, Non Sterile, 200 (4 x 50), 25 degrees C, Becton Dickinson and Company, Franklin Lakes, NJ |  |  | May 17, 2007 | 2007 | Becton Dickinson &amp; Company | Mislabeling : box and bag labels are incorrectly labeled to indicate the product contains Sodium Fluoride/Potassium Oxalate, however the instructions for use and the product itself are correct in the additive indicated and used (Sodium Fluoride/Na2EDTA). | Other | Product Recall letters dated May 1, 2007. were sent to all distributors by first class mail. The letter requested all distributors to examine inventory for specific catalog and lot numbers. It also requested distributors to send a copy of the enclosed customer letter to all customers as a subrecall. BD has contracted a third party (West Inc., Appleton, Wisconsin) to coordinate the mailing to all customers, if the firm chooses to use the third party, or the distributor can send the letters themselves. |  | 142800 | 142,800 worldwide | Worldwide; USA, Canada, Australia, Chile, Columbia, Japan, Malaysia, New Zealand, Singapore, Bangladesh and Indonesia | Terminated | August 15, 2007 | 90 | 142800 | Not_Computer
Z-0967-2007 | 37493 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Steris QFC1728 and QFC1729 Quick Connects used in connections and processing endoscopes. The QFC1728 is used to process the Fujinon 250/270/450/470/490 Group 5 and 530/590 series GI Endoscopes w/o Water Jet. The QFC1729 is used to process 250/450 Group 5 and 530/590 series Endoscopes w/ Water Jet. |  |  | February 23, 2007 | 2007 | Steris Corp | The center insert of the adaptor #6 restrictor could become loose and/or come out of the housing and may affect the sterilization process of the attached endoscope | Process control | On February 22, 2007, customers were contacted by phone and were notified by recall letter dated February 23, 2007. The recall notifications indicate that Steris Account managers will visit each consignee and replace affect product. |  | 146 | 146 | Worldwide-The product was distributed to USA domestic consignees located in MT, CA, NJ, IA, MA, WY, SC, MO, SC, TN, ID, MT, CO, PA, NV, and TX, and to foreign consignees located in Great Britian and Qatar. | Terminated | March 01, 2012 | 1833 | 146 | Not_Computer
Z-0956-2007 | 37948 | No | NPQ | Factor V Leiden DNA mutation detection systems. | Hematology | Hematology | 510(k) | MagNA Pure LC Total Nucleic Acid Isolation Kit - Large Volume; Catalog # 03264793001. |  |  | May 14, 2007 | 2007 | Roche Diagnostics Corp. | False Positive Results: Product labeling states cross contamination is not an issue for high viral load titer samples, but the firm has determined that a potential sample contamination may not be detected at these levels and could result in false positive results being reported. | Other | Consignees were sent an Urgen Medical Device Correction letter dated 5/14/07 which instructs them to define and establish a lower detection limit, to use UNG to prevent carryover whenever possible, and to confirm low positive results through an independent experiment. A revised device correction letter dated 6/15/07 was subsequently sent to each consignee updating the reason for recall and associated risks. |  | 4201 | 4,201 | Nationwide. | Terminated | December 17, 2009 | 948 | 4201 | Not_Computer
Z-0971-2007 | 36955 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Fisher &amp; Paykel FlexiFix Nasal Mask, Part Numbers: HC405A, HC405SL, HC405U, HCHC406A, HC406SL, HC407A, HC407SL &amp; HC407U. CPAP masks. |  |  | November 24, 2006 | 2006 | Fisher &amp; Paykel Healthcare Inc | Plastic tabs incorporated into the connector components of these CPAP masks may break off if they are not cleaned in accordance with the Instructions for Use.In the event of a tab breaking off, it may enter the CPAP system air path and there is an extremely remote possibility that the broken tab may then enter the patient's nasal cavity or mouth. Masks manufactured since April 2006 do not feat | Other | Fisher &amp; Paykel Healthcare Inc. commenced sending out letters by mail to consignees the week commencing 4 December 2006 advising consignees to quarantine and destroy affected product. A faxback form was also be provided for consignees to advise Fisher &amp; Paykel Healthcare of quarantined and destroyed stock.The firm will supply consignees with replacement product as determined from completed faxback information.The firm will collate responses from consignees and will follow up with a letter to consignees who do not respond within a month of initial mail-out. The firm will then make telephone contact with any consignees who do not respond within a month of the second mail-out.Fisher &amp; Paykel Healthcare Inc. estimated the completion of this voluntary recall by April 2007. |  | 815239 | 815,239 total for all models, all lots | Nationwide | Terminated | March 26, 2012 | 1949 | 815239 | Not_Computer
Z-0981-2007 | 37950 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Gyroscan MRI |  |  | April 30, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | Potential for gaseous helium to improperly vent, resulting in rupture of the magnet vessel and venting of helium into the MRI site. | Other | On 4/30/07 Philips sent an Urgent Device Correction letter dated April 25, 2007 via certified mail to customers. The letter advises customers of the issue and that that a representative will visit the sites, inspect and correct the problem if indicated. |  | 293 | 293 units | Units were distributed throughout the US to 1,078 hospitals and medical centers. | Terminated | September 21, 2011 | 1605 | 1372 | Not_Computer
Z-0968-2007 | 37300 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic U-CLIP Anastomotic Device Removal Tool, model REM1860FW. The U-Clip Anastomotic Device Removal Tool is designed to remove U-Clip Anastomotic Devices. |  |  | February 08, 2007 | 2007 | Medtronic Cardiac Surgery Technologies | Six lots of U-Clip Removal Tools, catalog REM1860FW were incorrectly assembled. The nose cone attachment to the body of the removal tool does not contain sufficient adhesive which can cause the nose to separate from the body of the device. | Other | Customers to be notified via a letter on 2/8/07. The letter will explain the situation and requests that the impacted product be returned to Medtronic for disposal. The customer will be requested to complete a "Recall Certificate" acknowledging the receipt of this information. |  | 187 | 187 (US 155, OUS 32) | Nationwide and Canada. | Terminated | February 24, 2008 | 381 | 187 | Not_Computer
Z-0985-2007 | 37965 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Toshiba Aplio 80 Digital Ultrasound System, SSA-770A with version 6 software |  |  | February 02, 2007 | 2007 | Toshiba American Med Sys Inc | When using the LV measurement function and making numerous measurements the value displayed on the image screen may on occasion not update correctly. The values that are stored in MDA (measurement display area) are correct and no issue has been reported with these values. | Other | Notification letters were sent to customers on February 2, 2007. Customers were informed of the safety issue regarding the possibility of incorrect measurement values being displayed on the active image screen while performing multiple measurements and instructed to verify actual results against the values stored in the MDA (measurement display area). Customers were also informed that a Toshiba Customer Service Engineer will contact them to schedule a software upgrade when software becomes available. If customers have any questions they were instructed to contact their Toshiba Customer Service Engineer or call (800) 421-1968 and ask for the Regulatory Affairs Department. Customers were also instructed to distribute this information to all users of the product and also with the Bio-medical Department at their facility. |  | 260 | 260 | Nationwide | Terminated | April 18, 2012 | 1902 | 305 | Software | Display/Image | Software update | Software Update
Z-0740-2007 | 38120 | No | KYG | Ocular surgery irrigation device. | Ophthalmic | Ophthalmic | 510(K) Exempt | Disposable Bimanuals Irrigation and Aspiration Handpiece Set: Item #1719 Disposable Bimanuals Irrigation/Aspiration Set 19 Gauge, Dual Port; Item #1719-A - Disposable Bimanual Aspiration Handpiece 19g; Item #1719-I Disposable Bimanuals Irrigation/Aspiration Set, 19g, Open End; Item #1719-S Disposable Bimanuals Irrigation/Aspiration Set, 19g, Dual Port, Smooth. For use in cataract surgical procedures. |  |  | April 09, 2007 | 2007 | Oasis Medical Inc | Plastic shavings are coming from the producthandle during irrigation. Use of this product may leave plastic shavings in the patient's eyes. | Other | An April 9, 2007, notification of the recall will be sent to distributors and users. They are requested to notify down to the user (physician) level. Included in the RecallNotification Letter are instructions: "In addition, if you may have further distributed this product,please identify your customers and notify them at once of thisproduct recall. Your notification to your customers may beenhanced by induding a copy of this recall notification letter.This recall should be carried out to the user level. Yourassistance is appreciated and necessary to prevent anyinconvenience." |  | 3050 | 3050 total for all item numbers | Worldwide, including USA, United Kingdom, Canada, Spain, Turkey, Kuwait, Colombia, Italy, Greece, Australia, Switzerland, South Africa, Germany, Portugal, and France. | Terminated | October 06, 2011 | 1641 | 3050 | Not_Computer
Z-0977-2007 | 37764 | No | NHG | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche Cortisol reagent for use in conjunction with Roche cobas e 411, cobas e 601, Elecsys 1010, Elecsys 2010 and Modular E 170 analyzers; Catalog no. 11875116122. in vitro diagnostic. |  |  | May 15, 2007 | 2007 | Roche Diagnostics Corp. | The test is not as sensitive, or as precise between runs, as claimed in the labeling. Low concentrations are observed only in saliva samples. | Other | The firm notified consignees via letter dated 5/15/07 and notifed them of the changes to functional sensitivity and between-run precision claims. |  | 6051 | 6051 | Nationwide. | Terminated | June 11, 2008 | 393 | 6051 | Not_Computer
Z-0976-2007 | 37651 | No | IZH | Mammographic x-ray system. | Radiology | Radiology | 510(k) | Diamond Mammographic X-Ray System, type MGX-2000, model 5111771. GE Healthcare P.O. Box 414, Milwaukee, Wisconsin 53201 U.S.A. |  |  | November 16, 2005 | 2005 | General Electric Med Systems LLC | GE Healthcare has become aware of the absence of thread-locking compound associated with the C-arm fasteners of your Diamond mammographic x-ray system that may impact patient safety. | Other | A Product Safety Notification, dated 05/24/07, was sent to Hospital Administrators, Managers of Radiology and Radiologists. This letter describes the safety issue, instructions to continue the use of the product and states that the device will be corrected by GE Healthcare service engineers. |  | 106 | 106 | Worldwide, including USA, Germany, France, Italy, South Africa, Finland, Norway, Sweden, Venezuela, Benelux, Japan, China, and the UK. | Terminated | February 24, 2008 | 830 | 106 | Not_Computer
Z-0980-2007 | 37872 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Proximate PPH, Procedure for Prolapse and Hemmorroids Set, REF PPH03. Each set contains: 1 hemorrhoidal circular stapler, 1 suture threader, 1 circular anal dilator, and 1 purse string suture anoscope. |  |  | April 02, 2007 | 2007 | Ethicon Endo-Surgery Inc | The adhesive used to bond the casing of the device may have migrated to the staple driver. While the device still fires, it may result in audible and tactile feedback at the midpoint of the firing cycle which may be interpreted by some surgeons as a completed firing cycle, which may cause the surgeon to prematurely stop firing the device and result in incomplete staple formation or cut line. | Other | The recalling firm sent a recall letter dated 04/02/07 to foreign consignees and a second letter dated 04/03/07 to their United States consignees. The letter asks that the customers: 1) Examine their inventories for the presence of the affected product and remove it from inventory if found; 2) Complete the Business Reply Form and fax it to Stericycle their representative and 3) Return all affected product in order to received credit, along with the completed Business Reply Form in the box identified with a pre-printed UPS shipping label supplied by the recalling firm. |  | 23966 | 23,966 units | Worldwide, including USA and Canada. | Terminated | December 09, 2009 | 982 | 23966 | Not_Computer
Z-0978-2007 | 37803 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | LINAC Scapel Treatment Planning System, incorporating FastPlan software versions 4.0, 5.0, 5.0.1, 5.1, 5.1.1. |  |  | March 09, 2007 | 2007 | Varian Medical Systems Inc | A software anomaly in a medical device used for patient treatment could result in accidental or unintended exposure radiation exposure. The tilt rotation compensation algorithm does not adequately compensate for the tilt rotation. | Software design | The recalling firm notified medical device end users by letter dated 03/09/07 and advised of the software anomaly and provided details for user corrective action pending software upgrade by the manufacturer. |  | 80 | 80 devices | The medical devices were distributed to medical centers nationwide and worldwide. | Terminated | February 28, 2012 | 1817 | 80 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0929-2007 | 37241 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 100V -120V, 15 A (circuit breaker), 50/60 Hz (20A power source required) base; Model 801763. |  |  | February 01, 2005 | 2005 | Terumo Cardiovascular Systems Corp | System 1 power supply may electrically short to the main console, disrupting normal system operation by causing unexpected re-setting or shutdown of pumps and modules, false alarms and/or overheating and failure of the central control monitor. | Other | The firm corrected the complaint units at the time of the complaints in 2005. The firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and advising them that their system will be corrected when replacement components become available. |  | 416 | 416 for both models | Worldwide: USA, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Malaysia, Mexico, Pakistan and Singapore. | Terminated | July 07, 2009 | 1617 | 416 | Hardware | Device Operation | Replace components | Remove or Replace
Z-0924-2007 | 37244 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 6 inch diameter Roller Pump; Model 801041. |  |  | April 01, 2004 | 2004 | Terumo Cardiovascular Systems Corp | A belt slip message may be displayed on the roller pump due to low belt tension or the presence of grease. | Other | The firm is visiting consignees and making corrections as complaints are received. These corrections, to date, have been made between 04/04 and 09/06. The firm will correct other units as complaints are received and issued a Device Correction Letter dated 6/6 /07 notifying all users of the problem and appropriate responses should it occur. |  | 525 | 525 total for both models | Worldwide: USA, Belgium, Canada, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Korea, Malaysia, Singapore and United Arab Emirates. | Terminated | June 11, 2008 | 1532 | 525 | Not_Computer
Z-0932-2007 | 37242 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 4 inch diameter Roller Pump; Model 801040. |  |  | August 01, 2005 | 2005 | Terumo Cardiovascular Systems Corp | The tube clamp mechanism on the roller pump may not open when the rotation knob is turned, making it difficult to remove and/or insert tubing. | Other | The firm is visiting the consignees as complaints are received and making corrections. Corrections were made in August 2005 and April 2006 to date. Other units will be repaired if complaints are received and the firm issued a Device Correction Letter dated 6/6/07 notifying all users of the problem and appropriate action to take should it occur. |  | 1982 | 1982 total for both models distributed | Worldwide; USA, Australia, Belgium, Canada, Dominician Republic, Egypt, Germany, Hong Kong, India, Indonesia, Japan, Korea, Malaysia, Mexico, Pakistan, Republic of Georgia, Singapore, South Africa, Taiwan, Thailand and United Arab Emirates. | Terminated | February 23, 2008 | 936 | 1982 | Not_Computer
Z-0852-2007 | 38279 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE OEC 2800 mobile Fluoroscopic X-ray System, GE Healthcare, Salt Lake City, Utah. |  |  | June 12, 2007 | 2007 | GE OEC Medical Systems, Inc | X-ray systems were not configured to display Air Kerma Rate and Cumulative Air Kerma. | Other | Consignees were notified by letter on 06/12/2007 and told that a service representative would arrange for an inspection and would modify their systems. |  | 25 | 25 units | Worldwide Distribution: USA including states of FL, GA, PA. CO, MD, VA and WA. and countries of Australia, Austria, Belgium, Brazil, Canada, Colombia, Finland, France, Germany, Hong Kong, India, Israel, Italy, Japan, Korea, Netherlands, New Zealand, Norway, Portugal, Switzerland, Taiwan, and the UK. | Terminated | November 01, 2011 | 1603 | 575 | Not_Computer
Z-0738-2007 | 38233 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Kodak Direct View DR7500 System, Catalog #s 855-1046, 879-1345, 161-9139 --- A stationary X-Ray System. Device Listing # E222512. |  |  | March 06, 2007 | 2007 | Carestream Health, Inc. | The x-ray systems do not indicate the field size when the system's detector is turned off and the collimator is in manual mode. This results in the systems being noncompliant with the field indication performance standards of section 1020.31(e)(2)&amp; (3) of 21Code of Federal Regulations (CFR). | Other | On 2/21/07, Eastman Kodak notified CDRH of its proposed Corrective Action Plan CAP). The CAP was approved by CDRH under the following conditions: (1) Kodak will visit each user facility, and apply dial indicator labels to the collimator knobs, apply a field size label to the collimator housing, and verify that each system complies with the field indication performance standard. (2) Kodak will prepare adequate instructions for the assembly/installation and testing of the dial indicators labels and field size labels. These instructions are to be furnished to all owners of the Kodak Direct View DR 7500 Systems. (3) Kodak will supplement the Kodak DR 7500 system product report with a copy of the adequate instructions. (4) Kodak will perform 100% verification of the recall effectiveness. (5) Kodak will not charge for the corrections and will reimburse the owner for any out-of-pocket expenses. (6) Kodak will notify the manufacturer of the Ralco 302 ACS collimators, Ralco, that the collimator is noncompliant with the performance standard. Kodak has reported that the CAP was completed as of 5/07/07. |  | 28 | Domestic: 28 units | Nationwide: CO, IL, MI, MN, MS, NY, OK, PA, SC, TN, WA | Terminated | June 29, 2007 | 115 | 28 | Not_Computer
Z-0928-2007 | 38272 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC Miniview 6800 mobile fluoroscopic x-ray system, GE Healthcare, Surgery, Salt Lake City, Utah. |  |  | May 24, 2007 | 2007 | OEC Medical Systems, Inc | X-ray systems did not have the required labeling to indicate Air Kerma Rate and cumulative Air Kerma during fluoroscopic procedures | Other | Consignees were notified by letter on 05/24/2007 and advised that a service representative would be contacting them for installation of the required label and to provide them with an update to their user's manuals. |  | 811 | 811 systems | Worldwide; USA including VA facilities in AK, CA, ID, MD, MI, NY and OR. Military distribution to CA, CO, CT, FL, GA, KS, LA, MS, OH, OK, TX, VA and Iraq. Foreign distribution to Austria, Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Hong Kong, Israel, Italy, Japan, Korea, Malaysia, Netherlands, Norway, Portugal, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan and UK. | Terminated | June 02, 2008 | 375 | 811 | Not_Computer
Z-0975-2007 | 37526 | No | GIO | Capillary blood collection tube. | Hematology | Hematology | 510(K) Exempt | SafeClinitubes, REF 942-892, D957P-70-100x1, contained in kits containing 250 SafeClinitubes, insert, mixing wires and end caps. Capillary samplers for drawing small amounts of blood for blood gas analysis. |  |  | June 22, 2006 | 2006 | Radiometer America Inc | The clinitubes are not able to be filled due to contamination with silicone during production. | Other | The firm notified their customers of this recall with a letter dated 6/22/2006. The letter informs the customers of the problem with the affected lots of SafeCLINITUBES and asks that the customers discard any remaining boxes of affected product on-site. The recall letter is accompanied by a FAX FORM (response form) which the customer is asked to complete in order to request replacement product to be sent to their individual locations. |  | 175 | 175 boxes | Nationwide and Canada. | Terminated | May 08, 2012 | 2147 | 175 | Not_Computer
Z-0853-2007 | 37746 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | STA Liatest VWF:Ag, Product Catalog Number : 0518. This is a Immuno-Turbidimetric Assay of von Willebrand Factor. |  |  | March 13, 2007 | 2007 | Diagnostica Stago, Inc. | A dose-hook effect is observed starting from VWF:AG levels of 300%, while the package insert indicates that no dose-hook effect has been observed with VWF:Ag levels up to 800%. | Other | Customers were notified on March 13, 2007 by letter, first class mail. |  | 1098 | 1,098 units | Hospitals and laboratories Nationwide | Terminated | January 16, 2009 | 675 | 1098 | Not_Computer
Z-0979-2007 | 37834 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | STA Neoplastine CI Plus (10), (cat. 0667). |  |  | March 30, 2007 | 2007 | Diagnostica Stago, Inc. | Based on the confirmed clotting times in the upper part of the assigned ranges of STA-Coag Control ABN control plasmas when used with STA-Neoplastine CI Plus, use of this lot may cause difficulty during the validation of the STA-Coag Control ABN. | Other | Recall notification letters were sent on March 30, 2007 by first class mail. The letter advises customers that lot 061781 of STA-Neoplastine Cl Plus may cause difficulty during the vailidation process of STA-Coag Control ABN control plasmas, and asks that the lot number be discontinued. |  | 170 | 170 units | Nationwide. | Terminated | November 30, 2009 | 976 | 170 | Not_Computer
Z-1041-2007 | 37969 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ACCELERATOR APS System Input/Output Module (IOM); List Number: 7L01-01; manufactured by Inpecco SpA, Segrate (MI), Italy. |  |  | May 15, 2007 | 2007 | Abbott Laboratories, Inc | Product's hinged lid can fall, if not fully opened as designed, representing a potential injury hazard to the operator. | Other | Firm notified consignee via letter on 05/15/07. Consignee informed of potential for injury and advised that a label would be affixed to the device lid by an Abbott representative. |  | 1 | 1 unit. | Device distributed to a hospital in UT. | Terminated | November 27, 2007 | 196 | 1 | Not_Computer
Z-0958-2007 | 37957 | No | LER | Theophylline test system. | Toxicology | Toxicology | 510(k) | IMx Theophylline Reagent Pack (List 1A81-20) In Vitro test. |  |  | April 11, 2007 | 2007 | Abbott Diagnostic International, Ltd. | A decrease in mP values (calibration curve Net Polarization value) has been observed with this reagent lot at all calibrator levels and may produce the following results: Calibration Error Code 148 (Polarization Too Small), Calibration Error Code 172 (Calibration Not Stored or Calibration Not Accepted) Controls out of range (when using a stored calibration curve). | Other | An April 11, 2007, Product Recall Letter was sent to all customers who received the affected lot. All customers who received the affected lot will be contacted. The following communication types were used: Product Recall Letter &amp; Customer Reply Form. Letters were sent to customers by Federal Express Priority mail. The Customers were instructed the following: To discontinue use of and destroy IMx Theophylline reagent lot 41610QI00. If they had forwarded the lots listed in the Product Recall letter to other laboratories, they were instructed to inform them of this Product Recall and provide a copy of this letter to them. They also we instructed to retain a copy of the communication for their records |  | 17 | 17 Kits | Worldwide: CA, IL, MA, NY, OR, PA, VA, and WV and countries of Germany and Singapore | Terminated | June 10, 2011 | 1521 | 17 | Not_Computer
Z-1006-2007 | 37832 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | ResMed S8 Continuous Positive Airway Pressure (CPAP) Flow generator with the following model names and product codes: S8 AUTOSET VANTAGE, #33112; S8 COMPACT, #33030; S8 ELITE, #33021; S8 ESCAPE, #33007. |  |  | April 23, 2007 | 2007 | Resmed Corporation | S8 devices manufactured between July 2004 and May 15, 2006 are susceptible to fatigue in the connection of the AC appliance inlet connector and the power supply (power cord) leading to possible failure (short circuit, electrical failure, sparks, smoke). | Other | ResMed began to broadly distribute information regarding the recall on 4/23/2007. A press release was issued on 4/23/2007 with information regarding the recall and information was posted at ww.resmed.com/s8program. The firm is providing a Consignee notification package and press release to include(1) specific device names subject to the recall and affected serial numbers; specific information to be provided in each package to the end-customer. The specific information to the end user customers will include: (1) specific device names subject to the recall; (2) affected serial numbers; (3)continuing use and safety information; and (4) contact information for requesting additional details. Patients are advised that they may continue to use their S8 flow generators until a replacement unit is provided. ResMed is instructing patients, sleep labs, and hospitals to continue using the device, placing it on a hard clean surface, with a clear area around the device. ResMed intends to replace the affected products with either a new device or a device in which the affected components have been replaced. |  | 300000 | 300,000 | Worldwide, USA, Canada, and Mexico. | Terminated | December 06, 2010 | 1323 | 300000 | I/O | Device Operation | Replace device | Remove or Replace
Z-1001-2007 | 38118 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Ascensia Contour Blood Glucose Monitoring System ( Bayer )Diabetes Care System; Product 7152A.---NO USA Units are affected---. |  |  | June 06, 2007 | 2007 | Bayer Healthcare LLC | Incorrect Unit of Measure displayed : The meters report results in the wrong unit of measurement for Canadian users. | Other | The firm began calling Canadian customers on or about 6/06/07 to inform them of the problem, to tell them that a meter with the correct measurement units will be sent to them, and to request that they return the meter currently in their possession. |  | 16 | 16-Canadian Units only are affected | Canada. | Terminated | November 30, 2007 | 177 | 16 | Other | Output/Calculation | Return | Remove or Replace
Z-0987-2007 | 37888 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | HD11XE Diagnostic Ultrasound System |  |  | April 25, 2007 | 2007 | Philips Medical Systems | Potential for erroneous measurements which may lead to incorrect diagnosis based upon the presence of an inaccurate, undetectable error in clinical measurement or calculation. If the user exits an imaging mode while Active Annotation is in use, the system can enter an undefined state where the image scale calibration may not match the scale of the on-screen measurements and graphics and possib | Other | On 04/25/2007 firm sent letters dated 04/23/2007 to consignees informing them of the issue and that a new software release to correct the issue became available 04/19/2007. The firm's Field Service Representatives will contact consignees to install the new software at no charge to the consignee. |  | 566 | 219 units in US and 347 units distributed internationally | Worldwide. | Terminated | January 24, 2011 | 1370 | 566 | Software | Device Operation | Software update | Software Update
Z-0717-2007 | 38248 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Kodak Direct View DR3000 System, Stationary X-ray System, Catalog #s 655-1345, 655-1386. |  |  | March 06, 2007 | 2007 | Carestream Health, Inc. | The x-ray systems do not indicate the field size when the system's detector is turned off and the collimator is in manual mode. This results in the systems being noncompliant with the field indication performance standards of section 1020.31(e)(2)&amp; (3) of 21Code of Federal Regulations (CFR). | Other | On 3/06/07, Eastman Kodak notified CDRH of its proposed Corrective Action Plan (CAP). The CAP was approved by CDRH under the following conditions: (1) Kodak will visit each user facility, and apply dial indicator labels to the collimator knobs, apply a field size label to the collimator housing, and verify that each system complies with the field indication performance standard. (2) Kodak will prepare adequate instructions for the assembly/installation and testing of the dial indicators labels and field size labels. These instructions are to be furnished to all owners of the Kodak Direct View DR3000 Systems. (3) Kodak will supplement the Kodak DR3000 system product report with a copy of the adequate instructions. (4) Kodak will perform 100% verification of the recall effectiveness. (5) Kodak will not charge for the corrections and will reimburse the owner for any out-of-pocket expenses. (6) Kodak will notify the manufacturer of the Ralco 302 ACS collimators, Ralco, that the collimator is noncompliant with the performance standard. |  | 9 | Domestic: 7 units, Canada: 2 units | FL, NY, MO, MI, TX, and Canada | Terminated | July 02, 2007 | 118 | 9 | Not_Computer
Z-1010-2007 | 38170 | holder, head, radiographic | IWY | Radiographic head holder. | Radiology | Radiology | 510(K) Exempt | BrainLAB Target Positioner for Leksell Headring used with the iPlan RT Dose (versions 3.0, 3.0.1 and 3.0.2) and BrainSCAN (version 5.31) Radiotherapy Treatment Planning Software, ; BrainLab AG, Ammerthalstrasse 8, 85551 Heimstetten, Germany; catalog/version number 40700-3A |  |  | June 05, 2007 | 2007 | Brainlab AG | Target Alignment Error : A mechanical component/software incompatibility caused by a specific combination of the BrainLAB Target Positioner version 40700-3A for Leksell headrings and BrainLAB planning software resulted in a 1.25 mm shift in target area alignment during radiosurgery treatment. | Other | BrainLAB immediately notified all customers with the BrainSCAN or iPlan RT Dose software in combination with the BrainLAB Target Positioner for Leksell Headring, via e-mail on June 5, 2007. The Urgent Medical Device Product Notification letter dated June 4, 2007 included in the e-mail informed the customers that both the BrainSCAN and iPlan RT Dose treatment planning software calculates the isocenter position printed on the Target Positioner Overlays shifted 1.25 mm in cranial direction from the intended position, when the BrainLAB Target Positioner for Leksell Headring Revision A is used. The customers were instructed to cease using the BrainSCAN or iPlan RT Dose software in combination with the BrainLAB Target Positioner for Leksell Headring for treatment planning and patient set-up. An updated target positioner will be provided to the customers to correct the problem, estimated to be available in September 2007. Any questions were directed to their local BrainLAB Customer Support Representative or CAPA Manager Markus Hofman at 49-89-99-15-68-44 or 1-800-597-5911 in the U.S. |  | 7 | 7 units | Ohio, Washington and France | Terminated | February 24, 2008 | 264 | 7 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0988-2007 | 37954 | No | MWM |  |  | Obstetrics/Gynecology | PMA | MediSpectra Luma Cervical Imaging System |  |  | May 15, 2007 | 2007 | Medispectra Inc. | Power supply may fail and render the unit inoperable. | Other | MediSpectra notified users by telephone on 5/15/07 and follow-up letter. Users are requested to select a date for field servce to replace the power supply. |  | 8 | 8 units | AZ, CA, FL, &amp; NY. | Terminated | July 09, 2007 | 55 | 8 | Battery | Device Operation | Replace power supply | Remove or Replace
Z-0887-2007 | 37692 | No | MJK | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Batteries used with the LIFEPAK 20 defibrillator/monitor |  |  | March 30, 2007 | 2007 | Medtronic Emergency Response Systems, Inc. | Defibrillator internal battery cell may contain a short which could reduce battery capacity and cause the device to shut down without warning while operating on backup (DC) battery power. | Other | Firm sent letters on 03/30/2007 dated 03/2007 via certified return receipt to consignees with affected LIFEPAK 20 devices. The firm advised consignees to use the Operator's Checklist in the Operating Instructions to perform daily user tests and function checks while operating the defibrillator on battery power. In addition consignees are advised to replace their battery every 1 to 2 years. |  | 3002 | 3002 | Distribution is worldwide. The device is designed to be used in healthcare provider facilities such as hospitals and clinic settings. | Terminated | May 04, 2011 | 1496 | 29420 | Not_Computer
Z-0991-2007 | 37968 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Pneupac VR1 Standard and VR1 Responder is a hand held portable, time cycled, gas powered, flow generator ventilatory resuscitators that are intended for emergency resuscitation by medical personnel, paramedics and ambulance technicians inside and outside hospital and for ventilation by medical personnel inside and outside hospital in emergency situations and for intra- and inter- hospital transport. |  |  | May 16, 2007 | 2007 | Smiths Medical PM, Inc. | Smiths Medical is conducting a recall of 93 Pneupac VR1 Emergency &amp; Transport Ventilators following the identification of a problem which may exist with the demand valve that could cause the ventilator to stop cycling &amp; to deliver continuous flow output to the patient. | Other | Distriblutors and end users were contacted by phone (May 10 - May 24, 2007) and end users were additionally sent a Safety Action Bulletin May 21, 2007. The Bulletin was used as a script. |  | 93 | 93 | Nationwide. | Terminated | October 10, 2008 | 513 | 93 | Not_Computer
Z-0966-2007 | 37929 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Invacare Storm TDX Power Wheelchair, Models: #TDX3-PS, TDX3-SE-PS, TDX4-PS. |  |  | March 19, 2007 | 2007 | Invacare Corporation | Software error may permit the chair to recline too far. | Other | The recalling firm sent a recall letter dated March 19, 2007. The recalling firm also posted information regarding the recall on their website. |  | 227 | 227 devices | Nationwide and Canada. | Terminated | September 27, 2012 | 2019 | 227 | Software | Physical Safety Hazards | Notification | Safety Notice/Insructions
Z-0996-2007 | 38130 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Premier Edge MVR 20g Microsurgical Knife, Item PE-3520 |  |  | June 08, 2007 | 2007 | Oasis Medical Inc | This voluntary recall has been initiated due to the product puncturing the packingduring rough handling. Once the packaging is damaged, the product is no longersterile as labeled. | Device Design | Notification of the recall sent to distributors and users. The 06/08/07 letter requests customers to notify down to the user level. Included in theRecall Notification Letter are instructions: "In addition, if you may have further distributed this product, please identify your customers and notify them at once ofthis product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter. This recall should be carried out to the user level. Your assistance is appreciated and necessary to prevent any inconvenience." |  |  |  | Worldwide, including USA, Switzerland, Germany, England, UK, Turkey, Czech Republic, and Lithuania. | Terminated | September 25, 2009 | 840 |  | Not_Computer
Z-0993-2007 | 38116 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Graft Preparation System, sterile, single use, disposable, Distributed by Biomet Biologics, Inc., Warsaw, Indiana; REF 800-0300. |  |  | June 11, 2007 | 2007 | Biomet, Inc. | The syringe assembly has an obstruction in the tube that prevents the plunger from being depressed. | Other | Consignees were notified by letter dated either 6/11/07 or 6/13/07 and requested to return the product. |  | 46 | 46 to U.S. and international | Worldwide, including USA, Chile, Italy, Spain and United Kingdom. | Terminated | December 05, 2007 | 177 | 51 | Not_Computer
Z-1005-2007 | 37367 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Perfusion System 8000 Roller Pump, 1040V; Model 8000; Catalog No. 16411. (Not distributed within the United States). |  |  | July 01, 2004 | 2004 | Terumo Cardiovascular Systems Corp | Due to a circuit board malfunction the roller pumps may stop during use without displaying an error message, may stop during use with an overspeed error posted on the local display, or may fail to start/restart resulting in the pump(s) being non-functional. | Other | As the complaints were received (between 07/04 and 01/06), the firm visited the customer filing the complaint and repaired the pumps demonstrating the problem. The firm will correct other units as complaints are received. The firm modified its recall strategy and issued a Device Correction Letter dated 6/22/07 notifying all users of the problem |  | 728 | 728 pumps of all models | Worldwide: USA and Bangladesh, Belgium, Bolivia, Canada, Chile, Colombia, Costa Rica, Egypt, Equador, Hong Kong, India, Indonesia, Japan, Latin America, Malaysia, Nicaragua, Philippines, Saudi Arabia, Taipe, Taiwan, United Arab Emirates, Uruguay, and Viet Nam. | Terminated | December 12, 2007 | 1259 | 728 | Hardware | Device Operation | Correction | Repair
Z-0999-2007 | 37587 | Yes | NVZ |  |  | Cardiovascular | PMA | Vitatron C-Series C60DR DDDR Dual Chamber Rate Responsive Pacemaker includes the following models: C60A1 and C60A3. The Vitatron C60 DR is a dual chamber rate responsive pacemaker (activity sensing using an accelerometer) for permanent atrial and ventricular pacing. C-series software is VSF11/VSF12 Version 1.0. Vitatron is a division of Medtronic, Inc., Sales Office is Vitatron USA, 7000 Central Ave., NE, Minneapolis, MN 55423-3576. The head office is Vitatron B.V., P.O. Box 5227, 6802 EE Arnhem, The Netherlands. |  |  | February 05, 2007 | 2007 | Medtronic Inc. Cardiac Rhythm Managment | Pacing rate; Software anomaly affecting Vitatron dual chamber C-series and T-series pacemakers, related to automatic retrograde conduction testing, can inhibit the device from pacing if the patient's intrinsic rate falls below the programmed lower rate. | Software design | Advisory letters to physicians were mailed in February 2007. Sales respresentatives will follow-up with physicians by visit. A programmer software update will be released upon regulatory approval. The Advisory Letter provided short-term direction to reduce problems pending software upgrades. |  | 4294 | 4,294 | within the US to include: AL, AR, AZ, CA, CT, FL, GA, IL, IN, KY, MA, MD, MI, MN, NC, ND, NJ, NY, OH, OK, SC, TX. | Terminated | October 16, 2008 | 619 | 10553 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1007-2007 | 37944 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Insulin Infusion Pumps for the following Paradigm models: MMT-511, MMT-512, MMT-712, MMT-515, MMT-715, MMT-522 and MMT-722 |  |  | April 27, 2007 | 2007 | Medtronic MiniMed | Exposure to Magnetic Resonance Imaging (MRI) has resulted in damage to the component that monitors and controls movement of the motor in the MiniMed Paradigm insulin infusion pump. Although there were alarms as a result of the damage, some users cleared these alarms and continued using the pump. Under such conditions, the pump will significantly over-deliver; potentially causing severe hypoglycemi | Other | Firm mailed Medical Device Safety letters on April 24, 2007, to healthcare professionals, existing pump users and MRI facilities reiterating the existing warning in the pump User Guide to avoid exposing the pump to MRI (or similar high strength electromagnetic fields) and strengthening previous warning by specifically mentioning the potential for over-delivery and severe hypoglycemia. Firm is also including an insert with this information with any Paradigm infusion pumps shipped to new customers. |  | 334000 | 334,000 Paradigm insulin pumps (approx) | Worldwide: USA, US Virgin Islands, Antigua, Australia, Austria, Bahamas, Bahrain, Bermuda, Brazil, Canada, Chile, China, Columbia, Czech Republic, Denmark, Dominican Republic, Egypt, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Puerto Rico, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela. | Terminated | September 19, 2012 | 1972 | 334000 | Hardware | Treatment/Delivery/Therapy | Update manual | Other
Z-1043-2007 | 37733 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova 2100 IQ Digital Fluoroscopic Imaging System (Versatile cardiac and vascular). |  |  | March 26, 2007 | 2007 | General Electric Med Systems LLC | Caution on the non-detection of small invasive objects/devices, which project out from the skin surfaces (e.g. biopsy or other needles) while using the Innova Sense software option designed for patient contouring. Usage of Innova Sense patient contouring for such devices could potentially lead to patient injury. | Other | A "GE Healthcare Urgent: Safety Notice" dated January 24, 2007 was sent to consignees describing the product, problem, safety instructions, product correction and contact information. |  | 205 | 205 | Worldwide, including USA, Guam, Canada, Australia, China, India, Japan, Korea, Malaysia, New Zealand, Algeria, Czech Republic, Denmark, Finland, France, Germany, Greece, Hungary, Italy, Jordan, Morocco, Norway, Poland, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Syria, Turkey, and United Kingdom. | Terminated | October 21, 2008 | 575 | 373 | Software | Output/Calculation | Software update | Software Update
Z-0888-2007 | 37870 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Patrol enteral pump set with piercing pin, list #520540. |  |  | May 02, 2007 | 2007 | Ross Products Division Abbott Laboratories | The enteral pump set may have a 1-3 mm slit in the silicone tubing located between the blue connector and the stretch band Safe-T-Valve. | Other | The recalling firm sent a recall letter dated April 27, 2007 to all customers. |  | 1042050 | 1,042,050 units | Nationwide | Terminated | March 09, 2012 | 1773 | 1042050 | Not_Computer
Z-0885-2007 | 37996 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Terumo TenderFlow Pediatric Arterial Cannulae, 5.3mm (16 Fr) OD, with 1/4 inch non-vented connector, 9 inch (22.9 cm) long; Catalog no. 813578. |  |  | May 22, 2007 | 2007 | Terumo Cardiovascular Systems Corp | The bond may fail and a leak or separation of the device occur at the distal end (where this device connects to the arterial line). | Other | Consignees were notified via phone beginning on 5/22/07 and by follow-up letters sent between 5/23/07 and 6/14/07. |  | 39 | 39 | Nationwide. | Terminated | July 07, 2009 | 777 | 912 | Not_Computer
Z-1054-2007 | 37769 | Yes | NIK |  |  | Cardiovascular | PMA | VITALITY DR HE Implantable Cardioverter Defibrillator (model T180). Sterile EO. Guidant Corporation, Cardiac Rhythm Management, 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. This ICD is designed to detect and terminate ventricular tachycardia (VT) and ventricular fibrillation (VF) and provide bradycardia therapy (atrial and ventricular pacing). Therapies include both low- and high-eneergy shocks using either a biphasic or monophasic waveform. Bradycardia pacing, including adaptive-rate features, is available to detect and treat bradyarrhythmias and to support the cardiac rhythm after defibrillation therapy. The device offers dual-chamber bradycardia features (atrial and/or ventricular pacing and sensing). |  |  | March 23, 2007 | 2007 | Boston Scientific CRM Corp | Boston Scientific has recently identified low-voltage capacitors from a former supplier that may be subject to degradation at a low rate of occurrence. These capacitors may cause accelerated battery depletion and may reduce the time between elective replacement indicator (ERI) and end of life (EOL) to less than 3 months. | Other | An Urgent Medical Devivce Information letter dated April 05, 2007 was sent to implanting and following physicians. The letter describes the issue, identifies patients with affected devices and provides recommendations. |  | 3041 | 3,041 | Worldwide, including USA, Argentina, Aruba, Australia, Austria, Barbados, Belgium, Bermuda, Brazil, Canada, Chile, Costa Rica, Cyprus, Czech Republic, Czechoslovakia, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guam, Hong Kong, Hungary, Iceland, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Polynesia, Portugal, Puerto Rico, Qatar, Saint Lucia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan R O C, Thailand, Tunisia, Turkey, U.S. Virgin Islands, United Arab Emirates, United Kingdom, United States, Uruguay, Venezuela, and West Indies. | Terminated | March 02, 2008 | 345 | 73008 | Not_Computer
Z-1023-2007 | 37887 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | AxSYM Troponin-I ADV Reagent Pack; list 02J44; 100 test pack; in-vitro diagnostic; Abbott Laboratories, Abbott Park, IL 60064 |  |  | April 23, 2007 | 2007 | Abbott Laboratories | Failure of the AxSYM instrument to open (actuate) some reagent packs from certain lots. The defects can result in probe crashes if not detected prior to placement of the reagent kit on the AxYM instrument. | Other | Abbott initiated a field correction recall by letter dated 4/23/07 to all customers receiving the affected reagents. The accounts were informed of the failure of the AxSYM to open (actuate) some reagent packs with the affected lots. The accounts were requested to inspect each reagent pack for broken or improperly sealed caps; manually open each bottle to ensure the flipper bar raises the reagent cap correctly; close each flipper bar prior to loading the reagent pack on the instrument; and discard any reagent packs that do not open properly. If the user observes a probe crash, prior to testing samples again, the user must perform a probe crash recovery following the procedure in the Operations Manual. If necessary, they need to replace the probe. If a probe crash recovery is not performed, subsequent results may be impacted. The customers were requested to retain a copy of the letter for their records. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. |  | 9277 | 9,277 packs | Worldwide, including USA, Puerto Rico, Argentina, Australia, Brazil, Canada, Cayman Islands, Chile, Colombia, Ecuador, Germany, Hong Kong, Mexico, New Zealand, Singapore, South Korea, Taiwan, Thailand and Venezuela. | Terminated | May 18, 2008 | 391 | 28061 | Not_Computer
Z-1008-2007 | 37977 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Anspach Single Use, Sterile Bone Cutting Burrs |  |  | May 18, 2007 | 2007 | The Anspach Effort, Inc. | The white film portion of the sterile pouch may become brittle prior to the expiry date compromising the sterility of the product. | Nonconforming Material/Component | The firm issued a recall letter on May 18, 2007, to its consignee, advising them of the recall. Consignees were asked to screen their inventory for the recalled devices and complete an accompanying product replacement form. The firm will replace all returned product within 30 days. |  | 1962123 | 1,962,123 | World wide-Distributed throughout the United States and to Puerto Rico, Canada, Mexico, Japan, Lithuania, Cyprus, Austrailia, England, Sweden, Bahamas, Korea, Germany, Switzerland, Malaysia, Iceland, Portugal, Singapore, Poland, Hungary, Thailand, Austria, Israel, Belgium, Egypt, South Africa, Netherlands, India, Syria, Kenya, Italy, Czech Republic, Jamaica, Columbia, Argentina, Honduras, Saudi Arabia, Greece, Finland, Taiwan, Republic of China, Lebanon, Denmark, France, Hong Kong, Philippines, Venezuela, Argintina, and the United Kingdom. | Terminated | January 17, 2009 | 610 | 1962123 | Not_Computer
Z-0992-2007 | 38052 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Amplatz Goose Neck Snare Kit, 6 Fr, Lot 1650307, REF: GN3000, Use By 2009-09-01, Sterile EO |  |  | May 18, 2007 | 2007 | Ev3, Inc | Amplatz Goose Neck Snare Kit's catheter contained in the kit may be 4 Fr instead of the intended 6 Fr. | Other | Letter May 18, 2007. |  | 23 | 23 units | Worldwide, including USA, Belgium, France, Germany, Italy, Jordan, Spain, United Kingdom, and Yemen. | Terminated | October 17, 2008 | 518 | 23 | Not_Computer
Z-1045-2007 | 38072 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Wallach Cryotip T-00337 Anorectal Tip used with the Cryosurgical LL100 or LL50X Cryosurgical Systems |  |  | June 06, 2007 | 2007 | Wallach Surgical Devices, Inc. | The tip may separate from the probe. | Device Design | Wallach Surgical notified accounts by fax on 6/6/07. A follw-up telephone call will be made to the 70 domestic consignees. Foreign consignees notiifed by fax or email with follow-up telephone calls. Accounts are requested to return product and replace with an alternate product. |  | 315 | 315 units | Worldwide, including USA, Canada, Australia, England, Greece, Jordan, Turkey, Hungary, Netherlands, Saudi Arabia, Spain, and Lithuania | Terminated | July 15, 2015 | 2961 | 329 | Not_Computer
Z-1070-2007 | 37920 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Total B-hCG Speciment Diluent is composed of 1 bottle of Total B-hCG Specimen Diluent containing animal serum. |  |  | April 10, 2007 | 2007 | Abbott Diagnostic International, Ltd. | An increase in background counts elevates the results of low level samples up to approximately 6mIU/mL. This is not enough to shift a negative result to positive on an undiluted sample. If low level samples are diluted then the background effect is magnified potentially yelding false positive results for pregnancy and or Dilution Inaccuracy. | Other | The 4/10/2007 Product Recall Letter was sent to all customers who received the affected lots. |  | 1853 | 1853 | Worldwide, including USA, PR, USVI, CANADA, MEXICO, GUATEMALA, COSTA RICA, COLOMBIA, ECUADOR, PERU, CHILE, URUGUAY, ARGENTINA, BRAZIL, DOMINICAN REPUBLIC, EL SALVADOR, GERMANY, JAPAN, SINGAPORE, HONG KONG, TAIWAN, THAILAND, SOUTH KOREA, AUSTRALIA, NEW ZEALAND, HONDURAS, TRINIDAD &amp; TOBAGO, GRENADA, ST. VINCENT, ST. LUCIA, ANTIGUA, BARBADOS, CURACAO, BRITISH VIRGIN ISLANDS, BAHAMAS, PANAMA, and CAYMAN ISLANDS. | Terminated | February 14, 2012 | 1771 | 30275 | Not_Computer
Z-1076-2007 | 38115 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Ossur Total Knee Model Number 2100, external knee prosthesis. |  |  | April 02, 2007 | 2007 | Ossur | Some units of the Total Knee Model 2100 produced from August 1, 2006 through October 4, 2006 may contain a faulty linkage component. | Other | On 4/2/07 a letter was mailed to all affected customers via certified mail, return receipt requested. All mailing will be followed by telephone calls from an Ossur Customer Service Representative. Customers were notified that Ossur will replace the affected knees with units carrying a two-year warranty from the date of replacement.Ossur will undertake a sub-recall by notifying the one affected distributor by letter on 4/2/07. Ossur will follow up with this distributor by telephone and request a list of consignees that the distributor sold the product to so that Ossur can directly contact each of those consignees. Upon receipt of this list, Ossur will then communicate the recall information directly to these consignees. |  | 44 | 44 | Nationwide. | Terminated | May 28, 2008 | 422 | 44 | Not_Computer
Z-1075-2007 | 38062 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | AirLife nCPAP System Driver; a nasal continuous positive airway pressure device, which is part of a larger system that includes an nCPAP generator, fixation device and patient interface (prongs and.or mask); Cardinal Health, McGaw Park, IL 60085; cat. 006900 |  |  | June 01, 2007 | 2007 | Cardinal Health | Under isolated conditions, the device may experience issues with O2 Calibration, which may result in either a nuisance alarm or a delivered FiO2 of up to 5% below the specified tolerance. | Other | Cardinal Health sent the accounts recall letters dated 6/1/07 via UPS next day delivery, which listed the catalog number and affected serial numbers of the CPAP Driver, and informed them that under certain isolated conditions, the CPAP Driver may experience issues with O2 Calibration, which may result in either a nuisance alarm or a delivered FiO2 of up to 5% below the specified tolerance. Customers have experienced this under the following conditions: 1) air inlet pressure is greater than O2 pressure by 5 to 10 psi and 2) FiO2 settings less than 45%. Separate recall letters were sent to those accounts who had filed complaints of O2 Calibration Errors, stating that their device will be replaced by the end of June, 2007; those accounts who purchased the devices, but had not reported O2 Calibration Errors, stating that a Cardinal Health representative will upgrade the firmware on the devices at their facility by the end of August, 2007; and those accounts who had had demonstration units in their possession, advising them to notify their sales representative if any of the conditions mentioned in the letter occurred. The accounts were requested to complete and return via fax at 847-689-9101, the enclosed acknowledgment form, indicating the number of affected drivers in their possession, and were advised that they would be contacted by a Cardinal Health representative to arrange for the replacement/service of their drivers. Any questions were directed to Quality Systems-Professional Services at 1-800-292-9332 or 1-800-327-7503. |  | 48 | 48 units | Nationwide. | Terminated | December 12, 2007 | 194 | 48 | Software | Alarm/Message | Firmware update | Software Update
Z-1071-2007 | 37921 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | NEFA Standard Solution Calibrator, Oleic Acid, 1 mmol/L, 4x10mL bottles per carton, Cat. 276-76491, labeled in part ***For in vitro diagnostic use ***Distributed by: Wako Chemicals, USA, Inc. 1600 Bellwood Road Richmond, VA 23237*** |  |  | March 30, 2007 | 2007 | Wako Chemicals USA, Inc. | Diagnostic reagent for clinical laboratory testing was incorrectly labeled. | Other | The recalling firm notified consignees by letter 03/30/07 and advised of labeling error. Consignees were advised that they could return the product for relabeling or affix correct label enclosed with the notification. Distributors were requested to forward notification to end users. The firm subsequently notified consignees on 04/26/07 by fax that product should be returned for label correction. |  | 94 | 94 cartons | Worldwide, including USA, Germany and Canada. | Terminated | July 24, 2007 | 116 | 94 | Not_Computer
Z-1074-2007 | 38048 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Clinac Medical Linear Accelerator. Clinac 600 C &C;/D, Clinac 2100 C, C/D &amp; EX and Clinac 2300 C &C;/D |  |  | May 02, 2007 | 2007 | Varian Medical Systems Inc | The wedge tray may slip out of the Clinac interface mount during gantry rotation. | Process control | The firm sent out a Product Notification Letter on 5/2/07 to customers who have ordered this guide block through a Varian Service Center. A Varian Field Engineer will visually inspect guide blocks at-site that have been installed from Service stock for all affected customers, and replace if necessary. |  | 209 | 209 units | Product was distributed and manufactured between 1999 and 2006. Distributed worldwide, including USA, Ecuador, Columbia, Taiwan, Austrailia, Austria, Brazil, Canada, Chile, China, Finland, France, Germany, India, Italy, Japan, New Zealand, Poland, Portugal, Serbia and Montenegro, Singapore, Spain, Switzerland, UAE, and UK. | Terminated | November 27, 2010 | 1305 | 209 | Not_Computer
Z-0989-2007 | 38339 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GEHC OEC 9900 ESP-15 Elite C-Arm fluoroscopic X-ray system, GE Healthcare, Salt Lake City, UT. |  |  | July 06, 2007 | 2007 | GE OEC Medical Systems, Inc | The Air Kerma Rate displayed value and cumulative air kerma displayed value could exceed the allowable error under certain imaging conditions. | Other | Consignees were notified by certified letter on 07/06/2007. They were told that a service rep would contact them to provide a label and an addendum to their operator's manual. A permanent software solution is under development and will be provided when it becomes available. |  | 22 | 22 units | Worldwide, including USA, France, Singapore, Sweden, and Switzerland. | Terminated | February 01, 2013 | 2037 | 26 | Software | Display/Image | Software update | Software Update
Z-1077-2007 | 38169 | No | EAZ | Dental operating light. | Dental | Dental | 510(K) Exempt | DCI Equipment Track Mount Light Assembly, 110v |  |  | May 01, 2007 | 2007 | Dental Equipment LLC | Potential for light shield to melt due to light assemblies not properly rated to handle the VAC. | Other | On 5/1/07 the dental office contacted the manufacturer via telephone regarding the incident. On 5/2/07 the firm's service representative visited the dental office and replaced the lights. |  | 4 | 4 dual track lights | The lights were distributed to one distributor in NY who in turn distributed the lights to one dental office. | Terminated | July 12, 2007 | 72 | 4 | Not_Computer
Z-1042-2007 | 38016 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OneTouch Ultra Test strip, a component of the following OneTouch Ultra Glucose monitoring systems: a)OneTouch Ultra, b) OneTouch InDuo, c) OneTouch Ultra2, d) OneTouch UltraMini, e) OneTouch UltraSmart. 25 strips per vial. Product comes in packs of 100 and 50. (Part numbers: 020-245-07, 020-244-07, 020-963-01) |  |  | March 22, 2007 | 2007 | Lifescan Inc | Damaged Vials/Inaccurate test results: a manufacturing equipment problem on one of the packaging lines had the potential to puncture vials, causing a single hole in the side of the vial below the rim. This result of this defect may produce inaccurately high or low blood glucose test results or error message readings. | Other | March 2007 via letters to consumers and consignees and web postings. The firm has also set up verbal communication plans for customers who contact Lifescan by phone. |  | 33000000 | ~33 million vials (25 strips per vial in the US, ~19 million worldwide. | Worldwide; USA, Europe, Middle East, Africa, Latin America and Canada. | Terminated | August 15, 2008 | 512 | 33000000 | Not_Computer
Z-1019-2007 | 38220 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Various Sizes of Sterile Polyglycolic Acid Synthetic Surgical Suture-NDR, POLYGLY 2-0 C-2/CST43 2N36, Catalog #833-135 |  |  | June 06, 2007 | 2007 | Teleflex Creek Dr | Suture Diameter/Strength specification failure: The products may not meet their final specifications for suture diameter and/or strength which could lead to premature breakage. | Other | Consignees were notified by letter on/about 06/07/2007. They were advised to cease use, remove and quarantine all affected product immediately. Affected products are to be returned to Teleflex Medical. An Acknowledgement &amp; Stock form was included to be completed and returned via fax to Teleflex Medical. |  | 1920 | 1920 units total for all types of sutures | Worldwide-, USA including states of CO, FL, IL, IN, MA, and OH, and Countries of Canada, France and Germany | Terminated | March 14, 2008 | 282 | 1920 | Not_Computer
Z-1091-2007 | 38219 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Triple Channel Volumetric Infusion Pumps; Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015 U.S.A., Made in Singapore; product codes 2M8153, 2M8163, 2M9163 |  |  | June 20, 2007 | 2007 | Baxter Healthcare Corp. | A software anomaly is causing newly upgraded Colleague Triple Channel Infusion Pumps to alarm, display an error code (16:310:867:0002) and stop the infusion. This occurs during user programming with all three channels simultaneously infusing fluids. In reported cases, the pump stopped infusing and both an audible and a visual alarm notification were activated. | Software design | Baxter telephoned the direct accounts on 6/20-21/07, informing them that Baxter has received multiple reports of newly upgraded Colleague Triple Channel Infusion Pumps alarming, displaying an error code and stopping infusion. These reported events occurred during user programming with all three channels simultaneously infusing fluids. In all cases, the pump stopped the infusions and both an audible and visual alarm notification were activated. The accounts were requested to immediately remove the affected triple channel pumps, identified by two attached hangtags entitled "Battery" and "Power On", from service. The accounts were informed that their Baxter representative would be in contact with them to coordinate the replacement of the affected pumps with upgraded single channel pumps or non-upgraded triple channel pumps, if available. Any questions were directed to Baxter's Medication Delivery Services at 1-800-843-7867. Baxter sent follow-up Urgent Product Recall letters dated 6/22/07. |  | 4512 | 4,512 pumps | Worldwide, including USA, Canada, Australia, Belgium, New Zealand, South Africa, Thailand, and Singapore. | Terminated | December 20, 2011 | 1644 | 4512 | Software | Treatment/Delivery/Therapy | Remove | Remove or Replace
Z-1092-2007 | 38019 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | COMPLETE MOISTUREPLUS MULTIPURPOSE SOLUTION, Formula 9451X; |  |  | May 25, 2007 | 2007 | Advanced Medical Optics, Inc. | In response to information received 05/25/2007 from the U.S. Centers for Disease Control and Prevention (CDC) regarding eye infections from Acanthamoeba, a naturally occurring water-borne organism which can contribute to serious corneal infections. CDC estimates a risk of at least seven times greater for those who use Complete MoisturePlus solution versus those who did not. | Labeling design | On Friday, May 25, 2007 at 11:16 ET, AMO issued a press release to the market related to the voluntary recall of AMO's COMPLETE MoisturePLUS TM Multi-Purpose Solution from the market. Similar press releases, specific to the geographic regions, were issued in the other markets beginning on Saturday, May 26, 2007.A copy of the United States press release is present on the AMO website (www.amo-inc.com) and was forwarded to the FDA for posting on the FDA web site. In addition, the FDA also completed a press release related to the AMO recall and it is available on the FDA web site.Regulatory agency notifications in all markets affected outside the United States were initiated beginning on Monday, May 28. Beginning on Thursday, May 31, AMO (via Stericycle) distributed RetaiI Customer notification packages via overnight UPS delivery to the 485 customer accounts in the United States. |  | 27807198 | 27,807,198 units in US (57,252,581 Worldwide) | Worldwide including USA, Canada, Europe, Japan, China and various Asia Pacific countries (to include Hong Kong, Australia, New Zealand, Korea, Vietnam, Singapore, India, Indonesia, Sri Lanka, Taiwan, Malaysia, Pakistan and the Philippines). | Terminated | January 10, 2012 | 1691 | 31602188 | Not_Computer
Z-1056-2007 | 37728 | No | LAO |  |  | Toxicology | 510(k) | TDx/TDxFLx Methotrexate II Reagent-List Number 07A12-60 - a reagent system for In Vitro diagnostic use. This reagent system is composed of one bottle of Wash Solution, one bottle of Methotrexate Antibody, one bottle of Methotrexate Fluorescein Tracer and one bottle of Pretreatment Solution. |  |  | March 20, 2007 | 2007 | Abbott Diagnostic International, Ltd. | Incorrect configuration: kits contain reagent bottles in the order S-W-T-P instead of the correct order, W-S-T-P. | Other | The Product Correction Letter dated 3/20/2007, was sent to all customers who received the affected lots. The Customers were instructed to: Determine if they had the lot listed in the Product Correction letter. Inspect each kit of the affected lot and discard any kits incorrectly configured. If they had forwarded the lots listed in the Product Recall letter to other laboratories, they were instructed to inform them of this Product Recall and provide a copy of this letter to them and retain Product Correction letter for their laboratory records. |  | 1026 | 1,026 | World wide Distribution : USA including states of AR, CA, CO, CT, DE, FL,GA, ID, IL, IA, KS, KY, LA, MA, ME, MD, MI, MN, MS, MO, NJ, NY, NC,NV, OH, SC, PA, RI, TN, TX, UT, WA, WI, WV, Hawaii, Washington DC, and Puerto Rico. Countries of :Canada, Mexico, Venezuela, Columbia, Chile, Brazil, Germany, Japan, Hong Kong, Australia and New Zealand | Terminated | June 10, 2011 | 1543 | 1026 | Not_Computer
Z-1115-2007 | 38055 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | ThromCat Thrombectomy Catheter System. Catalog number 63000-01 |  |  | May 21, 2007 | 2007 | Kensey Nash Corp | Defective component (power supply) | Other | New power supplies were sent directly to Sales Representatives who personally replaced the defective power supplies and returned them to Kensey Nash. A voice mail to all affected KNC Sales Representatives went out on 5/21/07 instructing them to contact the affected accounts immediately. |  | 11 | 11 units | Nationwide:The products were shipped to medical facilities in FL, MO, NJ, and TX. | Terminated | August 29, 2007 | 100 | 11 | Not_Computer
Z-1116-2007 | 38105 | No | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Thermodilution Right Heart Catheters, Catalog Numbers 604028 and 604165. The product is shipped 5 units per carton.-B. Braun Medical, Allentown, Pa. |  |  | June 05, 2007 | 2007 | B. Braun Medical, Inc. | Cracks at the Catheter Tip | Other | On 6/5/07, a Product Removal Notice was issued to the only distributor that received the product in Japan |  | 4695 | 4,695 | International ONLY- Japan | Terminated | October 31, 2007 | 148 | 4695 | Not_Computer
Z-1089-2007 | 38128 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | BCI Advisor Vital Signs Monitor (Model/Catalog Number: 9200) |  |  | June 05, 2007 | 2007 | Smiths Medical PM, Inc. | Turn off unexpectedly/ Not turn on : monitors turned off unexpectedly sounding the critical failure alarm. Advisor Vital Signs monitors that were off, would not turn back on. | Other | Consignees will be contacted via a Safety Action Bulletin 07-SAB02 dated June 5, 2007. Technical Worksheet, 07-TW01 including a kit will be sent to the International consignees whom Technical Service has deemed qualified to perform the repair. The repair will be performed at the consignee site. International consignees who require or request the repair be done at the Smiths Medical PM, Inc. Technical Service department will return their monitors for repair. |  | 57 | 57 | International Distribution Only: Czech Republic, Peru, Argentina, Pakistan, Greece, Romania, The Netherlands, Germany, Puerto Rico, United Kingdom, Brazil, Uruguay, Spain, Australia, and Israel | Terminated | October 29, 2008 | 512 | 57 | Other | Device Operation | Repair | Repair
Z-1090-2007 | 38154 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Performer CPB-Universal R2100375. |  |  | June 04, 2007 | 2007 | Medtronic Perfusion Systems | Device failure requiring use of back up systems (malfunctioning printed circuit board (PCB)) | Other | There are two separate letters dated June 4, 2007 for the customers, based on the configuration of the unit being inspected/updated in the field. These letters explains the situaiton and informs the customers that Medtronic will be visiting each customer site to inspect/update the units. Medtronic Service Technicians will complete a Field Action Certicate acknowledging the successful completion of the update. |  | 24 | 24 units | Worldwide Distribution: USA including states of PA, ND, MN, CA, AL, CO, MI, NY, AZ, and OH and countries of Sweden, Iceland, Latvia, Germany, Italy, Austria, Belgium, and France. | Terminated | March 02, 2008 | 272 | 24 | Hardware | Device Operation | Repair | Repair
Z-1117-2007 | 38194 | No | FRW | Patient scale. | General Hospital | General Hospital | 510(K) Exempt | Liko Universal Sling Bar 450 Quick Release Hook (QRH) |  |  | June 13, 2007 | 2007 | Liko, Inc. | Sling bar may have a center bolt break at the adapter while in use | Component design/selection | Liko Inc notified distributors and end users by letter on 6/13/07. |  | 22 | 22 | Nationwide | Terminated | October 20, 2015 | 3051 | 25 | Not_Computer
Z-1065-2007 | 37881 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ALFA II Modular Femoral Left Modular Stem Porous Coated (Bowed); Size 6 Left; Length: 250mm; Distal Diameter: 18.0mm; Catalog #652-06-180L; component of the ENCORE Hip System; Single Use, Sterile; distributed by Encore Medical, L.P., Austin, TX 78758. |  |  | April 12, 2007 | 2007 | Encore Medical, Lp | Sterility/Package Integrity: Device packaging may potentially be unable to ensure sterility when exposed to shipping conditions. | Other | Urgent Device Recall Notification letters distributed to consignees on 04/14/07. Consignees requested to return product to Encore Medical, L.P. |  | 2 | 2 units. | Product distributed to surgical centers in CT, KS and MO. | Terminated | July 25, 2007 | 104 | 8 | Not_Computer
Z-1119-2007 | 38354 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Biomedicare Implantium Fixtures, Model numbers : FX3408, FX3410, FX3412, FX3414, FX3808, FX3810, FX3812, FX3814, FX4308, FX4310, FX4312, FX4314, FX4808, FX4810, FX4812, FX4814, FX4814W |  |  | June 29, 2007 | 2007 | Biomedicare, Inc. | Lack of validation for sterilization cycle: This recall was initiated because the validation of the sterilization process for these devices has not been completed. As a result, Biomedicare cannot assure the sterility of fixtures distributed to date. | Other | The firm initiated the recall notification with telephone calls to all consignees on 06/30/2007. Recall letters with attached Recall Response Form were mailed to all customers on 06/30/2007 by certified US mail. Customers were asked to examine their inventory and immediately remove all Implantium Fixtures from their inventory and quarantine until arrangements could be made for product return. |  | 8844 | 8,844 units | In US to AZ, CA, IL, TX &amp; WA. Internationally to Turkey | Terminated | July 25, 2012 | 1853 | 8844 | Not_Computer
Z-1061-2007 | 37778 | No | MLY |  |  | Physical Medicine | 510(k) | Gebauer''s Pain Ease Medium Stream -Topical Anesthetic Skin Refrigerant is a vapocoolant (skin refrigerant) packed in an aerosol container. Rx only. 103.5ml (3.5 fl. oz.) can-(Product Number-0386-0008-03) |  |  | April 17, 2007 | 2007 | Gebauer Company | Mold contamination | Other | The firm sent a recall letter dated April 17, 2007 to all customers. |  | 9732 | 9,732 cans of product | Nationwide | Terminated | November 30, 2009 | 958 | 41499 | Not_Computer
Z-1095-2007 | 36812 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL 800 FLEX series blood gas analyzers containing software version 5.24. |  |  | November 13, 2006 | 2006 | Radiometer America Inc | Software anomaly. Protein removing program incorporated in the operating software for Versions: V3.833 and V5.24, of the ABL700 and ABL800 Blood Gas Analyzers, may allow for protein build up on the electrode membranes of blood gas analyzers. This condition can influence analyzer results. When the protein removal program is run too frequently, more than once a week, or when running the protein rem | Software design | The firm issued a 'Safety Alert Notice' [FAN 915-248] to all customers which received the recalled ABL700 and ABL800 Analyzers equipped with the referenced software. The 11/13/2006 notice describes the problem concerning the Protein Removal program to the user/customers and emphasizes that the frequency of running the program in these software versions must be no more than once a week in order to prevent the protein build-up problem. It also contains additional recommendations for the user/customers for operation of these analyzer and informs the customers that a Radiometer Service Rep. will be coming out and making a software upgrade to these systems within the next few weeks. The alert notification is accompanied by a 'FAX FORM' to be signed and completed by each customer in order to verify that the alert notification has been received and that the staff members responsible for operating and maintaining the analyzer(s) on-site have been informed of this situation and that the necessary changes in the Operator's Manual concerning the Cleaning, Protein Removal, and Decontamination programs have been made. |  | 464 | 464 | The recalled analyzers were distributed to customers nationwide throughout the U.S. and in some parts of Canada. | Terminated | May 09, 2012 | 2004 | 481 | Software | Output/Calculation | Software update | Software Update
Z-1099-2007 | 37982 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Power Linear Cutter Reusable Digital Loading Units with Reloads. Common name - Implantable staple. Model number PLC60 |  |  | May 12, 2007 | 2007 | Power Medical Interventions, Inc | Malfunction - improper staple formation | Other | The field sales representatives were telephoned 5/14/2007, and instructed to return the affected devices from the customer sites. |  | 9 | 9 units | The products were shipped to medical facilities in DC, FL, GA, and NY. The product was also shipped to Germany | Terminated | August 29, 2007 | 109 | 9 | Not_Computer
Z-1101-2007 | 38075 | No | FTA | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Tri-Star Surgical Exam Light MR-16 ( Lamp Holder) used in the following: Model Numbers: TS2015W (wall mount) , TS2015M (mobile) , TS2015W (wall mount) , TS2035M (mobile) |  |  | February 13, 2007 | 2007 | Sunnex Inc. | Surgical and/or exam lamp holder may malfunction, discolor, melt and smolder. | Other | Sunnex notifiied distributors by letter March 9, 2007. Sunnex is providing replacement product to end users |  | 30 | 30 lamps | Nationwide | Terminated | January 06, 2009 | 693 | 376 | Not_Computer
Z-1098-2007 | 37961 | No | LOM |  |  | Microbiology | PMA | ARCHITECT AUSAB Controls list number 1L82-10 have two bottles with the following components: 1) Negative Control and 2) Positive Control. in vitro diagnostic. |  |  | May 14, 2007 | 2007 | Abbott Diagnostic International, Ltd. | During an internal study, the Architect AUSAB assay was found to read preparations of the WHO standard up to 53% higher than the WHO assigned value. This has the potential to cause elevated results and may impact the interpretation for a small percentage of results with grayzone or low-level reactive concentrations and higher proficiency sample results. | Other | All ARCHITECT AUSAB U.S. customers will be contacted. New customers will be contacted through a Product Information Letter that will be included in all ARCHITECT AUSAB Reagent Kits until assay is restandardized. All ARCHITECT AUSAB U.S. customers will be contacted. The following communication types were used: Product Correction Letter Customer Reply Form &gt; Product Information Letter to be included in new kits to be distributed. Letters were sent to customers by Federal Express Priority mail. The Customers were instructed to: Please follow their laboratory procedures for evaluation of these results Disregard the automated interpretation of results for low-level reactive samples.A customer reply form will accompany the Product Correction letter. Effectiveness Check Depth: 100% of customers receiving the Product Correction letter. The start date of communication was May 14,2007. |  | 393 | 393 | Nationwide and Puerto Rico | Terminated | June 10, 2011 | 1488 | 2134 | Not_Computer
Z-0240-2008 | 38256 | No | FIN | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Hill-Rom VersaCare Hospital Bed System; Models P3200 and P3201, Hill-Rom, Batesville, IN 47006 |  |  | November 12, 2006 | 2006 | Hill-Rom, Inc. | False Latch/Failure to Latch-The side rail may not latch, or the latch may not fully engage, resulting in the siderail being able to be lowered without pulling on the release handle. | Device Design | The firm notified consignees via an Urgent Device Correction letter dated 11/12/07 informing them of the problem and informed them they would be receiving upgrade kits which they could have their qualified personnel install themselves following video instructions also provided, or they could contact Hill-Rom for assistance. |  | 43731 | 43,731 beds total for both models were distributed; 3,282 have been corrected to date. | Worldwide, including USA, Argentina, Australia, Brazil, Canada, Dominican Republic, Emirate of Sharjah, England, Greece, Indonesia, Japan, Jordan, Kuwait, Lebanon, Malaysia, Mexico, Oman, Pakistan, Philippines, Qatar, Saudi Arabia, Selangor, Singapore, South Korea, Spain, Taiwan, Thailand, Turkey, United Arab Emirates and Venezuela. | Terminated | October 22, 2010 | 1440 | 43731 | Not_Computer
Z-1080-2007 | 38261 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Allez Medical Applications, Inc. 4Closure Surgical Fascia Closure System , Catalog ref. 4-10GPS. Sterile, One Time Use, Contents 4 sets 1- suture passer instrument, 1-guide rod, 10 mm, 250 mm long. |  |  | June 19, 2007 | 2007 | Synovis Surgical Inovation Div. of | There is a potential for the pouch to become damaged during shipping, compromising the sterility of the device inside. | Other | Consignees were sent a letter on June 19, 2007. The letter identified the product and problem. The letter also included a form for inventory purposes which is to be completed and faxed to Synovis. |  | 2 | 2 boxes (8 units) | Nationwide. | Terminated | May 08, 2012 | 1785 | 63 | Not_Computer
Z-1102-2007 | 38206 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | Cell Search Circulating Tumor Cell Kit (Epithelial) |  |  | April 17, 2007 | 2007 | Veridex, LLC | An increased incidence of aggregates was observed in Lot 0036 and 0036A. The presence of aggregates will cause customers to invalidate patient results. | Other | All Direct customers and customers of distributors received a customer notification informing them of the increased incidence of aggregates in cartridges on April 30, 2007 by Fed Ex overnight. This letter instructed them to discontinue use and to discard any remaining inventory. This was followed by a notification on May 24, 2007 by Fed Ex overnight to all US Distributors which also requested that they immediately stop using and/or distributing kits and to destroy any remaining inventory. Distributors were also asked to notifiy customers who received product. All foreign affiliate consignees were also informed on this date by e-mail and instructed to notify their customers. |  | 665 | Lot 0036: 468 kits; Lot 0036A: 197 kits | The kits were distributed to hospitals and laboratories nationwide. They were also sent to 7 J&J; affiliated companies in France, Japan, Canada, England, Germany, Italy and Spain. | Terminated | December 06, 2007 | 233 | 665 | Not_Computer
Z-1086-2007 | 38267 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | SPI EASY Abutment. Endosseous dental implant abutment. |  |  | April 19, 2007 | 2007 | Thommen Medical, AG | The undercut on the SPI EASY abutments which engages the stabilization ring of the implant is missing. Although the abutments can still be placed on the implant by the user the parts do not fulfill the quality expectations of the firm. | Other | Letter 5/14/2007. Customers are requested to check their current inventory for the recall product. If found they are to return the product for replacement. If they have implanted the product they are to check patient x-rays for the defect. If defects are found they are to do a restoration. |  | 84 | 84 | Nationwide. | Terminated | May 12, 2008 | 389 | 84 | Not_Computer
Z-1104-2007 | 38313 | No | MID | Multiple autoantibodies immunological test system. | Immunology | Immunology | 510(k) | american diagnostica inc. IgG Conjugate, 15mL; component of the american diagnostica, inc. IMUCLONEÂż aPL IgG ELISA Kit, Product #649G; manufactured by Louisville APL Diagnostics, Inc., Seabrook, TX 77586. |  |  | August 10, 2004 | 2004 | Louisville Apl Diagnostics Inc | Product mislabeled; conjugate IgG placed in bottles labeled as conjugate IgM and vice versa. | Other | Firm notified consignee of recall via phone on 08/10/04. Consignee asked to return conjugates for properly labeled replacements. |  | 45 | 45 | Product distributed to a single direct account in CT. | Terminated | November 27, 2007 | 1204 | 80 | Not_Computer
Z-1110-2007 | 37273 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept Filtered Cardiotomy Reservoir, Model 1351, Non-pyrogenic, Sterile, Manufacturer: Medtronic, Inc., Minneapolis, MN |  |  | April 18, 2007 | 2007 | Medtronic Perfusion Systems | Pyrogen Levels detected above specification: The product contains a pyrogen level above Medtronic''s internal specifications while the product is labeled as nonpyrogenic. | Other | The consignees were notified by personal contact and requested to stop using the devices in question. A customer letter (April 24, 2007) will be sent by e-mail which explains the situation and requests that the impacted product be returned to Medtronic for disposal. The customer will be requested to complete a Recall Certificate acknowledging the receipt of this information. |  | 11 | 11 units | Nationwide: Distributed to states of CO and PA | Terminated | November 01, 2007 | 197 | 11 | Not_Computer
Z-1108-2007 | 38325 | No | DSQ |  |  | Cardiovascular | PMA | Thoratec TLC-II Portable VAD Driver. Component (pneumatic) driver of Thoratec Ventricular Assist Device system |  |  | June 08, 2007 | 2007 | Thoratec Corp | Component Failure; Device Failure without warning; due to premature wear-out of the compressor motor, VAD support for the patient's circulatory system may fail without warning. | Nonconforming Material/Component | Urgent Medical Device Recall letters, dated June 08, 2007, were issued to consignees, with instructions to take the following actions: 1) check the current number of service hours on all units. 2) do not use TLC-II drivers exceeding 1500 hours until they are serviced by thoratec. 3) note the indicated hours on the product inventory form and return the form to Thoratec. 4) ensure that all patients have a back up driver and hand pumps at all times and have been trained in backup procedures. The firm is also arranging for the return and servicing of all affected drivers, with highest priority to those drivers already exceeding 1500 hours of life. A July 19, 2007, Press Release was issued by the firm. |  | 378 | 378 drivers | Worldwide: USA including states of AL, AR, AZ, CA, CO, CT, FL, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, NY, OH, OR, PA, TN, TX, UT, VA, and WA, as well as, countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, France, Germany, Ireland, Israel, Italy, Netherlands, Qatar, Singapore, Switzerland, Taiwan, UAE, and UK. | Terminated | July 17, 2008 | 405 | 378 | Not_Computer
Z-1109-2007 | 38342 | No | DSQ |  |  | Cardiovascular | PMA | Thoratec Paracorporeal Ventricular Assist Device (PVAD) component of the Thoratec VAD system. Catalog number 14086-2550-000. |  |  | March 30, 2007 | 2007 | Thoratec Corp | Incorrect collect nut assembly (inadvertent use of the incorrect (Black) collect nut) during surgical implantation (wrong size collect nut) to secure the inflow cannula to the PVAD has resulted in patient deaths | Device Design | The firm has issued an Urgent Medical Device Correction notice dated March 30, 2007, to all VAD customers alerting them to the hazard. The letter instructs them that a Thoratec representative will be contacting them to make arrangements to apply a warning label to the affected devices. The letter also informs them to dispose of any expired original design caged tip atrial cannula and to 1) not use the black collet and nut with any current design VAD, 2) make sure that all current patients using PVADs are using the appropriate collet nut, and 3) provide a copy of the product safety bulletin to all outlying hospitals and/or physicians that may be following patients implanted with PVADs and provide contact information with the acknowledgement form. |  | 2484 | 2,484 units | Class I Recall - Worldwide Distribution --- USA, Canada, Australia, Mexico, Singapore, Taiwan, Thailand, Austria, France, Germany, Israel, and Switzerland. | Terminated | April 24, 2008 | 391 | 2484 | Not_Computer
Z-1112-2007 | 38221 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Flo-Gard 6301 Dual Channel Volumetric Infusion Pump; Baxter Healthcare Corporation, Deerfield, IL 60015; Product code # 2M8064 |  |  | June 21, 2007 | 2007 | Baxter Healthcare Corp. | Falsified Repair/Test and Inspection records: Pumps submitted for repair or service may not have been properly repaired or serviced before being returned to the customer. | Other | Baxter telephoned the direct accounts starting on 6/21/07, informing them that functional verification testing may not have been performed on their Flo-Gard and/or Colleague infusion pumps before they were returned to them. Since Baxter has no evidence that this quality control testing was conducted on these pumps, the pumps may not function as designed. The accounts were requested to locate the affected pumps and take them out of service as soon as possible, and send them back to Baxter for the required functional verification testing.*** Baxter expanded the recall and telephoned the direct accounts starting on 7/30/07, informing them that the electrical safety testing may not have been performed on their Colleague infusion pumps before they were returned to them. Since Baxter has no evidence that this quality control testing was conducted on these pumps, the pumps may not function as designed. The accounts were requested to locate the affected pumps and take them out of service as soon as possible, and send them back to Baxter for the required electrical safety and all functional verification testing. Follow-up recall letters were sent to the accounts on 8/7/07 reiterating the telephone conversations and providing a list affected pump serial numbers to be removed from service and returned to Baxter for testing. They were also requested to report any malfunction or adverse event related to the Colleague or Flo-Gard Infusion Pumps to Baxter Corporate Product Surveillance at 1-800-437-5176. ****Recall 2 Extension******* Baxter expanded the recall for a second time via letter on 8/20/07 and telephone calls starting on 8/21/07, informing the direct accounts that Baxter's investigation found documentation inconsistencies by three more of their service technicians, representing the possibility of incomplete testing and/or documentation by these technicians who serviced Colleague infusion pumps at the account's facility. Without complete and properly do |  | 144 | 144 pumps | Nationwide: Alaska, Arizona, California, Colorado, Florida, Georgia, Hawaii, Idaho, Illinois, Iowa, Kansas, Kentucky, Louisiana, Maryland, Michigan, Mississippi, Montana, Nebraska, Nevada, New Jersey, New Mexico, New York, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Washington and Wyoming | Terminated | March 16, 2010 | 999 | 2013 | Not_Computer
Z-1105-2007 | 37614 | No | DWQ | External vein stripper. | Cardiovascular | Cardiovascular | 510(k) | Smith &amp; Nephew TRIVEX System Pneumatic Footswitch, Part #7209791; distributed by Smith &amp; Nephew, Inc. - Endoscopy, Andover, MA 01810. |  |  | October 12, 2004 | 2004 | Smith &amp; Nephew, Inc Endoscopy Division | Build-up of static charge in Pneumatic Footswitch tubing may cause control unit to reset to default settings in the operating room during surgery. | Other | Firm sent notification letters to consignees via overnight FedEx on 10/12-13/04. International customers were notified via letter on 11/10/04. Customers were asked to return footswitches for a replacement. |  | 274 | 274 units. | Product sent to distributors nationwide and to the following countries: Australia, Austria, Canada, China, Germany, Italy, Japan, and Taiwan. | Terminated | August 02, 2007 | 1024 | 274 | Not_Computer
Z-1106-2007 | 37963 | No | JJF | Micro chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Amplilink Software versions 1.1, 1.3, 1.4, 2.41. |  |  | May 21, 2007 | 2007 | Roche Molecular Systems, Inc. | For select COBAS AMPLICOR Tests run on the COBAS AMPLICOR Analyzer in conjunction with AMPLILINK software, a discrepancy has been identified between the onboard working reagent stability information reported by AMPLILINK Software (versions 1.1, 1.3, 1.4 and 2.41) and the stability information provided in the test kit package inserts/method manual. | Other | Urgent Medical Device Correction letters were mailed on 5/21/07 by first class mail. Letters describe the discrepancy between software reagent expiration dates and package insert expiration dates. |  | 5806 | 5,806 total | Nationwide to hospitals and laboratories. | Terminated | February 18, 2010 | 1004 | 5806 | Computer
Z-1020-2007 | 38225 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | All IV Administration Infusion Sets intended for use with the Gemini Infusion Pump and Alaris Pump module (a.k.a. Medley Pump module) including: Alaris SmartSite Low Sorbing Set, Alaris Vented Syringe Adapter Set, Alaris Burette Set, Cardinal Health SmartSite Infusion Set, Alaris Infusion Set, Alaris SmartSite Burette Set, Cardinal Health SmartSite Burette Set, Alaris SmartSite Administration Set, Alaris Infusion Burette Set, Alaris SmartSite Blood Set, Cardinal Health SmartSite Combination Infusion Set, Cardinal Health Blood Set, Alaris Medley/Gemini Infusion Set, Alaris Medley/Gemini Blood Set, Alaris Latex-Free Infusion Set, Alaris Light Sensitive Solution Set, Alaris Medley/Gemini Enteral Set, Imed Gemini 20 Vented/Nonvented Primary Administration Set, Imed Vented/Nonvented Gemini Prime-Saver Set, Imed Vented/Nonvented Gemini Short Set, Alaris Half Set, Cardinal Health Half Set, Alaris Latex-Free Burette Set, Alaris Opaque Set, Alaris Medley/Gemini Combination Infusion Set |  |  | June 18, 2007 | 2007 | Cardinal Health 303 Inc DBA Alaris Products | It has been determined that with low probability, over-infusion may occur as a result of an open safety clamp fitment (a.k.a. Flo-Stop fitment) on the pumping segment, whether opened intentionally in the course of expected clinical practice, or as a result of inadvertent action. The over-infusion occurrence is dependent on the user not closing the roller clamp first, as required by clinical practi | Other | The field action will consist of a mailing and will be conducted through consignees (distributors) in the affected countries as well as direct customers for all IV Administration sets intended for use with Gemini Infusion Pumps and Alaris/Medley Pump modules. The customers and distributors will be notified by registered return receipt mail. For direct customers, separate letters will be sent to the Director of Nursing, Director of Materials and Director of Risk Management at each facility. Distributors will be sent multiple copies of the Safety Alert Notification and Tip Sheets; along with a cover letter outlining their responsibilities to complete the Safety Alert plan and notify their customers as requested. The domestic letters (direct customers and distributors) were mailed beginning on June 18, 2007 to all of the affected accounts via US Postal Service Certified Mail. |  | 2000000 | &gt;2,000,000 units produced annually | Worldwide including USA, Argentina, Australia, Bahrain, Belgium, Canada, China, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malta, Mexico, Myanmar, Netherlands, New Zealand, Norway, Oman, Paraguay, Peru, Philippines, Portugal, Puerto Rico, Qatar, Saudi Arabia, Serbia/Montenegro, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad &Tobago;, United Kingdom, United Arab Emirates, Uruguay &amp; Venezuela. | Terminated | March 28, 2012 | 1745 | 2000000 | Not_Computer
Z-1021-2007 | 38230 | No | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb Zyoptix XP Epi Separator, packaged 1 per sterile pouch, 10 per carton, Reorder No. 507-0068, RX, for single use only. Ophthalmic surgical accessory. The product is labeled as Manufactured by: Bausch &amp; Lomb Incorporated, St. Louis, MO. |  |  | June 08, 2007 | 2007 | Bausch &amp; Lomb Inc | Reports have been received of stromal incursions after use of the product. | Other | The firm notified their customers via telephone between 6/8-12/07. During the call, the customer was informed reports had been received of stromal incursions and they were to quit using the product and return it. |  | 243 | 243/10-pouch cartons | Distribution was limited to "key" physicians in MN, SC, MD, and NJ. Foreign distribution was made to the firm's entities in The Netherlands, Singapore, Mexico, Canada, Hong Kong, and India. | Terminated | April 28, 2009 | 690 | 544 | Not_Computer
Z-1078-2007 | 38215 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Medical Systems IV Administration Set, Low Sorbing Extension Set, Model C20014, Rx Only, Alaris Medical Systems, Lot Number 07036013 |  |  | June 05, 2007 | 2007 | Cardinal Health 303 Inc DBA Alaris Products | The sterility of the product may be compromised as the packaging used for the sterility barrier was not designed for the sterility process used by the firm. | Other | Recall Notification consists of a phone call and Recall Notification letters to consignees (distributors) as well as direct customers. On June 5, 2007, Cardinal Health began calling customers and distributors to inform them of a recall affecting IV Administration Set, Model C20014, Lot Number 07036013 (Low Sorbing Extension Set). A follow-up Recall Notification Letter was sent on June 11, 2007.The customers and distributors were notified by registered return receipt mail. For direct customers, letters were sent directly to the point of contact at the hospital. For distributors, they were sent multiple copies of the Recall Notification; along with a cover letter outlining their responsibilities to complete the Recall plan and notify their customers as requested. |  | 4248 | 4,248 sets (59 boxes) | Nationwide. | Terminated | February 10, 2012 | 1711 | 4248 | Not_Computer
Z-1148-2007 | 38232 | No | MKT |  |  | Microbiology | PMA | biphasic LIFEPAK 500 automated external defibrillators (AEDs) with software version 4.4 or lower; Part Nos. 3011790-XXXXXX, D3100790-XXXXXX, and U3011790-XXXXXX . (XXXX indicates various non-contiguous numbers) |  |  | June 13, 2007 | 2007 | Medtronic Emergency Response Systems, Inc. | Reduced Shock- the device may deliver 100 Joule (J) to a patient, which is less than the recommended factory minimum default setting of 200 J. | Other | On June 13, 2007 the recalling firm sent a letter dated June 11, 2007, titled URGENT - MEDICAL DEVICE CORRECTIVE ACTION. The letter advises of the situation, recommends to the consignee to set defibrillator to FIXED Sequence Protocol and to contact the firm for a FLEXIBLE Sequence Energy Protocol. The letter provides an instruction sheet and a response sheet to be faxed back to the company. |  | 59996 | 59,996 devices | Worldwide | Terminated | March 20, 2011 | 1376 | 59996 | Not_Computer
Z-1132-2007 | 37640 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE Medical Systems Mac-Lab System, cardiac catheterization data system. Software Version 6.5 2021423-015 Revision D GE Medical Systems, Information Technologies, Inc., 8200 West Tower Avenue, Milwaukee, WI 53223 USA. |  |  | February 16, 2007 | 2007 | General Electric Medical Systems Information Technology | GE Healthcare has identified conditions that could cause a potential loss of real-time monitoring with the CardioLab/Mac-Lab/ComboLab Electrophysiology and Hemodynamic Monitoring Systems. These issues are associated with an audio driver, and the use of the application's log and imaging windows. | Other | Consignees were sent a letter dated 2/16/07. The letter described the product and problem, made recommendations and gave information on software update. |  | 1513 | 1513 (1067 US and 446 OUS) | Worldwide, including USA, Argentina, Australia, Bangladesh, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Republic of Korea, Kuwait, Malaysia, Mexico, Netherlands, New Zealand, Nicaragua, Pakistan, Panama, Puerto Rico, Saudi Arabia, Singapore, South Africa, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Kingdom, and Venezuela. | Terminated | December 11, 2011 | 1759 | 2909 | Software | Device Operation | Software update | Software Update
Z-1137-2007 | 37878 | No | LKH | Chlamydia serological reagents. | Microbiology | Microbiology | 510(k) | Femoral Speed Block Size 6, Part/Catalog #800-01-367; positioning device included as part of the M.I.K.A. (Minimally Invasive Knee Arthroplasty) surgical kit; distributed by Encore Medical, L.P., Austin, TX 78758. |  |  | August 15, 2006 | 2006 | Encore Medical, Lp | Complaints received that device used to prepare femur to receive femoral implant would split during surgery. | Other | Firm sent notification letter to consignees on 08/25/2006, along with replacement speedblocks. Consignees asked to trade out old speedblocks with the replacements. |  | 24 | 24 units. | Product distributed to direct accounts in USA and one consignee in Japan. | Terminated | March 27, 2008 | 590 | 24 | Not_Computer
Z-1151-2007 | 38121 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Straight Magnum Inserter Handle, a preparatory instrument used for acetabular component insertion; REF # or Catalog Number: S313141. (systems affected by this action include M2a-Magnum, Regenerex, Acetabular Revision Shell and M2a-38) |  |  | June 14, 2007 | 2007 | Biomet, Inc. | The instrument was manufactured incorrectly and use could result in the instrument binding to the acetabular component and cause implant to not release following inpaction. | Other | On 6/13/07, the firm sent a Urgent Medica Device Correction Notice to its sales force instructing them to retrieve affected units from their hospital accounts and to return them to the firm. International accounts were notified via email on 6/14/07. |  | 94 | 94 | World wide: USA, Australia, and Finland. | Terminated | December 12, 2007 | 181 | 94 | Not_Computer
Z-1127-2007 | 38234 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | Instrumentation Laboratory ACL TOP Coagulation Analyzer (Base Model only). Blood analyzer. |  |  | June 19, 2007 | 2007 | Instrumentation Laboratory Co. | ACL TOP unit, when in emergency stop may release liquid during recovery sequence and contaminate sample tube. | Other | Instrumentation Laboratory notiifed the distributor Beckman Coulter on 6/19/2007. Beckman Coulter notified their customers and provided a copy with the Instructions for Recovery Procedure on how to avoid contamination from probe leaks. |  | 94 | 94 | Nationwide. | Terminated | May 23, 2012 | 1800 | 94 | Not_Computer
Z-1142-2007 | 38223 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | The Companion ClearStar Pump System, for enteral feeding. |  |  | June 21, 2007 | 2007 | Ross Products Division Abbott Laboratories | The transducer adapter bracket on the pump could crack or break and potentially result in under delivery of the enteral product. | Equipment maintenance | Consignees were notified by letter dated 6/25/2007. A revised letter was sent 07/26/2007. |  | 27814 | 27,814 | Worldwide, including USA and Canada. | Terminated | September 26, 2012 | 1924 | 27814 | Not_Computer
Z-1130-2007 | 38373 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Cleaning probe included as an accessory to aspirators that are sterile packed single use disposable accessories to the GE InstaTrak 3500 Plus x-ray system. Probes are packaged and shipped together with the following aspirators: Straight Aspirator - 1000370-NAV, 7 French Aspirator - 100936, Keane - Moses Extended Straight Aspirator - 1000748, 15 degree aspirator - 1003895, Fried 45 Degree Aspirator - 1000541, Kuhn 90 degree Aspirator - 1000543-NAV. |  |  | May 18, 2007 | 2007 | GE OEC Medical Systems, Inc | Cleaning probes used to unclog disposable aspirators may shed small particles of nylon material which may be transferred into the nasal cavity. | Other | Consignees were notified by certified mail on 05/18/2007 and were told to stop using and to dispose of the affected probes. |  | 67625 | 67,625 units | Worldwide, including USA, Canada and France. | Terminated | August 22, 2011 | 1557 | 67625 | Not_Computer
Z-1128-2007 | 38239 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips Medical HeartStartXL Defibrillator/Monitor |  |  | June 11, 2007 | 2007 | Philips Medical Systems | The ECG signal may become unusable to assess patient condition and need for therapy. | Nonconforming Material/Component | Philips Medical contacted accounts by letter dated June 11, 2007. Users are advised to remove product from service, if possible. If not, a back up unit should be available to use. Service reps will contact accounts to replace units. |  | 312 | 312 units | Worldwide, including USA, Canada, Australia, Brazil, China, Czech Republic, France, Germany, Hong Kong, India, Ireland, Italy, Japan, Korea, Kuwait, Macau, Poland, Portugal, Romania, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, UK, and Venezuela. | Open, Classified |  |  | 312 | Not_Computer
Z-1129-2007 | 38341 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Curlin Medical Administration Sets: Model Numbers 340-4111, IV Administration Sets. |  |  | February 09, 2007 | 2007 | Curlin Medical Llc | The Administration Sets may leak at the reservior bag. | Other | A Field Correction Advisory (Recall Letter) was issued on February 9, 2007 to the firm's one US Customer (distributor): B. Braun Medical Inc. recalling Curlin Medical, Inc. Administration Sets 340-4111- Lot # D628223 and D627616 because the affected lots do not meet acceptable performance levels.A second Field Correction Advisory (Recall Letter) was issued on March 15, 2007 to B. Braun Medical Inc. recalling Curlin Medical, Inc. Administration Sets: 340-4111- Lot # D628223, D6276l6 ,D620544, D630503; 340-4127- Lot# D620104; 340-4119- Lot# D623716 because the affected lots do not meet acceptable performance levels. |  | 25280 | 25, 280 | One US Customer (distributor): B. Braun Medical Inc. 901 Marcon Blvd. Allentown, PA 18109 | Terminated | May 28, 2008 | 474 | 25280 | Not_Computer
Z-1149-2007 | 38289 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | VS-1000 Senator Prosthetic Foot, (Categories 5 and 6, sizes 22-30 cm, both Left and Right) Applied Composite Technology, Gunnison, UT. |  |  | June 20, 2007 | 2007 | Applied Composite Technology | Prosthetic foot device may loosen or separate during use; due to structural inconsistency that may cause a pyramid failure. | Other | On 06/20 and 06/21/2007 all recipients of the affected Senators were notified by telephone. They were asked to return all suspect prostheses for replacement. Follow-up emails and fax letters (6/21/07) were also sent. |  | 97 | 97 units | Worldwide Distribution- USA including states of CA, CO, FL, IL, MI, MN, MS, NC, NY, OR, PA, SC, TN, TX, VA, WA. VA facilities in CA and MI. Foreign distribution to Australia, Canada, Germany, Russia, South Africa, South Korea, and UK. | Terminated | June 05, 2008 | 351 | 97 | Not_Computer
Z-1143-2007 | 38226 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Kimberly-Clark Microcuff Endotracheal Tube, Pediatric Oral/Nasal Magill, sizes 3.0mm, 3.5mm, 4.0mm, and 4.5mm, Rx Only, Sterile, Manufacturer: Kimberly-Clark NV Belgicastraat 13, 1930 Zaventem, Belgium, Made in Malaysia |  |  | June 04, 2007 | 2007 | Kimberly-Clark Corporation | The tubes may kink during use which may impede or prevent ventilation. | Other | Consignees were notified by letter sent via overnight mail on/about June 4, 2007. A follow-up letter was sent to consignees via fax on June 5, 2007. Consignees were instructed to cease distributing the product and return all affected product in inventory to Kimberly-Clark. Distributors were instructed to notify their accounts of the recall. A Product Recall Response Sheet (Packing Slip) was attached to the recall notice. |  | 4376 | 4376 cases (43760 units/total for both products) | Worldwide, including USA, Canada, Europe, Australia, Chile, Hong Kong, Singapore, Thailand, India, Malaysia and Saudi Arabia. | Terminated | June 12, 2008 | 374 | 4376 | Not_Computer
Z-1141-2007 | 38222 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | Embrace Enteral Pump |  |  | June 21, 2007 | 2007 | Ross Products Division Abbott Laboratories | Continued use of the pump may cause the door pivot point on the back of the pump to crack or break. This will cause the pump to shut down and stop the initiation of enteral feeding. | Equipment maintenance | Consignees were notified by letter dated 6/25/2007. The firm sent a revised letter on 07/26/2007. |  | 9600 | 9,600 units | The product was distributed throughout the US. | Terminated | September 20, 2012 | 1918 | 9600 | Not_Computer
Z-1147-2007 | 38358 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC 4.2 Software Part Number NT-NW-425-520. Image guided radiation treatment software. |  |  | May 31, 2005 | 2005 | Northwest Medical Physics Equipment Inc | When the ISOLOC (Version 4.2) profile distance unit is set to "centimeter" and the user clicked the "back" button in the digitizer window, the marker and target coordinates in the localization window erroneously display values ten times greater than the correct values. | Software design | Software to fix the described problems was written and customers contacted by e-mail on 05/31/2005, that a software patch, titled 05312005 was available on the firm's internet website. |  | 85 | 85 | Worldwide, including USA, Korea, Spain, and UK. | Terminated | April 19, 2011 | 2149 | 85 | Software | Output/Calculation | Software update | Software Update
Z-1146-2007 | 38355 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC 6.5 Software Part Number NT-NW-425-520. Software for radiotherapy treatment. |  |  | April 07, 2006 | 2006 | Northwest Medical Physics Equipment Inc | This correction was for several software bugs which caused image format problems; incorrect ruler display; software crashes; failure to update pixel spacing; display of unattainable treatment table positions; failure to update gantry angles; wrong shifts for two of three axes of treatment table movement. | Software design | Software to fix the problems was written and software users were notified of the availability by e-mail on 04/07/2006. They could download the fixed software from the firm's internet website. |  | 141 | 141 copies | Customers in USA, Korea, Spain and UK. | Terminated | April 19, 2011 | 1838 | 141 | Software | Display/Image | Software update | Software Update
Z-1153-2007 | 38361 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC Software Version 6.0, Localization software for determining treatment machine position (to position patients) for high accuracy radiotherapy. |  |  | April 11, 2005 | 2005 | Northwest Medical Physics Equipment Inc | When ISOLOC version 6.0 was used in the centimeter mode, the display in the couch moves window had a rounding error relative to the Portal Image Marker window of up to 0.5 mm. | Software design | Software to fix the described problem was written and customers contacted by e-mail on 04/11/2005, that a software patch, titled 04112005, was available to their customers on the firm's internet website. |  | 84 | 84 | Worldwide, including USA, Korea, Spain, and UK. | Terminated | April 19, 2011 | 2199 | 84 | Software | Display/Image | Software update | Software Update
Z-1145-2007 | 38319 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Isoloc software Versions 6.5 and 4.5, Part Number NT-NW-425-520. Software for image guided radiation therapy. |  |  | January 02, 2006 | 2006 | Northwest Medical Physics Equipment Inc | Isoloc version 6.5 and 4.5 software issued a false warning message regarding Gantry angles when set to "Anatomic Landmark" setting. | Software design | Software to fix the described problem was written and customers contacted by e-mail on 01/06/2006 that a software patch, titled 01022006 was available to their customers on the firm's internet website. |  | 132 | 132 | Nationwide, Korea, Spain, and UK. | Terminated | April 19, 2011 | 1933 | 132 | Software | Alarm/Message | Software update | Software Update
Z-1154-2007 | 38363 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC Software Version 6.0, Part Number NT-NW-425-520. Localization software for radiotherapy treatment system. |  |  | December 30, 2004 | 2004 | Northwest Medical Physics Equipment Inc | When ISOLOC version 6.0 software was used in the Anatomic Landmark mode, the couch moves display in the Couch Translator window was incorrect. | Software design | Software to fix the described problem was written and customers contacted by e-mail on 12/30/2004 that a software patch, titled 12302004 was available to their customers on the firm's internet website. |  | 79 | 79 | Worldwide, including USA, Korea, Spain, and UK. | Terminated | April 19, 2011 | 2301 | 79 | Software | Display/Image | Software update | Software Update
Z-1152-2007 | 38360 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC Software Versions 6.0 and 4.2. Part Number NT-NW-425-520. Image guided radiotherapy treatment software. |  |  | April 28, 2005 | 2005 | Northwest Medical Physics Equipment Inc | When ISOLOC version 6.0 and 4.2 distance unit is set to centimeter, the localization coordinates were rounded to the nearest millimeter which could affect positioning by 0.5 mm. | Software design | Software to fix the described problems was written and customers contacted by e-mail on 04/28/2005, that a software patch, titled 04282005, was available on the firm's internet website. |  | 84 | 84 | Worldwide, including USA, Korea, Spain, and UK. | Terminated | April 20, 2011 | 2183 | 84 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1072-2007 | 37952 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | STA Liatest Free Protein S |  |  | May 08, 2007 | 2007 | Diagnostica Stago, Inc. | Specific lots of STA Liatest Free Protein S may result in low results in the assigned value ranges and sometimes outside the lower limit. | Other | Customers were sent a Product Recall Notification on May 8, 2007 by first class mail labeled " Recall notice, Open Immediately" |  | 365 | 365 units | Medical facilities nationwide | Terminated | October 13, 2010 | 1254 | 365 | Not_Computer
Z-1073-2007 | 37999 | No | FFK | Electrohydraulic lithotriptor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gyrus CyberWand Ultrasonic Lithotripter System includes: |  |  | May 23, 2007 | 2007 | Gyrus ACMI Corporation | Sterilization Instructions not adequate for sterilizing. | Other | Gyrus has informed their sales force to visit each customer that has a CyberWand system(s) and hand deliver a letter, dated May 23, 2007, that explains the situation and requests that the customer cease using the unit and segregate it to prevent accidental use. The letter also explains that the firm is collecting any sterilization trays with printed sterilization instructions and replacing them. The rep will assist the customer in complying. The customer will also be required to complete a reply form to track compliance. |  | 53 | 53 units | Nationwide and Canada. | Terminated | May 25, 2010 | 1098 | 53 | Not_Computer
Z-1157-2007 | 38005 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The ClearGlide EVH small with Scissors. |  |  | May 07, 2007 | 2007 | Datascope Corporation | Datascope has experienced a higher than usual number of reports involving the Precision Bipolar Device. Reports include the locking or sticking of the buttons of the device which operate the jaws, particularly when a side load is placed on the shaft of the device. | Other | Letters were sent to all hospitals, (contact: Hospital Administrators) the week of May 7, 2007 (certified mail, Return Receipt requested) with Urgent Medical Device Field Correction. The letter requests that all potential users of the devices at the institution be mdae aware of the situation. The letter instructs clinicians what to do if they experience problems with locking buttons of the Precision Device and also if the device is clamped closed on a vein branch and cannot be opened.Institutions are also requested to sign and return a postal card acknowledging receipt of the Field Correct notification and Datascope provides telephone numbers for customer service. |  | 549 | 549 | Worldwide, including USA, Belgium, the Netherlands, and Germany. | Terminated | March 04, 2008 | 302 | 1671 | Not_Computer
Z-1160-2007 | 38112 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Impulse Flextrusion Shaft 5F FL4 16391-22 H74916391222, (5-Pack Carton Label), Angiographic Catheters. |  |  | June 04, 2007 | 2007 | Boston Scientific Corporation | Labels on the carton from two batches/lots of Impulse Angiographic Catheters may indicate a different curve style than the device that is actually in the carton. | Process control | Sales Force were contacted of the product and problem via voice mail and a memo dated June 4, 2007. Customers were contacted via a letter on June 4, 2007. |  | 233 | 233 | Worldwide, including USA, Austria, Australia, Belgium, Chile, Germany, Spain, Finland, France, Great Britain, Greece, Ireland, Panama, Poland, Sweden, Thailand, and South Africa. | Terminated | December 19, 2007 | 198 | 468 | Not_Computer
Z-1155-2007 | 38366 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ISOLOC Software Version 4.2, Part Number NT-NW-425-520. Image guided radiation treatment software. |  |  | November 30, 2004 | 2004 | Northwest Medical Physics Equipment Inc | When ISOLOC Version 4.2 unit is set to centimeter and the user clicks the Reenter button in the digitizer window, the coordinates of the markers and target in the anatomic imaging system erroneously displays values ten times greater than the correct values. | Software design | Software to fix the described problem was written and customers contacted by e-mail on 11/30/2004, that a software patch, titled 11302004, was available to their customers on the firm's internet website. |  | 83 | 83 | Worldwide, including USA, Korea, Spain, and UK. | Terminated | April 19, 2011 | 2331 | 83 | Software | Output/Calculation | Software update | Software Update
Z-1161-2007 | 38117 | Yes | NIM |  |  | Cardiovascular | PMA | Boston Scientific NexStent Monorail 5F Carotid Stent and Monorail Delivery System, Manufactured by: Endotex International Systems, Inc. 10231 Bubb Road Cupertino, CA 95014 USA |  |  | May 01, 2007 | 2007 | Boston Scientific Corporation | Boston Scientific NexStent Carotid Stent device failed to deploy when the outer catheter (proximal outer and distal sheath) did not pull back and expose the self-expanding stent. | Other | On May 1, 2007, Boston Scientific initiated a Customer Notification worldwide. On May 21, 2007 Boston Scientific expanded this field action to a recall/removal status for Germany. Affected customers were notified in Germany on May 24, 2007. |  | 2487 | 2487 | Worldwide, including USA, France, Italy, Germany, Belgium, Cyprus, Austria, Spain, Finland, Netherlands, Luxembourg, Romania , Czech Republic, Poland, Israel, Hungary, New Zealand, and Switzerland. | Terminated | March 02, 2008 | 306 | 2487 | Not_Computer
Z-1162-2007 | 38283 | No | EZB | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Cook PTFE Wire Guide with 3 cm flexible tip, to establish a tract for ureteral/renal access, .038", 145 cm; sterile; Order/Ref number 638413. |  |  | June 12, 2007 | 2007 | Cook Urological, Inc. | Lack of assurance of sterility, in that the product is labeled as sterile but the package may not have been sealed. | Process control | Consignees were notified via letter dated 6/12/07 and instructed to cease using the product and to inform the firm if they have product to return. |  | 100 | 100 | Nationwide and Canada. | Terminated | November 25, 2009 | 897 | 100 | Not_Computer
Z-1131-2007 | 38244 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Possis AngioJet XVG Thrombectomy Sets, |  |  | June 19, 2007 | 2007 | Possis Medical, Inc | Mislabeling : Possis AngioJet XVG Catheter Thrombectomy Sets display incorrect information on the outer packaging box and also on the inner sterile tray. The labeling incorrectly states that the unit is indicated for use in coronary vessels. These units should indicated use only for peripheral arteries. | Other | A Urgent Medical Device recall letter dated 19 June 2007, was sent to consignees advising that all unused product be removed and returned to Possis. An acknowledgement of the disposiiton of the product will be requested from consignees. |  | 194 | 194 | AR, AZ, CA, FL, GA, IL, IN, KS, KY, MA, ME, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, WA, WI, WV | Terminated | March 02, 2008 | 257 | 194 | Not_Computer
Z-1163-2007 | 38309 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Sunrise Medical Quickie Power Wheelchairs: Three models -Rhythm, Groove, and Zippy Z-Bop (with Quickie IQ Enhanced Displays) |  |  | April 23, 2007 | 2007 | Sunrise Medical Inc | The wire harness connected to the display housing (Quickie IQ Enhanced Displays) may separate and cause sparks from a short circuit. | Component design/selection | Firm contacted consignees by phone and mailed May 1, 2007, Urgent Medical Device notification letters. |  | 229 | 229 Enhanced Display units (affecting 220 wheelchairs) | Product was distributed Worldwide. | Terminated | March 25, 2008 | 337 | 229 | Not_Computer
Z-1150-2007 | 38458 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bayonet Electrosurgical Pencil. P/N 9560575. The product is shipped in cases containing 100 blades/pencils. |  |  | July 12, 2007 | 2007 | Zinnanti Surgical Design LLC | Bubble Leak Test Failure of sterility barrier, therefore sterility of the device may be compromised | Other | The recalling firm issued an Urgent Medical Device Recall Letter dated 7/24/07 to its single wholesaler requesting recall and sub-recall to all their consignees. |  | 8860 | 8, 860 units | Nationwide | Terminated | March 07, 2008 | 239 | 8860 | Not_Computer
Z-1166-2007 | 38209 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Orbit 90 Subcutaneous Infusion Set, 9 mm Catheter, 42" Tubing - Blue, List #P4291 |  |  | September 26, 2006 | 2006 | ICU Medical, Inc. | The product contains a female luer which has been reported to crack in use allowing fluid to leak from the device. | Other | Firm mailed a letter dated September 26, 2006 to notify customers of the recall. The letter advises the product contains a female luer which has been reported to crack in use allowing fluid to leak from the device. Customers are asked that if they have inventory of the lots identified to quarantine these lots and contact your consignees of this recall and return the product immediately for destruction. They are asked to complete and return the attached product inventory summary sheet in the enclosed self-addressed stamped envelope and return product to ICU Medical. |  | 26361 | 26,361 units | VA, OH, TX, CA, GA | Terminated | March 22, 2012 | 2004 | 26361 | Not_Computer
Z-1167-2007 | 38217 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROSÂż 5,1 FS (Fusion Series) Chemistry System Software Version 2.2.1, only when using VITROSÂż Chemistry Products THC (Cannabinoids) or AMPH (Amphetamine) Reagents, Cat. No. 680 1375. |  |  | May 17, 2007 | 2007 | Ortho-Clinical Diagnostics | Software design. The firm received a customer complaint regarding a qualitative positive THC (Cannabinoids) result obtained from a THC quality control fluid that was known to be negative when using VITROSÂż Chemistry Products THC Reagent. A second customer complaint was received a month later. OCD's investigation concluded that under specific circumstances, incorrect qualitative results for Cannabi | Other | On 5/17/07, OCD informed its foreign affiliates of the field correction by e-mail, and instructed them to inform their customers who received the VITROS 5,1 FS Chemistry Systems as well as the VITROS Chemistry Products THC (Cannabinoids) and/or AMPH (Amphetamine) Reagents of this issue. On 5/18/07, all US OCD direct customers and end-use customers of consignees that received the VITROS 5,1 FS Chemistry Systems as well as the VITROS Chemistry Products THC (Cannabinoids) and/or AMPH (Amphetamine) Reagents, were informed of the "field correction" by US Postal Service letter, provided with a new Assay Data Disk (ADD #5465), and were instructed to load it and recalibrate these assays, if necessary. |  | 1202 | Domestic: 604, International: 598 | Nationwide and Worldwide | Terminated | May 02, 2012 | 1812 | 1202 | Software | Output/Calculation | Software update | Software Update
Z-1165-2007 | 38035 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | One Touch SureSoft Lancing Device in carton. Part Number 021-140-01, aka OT SureSoft Hospital, Gentle 200 |  |  | March 30, 2007 | 2007 | Lifescan Inc | Lancet tip may not fully retract and cause needle stick and exposure to used lancets. | Nonconforming Material/Component | The firm has conducted a health hazard evaluation, and has initiated a corrective action plan. The firm initated recall on March 30, 2007, and notifications were sent to hospital end users known to Lifescan and its direct accounts. In addition, a customer service script has been developed to address any questions from incoming users. The firm has requested direct accounts to perform subrecalls on Lifescan's behalf to notify customers of the recall as needed. |  | 632 | 632 cartons (200 units per carton) | Product was distributed nationwide to distributors and wholesale dealers. | Terminated | December 03, 2007 | 248 | 632 | Not_Computer
Z-1169-2007 | 38352 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Microlab F.A.M.E. reagent containers, Part number 148482, box containing 48 containers each. For in vitro diagnostic use. |  |  | March 29, 2007 | 2007 | Hamilton Co | Some product containers may be leaking. | Other | The firm contacted consignees via certified mail and e-mail, initated on March 29, 2007. |  | 797 | 797 boxes | Product was distributed to 3 US consignees in NJ, WA and TX, with one consignee in France. | Terminated | December 21, 2007 | 267 | 797 | Not_Computer
Z-1172-2007 | 38096 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Non-AC-powererd Patient Lifts for assisting patients when bathing and toileting-sold under names "Bath Helper" and "Shower Helper" , Sisus Corporation, Santa Cruz, CA 95062. |  |  | June 01, 2007 | 2007 | The Sisus Corporation | Poor welding may cause the device(s) to fail, which may result in the user, falling in the bathroom and/or a lack of assistance with bathing or toileting. | Device Design | On 6/8/2007, a letter was sent to all dealers and each direct account customer/user followed by a phone call to ensure the notification letter was received. Customers were instructed to discontinue use of the devices pending receipt of modification kits/ instructions to be shipped priority overnight. The firm plans to call customers/dealers to support them with the installations and ensure smooth implementation. Dealers were instructed to conduct recall to the user level. |  | 100 | approximately 100 units have been distributed. | Nationwide Distribution including states of PA, TX, NY, MI, AL, FL, CA, OH, CT, WA, MD, OR, NC, NH, VT, IL, CO, HI, LA, GA, TN, OK, MA, KS, SC, ME, GA, MN, IN, and DE. | Terminated | August 19, 2008 | 445 | 100 | Not_Computer
Z-1175-2007 | 38396 | No | MID | Multiple autoantibodies immunological test system. | Immunology | Immunology | 510(k) | APhLÂż HRP IgG &amp; IgM ELISA Kit; Ref. #LAPL-K-HRP-00GM; distributed by Louisville APL Diagnostics, Inc., Doraville, GA 30360- Quality Control Fact Sheets for the Louisville APL Diagnostics, Inc. (address now in Seabrook, TX 77586). |  |  | October 06, 2005 | 2005 | Louisville Apl Diagnostics Inc | Mislabeing : Errors (control values) identified in the QC Fact Sheet for a specific lot of APhL HRP IgG and IgM test kits. | Other | Firm notified consignees of recall via fax on 05/05-06/2005. Consignees asked to replace Quality Control Fact Sheets shipped with test kits with the corrected Quality Control Fact Sheets provided with fax. |  | 29 | 29 test kits (ea. kit containing 1 IgG and 1 IgM test kit) | World Wide: Product distributed to direct accounts in CA, FL and GA. Foreign consignees located in Canada and Turkey. | Terminated | October 11, 2007 | 735 | 70 | Not_Computer
Z-1171-2007 | 38042 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cell-Dyn 1700CS (Closed Sampler) Hematology Analyzer, List Numbers 03H57-01 and 03H57-03, software controlled medical device for diagnostic use and CELL DYN Control Assay Disk, LIst Numbers 01H91-01 and 01H92-01 when used with above CELL DYN 1700, Abbott Laboratories Diagnostics Division |  |  | May 16, 2007 | 2007 | Abbott Laboratories | Abnormal results may not be "flagged" appropriately because of a software anomaly. (Flags for the Hemoglobin (HGB), MCH and MCHC parameters when the upper assay paramter limit has been exceeded) | Other | The firm initially sent reports (Product Information Letters) of this issue to the field on May 7, 2007. The firm sent a notice with customer notification response form on May 16, 2007 (Product Correction Letter). In addition, product information letters are being packed with the instruments accessory kit and quality control products used with the Cell-Dyn 1700 system. |  | 4197 | 4,197 units total (both models) | WORLDWIDE: Product was distributed nationwide to all 50 states, as well as worldwide to Anguilla, Argentina, Australia, Bermuda, Bolivia, Brazil, Canada, Cayman Islands, Chile, Columbia, Curacao, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Honduras, Hong Kong, Jamaica, Japan, Malaysia, Mexico, Panama, Paraguay, Peru, Puerto Rico, Singapore, South Korea, Taiwan, Thailand, Uruguay, and Venezuela. | Terminated | January 11, 2011 | 1336 | 4197 | Computer
Z-1173-2007 | 38359 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL77 Cal Pack, Model Number 944-069, Lot Number 17323, SenDx Medical Inc., A Radiometer Company |  |  | June 19, 2007 | 2007 | Sendx Medical Inc | Error in the Ca12 calibration value for pCO2 recorded in the cal pack barcode for lot 17323. This error will result in falsely elevated pCO2 results at the high end of the measuring range. | Other | A total of four (4) distributors were sent the recall package (Field Action Notice) by email on 6-29-07. The distributors were provided background information and specific instruction that all Cal packs from lot 17323 should be removed from customer sites and destroyed. The Recall package included a Management Cover for Field Action Notes, a Customer Letter and the Field Action Notice #915-261.All communications regarding field actions associated with this recall were distributed using the Radiometer Field Action Note (FAN) system. |  | 69 | 69 | Worldwide Distribution --- USA including state of OH and countries of China, Denmark Latvia, Mexico &amp; Spain. | Terminated | November 07, 2011 | 1602 | 69 | Not_Computer
Z-1184-2007 | 38409 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Arizant Healthcare Ranger Standard Flow Blood/Fluid Warming Disposable Sets with injection port and extension set, Part Number 24250. For use with a Ranger warming unit. |  |  | June 26, 2007 | 2007 | Arizant Healthcare, Inc | Incomplete seal on the product pouch; with the seal incomplete, the sterility of the product inside the pouch cannot be guaranteed. | Other | All consignees will be contacted by Urgent Medical Device Recall letter on 6/26/07. |  | 12000 | 12,000 | Worldwide Distribution --- USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL., IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and countries of United Kingdom, France, Switzerland, Germany, India, Canada, Australia, Netherlands, Italy, Sweeden, South Korea, and Greece. | Terminated | March 02, 2008 | 250 | 97710 | Not_Computer
Z-1187-2007 | 38167 | No | LGI | Neuroleptic drugs radioreceptor assay test system. | Toxicology | Toxicology | 510(k) | XSYSTEMS Quinidine Calibrators consists of six bottles of Cal A-F (2.5 mL each).-Product List Number-9506-01 |  |  | May 31, 2007 | 2007 | Abbott Diagnostic International, Ltd. | Controls Out of Range/Specification. | Other | May 31, 2007, Product Correction Letter to Customers were instructed to: "Revise the expiration date for all lots identified in the customer letter. "Discontinue use of the expired lots, and discard the material according to their laboratory procedures. "An additional option is to use the Abbott Immunoassay Multiconstituent Control (List Number 06E20), manufactured by BIO-RAD. |  | 250 | 250 | Worldwide: USA including states of AR, AZ, CA, CT, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NY, NV, OH, OK, OR, SC, TN, TX, VA, WA, WI, WV, Hawaii and Puerto Rico. International: CANADA, GERMANY, JAPAN, SINGAPORE, AUSTRALIA, and BARBADOS | Terminated | June 10, 2011 | 1471 | 729 | Not_Computer
Z-1178-2007 | 38251 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | PhaSeal, Protector 14, Item No. P14, US Order No. 21320 |  |  | June 26, 2007 | 2007 | Carmel Pharma, Inc. | Risk of Leakage: When the PhaSeal Protector and Injector are used with drug vials with thick rubber stoppers which have a concave upper surface and a flat bottom surface there is a potential for leakage. These rubber stoppers are made of a harder rubber formulation and may be more difficult to penetrate. | Other | On or about 6/15/2007, the firm sent advisory notice letters to the end users and the instructions for use have been updated by adding a precaution statement regarding the use of thick rubber stoppers. The firm is also updating the customer training program to include how to handle drug vials with thick concave rubber stoppers. |  | 6100000 | 6.1 million for Z-1178-1181-2007 recalls | Product was distributed throughout the US and Canada | Terminated | March 05, 2013 | 2079 | 6100000 | Not_Computer
Z-1191-2007 | 38351 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Polaris Dx Diagnostic Catheter, for cardiac electrophysiology studies, Model M0049663S0 |  |  | June 26, 2007 | 2007 | Boston Scientific Corp | Product is mislabeled as hexapolar (6 electrodes), but is actually quadripolar (4 electrodes). | Labeling mix-ups | The firm issued notification letters to its consignees, sent 06/26/2007, via FedEx. The firm intends to remove the product and destroy it. |  | 30 | 30 units | Product was distributed to 3 consignees in Florida and Tennessee. | Terminated | November 27, 2007 | 154 | 30 | Not_Computer
Z-1174-2007 | 38384 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Informatics System, software version 3.0, Part Number 1000034941, DADE BEHRING INC., Newark, DE. |  |  | July 03, 2007 | 2007 | Dade Behring, Inc. | Sample ID and Patient results could be associated with the wrong patient | Other | The recalling firm issued a Field Corrective Action letter dated 7/3/07 to inform their customers of the problem and that a software patch would be sent for them to download. The recalling firm plans to telephone the customers after the patch has been released to verify that it has been installed. |  | 72 | 72 units | Worldwide: USA including states of AZ, CA, FL, IL, MA. MI, MO, NC, OH, PA, TX, and WV. The product was also shipped to other Dade facilities in MO, DE, IL, CT, and Belgium. | Terminated | November 07, 2007 | 127 | 72 | Software | Output/Calculation | Software update | Software Update
Z-1189-2007 | 38433 | No | MBD |  |  | Orthopedic | PMA | BacT/ALERTÂż FN Culture Bottles, Product Number 259793 |  |  | February 16, 2007 | 2007 | bioMerieux, Inc. | Microbial Contamination: Bacillus sp. contamination was detected in inoculated bottles. | Other | Consignees were notified by Urgent Product Recall letter on 02/19/2007. They were instructed to destroy any unused inventory of this product lot that they have on hand. |  | 2403 | 2,403 cases (100 bottles per case) | Worldwide; USA, Austria, Bulgaria, Caledonia, China, Croatia, Curacao, Finland, France, Germany, Greece, Guadeloupe, Italy, Latvia, Mexico, Netherlands, Poland, Portugal, Spain, Saudi Arabia, Sultanate of Oman, Switzerland, Sweden, and United Kingdom | Terminated | December 08, 2009 | 1026 | 2403 | Not_Computer
Z-1193-2007 | 38031 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | Amsco Century Medium Steam Sterilizer (26" x 37.5") (660 x 950mm). |  |  | March 22, 2007 | 2007 | Steris Corp | The cables inside the pump conduit box may be exposed (not covered with wire nut and/or electrical tape) that could cause a potential electrical shock hazard. | Employee error | By telephone 3/22/2007, customers were instructed to cease using the product and a service technician would be dispatched immediately to repair the unit. |  | 78 | 78 | Worldwide, including USA, Denmark, Belgium, Greece, India, and Russia.There are no Canadian accounts. | Terminated | September 28, 2010 | 1286 | 78 | Not_Computer
Z-1194-2007 | 38462 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Anatomical Shoulder Prosthesis, Glenoid, Cemented, sterile; REF 01.04214.370. |  |  | June 25, 2007 | 2007 | Zimmer Inc. | The package actually contains a knee patellar implant. | Other | The firm notifed the distributor via letter dated 6/25/07, and sent by email, to retrieve the product from their customers and to return the product. |  | 10 | 10 | Maryland. | Terminated | December 05, 2007 | 163 | 10 | Not_Computer
Z-1190-2007 | 38324 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano s5 Imaging system, Ultrasound pulsed echo imaging system. Product consists of the CPU, control console, monitor, printer, and patient interface module (PIM) |  |  | April 02, 2007 | 2007 | Volcano Corporation | During emissions testing, it was discovered that the product exceeds emissions limits for electromagnetic compatibility. | Other | The firm notified its customers by letter beginning April 2, 2007, and started replacing the monitor with appropriate replacement parts. |  | 351 | 351 units worldwide | Product was distributed to health care facilities nationwide and to distributors in Europe (via subsidiary), Japan (via subsidiary), Latin America and Asia (third party distributors). | Terminated | June 03, 2008 | 428 | 351 | Not_Computer
Z-1196-2007 | 38172 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL80 FLEX Analyzer, Model #393-839 |  |  | March 27, 2007 | 2007 | Sendx Medical Inc | There have been reported experiences where waste fluids have sprayed out from the waste drain at the front of the ABL80 analyzer when excessive pressure builds within the waste line. | Other | The recall was initiated 3/27/2007. All distributors were notified via the Radiometer FAN (field action notice) system which provides a communication mechanism for these types of actions. The notice advises that there have been a few reported experiences where waste fluids have sprayed out from the waste drain at the front of the ABL80 analyzer which occurs when excessive pressure builds within the waste line. To prevent any spraying of fluids from the waste drain a check valve has been added to the waste tubing, between the waste drain and the waste pump which allows fluids to flowin one direction thus making it impossible to spray fluids out the waste drain. The letter further advises that all analyzers shipped from SenDx on or after March 16, 2007 will include this check valve. All analyzers already in the field must be updated to include this check valve. Distributors are to distribute a copy of the Customer Advisory Letter associated with this case to all their customers and update all ABL80 analyzers to include the waste drain check valve as per the instructions that accompany the check valve. The Advisory Letter informs the customer of the problem and the new waste drain check valve and provides precautionary instructions such as appropriate hand, eye, face and clothing shields. In addition, it advises that the analyzer should never be used if there is any indication of a waste drain occlusion or blockage. The most typical indication of an occlusion or blockage is fluid overflowing out of the waste drain. |  | 334 | 334 units worldwide | Class II Recall - Worldwide Distribution --- USA including state of Ohio, and countries of Austria, Belgium, Denmark, Finland, France, Germany, Italy, Japan, The Netherlands, Norway, Poland, Spain, Sweden, Switzerland, United Kingdom, Australia, New Zealand, Canada, and Denmark. | Terminated | October 13, 2011 | 1661 | 334 | Not_Computer
Z-1197-2007 | 38382 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Atkinson Retrobulbar Anesthesia Needles, 25 gauge 1 1/2". Catalog Number 8065420920. The product is sold in boxes that contain 10 needles each. |  |  | July 11, 2007 | 2007 | Alcon Research, Ltd | Mislabeled: boxes actually contained Tri-Fact needles | Process control | The recalling firm telephoned the hospital on 7/11/07 to inform them of the problem and the need to return the product. As a follow up to the telephone call, the recalling firm issued a recall letter via FedEx on 7/11/07. The firm's representative is making arrangements to pick up the recalled product. |  | 2 | 2 boxes | CA. | Terminated | August 28, 2007 | 48 | 2 | Not_Computer
Z-1202-2007 | 37904 | No | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | Abbott AxSYM System Digoxin III Reagent Pack |  |  | March 30, 2007 | 2007 | Abbott Diagnostic International, Ltd. | Customers are receiving error codes (1062, 1063, 1113, &amp; 1118) when running patient samples on the AxSYM Digoxin III assay. | Other | On Abbott letterhead dated March 30, 2007 - The instructions are to provide a copy of the accompanying Product Correction Letter to the laboratory manager/supervisor responsible for AxSYM Digoxin II or AxSYM Digoxin III testing. The laboratory manager/supervisor should complete the information on the letter acknowledging receipt of the product Correction Letter and Fax the Customer Reply Form, prior April 12, 2007 to Fax number 1-800-777-0051. |  | 2579 | 2,579 kits | Worldwide, including USA, Puerto Rico, Canada, Uruguay, Germany, Japan, Hong Kong, South Korea, Australia, and New Zealand. | Terminated | June 10, 2011 | 1533 | 2579 | Not_Computer
Z-1201-2007 | 38350 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL8O FLEX Software Version 1.10 included in the ABL8O Flex Analyzer, Model Number 914863 |  |  | June 04, 2007 | 2007 | Sendx Medical Inc | The ABL8O FLEX software version 1.10 contains errors in the derived parameter calculation for cBase(B,ox) and cBase(Ecf,ox). These incorrect calculations result in a mild to moderate lowering of the base excess value, the severity depending on the concentration of hemoglobin and the oxygen saturation in the sample. Elevated hemoglobin concentrations have the greatest impact on this calculation err | Other | A total of eight (8) distributors and the Radiometer main office in Denmark were sent the recall package (Field Action Notice) by email on 6-6-07. The distributors were provided background information and specific instruction on the steps to perform to correct the problem. The Recall package included a Management Cover for Field Action Notes and the Field Action Notice.Radiometer has released a subsequent software upgrade (v.1.11) to correct the two problems noted in the Recall Notice. |  | 165 | 165 software upgrade CDs | In US to Cleveland, OH and Worldwide to Belgium, Denmark, France, Germany, The Netherlands, Sweden &amp; United Kingdom | Terminated | February 29, 2012 | 1731 | 165 | Software | Output/Calculation | Software update | Software Update
Z-1218-2007 | 38512 | No | IZX | Diagnostic x-ray beam-limiting device. | Radiology | Radiology | 510(K) Exempt | PCS125 Swivel Collimator, Model F100; Manual Radiographic Collimator; Summit Industries, Inc., Chicago, Illinois |  |  | August 10, 2007 | 2007 | Summit Industries Inc. | The F100 Swivel Collimator may come loose from the tubestand and fall onto the x-ray table. | Other | The affected dealers were notified of the field correction recall by letter dated 8/3/07, informing them of the problem with the stripped screws and requesting that they notify their affected end user locations. The end user notification letters inform them of the problem and state that the dealer will be contacting them soon to arrange a time to visit the end user location and conduct the rework on the collimator, which should take less than 30 minutes. In the mean time the end user was instructed to examine their swivel collimator for looseness. If it is found to be loose, the end user was instructed to take the unit out of service and call the dealer. |  | 3034 | 3034 for all units | Worldwide, including USA, Australia, Cambodia, Canada, Chile, Colombia, Dominican Republic, Ecuador, France, Germany, Haiti, Mexico, Nicaragua, Singapore, and South Africa. | Terminated | March 18, 2010 | 951 | 3034 | Not_Computer
Z-1195-2007 | 38349 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian On-Board Imager Device versions 1.2 and 1.3, when used with Cone Beam Computed Tomography. Medical Linear Accelerator. |  |  | April 18, 2007 | 2007 | Varian Medical Systems Inc | Patient positioning errors may occur because of mechanical variations during product installation. | Device Design | The firm issued a notification letter to all customers April 18, 2007. The firm intends to arrange for service personnel to contact affected sites to verify the calibration of the OBI system, and make arrangements to correct any issues. The firm will also update installation manuals to clearly describe the OBI installation and calibration requirements. |  | 291 | 291 units | Worldwide, including USA, Puerto Rico, Australia, Belgium, Canada, China, Colombia, Czech Republic, Denmark, Finland, France, Germany, India, Ireland, Italy, Japan, Luxembourg, New Zealand, Norway, Poland, Portugal, Spain, Sweden, Switzerland, Netherlands, and United Kingdom. | Terminated | August 12, 2010 | 1212 | 291 | Not_Computer
Z-1200-2007 | 38322 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott CELL-DYN 3700 Hematology Analyzer |  |  | May 11, 2007 | 2007 | Abbott Laboratories | The product failed electrostatic discharge testing, which could result in product failure. | Other | Mandatory Technical Service Bulletins (TSB) are being issued to Field Service to upgrade all instruments in the field. All customers will be contacted by Field Service to schedule this mandatory upgrade. The upgrades are expected to take 12 months. |  | 3429 | 3,429 units, all models | Worldwide, including USA, Puerto Rico, Canada, Mexico, Argentina, Australia, Bahamas, Brazil, Chile, China, Colombia, Costa Rica, Dominican Republic, El Salvador, Ecuador, Germany, Guatemala, Hong Kong, Jamaica, Japan, New Zealand, Panama, Paraguay, Peru, Singapore, South Korea, Taiwan, Thailand, Uruguay, and Venezuela. | Terminated | June 19, 2008 | 405 | 3429 | Not_Computer
Z-1205-2007 | 39374 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Assorted non- sterile Tracecarts (with sterile component(s)); |  |  | July 03, 2007 | 2007 | DeRoyal Lafollette | Packaging failure. Lack of sterility assurance | Other | Written notification of the recall was sent (via UPS Ground) beginning 7/03/2007 to consignees to advise them of problem and the recall. The letter stated that the recall was to the User level (which in this case, are the hospitals) and requested return of the product. |  | 112 | 112 cases | Nationwide | Terminated | June 12, 2008 | 345 | 2135 | Not_Computer
Z-1206-2007 | 38898 | No | LZK | Chin prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Spectrum Designs Medical Anatomical I (Silhouette) Malar Facial Implant, Catalog Number 120-224S. |  |  | June 15, 2007 | 2007 | Spectrum Designs Inc | The product is being recalled because it was discovered the at least one unit in the lot contained a mismatched pair (two rights and no left). Since the lot was produced with equal numbers of pairs, it is assumed that one other unit was packaged with two lefts and no right. This was discovered through a physician complaint. | Other | The firm initiated the recall on 06/15/2007 by contacting their customers by telephone. The firm followed-up by sending recall letters via FedEx or certified mail with signature confirmation to each direct account who received the recalled product on 06/18/2007. |  | 15 | 15 pairs (left/right) | Nationwide. | Terminated | March 07, 2012 | 1727 | 15 | Not_Computer
Z-1211-2007 | 39147 | No | MAF |  |  | Cardiovascular | PMA | Multi-Link Vision RX Coronary Stent System, 3.0 x 15mm, Part number 1007848-15, lot number 7052951 |  |  | July 16, 2007 | 2007 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | The affected single lot did not meet an internal specification. One unit failed the test criteria for stent dislodgment testing. | Other | The firm initiated the recall on 07/16/2007 by sending recall letters with attached Customer Reply form via FedEx to each direct account that received the recalled product. Customers were informed of the issue and given instructions to work with their local account representative to review their inventory, remove the recalled product, complete the Customer Reply form and return the recalled product to Abbott Vascular. |  | 43 | 43 units | Nationwide | Terminated | June 14, 2011 | 1429 | 43 | Not_Computer
Z-1215-2007 | 38306 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Smiths CADD Medication Cassette Reservoir, With clamp and female Luer, Nonvented stopper included, 100 ml PC |  |  | June 18, 2007 | 2007 | Smiths Medical MD, Inc. | Smith Medical has become aware that an increase in pH may occur with some medications when instilled in the Affected Cassettes. Also, a small number of catheter occlusions have been reported from one customer when these cassettes have been used for delivering Flolan. | Other | All customers who have received the CADD Medication Cassette reservoirs have been sent the Product Recall Notification Letter on June 18, 2007 making them aware of the potential issue. Customers will be required to return the affected devices to Smiths Medical. All returned recall affected product will be quarantined and then destroyed. The consignees are to fill and fax "Attachment 1" form in regard to affected product inventory to Smiths Medical MD. |  | 113143 | 113,143 | Worldwide, including USA, Canada, Mexico, Sweden, Germany, Japan, Australia, Taiwan, Poland, Jordan, Czech Republic, Denmark, Paraguay, Hong Kong, Singapore, Argentina, Italy, Venezuela, Greece, Norway, Pakistan, Puerto Rico, Thailand, United Arab Emirates, Netherlands Antilles, Malaysia, Panama, Hungary, Saudi Arabia, Portugal, South Africa, Denmark, the Netherlands, France, United Kingdom, Great Britain, Spain, Finland, Guam, Virgin Islands, China, Israel,Austria, and Switzerland. | Terminated | February 10, 2011 | 1333 | 153798 | Not_Computer
Z-1207-2007 | 38308 | No | GZJ | Transcutaneous electrical nerve stimulator for pain relief. | Neurology | Neurology | 510(k) | Avazzia Tennant BiomodulatorÂż; Model: BEST AV-1; P/N #TB04303-103; manufactured by Avazzia, Inc., Dallas, TX 75240. |  |  | June 29, 2007 | 2007 | Avazzia, Inc | Plastic Cases may crack due to manufacturing by the supplier not meeting specifications | Other | Firm notified consignees by letter on June 29, 2007 requesting that all affected product be returned to the firm for repair or replacement. The firm further asked for distributors who had shipped affected product to notify receiving patients so they could arrange for repair or replacement. |  | 134 | 134 units. | Nationwide including states of Texas, Pennsylvania, Louisiana, Wyoming, and California | Terminated | April 04, 2008 | 280 | 556 | Not_Computer
Z-1168-2007 | 38273 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Urological Black Silicone Filiform Double Pigtail Ureteral Stent Set, 7.0Fr/24cm, sterile, Order Number (REF) 133724. |  |  | June 29, 2007 | 2007 | Vance Products Incorporated | The radiographic band may detach from the stent and remain in the patient after removal of the stent. | Other | Consignees were notified of the recall via letter dated 6//29/07, which instructed them to quarantine stock on hand and to fax number on hand back to Cook for return authorization. |  | 75 | 75 | Worldwide, including USA, Canada, and Japan. | Terminated | June 11, 2008 | 348 | 75 | Not_Computer
Z-1182-2007 | 38400 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Intellivue Central Station Software for the: Intellivue Information Center; Philips Information Center, Central Station |  |  | July 16, 2007 | 2007 | Philips Medical Systems | Incorrect Patient Assignment: the information center will in most cases be associated with the "incorrect" patient. This occurs on both standalone and networked systems running fixed mode monitoring. (This does not occur on systems that run flex mode monitoring) | Other | Philips Medical Systems issued a letter notification dated July 2007 " Urgent Device Correction". Customers with networked information centers will be corrected through a database server. Stand-Alone units will have a software correction. Software version J.00.25 will correct the problem. Users are advised not to use th 12 Lead Analysis and Export Option in fixed or flex mode monitoring unitl the software correction is made. |  | 76 | 76 units | Worldwide Distribution: USA including states of NJ, NY, OK, MI, and UT and Foreign: Australia, Finland, France, Netherlands, and the UK | Open, Classified |  |  | 76 | Software | Output/Calculation | Software update | Software Update
Z-1203-2007 | 38428 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Kion Anesthesia Workstation |  |  | July 10, 2007 | 2007 | MAQUET Inc. | Overcharging of batteries and the release of acid fumes | Other | Safety Alert letters were sent out by FedEx tracking.A Service Express letter will be sent out emphasizing that only qualified batteries must be used in KION equipment and will list all qualified battery types.A Field Action - update of existing operating manuals will include warning on battery charging on KION. This action is currently scheduled to start 9/1/07 and be completed in US by 10/1/07. After FDA consultation the recalling firm issued a 9/5/2007 URGENT DEVICE CORRECTION Letter stating that KION will visit consignees and replace batteries older than three years, verify that the batteries are ones of the approved type, and verify battery condition. |  | 9 | 9 units in US | Nationwide: 2 units at the National Institute of Health in Bethesda, MD and 7 units at Oconee Regional Medical Center, Milledgeville, GA | Terminated | December 07, 2007 | 150 | 9 | Battery | Physical Safety Hazards | Replace | Remove or Replace
Z-1219-2007 | 39362 | [6 empty] | Biogel Skinsense Polyisoprene Surgical Gloves, Non-Latex, Sterile, (Size 8) Made in Malaysia |  |  | July 20, 2007 | 2007 | Molnlycke Health Care, Inc | Sterility compromised (package integrity): The primary packaging for a portion of the lot may contain gloves that were not sealed properly. | Packaging process control | Consignees were notified by registered mail on 07/23/2007. The notice contained a Recall Response Form to confirm receipt of the notice and to verify if the product is still in inventory. The notice instructed to consignee to inspect their inventory for Biogel Skinsense Polyisoprene Surgical Gove, lot #07B026, Size 8. If this lot and size are in their inventory, they were instructed to discontinue use, quarantine the product and return the product to Molnlycke Healthcare. |  | 6000 | 6,000 pair | Nationwide including states of AZ, CA, CT, KY, MD, MN, MO, TN, and WI | Terminated | March 17, 2010 | 971 | 6000 | Not_Computer
Z-1234-2007 | 44721 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The product is an isolation transformer, an accessory to Computed Tomography Scanner. It consists of electronic circuitry enclosed in a metal case. According to Philips Medical Systems, the metal case is manufactured from self-extinguishing materials. The unit is labeled in part, "Power Conditioner *** Part Number 4450028-XX *** Frequency: *** S/N: *** Rev: *** MFG date: *** INPUT RATINGS *** Model: MCT 100/480 ISOTRAN PLUS *** OUTPUT RATINGS *** MEDICAL POWER SUPPLY UL *** TEAL Electronics Corporation, San Diego, California ***". |  |  | April 27, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | The TEAL Power Conditioning units may fail and result in the unit emitting fire and/or smoke. Note: the TEAL Power Conditioning Unit has the potential to fail when the scanner that is attached to the TEAL is powered and in use as well as when not in use. | Component design/selection | On 9/04/2007, Philips Medical Systems (PMS) sent a letter via courier to all affected customers. The letter notified customers about the failures of the TEAL power conditioner and the possible emisson of fire and/or smoke from the units. In addition, the letter notes that PMS implemented a Field Change Order in May, 2007 that directs onsite repair of all affected units by PMS service technicians. |  | 491 | 491 | Worldwide, including, USA, Canada, China, Italy, Russia, Brazil, Belgium, Germany, Austria, Australia, Netherlands, France, Korea, Denmark, Israel, England, Argentina, Mexico, Spain, Yemen, Greece, Egypt, Japan, UAE, &amp; Chile. | Terminated | March 02, 2012 | 1771 | 491 | Not_Computer
Z-1232-2007 | 38635 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Clinical Chemistry Lipase, List Numbers: 7D80-20 and 7D80-30 |  |  | June 01, 2007 | 2007 | Abbott Laboratories Inc. | Lipase results are falsely elevated when a MULTIGENT Acetaminophen test is pipetted immediately before a Lipase test from the same sample cup or tube on the ARCHITECT c8000. | Other | The firm initiated the recall on 06/01/2007 by sending recall letters with attached Customer Reply form via FedEx to each direct account that received the recalled product. Customers were informed of the issue and given instructions for modification to the SmartWash feature in the Lipase Assay Configuration in order to prevent Acetaminophen from contaminating Lipase. |  | 7260 | 7260 | Worldwide, including USA, Australia, Brazil, Canada, Chile, El Salvador, Dominican Republic, Ecuador, Germany, Hong Kong, Japan, Mexico, New Zealand, Panama, Puerto Rico, Singapore, South Korea, Thailand, and Venezuela. | Terminated | April 05, 2012 | 1770 | 7602 | Not_Computer
Z-1222-2007 | 38439 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Prolactin MasterCheck; a single analyte control - each kit contains 1 - 2 mL bottle of Architect Prolactin MasterCheck Level 0 containing HEPES buffer with protein (bovine) stabilizers, and 4 - 2 mL bottles of Architect Prolactin MasterCheck Levels 1, 2, 3 and 4 containing prolactin (human) prepared in HEPES buffer with protein (bovine) stabilizers; Abbott Laboratories, Abbott Park, IL 60064, list 6C27-05 |  |  | July 06, 2007 | 2007 | Abbott Laboratories | The Architect Prolactin MasterChecks may not be stable at the recommended storage conditions of 2 to 8 degrees C through the expiration dating (15 months). The MasterChecks may exceed the suggested tolerance limits listed in the Reportable Range/Calibration Verification section in Appendix B of the Architect iSystem Operations Manual. | Other | Abbott sent a Product Correction letter dated 7/6/07, to all customers who received the affected product lot, informing them that the ARCHITECT Prolactin MasterChecks may not be stable at the recommended storage condition of 2 to 8 degrees Celsius (C) through expiration dating (15 months). Labeling is being permanently updated to indicate that ARCHITECT Prolactin MasterChecks are to be stored frozen until first use. After thawing, ARCHITECT Prolactin MasterChecks can be stored unopened at 2-8 degrees C for thirty days or until expiration which ever occurs first. All future shipments will contain updated storage and handling conditions of the ARCHITECT Prolactin MasterCheck. They were requested to check their inventory to see if they have any remaining inventory of lot number 82900M. If they know the date the lot was received or the date it was thawed, they may continue to use the ARCHITECT Prolactin MasterChecks if it has been stored at 2 to 8 degrees C unopened for less than 30 days. If the material has been stored at 2 to 8 degrees C for greater than 30 days or they do not know the number of days it has been thawed, they were instructed to discard the material. Reimbursement for discarded kits will be provided based on information provided in the Customer Reply form. Until labeling is updated, it will be necessary to use an alternate material for verification of sensitivity, calibration linearity, and the reportable range of the ARCHITECT Prolactin assay. They were requested to forward a copy of the letter to any laboratory to whom they may have forwarded the kits, and to retain a copy of the letter for their records. Any questions were directed to Abbott Customer Support at 1-877-4ABBOTT in the U.S. The accounts were requested to complete and fax back the customer reply form, acknowledging the receipt and understanding of the correction letter. Customers outside of the U.S., were requested to contact their local customer support representative. |  | 57 | 57 kits | Nationwide and Canada. | Terminated | May 18, 2008 | 317 | 57 | Not_Computer
Z-1221-2007 | 38983 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Product is labeled as LTX XP Miniplant, Dental Implant, Sterile. |  |  | August 03, 2007 | 2007 | Biomet 3i | The LTX XP Miniplant, BIOMET 3i dental implant, was labeled incorrectly. The 13MM implants were labeled and packaged as 11.5MM implants and 11.5MM implants were labeled and packaged as 13MM implants. | Labeling mix-ups | International customers were contacted first by e-mail , then by fax letter 08/03/2007. |  | 240 | 240 (120 each model / lot) | Product was not distributed within the U.S. Product was distributed Internationally - Brazil, Argentina, Israel, Mexico, Colombia, Korea, Taiwan, Uruguay, Italy, Chile. | Terminated | September 11, 2009 | 770 | 240 | Not_Computer
Z-1220-2007 | 38348 | No | MQK | Assisted reproduction labware. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | IVF Pasteur Pipettes, glass, cotton plugged, length 9", in boxes of 90 consisting of 30 nylon tube packages of 3 pipettes each. Item # PP-9-90PL. Boxes are labeled in part ***Humagen Fertility Diagnostics, Inc. 2400 Hunter's Way Charlottesville, VA 22911*** |  |  | May 08, 2007 | 2007 | Humagen Fertility Diagnostics, Inc. | Pasteur pipettes for laboratory use were labeled as sterile but may have become contaminated. | Other | The recalling firm notified consignees by phone and e-mail on 05/07/07. International consignees were advised to discard product for replacement. |  | 160 | 160 boxes | The recalled product was distributed to 2 direct accounts in IL and to 6 international distributors in United Kingdom, Greece, Australia, United Arab Emirates, the Netherlands, and Israel. | Terminated | October 29, 2007 | 174 | 160 | Not_Computer
Z-1192-2007 | 38630 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT50 Ventilator (all Models) |  |  | July 31, 2007 | 2007 | Newport Medical Instruments Inc | Battery Life: With use over time, the internal battery begins losing its ability to hold a charge resulting in shortened ventilator operation time. Specifically, Newport has received a number of reports concerning the internal battery not lasting as long as expected followed by ventilator shut down with less advance warning than when the battery was new. | Other | An Important Medical Device Correction Letter was issued on September 14, 2007 as a short term fix. The firm has submitted a long term correction as a supplement to its 510 K which is pending review by FDA. |  | 8160 | 8,160 | Worldwide including USA, Algeria, Argentina, Armenia, Australia, Belgium, Bolivia, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cote D' Ivoire, Croatia, Cyprus, Dominican Republic, Ecuador, Egypt, El Salvador, England, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kenya, Korea, Kuwait, Liechtenstein, Lithuania, Malaysia, Mexico, Netherlands, New Caldonia, New Zealand, Oman, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Switzerland, Syria, Taiwan, Thailand, Turkey, UAE, Ukraine, United Kingdom, Venuzuela, Yemen, and Zambia | Terminated | March 21, 2012 | 1695 | 8160 | Battery | Device Operation | Repair | Repair
Z-1224-2007 | 39415 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague CX Single Channel Volumetric Infusion Pumps; Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015 U.S.A., Made in Singapore; product code 2M8161 |  |  | August 15, 2007 | 2007 | Baxter Healthcare Corp. | Some of the pumps were upgraded by one service technician who failed to replace prisms and batteries as required in the upgrade. | Other | The affected hospitals were sent an Urgent Colleague Deployment Correction letter dated 8/15/07 via overnight mail. The hospitals were informed that some of the Colleague pumps at their facility may not have been completely upgraded to include all the necessary components required as part of Baxter's Corrective Action Plan (CAP). The hospitals were provided a list of affected pump serial numbers and were requested to remove the pumps from clinical use and notify all end users of the action. The accounts were requested to complete the enclosed customer reply form, acknowledging the receipt and understanding of the letter and fax it back to Baxter at 1-847-270-5457. The hospitals were also requested to report any malfunction or adverse event related to the Colleague Infusion Pumps to Baxter Corporate Product Surveillance at 1-800-437-5176. |  | 42 | 42 pumps | Florida and California | Terminated | February 21, 2008 | 190 | 42 | Not_Computer
Z-1223-2007 | 39368 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 20 Automated External Defibrillator; a portable, 12V internal battery powered defibrillator; Manufactured by MRL Inc., A Welch Allyn Company, Buffalo Grove, Illinois; Model AED 20, Catalog Numbers 972213, 972214, 972215, 972216 (Part Numbers 972200E, 972211E, 972212E, 972213E, 972214E, 972215E, 972216E) |  |  | August 23, 2007 | 2007 | Welch Allyn Protocol, Inc | Failure to deliver therapy: The Welch Allyn AED20 Defibrillators may display a "Defib Comm" error message on the device display during use which may result in failure of the device to analyze the patient's ECG and deliver the appropriate therapy. | Process control | On 8/23/07 the firm sent a press release to the Associated Press for release. On 8/29/07 the firm began sending URGENT MEDICAL DEVICE RECALL letters to their consignees informing them of the possibility of the defibrillator to fail to resuscitate the patient. The accounts were asked to respond to the notification within 5 working days by faxing or mailing an enclosed response. End users will be provided with a loaner device while their unit is being serviced. Distributors were sent a separate letter request they provide their customer names to the firm so that the firm can send letters to those customers. |  | 1637 | 1,637 devices | Worldwide to include USA, Australia, New Zealand, South Africa, China, Hong Kong, Taiwan, Malaysia, Singapore, Phillippines, South Korea, Viet Nam, Egypt, France, Germany, Ireland, India, Israel, Pakistan, Poland, Russia, Saudi Arabia, Switzerland, United Arab Emirates, United Kingdom, Canada, Ecuador, Puerto Rico, and South America. | Terminated | September 26, 2011 | 1495 | 1637 | Other | Treatment/Delivery/Therapy | Return | Remove or Replace
Z-1226-2007 | 38399 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | ReliOn 0.5cc 30 g x 8mm Blister Pak U-100 syringe |  |  | June 21, 2007 | 2007 | Becton Dickinson &amp; Company | Complaints received indicating open seals on blister paks, potentially compromising the sterility of the product. | Other | Urgent Product Recall notifications were sent on 6/21/07 asking them to discontinue use of product, remove from inventory and complete Tracking Verification form. |  | 644500 | 644,500 units | Nationwide. | Terminated | June 03, 2008 | 348 | 644500 | Not_Computer
Z-1228-2007 | 38546 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage Cartridge Express, CAR-170, for hemodialysis. |  |  | July 27, 2007 | 2007 | NxStage Medical, Inc. | Potential for dialysate leaks from the cartridge during treatment. | Process change control | NxStage initiated the recall by letter on 8/24/07 advising users of the specific lots to be removed from use. Prior notifications were issued on July 27, 2007, July 30, 2007 and July 31, 2007 notifying users of the potential for leaks and instructions to detect leaks prior to use. |  | 12085 | 12,085 cases | Nationwide | Terminated | September 13, 2010 | 1144 | 12085 | Not_Computer
Z-1225-2007 | 38424 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | BMSI "MicroJack" PN 672-510000 patient interface electrode connection device used during electroencephalography (EEG) /Epilepsy Monitoring evaluations with various EEG monitoring systems. The MicroJack is used with BMSI Models 2000, 3000, 4000, 5000 &amp; 6000, and is used with the BMSI kit upgraded with NicoletOne. Nicolet Biomedical, Inc. |  |  | July 17, 2007 | 2007 | Nicolet Biomedical Div of Viasys Healthcare | The BMSI "MicroJack" patient interface electrode connection device used with the EEG monitor systems was found to be mis-wired. A complainant alleged that during depth electrode EEG study of a patient in preparation for brain surgery there was an observation that was perceived as non-logical data output. | Other | VIASYS will initiate a field correction for units that have been distributed. The correction consists of testing the units and replacing those that have been found to contain the defect. Method of communication will be by certiified mail to identify consignees and receipt of the communication will be tracked. Letters "Viasys HealthCare Urgent Medical Device Field Correction (Recall)" were sent 7/17/07 to both domestic and international customers. |  | 2102 | 2102 | Worldwide, including USA, ISRAEL, LITHUANIA, RUSSIA, FRANCE, IRELAND, UNITED KINGDOM , SPAIN, YEMEN, JAPAN, TAIWAN, KOREA, CHINA, SAUDI ARABIA, HONG KONG, INDIA, COLOMBIA, BRAZIL, GERMANY, CANADA, DENMARK, ITALY, MEXICO, and THAILAND. | Terminated | December 10, 2011 | 1607 | 2102 | Not_Computer
Z-1236-2007 | 39386 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Normal Saline 10ml in 12ml Syringe. The product is shipped in 120 units per carton. |  |  | July 30, 2007 | 2007 | B. Braun Medical, Inc. | Particulate matter in product. | Other | The recalling firm issued a Product Removal Notice to the distributors on 7/31/07 informing them of the problem and the need to contact their customers. Product should be returned to the recalling firm for destruction. On 9/14/07, the recalling firm issued a press release to alert any consumer that may have the product of the particulate problem. |  | 1292520 | 1,292,520 units | The products were shipped to distributors nationwide. | Terminated | April 01, 2008 | 246 | 1325760 | Not_Computer
Z-1230-2007 | 38623 | No | MUU | Suction lipoplasty system. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 5mm x 30cm Hollow "Golf" Tip Ultrasonic Aspiration Probe, Model 5620-530, Catalog Number 5620-530. |  |  | July 27, 2007 | 2007 | Mentor Texas, Inc | Mis-identified and mis-labeled surgical probes were distributed. 5mm x 30cm Hollow "Golf" Tip Ultrasonic Aspiration Probe, Model 5620-530, Catalog Number 5620-530 was shipped, instead of LySonix 5mm x 30cm Hollow Bullet Tip Ultrasonic Probe, Catalog Number 5610-530. | Other | Firm notified consignees by letter on July 27, 2007. |  | 4 | 4 units | Nationwide, including VA, IN, CA, and IL. | Terminated | April 04, 2008 | 252 | 4 | Not_Computer
Z-1237-2007 | 39413 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | 25-Hydroxyvitamin D 125 I RIA Kit, in vitro diagnostic. REF:68100E DiaSorin Stillwater, Minnesota 55082-0285, U.S.A. |  |  | August 21, 2007 | 2007 | Diasorin Inc. | 25 Hydroxyvitamin D RIA kits may contain the incorrect lot of tracer. | Other | Consignees were notified by phone on 8/21/07 followed by a written notification by Fax or email. The letter described the product and problem and recommended to inspect the kits and to destroy all tracer vials of lot number 726023. They also requested the consignee to return the findings by fax or mail. |  | 117 | 117 kits | Worldwide, including USA, CANADA, BRAZIL, INDIA, QATAR, SWEDEN, URUGUAY, ARGENTINA, and JORDAN. | Terminated | October 17, 2008 | 423 | 117 | Not_Computer
Z-1229-2007 | 38621 | Yes | MPB | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Guidewire, 70cm x 1mm, REF 5583705, Sterile, Single Use, BARD Access Systems, Hemodialysis Catheter, Salt Lake City, UT 84116. |  |  | August 10, 2007 | 2007 | C.R. Bard Access Systems, Inc | Metal guidewires used in Hemodialysis Catheters may have been reverse loaded within the dispenser. Potential exists for the stiff end of the guidewire to be inserted into the patient instead of the flexible end. | Other | Consignees were notified by letter on 08/10/2007 and asked to remove and return all unused affected product to Bard. |  | 7225 | 7225 units | Worldwide, including USA, Canada, Latin America, and Europe (Austria, Belgium, Finland, France, Germany, Great Britain, Holland, Italy, Norway, Spain, and Sweden). | Terminated | April 25, 2008 | 259 | 7225 | Not_Computer
Z-1009-2007 | 37978 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Medical -Wheeled Electric Stretcher, Model 1550 (with and without scales) |  |  | May 07, 2007 | 2007 | Stryker Medical Div. of Stryker Corporation | The internal power cord may become damaged due to wear and, if this occurs, a shock hazard exists. | Other | Consignees were notified of the problem via letter dated 5/7/07 instructing them to immediately unplug the affected stretchers with scales, affix a warning label enclosed with the letter to the power cord of stretchers with scales, and that they will be contacted so that the firm can schedule a visit to upgrade all affected stretchers. A second notification letter was issued on August 7, 2007 to cover all 1550 stretchers with and without scales. |  | 1048 | 1,048 (1009 in the U.S. plus 39 international) | Worldwide: USA, Latin America, Mexico and Spain. | Terminated | June 11, 2008 | 401 | 1048 | Not_Computer
Z-1240-2007 | 38472 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | BCI Autocorr Plus Pulse Oximeter/ECG/Respiration Monitor, Catalog Numbers, 3404001 (including S) and 3404002 (including E,F, G, and S) Smiths Medical- a part of Smiths Group plc |  |  | July 18, 2007 | 2007 | Smiths Medical PM, Inc. | Phantom pacer spikes caused by a damaged resistor. | Other | Customers will be notified via a Customer Information Bulletin CIB Tracking Number: 07-CIB 12, July 18,2007 requesting them to evalute the functionality of their monitor (s). If the monitor does not exhibit any false pace detect makers when pace detect is enabled and respiration is disabled, no further action is required. If it does exhibit false pace detect they will be instructed to contact the Smiths Medical PM, Inc service department. At that time Technical Worksheet, 07-TW02, will be provided. The consignee will need to verify they are capable of perfoming the Technical Worksheet provided. The Smiths Medical PM. Inc. Service Manger will veirfy if the consignee is capable of performing the rework. |  | 178 | 178 | World wide: USA including states of KY, MS, PA, LA, and PR and countries of Poland, Mexico, Kuwait, Taiwan, Singapore, Hon Kong, India, Bolivia, Colombia, Chile, Argentia, Brunei Darssalam, Israel, Greece, Romania, Germany, The Netherlands, France, Great Britain, Spain, Brazil, Russia, China. | Terminated | October 31, 2008 | 471 | 178 | Not_Computer
Z-1241-2007 | 38814 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Duracon Total Knee Modular Femoral Component; |  |  | October 17, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Loose fixation pegs: One lot of femoral components may have loose fixation pegs. | Other | Notification letters, dated 10/17/2006, were sent via Federal Express on 10/17/2006 to all consignees requesting return of product. |  | 3 | 3 total, 2 in USA | Worldwide : One Hospital in Georgia (USA), Stryker Branch in Indiana and Stryker Branch in Sweden. | Terminated | July 07, 2009 | 994 | 3 | Not_Computer
Z-1239-2007 | 38406 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Radius Locking Caps, Catalog Number: 486610000, Manufactured by Stryker Spine SAS Z.I. Marticot - 33610 Cestas France.,Distributed by Howmedica Osteonics Corp. 325 Corporate Drive, Mahway, NJ 07430 |  |  | July 10, 2007 | 2007 | Stryker Spine | Product dissembled in the tray or prior to being loaded onto the inserter or tightening on the rod | Other | Product recall letters were sent to branches/agencies and hospitals on 7/10/2007 requesting product return. |  | 1423 | 1,300 in US, 123 in Italy | Worldwide- USA: Distributed to Stryker Branches/Agencies in MA, NY, NC. IN, MI, OK, TX,CO, AZ, WI, KS, CA, OR and hospitals in NC IN, MI, MO. KS, TZ, AZ, CA. OR and WA and Country of Italy | Terminated | December 11, 2007 | 154 | 1423 | Not_Computer
Z-1256-2007 | 37874 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Reverse Shoulder Prosthesis (RSP) device component; RSP Stem; Device manufactured by Encore Medical, L.P., Austin, TX 78758. |  |  | March 22, 2007 | 2007 | Encore Medical, Lp | Marketed without approval: Investigational devices were implanted during an IDE study prior to receiving approval for the devices in the study; patients were implanted with investigational devices outside of the IDE study; and devices implanted as "custom devices" by the manufacturer did not meet the definition of "custom devices" and were, therefore, unapproved devices. | Other | Letters were sent via Certified Mail on 03/22/07 to the surgeons and their associated IRBs. Each letter included a list of the patients to be notified and a form letter that the surgeons could use to notify the patients. Surgeons asked to notify their patients and to provide Encore with documented evidence that patience were notified. |  | 120 | 120 units. | Nationwide: Devices sold to and implanted by physicians in CA, CO, FL, GA, NY, SC and TX. | Terminated | March 27, 2008 | 371 | 678 | Not_Computer
Z-1124-2007 | 44899 | [6 empty] | Laser Distance and Level Transmitters, Model No. LT200, ASI, Housotn, TX 77024 |  |  | November 03, 2006 | 2006 | K Tek | Failure to comply with Federal Laser Product Performance Standards (21 CFR 1040.10, 1040.11, 1002.10-improper classification, incomplete labeling, incomplete user information and failure to submit required product reports) | Other | The firm notified its consignee of problem and the field correction by letter dated 11/03/2006. |  | 21 | 21 units | Nationwide | Terminated | December 10, 2008 | 768 | 1008 | Not_Computer
Z-1231-2007 | 38624 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon X3 |  |  | July 31, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Stabilizer posts, packaged with the Triathlon TS Plus Tibial inserts have a specification anomaly related to length, resulting in a potential for the post to slightly protrude above the top surface of the plastic post. | Other | Recall notification letters and product acknowledgment forms were send via Federal Express on August 2, 2007. |  | 14 | 14 | Product distributed to 5 hospitals in GA, OH, MN, TX. | Terminated | December 05, 2007 | 127 | 14 | Not_Computer
Z-1227-2007 | 38485 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Intersept 3608 Sterile tubing (8') 1/2 x 3/32 in. Quantity 10, Lot 7626621, Use by 2009-07-31. Manufacturer: Medtronic, Inc. Minneapolis, MN 55432-5604 USA. |  |  | August 01, 2007 | 2007 | Medtronic Perfusion Systems | Medtronic Custom Tubing Packs might have open header bag seals. | Other | Certified letters dated 8/1/07 were sent to customers. The letter explains the situation and requests the impacted product to be returned to Medtronic for replacement or credit. The customers are being requested to complete a Recall Certificate acknowledging the receipt of the information. |  | 853 | 853 packs | Nationwide. | Terminated | October 10, 2008 | 436 | 853 | Not_Computer
Z-1212-2007 | 39420 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | The Pinnacle3 Radiation Therapy Planning (RTP) system and its associated software; composed of several modules including the core Pinnacle3 functionality, Syntegra (TM), P3IMRT, P3MD and AcQSim3 (TM). The Pinnacle3 RTP system is composed of a Sun UNIX workstation (or UNIX-complaint computer) running the Solaris operating system and software, and a UNIX terminal emulation package on a personal computer, which allows the user to enter patient data into the system, use that data to construct a plan for radiation therapy, and evaluate the plan- Philips Medical Systems-Model #9200-0652A-ENG, 9200-0700A-ENG, 9200-0700B-ENG, 9200-0700C-ENG (with software versions 7.6C, 8.0d and 8.0h) |  |  | August 21, 2007 | 2007 | Philips Nuclear Medicine, Inc. | Incorrect Absolute Marking Coordinates when used with PQ-Series scanners due to a software malfunction | Other | All customers will receive an August 21, 2007, Customer Advisory Notification letter notifying them of the issue and the steps required to avoid the issue. Return receipts will be tracked. This is a short-term fix, a long term software fix will be issued shortly by the firm. |  | 1407 | 1,407 | Worldwide Distribution; USA including states of AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, and WI and Country of CANADA | Terminated | December 17, 2011 | 1579 | 1407 | Software | Display/Image | Software update | Software Update
Z-1243-2007 | 44739 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Cerner Millennium RadNet Auto Launch Study and Auto Launch Report software functionalities. |  |  | August 23, 2007 | 2007 | Cerner Corp | Defects in the Auto Launch functionality make it possible for a mismatch of patient data. | Software design | The recalling firm initiated recall via telephone on 8/23/07 to their customers who had the 2007 version in their production environment, explained the reason for recall, and informed them they should turn off the affected functionality. The firm posted details of the situation on their website in two different documents (referred to as a Priority Review Flash), one dated 8/29/07 and the other dated 8/31/07. The 8/31/07 document was also mailed to their customers via certified mail on 9/4/07.The corrected software (EP28822 ver.1) was made available to the firm's clients on 9/7/07 via their website. Beginning 9/18/07, workarounds were discussed via telephone with the clients who had downloaded the software correction which contained a new bug that affected the dictation file. A new software correction to fix the dication file problem was posted on the recalling firm's website on 9/28/07 for clients to download and validate. A Priority review Flash #PR07-0176-2 client notification dated 10/4/07 was also posted on their website and mailed via certified mail return receipt. |  | 9 | Downloaded by 9 customers nationwide. | Distribution was made to health care facilities located in LA, MO, NJ, MS, MI, NE, ND, and AR. | Terminated | November 29, 2007 | 98 | 9 | Software | Output/Calculation | Software update | Software Update
Z-1262-2007 | 38950 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Merit Medical Surgical Convenience Kits, Custom and SPPT, Merit Medical, South Jordan, UT 84065. Part numbers: K05-01767A, K05-01769, K05-01769A, K09-03843M, K09-03951M, K09-05894F, K09-08610, K10-00033Q, K10-00382M, K10-00685K, K10-00984F, K10-01531F, K10-01826D, K10-01826F, K10-02902B, K10-03080D, K10-03378C, K10-03397B, K10-03590A, K10-03590B, K10-03659D, K10-03659F, K10-03659G, K10-03689, K10-03907B, K12-01342D, K12-01444B, K12-01468, K12-01484A, K12-01487, K12-01490A, K12-01593, K12-01638, K12-01661, K12-01667, K12-01667A, K12-01964, K12-01968, K12-02098, SPPT-100/C, SPPT-5F-10C, SPPT-5F-10L/A, SPPT-5F-7/C, SPPT-5F-7L/A. |  |  | August 15, 2007 | 2007 | Merit Medical Systems, Inc | Expiration Dating; Some Convenience Kits contain medications and/or skin preparations that may expire before the expiration date on the outer package label. | Error in labeling | An August 22, 2007, Product Correction/Removal letter was issued to all distributors requesting return of affected lots and sub recall notification to their consignees. |  | 16333 | 16,333 units | Nationwide | Terminated | May 13, 2011 | 1367 | 16333 | Not_Computer
Z-1087-2007 | 38633 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Proteus XR/a General Purpose Radiographic System, Models 2259973-4 and 2380179 containing Proteus Operator Console model# 2259976. |  |  | May 31, 2006 | 2006 | General Electric Med Systems LLC | Labeling; failure to properly label per 21 CFR 1010.2 and 1010.3 -The Operator Console of the Proteus XR/a General Purpose Radiographic Systems were not properly labeled. | Other | Consignees were notified of the problem by a visit. To address the Proteus System shipped to the field, a Field Modification (FMI) was created to correct units in the field. |  | 22 | 22 units | Nationwide; including states of WA,KS, CO, NM, NC, FL, MT, AZ, TX, IL, NY, MD, and GA | Terminated | December 11, 2007 | 559 | 22 | Not_Computer
Z-0911-2007 | 44910 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Omni Diagnost Eleva |  |  | July 23, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | Units do not comply with the requirement to provide automatic beam limitation within +/- 3 degrees of horizontal, vertical, and perpendicular to the plane of the image receptor during radiographic procedures. | Other | On 7/23/07 Philips mailed a letter dated July 9, 2007 titled Electronic Product Radiation Warning. The letter advises customers of the issue and asks that customers apply the enclosed cautionary label to the device. The letter also explains that there will be a system modification to be installed. On 8/20/07 the firm sent the same letter, along with a letter dated 8/20/07 to customers thought to have either scrapped their units and the customer had previously indicated they were no longer using the unit. The 8/20/07 letter requested the customer contact Philps if they are currently using the device. |  | 14 | 14 units | Units were distributed to 1543 hospitals and medical centers throughout the US. No international distribution from this firm. | Terminated | September 17, 2012 | 1883 | 3244 | Not_Computer
Z-0764-2007 | 37844 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Practix 300 Mobile X-ray systems. The systems are used in general purpose diagnostic imaging, and are specifically used as mobile X-ray systems for bedside radiological examinations. |  |  | March 23, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | X-ray systems lack product report certification demonstrating compliance, and were illegally introduced into the United States commerce. | Other | The firm sent letters out via certified mail from 04/03/2007 through 04/09/2007 to consignees informing them of the issue. The letter indicates that the firm plans to replace consignee's systems with the Practix Convenio Mobile X-ray System at no charge by the end of May 2007. |  | 9 | 9 units | Nationwide: Distributed to Puerto Rico. | Terminated | September 16, 2010 | 1273 | 11 | Not_Computer
Z-1242-2007 | 38459 | No | JXK | Cranial electrotherapy stimulator. | Neurology | Neurology | 510(k) | HealthPax Cranial Electrotherapy Stimulator |  |  | July 19, 2007 | 2007 | Health Directions, Inc | Peak voltage test result not in acceptable limits. | Other | The recalling firm sent a recall letter 07/30/2007 to all their customers. The letter indicated that it is necessary to return the product for testing and the device will either be returned or if necessary replaced. Each unit will be tested with an oscilloscope device to determine whether it meets or fails design specifications. The recalling firm will reimburse shipping costs. |  | 212 | 212 units | Worldwide, including USA, Canada, Ireland, and India. | Terminated | June 03, 2008 | 320 | 212 | Hardware |  | Replace | Remove or Replace
Z-0014-2007 | 37567 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Brilliance(a) CT 6 slice, Model #4535 670 73191. |  |  | November 06, 2006 | 2006 | Philips Medical Systems (Cleveland) Inc | A potential anomaly was identified on the CT subsytem, which could result in the radiation dose being different than the calculated dose reported to the operator. | Other | The customers were sent Public Safety Notification November 3, 2006, indicating there was a problem with the software. However, since their unit was recently upgraded the problem did not impact their unit. |  | 20 | 20 | Worldwide, including USA, Australia, Canada, China, England, Finland, France, Germany, Ireland, Israel, Italy, Japan, Jordan, Korea, Netherlands, Poland, Puerto Rico, Spain, and Taiwan. | Terminated | October 05, 2010 | 1429 | 90 | Software | Output/Calculation | Software update | Software Update
Z-1261-2007 | 38815 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident PSL Acetabular Shell;Arc Deposited;Hydroxylapatite Coated;Cluster Screw Holes;Howmedica Osteonics Corp., |  |  | March 13, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Dimensional anomaly | Other | Recall Notification letters were sent on March 13, 2006 to all hospitals, branches/agencies via Federal Express. |  | 1220 | 1,220 units | Worldwide -USA, Romania, Chile and Brazil. | Terminated | January 18, 2008 | 676 | 1220 | Not_Computer
Z-1244-2007 | 38069 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Endoscopic Instrument Control System (da Vinci Surgical System) Model IS1200, A4.3 SW level, Manufactured by Intuitive Surgical, Inc, Sunnyvale, CA 94086 |  |  | May 03, 2007 | 2007 | Intuitive Surgical, Inc. | Under certain circumstances, the product's software may crash and require a manual override or restart before functioning again. In addition, there is an unrelated addendum to the User Manual to clarify an icon not described in the previous User Manual. | Software design | As of 9/7/2007, Intuitive Field Service Engineers have performed the upgrade via service visit on 292 of 293 U.S. systems listed on the affected sites list. A notification letter that the upgrade has been performed was left with each site after the upgrade has been completed. The firm's distributors received advance notice of the upgrades 9/13/2007. |  | 405 | 405 | Worldwide, including USA, Canada, Australia, Austria, Belgium, China, Czech Republic, Denmark, France, Germany, Greece, India, Italy, Japan, Korea, Malaysia, Netherlands, Norway, Romania, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Venezuela. | Terminated | July 02, 2008 | 426 | 405 | Software | Device Operation | Software update | Software Update
Z-0555-2007 | 39378 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Pyxis Anesthesia System 3500, Cardinal Health, Clinical Technologies and Services, San Diego, CA |  |  | August 13, 2007 | 2007 | Cardinal Health 303 Inc DBA Alaris Products | System lock up; mini-drawers -under certain circumstances during use, are not accessable, delaying the user from accessing critical medications. | Other | The field action consisted of phone calls and a mailing conducted to all customers for the Pyxis Anesthesia System 3500 beginning on 08/13/2007. The customers will be notified by registered return receipt mail by Recall Notifications dated August 14, 2007. Separate letters will be sent to the Director of Pharmacy, Chief of Anesthesia, Director of Risk Management, and Director of Surgery at each facility. Recommended actions will be provided as a short-term fix and a long-term fix of the software will be issued shortly. |  | 369 | 369 | Nationwide to Arizona, California, Colorado, Florida, Georgia, Hawaii, Iowa, Illinois, Indiana, Louisiana, Massachusetts, Missouri, Mississippi, Montana, New Jersey, Ohio, Oklahoma, Pennsylvania, Rhode Island, South Carolina, Tennessee, Texas, Virginia, Vermont, and Wyoming. | Terminated | March 28, 2012 | 1689 | 369 | Software | Device Operation | Software update | Software Update
Z-1263-2007 | 44702 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Altaire Shoulder Coil, an accessory to the Altaire MRI system. The shoulder coil is a two-channel, receive only coil that functions at 29.641 MHz band. It consists of a cable and a coil (imaging) body. A portion of the cable is referred to as the balun assembly or balun trap. The shoulder coil assembly is packaged in a cardboard box with foam inserts. |  |  | July 30, 2007 | 2007 | Hitachi Medical Systems America Inc | Overheating; A portion (balun assembly) of the cable that connects the shoulder coil to the actual MRI unit may malfunction and result in overheating of the balun assembly. The overheating of the balun assembly could result in injury to the patient or operator if either came into contact with the assembly. | Other | On 7/30/07, Hitachi sent a Device Correction letter via courier with receipt tracking to all affected customers outlining the risk of overheating of the cable balun assembly and subsequent risk of patient/operator injury. The letter notifies customers that (1) Hitachi will be contacting them in the future to schedule an on-site visit for replacement of the Altaire Shoulder Coil; (2) the replacement will begin in September, 2007 and continue until all affected coils are replaced; and (3) customers that have noticed overheating of the cable balun assembly are urged to notify Hitachi for immediate replacement service. |  | 273 | 273 | Nationwide: U.S. A. to include: AL, AR, AZ, CA, CO, CT, DE, FL, HI, ID, IL, IN, KS, KY, LA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, &amp; WY. | Terminated | October 05, 2010 | 1163 | 273 | Not_Computer
Z-1245-2007 | 38301 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | daVinci S Surgical system, patient side cart Auxiliary power board (APB) (printed circuit board assembly), endoscopic instrument control system. |  |  | May 14, 2007 | 2007 | Intuitive Surgical, Inc. | Product may malfunction and fail to start up on AC power. | Process control | As of 9/12/2007, the firm has conducted a field inspection of any involved units via service visit. No other notice is planned or indicated by the firm's submission. The firm feels that the risk associated with this issue is minimal, and therefore, any advanced notification is unneccessary. If any system failed one or more of the inspection points, its APB (3510404-04) was replaced with a verified APB (351040-04 or 351040-05). |  | 38 | 38 units | Worldwide, including USA, France, China, Japan, Mexico, Spain, Italy, Belgium and Singapore. | Terminated | December 03, 2007 | 203 | 38 | Other | Device Operation | Replace component | Remove or Replace
Z-1238-2007 | 38372 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Exacta-Mix 2400, OS v1.07, Model No. 8300-0073, Pharmacy Compounding System, Baxa Corporation, Englewood, CO 80112. |  |  | June 27, 2007 | 2007 | Baxa Corporation | A software defect could allow up to 50mL of extra volume being added to a TPN solution. | Other | All customers were notified via telephone on 06/27/2007 with follow-up of a faxed/emailed Urgent Product Recall letter dated 27 June 2007, detailing the software defect. End users were given a work around procedure to use until conversion back to version 1.06 could be scheduled. An additional copy of the recall letter was sent by FedEx. |  | 30 | 30 units | Class I Recall-Worldwide Distribution: USA including states of CA, CO, FL, KY, MI, NY, OH, OR, and TX and the UK | Terminated | December 07, 2007 | 163 | 30 | Software | Treatment/Delivery/Therapy | Conversion back to previous version | Software Update
Z-1267-2007 | 38634 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Davol Salut Fixation Reusable 38cm Shaft (Laparoscopic) |  |  | July 16, 2007 | 2007 | Davol, Inc., Sub. C. R. Bard, Inc. | Malformed constructs (straight shots) may form causing injury to user/patient | Device Design | Davol issued recall letter notification on 7/16/07 via Federal Express. Sales of the disposable cartridges ceased July 31, 2007. Cartridges will not be requested to be returned to Davol. Customer requests to return cartridges will be managed on an account-by-account basis. A reminder Second Letter issued July 23,2007 as a follow-up to the letter of July 16, 2007.The second stage will be a series of three letters to customers following the cessation of disposable cartridge sales and to recall the re-usable handles. A Stock Status form will be provided to customers to inventory the product which they are returning.The Stock Status form will be used to document the handles being returned. In the event that the customer cannot locate one or more handles. All handles will be returned to Cranston and will be dispositioned. Account verification for the number of handles returned from the account will be performed. No verification testing will be performed on returned handles unless a formal complaint has been registered for a handle. |  | 4242 | 4,242 shafts | Worldwide-USA and Canada, Austria, Switzerland, Spain, Sweden, United Kingdom, Italy,Portugal, Greece,Holland, South Africa, Israel, Japan, Belgium, Denmark, Finland, France, and Germany | Open, Classified |  |  | 195104 | Not_Computer
Z-0676-2007 | 37498 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9900 Elite image-intensified fluoroscopic x-ray system, GE Healthcare, Salt Lake City, UT. |  |  | March 01, 2007 | 2007 | GE OEC Medical Systems, Inc | During routine service, primary and/or secondary collimators may not have been properly reinstalled. | Other | Customers were notified by letter on 03/01/2007. They were told to continue using the machines and that an inspection would be scheduled immediately to verify that collimators had been correctly installed. |  | 19 | 19 units | Nationwide and Germany. | Terminated | May 01, 2012 | 1888 | 1975 | Not_Computer
Z-1088-2007 | 44936 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 2100-IQ Digital Fluoroscopic Imaging System, Versatile cardiac and vascular imaging system. GE Healthcare Americas P.O.Box 414, Milwaukee, Wisconsin 53201, U.S.A. |  |  | June 06, 2006 | 2006 | GE Healthcare | The systems do not contain a certification label on the C2 Cabinet (Beam-limiting control). | Other | Starting on June 6, 2007, a GE service engineer will go to each affected site to affix certification labels on the affected devices. |  | 161 | 161 | World wide Distribution includes: AL, AR, AZ, CA,, CO, CT FL, GA, HI, ID, IL, IN, KS, MA, MD, MN, MO, NC, NE, NJ, NY, OH, OR, PA, SC, TN, TX, WA, WI,and WV, and countries of ; Algeria, Australia, China, Denmark, France, Germany, Hungary, India, Indonesia, Italy, Japan, Korea, Kuwait, Latvia, New Zealand, Norway, Poland, Saudi Arabia, Singapore, Switzerland, Syrian Arab Republic, Turkey, and United Kingdom. | Terminated | December 08, 2015 | 3472 | 161 | Not_Computer
Z-0036-2008 | 44827 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Heraeus Kulzer GLUMA Comfort Bond + Desensitizer light curing one component adhesive, Assortment Value Pack 3 x 4 ml, dental bonding agent. Distributed by Heraeus Kulzer, Inc., Armonk, NY; Reorder No. 66018222. |  |  | July 09, 2007 | 2007 | Heraeus Kulzer Inc., Dental Products Division | The material may solidify in the bottle, and thus be rendered unusable. | Process design | Consignees were notified via phone or undated letter sent via fax beginning on or about 7/9/07 and instructed to contact the firm for return authorization/instructions. |  | 736 | 736 | Nationwide and Canada. | Terminated | March 12, 2008 | 247 | 3944 | Not_Computer
Z-0005-2008 | 38449 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PathFinder Bone Drill-Awl-Tap, a part of the PathFinder spinal stabilization system, Part Number 1156-60. |  |  | July 18, 2007 | 2007 | Abbott Spine | Tip of PathFinder surgical instrument broke during procedure. | Device Design | Firm notified sales representatives with an approved phone script on 7/18/2007. |  | 11 | 11 units | Nationwide. | Terminated | April 02, 2008 | 259 | 11 | Not_Computer
Z-0012-2008 | 39369 | No | GES | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Shelf Pack: |  |  | August 10, 2007 | 2007 | Becton Dickinson &amp; Company | Complaints received of rust on the blades of BD Bard-Parker Surgical blades. | Packaging change control | Urgent recall letters were sent out 8/10/2007 via UPS second day air. |  | 4026 | 4026 units | Nationwide. | Terminated | March 07, 2008 | 210 | 121452 | Not_Computer
Z-1265-2007 | 44735 | No | LZK | Chin prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Spectrum Designs Medical Anatomical Chin Implant; Catalog Number 250-134S; Lot # 31140; Serial Numbers BAHO45 through BAH144 |  |  | July 24, 2007 | 2007 | Spectrum Designs Inc | Mispackaging/Mislabeing: a portion of the lot may be labeled as containing nasal implants | Other | The firm initiated the recall on 07/24/2007 by contacting the customer (distributor) who filed the original complaint of mislabeling by email. The firm followed-up by sending recall letters (dated July 24, 2007) via FedEx or certified mail with signature confirmation to the one other direct account who received the recalled product on 08/03/2007. Customers were instructed to check their inventory for mislabeled products and call Spectrum at (800)239-6399 to report the number of this product being returned. They were given an address for product return. The firm indicated they will examine returned products, and scrap after examination. |  | 100 | 100 | Nationwide: California and China | Terminated | July 29, 2011 | 1466 | 230 | Not_Computer
Z-0007-2008 | 38494 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | 4 glass vials per box. 2/4ml vials of Sample Diluting Solution &amp; 2/2ml vials of Calibration Verification Material. BetaHCG Calibration Verification Test Set, in vitro diagnostic. |  |  | June 22, 2007 | 2007 | Tosoh Bioscience Inc | Firm discovered that a deterioration value of approximately 20% was found when the product of this lot was measured. Tosoh AIA confirmed that deterioration of the product occurred over time and resulted in a deviation from the assigned value. | Nonconforming Material/Component | Letter, dated July 2, 2006. Immediately examine inventory and quarantine product. If product was further distributed, the direct accounts are to notify their customers including a copy of the original recall letter and Notification of Recall response card. |  | 47 | 47 Boxes | Nationwide. | Terminated | October 07, 2010 | 1203 | 47 | Not_Computer
Z-0003-2008 | 38430 | No | MSQ | Campylobacter fetus serological reagents. | Microbiology | Microbiology | 510(k) | BreathTek Urea Hydrolysis Rate Calculation CD, Version 1.3.5 - January 1,'03, in vitro diagnostic. Meretek Diagnostics, Inc., Lafayette, CO. |  |  | May 02, 2007 | 2007 | Meretek Diagnostics, Inc. | Diagnostic test kit for Helicobacter pylori was distributed for adult and pediatric use, but was not approved for pediatric use. | Labeling False and Misleading | Customers were notified by letter on 05/02/2007 and asked to remove the software program from any computers on which it had been installed and return the software disc and User's Guide to Meretek. |  | 92 | 92 units | Nationwide. | Terminated | November 26, 2007 | 208 | 92 | Not_Computer
Z-0006-2008 | 38484 | No | BSF | Carbon dioxide absorber. | Anesthesiology | Anesthesiology | 510(K) Exempt | GE Disposable Multi Absorber Canister used with EZchange Module for the Aisys, Avance, Aespire Anesthesia systems. The EZchange Module is an optional accessory for these anesthesia systems. |  |  | July 24, 2007 | 2007 | Datex - Ohmeda, Inc | Elevated inspired CO2 levels related to the use of the Datex-Ohmeda GE Healthcare Disposable Multi Absorber Canister with the EZchange Module. | Vendor change control | Because this issue only occurs when the Disposable Multi Absorber is used in conjunction with EZchange Module, they have contacted all EZchange Module customers. It is critical to contact these customers specifically, rather than all customers shipped Disposable Multi Absorbers as this may cause unwarranted confusion with customers who do not have the EZchange Module. Consignees were sent an letter dated 7/19/2007 describing the product and problem. The letter lists a number of interim actions to implement with the anesthesia systems. |  | 389070 | 389,070 | Worldwide, including USA, Australia, Belgium, Brazil, Canada, Chile, China, Denmark, Egypt, Estonia, Finland, France, Germany, India, Ireland, Italy, Jordan, Republic of Korea, Kuwait, Mexico, Netherlands, New Zealand, Norway, Peru, Poland, Qatar, Singapore, Slovak Republic, South Africa, Spain, Sweden, Switzerland, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | October 16, 2015 | 3006 | 389070 | Not_Computer
Z-0008-2008 | 38894 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | CVI OZ Palm Injector, used to perfuse contrast media into vessels for angiographic procedures. Ref : OZ 1035 (Cardiovascular Innovations LLC) |  |  | July 18, 2007 | 2007 | Lemaitre Vascular, Inc. | Sterility of device may be compromised due to breach in sterile barrier. | Package design/selection | The firm issued a "Adivsory Notice" to accounts 7/18/07, advising them to examine inventory and remove product with defective packaging. The firm issued an updated notification dated 9/24/07 advising users to return recalled lots instead of conducting an examination of the product at the user site. |  | 7860 | 7,860 units | Worldwide, including USA, Canada, Netherlands, and Germany. | Terminated | September 08, 2008 | 418 | 7860 | Not_Computer
Z-0034-2008 | 39412 | No | MID | Multiple autoantibodies immunological test system. | Immunology | Immunology | 510(k) | APhLÂż HRP IgG &amp; IgM ELISA Kit; Ref. #LAPL-K-HRP-00GM; in vitro diagnostic. Distributed by Louisville APL Diagnostics, Inc., Doraville, GA 30360 (address now in Seabrook, TX 77856). |  |  | July 18, 2006 | 2006 | Louisville Apl Diagnostics Inc | Product mislabeled; outer container labeled with incorrect lot number. | Process design | Firm sent replacement labels to consignee via FedEx on 07/18/06 with instructions for consignee to change the labels on the kit boxes before shipping kits to end users. |  | 3 | 3 test kits (each kit containing 1 IgG and 1 IgM test kit) | Product was shipped to a distributor in Canada. | Terminated | October 24, 2007 | 463 | 3 | Not_Computer
Z-0015-2008 | 39389 | No | JIF | Ammonia test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Ammonia Liquid Stable Reagent, Enzymatic Method. Catalog number TR60101. in vitro diagnostic. |  |  | July 24, 2007 | 2007 | Thermo Fisher Scientific | Method is susceptible to positive interference from endogenous ALT in patient samples. | Device Design | The firm has updated the Limitations section of the Ammonia Reagent Package Insert and circulated a Technical Bulletin dated July 24, 2007 to all customers that have purchased Ammonia Reagent that is still within expiration date. A Notification of Change has been introduced and shall be placed within subsequent Ammonia Reagent Kits with a copy of the updated package insert; the notification of change informs customers that a change has been made to the package insert. Notifications of Change are placed within kits for 6 months from the date of issue. |  | 6882 | 6,882 units | Worldwide, including USA, Ecuador, and Canada. | Terminated | December 18, 2007 | 147 | 6882 | Not_Computer
Z-0057-2008 | 44933 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | ON-Q PainBuster with OnDemand (270 ml, 2 ml/hr + 5 ml bolus /60 min refill), Model PMB01, I-Flow Corporation, Lake Forest, CA 92630 U.S.A. |  |  | March 27, 2007 | 2007 | I-Flow Corporation | Misassembly: The ON-Q PainBuster PMB01 package may contain a Soaker catheter instead of a Non-Soaker catheter. As a result, the bolus button may remain latched in the down position which may cause the the device to deliver continuously at 7 ml/hr instead of 2 ml/hr. | Process control | An Important Recall Notification Letter was faxed to affected customers on March 27, 2007. Affected customers were notified of the recall, the reason for recall, and instructed to immediately quarantine any remaining inventory. A Customer Response (and Fax back) form was included with the Recall Notification. Affected customers were requested to return the FAX back form to I-Flow within 48 hours of receipt. Replacement product will be sent out as applicable. |  | 2587 | 2,587 units | Worldwide, USA and Australia | Terminated | October 07, 2009 | 925 | 2587 | Not_Computer
Z-0009-2008 | 38951 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | EPIKÂż KNEE SYSTEM Angled Insert Impactor LM/RL; Catalog #801-03-051 Rev. A; non-sterile; knee prosthesis instrument. Distributed by Encore Medical, L.P., Austin, TX 78758. |  |  | August 09, 2007 | 2007 | Encore Medical, Lp | Angled Insert Impactor may scratch Tibial Insert on the articulating surface, causing increased wear at the damaged site. | Device Design | Firm notified consignees regarding recall via letter on 08/09/07. Consignees informed that Encore customer service representatives would be contacting them to exchange recalled units with replacements. |  | 5 | 5 units. | Products distributed to direct accounts in CA, FL and MD. | Terminated | March 10, 2008 | 214 | 5 | Not_Computer
Z-0014-2008 | 39382 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Brain Heart Infusion Agar labeled BHIA + Vancomycin 6 MCG/ML For Screening for Vancomycin Resistant Enterococci PML microbiologicals Wilsonville, Oregon 97070 USA |  |  | April 23, 2007 | 2007 | Pml Microbiologicals Inc. | Plates manufactured without vancomycin. Possibility for false resistant results. | Other | On 4/23/07 the firm began contacting customers via fax and telephone. |  | 204800 | 640 total plates, 320 per lot (lot 229957-1 was dist only in Canada; lot 229958-1 was dist in the US and Canada) | lot #229957-1: 120 plates were distributed to 9 Canadian consignees; 200 plates were shipped, quarantined, and destroyed at the Canadian distribution center.lot #229958-1: 110 plates were distributed to 9 US consignees; 210 plates were shipped, quarantined, and destroyed at the Canadian distribution center. | Terminated | June 07, 2011 | 1506 | 204800 | Not_Computer
Z-0011-2008 | 39367 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | Brain Heart Infusion Agar (BHIA) + Vancomycin 6 MCG/ML |  |  | August 17, 2007 | 2007 | Pml Microbiologicals Inc. | Marketed a regulated product without a 510(k). | PMA | Firm notified consignees by phone and fax on August 17, 2007, instructing them to destroy the product. |  | 2245 | 1205 plates distributed in US; 1040 distributed in Canada | US, Puerto Rico, &amp; Canada | Terminated | June 07, 2011 | 1390 | 2245 | Not_Computer
Z-0039-2008 | 39419 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | STA Compact and STA Compact CT, Haemostasis Analyzer, in vitro diagnostic. |  |  | July 30, 2007 | 2007 | Diagnostica Stago, Inc. | Incorrect results. Internal investigation determined the presence of a defective measurement head which can impact the results of clotting tests by either yielding no result or rendering an error message. | Process design | Recall notification letters were mailed to consignees on 7/30/07 by first class mail. |  | 41 | 41 units installed | Nationwide. | Terminated | December 17, 2007 | 140 | 41 | I/O | Output/Calculation |  | Other
Z-0037-2008 | 38371 | No | EZW | Implanted electrical urinary continence device. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | Medtronic InterStim II INS Model 3058 with N'Vision InterStim-B Software supplied on the Model 8870 Version MMB_01/NNB_01 N'Vision with a Clinic an Programmer Model 8840. Implantable Neurostimulator for Urinary Control. |  |  | June 25, 2007 | 2007 | Medtronic Neuromodulation | A software issue may cause a lower than expected battery capacity estimate when evaluating an InterStim II Implantable Neurostimulator (model 3058) with an N'Vision Clinician Programmer(Model 8840) before implant . | Software design | Begining on July 11, 2007 a letter to the physicians was sent informing them of the issue and patient mangement of the two issues. The letter includes a Cover Letter, A Reply Form, and two Recall letters explaining the issue in detail. |  | 1885 | 1885 | Nationwide and the Netherlands. | Terminated | October 08, 2010 | 1201 | 1885 | Software | Output/Calculation | Notification | Safety Notice/Insructions
Z-0041-2008 | 44791 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | PEMView Software used with Naviscan PEM Flex Solo I and PEM Flex Solo II PET Scanner, for nuclear medicine imaging. |  |  | August 01, 2007 | 2007 | Naviscan PET Systems | Naviscan has identified a potential safety issue with all versions of PEM View Software commonly used for analyzing patient images acquired on the PEMFIex product line.1. The Positron Emission Mammography Standardized Uptake Value (PEMSUV) and distance measurements for the product mentioned above may be inaccurate if a zoomed image has been opened, measurement tools are used, and the applet rema | Software design | Telephone notification regarding issues #1 and #2 were carried out on August 1 and 2, 2007 and included instructions for a validated work-around to avoid the software issues. Written notifications with these same validated work-around instructions were sent on August 2, 2007 via certified mail. Customers were notified that Naviscan has identified a potential safety issue with all versions of PEM View Software.Issue #3 was identified after the notification to customers for issues #1 and #2. Telephone notification regarding issues #3 was carried out via telephone and certified mailing beginning August 16, 2007. Customers were informed that Naviscan is preparing software to resolve these issues and that they would be contacted by Naviscan Customer Support to ensure they thoroughly understand the issues and the recommended procedure to prevent these issues from occurring. If customers have any questions they were instructed to contact Naviscan Customer Support at 1-888-628-4159 or service@navsicanpet.com. |  | 19 | 19 units | Nationwide. | Terminated | May 14, 2008 | 287 | 19 | Software | Output/Calculation | Software update | Software Update
Z-0038-2008 | 38511 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pump waste tube for blood gas analysis kit. One yellowish/white rubber tube per plastic bag. Used on ABL800 Flex Series blood gas analyzers. |  |  | June 22, 2007 | 2007 | Radiometer America Inc | Tube may break. The firm discovered that the yellowish/white tubes used in their blood gas analyzers (ABL800 Flex Series) were manufactured incorrectly. These tubes have the potential to break off at the black plastic connectors. The breakage may occur during the installation of the tubes or after a short time in use. | Process design | By telephone call or service representative visit beginning July 18, 2007, Consumers/End Users are asked to locate the analyzers and yearly service kits concerned and replace all pump tubes with new ones. |  | 289 | 289 pieces | Product was distributed throughout the United States and Canada. | Terminated | April 25, 2012 | 1769 | 289 | Not_Computer
Z-0055-2008 | 39383 | No | LGW |  |  | Neurology | PMA | Precision Implantable Pulse Generator (IPG), Model Number SC-110 |  |  | August 03, 2007 | 2007 | Advanced Bionics Corp | Mislabeled: Product labeled with the incorrect shelf-life information. | Error in labeling | There was one unit affected by this field action. An Advanced Bionics representative contacted and visited the facility on August 3, 2007. At this time, the affected device has been retrieved by the Company. |  | 1 | 1 unit affected | MD | Terminated | January 25, 2012 | 1636 | 1 | Not_Computer
Z-0040-2008 | 44740 | No | EZW | Implanted electrical urinary continence device. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | Medtronic InterStim INS (Model 3023) with N'Vision Inter Stim-B software supplied on the Model 8870 Version MMB_01/NNB_01 with N'Vision Clinician Programmer Model 8840 or InterStim iCon Patient Programmer Model 3037. Implantable Neurostimulator for Urinary Control. |  |  | June 25, 2007 | 2007 | Medtronic Neuromodulation | Software issue may cause false low battery alert to be displayed when interrogating an Interstim Implantable Neurostimulator (Model 3023) with an N'Vision Clinician Programmer (Model 8840), or with an Inter Stim iCon Patient Programmer (Model 3037). | Software design | Begining on July 11, 2007 a letter to the physicians was sent informing them of the issue and patient management for the two issues. The letter includes a Cover Letter, A Reply Form, and two Recall letters explaining the issue in detail. |  | 1885 | 1885 | Nationwide and the Netherlands. | Terminated | October 08, 2010 | 1201 | 1885 | Software | Alarm/Message | Notification | Safety Notice/Insructions
Z-0056-2008 | 39408 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Angiomat Illumena Injector with Auxiliary Console, Mallinckrodt, Inc., Hazelwood, MO. |  |  | August 24, 2007 | 2007 | Mallinckrodt Inc | Synchronization problems: units that have both a main programmable console and an auxiliary programmable console, when simultaneously, injection parameters are programmed into the unit's main console and the unit's auxiliary consoles, mis-synchronization could occur resulting in a variance between the intended injection parameters (those programmed) and the actual parameters of the injection deliv | Device Design | On 9/20/2007, a notification was sent from Mallinckrodt to all direct accounts via certified mail. The notification details the potential for mis-synchronization of injection parameters if the main and auxiliary consoles are simultaneously programmed. The notification reminds end-users that simultaneous programming should not be performed and to ensure it is not performed, a notification label (provided in the notification) is to be affixed on both consoles of the injector system. Once the label is affixed to the system, end-users are instructed to complete the notification survey and return the survey to the Mallinckrodt. In addition, the notification includes an updated chapter insert (Chapter 11.1) for the injector system's product manual that emphasize the separation between programming of the main and auxiliary consoles. Questions regarding this matter may be directed to Mallinckrodt's Product Monitoring Department at 1-888-744-1414, menu option 1 and then menu option 2. |  | 176 | 176 | Worldwide to include:U.S.A. states of - AL, AZ, CA, CO, CT, FL, GA, HI, IA, IL, KS, KY, LA, MA, MD, MI, MN, MO, NE, NJ, NY, OH, OK, PA, TN, TX, UT, VA, WA, &amp; WI. and Foreign countries of Argentina, Belgium, Brazil, Canada, Germany, Japan, Singapore, &amp; Zambia. | Terminated | August 17, 2010 | 1089 | 176 | Other | Device Operation | Update manual | Other
Z-0002-2008 | 38369 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | FP1000 Cell Preparation System Software versions 1.1 &amp; 1.2, Part Number 624922, for in vitro diagnostic use. |  |  | February 22, 2007 | 2007 | Beckman Coulter Inc | Pressure (from recapped tubes) or vacuum (from under-filled short-draw tubes) will cause inaccurate results if not properly vented. | Labeling design | A Product Corrective Action (PCA) letter was mailed on Feb 22, 2007, to all FP1000 customers to inform them of an issue associated with the FP1000 Cell Preparation System. Pressure (from recapped tubes) or vacuum (from under-filled short-draw tubes) will cause inaccurate results if not properly vented. Under these conditions the FP1000 will under-dispense specimen volume leading to lower than expected absolute counts. Specimen tubes that are recapped or underfilled must therefore be vented prior to running on the FP1000. If the initial state of the specimen tube is unknown, then it must be vented to ensure accurate results prior to running on the FP1000. Additional instructions are provided. A Fax Back Response Form was also provided. |  | 10 | 9 in the U.S.; 1 in Canada | Nationwide (MA, IN, IL, UT, RI, NJ), and Canada. | Terminated | April 26, 2012 | 1890 | 10 | Not_Computer
Z-0061-2008 | 44942 | Yes | MVK |  |  | Cardiovascular | PMA | Battery Pack portion of the LifeVest device. The LifeVest is a wearable cardioverter defibrillator worn by a patient at risk for sudden cardiac arrest (SCA). |  |  | September 14, 2007 | 2007 | Zoll Lifecor Corporation | Possible malfunction of the battery pack. A significant electrostatic discharge (ESD) can cause the battery pack to malfunction. The malfunction causes the battery pack to discharge more rapidly and prevents the pack from recharging. | Device Design | Patients currently using the LifeVest should continue to do so while the product removal is underway. Advisory notices will be sent via first class mail 9/14/2007 to all active patients using the LifeVest device. Product replacements began on 8/31/07 by shipping out revised battery packs for new patient use. Shipments of new battery packs will continue based on production capacity and material availability. Any patients expected to be using the device longer than 5 months will be provided with 2 replacement Battery Packs. They will be instructed to return their existing packs in the same shipping box. |  | 2006 | 2006 packs | The product was shipped nationwide to patients. The product was also shipped to distributors in Germany, United Kingdom, Italy, Spain, France, and Denmark. | Terminated | June 13, 2008 | 273 | 2006 | Battery |  | Replace | Remove or Replace
Z-0043-2008 | 44909 | No | DOF | Lead test system. | Toxicology | Toxicology | 510(k) | ESA LeadCare II Blood Test Kit, 6 test Kit containing Level 1 and Level 2 Control Vials with lot number: 0704A |  |  | September 20, 2007 | 2007 | Esa Biosciences, Inc. | Level 1 and Level 2 Control values were incorrectly assigned. | Process design | ESA notifed Distributors and End Users by telephone and letter on 9/20/07. Distributors and End Users will be informed of the Level 1 and 2 Control value assignment change. End Users will be sent a new Level 1 and 2 Control Viallabel for each vial they originally received. Distributors will be asked to return all test kits in their inventory to ESA Biosciences. |  | 2 | 2 kits | Nationwide and United Kingdom. | Terminated | March 25, 2008 | 187 | 367 | Not_Computer
Z-0046-2008 | 44940 | [6 empty] | Microtainer Safety Flow Lancets; Blue |  |  | September 26, 2007 | 2007 | Becton Dickinson &amp; Company | Potential for the blade and holder assembly to become loose and then separate from the main body of the lancet during the process of disposal. | Process design | Recall communication issued on or about 9/28/2007. |  | 4349000 | 4,349,000 units | Product distributed worldwide to Australia, Brazil, Canada, China, Colombia, Guatemala, Hong Kong, India, Indonesia, Mexico, Malaysia, New Zealand, Pakistan, Philippines, Singapore, Thailand, and United States. | Terminated | January 15, 2008 | 111 | 7803000 | Not_Computer
Z-0066-2008 | 38739 | No | HMY | Keratome. | Ophthalmic | Ophthalmic | 510(k) | BD Microkeratome K3000 and K4000;Catalog Number: 378660, Becton Dickinson &amp; Company (BD) |  |  | August 09, 2007 | 2007 | Becton Dickinson and Company | Microkeratome blade may stop oscillating and transverse the eye causing injury to the cornea | Component design/selection | BD Medical notified customers by Advisory Letter dated 8/09/07, advising users to return heads and bearing cartridges for replacement of the o-ring. |  | 256 | 256 units | Worldwide | Terminated | February 29, 2012 | 1665 | 256 | Not_Computer
Z-0058-2008 | 44938 | No | EXF | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Remington Medical, Inc, (RMI) -- Urinary Drainage Bags: 500ml Disposable Urinary Leg Bag (Order number: 500-D) and 1000ml Disposible Urinary Leg Bag (Order number 1000-D) -- Sterile Fluid Path, For Single Use Only, Remington Medical, Inc., Alpharetta, GA 30005. |  |  | July 23, 2007 | 2007 | Remington Medical Inc. | Sterility, inadequate -- Pathway may not be sterile as labeled. | Process control | Consignees were notified by Urgent Recall Letter (dated 07/23/07) via fax from 07/24/2007 through 07/27/2007. They were instructed to quarantine all affected product in inventory. If product was distributed they were instructed to identify their customers and notify them of the recall via copy of the notification letter. |  | 45000 | 45,000 units | Nationwide Distribution -- including states of AL, AR, AZ, CA, CO, CT, GA, HI, IL, IN, KS, KY, MD, MI, MT, NB, NC, OH, NY, PA, RI, TN, WA, WV, and WY. | Terminated | April 06, 2009 | 623 | 45000 | Not_Computer
Z-0101-2008 | 44810 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Steris QFC1728 and QFC1729 Quick Connects;(used to connect gastrointestinal endoscopes to processing trays) The malfunctioning piece of the quick connect component is adaptor #6, STERIS part number 201158. Adaptor #6 attaches to the suction port of the endoscope. |  |  | September 12, 2007 | 2007 | Steris Corporation Hopkins Facility | Loose or missing adaptor: The center insert of the adaptor #6 restrictor could become loose and/or come out of the adaptor's housing. The absence of adaptor #6's center insert may affect the sterilization process of the attached endoscope. | Component design/selection | STERIS has contacted affected direct accounts by phone to notify direct accounts of the recall. On 9/28/2007, STERIS will mail a recall notification to all affected direct accounts via certified mail. The notification informs direct accounts that STERIS representatives will be visiting direct accounts to retrieve affected products and provide replacements. Where direct accounts further distribute, STERIS representatives will inform the direct accounts to contact their sub-accounts and facilitate exchange of affected product for replacement product. Replacement product will be available beginning November 1, 2007. |  | 391 | 391 Quick Connects | Worldwide: United States, Australia, Bahrain, Canada, Germany, Great Britain, Hungary, Italy, Qatar, and Spain | Terminated | September 28, 2010 | 1112 | 391 | Not_Computer
Z-0033-2008 | 39402 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Smith + Nephew, Birmingham Hip Resurfacing, Acetabular Cup, 62 mm cup Ref.: 74120262. Hip implant component. |  |  | July 12, 2007 | 2007 | Smith &amp; Nephew Inc | The carton label may not reflect the correct size of the acetabular cup. | Packaging process control | The consignees were notified of the problem and the recall by initially by telephone on 07/12/2007 and followed by a letter dated 07/24/2007. |  | 1 | 1 unit | Worldwide, including USA, Australia, Belgium, Canada, Finland, France, Germany, Italy, The Netherlands, Norway, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | February 02, 2010 | 936 | 177 | Not_Computer
Z-0072-2008 | 44780 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident PSL HA Cluster Acetabular Shell; |  |  | August 30, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics had identified that specific lots of Trident PSL Acetabular shells may have a dimensional discrepancy. The deviation regarding the difference in wall thickness will increase the gap between the shell and liner on one side and will decrease the gap between shell and liner on the opposing side, resulting in interference. | Process control | Recall notification letters were sent out on August 30, 2007 with return receipt. |  | 142 | 142 units | Product was distributed nationwide and one unit to Brazil. | Terminated | September 05, 2008 | 372 | 142 | Not_Computer
Z-0062-2008 | 44751 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart XL Defibrillator/Monitor, Model: M4735A |  |  | September 04, 2007 | 2007 | Philips Medical Systems | Inadequate shipping package causing damage to control boards resulting in unit failure. Failure to deliver therapy (defibrillation, pacing or cardioversion) may occur as a result of undetected damage to the PCB assembly. | Package design/selection | Philips Medical notified accounts (Distributors/Healthcare Customers) by letter dated 9/4/07. Users are requested if possible to remove the unit from service. Firm will have Field representative conduct correction. |  | 147 | 147 boards | Puerto Rico, Canada, Argentina, Australia, Brazil, China, Domincan Republic, Czech Republic, Hong Kong, India, Japan, Korea, Macau, Singapore, Spain, South Africa, Slovenia, Sweden, Switzerland, Ireland, UK, France, Kuwait, Poland, and Germany | Open, Classified |  |  | 147 | Not Computer
Z-0067-2008 | 45403 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic Sprint Fidelis 6930 Steroid eluting, tripolar, ventricular lead with tined tip and right ventricular (RV) defibrillation coil electrode. Sterile. This lead is designed for pacing, sensing, cardioversion and defibrillation therapies. Medtronic World Headquarters, Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. |  |  | October 15, 2007 | 2007 | Medtronic Inc. Cardiac Rhythm Managment | Medtronic Cardiac Rhythm Disease Management has voluntarily suspended worldwide distribution of the Sprint Fidelis family of defibrillation leads because of the potential for lead fractures. In addition, the company recommends against new implants of the leads (Sprint Fidelis Models: 6930, 6931, 6948, 6949). Medtronic, its Independent Physician Quality Panel, and Bruce Lindsay, M.D., Professor of | Device Design | A press release was issued by Medtronic on 10/15/2007. An FDA Statement was also issued on 10/15/2007. An Urgent Medical Device Information letter, dated 10/15/2007, was sent to physicians. The letter states that Medtronic has suspended distribution of the leads, the leads should no longer be implanted and unused leads should be returned to Medtronic. This letter describes background information, performance and recommendations. Medtronic believes it is inappropriate to prophylactically remove Sprint Fidelis leads except in unusual individual patient circumstances. An Important Patient Information letter was addressed to Medtronic Heart Device Patients. Both letters can be found on Medtronic's website at www.medtronic.com/fidelis. Return of product will be achieved with the assistance of Medtronic Representatives. |  | 235000 | 235,000 for all models | Worldwide, including USA, Puerto Rico, Guam, and Northern Mariana Islands. | Terminated | June 21, 2011 | 1345 | 235000 | Not_Computer
Z-0102-2008 | 44812 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Orthopedics Apical Hole Plug, 3/8 - 24 threaded, TI-6AL-4V alloy, sterile; REF 123741, Biomet Orthopedics, Inc., Warsaw, IN 46581 |  |  | August 09, 2007 | 2007 | Biomet, Inc. | Incorrect or lack of threading: The plug component was manufactured incorrectly and does not fit with the screw hole as intended. | Process control | The distributor was notified via phone on or about 8/9/07 and by letter dated 8/31/07 and requested to return the product. Returns included subaccounts |  | 5 | 5 | Nationwide | Terminated | June 26, 2009 | 687 | 5 | Not_Computer
Z-0104-2008 | 44901 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific mach 1 Guide Catheter, VL3.5 SH Voda Left 3.5, REF 34358-78, UPN H74934358780, Sterile EO, Made in Mexico Avenida Norske Edicficio G1, Local B y H Parque Industrial La Mesa, Fracc. Rubio La Mesa, Tijuana, B.C. Mexico C.P. 22550 |  |  | September 11, 2007 | 2007 | Boston Scientific | Missing holes: Catheters are missing the catheter's side holes. | Process control | A voice mail was left to affected members of the domestic sales force on September 11, 2007 and an interoffice memo with the recall implementation instructions was also sent to the Sales Force. An Urgent Voluntary Medical Device Recall letter dated September 11, 2007, was sent to affected consignees describing the problem and products affected. The letter included the customer steps for recall, a Reply Verification Tracking Form and a return shipping label. |  | 7 | 7 | Worldwide: NY, PA, and Japan | Terminated | February 24, 2008 | 166 | 13 | Not_Computer
Z-0117-2008 | 44736 | No | FFF | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Boston Scientific Radial Jaw Large Capacity Biopsy Forceps with Needle, Sterile, 240 mm. Box of 20 |  |  | August 28, 2007 | 2007 | Boston Scientific Corporation | Device lacked manufacturing finishing process resulting in endoscope damage. | Process control | Boston Scientific notified accounts by letter on 8/28/07 by Federal Express Priority mail. Accounts are requested to remove product from inventory and return product. |  | 640 | 640 boxes | Worldwide, including USA, Sweden, and South Africa. | Terminated | July 10, 2008 | 317 | 1519 | Not_Computer
Z-0110-2008 | 38264 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT Systems, 40 slice configuration. Computed tomography scanner. |  |  | May 25, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | Two potential artifacts may appear in some images from head scans on the Brilliance 40/64 systems. These artifacts may lead to misdiagnosis. | Software design | Firm notified its customers via a Product Safety Notification letter dated May 25, 2007, warning the user of a possible appearance of artifacts which could lead to a misdiagnosis and improper treatment. |  | 404 | 404 | Worldwide, including USA, Canada, Argentina, Austria, Australia, Bangladesh, Belgium, Brazil, Chile, China, Colombia, Denmark, Ecuador, Egypt, England, Equatorial Guinea, Finland, France, Germany, Greece, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Korea, Lebanon, Libya, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Panama, Phillipines, Poland, Portugal, Puerto Rico, Russia, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Switzerland, Sweden, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, Uzbekistan, Venezuela, and West Indies. | Terminated | September 24, 2012 | 1949 | 1037 | Not_Computer
Z-0063-2008 | 44724 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Medtronic Export XT Aspiration Catheter, 6F (US) |  |  | August 29, 2007 | 2007 | Medtronic, Inc. | Additional Warning in Instructions for Use for guide wire management. Complaints of wire lumen tearing during use or band marker becoming dislodged. | Labeling design | Medtronic has notifed sales representatives to visit Cath lab site with the letter dated August 29, 2007 and educate the user, provide the Safety Alert Letter and document delivery of the letter with the confirmation form. |  | 13940 | 13,940 units | Worldwide, including USA, Canada, India, Austria, Belgium, France, Germany, Italy, Netherlands, Spain, Sweden, Switzerland, and UK. | Terminated | September 08, 2008 | 376 | 13940 | Not_Computer
Z-0065-2008 | 44881 | No | LOM |  |  | Microbiology | PMA | Bio-Rad Laboratories-MONOLISA Anti-HBc IGM EIA kit, an in-vitro diagnostic test kit, Catalog # 26174, manufactured by Bio-Rad Laboratories, Redmond, WA. |  |  | September 05, 2007 | 2007 | Bio-Rad Laboratories | Incorrect Instruction: the package insert states to add 100L of conjugate; it should state to add 100uL. | Error in labeling | On 9/5/07 a letter titled "IMPORTANT PRODUCT INFORMATION" was sent to customers advising them of the issue and requesting customers make a note of the correction in their laboratory procedures and/or any existing copies of the package insert. On 9/19/07 the firm sent a second letter (dated 9/18/07) via DHL titled "URGENT PRODUCT CORRECTION" and supplied a corrected package insert to replace the original insert sent. |  | 46 | 46 | Nationwide: Kits were distributed to 27 hospitals, clinics and testing labs throughout the US. | Terminated | November 15, 2007 | 71 | 46 | Not_Computer
Z-0076-2008 | 44825 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet --DCM AGC Molded Tibial Implant (12 mm x 80 mm)--Biomet Orthopedics , Interlok finish/posterior stabilized, 12 mm thick x 80 mm wide, UHMWPE/CO-CR-MO alloy, sterile; REF 155364, Biomet Orthopedic, Inc. |  |  | July 24, 2007 | 2007 | Biomet, Inc. | Incorrect Product -- Mislabeled as to size; a 18 mm x 65 mm tibula implant is enclosed in a package labeled as 12 mm x 80 mm. | Labeling mix-ups | The firm notified distributor by phone on or about 7/24/07 and followed up with an Urgent Medical Device Recall Notice/Letter dated 8/31/07. The letter requested that the product be discontinued and returned to Biomet. |  | 4 | 4 | Class II Recall - Worldwide Distribution --- USA including states of IA, SD, and countries of Brazil and Chile. | Terminated | July 24, 2008 | 366 | 5 | Not_Computer
Z-0078-2008 | 44915 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Spectra Optia apheresis system, a blood component separator. Part No. 61000, Gambro BCT Lakewood, CO USA 80215. |  |  | September 18, 2007 | 2007 | Gambro BCT, Inc. | Under specific conditions, apheresis machine could return up to 60mL of air to the patient. | Software design | All consignees were notified by telephone beginning 09/17/2007. They were given instructions on how to safely use the machine. Letters were also sent on 09/22/2007 restating the information. |  | 15 | 15 units | US, Canada, Australia, Europe. | Terminated | January 25, 2008 | 129 | 15 | Not_Computer
Z-0108-2008 | 44929 | No | LQC | Mechanical lithotriptor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Olympus Single Use Mechanical Lithotriptor, Model number BML-201Q, Olympus America, Inc. |  |  | September 20, 2007 | 2007 | Olympus America Inc. | Slider knob does not move smoothly causing the user difficulty in advancing the coil sheath | Process control | The recalling firm issued a recall letter dated 9/20/07 to the medical facilities informing them of the problem and the need to return the product. The letter also informed the hospitals that replacement products would be sent. |  | 37 | 37 units | Nationwide | Terminated | February 15, 2008 | 148 | 482 | Not_Computer
Z-0109-2008 | 44939 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific-Renegade Hi-Flo Microcatheter Kit, |  |  | September 10, 2007 | 2007 | Boston Scientific Corporation | Incorrect Guidewire: Product may incorrectly contain a 135 cm length Transend-18 guidewire instead of the kit labeled 165 cm length Transend-18 guidewire. | Packaging process control | Boston Scientific issued a Urgent Medical Device Recall Letter dated September 10, 2007 to Domestic accounts via Federal Express. The notification instructed the account to immediately discontinue the useof the product under this recall, check their inventory, remove any affected product , and return it to Boston Scientific. The notification also requested the completion of a Reply Verification Tracking Form |  | 153 | 153 | Nationwide | Terminated | December 20, 2007 | 101 | 153 | Not_Computer
Z-0080-2008 | 38775 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ReUnion Fracture Stem Humeral Component - PureFix HA; Shoulder prosthesis. |  |  | May 09, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The plasma spray may delaminate from the proximal portion of the humeral stem. | Process control | Market Withdrawal letters were sent on May 9, 2007 via Federal Express. Letters to implanting physicians will be mail on or about mailed November 1, 2007. Stryker is also reissuing original letters indicating that this action is a Recall and advising of possible health hazards. |  | 232 | 232 units | One hospital in Akron OH and 12 Stryker branches/agencies | Terminated | March 07, 2008 | 303 | 232 | Not_Computer
Z-0001-2008 | 38473 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Bio-logic System Corp --Telescoping Camera Pole Assembly (part number 585-CYBPSM and 580-VIDPOL) and assemblies (580IPCART and 580 PNCART; Bio-Logic Systems Corporation, Mundelein, IL 60060; The Model 580-VIDPOL is used with the 580-PNCART that is wired for use with a Panasonic Camera. The pole is also a subassembly of the 585-CYBPSM with wires used in the 580-IPCART wired for the use with an Ipela Camera. Both carts are a component of the Ceegraph/Sleepscan Netlink with Video. |  |  | July 26, 2007 | 2007 | Bio-Logic Systems Corp | Bio-logic System Corp Camera Pole may weaken and detach at the base of the telescoping camera pole during transport causing pole to break at the welded base and fall. | Device Design | Bio-logic sent a Safety Alert letters dated 7/26/07 to the affected customers, warning them that using the pole to transport the cart may cause the weld at the base of the pole to weaken, causing it to break. As a precaution to ensure customer safety, the accounts were instructed to place the enclosed red sticker on the pole, below the tightening knob. An Instruction sheet is enclosed . The warning label reads: "Caution: Do not use as a handle to move cart. Always lower camera when transporting. Avoid sudden stops or collisions." The accounts were requested to complete and return an enclosed reply sheet acknowledging the receipt of the letter and the placement of the warning labels on the poles. Any questions were directed to Customer Support at 1-800-332-3114. |  | 311 | 311 units | Worldwide Distribution -- USA and China, Egypt, Hong Kong, Israel, Malaysia and Saudi Arabia | Terminated | August 12, 2009 | 748 | 311 | Not_Computer
Z-0064-2008 | 39416 | No | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | Abbott Diagnostic International, Ltd.-- AxSYM Digoxin III Reagent Pack, List (6L07-20), For In Vitro Diagnostic use. This reagent system is composed of one bottle of Digoxin-Alkaline Phosphatase Conjugate, one bottle of Anti-Digoxin (Rabbit) Coated microparticles, one bottle of MEIA buffer, and one bottle of Digoxin Probe wash solution. |  |  | July 20, 2007 | 2007 | Abbott Diagnostic International, Ltd. | Incorrect Results -- Error codes 1062, 1063, 1113, and 1118 observed when running patient samples on the AxSYM Digoxin III assay, rather than a concentration, is reported for the patient sample. | Labeling design | Abbott Diagnostics International Limited, Ltd., mailed a Product Correction Letter to customers on 7/20/02 advising an additional option(s) has been identified to resolve Error Codes: - Perform one of the two options to resolve these error codes 1062, 1063, 1113, and 1118 observed when running patient samples on the AxSYM Digoxin III assay. - Dilute the sample one to one (1 part sample plus 1 part Calibrator A) and repeat analysis. - The final concentration is equal to the reported concentration of the sample multiplied by the dilution factor - The minimum final concentration that may be reported out is 0.6 ng/mL, which corresponds to the sensitivity of the AxSYM Digoxin III assay (0.3 ng/mL) multiplied by the dilution factor (2). OR- Centrifuge the sample at 8,000 to 10,000 RCF for 10 minutes and repeat analysis. |  | 3104 | 3,104 | Worldwide Distribution -- USA including States of AL, AR, AZ, CA, CT, FL, GA, IA, ID, IN, IL, KS, KY, LA, MA, ME, MD, MI, MN, MS, MO, MT, NV, NJ, NY, NC, OH, OK, OR, PA, SD, SC, TN, TX, UT, VA, VT, WA, WI, WY, HI, Puerto Rico, and countries of Canada, Venezuela, Uruguay, Germany, Hong Kong, South Korea, Australia, and New Zealand | Terminated | September 09, 2011 | 1512 | 3104 | Other |  |  | Other
Z-0133-2008 | 39354 | [6 empty] | Harleco |  |  | July 30, 2007 | 2007 | EMD Chemicals Inc | Out of Spec -- The Batch (Lot) was incorrectly released after being rejected for out of specification of high azure and eosin absorbance at the end of the run and particulate matter at the end of the run. | Employee error | EMD Chemicals via fax and mailed an Important Product Recall letter dated 7/30/07. The letter notified end use consignee that after conducting an investigation it has determined that Harleco May and Grunwald's Stain; 0.17% w/v Item #660/75 % 660/85 has unacceptable levels of azure and eosin absorbance and undissolved dye matter. Users were requested to discontinue shipping and return product. Wholesalers were advised this is a Customer Level Product Recall and requested to sub recall product sold after 7/5/07. |  | 26 | 26 units (1 liter), 1 unit (4 liter) | Product was released to 2 distributors; one in the US, one in Canada | Terminated | December 26, 2007 | 149 | 26 | Not_Computer
Z-0081-2008 | 39422 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Paceart System, 2006 Edition, Medtronic, Inc. 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. The Paceart System can act as a Practice Management software application and is intended for use by facilities that provide services such as event monitoring, trans-telephonic pacemaker evaluation, and in-office evaluation of implanted cardiac devices. |  |  | August 27, 2007 | 2007 | Medtronic Inc. Cardiac Rhythm Managment | An issue with the Paceart System - 2006 First Edition (Get Connected Edition) exists. Under certain circumstances, Paceart System Generic ICD reports may not accurately reflect patient VT/VF, SVT/NST and Mode Switch/AT/AF episode detection data. When episode detection data is imported into the Paceart System from certain sources, the Paceart System Generic ICD reports display a zero, suggesting | Software design | A Medtronic letter was sent to consignees on 8/27/07. The letter described the issue and the product involved. It also notified customers that Medtronic is developing a software update that will correct the issue. An Acknowledgement form was attached to be filled by consignees as confirmation and understanding of the information received. |  | 721 | 721 | Nationwide. | Terminated | November 03, 2008 | 434 | 721 | Software | Display/Image | Software update | Software Update
Z-0129-2008 | 45476 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10, Automated External Defibrillator, a portable 12V internal battery powered defibrillator; Model AED10; Part Numbers: 970302E,970308E,970309E, 970310E and 970311E; |  |  | October 26, 2007 | 2007 | Welch Allyn Protocol, Inc | Failure to Deliver Shock; a defective capacitor may cause the delay or non-delivery of the defibrillating shock which may result in failure to resuscitate the patient | Nonconforming Material/Component | The firm issued a press release on 10/30/07, and has set up a call center for customers, and plans to send three (3) certified mailings to notify affected customers, alert them of the issue, and provide instructions on how to return their device. The first letter dated 10/26/07 was sent via certified mail to customers on 10/26/07. The letter informs customers of the firm recalling the product to correct faulty capacitors on the circuit board, and which could result in a malfunction that prevents or delays delivery of a defibrillating shock. The letter indicates the faulty AED10 devices were manufactured between 3/29/07 and 8/9/07. The letter also indicates the serial numbers affected, and where the serial number is located - upper left hand corner of the back panel of the device. In addition the letter includes a response form, the firm's contact information, and indicates that the firm will exchange the recalled defibrillator with a replacement and new five (5) year warranty. |  | 3588 | 1,794 distributed - 1176 in US and 618 internationally | Worldwide. | Terminated | November 09, 2011 | 1475 | 3588 | Hardware | Treatment/Delivery/Therapy | Replace device | Remove or Replace
Z-0113-2008 | 39400 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | GE Healthcare Voluson E8 ultrasound scanner with Software version 6.2.0 and 6.2.1 and DICOM stations with these software versions; GE Healthcare 9900 West Innovation Drive, Wauwatosa, WI 53226, USA. |  |  | July 06, 2007 | 2007 | Ge Healthcare | When a spectrum of a Pulsed Wave Doppler shows an angle correction unequal to 0 is reloaded from the archive or the image is sent to a DICOM station, this angle correction is mistakenly disregarded resulting in an underestimation of the flow velocities. | Software design | GE Healthcare sent customer letters on July 15, 2007 to affected customers. The letter included a short term mitigation recommendation and a long-term correction. The long-term correction will be software upgrade provided free of charge to all affected customers. |  | 521 | 521 Units | Nationwide. | Terminated | October 21, 2008 | 473 | 521 | Software | Display/Image | Software update | Software Update
Z-0095-2008 | 44788 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | FoundationÂż Hip System Bipolar Modular, Size 53mm OD; Cat. #412-01-053; Sterile; assembly required by physician prior to implant; manufactured by Encore Medical, L.P., Austin, TX 78758. |  |  | August 02, 2007 | 2007 | Encore Medical, Lp | Device mislabeled; hip prosthesis components validated for 5 years were labeled as having a 6 year expiration date. | Packaging change control | Firm notified consignees of recall via phone, e-mail and overnight letters beginning on 08/02/07. Consignees asked to return devices for a properly labeled replacement. |  | 137 | 137 units. | Worldwide, including USA, Greece, Japan, and Saudi Arabia. | Terminated | January 09, 2008 | 160 | 5174 | Not_Computer
Z-0114-2008 | 39409 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic Sofamor Danek METRx (TM) System Bayonet Electrosurgical Pencil, sterile, Manufactured at Warsaw, Indiana; REF (P/N) 9560575. |  |  | September 04, 2007 | 2007 | Warsaw Orthopedic, Inc. dba Medtronic SOFAMOR DANEK | Packaging weakness may compromise sterility and instructions for use are inadequate. | Package design/selection | Consignees were notified via letter dated September 4, 2007 to cease using the product and to contact the firm for instructions on how to return the product. |  | 22632 | 22,632 | Worldwide, including USA, Hong Kong, India, Mexico, Netherlands, Qatar, Saudi Arabia and South Korea. | Terminated | July 14, 2009 | 679 | 22632 | Not_Computer
Z-0144-2008 | 44957 | No | HIG | Hysteroscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Storz Equipment, Fluid Management Stand, for Fluid Management Systems, Product Number 26601EQ-KT |  |  | August 17, 2007 | 2007 | Karl Storz Endoscopy America Inc | Outflow Difference: The accuracy of the Equimat displayed value could be influenced by external factors of the cart/stand and may result in a fluctuating or inaccurate reading. | Device Design | All affected customers were notified via a Customer Letter issued on September 10, 2007 via FedEx overnight mail. The letter describes the problem with the device and the health hazard risk, and provides instructions for users to return the old stand and make arrangements for replacement stands. |  | 287 | 287 | Nationwide | Terminated | February 29, 2012 | 1657 | 287 | Not_Computer
Z-0100-2008 | 44792 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Portex Perfit Percutaneous Tracheostomy Tube Replacement Kit, Sterile |  |  | September 18, 2007 | 2007 | Smiths Medical ASD, Inc. | Mislabeled. The product is labeled as a size 9.0 mm tracheostomy tube, but the package contains a 7.0 mm tracheostomy tube. | Process control | Smiths Medical contacted direct customers and US distributors by letter sent by UPS on Septmeber 18, 2007. US Distributors instructed to return their inventory and provide a customer list to Smiths Medical. |  | 40 | 40 units | Nationwide. | Terminated | May 08, 2012 | 1694 | 40 | Not_Computer
Z-0121-2008 | 44950 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance CT Big Bore is a whole body computed tomography X-Ray system. The malfunctioning software application is the Tumor Localization (Tumor LOC) application. |  |  | September 27, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | The maximum intensity projection (MIP) images produced from respiratory gated images by the Tumor Localization (Tumor LOC) software application on Brilliance Big Bore CT scanners will be labeled incorrectly for certain patient orientations and view conventions when saved for later viewing. | Software design | Notification letters were sent to the installed customer base informing them of the hazard on 10/1/2007. A mandatory corrective software release due out mid October 2007, will be available at no cost to the customer. The software will be installed by Philips Field Service Engineers. |  | 206 | 206 systems | Worldwide, including USA, Canada, France, Finland, Korea, Australia, Italy, Switzerland, China, Austria, Thailand, Malaysia, Denmark, Norway, Jordan, Netherlands, England, South Africa, Spain, Germany, Bahrain, Belgium, New Zealand, Puerto Rico, South Korea, Saudi Arabia, Czech Republic, Sweden, Israel, and Lithuania | Terminated | March 02, 2012 | 1618 | 206 | Software | Output/Calculation | Remove | Remove or Replace
Z-0147-2008 | 44961 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT c16000 Processing Module, for In Vitro Diagnostics; List Number 3L77-01; Affected software: v2.10, v2.11, v2.12, v2.20, v2.20DB, v2.30, v2.60, v3.00 and v3.10; Distributed by Abbott Laboratories, Irving, TX. |  |  | September 05, 2007 | 2007 | Abbott Laboratories, Inc | Incorrect Sample/Test Identification: Due to a software timing defect with certain versions of ARCHITECT cSystem Software, test results may be incorrectly associated to a Sample Identification during specific sample-processing conditions of the Robotic Sample Handler. | Software change control | Notification letters sent to consignees on 09/17/07. Consignees informed of the software defect and the conditions that would cause the software error to occur. Consignees told they would be contacted by an Abbott Field Service Rep who would install ARCHITECT System Software v3.11 and were asked to respond via fax using Customer Reply form acknowledging receipt and understanding of correction letter. Firm recalling to hospital/user level. |  | 89 | 89 units. | Worldwide Distribution: USA (including Puerto Rico), Argentina, Australia, Bahamas, Bermuda, Brazil, Cayman Islands, Canada, Chile, China, Colombia, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Hong Kong, Jamaica, Mexico, New Zealand, Panama, Singapore, Thailand, Trinidad and Tobago, Uruguay, and Venezuela. | Terminated | October 04, 2010 | 1125 | 2159 | Computer
Z-0170-2008 | 45491 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Signa MFO/I Magnetic Resonance System, Model Numbers : 2210339, 2262700, 2276937, 2282853, 2282860, 2282880, 2352129, 2388781, 5105123, 5113339, 2138300-32, 2189500-3, 2198999-4, 2259620-4, 2262700-2, 2269512-2, 2286849-2, 2352129-2, 2377062-5, 2377062-8, Manufactured by GE Hangwei Medical Systems Co., LTD, Beijing PR, China |  |  | December 31, 2006 | 2006 | GE Healthcare | Incorrect patient positioning; when moving the patient into the bore may lead to wrist and hand injury. | Labeling design | Field Modification Instruction : FMI 67631 (DEC 2006) was communicated directly to the GE Healthcare service engineers who provide the modification at the customer site. Corrections were performed in DEC 2006 |  | 163 | 163 | Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, FL, ID, IL, IN, KS, KY, LA, MD, MO, MS, MT, NJ, NY, OH, OK, PA, SD, VA, and WA, and countries of United Arab Emirats, Argentina, Belgium, Brazil China, Ivery Cost, Germany, Spain, France, United Kingdom, Hungary, Italy, Japan, Korea, Mexico, Oman, Panama, Puerto Rico, Russia, and Turkey | Terminated | December 11, 2007 | 345 | 163 | Not_Computer
Z-0075-2008 | 44813 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Sunisa Flexi-DT Table, Model Numbers: 51335678, 2269647, 2242088-2, 5141609-2, 5115763, and 5128145, |  |  | August 01, 2007 | 2007 | GE Healthcare | Design Problems Including: 1) Hot Brake (Longitudinal Overheating) potential risk of burns 2) Brake Locks and Sudden Releases 3) Emergency Stop Circuit Button; open instead of closed 4) Longitudinal Brakes Release does not engage properly 5) Tilting Movement causing tabletop to slide; may cause risk of pinched fingers to patient and/or operator 6) Lodging Pins on floor Uncovered; may cause potenti | Device Design | Recalling firm initiated correction by visit on August 2007. Corrections include replacement of one or more of the following: charger board, pin lodging, rear bumper. Also, replace the interface board and remove resistor mounted on chassis, remove brake pedal and assembly of cosmetic brake or apply caution labels. |  | 96 | 96 | Class II Recall - Worldwide Distribution --- USA including States of -- AL, CA, FL, GA, IA, IL, IN, LA, ND, NY, OH, PA, TX, and WI, and Countries of -- BELGIUM, CZECH REPUBLIC, FINLAND, FRANCE, GERMANY, GUAYANA, HONG KONG, HUNGARY, IRELAND, ISRAEL, ITALY, KUWAIT, LATVIA, LEBANON, SAUDI ARABIA, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Terminated | March 02, 2008 | 214 | 96 | Hardware | Device Operation | Replace component | Remove or Replace
Z-0213-2008 | 38376 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Ximatron Digital Imaging (Ximavision), V 7.5.51 with Service Pack (SP), Oncology Information System, Varian Medical Systems |  |  | June 25, 2007 | 2007 | Varian Medical Systems Inc | Unexpected loss of Treatement Field Add-Ons (Wedges, Blocks, etc) during Ximatron simulation. If an unapproved treatment field is created without an MLC, but with an add-on, the add on will be deleted without warning. If this is not detected and corrected, incorrect dose delivery will occur. | Software design | The firm issued an Urgent Medical Device Correction letter to its customers by mail on June 25, 2007. It informed consignees of the anomaly in the software and workarounds for the issue. In addition, the firm also stated that they are currently validating a new software modification to repair the problem. |  | 14 | 14 units | Worldwide Distribution-USA (MO and NY medical centers) and countries of Switzerland, Bahrain, Canada, France, Germany, Japan, Norway and Sweden.--- | Terminated | February 04, 2008 | 224 | 14 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0218-2008 | 44826 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Texium Closed Male Luer (CML), Model Number 10012241, Cardinal Health, SanDiego, CA. |  |  | August 30, 2007 | 2007 | Cardinal Health 303 Inc DBA Alaris Products | Leaks: there is a risk of leakage at the female end of the Texium closed male luer (CML) when attached to a spin luer. (NOTE: users who connect the female portion of the Texium closed male luer to fixed (one-piece) male luers on tubing or on syringes are not affected by this issue) | Finished device change control | Recall notification was made via a Safety Alert Notification Letter dated August 30, 2007. Recall will be conducted through consignees (distributors) in the affected countries, as well as, direct customers for the Texium closed male luer. The customers and distributors will be notified by registered return receipt mail. For direct customers, separate letters will be sent to the Director of Nursing, Director of Materials and Director of Risk Management, and Director of Pharmacy at each facility. Distributors were sent multiple copies of the Safety Alert Notification and Tip Sheets; along with a cover letter outlining their responsibilities to complete the Safety Alert plan and notify their sub-account customers. The domestic letters (direct customers and distributors) were mailed on August 30, 2007 to all of the affected accounts. The international notifications will be coordinated out of the offices located in the affected countries. The appropriate international regulatory authorities will be notified of this issue, which will work with the country managers at each location. |  | 76300 | 76, 300 units | Worldwide: USA including states of Arkansas, California, Colorado, Hawaii, Illinois, Indiana, Louisiana, Massachusetts, Maryland, Maine, Michigan, Mississippi, North Carolina, New England, New Hampshire, New Jersey, New York, Ohio, Oregon, Pennsylvania, South Carolina, Texas, Utah, Virginia, Washington, and Wisconsin. and countries of Australia, Canada, Japan and Switzerland | Terminated | March 28, 2012 | 1672 | 76300 | Not_Computer
Z-0159-2008 | 38516 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVATRON M-2 6300,Medical charge-particle radiation therapy system (Medical Digital Linear Accelerator),MODEL NUMBER: 9401316,MANUFACTURED BY: Siemens Medical Solutions USA, Inc, 4040 Nelson Ave, Concord, CA 94520 |  |  | March 22, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Mistreatment of radiation therapy due to inadvertent parameter changes or miss-calibration in Service Mode. The miss-calibration can affect the beam profile and Virtual Wedge angle. This leads to patient mistreatment and alteration of future treatment plans. A second reason is an unplanned Dosimetry calibration data set can be erroneously copied during "Service Mode." | Software design | Distribution of Update Instructions TH007/07/S was sent to all affected sites on March 26, 2007. This was intended to inform users that Service Mode is only for use by authorized personnel, and that unauthorized and/or incorrect softpot parameter changes performed in Service Mode can result in the risk of patient mistreatment.Current control system software versions do not allow unplanned Dosimetry calibration data sets to be erroneously copied during "Service Mode." |  | 6 | 6 units | Worldwide:USA including states of Alabama, Alaska, Arizona, Arkansas, California, Colorado, Delaware, District of Columbia, Florida, Georgia, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Puerto Rico, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Vermont, Virginia, Washington, West Virgina, Wisconsin, and Wyoming and FOREIGN COUNTRIES: Canada, Mexico, Serbia Montenegro, Angola, Argentina, Australia, Austria, Belarus, Belgium, Bosnia, Herzeg-Bosnia, Brazil, Bulgaria, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, DPR of Korea, Ecuador, Egypt, El Salvador, Estonia, France, Germany, Greece, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Lithuania, Malaysia, Morocco, Nepal, Netherlands, New Zealand, Norway, P.R. China, Pakistan, Palestine, Peru, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom, Uzbekistan, Venezuela, Vietnam, and Zambia | Terminated | December 18, 2010 | 1367 | 2120 | Computer
Z-0119-2008 | 44880 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OPTIÂż LION E-Plus Cassettes, Model BP7507, (25 per box), In Vitro Diagnostic, OPTI Medical Systems, Inc., Roswell, Georgia 30076 |  |  | July 18, 2007 | 2007 | Opti Medical Systems Inc | Incorrect results/stability failure; Actual device stability studies show sodium and chloride sensors age differently than anticipated when stored at the high end of labeled storage, therefore subsequent measurements give results outside the allowed error allowance before the 6 month expiration date shown on the device. | Device Design | An Urgent Medical Device Recall Letter dated 16 July 2007, was issued via E mail and Fax on/or about 7/18/2007 directing consignees to remove unused cassettes from inventory and destroy them. Replacement cassettes will be shipped upon receipt of information regarding destruction by calling customer service at 800-490-6784. Sub-recall by distributors was requested. |  | 7225 | 289 boxes of 25 cassettes | Worldwide Distribution-USA, Australia, Chile, China, Egypt, France, Great Britain, Guatemala, Indonesia, Italy, Korea, Lebanon, Mauritania, Morocco, Nepal, Russia, South Africa, Sweden, Tanzania, Thailand, and Venezuela. | Terminated | December 13, 2007 | 148 | 7225 | Not_Computer
Z-0256-2008 | 45332 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | BioPlex 2200 System, Cat. No. 665-1460A containing APF CD SW2_v1., Syphilis IgG IFU Packet contains the Syphilis IgG Assay Protocol File LAPF CD |  |  | September 25, 2007 | 2007 | Bio-Rad Labroatories Inc. | Mislabeling: Reporting terminology in the Syphilis IgG APF CD is not consistent with the distributed Instructions for Use. (Non-Reactive and Reactive rather than Negative and Positive) | Labeling Change Control | The recalling firm notified consignees by phone call on 9/25/2007, followed up with a fax notification. The firm plans to monitor the consignees that do not respond by tracking on a spreadsheet and following up with another phone call. Consignees are instructed to destroy the APF CD which will be replaced at a later date. |  | 36 | 36 | Nationwide | Terminated | April 16, 2008 | 204 | 36 | Not_Computer
Z-0059-2008 | 39392 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Medical System - BV Pulsera Mobile C-arms with Remote Control Option, (Release 2) |  |  | August 23, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | System Shut Down -- Unexpected Shut Down during use of the Optional Remote Control. | Software design | On 8/23/07, the firm sent out a Safety Notification letter, via certified mail. On 8/29/07, the firm sent out an URGENT - DEVICE CORRECTION letter via regular mail, dated 8/20/07. Both letters advised customers that if the device shuts down during use, it can be restarted within about 1 minute by switching the system off and then on again, using the system on/off buttons. The letter(s) states that software is being developed to correct the problem. |  | 596 | 596 units | Nationwide | Terminated | September 27, 2010 | 1131 | 633 | Other | Device Operation | Software update | Software Update
Z-0181-2008 | 37661 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE OEC, InstaTrak 3500 Plus FluoroTrak and Fluoro CAT Spinal Applications used with Regular and Super 9800 and 9600 Arm Systems. GE Health Medical Systems, Inc, Salt Lake City, UT 84116-2862 |  |  | March 23, 2007 | 2007 | GE OEC Medical Systems, Inc | Tracking Accuracy - Image intensifier supplier's change on the C-Arm could lead to patient injury, due to incorrect anatomical location of the instrument during a procedure. | Component change control | Consignees were sent an Urgent Recall Notice letter, via mail dated 10/11/07, advising users to stop usage of the FluoroTrak and FluoroCAT spinal applications on the InstaTrak 3500 Plus System. Once a solution is identified, a representative will contact consignees to arrange for installation of the solution at no charge. |  | 61 | 61 units | Class II Recall - Worldwide Distribution - USA including states of AZ, CA, FL, GA, IA, IL, KS, LA, MA, MD, MI, NC, NY, OK, PA, SC, TX, UT, VA, and WI, and countries of Canada, China, Germany, Hong Kong, Italy, Japan, Mexico, and Spain. | Terminated | April 16, 2012 | 1851 | 61 | Software | Physical Safety Hazards | Software update | Software Update
Z-0219-2008 | 37214 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Zoom (Atlas) powered patient transport stretcher, Model 660, Stryker Medical, Portage, MI 49002 |  |  | November 15, 2004 | 2004 | Stryker Medical Div. of Stryker Corporation | The brake cams may crack or develop excessive wear, resulting in the brakes not locking the wheels, or the stretcher not being locked firmly in position. | Component change control | The firm issued an Urgent Medical Device Correction Letter dated 9/24/07, to each consignee informing them of the problems, that the brakes on each affected stretcher needed to be inspected/tested and removed from service if brake problems were discovered, and that a Stryker service representative would be in contact to replace the brakes. |  | 81367 | 81,367 (61,708 U.S. and 19,569 ex-U.S.) for all models. | Worldwide-USA, Argentina, Australia, Brazil, Canada, Chile, China, Dubai, France, India, Italy, Japan, Germany, Greece, Hong Kong, Korea, Latin America, Malaysia, Mexico, Netherlands, Poland, Portugal, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, and United Kingdom. | Terminated | October 27, 2010 | 2172 | 81367 | Not_Computer
Z-0182-2008 | 44944 | No | FLG | Ultrasonic cleaner for medical instruments. | General Hospital | General Hospital | 510(K) Exempt | Clarisonic Skin Care System - A hand held cleansing device for the skin. The Unit comes with a charger cradle for charging the device. Pacific Bioscience Laboratories, Inc., 13222 SE 30th St., Bellevue, WA 98005 |  |  | July 03, 2007 | 2007 | Pacific Bioscience Lab, Inc. | Overheating: The handle or charger/cradle can overheat, melting the plastic, which potentially can lead to smoke and fire and has caused mild user burns. | Device Design | On August 30, 2007, the firm issued a press statement via PR Newswire and The Associated Press. They also posted a Recall Notice on their website (www.Clarisonic.com) requesting consumers to return their affected products for an exchange. All retailers and professional resellers were notified by telephone, E mail and/or letter and asked to return unsold inventory, as well as, post the provided point of sales notices and send provided post cards to known customers requesting exchange of recalled product. All users can contact the firm via the website or by phone at 1-866-378-7360 (6 am-6pm, M-F, Pacific Time) to arrange for the device exchange. |  | 97500 | 97,500 | Worldwide Distribution - including USA and countries of Japan and Canada. Units were distributed to retailers and professional skin care clinics throughout the USA. In addition, units were distributed to one distributor in Japan and to 4 clinics in Canada. | Terminated | December 13, 2010 | 1259 | 97500 | Not_Computer
Z-0217-2008 | 44962 | [6 empty] | Matrix MAX3 Actuator, Catalog Number: MAX30 A300515A254L000, SKF USA, Inc. Norristown, PA 19403 |  |  | July 24, 2007 | 2007 | Skf USA Inc | Sudden failure of the actuator during lifting operation | Labeling design | On or about 7/24/07, the firm sent an Important Safety Information letter and instructions to all known OEMs who purchase and sell SKF actuators for use in patient lifts. The letter reiterated and emphasized the importance of proper handling and maintenance for the actuator. The firm enclosed with this correspondence a proposed letter that the OEMs were asked to send to their customers, including the safety instructions and a new warning label. The firm also is making available for OEMs, dealers and their customers new warning labels to place on all actuators in service. The root cause of the problem is still under investigation by the recalling firm and additional instruction may follow later |  | 1298 | 1,298 units (Z-0215-0217-2008) | Worldwide Distribution-USA and Canada | Terminated | April 26, 2013 | 2103 | 1298 | Not_Computer
Z-0214-2008 | 45106 | [6 empty] | Stryker T2 Locking Screw, Fully threaded; 05 x 70mm, Catalog Number: 1896-5070S, Sterile, Qty: 1,Stryker Trauma GMbH, 24232 Schonkirchen, Germany |  |  | September 27, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Mislabeled: One lot of screws (05x70mm) is labeled as containing 70mm, but may actually contain 75mm length screws. | Packaging process control | Urgent Product Recall letters were sent on September 27, 2007, to all Stryker Branches/Agencies and hospitals who received the product requesting them to remove and quarantine the product for Stryker pick-up. |  | 50 | 50 | Nationwide | Terminated | May 08, 2008 | 224 | 50 | Not_Computer
Z-0176-2008 | 38464 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SmartSeal Wedge Femoral Pressuriser; Product Code/Catalog No: 3206002, DePuy Orthopaedics, Inc., 700 Orthopaedics Drive, Warsaw, Indiana 46581-0988-This product allows for effective and controlled pressurization of bone cement into the femur and the acetabulum during hip arthroplasty procedures |  |  | July 13, 2007 | 2007 | Depuy Orthopaedics, Inc. | Package Integrity: The outer peelable pouch seals may be compromised due to creasing, although the inner pouch seals are intact. | Packaging process control | Sales Force Managers were notified of this recall by an Urgent Information Recall Notice Letter, via email, on 7/13/07 and 7/23/07, to visit their Consignees. Letters for the Hospitals (Dear Materials Manager) and Dear Doctors Letters were provided as background by the Sales Force Managers during their visits. If any product remained in inventory, it should be quarantined immediately, and returned to DePuy. Hospitals should be notified immediately not to use the affected device. Replacement stocks will have a check mark on the bottom right front of package to indicate it has passed inspection. |  | 3530 | 3,530 | Nationwide Distribution -- including States of Arkansas, California, Colorado, Florida, Idaho, Illinois, Indiana, Kentucky, Maine, Michigan, Mississippi, New Mexico, North Carolina, Pennsylvania, South Dakota, Texas, Utah and Washington. | Terminated | June 23, 2009 | 711 | 3530 | Not_Computer
Z-0254-2008 | 44811 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | NEMIO XG Ultrasound System, Model SSA-580A , Software versions: VI .5.0, Vl .5.000A, VI .5.000B, V2.0.0,Toshiba America Medical Systems, Inc., Tustin, CA. |  |  | August 29, 2007 | 2007 | Toshiba American Med Sys Inc | Multiple problems including : (1) The surface temperature of the transducer tip can rise and exceed the limit set by the international standard. (2) Incorrect display of the TI indices on the display monitor. The displayed value is greater or smaller than the true TI value, and (3) Manuals contain incorrect values for MI/TI values. | Software design | Urgent Medical Device Notification letters were sent to customers on August 29, 2007. Customers were notified of the several issues related to the Nemio SSA-550A and Nemio SSA-580A recall and asked to share this information with all users of the device. Customers were informed that Toshiba is preparing software and corrected manuals for distribution to address these issues. If customers have any questions they were instructed to contact their Toshiba Customer Service Engineer or call (800) 421-1968 and ask for the Regulatory Affairs Department. |  | 128 | 128 units | Nationwide | Terminated | April 05, 2012 | 1681 | 1406 | Computer
Z-0174-2008 | 37879 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Electric Mobility Corporation (EMC)- Scooters and Turnabout Powerchairs with either an Electric Seat Lift Actuator or a Seat Post to attach the seat to the vehicle; Model Number(s) 300F, 600C, 600T, and 600F. Electronic Mobility Corporation, 599 Mantua Blvd, One Mobility Plaza, Sewell, NJ 08080 |  |  | March 22, 2007 | 2007 | Electric Mobility Corp | Seat and/or lift actuator problems: 1) Models equipped with a Seat Lift, the mast screw and washer holding the Seat Lift Actuator in place will loosen break off or fall out; 2) Models assembled with Seat Lift Actuator or Seat Post Mounting Bolts, the seat may wobble or the seat will break off from the vehicle; 3) Models using Seat-On Lifting brackets with a hoist to raise an EMC vehicle equipped | Process change control | Consumers were sent an Urgent Medical Device Recall Notification on October 12, 2007, requesting that they immediately respond to the Medical Device Recall by calling EMC at 1-800-814-4365, or by returning an enclosed Recall Response Form, or by responding on-line at www.EMCIT.NET/801B. The Notification address three important safety issues on 1) Seat Lift Actuator Mast Screws and Washers; 2) Seat Lift Actuator or Seat Post Mounting Bolts and; 3) Incorrect Seat-On Lift Bracket Installation. Corrections to the devices will be made by the nearest Service Representatives, by scheduled in-home repairs, or return of the rear section of the scooter with EMC providing shipping material and freight payment. |  | 12625 | 12,625 | Nationwide | Terminated | April 08, 2014 | 2574 | 12625 | Not_Computer
Z-0175-2008 | 37880 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Rascal Scooters with Electric seat lifts, Model 600 Series, distributed between 10/27/2005 - 3/31/2006, Electric Mobility Corporation (EMC), Sewell, New Jersey 08080 |  |  | March 29, 2007 | 2007 | Electric Mobility Corp | Seat and/or lift actuator problems: 1) Models equipped with a Seat Lift, the mast screw and washer holding the Seat Lift Actuator in place will loosen break off or fall out; 2) Models assembled with Seat Lift Actuator or Seat Post Mounting Bolts, the seat may wobble or the seat will break off from the vehicle; 3) Models using Seat-On Lifting brackets with a hoist to raise an EMC vehicle equipped w | Process change control | Consumers were sent an Urgent Medical Device Recall Notification on October 12, 2007, requesting that they immediately respond to the Medical Device Recall by calling EMC at 1-800-814-4365, or by returning an enclosed Recall Response Form, or by responding on-line at www.EMCIT.NET/801B. The Notification address three important safety issues on 1) Seat Lift Actuator Mast Screws and Washers; 2) Seat Lift Actuator or Seat Post Mounting Bolts and; 3) Incorrect Seat-On Lift Bracket Installation. Corrections to the devices will be made by the nearest Service Representatives, by scheduled in-home repairs, or return of the rear section of the scooter with EMC providing shipping material and freight payment. |  | 1756 | 1,756 | Nationwide: to consumers via direct purchase from Electric Mobility Corporation. | Terminated | April 08, 2014 | 2567 | 1756 | Not_Computer
Z-0180-2008 | 38276 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips, Voxel Q workstations, untilizing AcQSim or AcQPlan software version 4.2, 4.2.1, 4.9, 4.9.1, 4.9.2, 4.9.3, 5.0, 5.0.1, 5.0.2 or 5.0.3 and any Mx8000, Mx8001DT, or Brilliance 6, 10, 16, 16P, and Big Bore CT systems connected to the VoxelQ, Philips Medical Systems, Cleveland, OH 44143 |  |  | March 01, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | Incorrect simulation and/or beam placement: An anomaly has been identified with the Philips VoxelQ workstation utlizing AcQPlan software version 4.2, 4.2.1, 4.9, 4.9.1, 4.9.2, 4.9.3, 5.0, 5.0.1, 5.0.2 or 5.0.3 when computing absolute patient marking coordinates with data containing reconstruction offsets from the Mx800, the Mx8001DT, and the Brilliance 6, 10, 16, 16P, and Big Bore CT systems. Whe | Software design | On November 1, 2007, after telcon with FDA, the firm revised its recall strategy to notify consignees via a November 9, 2007, Urgent Device Notification letter which indicates that warning labels will be affixed to the units that alerts the user of the reconstruction offset defect by Philips employees. |  | 520 | 520 Devices | Worldwide Distribution- USA including states of AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV and WY and Canada | Terminated | March 01, 2012 | 1827 | 520 | Software | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-0245-2008 | 44894 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clearview hCG Combo Waived for urine &amp; Moderate for serum, Wampole Laboratories, ABI List No/Product No. 6026KCV; Contents 30 Individually pouched test cassettes with disposable pipettes, 1 Directional Insert; Inverness Medical Professional Diagnostics Group; Marketed by Wampole Laboratories, 2 Research Way, Princeton NJ 08540 USA |  |  | September 28, 2006 | 2006 | Applied Biotech Inc | Sensitivity: Kits contain test devices which may not meet the label claims for sensitivity, therefore use of such devices may produce false negative results with patient samples at or near the limit of detection of the product. | Process design | A Product Recall Notification dated 9/28/2006, was provided via fax, email or direct mail to all primary consignees. Customers were instructed to block all stock of affected lots in their warehouses, contact their primary consignees and advise them to discontinue use immediately and dispose of any unused materials; using a fax back form to communicate to the recalling firm the quantity blocked and disposed of so that arrangements could be made for providing replacement product |  | 3238 | 3,238 Kits (97,140 tests) | Worldwide: USA and Canada | Terminated | August 03, 2012 | 2136 | 3638 | Not_Computer
Z-0212-2008 | 45486 | No | DSQ |  |  | Cardiovascular | PMA | Thoratec Implantable Ventricular Assist Device (IVAD pump) |  |  | October 19, 2007 | 2007 | Thoratec Corp | Pneumatic Leaks: air leaks developing in the pneumatic driveline when subjected to excessive stress at the junction of the percutaneous line to the pump housing due to sharp bending when implanted (paracorporeal-external-position). | Device Design | An initial communication was sent to all affected IVAD customers on October 19, 2007, in the form of an "Urgent: Medical Device Correction" notice informing them of the problem. The correction notice requests that its customers/physicians take the following actions:1) Contact all ongoing patients with IVADs placed in the paracorporeal position toevaluate the percutaneous line adjacent to the pump housing to ensure that it is not subject to extreme flexing or bending at acute angles. 2) Remind implanting surgeons and support staff of proper percutaneous line management and provide the patient with a copy of the recall information provided. Complete and sign an Acknowledgement form and fax it to Thoratec Regulatory Affairs. This is a User Level Recall. The firm also issued a November 7, 2007, Press Statement- Urgent Thoratec Corporation Worldwide Recall of Implantable Ventricular Assist Device. |  | 766 | 766 | Worldwide: USA including states of AL,AR,AZ,CA,CO, FL,GA, IA,IL, IN,KS,KY,LA, MA, MD, ME,MI, MN, MO,NC, NJ, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA, and WI and to the following countries: Austria, Canada, France, Netherlands, Singapore, Switzerland, and The United Kingdom | Terminated | May 26, 2009 | 585 | 766 | Not_Computer
Z-0259-2008 | 38445 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific 6F Guider Softip XF Guiding Catheter, 6 French, 40 degree angle, 90cm usable length, Material Number (UPN) H965100420, Catalog Number : 10042, Made in USA: 5905 Nathan Lane, Plymouth, MN 55442 USA |  |  | July 11, 2007 | 2007 | Boston Scientific Target | Mislabeling: Catheters were found to be mislabeled (90 cm labeled devices actually contains 100 cm length catheter) | Labeling mix-ups | Firm notified consignees by Urgent Medical Device Recall Letter dated 11 Jul 07 requesting return of product. |  | 47 | 47 units | Worldwide Distribution-USA, EU, Japan, South Korea, Mexico, Colombia, and Hungary. | Terminated | December 05, 2007 | 147 | 96 | Not_Computer
Z-0255-2008 | 45327 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Interlok 67 mm Fixed Cruciate Tibial Plate with Locking Bar; CO-CR-MO/TI 6AL 4V Alloy; Product Number/REF 141232, Biomet Orthopedics, Inc., Warsaw, IN 46582 |  |  | September 18, 2007 | 2007 | Biomet, Inc. | Mislabeled as to size: the package actaully contained a 71 mm tray. | Labeling mix-ups | Consignees were notified via letter dated 9/18/07 to discontinue use of the product and to return it to the firm. |  | 7 | 7 | Nationwide : Distributors in Georgia, Kentucky, Texas and Washington. | Terminated | June 04, 2008 | 260 | 7 | Not_Computer
Z-0257-2008 | 45429 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott ACCELERATOR APS System Input/Output Module (IOM); List Number: 7L01-01; Abbott Laboratories, Abbott Park, IL 60064 |  |  | October 10, 2007 | 2007 | Abbott Laboratories, Inc | Sample problems; Five issues identified related to software utilized by the ACCELERATOR APS System Input/Output Module, to include: Under three specific conditions, the post-aspiration radio-frequency identification (RF ID) read of the tube carrier is not performed and the APS work cell may not appropriately generate Sample Presentation/Sample Queue errors with APS software version 1.1.1 or earlie | Software design | Firm notified consignee via letter on 11/10/07. Consignee notified of software issues and informed they would be contacted by an Abbott representative to schedule installation of software upgrade. |  | 1 | 1 unit. | Nationwide: including state of UT. | Terminated | March 10, 2008 | 152 | 1 | Software | Device Operation | Software update | Software Update
Z-0262-2008 | 45464 | No | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Edwards Vigilance II Continuous Cardiac Output/Oximetry/Volumetric (CCO/SvO2/CEDV) Monitor Model: VIG2/VIG2E Software Versions: 00.51,00.60, 00.61, 00.62, 00.63, 00.64, 00.67, 00.69, Edwards Lifesciences LLC., Irvine, CA 92614 |  |  | May 17, 2007 | 2007 | Edwards Lifesciences, Llc | MAP offset errors: If users are "slaving" mean arterial pressure (MAP) from the bedside monitor into the Vigilance II Monitor for the purpose of the continuous calculation of systemic vascular resistance (SVR), at times, on some bedside monitors, the Vigilance II monitor may display a MAP offset of as much as 25mmHg. Since the MAP is used in calculations of various patient parameters such as SVR, | Software design | An Urgent Customer notification letter dated May 17, 2007 was mailed to all Vigilance II customers advising that If they are "slaving" mean arterial pressure (MAP) from their bedside monitor into the Vigilance II Monitor for the purpose of the continuous calculation of systemic vascular resistance (SVR), to be aware that at times, on some bedside monitors, the Vigilance II monitor may display a MAP offset of as much as 25 mmHg. Since the MAP is used in calculations of various patient parameters such as SVR, these may also not be accurate. A table is provided showing those bedside monitors with the problem: GE Solar, Phillips (HP) M1006A, M1006B, as well as, those bedside monitors verified to not have the problem: SpaceLabs Model 90402 and Nihon Kohden Model BSM 8000. The letter also serves to remind the customer of the Warning in the Vigilance II Directions for Use which is also provided in the letter that instructs user to verify slaved MAP values with those displayed on bedside monitor. A temporary work around for this offset problem is to manually enter the MAP based on the readings from the bedside monitor. All Vigilance II calculations from the MAP have been re-verified to be correct when the MAP is accurate. The problem is in the transfer of the MAP between the bedside monitor and the Vigilance II. Customers are also advised that if they have the GE Solar or the Phillips (HP) models listed in the letter or any monitor not listed as verified to be compatible with the Vigilance II (Space labs and Nihon Kohden models listed above) Edwards recommends verification of MAP accuracy prior to treating based on parameters from a Vigilance II in a "slave" configuration. If there is a difference, they should enter the MAP into the Vigilance II manually for a static calculation of SVR. Edwards will be supplying a software update to correct this problem. |  | 2680 | 2,680 units in North America distribution | Worldwide: USA and Canada | Terminated | October 09, 2009 | 876 | 2680 | Software | Output/Calculation | Software update | Software Update
Z-0263-2008 | 45489 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Baxter Duploject Easy Prep, Fibrin Sealant Preparation and Application System ISO; Product Codes: 1500949, 1500950, 1500951, 1501250 &amp; 1501252, Baxter Healthcare Corporation, Deerfield, IL 60015 |  |  | October 25, 2007 | 2007 | Baxter Bioscience | Sterility (package integrity) compromised: This recall is being conducted due to incomplete seal of the product package at the corner of the outer tray and tyvek lid. | Packaging process control | An Urgent Product Recall Letter dated October 24, 2007, was issued with attached Product Recall Instructions and Customer Response forms to all consignees (Director of Materials Management, and Director of Risk Management). The recall letter requests sub recalls to the retail level and product quarantine and return. The recall notification letters were sent via commercial carrier (UPS Next Day Service). |  | 4782 | 4,782 units (6 units/case) | Worldwide-USA, Canada and European Union (France and UK) | Terminated | March 13, 2009 | 505 | 4782 | Not_Computer
Z-0264-2008 | 45539 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Vena Tech LP Vena Cava Filter System, Item No/REF. 05010024, Model No: 40DLP, B Braun. Manufacturer: B. Braun Medical 204, avenue du Marechal Juin BP 331 92107 Boulogne Cedex , France. (The VenaTech LP vena cava filter is pre-loaded in a syringe (cartridge) and provided as a system with introducer accessories and instructions to accommodate delivery and implantation either via the femoral or jugular approach.) |  |  | October 30, 2007 | 2007 | B. Braun Interventional Systems | Sterility (package integrity) compromised: package sealing problem leaving packaging being partially open on one of the four sealed Tyvek pouch edges. | Packaging process control | B. Braun sent an Urgent Medical Device Removal letter (10/10/07) by express mail to each of the hospital accounts that received the recalled lot of VenaTech LP. B. Braun sales represenatatives will also contact each account directly to inquire if any inventory of the suspect lot remains and to provide guidance for return and replacement of any existing inventory. |  | 60 | 60 units | Nationwide: States of AL, GA, PA, FL, NE, WI, IL, LA, TX, AZ, MN, MA, CA, NY, OH, IN, and MI. | Terminated | March 02, 2008 | 124 | 60 | Not_Computer
Z-0052-2008 | 44793 | Yes | LWP |  |  | Cardiovascular | PMA | Guidant CONTAK RENEWAL 3 (Models H170, H173, H175), Guidant CONTAK RENEWAL 3 HE (Models H177, H179). Guidant CONTAK RENEWAL 4 (Models H190,H195) *, Guidant CONTAK RENEWAL 4 HE (Models H197, H199)* Cardiac Resynchronization Therapy Defibrillator (CRT-D). Sterile EO. Guidant Corporation, 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. (* models are not available in the USA). |  |  | March 16, 2007 | 2007 | Boston Scientific CRM Corp | Atypical Charge Time Behavior: 1)-End of Life (EOL) or ERI is displayed during mid-life (typically 24-48 months), even though battery capacity remains available. 2)- Extended ERI charge time limits: Charge times during mid-life may remain below a normal extension of the ERI charge time . ****** Prophylactic replacement of these devices prior to ERI is not recommended. As stated in the Product Up | Device Design | A Product Update titled "ERI Charge Time Limit Extended During Mid-Life and Mid-Life Display of Replacement Indicators", dated 03/10/2007 was posted to Company's website on 03/13/2007. A CRM Product News letter was sent to physicians beginning 03/16/2007. The CRM Product News Letter gave an overview of the content of the Product Update and stated the web location for finding the complete Product Update. The Product Update identified affected devices, gave an overview of normal charge time behavior at mid-life for specific ICD and CRT-D devices. The letter also described the atypical charge time behavior, projected rate of mid-life display of replacement indicators and Patient Management Considerations. ****** Prophylactic replacement of these devices prior to ERI is not recommended. As stated in the Product Update, "Devices that have triggered charge time-based ERI or EOL during mid-life have several months, and in most cases more than one year of remaining battery voltage and capacity, which allows the devices in this pattern to continue to provide brady and LV pacing and maximum energy shocks. However, if ERI or EOL is triggered, device replacement should be scheduled."*****Recalled units only cover devices implanted prior to July 2005**** |  | 29375 | 22,761 within USA, 4,609 Foreign-Recalled units only cover devices implanted prior to July 2005 | Worldwide | Terminated | November 23, 2007 | 252 | 29375 | Not_Computer
Z-0291-2008 | 45526 | No | HWR | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Pioneer Quantum Spinal Rod System, MIS Yoke Manipulator, Long Slot, non-sterile; Cat # 51-YOKEMAN-L, Pioneer Surgical Technology, Marquette, MI 49855 |  |  | October 10, 2007 | 2007 | Pioneer Surgical Technology | The pins may come out of the instrument and, if this happens, they could fall into/remain in the patient. | Component change control | Consignees were notified via a Product Recall letter dated 10/10/07 to remove these instruments from their instrument cases and to return them to the firm. |  | 38 | 38 | Nationwide: including states of California, Florida, Michigan and North Carolina. | Terminated | March 31, 2008 | 173 | 38 | Not_Computer
Z-0294-2008 | 45492 | [6 empty] | Medtronic 0-arm Imaging System, Product Catalog Number: B1-700-00027, Medtronic Navigation, Inc., Littleton, MA 01460 |  |  | October 17, 2007 | 2007 | Medtronic Navigation, Inc | Reversed Instrument Position : a defect in the software version 3.0 causes an error in the position of surgical instruments and relative motion are reversed 180' on the StealthStation image. (If the StealthStation is not used, or if the patient orientation is prone or supine, the problem does not occur) | Software change control | Medtronic Navigation notified customers by letter on October 17, 2007. Medtronic Field service will perform upgrade (software version 3.0.1) at user site. |  | 7 | 7 units | Worldwide: USA including states of CO, IN, NE, and MN and countries of Canada and Germany | Terminated | November 27, 2007 | 41 | 7 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0290-2008 | 45462 | No | HWR | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Odyssey Femoral Holder/Driver, Part Number K001-7410, Non-sterile, 1 each, Wright Medical Technology, Inc., 5677 Airline Road, Arlington, TN 38002 |  |  | October 04, 2007 | 2007 | Wright Medical Technology Inc | Breakage: breaking of the connecting foot that is on the distal surface of the instrument | Device Design | The firm notified their distributors of the recall on October 4, 2007, with an Urgent Medical Device Recall letter explaining the problem and requesting return of the product from their inventory and from their consignees. The hospitals were notified of the recall on October 4, 2007 using the same Urgent Device Recall Letter. |  | 66 | 66 units | Worldwide-USA and Australia | Terminated | July 22, 2009 | 657 | 66 | Not_Computer
Z-0296-2008 | 45529 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT for ENT, Xoran Technologies, Incorporated, Ann Arbor, MI 48103. This device is a cone beam CT scan for sinue, temporal bones and maxillofacial imaging. |  |  | July 13, 2006 | 2006 | Xoran Technologies, Inc. | The hard drive may crash, resulting in loss of data from patient CT scans. | Finished device change control | Consignee accounts were visited between 7/25/06 and 8/10/07 and the hard drives on their systems were replaced. |  | 5 | 5 | Nationwide Distribution: California, South Dakota, Tennessee, Utah and West Virginia. | Terminated | March 13, 2008 | 609 | 5 | Other | Device Operation | Replace | Remove or Replace
Z-0293-2008 | 45567 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba-Interventional Angiography System, Infinix i, Model DFP-8000D, equiped with Digital Radiography System with Software version 3.4er000, Toshiba America Medical Systems, Inc., Tustine, CA 92780 |  |  | October 01, 2007 | 2007 | Toshiba American Med Sys Inc | System Lock-up: When performing DAS One Shot, the system may lockup due to a software timing error. | Software design | Notification letters (Urgent: Medical Device Correction-dated October 1, 2007) with an attached Customer Response form were sent to customers beginning on October 2, 2007. Customers were notified this action was being taken as a result of an investigation that resulted from a software issue that was found done by the manufacturer. Customers were informed that a Toshiba Customer Service Engineer would contact them to schedule a software upgrade when the software becomes available. If customers have any questions they were instructed to contact their or call (800) 421-1968 and ask for the Regulatory Affairs Department. |  | 23 | 23 | Nationwide-including states of AL, AR, FL, GA, MI, NY, OH, PA, TX &amp; WV | Terminated | April 05, 2012 | 1648 | 23 | Software | Device Operation | Software update | Software Update
Z-0171-2008 | 37912 | No | LYJ |  |  | Neurology | PMA | Cyberonics Neuro Cyberonics Prosthesis System NCP Bipolar Lead, Model #300; Sterile, Single use only; Distributed by Cyberonics, Inc., Houston, TX 77058. |  |  | August 02, 2007 | 2007 | Cyberonics, Inc | Dissolution/Fractures to the leads of the VNS Therapy System | Process control | Cyberonics issued a correction notice (Safety Alert) to physicians in June of 2007, informing them of the present labeling and including updated labeling to provide instructions on how to detect lead fractures and how to minimize risk to patients. |  | 21419 | 21,419 units. | Worldwide Distribution - USA and countries of Argentina, Australia, Austria, Belgium, Canada, Costa Rica, Cyprus, Czech Rep, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jordan, Korea, Malta, Mexico, Netherlands, Northern Ireland, Norway, Poland, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Kingdom, and Venezuela. | Terminated | October 05, 2009 | 795 | 56189 | Not_Computer
Z-0179-2008 | 45436 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Edwards Lifesciences LifeStent FlexStar Bilary Stent System Application, Model EX080801CD; 8 MM x 80 MM stent, 80 CM Delivery System, Lot FR6K0319, Edwards Lifesciences, LLC., Irvine, CA 92614 |  |  | March 19, 2007 | 2007 | Edwards Lifesciences, Llc | Failure to deploy. | Process control | This recall was executed by an Urgent-Product Recall letter dated March 15, 2007, which was delivered via overnight service to the affected accounts on 3/19/07. In addition, the affected sales force was notified by Email and given a copy of the recall letter. The sales force was asked to assist in retrieval of the product. The letter advises that one lot of the Edwards LifeStent FlexStar Bilary Stent System -- Model EX080801CD; 8 MM x 80 MM stent, 80 CM delivery system, Lot FR6K0319 may exhibit incidences of failure to deploy and is being recalled. Customers are advised to check their inventory for lot #FR6K0319 and remove all unused product from stock. Product is to be returned or it will be retrieved. The letter also advises that if the product has been implanted, there is no need to inform the patient of this notification as this is solely related to the delivery system and the characteristics of the stent are not affected in any way. Either the customer or the sales representatives will return written notification of the status of product from this lot which had been shipped. |  | 12 | 12 units | Nationwide including states of OK, GA, CA WI, MO, and LA | Terminated | March 26, 2012 | 1834 | 12 | Not_Computer
Z-0177-2008 | 44926 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Capnostream20 Monitor with Nellcor Pulse Oximetry Module--Part Number: CS 08657- Oridion Medical |  |  | July 28, 2007 | 2007 | Oridion Capnography Inc | Handle Cracking - Cracks in the handle will lead to disconnection of the handle from the monitor, causing the monitor to fall. | Process design | Oridion sent Dear Customer letters dated, August 21, 2007, informing users that a new monitor with the improved handle will be shipped by August 31, 2007. |  | 21 | 21 | Nationwide Distribution- including states of AZ, CA, FL, MA, MO, OH, TX, WA, and WI, | Terminated | September 08, 2008 | 408 | 27 | Not_Computer
Z-0265-2008 | 45514 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare Advantage Workstations, Advantage Worksation version 4.2 and 4.3 with Volume Viewer 2 (version 6.4-6.8) provided with Advance Vessel Analysis (AVA) and Advantage Worksation 4.4 with Volume Viewer 2 (version 6.8 or before) or Volume Viewer 3 (version 7.2 -7.3) provided with Advanced Vessel Analysis (AVA), Model 2378698-2,GE Healthcare, Waukesha, WI. 53188 |  |  | September 10, 2007 | 2007 | GE Healthcare | Aortic Length over-estimated: incorrect measurements when using the Aorta analysis protocol or customized protocols. Until a new software version is installed the aorta analysis protocol or the customized protocols for an aneurysm should NOT BE USED. | Software design | An Urgent Safety Notice to customers dated 9/10/07, was issued by the recalling firm informing them of the device defect, safety issues, warnings on future uses and promised new software updates to be installed by GE representatives, free of charge, beginning Nov. 2007. |  | 2518 | 2,518 | Worldwide Distribution - USA including states of AK, AL, AR, AZ, CA, CO, CT, CU, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WY, and Guam, and countries of Canada, Mexico. Argentina, Brazil, Chile, Colombia, Dominican Republic, Ecuador, Australia, China, Honk Kong, Indonesia, Japan, Korea, New Zealand, Singapore, Thailand, Taiwan, United Arab Emirates, Albania, Austria, Azerbaijan, Bosnia, Belgium, Belarus, Switzerland, Czech Republic, Germany, Denmark, Egypt, Spain, Finland, France, Great Britain, Guadeloupe, Greece, Hungary , Ireland, Israel, Italy, Kuwait, Lebanon, Morocco, Netherlands, Norway, French Polynesia, Poland, Portugal, Russian Federation, Saudi Arabia, Sweden, Slovenia, Tunisia, Turkey, and South Africa. | Terminated | December 17, 2011 | 1559 | 2518 | Software | Output/Calculation | Software update | Software Update
Z-0275-2008 | 45326 | No | KXA | Hip joint femoral (hemi-hip) metallic resurfacing prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Durom Hip Resurfacing Systems, Femoral Component 54 Code T; Catalog No.: 01.00211.154, Zimmer, Warsaw, IN, 46581-0708. |  |  | September 26, 2007 | 2007 | Zimmer Inc. | Mislabeled as to size; 54mm package contains 46 mm device and vice versa. | Packaging process control | Zimmer GmbH notified Distributors via a Product Recall Notification letter, dated 9/26/07, instructing them to return the products and to subrecall. |  | 1 | 1 | Nationwide | Terminated | June 04, 2008 | 252 | 6 | Not_Computer
Z-0269-2008 | 45154 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens ECAT ACCEL, Catalog No. 3554479, PET/CT Scanner, Manufactured by Siemens Medical Solutions USA, Inc., 810 Innovation Drive, Knoxville, TN 37932 USA, TEL: 865-218-2534 |  |  | July 23, 2007 | 2007 | Siemens Medical Solutions USA Inc | Incorrect Decay Correction Results - ECAT System's software versions 7.2.2 and 7.4 calculate the decay correction for emission data incorrectly, while performing whole-body studies with all transmissions followed by all emissions. | Software design | Siemans sent an Urgent Device Correction letter, dated June 15, 2007, to all affected consignees informing the user of the problem and requesting that they discontinue the practice of running all transmissions followed by all emissions until a corrective software patch is installed. The corrective software patch is expected to be installed by October 2007. |  | 105 | 105 units | Worldwide Distribution - USA and Countries of Argentina, Austria, Belgium, Brazil, Canada, Switzerland, Chile, China, Czech Republic, Germany, Denmark, Spain, Finland, France, Great Britain, Italy, Japan, South Korea, Netherlands, Norway, Saudi Arabia, Sweden, Slovak Republic, Turkey, and Taiwan. | Terminated | September 11, 2009 | 781 | 446 | Computer
Z-0266-2008 | 45515 | No | HTD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PillingÂż Center-Action Forceps with Insulated Shafts- Serrated Grasping Jaws, Catalog Number: 505015, Non Sterile, Rx Only, (ALLIG GRASPING FCPS 30 CM INSUL), Manufactured for : Teleflex Medical, Research Triangle Park, NC 27709 USA |  |  | October 31, 2007 | 2007 | Telefelx Medical | The specified insulation is not present on the shaft of the devices, leaving a risk of arcing and/or damage to surrounding tissue. | Process control | Consignees were notified by an Urgent Medical Device Recall letter on/about 11/1/2007. They were instructed to immediately cease use and distribution of the affected product and to return them to Teleflex Medical, Durham, NC. In addition, each were instructed to forward the letter if they had further distribution into institutional settings or to homecare patients and retrieve relevant product form those locations. An Acknowledgement &amp; Stock Status Form was enclosed to be completed and faxed back to Teleflex Medical. |  | 73 | 73 | Worldwide: USA including states of AL, AR, AZ, CA, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NY, NV, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI, and WV, and countries of Canada, France New Zealand and Singapore. | Terminated | August 31, 2009 | 670 | 235 | Not_Computer
Z-0384-2008 | 38345 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens, Mevatron M2 Primus, Part No.: 19 40 035, Mid-Energy Medical Charge particle radiation therapy system, with pre-R2.0 Optivue Flat Panel Positioners, Siemens Medical, Concord, CA 94520 |  |  | March 23, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Product's collision detection system may not be sensitive enough to trigger a stop action without a significant amount of force. | Device Design | The firm has issued a Customer Safety Advisory Notice dated 12/6/2006, to all sites that have the pre-R2.0 OPTIVUE Flat Panel Positioners installed. Users will have a new microswitch-based collision detection system, installed free of charge. |  | 9 | 9 units. | Worldwide Distribution- USA including MA, ME, NJ, NY, PA, MD, WV, NC, SC, GA, FL, AL, TN, KY, OH, IN, MI, IA, WI, SD, ND, MT, IL, NE, LA, OK, TX, WY, NV, CA, AK AND AR, and countries of Serbia, Australia, Austria, Belgium, Canada, Denmark, El Salvador, France, Germany, Hungary, India, Ireland, Italy, Japan, Kuwait, Malaysia, Mexico, Netherlands, New Zealand, Norway, China, Philippines, Poland, Portugal, Qatar, Republic of Korea, Russia, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, Turkey, Ukraine, UK. | Terminated | February 11, 2011 | 1421 | 516 | Not_Computer
Z-0323-2008 | 38467 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Integrated Centrifugal System Control Unit; Catalog Number 801046, Terumo Cardiovascular Systems, Ann Arbor, MI 48103 |  |  | September 10, 2007 | 2007 | Terumo Cardiovascular Systems Corp | Failure to operate: The unit display may go blank and reset when the start button is pressed, due to an internal electrical short, resulting in a centrifugal pump that fails to start as expected. | Component change control | Consignees were notified via an Urgent Medical Device Correction letter dated 9/10/07, given instructions to follow if the problem were to manifest itself, and informed the firm would contact them to schedule an upgrade. |  | 120 | 120 | Nationwide. | Terminated | July 07, 2009 | 666 | 120 | Hardware | Device Operation | Replace board | Remove or Replace
Z-0321-2008 | 38466 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System, 1 Roller Pump, 4 inch diameter; Catalog No. 801040, Terumo Cardiovascular Systems, Ann Arbor, MI 48103 |  |  | September 10, 2007 | 2007 | Terumo Cardiovascular Systems Corp | Failure to operate; The pump may fail to operate, and be difficult to hand crank, due to an internal electrical short condition. | Process control | Consignees were notified via an Urgent Medical Device Correction letter dated 9/10/07, which provided them guidance in the event the problem occurs and advised them they would be contacted to schedule replacement of the circuit boards. |  | 16 | 16 | Worldwide Distribution-USA including states of Alabama, Arizona, Florida, Illinois, Indiana, Mississippi, Nebraska, North Carolina, Pennsylvania, and Virginia, and countries of Belgium, Hong Kong and United Arab Emirates. | Terminated | February 23, 2008 | 166 | 184 | Hardware | Device Operation | Replace board | Remove or Replace
Z-0261-2008 | 45402 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | ARJO,Carendo Multi-Purpose Hygiene Chair; Model BIB2003-01, a non-AC powered patient lift; Manufactured by Arjo Hospital Equipment AB, Eslov, Sweden |  |  | October 15, 2007 | 2007 | Arjo, Inc. | Male patient genitals may be pinched between the seat and the frame of the patient lift chair when the Care Raiser seat is returned to its lowest position. | Device Design | Arjo sent Customer Field Correction Notification letters dated 10/15/07 to the end user accounts who received the affected chairs, advising them of the potential for the pinching of male patient genitals between moving parts of the care raiser function of the Carendo Hygiene Chair. They were instructed to review the enclosed revised Operating and Product Care Instructions manual, highlighting the warnings on pages 18 and 20; ensure all users receive refresher training in the chair's proper use and have access to the revised manual; and complete and return the enclosed customer response form, acknowledging the receipt and understanding of the letter. Any questions were directed to the Arjo Quality Department at 800-323-1245, ext. 6140 or 6165. |  | 1342 | 1,342 hygiene chairs | Nationwide | Terminated | July 27, 2009 | 651 | 1342 | Not_Computer
Z-0289-2008 | 44752 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | A1CNOW+ with sampler 10-Test Kits (Product No: 3022), Metrika, Inc., Sunnyvale, CA 94085. (This test provides quantitative measurement of the percent of glycated hemoglobin (%A1C) levels in capillary (fingerstick) or venous whole blood samples. The test is used by professional healthcare providers to monitor long-term glycemic control in people with diabetes. The product is sold as a kit with a disposable monitor and reagents for ten A1C tests and is suitable for use at the point-of-care. |  |  | August 31, 2007 | 2007 | Metrika Inc | Incorrect Test Results: Due to manufacturing errors regarding temperature compensation factors, the device will give erroneously high results if the test is performed below 70 degrees Fahrenheit. | Process control | A September 14, 2007, Urgent Device Recall letter was issued to all Distribution Center Managers to return recalled product for replacement. Customer notification was made via a Dear Healthcare Professional letter dated Sept. 14, 2007, also titled Urgent Medical Device Recall. |  | 4286 | 4,286 | Nationwide including states of AK, AL, CA, CO, CT, DC, FL, GA, HI, IL, IN, KS, KY, MA, MD, MI, MN, MS, NC, NJ, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI and Puerto Rico. | Terminated | March 25, 2008 | 207 | 4286 | Not_Computer
Z-0288-2008 | 45333 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott Clinical Chemistry Aeroset/Architect (c Systems)Bilirubin Calibrator, List Number (LN) 1E66-03 &amp; List Number (LN) 1E66-04, Abbott Laboratories, Abbott Park, IL 60064 |  |  | August 10, 2007 | 2007 | Abbott Laboratories Inc. | Incorrect Calibrator values: changes in a raw material used in the Bilirubin Calibrator value assignment process contributed to variability in the preparation of the standard. | Material/Component Contamination | An August 10, 2007, Product Correction letter was provided to all ARCHITECT cSystem and AEROSET System customers. The letter advises that changes in a raw material used in the Bilirubin Calibrator value assignment process contributed to variability in the preparation of the NIST standard 916a. Revised Bilirubin values for the calibrator lot numbers in this recall are attached to the recall notice. Six necessary actions were listed in the recall notice as: 1) Identify the lot numbers of Clinical Chemistry Bilirubin Calibrator(s) currently in use. 2) Discard the value sheet and replace with the appropriate values provided with the recall letter. 3) Enter the revised Bilirubin values. 4) Calibrate the assays. 5) Evaluate your quality control. 6) Evaluate your reference range and any shifts in patient results. |  | 7033 | 7,033 units total worldwide | Worldwide: USA, Canada, Mexico, Guatemala, Puerto Rico, Germany, Hong Kong, Singapore, Australia, New Zealand, Jamaica, Venezuela, Columbia, Ecuador, Chile, Uruguay, Argentina, Brazil, Thailand, Trinidad &amp; Tobago, Bahamas, Panama, Cayman Islands, and El Salvador | Terminated | April 19, 2012 | 1714 | 7033 | Not_Computer
Z-0278-2008 | 45420 | No | OAT | Endosseous dental implant. | Dental | Dental | 510(k) | Straumann Palatal Implant , 4.8 L 4.2 mm, |  |  | September 28, 2007 | 2007 | Straumann USA LLC | Sticking Transfer Piece: The Transfer Piece which accompanies the Palatal Implant may get stuck during implant requiring removal by using mechanical force. | Device Design | Straumann USA, LLC sent Urgent Device Recall letters, via FedFx on 9/28/07, requesting the return of all Straumann Palatal Implant, articles 042.335 with lot numbers 1001, FE3776, F4662, F6563, E7083, and article 042.336S lot number 1001. |  | 32 | 32 units | Nationwide Distribution - USA including states of AZ, CA, CO, ID, IL, IN, GA, KY, NH , MA, MD, MI, NM, PA, TX, UT, VA, and WA. | Terminated | April 16, 2010 | 931 | 90 | Not_Computer
Z-0122-2008 | 38310 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Target, Transend EX Soft Tip Guide Wire, .014in /205cm, UPN: M001468060, Catalog Number: 46-806, |  |  | June 14, 2007 | 2007 | Boston Scientific Target | Mislabeling: Product is labeled with the wrong expiration date: 2009-12. The correct expiration date is 2008-12. | Process control | Boston Scientific issued a letter to all consignees by mail on June 14, 2007, titled Urgent Medical Device Recall or Correction - Immediate Action Required. The letter describes that the expiration date is incorrect. Discontinue use of the product, segregate it, complete the Reply Verification Tracking Form and fax it to Boston Scientific. Call 888-272-1001 to obtain a Returned Goods Authorization (RGA) number. |  | 40 | 40 | Nationwide. | Terminated | December 21, 2007 | 190 | 40 | Not_Computer
Z-0123-2008 | 44814 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Orthopedics, Vanguard PS Open Box Femoral, 67.5 mm Left-Interlok, For Cemented Use Only, CO-CR-MO Alloy, REF 183130, Knee Prosthesis Component, Biomet Orthopedics, Inc., Warsaw, IN 46581 |  |  | August 31, 2007 | 2007 | Biomet, Inc. | Mislabeled as to size; labeled as 67.5 mm, actually contains 62.5 mm size femoral component. | Packaging process control | Biomet notified consignees via letter dated 8/31/07, titled Urgent Medical Device Recall Notice to immediately discontinue use of the product and to return units to Biomet. |  | 16 | 16 | Nationwide. | Terminated | June 04, 2008 | 278 | 16 | Not_Computer
Z-0385-2008 | 44893 | No | NJL |  |  | Orthopedic | PMA | Depuy P.F.C. Âż E Knee System Stabilized Tibial Insert (STAB), 67 mm M/L, 45 mm A/P, 8.0 mm, 2.5, sterile; Product Code/REF 96-0430. |  |  | September 27, 2007 | 2007 | Depuy Orthopaedics, Inc. | An incomplete seal on the inner pouch allows exposure to oxygen, which may result in oxidation and would affect long-term performance of the implant. | Packaging process control | Consignee hospitals and implanting physicians were notified of the recall via an Urgent Information-Recall Notice letter and Dear Doctor letters dated 9/27/07. The hospitals were instructed to determine whether they had any of the recalled products and to contact their sales representative to pick up the product. Depuy sales representatives were notified to collect the products from consignee hospitals via email on 10/1/07. The decision to discuss this product defect with patients who had already been implanted was left to the implanting physicians. |  | 10 | 10 | Nationwide. | Terminated | April 30, 2008 | 216 | 1160 | Not_Computer
Z-0279-2008 | 45432 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Imager II Angiographic Catheter, (Outer Carton Label): UPN M001315081, 5 Fr, 100 cm, Catalog No. 31-508, Lot 48951, Description 5 Fr Flush TR, Sterile EO, Made in Ireland, Unit 7, 8 &amp; 9, Annacotty Business Park, Annacotty, C. Limerick, Ireland ****Inner Pouch Label: Boston Scientific, Imager II Angiographic Catheter,UPN M001315080, 5 Fr, 100 cm, Catalog No. 31-508, Lot 48951, Description 5 Fr Flush TR, Sterile EO, Made in Ireland, Unit 7, 8 &amp; 9, Annacotty Business Park, Annacotty, C. Limerick, Ireland |  |  | October 12, 2007 | 2007 | Boston Scientific Corporation | Mislabeled: Units within the identified lot/batch were incorrectly labeled as 5F on the outer carton and pouch when the product was actually a 4F device. | Packaging | Boston Scientific sent an Urgent Medical Device Recall -- Immediate Action Required letter to consignees on 10/12/07. The letter described the product and the mislabeling problem (labeled outer carton and pouch as 4F when they are actually 5F devices). The firm recommended segregating the affected product and to return it to Boston Scientific PLEASE NOTE: The firm's recall letter incorrectly addressed the mislabeling of this product recall. The mislabeling involves the labeled outer carton and pouch were labeled as 5F when the device enclosed in the pouch was a 4F device. |  | 70 | 14 boxes each containing 5 Units | USA only including states of NJ, CA,TX,AZ | Terminated | December 19, 2007 | 68 | 70 | Not_Computer
Z-0328-2008 | 38471 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific ACCUSTICK II Introducer System with .038 inch Stainless Steel J Tip and .018 inch Nitinol Guidewire, sterile, REF/Catalog No. 20-710; UPN/Material Number: M001207100, Boston Scientific Corporation, Natick, MA 01760 |  |  | July 19, 2007 | 2007 | Boston Scientific Corp | Marker Detachment: Out of specifications for pull strength, thus the marker may detach from the introducer during use. | Process control | Consignees were notified via an Urgent Medical Device Recall Letter sent federal express on 7/19/07, which requested them to check their stocks, remove product from inventory, and return the product. |  | 249 | 249 (199 U.S.A. plus 50 international) | Worldwide: USA and The Netherlands. | Terminated | February 17, 2008 | 213 | 249 | Not_Computer
Z-0330-2008 | 44959 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper FD20/10, Release 1.2, Philips Medical Systems, Bothell, WA 98021 |  |  | September 24, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | Inaccurate anatomical measurements when using the optional SmartMask feature | Device Design | The recalling firm sent customers an Urgent Device Correction notification letter dated 09/21/07 via certified mail - return receipt requested. This notification provided a short term work around for the device defect. For a permanent fix, the recalling firm is developing software update to check for mismatch at the start of a subtracted fluoroscopy run, and will contact customers to schedule an install. Installation is expected to begin 10/2007 and be completed 06/2008. |  | 53 | 53 units | Nationwide | Terminated | May 03, 2012 | 1683 | 521 | Computer
Z-0331-2008 | 44963 | [6 empty] | Axya Bone Anchor , 3.0 mm Titanium with Nylon Suture |  |  | September 27, 2007 | 2007 | Axya Medical | Sterility Compromised (Package Integrity): Sterility of the device may be compromised due to holes in the sterile packaging | Package design/selection | Axya Medical notified customers by letter, Urgent-Product Recall, dated 9/25/07 , via UPS . Users are requested to return product and complete the reply form. |  | 1008 | 1,008 units | Nationwide | Terminated | June 21, 2012 | 1729 | 17956 | Not_Computer
Z-0358-2008 | 45427 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Haemonetics cardioPAT System, Product List Number 02050, Haemonetics Corporation, Braintree, MA 02184,(Cardiovascular Perioperative Autotransfusion) |  |  | October 15, 2007 | 2007 | Haemonetics Corporation | Labeling (additional) use instructions: to clarify proper position of cardioPAT device and post-op line to prevent fluid build up from patient's chest | Labeling design | Haemonetics issued a Labeling suppplement to users via a Field Notification letter on October 15, 2007. Users are provide additional Directions for use for the Operator's Guide and Quick Reference Guide. |  | 106 | 106 units | Nationwide: USA including states of CA, CT, FL, IN, KY, MI, NC,NE, NH, PA, RI, SC,TX, and WA | Terminated | March 19, 2012 | 1617 | 106 | Not_Computer
Z-0361-2008 | 45480 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VKS PE Tibia Insert Ultra-Congruent; Left, Size 2, 12mm; Model Number: 0214612; Plus Orthopedics, San Diego, CA 92121 |  |  | November 13, 2006 | 2006 | Plus Orthopedics USA | Labeling Error: The VKS PE Ultra-congruent insert package, may actually contain a Standard PE insert. | Employee error | Consignees were notified by telephone on 11/13/2006 to check their inventory and return any affected units. On 11/28/06 Physician Notifications and Patient Assessment forms were sent to the two physicians who had implanted a total of 4 devices. |  | 10 | 10 | Nationwide: USA including states of ID, NY, VA &amp; WI | Terminated | December 13, 2007 | 395 | 10 | Not_Computer
Z-0360-2008 | 45475 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Plus Orthopedics Gliding Nail 17mm proximal / 11mm distal, 125 degrees, 180mm; Model Number: 0132073, Plus Orthopedics USA, San Diego, CA 92121 |  |  | June 30, 2003 | 2003 | Plus Orthopedics USA | The device was produced without the required distal locking hole | Reprocessing Controls | The initial depth of the withdrawal action taken in June, 2003 ,was to the distributor (agent) level. Plus USA contacted individual agent sites showing inventory of the affected article numbers on hand and requested that those items be returned to Plus USA. |  | 3 | 3 | Nationwide: AZ and ID | Terminated | December 13, 2007 | 1627 | 3 | Not_Computer
Z-0359-2008 | 45465 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST Multi-Use Syringe Kit, Model A2000; SKU # : 014612; labeled as STERILE; Manufactured: Acist Medical Systems, Inc., 7450 Flying Cloud Drive, Suite 150, Eden Prairie, MN 55344, USA. Catalogue No: 800061-003. (The multiple use A20000 syringe kit is comprised of a terminally sterilized syringe barrel and a contrast spike assembly) |  |  | October 03, 2007 | 2007 | Acist Medical Systems | Syringe Fracture: the syringe may be fractured, causing contrast to be sprayed from syringe fracture. | Process design | The Urgent Medical Device Recall letter was sent by certified mail to all U.S. base customers on 10/3,4/07. Customers were advised to stop distribution of the product and immediately cease use of the product. Recalled product is to be returned using the FedEx Airbill that was enclosed with the recall letter. Customers were to complete the enclosed "Device Recall Response" form even if they did not have any impacted kits in inventory. |  | 177050 | 177,050 Kits | Nationwide Distribution - including states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Terminated | October 06, 2008 | 369 | 177050 | Not_Computer
Z-0357-2008 | 45445 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker inTouch Zoom Critical Care bed, Model 2140, Stryker Medical, Portage, MI 49002 |  |  | November 02, 2007 | 2007 | Stryker Medical Div. of Stryker Corporation | Multiple defects: 1) beds exhibit reverse motions when the Fowler, gatch or foot section is activated 2) a battery charging defect which may result in a depletion of battery power 3) the touch screen on the footboard freezing, resulting in controls to powered, pressure-relief mattresses being difficult to access 4) beds equipped with zoom drive systems may have an overdrive problem which could re | Process control | Consignees were notified via an Urgent Medical Device Correction letter dated 11/2/07, and instructed to remove the beds from service if they display any unintended motions and informed that one of the firm's field service representatives will visit the facility to make corrections to all beds under recall. |  | 329 | 329 units total for both models. | Worldwide-USA and Canada | Terminated | October 27, 2010 | 1090 | 329 | Not_Computer
Z-0362-2008 | 45484 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VKS Tibial Trial, Right/Left, Prep Guide; Model Number: 1214331, 1214332, 1214333, 1214334, 1214335: ( Model Number: 1214331 VKS Tibial Prep Guide, Right/Left, Size 1, Lot Code EM. 1 and EM.6; 1214332 VKS Tibial Prep Guide, Right/Left, Size 2 Lot Code EM.l and EM.6, 1214333 VKS Tibial Prep Guide, Right/Left, Size 3 Lot Code EM.1; 1214334 VKS Tibial Prep Guide, Right/Left, Size 4 Lot Code EM.5; 1214335 YKS Tibial Prep Guide, Right/Left, Size 5 Lot Code EM), Plus Orthopedics |  |  | November 27, 2006 | 2006 | Plus Orthopedics USA | The VKS "Tibial Trial" was found to have an incorrect fixation pin hole diameter. | Process control | The initial depth of the withdrawal action taken November 2006 was to the distributor (agent) level. On 11/27/2006 Plus Orthopedics USA contacted individual agent sites showing inventory of the affected article numbers on hand and requested that those items be returned to Plus USA. Agents were notified by telephone to check their inventory and return any affected units. A telephone script was used. |  | 19 | 19 | Nationwide to : AZ, CA, CO, FL, ID, NE, NY, SC, TX, UT, VA, and WA | Terminated | December 13, 2007 | 381 | 19 | Not_Computer
Z-0375-2008 | 44972 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Point of Care, i-STAT G Cartridge; (Glucose/Glu); List No: 06F09-01; Part No: 120100; Abbott Point of Care, Inc., East Windsor, NJ 08520, USA; Product of Canada. |  |  | October 18, 2007 | 2007 | Abbott Point of Care Inc. | Difficult to Close -- Certain lots of i-STAT cartridges are difficult to close, will not close or do not stay closed. | Packaging change control | Abbott Point of Care sent and Urgent Notice to all i-STAT Customers (US and Canadian consignees) notifying them that certain lots of i-STAT cartridges have snap closures that are difficult to close, will not close or do not stay closed. The customers were asked to review the lot numbers on the enclosure and, if the facility has any cartridges from the lot numbers listed, they should remove from service immediately. |  |  |  | Worldwide Distribution- including USA and country of Canada. | Terminated | July 18, 2008 | 274 |  | Not_Computer
Z-0441-2008 | 45456 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Excia Cemented Hip Stems, Size 14mm, Part Number : NJ314K, Aesculap Implant Systems, Inc., Center Valley, PA 18034 |  |  | October 10, 2007 | 2007 | Aesculap Implant Systems | Sterility (package integrity) Compromised: Device package damaged during shipment | Package design/selection | The recalling firm issued an Important Recall Notice letter dated 10/10/07 to all their customers informing them of the problem and the need to return the product. |  | 167 | 167 units | Nationwide-USA including states of AZ, GA, MD, OK, PA, TX, and VA (distributors) medical facilities in FL, NY, and PA; and sales representatives in CA and FL. | Terminated | February 07, 2008 | 120 | 167 | Not_Computer
Z-0363-2008 | 45523 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Ascent POCTS (Posterior Occipital Cervical Thoracic System) Part Numbers: 65-33 XX Series (65-3310,-3312, -3314, -3318, -3320-, -3322, -3324, -3326, -3328, and -3330), Multi-Axial Screw, 4.0mm x 26mm;Implant grade Titanium Alloy; Blacksone Medical Inc., 90 Brookdale Drive, Springfield, MA 01104. |  |  | May 15, 2007 | 2007 | Blackstone Medical, Inc. | Manufacturing Defect: The saddle of the screw's body relative to the keyway exceeds the tolerance outlined in the device specifications. The extent to which the tolerance is exceeded varies from part to part, in the extreme the tolerance is exceeded by (+) 50%. As a consequence, there is a potential for set-screw cross-threading during implantation. Cross-threading could in turn lead to a comprom | Nonconforming Material/Component | Blackstone Medical, Inc., field representatives were notified via phone, e-mail or fax on 4/30/07. Attention: Notification of Product Removal Letters were sent on May 15, 2007, to distributors and implanting surgeons requesting return of non-implanted stock. |  | 319 | 319 screws ( 212 domestic, 107 international) | Worldwide: USA and countries of China, Cyprus, Turkey and Switzerland. | Terminated | June 06, 2008 | 388 | 319 | Not_Computer
Z-0292-2008 | 45817 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Medrad Avanta Multi-Patient Disposable Hand Controller, Angiographic Injector and Syringe. Catalog number AVA 500 HC, For Use with MEDRAD Avant Fluid Management System, MEDRAD, Inc., Indianola, PA 15051 |  |  | October 25, 2007 | 2007 | Medrad Inc | Sterility (package integrity) compromised: Incomplete package seal; sterility may be compromised | Packaging process control | The recalling firm issued an Urgent Medical Device Field Correction Letter dated 10/25/07 via Federal Express to their customers informing them of the problem. Customers can either return the product or consistent with the indicated use and established pratice of sheathing, utilize sterile sheaths with the hand controllers for up to five patients. Each letter is being accompanied by a box of 25 sheaths. |  | 25977 | 25, 977 units | Worldwide-USA and countries of Germany, Israel, Finalnd, Italy, Netherland, Austria, France, Japan, Denmark, Ireland, South Africia, Australia, Saudi Arabia, Singapore, Canada, New Zealand, Spain, Sweden, Portugal, Korea, Great Britain, Turkey and The United Arab Emerites. | Terminated | April 22, 2008 | 180 | 25977 | Not_Computer
Z-0449-2008 | 45540 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon MIS Femoral Adjustment Block;Catalog Number: 6541-5-601;Stryker Orthopaedics, Mahwah, New Jersey |  |  | September 19, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Assembly problems: MIS instruments (Triathlon MIS AP Sizer Body - Left; Triathlon MIS Sizer Body - Right; Triathlon MIS Femoral Adjustment Block and Triathlon MIS Modular Capture) may not assemble properly. | Device Design | Important Product Correction letters and Product Accountability Forms were sent via Federal Express on 9/19/06 with return receipt. Foreign locations were notified by a Advisory Notice. Notification letters requested that all inventory of the affected items be examined to verify that it assemble correctly. Any product that does not assemble correctly was to be returned to Stryker, Mahwah, NJ. The Product Accountability form was to be completed by each location to notify Stryker that the letter ahd been received and inventory examined. Stryker representatives were to be contacted for product replacement. |  | 2868 | 2,868 units total for Recalls Z-0448-0450-2008 | Worldwide: USA including states of MI and IL. and countries of Australia, China, New Zealand, Singapore, France, Germany, Italy, Portugal, Spain, Sweden and UK. | Terminated | January 18, 2008 | 486 | 2868 | Not_Computer
Z-0054-2008 | 45525 | No | HIS | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Lubricated Latex Condom. Labeled in part: "COMING TOGETHER IN DC", Distributed by Global Protection Corp., #12 Channel Street, Boston, MA |  |  | September 11, 2007 | 2007 | Global Protection Corp. | Condoms fail the water leak test and package integrity | Process control | Global Protection Corp. recalled product by email on 9/11/07 |  | 12700 | 127,00 condoms | The District of Columbia Health Department | Terminated | March 18, 2008 | 189 | 12700 | Not_Computer
Z-0451-2008 | 45843 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Medical Systems, M3290A IntelliVue Information Center Software for M3170 Intellivue Patient Link, Catalog Number: 865007 and 865015; Running Software Versions J.00.23, J.00.24, J.00.25, and J.00.26, Manufactured and Distributed by; Philips Medical Systems, Andover, MA 01810 |  |  | November 13, 2007 | 2007 | Philips Medical Systems | Alarm failure : If changes are made to the Caregroup Editor configuration, Care Group alarm status, including alarm text and prompt tones, may fail to be displayed or annunciated at bedside monitors within the Care Group | Software change control | Philips Medical notified customers by letter dated November 2007 titled : Urgent Medical Device Correction Notice. Customers will be instructed to follow the Procedure to Mitigate Risk section * of the notice while they await the correction of their devices with a mandatory software upgrade to be performed by service representatives. * (no changes are to be made to the Caregroup Editor configuration on any system. If changes have already been made to the configuration prior to receipt of this notification, turn each bedside monitor "off " and then "on" in order to reset the system-Customer Care Contact telephone 1-800-722-9377) |  | 60 | 60 units | Worldwide Distribution -USA including states of CA, MA, TX, IN, MI, FL, MN, NC, MS, HI, RI, NH, and PA, and countries of Canada, Australia, Belgium, France, Germany, Ireland, Italy, Japan, Switzerland, United Arab Emirates, and The United Kingdom. | Open, Classified |  |  | 60 | Software | Alarm/Message | Software update | Software Update
Z-0452-2008 | 45867 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Conmed Accessory Electrode, Universal Electrodes, Part/REF Numbers : 138100 through 138-105, 138107, 138110, 138112 and 138113, and Ultra Clean Accessory Electrode, Part/REF Numbers: 139100, 139107, 139108, 139321, 139112, 139104 EXT. 139105EXT, 139110EXT, and 139112EXT, Sterile, Rx Only, Conmed Corporation, Utica, New York 13502. |  |  | November 02, 2007 | 2007 | ConMed Electrosurgery | Sterility (package integrity) Compromised: packages of the electrosurgical blades and needles have an inadequate packaging seal | Process design | Consignees were notified by Urgent Voluntary Device Recall letter on 11/02/2007, from ConMed Corporation. They were asked to return affected product for replacement. |  | 11319 | 11,319 cases (476,160 units) | Worldwide: USA and countries of Australia, Canada, China, Belgium, Chile, Denmark, Ecuador, France, Greece, India, Israel, Italy, Japan, Korea, Mexico, New Zealand, Norway, Philippines, Portugal, Saudi Arabia, South Africa, Spain, Thailand, Trinidad, and Uruguay. | Terminated | March 31, 2010 | 880 | 11319 | Not_Computer
Z-0454-2008 | 45887 | No | HNM | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD Visitec Nucleus Hydrodissector, .40 x 22 mm, (27 G x 7/8 in), Catalog #/REF #: 585793, Becton, Dickinson and Company, Waltham, MA 02452 |  |  | November 20, 2007 | 2007 | Becton Dickinson and Company | No Fluid Flow or Restricted Fluid Flow: Nucleus Hydrodissector(s) were found to have reduced or no fluid flow | Process control | The recalling firm notified consignees of the recall through a Voluntary Medical Device Recall letter and phone call on 11/20/2007. Instructions for product return were provided and a telephone number for questions was given (Customer Service 1-866-906-8080 Option 2) |  | 43000 | 43,000 | Nationwide-USA including the states of CA, CO, CT, FL, GA, HI, IL, IN, KS, LA, MA, MD, MO, NC, NJ, NY, OH, PA, SC, TN, VA, and WA | Terminated | February 29, 2012 | 1562 | 83450 | Not_Computer
Z-0364-2008 | 38346 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Accuray, Multiplan Treatment Planning Software, Versions 2.0, 2.0.1 .6, 1.6.1, 1.6..1 or 1.6.3., Accuray, Sunnyvale, CA 94089. |  |  | March 21, 2007 | 2007 | Accuray Inc | Software problem - Sagittal, oblique or coronal sliced MR Images imported from a non-Accuray manufactured device into the product may result in inverted image orientation. | Software design | The firm issued an Urgent Advisory Notification on March 20, 2007, notifying users to only import axial sliced MR image series for treatment planning. Users should not import sagittal, oblique or coronal images series until further notice. The firm intends to issue a software patch to correct the problem and will contact user regarding resolution. Users can contact Accuray Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 (non-USA) or customersupport@accuray.com. |  | 25 | 25 units | Worldwide Distribution - USA including states of MD, CA, FL, WA, LA, IL, IN, MI, MN, NC, NJ, OK, WI, and OH, and countries of Korea, Greece, Germany, France, and Turkey. | Terminated | December 26, 2007 | 280 | 25 | Software | Display/Image | Software update | Software Update
Z-0459-2008 | 38897 | No | JJF | Micro chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche Diagnostics LightCycler Capillaries (100uL) for use in conjunction with the LightCycler analyzers; Catalog #: 03337090001, plus any capillaries received at the time of installation of the LightCycler instruments or systems **, Roche Diagnostics Corporation, Indianapolis 46250.(**LightCycler 2.0 System, Catalog Number: 04391039001; Batch Number Range: 1415471-1415933; LightCycler 2.0 Instrument-Laptop; Catalog Number: 03603946001; Batch Number Range; 1415434-1415734; LightCycler 2.0 Instrument-Desktop; Catalog Number: 03531414201; Batch Number Range: 1415271-1415732 |  |  | August 30, 2007 | 2007 | Roche Diagnostics Corp. | False Negative Results: Air bubbles in the lens or the area surrounding the lens of the LightCycler 100 uL capillaries may cause false negative results. The affected capillaries showed up to a 40% reduction in signal | Process design | Consignees were notified by Urgent Medical Device Recall letter dated 8/30/07, to discontinue use of the 100uL capillaries, plus all of those received at installation, and to disposed of them per local guidelines. |  | 11424 | 119 boxes of 96 capillaries each | Nationwide | Terminated | September 08, 2008 | 375 | 11424 | Not_Computer
Z-0074-2008 | 44787 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Medtronic Sofamor Danek, TSRH SiLo 5.5 Spinal System, Sagittal Adjusting Screw Implants, Supplement Sets, REF: 8889245, Pedicle Screw Fixation System, Medtronic Sofamor Danek, Memphis, TN 38131 |  |  | August 24, 2007 | 2007 | Medtronic Sofamor Danek USA Inc | Disassembly: The movable cradle component could detach from the saddle under extreme manipulation of the associated rods prior to final tightening of the set screws. | Device Design | Medtronic sent letters, titled URGENT DEVICE RECALL, on 08/24/2007, to the Risk Managers of the hospital consignees explaining the potential for these sets to disassemble, and requesting the discontinuation of the use of and quarantine of the product. The product will be retrieved by the firm and replaced. |  | 20 | 20 units | Worldwide Distribution - USA, The Netherlands, Australia, and Canada. | Terminated | August 21, 2009 | 728 | 20 | Not_Computer
Z-0241-2008 | 38307 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | OSI, Allegro 6800 Imaging Table (Operating), single pedestal AC or battery operated table, Orthopedic Systems, Inc., Union City, CA |  |  | February 28, 2007 | 2007 | Mizuho Orthopedic Systems Inc | Failure to anchor; due to a failure in the locking mechanism, the product may not anchor properly. | Software design | The firm issued a Product Correction Notification letter by mail on 02/28/2007 and follow up with phone calls. The letter notified users that a required no-cost upgrade of the removal and replacement of a microprocessor that controls the logic for the floor locking mechanism will be conducted by the recalling firm. |  | 117 | 117 units | Worldwide: USA product was distributed to 16 wholesale dealers, 212 medical facilities, 14 distributors throughout the US in the following states: AL, AZ, CA, FL, GA, HI, IL, IN, KS, KY, MA, MD, ME, MI, MO, NC, NE, NH, NJ, NM, NY, OH, OR, PA, SC, TX, VA, WA, WI, WV and WY and to the following countries Taiwan, Australia, German, Japan, Canada, Saudi Arabia, and The UK. | Terminated | December 18, 2008 | 659 | 343 | Not_Computer
Z-0458-2008 | 45483 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Hypoguard Advance Micro-draw Blood Glucose Monitoring System, Product Number : 411100, For In Vitro Diagnostic Use, Manufactured for ARKRAY USA, Inc., Minneapolis, MN 55439 USA. (Package Contents -Advance Micro-draw Blood Glucose Meter, 10 test Strips, Codeing chip and Control Solution (3mL), Lancing device, 10 Lancets, 3V Lithium Battery, Logbook, User Instruction Manual, Quick Reference Guide) |  |  | October 18, 2007 | 2007 | ARKRAY USA INC. | Mis-labeled/coded strips : Meter kits from this lot contain mislabeled 10-count test strip vials. The code number printed on the Advance Micro-draw 10-count test strip bottle is incorrect. The code number on the vial is 867 when it should be 687. The 687-labeled code chip contains the correct calibration information for the "867" test strips, so the meter will function properly and give accurat | Process control | Distributors were sent a letter on 10/18/07 asking that product be quarantined. On 10/26/07, an Urgent Voluntary Medical Device Recall letter was sent asking that they return affected product and contact their customers to the retail level. |  | 2005 | 2,005 Kits | Nationwide: USA including states of CA, MI, MO, MS, NE, OH and TX. | Terminated | March 02, 2008 | 136 | 2005 | Not_Computer
Z-0463-2008 | 45334 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Hipstar V40 Femoral Stem-127 degrees; Catalog No: 7850-5-001, N degrees 3 high offset 31 mm;Sterile; IMPLANT SIZE 1, Howmedica,Benoist Girard dt Cie, 203 bd de la Grande Della B.P.8;14291 Herouville Saint Clair |  |  | September 28, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Mis-labeled: The warning label "Do not use with heads more than +10mm offset" does not appear on 127 degree stem boxes. | Packaging process control | Urgent Product Correction/Recall notification letters were sent by Fed Ex return receipt on 9/28/07, requesting quarantine for Styker representatives to return or re-label. |  | 59 | 59 | Nationwide- USA including states of OR, AZ, NV, MI, AL, MI, FL, MN, and CA | Terminated | October 27, 2008 | 395 | 910 | Not_Computer
Z-0475-2008 | 45982 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Hipstar TMZF Cementless Hip Stem Protocol, Product Number: LSP60. |  |  | September 28, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Mis-labeling: An error on page 4 of the Hipstar protocol was found. In the table for compatible Heads, the 40 mm diameter heads with a Neck Length of +12 should have been omitted. | Other | Urgent Product Correction Letters, dated 10/17/2007 were sent on 10/17/07by Federal Express to all Stryker Branches/Agencies who had received the protocol. The branches were instructed to quarantine and destroy the protocols. |  | 192 | 192 protocols | Nationwide:The protocol was distributed to 9 Stryker branches/agencies in NJ, MD, IN, MI, MN, IL and TX. | Terminated | September 05, 2008 | 343 | 192 | Not_Computer
Z-0132-2008 | 38381 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VNUS ClosureFAST catheters, Model Number: CL7-7-100, VNUS Medical Technologies, Inc., San Jose, CA 95138 |  |  | March 06, 2007 | 2007 | VNUS Medical Technologies Inc | Sterility (package integrity) Compromised: Product sterility may be compromised due to a pinhole under the pouch foldover. | Package design/selection | The recalling firm has issued customer notification letters by mail on 03/06/2007. The firm is requesting that consignees discontinue their use of the device and return recalled product . The firm will provide replacements. |  | 92 | 92 units | Nationwide: Product was distributed to 18 health care providers and medical centers throughout the US in the following states: TX, UT, CA, CT, MI, NY, MI, OR, WA, VA, PA, and NJ. | Terminated | December 17, 2007 | 286 | 92 | Not_Computer
Z-0317-2008 | 38468 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Minimally Invasive Solutions Technology Offset Rasp Handle, Right, 45 degrees, non-sterile; REF 7712-35-02, Zimmer U.K. Ltd., SN3, 4FP, UK |  |  | June 28, 2007 | 2007 | Zimmer Inc. | Clip Detachment: The retainer c-clip can detach from the instrument during use and may fall into the surgical wound | Process control | Zimmer notified their distributors via Product Recall Notification Letter dated 6/30/07 and instructed them to notify their customers of the recall and to pick up the product and return it to Zimmer. |  | 62 | 62 (total for Z-0316-0317-2008) | Worldwide: USA including states of Arizona, California, Florida, Georgia, Indiana, Michigan, New York, North Carolina, and Virginia, and countries of Germany and Japan. | Terminated | March 17, 2008 | 263 | 62 | Not_Computer
Z-0138-2008 | 44953 | No | MJO |  |  | Orthopedic | PMA | Medtronic Sofamor Danek, Prestige Cervical Disc System, Implant Trial, REF/Catalog #: 6972226, SIZE 6 X 12, MATL: Stainless Steel/Titanium, Rx only, Manufactured at: Bartlett, Tennessee USA |  |  | September 04, 2007 | 2007 | Medtronic Sofamor Danek USA Inc | Implant mis-seating; Variance in size between the trial and the implant could cause the implant to be improperly seated. | Device Design | The firm notified their consignees of the problem and the recall by Urgent Medical Device Recall letter, telephone calls and electronic mail on 09/04/2007. Instructions include requests to immediately quarantine the product and hold for recalling firm collection. |  | 196 | 196 units | Nationwide | Terminated | November 03, 2008 | 426 | 196 | Not_Computer
Z-0136-2008 | 44883 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | DRL System Trays comprised of DRLock Bone Fixation Plates, Pegs, and Screw Systems: Four Plate Styles include: Part Numbers: DRV-002-0L Fixed Short Left Plate, DRV-002-0R , Fixed Short Right Plate, DRV-002-1L, Fixed Standard Left Plate, and DRV-002-1R, Fixed Standard Right Plate; Distal Screw Part Numbers: DRV-021-24-10, -12, -14, -16, -18, -20, -22 and -24; Distal Peg Part Number: DRV-121-20-10, -12, -14, -16, -18, -20, -22, and -24; All screw caddies and any distal screw and peg inventory not yet added to screw caddies, Ortho Helix Surgical Designs, Inc., Akron, OH |  |  | September 04, 2007 | 2007 | OrthoHelix Surgical Designs Inc | Seat/Fit problems: The heads of the distal screws/pegs rest above the surface of the distal plate after the screws/pegs are fully tightened. | Component change control | On 9/4/2007, OrthoHelix mailed an Important Recall Notification letters via certified mail to all direct accounts. The notification instructed all accounts to examine their inventory and return all recalled product, identified within the notification, to OrthoHelix for replacement. |  | 13884 | 267 plates/52 DRL System Trays | Nationwide; including states of IL, CA, VA, TX, MD, UT, GA, CO, NJ, OH, &amp; WI. | Terminated | May 20, 2011 | 1354 | 13884 | Not_Computer
Z-0185-2008 | 44800 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | EncoreÂż ReverseÂż Shoulder Prosthesis (RSP) System, Standard Size 40 RSP Humeral Socket Insert; Catalog #: 508-00-040; Sterile; Manufactured by Encore, Medical, L.P., Austin, TX 78758. |  |  | August 08, 2007 | 2007 | Encore Medical, Lp | Devices Mislabeled; one lot of Size 36 Standard RSP Socket Inserts were labeled as Size 40 Standard RSP Sockets and vice versa. | Other | Encore notified their sales reps (direct accounts) via telephone and overnight letter on 08/08/07. Encore sales reps were asked to return all outstanding devices, for which they would be given replacements. Notification of implanting surgeons was done by phone and letter (Urgent : Device Recall) between 08/07 and 09/12/07. Implanting surgeons notified of recall and asked to closely monitor patients who had received and already been implanted with the affected lots. If explant was required decision of the surgeons, they were asked to notify the recalling firm. |  | 25 | 25 units. | Nationwide:Direct accounts and surgeons in the following states: CA, GA, IA, MA, MD, MI, MO, SC, TX and UT. | Terminated | March 10, 2008 | 215 | 49 | Not_Computer
Z-0145-2008 | 44958 | No | MDK | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Accuprobe Group B Streptococcus Culture Identification Test, Catalog Number 2820B, Batch Number 531379, Gen-Probe, Incorporated, San Diego, CA. 92121 |  |  | May 02, 2007 | 2007 | Gen-Probe Inc | Mispackaging: Kits may contain Haemophilus Influenza probe pouches in addition to Group B Streptococcus probe pouches | Other | Gen-Probe notified customers by phone on May 2nd and 3rd, 2007; a recall notice letter was faxed to each customer on May 3rd, 2007 and mailed on May 4th, 2007 via certified mail. The letter describes the packaging error and the remote health hazard risk, and provides instructions for users to return all kits of this lot number. |  | 106 | 106 kits | Nationwide: including states of AZ, CA, CO, FL, KY, MD, MO, NC, NJ, TX, and UT | Terminated | May 03, 2010 | 1097 | 106 | Not_Computer
Z-0137-2008 | 44928 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Freedom Designs TriPod w/Anti tipper Manual Wheelchair, Models Numbers: TriPod, Tripod LT and TriPod Pediatric LT with Trailing Anti-Tipper-Caster-Axle Link, Freedom Designs Incorporated, Simi Valley, CA 93065 |  |  | May 14, 2007 | 2007 | Freedom Designs Inc | Tip-over hazard; the Anti tipper may fail/break and the wheel chair could tip over. | Process control | Urgent Field Correction Notification letters were sent to customers on May 15, 2007 via US Postal return receipt requested. Consignees were notified via delivery of a field correction package consisting of:(1) A letter to the dealer (2) FAQ sheet clarifying their responsibility (3) A corrective/preventive action tracking sheet specific for the consignees affected units(4) Replacements of anti-tipper assembly for each unit shipped with an instruction sheet.A Notice was also posted on the Freedom Designs website.Correction will be made to the consumer level |  | 41 | 41 units | Nationwide: including states of AL, CA, FL, CO, GA, KS, LA, MI, MS, NC, OH, TN, TX, and VA | Terminated | July 14, 2008 | 427 | 41 | Not_Computer
Z-0304-2008 | 44913 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | c.cam Gamma Camera, Model No: 9VIR1200; Tomography System (Emission Computed); Siemens Medical Solutions USA, Inc., Molecular Imaging, 2501 N. Barrington Road, Hoffman Estates, IL 60195-5203; Made in Denmark |  |  | September 17, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Unintended Motion: A hardware fault in the c.cam system can cause the system to experience unintended motion, thereby causing a collision hazard. | Software design | Urgent Field Correction Recall letters dated 9/11/07, on Siemens letter head were sent to the end users via U.S. certified mail, informing them of the potential risk of unintended motion of the detector head resulting in a collision hazard. The end users were warned not to press any of the detection motion buttons while simultaneously holding the collision over ride button on the c.cam hand controller, and were provided steps to take if the system stops during an acquisition and displays the Collision Error message. They were told that a Siemens service representative will contact them within 90 days to schedule the repair of the c.cam at no cost to the end user. Any questions regarding the Important Safety Notice were directed to their service representative or Siemens office at: 1-800-888-7436 (USA). |  | 486 | 486 units | Worldwide: USA including Puerto Rico and Brazil, Canada, France, Germany, Great Britain, Israel, The Netherlands, and Portugal. | Terminated | January 09, 2009 | 480 | 486 | Hardware | Physical Safety Hazards | Repair | Repair
Z-0287-2008 | 45499 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Ellex Integre Duo, Model No: LP1RG - Ellex Medical PTY, LTD 82 Gilbert Street, Adelaide, S.A. 5000, AUSTRALIA. (The Integre Duo is a dual-wavelength surgical photocagulator laser intended for use in a clinic, outpatient clinic, or a retinal specialist's office. ) |  |  | June 13, 2007 | 2007 | Ellex USA | Electric Shock from Power Supply: The power supply has a defect which can cause an electrical shock to user. | Device Design | Ellex, Inc. called each customer on 6/13/07, about the problem (defect which can cause electric shock) and advised that they would be performing modification of the power supply in the laser system and scheduling a field service visit to complete the work. Upon completion of the modification, the service technician completed the service record and left a copy with each client and informed them of the completed work. A copy of each completed service record was sent to Ellex Medical Pty. Ltd. On 9/27/07, Ellex sent a letter to all consignees, informing them that all modifications by service technicians' have been completed, and that their Solitaire LP4532 laser and/or Intregre Duo LP1RG is up to the manufacturers specifications and is in good operating order. |  | 1 | 1 | Worldwide Distribution - USA including states of WI, PA, CO, CA, NC, NY, SC, MD, and MS, and countries of Canada and Brazil. | Terminated | October 17, 2008 | 492 | 11 | Battery | Physical Safety Hazards | Replace | Remove or Replace
Z-0204-2008 | 44967 | No | NVR |  |  | Orthopedic | 510(k) | Medtronic Sofamor Danek, Satellite System Sphere PEEK, Size 12 mm, Material PEEK-Optima, quantity 1, Sterile, Rx only, REF 9000212, Manufactured at: Warsaw, Indiana, USA |  |  | September 13, 2007 | 2007 | Medtronic Sofamor Danek USA Inc | Labeling, including indications for use, lack 510 (K) approval | Labeling design | On September 12, 2007, the recalling firm's representatives were instructed by electronic mail to remove any manuals in the field and contact surgeons who have evaluated, implanted or been trained to implant the Satellite System and to collect any guides they may have on site. The sales representatives that have accounts where the Satellite product was shipped directly were mailed the information and instructed to hand deliver the 9/13/07, Urgent Device Recall notices to each consignees risk management office. The recall was then expanded on 10/31/2007, to include the surgical manual, the implants and associated instruments. The delivery of the second Urgent Device Recall Notices, dated 10/25/07, were handled in three ways: 1) The letter was delivered to all initial consignees that received satellite products along with a questionnaire (response card) to be completed by the Risk Manager or equivalent administrative function., 2) The letter without the questionnaire was sent directly to target surgeons who are known to have implanted or have been trained to implant Satellite Product 3) a field action confirmation form was sent to Sales Representative with surgeon customers that are receiving the recall communication. The form will be completed by the Sales Representative to confirm the surgeons have received the recall communication and have returned any of the surgical techniques and instruments. |  | 179 | 179 units | Nationwide | Terminated | August 18, 2009 | 705 | 6071 | Not_Computer
Z-0533-2008 | 44734 | Yes | BYE | Positive end expiratory pressure breathing attachment. | Anesthesiology | Anesthesiology | 510(k) | Ambu, Positive End Expiratory Pressure (PEEP) Valves, PEEP Valve 10, 30 mm, REF 000 137 000, Manufacturer: Ambu International A/S, DK2750 Ballerup USA: Ambu Inc. (Ambu USA) Linthicum, MD 21090-1356 USA, Made in Denmark |  |  | August 21, 2007 | 2007 | Ambu Incorporated | Lower pressure than expected: Medical device for respiratory care may leak and not register accurate flow settings that affect patient respiration. | Process design | Ambu notified consignees with a letter dated 08/21/07 and flagged as Trouble Shooting Guide. The letter advised that some valves do not hold the set PEEP level under certain conditions and consignees were further advised to consider the guideline as an addendum to existing directions for use. Consignees were instructed to return (to Ambu) for replacement the identified lots listed in the notification and confirm receipt of the advisory by e-m or fax. Ambu subsequently identified additional lots of valves and replacement valve parts for recall and notified all consignees by "Urgent Device Recall" letter dated 11/29/07. The letter advised healthcare practitioners of the impact on patient treatment that may result in less than effective treatment. Consignees were further advised to cease use of the device and return for replacement. |  | 197 | 197 total | Nationwide. | Terminated | June 16, 2008 | 300 | 197 | Not_Computer
Z-0312-2008 | 44916 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER LH750 Analyzer, Part Number 6605632, Beckman Coulter, Brea, CA. (Beckman Coulter, Brea, CA. (Hematology analyzer) |  |  | April 18, 2007 | 2007 | Beckman Coulter Inc | Patient Mis-identification can occur:1)- When a positive identifier (Sample ID or Cassette/Position) is manually edited to a positive identifier that is already in the To Do list. The workstation will accept the entry and no error message will be generated, creating 2 samples with the identical positive identifier.2)- When manually entering a Patient ID, if a blank space is entered between th | Software design | Product Corrective Action (PCA) letters were mailedon April 18,2007, to all GENS, LH500, LH750 andLH780 accounts. They are informed that a potentialmisidentification can occur in 2 scenarios.1. When a positive identifier (Sample ID orCassetteIPosition) is manually edited to apositive identifier that is already in the ToDo list.The workstation will accept the entry and noerror message will be generated, creating 2samples with the identical positive identifier.2. When manually entering a Patient ID, if a blankspace is entered between the characters in thePatient ID field, the system will only accept thecharacters before the space.----Short term fix-----The letter explains the issues and provides thecustomer work around. (Letters were sent by USmail). A customer response form was provided. The root cause of the problem is still under investigation, software upgrades, when approved will be made to devices. |  | 2344 | 2,214 units in US; 130 units in Canada | World wide: USA and Canada | Terminated | April 16, 2012 | 1825 | 4294 | Software | Output/Calculation | Software update | Software Update
Z-0319-2008 | 44892 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Optigun Cement Gun, Ratchet, non-sterile; REF 419500 and REF 4195 (in Europe), Biomet, Inc., Warsaw, IN |  |  | September 21, 2007 | 2007 | Biomet, Inc. | Loose Screw Component: A screw may come loose and fall from the device during use, and result in implantation with the bone cement. | Device Design | Consignees were notified via Urgent Medical Device Correction letter dated 9/21/07 to assure the screw components are fully seated and secure, to not use any guns with loose or missing screws, and to contact Biomet for instructions to return all instruments for evaluation and refurbishment. |  | 6590 | 6,590 worldwide (2,301 to U.S.) for both recalls Z-0318-0319-2008 | Worldwide: USA, Canada and Japan. | Terminated | June 30, 2009 | 648 | 6590 | Not_Computer
Z-0523-2008 | 45814 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | life+cel, Replacement Battery for Cardiac Science Powerheart AED G3, Lithium Sulfur Dioxide, OEM P/N: 9146-001, voltage 12V, 7.5 amps, Mfg for Battery Zone, Inc ; Mfg by Nordix, Inc. |  |  | August 24, 2007 | 2007 | BatteryZone, Inc. | Marketed without 510 (K) approval | No Marketing Application | The original recall communication (important recall information) was sent out on 9/3/07 by Phoenix Solutions, a consultant to Battery Zone. A second recall communication, Urgent Device Recall, was issued by Battery Zone dated 10/25, and was sent out to all customers (even those who had returned product) on 11/5/07. Both communications were sent by first class mail. |  | 24 | 24 units | Nationwide; (primarily police/fire departments) | Terminated | July 01, 2013 | 2138 | 142 | Not_Computer
Z-0153-2008 | 45435 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Precept Medical Products, Tape Fog Shield Surgical Mask Green Fiberglass Free, REORDER #: 65-3322, non-sterile, 50/box, 5 box/case, Precept Medical Products, Inc., Arden, NC 28704, 1-800-438-5827, Mexico |  |  | September 14, 2007 | 2007 | Precept Medical Products, Inc. | Small slits under the folds of the masks; thus they may not provide the protection necessary between the patient and surgical staff | Process control | The firm notified its consignees on 09/14/2007 by Urgent Voluntary Medical Device Recall letter (USPS mail, certified, return-receipt requested). The recall was expanded on 10/04/2007 by Urgent Voluntary Medical Device Recall--Expansion Letter to include additional lots of product. Recall letters instruct users to cease use of devices and return them as well as conduct sub-recalls to the users level for any devices they have sub-distributed. |  | 11546350 | 11,546,350 (both products-Z-0153/0154-2008) | Worldwide; USA, Australia and Canada | Terminated | September 11, 2009 | 728 | 11546350 | Not_Computer
Z-0148-2008 | 45398 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Ellex, Eye Cubed, I 3 System-ABD Ultrasound Version 3.0 with Software Version 3.0 and Biometry A-scan modality, Ellex Innovative Imaging, Sacramento, CA 95827 |  |  | September 25, 2007 | 2007 | Ellex Inc. | Incorrect IOL calculations: software behavior with Version 3.0 Eye Cubed Ultrasound systems running Biometry A-scan enabled Version 3.0.0 software can cause incorrect IOL calculations. | Software design | A Global Customer Service Bulletin (GCSB)-Urgent Notification of Software Correction was issued on 25 September 2007, which provided short term workaround directions pending a free software undate promised in the near future. |  | 59 | 59 | Worldwide; including Canada and Mexico. U.S.A. distribution included states of AL, AZ, CA, CT, FL, GA, IL, MA, ME, MI, MN, NC, NE, NJ, NM, NY, OH, OR, PA, SC, TN, and TX. | Terminated | September 08, 2008 | 349 | 59 | Software | Output/Calculation | Software update | Software Update
Z-0149-2008 | 45404 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | VortexÂż VX, Port System, 9.6FR, REF/Part Number : P5455K, Catalog Number: P5455, Latex Free, Sterile, AngioDynamics Incorporated, Manchester, GA. 31816 |  |  | October 03, 2007 | 2007 | Rita Medical Systems, Inc. | Incorrect Size Catheter: The product may contain an incorrect size catheter. | Employee error | Notification of the AngioDynamics, Inc. sales force and the domestic and international distributors began on 10/3/2007. Starting on 10/09/2007, Urgent Medical Device Recall letters were sent to all customers requesting that device useage cease and immediate return begin as well as the completion of a fax back form (number of devices used, number to be returned). |  | 150 | 150 kits | Worldwide : USA including states of AL, AR, CA, CO, FL, GA, MD, NC, OH, TX, and VT, and countries of Australia, Korea and The Netherlands. | Terminated | July 14, 2008 | 285 | 150 | Not_Computer
Z-0211-2008 | 45426 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | HUDSON RCIÂż, Variable Concentration, Large Volume Nebulizer, Catalog Number (REF) 1770 and 41770, Rx Only, Teleflex Medical, Research Triangle Park, NC 27709. |  |  | August 22, 2007 | 2007 | Teleflex Creek Dr | Decreased in flow output; through the nebulizer -(when used at higher oxygen concentration settings, this decrease in flow, may lower the oxygen concentration inspired by the patient) | Release of Material/Component prior to receiving test results | Customers were notified by an Urgent Medical Device Recall letter on/about 08/22/2007. They were instructed to return the affected products to Teleflex Medical or acknowledge that the products no longer exist. In addition, each were instructed to forward the letter if they have further distributed into institutional settings or to homecare patients and retrieve relevant product from those locations. This recall is to the consumer/user level via sub-recall from distributors or direct accounts as per the Recalling firm. |  | 215656 | 215,656 units | Worldwide: USA, Brazil, Canada, Chile, Colombia, Costa Rica, Ecuador, France, Germany, Guatemala, India, Mexico, Panama, Peru, Philippines, South Korea, Thailand, United Arab Emirates, and The United Kingdom. | Terminated | August 13, 2012 | 1818 | 215656 | Not_Computer
Z-0366-2008 | 44790 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems, Cardiac Viewer or Pulmonary Viewer Application that can run on all Gemini TFs, GXLs and EBW Workstations. Philips Medical Systems, Cleveland, OH 44143. (The product is a software application that is installed in Philips Medical System CT scanners.) |  |  | September 10, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | Incorrect Measurement: Incorrect measurement will occur when the operator uses the distance or area measurement function, such as Graphic Tool line or ROI on zoomed and batch saved images from the Cardiac Viewer or Pulmonary Viewer and then views them in the CT Viewer or on a PACS System. | Software design | Consignees were notified of this recall by an Urgent Device Notification letter sent on 9/10/07. The letter instructed the users to not save any zoomed images to a batch file. Make all measurements of zoomed images in the CT viewer, which provides correct measurements. Do not perform measurements on any zoomed batched images that have been stored on a PACS System. Phillips in investigating the problem and will provide updates free of charge when completed. |  | 2544 | 2,544 | Worldwide Distribution - USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WI, WV, and WY, and country of Canada. | Terminated | September 19, 2012 | 1836 | 2544 | Software | Output/Calculation | Update | Other
Z-0456-2008 | 45442 | No | JTR | Quality control kit for culture media. | Microbiology | Microbiology | 510(K) Exempt | MicroBiologics; P. vulgaris, DuoPacksl Product Catalog Number: 0691P; ARCC 0427; Exp 2008-9; MicroBiologics, St. Cloud, MN 56303. |  |  | October 12, 2007 | 2007 | Microbiologics Inc | Mislabeled: DuoPacks of Proteus vulgaris contained Listeria monocytogenes instead of the labeled Proteus vulgaris. | Employee error | Consignees were initially notified of this recall by a telephone survey. An Urgent Product Recall letter was sent on 10/17/07 explaining the risk to patients. The firm sent replacements of the recalled product on 10/17/07 or 10/18/07, with request for acknowledgement of Recall Notification, and certification of disposal of incorrect product. |  | 20 | 20 | Worldwide Distribution - USA including states of OR, PA, NY, RI, GA, SC, IL, PR, CA, AZ, and NJ, and countries of Japan, France, Colombia, and Brazil. | Terminated | June 26, 2008 | 258 | 20 | Not_Computer
Z-0377-2008 | 45805 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche Ammonia/Ethanol/CO2 Calibrator for use in conjunction with Roche/Hitachi, COBAS INTEGRA, cobas c111 and cobas c 501 clinical chemistry analyzers; Catalog No. 20751995190. Roche Diagnostics Corporation, Indianapolis, IN 46250 |  |  | November 14, 2007 | 2007 | Roche Diagnostics Corp. | High Quality Control Recovery: Undetected high quality control recovery in bicarbonate assay. | Process design | Consignees were notified of this recall by an Urgent Medical Device Recall letter sent on 11/14/07. The letter instructed the users to immediately discontinue use of Ammonia/ Ethanol/CO2 Control N and Ammonia/Ethanol/CO2 Control A (with listed lots) for quality control assessment of the bicarbonate assay. They were informed that controls may still be used to assess ammonia and ethanol, but alternative controls must be used for assessment of the bicarbonate assay. Consignees are instructed to destroy products at their facility. Consignees were notified of the availability of replacement product via letter dated . |  | 1904 | 1,904 | Nationwide Distibution | Terminated | September 22, 2009 | 678 | 4011 | Not_Computer
Z-0126-2008 | 45328 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Viasys Healthcare NicoletOne ICU Monitor, Modular Neurodiagnostic System (containing Axiomtek Panel PC computer system), released and installed between December 2004 and March 2007 (Axiomtek Panel PC Model Number: 2147-671) VIASYS NeuroCare Part Numbers: 842-673800, 842-673900, 982A0421 and 982A0422. |  |  | September 11, 2007 | 2007 | Nicolet Biomedical Div of Viasys Healthcare | Display lock-up: VIASYS Healthcare NicoletOne ICU Monitor freezes up during operation when using the digital video option. | Component design/selection | Certified direct mail delivery of URGENT MEDICAL DEVICE FIELD CORRECTION (RECALL) were made to consignees on 9/11/07 by Nicolet Biomedical Div of Viasys Healthcare. The letter describes the product and problem. VIASYS requests that customers cease use of the video option of the device until the monitor can be replaced by VIASYS authorized personnel. Letter also requests consignees to fax or mail back the response form. |  | 33 | 33 | Worldwide Dsitribution - USA, Germany, Netherlands, Kuwait, Japan, Ireland, and China. | Terminated | December 10, 2011 | 1551 | 33 | Other | Device Operation | Replace | Remove or Replace
Z-0246-2008 | 45153 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Biograph 16, Biograph PET/CT Scanner, Catalog No. 872858, Manufactured by Siemens Medical Solutions USA, Inc., 810 Innovation Drive, Knoxville, TN 37932 USA, TEL: 865-218-2534 |  |  | July 23, 2007 | 2007 | Siemens Medical Solutions USA Inc | Visual asymmetric appearance; An asymmetry can be introduced into the attenuation-corrected PET images when using iterative reconstruction in combination with a specific number of subsets and reconstruction matrix sizes. (This problem only occurs when using protocols other than the default Biograph protocols) | Software design | A certified letter-Urgent Device Correction- dated June 15, 2007 was sent to all affected consignees on July 23, 2007 informing them of the problem (PET iterative reconstruction algorithm) and requesting as a short term fix that they discontinue the practice of operating using protocols other than the default Biograph. Software upgrades as the long term fix were promised in the Urgent Device Correction Notice. Updated information received October 4, 2007: a software patch to fix PETsyngo 5.0 was released, software patch to fix PETsyngo 4.x was released, and target date for the software patch to fix PETsyngo 3.2 and 3.4 is the end of November 2007. |  | 162 | 162 units | Worldwide-USA and countries of Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, Chile, China, Czech Republic, Germany, Denmark, Spain, France, Great Britain, Hong Kong, Hungary, Ireland, India, Italy, Japan, South Korea, Kuwait, Lebanon, Mexico, Malaysia, Netherlands, Norway, Poland, Portugal, Sweden, Singapore, Slovak Republic, Thailand, Turkey, and Taiwan. | Terminated | September 11, 2009 | 781 | 415 | Computer
Z-0301-2008 | 45446 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott Diagnostics CELL-DYN Sapphire Hemoglobin Reagent Analyzer Syringes, List Number 08H49-02, Abbott Laboratories, Abbott Park, IL (the syringe is packaged in a white box and sealed with a label wrapped around the center of the box-The main analyzer list number is 08H00-01, ) |  |  | July 25, 2007 | 2007 | Abbott Laboratories | Syringe Failure: Insufficient lubrication of the plunger tip may cause syringes to fail at installation or shortly thereafter | Nonconforming Material/Component | This Recall was initiated 07/24/07 via a Product Recall letter-Immediate Action Required-instructing users to discontinue use and destroy any units under recall. Replacements may be obtained by contacting the local Customer Support Center at 1-877-4Abbott |  | 302 | 302 units | Worldwide Distribution: USA including states of AR, AZ, CA, CT, FL, GA, ID, KY, LA, MD, MI, NC, ND, NV, NY, OH, OK, PA, SC, TX, UT, VA, and W I (4 distributors and 50 consignees)and countries of Australia, Brazil, Chile, Germany, and Hong Kong. | Terminated | November 17, 2009 | 846 | 302 | Not_Computer
Z-0460-2008 | 45504 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Cardinal Health Alaris Pump Module (formerly Medley Pump Module), Model 8100, Cardinal Health, San Diego, CA. (All Alaris Pump modules shipped prior to December 23, 2007 are subject to this recall) |  |  | October 29, 2007 | 2007 | Cardinal Health 303 Inc DBA Alaris Products | Inaccurate flow rate; related to misassembled (missing, bent or broken) springs during the manufacturing or servicing of the mechanism assembly | Process control | The field notification (Urgent Medical Device Recall Letter) will consist of a mailing conducted through consignees (distributors in the affected countries as well as direct customers for the Alaris Pump module (a.k.a. Medley Pump module). The domestic letters are expected to be mailed on November 5, 2007 to all of the affected accounts by return receipt mailing. A Customer Response Card will be mailed with each letter along with a Service Bulletin detailing the protocol to conduct the Occluder Pressure Test. Four (4) letters will be mailed to each account to the following titles/job functions: Director of Nursing, Director of Biomed, Director of Materials Management, and Director of Risk Management. Distributors will be sent multiple copies of the Recall Notification, Customer Response Cards, and Service Bulletin, along with a cover letter outlining their responsibilities to complete the Recall notification plan and notify their customers as requested. The letters will be the first source of communication regarding this issue. The Recall Center will be contacting customers shortly to make arrangements for inspection of the devices at the users site. The international mailings are being coordinated out of the offices located in the affected countries. The letter will be translated to the appropriate language and mailed to all of the affected customers. The appropriate international regulatory authorities will be notified of this issue, and will work with the country managers.---------A web site for all information on this recall is listed as: www.cardinalhealth.com/alaris/indexmodulealert.asp******************January 23, 2008 the firm notified the Los Angeles Districe that they will be sending a recall notification update to 267 customers and distributors who are affected by this 806 Notification Amendment. Cardinal Health intends to notify these consignees by mail on January 30, 2008. This update will affect a total of 267 customers, 5,681 units and 798 |  | 207257 | 207,257 units | Worldwide: USA including the following forty-six (46) states, along with Washington, D.C. Puerto Rico and Guam: Alaska, Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Florida, Georgia, Hawaii, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Maine, Michigan, Minnesota, Missouri, Mississippi, Montana, North Carolina, Nebraska, New Hampshire, New Jersey, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, and Wyoming and countries of Canada &amp; Saudi Arabia (Note: Devices were shipped to the Cardinal Health Distribution Center in the Netherlands. These devices were then shipped to Saudi Arabia.) | Terminated | June 06, 2011 | 1316 | 207257 | Not_Computer
Z-0457-2008 | 45799 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | bioMerieux, BacT/ALERTÂż FA Culture Bottles, bioMerieux, Inc., Durham, NC 27704 |  |  | November 05, 2007 | 2007 | bioMerieux, Inc. | Duplicate Bottle Indentification -- Three lot numbers of BacT/ALERTÂż FA Culture Bottles labels contained duplicate bottle identification (Bottle ID) barcode numbers. (Lots 1017685 and 1017788 contain the same series of Bottle ID's; Lot 1017693 contains duplicate Bottle ID's within the lot and may share Bottle ID's with both lots 1017685 and 1017788) | Process control | bioMerieux notified all Consignees with a Urgent Product Recall Notice letter on 11/05/2007. They were instructed to examine their inventory to determine if they have any remaining inventory of the 3 affected lots. If they have any of these bottles they were instructed to stop using them and destroy all remaining inventory. A Customer Acknowledgement Form was attached to be completed and returned to bioMerieux via fax or mail. There were also instructions attached for the use of Generic Bottle Barcodes and Creating a Unique Bottle ID. |  | 7025 | 7,025 | Worldwide Distribution-USA including countries of Australia, Austria, Bangladesh, Brazil, Bulgaria, Byelorussia, Chile, China, Colombia, Croatia, Denmark, Dijbouti, Estonia, Federation of Russia, Finland, France, Germany, Greece, Guadeloupe, India, Israel, Italy, Korea, Kuwait, Martinique, Mexico, Netherlands, New Caledonia, Norway, Panama, Poland, Portugal, Reunion, Saudi Arabia, Serbia Montenegro, Slovenia, Spain, Sweden, Switzerland, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Vietnam. | Terminated | December 08, 2009 | 764 | 7025 | Not_Computer
Z-0566-2008 | 45904 | No | MQX | Acupuncture needle. | General Hospital | General Hospital | 510(K) Exempt | HAENG LIM --Acupunture Needles- PRE-STERILIZED GAMMA-RAY; 100 PKGS (10 NEEDLES/PACKAGE-1,000 needles in each box). Sale Agency HAENG LIM SEO WON #45-9, Gunjang-ri, Ganam-myeon, Yeoju-gun, Kyungki-do, Korea |  |  | November 28, 2007 | 2007 | Marknew Products | Marketed without 510(k) | PMA | Marknew Products, LLC, sent an Important Recall Notice to their customers on November 27, 2007, advising them that all Haeng Lim Seo Won Acupuncture Needles that they have received, have not been cleared for distribution in the USA. All customers have been requested to return product back to Marknew Products for credit. A response form was included. If customers have questions concerning this product, they are asked to call 562-908-5000. |  | 32038000 | 320-380 boxes/100 packages (10 needles/package) | Nationwide Distribution | Terminated | January 18, 2012 | 1512 | 32038000 | Not_Computer
Z-0320-2008 | 45852 | No | HJM | Transilluminator. | Ophthalmic | Ophthalmic | 510(K) Exempt | Transilluminator, Model BFO-150, a Fiber Optic Light Source (Illuminator), available in 2 voltages: 115v for USA users &amp; 220v for foreign countries. --- Chiu Technical Corporation, 252 Indian Head Road, Kings Park, NY 11754. |  |  | November 09, 2007 | 2007 | Chiu Technical Corp | Device was marketed without a 510(k). | Process design | On 11/09/07, the recalling firm started faxing the Recall Letter and response form to its customers, informing them of the recall of the unapproved device to the user level. |  | 191 | Domestic: 170 units, Foreign: 21 units | Worldwide: USA, Canada and India | Terminated | October 17, 2012 | 1804 | 191 | Not_Computer
Z-0300-2008 | 45437 | No | KEY | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Ventana Medical System's Symphony Staining System, KPA, Model #: 2200000, Catalog #: 900-SYM1 &amp; 900-SYM3, Ventana Medical Systems, Tucson, AZ |  |  | September 06, 2007 | 2007 | Ventana Medical Systems Inc | Shock hazard: Devices released prior to full testing; the configuration of the Symphony instrument during Hi pot testing procedures did not completely evaluate all components in the test loop. In specific circumstances, a shock hazard exists. | Process control | On 9/6/07, the firm began visiting customers to perform the HiPot test. Prior to performing the High Potential Test of Symphony Instruments already in the Field, the Field Service Engineer or FSE ensures that the customer fully understands the purpose of the test and why it is necessary to perform the test at their site. The FSE conveys the following points regarding the HiPot test:- The HiPot test is a safety verification for the electrical path within the instrument.- The HiPot test was performed incorrectly during the manufacturing process.- The HiPot test is used to verify that if a high voltage overload condition were to occur it would not lead to a short/arc within the instrument, which could pose a risk to anyone that touches the instrument- The HiPot test places a voltage of 2,000VDC onto the high voltage path of the instrument and if any of the insulating components are compromised, the tester will notify the FSE of the condition.- If the instrument fails the HiPot test, the faulty component will be replaced and the instrument will be retested until it passes.- Due to the high voltage used the during the test, no one will be allowed to touch the instrument until the test is complete.The FSE performing the test documents that the customer was informed of the reason for the High Potential Test on their Field Service Report. |  | 45 | 45 units | Worldwide: USA, Japan and France | Terminated | July 08, 2008 | 306 | 45 | Not_Computer
Z-0299-2008 | 45397 | No | LOJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | ARCHITECT i System Assay CD-ROM - WW (excluding USA accounts), Versions prior to Version 24.0; for use on the ARCHITECT i System; List Number (LN) 6E59,Abbott Laboratories, Diagnostics Division, Abbott Park, IL 60064 USA |  |  | September 18, 2007 | 2007 | Abbott Laboratories | The lower limit flag for auto-dilution protocol 3 (amniotic fluid) was set too low (15 ng/mL) for the ARCHITECT AFP assay file for the ARCHITECT AFP non-U.S. product. The assay labeling indicates that the limit flag should be set no lower than 20 ng/mL | Software design | Abbott affiliates were e-mailed copies of the Product Correction (Immediate Action Required) Recall Letter on 9/18/07 for translation where needed and dissemination to their ARCHITECT accounts who have the affected assay CD-ROMs. Each Abbott international organization will be responsible for contacting customers and determining the method of effectiveness checks based on their country requirements. Abbott sent Product Correction, Immediate Action Required letters dated 18 September 2007 to all active non-U.S. ARCHITECT customers who have the affected assay CD-ROMs. The end users were informed that current AFP file has been updated on a new ARCHITECT I Assay CD-ROM (LN 6E59, version 24). The Low-Linearity flag threshold for the Amino 1:50 dilution has changed from 15 ng/mL to 20 ng/mL for consistency with overall assay sensitivity. The accounts were advised to install the updated version of the current ARCHITECT AFP assay when they receive the version 24 ARCHITECT I Assay CD-ROM. The accounts were requested to keep a copy of the letter on file. The accounts were requested to complete the customer reply form, indicating that the letter was received, understood and disseminated to the laboratory manager/supervisor responsible for Architect testing, and fax it to Abbott by 9/28/07 at 1-800-777-0051. |  | 2490 | 2,490 | International Only: Australia, Canada, Chile, Costa Rica, Dominican Republic, Germany, Guatemala, Honduras, Hong Kong, Jamaica, Japan, Mexico, New Zealand, Peru, Singapore, South Korea, Thailand, Uruguay and Venezuela. There were no USA accounts. | Terminated | November 05, 2008 | 414 | 2490 | Not_Computer
Z-0310-2008 | 45461 | No | JDK | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VariLift Bone Plug w/ End Cap; 13mm Dia, 24mm long; Ref #: CIE1-13-24-BP; Sterile; Encore Orthopedics, Austin, TX 78758. |  |  | September 19, 2007 | 2007 | Encore Medical, Lp | Device marketed without label warning required in the device's 510(k) approval letter. (In part-Warning: This device is not intended for any spinal indications.....) | Labeling Change Control | The recalling firm notified consignees of its recall via an Urgent Device Recall letter on 09/19/07. Consignees notified of a labeling error regarding device usage and were told they would be notified by a customer service rep to facilitate return of the recalled devices. |  | 61 | 61 units. | Nationwide: including direct accounts in CO, KS and TX. | Terminated | March 10, 2008 | 173 | 175 | Not_Computer
Z-0297-2008 | 45511 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens, ACUSON Antares, Ultrasound System (Standard and Premium Edition), 5.0 Systems with software version 200.0.054, P/N : 10032747 and 10032746 standard, 10032746, 10037591, and 10038202 for premium edition, Siemens, Ultrasound Division, Mountain View, CA 94039 |  |  | July 10, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Image errors: A software problem results in on-screen indications (graphics) that lead the user to believe that the patient's right and left and the transducer orientation are oriented the same. This is not the case and the actual patient's right and left are reversed from the on-screen indicators. | Software design | The firm sent a recall letter, Urgent Medical Device Notification, to consignees on 07/10/2007, to be followed by a service visit by Siemens personnel for a mandatory software update to be installed. Short term the users were instructed to avoid increasing the SieClear compounding levels to an Advanced SieClear compounding setting after an L/R invert action with Biopsy guidelines on-screen. |  | 418 | 418 units (Z-0297-0298-2008) | Worldwide: Product was distributed to medical facilities Nationwide including the states of OH, FL, IL, NY, CA, NH, TX, NJ, PA, MD, MO, GA, WA, WI, NE, WV, HI, MA, SC, MI, VA, AZ, UT, NH, RI, KY, SD, and IN, and internationally to Canada, Germany, Australia, Sweden, Saudi Arabia, Chile, South Korea, Mexico, France, UK, Taiwan, Singapore, Denmark, Brazil, New Zealand, Ireland, Belgium, Norway, Serbia, Vietnam, UK, Poland, China, Spain, South Africa, Austria, Croatia, Russia, Greece, Trinidad and Tobago, United Arab Emirates, India, Japan, Malaysia, Turkey, Thailand, Norway, Slovakia, Romania, Netherlands, Switzerland, Portugal, Italy, and Venezuela. | Terminated | December 20, 2010 | 1259 | 418 | Software | Display/Image | Software update | Software Update
Z-0305-2008 | 45811 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Toshiba Aplio XG Diagnostic Ultrasound System, Model: SSA-790A with version 1.1 software, Toshiba America Medical Systems, Inc., Tustin, CA. 92780 |  |  | October 02, 2007 | 2007 | Toshiba American Med Sys Inc | Panel Lock-up: When using dynamic image acquisition modes the panel may lock up, if the operator attempts to perform a still image capture. If the panel locks up the system will have to be rebooted to continue the examination which may cause a delay in completion of the patient examination. | Software design | Notification letters, Urgent Medical Device Correction, with an attached Customer Response form were sent to customers beginning on October 2, 2007. Customers were notified this action was being taken as a result of an investigation that resulted from a software issue that was found by the manufacturer. Customers were informed that a Toshiba Customer Service Engineer would contact them to schedule a software upgrade when the software becomes available. If customers have any questions they were instructed to contact their or call (800) 421-1968 and ask for the Regulatory Affairs Department.--A short term instruction included "Avoid using the Freeze Key when performing dynamic image acquisitions" |  | 69 | 69 | Nationwide: USA including the states of AK, AZ, CA, FL, IA, IL, IN, KS, MI, MN, MO, MT, NE, NV, NY, OH, OR, PR, VA, WI &amp; WV | Terminated | April 20, 2012 | 1662 | 69 | Software | Device Operation | Software update | Software Update
Z-0447-2008 | 45460 | No | CBH | Tracheal tube fixation device. | Anesthesiology | Anesthesiology | 510(K) Exempt | Dale 242 Blue (Neonate, Infants and Small Children) Tracheostomy Tube Holder, Product Number: H84102421, Dale Medical Products, Inc., Plainville, MA 02762 |  |  | September 27, 2007 | 2007 | Dale Medical Products, Inc. | Revised Labeling : Instructions For use: "Do not trim fastener hook tabs. " Trimming can cause tabs to disengage at the trach plate | Labeling design | Dale Medical notified Distributors by letter dated 9/27/07, identified as an Urgent Medical Device Correction Letter. The letter requested distributors to notify all their customers by letter. The Correction Letter reminded users to take notice of new labeling " Do not trim fasterner hook tabs" |  | 65000 | 65,000 | Worldwide-USA and countries of Canada, Austria, Belgium, Denmark, Finland, France, Germany, Greece, Iceland, Italy, The Netherlands, Norway, The UK, Spain, Switzerland, Australia, Hong Kong,India, Japan, Korea, Malaysia, New Zealand, The Philippines,Singapore, Taiwan, ROC, Brazil, Bermuda, Chile, Costa Rica,Ecuador, Mexico, Panama, Venezuela, Thailand, Saudi Arabia, and The United Arab Emirates. | Terminated | May 09, 2012 | 1686 | 665000 | Not_Computer
Z-0302-2008 | 45474 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Extron Electronics OEM, CAB, DVI-Male to DVI-Female Cable Assembly (14 inch.); Part #: 43-111-01 (Stryker Part #: 0100-224-131); Distributed by Stryker Communications, Flower Mound, TX 75028. (These cables are commonly used with the Switchpoint Infinity II) |  |  | October 03, 2007 | 2007 | Stryker Communications | Intermitten or Complete Loss of Signal: out-of-specification DVI cables causing intermitten or complete loss of signal on monitors used in surgery. | Nonconforming Material/Component | Firm notified consignees of recall via letter (Urgent Device Recall) on 09/21/07. Consignees were notified that a customer service representative would visit each facility to replace out-of-spec cables with properly shielded cables at no charge. |  | 2255 | 2,255 cables. | Worldwide: USA and Australia, Belgium, Canada, France, Germany, Great Britain, Hong Kong, Japan, Korea, Mexico, Netherlands, New Zealand, Portugal, South Africa, Spain, Sweden, Switzerland, Taiwan, and The U.A.E. | Terminated | June 16, 2008 | 257 | 3908 | Not_Computer
Z-0376-2008 | 45494 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Cepheid, Xpert GBS Assay Diagnostic Test Kit Part No: GXGBS-100N-10 (in foil pouch) for use with the GeneXpert DX System, Cepheid, Sunnyvale, CA 94089 |  |  | September 05, 2007 | 2007 | Cepheid | Labeling Error: Some cartridges may be mislabeled as "Xpert EV" instead of "Xpert GBS". (The information embedded in the barcode label on the cartridge is correct for Xpert GBS and brings up the GBS assay definition file) | Release of Material/Component prior to receiving test results | Consignees were sent a Correction and Removal letter on 09/04/2007, notifying them that there was a labeling error with some of the cartridges labels indicating Xpert EV instead of Xpert GBS on the cartridge label. They are asked to discontinue using it immediately and return to Cepheid. |  | 105 | 105 | Worldwide Distribution | Terminated | April 04, 2008 | 212 | 105 | Not_Computer
Z-0488-2008 | 45870 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Premier Edge, Bulk Microsurgical Knife, 3.0mm Slit Knife, Item Code B-PE3830, Oasis, Glendora, CA |  |  | July 18, 2007 | 2007 | Oasis Medical Inc | Sterility (package integrity) compromised: the device is puncturing the packing during handling. Once the packaging is damaged, the product is no longer sterile as labeled. THIS IS AN EXPANSION OF A PREVIOUS RECALL | Device Design | A recall notification, Urgent Medical Device Recall Expansion Letter, dated 7.18.2007, was sent to customers advising the letter is to inform them of a voluntary product recall expansion involving the OASIS Medical Premier Edge Knives. A product label for ease in identifying the product is included. The letter further advises this voluntary recall has been expanded due to the product puncturing the packing during rough handling. Once the packaging is damaged, the product is no longer sterile as labeled. The letter instructs customers to immediately examine their inventory and quarantine product subject to recall. Once the product has been identified, product is to be returned by FedEx to OASISMedical. The letter further instructs that if the customer has further distributed this product, to identify their customers and notify them at once of this product recall. This recall should be carried out to the user level. |  | 28319 | 28,319 individual knives (Z-0485-0516-2008) | Worldwide: USA, Kuwait, Lebanon, Lithuania, Mexico, Netherlands, Norway, South Africa, Saudi Arabia, Slovakia, Spain, Sweden, Switzwerland, Turkey, United Kingdom, Ecuador, Finland, France, Germany, Greece, Italy, Denmark, Columbia, Chile, Brazil, Belgium, and Australia | Terminated | September 25, 2009 | 800 | 28319 | Not_Computer
Z-0326-2008 | 38368 | [6 empty] | Actiwatch 64 AM, Bandless Part Number 198-0301-AM1, Respironics, Inc., Mini Mitter, a Respironics Company, Bend, Oregon |  |  | June 27, 2007 | 2007 | Mini-Mitter Company, Inc. | Communication Failure: the devices have a memory chip that has a limited number of read/write cycles which can trigger set up information and/or subject information anomalies or intermittent communication annomalies. If left unaddressed, the issues can lead to permanent communication failure. | Software design | A Voluntary Medical Device Recall letter dated June 26, 2007. The recall letter will instruct the customers to 1) fill out a business reply form to confirm contact information and devices serial numbers in their possession. Then FAX this info back immediately to Stericycle or Mini-Mitter. 2) Refer to the work scheduled rework date reserved for them and noted in their recall notification letter and return the device for service at that time; and 3) Package the device appropriately, insure it and return it via UPS using the provided UPS shipping label. Product returned to Mini-Mitter will be quarantined, decontaminated, and repaired. Min-Mitter will rework all the devices returned. Each will have a new memory chip installed, be re-calibrated, retested, and have a new battery installed if required. Min-Mitter will provide a new one year warranty. |  | 20 | 20 | Worldwide | Terminated | October 20, 2010 | 1211 | 1073 | Not_Computer
Z-0150-2008 | 45405 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Datascope, Panorama Patient Monitoring Network; Panorma Telepack 608;Software Versions 8.1.X, 8.3, 8.3.1, 8.4.1, and 8.5.X, Part Number: 0998-00-0191-04, Datascope Corp., Mahwah, NJ |  |  | October 09, 2007 | 2007 | Datascope Corporation | Display problems: If an ECG cable, which has been damaged due to the ingress of liquid or by mechanical trauma, is utilized with the Telepack, it may cause the Telepack to switch between the 3-lead and 5 lead-input modes. During the switching process, the digital heart rate displayed at the Panorama Central Station will be frozen, and subsquent arrthymia alarms will not be announced. | Software design | Urgent Product Field Correction notices were sent to hospital administrators on October 9, 2007, by certified mail, return receipt requested. The letter describes the scenario that causes the problem and its remediation It also describes how the software upgrade will fix the problem and alerts hospitals that a Datascope Service Representative will be contacting them to arrange for the upgrade. |  | 338 | 338 sites | Nationwide | Terminated | November 04, 2008 | 392 | 338 | Software | Display/Image | Software update | Software Update
Z-0295-2008 | 44941 | timer, general laboratory | JBS | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Nova 8 Analyzer Calibrator Pack (Catalog/REF # 15196), Nova Biomedical, Waltham, MA |  |  | September 17, 2007 | 2007 | Nova Biomedical Corporation | Elevated Normalized Ionized Calcium (nCa) and Normalized Ionized Magnesium (nMg) calculated values on patient samples | Device Design | Nova issued a Customer Advisory Notice notification by phoneon 9/17/07 to contact customers using a script . If the facility reports Normalized Ionized Calcium or Normalized Ionized Magnesium, Nova Biomedical will replace existing stock of Nova 8 calibrator packs. If the facility does not report Normalized Ionized Calcium or NormalizedIonized Magnesium, the account can continue to use the existing lot of calibrator packs.This contact will be documented within the Nova Biomedical complaint database. Intemational Customer Advisory Notice being sent to Nova Biomedical Subsidiariies and Distributors by Federal Express Envelope.The Federal Express envelopes tracking number will be used for delivery confirmation. |  | 2999 | 2,999 units | Worldwide, USA, Canada, Argentina, Bangladesh, Bolivia, China, Czech Republic, Denmark, Ecuador, Egypt, Finland, France,Germany,Greece, Hungary, India, Indonesia, Ireland,Italy, Japan, Korea, Mexico, Peru, Poland, Scotland, South Africa, Spain, Sri Lanka. Switzerland, Syria ,Taiwan, Turkey and The United Kingdom | Terminated | March 16, 2012 | 1642 | 2999 | Not_Computer
Z-0128-2008 | 45431 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Biomet 3i, Certain PreFormance Temporary Cylinder, Catalog Number: IPFTC42, Dental Implant, Biomet 3i, Palm Beach Gardens, FL 33410 |  |  | October 08, 2007 | 2007 | Biomet 3i | Incorrect assembly: The product does not allow for the screw to pass through the access hole and engage the implant properly. Therefore, the provisional restoration will not be seated on the implant platform. | Process control | Beginning on 9/27/07, Biomet 3i contacted its US consignees by telephone using a prepared script. Following telephone contact, Biomet 3i sent a recall letter, titled Urgent Medical Device Recall, by fax. The international consignees were contacted by email followed by a faxed letter. Each customer was requested to examine their inventory for the affected part and corresponding lot number, and return them to Biomet 3i. |  | 114 | 114 | Worldwide Distribution - USA, Italy, France, Spain and UK. | Terminated | July 10, 2008 | 276 | 114 | Not_Computer
Z-0484-2008 | 45822 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SERFAS Energy Suction Probe, Model 279-351-100, 3.5mm 90-S, Manufactured by Stryker Puerto Rico |  |  | April 26, 2007 | 2007 | Stryker Endoscopy | Device Shut off/on: Device may not turn off after taking finger/foot off of activating switch and may activate without pushing a button. (Patient burns may result from these device malfunctions) | Mixed-up of materials/components | Consignees were mailed an Urgent Device Removal Recall notice, dated April 27, 2007. International subsidiaries will receive the recall notice and they will distribute to their consignees. Product is to be returned as a recall. |  | 6037 | 6,037 units | Worldwide Distribution: USA including states of AL, AR, AZ, CA, CO, CT, DE, DC, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI, WY, and WV; INTERNATIONAL: Affiliated Stryker Distribution Companies: Sweden (for Scandinavia area), Australia, Canada, France, Germany, Greece, South Africa, Iberia, Italy, Mexico, Portugal, Switzerland (2 separate sites), and The UK | Terminated | January 16, 2008 | 265 | 6037 | Not_Computer
Z-0519-2008 | 45913 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Medical Systems Information Technologies Clinical Information Center (CIC) Pro v5.0.3 &amp; v5.0.6 systems, GE Healthcare, Clinical Systems |  |  | November 08, 2007 | 2007 | General Electric Medical Systems Information Technology | Potential for Loss of Audible Alarming and ECG Tab Setting Changes; from Clinical Informaiton Center (CIC) Pro v5.0.3 &amp; 5.0.6 when used in combination with Telemetry and bedside monitors. | Software design | A "GE Urgent Medical Device Correction" letter was sent to consignees on 10/16/2007. The letter described the problem and product involved. The letter also gave short term recommendations. A long term solution is being worked on by GE and will be announced when determined. |  | 3116 | 3, 116 | Worldwide-USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MN, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY, PR, and District of Columbia and countries of The United Kingdom, Turkey, Thailand, Tsiwan, Syria , Switzerland, Sweden, Spain, Slovenia, Singapore, Saudi Arabia, Russia, Poland, Qatar, Pakistan, Philippines, Palestinian Territory, New Zealand, Norway, Netherlands, Nigerie, Malaysia, Lebanon, Jordan, Italy, Israel , Ireland, India, Honk Kong, Greece, Germany, France, Finland, Egypt, Denmark, Czech Republic, Cyprus, Costa Rica, Colombia,China, Chile, Canada, Delgium, Bahamas, Austria, and Australia. | Terminated | December 10, 2011 | 1493 | 3116 | Software | Alarm/Message | Notification | Safety Notice/Insructions
Z-0733-2008 | 45906 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Sechrist Hyperbaric Chamber Gurney, Standard Gurney &amp; Stretcher Model, P/N : 21502, Sechrist Industries, Inc., Anaheim, CA |  |  | October 31, 2007 | 2007 | Sechrist Industries Inc | Brake Problems: The stretcher brake cams may crack while attempting to lock the stretcher brake (Sechrist is supplied the stretchers by the manufacturer Stryker Medical) | Device Design | Urgent Medical Device Correction letters were sent to customers beginning on November 9, 2007, via FedEx &amp; DHL overnight. Customers were notified this action was being taken to providing notification to their customers as a continued implementation of the Stryker Medical field correction.The Notification letter instructed their customers to:1. Locate the gurneys with the serial numbers listed in this advisory.2. If any of the affected gurneys cannot be located, determine if they have been sold, transferred, or donated to another facility and notify us regarding their new location.3. Verify the brakes are working properly. The brake function may be checked by engaging the brake pedal and pushing on each end of the gurney to make sure that casters are locked securely.4. Remove the gurneys with non-functioning brakes from service until brake cams have been replaced and brakes are fully functional.5. Contract Sechrist for the replacement parts package.6. Complete the retrofit of the Gurney.7. Verify correct function of the brake assembly following retrofit.8. Complete and send the Correction Verification Form contained in the package to Sechrist. |  | 559 | 559 Units Total (483 in US) for all models | Worldwide-USA, Brazil, Bulgaria, Canada, Dubai, U.A.E., Estonia, India, Indonesia, Japan, Malaysia, Saudi Arabia, Sweden, Taiwan, Thailand, and Vietnam. | Terminated | April 19, 2012 | 1632 | 559 | Not_Computer
Z-0664-2008 | 45413 | No | MQF | Assisted reproduction catheters. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Tucker Embryo Catheter with URM Tip, Embryo Transfer Catheter, 20 cm, Order #: ftr120, Manufactured by: Fertility Technology Resources, Inc., Marietta, GA 30062 |  |  | September 13, 2007 | 2007 | Fertility Technology Resources, Inc. | Expiration Date for sterility; The 5 year dating could not be validated/supported for sterility assurance. | Process design | All customers were contacted by telephone starting 9/13/2007. They were asked to return all Tucker family of catheters and all Morton IUI catheters that are on the shelves. A UPS call tag was sent to each customer to facilitate the return of the products. A follow up Urgent Device Recall letter was sent to all customers on 10/9/2007 reminding them to return all outstanding inventory when the UPS call tag arrives. |  |  | not reported | Worldwide: USA including states of CA, CO, FL, GA, LA, MD, ME, MI, MS, NC, NH, NJ, NY, OH, TN, and WA and The Netherlands. | Terminated | February 06, 2008 | 146 |  | Not_Computer
Z-0120-2008 | 44937 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AmniSure ROM Test, Fetal Membrane Rupture Test for |  |  | October 03, 2007 | 2007 | AmniSure International LLC | Weak true postive/false negative results: weak true positive or false negatives in ruptured fetal membranes may occur due to a "hook effect" | Device Design | AmniSure notified users by Medical Device Correction (Recall) letter dated October 1, 2007 issued via US Mail. The letter provides additional labeling instructions and a checklist for trouble shooting, including a diluation procedure that allows users to determine and solve the problem of the hook effect with this device. |  | 319976 | 319,976 units | Worldwide, USA, Australia, Austria, Italy, Korea, Bulgaria, Russia, Spain, Kuwait, Switzerland, Luexemburg, Ireland, Denmark, Netherlands, Jordan, Israel, Chile, UK, Singapore, Basil, New Zealand, Turkey, France, Greece, and Equador | Terminated | February 06, 2008 | 126 | 319976 | Not_Computer
Z-0529-2008 | 45482 | No | NFB | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Oxy Serve II Oxygen Conserving Regulator, Model Number: MCRB-870-6, Superior Products, Inc/SP Medical, LLC, Cleveland, Ohio 44144 |  |  | October 19, 2007 | 2007 | Superior Products Inc | Leakage: Device was leaking oxygen | Component design/selection | A Voluntary Medical Device Recall notification was delivered by certified mail on or about 10/19/2007, to all consignee's; "bill to" and "sent to." Consignee's are asked to contact their customers (if any), notify them of the recall, and return all implicated products to Superior Products. |  | 125 | 125 | Worldwide-USA including states of CA, FL, IA, IN, KY, OH, SC, TX, and WA. and countries of Singapore, Mexico City, MX, and Turkey | Terminated | April 25, 2012 | 1650 | 125 | Not_Computer
Z-0365-2008 | 39353 | No | LZS |  |  | Ophthalmic | PMA | Advanced Medical Optics (AMO), VISX WaveScan WaveFront System (software). Part Number 0070-1478. Advanced Medical Optics, Santa Clara, CA 95051. |  |  | May 30, 2007 | 2007 | Visx Inc | Erroneous Treatment Calculations - Two software caused errors; in combination, in the WaveScan System Software, will result in an erroneous treatment calculation (overcorrection) in patients. (1) Installation of software for a certain brand of computer (Dedicated Computing LLC) can cause a software registry setting for an algorithm that identifies image reflection to be erroneously set in the "off | Software design | AMO sent an Action Required notification on May 30, 2007, to AMO Certified Physicians. The letter included a CD software update (0070-1721) to be installed by users that provided correct registry parameter settings. Upon completion of installation, a confirmation form will print out. Customers are requested to fax back confirmation to AMO (408-773-7055). |  | 26068 | 196 systems are affected in the U.S. and 133 outside of the U.S. | Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MD, MI, MN, MO,MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK,OR,PA, SC, TN, TX, UT, VA, WA, WI, and WV and countries of Canada and Mexico. | Terminated | April 17, 2008 | 323 | 26068 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0183-2008 | 45441 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE OEC, Insta Trak 3500 Plus with the Cranial Application for Image Guided Surgery, GE Healthcare, Surgery, Salt Lake City, UT 84116 |  |  | October 18, 2007 | 2007 | GE OEC Medical Systems, Inc | Crack, Fracture or Detachment: The Transmitter Arm for the Mayfield-compatible Transmitter may become detached from the unit and enter into the surgical field. | Device Design | Consignees were sent an Urgent Recall Notice via mail dated 10/11/07, notifying the customer to examine the plastic end of the transmitter arm and to not use the transmitter arm, if any cracks or damage are observed. Also to stop using Steam Sterilization method and to utilize the other Sterilization methods indicated in the Operators Manual. The root cause is still under investigation GEHC Surgery will provide, at no charge, a solution to this issue when it becomes available. |  | 169 | 169 units | Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IL, MA, MD, MI, MS, NC, NY, OK, OR, PA, SC, TX, UT, VA, WI, and WV, and countries of Canada, China, France, Hong Kong, Japan, Mexico, Spain, and Switzerland. | Terminated | August 10, 2012 | 1758 | 169 | Not_Computer
Z-0481-2008 | 45528 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | StelKast Proven Knee - Tibial Half Block Augmentation and Screw. Part Number SC2287-3-5, Size 3, Thickness 5 mm, StelKast, Murray, PA 15317 |  |  | October 05, 2007 | 2007 | Stelkast Co | The screws may fracture | Device Design | The recalling firm telephoned their customers on 10/5/07 to inform them of the problem and the need to return the product. As a follow up to the telephone calls, the recalling firm issued a recall letter (Field Notification) on 10/9/07 via certified mail. |  | 41 | 41 units | Nationwide: shipped to medical facilities in IL, NC, PA, TN, TX, VA, and WV. | Terminated | January 23, 2008 | 110 | 207 | Not_Computer
Z-0588-2008 | 45919 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Single Day INFUSOR Portable Elastomeric Infusion System, Product Code: 2C1071KJP; 2mL/hr with Tubing Flow Restrictor; 12 units per case; Baxter Healthcare Corporation, Deerfield, IL 60015 |  |  | November 29, 2007 | 2007 | Baxter Healthcare Corp. | Leaking: leaks at the tubing flow restrictor connection to either the coupler or the male luer during filling. | Process control | Baxter on 11/30/07 sent an Urgent Product Recall letter to affected customers, to the attention of the Director of Nursing with copies to the Materials Manager, via first class mail to advise them of the potential leaks during filling or use of the infusor. The accounts were instructed to immediately discontinue use and segregate the listed affected lot numbers, and contact Baxter Center for Service at 1-888-229-0001 to arrange for the return and credit for the affected product. Any clinical/technical questions were directed to the Baxter Healthcare Product Information Center at 1-800-933-0303. |  | 86192 | 86,192 units | Worldwide: USA, Canada, Austria, Belgium, Denmark, Finland, France, Germany, Greece, Italy, Portugal, San Marino, Spain, Sweden, United Kingdom, South Africa, Colombia, Argentina, Hong Kong, Australia and New Zealand. | Terminated | April 29, 2010 | 882 | 86192 | Not_Computer
Z-0589-2008 | 45965 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Push Button Blood collection Sets with Pre-Attached Holder;Catalog Number : 367352: 21G x 3/4" x 12"; 0.8 x 19mm x 305mm, Wingset PBBCS, BD Diagnostics, Franklin Lakes, NJ |  |  | November 20, 2007 | 2007 | Becton Dickinson &amp; Company | Safety Mechanism Failure: the safety mechanism of the BD Vacutainer Push Button Collection Set with Pre-attached holder cannot be activated; thus the needle cannot retract into the rear barrel. | Material/Component Contamination | BD sent Recall notification letters, Urgent Class II Product Recall, to customers and distributors on 12/10/07 through a third party, West Inc., Appleton, Wi. Mailings were sent return receipt requested. Letters requested immediate inventory search and product return. Contact number for questions was given as 1-800-631-0174 |  | 105600 | 105,200 USA; 400 Canada | Worldwide-USA and Canada | Terminated | September 08, 2008 | 293 | 876100 | Not_Computer
Z-0591-2008 | 46149 | No | NWX |  |  | Cardiovascular | PMA | AngioSculpt PTCA Scoring Balloon Catheter, Catalog (REF) Number: 2027-2510, size 2.5mm x 10 mm, AngioScore, Inc., Fremont, CA. |  |  | November 13, 2007 | 2007 | AngioScore Inc. | Mis-labeling: Product labeling on product pouch indicates wrong size, the correct size is on the product carton | Process control | The firm initiated a field correction and issued Medical Device Field correction letters to its consignees on 11/09/2007, instructing consignees to review their inventory and return any mislabeled product or contact the firm. |  | 40 | 40 units | Nationwide-to 30 medical facilities in NC, SC, OH, IL, NY, PA, NJ, PA, LA, KS, FL, WI, MI, WA, and CA. | Terminated | March 25, 2008 | 133 | 40 | Not_Computer
Z-0115-2008 | 45855 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Straight Magnum Inserter Handle, a preparatory instrument used for acetabular component insertion; REF # or Catalog Number: S313141. (systems affected by this action include M2a-Magnum, Regenerex, Acetabular Revision Shell and M2a-38), Biomet Orthopedics, Inc., Warsaw, IN. 46581 |  |  | September 04, 2007 | 2007 | Biomet, Inc. | Binding: The instrument was manufactured incorrectly and its use may result in binding to the acetabular component and cause the implant to not release following impaction. | Other | Consignees were notified by Urgent Medical Device Correction Notice (letter) dated 9/4/07 to discontinue use of the instruments and to return them for evaluation. |  | 22 | 22 | Nationwide; including states of Alabama, Arizona, Arkansas, Georgia, Illinois, Indiana, Kentucky, Massachusetts, Michigan, Nevada, North Carolina, Ohio and South Dakota. | Terminated | September 22, 2009 | 749 | 22 | Not_Computer
Z-0532-2008 | 37989 | No | KRL | Cardiopulmonary bypass bubble detector. | Cardiovascular | Cardiovascular | 510(k) | Terumo Ultrasonic air sensor, 1/4" by 3/32" (6.4 mm x 2.4 mm), Gold, for use on Terumo perfusion systems 1, 8000 and 9000; Catalog No: 5791, Terumo Cardiovascular Systems, Ann Arbor, MI 48103 |  |  | September 19, 2007 | 2007 | Terumo Cardiovascular Systems Corp | False Alarm: The air sensor may malfunction and trigger false alarms, and may continue to alarm thus preventing the device from being reset. | Process control | Consignees were notified of the defect/issue and what to do if it were to occur (short term work around) via an Urgent Medical Device Removal letter dated 9/14/07. As the permanent fix users were informed the sensor will be replaced when replacements become available. |  | 175 | 175 | Worldwide: USA, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, Costa Rica, Dominican Republic, Equador, Germany, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Latin America, Malaysia, Mexico, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, United Arab Emirates, Uruguay, Venezuela and Vietnam. | Terminated | July 07, 2009 | 657 | 2826 | Not_Computer
Z-0258-2008 | 38312 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive Surgical EndoWrist, 8 mm, PK Dissecting Forceps for use with the da Vinci and da Vinci S Surgical systems, Part Numbers: 400214-02, 400214-03, 420214-02 and 420214-03, Intuitive Surgical, Inc., Sunnyvale, CA |  |  | April 13, 2007 | 2007 | Intuitive Surgical, Inc. | Mislabeling-electrical isolation requirements: devices were incorrectly labeled with a CF symbol not their proper BF Symbol on the instrument housing. (All instruments meeting the CF isolation requirements are suitable for direct cardiac application, while instruments meeting BF isolation requirements are not suitable for direct cardiac application) | Device Design | On January 9, 2008, via an Urgent Device Recall Letter the recalling firm notifed users of the defect (mislabeling with heart symbol) and to Immediately Discontinue any direct cardiac applications. The recalling firm will arrange with each consignee a time to relabel their devices. Customers on April 17, 2007 had earlier been warned of the device defect via letter but no correction was indicated. |  | 1136 | 1,136 units | Worldwide Distribution --- USA including states of MN, PA, AK, NY, CA, MD, IN, AZ, CO, AR, TX, TN, FL, VA, MA, NE, NC, IL, OH, ME, WI, MD, NJ, MN, MI, VA, IA, OR, MO, NC, MS, AL, WA, GA, KY, LA, AK, AR, SD, SC, DE, NV, HI, NH, WV, and CT, and countries of Canada, Australia, Belgium, Spain, Hong Kong, Italy, Switzerland, and Sweden. ( 192 USA hospitals, 3 foreign hospitals and 5 foreign distributors) | Terminated | August 01, 2008 | 476 | 1136 | Not_Computer
Z-0539-2008 | 45496 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Benchmark LT, Slide Staining System, Catalog Number: N750-BMKLT-FS, Ventana Medical Systems Inc, Tucson, AZ |  |  | September 29, 2006 | 2006 | Ventana Medical Systems Inc | Leaking Carboys/Safety Updates:The Ventana Staining Platform includes one or more carboys that may have the potential to leak if the spigot is either improperly installed or not fully closed. | Device Design | The firm has mailed three customer letters, a) the first letter dated 9/30/2005 describes the upgraded Voila CD that contains safety information; b) a second letter, dated 6/27/2006, identified as a "Customer Alert" was mailed describing an upgrade to the software system version 9.3/10.3 and the CD that includes safety improvements, c) the third letter dated 9/29/2006 and identified as a Safety Alert advises that as part of their effort to continuously improve laboratory safety, they are updating the labeling and instructions for use regarding installation and cleaning for the carboy as the Ventana staining platform includes one or more carboys that may have the potential to leak if the spigot is either improperly installed or not fully closed. Enclosed are pictogram labels that indicate "slip hazard" and "refer to instructions for use." The customer is instructed to apply these two labels to the carboys above the spigot as a reminder to laboratory personnel to keep the spigot in the full "off" position when not in use and to refer to the instructions for use for maintenance. In addition, the letter advises that there is an electrical shock warning for the power strip that is part of their system to indicate that it should be mounted off of the floor. This power strip is intended to connect the computer components of the staining system and should not be placed on the floor where fluid might contact it. Also included is a warning label for shock hazard to apply to their power strip. |  | 1712 | USA - 1,712 units; Foreign - 1,159 units (total for recalls-Z-0537-0541-2008) | Worldwide; USA, Australia, China, Canada, Germany, Great Britain, France, Italy, Hungary, Singapore, Korea, Taiwan, Spain, Poland, Greece, Austria, Sweden, Japan, Israel, Denmark, Turkey, Finland, Saudi Arabia, The Netherlands, Norway, Slovenia, Ireland, and Belgium | Terminated | June 06, 2013 | 2442 | 1712 | Computer
Z-0127-2008 | 45395 | No | LOM |  |  | Microbiology | PMA | Abbott AxSYM AUSAB Reagent Pack, List No: 3C74-20, 100 tests, an in-vitro test for Antibody to Hepatitis B Surface Antigen (Anti-HBs); Abbott Laboratories, Abbott Park, IL 60064 |  |  | October 01, 2007 | 2007 | Abbott Laboratories | Incorrect results: Reagents in these lots were manufactured with an incorrect ratio of 2 key components, which could result in both false reactive and false nonreactive specimens. | Process change control | Abbott sent a Product Recall letter, dated 10/01/07, to all AxSYM AUSAB customers that received the affected lots of product, informing them of the manufacturing error that resulted in both false reactive and false nonreactive specimens. The letter provided the assay bias for both lots. The accounts were requested to examine their inventory for the affected lots; destroy any packs found; record the number of packs destroyed on the Customer Reply Form; and fax the completed form back to Abbott at 1-800-777-0051 prior to 10/12/07. They were also requested to evaluate patient results generated with these lots in alignment with the information provided in the letter to determine if additional action is necessary. They were requested to forward a copy of the letter to any laboratory to whom they may have forwarded the lots; and to retain a copy of the letter for their records. Any questions were directed to Abbott Customer Support at 1-877-4ABBOTT. |  | 1722 | 1,722 packs | Nationwide Distribution, including Puerto Rico. | Terminated | May 09, 2008 | 221 | 1722 | Not_Computer
Z-0542-2008 | 45510 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VNUS 7F ClosureFAST Catheter, Order #: CF7-7-60 (60 cm working length), Manufactured by VNUS Medical Technologies, Inc., 5799 Fontanoso Way, San Jose, CA 95138 |  |  | August 08, 2007 | 2007 | VNUS Medical Technologies Inc | Instructions for Use change: Instructions for Use have been changed to include a modified placement recommendation. | Labeling design | The Recalling firm is updating the Instructions For Use, training materials, and procedural guides to incorporate the treatment recommendations. The recalling firm sent an Update to customers who received the product, Important Update Regarding the Use of the VNUS ClosureFAST Catheter, dated August 6, 2007 with supplemental device instructions. Future product distribution will include the updated labeling. |  | 16716 | 16,716 units | Worldwide | Terminated | May 22, 2008 | 288 | 22937 | Not_Computer
Z-0569-2008 | 45951 | No | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE Unitek Stainless Steel Permanent Molar Crowns , Product Code: 900321, Made in U.S.A by 3M ESPE Dental Products St. Paul, MN 55144-1000. |  |  | October 17, 2007 | 2007 | 3M Company / Medical Division | Mislabeled: Unitek Stainless Steel Permanent Molar Crowns were placed in packaging which incorrectly identified the product as Unitek Stainless Steel Primary Molar Crowns. | Packaging process control | 3M ESPE Dental Products initially contacted Distributors and Doctors on 10/22/07 through telephone calls, and sent a Device Recall letter via Fed Ex - Overnight on 10/22/07.The letter informed the users that Unitek Stainless Steel Permanent Crowns Lot # P060928, Product Code 900321 may be mislabeled as Unitek Stainless Steel Primary Molar Crowns. Users were asked to stop using and/or distributing the product and contact ESPE Dental Products at 1-800-237-1650 to initiate product return and replacement. |  | 975 | 975 | Worldwide Distribution USA including states of TX, IN, FL, NV, KY, TN, MN, WA, NJ, PA, IL, CA, and IA, and country of CANADA. | Terminated | October 17, 2008 | 366 | 975 | Not_Computer
Z-0592-2008 | 46154 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Hoana LifeBed Patient Vigilance System, (Intelligent Medical Vigilance System),Model Number: 010135-01, Product distributed by Hoana Medical, Inc., 828 Fort St. Mall, Suite 620, |  |  | November 07, 2007 | 2007 | Hoana Medical | Bed Exit Failure: Failure to recognize "Bed Exit" when the feature is active due to a software anomaly. | Software design | On 11/7/07, all consignees were notified via an untitled letter and on-site visit by the firm's sales representative, informing them of the affected product and providing instructions on the recall. The recalling firm will arrange for installation of a software fix at each consignee. |  | 216 | 216 units | Nationwide including states of HI, CA, TX and MA | Terminated | February 01, 2008 | 86 | 216 | Software | Device Operation | Software update | Software Update
Z-0675-2008 | 46250 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Organogenesis Apligraf, interactive wound and burn dressing, Organogenesis, Inc., Canton, Massachusetts 02021 |  |  | December 26, 2007 | 2007 | Organogenesis, Inc. | Contamination: Apligraf units were reported to have contamination in the agarose nutrient medium of some units retained at Organogenesis. Preliminary tests show a gram negative rod (bacillus) organism. | Material/Component Contamination | On 12/26/07, Organogenesis began to telephone the consignees to inform them of the recall. The firm sent a follow-up fax on 12/31/07 to all treating physicians who applied the product requesting patient status information. The Dear Doctor letter, dated December 27, 2007, was titled, "Subject: Potential contaminated Apligraf Units manufactured by Organogenesis, Inc." Consignees were instructed to return unused product, and monitor patients who received the Apligraf for adverse events. |  | 177 | 177 units | Nationwide Distribution. | Terminated | May 15, 2015 | 2697 | 177 | Not_Computer
Z-0435-2008 | 45478 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Depuy TK2 Compression Hip Screw trauma plate, Short Barrel, 135 DEG, 4H-92.6mm, sterile; REF 8315-35-104, DePuy Orthopaedics, Inc., Warsaw, IN 46581 |  |  | October 30, 2007 | 2007 | Depuy Orthopaedics, Inc. | Mis-etched/labeled as to barrel length: the lot with the short barrel was etched and labeled as a standard barrel and the lot with the standard barrel was etched and labeled as a short barrel. | Process control | Consignees were notified via Urgent Information, Recall Notice(letter) dated 10/30/07 to determine whether they have product on hand and to contact their sales representative to return the product. Dear Doctor letters will be issued to implanting physicians as they are identified. |  | 4 | 4 | Nationwide including the states of Florida, Massachusetts, Ohio, Pennsylvania, South Carolina, Tennessee and Virginia. | Terminated | April 21, 2008 | 174 | 13 | Not_Computer
Z-0546-2008 | 45809 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | ProTime Microcoagulation System Instrument, Catalogue Number: PROTIMEPRO, International Technidyne Corporation (ITC), Piscataway, NJ 08854 |  |  | November 19, 2007 | 2007 | International Technidyne Corporation | An increase in the frequency of non-conformance reports (NCMR) for certain displayed error messages. | Device Design | ITC sent to distributors, self test users and health care professionals on 11/19/07, an Urgent Medical Device Recall Letter requesting return of the devices. |  | 3482 | 3,482 units | Worldwide-USA and countries of China, Taiwan, Greece, Ireland, Chile, Finland, Italy, Spain, Norway, Germany, Switzerland., Singapore, Korea, Indonesia, Lichtenstein, Israel, Saudi Arabia, Latvia, Kuwait, and Slovenia. | Terminated | August 12, 2008 | 267 | 5910 | Not_Computer
Z-0549-2008 | 45871 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | ACAT 1 Series Intra-Aortic Balloon Pumps, Catalog # IAP-0100, Arrow International, Everett, MA 02149 |  |  | November 12, 2007 | 2007 | Arrow International Inc | Leak may occur in the helium drive system for the IABP | Material/Component Contamination | Arrow International notified customers by letter dated November, 2007 titled: Urgent Medical Device Field Correction. Customers are instructed to follow the Recommended Action included in the letter while awaiting replacement of the defective component to be performed by service representative. |  | 47 | 47 units | Worldwide- USA, including the states of CA, CO, FL, IN, KS, MA, MO, NV, NY, PA, SC, TN, TX, and WA and countries of Canada, Africa, Argentina, Austria, Bangladesh, Belarus, Belgium, Brazil, Canada, Columbia, Czech Rep, UK, Finland, France, Germany, Hungary, India, Indonesia, Italy, Japan, Korea, Mexico, Netherlands, New Zealand, Panama, Philippines, Portugal, Russia, Saudi Arabia, Italy, Spain, Sweden, Taiwan, and Turkey. | Terminated | January 11, 2010 | 791 | 272 | Not_Computer
Z-0551-2008 | 45873 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott, ACCELERATOR Device Manager (ADM); List Number: 08H74-01; Abbott Laboratories, Abbott Park, IL |  |  | October 25, 2007 | 2007 | Abbott Laboratories, Inc | Sample/Patient Mis-identification: Two issues identified related to the ACCELERATOR Decision Manager (ADM), to include: 1) Sample Identification (SID) and/or Patient Identification (PID) numbers that contain more than 12 characters are truncated to the final 12 characters by the ADM; 2) for SIDs using alpha characters that are case sensitive, SIDs with lower case characters are transmitted by the | Software design | Firm notified consignees via Product Correction, Immediate Action Required letter(s) on 10/25/07. Consignees were notified of software issues and informed they would be contacted by an Abbott representative to schedule installation of software upgrade. |  | 19 | 19 units. | Nationwide:Distributed to hospitals/laboratories in IL, PA, TN, TX, UT, VA and WV. | Terminated | April 01, 2008 | 159 | 19 | Software | Output/Calculation | Software update | Software Update
Z-0552-2008 | 45888 | No | IWX | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | AADCO Rayshield Frame Mounted Overhead X-Ray Barrier; |  |  | November 01, 2007 | 2007 | Aadco Medical Inc. | Screws may become loose from the mounted shield compromising stability | Device Design | AADCO Medical Inc., issued a Product Safety Notification/recall letter to Hospital Administrators, Managers of Radiology/Cardiology and Radiologists and Cardiologists dated 11/1/07. AADCO will perform on site inspections with device tightening or replacement of parts as needed; at no cost the the users |  | 324 | 324 units | Worldwide; USA, Canada, Mexico, Brazil, Venezula, Chile, Bermuda, Spain, and Panama | Terminated | May 02, 2012 | 1644 | 324 | Not_Computer
Z-0553-2008 | 45891 | No | KEY | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Symphony Staining System (High Volume Stainer), KPA, Model Number 2200000, Ventana Medical Systems Inc |  |  | October 30, 2007 | 2007 | Ventana Medical Systems Inc | Fluid Leaks-The T fitting located in the Symphony Clear system of the Symphony High Volume Slide Stainer needs to be larger to securely hold the tubing in place to eliminate any possible fluid leak due to the tubing movement. | Component design/selection | The end user customers were verbally contacted starting in November 2007 to set up a scheduled visit to replace the fitting. |  | 55 | 55 | Worldwide-USA, France and Japan | Terminated | April 14, 2008 | 167 | 55 | Not_Computer
Z-0477-2008 | 45485 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Blood Collection Reservoir, 40 micron,Catalog/Model # EL240, Medtronic, Inc., Minneapolis, MN 55432-5804 |  |  | October 01, 2007 | 2007 | Medtronic Perfusion Systems | Unnaproved Component Resin: Affinity NT Integrated CVR Membrane Oxygenators and EL 240 Autotransfusion Blood Collection Reservoirs contain components made from an unapproved resin | Nonconforming Material/Component | Medtronic sales force was notified and initiated calls with cutomers on 10/1/07. Additionally, certified Urgent Medical Device Recall letters dated 10/5/2007, were sent to the 32 customers. The letter explained the situaiton and requested that the impacted product be returned for replacement or credit. Medtronic is assessing patient use risks and will report those when determined. Standard post-surgical monitoring is advised at the present by Medtronic . |  | 413 | 413 | Nationwide including the states of IL, CT, OK, OH, MN, FL, NY, KS, MO, LA, PA, IN, CA, MI, KY, SC, NE, and AR. | Terminated | October 23, 2008 | 388 | 447 | Not_Computer
Z-0522-2008 | 45819 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | life+cel, Replacement battery for Cardiac Science FirstSave Survivalink; OEM P/N:9141-001, Part #: 2L561, Lithium Sulfur Dioxide, Voltage 12V, 7.5 amps., Mfg for Exocomm Technology Group; Mfg by Nordix, Inc. |  |  | September 03, 2007 | 2007 | Exocomm Tech Grp | Marketed without a 510K | PMA | The original recall communication was sent out on 9/3/07 by Phoenix Solutions, a consultant to Battery Zone, as a Dear Valued Customer notice. A second recall communication, Urgent Device Recall Letter, was issued by Exxocom Technology Group dated 10/25, was sent out to all customer (even those who had returned product) on 11/5/07. Both communication were sent by first class mail, requesting return of batteries |  | 266 | 266 units | Worldwide-Batteries were distributed to dealers or distributors in CT, GA, MN, NC, NY, OH, OK, OR, PA and TN plus one Canadian Consignee | Terminated | July 01, 2013 | 2128 | 312 | Not_Computer
Z-0527-2008 | 45856 | Yes | KPF | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter RenalSoft v.2.0 Patient Management Software Suite - HD Module; clinical data management software; Baxter Healthcare Corporation, McGaw Park, Illinois 60085 |  |  | November 14, 2007 | 2007 | Baxter Healthcare Renal Div | Pre-Treatment/Post-Treatment Report errors: 1) The Heparin Bolus Value is displayed on Pre and Post Treatment Reports for patients with "No Heparin" orders (RenalSoft v2.0, Renal Software Suite v.3.0 and v.3.1), and 2) Withheld Medications are incorrectly displayed as Administered on the Pre-Treatment report (RenalSoft v.2.0 only). | Software design | Baxter Healthcare Corporation sent an Urgent Device Correction letter, dated 11/14/07, to all of the affected customers, informing them of the Heparin Bolus and the Withheld Medication issues and provided them with interim mitigation instructions that they should implement while waiting for a software correction. A Baxter representative will contact the customer to schedule the software update when the upgraded software is released. Customers were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff, and indicating if they still have RenalSoft software and require an upgrade. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, option 4.A follow-up letter dated 3/6/08 was sent to the consignees on the same date, informing them that the upgraded RenalSoft Software v.2.0.1 was now available for installation. The consignees were requested to complete and fax back the customer reply form indicating that they received the letter, and were informed that they would be contacted by a Baxter representative to schedule an installation of the software upgrade at no charge. |  | 25 | 25 units | Worldwide Distribution-USA including Puerto Rico, and countries of Australia, Austria, Canada, Chile, Denmark, Finland, Greece, Ireland, Italy, Mexico, The Netherlands, New Zealand, Norway, Panama, Portugal, Spain, Sweden and the United Kingdom | Terminated | February 03, 2010 | 812 | 34 | Software | Display/Image | Software update | Software Update
Z-0478-2008 | 45513 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Hallmark Anterior Cervical Plate System: Primary Bone Screw, Semi Constrained, REF # 62-3140,14mm; Blackstone Medical, 90 Brookdale Drive, Springfield, MA 01104 (The Hallmark Anterior Cervical Plate Systems consists of an assortment of implantable titanium alloy plates and bone screws that are provided non-sterile) |  |  | May 08, 2006 | 2006 | Blackstone Medical, Inc. | Incomplete Design Validation: the laser marking on the head of the screw lacked complete process qualification | Device Design | On May 8, 2006, a communication to the field was sent via phone, e-mail and fax. Instruction was made to retrieve and return all product in the field. |  | 1991 | 1,991 semi-constrained primary and rescue screws. | Nationwide | Terminated | November 06, 2008 | 913 | 1991 | Not_Computer
Z-0554-2008 | 45519 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Real-Time Position Management System (RPM), A Respiratory Gating System; Version 1.7.3, All Models; Varian Medical Systems, Palo Alto, CA 94304 |  |  | September 26, 2007 | 2007 | Varian Medical Systems Inc | Erroneously Resetting: RPM System version 1.7.3, software failure can occur that affects the gated radiotherapy treatment when phase-based gating is used. RPM 1.7 erroneously resets the phase-based gating thresholds to system default during the treatment session, if the user accesses the Session options during the treatment session. This can lead to difference between the planned and delivered | Software design | Varian sent an Urgent Medical Device Correction letter on September 26, 2007, to all its direct consignees, informing them that the RPM version 1.7 may reset the phase thresholds if Sessions options are accessed during treatment. A Service Representative will be scheduling a visit to user site to perform the modification requested to address the problem, when a corrected software version is available. |  | 59 | 59 units | Worldwide including countries of: Ireland, Italy, Spain, UK, Denmark, Finland, France, and Canada | Terminated | September 30, 2009 | 735 | 59 | Software | Device Operation | Software update | Software Update
Z-0565-2008 | 45869 | No | GDI | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ethicon Endopath 5 mm Curved Dissector with Monopolar Cautery, Product Code: 5DCD, Ethicon Endo-Surgery, Inc, Cincinatti, OH 45242 |  |  | November 06, 2007 | 2007 | Ethicon Endo-Surgery Inc | Sterility (package integrity) compromised; Damaged packaging may have compromised the sterility of the devices. | Process control | Endopath sent consignees an Urgent Device Recall notification letter on or about 11/15/07, informing them that the possibility exists that damage to the packaging of this product may have compromised the sterility of the product. The firms were told to examine their inventory, and remove all affected product. All consignees were instructed to complete the Business Reply Form; even if they did not have any product. Recalled products are to be returned to Ethicon for a credit. |  | 2844 | 2,844 | Worldwide, USA including states of AL, AR, AZ, CA, CO, DE, FL, GA, IL, KS, KY, LA, MA, MN, MO, MS, NC, NJ, NV, NY, OH, PA, RI, TN, TX, AND WA, and countries of Austria, Belgium, Canada, China, Colombia, Czech Republic, Ecuador, Egypt, Greece, Hong Kong, Italy, Japan, Korea, Mexico, Portugal, Puerto Rico, Republic of Panama, South Africa, Spain, Sweden, Taiwan, Turkey, United Kingdom, and Venezuela. | Terminated | December 09, 2009 | 764 | 2844 | Not_Computer
Z-0568-2008 | 45958 | No | LYJ |  |  | Neurology | PMA | Cyberonics, VNS Therapy System Model 250 Handheld programming system using Dell x5 handheld and preprogrammed with software v7.1 and higher; manufactured by Cyberonics, Inc., Houston, TX. |  |  | October 25, 2007 | 2007 | Cyberonics, Inc | Screen Freezes-- The Dell X5 Handheld PC screen will freeze caused due to incompatibility between the Microsoft 2002 OS and the model Dell X5 handheld computer. Once frozen, the handheld device becomes non-responsive to user input. | Software design | Cyberonics sent an initial notification (Important Safety Alert) to physicians via Certified Mail on 10/01/07. Letter notified physicians that the Dell X5 handheld computer screen will freeze following an interrogation. Once frozen the handheld device becomes non-responsive to user inputs. The firm informed the user that in the event that the computer becomes non-responsive, the user has two options, 1) Reinitialized the flash memory and, 2) Perform a software reset of the handheld computer. A second notification addressing the approval and addition of labeling was sent to physicians on 11/15/07. Physicians were asked to respond to both letters via an enclosed response card, acknowledging receipt of the notification. Additional information will be posted at www.VNStherapy.com. |  | 4201 | 4,201 (US) 199 (International) | Worldwide Distribution- USA and the countries of; Austria, Belgium, Denmark, France, Germany, Ile de la Reunion, Ireland, Israel, Italy, Japan, Marocco, Mexico, Netherlands, Ireland, Norway, Poland, Portugal, S. Korea, Slovenia, Sweden, Switzerland, Turkey, United Kingdom, and Venezuela. | Terminated | October 05, 2009 | 711 | 4201 | Software | Display/Image | Software update | Software Update
Z-0708-2008 | 46142 | No | HCI | Aneurysm clip applier. | Neurology | Neurology | 510(k) | Caspar Titanium Aneurysm Clip Appliers, Catalog No/Part #: FT420T, Casper Applier F/Mini Ti-Clipsstr 146 mm, Aesculap, Inc., Center Valley, PA. 18034 |  |  | November 01, 2007 | 2007 | Aesculap Inc | Breakage: the inner rod of the applier may corrode and break | Device Design | The recalling firm issued an Import Recall Notice, on or about December 18, 2007, to all their customers to inform them of the problem and the need to return the devices . |  | 60 | 60 clips | Worldwide- USA hospitals in AZ, CA, CO, IL, KY, MI, MS, ND, NM, NY, OH, OK, PA, TX, VA, WA, and WI, to sales representatives in NC, OK, and MO and shipped to distributors in CA and Canada. | Terminated | March 28, 2008 | 148 | 144 | Not_Computer
Z-0124-2008 | 44824 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | Philips PageWriter Touch Cardiograph, |  |  | September 12, 2007 | 2007 | Philips Medical Systems | Incorrect display: When using the 16 lead Patient Interface Module (PIM), lead tracings of V3R and V4R are reversed in the extended lead modes of Pediatric and Balanced. | Software change control | Philips Medical Systems notified affected customers by Customer Letter and Urgent - Device Correction Notice dated 9/12/07. The correction will be a software upgrade. In the US, the Philips Response Center will contact US customers and arrange their upgrades, either via a software download, or by sending a Field Service Engineer to the customer site to perform the upgrade. Until new software is installed, customers are being advised to use the 12 lead module, and to disable or disconnect the 16 lead connection. |  | 266 | 266 units | Worldwide Distribution - USA, Canada, Australia, Hungary, New Zealand, Nigeria, Norway, Qatar, Taiwan, and UAE. | Open, Classified |  |  | 266 | Software | Display/Image | Software update | Software Update
Z-0828-2008 | 46182 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Sierra Pre-Filled, Heparin Lock Flush Solution USP, 100 units/ml in 0.9% Sodium Chloride 120/Case, Catalog #1003-02, NDC #64054-1003-02, 5 mL Fill in a 12mL Syringe, For Single Use only, Sterile Solution Pathway, For IV Flush, Sierra Pre-Filled Inc., 455 W. Depot St. Angier, NC 27501 |  |  | December 11, 2007 | 2007 | AM2 PAT INC | Lack of assurance of sterility by manufacturer | Process design | AM2PAT contacted their distributors by phone on or about January 17, 2008 and instructed them to stop distribution of their pre-filled syringes and to contact their customers to recall all AM2PAT manufactured lots and all sizes of prefilled, Heparin Lock Flush syringes and Normal Saline IV Flush syringes (B Braun and Sierra Brands) to the consumer level. A press release was issued on January 17, 2008. The firm had eariler on 12/20/2007 recalled a single lot of Heparin syringes, the January extension was to all lots/all sizes of its Heparin and Saline syringes. |  |  |  | Nationwide | Terminated | December 02, 2010 | 1087 |  | Not_Computer
Z-0659-2008 | 38315 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System 5 mm instrument cannula, Part #: 420011-01, Intuitive Surgical, Inc., Sunnvale, Ca. 94086 |  |  | March 09, 2007 | 2007 | Intuitive Surgical, Inc. | Devices may have a ridge on the side of the cannula, which has the potential to abrade instrument shafts and generate black particulate matter. | Nonconforming Material/Component | The recalling firm issued an Important Device Recall Notification on 3/9/2007, announcing the recall and exchange program for device replacement. A toll free number of 1-800-876-1310 was given in the recall notice for questions. |  | 11 | 11 units | Worldwide: 92 hospitals nationwide and abroad, as well as 6 distributors abroad, USA including the states of AL, AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, LA, MA, MD, MI, MN, MO, NC, NH, NJ, NV, NY, OH, OR, PA, SD, TN, TX, VA, WA, WI, and WV and countries of Belgium, China, France, Germany, India, Italy, Japan, Mexico, The Netherlands, Spain and Sweden. | Terminated | April 19, 2011 | 1502 | 896 | Not_Computer
Z-0462-2008 | 44934 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Philips Medical Systems, GEMINI GXL 6, 10, &amp; 16 slice PET/CT Systems, diagnostic imaging systems for fixed or mobile installations that combine Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems, Philips Medical Systems, Cleveland, OH 44143 |  |  | September 21, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | Artifacts: Artifacts may appear in some images from head scans on the CT subsystem of the GEMINI PET/CT system. Reading images with these artifacts may lead to misdiagnosis or improper treatment. | Software design | Philips Medical notified domestic customers via Product Safety Notification dated May 30, 2007. The product safety notification includes a Technical Bulletin notifying customers of the issue and how to avoid it. This information will be used until the software correction becomes available. The software correction will be released in October 2008. Once software is available, the firm will begin correcting the field. |  | 126 | 126 | Worldwide Distribution - USA, Australia, Belgium, Canada, China, Denmark, England, Finland, France, Germany, Hong Kong, Hungary, India, Ireland, Italy, Japan, Jordan, Korea, Netherlands, Poland, Portugal, Saudi Arabia, Slovenia, Spain, Sweden, Taiwan, Turkey, and Venezuela. | Terminated | March 02, 2012 | 1624 | 187 | Software | Display/Image | Software update | Software Update
Z-0660-2008 | 44965 | No | CBR | Nitrous oxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Spacelabls Medical, Ultraview Multigas Analyzer, Model 91518,Spacelabs Healthcare, Issaquah, WA, 98027 |  |  | October 08, 2007 | 2007 | Spacelabs Healthcare, Incorporated | Inaccurate partial pressure readings; readings are not compensated for ambient pressure in mmHg or kPa. | Software design | The firm sent out their Urgent Medical Device correction/recall letters on 10/5/2007, and an updated versions on 11/30/07 and 12/7/2007, forwarding an attached set of work around directions for use and associated tables for calculations. This short term fix will be updated to a permanent fix when a new software revision is released and installed at no cost to each user as noted in the recall notification. Telephone contact was provided at 1-800-522-7025. |  | 722 | 722 devices | Worldwide | Terminated | April 12, 2011 | 1282 | 722 | Software | Output/Calculation | Software update | Software Update
Z-0690-2008 | 45845 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 2100IQ Versatile Cardiac and Vascular Imaging System, GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | November 15, 2007 | 2007 | GE Healthcare | Frozen Image Problems: During an acquisition (fluoro and/or record), there have been cases reported where an image became "frozen" on the DL (digital leader acquisition system) live monitor screen. In such cases, the system continued to send out X-rays without reporting an error message. The result was that the live imaging screen displayed an older "frozen" image until the operator released the | Device Design | Customers were contacted by means of an Urgent Medical Device Correction Letter on November 9, 2007. The letter described the product and the safety issue. The letter also made short term recommendations and described a long term solution, which is pending as an expected hardware and/or software modification that will be made at no cost. Short Term Recommendation: Take extra care when working with motionless anatomy or devices where there is little evident motion to warn you of the frozen image occurrence. Always completely shut down and restart the system (not just reset) immediately after such an error occurs. Also, please immediately report this issue to your local GE Healthcare field service engineer. |  | 378 | 378 | Worldwide-USA including states of AR, CA, CO, FL, GA, IL, IN, LA, MA, MD, MO, NC, NJ, NM, NV, SC, TN, TX, WA, WI, AK, AL, AZ, ,CT, DE, HI, ID, KS, KY, ME, MI, MN, MS, MT, ND, NE, NH, NJ, NY, OH, OK, OR, PA, SD, UT, VA, WV, and Puerto Rico and countries of Hungary, Korea, Algeria, Argentina, Australia, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Columbia, Czech Republic, Dominican Republic, Finland, France, Georgia, Germany, Great Britain, Greece, Hong Kong, India, Indonesia, Ireland, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Mexico, Morocco, New Zealand, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Singapore, Spain, Sudan, Sweden, Switzerland, Syria, Taiwan, Turkey, and Venezuela. | Terminated | January 29, 2016 | 2997 | 869 | Software | Display/Image | Software update | Software Update
Z-0520-2008 | 45820 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vital Images ViTALConnect , Version 4.0 and 4.1, PET/CT Visualization option. Vital Images, Inc., 5850 Opus Parkway, Suite 300, Minnetonka, MN 55343. ( A web-based medical diagnostic tool that allows physicians to use PCs or notebook computers to gain remote access to 2D, 3D, and 4D advanced visualization. The software enables users to measure, rotate, and analyze images) |  |  | November 08, 2007 | 2007 | Vital Images, Inc. | Standard Uptake Values (SUV) display/update problems; in the way Standard Uptake Values (SUV) are displayed in 3D/MPR views. The problem arises when a user draws a selected region of interest in order to calculate the SUV statistics of the region. If the user then changes the slice shown by scrolling through additional slices, the original SUV measurements will remain on the screen and do not up | Process control | Customers were notified via telephone (11/7 -11/8/2007) and issued an Urgent Medical Device recall notification letter on 11/8/07. The letter described the problem and product involved. The letter also informed customers they they will provide a free software patch within the next 6-8 weeks. |  | 18 | 18 | Worldwide-USA including states of SC, MO, MI, WI, NY, FL, CO, CA, PA, and WA, and countries of England and The Netherlands | Terminated | December 17, 2011 | 1500 | 18 | Software | Display/Image | Software Patch | Software Update
Z-0285-2008 | 45495 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Si emans Evansi, with patient table, Part No: 44 81 586, and without patient table Part No: 44-81-669; Siemens Medical Solutions USA, Concord, CA 94520 |  |  | February 27, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Loose Mounting Screws -- Cassette holder may become loose and result in the device dislodging and colliding with the patient. | Labeling design | Customers were sent a Customer Safety Advisory Notice on February 27, 2007, notifying them that the mounting screws on the cassette holders on the SIMVIEW 3000 and NT Simulators, as well as some MEVASIM Simulators, could loosen and can cause damage to the system, or personal injury. The firm requested that they check of tightness on the mounting hardware. A STOP Modification notice was issued on 06/01/2007, to all affected regional service organizations that are responsible for implementing updates due to new information. Updated letters and User instructions will be released in August 2007. |  | 287 | 287 units total, all recalls Z-0281-285-2008 | Worldwide Distribution - USA including states of IN, VA, MA, VT, NJ, NY, PA, DC, NC, SC, GA, AL, TN, MI, KY, OH, IN MI, MN, ND, MT, IL, MO, KS, NE, AR, OK, TX, AZ, CA, SD, LA, Puerto Rico, and countries of Argentina, Australia, Austria, Bangladesh, Belarus, Belgium, Brazil, Bulgaria, Canada, Colombia, Costa Rica, Croatia, Denmark, Ecuador, Egypt, El Salvador, Estonia, France, Germany, Greece, Hungary, India, Iran, Ireland, Italy, Kuwait, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Norway, China, Israel/Palestine, Philippines, Poland, Portugal, Qatar, Republic of Korea, Russia, Singapore, South Africa, Spain, Thailand, Turkey, Ukraine, UK, and Vietnam. | Terminated | February 03, 2011 | 1437 | 287 | Not_Computer
Z-0695-2008 | 45889 | No | CBT | Arterial blood sampling kit. | Anesthesiology | Anesthesiology | Enforcement Discretion | Smiths Medical, Add-On Kids Kit , Blood Sampling System, Product Code: MX4037, Smiths Medical ASD, Inc, Dublin, Ohio 43016 |  |  | November 21, 2007 | 2007 | Smiths Medical Asd Inc | Misbranded; The product codes are labeled incorrectly with a green promotional label for "New split septum sampling site. Only use needle-free access devices .......", however this product code contains a sampling site which is accessed with a shrouded needle, NOT with a needle-free access device | Packaging process control | Smiths Medical mailed certified letters (Urgent Product Advisory Notifications) dated November 19, 2007, to all customers that received the affected lot. Firms are instructed to make the field correction (extra label removal) or mail the product back to Smiths for the correction. |  | 1440 | 1440 units | Nationwide including the states of CA, DC, IL, IN, MD, MN, NC, NE, SC, and TN and Puerto Rico. | Terminated | September 30, 2010 | 1044 | 1690 | Not_Computer
Z-0891-2008 | 45857 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Argon Medical Devices, ArgoGuide Hydrophilic NiTi Guidewire, K011905, Catalog Number: 00-AG387324, .038/300cm, stiff/angled/standard, sold as a single hydrophilic guidewire in both sterile and bulk non-sterile packaging configurations. Argon Medical Devices, Inc. Athens, TX 75751 |  |  | November 19, 2007 | 2007 | Argon Medical Devices, Inc | Degradation -- Hydrophilic guidewires may exhibit degradation of blue Pebax cladding causing particulate to dislodge potentially resulting in embolism. | Process design | Argon Medical sent an Urgent Product Recall letter to consignees on 11/15/2007. They requested that users immediately quarantine and return any product remaining at their facility, utilizing instruction on response card. Distributors were directed to issue the recall to their customers and return all retrieved product to Argon. |  | 13 | 13 | Worldwide Distribution -- USA including states of FL, KS, NV, IL, CA, NY, CO, VA, MD, NC, UT, GA, MI, NJ, TN, MA, and CT, and countries of The Netherlands, Germany, and Turkey. | Terminated | June 05, 2008 | 199 | 2951 | Not_Computer
Z-0935-2008 | 45903 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Impulse Angiographic Catheter, Sterile EO. Catheters are single lumen catheters with a single proximal hub for attachment to a dye delivery system. Catalog No. 16599-301. 6F MULTI PIT 124 (Single Pack). Distal Curve FL4, FR4, PIG 145. 100 and 110 cm Length. Made in Mexico: Parque industrial Le Mesa, Tijuana, Baja California 22050 MEX. Boston Scientific Corporation, Natick, MA 01760-1537 USA. |  |  | November 20, 2007 | 2007 | Boston Scientific | Flash may be protruding from the lumen of the catheter shaft. If this flash is present in the catheter, there is the potential of an embolization resulting from the flash detaching from the device during the procedure. | Release of Material/Component prior to receiving test results | The firm sent an Urgent Medical Device Recall letter dated 11/20/2007, to all affected hospitals. The letter describes the defect with the catheters and Product Table information including UPNs, Catalog Numbers, Lot/Batch Numbers and Product Descriptions. The firm requests that further distribution or use of any remaining product affected by the recall should cease immediately. ****IMPORTANT INFORMATION: If it is a practice of the hospital to remove product from the outer carton and store on shelves in the inner-pouch only. Users are to review the Product Table carefully and consider both the inner and outer packaging product codes when search for affected/recalled product.**** |  | 211 | 211 single units | Nationwide Distribution including states of AL, CA, FL, GA, KS, KY, LA, MN, NV, OH, OK, PA, SD, TN, TX, and WI. | Terminated | October 23, 2008 | 338 | 211 | Not_Computer
Z-0567-2008 | 45957 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Datascope CS 100 Intra-Aortic Balloon Pump and Datascope |  |  | November 19, 2007 | 2007 | Datascope Corp | Defective Circuit Board -- A defect in the Printed Circuit Board's (PCB) in the CS 100 and CS300 Intra-Aortic Balloon Pump may cause the pump to exhibit intermittent malfunctions, which include failure to start-up or reset of the display screen during therapy. | Other | Datascope sent a Product Field Correction notification letters to Hospital Administrators, via certified mail, on November 19, 2007. The letter informed them of a defect in the circuit board. It described the possible malfunction and provided instructions should malfunctioning occur. Datascope will be replacing the defective circuit boards in all affected IABP units. A Service Representative will be contacting the hospital to arrange for the replacement of the defective board. |  | 70 | 70 units US | Worldwide Distribution | Terminated | May 20, 2008 | 183 | 70 | Hardware | Device Operation | Replace board | Remove or Replace
Z-0706-2008 | 46141 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet CC Cruciate Tibial Tray, 67 mm.; Product No. 141232, Biomet, Inc., Warsaw, IN 46581 |  |  | December 12, 2007 | 2007 | Biomet, Inc. | Mislabeled as to size: Package labeled as containing a 67 mm tray actually contains a 75 mm tray and vice versa. | Packaging process control | Consignees were notified via an Urgent Medical Device Recall Notice dated 12/12/07 to cease use of the product and to return it. |  | 8 | 8 | Worldwide: USA including states of Georgia, Idaho, Indiana, Missouri, Ohio, and Virginia and country of Turkey. | Terminated | June 11, 2008 | 182 | 8 | Not_Computer
Z-0669-2008 | 46160 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive brand da Vinci S, 5 mm Instrument Cannula |  |  | November 19, 2007 | 2007 | Intuitive Surgical, Inc. | Sharp Edges: the 5mm Cannula may have a sharp edge on the inner diameter of the Cannula. The defective cannulae may cause particulate shavings to be skive (scraping) from the instrument shafts during surgery. | Other | On 11/19/07, the firm issued Urgent Device Recall Notification letters via Federal Express to all its direct consignees, informing them of the affected product and providing instructions on the recall. |  | 89 | 89 units | Worldwide: USA including states of AZ, CA, CT, DC, GA, lA, MI, MN, NC, NJ, NV, OH, PA, SD, TX, and VA and country of Hong Kong | Terminated | August 27, 2008 | 282 | 89 | Not_Computer
Z-0280-2008 | 45454 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CIC Pro Clinical Information Center, Central Station, Software Version 4.1. (CIC Pro v4.1). GE Medical Systems Information Technologies, Inc., Milwaukee, Wisconsin (System hardware consists of a processing unit and a monitor) |  |  | January 24, 2007 | 2007 | General Electric Medical Systems Information Technology | Two separate issues may occur with operating The CiC Pro v4.1 System with Telemetry. 1) Experience a loss of audible alarms when operating system memory resource reaches approximately 170MG. This will occur if CIC Pro continuously runs for average of 160 day, and 2) System may continuously reboot upon start up in an attempt to correct the issue. | Software design | GE Healthcare sent a certified letter titled "Urgent Medical Device Correction" to customers on 3/13/07. As a Short Term Solution, GE recommends that user reboot of system according to instructions and precautions in Service Manual at least once every three months. Also, take the necessary precautions to ensure that system is shut down properly. This will decrease the likelihood of a continuous reboot. If system continues to reboot, GE suggest that user contact Technical Support at 1-800-588-7044 (Press Option 2) and ask for Technical Support. As for a Long Term Solution, GE will provide a software update that will rectify the issue as soon as an upgrade is available. |  | 6217 | 6,217 Units (1,549 OUS, 4,668 within US) | Worldwide Distribution - USA (excluding DE, HI, ID, MT, ND, SD, VT), Peurto Rico, Washington DC and including the countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Greece, Hungaria, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Malaysia, Malta, Netherland, New Zealand, Norway, Pakistan, Philippines, Poland, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom | Terminated | May 16, 2012 | 1939 | 6217 | Software | Alarm/Message | Software update | Software Update
Z-0699-2008 | 45966 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Assure 3 Blood Glucose Test Strips, For in vitro Diagnostic use. Product Number: 555050, Control Solution Range (mg/dL). Manufactured for Hypoguard, Minneapolis, MN 55439 USA. Made in Taiwan. |  |  | November 16, 2007 | 2007 | ARKRAY USA INC. | Mislabeling: incorrect control range-The Level 1 control solution range printed on the 50-count test strip bottles is incorrect. The range printed on the bottle is 62-54 when it should be 62-94. | Process design | On Friday, November 14, 2007 the recalling firm notified their distributors via telephone and email to inform them of the device defect and replaced all devices still in inventory. The distributors were also instructed to inform all of their customers of the issue and offer replacement product. Additionally, replacement will be offered to any end user that calls customer service about the issue. To reach ARKRAY use 1-800-818-8877 extension 3121 for replacement product. |  | 4920 | 4,920 | Nationwide including the states of CA, FL, GA, IL, IN, MA, MI, MS, NC, NJ, OH, PA, and TX | Terminated | October 30, 2008 | 349 | 4920 | Not_Computer
Z-0738-2008 | 46292 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Optiflux 180NR Advanced Fresenius Polysulfone, Catalog No.: 0500318E, Fresenius Medical Care North America, Waltham, MA 02451. (Hemodialyzer) |  |  | December 03, 2007 | 2007 | Fresenius Medical Care North America | Leaking: the Dialyzer may leak at the header resulting in small amounts of blood loss. | Process control | Consignees were notified by telephone and certified letter , Urgent Device Recall, on 12/03/2007. They were told to discontinue use and to return all unused affected dialyzers to Fresenius. |  | 2791 | 2,791 cases (33,492 dialyzers) | Nationwide including the states of AZ, CA, HI, IA, ID, IL, IN, KS, KY, MI, MN, MO, NE, NV, NY, OH, OR, PA, WA, and WI. | Terminated | February 08, 2008 | 67 | 2791 | Not_Computer
Z-0739-2008 | 36595 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable Pump, Model 8626-10. Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA.-- (The implantable Medtronic SynchroMed EL Programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site) |  |  | August 03, 2007 | 2007 | Medtronic Neuromodulation | Pump motor stall due to gear shaft wear | Device Design | Beginning August 3, 2007, consignees were notified by letter. The letter "Medtronic Urgent: Medical Device Correction Aug 2007" described the problem and the Affected Devices as well as Patient Risk, Patient Management Recommendations, Next Steps and Physician and Patient Support. |  | 23 | 23, 895 total for recalls Z-0739-0746-2008 | Worldwide: USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and countries of HONG KONG, AUSTRALIA, CANADA, AUSTRIA, BELGIUM, CROATIA, The CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HUNGARY, IRAN, ICELAND, IRELAND, ITALY, THE NETHERLANDS, NORWAY, POTUGAL, SAN MARINO, SAUDI ARABIA, SPAIN, SWEDEN, SWITERLAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, IRAQ, BRAZIL, and MEXICO. | Terminated | June 06, 2011 | 1403 | 23 | Not_Computer
Z-0727-2008 | 45940 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic Cardioblate Gemini-s, model 49260. Surgical Ablation Device. Sterilized using Ethylene Oxide. Two, unique jaw curvatures are provided: a standard curve (model 49260) and extra curve (Model 49261). The device is intended for intermittent operation. Manufacturer: Medtronic, Inc. Minneapolis, MN 55432-5604. |  |  | November 21, 2007 | 2007 | Medtronic Cardiac Surgery Technologies | Mechanism failure: Medtronic has identified an issue with four lots of the Cardioblate Gemini-s Surgical Ablation Devices in which the polycarbonate distal coil retainer can fail during the course of a procedure, potentially resulting in an inability to apply the necessary jaw closure force. | Nonconforming Material/Component | Certified letters were sent on 11/21/07 to Consignees. The letter identifies the issue with the device and requests that customers return the device back to the firm. Customers will be requested to complete a Field Correction Action certificate acknowledging the receipt of the information. For additional information you can contact a Medtronic Sales Representative at 763-391-9167. |  | 28 | 28 units | Nationwide Distribution including states of FL, IL, IN, LA, MI, PA, TN, TX, UT, and WI | Terminated | December 17, 2011 | 1487 | 28 | Not_Computer
Z-0557-2008 | 45862 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | FRAZIER Surgical Suction Instrument, 7 French with Control Vent and Obturator (Bend with Obturator in Place), Product Code: 0031070, 50/C. STERILE, EO, Rx ONLY. CONMED CORPORATION 310 Broad Street, Utica, NY 13501 |  |  | September 12, 2007 | 2007 | Conmed Corporation | Sterility (package integrity) compromised -- Frazier and Poole suction instruments may have an inadequate seal. The product pouch seal made during assembly of this product, may be incomplete or not present. | Packaging process control | URGENT-Device Recall Letters (dated 9/10/07) and FIELD CORRECTIVE ACTION CERTIFICATION forms were sent via Certified Mail to affected customers the week of 9/09/07 (mailed on 9/12/07). The recall letters informed the consignees of the recall due to potentially inadequate seals of the instruments, and instructed them to review their inventory of the instruments checking them for inadequate seals. The letters also instructed the consignees to complete the form and fax it back by 9/21/07. For those consignees that have any of the recalled instruments or may have questions, please contact ConMed Quality Assurance at 1-800-552-0138 ext. 5225 between the hours of 7 am - 6 pm (MST). International customers were notified on 9/18/07. A second recall letter (dated 10/09/07) was sent to non-respondents of the initial mailing. ConMed plans to send a third recall letter to the non-responders during the week of 12/10/07 which addresses sub-recall notifications to the distributors' customers. |  | 53691 | 53,691 instruments | Nationwide and Worldwide Distribution | Terminated | August 30, 2012 | 1814 | 4700864 | Not_Computer
Z-0948-2008 | 39366 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | Functional Anaesthetic Discography (F.A.D.) Catheter |  |  | April 25, 2006 | 2006 | Kyphon Inc | Guidewire Breakages: during the Functional Anaesthetic Discography (F.A.D.) procedure. | Labeling design | A customer advisory letter was sent on 4/25/06 to consignees. In some cases some were issued in person. Follow up with field representatives was also conducted. Both contacts involved providing IMPORTANT CHANGES IN INSTRUCTIONS FOR USE that are to be attached to the on site users documents. The recalling firm is updating future distribution with the new instructions for use. |  | 23 | By Lot Number : J 6013102 = 147, Lot: J6033012 = 23 | Nationwide including states of AL, AZ, CA, FL, IN, NH, NY, DC, ME, MD, WI, and PA. | Terminated | June 10, 2008 | 777 | 66 | Not_Computer
Z-0693-2008 | 45847 | No | MHY |  |  | Neurology | PMA | Medtronic Kinetra 7428, Dual Program Neurostimulator for Deep Brain Stimulation, Rx only, Medtronic, Inc., Minneapolis, MN 55432-5604 USA. (The Medtronic Kinetra Model 7428 Neurostimulator is a multiprogrammable device and a component of the Kinetra Neurostimulator System for deep brain stimulation.) |  |  | October 01, 2007 | 2007 | Medtronic Neuromodulation | Cessation of Therapy: separation of internal connections (lifted wirebonds) between the electronic circuit and battery may lead to sudden cessation of therapy. | Packaging change control | A letter "Medtronic Urgent: Device Recall" dated October 2007 was sent to Healthcare Professionals. The letter described the device defect, product involved, root cause, probability of occurrence, recommendations and also provided information for assistance. At the time of the recall all of the product remaining in date for use (not expired) had been implanted, therefore Medtronic provided Important Patient Management Information for the implanting physicians to use in patient care and decisions regarding patient explants. |  | 1427 | 1,427 | Worldwide: USA including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NM, NY, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, WA, WI, WY, UT, VA, VT, and WV and countries of Australia , Austria , Belgium , Canada , Cyprus , Czech Republic , Denmark , Finland , France , Germany , Greece , Hong Kong , Hungary , Iceland , India , Iran , Israel , Italy , Japan , Latin America, Lebanon , The Netherlands , Norway , Poland , Portugal , Russian Federation , Saudi Arabia , South Africa , Spain , Sweden , Switzerland , Turkey , and The United Kingdom | Terminated | May 03, 2012 | 1676 | 3079 | Not_Computer
Z-0697-2008 | 45936 | No | DJZ | Breath-alcohol test system. | Toxicology | Toxicology | 510(K) Exempt | Maximo Concepts Alcohol Breath Tester, Model Number: MAX91721, Sakar International, Inc., Edison, NJ (LED 3 Step Alcohol Breath Tester) |  |  | November 30, 2007 | 2007 | Sakar International, Inc. | Marketed without a 510(k) | PMA | Recall notification letters , Urgent Device Recall, were sent by first class mail on 11/30/2007 requesting return of product and notification of accounts sold product. |  | 72581 | 72,581 | Nationwide to retail distribution centers in CA, CO, IL, KS, MA, MO, NE, NJ, NY, PA, VA and WI. | Terminated | June 30, 2010 | 943 | 72581 | Not_Computer
Z-0698-2008 | 45949 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Quick-Cross Support Catheters, .035in x 90cm, REF#: 518-036, Sterile, Manufactured by: Spectranetics, Colorado Springs, Colorado 80907. |  |  | December 10, 2007 | 2007 | Spectranetics Corporation | Mislabeling: Catheters were placed into mis-labeled inner pouches as REF Number 518-037, 135 cm--the exterior box labeling is correct as 90 cm. | Packaging process control | Consignees were notified by Dear Healthcare provider letter on 12/10/2007 and were told that a sales rep would remove and replace all affected catheters. |  | 64 | 64 | Nationwide including states of FL, IL, IN, MO, NC, NY, TX, and UT. | Terminated | February 08, 2008 | 60 | 64 | Not_Computer
Z-0710-2008 | 46276 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Horizon Cardiology Hemo Monitoring System;Horizon Cardiology Hemo External Control Unit;~100 - 230 VAC |  |  | November 12, 2007 | 2007 | McKesson Provider Technologies | Power supply reliability: McKesson received word from their manufacturer of the medical grade power supply that some power grade supplies may not provide reliable power output. | Nonconforming Material/Component | Recalled by letter sent to the firm's customers beginning on 11-12-2007, advising them that the field correction would be completed by the recalling firm service technicians. |  | 13 | 13 | Nationwide to hospital and medical centers in LA. CA, PA and IL, | Terminated | May 28, 2008 | 198 | 13 | Battery | Output/Calculation | Correction | Repair
Z-0518-2008 | 45986 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Magnetom Espree with Swiveling OR (Operating Room) Table, Magnetic Resonance Imaging System,Siemens Medical Solutions USA, Inc., Malvern, PA |  |  | November 05, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Table may experience a deadlock situation | Software design | The recalling firm issued a Customer Safety Advisory to affected customers per Update Instructions MR031/07/S. The letter informs customers of the potential issue and provides instructions to avoid its occurrence. The recalling firm has also released the software update to correct the issue via Update Instructions MR034/07/S. Notifications were issued on or about November 5, 2007. |  | 4 | 4 units | The products were shipped to medical facilities in TX and NY. | Terminated | February 05, 2008 | 92 | 4 | Computer
Z-0705-2008 | 45987 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet RINGLOC + Instrumentation Impactor Plate, Size 28; REF/Part#: 31-400628, Biomet Orthopedics, Inc., Warsaw, IN. |  |  | October 18, 2007 | 2007 | Biomet, Inc. | The instrument will not lock onto the impactor handle; the device was distributed without assembly of the Bal Seal component. | Process control | Distributors were notified of the device defect by letter (Urgent Field Correction Notice) dated 10/18/07, and provided with assembly instructions and parts, and instructed to make corrections at consignee locations. |  | 141 | 141 total for all recalls Z-0700-0705-2008 | Worldwide-USA, Germany, The Netherlands, Sweden and The United Kingdom. | Terminated | March 11, 2008 | 145 | 141 | Not_Computer
Z-0135-2008 | 44879 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | THSCF-35-260-3-AUS2, (Cook Vascular Wire Guide, PTFE coated curved Amplatz ultra stiff catheter exchange, Diameter .035 Inch, length 260 cm, curve radius 3 mm, Cook Incorporated, Bloomington, IN 47402 - not on label) |  |  | August 15, 2007 | 2007 | Cook, Inc. | The device is normally sold as sterile, however this lot was not sterilized and the device is not adequately labeled, as the label does not state whether the device is sterile or non-sterile and lacks the manufacturer's name and address. | Process control | Cook Incorporated notified consignees by Urgent Product Recall letter dated 8/15/07 instructing them to return the product. |  | 33 | 33 | Worldwide Distribution - USA, Australia, India, Trinidad and Tobago. | Terminated | April 21, 2008 | 250 | 33 | Not_Computer
Z-0525-2008 | 45918 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Beckman Coulter Flow-Count Fluorospheres - Part Number 7547053, Beckman Coulter, Inc., Fullerton, CA 92835 |  |  | May 21, 2007 | 2007 | Beckman Coulter Inc | Sporadic absolute count recovery failures with assayed control cell products when using Lot Number 754801F of the Flow-Count Fluorospheres. Additionally, this lot has experienced secondary fluorescent populations containing more than 20% of the total population, as stated in the Evidence of Deterioration section of the Package Insert. | Process design | A Product Corrective Action (PCA) letter was mailed on May 21, 2007, to all accounts that have ordered Flow-Count Fluorospheres, Part Number 7547053, Lot # 754801F. The letter advised that Beckman Coulter has confirmed absolute count recovery failures with assayed control cell products. Additionally, internal testing has indicated secondary fluorescent populations slightly above their indicated 20% of the total population as stated in the Evidence of Deterioration section of the Package Insert. The letter requests that users immediately discard any bottles of the product and call for a replacement. If no evidence of deterioration has been observed, the firm states that there should not be any negative impact to the results obtained. A response form is included and should be returned to the firm within 10 days. Questions regarding this recall can be addressed by contacting Beckman Coulter Customer Service at 1-800-526-7694. |  | 134 | 97 units in the US; 37 units Internationally | Worldwide Distribution --- including USA and country of Canada. ***CENTER COMMENTS -- The Center requests that the subject recall not be terminated until CDRH has determined that the firm has identified the root cause and performed an adequate CAPA to address the cause of the problem, i.e., quality control.*** | Terminated | October 15, 2012 | 1974 | 134 | Not_Computer
Z-0747-2008 | 46378 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | CORDIS "Dura Star" 2.25 x 10 Dilatation Catheter, Catalog # 70110225, Distributed by Cordis Corporation, Miami Lakes, FL 33014 |  |  | January 14, 2008 | 2008 | Cordis Corporation | Slow Deflation or No Deflation | Process control | Cordis initated the recall on 1/14/2008. For U.S. customers: An Urgent Medical Device Recall letter and an Acknowledgement Form was sent to the following 4 individuals in each account: Cath Lab Director, Operating Room Director, Materials Director, Risk Manager; to be followed up with direct Cordis representative contact to facilitate obtaining signature, faxing form, collecting and returning units. For non-US customers, email notification with return acknowledgement is used to notify the affiliate distributors, who then notify and reconcile the customers in those countries. All users were instructed to immediately set aside product and return it for replacement. |  | 133300 | 133,300 units total for recalls Z-0747-0824-2008 | Worldwide. | Terminated | October 20, 2009 | 645 | 133352 | Not_Computer
Z-0826-2008 | 45532 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT System for ENT (server and work station), Xoran Technologies, Ann Arbor, MI |  |  | January 04, 2008 | 2008 | Xoran Technologies, Inc. | Misbranding; These devices do not bear the manufacturer's name or address. | Labeling design | Consignees were notified via letter, Urgent Medical Device Correction, dated 1/4/08 of the need to apply proper labeling to the products, and were sent instructions and labeling to affix to their units. |  | 71 | 71 | Nationwide. | Terminated | October 22, 2010 | 1022 | 71 | Not_Computer
Z-0934-2008 | 45407 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Inotech brand Heparin Adsorbant, Catalog #: HA-1, in 10mm x 75mm polystyrene tubes, each containing 70mg of Heparin Adsorbent, 10 tubes per foil lined pouch, 10 foil pouches/box, Inotech Biosystems International, Inc., Rockville, MD 20855 |  |  | August 20, 2007 | 2007 | Inotech Biosystems Intl., Inc | Erroneous Coagulation Tests | Labeling design | The recalling firm notified consignees by phone on 08/20/07 and provided instructions for increasing the speed of centrifugation from 1,500 x g to 2,500 x g and decanting the supernatant aliquot carefully so as to not draw in resin fibers that may be stuck on the test tube walls during sample preparation for testing in the coagulation studies. The firm updated their recall notification on 10/01/07 as an Urgent Recall/Correction to alert healthcare practitioners to not use "lupus sensitive reagents" with the heparin absorbant product and that Platelet Factor 3 and any other negatively charged molecule in the plasma may interact with the resin causing false high readings due to the partial removal of the clotting initiating factors. |  | 37370 | 37,370 tubes | Worldwide; USA to Hospitals and Clinical Laboratories and to one domestic distributor for export to Canada. | Terminated | August 07, 2012 | 1814 | 37370 | Not_Computer
Z-0849-2008 | 45912 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | NexframeÂż Stereotactic System Kits, Model Number(s): DB-1021-MR, DB-1031, DB-1040-BL, DB-1040-ST, DB-1041, |  |  | October 04, 2007 | 2007 | Medtronic Image Guided Neurologics, Inc. | Sterilty (package integrity) Compromised: Some failures were for damage to the outer pouch, while another set of failures were for the seals on this pouch. The seal between the inner tray and lid has not been compromised and the contents remain sterile, however, the sterility of the outer surfaces of the inner tray and lid cannot be assured. | Package design/selection | On 10/4/2007, consignees were notified by Medtronic field personnel, either by personal visits or telephone. All wholesalers of this product are OUS. Medtronic OUS Regulatory personnel communicated this issue directly to the wholesalers, who provided the letters to their customers. Medtronic field personnel are responsible for completing and returning reply cards confirming all notifications (100%). |  | 385 | 385 | Worldwide, USA, Taiwan, Thailand, Japan, Spain, Sweden and The Netherlands. | Terminated | July 13, 2012 | 1744 | 385 | Not_Computer
Z-0125-2008 | 44854 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Medtronic Sofamor Danek COLORADO 2" SPINAL SYSTEM Break-Off Nut, non-sterile, RX; Catalog Number 8634111, Medtronic, Memphis, TN 38138. |  |  | September 10, 2007 | 2007 | Warsaw Orthopedic, Inc. dba Medtronic SOFAMOR DANEK | Assembly may be compromised: Two thread grooves are missing on the nut, which may cause damage to the bone screw and result in the need for the bone screw to be replaced. | Process control | Medtronic Sofamor Danek notified consignees by Urgent Device Recall letter dated 9/10/07 to immediately cease using the product and to return the devices with a Product Recall Questionnaire, recording quantity of devices and any patient injuries. |  | 2569 | 2,569 | Worldwide Distribution - USA, Australia, Hong Kong and the Netherlands. | Terminated | July 14, 2009 | 673 | 2569 | Not_Computer
Z-0844-2008 | 45846 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD Sterilizers ( Models 100S, 50, 200) Compatible Medical Device Reference List for US Customers, Document Numbers 15-53229 &amp; 15-53230; STERRAD NX Sterilizers Compatible Medical Device Reference List for NX Customers, Product Code 15-53231; STERRAD NX System Brochure, Product Code AD-53421-001, Advanced Sterilization Products, Irvine, CA |  |  | November 12, 2007 | 2007 | Advanced Sterilization Products | Outdated Compatibility Lists: Recalling firm has decided to discontinue dissemination of Compatible Medical Device Reference Lists and instrument assessment activities | Other | Notification letters, Urgent Product Correction, were sent to customers beginning on November 13, 2007 via UPS overnight. Customers were notified this action was being taken as a result of an investigation by the firm determined that all brand- and model-specific compatibility lists and associated instrument assessments provided to STERRAD Sterilization System customers are now out-of-date. User were instructed to discontinue their use and destroy them. |  | 9686 | 9,686 units | Worldwide: USA, France, Germany, Italy, The UK, Switzerland, Benelux,Uruguay,Venezuela,Argentina,Chile,Colombia,Ecuador,Belgium, Puerto Rico, Hong Kong,Indonesia,Mexico,Philippines,Australia,Austria,China,Greece,Ireland,Israel, Japan,Korea,Malaysia,Singapore,Taiwan,Thailand,Turkey, Middle East, Peru,Canada,The Czech Republic,Portugal,S Africa,Spain,Sweden,Hungary,Poland,Brazil,Russia,Egypt,India &Slovenia; and Canada | Terminated | December 19, 2011 | 1498 | 9686 | Not_Computer
Z-0967-2008 | 45902 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker LFIT COCR V40 Femoral Head; xx-Long (+16mm) Neck Length 32mm Diameter Sterile, Forged Vitallium Alloy Product Number 6260-9-532 |  |  | November 02, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Fracture of Stem Neck - The firm recently performed laboratory testing indicates that there may be a potential increased risk of stem neck fractures when a V40 +16mm offset femoral head is mated with certain Accolade TMZF femoral stems in patients classified as overweight or obese. | Other | Stryker Orthopaedics sent an Urgent Product Recall notification letters on November 9, 2007, by Federal Express with return receipt. The letter informed users of potential increased risk of stem neck fracture, when a V40 + 16mm offset femoral head is mated with certain Accolade TMZF femoral stems in patients classified as overweight or obese. The firm has asked that inventory be examined, removed and held in quarantine. A Stryker representative will contact the medical facilities to retrieve the product. If a representative has not contacted the facility to date, it is suggested that they contact Stryker at (201) 831-5972. |  | 5599 | 5,599 total units | Nationwide Distribution | Terminated | August 30, 2010 | 1032 | 5599 | Not_Computer
Z-0718-2008 | 38335 | No | JTC | Microtiter diluting and dispensing device. | Microbiology | Microbiology | 510(K) Exempt | Hamilton Sample Carriers for the ML STAR line (SMP-CAR-24). Product Code: 866.2500; Part Number: 173400. Hamilton Company, Reno, NV 89520 (Accessory on the Microlab STAR series instruments.) |  |  | June 25, 2007 | 2007 | Hamilton Co | Incorrectly labeled-- Product has the wrong barcode labeling. The barcode error will not influence the results of testing. The instrument will discover the error and post an error notice. | Error in labeling | The firm sent an Urgent: Product Recall letter to consignees on June 25, 2007. The firm also issued an alert via their User news distribution. They instructed consignees to examine their stock and return any affected carriers. A fax back form was supplied with the notification. The firm will be providing replacements. |  | 179 | 179 units distributed. | Worldwide Distribution --- USA including states of OH, CA, NJ, NC, and countries of Japan, France, Switzerland, Spain, Italy, Germany, and the UK. | Terminated | April 15, 2008 | 295 | 179 | Not_Computer
Z-0676-2008 | 45329 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Philips Medical Systems -- Cardiac Viewer or Pulmonary Viewer software application version 2.0 in use with Precedence SPECT/CT System and all CT EBW workstations using software version 3.5. Philips Nuclear Medicine, Milpitas, CA 95035 |  |  | September 28, 2007 | 2007 | Phillips Nuclear Medicine | Incorrect Measurement: The firm discovered a software defect in Cardiac and Pulmonary Viewer software application options used in the Precedence CT Sub-System. The defect causes incorrect measurements in derived images which are zoomed and saved in a batch file. | Software design | A Customer Advisory Notification was mailed to affected consignees on September 28, 2007. The firm informed consignees that incorrect measurement can occur when the Cardiac Viewer or Pulmonary Viewer is used to zoom and save derived images in batch files. The firm provided instructions for operators to prevent incorrect measurements. Phillips is currently investigating methods to correct this issue and will provide updates, free of charge, when they become available. If you have questions concerning this recall, please contact Phillips Medical at 1-800-722-9377, (option5: Diagnostic Imaging, option 5:NM). |  | 19 | 19 | Worldwide Distribution --- USA including states of FL, NY, VT, MD, TX, and MO and country of Canada. | Terminated | February 04, 2009 | 495 | 19 | Software | Output/Calculation | Software update | Software Update
Z-0847-2008 | 45851 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 200 Sterilizer, Product Code: 10201, Advanced Sterilization Products, Irvine, CA |  |  | November 12, 2007 | 2007 | Advanced Sterilization Products | User Guide Update: based on user reports about contact with residual hydrogen peroxide from instruments, pouches and trays after sterilization and subsequent light colored residue on these devices after sterilization updated user guidance has been issued. | Employee error | Urgent Product Correction letters were sent to customers beginning on November 13, 2007 via UPS overnight. Customers were notified this action was being taken to clarify instructions contained within the STERRAD System User Guides to reinforce appropriate use and to address two separate issues which have arisen from user reports: (1) User reports about contact with residual hydrogen peroxide from instruments, pouches and trays after sterilization. (2) User reports about light-colored residues observed on instruments, trays and pouches after sterilization. The Product Correction Letters instructed customers to refer to the revised excerpt of the User Guide (attached to the letter) for information about instrument preparation prior to sterilization, particularly with regard to the washing, rinsing and thorough drying of instruments prior to sterilization in the STERRAD Sterilization System. If customers have any questions they were instructed to call ASP Customer Care Center to answer additional questions at (888) 783-7723. |  | 9686 | 9,686 units total for recalls Z-0845-0848-2008 | Worldwide: USA, France, Germany, Italy, The UK, Switzerland, Benelux, Uruguay, Venezuela, Argentina, Chile, Colombia, Ecuador, Belgium, Puerto Rico, Hong Kong, Indonesia, Mexico, The Philippines, Switzerland, Australia, Austria, China, Greece, Ireland, Israel, Japan, Korea, Malaysia, Singapore, Taiwan, Thailand, Turkey, Middle East, Peru, Canada, The Czech Republic, Middle East, Portugal, S Africa, Spain, Sweden, Hungary, Poland, Brazil, Russia, Egypt, India &amp; Slovenia | Terminated | March 14, 2012 | 1584 | 9686 | Not_Computer
Z-0688-2008 | 45910 | No | JFJ | Amylase test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche LIPC Lipase Colorimetric Reagent; Catalog Number: 03029590322. Roche Diagnostics Corp., Indianapolis, IN 46250 |  |  | December 03, 2007 | 2007 | Roche Diagnostics Corp. | False Results- The reagent of the LDL-Cholesterol assay shows a carryover effect on the Lipase assay. Results may be falsely elevated by up to 50 units per liter when used on COBAS INTEGRA 400 or COBAS INTEGRA 400+ analyzers. | Process design | Consignees were sent and Urgent Medical Device Correction notice dated 12/03/07. The notice provided instructions for conducting an additional wash cycle whenever both the LDL-Cholesterol assays and Lipases assay are used on the COBAS INTEGRA 400 or COBAS INTEGRA 400+ analyzers. Contact Roche Technical Support at 1-800-428-2336 for questions regarding these instructions. |  | 9817 | 9,817 kits. | Nationwide Distribution | Terminated | June 11, 2008 | 191 | 9817 | Not_Computer
Z-0950-2008 | 45595 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable Pumps, Models 8626-10, Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. (The implantable Medtronic SynchroMed EL Programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site) |  |  | August 03, 2007 | 2007 | Medtronic Neuromodulation | Pump Motor Stall; pumps can stall due to gear shaft wear. | Device Design | Begining August 3, 2007, consignees were notified by letter. The letter "Medtronic Urgent: Medical Device Correction Aug 2007" described the problem and the Affected Devices as well as Patient Risk, Patient Management Recommendations, Next Steps and Physician and Patient Support. |  | 81023 | 81,023 total for recalls Z-0950-0957-2008 | Worldwide-USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, DS, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and countries of HONG KONG, JAPAN, AUSTRALIA, CANADA, BELGIUM, CROATIA, AUSTRIA, CYPRUS, The CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HUNGARY, INDIA, IRAN, ICELAND, IRELAND, ITALY, JORDAN, KWAIT, LEBANON, MALTA, The NETHERLANDS, NORWAY, POLAND, PORTUGAL, SAN MARINO, SAUDI ARABIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, The UNITED KINGDOM, CHILE, VENEZUELA, BRAZIL, and MEXICO. | Terminated | December 27, 2011 | 1607 | 81023 | Not_Computer
Z-0881-2008 | 46322 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | PERFEKTUM TOOMEY EVACUATING SYRINGES, REF/Part # 5047, DESCRIPTION: 30CC LOCK TIP, RX ONLY, NON-STERILE. Manufactured By: Popper &amp; Sons, Inc. New Hyde Park, NY. 11040. |  |  | January 02, 2008 | 2008 | Popper &amp; Sons Inc | Mis-labeling: the product name on the label was erroneously declared as PERFECTUM TOOMEY EVACUATING SYRINGES instead of POPPER INTERCHANGEABLE HYPODERMIC SYRINGES for one lot of the product. Note: the product information (part number, description, lot number) found in the white box portion of the label is listed correctly. | Process control | The recalling firm, Popper &amp; Sons, began contacting its customers by telephone on/about 1/02/08, informing them of the issue. This was followed up with formal recall letters, Urgent Recall, and response forms starting on 1/09/08 via Certified Mail, Return Receipt Requested to the US accounts, and Registered Mail to the two Canadian accounts. The firm has directed the customers to discontinue using or distributing the lot and return any remaining quantity of this lot of syringes. If the lot has been further distributed, the distributors were instructed to immediately contact their accounts, advise them of the recall situation and have them return their outstanding recall stock as indicated above. |  | 272 | Domestic - 259 units; Canada - 13 units | Worldwide-USA including the states of NY, NJ, PA, MI, IL, OH, FL, MO, TX, and CA, and country of Canada | Terminated | May 12, 2008 | 131 | 272 | Not_Computer
Z-0873-2008 | 46267 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RenalPureÂż Liquid Acid Concentrate R-259; One Gallon (3.78 liters); manufactured by Rockwell Medical Technologies, Inc., Grapevine, TX 76051. |  |  | December 03, 2007 | 2007 | Rockwell Medical Technologies, Inc | Foil Seal Deterioration; occuring on the tamper resistant foil seals on dialysis liquid concentrate and Dri-Sate product containers older than two years old. | Process control | Firm began notification of consignees by letter, Alert: Voluntary Recall, on 12/03/07, sent via e-mail, fax, and/or mail. Consignees asked to check inventory and to identify and call for the replacement of any product older than 2 years old. |  | 6 | 6 cases (4 one gallon containers/case). | Nationwide Distribution to the following states: AL, AR, FL, KS, LA, MO, MS, OK, TN and TX. | Terminated | March 31, 2008 | 119 | 49 | Not_Computer
Z-1040-2008 | 37183 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Anspach MicroMax Pneumatic Motor System; ENT Surgical Drill, including handpiece and Accessories Attached. All Models; Black Motors; Anspach Effort, Inc., Palm Beach FL 33410. |  |  | December 18, 2006 | 2006 | The Anspach Effort, Inc. | Incorrect Sterilization Process --The Operating Manual for the product contains an incorrect statement regarding the sterilization of the Autolube foot control. The manual states that the product can withstand ETO sterilization, but there is no data to support this process. | Labeling False and Misleading | Anspach issued an Urgent Device Correction letter on 11/27/06, to all manual holders advising them of the error and requested that they update the manual with the corrected instructions. A reply receipt confirming notification was included. |  | 7399 | 7,399 total units | Worldwide Distribution | Terminated | July 06, 2012 | 2027 | 7399 | Not_Computer
Z-0880-2008 | 46304 | No | GZT | Self-retaining retractor for neurosurgery. | Neurology | Neurology | 510(k) | Comfort T-Handle Hudson Connector with Impactor Cap; Part #1006-1; handle is component of the following consignment surgical kits: Harmony 2700-0001-PL, Harmony 2700-0009-PL, Harmony 2900-0001-PL, Harmony 2900-0002-PL, Cadence 1500-0002-PL and Cadence 1500-0008-PL; Distributed by Abbott Spine, Austin, TX 78727. |  |  | December 19, 2007 | 2007 | Abbott Spine | Dislodged Collar Spring; The collar spring binding on the shaft of generic handle used with multiple surgical instruments has the potential to release the device's internal bearings during use. | Device Design | All medical suppliers with consigned inventory from affected lots were contacted via phone beginning 12/19/07. Consignees notified of potential issue and were asked to inspect their inventory and contact Abbott Spine for return and replacement of the affected handles. |  | 230 | 230 handles. | Worldwide:Medical suppliers in the following states: AR, AZ, CA, CT, FL, IL, LA, MD, MI, MO, MN, NC, NJ, NM, NV, OH, OK, OR, PA, TX, VA and WI. and country of Spain. | Terminated | October 04, 2010 | 1020 | 230 | Not_Computer
Z-0879-2008 | 46288 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Smiths CADD Yellow Medication Cassette Reservoirs with Flow Stop, Clamp and Female Luer, 100 ml, Reorder No. 21-7300-24, Nonvented stopper included, Sterile, Smiths Medical, St. Paul, MN 55112 |  |  | December 14, 2007 | 2007 | Smiths Medical MD, Inc. | Leakage associated with the CADD Medication Cassette Reservoirs for use with the CADD Ambulatory Infusion Pumps | Process control | Consignees were mailed the recall letters on December 21,2007. The recalled letter "Smiths Mecical Urgent Product Recall Notification" makes them aware of the potential issue. Customers will be requiired to return the affected devices to Smiths Medical. |  | 4644 | 4,644 | Worldwide Distribution --- USA including states of AL, AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, GU, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND,NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY, and countries of AN, AR, AU, BR, CA, CL, CN, CZ, DK, EE, ES, FI, FR, GB, GR, HK, HU, IE, IL, IT, JO, JP, KW, LB, MY, NL, NO, PA, PL, PT, SA, SE, SG, TAIWAN, TH, VI, ZA, CH, AT, and BE. | Terminated | January 11, 2011 | 1124 | 1266792 | Not_Computer
Z-0112-2008 | 39371 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Blood glucose meter, marketed with multiple brand names including: Precision Xtra New and Improved meter, Precision Xceed meter, Optium Xceed meter, Optium meter, ReliOn Ultima meter, Rite Aid meter, Kroger meter, Omron HEA-214 meter, Optium Xido meter, Boots meter. |  |  | August 02, 2007 | 2007 | Abbott Diabetes Care, Inc. | Damage could lead to no display. Meters manufactured after January 31, 2007 could exhibit meter display damage if dropped on a hard surface. These meters could exhibit unreadable lot number fields and date/time fields, in addition to complete blanking of the numerical reading portion of the display. | Component design/selection | Abbott sent Urgent Medical Device Correction letters, dated 31Aug07, to registered users/customers via FedEx or mail. The letter instructs users to verify that their display is correct, and if not, to discontinue use and call Abbott Customer Care at 1-877-844-4404 for a new meter. Warranty kits will have fliers added to them to alert users. Email and telephone contacts may be made as needed. U.S. consignees and registered users will be contacted by letter, telephone call or visit starting 23Aug07. Worldwide, customer lists will be created and communications tracked according to local requirements. Abbott issued a Press Release August 31, 2007 also. |  | 279289377839 | 882,601 total meters distributed, 316,439 in the U.S. and 566,162 outside of the U.S. including Mexico and Canada | Worldwide Distribution - USA, Canada, Mexico, Australia, Malaysia, Philippines, Germany, New Zealand, China, India, Singapore, Hong Kong, Taiwan, Argentina, Brazil, Chile, Colombia, Costa Rica, Ecuador, Mexico, Peru, Puerto Rico, Uruguay, Venezuela, and Japan. | Terminated | March 28, 2011 | 1334 | 279289377839 | Not_Computer
Z-0619-2008 | 45469 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Allogenix; DBM in a Lipid Carrier with Calcium Granules; Model Numbers: 02-3110, Interpore Cross International, a Biomet company (Allogenix DBM products contain human tissue (allograft bone) and are intended for transplantation) |  |  | October 15, 2007 | 2007 | Interpore Cross International Inc | Unapproved Testing: Interpore Cross International was recently informed that one of its human tissue suppliers, LifeLink Tissue Bank ("LifeLink"), utilized unapproved diagnostic testing (HBsAg &amp; HBc) on certain lots of distributed donor tissue instead of the FDA approved screening test. As a result, Interpore Cross International has initiated a recall on specific lots of products containing human | Nonconforming Material/Component | The Recall Notification letters will be sent via commercial carrier (UPS) with signature for verification of receipt required on Friday October 26, 2007, to all consignees, including hospitals and sales representatives, which summarizes the recall details and instructs the consignee to contact Interpore Cross via mail or facsimile with the enclosed Customer Response form, and to immediately return the affected product. |  | 34 | 34 | Nationwide | Terminated | September 25, 2008 | 346 | 1290 | Not_Computer
Z-0677-2008 | 46157 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Mevatron KD2 Medical Charged-Particle Radiation Therapy System; (Interlock #51 Jaw Field Position) Model Number:08515520, Siemens Medical Solutions, Concord, CA 94520 |  |  | August 07, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Incorrect field size -- Products manufactured prior to 1999 may experience jaw field size errors during interlock resulting in potential mistreatment if a light field check is not performed. This problem is present in machines with serial numbers less than 3094, as they do not contain redundant sensors for jaw position. | Device Design | The firm has sent consignees a Customer Safety Advisory Notice on August 7, 2007 informing them that a misplacement of the X and/or Y jaws can occur if interlock #51 Jaw Field Position is asserted and then cleared with the left-hand key rest. The firm recommends that if an interlock issue is encountered tha treatment of the patient be discontined until this issue is resolved. |  | 55 | 55 units | Worldwide Distribution | Terminated | December 01, 2010 | 1212 | 642 | Not_Computer
Z-0545-2008 | 45520 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Clinac, Accelerator, Linear, Medical charged-particle radiation therapy system; |  |  | September 10, 2007 | 2007 | Varian Medical Systems Inc | The collimator drive chain may break or slip off of its drive track, allowing the collimator to rotate freely without motor control; if undetected resulting in a treatment with the wrong collimator angle. | Labeling design | On 9/10/07, the firm issued an Urgent Medical Device Correction letter to all its direct consignees, informing them of the device defect and providing instructions on the recall. Short term user instructions were provided and the firm will schedule follow-up visits to make the needed permanent device corrections at each device site. |  | 418 | 418 (total for Recalls Z-0544-0545-2008) | Worldwide- Canada, Mexico, and the USA | Terminated | November 27, 2010 | 1174 | 418 | Not_Computer
Z-0825-2008 | 45502 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | OsteoStim Cervical Allograft (Lordotic Cervical Spacer) |  |  | October 23, 2007 | 2007 | EBI, L.P. | Unapproved Testing: Tissue supplier (LifeLink) utilized unapproved diagnostic testing on certain lots of distributed donor tissue instead of the FDA approved donor screening test. (Note: tissue samples were later re-tested by an approved method and found non-reactive) | Nonconforming Material/Component | Urgent Recall Notice letters were mailed on October 23, 2007, requesting location of all product including sub-accounts and return for credit. |  | 19 | 19 units | Nationwide to distributors/customers in NY, NC, MO, CA, FL and KS | Terminated | April 15, 2008 | 175 | 19 | Not_Computer
Z-0720-2008 | 45527 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand VARiS 1.4g (Medical Charged Particle Radiation Therapy System, Record and Verify System); Linear Accelerator with RTP Exchange v6.2, v6.6, &amp; v8.0, All Models; Varian Medical Systems, Palo Alto, CA |  |  | October 02, 2007 | 2007 | Varian Medical Systems Inc | Incorrect Software Validation- The ARC treatment plans containing a Dose Dynamic MLC for IMRT delivery that will not be recognized properly by the system. Under these circumstances, an arc plan can be administered to a patient whereby the gantry will not rotate, resulting in delivery to a single location rather than through the gantry arc. | Software design | An Urgent Medical Device Correction letter was distributed October 2, 2007, to all direct consignees, informing them that the ARC treatment plans containing a Dose Dynamic MLC for IMRT delivery will not be recognized properly by the System. The firm provided corrective instructions and informed consignees that a product modification will be developed to resolve the issue. A Varian Service Representative will be scheduling a visit to upgrade the application. |  |  |  | Worldwide Distribution | Terminated | May 17, 2011 | 1323 |  | Software | Device Operation | Software update | Software Update
Z-0721-2008 | 45530 | No | JTR | Quality control kit for culture media. | Microbiology | Microbiology | 510(K) Exempt | PML -- LyfoCults Haemophilus parainfluenzae - micro organism. Product label includes LyfoCults Haemophilus parainfluenzae ATCC 7901 Lot No.: 233851-1 Exp 05-17-09. PML Micrcobiologicals, 27120 SW 95th Avenue, Wilsonville, Oregon 97070 |  |  | October 15, 2007 | 2007 | PML, Inc. dba Pml Microbiologicals Inc. | Incorrect micro-organism-- Quality control In-vitro diagnostic test was manufactured with Cryptococcus neoformans (ATCC 76484) instead of Haemophilus parainfluenzae (ATCC 7901). | Mixed-up of materials/components | On 10/15/2007, firm notified consignees by telephone. A follow up letter dated 10/26/07, was faxed to customers instructing them to discontinue use and discard remaining inventory on site as a biohazardous material. Consignees were requested to fax back response indicating the amount of product they had on hand at time of notification. |  | 150 | 30 packs (each pack contains 5 vials of Haemophilus parainfluenzae) | Nationwide Distribution | Terminated | April 07, 2011 | 1270 | 151 | Not_Computer
Z-1107-2004 | 45411 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemans- ACUSON Aspen Diagnostic Ultrasound System, general purpose Diagnostic Ultrasound system, Model Number: 8247887, Siemans, Mountain View, CA 94039 |  |  | September 18, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Software error: Error affects all Aspen customer systems using the PAL video standard configuration and have a Perspective PC installed. This can lead to a measurement error under this specific condition: On secondary capture images/clips stored from within any Perspective advanced display option (specifically: 3D Surface Rendering, 3D MultiPlanar Reconstruction, Color Doppler 3D MultiPlanar Reco | Software design | Siemens issued an Urgent Medical Device Notification to consignees 9/18/07, under Siemens update program US014/07/S. (Distribution on this product had ceased on 6/29/2005.) The firm has instructed the user to not take measurements or use measurements on these secondary captures under the conditions described in the notice, to place a copy of the instructions for use in the file, and to notify all users of the issue with the software. The Siemens Service organization has coordinated the implementation of the field correction with the distributors for all countries. Non-responding consignees will be visited or called by telephone for follow up. |  | 280 | 280 units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Austria, Belgium, Brazil, China, Czech Republic, Denmark, Egypt, France, Germany, Great Britain (UK), Greece, Hong Kong, India, Italy, Jordan, Malaysia, Norway, Poland, Portugal, Romania, Russian Federation, Sweden, Switzerland, Thailand, and Turkey. | Terminated | April 21, 2011 | 1311 | 280 | Software | Output/Calculation | Correction | Repair
Z-0724-2008 | 45815 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova - EZ CHEM Creatinine Meter: SRS Versions 5.1 and 5.2 |  |  | November 02, 2007 | 2007 | Nova Biomedical Corporation | Incorrect Results - Creatinine results were lower than the laboratory reference method. | Device Design | Nova Biomedical issued an Urgent: Medical Device Recall notification letter to consignees dated November 19, 2007. The letter addressed the issues with the product(s) and requested that if users had any remaining stock that it must be returned. The firm also provided a Customer Response form to be returned. Questions concerning this recall will be addressed by Gina Vogelsberg at 1-800-544-4624 or by email at gvogelsberg@ezem.com. |  | 133 | 133 units | Nationwide Distribution, including the state of NY | Terminated | December 09, 2008 | 403 | 719 | Not_Computer
Z-0728-2008 | 45953 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Invictus-- Bionime Rightest Blood Glucose Monitoring System |  |  | November 26, 2007 | 2007 | Invictus Scientific Inc | Abnormal Memory Recall: When the meter is fixed in 12 hour mode, the time frame display of all memory data will be shown incorrectly as AM or PM depending on when the user is recalling the memory stored in the meter. This will cause confusion as to when the last test was taken. | Software design | An Urgent: Device Recall Notification letter were sent to customers (direct accounts) beginning on November 20, 2007, via US Postal Service, with a return receipt requested. The distributors were instructed to inform consumers/users to fix the time format to the 24 hrs mode and to contact the distributors or the firm for replacement, if they prefer to use the time format in the 12 hrs mode. They were also instructed that if they have further distributed any of the recalled products, to please immediately contact those accounts and advise them of the recall situation, and have them return their outstanding recalled stocks. If customers have any questions, they were instructed to contact at 858-481-8485. |  | 21070 | 21,070 | Nationwide Distribution | Terminated | October 13, 2011 | 1417 | 21070 | Hardware | Display/Image | Change time format/Replace | Remove or Replace
Z-0850-2008 | 45991 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Primus Linear Accelerator System, PRIMART MX Linear Accelerator System; PRIMUS Mid, Part Number:1940035 and PRIMART MX, Part Number: 5500371, Siemens Medical Solutions USA, Inc., Concord, CA |  |  | July 27, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Door Hinge failures; loose or falling off (Stationary Structure Doors) | Component design/selection | Stationary Door Hinges will be replaced on each device by the recalling firm's technician visits. |  | 1284 | 1,284 total for recalls Z-0850-0852-2008 | Worldwide- USA including the states of Missouri, Georgia, Delaware, Texas, Ohio, Alabama, Kentucky, Florida, California, North Carolina, New Jersey, New York, Illinois, Pennsylvania, Colorado, Tennessee, Arizona, Pennsylvania, Virginia, New Jersey, Michigan, Wisconsin, Indiana, Nebraska, Montana, Kansas, Mississippi, Rhode Island,Washington, South Carolina, Maryland, New Hampshire, Maine, Arkansas, Idaho, Missouri, Louisiana, Oklahoma, North Dakota, Minnesota, Massachusetts, and The District of Columbia and countries of Canada and Mexico. | Terminated | April 29, 2011 | 1372 | 1284 | Not_Computer
Z-0853-2008 | 46198 | No | LRR |  |  | General &amp; Plastic Surgery | Enforcement Discretion | First Aid Kits, soft canvas, plastic, and metal containers holding various products to include ibuprofen and other first aid supplies, in various sizes. Kits include: |  |  | December 10, 2007 | 2007 | First Aid Only Inc | Failed USP Impurity testing: the Ibuprofen supplied by its own manufacturer, used in these first aid kits, tablet refills, and cabinets, is being recalled for failing its USP limits for impurities during stability testing. | Nonconforming Material/Component | The firm issued an Urgent Device Recall letter dated 10/10/07 to consignees receiving any of the products. The letter instructs consignees to return any recalled product to the firm. If any recalled prdoucts were further distributed, the consignee was requested to notify those consignees and have product returned. |  | 72904 | 72,904 kits | Worldwide: USA and to one distributor in Dubai, United Arab Emirates. | Terminated | April 08, 2011 | 1215 | 137211 | Not_Computer
Z-0079-2008 | 38311 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive Surgical, daVinci S Surgical System, Model Number: IS2000, A5.1 P1 and A5.0 P5 software level, Surgical Endoscopic Instrument Control System, Intuitive Surgical, Sunnyvale, CA 94086. |  |  | March 30, 2007 | 2007 | Intuitive Surgical, Inc. | System Lock-Up: Software anomalies could cause product failure during use, or on start-up. System transitions to a safe "soft-lock" state. | Software design | Intuitive Field Service Engineers installed a software upgrade beginning on 03/30/2007, and left a notification letter that the upgrade had been performed at each site. New software level is A5.1 P2. |  | 159 | 159 | Worldwide Distribution - USA, Belgium, China, France, India, Italy, Japan, Korea, Spain, Sweden, Switzerland, UK and Mexico. | Terminated | September 05, 2008 | 525 | 159 | Software | Device Operation | Software update | Software Update
Z-0151-2008 | 45518 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive brand da Vinci S Surgical System, Endoscopic Instrument Control System, Model IS1200 with Isolation Transformer (P/N: 952012) and Vision Cart (Model VS1000), Distributed and Manufactured by: Intuitive Surgical, Inc. 950 Kifer Road Sunnyvale, CA |  |  | October 04, 2007 | 2007 | Intuitive Surgical, Inc. | Under-Rated Fuses--The Firm may have installed incorrect fuses. The daVinci Surgical System Vision Cart's isolation transformer requires a 10A fuse for 115V operation or a 5A fuse for 230V operation. Operation of this product with the use of an under-rated fuse will result in failure of the fuse and a loss of power to the vision cart, including the loss of power to any ancillary equipment that m | Mixed-up of materials/components | A product recall was initiated on October 4, 2007. Direct and Distributor consignees were sent an Important Product Notice, via electronic mail on October 5, 2007. The firm plans on sending a Field Service Engineer to inspect the system and make any necessary corrections. |  | 63 | 63 total units | Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, IL, IN, KS, LA, MA, MD, MN, NC, NJ, NV, OH, PA, PR, SD, TN, TX, WA, and WI, and countries of Belgium, Ireland, Italy, Korea, Russia, and the United Kingdom. | Terminated | March 31, 2011 | 1274 | 63 | Not_Computer
Z-0583-2008 | 45807 | No | KNC | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Hill-Rom Affinity Three Birthing Bed, Model P3700A. |  |  | November 21, 2007 | 2007 | Hill-Rom, Inc. | Hole/opening on the labor bar -- Amputation of a finger could result, if a finger is placed into the opening, as the hole/opening creates a shear point when the mechanism is activated. | Device Design | Consignees were notified of the problem via an Urgent Medical Device Correction letter dated 11/21/07. Users were informed that they would be receiving a free upgrade kit for all of their Affinity Birthing Beds. The Kit includes two plugs per bed, an insertion tool, and full written instructions. An instruction video is available at www.Hill-Rom.com. |  | 32379 | 32,379 for all model numbers | Worldwide Distribution-- including USA and countries of Argentina, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Brazil, Brunei, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Dominican Republic, Egypt, Finland, France, Germany, Greece, Guam, Honduras, Hong Kong, Hungary, Iceland, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Malaysia, Mexico, Morocco, The Netherlands, New Calidonia, New Zealand, Nigeria, Northern Mariana Islands, Norway, Oman, Panama, Philippines, Poland, Portugal, Qatar, Republic of Korea, Saudi Arabia, Singapore, South Korea, Romania, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguey, Venezuela, and Virgin Islands . | Terminated | September 22, 2009 | 671 | 32379 | Not_Computer
Z-0729-2008 | 46155 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Centurion Healthcare Products Laceration Tray, sterile; (Reorder SUT8805A) Tri-State Hospital Supply Co., Howell, MI 48844. The primary kit is labeled as containing 1 needle, 21G x 1-1/2", among other components, however some kits may contain a 21Gx1" needle. The label on the Kendall Monoject Magellan 21G Safety Needle itself is the correct size. |  |  | October 26, 2007 | 2007 | Tri-State Hospital Supply Corporation | Incorrect Needle size - Some of the kits contain a 1 inch 21G needle instead of the 1-1/2 inch 21G needle specified on the kit label. | Mixed-up of materials/components | An Urgent Product Recall Notice was sent on October 25, 2007 to Distributors, Distributors customers, and all third party consignee. The letter provided information concerning the product recall, and requested removal of all product inventory. A product return form requesting information on products still in the facility was enclosed. A Tri-State Representative will be in contact regarding retrieval of the product. |  | 40 | 40 kits. | Nationwide Distribution, including states of Massachusetts and Washington. | Terminated | March 13, 2008 | 139 | 436 | Not_Computer
Z-0732-2008 | 46318 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens Syngo Imaging System with VB30A version of software. Model Number: 10014063, Siemans Medical Solutions, Malvern, PA 19355 |  |  | November 27, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Inconsistent data - A buffer overflow may occur between information system and picture system if the system experiences a high load. (e.g. RIS/PACS communication is unusally high) | Software design | The recalling firm issued a Customer Safety Advisory, via certified mail. to the affected customers on 11/27/07. The letter informs the customer of the issue and also provides instruction to detect and avoid its occurrence. In addition, the firm is developing a software update to correct the issue, which will be installed on the affected systems when it is released. |  | 10 | 10 units | Nationwide Distribution, including states of AL, MO, NC, NY, OH, OR, TN, and TX. | Terminated | September 17, 2008 | 295 | 10 | Software | Output/Calculation | Software update | Software Update
Z-0731-2008 | 46266 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Kodak DirectView Model DR 3000 System, Stationary X-ray System; Catalog #s 6551360, 6551386, 6552152, 6552160, 6551345, 6552319, 6552335. The units are labeled under the following firms: EASTMAN KODAK COMPANY (now known as Carestream Heath Inc.) Rochester, NY 14608, or CARESTREAM HEALTH, INC., Rochester, NY 14608. |  |  | December 21, 2007 | 2007 | Carestream Health, Inc. | Unintended Positioner Movement - The U-Arm Positioner starts to move 1) without command from the operator and/or 2) can unexpectedly start movement if the system is powered down and re-energized after a collision has occurred with the Quantum Table. | Software design | On 12/21/07, Service Bulletins (dated December 2007) were sent to consignees informing them of design modification, and providing modification instructions. The firm followed up with telephone calls to the direct user sites to schedule appointments to visit each site, in order to make the necessary modifications. The firm will be installing Safety Adaptation Board and software into all the affected DR3000 systems. |  | 19 | Domestic - 19 units | Nationwide Distribution including states of NY, PA, MA, MI, IL, NC, GA, FL, KS, MO, TX, IA, and WA. | Terminated | April 10, 2008 | 111 | 19 | Software | Physical Safety Hazards | Software update | Software Update
Z-0131-2008 | 38377 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | American Medical Systems, Inc., (AMS) GreenLight HPS Fiber Optic (ADDStat) Part Number 0010-2090B. Fiber optic laser for treatment of BPH (Benign Prostatic Hyperplasia). AMS Innovative Center, San Jose, CA 95134 |  |  | June 18, 2007 | 2007 | AMS Innovative Center - San Jose | Control knob break: The knob used to finely control the direction of the laser energy may break. | Process design | American Medical Systems, Inc. (AMS) issued Important: Medical Device Recall AMS Greenlight ADDStat Fibers notification letters to customers who have purchased the product on June 28, 2007. The customer letter requests recipients that the product use be discontinued, and the fibers be returned. The firm will send replacements. |  | 611 | 611 units | Worldwide Distribution - USA and Singapore. | Terminated | March 23, 2010 | 1009 | 611 | Not_Computer
Z-0134-2008 | 44799 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical CADD-Sentry Pro Medication Safety Software - electronic pump protocol programming software - Administrator |  |  | August 14, 2007 | 2007 | Smiths Medical MD, Inc. | Incorrect dosage possible: A software anomaly in the CADD-Sentry Pro Administrator software (version B) where the values of a weight based protocol go to zero or round up to the next higher hundredths place, even if Administrator displayed the values correctly at data entry. | Software design | A Smiths Medical MD employee visited the consignees between August 14 and 31, 2007, to verbally notify them of the recall, collect the recall CD-ROM containing the affected software and replace with the updated software. |  | 3 | 3 CD-ROMs | UT, MD, CO | Terminated | October 31, 2008 | 444 | 3 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0674-2008 | 45878 | No | LXT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Orthofix TrueLok External Ring Fixation System, Telescopic Linear Distractor, Threaded in the following sizes: 75mm (Catalog Number: 50-1101), 100mm (Catalog Number: 50-1014), 150mm (Catalog Number: 50-1015) and 200mm (Catalog Number: 50-1016), distributed by Orthofix, Inc., McKinney, TX 75069. |  |  | September 28, 2007 | 2007 | Orthofix, Inc | Markings reversed: Graduated markings on telescoping bone distractors were reversed, causing the distractors to compress rather than distract the treatment site when adjusted. | Process control | Orthofix notified sales agents by phone starting on 09/28/07. Sales agents were to return product for replacements. |  | 69 | 69 units. | Nationwide Distribution - USA states: CA, FL, GA, IL, PA, MI, NY and WI. | Terminated | March 20, 2008 | 174 | 69 | Not_Computer
Z-0920-2008 | 45947 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Modular Replacement System Curved Cemented Stem; |  |  | November 15, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Curvature Discrepancy -- The radius of curvature may be out of specification to varying degrees causing a discrepency in the stem curvature between trials and implants. | Process design | On November 26, 2007, Stryker Sent Urgent Product Recall notification letters to Stryker Agencies/Branches and hospitals (OR Supervisor) by Federal Express. The notification requested that all inventory be examined and returned. Returning instructions were provided in the letter. Questions concerning this recall are being addressed by contacting the Regulatory Compliance Hotline at (201) 831-6633. |  | 803 | 803 total, all sizes | Nationwide Distribution | Terminated | September 22, 2008 | 312 | 803 | Not_Computer
Z-0930-2008 | 46146 | No | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | Abbott AxSYM Gentamicin Reagent - Is a reagent for in vitro diagnostic use. This reagent system is composed of one bottle of gentamicin antiserum, one bottle of pretreatment solution and one bottle of gentamicin flourescein tracer. Abbott Laboratories, Abbott Park, IL 60064 |  |  | October 26, 2007 | 2007 | Abbott Diagnostic International, Ltd. | Incorrect Results -- Through evaluation of stability, Abbott have found a decrease in the calibrator A to F span with these reagents lots, which may produce the following results: - Printed Error Code "SPAN LESS THAN MIN SPAN" - Controls out of range. When a control is out of range, patient results should not be reported. | Other | Letters were sent to customers by Federal Express Priority mail on October 26, 2007. Two unique customer letters were sent. One letter was delivered to TDxlTDxFLx customers who received shipment oflots 5ll87Q100 and 51189QI00 and a second letter was delivered to AxSYM customers who received lots 53190QI00 and 53190QI01. Two letters were written because the actions required by the customer differ for the AxSYM and TDxlTDxFLx platforms. TDx/TDxFLx customers "Customer Letter (Product Recall) "Customer Reply Form AxSYM Customers "Customer Letter (Product Correction) "Customer Reply Form The Customers were instructed the following: TDx/TDxFLx Gentamacin Reagents " Discontinue use of and destroy TDx/TDxFLx reagent lots 51187Q 1 00 and 51189QI00. " Document the number of kits discarded for your records. Abbott will provide reimbursement for your destroyed kits. " If customers have forwarded any of the affected lots to another laboratory, to provide the laboratory with a copy of the communication. AxSYM Gentamicin Reagents " Revise the expiration date for lots 53190Ql 00 and 53190QI0l from 4/12/08 to 11/30/2007. " Run each level of control with each AxSYM Gentamicin sample or batch of samples. " If customers have forwarded any of the affected lots to another laboratory, to provide the laboratory with a copy of the communication. |  | 3787 | 3787 kits | Worldwide, USA including states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, Hawaii and Puerto Rico, an countries of Canada, Mexico, Guatemala, Germany, Australia, Japan, Singapore, Hong Kong, Taiwan, South Korea, and New Zeland. | Terminated | July 30, 2012 | 1739 | 5186 | Not_Computer
Z-0578-2008 | 46158 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TENET Medical Engineering Wrist Stabilization Kit, Medium; REF (Part) #7692; Sterile; manufactured by TENET Medical Engineering, Calgary, Alberta, Canada T2Z 4M3. |  |  | November 23, 2007 | 2007 | Tenet Medical Engineering Inc | Sterility compromised: Sterility seals on some of the firm's wrist stabilization kit pouches may be compromised. | Packaging process control | TENET notified US distributor consignee via Urgent Product Recall Notification letter sent FedEx on 11/23/07. Consignee asked to return product and to contact their consignees to also return product. |  | 360 | 360 units. | Nationwide Distribution. | Terminated | May 01, 2009 | 525 | 820 | Not_Computer
Z-1135-2008 | 46210 | No | FRW | Patient scale. | General Hospital | General Hospital | 510(K) Exempt | BHM Medical Scales, 800 lb. scale weight limit; Part Number 700.00510, Manufactured by BHM Medical Inc., Magog, QC, Canada J1X 5Y5; |  |  | December 19, 2007 | 2007 | Arjo, Inc. | Unintended Hanger Bar Detachment: The locknut on the scales may loosen during use and allow the hanger bar assembly of the patient lift to detach and fall. | Device Design | Arjo sent Urgent Device Field Correction Notification letters dated 12/19/07 to the end user accounts who received the affected scales, advising them of the potential for the hanger bar assembly to detach. The accounts were given the option immediately stopping use of the lifts with the affected scales or removing the scales from the lift so the lift can continue to be used. Instructions for removal of the scales were provided with the letter. The accounts were requested to complete and return the enclosed customer response form, acknowledging the receipt and understanding of the letter, and indicating whether they will complete the correction upon receipt of the correction parts or if they require the assistance of an Arjo Service Technician. |  | 772 | 772 scales | Nationwide including the states of Alabama, Arizona, California, Delaware, Florida, Hawaii, Iowa, Illinois, Indiana, Massachusetts, Michigan, Minnesota, Mississippi, North Carolina, North Dakota, New Hampshire, New Jersey, New York, Ohio, Pennsylvania, South Dakota, Tennessee, Washington and Wisconsin. | Terminated | November 18, 2009 | 700 | 798 | Not_Computer
Z-0961-2008 | 45853 | No | LGW |  |  | Neurology | PMA | Advanced Bionics Precision Implantable Pulse Generator Kit, Model Number SC-1100 with firmware version prior to Revision 3.02, Advanced Bionics Corportation, Sylmar, CA 91342 |  |  | October 05, 2007 | 2007 | Advanced Bionics Corp | Incorrect Data -- Corruption of internal memory component results in an inability for the physician to reprogram the IPG with firmware version prior to Revision 3.02. When this occurs, the IPG will report an error code of "10h0" or "00h0" through the Remote Control. Under this condition, the IPG will cease to log in some data that could be used for informational purposes. ** It should be noted t | Device Design | Advanced Bionics sent an Important Notification letter on October 5, 2007, via FED-Ex or US Postal Service Certified mail. (International notification activity is in progress.) The recall letter advises that the firm has become aware of an issue with Precision Implantable Pulse Generators (IPG model numbers SC-1100, Sc1110, and ONS-1110) manufactured prior to July 2007. The letter describes the issue in detail. The firm is requesting that customer immediately locate the affected IPGs and set them aside so that they are not used. A Boston Scientific/Advanced Bionics sales representative will contact your facility to confirm these IPGs and arrange for the exchange of these devices.The firm enclosed Reply Verification Tracking Form listing the IPGs that are at your Clinic according to the firm's records. In addition, the firm recommends that if error code 10h0 or 00h0 is encountered in Implanted IPGs, it may be corrected by upgrading to a newer version of IPG firmware (software that runs in the IPG). The new version of firmware corrects this condition and helps prevent it from occurring in the future. This firmware upgrade is performed non-invasively, through wireless RF telemetry.If any of patients encounter error code "10h0" or "00h0" on their Remote Control, consignees are advised to contact their Advanced Bionics representative or Technical Services team at (866) 566-8913 to help identify whether the condition can be corrected through a firmware upgrade and, if so, they will help arrange for the upgrade. Please note that some devices may not be upgradeable. The firm request that all consignees return the notification acknowledgement form is included with the letter. |  | 12000 | Over 12,000 units | Class II Recall - Worldwide Distribution --- including USA and countries of EU, Canada, Australia and Israel. | Terminated | February 29, 2012 | 1608 | 12000 | Computer
Z-0570-2008 | 45447 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Biosciences MultiSET Flow Cytometry Software Versions 2.1, 2.1.1, and 2.2 with OS X operating systems, BD Biosciences, San Jose, CA 95131. |  |  | September 06, 2007 | 2007 | BD Biosciences | Inaccurate results: Software error results in inaccurate display result statistics. If the user adjusts the lymph gate or attractors in the Lab Report view without using the Manual Gate function (as described in the system manual), the statistic results will not be updated to match the adjusted gate. | Software design | Becton Dickinson sent an Important Product Information letter by mail on 09/14/2007 informing customers of the need to follow the instructions for adjusting the lymph gate or attractors. The next software revision will alleviate this limitation. |  | 1136 | 1136 units | Worldwide Distribution - USA, Belgium, Brazil, Peru, Singapore, Canada, Jamaica, Chile, Ecuador, Hong Kong-China, Thailand, Mexico, Venezuela, Barbados, Philippines, Puerto Rico, Guatemala, Japan, Uruguay, Argentina, China, Trinidad and Tobago, Honduras, Australia, New Zealand, and India. | Terminated | April 15, 2008 | 222 | 1136 | Software | Output/Calculation | Software update | Software Update
Z-0576-2008 | 45984 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS 6000 Clinical Chemistry analyzer; Catalog Number: 04745914001, Roche Diagnostics Corporation, Indianapolis, IN 46250. |  |  | December 04, 2007 | 2007 | Roche Diagnostics Corp. | Incorrect results: CRP assays may be calculated using incorrect calibration parameters which could result in falsely high or falsely low patient results being reported. | Software design | Roche Diagnostics contacted the consignee by phone on 12/4/07 and by follow-up letter dated 12/4/07. The user was instructed to discontinue use of two CRP assays on this analyzer. The letter described a workaround and that a new revision of software, due out first quarter 2008, will fix this. |  | 1 | 1 | USA Distribution in the state of Nebraska. | Terminated | August 04, 2008 | 244 | 1 | Software | Output/Calculation | Software update | Software Update
Z-0574-2008 | 45881 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Disposable StrykeProbe 32cm Spatula Tip, Ref: 250-070-551, Electrosurgical Suction Tip, Stryker Endoscopy, San Jose, CA 95138. |  |  | November 16, 2007 | 2007 | Stryker Endoscopy | Cannot connect: Tips have electrodes that are too large to fit monopolar cables and are unusable. | Device Design | Stryker notified their sales representatives by phone. Stryker sent all consignees a Product Advisory Notice dated November 19, 2007, and verbally notified them of the recall via telephone. A response form with return postage was sent with the notice. The letter requested the hospitals to quarantine and return the product to Stryker. |  | 48 | 8 boxes of 6 units (48 total) | Nationwide Distribution. | Terminated | March 26, 2008 | 131 | 48 | Not_Computer
Z-0865-2008 | 46206 | No | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | Pulsavac Plus Low Pressure Fan Spray Kit, Catalog #00-5150-975-00, 10 kits/package, Zimmer Orthopaedic Surgical Products, Dover, Ohio |  |  | December 18, 2007 | 2007 | Zimmer Orthopaedic Surgical Products | Sterility (package integrity) compromised; The effect of a silicone stain produced during assembly operations on the sterility barrier properties of the Tyvek lids has not been validated by the firm. | Other | Zimmer notified their customers via an Urgent Device Recall letter dated January 8, 2008. The letter requests each customers to inventory their stock. If any recalled product is located the customer is to notify Zimmer via fax and a replacement part will be shipped to them. |  | 607 | 607 kits | Worldwide-USA, forty-two international and one Canadian account, including countries of Singapore, Hong Kong, Lebanon, Finland, Australia, Japan, Israel, Korea, Saudi Arabia, Dubai, Denmark, Czech Republic, Guatemala, Brazil, Taiwan, China, Nicaragua, Virgin Island, Switzerland, Austria, Spain, Belgium, The Netherlands, France, Germany, Portugal, England, Sweden, India, Argentina, Italy, Chile, Thailand &amp; Saudi Arabia | Terminated | May 03, 2012 | 1598 | 84644 | Not_Computer
Z-1182-2008 | 46877 | [6 empty] | System 83 Plus Washer-Disinfector, Endoscope use, DOS based system, Custom Ultrasonics Inc., Buckingham, PA |  |  | February 11, 2008 | 2008 | Custom Ultrasonics, Inc. | Inadvertent selection of the "wash" cycle | Labeling design | The recalling firm issued a Product Correction letter dated 2/11/08 to its customers informing them of the problem and the correction choices. The letter states that a key cover is available to reduce the error in the future and will be provided to each users upon return of the recall letter notification questionnaire. |  | 3190 | 3,190 units | Worldwide; USA medical facilities and physicians and Canada | Terminated | September 03, 2009 | 570 | 3190 | Not_Computer
Z-0947-2008 | 45861 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HomeChoice PRO Automated Peritoneal Dialysis Systems; Catalog Numbers; 5C8310, 5C8310R, R5C8320, R5C8320R, T5C8300, and T5C8300R, Manufactured for Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A., Made in U.S.A |  |  | November 20, 2007 | 2007 | Baxter Healthcare Renal Div | Unrecorded Infusion; If the system is powered down or a power failure occurs during a fill cycle, the HomeChoice system may not record the last 1-4 pump strokes infused into the patient when the power is restored. This situation may create a potential unrecorded delivery of approximately 15-60 mL of fluid being infused into the patient. | Device Design | Urgent Product Recall letters dated 11/20/07 were sent via first class mail to all of the affected customers, informing them of the potential for overfill if the system is powered down or a power failure occurs during a fill cycle. A Baxter representative will contact the customer to schedule the software update which corects this anomaly. The accounts were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff and home patients. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 1.Baxter sent a follow-up letter dated 3/17/08 to the consignees, accompanied by a copy of the 11/20/07 letter. The letter reiterated the information in the 11/20/07 letter, emphasizing the serious consequences of an overfill for neonates and young pediatric patients, and providing information describing the symptoms of excess fluid in the abdomen and steps to be taken if the user suspects overfill of a neonate or young pediatric patient. Baxter requested that Baxter be contacted if an overfill should occur so Baxter can perform a complete investigation into the event. The accounts were requested to complete and return the enclosed Customer Reply Form, indicating that they have disseminated the information to their staff and home patients. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 1. |  | 32229 | 32,229 units | Worldwide; USA, Australia, New Zealand, China, Hong Kong, the Philippines, Thailand, Asia, Canada, Algeria, Austria, Bahrain, Belgium, Bulgaria, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Hungary, Iceland, Iraq, Ireland, Israel, Italy, Jordan, Kuwait, Latvia, Lebanon, Lithuania, Malta, Morocco, The Netherlands, Norway, Oman, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Senegal, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Ukraine, Latin America and Japan. | Terminated | October 26, 2009 | 706 | 103490 | Computer
Z-1196-2008 | 46784 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Icy Hot Heat Therapy , Air Activated Heat Patch, Back and Large Areas, UPC Code 0-41167-08330-7, Product Code 08330, 4.1" x 8.4" (10.5 cm x 2.1 cm) each, 1 and 3 count patches, Distributed by Chattem, Inc. P.O. Box 2219, Chattanooga, TN 37409-0219 |  |  | February 12, 2008 | 2008 | Chattem Inc | Skin irritation and burns with product usage | Device Design | The firm initiated its recall on 02/08/2008 by issuing a press release warning the public of the hazard and directing them to stop its use and either return or discard the products. The firm initiated the recall at the retail level on 02/12/2008 by notifying its employee sales force by email, containing a Urgent Medical Device Recall Letter, with instructions to immediately email a copy of the recall letter to each Chattem consignee (wholesalers and direct retail accounts) with sub recall requests. |  | 666636 | 666,636 units (total for Recalls Z-1196-1197-2008) | Worldwide - USA, Canada, and Honduras | Terminated | October 05, 2009 | 601 | 712656 | Not_Computer
Z-0186-2008 | 44911 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Size 8 Pedicle Screws of the Allez Spine Laguna Pedicle Screw System, Screw Size (diameter x length) 8 x 35, Catalog Number LG-PR0835, Allez Spine, LLC., Irvine, CA |  |  | July 30, 2007 | 2007 | Allez Spine, LLC | Screw Shank Failure; separation of the screw shank portion from the pedicle head portion of the Size 8 Pedicle Screw, when torquing down the single piece locking nut during implantation. | Device Design | On July 30, 2007 Allez Spine notified all its field representatives and distributors via email of a potential for failure of the Size 8 Pedicle Screws when used in combination with the single piece locking nuts and instructed them to return all unused inventory of Size 8 Screws and the two piece Locking Nuts with saddle. This action was to be carried out to the retail level-hospital, clinic, lab, and physicians sold the product. The email also had an attachment of a list of expected frequently asked questions (FAQ) from surgeons and answers to help them respond. A second Notice to all users was sent by the recalling firm on Feb 29, 2008, to request follow-up of implanted units. |  | 317 | 317 | Nationwide: including the states of AZ, CA, CO, GA, NV, TX and UT | Terminated | April 09, 2012 | 1715 | 1578 | Not_Computer
Z-1193-2008 | 46726 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker VertePort, 10 G Cement Cannula, REF 306-410, Qty 18; For use with the Stryker PCD Precision System and ACM Delivery System, Rx only, sterile, Stryker Instruments, Kalamazoo, MI. |  |  | February 13, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | Mis-labeling; The instructions for use include a use for which the device is not intended. (Device not approved to capture a bone biopsy) | Labeling Change Control | Consignees were notified via Medical Device Recall Notification letter dated 2/13/08 to read the enclosed instructions for use, remove all old instructions for use at the consignee location and replace them with the new instructions for use, and to discard the old instructions for use. |  | 953 | 953 boxes Recalls Z-1192-1195-2008 | Nationwide. | Terminated | December 30, 2008 | 321 | 953 | Not_Computer
Z-0572-2008 | 45859 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Exactech 12/14 Biolox Alumina Femoral Head 28MM +3.5, hip prosthesis component, Exactech, Gainesville, FL 32653 |  |  | November 01, 2007 | 2007 | Exactech, Inc. | Mislabeled: A ten piece lot of 140-32-03, 12/14 BIOLOXforte Alumina Femoral Heads 32mm (1069020 - 1069029), from Ceramtec were incorrectly labeled as 140-28-03 12/14 BIOLOXforte Alumina Femoral Heads, 28mm | Process control | Exactech issued a recall letter, titled Important Product Recall Notice on 10/31/2007 to the Distibutors that received the recalled product. It requested that all device recipients be notified and all product be returned to the firm. |  | 10 | 10 | Distributed to NY, OH and TN. | Terminated | May 06, 2012 | 1648 | 10 | Not_Computer
Z-1206-2008 | 45901 | No | MCY | Elastic bandage. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Smith &amp; Nephew, Inc., Exu-Dry Wound Burn Dressing, Product Number: 5999M36, Exu-Dry Pad 24"x36" Non-Permeable Heavy Absorbency, Smith &amp; Nephew, Inc.,Largo, Florida. |  |  | November 21, 2007 | 2007 | Smith And Nephew, Inc. Wound Management Division | Sterility (package integrity) compromised: Distribution of product that did not meet specifications; open seals. | Package design/selection | Consignees were notified of recall by certified mail, Urgent Device Recall letter dated 11/21/2007, which was sent to each consignee. The recalled product will be returned to the firm by pre-paid FedEX return labels. A second letter is to be issued to each distributor for recall to the retail level, as the firm's original recall strategy of distributor level recall, was deemed inappropriate by CDRH |  | 5100 | 5,100 Dressings | Worldwide-USA, England, Japan, South Africa, New Zealand, Australia and Canada. | Terminated | December 15, 2008 | 390 | 47554 | Not_Computer
Z-0455-2008 | 38465 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Abbott - CELL-DYN 1800 System with CELL-DYN 22 Calibrators and the CELL-DYN Calibrators and CELL-DYN 16 Controls. Part/Model Number: 07H77-01. Abbott Laboratories, Santa Clara, CA 95054-1113 |  |  | October 27, 2006 | 2006 | Abbott Laboratories | Incorrect Results: Hemoglobin results in donor samples assayed on the CELL-DYN 1800 System vary more than expected across the range of claimed operating temperatures of 18 degrees Celcius to 30 degrees Celcius. | Software design | A Product Correction letter was sent to all customers with a reply form on October 27, 2006. Product information letters are being packed with the instrument's accessory kit and with calibrators and quality control products used with the CELL-DYN 1800 System. Correction letters with reply forms have been resent to nonresponding customers. |  | 4224 | 4,224 Systems | Worldwide Distribution, USA and countries of Argentina, Bolivia, Brazil, Canada, Chile, Columbia, Costa Rica, Dominican Republic, El Salvador, Germany, Guatemala, Honduras, Hong Kong, China, Japan, Mexico, Paraguay, Puerto Rico, Singapore, South Korea and Venezuela. | Terminated | January 06, 2011 | 1532 | 4224 | Not_Computer
Z-0933-2008 | 46293 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Linemaster Wireless Footswitch Accessory, Model SP-9970214-005112010 and Model SP-9970214-005112232. Sound Surgical Technologies, LLC, Louisville, CO 80027. |  |  | December 28, 2007 | 2007 | Sound Surgical Technologies, LLC | Failure to shut-off: Surgical equipment, activated by a wireless foot switch with VASER system, may remain powered on when switch is no longer depressed. | Device Design | Consignees were sent an Advisory Notice on 12/28/2007. The notification describes the defect. Sound Surgical will be contacting facilities to arrange a service call to install new cables between the footswitch receiver and the VASER system. Facilities may contact the Technical Service Center at 1-888-677-8765. |  | 127 | 127 switches | Worldwide Distribution including USA states of AL, CA, CT, FL, GA, IN, KY, LA, MI, NJ, NY, OH, OK, PA, TX, and VA. Foreign Distribution to Colombia, Costa Rica and Switzerland. | Terminated | April 01, 2008 | 95 | 127 | Not_Computer
Z-0931-2008 | 46207 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Bayer Contour TS blood glucose monitoring system with 10 blood glucose test strips enclosed; Bayer Healthcare, Mishawaka, IN. Products 1802 and 1802N (U.S.), 1810 (Korea), 1811 (India, Korea and Mexico) and 1825 (Mexico). |  |  | December 21, 2007 | 2007 | Bayer Healthcare LLC | InaccurateTest Strip Results: Results in blood glucose readings with a positive bias are outside the product specifications. Patient test results may demonstrative results 5 percent to 17 percent higher. | Process design | The firm issued a Press Release on 12/21/07. A Test Strip Recall letter was also sent to the firm's consignees and health care providers instructing them that use of the product should cease and that health care providers contact the firm for replacements. Additional information can be found at www.bayerdiabeties.com or by contacting the Bayer at 1-800-348-8100. |  | 97000 | 97,000 bottles of test strips | Worldwide Distribution --- including USA and countries of Austria, France, India, Korea, Mexico, Philippines and Turkey. | Terminated | July 07, 2009 | 564 | 97000 | Not_Computer
Z-0717-2008 | 46349 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Mobile Fluoroscopy X-Ray System, OEC, Model 7700, GE Healthcare, Surgery, Salt Lake City, Utah |  |  | September 28, 2007 | 2007 | GE OEC Medical Systems, Inc | Unnecessary operator exposure; X-ray machine may allow unwanted x-ray exposure to the operator. | Other | On or about 9/6/2007, by Customer Notification Letter, users were notified that all affected units will be corrected by GE personnel with the addition of an additional shield. |  | 309 | 309 systems (Domestic), 101 (Foreign) | Worldwide- USA including states of CO, DE, FL, NJ, MO, NY, PA, and TX (with VA and Military facilities) and countries of Australia, Austria, Belgium, Brazil, Chile, China, The Czech Republic, Denmark, France, Germany, Hungary, Israel, Italy, Norway, Spain, Sweden, Switzerland, Thailand, and The UK. | Terminated | November 01, 2011 | 1495 | 466 | Not_Computer
Z-1051-2008 | 46742 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Draeger TI500. Isolette Infant Incubator. Catalog Number: MU20505. Draeger Medical, Inc., Telford, PA 18969 |  |  | January 31, 2008 | 2008 | Draeger Medical, Inc. | Heating failure- The power board which controls the incubator heater may not regulate the temperature properly resulting in a high temperature alarm and a loss of temperature control within the patient compartment. | Nonconforming Material/Component | The firm notified consignees on 2/1/08, by an Important Safety Notice. The letter instructed users to remove any units from use until a Draeger Service Representative can replace the power board. A Service Representative will be in contact to schedule a service visit. For questions concerning this recall please contact 215-660-2418 or Anbu.arasu@draeger.com. |  | 18 | 18 units | Worldwide Distribution-- USA including states MO, OH, and PR, and countries of Australia, Canada, Columbia, Israel, UAE, France, China, Italy, Denmark, Switzerland, and UK. | Terminated | May 08, 2008 | 98 | 24 | Hardware | Device Operation | Replace board | Remove or Replace
Z-0571-2008 | 45798 | No | NQN | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | Ventana Image Analysis System (VIAS) - PATHWAY anti-HER-2/neu (4B5) primary antibody with software version 2.3.0.2, Model Number: VIAS-799-10000-00, Pathology Image Analysis System, Ventana Medical Systems, Inc., Tucson, AZ 85755 |  |  | June 26, 2007 | 2007 | Ventana Medical Systems Inc | Incomplete upgrade: The software upgrade from February 1, 2007 may not have beencompleted. This software upgrade adds the algorithm (Slidetype) for thePATHWAY anti-HER-2/neu (4B5) antibody to the VIAS system. | Software change control | Beginning in 6/07, Ventana Medical Systems, Inc., made phonecalls to labs asking to confirm software version, date of upgrade if available and confirmation of whether site uses 4B5 and Imaging together. A phone script was provided. Customers with systems not ugraded were to upgrade before further use of system. A notification letter dated 8/17/2007 was later sent to VIAS customers advising that the software upgrade from February 1 , 2007 may not have beencompleted on their system. This software upgrade adds the algorithm (Slidetype) for thePATWAY anti-HER-2/neu (4B5) antibody to their VIAS system . This Slidetype is cleared for marketingand is required prior to imaging slides stained with PATHWAY anti-HER-2/neu (4B5).The letter also advises that their VIAS and NEXES systems are conneded to the internet which enables this upgrade to beperformed remotely with the support of the Ventana Medical Systems Customer Support Center(CSC). customers are instructed to call the CSC to carry out the upgrade or to schedule aTechnical Applications Specialist (TAS) to assist with the upgrade in their lab. Arrangements for the upgrade are to be made prior to imaging slides stained with PATHWAYanti-HER-2/neu (4B5). |  | 77 | 77 devices | Nationwide and Canada | Terminated | March 26, 2012 | 1735 | 77 | Software |  | Software update | Software Update
Z-0603-2008 | 44704 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew, CALAXO Bioabsorbable lnterference Screw, 9 mm x 25 mm, |  |  | August 21, 2007 | 2007 | Smith &amp; Nephew, Inc. Endoscopy Division | Patient Reaction: Post operative condition sterile fluid pocket has been identified, including the potential for graft failure and premature material degradation. | Nonconforming Material/Component | Smith &amp; Nephew issued an Urgent Product Recall letter via Federal Express to User Facilities, Sales Reps and Distributors on 8/21/07. A second letter was issued on 8/23/07 to correct two errors in model numbers listed in the original letter. Letter issued by FedEx and/or email to OUS contacts. |  | 6925 | 6925 units | Worldwide Distribution - USA, Austria. Belgium, Germany, Denmark, Spain, Finland, France, England, Italy, Portugal, Netherlands, Norway, New Zealand, Sweden, Turkey, Switzerland,China, Hong Kong, Malaysia, Taiwan, Singapore, Thailand, Australia, Canada, Mexico, Colombia, South Africa, Dubai, Venezuela, and Chile. | Terminated | May 09, 2012 | 1723 | 66263 | Not_Computer
Z-0620-2008 | 45479 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Trilogy Acetabular System Shell with Cluster Holes, Porous, 60mm O.D., Tivanium TI-6AL-4V Alloy, Hip Prosthesis Component, REF 6200-060-22, Catalog # 00-6200-060-22, Zimmer, Inc., Warsaw, IN, 46580. |  |  | October 19, 2007 | 2007 | Zimmer Inc. | Parts missing and incomplete etch: The units may not contain the locking ring or the etched alignment marks on the rim. | Process control | Zimmer sent consignees a Product Recall Notification letter dated 10/19/07, requiring them to return the product. |  | 31 | 31 | Worldwide Distribution - USA (Arizona, California, Delaware, Indiana, Iowa, Kansas, Massachusetts, New York, Ohio, Pennsylvania, South Carolina, Texas, and Wisconsin), Australia, Canada, Germany, Italy, and Spain. | Terminated | May 21, 2008 | 215 | 31 | Not_Computer
Z-1108-2008 | 46362 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UniCel Dxl Access Immunoassay Systems Reaction Vessels for Use with UniCel Dxl Access Immunoassay Systems, Part Number : 386167, Beckman Coulter, Inc, Brea CA 92822 |  |  | July 19, 2007 | 2007 | Beckman Coulter Inc | Erroneous Results - UniCel Dxl Access Immunoassay Systems Reaction Vessels have been found to be defective and there is a potential for an increase in signal leading to an erroneous result. | Process design | A Product Corrective Action (PCA) letter was mailed on July 19, 2007, to the Customers that received UniCel Dxl Accesslmmunoassay System Reaction Vessels lot 1201291 and1201431. The customers were informed that a smallpercentage of the reaction vessels (Lot #s: 1201291and1201431 ) have been found to be defective and there is apotential for an increase in signal leading to an erroneousresult. The customers were instructed to discontinue use anddiscard these lots. A response form was included. |  | 273 | 264 units in the US; 9 units in Canada | Nationwide &amp; Canada | Terminated | May 04, 2012 | 1751 | 273 | Not_Computer
Z-1109-2008 | 46387 | No | CGS | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Abbott Clinical Chemistry Creatine Kinase, Model LN 7D63-30, Lot Number: 52044HW00; Expires October 19, 2007, RCR#2018433-10/25/07-006R Abbott Laboratories, Inc., South Pasadena CA 91030 |  |  | October 12, 2007 | 2007 | Abbott Laboratories Inc. | Three potential issues with Creatine Kinase Lot Number: 52044HW00: (1) Decrease in quality control and/or patient results (2) An iIncreased imprecision, or (3) Error Code 1054 (Unable to calculate results, Reaction Failure) on Architect cSystems. | Nonconforming Material/Component | On October 12, 2007, A Product Recall letter was provided to all Clinical Chemistry Creatine Kinase Reagent customers that have received lot number 52044HW00 . Customer reply form were included in all US customers letter. The letter instructed customers to discontinue use of the kits and destroy them following their laboratory procedure, and document the number of kits discarded. Abbott will provide reimbursement for destroyed kits. |  | 498 | 334 Distributed Worldwide; 164 Domestically | Worldwide Distribution --- including USA and countries of Canada, Chili, Uruguay, Germany, Hong Kong, and New Zealand. | Terminated | March 28, 2012 | 1629 | 498 | Not_Computer
Z-1054-2008 | 46743 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Latex-Free Needle Guide, Sterile transrectal needle guide for use with Shimadzu UB1OR-065U transducer, Part #677-057, packaged 1 needle guide per tyvek pouch, 20 pouches per box. The label shows the product is Mfg by CIVCO, Kalona, IA. |  |  | January 16, 2008 | 2008 | Civco Medical Instruments Inc | Incorrect Needle Guide- The individual pouches contained V53W Endocavity Needle guide(p/n/034-054) instead of the Shimadzu T/R Needle guide (p/n134-054). | Packaging | On 1/16/08, the recalling firm contacted their customers and notified them of the problem. Pictures of the problems were provided for identification. The firm requested that all products be returned. |  | 207697 | Lot M102250 - 5 boxes; Lot M105440 - 2 boxes | Nationwide Distribution- including states of TX and GA | Terminated | April 03, 2008 | 78 | 207697 | Not_Computer
Z-0130-2008 | 38336 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GreenLight HPS System, Surgical Laser System and Accessories. Part Numbers 0010-0070 and 0010-0075 (international units), AMS Innovative Center, San Jose, CA 95134. |  |  | April 17, 2007 | 2007 | AMS Innovative Center - San Jose | Component missing: Some products were issued without thermal protection switches. | Process design | AMS issued Important: Medical Device Correction letters, dated April 17, 2007, to its US and international consignees by mail, instructing them to return an acknowledgement form. The letter informs customers that some HPS systems may lack a motor overload switch, that the firm believes that there is no immediate safety issue, and installation of the MOL switch will be conducted during normal preventive maintenance. AMS expects to complete installation of the switches by end of August 2007. |  | 169 | 169 units | Worldwide Distribution - USA, Italy, France, Germany, Spain, UK, Denmark, Sweden, Australia, India, Japan, Greece, Switzerland, Saudi Arabia, Thailand, Turkey, Singapore, Taiwan/Hong Kong (China), and Austria. | Terminated | March 23, 2010 | 1071 | 169 | Not_Computer
Z-0646-2008 | 45854 | No | KWL | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics; UHR Universal Bipolar Component; 28 x 56 mm, Hip Prosthesis Component, Stryker Howmedica Osteonics Corp., Mahwah, NJ 07430 |  |  | October 26, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Incorrect part: Device may have an incorrect locking ring assembled with the UHR bipolar head. | Process control | Stryker sent Urgent Product Hold letters to US branches/agencies and hospitals on October 26, 2007 advising them to quarantine the product from these lots, and Urgent Product Recall letters on November 2, 2007 advising them to return the devices to Stryker. International distributors were sent an advisory on November 6, 2007. |  | 360 | 360 total | Worldwide Distribution - USA, Canada, China, India, Australia, Korea, Singapore, Taiwan, Sweden, Germany, Netherlands, France, Spain, Italy, Poland, South Africa, and England. | Terminated | January 27, 2010 | 824 | 360 | Not_Computer
Z-1142-2008 | 46685 | No | LKK |  |  | General Hospital | PMA | Medtronic - SynchroMed EL Implantable Infusion Pump, Model Number: 8626-10. The contents of the inner package have been sterilized by ethylene oxide gas. 10 mL pump with reservoir valve. Medtronic, Inc., 710 Medtronic Parkway NE, Minneapolis, MN 55432-5604, USA. The implantable Medtronic SynchroMed EL Programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site. The implantable components of the SynchroMed EL Infusion System include the pump with or without a side catheter access port, catheters, and catheter accessories. |  |  | January 16, 2008 | 2008 | Medtronic Neuromodulation | Device/Drug Interaction - The company updated the labeling for the devices to include current patient management and treatment recommendations. The company received reports of inflammatory mass formations at or near the distal tip of intrathecal catheters which infuse opioids, baclofen, or chemotherapy drugs into patients. On January 16, 2008, Medtronic sent a letter to doctors who implant these d | Device Design | An Urgent Medial Device Correction letter was sent January 16, 2008, to Health Care Professionals. The letter describes the incidences, symptoms and recommendations for patient management. Excerpts from the approved Medtronic Professional Labeling are also included with the letter. The firm has requested unreported inflammatory mass in a patient with a Medtroinc device to the firm and to the FDA MedWatch Program by phone at 1-800-FDA-1088. Additional assistance may be obtained by contacting Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 102792 | ~102,792 worldwide. 90,330 within US and 12,462 OUS | Worldwide Distribution -- USA including states of Washington D.C., Puerto Rico, and countries of Aruba, Australia, Austria, Belarus, Belgium, Brazil, Canada, China, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Faroe Islands, Finland, France, French Polynesia, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Luxembourg, Malta, Mexico, Netherlands, Netherlands Antilles, New Caledonia, Norway, Pakistan, Poland, Portugal, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia and Montenegro, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Vatican City State. | Terminated | June 06, 2011 | 1237 | 925128 | Not_Computer
Z-1044-2008 | 46245 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Siemens Magnetom Harmony sygno. Model Number: 7104693, Siemens Medical Solutions USA, Malvern PA, 19355 |  |  | December 01, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Magnet Quench, unintended- magnet of a mobile MRI system quenched into the exam room. | Device Design | The firm issued an Update Instruction package to consignees 12/07. As an immediate measure, the recalling firm is initiating field checks of mobile MRI systems per Update Instruction MR039/07/S. This update defines a safety inspection for mobile systems, including checks for potential ice formation and proper configuration of venting assemblies. |  | 10 | 10 units | Nationwide Distribution | Terminated | October 01, 2012 | 1766 | 210 | Not_Computer
Z-1367-2008 | 45950 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gambro cartridge blood tubing set used with Gambro Dasco Phoenix and Cobe Centrysystem 3 (and 3+) Hemodialysis System, Gambro, Lakewood, CO 80401 |  |  | September 10, 2007 | 2007 | Gambro Renal Products, Inc. | Kinked tubing on hemodialysis device may cause hemolysis. | Device Design | On September 10, 2007, the firm sent an Advisory Notice to all consignees alerting them that a kink in the tubing can result in physical hemolysis. The firm included with the Notice a Quick Reference Guide that provides drawing of the proper installation of the Phoenix and describes how to prevent bloodline kinking. A customer reply form was also included. The firm is in the process of making labeling changes that will be released with the updated software revisions. |  | 41000000 | 41,000,000 | Nationwide, including 17 VA facilities and 1 military facility. Foreign distribution to Canada, Korea, Mexico. | Terminated | January 03, 2012 | 1576 | 41000000 | Not Computer
Z-1297-2008 | 46746 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics, Locking Screw, T7, Catalog Number: 40-27622, Size: 2.7X22MM, Non Sterile,Titanium Alloy; Manufactured and Distributed by: Stryker Leibinger GmbH &amp; Co. KG;Boetzinger Strasse 41; 0-79111 Freiburg, Germany |  |  | January 28, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Screws do not have locking threads | Process control | Recall notification letters (Urgent Product Recall) were sent on January 28, 2008 to all Stryker Branches/Agencies by Federal Express. All product distributed was requested to be located and returned, including from hospitals. |  | 100 | 100 units | Nationwide-USA including states of MA, PA, GA, OH, and TX (hospital consignees) and Stryker Branches/Agencies in MA, NJ, PA, MD, VA, NC, GA, TN, KY, OH, MI, LK, OK, TX, and CA. | Terminated | June 04, 2008 | 128 | 100 | Not_Computer
Z-1298-2008 | 46780 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | AlloCraft DBM with Syringe Assembly, 5CC, Catalog/Reorder Number: 653005, LifeCell Corporation, Branchburg, NJ 08876 , Distributed by Stryker Spine |  |  | October 17, 2007 | 2007 | LifeCell Corporation | Improper donor screening test; diagnostic test kits were used in lieu of a donor screening test kits for HBsAG and HBcAB (later re-testing with FDA-approved donor screening test found samples non-reactive for both Hepatitis B Surface Antigen and Hepatitis B Core IgM/IgG Antibody) | Release of Material/Component prior to receiving test results | Consignee was notified by telephone, followed by a Recall letter sent by FAX on October 17, 2007 requesting return of any unused product. |  | 46 | 46 kits | All syringe assemblies were distributed to Stryker Orthopaedics, Mahwah, NJ. | Terminated | May 14, 2008 | 210 | 156 | Not_Computer
Z-1352-2008 | 46307 | No | MCB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Triage Micro ClostrIdium difficile (C.difficile) Panel, Catalog No: 96000, BIOSITE INCORPORATED, San Diego, CA 92121 |  |  | December 04, 2007 | 2007 | Biosite Inc | False Lower Results: Toxin A results in the Triage Micro C. difficile Panel have shown falsely lower visual signals (color strength) in certain lots compared to results observed with previous lots. | Process design | Consignees located in the United States received an Urgent Medical Device Recall Letter notification by United Parcel Service on 12/4/2007 and 12/14/2007. Consignees outside of the United States were notified by Federal Express on 12/4/2007 and 12/14/2007. The Customer Recall Letter with attached fax-back Verification form informed customers of the recall and instructed customers to dicontinue use, discard appropriately all affected kits and complete and FAX the enclosed Verification form. Contact information was also provided form the recalling firm. (USA-(800) 640-0640 ) |  | 4549 | 4,549 kits | Worldwide- USA, Argentina, Australia, Belgium, Canada, Switzerland, Colombia, The Czech Republic, Germany, France, Greece, Guatemala, Hong Kong, Croatia, India, Italy, Mexico, Portugal, Sweden, Turkey &amp; The United Kingdom | Terminated | March 20, 2012 | 1568 | 4549 | Not_Computer
Z-1351-2008 | 46214 | No | JJM | Osmometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Osmocoll, COP Control Reference, SS-038, Contents: 6 unit packages of 1 mL vials (3 each High and Low), Wescor, Inc., 459 South Main Street, Logan, Utah 84321. |  |  | December 26, 2007 | 2007 | Wescor, Inc | Inaccurate readings; The control does not provide an accurate calibration value/readings when performing Colloid Osmotic Pressure (COP) testing. | Nonconforming Material/Component | Consignees were notified by Recall Notification letter on 01/11/2008, and asked to arrange for return of affected product. |  | 75 | 75 vials | Worldwide- USA, Argentina, Australia, Brazil, Canada, Denmark, England, France, Germany, Greece, Hungary, Japan, Malaysia, Mexico, The Netherlands, New Zealand, Paraguay, Poland, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey, and The United Kingdom. | Terminated | March 27, 2008 | 92 | 1025 | Not_Computer
Z-1360-2008 | 46783 | No | CJE | Alkaline phosphatase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Liquid Alkaline Phosphatase Reagent Set; Catalog No. A7516-150-SAM, In vitro diagnostic; for the quantitative determination of alkaline phosphatase in serum, Sold by Pointe Scientific, Inc |  |  | February 07, 2008 | 2008 | Pointe Scientific, Inc. | May be contaminated with microorganisms. | Other | Consignees were notified via recall letter, Urgent Medical Device, dated 2/7/08, to cease using the product, conduct a recall to the user level, destroy product on hand and to notify the firm via the response form of the amount destroyed for replacement purposes. International consignees were notified via email. |  | 6 | 6 | Worldwide-USA, Canada, Chile, Cyprus, Ecuador, Greece, India, Italy, Jamaica, Korea, Nigeria, Pakistan, Philippines, Poland, Turkey and The West Indies. | Terminated | June 12, 2009 | 491 | 302 | Not_Computer
Z-1361-2008 | 46706 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Vanguard Width Checker, Part Number: 32-487062, 67.5/75 Sizing Wing, Biomet Orthopedics, Inc., Warsaw, IN; The device is a preparatory instrument used for femoral knee insertion. |  |  | February 12, 2008 | 2008 | Biomet, Inc. | The instrument's sizing line is in the wrong place. | Process control | Distributors were notified via Urgent Medical Device Removal Notice letter dated 2/12/08 to locate the instruments in their possession and at hospital sites and to return them to Biomet. Hospitals with the instruments on hand will be provided with a recall letter dated 2/12/08 which describes the nature of the problem and the removal of these instruments. The Puerto Rico consignee was notified by Biomet UK, Ltd. via letter on 2/26/08. |  | 35 | 35 (31 to U.S. and 4 international). | Worldwide: USA, Canada, China and Finland. | Terminated | June 26, 2009 | 500 | 36 | Not_Computer
Z-1311-2008 | 46869 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Cardiovascular Procedure Kit, P/N 64789-01, Wisconsin Cardiovascular Services Pump and Table Pack, Terumo Cardiovascular, Ashland, MA. (the Pump &amp; table pack contains a weighted flexible sucker, Catalog No. 804113) |  |  | January 29, 2008 | 2008 | Terumo Cardiovascular Systems Corp | Loose Flash at tip of Weighted Flexible Sucker: A loose strand of stainless steel left on the device during manufacturing by an outside supplier may result in an unretrieved metal fragment remaining in the patient. | Nonconforming Material/Component | Consignees were notified via recall letter, Urgent Medical Device Removal, dated 1/29/08 via fax or email to cease using the weighted sucker in the packs and to use an alternative sucker until replacement product arrives from Terumo. |  | 110 | 110 | Nationwide-USA including states of Missouri, Tennessee and Wisconsin-Custom Kits for 3 hospitals. | Terminated | July 07, 2009 | 525 | 1121 | Not_Computer
Z-1365-2008 | 47341 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | BIOMET 3i Certain MicroMiniplant Straight Healing Abutment, Model/Item Number: ISMHA32, 3.4mm(D) X 0mm(P) X 2mm(H), BioMet 3i, Palm Gardens, FL. |  |  | February 06, 2008 | 2008 | Biomet 3i, Inc. | Seating Problems: The Micro-miniplant Straight Healing Abutments do not fully seat onto the restorative platform of the implant leaving a gap between the abutment and the implant. | Device Design | USA customers were contacted by telephone with a follow-up fax. International customers were first contacted by E-mail then by fax letter. Each customer was instructed to check their respective inventory for the affected part and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory. They were also instructed to return the affected product to Biomet 3i. |  | 270 | 270 | Worldwide | Terminated | July 07, 2008 | 152 | 270 | Not_Computer
Z-1050-2008 | 46351 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Physio-Control LIFEPAK 20 defibrillator/monitor, AC powered external defibrillator with battery backup. Part Numbers: 3202487-073, 3202487-079, 3202488-036, Physio Control, Inc, Redmond, WA 98052 |  |  | January 08, 2008 | 2008 | Physio Control, Inc. | Inoperable Device - Potential for delay in therapy or prevention of defibrillation therapy due to corrosion of the Printed Circuit Board Assemblies (PCBA). | Nonconforming Material/Component | The firm contacted Customers by phone and by Urgent-Medical Device Recall letters sent via return receipt requested to their consignees on 1/8/08. The letter advised consignees that the Printed Circuit Board Assemblies (PCBA) might not meet quality requirements. This manufacturing deviation could lead to eventual PCBA corrosion that could render the device inoperative and not ready for use in an emergency under battery power. The firm recommended keeping the unit in service until it could be replaced. Representatives will visit each consignee and replace units. |  | 103 | 103 total devices - 84 domestic, 19 international | Worldwide Distribution --- USA including states of OK, SC, TX, and WA, and country of Germany. | Terminated | September 04, 2013 | 2066 | 103 | Hardware | Device Operation | Replace device | Remove or Replace
Z-1366-2008 | 47342 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | BIOMET 3i, EP Healing Abutment, REF/Item Number: THA53, 4.1mm(D) X 5mm(P) X 3mm(H), Sterile, BIOMET 3i, Palm Beach, FL |  |  | February 14, 2008 | 2008 | Biomet 3i, Inc. | Mis-labeling: Labeled THA53 device packaging may contain ITHA53 devices (The ITHA53 is designed to be used with an internally hexed implant--The THA53 is used with an external hex platform) | Process control | Each USA customer was first contacted by telephone and then with a faxed, Urgent Medical Device Recall letter, dated 2/25/2008. Each International customer was first contacted by E-mail and then with a faxed letter. They were instructed to return the affected product to Biomet 3i. |  | 254 | 254 | Worldwide | Terminated | July 10, 2008 | 147 | 254 | Not_Computer
Z-1349-2008 | 46889 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa IMPAX CardioVascular Suite Results Management Nuclear Reporting Module, Outbound Report, Model/Catalog No.: L7W1P, Software Versions: 2.04.17, 2.04.19 and 2.04.20, AGFA HealthCare, Westerly, RI |  |  | February 22, 2008 | 2008 | AGFA Corp. | Reporting Differences: The Outbound Report software produces an Outbound Report (text version) which does not contain all the clinical content elements present in the PDF version of the report, also generated on the Nuclear Cardiology Reporting Module. | Software design | All five (5) sites utilizing this product were notified about the issue directly via telephone. An "Urgent Safety Notice" was sent via FED-EX to all the sites using the product on February 22, 2008. The letter describes the potential issue and mitigation. The recalling firm will also visit each site and make configuration changes to resolve this issue. |  | 5 | 5 units | Nationwide to five hospitals in DC, NY, OH, TN, and VA. | Terminated | July 27, 2010 | 886 | 5 | Software | Output/Calculation | Change configurations | Other
Z-1383-2008 | 46882 | No | ODW | Drug metabolizing enzyme genotyping system. | Toxicology | Toxicology | 510(k) | Warfarin Metabolism Nucleic Acid Test Cartridge; Part Number: 20-006-002; an in vitro diagnostic drug metabolizing enzyme genotyping system; Nanosphere, Inc., 4088 Commercial Ave., Northbrook, IL 60062 |  |  | February 25, 2008 | 2008 | Nanosphere, Inc. | Aberrant Results: A warfarin 2C9*2 mutant capture signal was found to be high enough to result in an aberrant result (i.e., wild-type called heterozygous). | Other | Nanosphere sent Urgent Device Recall letters dated 2/25/08 to the direct accounts on the same date, informing them of the potential for aberrant results with the use of the affected lots of cartridges. The accounts were requested to examine their stocks for the affected lots of cartridges, segregate those found and return them to Nanosphere for replacement. Any questions can be directed to Dr. Gregg Shipp at 847-400-9115. |  | 624 | 624 units | Nationwide; USA including states of Kentucky, Michigan, Tennessee, Texas, Louisiana, Connecticut, Illinois, and California. | Terminated | May 17, 2010 | 812 | 624 | Not_Computer
Z-0617-2008 | 45444 | No | HDD | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Hill-Rom Affinity 4 Birthing Bed; Model P3700B, Hill-Rom, Inc., Batesville, IN 47006 |  |  | September 06, 2007 | 2007 | Hill-Rom, Inc. | Brake failure: The brakes may not hold or lock properly. | Component design/selection | Hill-Rom notified consignees via Urgent Medical Device Correction letter dated 9/6/07. The letter instructed facilities to continue using the beds if they pass a test of the brake mechanism; and a Hill-Rom representative will contact the facility to schedule a replacement of the detent mechanism and adjustment of the casters. |  | 1796 | 1,796 | Worldwide Distribution - USA, Argentina, Australia, Bahamas, Belgium, Belize, Brazil, Canada, China, Denmark, Dominican Republic, France, Germany, Guam, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Jordan, Kenya, Malaysia, Mexico, Netherlands, Norway, Philippines, Russia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Syria, Turkey, United Arab Emirates and United Kingdom. | Terminated | October 15, 2010 | 1135 | 1796 | Not_Computer
Z-1107-2008 | 46350 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Stryker ACL Workstation slider, Model Number: 234-010-503, Stryker, San Jose, CA 95138 |  |  | December 07, 2007 | 2007 | Stryker Endoscopy | Slider Lock Failure - The ACL Workstation slider, a component of the ACL workstation, does not consistently lock into the ACL Workstation Base. | Component design/selection | On January 9, 2008, sales representatives were notified via telephone. All consignees were mailed an Urgent: Device Removal notice, dated January 10, 2008. The firm requested that any affected inventory in stock be quarantined and returned to Stryker Endoscopy for destruction. A response form with return postage was included with notification mailed to all consignees. |  | 142 | 142 sliders were distributed. | Worldwide Distribution - including USA states of South Carolina, Oregon, Michigan, Nebraska, Maryland, California, New Mexico, Ohio, Texas, West Virginia, Massachusetts, and Florida and countries of Canada, Mexico, Hong Kong, Argentina, Brazil, France, Germany, Greece, Spain, India, Italy, Columbia, Switzerland, Poland, Sweden, Australia, and Netherlands. | Terminated | December 29, 2008 | 388 | 142 | Not_Computer
Z-1400-2008 | 46940 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncor Expression Medical Linear Accelerator, Material Number: 07360717, with control console software versions 6.5/7.2, 8.0, 9.0 and 9.1, Manufactured by Siemens Medical Solutions USA, Inc, Concord, CA |  |  | January 11, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Linear accelerator gantry may rotate in an unexpected direction and may cause patient injury on impact. | Software design | This is a mandatory software upgrade for consignees. The firm's representatives have contacted consignees via Urgent Device Notification letters dated February 14, 2008, notifying them of the new update and making arrangements to install it. |  | 5 | 5 units | Worldwide-USA including states of AL, AK, AZ, AR, CA, CO, DE, DC, FL, GA, IL, IN, KS, KY, LA, MD, ME, MI, MN, MS, MO, MN, NE, NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, VT, VA, WA, WV, WI, and WY and countries of Angola, Australia, Austria,Belgium,Brazil,Canada,Chile,Colombia,Costa Rica, Croatia,Cyprus,The Czech Republic,Denmark,DPR of Korea,El Salvador,France,Germany,Greece,Hungary,India,Iran,Ireland,Italy,Japan,Jordan,Malaysia,Mexico, The Netherlands,New Zealand,Norway,P.R. China,Pakistan,Peru,Philippines,Poland,Portugal,Republic of Korea,Romania,Russian Fed.Saudi Arabia,Singapore,Slovakia,South Africa,Spain,Sweden,Taiwan,Thailand,Trinidad,Tobago,Turkey, U.A.E.,Ukraine, The United Kingdom, and Venezuela | Terminated | May 06, 2011 | 1211 | 3292 | Computer
Z-0673-2008 | 45800 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | Parks Medical Dual Frequency Doppler Vascular Flow Detector, Model 915-BL (finished assembly without accompanying probes), Parks Medical Electronics, Inc., Aloha, OR 97007 |  |  | November 08, 2007 | 2007 | Parks Medical Electronics Inc | Premature battery failure: Wrong acid flux used on circuit board which may lead to premature battery failure of ultrasonic non-directional doppler flowmeter. | Process control | 811-B: On 9/28/07 Parks Medical Electronics notified their related U.S. distributor, Parks Medical Electronics Sales, Inc., Las Vegas, NV by email instructing them to stop further shipments. On 10/1/07 firm followed up with phone call instructing all units be returned for rework. On 11/8/07 firm sent replacement units accompanied with a Warranty Exchange Letter to their other U.S. customers notifying them of affected units. Letters were customized for each customer with a fill-in-the-blank of the model and serial numbers. Letters and replacement units were sent UPS 2nd Day Air and informed customer about the assembly error and instructions for returning the defective units. On 11/8/07 emails were sent to Brazilian and Australian distributors informing about assembly error and affected units. Rework instructions and a schematic were provided by email for each distributor to perform their own rework. On 11/28/97 Canadian customers were sent Warranty Exchange Letters and replacement units via UPS Ground.Firm is retrieving units from US and Canadian firms and will rework them. Firm is waiting to hear whether Brazilian or Austraian distributors need to sub-recall. Canadian distributors instructed to conduct subrecall if needed.All returned units will be reworked by disassembly of unit, removal of affected printed circuit board, washing board to remove acid flux, and testing. |  | 5 | 4 units distributed to 1 U.S. hospital | Worldwide Distribution - USA, Brazil, Australia, and Canada | Terminated | October 13, 2011 | 1435 | 3176 | Not_Computer
Z-1364-2008 | 47340 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | BIOMET 3i LOCATOR ABUTMENT, Catalog/Item Number: LOA003, 3.0mm, BioMet 3i , Palm Beach, FL. |  |  | February 07, 2008 | 2008 | Biomet 3i, Inc. | Mispackaging: The Locator Abutment labeled as LOA002 may contain Catalog/Item # LOA003 product and visa versa | Other | Each USA customer was contacted by telephone and then by letter, Urgent Medical Device Recall, dated 2/15/2008, sent by fax. International customers were first contacted by E-mail then by fax letter also. Each customer was instructed to check their respective inventory for the affected part and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory for product return. They were also instructed to return the affected product to Biomet 3i. |  | 97 | 97 | Worldwide. | Terminated | July 10, 2008 | 154 | 197 | Not_Computer
Z-0713-2008 | 45907 | No | FQH | Jet lavage. | General Hospital | General Hospital | 510(K) Exempt | Coloplast Elefant Suction/Irrigation Cannula for Laparoscopy |  |  | November 12, 2007 | 2007 | Coloplast Corp | Lack of sterilization assurance: Sterilization of Elefant Suction-Irrigation Devices is not performed according to ISO 11135 standard. | Process design | Coloplast sent customers an "Urgent: Recall Notification" letter on 11/12/07. The letter described that the sterilization of the products was not performed according to ISO 11135 standard. The notification requested the customers to return the product immediately and to send a fax or contact the company via email in regard to inventory at hand. A phone script for distributors also included an advice to discontinue dispensing the product and to fax or email in regard to inventory at hand. |  | 154 | 154 | Nationwide, including Puerto Rico. | Terminated | November 04, 2008 | 358 | 717 | Not_Computer
Z-1384-2008 | 46912 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Baxter HEP LOCK (Heparin Lock Flush Solution, USP) 10 u/mL-1mL vial, 25 Pack, NDC 0641-0392-25 (NDC 0641-0392-21; on vial), Baxter Healthcare Corporation, Deerfield, IL 60015 |  |  | February 28, 2008 | 2008 | Baxter Healthcare Corp. | Due to an increase of adverse events in Baxter single and multi-dose Heparin for injection products, which used the same active pharmaceutical ingredient (API) source as the HEP LOCK products | Nonconforming Material/Component | A Press Release was issued on February 28, 2008. Recall notification letters , Urgent Product Recall, were sent on February 29, 2008, by overnight mail to customers/wholesalers/ distributors and dialysis center/renal home patients. Recall letters were also sent on March 7, 2008, to health care professionals |  | 335860 | 335,860 vials | Worldwide Distribution --Worldwide: USA including Puerto Rico, and countries of Germany, Guam, America Samoa, Marshal Islands, Northern Mariana Island/Virgin Islands, and Qatar . Accounts included wholesalers, pharmacies, hospitals, and clinic/renal home patients | Open, Classified |  |  | 4603297 | Not_Computer
Z-0711-2008 | 38631 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ConMed Linvatec, Ultrapower Burs, Linvatec Corporation, Largo, FL 33773 |  |  | August 09, 2007 | 2007 | Linvatec Corp. | Compromised sterility: ConMed Linvatec medical devices packaged within a certain Peel Pouch Sterile Package may have compromised sterility due to an improper seal. | Packaging process control | Conmed Linvatec sent their consignees Urgent New Recall Medical Device Recall Notification letters On August 9, 2007 by UPS overnight delivery. The letter requests the customer return the devices to Linvatec for credit. Domestic Distributors / Sales Representatives were notified by telephone and email. |  | 1181 | 1181 units | Worldwide Distribution. | Terminated | October 01, 2009 | 784 | 1181 | Not_Computer
Z-0712-2008 | 45503 | No | JXK | Cranial electrotherapy stimulator. | Neurology | Neurology | 510(k) | CES Ultra, Cranial Electrotherapy Stimulator, Serial Numbers: 1000 through 3999, Manufactured by Arco Electronics Limited, Distributed by Neuro-Fitness LLC, Fall City, Washington 98024 |  |  | October 29, 2007 | 2007 | Neuro-Fitness LLC | No FDA clearance for the .35Hz/.45Hz frequency option on device. | Finished device change control | Consignees were notified of this recall by an Urgent Device Recall Letter sent via U.S. First Class mail on 10/29/07. The letter instructs users to contact the recalling firm to make arrangements to return the device. The letter also instructs users to return an enclosed response card. Consignees who do not respond will receive a follow-up telephone call two weeks after initial mailing of recall notice. For more information, contact firm at 1-866-937-4237. |  | 3000 | 3,000 units | Worldwide Distribution | Terminated | September 21, 2011 | 1423 | 3000 | Not_Computer
Z-0689-2008 | 46353 | No | MIH |  |  | Cardiovascular | PMA | Medtronic brand AneuRx AAAdvantage Endovascular Stent Graft with the Xcelerant Delivery System; |  |  | December 10, 2007 | 2007 | Medtronic CardioVascular | Sterility compromised: Device sterility may be compromised as evidenced by a loss of outer package integrity. | Package design/selection | Medtronic sent Urgent Communication - Medical Device Correction Advisory Notice letters to all its direct consignees on 12/7/07, via Federal Express, informing them of the affected product and providing instructions on avoiding contamination of the sterile surgical field with these devices. |  | 7000 | 7,000 | Nationwide Distribution. | Terminated | July 28, 2008 | 231 | 7000 | Not_Computer
Z-1317-2008 | 46398 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | LifeScan OneTouch Data Management Software Version 1.0, Part Number: 021-722-01(for OneTouch Select and UltraMini (Download) Blood Glucose Meters), LifeScan Inc, Milpitas, CA 95037. |  |  | December 18, 2007 | 2007 | Lifescan Inc | Meter Temporarily Freezes-- A software compatibility issue may cause the blood glucose meter to cease operations and freeze temporarily. | Software design | The firm sent an Urgent Recall letter to Health Care Professionals on 12/18/2007. The letter advised user to uninstall software drivers and provided options on how to uninstall. For further assistance, please contact LifeScan Customer Service at 1-866-679-9250. |  | 260 | 260 units | Nationwide Distribution including the states of AL, CA, IN, IO, MA, MD, MI, MS, NE, NY, NC, OR, PA, TN, TX, UT, VT, VA, WV, and WI. | Terminated | July 18, 2008 | 213 | 260 | Software | Device Operation | Uninstall drivers | Other
Z-0585-2008 | 46251 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific SlimLine 200 Micron Holmium Fiber Delivery Device, Catalog No. 840-840, Sterile, Made In Israel, Manufactured By LUMENIS, P.O.Box 240, Yokneam, ISRAEL. |  |  | December 18, 2007 | 2007 | Lumenis, Inc. | Incorrect Label: The Units outer pouch is labeled 200 micro fiber however, the inner pouch's label affixed to the product indicated that it is a 365 micron fiber. The micron fiber contained in the inner pouch is 200 micro fiber. The inner product is mislabeled with the wrong size. | Labeling mix-ups | The firm notified Consignee by certified letter on 12/18/2007, and asked that all affected products be return. Any questions concerning this recall can be addressed by contacting the Manager, Regulatory Compliance, at 1-877-586-3646 ext. 3355. |  | 59 | 59 units. | MA | Terminated | April 09, 2008 | 113 | 59 | Not_Computer
Z-1449-2008 | 46792 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Hemovac Wound Drainage Device Infection Control Kits, Catalog Number/ REF 00-2550-002-10, 400 ml-INF. Cntrl. Kit w/ 1/8" drain, 3.2 mm, 10Fr, PVC, 107 cm, medium, sterile, latex free, Rx only; Sold in packages of 10, which contain packages of 1 and 5; Zimmer U.K. Ltd., Swindon, Wiltshire, United Kingdom; Zimmer, Dover, Ohio |  |  | January 31, 2008 | 2008 | Zimmer Inc. | Leakage: The kits may disassemble at the fluid collection port, which would present a risk of exposure to blood borne pathogens to health care providers should it occur during use. | Process control | Consignees were notified via Urgent Infection Control Recall letter dated 1/31/08 to examine their inventory and return the recalled lots. Custom kit packers were instructed in the same letter to recall these products/lots from their own customers. |  | 34354 | 34,354 (Recalls Z-1448/1458-2008) | Worldwide-USA, Australia, Canada, Colombia, Denmark, Dubai, Italy, Japan, Mexico and Singapore. | Terminated | November 12, 2009 | 651 | 34354 | Not_Computer
Z-0526-2008 | 45880 | [6 empty] | BD FACS Sample Prep Assistant II, Part Number: 337170; |  |  | October 02, 2007 | 2007 | BD Biosciences | Mixing Performance: SPA II devices manufactured with the affected product motor control board may exhibit suboptimal mixing performance. | Software change control | On 10/1/07, all consignees were notified via Important Product Information letter and also an on-site visit by the firm's sales representative. The letter informed users of the affected product and provided instructions on the recall. For additional concerns or questions contact BD Customer Support Center at 1-877-232-8995. |  |  |  | Worldwide Distribution --- USA including states of: NJ, IL, TX, PA, NY, MN, and AZ, and countries of Australia, Belgium, Hong Kong, Trinidad and Tobago.Australia, Belgium, Hong Kong, Trinidad and Tobago. | Terminated | December 13, 2010 | 1168 |  | Not_Computer
Z-0587-2008 | 46317 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | MyelotecÂż Steerable Video Guided Catheter, 3.0mm, Label Insert contained in LB0020-0, LB0020-1, LB0020-2, LB0020-3, and LB0020-4 (Product Code: 2010). Myelotech, Roswell, GA 30076. |  |  | December 19, 2007 | 2007 | Myelotec, Inc. | Incorrect Labeling information: Product may contain inappropriate information in its label insert that refers to the product's use with energy delivering instrumentation. | Error in labeling | Consignees were first notified by phone on 12/19/2007. A follow up Urgent Label Insert Recall Notification letter was sent via certified mail on/about 12/28/2007. Consignees were instructed to remove the subject label inserts from the affected product in inventory, replace with the enclosed revised label inserts and to complete the attached Recall and Accountability Form. The incorrect label inserts were to be returned, via package provided to Myelotec. |  | 1252 | 1,252 shipping containers | Worldwide Distribution --- USA including states of AL, CA, CT, FL, GA, ID, IN, ME, MN, NH, NJ, NY, OH, and OR, and countries of Aurba, Korea, Japan, Switzerland, South Africa, UK. | Terminated | December 03, 2008 | 350 | 1252 | Not_Computer
Z-0916-2008 | 46380 | No | DYH | Bracket adhesive resin and tooth conditioner. | Dental | Dental | 510(k) | Grengloo (Unidose Tips), Bracket adhesive and tooth conditioner, Part Number 740-0322, Lot Number 2792584, Ormco Corporation, Glendora, CA 91740 |  |  | October 31, 2007 | 2007 | Ormco Corporation | Mislabeled: Product mislabeled with incorrect expiration date. Expiration date labeled as 2010/07 instead of 2009/07. | Process design | Ormco sent fourteen (14) consignees Urgent Device Recall Notification via U.S. First Class mail on October 31, 2007. There were two (2) additional letters sent to the consignees in France that were translated into French. The consignees were instructed to complete the Return Form and return any affected product in their inventory. Periodic effectiveness checks will be conducted in order to follow up with non-responders. For more information, contact firm at 1-800-854-1741. |  | 4 | 4 kits (Total US and Internationally) | Worldwide Distribution-USA states of IL, NY, LA, CA, NJ, NV and SC and countries of Canada, France and South Africa. | Terminated | March 22, 2012 | 1604 | 4 | Not_Computer
Z-0843-2008 | 45812 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Nonin Model 2500 PalmSAT handheld pluse Oximeter , Nonin Medical, Inc., Plymouth, MN USA |  |  | October 11, 2007 | 2007 | Nonin Medical, Inc | Incorrect Label: The label on the back of the device may show Model 2500A rather than Model 2500. | Error in labeling | An Urgen Field Action letters was sent (10/11/07) to Nonin distributors who received the potentially mislabeled product. The letters provided labeling differences and included instructions for notifying their customers, with a response form to be returned to Nonin. For additional information please contact Kim Aves at 763-577-3196. |  | 240 | 240 | Worldwide Distribution ---- USA including states of ME, SC, SD, CA, MO, NY, MI, FL, IL, PA, NC, CT, and TX, and countries of Canada, Greece, Russia, Italy, Sweden, Australia, &amp; Japan. | Terminated | December 18, 2011 | 1529 | 240 | Not_Computer
Z-0892-2008 | 46151 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Peel-Away Introducer set, 7.0 FR; Catheter introducer, includes sheath: 7.0 FR., 15.5 CM long, Wire: .035" DIA., 50 CM long, 3 MM J needle: 18 ga., 7.0 cm long. Accepts devices up to 7.0 French,sterile, Catalog Number: C-PLI-7.0-35. Cook Inc., Bloomington, IN 47404 |  |  | December 07, 2007 | 2007 | Cook, Inc. | Packaging Problem: The peel-away sheath does not peel uniformly or completely. | Package design/selection | Cook Inc. initially notified consignees via FEDEX by an Urgent: Product Recall on 12/7/07. The recall was then expanded and consignees were notified on 12/18/08 by another Urgent: Product Recall. The letters instructed any distributors to inform their customers of recall, and all users to return the recalled products and enclosed Recall Response Form to the firm. For additional information, contact 1-800-346-2686. |  | 1881 | 1,881 of total products | Worldwide Distribution- including USA and countries of Argentina, Australia, Canada, Chile, Denmark, Japan, New Zealand and Singapore. | Terminated | April 15, 2008 | 130 | 1881 | Not_Computer
Z-0915-2008 | 46347 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 10 ml Syringe Luer Lok Tip, Reorder #3039605, BD, Franklin Lakes, NJ 07417. |  |  | January 07, 2008 | 2008 | Becton Dickinson &amp; Company | Defective Seal: The convenience trays may have open seals, which can adversely impact tray sterility. | Packaging | BD notified distributors by an Urgent Product Recall Letter sent via UPS on January 7, 2008. In addition, the firm sent an Urgent Product Recall Letter on January 11, 2008 to customers identified by the distributors. The letter instructed the users to remove the product from its inventory and to return affected products to the firm using an enclosed customer packing list. For additional information, contact 1-888-237-2762. |  | 2926560 | 2,926,560 | Nationwide, including the states of CA, NY, MO, MA, IL, OH, TN, TX, MI, LA, MO, ID, GA, NE, SC and FL. | Terminated | May 19, 2008 | 133 | 4680660 | Not_Computer
Z-0668-2008 | 46275 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker GMRS Proximal Femoral Surgical Protocol; |  |  | December 18, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The GMRS Proximal Femoral Surgical Protocol states " The GMRS Proximal Femoral components are fully compatabile with all Stryker V40 femoral heads." This statement is incorrect and conflicts with the Alumina Head label and packaging insert. | Labeling design | Stryker sent Urgent Product Recall notification letters and Product Accountability forms via FedEx on 12/21/2007, return receipt requested. An Urgent Product Recall letter dated 1/7/2008, was sent to surgeons who are GMRS users, advising that the GMRS Proximal femoral components are not compatible with Alumina ceramic heads. |  | 5869 | 5869 | Worldwide Distribution - USA, Canada, Mexico, China, Australia, Hong Kong, Singapore, Sweden, Germany, Brussels, the Netherlands, Italy, Poland, Dubai, United Kingdom, Spain, and South Africa. | Terminated | October 15, 2008 | 302 | 5869 | Not_Computer
Z-0575-2008 | 45892 | No | NVI | Amphetamine test system. | Toxicology | Immunology | 510(k) | BioRad BioPlex 2200 Software, for BioPlex 2200 Instrument, an automated immunodiagnostic analyzer, Bio-Rad Laboratories, Inc., Hercules, CA 94547 |  |  | November 13, 2007 | 2007 | Bio-Rad Laboratories Inc | Incorrect values: An error was found in the BioPlex 2200 Software, Version 2.0 that results in assignment values being utilized for the BioPlex assay Calibrator Set that may be slightly different from those printed on the Value Assignment Data Sheet. Some values may be identical between the lot assignment combinations with the ANA Screen with MDSS kit. | Software design | Bio-Rad sent a Voluntary Field Correction notice dated November 1, 2007. Consignees were notified by phone call on Nov. 13, 2007, followed up by facsimile (domestic). Sub-recall to be by e-mail message for consignees outside the US. Distribution of the product ceased on Nov. 2, 2007. Bio-Rad will provide Calibrator Set CDs with only one lot of calibrator assignment values until a new version of software is available. |  | 7333800 | 85 units of calibrator lot 86280 and 16 units of calibrator lot 86290 were distributed. | Worldwide Distribution - USA, Canada, France, and Germany. | Terminated | July 23, 2008 | 253 | 7333800 | Software | Output/Calculation | Software update | Software Update
Z-0988-2008 | 46341 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Stryker Round Fluted Bur (6.0mm) Extra-Long for use with TPS Universal Drill Extra Long "R" Attachment, sterile, Stryker Instruments, Kalamazoo, MI; REF 5160-10-60. |  |  | January 10, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | Lack of assurance of sterility. | Package design/selection | Stryker sent a Medical Device Recall Notification letter dated 1/10/08 to consignees. The letter instructed them to check their facility for product, cease using it, and to return the product using a mailing label provided. International subsidiaries were notified via email. |  | 10588 | 10,588 of all products | Worldwide Distribution - USA, Australia, Brazil, Canada, England, France, Germany, Hong Kong, India, Italy, Japan, Mexico, the Netherlands, New Zealand, Portugal, Saudi Arabia, Poland, Romania, South Africa, Spain, Sweden, Switzerland and United Arab Emirates. | Terminated | December 01, 2009 | 691 | 10588 | Not_Computer
Z-0661-2008 | 45410 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON Sequoia Diagnostic Ultrasound System, general purpose Diagnostic Ultrasound system. Siemens Medical Solutions USA, Inc., Mountain View, CA |  |  | September 25, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Measurement Error affecting all Sequoia systems using the PAL video standard configuration and have a perspective PC installed: 1) The Sequoia" system is operating at a software revision below 8.0. 2) The Sequoia" system is configured to PAL video mode (normal for systems running at 240V 50 Hz.). 3) While using any of the Perspective" advanced display options listed above, an image or clip is t | Software design | Consignees were notified via a Urgent Medical Device Notification letter/customer letter sent out on 9/25/2007 under Siemens update program US013/07/S. The Siemens Service organization coordinates the implementation of the field correction with the distributors for all countries. Consignees who do not respond to the recall communication will be visit ed and/or telephoned for follow up. |  | 280 | 280 | Worldwide Distribution including the countries of: Australia, Austria, Belgium, Brazil, Brunei, China, Denmark, Finland, France, Germany, Great Britain (UK), Greece, Hungary, India, Iran, Ireland, Italy, Latvia, Netherlands, New Zealand, Romania, Russian Federation, Singapore, Spain, Switzerland, Thailand. U.S customers were not affected. | Terminated | December 20, 2010 | 1182 | 280 | Software | Output/Calculation | Software update | Software Update
Z-1073-2008 | 45545 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Non-continuous ventilator accessories under the following brand names: a) RP-P ComfortGel Cush and Flap model number 1031391 and b) RP-S ComfortGel Cush and Flap model number 1031392. Filter, mask, and/or accessories for sleep therapy and ventilation devices, Respironics, Inc., Murrysville, PA 15668. |  |  | October 15, 2007 | 2007 | Respironics, Inc. | Foreign material: Glass fragments may present in the plastic bag material used to ship components. | Environmental control | The recalling firm issued a Recall Notification dated 10/18/07 to its customers. Replacement products were to be sent. |  | 20950 | 20950 units | The products were shipped nationwide to medical facilities. | Terminated | September 01, 2009 | 687 | 20950 | Not_Computer
Z-1165-2008 | 46368 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Hemoglobin Reagent Syringe used on the CELL-DYN Sapphire Automated Hematology Analyzer. Abbott Labs, Abbott Park, IL. |  |  | January 15, 2008 | 2008 | Abbott Laboratories | Count may be out of specification--Hemoglobin background count may be out of specification (high) after installation of new CELL-DYN Sapphire Hemoglobin Reagent Syringes with packaging dates between May 8, 2007 and November 29, 2007. | Process control | On January 11, 2008, an Urgent Product Recall Letter and Customer Reply Form were sent to all currently active CELL-DYN Sapphire customers (end-users). The firm provided information for verification of replacement syringes and instructions if hemoglobin background is not within specification. A second Urgent Product Recall letter was sent on January 15, 2008. If this second Customer Reply Form is not faxed back to Abbott, customers will be contacted via an approved telephone protocol. For further information contact customer support at 1-877-4ABBOTT. |  | 2746 | 2,746 syringes were distributed. | Worldwide Distribution-USA including states of AR, AZ, CA, GT, FL, GA, ID, IL, KY, LA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NV, NY, OH, OM, OR, PA, SC, TN, TX, UT, VA, WI, and WV and countries of Canada, Mexico, Australia, Brazil, Canada, China, Chile, Columbia, Germany, Hong Kong, Japan, Mexico, Peru, Singapore, and South Korea. | Terminated | December 23, 2010 | 1073 | 2746 | Not_Computer
Z-1102-2008 | 46150 | No | LEH | Vancomycin test system. | Toxicology | Toxicology | 510(k) | Seradyn QMS Vancomycin reagents, Part Number: 0373589. Seradyn, Inc, Indianapolis, IN. |  |  | December 20, 2007 | 2007 | Seradyn, Inc. | False negative test results due to Interfering substances (heterophile antibodies) in the patient's blood may produce erroneously low results. | Device Design | Consignees were notified of the problem via Urgent Medical Device Correction letter dated 12/20/07. The letter provided information on the issues concerning the recall. The firm will revise the package insert to address the potential occurrence of false negative bias. If inconsistencies with Vancomycin level exist, confirm with an alternative test method. For further information please contact 1-800-428-4072. |  | 3682 | 1,819 kits plus bulk material for 1,863 kits. | Worldwide Distribution --- USA including states of California and New York, and countries of Australia, Belgium, Canada, Finland, Germany, Ireland, Japan, Spain and United Kingdom. | Terminated | December 30, 2008 | 376 | 3682 | Not_Computer
Z-1038-2008 | 46360 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Leonardo Workstation Picture Archiving and Communication System, Model Number: 7129534, Software Version: VB30B, Siemens Medical Solutions, Inc., Malvern, PA 19355 |  |  | December 21, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Incorrect display: The orientation labels will be incorrectly displayed on the reconstructed InSpace 3-D image if the orientation was not originally HFS on the acquisition system. | Software design | Siemens Medical Solutions issued a Customer Safety Advisory Notice dated 12/21/07 to its customers to inform them of the problem and instructions to avoid the problem until the firm's representative is available to upgrade the software. |  | 47 | 47 Units | Nationwide Distribution. | Terminated | September 30, 2008 | 284 | 47 | Software | Display/Image | Software update | Software Update
Z-0958-2008 | 45945 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 MPL Flow Cytometry System with MXP Software, Software Version 2.1, Automated differential cell counter, Part Number: 626554, Beckman Coulter, Fullerton, CA 92834 |  |  | June 29, 2007 | 2007 | Beckman Coulter Inc | Data generation error: If the cytosettings is not refreshed or restarted during data importation, the output will include old and new data. | Software design | Consignees were notified by an Urgent Product Corrective Action letter sent on 6/29/07. The letter provided details on the data error and instruction on how to avoid the error. Users were asked to retain notification as part of the Quality Systems documentation and return a response form acknowledging receipt. For additional information, contact 1-800-526-7694. |  | 589 | 516 units in the US; 73 units in Canada | Worldwide Distribution- USA and country of Canada. | Terminated | December 13, 2012 | 1994 | 589 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0973-2008 | 46200 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Coherence Impression Therapist system, Part No: 7341410, medical charged-particle radiation therapy system with software version 2.1. Siemens AG, Concord, CA. |  |  | October 02, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Image Orientation Incorrect (Image Shifts) - 1) if the reference image has been calibrated for centering and the reference image is used for patient positioning, the image will shift when the positioning tools are used. 2) This problem occurs only when non-square reference images are used, such as when images are cropped. At this point, the images will refresh, and the center of the reference im | Software design | On 10/03/2007, the firm sent Consignees a Customer Safety Advisory Notice informing them of the two potential issues with the display of reference images used for position correction. Consignees will be informed when a solution to this issue has been validated. |  | 16 | 16 units | Worldwide Distribution-USA including 48 consignees in the US and 39 outside the US. | Terminated | December 24, 2008 | 449 | 94 | Not_Computer
Z-0905-2008 | 46330 | No | EZN | Ureteral dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Pathway Balloon Expandable Ureteral Access Sheath- 11/14/28, Ureteral dilator, Catalog Number: 255100; Product Number: M0062551000, Onset Medical, Boston Scientific Corp., Natick, MA 01760 |  |  | November 30, 2007 | 2007 | Boston Scientific Corporation | Sheath Removal Difficulty: Users experience difficulty in removing the sheath. | Device Design | Consignees were sent an Urgent Medical Device Recall letter on November 29, 2007. The letter instructed users to check their inventory, remove any affected product and place them in quarantine. It also included instructions for returning the product. Users were asked to return a Reply Verification Tracking Form. Follow-up communication will be made to users who do not return the tracking form. |  | 89 | 89 | Nationwide Distribution | Terminated | October 22, 2008 | 327 | 3996 | Not_Computer
Z-1037-2008 | 46358 | No | EGG | Precision attachment. | Dental | Dental | 510(K) Exempt | Straumann NC Closure Screw, Titanium, Precision attachment, Article Number: 024.2105.04, Straumann, Andover, MA 01810 |  |  | January 10, 2008 | 2008 | Straumann Usa, Llc | Mislabeled package: Narrow Connection (NC) Closure Screws contain closure screws for the regular connection (RC) bone level implant. The RC closures screw is larger (3.5mm diameter) and therefore, will not fit into the NC implant (3.0mm diameter) narrow connection. | Mixed-up of materials/components | Consignees were notified by telephone on 1/10/08 followed by an Urgent Device Recall letter on January 10, 2008. The letter instructs users to return any affected product to the firm along with a recall acknowledgement form. For additional information, contact 978-747-2514. |  | 26 | 26 units | Nationwide Distribution including states of CA, DE, FL, IL, ME, NY, OH, OR, and TX. | Terminated | November 06, 2009 | 666 | 26 | Not_Computer
Z-1106-2008 | 46340 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Advanced Cement Mixer (ACM) Bowl and Base without 180 Gram Cement Cartridge for use with Stryker High Vacuum Cement Injection System, sterile; REF 306-554, Stryker Instruments, Kalamazoo, MI 49001. |  |  | January 10, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | Sterility may be compromised, as the packaging may have channels in the packaging seal. | Nonconforming Material/Component | Stryker notified consignees via Medical Device Recall Notification letter dated 1/10/08 to quarantine the product and to return it. |  | 150 | 25 boxes of 6 each | Nationwide Distribution - Illinois, Michigan, New York, Ohio and Wyoming. | Terminated | July 31, 2008 | 203 | 150 | Not_Computer
Z-1104-2008 | 46306 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation; for diagnostic image analysis. GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | December 28, 2007 | 2007 | GE Healthcare Integrated IT Solutions | Report missing characters: Four special characters, when entered into the exam notes, are not transferred to the preview panel or hard copy printout (greater than [&gt;], less than [&lt;], quotation mark ["], and apostrophe [']). This may not be evident to the end user, as the symbols only appear when the Exam Note window is open. | Software design | GE Healthcare IITS sent Product Safety Notification letters dated 12/28/07 to all customers who have the GE Centricity PACS RA1000 Workstation software versions 2.1.X, or 3.0.X, informing them of the potential patient safety issue involving the exam notes window, which may be used to provide various patient information including lab results. To mitigate this problem, the users were instructed not to use the following four special characters (greater than sign &gt;, less than sign &lt;, double quotation mark " and apostrophe ') in the exam notes field in order to prevent any adverse patient events. For example, if a lab result was entered as GFR &lt;60, it should be typed as GFR less than 60. The customers were requested to make sure that all necessary physicians and personnel at their location are made aware of the problem and the recommended actions provided. A software patch to permanently resolve the issue will be installed at no charge when it becomes available. Any questions were directed to the GE Customer Care Center at 800-437-1171. |  | 557 | 557 units | Worldwide Distribution - USA, Australia, Belgium, Canada, Denmark, France, Germany, Hong Kong, Israel, Italy, Japan, Korea, Kuwait, Malaysia, New Zealand, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | May 17, 2010 | 871 | 557 | Software | Output/Calculation | Software update | Software Update
Z-1118-2008 | 45911 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Fisher Healthcare Sure-Vue Serum/Urine hCG-STAT; in vitro diagnostic test; Catalog Number: FHC-A202; Product Codes: FHC-A202-KFI3O &amp; FHC-A202-KFI5O; Inverness Medical Professional Diagnostics, San Diego, CA 92121. |  |  | October 01, 2007 | 2007 | Innovacon Inc | Incorrect results: The Sure-Vue Serum/Urine hCG-STAT test, when interpreted at extended read times, may exhibit sensitivity to patient samples containing hCG at levels well below the cut-off, potentially resulting in a false positive interpretation by the user. | Labeling design | Inverness Medical sent an Urgent Medical Device Alert letter, dated October 1, 2007, to all Innovacon customers for the product via US Postal Service return receipt requested. The field correction notice removed the extended read time of 10 minutes. Customers were instructed to note that although the package insert currently allows interpretation up to 10 minutes, do not read at times greater than 3 minutes for urine and 5 minutes for serum. |  | 1825480 | 1,825,480 units | Nationwide | Terminated | August 12, 2011 | 1411 | 1825480 | Not_Computer
Z-1128-2008 | 46693 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 IntraStent LD Mega Large Diameter Biliary Stent, Model Number: S17-36, |  |  | January 18, 2008 | 2008 | Ev3, Inc | Incorrect information on label: Specific lots of ev3 IntraStent Unmounted Balloon Expandable Stents have the two symbols for length and diameter on the side and end flaps reversed. | Labeling Change Control | ev3 notified all consignees via "Medical Device Recall" letter on 1/18/08. The letter was addressed to the Hospital and it described the device affected and the problem. The letter provided recommendations to segregate the affected devices and requested the return of the "Device Recall Field Action Reconciliation Form." |  | 70 | 70 | Worldwide Distribution - USA, ARGENTINA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CHILE, DENMARK, GERMANY, GREECE, HONG KONG, IRAN, IRELAND, ITALY, JORDAN, KUWAIT, LEBANON, NORWAY, POLAND, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | November 04, 2008 | 291 | 942 | Not_Computer
Z-1119-2008 | 46346 | No | LOL | Hepatitis A virus (HAV) serological assays. | Microbiology | Microbiology | 510(k) | ADVIA Centaur HAV IgM, 100 Test Kit, used on the ADVIA Centaur or XP, REF: 05004126, in vitro diagnostic Hepatitis A test, Siemens Medical Solutions Diagnostics, Tarrytown, NY 10591. |  |  | January 12, 2008 | 2008 | Siemens Medical Solutions Diagnostics | Incorrect results: False reactive specimens on the ADVIA Centaur HAV IgM assay. | Nonconforming Material/Component | Siemens issued a Urgent Field Safety Notiifcation (078D0709-01) to all Siemens Diagnostics Branches on 1/12/08 for communication to their customers. This Customer Bulletin/Urgent Field Safety Notification advises the customer of the problem and the actions to be taken to mitigate risk. A Confirmation fax-back form is included to ensure that all customers worldwide have been notified and have complied with the mandatory action. |  | 49533 | 49,533 | Worldwide Distribution - USA, Puerto Rico, Argentina, Australia, Brazil, Canada, Chile, Colombia, Czech Republic, Egypt, France, Germany, Greece, Hong Kong, India, Israel, Italy, Malaysia, Mexico, the Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Terminated | June 20, 2012 | 1621 | 49533 | Not_Computer
Z-1036-2008 | 46348 | Yes | FJK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Diapact CRRT Assembled Kit HF/HD with Manifold and Standard Spike, item number 7210517. The product is shipped 4 units per shipping carton. Hemodialysis kit, B. Braun Medical, Inc., Allentown, PA 18109 |  |  | December 10, 2007 | 2007 | B. Braun Medical, Inc. | Faulty tubing does not prime machine as intended. | Process control | B. Braun Medical sent an Urgent Medical Device Reall letter dated 12/12/07. Customers were told to discontinue use of the tubing. |  | 956 | 956 units | Nationwide Distribution. | Terminated | October 28, 2008 | 323 | 2500 | Not_Computer
Z-0517-2008 | 46044 | No | FFT | Stomach pH electrode. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Medtronic Bravo pH Capsule with Delivery System, Stomach pH electrode, Product Numbers 9012B1011 and 9012B1001. Medtronic, Inc., Minneapolis, MN 55432 |  |  | December 03, 2007 | 2007 | Medtronic Neuromodulation | Unable to detach from source: The capsule may not detach from the delivery system following attachment to the esophageal wall. | Labeling Change Control | Consignees were notified by an Urgent Device Recall in December 2007. The letter emphasizes to users the current Instructions for Use regarding detachment of the capsule from the deliver system. It advises that users may continue to use the Bravo capsules and delivery systems. It instructs users to discard all copies of the Âż2005 Bravo pH Monitoring System Interactive Training CD and to use the enclosed Âż2007 Bravo pH Monitoring System Interactive Training. A "Bravo Delivery System Disassembly Procedure" was also provided with the letter. The letter requests consignees to complete and return the reply form included with the notification. For additional information, contact 1-800-707-0933. |  | 169365 | 169,365 individual Bravo capsule and delivery systems have been sold worldwide | Worldwide Distribution -- USA states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY and countries of Australia, Austria, Belgium, Brazil, Chile, Czech Republic, Denmark, France, Germany, Greece, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan , Korea, Kuwait, Lebanon, Liechtenstein, Luxembourg, Malaysia, Mexico, Netherlands, Norway, Panama, Portugal, Russian Federation, Saudi Arabia, Serbia and Montenegro, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates and United Kingdom. | Terminated | May 03, 2012 | 1613 | 169365 | Not_Computer
Z-1305-2008 | 46785 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Mach 1 peripheral Guide Catheter 8F, RDC Renal double Curve, .091 in (internal dia), Shaft lenght 55cm, Catalog Number: 19-866, UPN M001198660, Sterile EO, Availmed S.A. de C.V. (Mexico), Boston Scientific Corp., Maple Grove, MN 55311 |  |  | February 07, 2008 | 2008 | Boston Scientific Corporation | Mislabeled: Seven lots/ batches of Boston Scientific Mach 1 Guide Catheter may be mislabeled with the incorrect device length. | Labeling mix-ups | Consignees were sent a Boston Scientific Urgent Voluntary Medical Device Recall letter on 2/7/08. The letter included the customer steps for recall, a complete listing of all product batches being recalled, the Reply Verification Tracking Form(s) and a return shipping label. |  | 4 | 4 | Worldwide Distribution including USA states of PA, AZ, NY, KY, CA, WV, OH, WA, MI, TN, IN, IL, NJ, KY, FL, AR, FL, NC, CO, TX, WV, WI, OR, AL, MN and ND and countries of BE, FR, GB, DE, IE, IT, JP, AU, SOUTH KOREA, NZ, TR, TW and ZA. | Terminated | November 03, 2008 | 270 | 94 | Not_Computer
Z-1240-2008 | 46832 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | LIAISONÂż 25-OH Viatamin D Kit, Vitamin D test system, Model Number 310900, DiaSorin Inc, Stillwater, MN 55082 |  |  | February 01, 2008 | 2008 | Diasorin Inc. | Inaccurate measurements: Kits may sporadically recover kit and external control values out of range high. | Nonconforming Material/Component | Consignees were notified by a Diasorn Customer Notification Letter sent by email on 2/1/08. The letter instructs users to stop using the affected kits, destroy any kits remaining in inventory and to return a response form in order to receive replacement product. For additional information contact 1-800-328-1482. |  | 2267 | 2,267 kits | Worldwide Distribution including USA states of AR, CA, GA, IA, IL, MA, MD, MI, MN, MO, NC, NJ, NM, NY, OR, PA, TN, UT and WA and country of Austria. | Terminated | October 31, 2008 | 273 | 2267 | Not_Computer
Z-1179-2008 | 46802 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Active Life Little Ones One-Piece Custom Urostomy Pouch 5/16 - 1 inch, 8-25mm; Ostomy pouch and accessories, Product Order Number 020917, Nypro Dominican Republic Inc., ConvaTech, Skillman, NJ 08558 |  |  | February 14, 2008 | 2008 | ConvaTec | Mislabeled: The market unit carton label reads 5/8 - 1 inch, and it should read 5/16 - 1 inch. The metric units (8-25mm) on the label are correct. | Labeling mix-ups | Consignees were notified by an Urgent Product Recall Letter on 2/14/08. Updated letters were issued on 2/21/08 and 4/10/08 with expanded lot numbers. The letters instruct users to return any affected products and a product recall response form. For additional information, contact (908) 904-2199. |  | 3986 | 3986 market units | Nationwide Distribution. | Terminated | October 21, 2008 | 250 | 3986 | Not_Computer
Z-1141-2008 | 46414 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iSite PACS image management system; Picture Archiving and Communications System; Radiology software package, Philips Healthcare Informatics, Foster City, CA 94404. |  |  | January 04, 2008 | 2008 | Philips Healthcare Informatics, Inc. | Incorrect display of one of the patient's images. This malfunction could cause an image to be incorrectly displayed, for example flipped left to right and anterior to posterior. The original image is still presented in the correct orientation next to the incorrect image when displayed on the monitor. | Software design | Philips Healthcare Informatics mailed customers notification: Urgent Product Correction regarding FCO#2954704-01/04/08-001-C. Firm released a corrected version of software (iSite PACS v.3.5.69) on January 21, 2008. Firm is scheduling customers to upgrade to the corrected version. |  | 197 | 197 units | Worldwide Distribution: USA, United Arab Emirates, Germany, the Netherlands, Denmark, Australia, France, Italy, Belgium, Ireland, Italy, and Canada. | Terminated | October 07, 2009 | 642 | 197 | Software | Display/Image | Software update | Software Update
Z-1543-2008 | 47378 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Medifil Heparin I.V. Flush Syringe 100 units/mL; a polypropylene luer locking syringe containing Heparin Lock Flush Solution, USP, a sterile non-pyrogenic, isotonic injection; each mL in a Heparin I.V. Flush Syringe contains 100 units of heparin sodium derived from porcine intestinal mucosa and 9 mg of sodium chloride in water for injection; Medefil, Inc., Glendale Heights, IL 60139 |  |  | March 20, 2008 | 2008 | Medefil Incorporated | The heparin lock flush solution was manufactured from contaminated Heparin Sodium USP Active Pharmaceutical Ingredient (API). The FDA has received reports of serious injuries and/or deaths in patients administered finished heparin injectable products manufactured from heparin API containing this contaminant. | Material/Component Contamination | Medefil telephoned the direct accounts on 3/20/08, e-mailed them a copy of the 3/20/08 recall letter on 3/21/08, and sent the accounts a hard copy of the recall letter on 3/24/08. The accounts were informed of Medefil's recall of 19 lots of Heparin I.V. Flush Syringes produced from the active pharmaceutical ingredient (API) Heparin Sodium, USP, lot 1035-0773, that has been recalled by the manufacturer, Scientific Protein Laboratories, due to the presence of a contaminant in the recalled lot of API. The accounts were instructed to stop use of the 19 lots of Heparin I.V. Flush Syringes immediately, place them under quarantine, and contact all customers to whom they further distributed the product, telling them to stop use of the product and return the product to Medefil through their supplier. The accounts were requested to complete and return to Medefil via fax the enclosed response form, indicating the quantity on hand of each recalled lot. Any questions were directed to Pradeep Aggarwal, President at 630-682-4600. |  | 3151080 | 3,151,080 syringes | Class II Recall - Nationwide Distribution --- including states of Arizona, Florida, Illinois, Massachusetts, Pennsylvania, and Texas. | Terminated | April 15, 2010 | 756 | 3713040 | Not_Computer
Z-1602-2008 | 47436 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Tyco Healthcare Monoject Prefill 100UlmL Heparin |  |  | March 28, 2008 | 2008 | Covidien LP | Scientific Protein Laboratories (SPL), disclosed that two lots of Heparin Sodium, USP Active Pharmaceutical Ingredient has a heparin-like contaminant | Material/Component Contamination | Tyco Healthcar notifed customers by letter sent via US Postal Service beginning 3/28/08, through 4/1/08. The letter requests that they immediately stop using and return the product. Distributors are requested to in-turn notify their customers. An instruction sheet was attached for product return and replacement.All customers are required to respond to the notice regardless of remaining inventory.. |  | 21600 | 21,600 units | Nationwide Distribution | Terminated | June 25, 2012 | 1550 | 11054984 | Not_Computer
Z-0987-2008 | 46294 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | AirLife Infant nCPAP System Driver (Catalog No: 006900;) a nasal continuous positive airway pressure device, which is part of a larger system that includes an Infant Nasal CPAP generator, fixation device and patient interface (prongs and/or mask); Cardinal Health, McGaw Park, IL 60085 |  |  | January 04, 2008 | 2008 | Cardinal Health | Oxygen fluctuations - The AirLife Infant Nasal CPAP System exhibited O2 fluctuations when used with low O2 flow rates when the PTO/Auxillary port on the driver is used as a blender in administering O2 therapy via nasal cannula at flow rates of less than 4L/min. | Device Design | Cardinal Health sent an Urgent Medical Device Correction Notice dated 1/4/08 with a list of catalog number and affected serial numbers of the nCPAP Driver, and informed them that the nCPAP Driver may exhibited O2 fluctuations when used with low O2 flow rates when the PTO/Auxiliary port on the driver is used as a blender in administering O2 therapy via nasal cannula at flow rates of less than 4L/min. Customers were instructed to discontinue use of the PTO/Auxiliary port and were provided with a sticker reading "WARNINIG! DO NOT USE PTO/AUXILIARY PORT" that must be affixed to each unit in their possession. Sticker placement instructions and photo showing proper placement were provided with the letter. The accounts were informed that future upgrades to this device will eliminate the PTO/Auxiliary port and are expected to be completed in the field by June1, 2008. The accounts were requested to complete and return via fax at 847-689-9101, the enclosed acknowledgment form, indicating the number of affected drivers in their possession. Any questions were directed to Quality Systems-Professional Services at 1-800-292-9332 or 1-800-327-7503. |  | 314 | 314 units | Nationwide Distribution | Terminated | January 09, 2009 | 371 | 314 | Not_Computer
Z-0982-2008 | 46215 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | VIASYS AVEA Ventilator, Catalog numbers: 17210-00 (USA Comprehensive Version) |  |  | December 17, 2007 | 2007 | Viasys Respiratory Care, Inc.dba Bird Products | Alarm Failure -- There is a potential of alarms not being activated subsequent to specific electronic faults that may cause INOP condition. | Component design/selection | On December 18, 2007, a product notification letter was mailed to all customers in the United States via overnight carrier (registed mail) and a product notification letter was mailed to all international customers via Federal Express. The recall letter informed the customers of the problem and also inform them that a Cardinal Health (Viasys) service representative will be contacting them immediately to arrange &amp; service their ventilator. For futher information please contack Andrea von Muller at 760-778-7323. |  | 2 | 2 | Worldwide Distribution --- USA including states of AL, CA, MA, NM, NV, and NY and country of Italy. | Terminated | May 13, 2008 | 148 | 41 | Not_Computer
Z-1105-2008 | 46331 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Triathlon Baseplate Impactor Extractor, Orthopedic manual surgical instrument, Catalog Number 6541-4-805, Non-Sterile, Stryker Howmedica Osteonics Corp, Mahwah, NJ 07430 |  |  | February 23, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | Insertion/Removal Difficulties: The Baseplate Impactor/Extractor may not assemble/disassemble easily to the baseplate. | Device Design | Consignees were notified via an Important Market Withdrawal Letter sent 2/25/05. The letter instructed users to fax a Product Accountability Form to the recalling firm identifying any affected products and to return them to the firm. The letter informed users that the products will be reworked and returned at a later date. The letter also identified other products that could be used as replacements. For additional information, contact 201-831-5825 or 201-831-5110. |  | 293 | 293 total units | Worldwide Distribution | Terminated | May 30, 2008 | 1192 | 293 | Not_Computer
Z-1103-2008 | 46296 | No | GGK | Red cell lysing reagent. | Hematology | Hematology | 510(K) Exempt | White Blood Cell (WBC) Reagent-Part A used with the CELL-DYN 4000 and CELL-DYN Sapphire Systems, Red cell lysing reagent, List Number 01H77-01, Supplied in 3.8L bottles, Abbott Laboratories, Abbott Park, IL 60064 |  |  | December 19, 2007 | 2007 | Abbott Laboratories | Microbial Contamination: The WBC (White Blood Cell) Reagent Part A was confirmed positive for Pseudomonas contamination. | Material/Component Contamination | Consignees were notified by a Product Recall letter sent 12/19/07. The letter instructs users to discontinue use of affected product and destroy any remaining inventory. Users were requested to return a customer reply form. For additional information, contact 1-877-4ABBOTT. |  | 1378850 | Product was distributed in the follow amounts: 1,199 botttles of lot code 51749i2; 1,150 bottles of lot code 1574i2, 1,150 bottles of lot code 56675i2 | Worldwide Distribution including USA states of : AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WI and WV, and countries of: Argentina, Australia, Brazil, Canada, Germany, Hong Kong, Japan, Singapore, South Korea and Venezuela. | Terminated | July 02, 2008 | 196 | 1378850 | Not_Computer
Z-1462-2008 | 46740 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | AxSYM Drugs of Abuse/Toxicology Assay Disk Version 8.0; List Number: 3D54-08; Abbott La boratories, Diagnostics Division, Abbott Park, IL 60064 |  |  | January 21, 2008 | 2008 | Abbott Laboratories | Error Message: If the Positive and Negative Interpretation Cutoff parameters (116/117) are edited after installation of the revised assay file, VRTX error #0002 in Task 40 is generated and the instrument locks up when it attempts to report an AxSYM Amphetamine/Methamphetamine II patient result. | Software change control | Consignees were sent a Product Correction-Immediate Action Required letter on 1/21/08. The letter provided users with the steps to take to temporarily correct the problem until the recalling firm is able to send a corrected assay file. Users were requested to return a customer reply letter acknowledging receipt of the letter. For additional information, contact 1-877-4ABBOTT. |  | 1270 | 1,270 disks | Worldwide Distribution including USA territory of Puerto Rico, and countries of Australia, Canada, Chile, Colombia, Germany, Mexico, New Zealand, South Korea, Taiwan, Thailand and Venezuela. | Terminated | December 01, 2009 | 680 | 1270 | Not_Computer
Z-1624-2008 | 46801 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Remel RapID Inoculation Fluid, Microorganism differentiation and identification device, Packaged in 1-ml tube, Catalog Number: R8325102, 20/pack, Remel, Lenexa, KS 66215 |  |  | October 23, 2007 | 2007 | Remel Inc | Failure to produce reaction: Use of the product with various RapID Identification systems panels may exhibit poor or no reactions. | Other | Consignees were notified by an Important Medical Device Customer Advisory Notice on 10/23/07. The letter instructs users to review inventory and discard any affected products. The letter also instructs users to return a product inventory checklist by 11/23/07. For additional information, contact 1-800-447-3641. |  | 500 | 500/1-ml packs | Worldwide Distribution including USA and countries of Taiwan, Japan, and U.K. | Terminated | July 16, 2008 | 267 | 1054 | Not_Computer
Z-1461-2008 | 47401 | No | HNR | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Alcon/Grieshaber Morris ILM Forceps, Direct Action; Manual ophthalmic surgical instrument, Catalog Number: 641.22; Non-sterile/Reusable; Alcon Laboratories, Inc., Fort Worth, TX 76134. |  |  | March 07, 2008 | 2008 | Alcon Research, Ltd | Device corrosion: Forceps are prone to corrosion and possibly premature fracture and/or malfunction. | Device Design | Consignees were notified by an Urgent Medical Device Recall that was sent via fax on 03/07/08. Users were instructed to cease using the affected products and to return a Recall Acknowledgement Form to the recalling firm. The sales representative will pick up the affected product and discuss replacement alternatives. For additional information, contact 1-800-757-9780. |  | 24 | 24 units | Nationwide Distribution including states of AL and CA. | Terminated | July 18, 2008 | 133 | 24 | Not_Computer
Z-1138-2008 | 46373 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo US Workplace Picture Archiving and Communication System, software version 2.0, Siemens Medical Solutions, Inc., Mountain View, CA 94043 |  |  | December 06, 2007 | 2007 | Siemens Medical Solutions USA, Inc. | Inaccurate results: A software bug may cause inaccurate wall motion abnormality scoring results to be displayed. | Software design | Siemens Medical Solutions sent an Urgent Medical Device Notification letter, dated November 2007, to its consignees. The letter described action required by users, and that Siemens will contact the user with a new software version correcting this error. A field correction to update software is planned to be underway, designated to start on 12/21/2007, and the correction is expected to be completed within 45 days. |  | 17 | 17 units | Worldwide Distribution: South Korea, Malaysia, Lithuania and China, and 1 unit to New Jersey in the USA. | Terminated | December 20, 2010 | 1110 | 17 | Software | Output/Calculation | Software update | Software Update
Z-1616-2008 | 47205 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | 9051 Electrode Adapter; Automated external defibrillator, Used with QUIK-COMBO System; Cardiac Science Corp, Bothell, WA 98021. |  |  | January 23, 2008 | 2008 | Cardiac Science Corporation | Adaptor Cable Incorrectly Manufactured: The red and white connector ends of the adapter are switched. Therefore, the cable will not connect to the electrodes as indicated in the instructions. | Process control | Consignees were notified by an Urgent Medical Device Recall letter sent on 01/23/2008. The letter instructs users to check each cable of the affected lot to ensure that it properly connects to the electrode. Users were also instructed to return any defective product to the recalling firm for a replacement. For additional information, contact 1-888-466-8686 x2041. |  | 58812 | Lot W029275 - 24; Lot W029475 - 38 | Nationwide Distribution including states of CT, IA, IN, NC, NY, OH, PA, WA, and WI. | Terminated | July 21, 2011 | 1275 | 58812 | I/O | Device Operation | Check/Replace | Remove or Replace
Z-1614-2008 | 46905 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Fiducial Markers (0.9 x 3 mm) Soft tissue gold markers, Medical charged-particle radiation therapy system; Part Number: MT-NW-887-809, CIVCO Medical Solutions (MEDTEC), Orange City, IA 51041. |  |  | February 28, 2008 | 2008 | Med Tec Inc dba CIVCO Medical Solutions | Mislabeled: Fiducial markers were incorrectly labeled. They are marked (0.9x3mm cylindrical) Soft tissue gold markers, but contain (2mm spheres) Bone gold markers. | Labeling mix-ups | Consignees were notified via telephone beginning 2/28/08 and a Recall Notification letter sent on 2/29/08. Users were requested to return any affected products for a replacement. They were also requested to return a recall notification form indicating the amount of affected product in their inventory. For additional information, call 800-842-8688. |  | 33 | 33/3-marker pouches | Nationwide Distribution including states of WA, PA, DE, AL, CT, WI, FL, TX, MN, and IL. | Terminated | July 02, 2008 | 125 | 69 | Not_Computer
Z-1608-2008 | 46863 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MT-APSD-2.4 Type-S Thermoplastic Mask, Medical charged-particle radiation therapy system; Disposable non-sterile (43.2cm) standard perforated thermoplastic head, neck and shoulder mask (2.4mm), 1 device per package; CIVCO Medical Solutions (MEDTEC), Orange City, IA 51041. |  |  | February 25, 2008 | 2008 | Civco Medical Instruments Inc | Mislabeled: The mask is mislabeled with the incorrect part number but the outer packaging is correct. | Labeling mix-ups | Consignees were notified by a Recall Notification letter on 2/25/08. The letter informed users that the recalling firm will replace the affected product. It also requested that users return a recall notification acknowledging notification of recall. Users were instructed to contact their sales representative to arrange servicing or for additional information at 1-800-842-8688. |  | 160 | 160 | Worldwide Distribution including USA states of IL, PA, NC, TN, SD, CA, MI, WI, MA, and CO; and country of Japan. | Terminated | July 02, 2008 | 128 | 160 | Not_Computer
Z-1132-2008 | 46741 | No | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(K) Exempt | Keeler All Pupil II Indirect Ophthalmoscope, Keeler Instruments, Inc., Broomall, PA 19008 |  |  | January 08, 2008 | 2008 | Keeler Instruments Inc | Good Manufacturing Practices (GMP) deficiencies may compromise the safety and effectiveness of the device. Firm is on import alert 89-04 and released product. | Process design | Keeler Instruments issued an email dated 1/8/08 to the distributor requesting that they return the product. |  | 25 | 25 units | The products were shipped to a distributor in NJ. | Terminated | August 12, 2008 | 217 | 25 | Not_Computer
Z-1110-2008 | 45909 | No | CKA | Alanine amino transferase (ALT/SGPT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific Liquid ALT (SGPT) reagent set for the quantitative determination of alanine aminotransferase in serum, 150 ml, Catalog no. A7526-150, in vitro diagnostic, Pointe Scientific, Inc.; Canton, MI 48188 |  |  | December 14, 2007 | 2007 | Pointe Scientific, Inc. | Failure of the reagent to produce test results. The R1 reagent may be contaminated with microorganisms. | Process control | Pointe Scientific notified consignees via Urgent Medical Device Recall letters dated 12/14/07, instructed to dispose of the product and to send back a reply form to request replacement product. |  | 955 | 955 for all products. | Worldwide Distribution - USA, Bolivia, Canada, Chile, China, Cyprus, Greece, Indonesia, Israel, Korea, Mexico, Pakistan, Panama, Peru, Philippines, Republic of Yemen, Russia, South Africa, Trinidad, Ukraine and West Indies. | Terminated | September 15, 2008 | 276 | 955 | Not_Computer
Z-1094-2008 | 45960 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron LX20 Clinical Chemistry System, Part Numbers: 466200,473032, Beckman Coulter, Inc. |  |  | April 05, 2007 | 2007 | Beckman Coulter Inc | Intermittent Failure of Stirrer Motor-Beckman Coulter has determined that incorrect Glucose results on LX20 orDxC instruments have been related to an intermittent failure of the stirrermotor.-Stirrer motors can stall without any flags or motion errors.-If a stall occurs during a GLUm test near the time of sample inject,results may be affected.-Affected GLUm results may be high or low. | Software design | GLUm stir motors greater than one year old were replaced. Replacements began April 5, 2007. Based on the Risk Assessment, BCI decided that Customer Notification is not needed. GLUm stir motors greater than one year old will be replaced on high volume LX/DxC systems by 3rd quarter of 2007. Replacements began April 5, 2007. A new stirrer design is being developed and is expected to be available before 4th quarter of 2007. Implementation Strategy of Stock Control: 1. A Stop Ship was not necessary based on risk assessment. The possibility of stirrer motor failure may increase with time and duty cycles. Systems running a high volume of tests may have a greater likelihood of exhibiting stirrer motor failure with motors greater than one year old. 2. New Customers were placed on a preventative maintenance plan to have their motor replaced at one year cycle. Completed 3. A Technical Update was issued to notify Hotline and Field Service Personnel of the issues. |  | 370 | 344 in the US; 26 in Canada | Worldwide Distribution - USA including Canada. | Terminated | January 15, 2013 | 2112 | 2599 | Not_Computer
Z-1166-2008 | 46389 | No | CGZ | Chloride test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinical Chemistry ICT Calibrator, List Number: 1E46-02, Abbott Labs, Abbott Park, IL 60064 |  |  | September 26, 2007 | 2007 | Abbott Laboratories, Inc | Quality Control (Out of Range)- ICT Calibrator lot number: 0505017, may contain an interferent. Not all vials within a kit are affected.The investigation determined that the interferent affects Potassium (K+) only. Use of the affected calibrator vials will generate an acceptable, but low calibration slope. When Quality Control (QC) is run to verify the calibration, K+ QC results may be bel | Component design/selection | An Immediate Action Product Recall dated 9/26/07, was provided to all AEROSET and ARCHITECT cSYSTEM customers that have received Lot Number 0505017. The effectiveness check will be completed through a customer reply form for 100% of US customers. Outside of the U.S. each individual country will determine the methods of effectiveness check.Necessary Actions:If the ICT calibration generates an acceptable slope for (K+), and the calibration is verified by generating QC results within the acceptable range, then the customers may continue to use the lot until a new one is available.If the ICT calibration generates an acceptable slope for (K+) with results below QC range, and (Na+) and (Cl-) are not affected, then they should discontinue use of the Calibrator vials that generate the unacceptable calibration and destroy them following their laboratory procedures.If (Na+) and/or (Cl-) calibration slopes or QC results are not acceptable, additional troubleshooting is required.Questions may be addressed to Abbott Customer Support at 1-877-4ABBOTT. |  | 1796 | 1,377 Distributed Worldwide; 419 Domestically | Worldwide Distribution --- including USA, Puerto Rico and countries of Mexico, Ecuador, Brazil, Canada, Hong Kong, New Zealand, Venezuela, Chile, , Germany, Thailand, Trinidad/Tobago, Columbia, Argentina, El Salvador, Singapore, Australia, and the Bahama. | Terminated | March 27, 2012 | 1644 | 1796 | Not_Computer
Z-1644-2008 | 47526 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | EVMS (Enterprise Visual Medical System); Picture archiving and communications system; Model Numbers: 5.30.4, 5.30.5, 5.30.6, 5.30.7, 5.40.0 and 7.0.0.1. Emageon, Inc., Hartland, WI 53029. |  |  | January 01, 2008 | 2008 | Emageon | Miscalculation: The calculation of the standard uptake value does not include the required calibration factor/decay correction factor. The resulting value is incorrectly calculated, which may lead to incorrect interpretation of tissue density. The calculation error exists in both the software and user guide. | Device Design | Consignees were notified by Urgent Device Correction letters that were sent on 1/9/08 and 1/18/08. The letters instructed users not to use the standard uptake value in EVMS for diagnostic interpretation when viewing PET studies unless they received service bulletin 445, 446 or 7.0.1. Users were also instructed to replace pages in the user's guide with the pages provided and to affix the enclosed warning sticker on all workstations. Users will be contacted by the recalling firm to schedule a software patch. |  | 99 | 99 | Nationwide Distribution including states of: MI, WI, CA, MN, AZ, CO, MN, MD, DC, MO, TX, IL, GA, FL, WI, SC, PA, NY, KS, TN, AL and OK. | Terminated | June 11, 2009 | 527 | 99 | Software | Output/Calculation | Software/manual update | Software Update
Z-1546-2008 | 47386 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Selenia Full Field Digital Mammography System with Software v. 3.3.1.1; Full field digital, system, x-ray, mammographic; Model Number: SEL-00002; Hologic, Danbury CT 06810 |  |  | February 28, 2008 | 2008 | Lorad, A Hologic, Inc. | Marker Misalignment: The new software version contained a magnification factor that shows CAD markers misaligned with the identified indications. The marked area is misaligned up to 2cm from the indication. | Software design | Consignees were notified by an Urgent Recall Notice sent on 3/21/08. The letter informed users that representatives were dispatched to all locations and they have corrected the magnification factor problem. The letter instructed users to re-evaluate any cases read using CAD from the period of installation until the date of system correction. The letter also instructed users that only images read with CAD markers are affected. The letter requests users to return a confirmation of notification. For additional information, contact 1-877-371-4372. |  | 73 | 73 units | Worldwide Distribution including USA and countries of Canada and the Republic of Korea. | Terminated | February 03, 2009 | 341 | 73 | Software | Display/Image | Software update | Software Update
Z-1137-2008 | 45959 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | MultiDiagnost Eleva Flat Detector X-Ray System; Part Number: SW 9896 001 43361; Philips Medical Systems, |  |  | November 29, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | Delay in Imaging: MultiDiagnost Eleva with Flat Detector x-ray system cannot perform fluoroscopy procedures while certain image file housekeeping operations are performed. Fluoroscopy cannot be performed and new exposures are not able to be acquired until the housekeeping operations are complete which may take several minutes. The loss of system availability could occur during an interventional | Software design | On 11/29/07 Philips Medical Systems sent a letter dated 11/28/07 to customers stating that if users perform certain image file housekeeping operations, fluoroscopy cannot be performed and new exposures cannot be acquired until those operations are complete. It also instructs that the unavailability of fluoroscopy and exposure can be avoided by not pressing the "Delete All Images" button until any pending Autopush operation is complete. FCO 70800067 will be issued to service representatives to install new software. |  | 18 | 9 units were distributed to hospitals in the US, 6 were installed, 3 were not installed or in use. | USA Distribution - CO, MN, NC, NJ and TX. | Terminated | September 17, 2010 | 1023 | 18 | Software | Display/Image | Software update | Software Update
Z-0580-2008 | 44725 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur TnI-Ultra Assay, Troponin Immunoassay, 500 tests - Part/Catalog # 02790309, and 100 tests - Part/Catalog # 02789602. For in vitro diagnostic use, Siemens Medical Solutions Diagnostics, Tarrytown, NY 10591 |  |  | June 07, 2007 | 2007 | Siemens Medical Solutions Diagnostics | Incorrect results: Falsely elevated Troponin values have been reported which are inconsistent with the patient's clinical picture and test negative by other Troponin assays. | Material/Component Contamination | Two Support Bulletins (078N0361-01 &amp; 078N0361-02) were sent by e-mail to all affected Siemens Diagnostics Branches to inform them of this issue. The first one was sent on 6/07/07 and the second one was sent on 7/05/07. Two Customer Bulletins (078D0647-01 &amp; 078D0651-01) were sent by e-mail to the Siemens Diagnostics Branches for communication with the affected customers. The first one was sent on 6/07/07 and the second one was sent on 7/05/07. These bulletins state that if a false positive is suspected, the sample should be diluted and retested. |  | 58818 | Domestic: 31,497 kits; Foreign: 27,321 kits | Worldwide Distribution - USA, Argentina, Austria, Australia, Brazil, Canada, Chile, Colombia, Egypt, France, Germany, Greece, Hong Kong, India, Israel, Italy, Malaysia, Morocco, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Venezuela | Terminated | August 28, 2012 | 1909 | 58818 | Not_Computer
Z-1139-2008 | 46393 | No | LZS |  |  | Ophthalmic | PMA | AMO VISX STAR Excimer Laser System, manufactured by VISX, Santa Clara, CA 95051 |  |  | December 13, 2007 | 2007 | VISX INCORPORATED, A SUBSIDIARY OF AMO INC | Firm received two complaints (including one injury complaint) associated with the movement of the VISX chair while in the patient loading ("fully swung out" unlocked) position under an IntraLase FS laser. | Component design/selection | On December 13, 2007, firm issued an Urgent Safety Notice to all VISX and IntraLase laser owners related to recall. The Safety Notice contains information related to the measures that can be taken to prevent unintentional movement of any chair during other procedures. AMO is initiating a field corrective action, whereby Field Service Engineers will visit each practice that owns both systems to disable the VISX chair back switch. For additional information, please contact AMO at (714) 247-8656. |  | 1853 | 1,853 units | Worldwide Distribution - USA, including the states of: AKALARAZCACOCTDCDEFLGAHIIAIDILINKSKYLAMAMDMIMNMOMSMTNCNDNENHNJNMNVNYOHOKORPAPRRISCSDTNTXUTVAVTWAWIWVand countries of: ArgentinaAustraliaAustriaBahrainBelgiumBoliviaBrazilBulgariaCanadaChileChinaColombiaCosta RicaCyprusCzech RepublicDjiboutiDominican Rep.EcuadorEgyptFinlandFranceGermanyGreat BritainGreeceHong KongIndiaIndonesiaIrelandIsraelItalyJapanJordanKasachstanKuwaitLebanonMalaysiaMartiniqueMexicoMongoliaNetherlandsNew ZealandNorwayOmanPakistanParaguayPhilippinesPolandPortugalRussian Fed.Saudi ArabiaSingaporeSlovakiaSouth AfricaSouth KoreaSpainSwedenSwitzerlandTaiwanThailandTrinidad,TobagoTurkeyUnit.Arab Emir.VenezuelaVietnam | Terminated | May 05, 2010 | 874 | 1853 | Not_Computer
Z-1140-2008 | 46394 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | B-K Medical Intraoperative Transducer Type 8814, diagnostic ultrasound transducer, B-K Medical Systems, Inc., Wilmington, MA 01887 |  |  | January 24, 2008 | 2008 | B &amp; K Medical Systems, Inc. | Crack: Incomplete glue joint in the device housing may affect electrical safety or sterilization. | Process control | B-K Medical issued letter notification on 1/23/08 to accounts via Fed Ex, advising them not to use the unit until a representative can inspect and exchange, if unit is found defective. Instructions are also provided to continue use of the device with a sterile cover until site visit. |  | 7 | 7 | USA Distribution: MA, MI, PA, TX, and OR. | Terminated | June 23, 2008 | 151 | 7 | Not_Computer
Z-0842-2008 | 46859 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Proteus XR/a Radiographic X-ray System, Model Number 2259976; previously marketed as the Silhouette FC, GE Healthcare, Waukesha, WI 53188. |  |  | November 01, 2007 | 2007 | GE Healthcare | The warning label that is required by 21 CFR 1020.30(j) was not on the control console of the Proteus XR/a Radiographic X-ray System, Model 2259976. | Process control | GE issued a Field Modification Instruction to their field service representatives, released in December 2007, with instructions to replace the labels on the console with the appropriate labeling. Consignees will be visited by GE field service representatives beginning December 2007. All consignees will have the required labels affixed to their units by July 31, 2008. |  | 400 | 400 Units | Worldwide Distribution, including USA, Canada, Mexico, France, Hong Kong, Italy, Korea, China, Poland, UK, New Zealand, Saudi Arabia, Taiwan, Turkey, and Australia. | Terminated | December 18, 2011 | 1508 | 400 | Not_Computer
Z-1711-2008 | 47335 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Grey Inter-Unit Interface Connectors ("IUI") on the following Alaris System modules built on or before March 26, 2004: (1) Alaris PC point-of-care unit ("PC Unit") (Model 8000)* (a.k.a. Medley PC Unit), |  |  | March 05, 2008 | 2008 | Cardinal Health 303 Inc DBA Alaris Products | A channel error on an affected Alaris System module will stop an active infusion (or monitoring) with an audible and visual alarm. Channel errors with the Alaris System due to the Inter-Unit Interface (IUI) connectors on all Alaris System modules (manufactured on or before March 26, 2004). Grey Inter-Unit Interface Connectors ("IUI") on the following Alaris System modules built on or before March | Device Design | An Urgent Medical Device Recall Notification letter was sent to consignees on April 10, 2008. Four customer letters were issued. The U19 IC issue affects only the Pump module and the U9 IC issue affects only the Syringe and PCA modules. Therefore there will be one letter for customers affected by all three issues Ul9 IC, U9 IC, and the IUI, one letter for affected Ul9 IC and U9 IC customers, one letter for U19 IC-only customers and one letter for U9 IC-only customers. All mailings were conducted through consignees (distributors as well as direct customers for the affected devices). Distributors will be sent multiple copies of the Recall Notification and Customer Response Cards. The Recall letters all include Required Actions sections which give their customers specific information about the issue and what the customers are to do and includes a web site link for each specific problem. Each recall letter also has chart of Cardinal Health contacts support. |  | 77515 | 77,515 Units | Class II Recall - Worldwide Distribution --- USA including states of AR, AZ, CA, CT, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD MI, MN, MO, MS, NC, NE, NH, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI and WY, Guam and country of Canada.****CDRH Comments: "Note for District Office to follow-up" The firm has not provided a grand total of devices affected by these problems. It is possible that devices suffering from U19 &amp; U9 problems may also experience IUI problems. This recall has 3 problems with 6 modules. One problem affects only 2 modules, one problem affects 6 modules. For simplicity and expediency, CDRH agrees to keep the file as is, rather than separate into 6 products and have complicated explanations of the problems.77,515 devices affected by IUI problem.85,286 devices affected by U19 problem.8,286 devices affected by U9 problem. **** | Terminated | September 13, 2011 | 1287 | 171087 | I/O | Device Operation | Instructions | Safety Notice/Insructions
Z-1178-2008 | 46795 | No | LRW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Presource Standard Sterile OR Scissors; straight scissors, S/B, 5.5", made in Pakistan; individually packaged in rigid styrene roll stock with a tyvek top, 50 units per case; Cardinal Health, McGaw Park, IL 60085; catalog #SSI-0003 |  |  | February 12, 2008 | 2008 | Cardinal Health | The product has the potential of being non-sterile. | Process control | Cardinal Health notified the accounts who ordered the affected scissors via letter dated 2/12/08, informing them that the scissors that were labeled sterile were subject to a sterility test failure. Upon review of the letter, the accounts were requested to inspect their inventory and immediately segregate any affected product and to telephone Cardinal Customer Sales Operations Group at 1-800-766-0706 for return instructions and to obtain credit/replacement. Any questions were directed to Cardinal's Quality Systems, Professional Services at 800-292-9332. |  | 16 | 16 cases | USA Distribution: Kansas and Missouri | Terminated | October 16, 2008 | 247 | 16 | Not_Computer
Z-1156-2008 | 45875 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | BIOPLEX 2200 SYPHILIS IgG Kit: 100 tests - Catalog No. 665-1450 - manufactured by Bio-Rad Laboratories in Redmond, WA 98052. The product is a multiplex flow immunoassay intended for the qualitative detection of Treponema pallidium IgG antibodies in human serum. |  |  | November 13, 2007 | 2007 | Bio-Rad Laboratories | False negative results due to reagent packs exhibiting low signal. | Nonconforming Material/Component | On 11/13/2007, the firm faxed Voluntary Field Correction Letters (which are dated 11/9/2007) to its consignees and also contacted them via telephone. The Voluntary Field Correction Letters: a) informed them of the potential for false negative results; b) recommended that they perform QC testing daily with each reagent pack; c) provided a new insert for the Instructions for Use; and d) provided a Response Sheet for each consignee to complete and return to the firm. |  | 796 | 796 kits | Worldwide Distribution -- USA, Canada, France, United Kingdom, Italy, Germany, South Africa, and Switzerland. | Terminated | May 23, 2011 | 1287 | 11616 | Not_Computer
Z-1176-2008 | 46745 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | CalMed Pericardial Sump, with 1/4"/0.635 cm Connector, Single Use Only. Do Not Reuse; REF SU-20602; CalMed Laboratories, Costa Mesa, CA 92626 |  |  | January 24, 2008 | 2008 | California Medical Laboratories Inc | Fragment of material: The product may have a loose, stainless steel stringer flash located at the distal tip of the product. | Process control | CalMed emailed letters to consignees on 1/24/2008 notifying customers of the recall. The letter advises that that the following Medical Devices are subject to an immediate Voluntary Recall forCorrection: Pericardial Sump (Catalog #SU-20601), Pericardial Sump (Catalog #SU-20602), and Malleable Pericardial Sump (Catalog #SU-20802). It is asked that further distribution or use of any remaining product should cease immediately. Where appropriate, the firm asks that consignees notify their own customers who received the product of the same.The Product is being recalled due to a potentially loose, stainless steel stringer flash located at the distal tip of the Product. Customers are asked to contact the firm's Quality Assurance Manager immediately for purposes ofissuing a Returned Goods Authorization and coordinating the return of the Subject Medical Devices such that they may be inspected and nonconforming product removed. A response is requested via e-mail. |  | 31124 | 31,124 units of all models | Worldwide Distribution - USA, Turkey, Brazil, Germany, Italy, Spain, Philippines, South Africa, Japan, Israel, Saudi Arabia, United Arab Emirates, India, Ireland, China, Hungary, Poland, Taiwan, Belgium, Pakistan, Lithuania, Greece, Canada, Syria, Israel, Austria, Denmark, Bangladesh, Costa Rica, Mexico, Spain, and Sultanate of Oman. | Terminated | March 05, 2013 | 1867 | 31124 | Not_Computer
Z-1155-2008 | 44884 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE LOGIQ 3 Expert ultrasound scanner with software versions 4.0.5, 4.1.1 and 4.1.2. GE Healthcare P.O. Box 414, Milwaukee, Wisconsin 53201 |  |  | September 04, 2007 | 2007 | GE Healthcare Clinical Systems | Inaccurate guidelines: The BE9C is an optional, small radius Bi-planar convex probe intended for use with the LOGIQ 3 Expert system in Urology application. An attempt by the user to activate the biopsy guidelines while this probe is in use will cause incorrect guidelines to be displayed on the image. The incorrect guidelines are recognized by their orientation that is left/right reversed from t | Software change control | GE sent consignees a "GE Urgent Medical Device Correction" letter dated 8/14/07. The letter informed the customers about the problem, offer a short-term workaround and informed them that a representative will schedule a visit to upgrade the software. |  | 62 | 62 | Worldwide Distribution - USA (NY &amp; AL), Korea, China, India, Brazil, Mexico, Germany, Greece, Italy, Poland, Russia, Saudi Arabia, Spain, and Czech Republic. | Terminated | May 29, 2012 | 1729 | 62 | Software | Display/Image | Software update | Software Update
Z-1173-2008 | 46691 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident PSL Acetabular Shells; Trident PSL HA Solid Back, |  |  | January 21, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Foreign material: Some of the parts tested exceeded Stryker Orthopaedics internal acceptance criteria for manufacturing residuals. | Process design | Recall notification letters were sent to Stryker Branches/agencies, OR Supervisors and Chief of Orthopaedics on 1/21/08. A Patient information sheet was sent on 2/4/08 to surgeons and hospitals. Per call with Center 2/4/08, recall letter to be revised and RES updated upon receipt.The letter was revised and sent on February 28, 2008 to include Trident PSL Acetabular shells. This letter was sent to Risk Management at hospitals and included the scope of the recall , the potential hazard and recommendation to physicians to monitor patients consistent with care for those receiving total hip replacement. It was also sent to surgeons and a revised letter to Stryker branches. A product acknowledgement form was included in all letters to indicate receipt of letter and quantity of product on hand, if applicable. |  | 10191 | 10191, both PSL types | Nationwide distribution. | Terminated | September 25, 2012 | 1709 | 10191 | Not_Computer
Z-1180-2008 | 46716 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System, Model number IS2000 Revision number A51 P5; endoscopic surgical control system, Intuitive Surgical, Inc., Sunnyvale, CA 94086 |  |  | December 27, 2007 | 2007 | Intuitive Surgical, Inc. | Delay in responding: In certain circumstances, the device may not respond immediately to a user's command such as master clutch or camera control. | Software design | Beginning 12/28/2007, Intuitive Surgical Field Service Representatives visited each site to correct the issues. A notification letter was provided to all sites. |  | 9 | 9 units | USA Distribution: TX, NJ, WA, NV, HI, MI, NH and CA | Terminated | August 15, 2008 | 232 | 9 | Not_Computer
Z-1161-2008 | 46204 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | S/5 iCentral, Cardiovascular Monitoring Devices, Versions L-NET03, L-NETC03, L-NET05 and L-NETC05, GE Healthcare, Wauwatosa, WI 53226. |  |  | December 13, 2007 | 2007 | GE Healthcare | Failure to alarm: An active monitor could become disconnected from the unit without any notification or alarm. | Device Design | Consignees were notified via a Medical Device Correction letter on 12/13/07. The letter explained the circumstances in which the "connection to monitor loss" alarm would not display, if a monitor was disconnected. The letter also included an attachment entitled "How to ensure a 'Connection to monitor lost' alarm on iCentral." For additional information, contact 1-800-558-7044. |  | 2146 | 2146 | Worldwide Distribution including USA states of AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, ME, MI, MN, MO, MS, NC, NH, NJ, NY, PA, SC, TN, TX, VA, VT, WA and WI; and countries of Algeria, Australia, Austria, Barbados, Belgium, Bolivia, Brazil, Bulgaria, Canada, Chile, China, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Guans, Hong Kong, Hungary, Iceland, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Libyan, Macedonia, Malaysia, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Syrian Arab Republic, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela and Yemen. | Terminated | August 13, 2013 | 2070 | 2146 | Software | Alarm/Message | Instructions | Safety Notice/Insructions
Z-0189-1 | 2263 | [6 empty] | Immunosimplicity brand anti-SSA Test kit, Catalog 720-250, Each Kit |  |  | October 31, 2000 | 2000 | DIAMEDIX CORP. | Unexpected equivocal/weak positive results |  | The firm notifed both clients on 3/21/03 during their service telephone/e-mail communications to correct the problem, and they were given a correction work around. A software patch was installed at both clients on 6/25/2003. |  | 113 | 113 | Product was distributed to 18 domestic (user accounts, e.g. labs) in TN, GA, NY, MS, OH, PA, AL, CA, VA, KY, CA, MO, HI, AND TO 5 international accounts as follows: Delta Biologicals, Cipollia and Zanetti, Fiumicino Airport, Roma, Italy; Movaco, S.A.,Iberconco S.A., Loading terminal office NR-9, Barcelona Airport, Spain; Rochem Biocare Columbia S, Deposito Snider y Cia, S.A., Sant Fe de Bogota, Colombia; | Completed |  |  | 113 | Computer
Z-1160-2008 | 45955 | No | CFP | Immunoreactive insulin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Ultrasensitive Insulin Assay Kit, Part Number 33410. The product is a paramagnetic particle, chemiluminescent immunoassay for the quantitative determination of insulin levels in human serum and plasma (EDTA) using the Access Immunoassay Systems. Beckman Coulter, Inc., Chaska, MN 55318. |  |  | September 19, 2007 | 2007 | Beckman Coulter Inc | False negative results when used to test serum samples (as compared to plasma samples). | Other | On September 19, 2007, Beckman Coulter mailed to its consignees an Urgent Product Corrective Action Letter informing them that the affected lots of the product were giving false negative results when used to test serum samples (as compared to plasma samples). The letter instructed consignees to discontinue using the affected lots of the product, and to re-test using an EDTA plasma sample if there was any discordance between the results of the serum sample (if it was obtained using an affected lot of the product) as compared with the clinical presentaion of the patient and the results of other diagnostic tests. The letter also instructed consignees to return, within 10 days, the Response Sheet . |  | 8537 | 8,205 units in the US; 332 in Canada | Worldwide: including USA and Canada. | Terminated | May 09, 2012 | 1694 | 8537 | Not_Computer
Z-1162-2008 | 46280 | No | BTC | Reservoir bag. | Anesthesiology | Anesthesiology | 510(K) Exempt | 1 Liter Latex-Free Breathing Bags; Reservoir Bag; Vendor Part Number: 720059B, Medline Item Number: 40056; Available in various Anesthesia Circuit Kits and Complete Delivery System kits including: 1) Item Number CDS980230I, NICU-OR CDS; 2) Item Number DYNJ22691, Basin Supply Pack; 3) Item Number DYNJ24724, Spine-Crani Anesthesia Kit; 4) Item Number DYNJAAHY11, Hyperinflation System-UHS-1L; 5) Item Number DYNJAP0010, Anesthesia Circuit Ped-LF; 6) Item Number DYNJAP0011A, Pediatric Circuit-LF; 7) Item Number DYNJAP0012, Pediatric Anesthesia 40 In. CI; 8) Item Number DYNJAP0100, Pedi Breathing Circuit-LF; 9) Item Number DYNJAP0101, Anesthesia Circuit-Ped; 10) Item Number DYNJAP4012, Anesthesia Circuit Ped; 11) Item Number DYNJAP4013, Anesthesia Circuit; 12) Item Number DYNJAP4702A, Anesth Circuit Ped GS Line; 13) Item Number DYNJAP4707, Anesth Circuit 40in Ped-LF; 14) Item Number DYNJAP4722, Anes Cir 40 Ped GSEPW HYBABYH; 15) Item Number DYNJAP4729A, Circ Ped Anes Super-LF; 16) Item Number DYNJAP4730A, Circuit Pediatric Anesthesia; 17) Item Number DYNJAP4732, Anesth Circ Ped 40in 2B/V; 18) Item Number DYNJAP4740, Anesth Circ 40in Ped; 19) Item Number DYNJAP6101, Anesthesia Circuit; 20) Item Number DYNJAP6105, Anesth Circuit 60in Ped; 21) Item Number DYNJAP6106, Anes Circ Ped 60in BI GSE-LF; 22) Item Number DYNJAP6107, Anesth Circ Ped 60in ANG HME; 23) Item Number DYNJAP6108, Anesth Circuit 60in Ped; 24) Item Number DYNJAP6109, Anesth Circ Ped 60in B/V 1L-LF; 25) Item Number DYNJAP6111, Anes Cir Ped 60 GSEPW PALL FIL; 26) Item Number DYNJAP6112, Anesthesia-Peds Circuit; 27) Item Number DYNJAP6113, Circuit Pedi 60in; 28) Item Number DYNJAP6114, Anesth Circ 60in Ped; 29) Item Number DYNJAP6116, Anesth Circ 60in Ped; 30) Item Number DYNJAP6120A, Pediatric Circuit 40in LF; 31) Item Number DYNJAP6124, Anesth Circ Pedi 60in 1L GSL; 32) Item Number DYNJAP6127, Anes Circ Pedi 60in-LF; 33) Item Number DYNJAP6130, Anesth Circ Ped 60in 1L-LF; 34) Item Number DYNJAP6132A, Circuit Pediatric Anesthesia; 35) Item Number DYNJAP6133, Anesth Circuit Peds 60in-LF; 36) Item Number DYNJAP6136, Circuit Peds Anesthesia; 37) Item Number DYNJAP6137, Circuit Pediatric Anesthesia; 38) Item Number DYNJAP7217A, Anesth Circuit 72in Exp Ped BI; 39) Item Number DYNJAP7225A, Circuit Pediatric Anesthesia; 40) Item Number DYNJAP7227, Anes. 90in Exp Pedi -LF; 41) Item Number DYNJAP7233, Anes. Circuit 90in Exp Pedi-LF; 42) Item Number DYNJAP7234A, Anesthesia 72in Pediatric Exp; 43) Item Number DYNJAP7235A, Anesth Circ 72in Ped; 44) Item Number DYNJAP7236A, Anesthesia Pedi 72in; 45) Item Number DYNJAP7237, Circuit Pedi 90in EXP; 46) Item Number DYNJAP7241A, Circuit Pediatric Anesthesia; 47) Item Number DYNJAP7243B, Circ Anesth Ped 2 B/V LFree; 48) Item Number DYNJAP7246C, Circuit Pediatric Anesthesia; 49) Item Number DYNJAP7247, Anes Circ Peds 90in Exp; 50) Item Number DYNJAP7248, Anesth Circ Ped 72in Exp; 51) Item Number DYNJAP7261, Circuit Pediatric Anesthesia; 52) Item Number DYNJAP7262, Circuit Pediatric Anesthesia; 53) Item Number DYNJAP7263, Pediatric Anesthesia Circuits; 54) Item Number DYNJAP9005, Anesthesia Circuit; 55) Item Number DYNJAP9721A, Anesth Circuit Pedi HMEF LF; 56) Item Number DYNJAP9725, Pedi Expand Circuit; 57) Item Number DYNJAP9726, Anesth Circ Ped 90in; 58) Item Number DYNJAP9729, Anesth Circ 90in Exp Ped; 59) Item Number DYNJAP9730A, Circuit Pediatric 60in CORR-LF; 60) Item Number DYNJAP9732A, Circuit Pediatric Anesthesia; 61) Item Number DYNJAP9735, Peds Circuit; 62) Item Number DYNJAP9739, Circuit Pediatric Anesthesia; 63) Item Number DYNJAP9740, Circuit Pediatric Anesthesia; 64) Item Number DYNJAP9743, Circuit Pediatric Anesthesia; 65) Item Number DYNJAPF6003, Anesth Circ 60in Ped Unilimb; 66) Item Number DYNJAPF6005A, Anesth Circ 60in Uni Peds; 67) Item Number DYNJAPF6006, Anes Circ Peds 60in Uni; 68) Item Number DYNJAPF6009, Peds Anesthesia Circuit-LF; 69) Item Number DYNJAPF6011, Circuit Pediatric Anesthesia; 70) Item Number DYNJAPF601 |  |  | December 07, 2007 | 2007 | Medline Industries Inc | Separation from Bushing: The breathing bag may separate from the bushing during use. | Device Design | Consignees were notified by a Sub-Recall Notice letter sent on 12/5/07. The letter instructed users to check their inventory for any affected product and to return a response card indicating the quantity of affected product. The users were instructed to return the affected products to the recalling firm so that they can be reworked. Users who did not wish to return the affected products could request a replacement product and were instructed to destroy the affected products. The letter also informed users that affected products in the manufacturer's inventory were being replaced and these replaced products would contain the label "INSPECTED." For additional information, contact 1-866-359-1704. |  | 77750 | 77,750 bags | Nationwide Distribution. | Terminated | October 16, 2008 | 314 | 571500 | Not_Computer
Z-1220-2008 | 46281 | [6 empty] | VECTRA GENISYS, Model No. 2761, Chattanooga Group, Hixson, TN 37343, Made in the USA |  |  | December 17, 2007 | 2007 | Chattanooga Corp Inc | Device malfunction: causing electric shock and burn. | Device Design | Beginning on 12/17/2007, Chattanooga Corp. sent two recall notification letters, via UPS mail, to its consignees informing them of the situation. The letters stated that the recall was to the user level and requested that the user perform the software upgrade, which will eliminate the possibility of shock and burn. |  | 4047 | 4047 units | Worldwide Distribution - including USA | Terminated | February 19, 2010 | 795 | 22000 | Computer
Z-1236-2008 | 45946 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Solar HA Humeral Stem with Plasma Coating 7mm to 17 mm Product Catalog Number: 5351-4307 to 5351-4317. For use as a Shoulder Replacement. Stryker Orthopaedics Made in USA |  |  | November 15, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Delamination: Loosening of the titanium plasma coating on the product was observed on twelve non-distributed units. | Process control | On 11/28/2007, Stryker Orthopaedics sent, via FedEx, Urgent Product Recall notification letters to distributors, hospitals, surgeons and hospitals' Chiefs of Orthopaedics informing them of the firm's voluntary recall of the affected product. They were instructed to discontinue use of the product and to have the units returned to the firm. Distributors were further instructed to notify their customers using the sub-recall letter provided by the firm. The letter to surgeons also informed them to monitor their patients who had received an implant of the product. |  | 915 | 915 units | Nationwide - USA | Terminated | January 25, 2010 | 802 | 915 | Not_Computer
Z-1188-2008 | 46397 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Reusable Ohmeda Compatible Finger Clip Sensor, Model Number: E403-02, Product Code: PS2225. To monitor oxygen saturation level of patients under doctor's care. Kentec Medical Inc., Irvine, CA 92614 |  |  | May 18, 2007 | 2007 | Kentec Medical Inc | Incompatible components: The Finger Clip Sensors - with a Serial Number prefix of PA514, PA515, YE508, ZC502 and ZG512 - are missing an additional diode and higher resistor in its existing configuration making them inoperable when used to interface with those Ohmeda Tuffsat monitors previously re-configured by Ohmeda to include a detection circuit for diodes. | Process change control | Consignees were notified via telephone on 5/18/07 and an "Important Voluntary Market Withdrawal Notice" letter on 5/23/07. The letter instructed consignees to segregate and return all affected devices. The letter included an "Acknowledgement of Market Withdrawal Notification" form for customer to complete and return to the recalling firm. For additional information, contact 1-800-825-5996. |  | 41 | 41 | Nationwide Distribution including states of AZ, CA, CO, FL, KY, LA, MA, NC, NJ, TX and UT. | Terminated | February 29, 2012 | 1748 | 41 | Hardware | Device Operation | Return | Remove or Replace
Z-1241-2008 | 46381 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa HeartLab Cardiovascular DICOMstore Version 2.04.40.00 software, Heartlab, Inc., Westerly, Rhode Island 02891. The software is used in the Agfa HeartLab Cardiovascular System. |  |  | June 21, 2007 | 2007 | AGFA Corp. | MIsidentification: The electrocardiogram (ECG) data of one patient is misidentified as the data of another patient. | Software design | On 6/21-26/2007, AGFA Healthcare notified, via telephone, the affected consignees of the situation. On 7/2/2007, affected consignees were sent, via registered mail, an "Urgent Medical Device Correction" letter instructing them to: a) remove the affected software DICOMstore 2.04.40.00 and replace it with the prior software version, DICOMstore 2.04.34.02; and b) review patient ECG reports to determine if any patients were actually affected and to notify the firm. AGFA Healthcare also provided a customer service contact number (877) 777-2432. |  | 43 | 43 units | Worldwide Distributution - USA, including states of AZ, CA, DC, FL, IL, IN, KY, MA, ME, MN, MS, OH, OK, OR, NH, PA, SC, TN, TX, WI, and country of Canada. | Terminated | August 04, 2008 | 410 | 43 | Software | Output/Calculation | Software update | Software Update
Z-1242-2008 | 46382 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa Heartlab Cardiovascular Results Management Product: Versions 2.03 and 2.04. Heartlab, Inc., Westerly, Rhode Island 02891 |  |  | July 27, 2007 | 2007 | AGFA Corp. | Incorrect Results: Erroneous echocardiographic measurement values due to mathematical formula being mis-configured. | Software design | On 7/27/2007 to 8/16/2007, AGFA Healthcare notified, via telephone, consignees of the situation. On 7/27/2007, a formal customer notification letter was sent to affected consignees, providing them with a respective listing of patients that were potentially affected. AGFA informed the consignees that the firm would correct any incorrect patient ECG data. AGFA also informed their consignees that they would provide a configuration update to correct the problem. A notification confirmation form to be return to the firm was also attached to the notification letter, which also provided the firm's customer service number: (877) 777-2432. |  | 6 | 6 units | Nationwide Distribution -- USA including states of GA, NH, NY and TX | Terminated | October 10, 2008 | 441 | 6 | Other | Output/Calculation | Upgrade | Repair
Z-1167-2008 | 46408 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll E Series Defibrillator with 12 Lead ECG, Automated external defibrillator, Software Version 3.00 and higher; Zoll Medical Corporation, Chelmsford, MA 01824. |  |  | January 15, 2008 | 2008 | ZOLL Medical Corporation, World Wide Headquarters | Incorrect Patient Records: Patient records that have been stored, then later transmitted or printed may not be the correct record. Records that are transmitted or printed immediately following acquisition are correct. | Software design | Consignees were notified by an Urgent Device Correction letter sent on 1/14/08. The letter instructed users to complete and return a response form. Upon receipt of the response form, the recalling firm will have sent a software update. The letter also provided instructions for users to follow in order to avoid the software error before the software update was received. For additional information, contact 1-800-348-9011. |  | 5018 | 5,018 units | Worldwide Distribution including USA and countries of Canada, AE, Australia, Belgium, Chile, Germany, Spain, France, UK, Hungary, India, Italy, Jordan, Kuwait, Netherlands, Russia, Saudi Arabia, Singapore and ZA. | Completed |  |  | 5018 | Software | Output/Calculation | Software update | Software Update
Z-1606-2008 | 47518 | No | MTF |  |  | Immunology | PMA | Dimension Total Prostate Specific Antigen (TPSA) Flex Reagent Cartridge (RF451). Siemens, Newark, DE 19714 |  |  | March 24, 2008 | 2008 | Siemens Healthcare Diagnostics Inc. | False Results: The product may exhibit falsely elevated results. The affected lots demonstrate sporadic variability in recovery of calibrator and quality control; therefore this issue may not be readily detected upon calibration or during routine processing of Quality Control. | Device Design | The recalling firm issued a letter (Urgent Field Safety Notice) dated 3/25/08 to inform their customers to discontinue use of any remaining inventory, retest using an alternate lot for those samples from post-prostatectomy patients which recovered positive TPSA results less than 1.0 ng/nL. The firm has requested that you call the them for no-charge replacement product at 1-800-441-9250 and/or for technical assistance. |  | 1631 | 1631 cartons | The product was shipped to medical facilities nationwide. The product was also shipped to Germany, Singapore, Canada, Japan and AE | Terminated | October 15, 2008 | 205 | 1631 | Not_Computer
Z-2196-2008 | 48655 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Update Instructions "TH003/08/S" applies to Siemens brand SIMIVIEW 3000, SIMVIEW NT, and MEV ASIM S systems; |  |  | March 10, 2008 | 2008 | Siemens Medical Solutions USA, Inc | The firm sent an Update Instruction, "TH003/08/S" for Motion Enable Switch/holder assembly. The switch holder allows the motion enable switch to be activated if inserted partially; thereby enabling inadvertent motion and serious injury may result. | Software design | On June 25, 2008 all consignees were notified via written letter and on-site visit by the firm's sales representative, informing them that the switch holder allows the motion enable switch to be activated if inserted partially; thereby enabling inadvertent motion and serious injury may result and providing instructions on the recall. A new software version will be provided. For assistance contact Siemens at 1-925-246-8407. |  | 300 | 300 units | Worldwide and Nationwide | Terminated | December 22, 2010 | 1017 | 300 | Software | Physical Safety Hazards | Software update | Software Update
Z-1306-2008 | 46812 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | IMPAXÂż 6.2.1 Network Gateway Server, License Key. Catalog # ABC code: EXEGE000. The product is used to transfer medical imaging data. AGFA Corp., Greenville, South Carolina 2960. |  |  | December 11, 2007 | 2007 | AGFA Corp. | Failure of Data Transmission: No imaging data could be transmitted by the device, because a key value needed for data transmission was incorrectly entered in the affected units by AGFA Healthcare. | Release of Material/Component prior to receiving test results | On 12/11/07, AGFA Healthcare sent out an e-mail to all affected consignees informing them of the situation. The firm also arranged to install the corrected License Key on all affected units of the product. |  | 12 | 12 units | Nationwide Distribution - USA, including states of CA, GA, IA and SD. | Terminated | July 21, 2010 | 953 | 12 | Other |  |  | Other
Z-1248-2008 | 46679 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Edwards LifeStent FlexStar Biliary Stent System; Biliary catheter and accessories; Model Number: EX060803CD; Edwards Lifesciences LLC, Irvine, CA 92614. |  |  | December 05, 2007 | 2007 | Edwards Lifesciences, Llc | Incorrect Expiration Date: Products were labeled with an incorrect shelf-life expiration date. Therefore, the sterility of the device may be compromised. | Packaging process control | Consignees were notified by an Urgent Product Recall letter sent on 12/5/07. The letter instructed users to return any affected product to the recalling firm for replacement products. The letter also advised users that the likelihood of patient complications is extremely low for products that have already been implanted in patients. Users were requested to return a confirmation form after checking the inventory for any affected products. For additional information, contact 1-949-250-3779. |  | 678 | Lot #FR6G0122 - 3 UNITS, FR6G0123 - 2 UNITS, FR6H0193 - 4 UNITS, FR6H0201 - 6 UNITS | Nationwide Distribution | Terminated | January 11, 2012 | 1498 | 10320 | Not_Computer
Z-1307-2008 | 46813 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Impax 4.5, Impax 5.0, Impax 5.2 and Impax 5.2x Picture Archiving Systems (PACs) with integrated Voxar 3D software; and Impax 5.0 and Impax 5.2 PACs with integrated MediCAD software. The softwares are used to provide digital X-ray images for pre-surgical orthopedic planning. AGFA Corp., Greenville, South Carolina 29601 |  |  | February 15, 2007 | 2007 | AGFA Corp. | Misidentification: The orthopedic planning X-ray images for one patient are misidentified as the images for another patient. | Software design | On 2/15/07, AGFA notified their consignees of the situation by sending them "Urgent Medical Device Correction" letters, instructing them on how to work around the problem until the firm was able to install the upgraded software that corrected the problem. On 12/11/2007, the firm reported that the upgraded software had been installed at all affected sites. |  | 879 | 879 units | Worldwide Distribution - USA including states of AR, AZ, CA, CO, DC, FL, GA, IL, IN, KY, KS, MA, MD, ME, MI, NC, NJ, NV, NY, NV, OH, OK, PA, SC, TN, TX, UT, VA and WA, and country of Canada. | Terminated | June 30, 2009 | 866 | 879 | Software | Output/Calculation | Software update | Software Update
Z-1314-2008 | 47210 | No | EXB | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | GentleTouch System-Colostomy/Ileostomy Postoperative Kit; Ostomy pouch and accessories; UPC Code Number: 30003-020938; 2 3/4 in., 70mm; ConvaTec, Skillman, NJ 08558. |  |  | March 07, 2008 | 2008 | ConvaTec | Mislabeled: The carton label reads "For postsurgical use following a urostomy procedure." It should read "For postsurgical use following a colostomy or ileostomy procedure." | Labeling mix-ups | Consignees were notified by an Urgent Product Recall letter sent on 3/7/08. The letter instructed users to return any unused affected product to the recalling firm for a credit. For additional information, contact 1-800-422-8811. |  | 9299 | 9,299 units | Worldwide Distribution including USA and country of Canada. | Terminated | October 21, 2008 | 228 | 9299 | Not_Computer
Z-1779-2008 | 47833 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips defibrillator, model M5066A. Sold in the US by Philips as an OTC under the brand name Philips "HeartStart On Site Defibrillator" and outside the U.S. by Philips with the brand name Philips "HeartStart Defibrillator". The HeartStart is an automated external defibrillator for use on site and consists of a computer with digital interface and cartridge containing pads. |  |  | April 30, 2008 | 2008 | Philips Medical Systems | Buttons Sticking - The On/Off and Shock buttons on the front panel of the device may stick in place when pressed and fail to respond to the button press, causing a delay in the delivery of a defibrillation shock when the system advises a shock. | Component design/selection | Three separate Urgent - Device Recall letter were sent out on April 30, 2008 via tractable letter using UPS. 1) The end-customer letter advises of the recall, contact information to arrange a replacement of the device and instructions for checking the equipment to insure it is working correctly. "If your HeartStart Home Defibrillator is one of the affected devices or if you have any questions about this recall, please contact Philips Healthcare at 1.800.263.3342, option 5, or visit www.HeartStartHome.com to arrange for a free replacement defibrillator. If you have given your HeartStart to someone else as a gift, please call us promptly and let us know how to contact that person to arrange for the replacement. While waiting for your replacement unit, Philips recommends that you perform a battery insertion test (BLT) of your HeartStart, as described in the Owner's Manual provided with it. Simply remove the battery, wait 5 seconds, reinstall the battery, then press the shock button when directed. After the BIT has finished, press the green On/Off button twice. If either button does not seem to work, please contact Philips immediate." 2) The dealer/distributor and retail letters advises of the recall; request that the dealer/distributor provide end-customer information. "As part of a coordinated effort involving notification of regulatory authorities, other Philips entities, and support personnel, Philips will be contacting all affected customers directly. It is not necessary for you or your sub-distributors to contact customers." 3) The international customer contacts will be made via established Philips sales and distribution channels. Philips will be replacing international units under a Field Change Order Process (to be provided by the firm). In addition, Philips will provide a website that will provide customers with information related to the recall (www.HeartStartHome). |  | 3959 | 3,959 units | Worldwide including USA and countries of Austria, Australia, Canada, Colombia, German, Denmark, Dominican Republic, Spain, France, England, Ireland, Italy, Japan, Netherlands, Norway, New Zealand, Panama, Sweden, Singapore, and South Africa. | Terminated | October 23, 2012 | 1637 | 4596 | Not_Computer
Z-2019-2008 | 46287 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with CXP Software Version 2.1 and 2.2 - CXP analysis SW V2.1 single user upgrade, Automated differential cell counter; Part Number 722395; Beckman Coulter; Miami, FL 33196. |  |  | May 04, 2007 | 2007 | Beckman Coulter Inc | Incorrect Display: In certain modes, the protocol will not initially display the correct data in some plots. In addition, when running EWL files, any edited SampleID1 is not automatically updated in all tubes in the panel. | Software design | Consignees were notified of this safety recall by an Urgent Product Corrective Action letter sent on 5/3/07. The letter provided instructions on how to ensure correct data acquisition. Users were requested to return a response form acknowledging receipt of the letter. For additional information contact 1-800-526-7694. |  | 501 | 445 units in the US; 56 units in Canada | Worldwide Distribution including USA and country of Canada. | Terminated | July 11, 2011 | 1529 | 501 | Not_Computer
Z-1235-2008 | 46833 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | AMS In-Fast Ultra Kit with Polypropylene Suture, REF 72403885. Sterile Transvaginal Bladder Neck Support System. |  |  | January 28, 2008 | 2008 | American Medical Systems | Missing part: The units in two lots of In-Fast Ultra Transvaginal Bladder Neck Support System may contain only one bone screw, but should contain two as indicated on the labeling. Two screws are required to complete the procedure. The missing screw will result in the inability to complete the procedure as expected if there are no additional screws readily available. | Process control | All customers received phone calls on 1/28/08 explaining the error and associated risks and were followed up with detailed letters. The letter "AMS Important Product Recall" dated 1/28/08 described the product and problem. Customers were advised to return the enclosed "Recall Acknowledgement Form" to AMS along with any affected product that they have in inventory. |  |  |  | Nationwide Distribution. | Terminated | October 13, 2010 | 989 |  | Not_Computer
Z-1315-2008 | 46806 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Fabius GS Anesthesia Machine; Gas machine for anesthesia or analgesia; Catalog Number: 8604700; Drager Medical Inc, Telford, PA 18969. |  |  | February 05, 2008 | 2008 | Draeger Medical, Inc. | Detachment of Component: The chassis on some machines were manufactured out of specifications. As a result, the caster may break loose from the chassis. | Nonconforming Material/Component | Consignees were notified by an Urgent Medical Device Recall letter sent on 2/7/08. The letter informed users that the recalling firm will contact them to schedule appointments to inspect any potentially affected machines. The letter instructed users to continue using the machines, before the inspection, making sure to carefully follow the warnings in the instruction manual. If a caster on the machine appears unstable before the inspection is conducted, users were advised to contact the recalling firm. For additional information, contact 1-800-543-5047. |  | 196 | 196 UNITS | Nationwide Distribution. | Terminated | September 09, 2008 | 217 | 196 | Not_Computer
Z-1379-2008 | 46947 | Yes | BYE | Positive end expiratory pressure breathing attachment. | Anesthesiology | Anesthesiology | 510(k) | Smiths Medical FLO2 Emergency Non-rebreather High Flow O2 System; Positive end expiratory pressure breathing attachment; Reference Number: P60000; Smiths Medical, Rockland, MA 02370. |  |  | March 04, 2008 | 2008 | Smiths Medical ASD, Inc. | Incorrect Oxygen Level: Device may not provide expected oxygen concentration or the expected flow rate due to incorrectly molded part. Patient will not receive proper amounts of air/oxygen when in need. | Nonconforming Material/Component | Consignees were notified on 2/26/08 via telephone to place the affected products on hold pending further instruction from the recalling firm. A follow-up letter was issued on 3/4/08 advising consignees to return any product in stock for a replacement or credit. Distributors were requested to provide their customer lists to the recalling firm in order to notify their customers of this recall. |  | 650 | 650 | Worldwide Distribution including USA states of CA, FL, IN, KS, MN, MO, NC, NH, NY, PA and TN, and country of Canada. | Terminated | August 20, 2012 | 1630 | 846 | Not_Computer
Z-1381-2008 | 47219 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Scorpio Total Knee Cruciate Retraining Femoral Component; Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis; Product Number: 70-300-7R; Stryker Corp, Howmedica Osteonics Corp, Mahwah, NJ 07430. |  |  | February 08, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | Mislabeled: The labeling indicates a Scorpio CR Femoral Component, however, the device inside the package is actually a Scorpio PS Femoral Component. | Labeling mix-ups | Consignees were notified by letter on 2/8/05. For additional information contact 201-831-5825. |  | 7 | 7 units | Worldwide Distribution including USA and countries of Canada, Mexico, Italy, The Netherlands and the UK. | Terminated | July 16, 2008 | 1254 | 12 | Not_Computer
Z-1233-2008 | 46800 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BHM Medical Infrared (IR) Hand Control System, an optional wireless control component of the Arjo Maxi Sky 600 and BHM V4 Ceiling Patient Lifts; the hand control system is composed of two components, part 700.13640.33 - IR hand control, and part 492.00038 - IR receiver board |  |  | February 20, 2008 | 2008 | Arjo, Inc. | The ceiling lift may not stop lateral movement after releasing the left (or right) action button on a ceiling lift equipped with an infrared hand control. | Device Design | Arjo sent Urgent Device Field Correction Notification letters dated 2/15/08 on the BHM Medical letterhead to the end user accounts who received the affected IR remotes, advising them of the potential for the remote to become difficult to operate or to create uncontrolled movement of the ceiling lift. The accounts were instructed to remove the IR remote from service until the field correction is issued by BHM Medical. Instructions for removal of the IR equipment were provided in the Safety Advisory Notice included with the letter. The accounts were requested to complete and return the enclosed customer response form, acknowledging the receipt and understanding of the letter, and indicating the number of units requiring correction. |  | 25 | 25 units | Nationwide Distribution. | Terminated | September 11, 2010 | 934 | 25 | Not_Computer
Z-1917-2008 | 48084 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Celsus Laboratories Heparin Lithium Lyophilized; Lithium Heparin for use in blood collection tube devices and in vitro diagnostic devices. |  |  | March 31, 2008 | 2008 | Celsus Laboratories Inc | Contamination: The firm learned that up to 2.2% oversulfated chondroitin was found in a batch of crude heparin they received from their supplier. | Nonconforming Material/Component | Celsus notified their customers by e-mail on April 24, 2008, explaining that oversulfated chondroitin was found in a batch of crude heparin they received from their supplier. Celsus recommended the device manufacturers quarantine material from the specified lots. |  | 211818 | HL-31707=11,542g, HL-31807=12,836g, HL-32708=8,374g, HL-32808=12,128g &amp; HL-32908=5,000g | Worldwide Distribution. The Lithium Heparin was sold to companies in USA (Nebraska, New Jersey, Ohio), Puerto Rico, Austria, and Denmark. | Terminated | September 21, 2012 | 1635 | 211818 | Not_Computer
Z-1308-2008 | 46814 | No | LKK |  |  | General Hospital | PMA | Medtronic Intrathecal Catheter Pump Segment Revision Kit; Programmable implanted infusion pump; Model Number: 8596SC; Packaged with catheter model number 8731 or 8731SC. |  |  | January 09, 2008 | 2008 | Medtronic Neuromodulation | Incorrect Connector Pin: The Medtronic 8596SC Intrathecal Catheter Pump Segment Revision kits were packaged with the incorrect connector pin. | Packaging process control | Consignees were notified by an Urgent Medical Device Recall letter sent in January 2008. The letter informed users that the recalling firm would contact them regarding removing the affected products from their inventory. Users were requested to return a response form indicating the number of affected product. For additional information, contact 1-800-707-0933. |  | 18 | 18 | Nationwide Distribution including states of MI, OR, IN, NY, FL, PA, MN, TX, CO and UT. | Terminated | October 08, 2010 | 1003 | 18 | Not_Computer
Z-1479-2008 | 46888 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Sequoia Spinal System Surgical Kits; Pedicle screw spinal system; Non-sterile (kits intended to be sterilized prior to use); Rx Only; |  |  | February 13, 2008 | 2008 | Abbott Spine | Screw Head Dissociation: The polyaxial screw head may dissociate from screw shaft during surgery. Also, the polyaxial screw driver is experiencing difficulty (broken/bent driver shafts, difficulty attaching screws to the driver and jamming of the collet.) | Device Design | The recalling firm notified sales reps regarding the recall by phone on 02/13/08. The sales reps were instructed to discontinue kit use in surgeries and to return inventory. Surgeons who had used the Sequoia System in surgery were notified of the recall via letter on 02/20/08. Surgeons were asked to monitor for post-operative dissociation in implanted patients and to notify Abbott Spine sales reps if any incidences occur. |  | 38 | 38 surgical kits. | Nationwide Distribution including states of AR, AZ, CA, CT, CO, FL, IL, MD, MI, MN, NC, OH, OK, OR, PA, TX, VA and WI. | Terminated | July 28, 2008 | 166 | 38 | Not_Computer
Z-1474-2008 | 47225 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | Medline EndurO2 Oxygen Concentrator without Oxygen Monitor; Portable oxygen generator; Reorder Number: HCS021; Packaged in 1 unit per case; Product is an oxygen generator designed to provide oxygen to patient who has been prescribed oxygen therapy by a physician. |  |  | March 04, 2008 | 2008 | Medline Industries Inc | Overheating of Components: There is the potential for an electrical component in the oxygen concentrator to overheat, causing damage to the unit and making it nonfunctional. | Device Design | Consignees were notified by a letter sent on 3/4/08. The letter informed users that the recalling firm will replace the affected units with new units or issue a credit for the returned recalled units. Distributors were requested to notify their accounts of the recall. The letter requested users to complete and return an enclosed response sheet, indicating the number of units to be replaced or returned for credit. For additional information, contact 866-359-1704. |  |  |  | Nationwide Distribution | Terminated | June 17, 2010 | 835 |  | Not_Computer
Z-1902-2008 | 47409 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Levitronix CentriMag Back-Up Console |  |  | March 17, 2008 | 2008 | Levitronix, Llc | Stops pumping: Interruption of CentriMag System support may occur when using a Valleylab Force FX-C Electrocautery Unit | Device Design | Levitronix issued an Urgent - Voluntary Device Correction letter on 3/17/08 mailed to the US distributor and requested to contact their accounts. The letter revises the warning in the CentriMag Console Manual. Levitronix is issuing revised manuals.On March 13, 2008 the firm emailed and mailed (FedEx) to the one Canadian Hospital and Health Canada was informed of the Voluntary Advisory Notice via email on Wednesday, March 19, 2008.Rest of the World: Wednesday, March 12, 2008 throughMonday, March 17, 2008. The Notice was issued as aVoluntary Advisory Notice after review and follow up with theUK MHRA. On 7/24/08 the firm issued a revised Recall Letter to advise users not to use Valley Lab Force FX-C or SSE2L Electrosurgery device with the firm's CentriMag Blood Pumping System and include the significant risk to health. The firm revised their letter and Black Box Warning and reissued the letters to the hospital, surgeons, etc. on 7/24/08.On 2/5/09 the firm issued a "Follow-Up to Urgent Device Correction of 3/17/08 " to advise users of the software upgrade to be installed in the CentrigMag Primary software Version 3.0 and CentriMag Back Up console , software version 2.0. Site visits will be made by Levitronix between Feb 2009 and April 2009 , software upgrade installed and new Operating Manuals provided. |  | 237 | 237 units | Worldwide Distribution: USA, Canada, Switzerland, UK, Italy, France, Greece, Spain, Slovenia, Czech, Turkey, South Africa, Argentina, Chile, Panama, Singapore, Taiwan, Israel, and Saudi Arabia. | Terminated | July 06, 2010 | 841 | 867 | Not Computer
Z-1372-2008 | 46212 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Vascular Drape, Surgical drape and drape accessories; Reference Number: EF 00-902776-01; Packaged in a box of 10. |  |  | December 20, 2007 | 2007 | OEC Medical Systems, Inc | No Premarket Notification: Product does not have FDA 510(k) approval for use. | PMA | Consignees were notified by an Urgent Recall Notice letter on 12/20/2007. The letter instructed users to stop using and to dispose of the recalled drapes. For additional information, contact 1-800-874-7378. |  | 158 | 158 units | Nationwide Distribution including states of AL, CA, FL, MA, MD, MO, NJ, NY, PA, OR, SC and TX. | Terminated | July 22, 2008 | 215 | 158 | Not_Computer
Z-1928-2008 | 48100 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Bio-Medicus Femoral Cannula/Introducer, for Cardiopulmonary Bypass. Catalog Numbers: CB96345-023, CB96345-023, and CB96345-029. |  |  | May 09, 2008 | 2008 | Medtronic Perfusion Systems | Medtronic is recalling selected Carmeda-coated cardiopulmonary bypass (CPB) products (e.g., oxygenators, reservoirs, pumps, cannulae and tubing packs) that were manufactured with contaminated heparin. Medtronic's heparin supplier recently reported that it had provided some batches to Medtronic that were contaminated with oversulfated chondroitin sulfate (OSCS). | Nonconforming Material/Component | Medtronic issued a Press Release on 05/07/2008. An Urgent Medical Device Recall letter was sent to customers on May 07, 2008. The letter identifies affected devices, explains the situation and requests that the impacted product be returned to Medtronic for disposal. The end user will be requested to complete a Recall Certificate acknowledging the receipt of this information. |  | 46 | 17 US, 29 OUS | Worldwide Distribution: USA, Australia, Belgium, Canada, Czech Republic, France, Italy, Japan, Norway, and Taiwan. | Terminated | May 08, 2012 | 1460 | 844 | Not_Computer
Z-1566-2008 | 47203 | No | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AxSYM Ultrasensitive hTSH II Master Calibrators |  |  | March 06, 2008 | 2008 | Abbott Diagnostic International, Ltd. | High Control value out of range after calibration. --AN INCREASED FREQUENCY OF HIGH CONTROL VALUES BEYOND THE UPPER LIMIT OF THE RANGE SPECIFIED BY THE PACKAGE INSERT. | Nonconforming Material/Component | The firm issued Product Recall letters dated February 22, 2008, requesting remaining units of the affected product be discontinue using and destroyed. A Customer Reply Immediate Action Required letter was enclosed to indicate acknowledging receipt of the product recall letter and to provide information on the number of kits destroyed before March 7, 2008 and faxed to 1-800-777-0051. |  | 882 | 882 total units | Worldwide Distribution --- USA including states of AZ, CA, CO, FL, GA, ID, IL, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NY, OH, OK, PA, RI, SC, SD, TN, TX, WA, WI, WV, and WY, and Puerto Rico, and countries of Canada, Mexico, Brazil, Chile, Germany, Japan, and Hong Kong. | Terminated | January 25, 2012 | 1420 | 882 | Not_Computer
Z-1571-2008 | 47253 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | EUB-5500 VISION 5500 Diagnostic Ultrasound Scanner w/Step 5 (V05-7) Software. Imaging device intended to provide the physician with physiological and clinical information. |  |  | February 25, 2008 | 2008 | Hitachi Medical Systems America Inc | Miscalculation reading. A software error in the firm's EUB-5500/HI VISION diagnostic scanning system causes a miscalculation of the left ICA/CCA ratio when using the Carotid Calculation package for patient scans. | Software design | A Device Correction Letter was issue by the firm on 02/25/2008. The letter asks the affected customers not to use the value of ICA/CCA on the Measurement Report for diagnosis, but instead to follow the Immediate countermeasure detailed in the 'Note in Use' document which accompanies the Device Correction letter. The customers are asked to follow the Immediate Countermeasure until they receive a Temporary Countermeasure to correct the ICA/CCA ratio calculation. The Temporary Countermeasure will include: 1) a Disk to import a new carotid measure package; 2) an Installation booklet detailing how to import the new carotid measurement package and optional customization tools; and 3) an Operator manual addendum Upon release of the countermeasure, the firm's Applications or Servcie personnel shall perform the corrective action as identified in the Service Pack (A) Installation Instructions. Finally, upon release of the next software revision, service personnel will replace the software at each customer location. |  | 20 | 20 units | Nationwide Distribution --- including states of OH, FL, CA, TX, DE, VA, NY, MO, KS, NJ, MD and ID. | Terminated | February 24, 2012 | 1460 | 20 | Software | Output/Calculation | Software update | Software Update
Z-1376-2008 | 46708 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Vanguard DCM CR Tibial Bearing, 14MM x 79/83MM, CR STD bearing, sterile, Biomet Orthopedics, Inc., Warsaw, IN; Part 183464. |  |  | February 12, 2008 | 2008 | Biomet, Inc. | Mislabeled: The packaging of the two products was mixed up, resulting in the products being labeled with an incorrect size. If implanted, the product may not fit properly, which could require repeat surgery. The problem could also result in a delay in the procedure. | Packaging process control | Distributors were notified via recall letter (Urgent Medical Device Removal Notice) dated 2/12/08 to locate the devices in their possession and at hospital sites, discontinue use, and return them to Biomet. Hospitals having received the devices will be provided with a recall letter dated 2/12/08 which describes the nature of the problem and the removal of these instruments. A third letter dated 2/12/08 was provided for implanting physicians. |  | 8 | 8 | Worldwide: USA and Australia | Terminated | December 22, 2009 | 679 | 12 | Not_Computer
Z-1587-2008 | 47400 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Triathlon TS Femoral Trial; |  |  | March 05, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Sharp edges will cause a glove to tear and cut someone.Sharp edges on the slots machines in select lots of the Triathlon TS Femoral Trials, sizes 1-8, have been discovered. | Process control | Stryker Branches/Agencies were notifed by letter (Urgent Product Recall) on March 5, 2008. Reps were asked to examine inventory and visit hospitals that were provided with the device to retrieve product. Reps were instructed to retrieve all of the affected product lots and return them to their branch or agency warehouse for reconciliation. All products should be reconcilied on the Product Accountability Form that is attached to the letter and sent via Inter-Org Transfer back to Stryker Orthopaedics. Questions and re-orders of the product that is being returned should be directed to 201-831-5825. |  | 2 | 2 | Nationwide Distribution --- including states of CT, NJ, PA, VA, NC, GA, FL, AL, MI, MN, IL, TX and AZ. | Terminated | October 25, 2010 | 964 | 104 | Not_Computer
Z-1373-2008 | 46751 | Yes | MRM |  |  | Cardiovascular | PMA | Lumax Cardiac Resyncronization Therapy Defibrillator (CRT-D); Implantable, dual-chamber defibrillator; Model Number: HF-T; Catalog Number 355263; Software Version 603.U; Catalog Model Number 359251; ICS 3000 units are composed of three Catalog/Model Numbers: 349527, 349530, and 355547. The model 349530 is the component which is loaded/updated with software. |  |  | January 10, 2007 | 2007 | Biotronik Inc | Timing Feature Error: An implementation error allows the usage of the V-V timing feature, which is not approved for use in the U.S. | Software design | Consignees were notified by a letter sent 1/10/07. The letter advised users to stop usage of the affected software. Users were instructed not to implant additional devices until the software program has been corrected. The recalling firm requested that all affected software be scrapped or returned to them. Users were requested update the software with a corrected version as soon as it becomes available. |  | 53 | 53 of the 603.U discs distributed. | Nationwide Distribution including states of AL, AZ, CA, CO, DE, FL, GA, IL, LA, MI, MO, NC, NJ, NV, NY, OH, OR, PA, SC, TX and WA. | Terminated | May 05, 2011 | 1576 | 53 | Software |  | Software update | Software Update
Z-2005-2008 | 48101 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Trillium AFFINITY NT, 511T, Hollow Fiber Oxygenator with Trillium Biopassive Surface, with Plasma Resistant Fiber. MODEL Numbers: 95215, 95216, 95217, 95249, 95250, 95252, 0H46R28, 0M81R41, 0R35R13, 0Y95R22, 1702R51, 1A11R16, 1A19R26, 1B22R5, 1F66R12, 1G47R4, 1K86R3, 1K93R5, 1L71R20, 1M05R12, 1P07R6, 1X24R, 2892R18, 2D64R12, 2E55R9, 2N90R8, 2V85R4, 3B89R12, 3C98R7, 3P11R4, 3R39R2, 3R42R14, 4B46R16, 4E58R16, 4F13R2, 4J01R7, 4Q39R1, 4R68R5, 4R68R6, 4W65R, 4X44R7, 5633R30, 5C89R2, 5D83R, 5R69R1, 5U35R, 9059R20, 9243R19, CB0A12R15, CB2T58R4, CB3L15R5, CB4D99R4, CB4E05R6, CB4J24R2, CB4J67R1, CB4K47R1, CB4K64R5, CB5B53R, CB5C01R2, CB5C82R9, CB5P29R4, CB5Q88R, CB5S97R2, DP38T, HY0W55R17, HY0Y86R17, HY1709R34, HY1C73R16, HY2F13R22, HY2F13R23, HY2J04R2, HY2K12R5, HY2K16R5, HY2K51R14, HY2K60R3, HY2Q31R1, HY2Q32R2, HY2T49R3, HY2U43R5, HY2U98R5, HY2W85R9, HY3B29R13, HY3C53R1, HY3E77R13, HY3H52R13, HY3L42R12, HY3L78R1, HY3M41R4, HY3M43R1, HY3M45R1, HY3Q11R14, HY3T77R2, HY3U59R4, HY4A96R3, HY4B44R, HY4B50R, HY4E05R8, HY4E66R5, HY4V82R2, HY4X31R7, HY4Y15R11, HY4Z59R7, HY4Z62R5, HY5C74R15, HY5D07R, HY5D70R1, HY5L38R1, HY5N41R2, HY5N42R3, HY5P52R, HY5Q00R, HY5S13R, HY5S33R2, JP1Z53R11, M310113Q, M310319A, M310523D, M310525C, M310531G, M310538A, M310543F, M310824V, M310852C, M310854A, M311028K, M311224A, M311818C, M320234E, M320410V, M320608K, M321014K, M321505L, M322318C, M330112G, M330344D, M330607H, M332108C, M332607G, M333405W, M333809B, M334007C, M334007C, M334008D, M334209A, M335615B, M336202V, M337412A, M337414A, M337415A, M337417A, M337418B, M338001P, M340305C, M340615E, M343105G, M380324D, M390206E, M390322C, M390508D, M390824C, M391817L, M391817M, M391908B, M392814A, M392815A, M395216B, M395217B, M395319B, M396010D, M397013G, M397912B, M398309C, M399510A, M399511A, M410811K, M430737A, M443307B, M443822B, M443823B, M447005R, M460524D, M460530B, M460531B, M460532B, M462007B, M470233C, M470237C, M470945S, M470945T, M490750B, M495512K, M580211B, M580416B, M580418E, M720407F, M720409F, M720714C, M720718D, M720720B, M930401B, M930402B, M960116A, M960119A, M973502E, M973901C, M999308D, P391401B, TL0A42R21, TL0W56R7, TL1E95R12, TL1J50R3, TL2C82R4, TL2H89R7, TL2J77R, TL2K45R6, TL2L66R5 TL2L67R, TL2M25R6, TL2P63R1, TL2P93R5, TL2P94R9, TL2R00R3, TL2R66R, TL2S03R4, TL2S21R1, TL2S24R3, TL2S80R24, TL2S80R30, TL2S80R30, TL2T46R7, TL2U42R13, TL2U59R, TL2W21R, TL2W22R5, TL2W90R11, TL2Z00R4, TL2Z34R, TL2Z35R, TL3A08R16, TL3B47R8, TL3C66R4, TL3C69R16, TL3E97R5, TL3E98R7, TL3E99R2, TL3F17R2, TL3F33R1, TL3G03R2, TL3G89R1, TL3H76R12, TL3H85R1, TL3H86R6, |  |  | May 07, 2008 | 2008 | Medtronic International Trade Inc | Medtronic is advising that certain Trillium-coated Medtronic cardiopulmonary bypass (CBP) products (e.g., oxygenators, reservoirs, pumps and tubing packs) were manufactured with heparin batches contaminated with oversulfated chondroitin sulfate (OSCS). Based upon currently available information, Medtronic has determined that Medtronic Trillium-coated products can continue to be used in cardiopulm | Nonconforming Material/Component | Medtronic Cardiovascular Customers were sent the May 05, 2008 Urgent Medical Device Corrective Action letter to advise them that certain Trillium-coated cardiopulmonary bypass products were manufactured with heparin which may be contaminated with oversulfated chondroitin sulfate. The letter discussed the risk evaluation and gave recommendations to healthcare professionals for patient care. Affected devices can continue to be used in cardiopulmonary bypass procedures. |  | 59500 | 59500 | Worldwide Distribution, including USA. | Terminated | December 01, 2008 | 208 | 59500 | Not_Computer
Z-2037-2008 | 48755 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | GE Healthcare Centricity Perinatal (formerly QS) System - PFILS (Patient File Server) software; automatic patient data management providing clinical information at the bedside in Labor &amp; Delivery, Mother-Baby and the Neonatal Intensive Care Unit. |  |  | June 30, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Patient information mix-up: When the PFILS application, configured with a backup server has experienced a network interruption, PFILS verify could result in the Centricity Perinatal application recording non-identifiable patient information to the incorrect patient file. | Software design | GE Healthcare notified their Centricity Perinatal customers via letter dated 7/20/07 of the problem with the PFILS application that could result in the Centricity Perinatal application recording non-identifiable patient information to the incorrect patient file. The letter provided the users with a workaround and provided them with a Frequently Asked Questions document which listed the availability of a software patch to fix the problem. A follow-up letter dated 6/27/08 was sent on 6/30-7/2/08 to those accounts who had not requested a PFILS Hotfix CD-ROM or had not had their software recently upgraded. The letter a copy of the 7/20/07 letter, the most recent FAQ, the PFILS Hotfix CD (PFIL67_66) and Installation Procedure for the PFILS Patch instructions for installation of the software patch. The accounts were advised that they can request an analysis of existing patient records to determine whether they have been affected by the PFILS issue, by submitting a non-critical Technical Support Request on the GE Centricity Perinatal Web site: http://www.gehealthcare.com/usen/security/cis/stupgrades.html, indicating "PFILS ANALYSIS REQUEST" in the Description of Issue field. Any questions were directed to Technical Support at 800-433-2009. |  | 1031 | 1,031 units | Worldwide: USA, Puerto Rico, Canada, and China. | Terminated | September 06, 2010 | 798 | 1031 | Software | Output/Calculation | Software update | Software Update
Z-1557-2008 | 47456 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Medi-Lifter III Plus Total Lift, Model 7000-2000 and Summit Total Lift, Model 4800-2000; patient lifts with mast scale (600 lbs. capacity); |  |  | April 10, 2008 | 2008 | Arjo, Inc. | Scale could break and fall potentially causing injuries to someone who is on it. There is a potential for the mast scale load cell ssembly to fracture and allow the mast/boom to fall. Should a patient be in the lift at the time, the patient could fall and sustain injuries. | Device Design | Arjo sent Urgent Device Recall Customer Notification letters dated 4/1/08 on the BHM Medical letterhead to the end user accounts who received the affected patient lifts with mast scales, advising them of the potential of fracture at the load cell assembly which might cause the scale to fall. Should a patient be in the lift at the time, the patient could fall, potentially sustaining injuries. The accounts were instructed to immediately stop using the patient lifts. A BHM representative will contact the accounts within 5 days to arrange for the correction of the lifts. The accounts will be given the option of returning the lifts for repair (with shipping costs covered by BHM), or purchasing a replacement lift at a significantly reduced price. Any questions or concerns were directed to the BHM Service Department at 1-819-868-0441 or service@bhm-medical.com. |  | 37 | 37 lifts | Nationwide Distribution --- including states of California, Florida, Illinois, Michigan, Minnesota, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Texas, Washington, and Wyoming. | Terminated | November 18, 2009 | 587 | 37 | Not_Computer
Z-1549-2008 | 47474 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medex MX821L 72 inch (182.9cm) Administration Sets, Standard Drip, Sterile Single Use, Latex Free device packaged 25 units per case. |  |  | March 24, 2008 | 2008 | Smiths Medical Asd Inc | Misbranded/mispackaged device was distributed. Product labeled to contain 72 inch Standard Drip ( IV) Administration Sets, actually contained 60 inch (IV) administration sets with a' Y' site. | Packaging process control | On 3/24/2008, the firm issued separate URGENT: PRODUCT RECALL NOTIFICATIONs sent via certified mail to their distributor and medical customers. Both recall notifications inform the customers of this labeling error and ask that the customers: 1) check their inventories for the presence of the suspect product; 2) complete and return the attached 'Product Recall Notification Response Form'; and 3) make arrangements to return any unused affected product to the recalling firm in exchange for a credit or replacement product. In addition, the distributors are asked to either provide the recalling firm with a list of their sub-account customers so that Smith Medical can contact the enduser customers and notify them of the recall, or as an alternative measure, the distributor can notify their branches and customers directly, via their own certified letter. |  | 350 | 350 units | Worldwide Distribution --- USA including states of PA, NJ, IL, MN and ND and country of Canada. | Terminated | September 30, 2010 | 920 | 350 | Not_Computer
Z-1460-2008 | 46919 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Acuity with Conebeam Computed Tomography; Radiation therapy simulation system; Model Number: H77. |  |  | January 03, 2008 | 2008 | Varian Medical Systems Oncology Systems | Incorrect Image Orientation: The image orientation tag may not be set correctly if the patient has been scanned with an orientation other than Head First Spine. When such images are imported to a treatment planning system, they may appear with either a mirrored or a rotated view. | Software design | Consignees were notified by an Urgent Medical Device Correction letter issued on 02/05/2008. The letter provided recommended actions for users to follow in order to avoid the software error. The letter informed users that a service pack is being created to correct the issue and that the recalling firm will notify users when it becomes available. For additional information, contact 888-827-4265. |  | 16 | 16 UNITS | Worldwide Distribution--USA including states of CA, MA and PA, and countries of UK, Chile, Belgium, Canada and Germany. | Terminated | September 30, 2009 | 636 | 16 | Software | Display/Image | Software update | Software Update
Z-1550-2008 | 47334 | No | FZW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Dermacarriers II skin graft carriers, 1.5 to 1 dermacarrier, Rx, latex free, sterile, Zimmer, Dover, Ohio; REF 00-2195-012-00. |  |  | April 03, 2008 | 2008 | Zimmer Inc. | Sterility compromised. Lack of assurance of sterility, as the packages may not have been sealed on the manufacturer's end. | Packaging process control | Consignees were notified via letter (Urgent: Medical Device Correction) dated 4/3/08 to identify their stocks on hand, verify package sealing, to discard any non-sealed packages, mail the Business Reply Card regardless if there are any remaining affected products. |  | 99345 | 99345 of all units | Worldwide Distribution --- including USA and countries of Algeria, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Dubai, Ecuador, England, Finland, France, Germany, Ghana, Greece, Guatemala, Hong Kong, Hungary, Iceland, India, Iraq, Israel, Italy, Japan, Jordan, Kenya, Largos, Lebanon, Lithuania, Malta, Mexico, Morocco, Netherlands, Nigeria, Panama, Peru, Poland, Saudi Arabia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunis, Turkey and Venezuela. | Terminated | December 17, 2009 | 623 | 99345 | Not_Computer
Z-1569-2008 | 47236 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Architect Free T4 MasterCheck; a single (specified) analyte control (assayed and unassayed); list 6C50-05. |  |  | February 27, 2008 | 2008 | Abbott Laboratories | Testing material becomes unstable at the recommeded storage temperature. Some of the Architect Free T4 MasterCheck and Total T4 MasterCheck lots are not stable at the recommended storage condition of 2 to 8 degrees Celsius through expiration dating. These lots may exceed the suggested tolerance limits listed in the Reportable Range/Calibration Verification section in Appendix B of the Architect | Nonconforming Material/Component | Abbott sent a Product Recall letter dated 2/27/08, to all customers who received the affected product lots, informing them that the storage conditions are being changed to -10 degrees C or colder for all lots of material manufactured after lot numbers 82130M and 84110M until the cause has been determined and a corrective action has been implemented. All new shipments will contain additional instructions indicating that the ARCHITECT Free T4 and Total T4 MasterChecks are to be stored frozen until first use. Customers were requested to check their inventory for lot numbers 82130M and 84110M, discontinue use of the lots and destroy any remaining inventory. For lots 82140M, 82150M and 84120M, if they know the date the lot was received or the date it was thawed, they may continue to use lots if they have been stored at 2 to 8 degrees C unopened for less than 60 days. If the material has been stored at 2 to 8 degrees C for greater than 60 days or they do not know the number of days it has been thawed, they were instructed to discard the material. They were requested to forward a copy of the letter to any laboratory to whom they may have forwarded the kits. Any questions were directed to Abbott Customer Support at 1-877-4ABBOTT in the U.S. Customers outside of the U.S., were requested to contact their local customer support representative. |  | 197 | 197 units | Worldwide Distribution --- USA including state of Puerto Rico and country of Canada. | Terminated | November 05, 2008 | 252 | 325 | Not_Computer
Z-1556-2008 | 47347 | Yes | NIO |  |  | Cardiovascular | PMA | CORDIS S.M.A.R.T.Âż Control" Nitinol Stent System |  |  | March 10, 2008 | 2008 | Cordis Corporation | Data transcription error. The 10 affected lots of Catalog Number C10040SL were manufactured with a prior version of the guidewire lumen material, which is not validated with the current manufacturing process. This condition may result in a separation of the device. If a separation was to occur, and the separated fragment of lumen material was not retrieved on the guidewire, blockage of the vessel | Mixed-up of materials/components | A letter (RE: Medical Device Recall) dated March 11, 2008, and Acknowledgement Form was sent overnight to the following 6 individuals in each account: Dir. Cardiac Cath Lab, Dir. Radiology Services, Dir. Vascular Services, Risk Manager, Materials Dir., Hospital Administrator; requesting that the affected product be set aside, and that the form be reviewed, signed, and faxed back to Cordis. Affected product is to be either returned or the local sales rep should the contacted to facilitate the return of the affected product. All product will be replaced. Customer Service can be contacted if there are additional questions relative to product return, replacement, or availability. |  | 398 | 398 units | Worldwide Distribution --- including USA and countries of Japan and Canada. | Terminated | October 01, 2009 | 570 | 398 | Computer
Z-2138-2008 | 48974 | No | MVR | Implantable clip. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | GEM 2753 Microvascular Anastomotic Coupler, 2.5 mm, REF GEM2753/I, P/N 5111-00250-010, Rx only, Irradiation Sterilized |  |  | May 14, 2008 | 2008 | Synovis Micro Companies Alliance Inc | Coupler rings may slip out of the delivery toolcaused by jaws larger than specification. | Device Design | Urgent Recall Notification was sent via Fed Ex on May 15, 2008 to 432 hospitals. The letters were addressed to the attention of the Risk Manager. The recall notification instructed the hospital to immediately discontinue use, examine their inventory and, if found, to remove the product and quarantine it for return. A second notification was sent on July 21, 2008 to non-responding consignees. |  | 3812 | 3,812 units | Worldwide Distribution --- including USA and countries of Canada, Japan, UAE, Korea, South Africa, Australia, Taiwan, Norway, Netherlands, Sweden, Greece, Norway, Turkey, Portugal, Taiwan, Denmark, Italy, Germany, Japan, Spain, Austria, and Great Britain. | Terminated | April 09, 2010 | 695 | 3812 | Not_Computer
Z-0274-2008 | 47235 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Precision 500D Classical R&F; System with control room PC; Image-intensified fluoroscopic x-ray system; ; Model Number: 5117866-27. |  |  | August 01, 2007 | 2007 | GE Healthcare | Inaccurate cassette size display reading/ Inaccurate mAs reading: 1. When the device is in PBL mode and a larger film cassette is inserted into the table Bucky (replacing a smaller cassette) after a protocol is selected on the console, the collimator will not adjust to the larger cassette size. The collimator stays at the original smaller cassette size and allows the operator to take an exposure. | Software design | Consignees were notified by two "Product Safety Notification" letters on July 10, 2007. Each letter addressed one of the two safety issues with the affected product. The letters provided safety instructions on how avoid the software error and informed users that their representative would contact them to schedule a visit for a software correction installation. |  | 101 | 101 | Worldwide Distribution --USA including states of: AK, AL, AR, CA, CO, CT, FL, GA, IA, ID, IL ,IN, LA, MA, MD, MI, MO,MS, NC, ND, NE, NH, NJ, NM, NV, OH,OK, OR, PA, SC, SD, TN, TX, UT, VA, VT and WA; USA territory of: PR and countries: EGYPT, SAUDI ARABIA, MACEDONIA, FRANCE and GREAT BRITAIN. | Terminated | September 10, 2008 | 406 | 101 | Software | Display/Image | Software update | Software Update
Z-1459-2008 | 46852 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | High-Definition 120 MultiLeaf Collimator, Medical charged-particle radiation therapy system; Software Version 7.0; Product Code: H56; Accessory to Medical Linear Accelerator. |  |  | January 30, 2008 | 2008 | Varian Medical Systems Oncology Systems | Dose Calculation Error: Software anomaly may result in failure of a MLC leaf to reach planned position, potentially resulting in misadministration of dose to a patient. | Software design | Consignees were notified by a Medical Device Notification letter sent on 01/29/2008. The letter advised users to cease use of the device until a corrected version of the software is available. A representative will contact users to upgrade the application once the revised software has been validated and released for distribution. |  | 3 | 3 units subject to correction | Nationwide Distribution including states of NC, FL and CA. | Terminated | August 18, 2010 | 931 | 3 | Software | Output/Calculation | Software update | Software Update
Z-1190-2008 | 46753 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Steris C1160 Universal Flexible Processing Tray, for use with STERIS SYSTEM 1 Processor, used for sterilization of flexible endoscopes. |  |  | February 19, 2008 | 2008 | Steris Corp | Failure to effectively sterilize: Under certain conditions, the design/operation of the C1160 Tray may compromise the ability of the SYSTEM 1 processor to correctly assess the state of the high pressure pump during the Sterile Cycle. The L6 sensor fails to alarm when the high pressure pump is not working properly. | Device Design | Steris sent an Urgent Safety Alert/Recall letter dated 2/8/2008, to all customers. The letter stated that any device sterilized in the C1160 Tray should not be used clinically prior to a completion of a successful diagnostic cycle. Until further notice, you must run a diagnostic cycle after each sterile cycle involving a C1160 tray or alternately you must discontinue use of these trays. Steris service technicians will schedule visits to correct the problem. |  | 28766 | 28,766 units | Worldwide Distribution, including USA, Puerto Rico, Canada, and Japan. | Terminated | February 24, 2012 | 1466 | 28766 | Not_Computer
Z-1191-2008 | 46788 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | INTEGRA NeuroSensor Probe, REF NS-P, component of Cerebral Blood Flow and Intracranial Pressure Monitoring Kit, Integra NeuroSciences |  |  | January 28, 2008 | 2008 | Integra Neuro Sciences | Incorrect pressure reading: Due to a manufacturing error, the ICP reading could be inaccurate to the extent that they exceed label claims. Specifically, the Insertion Error label claim of +/- 3mmHg could be exceeded by up to -3mmHg, that could result in a maximum Insertion Error of +/- 6mmHg. | Process design | Integra NeuroSciences sent a Recall Notification with a Recall Acknowledgement form to all customers with product currently in their possession through an Integra Field Sales Repesentative on January 28, 2008. The notification advises that the firm has determined that due to a manufacturing error, the ICP reading could be inaccurate to the extent that they exceed label claims. Specifically, the Insertion Error label claim of +/- 3mmHg could be exceeded by up to -3mmHg that could result in a maximum Insertion Error of +/- 6mmHg. Customers are asked to examine their inventory immediately to determine if they have any NeuroSensor Model NS-P's lot #100063 on hand and if so to not use the product. Customers are advised to contact their local Integra representative for further assistance. The letter advises that the lot/serial number is located on a white 1/2" x 1/2" label on the large beige optical connector, and can readily be seen through the bottom of the product tray packaging. The letter also advises that this recall pertains to NeuroSensor Model NS-P's, lot # 100063 -all serial numbers. A response form is included.In addition, the Field Sales Representatives have been instructed to physically remove all lot 100063 of NeuroSensor Probes from inventory and provide replacement product. |  | 43 | 43 units | USA Distribution: CA, GA, NM, TX, NE, WA, and NY. | Terminated | April 22, 2010 | 815 | 43 | Not_Computer
Z-1638-2008 | 47428 | No | IMG | Ultrasound and muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | VECTRA GENISYS, electromechanical therapeutic ultrasound and powered muscle stimulation device, Model No. 2761, Chattanooga Group. |  |  | February 28, 2008 | 2008 | Chattanooga Corp Inc | Device may fail to administer therapy during use. | Component design/selection | Chattanooga Group sent an Urgent Medical Device Recall letter dated 02/29/2008 to consignees to advise them of problem and the recall. The first group of letters was sent to the distributors via certified mail. The letter stated that the devices were failing and the recall was to the User level. The firm is recalling the devices for repair and calibration. The firm followed with a user notification, by US Mail not certified beginning on March 11, 2008. |  | 479 | 479 devices | Nationwide Distribution. | Terminated | February 19, 2010 | 722 | 1565 | Not_Computer
Z-1181-2008 | 46846 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira LifeCare PCA Label Utility, Version 1.0 CD. The installation/user guide is entitled "LifeCare PCA Barcode Label Utility Program." This software program is designed to allow entry of each drug concentration into a formulation library on a pharmacy computer for standardized printing of barcode labels for compounded drug product vials. |  |  | January 03, 2008 | 2008 | Hospira Inc. | Incorrect dosage labels created: Customers have experienced independent changes in drug concentration (mg/mL to mcg/mL and vice versa) and expiration dates, and the addition of trailing zeros to the barcodes. | Software design | Hospira sent letters dated 1/3/08 to the immediate recipients of the product via FedEx overnight letter on the same date. The letters informed the accounts of the reports of independent changes in drug concentration while editing a saved formulation using the Hospira LifeCare PCA Label Utility Software CD that was provided to some customers in a Hospira MedNet Drug Library Creation Team Introductory Kit-LifeCare PCA kit. The accounts were requested to remove the utility from any PC in their facility on which it is currently installed, destroy the CD and sign and return the enclosed removal verification form verifying that they have removed the software from their system. The letter listed commercially available software packages that the account could validate and implement in their facility. Any questions concerning the utility software withdrawal were directed to Hospira at 1-800-241-4002, option 4, then option 1. |  | 127 | 127 CDs | Nationwide Distribution and Canada | Terminated | March 09, 2010 | 796 | 127 | Computer
Z-1538-2008 | 46916 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Wedge HS Suture Anchor II with Needles, 6.5 mm, Model number 3910-100-145 |  |  | February 07, 2008 | 2008 | Stryker Endoscopy | Product package may crack and cause loss of sterility. | Packaging | Stryker Endoscopy initiated the recall on 02/07/2008 by notifying its consignees via telephone, as well as sending out Urgent: Device Removal notices. Recipients of affected lots are to inspect their inventory for part/lot numbers and quarantine any affected product. Affected products are to be returned to Stryker Endoscopy and Customer Service should be contacted for replacement alternatives. Also, the "Acknowledgment of recipt" card needs to be signed and returned to Stryker Endoscopy. |  | 53 | 53 units | Product was distributed to 8 consignees in CA, OH, PA, CO, MA, and NJ. | Terminated | December 21, 2010 | 1048 | 927 | Not_Computer
Z-1348-2008 | 46890 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Da Vinci S 8mm Instrument Cannula with outlet for use with the Da Vinci Surgical System; |  |  | January 04, 2008 | 2008 | Intuitive Surgical, Inc. | (1) Incorrect dimension on Luer on smoke evacuation cannulae not allowing for secure attachment function.(2) Incorrect labeling. External labeling lot number differs from lot number etched on the cannula. | Process control | On 1/4/08, the firm initiated the recall and its notification was via Urgent Device Recall letters explaining the reason for the recall, requesting immediate discontinued use, segregate non-conforming product in a secure area, and a Customer Service Rep will make contact in order to arrange for the retrieval of the material. |  | 39 | 39 units | Nationwide | Terminated | March 31, 2011 | 1182 | 39 | Not_Computer
Z-1620-2008 | 47521 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FLAT DRAIN -10MM, CHARIKER-JETER DRESSING KITS (LARGE AND SMALL) -Product numbers 6680217 and 6680221 - Wound Management Product. |  |  | March 31, 2008 | 2008 | Smith And Nephew, Inc. Wound Management Division | Incorrect labeling. Kit labels may incorrectly state that a flat drain is contained as a kit component while some kits may contain a round drain. | Labeling mix-ups | Consignees were notified of recall by letter (Correction and Removal Notification) sent to customers on 3/31/08. They were asked to check their stocks, identify, and return all affected units to Smith &amp; Nephew according to the instructions. Consignees are to contact the Customer Care Center at 1-800-876-1261, if they have additional questions. |  | 4796 | 4796 | Nationwide Distribution. | Terminated | November 18, 2008 | 232 | 4796 | Not_Computer
Z-1239-2008 | 46773 | No | JKS | Carbon monoxide test system. | Toxicology | Toxicology | 510(K) Exempt | International Technidyne Corp. (ITC) AVOXimeter 4000 Co-oximeter, Catalog Numbers: AVOX4000-110 and AVOX4000-220 |  |  | February 13, 2008 | 2008 | International Technidyne Corp. | Incorrect reading: ITC discovered that instruments built or repaired since approximately April 2007 were inadvertently left with the diagnostic mode enabled when shipped to customers. | Process design | ITC sent Urgent - Medical Device Field Recall - ITC No. 08-001 letters dated 2/08/2008 to hospitals and/or medical facilities identified as having affected instruments. The letter instructed users not to use the instrument if the diagnostic mode is enabled. ITC will reset affected devices and return them to the customers. |  | 63 | 61 US, 2 International | Worldwide Distribution: USA, United Arab Emirates, Malaysia, Saudi Arabia, India, Germany, United Kingdom, Greece, Romania, Singapore, ROC, Pakistan, Netherlands, Korea, France, Belgium, Switzerland, Canada, Thailand, Slovenia, Turkey and Australia. | Terminated | July 24, 2009 | 527 | 273 | Not_Computer
Z-1529-2008 | 46370 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with Data Innovations Instrument Manager, Part Numbers 626553 and 6605628 |  |  | August 20, 2007 | 2007 | Beckman Coulter Inc | Labeling provided by two integrated software systems are not clear enough to avoid potential demographic and sample type mismatches under certain conditions. | Software design | A Product Corrective Action (PCA) letter, dated August 16, 2007, informed the customers that the User documentation and Online Help for the FC500 with Data Innovations lnstrument Manager are not clear enough to avoid potential demographic and sample type mismatches under certain conditions. The letter asks the customers to share the information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. They are also to review all sample identification, demographics, data and results before reporting results out. A response form is included. Questions and concerns regarding the notification should be directed to Beckman Coulter Customer Service at 1-800-526-7694. |  | 141 | 141 units | Worldwide Distribution --- including USA and country of Canada. | Terminated | October 15, 2012 | 1883 | 141 | Computer
Z-1530-2008 | 46372 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACSDiva Software v6.0 and 6.1 with the HP XW4400 Computer System |  |  | December 21, 2007 | 2007 | BD Biosciences | Values might not update in statistics views after certain functions are performed in a worksheet. | Software design | The firm issued an Important Product Information notification letter on 12/21/2007. The letter instructs users to reload data by closing and reopening an experiment before reporting results in order to determine if their statistics have been impacted. They are asked to fill out and return a form pertaining to this problem. BD Biosciences will release a version of BD FACSDiva software to correct this problem. The letter also stated that BD Biosciences rep will contact the users when the new software is available. User are asked to contact BD Customer Support Center at 1-877-232-8995 if they have any additional questions. |  | 74 | 74 units | Nationwide Distribution. | Terminated | September 24, 2008 | 278 | 74 | Software | Output/Calculation | Software update | Software Update
Z-1561-2008 | 46750 | No | GTJ | Neisseria spp. direct serological test reagents. | Microbiology | Microbiology | 510(k) | BD DifcoTM Neisseria Meningitidis Antisera Group Y 241113, 1mL, in vials, which were labeled in part, 8019456 6250890 N. MEN. GP Y UNLABELED LYOPH, Becton, Dickinson and Company, Sparks, MD 21152. The product is packaged in 100 vial units |  |  | September 06, 2007 | 2007 | BD Diagnostic Systems / Lee Laboratories | Exhibits cross reactivity with Neisseria meningitides Group W135 | Mixed-up of materials/components | BD Diagnostics sent an Urgent Product Recall letter, dated September 2007, to all accounts, requesting the discontinue of distribution and that any remaining inventory of the recalled product be discarded. BD will issue credit or send replacement. A receipt of notification needs to be completed and sent back whether or not customers have any remaining inventory. Also, a customer list of end-users who were shipped this product needs to be provided so that the firm may inform them of the issue. |  | 200 | 200 units | Worldwide Distribution --- including USA and country of New Zealand. | Terminated | March 24, 2009 | 565 | 1240 | Not_Computer
Z-1368-2008 | 46906 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EvoTech Endoscope Cleaner &amp; Reprocessor, Product |  |  | February 13, 2008 | 2008 | Advanced Sterilization Products | Cleaning cycles are being cancelled. Specific issues are:a) Water heater tank is rusting resulting in the in-line 0.2 micron filter prematurely clogging,b) Interface board malfunction resulting in a premature lid opening, andc) Minor fluid leaks within the system (commonly at the 0.2 micron filter attachment points). | Other | Feburary 13, 2008 ASP issued an Urgent Product Recall Letter. ASP recommends that customers use an alternative reprocessing method until they return the system. However, if they do not have an alternate method for processing endoscopes and choose to continue using the EvoTech ECR System during this time, endoscopes processed through completed cycles in accordance with the User's Guide will be appropriately cleaned and highlevel disinfected. ASP is prepared to tailor an alternative solution to minimize the interruption to the customer's processing needs. |  | 64 | 64 units | Worldwide Distribution --- including USA and countries of Canada and Japan. | Terminated | November 29, 2011 | 1385 | 66 | Not_Computer
Z-1568-2008 | 46950 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | 3M Comply EO Chemical Indicators Strips, Catalog #1251 |  |  | March 19, 2008 | 2008 | 3M Company / Medical Division | A production specification discrepancy may cause the indicators to show an inaccurate result, which could incorrectly lead customers to conclude the sterilization cycle was adequate. | Nonconforming Material/Component | Customers and Distributors were each sent a "Voluntary Product Recall" letter dated 3/17/08 on 3/19/08. The letter states that use of the affected product should immediately be discontinued, stocks should be checked for any remaining inventory, and reprocess any unused packs using the product not from the affected lot or from an alternative supplier. The letter requested consignees to notify 3M whether the facility has affected product or not by sending the "Product Return Form". Replacements will be provided to those with affected products remaining in their stock. |  | 51992 | 12,998 cases ( 4 boxes per case &amp; 480 strips per box)) | Nationwide Distribution --- including states of AL, AR, AZ, CA, CO, CT,DC, DE, FL, GA,HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS,MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Terminated | March 01, 2012 | 1443 | 51992 | Not_Computer
Z-1628-2008 | 47355 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100S Sterilization System, Product Code 10101 |  |  | February 25, 2008 | 2008 | Advanced Sterilization Products | Oil mist filter fails allowing oil mist to be emitted into the vicinity of the sterilizer, which may result in the release of a mist, "haze," or "smoke." | Device Design | The firm sent an Urgent Device Correction letter dated 4/23/08, notifying customers of a situation that may occur with all models of STERRAD Sterilization Systems (100S, 50, 200, NX and 100NX). As a precaution, personnel should leave the room if mist is observed and avoid returning until the mist has cleared. ASP asks that customers circulate this letter to any STERRAD system users within their facility and also to the Chief Administrative Officer, the Director of Biomedical Engineering, Director of Materials Management and the Chief of Nursing. The letter instructs customers on what to do if they experience a filter failure - a) cancel the cycle if the sterilizer is running [and as always, do not use instruments from an incomplete cycle], b) leave the room, c) discontinue use of the STERRAD Sterilizer until the system is repaired. d) Personnel should avoid working in the room until the mist has cleared. The letter advises that ASP service will replace the sterilizer's filter and increase the frequency of filter changes by adjusting the Planned Maintenance (PM) schedule within the software on their STERRAD 100S and STERRAD NX Sterilizer. Per the notice, for each STERRAD Sterilizer model, ASP will do the following: a) STERRAD 200 Sterilizer - beginning April 28, 2008 through October 2008, ASP will install a new filter and gasket assembly at no charge. b) STERRAD NX Sterilizer - beginning April 28, 2008 through October 2008, ASP will install a new filter assembly at no charge and adjust the PM schedule to 650 cycles or 6 months (whichever occurs first). c) STERRAD 100S Sterilizer - beginning April 28, 2008 through October 2008, ASP will adjust the PM schedule to 750 cycles or 6 months (whichever occurs first). d) STERRAD 50 and 100NX Sterilizers - No adjustments to the oil mist filters or PM schedules are required at this time. PMs will be performed as currently scheduled. Information for MedWatch reporting is included in the letter. |  | 9950 | 9,950 | Worldwide Distribution -- including USA and countries of France, Germany, Italy, United Kingdom, Switzerland, Benelux, Uruguay, Venezuela, Argentina, Chile, Colombia, Ecuador, Belgium, Puerto Rico, Hong Kong, Indonesia, Mexico, Philippines, Australia, Austria, China, Greece, Ireland, Israel, Japan, Korea, Malaysia, Singapore, Taiwan, Thailand, Turkey, Peru, Canada, Czech Republic, Middle East, Portugal, S. Africa, Spain, Sweden, Hungary, Poland, Brazil, Russia, Egypt, India, and Slovenia. | Terminated | November 21, 2011 | 1365 | 9950 | Computer
Z-1463-2008 | 46850 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Viasys Healthcare, Pulmonetic Systems LTV 1000 Ventilator |  |  | February 22, 2008 | 2008 | Cardinal Health 203, Inc | The failure of an electronic component could result in failure of the ventilator to breathe for the patient and/or failure of the ventilator to properly alarm to alert the caregiver. | Device Design | An Urgent Medical Device Recall letter dated February 22, 2008 was sent to ventilator consignees/owners and distributors. The letter describes the issue, identifies the affected devices, states that the devices will be corrected by Cardinal Health and provides direction to verify ventilator operation and audible alarm function daily. Service Centers were provided information to identify affected printed circuit board assembly replacement kits' lot numbers and asked to return them. An "Update" was mailed 03/04/2008 to identify additional ventilator serial numbers that also need correction. |  | 4809 | 4809 (2427 US, 2382 OUS) | Worldwide Distribution --- USA including Puerto Rico and countries of Algeria, Argentina, Australia, Austria, Bahamas, Belgium, Brazil, Canada, Chile, China Peoples Republic, Colombia, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Indonesia, Israel, Italy, Japan, Malaysia, Mexico, New Zealand, Norway, Panama, Paraguay, Philippines, Portugal, Saudi Arabia, Serbia &amp; Montenegro, Singapore, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Vietnam, and West Indies. | Terminated | August 16, 2012 | 1637 | 20603 | Not_Computer
Z-1633-2008 | 47361 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | 6F Viking Guiding Catheter, Percutaneous catheter, Shape ALR 1.2, Part Number: 1001934-06, Abbott Vascular, Santa Clara, CA 95054. |  |  | March 17, 2008 | 2008 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Mislabeling: The part number on the product label packaging may not match the product. | Labeling mix-ups | Consignees were notified via an Urgent Device Recall letter dated and hand delivered starting 3/18/08. The letter requests users to remove and quarantine the affected lots, which will be picked up and replaced by the recalling firm. Users were also requested to return the enclosed Efficiency Check Reconciliation Form. Abbott Vascular will exchange all returned devices. For additional information contact 1-800-227-9902. |  | 10 | 10 units | Nationwide Distribution. | Terminated | July 07, 2011 | 1207 | 17 | Not_Computer
Z-0866-2008 | 47371 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche/Hitachi Modular Analytics System, Modular D2400 Module GMMI clinical chemistry analyzer; Catalog number 04998618001. |  |  | March 21, 2008 | 2008 | Roche Diagnostics Corp. | Possible mismatch between patient and result. | Process change control | Consignees were notified via Urgent Medical Device Correction letter dated 3/21/08, which instructed them to not manually order QC test selections while the system is in operation mode and that they can only manually order QC when the system is in Stand-by mode. Consignees are instructed to consult with their physicians or pathologist to determine specific clinical implications for their patients. The firm states that this issue will be resolved in the next version of software. |  | 11 | 11 of all models | Nationwide Distribution --- including states of Florida, Georgia, Indiana, New Jersey and North Carolina. | Terminated | November 18, 2009 | 607 | 11 | Software | Output/Calculation | Software update | Software Update
Z-2141-2008 | 48697 | Yes | NIM |  |  | Cardiovascular | PMA | Boston Scientific NexStent Monorail, NexStent Carotid Stent and Monorail Delivery System. Material Number (UPN) M001553030, Catalog Number 55-303. [OUS Material Number (UPN) M001553000, Catalog Number 55-300] |  |  | June 06, 2008 | 2008 | Boston Scientific Corporation | Tip detachment: Boston Scientific has received complaints involving a detachment of the tip from the NexStent Carotid Stent delivery system. Potential health hazards resulting from this type of failure include increased procedure time, vessel wall injury, stroke and/or emergency surgery to remove the detached tip. Since the potential for the tip to detach occurs during delivery of the sten | Component design/selection | Boston Scientific sent an urgent Voluntary Medical Device Recall letter, dated 06/06/08, to affected customers. The letter described the issue, identified affected product, requested to cease using or distributing the product and provided direction on how to return the recalled devices to Boston Scientific, Quincy, MA. A Reply Verification Tracking Form was asked to be returned to Boston Scientific. Since the issue is with the delivery system, patients already implanted with the stent are not affected. |  | 1946 | 1946 units (1570 U.S., 376 OUS) | Worldwide Distribution: USA, Austria, Belgium, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Hungary, Italy, Kuwait, Luxembourg, Netherlands, Norway, Panama, Poland, Portugal, Romania, Spain, Switzerland, and Ukraine. | Terminated | February 16, 2011 | 985 | 1946 | Not_Computer
Z-2152-2008 | 48435 | No | GIO | Capillary blood collection tube. | Hematology | Hematology | 510(K) Exempt | Vitrex Blood Gas Capillary Tubes, MC0024, REF 189301, 250 pcs, 1,45 x 1,85 x 125 mm app. 200Âżl, Heparinized: 6 IU Na &amp; 9 IU Li/100 Âżl, |  |  | May 08, 2008 | 2008 | Opti Medical Systems Inc | This lot of Heparinized Capillary Tubes contain contaminated lithium heparin. | Nonconforming Material/Component | Opti Medical sent a Recall Notice, dated May 8, 2008 to all consignees. Consignees were notified via email on/about May 5th thru 7th , 2008. They were instructed to discontinue use of the affected lot and to immediately return all remaining product to OPTI Medical for replacement. |  | 31750 | 31750 tubes | Worldwide Distribution: USA, Puerto Rico, Austria, China, Chile, Cyprus, Great Britain, Japan, Netherlands, and Senegal | Terminated | January 22, 2010 | 624 | 31750 | Not_Computer
Z-1635-2008 | 47373 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Integra LifeSciences Corporation Disposable Convenience Kit, PICC Insertion Tray, CUS791 |  |  | March 18, 2008 | 2008 | Integra Pain Management | Hep-Lock Flush vials included in Integra's PICC Insertion Trays contain a heparin component that has been recalled by Baxter for contaminated heparin. | Nonconforming Material/Component | Integra notified the consignee by letter on 03/18/2008. Instructions were to cease distribution/use of the affected kits and arrange for return to Integra. Also, consignee was instructed to notify all customers who received the kits to cease use and return any affected product for further disposition. |  | 500 | 500 kits | CO | Terminated | November 05, 2008 | 232 | 500 | Not_Computer
Z-1833-2008 | 48124 | No | GBS | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | HydraGlide Heparin coated PVC Straight Thoracic Catheters with Tapered Connector Tip, |  |  | May 12, 2008 | 2008 | Atrium Medical Corporation | Heparin-coated Thoracic Catheters were manufactured with heparin allegedly contaminated with oversulfated chondroitin sulfate | Nonconforming Material/Component | Atrium Medical issued a press release on 5/12/2008 and notified consignees by facsimile on 5/12/2008. Customers were instructed to contact the firm for a return goods authorization and to notify the firm upon receipt of the recall notice. |  | 41 | 41 cases (10 individual/case) | Worldwide Distribution: USA, US Virgin Islands, South Africa, Austria, Taiwan, Venezuela, and Saudi Arabia. | Terminated | August 20, 2009 | 465 | 1842 | Not_Computer
Z-1442-2008 | 46920 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems On-Board Imager, Versions 1.0.15 and 1.2.05 Imaging Accessory to Medical Linear Accelerator. The On-Board Imager Device is used for verification of correct patient position in relation to isocenter and verification of the treatment fields in relation to anatomical and/or fiducial landmarks. |  |  | January 03, 2008 | 2008 | Varian Medical Systems Oncology Systems | If used with a third party radiation therapy treatment planning software system, mistreatment may occur because of a misalignment. | Software design | The firm has upgraded all user sites through normal course of service and installation activities. Notification letters were distributed starting January 03, 2008. |  | 291 | 291 units | Worldwide Distribution -- USA including states of GA, MI, NY, MN, TX, TN, CA, MD, KS, AL, KY, CT, WI, LA, NC, NJ, PA, IL, RI, FL, NE, AK, VA, WV, NV, MO, TN, OR, IN, MA, CT, AL, AZ, NH, ID, OH, SC, CO, CT, ND, WA, and UT, and countries of Sweden, Switzerland, Germany, Denmark, Japan, UK, France, Italy, Norway, Taiwan, New Zealand, Canada, Poland, Australia, Colombia, China, Spain, Belgium, Russian Federation, India, Iceland. | Terminated | November 27, 2010 | 1059 | 291 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1445-2008 | 47244 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | QWIX Screw 4.3mm diam; length 26 mm; The stabilization screw is indicated for fixation of bone fractures or for bone reconstruction. |  |  | February 27, 2008 | 2008 | Integra LifeSciences Corp. | Certain QWIX Fixation Screws (Reference numbers 111426S and 111428S) have been etched and labelled with an incorrect length. | Labeling mix-ups | Integra Reps were sent recall notification with a recall acknowledgement on 2/27/2008. The customers who were shipped affected product were also notified on that date. |  | 42 | 42 units | Nationwide Distribution --- including states of NC, IL, CA, NC and TX. | Terminated | August 21, 2008 | 176 | 62 | Not_Computer
Z-1619-2008 | 47495 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | BCI Âż3180 Oximeter - Patient monitoring medical device, Manufactured by Smiths Medical PM, Inc. For continuous patient monitoring of blood oxygen levels, pulse rate and pulse strength measurements. It is only used in hospital or clinical settings or emergency response environments. |  |  | February 27, 2008 | 2008 | Smiths Medical PM, Inc. | The Smith Medical PM, Inc. BCI ÂżÂż3180 Pulse Oximeter tantalum capacitors C38, C72, C74, and C98 were installed backwards during assembly at the board manufacturer. This exposes the components to reverse polarity voltages that may lead to their pre-mature failure. As a result, the identified monitors may exhibit failure modes such as the monitor may shut down and not produce alarms or signals to | Vendor change control | Consignees were contacted by an URGENT: Product Safety Information Voluntary Recall Notification letter dated March 11, 2008 with attached Safety Action Bulletin 08-SAB01 dated February 27, 2008 and email between March 7 and March 13, 2008. The consignees will need to return the device to be evaluated in the Smiths Medical PM, Inc. |  | 900 | 900 | Class II Recall - Worldwide Distribution --- USA including states of GA, AL, IL, TX, MI, VA, IA, TN, PA, NC, HI, NY, FL, AZ, OH, CA, MS, UT and Countries: Argentina, Bangladesh, Iran, Thailand, Mexico, India, Germany, Guatemala, Pakistan, Chile, Colombia, Brazil, Saudi Arabia, Morocco, Kuwait, Trinidad and Tobago, Ecuador, Paraguay, Uruguay, Canada, Poland, Portugal, Peru, Qatar, Cyprus, Greece, Italy, Spain, United Arab Emirates, Israel. | Terminated | December 17, 2011 | 1389 | 900 | Hardware | Device Operation | Return | Remove or Replace
Z-1622-2008 | 47753 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Deltec CozmoÂż Insulin Pump, Model 1700, Smiths Medical International Ltd WD24 4LG UK, Smiths Medical MD, Inc., St Paul, MN 55112. This is a syringe infusion pump intended for Continuous Subcutaneous Insulin Infusion for the control of diabetes. |  |  | March 10, 2008 | 2008 | Smiths Medical MD, Inc. | Smith Medical discovered an issue with a motor component in specific serial numbers of the Deltec Cozmo ÂżInsulin Pump, Model Numbers 1700 and 1800. This issue affects how the motor operates and may cause an over-delivery of insulin which could result in injury to the user. | Component change control | Consignees were mailed an Urgent Medical Device Recall letter dated March 10, 2008. The Urgent letter addressed to "Distributors, Clinicans, Patients and Other Customers" contains Description of the Problem, Results of Investigation, and Advise on Action to be taken by the User. The patient users will be required to return a Recall Confrirmation Form. The USA distributors, USA prescribing physicians and the USA managed Care Accounts (Insurance Companies) will be mailed also a cover letter. |  | 436 | 436 | Class II Recall - Worldwide Distribution --- USA including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME. RO, MN, MO, MS, ME, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, and countries of Greece, Netherland, Sweden, France, Germany, Turkey, China, Great Britain, Israel, Italy, Kuwait, New Zealand, Australia. | Terminated | January 13, 2011 | 1039 | 1502 | Not_Computer
Z-1694-2008 | 47766 | No | IZQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | The OptiVantage DH Injection System is a contrast delivery system and is designed to inject radiopaque contrast media into a patient's vascular system to obtain diagnostic images when used with computed tomography equipment. The J-Bow component is used to connect the powerhead to a suspension system. |  |  | April 07, 2008 | 2008 | Mallinckrodt Inc | Screws can begin to loosen and shear off to the point of the J-Bow falling from the suspension system. This may cause injury to patients or hospital staff. | Component design/selection | On April 18, 2008, Mallincrodt Inc. issued an Urgent Device Correction letter informing customers to check their inventory, and determine if they have any defective system units. If they observe any loose screws, sheared screws and/or elongated opening(s) around the screws, they are to immediately take the system out of service and contact the firm's service representative. Covidien's customers were all asked to inspect their inventories and to complete an 'Effectiveness Check Return Form' whether or not the product is defective. However, only those system units observed to have loose screws, sheared screws, and/or elongated openings around the screws, will be scheduled to receive correction by a Covidien Service Representative. |  | 783 | 783 Units | Product is distributed in the US throughout the following states: AL, AR, AZ, CA, CO ,CT, DC, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, and WV. | Terminated | May 20, 2011 | 1138 | 783 | Not_Computer
Z-1621-2008 | 47767 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD-MS ÂżÂż3 Ambulatory Infusion Pump, Model 7400 Smiths Medical MD, Inc., St. Paul, MN 55112 USA. Electromechanical pumps used for general drug delivery infusion therapies. Used mainly by home care patients but can also used in health care facilities. |  |  | March 07, 2008 | 2008 | Smiths Medical MD, Inc. | Smiths Medical discovered an issue with a motor component in specific serial numbers of the CADD-MS Âż3 Ambulatory Infusion Pump. This issue affects motor operation and may cause an over-delivery of medication which could result in injury to the user. | Component change control | Consignees were mailed an Urgent Medical Device Recall letter dated March 7, 2008. (Urgent Field Safety Notice dated March 7, 2008 for International Consignees.) The Urgent letter was addressed to "Risk &amp; Safety Mangers, "Distributors", "Clinicians", "Patients" and "other Customers". The letter contains a Description of the Problem, Results of Investigation, and Advice on Action to be taken by the User which includes contacting Smiths Medical MD, Inc. Customer Service Department to make arrangements for a replacement pump. The letter also requests consignees to return the confirmation form. |  | 646 | 646 | Class II Recall - Worldwide Distribution --- USA including states of PA, TX, KY, FL, MN, GA, HI, NM, MD, and countries of CZ, GB, IT. | Terminated | December 17, 2011 | 1380 | 646 | Not_Computer
Z-1526-2008 | 44927 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Datex-Ohmeda Neonatal Enhancement for the Engstrom Carestation, Model # 1505-3272-00, GE Healthcare P.O. Box 7550, Madison, WI 53707-7550, USA |  |  | July 02, 2007 | 2007 | Datex - Ohmeda, Inc | Neonatal Flow Sensors supplied for use with the Neonatal Enhancement for the EngstrÂżm Carestation may result in the sensors providing invalid values and alarm messages when exposed to high flow rates. | Nonconforming Material/Component | Consignees were sent a Medical Device Correction Notification letter on July 2, 2007. The firm's short term recommendation is that Neonatal Flow Sensors may continue to be used with added viligance around potential patient circuit disconnections and the associated alarms. Ther firm recommends that the MV exp high and Circuit leak alarms be set appropriately for each patient in order to detect a patient disconnect. The long term recommendation was to order the replacement part when available and dispose of the old part when the replacement is received. |  | 394 | 394 | Worldwide Distribution --- USA including state of TX and countries of Austria, Bangladesh, Belgium, Costa Rica, Denmark, Finland, France, Germany, India, Italy, Mexico, Saudi Arabia, South Africa, Sweden, Switzerland, United Kingdom, United Arab Emirates. | Terminated | July 03, 2012 | 1828 | 394 | I/O | Output/Calculation | Replace component | Remove or Replace
Z-1528-2008 | 46323 | No | GNB | Shigella spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Lee Laboratories brand Shigella Antiserum Poly Group B, catalog number 210722, 10 mL vials and labeled in part ***Lee Laboratories, Inc Grayson, Georgia 30017*** |  |  | November 30, 2007 | 2007 | Becton Dickinson &amp; Co. | In vitro diagnostic test reagent for identification of Shigella bacteria in patient samples may cause false negative results. | Nonconforming Material/Component | The recalling firm notified distributors by faxed letter flagged as "Urgent Product Recall" and e-mail on 11/30/07. Distributors were instructed to cease distribution and discard identified product and return a list of end user consignees for direct notification by Becton Dickinson. End users were notified by letter flagged as an "Urgent Product Recall" on 11/30/06 through 12/06/07. Customers were instructed to discontinue and discard affected product for replacement. The notification included a reply form that BD requested users to fax back with the quantity of replacement vials needed. |  | 643 | 643 (recall total) vials | Worldwide Distribution | Terminated | November 12, 2008 | 348 | 643 | Not_Computer
Z-1515-2008 | 46941 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER, MOSAIQ versions 1.3 and 1.4 with Elekta and Siemens Machines, Model number 40003-001000-RO, Manufactured by IMPAC Medical Systems, Sunnyvale, CA. |  |  | February 14, 2008 | 2008 | Impac Medical Systems Inc | Software issue may result in change to intended treatment field, potentially resulting in mistreatment. | Software design | Firm issued field advisory to its consignees on 02/14/2008. Phone calls will also be made to affected customers to ensure reception and understanding of the field notification. Firm advises users with affected product to review the jaw positions of IMRT fields to ensure that they match the RT Plan after an Apply Limits or Machine Change occurs, until software at their site can be updated to the corrected version. A software patch is under construction. |  | 36 | 36 units | Worldwide Distribution --- USA including states of OH, OK, NJ, MN, OR, TN, NY, CA, AZ, NV, IA, GA, AL, MD, MI, PA, IN, WA, FL, TX, CO, and countries to Canada, Australia, Germany, Greece and Sweden. | Terminated | June 24, 2009 | 496 | 36 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1444-2008 | 46942 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips Medical Systems iSite PACS (Picture Archiving and Communications Systems) version 4.1.17.0 - Radiology Image management system. |  |  | January 29, 2008 | 2008 | Philips Healthcare Informatics, Inc. | When using the Freehand Region of Interest (ROI) tool, errors may occur in area calculation yielding incorrect results. The firm discovered that a descrepancy exists in the measurement provided by the tool. | Software design | The firm issued product notification letter February 20, 2008, to its consignees. The notification alerts customers to the behavior, requires them to post the notification and notify users not to use the Freehand ROI feature until they are upgraded to a corrected software version. Philips will deliver a software upgrade to correct the problem. |  | 2 | 2 units | Worldwide Distribution --- USA including states of MN, CO, and PA, and country of Germany. | Terminated | September 25, 2008 | 240 | 5 | Computer
Z-1525-2008 | 38337 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | FACSDiva Software Versions 4.0.1 to 5.0.2, Flow Cytometry Acquisition &amp; Analysis software. Catalog number 641609 |  |  | April 05, 2007 | 2007 | BD Biosciences | When data file containing one or no fluoresence parameters is exported, the software will automatically apply compensation to this file and all subsequently exported files. | Software design | The firm has issued a customer Important Product Information letter via mail on 04/20/2007 to inform them of the software issues. The letter states that to avoid this problem, make sure each tube or well has at least two fluorescence parameters before exporting FCS 3.0 data files. Manually exporting tubes or wells one at a time or restarting the software after the export of an FCS data file containing one or no fluorescence parameters will prevent the problem from occurring. Customers are requested to fill out and return the response form and to contact BD Customer Support Center if there are any additional questions. A software fix is being developed in version 6.0 and a free maintenance release with no product enhancments (version 5.0.3). |  | 1074 | 1074 units | Worldwide Distribution --- USA including states of NY, NM, MI, GA, MD, MA, AL, CA, VA, NC, OH, TX, CO, NJ, PA, MI, CT, IA, MO, NV, KY, IL, TN, FL, MN, FL, UT, WA, OK, CT, WI, SD, AZ, and IN, and countries of Buenos Aires, BE, Brazil, Canada, Chile, China, Hong Kong, Japan, Singapore, Taiwan, Venezuela, Mexico, Trinidad and Tobago, and Columbia. | Terminated | September 30, 2009 | 909 | 1074 | Software | Output/Calculation | Software update | Software Update
Z-1750-2008 | 47764 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Omnifit EON Surgical Protocol; |  |  | May 17, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The Table 2 Broach and Neck Trial Sizing (Literature # LSP48 04/04) on page 3 is incorrect, reading 5mm less than what it should be for certain sizes. | Labeling design | Styrker Branches/Agencies were notified of the incorrect information printed in the protocal on May 17, 2007 via an Important Product Correction letter. The letter requested that all incorrect copies of the protocol be located and destroyed immediately. They requested that the Acknowledgement Form be completed and faxed, and that a Stryker Orthopaedics Customer Service rep be contacted in order to reorder the destroyed product. |  | 2501 | 2501 units | Worldwide Distribution --- including USA and countries of Canada, Australia, Latin America and Greece. | Terminated | August 19, 2008 | 460 | 2501 | Not_Computer
Z-1764-2008 | 47459 | No | LGC | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | HerpeSelect 1 and 2 Immunoblot IgG Kit, IB0900G |  |  | July 06, 2007 | 2007 | Focus Diagnostics Inc | A small percentage of samples tested in specific lots, the HSV common antigen band appears lighter (less reactive) than the reading control band. Reduced reactivity of the common antigen band may lead to inaccurate interpretation of a sample as negative. | Process design | Focus notified all clients which received the recalled lots starting on July 6, 2007 via an Important Medical Device Information letter by fax or email. The letter explains how to identify suspect results and the steps to be taken if suspect results are obtained. If customers find the interpretation of results with the indicated kit lots to be problematic, customers have the option of having their kits replaced at no charge within 90 days of receiving the notification. |  | 190065 | Lot Code 063169 - 97 Kits, Lot Code 063354 - 16 Kits, Lot Code 063355 - 74 Kits | Worldwide Distribution --- USA including states of OH, CA, CT, IL, NY, and countries of Barbados, Canada, Germany, Israel, Slovakia, France, and England. | Terminated | June 25, 2009 | 720 | 190065 | Not_Computer
Z-1370-2008 | 46771 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Exacta-Mix 2400, Pharmacy Compounding System is used in the pharmacy to compound multiple source ingredients into on final solution based on user-specified sequences and volumes. |  |  | February 11, 2008 | 2008 | Baxa Corporation | Flush solution from TPN compounder may be added to patient's TPN bag. | Software design | Consignees were notified by letter (Exacta-Mix Compounder Software Safety Alert) on 02/11/2008 advising them of the necessity to install a waste bag prior to a machine flush. |  | 1000 | 1000 units | Nationwide, including military facilities in CA, FL, HI, MS, NC, TX, VA and VA facilities in NY, OR, SC, UT. Foreign distribution to Australia, Canada, China and New Zealand. | Terminated | March 16, 2012 | 1495 | 1221 | Not Computer
Z-1322-2008 | 46870 | [6 empty] | Implantable Cardioverter Defibrillator |  |  | January 16, 2008 | 2008 | St Jude Medical CRMD | A condition that could lead to a ventricular sensing anomaly in Epic and Atlas family of lmplantable Cardioverter Defibrillators(ICDs) has been identified. A loss of ventricular sensing would prevent the ICD from being able to detect an arrhythmia. If an arrhythmia occurred and the device was unable to detect it, then a patient would not receive potentially life saving therapy. | Software design | A January 16, 2008, letter entitled, "Important Device Upgrade Information" was sent to physicians by FedEx. In the US, sales representatives will follow-up with the physician and will upgrade all programmers. For the rest of the world, it will be monitored by various St. Jude offices and local distributors with this information being fed back to St. Jude Medical CRMD in Sylmar. In addition, St. Jude Medical issued a press release on January 18, 2008 and information concerning this advisory can be found on St. Jude's website, www.sirn.com. |  |  |  | Nationwide and Internationally: Argentina, Australia, Austria, Barbados, Belgium,Bolivia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Denmark, Dominican Republic,Ecuador, Finland, France, Germany, Hong Kong, Hungary, India, Italy, Japan, Korea, Malaysia,Mexico, Netherlands, New Zealand, Norway, Peru. Poland, Portugal, Republic of Panama,Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, United Kingdom, Uruguayand Venezuela. | Terminated | April 03, 2012 | 1539 |  | Not_Computer
Z-1772-2008 | 47781 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Bariatric Wheelchair, 28" Wide and 20" Deep Seat, 700 lb. Weight Capacity; Manufactured in Taiwan for Medline Industries, Inc., Mundelein, IL 60060 USA; |  |  | March 31, 2008 | 2008 | Medline Industries Inc | Injuries to fingers. There is a potential for users to pinch their fingers between the seat bars when opening the wheelchairs. The injuries can be severe from fracture to severing. | Component design/selection | A Voluntary Field Correction letter dated 3/31/08 was issued to customers. The letter informed them of reports of users catching their fingers under the seat bars, which are identified on the chair with warning stickers. The customers were requested to change the backrest with the enclosed replacement upholstery kit, properly dispose of the old back rest, and place the enclosed cautionary tags on both sides of the chair seat. Instructions for installation of the upholstery and cautionary tags were included, as was a sheet entitled "How to Open a Wheelchair" which gives step by step instructions with illustrations for opening the wheelchair, as well as the warning "CAUTION KEEP YOUR FINGERS ON TOP OF THE UPHOLSTERY AT ALL TIMES. DO NOT PLACE YOUR FINGERS ON THE SIDES OR BOTTOM OF THE CHAIR SEAT." The accounts were requested to complete the enclosed response sheet and fax it back to Medline at 847-643-4482, indicating the serial numbers of chairs corrected. Any questions were directed to 800-633-5463, extension 5349, between 8:00 AM and 5:00 PM CST. Distributors were requested to notify their accounts of the recall. |  | 201 | 201 wheelchairs | Class II Recall - Worldwide Distribution --- including USA and Canada. | Terminated | May 11, 2010 | 771 | 1138 | Not_Computer
Z-1378-2008 | 46938 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Radius Capspin is a non-sterile instrument for use with other devices in spinal surgery. |  |  | February 19, 2008 | 2008 | Stryker Spine | The Radius Capsin has separation of the tip from the inserter. | Device Design | Recall notification was sent via FedEx on February 19, 2008 to all branches/agencies. All questions are to be directed to Michelle Klein, Regulatory Compliance Specialist - 201-760-8256. |  | 52 | 52 | Product was distributed to a limited population during a product launch for the purposes of gathering feedback on the product. | Terminated | August 19, 2008 | 182 | 117 | Not_Computer
Z-1183-2008 | 46316 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Power Supply Assembly Sled |  |  | October 17, 2007 | 2007 | Beckman Coulter Inc | Electrical grounding failure: During manufacturing, internal testing found one instrument wlth an intermittent failure ofthe electrical grounding. The source of the problem was traced to inadequate soldering. | Process control | Beckman Coulter mailed a Product Corrective Action (PCA) letter on October 17, 2007 to all Access and Access 2 lmmunoassay Systems, SYNCHRON LXi 725 System, and UniCel DxC 600i System customers.The customers were informed that during manufacturing, internal testing found one instrument wlth an intermittent failure ofthe electrical grounding. The source of the problem was traced to inadequate soldering. The letter referenced aproactive initiative to inspect the power supply sled assemblieson all instruments referenced. A response form was included. |  | 3247 | 3055 instruments in the US; 192 instruments in Canada | Nationwide Distribution and Canada | Terminated | April 27, 2012 | 1654 | 3247 | Not_Computer
Z-1184-2008 | 46336 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter UniCel Dxl 800 Access Immunoassay System; |  |  | October 08, 2007 | 2007 | Beckman Coulter Inc | Incorrect results: There is an increased variance in assay performance when the UniCel Dxl 800 Modification (MOD ID 10027)and/or UniCel Dxl 800 Modification (MOD ID 6733) have been implemented and there are room fluctuations after assay calibration between the extremes of the 18ÂżC and 32ÂżC operating range. | Device Design | Beckman Coulter mailed a Product Corrective Action (PCA) letter the week ofOctober 8, 2007 to the customers who have the affected UniCel Dxl 800Access lmmunoassay systems advising that there is an increased variance in assay performance when the UniCel Dxl 800 Modification (MOD ID 10027)and/or UniCel Dxl 800 Modification (MOD ID 6733) have been implemented and there are room fluctuations after assay calibration between the extremes of the 18ÂżC and 32ÂżC operating range. The letter further advises that Beckman Coulter will contact the customer to arrange a date to remove the modificatlons but until the modification is completed, the system should be operated in aroom temperature of 20ÂżC to 30ÂżC. This can be acomplished by monitoringand recording room tamperature in the area of the instrument. Customers are also asked to review historical quality control records to ensure assay performance is within their laboratory specification and that patient results should be reviewed in context of other diagnostics tests and the clinical presentations of the patient. A response form is included. |  | 120 | 111 units in the US; 9 units in Canada | Nationwide and Canada | Terminated | January 11, 2012 | 1556 | 120 | Not_Computer
Z-1189-2008 | 46728 | [6 empty] | Bio-Rad MONOLISA Anti-HBc EIA, Catalog Number 26186. In vitro diagnostic. |  |  | January 07, 2008 | 2008 | Bio-Rad Laboratories | Incorrect result: Elevated reactivity with a kit lot 232BGG. Patient samples that are negative for antibodies to anti-HBc could be assigned a positive result. | Process control | Bio-Rad called each customer and alerted them to the issue on 1/7-8/2008. A letter was faxed and sent via overnight carrier (DHL) to each customer site on 1/8/2008. Customers were told not to use the affected lot of the product and to discard it. The firm will provide another lot of the product to customers. |  | 31 | 31 distributed | Nationwide Distribution. | Terminated | March 07, 2011 | 1155 | 31 | Not_Computer
Z-1478-2008 | 46781 | No | CCX | Patient position support. | Anesthesiology | Anesthesiology | 510(K) Exempt | Allen Medical C-Flex, Polar Head Positioner |  |  | November 27, 2007 | 2007 | Allen Medical Systems, Inc. | Handle jamming in the open position, preventing the device from adequately supporting the head during spinal surgery. Possible delay in surgery, or injury. | Device Design | Allen Medical notified accounts by telephone on 11/08/07 and the Sales Representative. The Hospitals confirmed that the device would be sent back for evaluation and upgraded. |  | 2 | 2 | MA, MS | Terminated | August 26, 2008 | 273 | 2 | Not_Computer
Z-1447-2008 | 46388 | No | NEZ | Ingestible telemetric gastrointestinal capsule imaging system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | GivenÂż Diagnostic System with PillCamÂż SB Capsule - Gastrointestinal capsule imaging system. The device is an ingestible telemetric gastrointestinal capsule imaging system is used for visualization of the small bowel mucosa as an adjunctive tool in the detection of abnormalities of the small bowel. The device captures images of the small bowel with a wireless camera contained in a capsule. This device includes an ingestible capsule (containing a light source, camera, transmitter, and battery), an antenna array, a receiving/recording unit, a data storage device, computer software to process the images, and accessories. |  |  | December 13, 2007 | 2007 | Given Imaging Inc | The cards included in the product's education kit (proposing a clinical algorithm for capsule endoscopy in patients with suspected Crohn's disease), were printed incorrectly. | Employee error | Consignees were notified by letter on 12/13/2007. The letter contains a diagram illustrating the error on the incorrect card and a diagram illustrating the corrected card. Consignees were instructed to discard the incorrect card if they still have it and to complete a response form and return it via fax to Given Imaging. |  | 1392 | 1392 units | Worldwide Distribution --- US including Puerto Rico and country of Canada | Terminated | November 15, 2010 | 1068 | 1392 | Not_Computer
Z-1695-2008 | 47824 | No | LZS |  |  | Ophthalmic | PMA | Excimer Laser System; WaveScan WaveFront System, Software Version 3.65, for use with the VISX STAR S4 and STAR S4 IR Excimer Laser Systems. System is a diagnostic instrument intended for the automated measurement, analysis and recording of refractive errors and wavefront aberrations to the target shape and then generates the commands for the VISX STAR laser to create the target shape on the cones. |  |  | November 17, 2006 | 2006 | VISX INCORPORATED, A SUBSIDIARY OF AMO INC | Firm became aware of a significant overcorrection in one eye of a patient who had received a wavefront-guided LASIK treatment at a facility in Germany. The probability of a recurrence of this error was determined to be extremely remote. To ensure that this type of incident will not happen again, a software update was developed. | Software design | VISX issued a notification letter on November 27, 2006, to all distributors and direct accounts. The letter informed the recipients that they were performing a field correction in the form of a mandatory upgrade on all VISX WaveScan WaveFront Systems and Offline Programming Modules to correct an issue. The field upgrade should be made by February 16, 2007. A Dear Doctor letter was also issued instructing surgeons to upgrade all Offline Programming Modules using an enclosed CD. |  | 1310 | 1310 units | Class 2 Recall - Worldwide Distribution: USA, Venezuela, Colombia, Italy, Saudi Arabia, Kuwait, Egypt, UAE, Jordan, Lebanon, Australia, Canada, France, Germany, India, Ireland, Japan, Sweden, Vietnam, Czech Republic, Slovak Republic, Bulgaria, Philippines, South Korea, Turkey, UK, Portugal, Spain, Argentina, Israel, Cyprus, India, Mexico, Greece, Singapore, Chile, Poland, Brazil, China, Thailand, Taiwan, Hong Kong and Russia. | Terminated | June 08, 2010 | 1299 | 1310 | Software | Output/Calculation | Software update | Software Update
Z-1769-2008 | 47773 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Medical charged-particle radiation therapy system. Siemens Linear Medical Accelerators equipped with COHERENCE Therapist R2.0 and Oncologist R2.0. Product update, TH 030//07/S. The COHERENCE Therapist R2.0 software numbers are: 7358299, 7360857, and 8141835 and applies only to those parts used with the following devices: Coherence AG Therapist; Coherence Therapist System; Coherence Impression Therapist; PRIMEVIEW 3i; AG Therapist 3rd party V&R; Coherence Therapist 2.0 and PRIMEVIEW 3i 2.0. |  |  | October 03, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Software. The firm issued an Advisory Letter Update instruction TH030/07/S as a result of a complaint that images viewed in the device may shift, which could result in mistreatment. The advisory alerts all its affected customers of the image issue to prevent potential mistreatment. | Software design | Customer Safety Advisory Notice was issued on 10/03/2007 to consignees, via certified mail. Consignees are strongly cautioned that they should not further correct the image alignment at this time as doing so will result in incorrect offset values which could lead to incorrect repositioning of the patient. Consignees are instructed to check the "Filter Images on Loading" and "Saving of Filtered Images" config settings, and apply any additional filters to the image BEFORE using the interactive shift or landmark registration tools. Do not correct any perceived issues. A software fix is being validated and will be available soon. |  |  |  | Class II Recall - Worldwide Distribution --- including USA and countries of Angola, Australia, Brazil, Canada, Croatia, Denmark, Egypt, France, Germany, Greece, Hungary, India, Iran, Italy, Japan, Malaysia, Netherlands, New Zealand, P.R. China, Poland, Portugal, Republic Korea, Russian, Fed. Saudi Arabia, Slovakia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Thailand, Trinidad, Tobago, Turkey, United States, Vietnam. | Terminated | April 29, 2011 | 1304 |  | Software | Display/Image | Software update | Software Update
Z-1773-2008 | 47785 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Infusion System. EKOS EndoWave Infusion System. EKOS EndoWave Infusion System Drug Delivery Catheter with Ultrasound Core. Labeled as Sterile, Single Use Only. Connector clip is CIC Connector Clip. The infusion system is intended for the controlled and selective infusion of physician-specified fluids, including thrombolytes, into the peripheral vasculature. |  |  | April 01, 2008 | 2008 | EKOS Corporation | Catheter disconnection. There is a potential for the catheter to disconnect from the Catheter Interface Cable (CIC) pod which connects to the Control Unit. | Other | Firm representatives will visit and distribute connector clips and instructions for use to consignees beginning 04/01/2008 at no cost to the consignee. |  | 218 | 218 units | Class II Recall - Nationwide | Terminated | May 13, 2011 | 1137 | 218 | Not_Computer
Z-1186-2008 | 46345 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet brand Modular Microplasty Cup Impactor, 3/8" - 24 Thread Insert; Model 31-400603. Instrument for hip prosthesis acetabular cup component. |  |  | January 16, 2008 | 2008 | Biomet, Inc. | The weld at the lock location may fracture during impaction. | Device Design | Biomet notified distributors and medical facilities by letter dated 1/16/08. Distributors with distributor-owned instruments at medical facilities were notified by email on 1/17/08, and provided with a letter to present to their associated medical facilities describing the recall. |  | 59 | 59 | Worldwide Distribution: USA, Germany, Mexico, Netherlands, Sweden and United Kingdom. | Terminated | September 22, 2009 | 615 | 1004 | Not_Computer
Z-1539-2008 | 46960 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKNEE Tibial Insert Trial, Size 2, 13mm Right; Part #801-01-640; manufactured by Encore Medical, L.P., Austin, TX 78758. |  |  | February 20, 2008 | 2008 | Encore Medical, Lp | Tibial insert trial incorrectly marked for thickness; device thickness marked correctly as 13mm on the underside of the trial, but is incorrectly marked as 15mm on the side. | Error in labeling | Firm notified consignee by phone and asked them to return any outstanding product on 02/20/08. |  | 6 | 6 units. | Distribution --- country of Germany. | Terminated | September 02, 2008 | 195 | 6 | Not_Computer
Z-1514-2008 | 46297 | Yes | LWP |  |  | Cardiovascular | PMA | VITALITY EL, models T127, T167, T177 Implantable Cardioverter Defibrillator (ICD), pulse-generator, dual Chamber |  |  | January 07, 2008 | 2008 | Boston Scientific CRM Corp | Repetitive mechanical stress applied to the titanium case can induce component damage and device malfunction in certain ICD and CRT-D devices implanted subpectorally with the serial number facing the ribs. Devices implanted subcutaneously or in a subpectoral position with the serial number facing away from the ribs are not included in this advisory. This update identifies additional VITALITY ICD | Device Design | An Urgent Medical Device Information Update to May 2006 Product Advisory, dated January 04, 2008, was sent to implanting physicians and physicians following implanted devices beginning Monday, January 07, 2008. The letter described a malfunction that can occur when the ICD is implanted subpectorally with the serial number facing the ribs. Devices implanted subcutaneously or in a subpectoral position with the serial number facing away from the ribs are not included in this advisory. The physicians were asked to review records to determine if an affected device was implanted subpectorally. An AP radiograph should be done to determine specific device orientation for subpectoral implants. If a subpectoral implant is in a susceptible orientation, consider repositioning or replacement for physically active patients who regularly need device therapy. Follow patients with susceptible devices at 3-month intervals. |  | 25000 | BSC CRM estimates that the number of devices implanted in a susceptible orientation is likely less than 1% of the total active implants 25,000 (w.w.) | Nationwide and Worldwide: Bolivia, Panama, Iran, Saudi Arabia, Kuwait, United Arab Rep., Austria, Belgium, Cyrus, Czech Republic, Denmark, Finland, France, Germany, Greece, Iceland, Ireland, Italy, Netherlands, Norway, Portugal, Spain, Sweden, Switzerland, UK, India, Australia, New Zealand, Malaysia, Singapore, Hong Kong, China, Israel, Turkey, Thailand, and Canada. | Terminated | December 30, 2008 | 358 | 25000 | Not_Computer
Z-2139-2008 | 47801 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Sapphire Hematology Analyzer, List number 08H00-01, LIS/ Middleware Interface Specification, Abbott Diagnostics Division, Santa Clara, CA |  |  | April 11, 2008 | 2008 | Abbott Laboratories | Table in LIS Specification depicting the association between Record ID and Numerical Result Label is incorrect. If a record ID is used to configure the system for mapping, results will come out nonsensical. | Software change control | Recall initiated on April 11, 2008. A Product Correction letter and customer reply form were sent to all currently active CELL-DYN Sapphire Customers. The letter requests that users be certain that the transmission of test results receipt by the host interface is mapped according to the Numerical Result Label, when configuring their CELL-DYN Sapphire Hematology Analyzer for results transmission to a Clinical Laboratory Information System or Middleware. Affected product was distributed to a total of 4 distributors. However, notification was made to all 154 customers that received the CELL-DYNE Sapphire product. If the customer or the health care providers that they serve have any questions regarding this information, U.S. customers should call Customer Support at 1-877-4ABBOTT. |  | 15 | 15 copies of the specification were sent worldwide | Nationwide Distribution. | Terminated | March 12, 2010 | 700 | 15 | Not_Computer
Z-1471-2008 | 46903 | No | KEY | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Xmatrx FISH (fluorescent in situ hybridization) Automated Staining System, Catalog number AS4010B |  |  | January 24, 2008 | 2008 | Biogenex Laboratories | Bulk reagent dispensing mechanism may break resulting in reagent spill and possible exposure to chemicals. | Component design/selection | The firm initiated written notification on 01/24/2008. The firm has asked its consignee to inform its consignees of the recall. |  | 10 | 10 units distributed. | Product was distributed to a single distributor, which further distributed the product. | Terminated | January 08, 2011 | 1080 | 10 | Not_Computer
Z-2150-2008 | 48184 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Equal Diagnostics Lipase Color Reagent Kit |  |  | April 14, 2008 | 2008 | Genzyme Corporation | Calibration and Quality Control outside of established range resulting from an increase in initial OD readings with the reagent contained within the kits. | Nonconforming Material/Component | Genzyme notified accounts by Fed'X on 4/14/08 via an Imprtant Field Correction letter, requesting users to discard kits and follow the "User Update" instructions previously provided to minimize calibration error. Customers are to complere and sign the Confirmation of Notification form and fax it back to the firm regardless if they have any affected product or not. Also, requires that the customer confirms that the REagent OD limits for First and Last H are set at 0.15 for Lipase on the Olympus AU400/400e, AU600, AU640/640e and AU2700/AU5400 Systems. |  | 3443 | 3,443 | Nationwide Distribution | Terminated | September 16, 2008 | 155 | 3443 | Not_Computer
Z-2140-2008 | 48465 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Osteo-Clage Cable System Package Assembly (01-0001-S). Product labeled in part, "2.0mm x 850 mm CoCr Cable / Sleeve... REF 01-0001-S LOT W59192... Acumed 5885 NW Cornelius Pass Road Hilsboro, OR 97124-9432 ". |  |  | March 28, 2008 | 2008 | Acumed LLC | Cables of incorrect diameter were used in product assembly, resulting in the sleeve not being able to bind to the cable after crimping. | Employee error | On 3/28/08, the firm sent an Urgent Notification letter to their consignees. The letter explained the issue and requested the customers return the product to Acumend. On 7/16/08, the firm sent a 2nd notification, via mail, to thier customers. The Urgent Notice: Device Recall 2nd Notification, advised customers of the recall, the hazard and requested customers to return existing product to Acumed. |  | 50 | 50 units | Nationwide Distribution --- including states of AR, CA, PA, GA, KS, MO, and TN. | Terminated | February 03, 2010 | 677 | 50 | Not_Computer
Z-1541-2008 | 46961 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | IntelliVue Clinical Information Portfolio |  |  | March 04, 2008 | 2008 | Philips Medical Systems | Medications prescribed for one patient were printed on the record of another patient. | Software design | Philips Medical issued letter notification, dated 29 Feb 2008 to customers titled; "Urgent Medical Device Correction Notice".Customers were asked to not solely rely on the Worklist to remind them when to administer orders or interventions until the software has been upgraded. Customers should review the reminders and verify orders and interventions on the Medication Administration Record and Administration record. The correction will be in the form of a software upgrade and will be performed remotely by Philips service personnel. |  | 138 | 138 for products | NationwideCanada, Australia, Austria, Belgium, Denmark., France, Finland, Germany, Great Britain, Italy, Malaysia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, and UAE | Open, Classified |  |  | 138 | Computer
Z-1532-2008 | 46908 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Quickie Groove Power Wheelchair with QC (non-expandable) electronics, Model number EIPW13. |  |  | January 11, 2008 | 2008 | Sunrise Medical Inc | Product experiences a longer than expected stopping distance, which may cause injury to the user, objects, or the chair itself. | Software design | Initial phone contact began on January 11, 2008. Medical device correction letters were also sent to affected dealers. |  | 141 | 141 units total | Product was distributed to 126 retailers nationwide throughout the USA, as well as Canada and Australia. | Terminated | November 16, 2009 | 675 | 141 | Not_Computer
Z-1542-2008 | 47252 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Altaire MRI System Emergency Rundown Switch Unit (ERDU) is intended to provide the physician with physiological and clinical information obtained non-invasive and without the use of ionizing radiation. |  |  | February 11, 2008 | 2008 | Hitachi Medical Systems America Inc | The firm's Altaire MR Scanner's Emergency Rundown Switch Unit (ERDU), that is used to shut down the device in an emergency, was found to be defective. An ERDU failure could delay or prevent patient access leading to possible patient death or serious injury. | Component design/selection | The firm issue a DEVICE CORRECTION letter to their customers on 2/11/2008. The letter informs the customers of the Emergency Rundown Switch Unit (ERDU) defect and provides a warning which lists instructions for the safe use of the ERDU Unit. The letter informs the customer(s) that their local Hitachi Field Service Engineer (FSE) will be scheduling a visit to test the ERDU in order to verify its state of operation. The letter also states that This test requires specific tools in order to verify the operation without actually quenching the magnet and that Hitachi will begin the test ING in March, 2008 and will continue the program until all installed ERDU units are verified.Should an emergency arise requiring you to activate the ERDU, verify the Emergency Run Down light illuminates and that the magnet does quench (a very loud noise will be audible) before entering the scan room without screening personnel and equipment. Notify the Hitachi Customer Service Department at (800) 800-4925 immediately should you initiate an Emergency Run Down. |  | 235 | 235 units | The recalled device units were distributed to customers located in the following states nationwide: CA, OH, IL, NJ, NY, FL, TX, NE, AR, MD, OR, MN, NV, MS, OK, IN, TN, MO, SC, VA, HI, LA, NC, NM, WY, WA, AL, PA, UT, WA, and KY | Terminated | September 24, 2010 | 956 | 235 | Not_Computer
Z-2156-2008 | 48472 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Medtronic Midas Rex Legend tapered dissecting tool, Model F2/8TA23 |  |  | May 30, 2008 | 2008 | Medtronic Powered Surgical Solutions | Cutting flute geometry of surgical dissecting tools may be outside of established tolerance. This may lead to reduced cutting performance, inability to perform desired dissection, increased risk of overheating, and potential need for revision surgery. | Process control | Firm notified consignees by Urgent: Medical Device Recall letter sent May 30, 2008. The letter requires consignees to immediately compare their inventory of products against the list provided and quarantine any unused dissecting tools. The letter also states that Medtronic would provide replacement product to match the quantity and type returned by the consignees' institution. |  | 1306 | 1,306 units | Nationwide Distribution. | Terminated | August 22, 2009 | 449 | 1306 | Not_Computer
Z-2157-2008 | 48596 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Stem Trial, Triathlon Revision Instruments; Non-Sterile; 100 mm length |  |  | May 14, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Triathlon TS Stems Trails may crack, potentially causing fragments of the plastic handle to fragment and break off completely. | Process control | Urgent Product Recall notification letters were sent to all Stryker Branch/Agency Manager/Quality contacts on May 14, 2008 via Federal Express. The letter requests that consignees examine their inventory and hospital consignment locations to identify the product; and retrieve all affected product lots and return it to their branch or agency warehouse for reconciliation. The letter also states that a Stryker Orthopaedics Customer Service rep should be contacted for alternative products available. |  | 16862 | Catalog Number 5565 T-009 to 025: 2629 US, 238 international: Catalog 5566 T-009 to 025: 2625 US, 239 international | Worldwide Distribution --- including USA and countries of Australia, Canada, Sweden, Germany, Spain, Portugal, United Kingdom. | Terminated | January 26, 2010 | 622 | 25292 | Not_Computer
Z-2234-2008 | 48685 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | StatSampler 200 uL, Li Heparin (micro blood collector) |  |  | June 20, 2008 | 2008 | StatSpin, Inc., dba Iris Sample Processing | EDTA inadvertently mixed with lithium heparin solution contained in capillary tubes. | Mixed-up of materials/components | Iris Sample Processing notified customers by telephone followed by an Urgent: Product Recall letter dated 6/19/08. Users are requested to return inventory, to not use product from Lot Number 7296, and to repeat tests where the recalled tubes were used.The firm identified an additional lot : 8031 was effected and recalled the lot to two consignees. |  | 171 | 171 vials | Worldwide Distribution --- USA including states of CA, CO, FL, KY, KS, GA, NC, NJ, OR, PA, and countries of Canada, Hong Kong, UK. | Terminated | May 30, 2012 | 1440 | 171 | Not_Computer
Z-1234-2008 | 46383 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Nautica (GK) Mattress Stretcher, manufactured by Georg Kramer GES.M.B.H. &amp; CO. KG A-4061 PASCH ING, Neuhauserweg 14 Austria. These mattress stretchers were sold for installation on or as a component of the following patient bath lifts: a) Manufactured by BHM Medical Inc., Magog, QC, Canada J1X 5Y5: Part number 88004051, Nautica Bath Lift System; b) Distributed by Whitehall Manufacturing, P.O. Box 3527, City of Industry, CA 91744, Manufactured by BHM Medical Inc., Canada: Part number 88004051.12, Nautica Bath Lift System |  |  | January 23, 2007 | 2007 | Arjo, Inc. | Repeated use of the ratchets that lock the position of the back and leg rest of the stretcher may damage the retention pin that prevents the latch from disengaging from the anchor system. This would allow the backrest to pivot freely down to the floor. | Device Design | BHM Medical initially notified the end users on 1/23/07 of the potential for the back and leg rest latch ratchets to become damaged, allowing the rests to pivot freely. Should a patient be in the stretcher at the time, the patient could fall, potentially sustaining cuts, bruises or even more severe injuries. The letters provided safety instructions to follow until a permanent fix is available, and requested the accounts to complete and return the enclosed response form acknowledging receipt of the letter and dissemination of the safety instructions to all employees operating the lifts with the affected GK stretchers.BHM sent Customer Field Correction Notification letters dated 1/23/08 to the accounts, along with the Technical Advisory Notice - TAN A dated 1/3/08, reiterating the safety instructions from the 1/23/07 letter and requesting the accounts to, if at all possible, remove the equipment from service until an authorized engineer can replace the ratchets in the backrest and leg rest latches.All consignees have been contacted. |  | 7 | 7 units | Nationwide Distribution --- including states of California, Connecticut, Florida, Mississippi and Texas. | Terminated | August 25, 2008 | 580 | 7 | Not_Computer
Z-1564-2008 | 46807 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision RXi Digital Remote X-Ray Imaging R&F; System. It is a remote R&F; system consisting of a tilting table with integrated spot film device, 65 or 89 kW pulsing generator and high heat capacity X-Ray tube. The system can be configured to meet a wide variety of digital record and fluoroscopy exams for various clinical applications with either 32 cm or 40 cm image intensifier. |  |  | December 18, 2007 | 2007 | GE Healthcare | GE Healthcare Precision RXi tilt joystick became stuck in the Trendelenburg direction causing the patient to slide off the table. Also recall conducted to investigate the possibility of accidentally activation of the remote console joystick where a fatal accident occurred on a non-GE Remote R&F; System. . | Device Design | Two letters "GE Product Safety Notification" dated 12/18/07 and 3/7/08, addressed to Hospital Administrators, Manager of Radiology and Radiologists, were sent to each consignee. Each letter addressed a separate safety issue. Letters described the affected product, safety instructions, corrective action and contact information. The joystick can become stuck causing unintentional motion of the tube.Field Modification Instructions (FMI 10802/10803) have been issued to the field to install additional warning labels on the patient table, install a protector for the console joystick and inspect and correct the clearance between the console joystick and its mounting location.If your facility has a Precision RXi or Precision e system manufactured before March 2007, a GE Healthcare Field Engineer will schedule a field visit to install the added joystick protection. This upgrade will be performed at no charge.If you have any questions/concerns regarding this issue, please contact your local GE Healthcare sales or service representative. |  | 193 | 193 | (USA)AKCOILLAMAMIMSNCNJNYPASCTNUTVAWAWI(OUSA)BR CA MX AU C3 KR PH TW AE AT BE BF BG CM CY CZ DE ES FI FR GB GE GR GY HR HU IT LB LV MT NA NG PL PT RO RU SA SD SE SI SK SY TN TR UA YU ZA | Terminated | August 09, 2012 | 1696 | 265 | Not_Computer
Z-1567-2008 | 46921 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | FASTPAK batteries labeled Manufactured for Medtronic Physio-Control 11811 Willows Rd, Redmond, WA USA. Batteries are Nickel-Cadmium (NiCd) battery packs used to provide DC power to LIFEPAK 5, 10, 11, or 12 defibrillator/monitors. The FASTPAK battery has a capacity of 1 amp/hour (1.0AH). The FASTPAK battery can be charged in a Battery Support System, Battery Support System 2, Mobile Battery Service Station, or the LIFEPAK 10, 11 or 12 defibrillator/monitor when the device is powered by an external power adapter. The FASTPAK battery does not have a pushbutton fuel gauge like the LIFEPAK NiCd. The fuel gauge provides a visual indication of battery capacity. |  |  | February 27, 2008 | 2008 | Physio Control, Inc. | Batteries mislabeled - batteries labeled as 2.4 amp/hour batteries, however are 1.0 amp/hour batteries. | Labeling mix-ups | On February 27, 2008, the firm sent a "URGENT MEDICAL DEVICE - ACCESSORY RECALL" letter via certified return receipt to their consignees. The letter states the batteries are fully functional 1.0 amp/hour batteries mislabeled as 2.4 amp/hour batteries and that the batteries will provide half the monitoring time and half the number of defibrillation shock as of a 2.4 amp/hour LIFEPAK NiCd battery. Enclosed with the letter is a confirmation sheet for the consignee to indicate the number of recalled batteries they have and fax back to the firm. The firm will send out replacement batteries and the consignee is to then send back the mislabeled batteries.If there are any questions, contact the technical support group at 1-800-442-1142 - option 5 or visit www.physio-control-notices.come/fastpak. |  | 811 | 811 batteries | Batteries were distributed nationwide and to Canada, Netherlands, Asia Pacific and Latin America. | Terminated | March 07, 2011 | 1104 | 811 | Not_Computer
Z-1615-2008 | 46965 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art System with software versions 3.0 and 3.1 - Radiation Therapy System. The device is intended to be used as an integrated system for the planning and delivery of intensity modulated radiation therapy. |  |  | February 22, 2008 | 2008 | TomoTherapy Incorporated | Registration adjustment field numbers may appear on printed pages of the Register Tab as nonsense characters or values different than the numbers shown on the Operator Station (OS) display. Setup will move the couch to the correct offset position. If the couch is manually adjusted to the incorrect value, the patient may not be the correct position during treatment. | Software design | Consignees were sent a letter "TomoTherapy Incorporated Urgent Medical Device Correction Field Safety Notice" dated February 22,2008. The letter described the Issue, Product Affected, Cause, Recommended Customer Action and Resolution.If you have not received this letter or have questions, contact the firm as outlined below:North America: 1 866 368 4807Belgium: 0800 38783France: 0805 631 565Germany: 0800 000 7541Italy: 800 986 399Netherlands: 0800 0201364Spain: 800 300049Switzerland: 0800 001927United Kingdom: 0808 238 6035China/CNC (Northern): 10 800 712 1701China (Southern): 10 800 120 1701Hong Kong: 800 967912Japan: 0044 22 132374Singapore: 800 1204 683South Korea: 0079 81 4800 7204All other locations: +1 608 824 2900 or +32 2 400 44 44 |  | 32 | 32 | Worldwide Distribution --- USA including states of CA ,CT, DC, FL, IL, LA, MD, MI, MO, NE, NJ, NY, SC, TN, VA, WI , and countries of Canada, Germany, Taiwan, and Turkey. | Terminated | December 02, 2008 | 284 | 32 | Software | Display/Image | Instructions | Safety Notice/Insructions
Z-1558-2008 | 47463 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 24mm Rotating C-Mount Coupler, Model Number 1088-020-121S2. The device in indicated for use in general laparoscopy, nasopharyngoscopy, ear endoscopy, sinuscopy, and plastic surgery wherever a laparoscope, endoscope or arthroscope is indicated for use. |  |  | March 31, 2008 | 2008 | Stryker Endoscopy | The coupler on this device was manufactured without set-screws that are used to hold and center the rear assembly and the endobody. The coupler is part of the camera system and affects the ability to focus. | Nonconforming Material/Component | Only one customer had all of the defective couplers. The firm called this customer on March 27, 2008, and asked that the 28 couplers remaining in their possession be returned. |  | 28 | 28 units | Product was delivered to one consignee in Iowa. | Terminated | September 17, 2008 | 170 | 28 | Not_Computer
Z-1640-2008 | 47441 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Logix TPN Software Solution, catalog 2M8410, consists of the Logix Order Entry (OE) Installation CD, version 1.0.28 and Logix Compounder Manager (CM) Installation CD, version 3.1.8. The software is included in a number of Baxter's pharmacy compounders. The software assists the pharmacist in management of compounding, creating and authorizing nutritional solutions intended for IV administration. |  |  | March 28, 2008 | 2008 | Baxter Healthcare Corp. | Software anomaly results in inaccurate information being printed out on the Delivery Report, even though the compounder performed the compounding correctly. | Software design | Baxter sent the 3/28/08 Urgent Device Correction letter to affected customers, to the attention of the Director of Pharmacy, via first class mail to advise them of the potential inaccuracy of the Delivery Report due to a software anomaly. The accounts were instructed to review each delivery report for accuracy. If the report contains zero actual volumes the user must manually reprint the report. The letter outlined the process for reprinting the Delivery Report. Any questions were directed to the Baxter Nutritional Professional Services at 1-800-422-2751. The accounts were requested complete the enclosed reply form acknowledging receipt and understanding of the letter and dissemination of the information to their staff. A software upgrade will be provided.If you have questions about this letter or you did not receive this letter contact Baxter Nutrition Professional Services at 1-800-422-2751. |  | 5 | 5 units | Arizona, Minnesota, Wisconsin, Puerto Rico and Canada | Terminated | August 25, 2008 | 150 | 5 | Software | Output/Calculation | Software update | Software Update
Z-1641-2008 | 47471 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD-MS"3 Ambulatory Infusion Pumps, Model 7400, Rx only, CE 0473, SN 499000. The device is designed for subcutaneous, intravenous, epidural and intrathecal infusion of medication. |  |  | February 22, 2008 | 2008 | Smiths Medical MD, Inc. | Smiths Medical has received reports that the device powers down without an alarm. Investigation concluded that if the battery cap is not fully tightened, the pump may power down and a brief "chirp" will sound. This may occur if the user has not sufficiently tightened the battery cap or if the battery cap is damaged. If a patient is asleep or does not regularly check their Pump, there is a potentia | Labeling design | A "Smith Medical Urgent Medical Device Correction" letter dated February 22, 2008 was distributed to consignees. The letter addressed to Risk/Safety Managers, Distributors, Clinicians, Patients and other Customers, instructs them to fully tighten the battery cap and to routinely monitor the pump for damage that may result in power down situations. Consignees are to return the "Urgent Medical Device Correction" form.If you have questions about the letter or did not received the letter contact Mike Herbert at 651-628-7049 or Jeff Esselman at 651-628-7920. |  | 2303 | 2303 | AZCAFLGAHIILKYMDMIMNNCNENJNMNYPATNTXUTCANADACHILECZECH REPUBLICGERMANYGREAT BRITAINITALYJAPANNETHERLANDSNORWAYSWEDEN | Terminated | January 11, 2011 | 1054 | 2303 | Not_Computer
Z-1643-2008 | 47524 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syn go Imaging XS is a picture archiving and communication system intended to display, process, read, report, communicate, distribute, store and archive digital radiological images, including digital mammography images. |  |  | March 10, 2008 | 2008 | Siemens Medical Solutions USA, Inc | The firm discovered a problem where the selected patient images on the device may display an additional image from another patient or study. | Software design | The recalling firm issued a Customer Safety Advisory via certified mail to affected customers per Update Instructions IM011/08/S. The letter informs customers of the potential issue and provides instructions to detect its occurrence. A software update to correct the issue is already available. A service representative will visit affected sites to install the software update.If you have questions concerning the advisory letter or you did not receive a letter contact Alexander Vogel at +49 (9131) 84-3840 or alexander.vogel@siemens,com and make reference to complaint 08-1KM-0073 date 2008-03-06. |  | 236 | 236 units | The product was shipped to medical facilities in AL, AZ, CA, CO, CT, FL, GA, IA, KS, KY, MD, MI, MN, MO, NC, ND, NE, NJ, NM, NV, NY, OH, OK, PA, TN, TX, WI, and WV. | Terminated | November 19, 2008 | 254 | 236 | Software | Display/Image | Software update | Software Update
Z-1648-2008 | 47758 | [6 empty] | Seprafilm Adhesion Barrier (chemically modified sodium hyaluronate/carboxymethylcellulose absorbable adhesion barrier). Japan Distribution Only |  |  | April 10, 2008 | 2008 | Genzyme Corporation | Sterility may be compromised | Employee error | Genzyme notified Distributors and hospitals by letter dated 4/10/08, requesting ceasing distribution and return recalled product. |  | 517 | 517 (10 packs) | NationwideKorea, Taiwan, Japan | Terminated | September 29, 2015 | 2728 | 3149 | Not_Computer
Z-1697-2008 | 47871 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Dental Cement-Bisco BISCEMÂż Dual-Cured Self-Adhesive Resin Cement, Opaque; The product is intended to serve as a temporary tooth filling or as a base cement to affix a temporary tooth filling, to affix dental devices such as crowns or bridges, or to be applied to a tooth to protect the pulp. |  |  | April 28, 2008 | 2008 | Bisco Inc | During product stability testing, Bisco determined that after 15 months of storage, BISCEMÂż Self-Adhesive Luting Cement used in self-cure mode only without light-curing may take longer than specified to set, resulting in reduced bond strength. Bisco also determined that product component(s) may thicken slightly over time. | Incorrect or no expiration date | Bisco sent Urgent Medical Device recall letters dated 4/25/08 to the U.S. consignees via 1st class mail on 4/28/08. The letters informed the users that after 15 months of storage, BISCEMÂż Self-Adhesive Luting Cement used in self-cure mode only without light-curing may take longer than specified to set, resulting in reduced bond strength. Future production lots will be labeled with a shelf life of 1 year with a requirement to refrigerate the product when not in use. The accounts were requested to destroy their remaining stocks of the product and record the number of used or partially used syringes destroyed, indicating the lot numbers of the units destroyed on the enclosed response form and fax it back to Bisco at 847-534-66101. Questions were directed to Bisco Dental Products Customer Service at 1-800-247-3368. |  |  |  | Nationwide including Puerto Rico and internationally to Argentina, Australia, Bulgaria, Brazil, Canada, Chile, China, Colombia, Costa Rica, Denmark, Ecuador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Iceland, India, Indonesia, Israel, Italy, Jordan, Korea, Lebanon, Malaysia, Netherlands, Norway, New Zealand, Peru, Philippines, Portugal, Romania, Saudi Arabia, Serbia, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, Vietnam | Terminated | September 14, 2010 | 869 |  | Not_Computer
Z-1703-2008 | 47522 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Circular Mapping Catheter-Biosense Webster LASSO 2515 Variable Circular Mapping Catheter- The catheter is indicated for multiple electrode electrophysiological mapping of the cardiac structures of the heart, i.e., recording and stimulation only. The device is designed to obtain electrograms in the atrial regions of the heart. |  |  | March 25, 2008 | 2008 | Biosense Webster Inc | This recall was initiated after the firm received several complaints reporting issues with deflection. The catheter may become caught in the mitral valve, requiring surgical procedures, valve may be torn while trying to dislodge, unable to retract the catheter through the sheath and require surgical intervention to remove, pulling on catheter to remove may lead to large atrial septal tear, and sep | Device Design | A letter entitled, "Urgent Safety Field Notice" dated March 25, 2008, and certification form were sent to all customers and Johnson &amp; Johnson affiliates who purchased the product in the past year as of March 25, 2008, due to the expiration date. For customers in the US, two letters and forms were sent out - one addressed to the EP/Cardiology Lab and a second to the Risk Management Office.If you have questions or have not received a letter contact Customer Service at 1-800-729-9010. |  | 14295 | 14,295 units for both products | Nationwide and Worldwide to:Argentina,Australia,Austria,Belgium,Brazil,Canada,Chile,China,Columbia,Czechia,Denmark,Egypt,Finland,France,Germany,Greece,Hong Kong,Hungary ,Iran,Ireland,IsraÂżÂżl,Italy,Japan,Lithuania,Luxembourg,Mexico,Middle East,Netherlands,Norway,Panama,Poland,Portugal,Russia,Serbia,Slovenia, South Africa,Spain,Sweden,Switzerland,Taiwan,Turkey,United Kingdom | Terminated | May 03, 2012 | 1500 | 14295 | Not_Computer
Z-1649-2008 | 47800 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | BCI Non Invasive Blood Pressure Monitor |  |  | April 02, 2008 | 2008 | Smiths Medical PM, Inc. | Due to a component that may be installed backwards, the device may turn off or may immediately reset. | Employee error | Consignees were contacted on 4/2/08 with a Safety Action Bulletin, 08-SAB02, and Technical Worksheet 08-TW01 dated April 1, 2008, which included information regarding the potential safety issue and the proper inspection method for the monitor(s).If you have not received the Safety Action Bulletin or the Technical Worksheet contact Smiths Medical PM, Inc. at Telephone: (262) 542-3100 or Toll Free: (USA): (800) 558-2345 or email info.pm@smiths-medical.com. |  | 195 | 195 | USA: IL, FL, TX, IA, CA, ID, MD, AZ, MN, PA, CO, IN, KS, WA, TN, AL, WI, NM, LA, and PR.OUSA: Argentina, Canada, Ecuador, Greece, Hong Kong, Italy, Pakistan, Peru, Russia, saudi Arabia, Slovakia, Thailand, United Kingdom | Terminated | December 17, 2011 | 1354 | 195 | Hardware | Device Operation | Instructions | Safety Notice/Insructions
Z-1702-2008 | 47908 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Opthalmic microsurgical system-AMO WHITESTAR SIGNATURE Phacoemulsification System, Product Number NGP680300. The device is a modular ophthalmic microsurgical system that facilitates both anterior segment (i.e., cataract) and posterior segment (i.e., vitreoretinal) ophthalmic surgery. |  |  | April 02, 2008 | 2008 | Advanced Medical Optics, Inc. | AMO initiated this field correction event after becoming aware of a trend in complaints associated with AMO Vitrectomy Cutter used in conjunction with the AMO WHITESTAR SIGNATURE Phacoemulsification System. The physician may be unable to make a cut during cataract surgery. | Device Design | On April 2, 2008, AMO began distribution of an Urgent Safety Notice to all customers in the United States via overnight carrier (Federal Express). The U.S safety notice was used as a template for all global AMO communications regarding the Urgent Safety Notice , many of which required translation into local language prior to distribution. Included in this correspondence is a listing of the serial numbers of the affected products.The Notice informed customers of the problem and gave the customers modifications to the Vitrectomy Cutter Priming Instructions. In addition, the notice informed the customers that an AMO Field Service Engineer will be scheduling a site visit to modify a pressure setting on the AMO Whitestar Signature system. Contact the AMO Phaco Technical support at 1-877-AMO-4LIFE for assistance. AMO will provide updates to the FDA, including all information regarding the recall effectiveness as part of the recall progress reports. |  | 191 | 191 units | Nationwide and to: AL, AR, CA, CO, FL, HI, IL, IN, MD, NC, ND, NJ, NY, OH, OR, PA, PR, SC, TX &amp; WA and worldwide to: Azerbaijan, Belgium, Germany, Denmark, Egypt, Spain, France, Great Britain, Croatia, Hungary, Israel, Italy, Jordan, Kuwait, Lebanon, Liechtenstein, Netherlands, Norway, Poland, Portugal, Saudia Arabia, Sweden, Turkey, South Africa, Australia, India, New Zealand and Singapore | Terminated | January 21, 2010 | 659 | 191 | Not_Computer
Z-2160-2008 | 48602 | No | IZJ | Cine or spot fluorographic x-ray camera. | Radiology | Radiology | 510(k) | General X-Ray System types with Digital Spot Imaging (DSI) software |  |  | May 23, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | X-ray images may be stored in the wrong patient file or corrupted, images from two different patients may be combined in one patient's folder or may be combined on one image (corrupted). | Software design | On 5/20/08 the firm issued an URGENT - DEVICE CORRECTION letter to the consignees advising of the problem, requesting consignees to check the stored or printed images to ensure that it represents the data and images captured from the examination. The firm will issue a Field Change Order to their service representatives providing new software, when software is available. This should be issued by the end of June. |  | 41 | 41 units | Nationwide Distribution. | Terminated | September 28, 2010 | 858 | 41 | Software | Output/Calculation | Update software | Software Update
Z-2162-2008 | 48604 | No | FZT | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ENDOPATHÂżETS Endoscopic Linear Cutter, TSW35. |  |  | April 24, 2008 | 2008 | Ethicon Endo-Surgery Inc | The firm discovered a component production issue which may cause the cartridge to deploy an incomplete staple line. | Other | Firm issued an Urgent: Device Recall notification instructing the user to not use products from these lot numbers. Examine their inventory immediately to determine if they have affected product on hand and remove them from inventory. Complete the Business Reply Form and fax it back. To return affected product, copy the BRF, place in box with your product, and affix the pre-printed UPS shipping label. |  | 10734 | 10,734 pieces | Worldwide Distribution --- including USA and countries of ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CAYMAN ISLANDS, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, EGYPT, EL SALVADOR, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEBANON, MEXICO, PAKISTAN, POLAND, PORTUGAL, PUERTO RICO, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM &amp; VENEZULA. | Terminated | April 03, 2012 | 1440 | 42434 | Not_Computer
Z-2159-2008 | 48601 | formalin, neutral buffered | IFP | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | 10% Buffered Formalin Phosphate; Certified |  |  | May 01, 2008 | 2008 | ThermoFisher Scientific Co., LLC | Assay tests on retention sample of 10% Buffered Formalin Phosphate showed it to be 3.417%, which is below the specification of 4-5%. | Nonconforming Material/Component | Urgent: Product Recall notification letters were sent by US mail on May 1, 2008. The letters asked that those in receipt of the product check their stock for this product, destroy any of the material that remains from the referenced lot, and then contact their local customer service center for a replacement. The letter also states that the appropriate corrective measures have already been instituted to prevent a recurrence. |  | 1640 | 1052 x 4 liters, 564 x 20 liters | Worldwide Distribution --- including USA and Canada. | Terminated | August 13, 2012 | 1565 | 1640 | Not_Computer
Z-2167-2008 | 48611 | shield, eye, ophthalmic (including sunlamp protective eyewear and post-mydriatic eyewear) | HOY | Ophthalmic eye shield. | Ophthalmic | Ophthalmic | 510(K) Exempt | Radiation Protective Eyewear (only involving multifocals. |  |  | June 05, 2008 | 2008 | Phillips Safety Products | Radiation Protective Eyewear does not meet the radiation protection levels stated on the labeling. | Process control | Customers were all notified by phone, followed by an Urgent Recall letter dated June 3, 2008. The letter asked that customers examine their stock immediately to determine if they had any of these radiation protective bifocals on hand, discontinue use or distribution of the lot, and set the products aside for further instructions. Also, if customers distributed these items, they needed to contact their accounts and advise them of the recall and have them return their outstanding recall stock to them. The firm will send replacement products in the weeks to follow. |  | 845 | 845 pairs of eyewear | Worldwide Distribution --- including USA and countries of Spain, Norway, Great Britain and the Netherlands. | Terminated | December 24, 2008 | 202 | 845 | Not_Computer
Z-1716-2008 | 47468 | No | BYJ | Portable liquid oxygen unit. | Anesthesiology | Anesthesiology | 510(k) | Generation 4 electronic liquid gauges used on the Caire Liberator, Stroller, Sprint, and Hi-Flow Portable oxygen unit. |  |  | October 23, 2007 | 2007 | Caire, Inc. | Moisture ingress into the electronic liquid level indicator may cause inaccurate liquid oxygen level readings. The oxygen could run out without the patient knowledge. | Device Design | A Technical Service Bulletin was sent to consignees on/about October 23, 2007 to notify customers of the potential issue. Technical support and information for ordering replacement Generation 4 liquid gauges were offered.For technical concerns or if you have not received the notification from the firm, contact U.S. Technical Service at 1-800-482-2473 or European Technical Service at +44 (0) 797 1260077. |  | 3470 | 3470 | U.S. Nationwide and the following foreign countries: Argentina, Austria, Brazil, Canada, China, Chile, Central America/Caribbean, Colombia, Costa Rica, Europe, Austria, Belgium, Denmark, France, Germany, Greece, Hungary, Irish Republic, Italy, Netherlands, Norway, Portugal, Spain, Sweden, Switzerland, United Kingdom, Mexico, South America, Peru and Uruguay. | Terminated | December 17, 2009 | 786 | 3470 | Not_Computer
Z-1868-2008 | 48427 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Encore Modular Femoral Neck, 12 degree/35 mm; Ref #410-35-112; distributed by Encore Medical, Austin, TX 78758. |  |  | May 08, 2008 | 2008 | Encore Medical, Lp | Product not manufactured to specifications; modular femoral necks found to contain 24 scallops, instead of the standard 12. | Nonconforming Material/Component | Firm notified sales agents of recall on 05/08/08 via an Urgent: Medical Device Recall letter. Sales agents asked to immediately examine their inventory and quarantine product subject to recall. Agents are asked to notify their customers if they had further distributed the product. All product must be returned to Encore along with the response form. |  | 20 | 20 units | Worldwide Distribution --- USA including states of IN, KS, GA, WA, SD, and countries of Greece and Australia. | Terminated | September 08, 2009 | 488 | 20 | Not_Computer
Z-2154-2008 | 48467 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CorometricsÂż MODEL 250 Series Maternal/ Fetal Monitor, Models 256 and 259, part number 2024489-002. |  |  | May 05, 2008 | 2008 | Datex Ohmeda, Inc. dba GE Healthcare | Communication connector within monitor is defective and may cause potential loss of telemetry. | Component design/selection | The recalling firm notified healthcare professionals by an Urgent Medical Device Correction letter dated 07/01/2008 to Healthcare Administrators, Risk Managers, Nurse Managers and Directors of Clinical Engineering. The letter advised of a defect in a communication board that may cause a potential loss of telemetry monitoring. The firm recommended that the device should not be used with the firm's 340 Series telemetry system pending replacement of the defective board by service engineers. The Healthcare Professionals are asked to contact the GE Healthcare Technical Support center at 1-800-558-7044. |  | 133 | 133 units | Nationwide Distribution. | Terminated | August 07, 2012 | 1555 | 133 | I/O | Device Operation | Replace component | Remove or Replace
Z-2165-2008 | 48607 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Scimed Maverick 2 Monorail" PTCA Catheter, 3.0 mm 15 mm |  |  | May 01, 2007 | 2007 | Boston Scientific Corporation | The units within this lot may not meet thickness specifications surrounding the wire exit port wall, which may result in an air embolization if the Port Septum Wall burts. | Nonconforming Material/Component | Consignees were sent a "Boston Scientific Urgent Medical Device Recall or Correction-Immediate Action Required" letter dated May 11, 2007. The letter addressed to Risk mangers, described the issue and product. The letter requested the immediate discontinued use of and segregation of recalled product, A Reply Verification Tracking Form was requested to be completed and returned. |  | 24 | 24 Units | International Distribution --- including countries of Greece, Germany, Great Britain, Spain, Sweden, Belgium, and Algeria. | Terminated | September 10, 2008 | 498 | 24 | Not_Computer
Z-2170-2008 | 48615 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ev3 Primus GPS Biliary Stent System |  |  | April 22, 2008 | 2008 | Ev3, Inc | A Primus GPS Biliary Stent System was mislabeled. The mislabeling resulted in a 8 mm diameter, 27 mm length stent being labeled as a 8 mm x 37 mm stent. | Employee error | The single affected consignee was visited by an ev3 representative to verify the use of the device or exchange it for an unaffected device. Information recorded on an ev3 Device Reconciliation Form. |  | 1 | 1 Unit | Nationwide Distribution --- state of NJ | Terminated | December 10, 2011 | 1327 | 1 | Not_Computer
Z-2166-2008 | 48608 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Scimed Maverick 2 Monorail 12mm Balloon Catheter, 2.5 mm x 12 mm |  |  | November 01, 2005 | 2005 | Boston Scientific Corporation | Two lots may not be correctly labeled. Specifically , a 2.5 mm x 12 mm unit was incorrectly labeled as a 3.5 x 20 mm unit on the box and tyvek label. | Packaging process control | Consignees were sent a "Boston Scientific Urgent Medical Device Recall or Correction-Immediate Action Required"letter dated November 7, 2005. The letter addressed to Risk mangers, described the issue and product. The letter stated that immediate discontinued use of and segregation of the recalled product was required. A Reply Verification Tracking Form was requested to be completed and returned. |  | 67 | 67 Units | International Distribution --- including countries of Czech Republic, France, Germany, Great Britain, Italy, Spain, Sweden, India, and Iran. | Terminated | September 10, 2008 | 1044 | 67 | Not_Computer
Z-2169-2008 | 48614 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Ev3 Protege GPS Biliary Stent Systems |  |  | April 22, 2008 | 2008 | Ev3, Inc | A Protege PS Biliary stent system was mislabeled. The mislabeling resulted in a 9 mm diameter, 60 mm length stent being labeled as a 6 mm x 60 mm stent. | Employee error | The single affected consignee was visited by an ev3 representative to verify the use of the device or exchange it for an unaffected device. Information recorded on an ev3 Device Reconciliation Form. |  | 1 | 1 | Nationwide Distribution --- state of LA. | Terminated | December 10, 2011 | 1327 | 1 | Not_Computer
Z-1776-2008 | 47807 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Q-Logic Controller Software used on the following power wheelchairs: Q6000, Q6000XL, Q6000Z, R4000, Q600, Q600XL, Q610, Q1107, Q600E, and Dynamo 1107 |  |  | April 11, 2008 | 2008 | Pride Mobility Products Corp | Software - Watch dog timer feature was disabled. | Software design | The recalling firm issued an Urgent Device Correction letter on 4/11/08. This letter instructs the dealers to contact the end users so that the units can be upgraded to enable the watch dog feature. |  | 4129 | 4129 | Class II Recall - Worldwide Distribution --- including USA and Canada. | Terminated | November 13, 2009 | 581 | 4129 | Software |  | Software update | Software Update
Z-1774-2008 | 47795 | [6 empty] | Hall Easy-FitÂż Prosthetic Heart Valve - Aortic A7700. The device uses a single pyrolytic carbon disc occluder pivoting within a titanium housing. The housing is machined into final configuration from a single piece of solid titanium and has no welds or introduced bends. It is an open-membered design in which all members project centrally without distal attachment. A rotatable fabric sewing ring is affixed to the valve. The valve is provided on a holder (collect) to facilitate insertion and to provide protection during shipping and handling. The valve is sterile and non-pyrogenic if seal and package are not opened, damaged, or broken. |  |  | March 25, 2008 | 2008 | Medtronic Heart Valve Division | Medtronic has identified an issue with 685 Hall Easy-FitÂż valves Models A7700 and M7700 in which the holder of the valve was incorrectly secured to the valve. This could potentially make it difficult to easily disengage the holder from the valve. | Process control | A Medtronic "Urgent Communication" letter dated March 24, 2008 was sent to consignees on 3/25/08. The letter describes the product and problem. The letter requests that they return the impacted devices and to return the reconciliation form "Medtronic Recall Certificate (FCA0807)" confirming whether they have the valves in inventory or if they have been implanted. |  | 69 | 69 (USA), 117 (EUROPE), 5 (CANADA), 179 (HONG KONG) | Class II Recall - Worldwide Distribution --- including USA states of TX, WY, CA, MI, WI, MO, IL, NY, TN, AR, LA, KY, WA, WV, and countries of Canada, Hong Kong, Australia, Albania, Belgium, Croatia, Germany, India, Italy, Netherlands, Poland, Romania, Russian Federation, South Africa, and Turkey. | Terminated | May 15, 2012 | 1512 | 99 | Not_Computer
Z-1777-2008 | 47819 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom Century+ Bed Model P1400 HBSW siderail upgrade kits; accessory model numbers P9964-11, P9964-12, P9964-13, P-9964-14, P9964-15, P9964-16, P9964-41 and/or P9964-42. |  |  | April 28, 2008 | 2008 | Hill-Rom, Inc. | A risk of entrapment will exist between the head and foot rails if installation instructions are followed. | Other | Hill-Rom issued an Urgent Medical Device Correction letter on 4/28/2008 notifying consignees that they will be visited by a qualified Hill-Rom service technician or designee and provide a revised copy of the instruction sheet. The technician will either assist the facility to ensure proper installation or verify by signature that the facility understands the issue and will verify their beds had the upgrade kits installed in such a way that an entrapment gap does not exist. |  | 69 | 69 | Class II Recall - Nationwide Distribution --- including states of Florida, Illinois, New Mexico, Texas and Virginia. | Terminated | September 22, 2009 | 512 | 69 | Not_Computer
Z-1778-2008 | 47829 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Xcelerate Patella System Reamer Shaft Assembly; |  |  | April 15, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | The Xcelerate Patella Reamer Shaft Assembly and the Reamer Adapter do not mate properly. This could result in the removal of excessive bone from the patella leading to a thinner patella; a patella that does not have enough bone for adequate engagement with pegs or patella fracture. | Process control | An IMPORTANT MARKET WITHDRAWAL letter was issued on April 19, 2005 and sent Federal Express to those consignees who received the device. The letter instructed them to examine their inventory to identify instruments, reconcile these instruments on an attached accountability form; fax a copy of this form within 2 days of receipt of the letter; retrieve and return affected product; contact Stryker Customer Service to re-order. |  | 24 | 24 total | Class II Recall - Worldwide Distribution --- including USA and Canada, United Kingdom. | Terminated | September 24, 2008 | 1258 | 24 | Not_Computer
Z-2174-2008 | 48616 | No | LKK |  |  | General Hospital | PMA | Model 8731SC Intrathecal Catheter |  |  | April 14, 2008 | 2008 | Medtronic Neuromodulation | The SC catheters cannot completely engage with the Model 700-04M portal connector. The result of this condition is a possible leakage or disconnect of the SC catheter from the portal during the continuous therapy trial period. | Device Design | Medtronic informed all Belgium implanting and managing physicians of the SynchroMed II infusion system. Each customer was mailed a Dear Healthcare Professional letter (Medtronic Urgent: Field Safety Notice) on 4/3/08 providing important safety information and patient management recommendations regarding the device incompatibility. The firm recommends that healthcare professionals do not use SC catheters if a connection with an implantable access port is required. Also, if a continuous infusion trial is required with use of an implantable access port, they should use a model 8711 intrathecal catheter. The firm states that for patients who are currently undergoing a continuous trial with an implantable port, replacing the catheter connector before connecting to the SynchroMedd II using the Medtronic model 8578 revision kit is recommended. Medtronic also recommends discussing the risks of connection failure with patients. Local field reps will personally follow-up with affected customers to confirm and document notification. |  | 232 | 232 total from all models | International Distribution --- country of Belgium. | Terminated | September 10, 2008 | 149 | 232 | Not_Computer
Z-1309-2008 | 46834 | No | FTC | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HOUVA III Phototherapy System with DermaSense Technology comes on one skid within 10 cardboard boxes. |  |  | February 19, 2008 | 2008 | National Biological Corp | Software allows operator to override "low line voltage" error warning and store light intensity value. | Software design | National Biological sent an "Important safety reminders for your Houva DermaSense phototherapy device" letters, dated 2/11/08 to users. It states that the user should collect a "Profile" of the light energy and record it on a daily/weekly flow chart. The chart should be in plain view so each operator can clearly determine the units output. If an error message appears in the display panel, treatment should not be continued unless the phototherapy booth operator was able to identify and correct the cause of the error. Make sure both breakers are in the "ON" position before starting up the unit each morning and before giving treatment. |  | 48 | 48 Units | Worldwide distribution --- USA including states of: CA, DE, FL, GA, ID, IL, IN, MA, MI, MN, NC, NH, NY, OH, OR, PA, TX, VA, WA and WI, and country of Canada. | Terminated | September 20, 2011 | 1309 | 48 | Software |  | Instructions | Safety Notice/Insructions
Z-1650-2008 | 47885 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Ambulift, standard basic model manually powered patient lift. The Ambulift is used to move and lift patients from their beds, transport them to the bathroom, and lift them in and out of the bath. |  |  | May 05, 2008 | 2008 | Arjo, Inc. | The armrests of the lift chair may push up, allowing the patient to slip down and out of the seat. The resultant fall could result in serious or fatal patient injuries. | Device Design | Arjo sent Customer Field Correction Notification letters dated 5/2/08 on the Arjo letterhead to the end user accounts who received the affected Ambulift Patient Lifts, to the attention of the Administrator/Risk Manager, via UPS 2nd day delivery on 5/5/08. The letters informed the consignees of the incident with the patient slipping from the seat of the lift and sustaining fatal injuries and advised them to always ensure that a clinical assessment to verify that the resident has sufficient upper body strength to maintain a steady position on the chair, and that the resident has sufficient mental capacity to adhere to the caregiver's instructions, is carried out by a qualified nurse or therapist before lifting residents with the Ambulift. To further enhance safety, a seat belt is being made available, which must be fitted to the chair, enabling it to be used when required by clinical assessment. Instructions on how to apply, use, and maintain the seat belt will be provided with the belts. The belts will be provided at no cost. The accounts were requested to review, complete and return the enclosed customer response form, indicating the number seat belts needed. Any questions were directed to Arjo's Quality Department at 1-800-323-1245, ext. 6140 or 6165. |  | 11 | 11 lifts | Worldwide Distribution-USA includes the states of AL, AZ, DE, FL, KY, MO, NY, OK, and VA | Terminated | February 03, 2010 | 639 | 14 | Not_Computer
Z-2153-2008 | 48466 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Acu-Loc Distal Radius Plate, Standard, Left. Product labeled in part, " Acu-Loc Distal Radius Plate, Standard, Left... PL-DR50L W48003... Manufactured or Distributed by: Acumed LLC Hillsboro, OR ". |  |  | June 29, 2007 | 2007 | Acumed LLC | Plates are incorrectly laser marked. The PL-DR50L plates are incorrectly laser marked with PL-DR70L. | Process control | The firm issued an Urgent Notification: Product Removal letter on 6-29-2007. The notification was reportedly sent by either FAX or e-mail. The notification informed consignees to remove the plates from their inventory and return them to Acumed. The notification did not include any instructions for a sub-recall. On 6/13/08, the firm sent a second Urgent Notification: Product Removal letter. This letter describes the product as PL-DR50L - Lot Code - W48003, explains the problem and asks customers to return product to Acumed. |  | 51 | 51 pieces | Worldwide Distribution --- including USA and countries of South Africa, Italy, Finland, and Brazil. | Terminated | February 03, 2010 | 950 | 51 | Not_Computer
Z-2155-2008 | 48468 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.5mm X 32.5mm Cortical Screw used with the Polarus Humeral Fixation System,. Product labeled in part, "Polarus Humeral Fixation System&amp; SIZE: 3.5mm X 32.5mm Cortical Screw... REF HCO3325-S... Acumed 5885 NW Cornelius Pass Road Hillsboro, OR 97124-9432 ". |  |  | December 27, 2005 | 2005 | Acumed LLC | Package of 3.5mm X 32.5mm cortical screw may contain 3.5mm X 22.5mm cortical screw. | Process change control | The firm issued a Product Removal notification on 12-27-2005 to their customers. The notification was sent by either FAX or by e-mail. The Product Removal Notification informed their customers of the issue, asked them to remove the screws from inventory and return the screws to Acumed. A reminder notification, dated 1/20/08 was issued to customers identifying the product and the problem and requesting the product be returned. |  | 112 | 112 pieces | Worldwide Distribution --- USA including states of CA, KY, FL, TX, OH, OR, NY, ID NJ, NE, LA, WI, UT, MO, MS, GA, AZ and countries of United Kingdom, Italy, Puerto Rico, Turkey, Spain, Japan, South Africa, and Belgium. | Terminated | February 03, 2010 | 1499 | 112 | Not_Computer
Z-2201-2008 | 49030 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | Pacific Consolidated Industries Mobile Oxygen Storage Tank, MOST 20-2B, Part Number: 793080-002 |  |  | September 08, 2007 | 2007 | Pacific Consolidated Industries LLC | The recall was initiated after the firm received two customer complaints of ruptured bourdon tubes that had resulted in bellowed out face gauges. With one unit the operator was able to shut off the oxygen flow, however one rupture resulted in a fire. No injuries were reported. | Vendor change control | PCI sent a Recall Notification letter to the Spanish Air Force September 8, 2007. The Recall Notification letter was mailed to the US Defense Supply Center on July 24, 2008 after a recall agreement was signed on July 8, 2008 between Pacific Consolidated Industries and the Defense Supply Center. The recall notification to Spanish Air Force informed the consignee of the problem and asked the consignee to remove the affected units from service until a Technical Service Representative could be sent to Spain to inspect the affected units and perform the corrective action. The recall notification to Defense Supply Center instructed the consignee to keep the affected units in quarantine (units placed in quarantine by US Air Force August 3, 2007) until they have been refurbished under recall agreement that was signed on July 8, 2008. |  | 1032 | 1032 units | Worldwide to US Air Force and also to Spain | Terminated | November 01, 2012 | 1881 | 1032 | Not_Computer
Z-1617-2008 | 47388 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific SteeroCath-T Ablation Catheter. This catheter is indicated for general indications, treatment of supraventricular tachycardia (svt) and ventricular tachycardia (vt). This product in not indicated for use as an ablation catheters indicated for treatment of atrial flutter or for atrial fibrillation. |  |  | February 14, 2008 | 2008 | Boston Scientific Corp | Package sterile barrier may be breached, compromising sterility. | Nonconforming Material/Component | Recall initiated on Feb. 14, 2008, via Urgent Medical Device Recall notification being sent to consignees, along with a reply verification tracking form.If you have questions or need a copy of the recall notification, contact Deborah J. Springer at 1-408-935-4944 or springed@bsci.com. |  | 827 | 827 units total | Worldwide Distribution --- including USA and countries of Saudi Arabia, Spain, Italy, Israel, Denmark, Canada and Australia. | Terminated | January 08, 2011 | 1059 | 827 | Not_Computer
Z-1705-2008 | 47774 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OEC 9800 Image Intensified fluoroscopic x-ray system (includes OEC 9800, OEC FluoroTrak 9800 Plus, OEC 9800 Plus, and OEC 9800MD Motorized C-arm System), GE Healthcare, Surgery, Salt Lake City, UT. |  |  | April 15, 2008 | 2008 | OEC Medical Systems, Inc | Use of existing four-pedal footswitch on a different machine may cause various operational errors. | Device Design | Consignees were notified by letter on 03/21/2007 and advised that GE would be replacing the 4-pedal switches at some time in the future. Replacement began on 04/15/2008. Contact GE Healthcare at 1-800-874-7378 for assistance. |  | 78 | 78 units | Nationwide, including one VA facility in FL. No military or foreign distribution. | Terminated | December 21, 2011 | 1345 | 103 | Not_Computer
Z-1642-2008 | 47487 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | AD 7 Patient table supplied with Allura Xper FD10 and Allura Xper FD10/10. Product is non-tilting but can be purchased with a tilt option. |  |  | March 14, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Potential for radiologic patient table to become immobile and unable to move again due to force sensor sensitivity to electromagnetic radiation. | Device Design | On 03/22/2004, firm sent letters via certified mail to notify customers of issue. Letter advises customers of potential risk for patients and users, and states to not use system for critical examinations or treatments if table movements suddenly stops or makes a freeing movement. The letter indicates a representative will contact customer for installation of a Hardware filter, free of charge, to prevent Electromagnetic interference (EMI) from affecting the force sensor. Please contact Philips at 1-800-722-9377 #5, #4, #1. |  | 39 | 39 units | Product is distributed to hospitals and medical centers nationwide to the following states: ND, FL, KY, MN, VA, ME, PA, NJ, CT, TX, CA, OH, TN, OR, GA, NY, SC, MO, WA, IL. | Terminated | July 18, 2012 | 1587 | 39 | Not_Computer
Z-1715-2008 | 47384 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | In vitro Diagnostic-Clinitest hCG Cassette Pregnancy Test, Radio immunoassay for Human Chorionic Gonadotropin (hCG) test system Product Number: 1760 The Clinitest hCG Pregnancy Test is for in vitro diagnostic use as a qualitative method in the rapid detection of human chorionic gonadotropin (hCG) in urine specimens. |  |  | April 04, 2008 | 2008 | Siemens Medical Solutions Diagnostics | False negative hCG results due to decreased sensitivity | Component design/selection | Siemens notified Distributors ( Distributor Bulletin Urgent Field Safety Notice) and end-users (Customer Bulletin Urgent Field Safety Notice) Clinitest hCG Cassette Recall by letter dated 4/3/08. Users are requested to, discontinue use, discard the product, and immediately contact Siemens technical support representative or distributor for replacement cassettes from a different lot.If you have any question or have not received the recall notification, contact your local Siemens Distributor Relations Representative or Siemens Technical Support at 877-229-3711 option 13 then 5. |  | 8180 | 8180 | Nationwide | Terminated | March 15, 2012 | 1441 | 8180 | Not_Computer
Z-1524-2008 | 47517 | No | MIE |  |  | Cardiovascular | Contact ODE | CryoValve, Pulmonary Valve &amp; Conduit |  |  | March 26, 2008 | 2008 | CryoLife, Inc. | Tissues taken from a donor whose blood culture was found positive for Serratia marcescens, was distributed. | Material/Component Contamination | Hospitals were notified by letter on 3/24/2008. They were advised not to implanted the tissue and if implanted, they do not recommend explant unless medically necessary. A self addressed, stamped post card was included to confirm receipt of the notification. Serial number 8895391 was reported implanted. Serial number 8895131 was returned to CryoLife for destruction. |  | 1 | 1 tissue | Hospitals in DE, TN | Terminated | January 12, 2010 | 657 | 1 | Not_Computer
Z-1593-2008 | 47424 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PathFinderÂż Spinal Fixation System T-Handle Driver; Part 2151-7 (13" shaft); Non-sterile, Reusable Instrument; manufactured by Abbott Spine, Austin, TX 78727. |  |  | March 24, 2008 | 2008 | Abbott Spine | Reports received regarding disassembly of T-Handle Driver components during surgery dropping the closure top into the surgical site. | Error in labeling | Firm sent notification letters to sales agents via letter on 03/24/08. Enclosed with the notification letter, the firm provided updated product inserts for the InCompassÂż and PathFinderÂż Spinal Fixation System (Cat. #2199-0002-LB Rev. H and #1199-0002-LB Rev. H, respectively), as well as the "Disassembly and Assembly Instructions for PathFinderÂż &amp; InCompassÂż T-Handle Drivers" (Cat #1199-0040-MKC Rev. A).Sales Agents instructed to distribute updated inserts and instructions to all locations that handle the drivers for cleaning and to destroy all copies of old inserts.If you have questions or have not received the recall notice from the firm, contact Kirstin Widding at 512-533-1852. |  | 224 | 224 units. | Nationwide Distribution --- including states of AR, AZ, CA, CT, CO, FL, IL, LA, MD, MI, MO, MN, MT, NC, NM, NJ, NY, OH, OK, OR, PA, SD, TX, VA and WI. | Terminated | December 02, 2009 | 618 | 502 | Not_Computer
Z-1475-2008 | 47352 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Clinitest hCG Cassette Pregnancy Test, Radio immunoassay for Human Chorionic Gonadotropin (hCG) test system |  |  | October 17, 2007 | 2007 | Siemens Medical Solutions Diagnostics | False positive hCG results | Employee error | Siemens issued a Service Bulletin to Siemens Healthcare Diagnostics Branches via email on 10/17/007, and a customer bulletin to outside the US . Users are requested to discontinue use and discard the product. A Completion Notification was requested to be completed and returned by the faculties. |  | 3065 | 3065 kits | NationwideCanada, France, Korea, Netherlands, UK | Terminated | March 15, 2012 | 1611 | 3065 | Not_Computer
Z-1749-2008 | 47536 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UBiTÂż-IR300 Infrared Spectrophotometer, Software Versions 7.0, 7.01 and 7.11. |  |  | January 18, 2008 | 2008 | Otsuka America Pharmaceutical, Inc. | Medical device power supply may overheat and cause smoke to be emitted from the device. | Nonconforming Material/Component | The initial distributor in the US notified consignees 03/14/08 and 03/26/08 by letter. Consignees were provided two different letters depending on the hours of use of the spectrophotometer. The notifications advised of the firm's preventive service and upgrade program for the instrument's power supply, in which consignees would receive a temporary replacement instrument while the service is completed. Consignees with spectrophotometers having &lt;15000 light source hours were notified similarly, and were provided filters and instructions for cleaning and replacing the filters to protect the power supply from unnecessary accumulation of debris. Contact firm at 1-888-MERETEK for assistance. |  | 517 | 517 units | The affected medical devices were distributed to medical facilities and physician offices nationwide and to distributors in Mexico. | Terminated | February 18, 2011 | 1127 | 517 | Battery | Physical Safety Hazards | Replace/Instructions | Remove or Replace
Z-1646-2008 | 47535 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | i-CAT 17-19 Cone Beam 3D Dental Imaging System. This product is also sold under a private label as KaVo 3D eXam |  |  | March 19, 2008 | 2008 | Imaging Sciences Intl Inc | Incorrect scatter information in the manual. | Error in labeling | Firm issued Urgent Medical Device Recall letters dated 3/19/08 to their customers. The distributors were provided background information and specific instruction on the steps to perform to correct the problem.If you have questions or have not received this letter, contact Francine McCoonse at 215-997-5666 ext 1340. |  | 1203 | 1203 units | The products were shipped to distributors and dental offices nationwide. The products were also shipped to Canada, Austria, Belgium, Czech Republic, England, Finland, France, Germany, Hungary, Ireland, Italy, Latvia, Lithuania, Netherlands, Norway, Portugal, Lavenia, Spain, Sweden, Switzerland, Turkey, United Kingdom, Australia, New Zealand, Chile, Israel, South Korea, UAE, India, Argentina, Brazil, China, Columbia, Hong Kong, Jamaica, Korea, Lebanon, Malaysia, Morocco, Russia, Scotland, Taiwan, Thailand, and Singapore. | Terminated | May 05, 2009 | 412 | 1203 | Not_Computer
Z-1748-2008 | 47527 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Zimmer Spine Cyclone anterior cervical plate, 4 Level, 88MM (Ti-6Al-4V); REF 600-04-188. |  |  | March 26, 2008 | 2008 | Zimmer Inc. | The screw locking cap may fracture when the surgeon rotates it into the locked position. | Device Design | Consignees were notified via voicemail on 3/26/08 and a letter was sent to all distributors on 3/27/08 informing them that the plate should no longer be used and that all lots are being recalled. A letter dated 6/18/08 was issued to implant physicians recommending patient monitoring. Contact Zimmer Spine at 1-952-857-5636 for assistance. |  |  |  | Arizona, California, Florida, Georgia, Louisiana, Michigan, Minnesota, Montana, New Hampshire, North Carolina, Oklahoma, Pennsylvania, Tennessee, Texas, Utah, Washington and Wisconsin. | Terminated | November 12, 2009 | 596 |  | Not_Computer
Z-1591-2008 | 47423 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Deltec CozmoÂżÂż Insulin Pump, model 21-1800. |  |  | February 22, 2008 | 2008 | Smiths Medical MD, Inc. | Smith Medical has received reports of Deltec CozmoÂżÂż Insulin Pumps powering down without an alarm. Investigation concluded that if the battery cap is not fully tightened, the Pump may power down and a brief "chirp" will sound. This may occur if the user has not sufficiently tightened the battery cap or if the battery cap is damaged. If a patient is asleep or does not regularly check their blood gl | Device Design | Consignees and the USA demonstration pump physicians were mailed an "Urgent Medical Device Correction" letter on 2/22/08. The letter stated Details on Affected devices, Description of the problem, Advice on action to be taken by the user, Additional Information and New Warning. These consignees will be required to return the Correction Confirmation Form. The 17 countries with international mailing addresses were sent the Urgent Medical Device Correction letter (Urgent Field Safety Notice) via the Smiths Medical International Regulatory department based in Watford UK. The Smiths Medical Distributors and Affiliates will be required to mail the Safety Notice to their direct users. The USA distributors were mailed the Urgent Medical Device correction letter with a cover letter. These consignees will be required to add the letter to in-house inventory. The USA prescribing physician and the the USA Managed Care Accounts (Insurance Companies) were mailed the Urgent Medical Device Correction letter with a cover letter. These consignees will not be required to respond and no follow-up will be done by Smiths Medical.Contact Mike Herbert at 651-628-7049 if you have any questions or did not receive the recall information from the firm. |  |  |  | Worldwide Distribution --- USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, and countries of AU, CA, CL, CN, DE, EG, FR, GB, GR, IL, KW, MX, NL, NZ, ON, PL, SA, SE, TR, UK. | Terminated | January 19, 2011 | 1062 |  | Not_Computer
Z-1766-2008 | 47529 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON Antares 5.0 ultrasound system at SW revision 200.0.054 or 200.0.059A, with one of the following options installed: 3-Scape Imaging, Factory, PN 05937011 or 3-Scape Real Time Imaging, UPG ACU, PN 10033688. Material number 10032746, 10037592, 10032746, 10037591 or 10038202. |  |  | March 21, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | Software issue may result in distorted images and inaccurate measurements, which could lead to a misdiagnosis. | Software design | The firm initiated a field correction on 03/21/2008. Site visits to correct the problem and install the revised software version 200.0.063A are planned for all consignees. No other notification was made. The firm can be contacted at 650-694-5993. |  | 488 | 488 active units total | Product was distributed worldwide to 539 consignees, in the USA to MD, SC, AZ, FL, IL, WI, NJ, NC, CA, OH, VA TX, PA MO, NH, IA, NE, MN, NY, GA, MI, WA, HI, MA, SC, WV, VA, ND, IN, WY, UT, MS, TN, PR, SD, KY, OK, OR, and worldwide to Sweden, Saudi Arabia, Austria, New Zealand, Italy, Switzerland, Mexico, Germany, Australia, Canada, Great Britain, France, China, South Korea, Singapore, Denmark, Brazil, Poland, Norway, Faroe Islands, South Africa, Russian Federation, Romania, Trinidad and Tobago, Greece, India, Spain, Thailand, Ukraine, Nicaragua, Japan, India, Czech Republic, Egypt, Taiwan, Turkey, Malaysia, Qatar, Bosnia an dHerzogovina, Vietnam, Brunei Darussalam, Italy, Kuwait, Turkey, Portugal | Terminated | September 30, 2009 | 558 | 488 | Software | Output/Calculation | Software update | Software Update
Z-1823-2008 | 47834 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Headset Model Number RTC-2K for use with the Serene Sound Digital MRI Compatible High Fidelity Stereo Sound System. |  |  | May 14, 2007 | 2007 | Resonance Technology Inc | Potential safety concerns - Use of this product with a MRI may result in the headset's cord overheating and patients experiencing burns to the skin. | Device Design | The strategy was notifying customers by written communication. The letter advises Resonance Technology, Inc. is voluntarily recalling the MRI Headset, Model No. RTC-2K and RTC- 551-070 because of potential safety concerns. Use of this product with a MRI may result in the headset's cord overheating and patients experiencing burns to skin. The letter also advises that Resonance Technology Inc. has designed and developed the significantly improved MRI Headset, Model No. RTC-661-060-09P-001 (upgrade) and began shipping this product in January 2007. Customers who received the Headset Model; RTC-2K and RTC- 551-070, are asked to upgrade to MRI Headset Model No. RTC-661-060-09P-001. Response is requested as to the number of headsets to be replaced. In addition, consignees are advised that if they further distributed this product, to identify their customers and notify them at once of this product recall. If you have questions, contact Olivia Shayesteh at 818-882-1997 or via email at Olivia@mrivideo.com |  | 20 | 20 units | Nationwide and Internationally: Canada, Japan, Taiwan, United Kingdom, Netherlands, Italy, Spain, Germany, Switzerland, France, | Terminated | July 07, 2011 | 1515 | 49 | Not_Computer
Z-1826-2008 | 48078 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation; for diagnostic image analysis. GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | May 06, 2008 | 2008 | GE Healthcare Integrated IT Solutions | A software anomaly results in incorrect study date and time information being displayed in the report screen and title, which may result in a potential patient misdiagnosis. | Software design | GE Healthcare IITS sent Product Safety Notification letters dated 5/5/08 to all customers who have the GE Centricity PACS Software Versions 3.0.0.X, 3.0.1.X, 3.0.2.X and 3.0.3.X informing them of the incorrect study date and time information being displayed in the report screen and title, which may lead to a potential patient misdiagnosis. The users were instructed to refer to the image overlay for the correct date and time until the new software version can be installed. A software patch to permanently resolve the issue will be installed at no charge. Any questions were directed to the GE Customer Care Center at 800-437-1171. |  | 177 | 177 units | Nationwide and internationally to Australia, Belgium, Canada, Denmark, France, Germany, Italy, Japan, Malta, South Korea, Spain, Switzerland and United Kingdom. | Terminated | February 03, 2010 | 638 | 177 | Software | Display/Image | Software update | Software Update
Z-1821-2008 | 48123 | No | MVV |  |  | Neurology | 510(k) | Psi Bands acupressure/acustimulation wrist bands in the following colors: Racer Black (UPC 859570001036), Daisy Chain (UPC 859570001029), Color Play (UPC 859570001043) and Cherry Blossom (UPC 859570001012). |  |  | April 26, 2007 | 2007 | Psi Health Solutions | Product was marketed before its 510(k) application cleared. | PMA | Recall imitated 04/26/07. Retailers were informed via phone and email and mail.The product was marketed prior to the firm receiving clearance to market from the FDA. Contact Carla Falcone at 831-373-7712 if you have questions. |  | 21000 | 21000 units | Product was distributed to approximately 21,000 consumers and retailers throughout the US. | Terminated | December 18, 2008 | 602 | 21000 | Not_Computer
Z-1825-2008 | 47840 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity AW Suite 2.0 Software running on the Centricity RA 1000 or RA600/CA1000 Workstation; for diagnostic image analysis; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | May 09, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Incorrect measurement of the aortic length in certain cases of aneurysm during the use of the aorta protocol or customized protocol of the Centricity AW Suite 2.0 running on the Centricity RA 1000 or RA600/CA1000 workstations. | Software design | GE Healthcare IITS sent Urgent Safety Notice letters dated 5/9/08 to all customers who have the AW Suite Software Version 2.0, informing them of the incorrect measurement of the aortic length in certain cases of aneurysm during the use of the aorta analysis protocol or customized protocols. The users were instructed not to use the aorta analysis protocol or the customized protocols for an aneurysm until the new software version can be installed. A software patch to permanently resolve the issue will be installed at no charge. Any questions were directed to the GE Customer Care Center at 800-437-1171. |  | 207 | 207 units | Nationwide and internationally to Australia, Belgium, Canada, Denmark, France, Germany, Great Britain, Hong Kong, Ireland, Israel, Italy, Japan, Malaysia, Malta, Portugal, Saudi Arabia, South Korea, Spain, Sweden, Switzerland and Turkey. | Terminated | September 24, 2010 | 868 | 207 | Software | Output/Calculation | Software update | Software Update
Z-1752-2008 | 47245 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | PT-3B Control Unit - component of the EndoWave Infusion Catheter System. EndoWave Infusion System consists of three main components, a single use EndoWave Infusion Catheter System which is comprised of a Drug Delivery Catheter (CD) and removable Ultrasound Core (USC) and a reusable EKOS PT-3B Control System. Box is labeled EndoWave Infusion System Drug Delivery Catheter with Ultrasound Core |  |  | March 10, 2008 | 2008 | EKOS Corporation | Potential for unit to burn patient. | Software design | In March 2008, the firm planned to visit each consignee to install the new software version 01.05.02 and hand deliver letters, dated 02/20/2008, notifying consignees of the reason for the action. The old software version is 01.04.02. The software was upgraded because the Ultrasound Core heats very quickly and if removed from the DDC while operating it may result in burns to the patient or user. The software will turn off the Ultrasound Core automatically if removed from the DDC during operation, and will prompt to restart therapy if ultrasound is turned off during treatment.On 04/17/2008, the firm issued a new letter with the date of April 2008 informing customers of the issue. In addition the letter indicates that the corrective action, release of the new software is delayed to incorporate further changes, and that customers will be notified when the software is available for their unit. Customers are advised that the device is safe to use when operating according to the Instructions for Use. EKOS contact number 1-888-356-7435 |  | 218 | 218 units | Products were distributed nationwide to hospitals. | Terminated | May 13, 2011 | 1159 | 218 | Software | Physical Safety Hazards | Software update | Software Update
Z-1820-2008 | 48116 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Allez Spine Del Mar Pedicle Screw System: IT-FN 1007, Del Mar Sterilization Tray (Part Number: 30013) |  |  | April 14, 2008 | 2008 | Allez Spine, LLC | This field action was initiated because the cleaning, decontamination and sterilization procedures specified in the "Instructions for Use" for the impacted systems are incorrect.The use of these parameters in performing cleaning, decontamination and sterilization procedures may not result in clean and sterile product. | Error in labeling | Allez Spine issued a safety notice with attached fax-back attachment on April 14, 2008 via FedEx to all its hospital users, distributors and field representatives informing them of the failure of the sterilization validation and requesting them to use the new recommended sterilization parameters. The notice included copies of the updated Instructions for Use for the Allez Spine Laguna Pedicle Screw System and the Allez Spine Del Mar Pedicle Screw System. The safety notice requested that the revised recommended instructions be used for all future cleaning, decontamination and sterilization. Copies of the revised instructions were also provided with the safety notice along with links to the company website where additional copies may be downloaded.Contact Mark De Baca at 415-357-9908 if you have any questions. |  | 2 | 2 systems | Nationwide to AZ, CA, CO, GA, IL, NV, TN, TX, UT | Terminated | January 25, 2010 | 651 | 66 | Not_Computer
t | o | o |  | b | i | g | [18 empty] | Not_Computer
Z-0156-2009 | 49452 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Weck Hemoclip Traditional Ligating Clips, Rx Only, Sterile, Teleflex Medical, RTP, NC 27709 USA. The use of the reusable Teleflex Medical metal ligating clips is transient, invasive and intended for the application of hemostatic clips during general open surgery. Ligation clips are non-absorbable to thermostatically restrict flow of fluids within vessels. Clips are removed from the cartridge and applied in such a manner as to engulf vessels. Teleflex Medical metal ligating clips are intended for use in procedures involving vessels or anatomic structures for which the surgeon determines ligating clips are the best choice. Surgeons should select the size type, and material of the clip based upon their experience, judgment and needs. |  |  | August 22, 2008 | 2008 | Telefelx Medical | A hole in the sterile unit blister pack was detected that would compromise sterility. | Packaging process control | Consignees were notified via UPS next day on/about August 22, 2008. They were instructed to cease use and distribution and quarantine all affected product immediately. If they have any of the affected product they were instructed to return all it to Teleflex Medical via provided UPS-Ground. An Acknowledgement form was provided to be completed and returned via fax to Teleflex Medical to confirm receipt of the notification and amount of product on hand for return. A second mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks from the second mailing. Consignees also may be contacted by phone, e-mail, or fax. A record of these contacts will be maintained. Upon return to Teleflex Medical, product will be quarantined, assessed for recall effectiveness, and either reworked or destroyed. Contact Telefex Medical at 1-800-334-9751 ext. 8066 for assistance. |  | 786036 | 786036 (total pieces) | Nationwide | Terminated | March 15, 2012 | 1301 | 786036 | Not_Computer
Z-0030-2009 | 49462 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Platinum Plus" LT , REF 46-601, UPN M001466010, LT/0.018/180, Sterilized with ethylene oxide gas, Rx only, Legal Manufacturer Boston Scientific Corporation, One Boston scientific Place, Natick, MA 01760-1537, Made in USA: 8600 NW 41st Street, Miami, FL 33166-6202. The Platinum Plus" Guidewire facilitates placement of a catheter during diagnostic or interventional intravascular procedures. The wire can be torqued to facilitate navigation through tortuous arteries and/or avoid unwanted side branches. |  |  | August 25, 2008 | 2008 | Boston Scientific Corporation | Boston Scientific is initiating a voluntary recall of the Platinum Plus Guidewire (peripheral). BSC has received reports that product may be labeled as a 260 CM long guidewire when the actual packaged device is a 180 CM long guidewire and the product may be labeled as a 180 CM long guidewire when the actual packaged device is a 260 CM long guidewire. There is no expected potential for injury to | Labeling mix-ups | A "Boston Scientific Urgent Voluntary Medical Device Recall or Correction - Immediate Action Required" letter dated 8/25/08 was sent to each affected account (To The Hospital Field Action Contact). The letter includes the customer steps for recall, a complete listing of all product batches/lots being recalled, the Reply Verification Tracking Form(s), and a return shipping label. Contact Boston Scientific Corporation at 1-763-494-1700 for assistance. |  | 47 | 47 | USA: CAFLGAILINLAMNMONJNVNYOHPATNTXUTWAWI andWVOUS: France, Germany, Great Britain, Italy, and the Netherlands. | Terminated | December 10, 2011 | 1202 | 82 | Not_Computer
Z-0034-2009 | 49494 | No | GAR | Nonabsorbable polyamide surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 3-0 Black Mono Nylon Suture, 12 units per box. product number 925B. The product intended for use in general soft tissue approximation and/or ligation including use in ophthalmic procedures. |  |  | September 09, 2008 | 2008 | Surgical Specialties Corp | mislabeled product - labeled as 3-0 black mono nylon contains 4-0 black braided silk suture | Employee error | The recalling firm issued a Recall letter dated 9/9/08 to its customers. The letter informed the customers of the problem, the need to conduct a sub-recall, and return the product. Contact Surgical Specialties Corp. at 1-610-404-1000 for assistance. |  | 956 | 956 boxes | USA--AZ, CA, CT, FL, GA, IA, IL, KS, KY, LA, MD, MI, MS, NC, NJ, NY, OH, TN, TX, VA, and WA. OUS: Australia, Canada, and Fiji Islands. | Terminated | March 27, 2009 | 199 | 956 | Not_Computer
Z-0021-2009 | 49435 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | VICRYL RAPIDE (polyglactin 910) Braided Coated Synthetic Absorbable Suture, Undyed, Non-USP. |  |  | August 11, 2008 | 2008 | Ethicon, Inc. | Package defect compromised the integrity of the primary seal, which could lead to premature suture degradation and/or impair the sterile barrier of the product. | Packaging | Letters were sent to distributors, Director of Surgery, Director of Risk Management, Director of Materials Management and Director of Operating Room on August 11, 2008, requesting that all affected lot numbers of Vicryl Rapide be discontinued immediately from sale and/or use. A separate correspondence was sent to hospitals. Letters to Director of Surgery, Director of Risk Management, Director of Materials Management and Director of Operating Rooms were advised of the recall and asked to discontinue use of any of the affected lot codes of sutures. Return and credit instructions were provided to the Director of Material Management. Contact Ethicon, Inc. at 1-908-218-2553 for assistance. |  | 6912 | 6912 | Worldwide: US and Canada OUS: Australia, New Zealand and Indonesia. | Terminated | June 05, 2012 | 1394 | 6912 | Not_Computer
Z-0014-2009 | 49022 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Mobetron Model 1000, mobile electron linear accelerator, a device system which delivers radiation to an intraoperative site. |  |  | June 20, 2008 | 2008 | Intraop Medical Corporation | Docking screw/nut failure, which may result in a treatment head dropping, potentially impacting the patient. | Component design/selection | Recall initiated on 06/20/2008, advising users via Advisory Notice that a safety upgrade would be available. Two options are given to either self inspect the product or schedule an Intraop Authorized Field Service Engineer to visit their facility to perform the inspection. The letter notes that if no nut failure was evident, the machine could still be used. The firm continues to send retrofit kits. |  | 24 | 24 units | Worldwide Distribution | Terminated | April 05, 2011 | 1019 | 24 | Not_Computer
Z-0013-2009 | 49003 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences***Fogarty Fortis***Arterial Embolectomy Catheter |  |  | June 23, 2008 | 2008 | Edwards Lifesciences, LLC | Potential for tubing fracture near the tip of the catheter. | Process control | The firm sent Urgent Field Safety Notices dated June 23, 2008 to customers. The letter gave details on affected products, a description of the problem, and advise on action to be taken. Edwards has decided to recall these lots of Fogarty Fortis catheters which were manufactured between 9 June 2007 and 20 June 2008. The expiration date for these lots is between 11 September 2009 and 26 September 2009. The attached sheet lists product by model and lot number which you have ordered during that time frame. Please stop using these lots of catheter products immediately. Edwards requests that you return all Fogarty Fortis catheters from these lots that are still in your inventory. Please call Edwards Customer Service to arrange for return of any unused product and obtain information about replacement product. After you have verified your inventory, please complete the attached Confirmation Form, even if you no longer have any affected units at your facility. This will allow us to verify the completion of this recall action. Please return the Confirmation Form by FAX to: +31.412.47.66.99 , Attention: Mark vanSchijndel Transmission of this Field Safety Notice: This notice needs to be passed on all those who need to be aware within your organization. Please transfer this notice to other organizations on which this action has an impact and where or to any organization where the potentially affected devices have been transferred. Please maintain awareness on this notice and resulting action for an appreciate period to ensure effectiveness of the corrective action. We sincerely regret the inconvenience caused by this action and greatly appreciate your immediate attention to this matter. If you have specific questions about this notification, please contact Customer Service |  | 2975 | 2,975 units | International Distribution --- including countries of Austria, Belgium, Bulgaria, Switzerland, Czech republic, Germany, Denmark, United Kingdom, Italy, Luxembourg, Netherlands, Sweden, Slovak Republic, South Africa, and Taiwan. | Terminated | September 19, 2012 | 1549 | 2975 | Not_Computer
Z-0032-2009 | 49465 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet brand Discovery Elbow, 3 mm x 75 mm right ulna with bearing/bond coat, TI 6AL 4V Alloy/ARCOM UHMWPE, For cemented use only, sterile, Biomet Orthopedics, Inc., Warsaw, IN; REF 114813. The product intended use is for Orthopedic implant. |  |  | May 23, 2008 | 2008 | Biomet, Inc. | The component in the package is not the correct size. | Packaging process control | Consignees were notified by letter dated May 28, 2008 to locate and return the recalled lots. Contact Biomet Orthopedics at 1-800-348-9500, ext. 3983 for assistance. |  | 6 | 6 | USA--California, Illinois, Indiana, Massachusetts, Ohio Puerto Rico, and Canada, OUS: Germany and Great Britain. | Terminated | September 22, 2009 | 487 | 11 | Not_Computer
Z-0022-2009 | 49436 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD Visitec High Viscosity Injector, 4mm |  |  | September 15, 2008 | 2008 | Becton Dickinson and Company | Product labeled as 4 mm High Viscosity Injector tip contains a 6 mm tip. | Employee error | BD contacted accounts by telephone on 9/15/08 requesting users to discard units and complete the response form. Contact Becton Dickinson and Co. at 1-781-906-7950 for assistance. |  | 1485 | 1485 units | AZ, CA, CO, IL, OH, FL, PA , TN, TX,and VAForeign: Australia and Japan. | Terminated | February 29, 2012 | 1262 | 1485 | Not_Computer
Z-0020-2009 | 49379 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Datex-Ohmeda Aisys, Datex-Ohmeda, Inc., a General Electric Company going to market as GE Healthcare, GE Healthcare PO Box 7550, Madison, WI. 53707-7550 |  |  | August 28, 2008 | 2008 | Datex - Ohmeda, Inc | The electronic vaporization system in the GE Aisys Anesthesia machine contains the following components: Backpressure Valve, Inflow Check Valve and Cassette Interface Board, that can contribute to independent failures. | Manufacturing material removal | GE Healthcare "Urgent Medical Device Correction" letters dated 8/12/08 were sent on 8/20/08 individually to Health Administrator/Risk Manger, Head of Anesthesia and Director of Clinical Engineering. The letters described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. The Safety Instructions state that the machines can be safely used as long as the following guidelines are followed:1. Carry out the daily system checkout for any system in use.2. The Aisys Anesthesia will alert the user both visually and audibly if it detects a fault condition.3. Have an alternative method of anesthesia available at all times.Once each machine has been corrected an "Update" will be provided by the service technician responsible for the correction. Each customer/machine will receive a "GE Healthcare Urgent Medical Device Correction Update" letter which informs the customer that this correction is a follow-up activity to the letter that was sent earlier. |  | 1571 | 1571 | Worldwide Distribution --- USA including states of AL, AK, CA, CO, FL, GA, IL, IN, KS, KY, ME , MD, MA, MI, MO, MT, NE, NV, NH, NY, NC, OH, OK, SC, SD, TX, VA, WV, WI, and DC, and countries of SOUTH AFRICA, TURKEY, THAILAND, SINGAPORE, SWEDEN, SAUDI ARABIA, QATAR, PORTUGAL POLAND, NEW ZEALAND, NORWAY, NETHERLANDS, MEXICO, LIBYA, KUWAIT, KOREA, JAPAN, JORDAN, ITALY, UNITED KINGDOM, FRANCE FINLAND, SPAIN, ESTONIA, DENMARK, GERMANY, COLOMBIA, CHINA, SWITZERLAND, CANADA, BRAZIL, BAHRAIN, BELGIUM, AUSTRALIA, AUSTRIA, and UNITED ARAB EMIRATES. | Terminated | March 08, 2016 | 2749 | 1571 | Hardware |  | Instructions to verify | Safety Notice/Insructions
Z-0028-2009 | 49445 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical, Lap Banding Kit, REF: K0406 Contains: (1) C0658, 11x100mm Threaded Shielded Trocar w/Universal Seal (1) C0522, 5x100mm Threaded Premium Flat Blade (1) C2202, 150mm Insufflation needle (1) C0605, 15x100mm Threaded Separator Access System (1) CB030, 5mm x 35cm Direct Drive Disposable Scissors (2) C4120, 5mm x 38cm Direct Drive Grasper Reposable Cartridge, for laparoscopic procedures. |  |  | June 25, 2008 | 2008 | Applied Medical Resources Corp | Potential inability to insufflate through the stopcock. Under certain circumstances, a flexible elastomeric component inside the 15mm trocar canstretch into a configuration that blocks the flow of insufflation gas (C02). | Process control | Customers were notified via FedEx or UPS on June 25,2008 with an overnight letter. The letter advises that Applied Medical is conducting a voluntary recall of the 15mm Separator Access Systems, Models C0604, C0605, C0606 and C0607 due to the potential inability to insufflate through the stopcock. Per the letter, the firm is asking that all 15mm product listed in the letter be returned. A response form is provided. |  | 14113 | 14,113 total | Nationwide and Canada distribution. | Terminated | March 26, 2012 | 1370 | 14113 | Not_Computer
Z-0088-2009 | 49490 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | Envision E700 Low Airloss Therapy Surface, Made By Hill-Rom Manufacturing, Inc., Charleston, SC. |  |  | June 19, 2008 | 2008 | Hill-Rom Manufacturing, Inc. | A defect in the software of the device may not allow the patient bed exam alarm or the patient movement alarm to function correctly. | Software change control | An internal Modification Notice dated 9/1/08 was sent via email to Service Centers (because the device is owned by the manufacturer, Hill-Rom, each Service Center in possession of the device will be the consignee). Each Service Center will perform the correction on the units affected. Each device in the affected serial number range will be tested for the software defect. If the defect is found, the circuit board that contains the defective software will be replaced. The Field Action Plan is expected to be completed on July 24, 2008. Expected completion date of the Correction will be 12 months from July, 2008. |  | 779 | 779 units | Worldwide Distribution --- including USA and country of Canada. | Terminated | May 21, 2009 | 336 | 779 | Software | Output/Calculation | Replace circuit board containing software | Software Update
Z-0087-2009 | 49442 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Draeger Medical Apollo Anesthesia Machine, catalog number 8606500 |  |  | August 22, 2008 | 2008 | Draeger Medical, Inc. | Sporadic errors in various device functions, including low readings of the single gas flows for oxygen, nitrous oxide, or air delivery, or mechanical ventilation failure. Alarms and error messages function as designed. | Nonconforming Material/Component | Draeger Medical sent a Medical Device Recall letter to their customers via certified mail on 8/22/08 informing them of the problem. Until the exchange of affected printed circuit boards has been implemented, the firm is recommending that users closely monitor all device displays and error messages and if applicable act in accordance with the advisory statements and remedial information provided in the Instructions for Use as well as making provisions for alternative patient care measures. |  | 214 | 214 units | Nationwide Distribution. | Terminated | September 01, 2009 | 375 | 214 | Other | Output/Calculation | Replace | Remove or Replace
Z-0086-2009 | 49411 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens Ultrasound, Acuson/Sonovista X300 ultrasound systems, software revisions Sonovista X300 software revision 1.0.06 and 1.0.07, Acuson X300 sofware revision 1.0.08 ultrasound system with onscreen display. |  |  | August 19, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | Incorrect value calculations by the device may result in inaccurate aortic stenosis estimates. | Software design | Siemens initiated a Field Correction 08/19/2008. Service personnel will conduct a site visit and distribute a Customer Safety Advisory Notice, correct the problem and install a revised software version. |  | 169 | 169 units total | Worldwide Distribution: USA, Singapore, Germany, Denmark, Romania, Sweden, Poland, Portugal, Australia, Chile, India, Spain, Thailand, Malaysia, South Africa, Hungary, Saudi Arabia, Qatar, Peru, Italy, Azerbaijan, Greenland, Austria, Venezuela, Puerto Rico, Brazil and Japan. | Terminated | April 21, 2011 | 975 | 169 | Software | Output/Calculation | Software update | Software Update
Z-0012-2009 | 48638 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche Diagnostics NAPA2, N-Acetyl-Procainamide, COBAS INTEGRA, Roche Diagnostics, Indianapolis, IN.; 20737852322. |  |  | June 11, 2008 | 2008 | Roche Diagnostics Corp. | Some of the labels are glossy, resulting in the bar code being difficult to read by the bar code reader. | Labeling Change Control | Consignees were notified via an Urgent Medical Device correction letter dated 6/11/08 informing users of the problem and providing them with work-around instructions. The instructions suggests 1) inserting the cassette rack slowly; 2) using mainly reagent slots A, B, or C; 3) tilting the rack slightly clockwise so the barcode scanner does not read the barcode at an exact 90 degree angle; and 4) confirming the rack is properly seated on the track and displayed on the Cassette Status screen. If the cassette barcode does not register after following the Work-Around instructions, then customers are required to contact Diagnostics Technical Support. |  | 2 | 2 | Nationwide Distribution | Terminated | September 22, 2009 | 468 | 3624 | Not_Computer
Z-0015-2009 | 49107 | No | OHM | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | Enforcement Discretion | Medline Latex Free C-Section CDS-LF and C-Section Supplemental CDS-LF; a single patient prescription procedure pack and disposables, including a Gerber Nuk Pacifier; Made in USA by Medline Industries, Inc., Mundelein, IL 60060 |  |  | August 04, 2008 | 2008 | Medline Industries Inc | The latex free labor &amp; delivery packs contain a latex Nuk pacifier. The pacifier is packaged separately and offers no risk of exposure of latex particles to other components of the pack. | Employee error | Recall letters dated 8/4/08 were sent to the direct accounts on the same date, informing them that the labor and delivery packs were labeled Latex Free, but contained a Latex Nuk pacifier. The letters were customized for each account, listing the specific affected reorder numbers and lot numbers shipped to each account. The accounts were requested to immediately check their stocks for the affected lot numbers and quarantine the affected product found. The accounts were given three options: A) Attach the enclosed stickers to the packs so that the user is aware of the latex Nuk pacifier, B) Call Medline to receive replacement non-latex pacifiers, but also place the stickers on the packs, C) Return the packs to Medline for replacements. The accounts were requested to complete the enclosed response form, indicating which option the account selected and whether they needed any additional stickers, and fax the form back to Medline at 847-643-4482 by 8/18/08. Any questions were directed to Kathy Dunne at 1-866-359-1704. |  | 579 | 579 cases | Nationwide Distribution --- including states of Alabama, Arizona, Georgia, Missouri, Nevada, New Jersey, and Washington. | Terminated | October 20, 2008 | 77 | 2115 | Not_Computer
Z-0101-2009 | 49417 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM Luminos TF, model number 10093902 |  |  | August 12, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Liquids may enter the system and cause potential malfunction and possible hazard to patients, user or other persons. | Device Design | The recalling firm was made aware of a potential problem with the device from the manufacturer in Germany. The recalling firm sent a copy of the Customer Safety Advisory Notice, dated June 18, 2008, received from Germany to their customers. The recalling firm has representatives visiting each location to modify the units. |  | 211 | 211 units | Nationwide Distribution | Terminated | October 05, 2009 | 419 | 211 | Not_Computer
Z-0036-2009 | 49405 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Edwards Lifesciences LifeStent FlexStar Self-Expanding Biliary Stent System, Edwards Lifesciences LLC Irvine, CA 92614, Model EX060203CD. The LifeStent FlexStar self-expanding biliary stent system is intended for use in the palliation of malignant strictures (neoplasms) in the biliary tree. |  |  | August 29, 2008 | 2008 | Bard Peripheral Vascular Inc | Some LifeStent FlexStar Systems may exhibit a gap between the tip of the delivery system and the primary sheath such that the guidewire lumen is visible. | Process control | The customer letters for the tip to sheath gap were sent on 8/29/2008 with return receipt. The letter advises that C.R. Bard has determined that some LifeStent FlexStar Systems may exhibit a gap between the tip ofthe delivery system and the primary sheath such that the guidewire lumen is visible. Per the letter, the firm's analysis has determined that a gap between the tip of the delivery system and the delivery system sheath may be associated with difficult advancement to the treatment site. Users are instructed to inspect all devices to verify that there is not a gap between the tip and sheath (orange colored guidewire lumen is visible), prior to use and to not use any devices that exhibit a tip to sheath gap condition. If devices that exhibit a tip to sheath gap are identified, customers are advised to call Customer Service at 1-800-321-4254 to return the device and a replacement will be provided at no cost. The letter also advises that the Instructions for Use for the LifeStent FlexStar System will be updated to include the inspection step which is provided in the letter. |  | 92 | 92 units | Nationwide | Terminated | April 06, 2012 | 1316 | 9889 | Not_Computer
Z-0111-2009 | 49412 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens Ultrasound, Acuson/Sonovista X300 ultrasound systems, software revisions 2.0.1 to 2.0.05, 3.0.01 and 3.0.02, manufactured by Siemens Ultrasound |  |  | August 19, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | Thermal Index cranial (TIC) is not displayed for the Neo-Head exam type with the C8-5 transducer. This is a required display to support the alara principle when imaging cranial structure. | Employee error | Field correction was initiated on 8/19/2008. Under Siemens update program US020/08/S &amp; US021/08/S Siemens service personnel make a site visit, deliver a customer safety advisory notice to the customer and correct the problem by installing the revised Software versions. After the planned field correction is performed at the customer site the Service personnel log that information through Siemens internal web or reporting tools into the Siemens field update database. |  | 72 | 72 units | Worldwide Distribution --- including USA and countries of France, Colombia, Poland, Taiwan, Saudi Arabia, Austria, Malaysia, India, Czech Republic, Croatia, Argentina, Venezuela, Italy, Philippines, Mauritius, Spain, Bolivia, Germany, Ukrain, Russia, United Kingdom, Brazil, South Africa and Canada. | Terminated | December 20, 2010 | 853 | 72 | Software | Display/Image | Software update | Software Update
Z-0100-2009 | 49402 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | M3811B Philips Telemonitoring Clinical Review Software part of Philips Telemonitoring System, M3810A, software revision identified as Build 1.1.2.11d, B.02.07. Prescriptive medical device used to automatically collect and transmit medical information (weight, blood pressure, non-diagnostic ECG) over phone lines between provider and patient. |  |  | September 11, 2008 | 2008 | Philips Medical Systems | Multiprint report may contain incorrect vital data for patients | Software design | Philips contacted customers by sending an Urgent Medical Device Correction Notification on 9/11/08. Letters to customers sent via certified mail and receipt of delivery will be tracked.Customers are asked to follow the Action To Be Taken By Customer/User section of the Urgent Medical Device Correction Notification. This section instructs users to allow the upgrade of the software upon logging in to the system. Contact Philips at 1-866-246-7316 for assistance. |  | 52 | 52 units | Nationwide | Completed |  |  | 52 | Software | Output/Calculation | Software update | Software Update
Z-0105-2009 | 49061 | No | NTE | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific FilterWire EZ" Embolic Protection System, Model FW EZ 190cm, 3.5-5.5 mm, US, Material H749201001900, Catalog # 20100-190, Sterilized with irradiation, Made in USA: 2011 Stierlin Court, Mountain View, CA 94043-4655. Filter wire EZ" Embolic Protection System is indicated for use as a guide wire and embolic protection system to contain and remove embolic material (thrombus/debris) while performing angioplasty and stenting procedures in coronary saphenous vein bypass grafts and carotid arteries. The diameter of the vessel at the site of filter loop placement should be between 2.25 mm and 5.5 mm for coronary saphenous vein bypass graft procedures and between 3.5 mm and 5.5 mm for carotid procedures. |  |  | May 20, 2008 | 2008 | Boston Scientific Corporation | Boston Scientific is initiating a field correction for 48 lots/batches of the Filter Wire EZ" Embolic Protection System. Boston Scientific has determined that the Directions for Use (DFU), which accompanies each packaged device, may be missing. If the product does not include a DFU and the user is unable to find a copy, the user will have to seek another product box containing the DFU. This may | Employee error | Consignees were sent a "Boston Scientific Urgent Medical Device Recall or Correction-Immediate Action Required" letter dated May 20, 2008. The letter addressed to Risk Mangers, described the issue and product. A Reply Verification Tracking Form was requested to be completed and returned. Contact Boston Scientific Corporation at 1-800-272-3737 for assistance. |  | 4338 | A total of 4338 units distributed. | US: AL AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD. TN, TX. UT, VA, VT, WV, WA, WI, WY, and PR.European Countries: Austria, Belgium, Cyprus, Czech Republic, Denmark, Finland, France, Great Britain, Germany, Greece, Hungary, Italy, Latvia, Luxembourg, Netherlands, Norway, Poland, Portugal, Russia, Slovakia, Spain, Sweden, and Switzerland, Intercontinental: Taiwan, Korea, Australia, China, Singapore, Canada, Columbia, and Mexico. | Terminated | December 10, 2011 | 1299 | 4338 | Not_Computer
Z-0114-2009 | 49416 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Safety Lancet for capillary blood sampling, 1.6 mm x 28G, Mini, Sarstedt. (Blood lancet) |  |  | August 27, 2008 | 2008 | Sarstedt Inc | The product from this lot may not retract the lancet into the safe position inside the lancet body after the trigger is actuated. | Process control | The consignee was notified via letter that was hand delivered on 08/27/2008. All of the distributed product was picked up from the consignee and returned to Sarstedt, Inc. None of the product had been used. All action was completed on 8/27/2008. Contact Sarstedt at 1-828-465-4000 for assistance. |  | 1000 | 1000 pieces | NC | Terminated | April 24, 2012 | 1336 | 1000 | Not_Computer
Z-0113-2009 | 49518 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Zilver Expandable Metal Biliary Stent System, ZILBS-10-6, Stent Diameter: 10mm, Stent Length: 6cm, Introduction System: 7 FR., Disposable-Single Use Only, Rx Only, Cook GI Endoscopy, Winston-Salem, NC. The device is used in palliation neoplasms in the biliary tree. |  |  | September 11, 2008 | 2008 | Cook Endoscopy | A section of the introduction system may detach after the stent has deployed. | Process control | The consignee was notified by letter on/about 09/11/2008. The distributor was instructed to contact their customers and coordinate the return of all affected product to Cook Endoscopy. A responds form was attached to the letter to be completed and returned to Cook Medical to confirm receipt of notification and indicate quantity to be returned. Contact Cook Endoscopy at 1-336-744-0157, ext. 6325 for assistance. |  | 12 | 12 | USA: IL and OUS: China | Terminated | April 01, 2009 | 202 | 12 | Not_Computer
Z-0102-2009 | 49496 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging with versions VB20B, VB20D, VB20F, VB20G, VB25B, VB30A, VB30A_SP1, and VB30A_SP2. |  |  | August 22, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Image may not be visible after merge. | Software design | The recalling firm issued Customer Safety Advisory Notice dated 8/22/08 to inform the consignees of the potential issue and provide instruction to avoid its occurrence. To avoid potential risk of this issue, users are asked to check whether the extended HL7 messages as stated in the letter are mandatory for their RIS-PACS workflow, and disable these messages if they are not. Siemens is preparing a modification under the UI IM086/08/S for the affected versions of syngo Imaging that will resolve the potential malfunction. The field modification will be immediately distributed and installed by the firm's Service as soon as it is available. |  | 59 | 59 units | Nationwide Distribution --- including states of AL, AZ, CA, CO, CT, FL, IA, KS, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, TN, TX, and WI. | Terminated | November 19, 2008 | 89 | 59 | Software | Display/Image | Software update | Software Update
Z-0112-2009 | 49459 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Centricity Perinatal (formerly Quantitative Sentinel) System - Fluid Total Precision software; automatic patient data management providing clinical information at the bedside in Labor &amp; Delivery, Mother-Baby and the Neonatal Intensive Care Unit; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | September 25, 2008 | 2008 | GE Healthcare Integrated IT Solutions | On the I&O; chart, the IN, OUT and NET fluid totals values will not honor the numeric precision configuration, always displaying a whole number, losing decimal point accuracy. | Software design | GE Healthcare notified their Centricity Perinatal customers via letter dated 4/20/07 of the problem with the Fluid I &amp; O Module when fluid precision is required. The letter instructed those accounts that require fluid precision to contact GE Technical Support at 1-800-433-2009 to obtain a software patch to fix the problem. A follow-up letter dated 8/29/08 was sent on the same date to the accounts informing them that GE had re-classified the issue as a class II safety issue and that all customers need to have their software updated. The accounts were requested to contact GE Tech Support at 800-433-2009 to schedule a time to upgrade their system to 6.70.6 P02 as soon as possible. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax it back to 847-277-5240. |  | 470 | 470 units | Worldwide Distribution --- including USA and country of Canada. | Terminated | May 17, 2010 | 599 | 470 | Software | Output/Calculation | Software update | Software Update
Z-0092-2009 | 48432 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PRIMEVIEW 3i System equipped with a Coherence Therapist R2.0 or R2.1, and PRIMEVIEW 3iR2.0 or R2.1, part number 7341428 |  |  | November 06, 2007 | 2007 | Siemens Medical Solutions USA, Inc | Flat panel positioning calibration could be off by as much as 4 mm without the machine discovering detail. | Software design | Service calibration inspection/correction Update instructions TH038/07/S was issued on 11/06/2007 to all customers. The instructions identified the affected products, stated the reason for the update, and listed the steps to update the software. Customer notification was also issued. Software patch planned for release in June 2008. |  | 13 | 13 units | Worldwide Distribution | Terminated | December 18, 2010 | 1138 | 251 | Computer
Z-0161-2009 | 49545 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Accessory Head Rest for Alphastar Operating Tables |  |  | September 26, 2008 | 2008 | MAQUET Inc. | During installation or cleaning there is a risk of crushing the finger by handling the head rest in the area of the gas strut while simultaneously actuating the release lever. | Other | Urgent Device Correction letters were sent beginning September 26, 2008 via Fed Ex. Letters identified the potential for finger crush in conjunction with the head rest. Proper operating instructions are provided as well as a warning sticker to be applied to the head rest indicating proper positioning. All letters included a customer response form to be completed by each institution. Information requested is: that device correction letter has been read and instructions followed; that the corrected operating instructions have been received; that the warning sticker has been properly placed on the headrest, or that they do not have the head rest identified in the device correction letter. Questions should be addressed to USA Maquet Inc at 1-888-627-8383 #1. |  | 1972 | 88 units in US; 1884 units worldwide | Nationwide Distribution | Terminated | April 16, 2009 | 202 | 1972 | Not_Computer
Z-0116-2009 | 49024 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV system, Part number 8139789, medical linear accelerator equipped with COHERENCE therapist RTT4.1. Medical Linear accelerator for radiation therapy. The device utilizes a console system for selecting various treatment sequences and manages commands for all motorized movement. |  |  | July 04, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Three issues: unexpected rotation of gantry, unexpected movement of table between beams during patient setup, and single exposure images will be overwritten by double exposure images. | Software change control | A letter marked "Urgent" dated July 4, 2008 was sent to customers. The letter provided a detailed description of the issues, the conditions under which these issues may occur, as well as recommendations to avoid these issues. Customer notification TH022/08/S was released on July 11, 2008. A software fix is currently being validated. All consignees notified by certified mail. Contact Siemens Medical Solutions USA, Inc. at 1-925-246-8407 for assistance. |  | 16 | 16 | Nationwide in LA, WI, DE, PA and OH, and internationally in Netherlands, Germany, Poland, Norway, Australia, Hungary, Belgium and France | Terminated | December 22, 2010 | 901 | 16 | Software | Physical Safety Hazards | Software update | Software Update
Z-0115-2009 | 48995 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Scorpio Series 7000 Tibial Impactor/Extractor, Catalog number 3770-000 |  |  | July 02, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that there is a potential inability to release the Scorpio Series 7000 Tibial Impactor/Extractor from a tibial baseplate during use. | Component design/selection | US recall notification (product correction) letters were send out by Federal Express with return receipt on July 2 and 9, 2008.Advisory notices were sent to foreign customers on July 2, 2008.Both letters include a Product Correction Bulletin that instructs all used to lubricate the divide with an approved medical lubricant, especially the extractor shaft and threads. It is recommended that all moving parts be evaluated for functionality and useful service prior to each surgery and instructions on how to perform this process. Contact Stryker Howmedica Osteonics Corp. at 1-201-831-5918 for assistance. |  | 5126 | 5126 | Worldwide including the US and Asia Pacific, Canada, Europe, Middle East, Africa, Japan and Latin America. | Terminated | August 17, 2012 | 1507 | 5126 | Not_Computer
Z-0117-2009 | 49399 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Exact Couch with Exact Couch Top, a component of the C-Series Clinacs, Acuity and Ximatron treatment systems. Model numbers H14, H18, H27, H72, H77. A powered radiation therapy patient support assembly is an electrically powered adjustable couch intended to support a patient during radiation therapy. |  |  | August 11, 2008 | 2008 | Varian Medical Systems Oncology Systems | During couch movement, a patient's fingers may get pinched at certain points which could result in broken, disengaged or pinched fingers. | Device Design | Recall initiated on 08/11/2008. Notification will be sent to Varian Sales, Marketing and Service organizations as well as to all affected users. Contact Varian Medical Systems Oncology Systems, at 1-650-424-5731 for assistance. |  | 3296 | 3296 units | Worldwide including to the US and to Austria, Belgium, Czech Republic, Denmark, Finland, France, Germany, Greece, Hungary, Iceland, Ireland, Italy, Latvia, Lithuania, Luxembourg, Norway, Poland, Portugal, Romania, Slovakia, Slovenia, Spain, Sweden, Switzerland, The Netherlands, United Kingdom, Algeria, Australia, Bahrain, Bangladesh, Belorussia, Bolivia, Brazil, Canada, Chile, China, Colombia, Ecuador, Egypt, Guadeloupe, Hong Kong, India, Indonesia, Iran, Israel, Jamaica, Japan, Kazakhstan, Lebanon, Macau, Macedonia, Malaysia, Martinique, Mexico, Morocco, Nepal, New Zealand, Oman, Pakistan, Panama, Peru, Philippines, Russian Fed., Saudi Arabia, Serbia, Montenegro, Singapore, South Africa, South Korea, Sudan, Syria, Taiwan, R.O.C., Thailand, Turkey, U.A.E, Ukraine, Uruguay, Venezuela, and Vietnam. | Terminated | November 27, 2010 | 838 | 3296 | Not_Computer
Z-0151-2009 | 49439 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Ray TFC Device with End Caps: |  |  | August 18, 2008 | 2008 | Stryker Spine | The label on one side on the Ray TFC Device (Ray Cage) is incorrect. It reads 20 mm x 26 mm instead of 20 mm x 21 mm. | Labeling mix-ups | Urgent Product Recall notification letters were sent via Fed Express on August 18, 2008 to Stryker Branches and Hospital Risk Managers. The firm asked customers to examine their inventory and hospital locations to identify the product; reconcile all products on the Product Accountability Form and attached distribution sheet; and fax a copy of the Accountability Form and distribution sheet to Sarah Thompson within 5 days of receipt of the notice. All questions are to be directed to Susan Krasny at 201-760-8150. |  | 69 | 69 units | Nationwide Distribution | Terminated | November 04, 2008 | 78 | 69 | Not_Computer
Z-0119-2009 | 49407 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Sonesta 6210 Fluoroscopy Procedure Table; a motorized procedure table with adjustments (via a remote control) for height, tilt, back and seat cushion position; Stille Surgical AB, Sundbybergsvagen 1A, SE-17173 Solna, Sweden |  |  | August 27, 2008 | 2008 | Stille AB | The table contains a battery backup, not mentioned in the User Manual, which automatically will provide power to the electrical actuators/motions in case of a power loss from the main power supply. However, if a short circuit occurs, it can result in an uncontrolled movement of the table which cannot be stopped by disconnecting the device from the main power supply, as the battery automatically w | Component design/selection | Stille sent Field Safety Notices dated 8/27/08 to the direct accounts, informing them that the table contains a battery backup, not mentioned in the User Manual, which automatically will provide power to the electrical actuators/motions in case of a power loss from the main power supply. However, if a short circuit occurs, it can result in an uncontrolled movement of the table which cannot be stopped by disconnecting the device from the main power supply, as the battery automatically will supply continuous power to the control box. The accounts were requested to complete and return the enclosed acknowledgement card to Stille by fax or e-mail as soon as possible and no later than 9/26/08. The accounts were advised to not further use the table in conjunction with a C-arm or any other heavy object until a service technician can permanently remove the battery from the control box according to the enclosed Service Manual Amendment No. 1. Stille will make a permanent correction to the table once it has been validated and approved. Contact Stille Surgical Inc. at 1-800-655-1614 for assistance. |  | 114 | 114 units | Nationwide and Canada | Terminated | December 10, 2010 | 835 | 114 | Battery | Physical Safety Hazards | Instructions until upgrade | Repair
Z-0163-2009 | 49567 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Aristos FX Multipurpose Radiography System, stationary X-Ray system, Model number: 7414803. |  |  | August 25, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Tube support arm or the detecting mounting hardware may become loose at the joint with the telescope. | Equipment maintenance | Updated Instructions were issued in August 2008 (AX049/08/S) to Siemens Service Engineers to have them inspect and tighten any loose hardware on the affected systems. |  | 7 | 7 units | Nationwide distribution including states of CA, IN, KY, UT, VA and WI. | Terminated | October 20, 2008 | 56 | 7 | Not_Computer
Z-0152-2009 | 49443 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Auto Syringe AS50 Infusion Pump, product code 1M8550; Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015 U.S.A. |  |  | September 05, 2008 | 2008 | Baxter Healthcare Corp. | Non-conforming Electrostatic Discharge (ESD) grounding squares may cause the pump to be susceptible to shorting-out of circuitry, resulting in a loss of audio, and/or interruption of therapy. | Nonconforming Material/Component | Urgent Product Recall letters dated 9/5/08 were sent on the same date via first class mail to the direct accounts, to the attention of the Director of Biomedical Engineering and Facility Risk Manager. If the account had device(s) that had been serviced by Baxter with a part affected by this action, a serial number list of the affected device(s) was provided with the notification. Customers were instructed to remove the affected device(s)/spare parts from service and contact Baxter to arrange for their return and repair. Alternatively, if removal of the device is not a viable alternative for the customer and they wish to perform self-service, the customers will also be provided with a Service Bulletin to provide instruction on replacement of the rear case assembly. |  | 509 | 509 units | Worldwide Distribution --- USA including Puerto Rico, and countries of Canada and Korea. | Terminated | February 03, 2010 | 516 | 509 | Hardware | Alarm/Message | Repair | Repair
Z-0103-2009 | 47839 | No | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | Zyoptix XP Microkeratome maxon motor, Ref #ZXP09183. The label shows the product is manufactured by Bausch &amp; Lomb Incorporated, Rochester, NY. |  |  | April 23, 2008 | 2008 | Bausch &amp; Lomb Inc | The plastic tip on the drive shaft may come off the shaft and cause the blade to stop oscillating. | Device Design | Phone calls were initiated on 4/28/08 and an Urgent Recall letter dated 4/23/08 was issued 4/28/08 via overnight mail to U.S. customers explaining the reason for recall. The letter also requested that customers stop using the motors identified and return of the motors for evaluation. A return form was enclosed. |  | 69 | 69 motors | Worldwide Distribution --- including USA and countries of Brazil, The Netherlands, Canada, India, Malaysia, Mexico, and Singapore. | Terminated | April 02, 2012 | 1440 | 113 | Not_Computer
Z-0164-2009 | 48661 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian RV Software |  |  | June 02, 2008 | 2008 | Varian Medical Systems Oncology Systems | A malfunction within the Varis software might result in a misadministration (underdose). The software does not correctly interpret the number, causing field parameters having decimal values to be incorrect. | Software design | Recall initiated on June 2, 2008. A product notification letter was distributed by certified mail to all affected consignees. |  | 18 | 18 | Worldwide Distribution --- including USA and countries of Canada, Austria, France, Germany, South Korea, Sweden, and the UK. | Terminated | May 03, 2011 | 1065 | 18 | Software | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-0165-2009 | 48711 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche/Hitachi Modular E Module immunoassay analyzer, Roche Diagnostics, Ind., IN; GMMI Nos. 04998642001 and 03617505001. |  |  | June 24, 2008 | 2008 | Roche Diagnostics Corp. | A software bug may result in pipetting from an incorrect reagent pack and/or assigning calibration curve parameters incorrectly. | Other | The consignees were notified of the problem via letter dated 6/24/08 and were provided instructions for a workaround. |  |  |  | Nationwide Distribution --- including states of Florida, New Jersey and North Carolina. | Terminated | October 22, 2009 | 485 |  | Software | Output/Calculation | Instructions for workaround | Safety Notice/Insructions
Z-0170-2009 | 49128 | No | IWJ | Manual radionuclide applicator system. | Radiology | Radiology | 510(K) Exempt | Alpha-Omega Services, Inc. Product No:NFW0001-010 |  |  | November 15, 2007 | 2007 | Alpha Omega Services Inc | There is a possibility that the Flexiguide Needle tip is defective. There is the potential for the needles tip to leak, thereby breaking the sterile boundary and possibly contaminating other devices with body fluids. | Nonconforming Material/Component | Notification of customers initially began November 15, 2007 based on a cursory visual inspection of the defective needle tips. The initial letter advised that there is a possibility that the Flexiguide Needle tip is defective in that there is the potential for the needle's stylet or an HDR source wire/cable to puncture through the defective needle tip during implant or treatment and penetrate the surrounding tissue thereby injuring the patient and/or breaking the sterile boundary. Further distribution and/or use of these products should not take place. The recall notification was updated on November 27,2007 once results from actual testing were available. The letter advises that there is a potential for the needle's tip to leak, thereby breaking the sterile boundary and possibly contaminating other devices with body fluids. Further distribution and/or use of these products should not take place. Customers are to return the products for credit. A response form is provided. Contact Alpha-Omega Services, Inc. at 1-800-346-7894 for assistance. |  | 40 | 40 pkgs (10 each pkg) | NV, AZ, CA, MI, UT, MO, CO, and OH | Terminated | December 09, 2008 | 390 | 198 | Not_Computer
Z-0171-2009 | 49131 | No | CIT | Aspartate amino transferase (AST/SGOT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid AST (SGOT) Reagent Set, Manufactured by Pointe Scientific, Inc., Canton, MI in 150 ml (Product no. A7561-150) and 450 ml (Product no A7561-450). For the quantitative determination of aspartate amoinotransferase in human serum |  |  | August 25, 2008 | 2008 | Pointe Scientific, Inc. | The R1 component may be contaminated with Serratia liquefaciens, resulting decreased absorbances and failure of the reagent to produce test results. | Material/Component Contamination | Consignees were notified via recall letter dated 8/22/08 and mailed 8/25/08 to examine their inventory, dispose of any recalled product found, and to notify the firm for replacements. Contact Pointe Scientific Inc., at 1-800-757-5313 for assistance. |  | 109 | 109 kits | Nationwide, Bolivia, Ecuador, Greece, Indonesia, Mexico, Nepal, Philippines, Romania and Serbia. | Terminated | November 27, 2009 | 459 | 134 | Not_Computer
Z-0183-2009 | 49540 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Vacora Biopsy Vacuum Assisted Biopsy Probe; Catalog numbers VB10118 and VB10140. |  |  | September 12, 2008 | 2008 | Bard Peripheral Vascular Inc | Some of the Vacora Biopsy Vacuum Assisted Biopsy Probe thumb wheels from these lots may fracture when fired using the prime/pierce option. | Device Design | Product Recall Notification Letter sent dated September 11, 2008 Attention: Risk Managers "All VB10118 and VB10140 catalog numbers with expiry dates prior to September 2008 are now beyond their expiration date. In accordance with accepted medical practice, they should have been discarded and not used. Do not use or further distribute any affected products. Once the product affected by the recall has been removed from your inventory; fill out the recall &amp; effectiveness check form, call Silvia Carrillo at 1-800-321-4254 ext 2727 to issue your with either a Return Authorization Number or Consignment Recall Number. fax the form to BVP at 1-800-440-5376 |  | 225 | 225 | Nationwide Distribution --- including states of AL, AR, CA, FL, KY, MA, NY, OH, SC, TN, TX, VA, and WA. | Terminated | September 19, 2012 | 1468 | 225 | Not_Computer
Z-0182-2009 | 49514 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Circlip component of the Stryker Flat Panel and Navigation Arm System. Stryker Flat Panel and Navigation Arm System are intended for use as a ceiling-mounted device to support a medical grade flat panel monitor or camera in the patient area. The Circlip within the flat panel and navigation arms is part of the mechanical stop designed to prevent the arm from rotating infinitely. It also prevents the cables routed through the arm from becoming twisted during routine movement and rotation of the arm. |  |  | September 05, 2008 | 2008 | Stryker Communications Corp | Circlip component used to suspend flat panel and navigation arm system may become dislodged. | Device Design | Notification letter was sent to consignees September 5, 2008. Field service will be performed at all affected sites. Old Circlips will be replaced with new clamps that do not allow the arm to fall. Contact Stryker Communications Corp. at 1-972-410-7310 for assistance. |  | 705 | 705 units | Nationwide and International Distribution | Terminated | October 02, 2009 | 392 | 705 | Not_Computer
Z-0181-2009 | 49464 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet brand Modular Microplasty Cup Inserter, 3/8" thread, Biomet Orthopedics, Inc., Warsaw, IN; Model 31-400600. Non-powered, hand-held instrument intended for use as an impactor during orthopedic surgery. |  |  | July 28, 2008 | 2008 | Biomet, Inc. | The pin and clip may fracture during use. | Device Design | Consignees were notified by recall letter dated 7/28/08 to remove the recalled lots from use and to return them. Contact Biomet, Inc. at 1-800-348-9500 or 1- 574-372-3983 for assistance. |  | 85 | 85 | Nationwide, Australia, Belgium, Canada, Denmark, New Zealand and Sweden. | Terminated | September 22, 2009 | 421 | 85 | Not_Computer
Z-0173-2009 | 49393 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IDS 25-Hydroxy Vitamin D EIA Enzymeimmunoassay for the quantitative determination of 25-hydroxyvitamin D and other hydroxylated metabolites in serum or plasma. Ref: AC-57F1 Lot 60897 Exp 2009-08; Immunodiagnostic Systems Ltd, Unit 10, Boldon Business Park, Boldon, Tyne &amp; Wear, NE35. 9PD, UK. Enzymeimmunoassay for the quantitative determination of 25-hydroxyvitamin D and other hydroxylated metabolites in serum or plasma |  |  | August 02, 2008 | 2008 | Immunodiagnostics Systems Inc | A small number of customers have experienced low absorbance values when using kits from this batch. | Nonconforming Material/Component | All customers who received 25-Hydroxy Vitamin D EIA lot # 60897 were contacted by mail. Customers were instructed to sign a confirming document for having received the notification letter and return via fax. Customers are requested to contact Immunodiagnostic System Inc's office if they have any kits of 60897 remaining in stock. Rather than replacing only the biotin buffer as stated in the letter dated August 2, 2008, Immunodiagnostic Systems Inc. will replace the kit per the letter dated August 8, 2008. Contact Immunodiagnostic System Inc. at 1-480-836-7435 for assistance. |  | 691 | 691 kits | Nationwide | Terminated | August 13, 2012 | 1472 | 691 | Not_Computer
Z-0196-2009 | 49549 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 26" Vision Elect HDTV Surgical Viewing Monitor, model number 0240030960, Manufactured by Stryker Endoscopy San Jose |  |  | September 22, 2008 | 2008 | Stryker Endoscopy | Knob parts stick together, resulting in poor functionality of the knob. | Device Design | Urgent: Device Correction letters dated 9/22/2008 were sent to Sales representatives and all accounts. Sales reps were also notified via email. Each Sales Rep was provided rework instructions, a list of affected accounts and serial numbers, and new knobs to install on affected units. Customers were asked to inspect their inventory for the part number, they were informed that their sales rep would be responsible for replacing the rotary knobs on all affected units, and that they should contact their sales rep to schedule the knob replacement. |  | 2265 | 2265 devices | Worldwide Distribution --- USA including states of Alabama, Arkansas, Arizona, California, Colorado, Connecticut, DC, Delaware, Florida, Georgia, Hawaii, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Michigan, Minnesota, Missouri, Montana, North Carolina, North Dakota, Nebraska, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, West Virginia, and Wyoming, and countries of Australia, Canada, China, England, France, Germany, Greece, Iberia, India, Israel, Japan, Korea, Mexico, Netherlands, Poland, Portugal, Scandinavia, Singapore, South Africa, Switzerland, and Taiwan. | Terminated | December 21, 2010 | 820 | 2265 | Not_Computer
Z-0199-2009 | 49565 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Smallbore T-Port Extension Set with Female Luer Lock Connector and Male Luer Lock T-Port with Ultrasite Injection Site 5 in. The product is shipped 100 units per carton. |  |  | September 17, 2008 | 2008 | B. Braun Medical, Inc. | Product was assembled incorrectly and connected to the wrong part. | Employee error | The recalling firm issued an Urgent - Medical Device Recall letter dated 9/17/08 to the distributors informing them of the problem and the need to return the product. The letter states that further distribution of the product should be immediately discontinued. Utilizing the "Product Removal Acknowledgement" form, customers should combine their inventory with the inventories reported by their customers and record them on the form. If affected products and associated lot numbers are currently in their possession, customers should call Customer Support Department to arrange for shipping and replacement product. A Customer Support Rep will provide instructions for handling affected product. |  | 3900 | 3900 units | Nationwide Distribution --- including states of CA, FL, NC, NJ, and TX. | Terminated | May 05, 2009 | 230 | 3900 | Not_Computer
Z-0186-2009 | 49446 | No | MHY |  |  | Neurology | PMA | Medtronic DBS" 3387 Lead Kit for Deep brain stimulation, DBS Leads: Proximal 40 mm, Distal 12.0 mm, Medtronic, Inc. |  |  | June 14, 2008 | 2008 | Medtronic Neuromodulation | Medtronic has received reports of DBS" leads being damaged at the proximal connector end of the lead when the lead cap is used in the implant procedure. The lead damage is the result of excessive torsion and / or tensile lading of the proximal connector end of the lead. Data also suggest that excessive torque while tightening or loosening the setscrews may twist the setscrew connector block and d | Process design | Consignees were sent a "Medtronic Urgent: Medical Device Correction" letter 8/2008. The letter addressed to Healthcare Professionals, described the product involved, explanation of the issue, recommendations for Implant and included an attachment "Updated Instructions for Using the Lead Cap". Contact Medtronic Neuromodulation at 1-800-707-0933 for assistance. |  | 3939 | 3939 | AL, AZ, AR, CA, CO, CT, FL, DC, GA, HI, ID, IL , IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE , NV, NH, NJ, NM, NY, NC, ND, OH, OR, OK, PA, RI, SC, SD, TN, TX, UT , VT, VA, WA, WV, and WI.ALGERIA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CROATIA, CYPRESS, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JORDAN, LEBANON, LUXEMBOURG, MEXICO, MOROCCO, NETHERLANDS, NORWAY, POLAND, PORTUGAL, PUERTO RICO, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEEDEN, SWITZERLAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | December 31, 2011 | 1295 | 26313 | Not_Computer
Z-0174-2009 | 49360 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Premier Value Lancing Device. UPC Code: 840986016707, The Package Includes: 1 Lancing Device, 1 Clear Cap and 5 Premier Value Ultra Thin Lancets, Distributed by: Chain Drug Consortium, LLC, Boca Raton, FL 33431. |  |  | July 18, 2008 | 2008 | Facet Technologies | The Alternate Site Cap was not included in the package as indicated on the labeling. | Error in labeling | The consignees were notified by letter on/about 7/18/2008. The Return Material Authorization number was communicated to the customer. They were asked to place the items on hold and Can Am will obtain any quantities. The Returned Product Mechanics Process will: Inventory will be sent back to Facet in two stages, Inventory on hand at Cimetra will be sent back to Facet, Inventory received back from customers will be sent back to Facet in one shipment 30-60 days after notice. Inventory received back after this second shipment will be destroyed. |  | 2976 | 2976 units | Nationwide--GA and FL. | Terminated | August 24, 2009 | 402 | 2976 | Not_Computer
Z-0189-2009 | 49456 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia T6 System; a dual-detector variable angle gamma camera with a six slice CT scanner; Siemens Medical Solutions USA, Inc., Molecular imaging, Hoffman Estates, IL 60192. The system is used to perform CT scans and nuclear imaging studies with the same instrument, obtaining attenuation corrected images and providing registration of anatomical and physiological images within the patient's anatomy. The Symbia T6 System is used for advanced oncology, neurology and cardiology applications. |  |  | September 29, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | The protective plastic cap over the CT gantry power switch on the Line Connection Box may come loose, exposing the energized electrical contacts within the switch, thereby causing an electric hazard. | Nonconforming Material/Component | Siemens sent Customer Safety Advisory Notice/Urgent Field Correction Recall letters dated 9/29/08 to all users of the Symbia T6 and Symbia T16 Systems. The accounts were informed of the potential risk that the protective plastic cap over the CT gantry power switch could come loose, causing an electrical hazard. The customers were advised to avoid contact with the gantry power switch until the protective plastic cap has been replaced, and to leave the power switch on at the line connection box and cycle the CT gantry on and off by toggling the CT gantry circuit breaker within the circuit box. They also were advised that a Siemens Customer Service Representative would contact them within 90 days to schedule a corrective update to their Symbia system free of charge. The accounts were requested to complete and return the enclosed mandatory fax-back form to 1-865-218-3019, acknowledging the receipt and understanding of the 9/29/08 letter and providing their Symbia System model number and serial number. Contact Siemens Medical Solutions USA, Inc. at 1-800-767-2313 for assistance. |  | 11 | 11 units | Nationwide to the District of Columbia, Florida, Georgia, Illinois, Indiana, Missouri, Ohio, and Texas, and internationally to Australia, Brazil, Canada, Great Britain, India, Italy, Japan, Norway and Switzerland. | Terminated | December 09, 2009 | 436 | 21 | Not_Computer
Z-0162-2009 | 49566 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GlobTek, Inc. Power Supply for Medical Use Part No: TR9CE4000LCP-Y-MED MODEL: GTM21097-5012, Made in China. |  |  | August 07, 2008 | 2008 | Smith And Nephew, Inc. Wound Management Division | External battery chargers used with the Version 29 VISTA Negative Pressure Wound Therapy pumps are failing to properly charge the pump's battery. | Nonconforming Material/Component | Each consignee was notified by recall letter (Field Correction Notification) dated 8/7/08 via FEDEX overnight. Within five (5) working days of the date of the customer notification a Smith &amp; Nephew representative will contact the customers to assess their inventory and schedule delivery of replacement battery chargers. Contact Smith &amp; Nephew, Inc. (Customer Care Center) at 1-800-876-1261 for for assistance. |  | 2550 | 2,550 | Nationwide within U.S. and internationally to United Kingdom, New Zealand, Germany, Canada and Australia. | Terminated | October 01, 2009 | 420 | 2550 | Battery |  | Replace chargers | Remove or Replace
Z-0200-2009 | 49581 | No | IZO | Diagnostic x-ray high voltage generator. | Radiology | Radiology | 510(K) Exempt | Monitor suspensions used with Advantx Legacy, RFX Classical, SFX-90, Model # 46-240838G2, manufactured after March 8, 1985 for Advantx Legacy, RFX and SFX systems. |  |  | March 09, 2007 | 2007 | GE Healthcare | GE Healthcare has recently become aware of a potential risk of monitor dropping associated with the monitor suspension of the Advantx system that may impact patient safety. It has been reported that suspensions manufactured after March 8, 1995 did not have the required thread locking agent applied to the setscrew, allowing the setscrew to back out over time, which can lead to the monitor dropping | Component change control | Consignees were sent a GE Healthcare "Product Safety Notification" letter dated September 13, 2007. The letter was addressed to Hospital Administrators/Managers of Radiology/Radiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Corrective Action and Contact Information. The letter recommends that customers move the monitor suspension to an open area in the room where viewing is still possible and make sure to minimize the movement of the monitor until a Qualified GE Healthcare Field Engineer has inspected the suspension. The letter also states that a Qualified GE Healthcare Field Engineer will schedule a field visit to inspect the monitor suspension. Customers should contact their sales or service rep is they have any questions. |  | 21 | 21 | Worldwide Distribution --- USA including states of AZ, CA, FL, GA, IA, MD, MO, NC, NY, OH, OK, TN, TX,WI, and LA, and country of Slovak Republic. | Terminated | August 24, 2012 | 1995 | 21 | Not_Computer
Z-0202-2009 | 49578 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Oxylog 3000 Emergency and Transport Ventilator Continuous Respirator with software version 1.10 pre-installed; Catalog number 2M86300. |  |  | October 03, 2008 | 2008 | Draeger Medical, Inc. | May experience an interruption of ventilation for approximately 5 seconds | Software design | The firm issued an Urgent - Medical Device Recall letter dated September 2008. Software version 1.10 may under specific conditions experience an interruption of ventilation for approximately 5 seconds. During that time period, the patient airway pressure may drop to zero. This phenomenon was identified by Draeger engineers and to date has not been reported by any user. The firm's investigations revealed that when ventilating in BIPAP mode with the pressure rise time (Ramp) set to the highest setting and the expiration time set to exactly 0.5 seconds, the inspiration phase can coincide with the internal auto zero function of the device. Under these circumstances a pressure spike will occur upon reaching Pmax +5mbar (cmH2O) which with software version 1.10 will lead to the inspiratory phase stopping. Ventilation will resume after approximately 5 seconds. The device will generate a Paw Measurement INOP alarm. Until the new software is available, customers are being instructed to avoid the use of the following combination of settings: 1) BIPAP mode; 2) Ramp = highest, and 3) Texp = 0.5 seconds. |  | 50 | 50 units | Nationwide Distribution --- including states of AZ, CA, FL, IL, IN, KY, LA, ME, MI, MO, NC, NE, NJ, NY, OH, PA, TX, UT, VT, WA, WV, and WY. | Terminated | October 16, 2009 | 378 | 50 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0153-2009 | 49444 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Allez Spine Laguna Pedicle Screw System; Catalog Number: IT-FN1003 |  |  | August 21, 2008 | 2008 | Allez Spine, LLC | This action is being taken voluntarily as a result of a Warning Letter issued to AlIez Spine LLC by the Food &amp; Drug Administration. The Company decided it could better address the FDA's concerns regarding Current Good Manufacturing Practice requirements of the Quality System without any product in the market so it could focus all of its efforts on responding to the FDA. | Pending | Allez Spine issued a field notification beginning August 21, 2008 via FedEx to all its hospital users, distributors and field representatives informing them of their decision to recall the Allez Spine Laguna Pedicle Screw System, Allez Spine Del Mar Pedicle Screw System and the Allez Spine Cross Connectors as a result of a Warning Letter issued to AlIez Spine LLC by the Food &amp; Drug Administration.The recall letter requested their customers to return their entire inventory of AlIez Spine manufactured products as soon as possible, gave instruction for product return and provided a contact person and telephone number for information regarding the recall. |  | 6472 | 6472 units for all systems | Nationwide Distribution --- including states of AZ, CA, CO, GA, IL, NV, TN, TX, &amp; UT. | Terminated | January 19, 2010 | 516 | 6472 | Not_Computer
Z-0118-2009 | 49403 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iSite PACS (Picture Archiving and Communication System), software versions 3.6.28.x and 4.1.x. |  |  | August 29, 2008 | 2008 | Philips Healthcare Informatics, Inc. | Two defects have been identified: 1) potential to display a patient on the canvas page that is different from the patient whose images are displayed on diagnostic monitors when using Conference Presentation States and when opening two or more studies from a folder or an exam worklist and 2) potential to miscalculate measurements when Pixel Spacing and Imager Pixel Spacing DICOM tag values are bot | Software design | A notification letter (Urgent Product Correction) dated 8/29/08 was sent via certified mail to all customers alerting them of the issue, requiring posting of notification. Philips Healthcare Informatics, Inc. strongly recommend all customers advise their iSite Radiology users not to open multiple patient exams simultaneously from a folder or worklist using Conference Presentation States. A software correction is being worked on, and should be available by end of Q3 2008 for version 4.1.45.2. A mandatory upgrade is required for all customers who are currently using iSite version 4.1.45.2 and former versions. For specific questions regarding these issues, contact the Customer Advocate (CARE) or Customer Care at Philips Healthcare Informatics, Inc. at 1-877-328-2808. In your communication with Customer Care at Philips Healthcare Informatics, Inc. , please refer to this notification and associated FCO 22945704-08292008-004. |  | 27 | 27 units. | Worldwide Distribution including USA (MN, CO and PA) and countries of Italy, Germany, Netherlands, Sweden, France, Portugal, Switzerland, Thailand, Japan and Australia. | Terminated | November 30, 2010 | 823 | 27 | Software | Output/Calculation | Software update | Software Update
Z-0132-2009 | 49520 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Respiratory Gating System which includes: Pulmonary Digital Assembly Kit and Pulmonary Toolkit Oncology Pro, Model #455011203131 and Model #455011203191. |  |  | September 22, 2008 | 2008 | Philips Medical Systems (Cleveland) Inc | Images not correlated: A leak between the Tube Interface and the Outlet Tube of the transducer may result in a failure to produce respiratory correlated images. | Process design | On 9/8/2008 Philips Medical Systems sent recall notification letters to their customers requesting they inventory their stock. They sent instructions to verify that the product is working properly. They are also sending Field Service Engineers to each customer to conduct field test. |  | 507 | 507 units | Worldwide distribution: USA, Australia, Austria, Bahrain, Belgium, Canada, China, Czech Republic, Denmark, England, Finland, France, Germany, Israel, Italy, Japan, Jordan, Korea, Lithuania, Malaysia, Netherlands, Norway, New Zealand, Puerto Rico, South Africa, Saudi Arabia, Spain, Switzerland, Sweden, and Thailand. | Terminated | March 01, 2012 | 1256 | 507 | Not_Computer
Z-0180-2009 | 49406 | Yes | LFK | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | VCARE (Vaginal-Cervical-Ahluwalia's-Retractor-Elevator) LARGE, Uterine Manipulator, REF 60-6085-102. Made in USA, CONMED CORPORATION, UTICA, NEW YORK USA 13502. The ConMed VCAREÂż Retractor/Elevator is indicated for manipulation of the uterus and injection of fluids or gases during laparoscopic procedures such as Laparoscopic Assisted Vaginal Hysterectomy (LAVH), Total Laparoscopic Hysterectomy (TLH), minilap, laparoscopic tubal occlusion, or diagnostic laparoscopy and also maintains pneumoperitoneum by sealing the vagina once colpotomy is performed. |  |  | September 08, 2008 | 2008 | ConMed Corporation | The firm received several complaints sighting detachment of the balloon at the distal end of the shaft. It was also noted that the incidence of the forward cup slipping off the shaft and being retained in the vaginal canal had increased. | Process change control | The domestic recall notices were mailed on 9/08/08 by USP Priority Mail pouches. The international notices were sent via FedEx International Priority on 9/08/08. The letter described the problem, the health risk and requested that customers inventory the product and return them back to the firm. Customers are requested to fax the reply form to the firm. Contact ConMed Corporation at 1-800-765-8375, ext 3237 for assistance. |  | 32453 | 32,453 (total devices) | Nationwide, Australia, Belgium, Brazil, Chile, Ecuador, France, and Switzerland | Terminated | August 29, 2012 | 1451 | 32453 | Not_Computer
Z-0191-2009 | 49497 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Osteosynthesis Reduction Spoon; |  |  | August 14, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker has become aware of the potential for breakage of the Reduction Spoon designed prior to 2003. | Device Design | Recall notification letters were sent on August 14, 2008 to all branches and hospitals via Federal Express. Questions are to be directed to 201-831-5718 or a Trauma team member listed in the letter. |  | 1070 | 1070 units | Nationwide | Terminated | May 29, 2012 | 1384 | 1070 | Not_Computer
Z-0130-2009 | 49558 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Oscor Adelante Luer-Lock Peel Away Introducer Set, for introduction of diagnostic or therapeutic devices into the body. |  |  | July 03, 2008 | 2008 | Oscor, Inc. | Difficulty breaking the sheath hub and subsequently to peel the sheath off for Introducer Set, Adelante size 7F. | Process design | Oscor notified the Direct Account by letter on 07/03/2008. Direct account was asked to return all unused Introducers back to Oscor. All returned products should be labeled with the Return Goods Product number. |  | 396 | 396 | Product was distributed to one direct account in MN. | Terminated | December 09, 2008 | 159 | 396 | Not_Computer
Z-0150-2009 | 49056 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK Express Automated External Defibrillator (AED). Labeling on the device states LIFEPAK CR Plus/LIFEPAK EXPRESS Defibrillators with ADAPTIV Biphasic Technology; Medtronic Emergency Response Systems, Inc. 11811 Willows Road NE Redmond, WA. The battery operated, small portable device is primarily deployed in Public Access Defibrillator (PAD) environments where the expected use for any one device is once every 20 months. Product is capable of 30 full discharges. |  |  | July 24, 2008 | 2008 | Physio Control, Inc. | The device may not power on although it indicates it is ready for use and would not be able to provide defibrillation therapy. | Nonconforming Material/Component | On 07/24/08, the firm sent an "URGENT MEDICAL DEVICE RECALL" letter with a confirmation sheet via certified return receipt to customers. The letter notified the customer of the recall and the hazard associated with the recall and provided instructions to customers: 1) Immediately verify the AED powers on and voice prompts begin. After verification, press and hold the On/Off button for approximately two seconds to turn the AED off. 2) Keep the AED in use and perform monthly inspections that include the additional AED power on test. Perform a monthly inspection which is consistent with the device labeling. 3) If "OK" is visible on the Readiness Display and AED powers on, it is ready for use. 4) If, at any time, the AED does not power on or if any other indicator displays (i.e. battery or CHARGE-PAK symbol, attention symbol or wrench symbol) immediately call Physio-Control Technical Support (1-800-442-1142). The letter advises customers a Physio-Control representative will contact them within 60 days to arrange a replacement. A confirmation sheet was sent with the letter; to be returned to Physio-Control, Inc. For specific questions or additional information, contact Physio-Control, Inc. at 1-800-442-1142 weekdays between 6:00 a.m. and 4:00 p.m. (PST) or at www.physio-control.com. |  | 2230 | 2,230 Units | Worldwide including USA (and Puerto Rico) and countries of Argentina, Australia, Bahamas, Barbados, Brazil, Canada , Chile, Colombia, Dominica, Germany, Hong Kong, Hungary, Jamaica, Japan, Mexico, Netherlands, New Zealand, Singapore, Uruguay, Venezuela, and Virgin Islands. | Terminated | August 08, 2013 | 1841 | 39020 | Not_Computer
Z-0205-2009 | 44852 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 base, 110/120V; Model 801763. For use for up to 6 hours in the extracorporeal circulation of blood for arterial perfusion, regional perfusion and cardiopulmonary bypass procedures. |  |  | July 18, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | Local pump controls including start/stop button, select button and manual button may not function. | Component design/selection | Consignees were notified by recall letter dated 7/18/08 and were provided with revised operating instructions telling them to revert to the central control module should these problems occur. |  | 626 | 626 | Nationwide, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Viet Nam. | Terminated | July 07, 2009 | 354 | 926 | I/O |  | Instructions | Safety Notice/Insructions
Z-0207-2009 | 48174 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory versions 5.3, 5.4, 6.1, 6.2 and 6.3. |  |  | April 17, 2008 | 2008 | Sunquest Information Systems, Inc. | Sunquest has identified a problem with IGO interfaces when performing an Online File Cleanup (OFC) that can lead to the test results of one patient filing to another patient when the order codes are the same. Performing an OFC prior to all results being assigned an accession number may cause the patient results to remain in a temporary file where they can be incorrectly associated to the wrong pat | Component design/selection | Customers were sent a preliminary notice on 4/17/2008 followed by customer notice PSN-08-04 on 05/07/2008 via email. The notice advises that this notice is follow-up to the Preliminary notice distributed on 4/17/08 and applies to all clients who have an Instrument Generated Order (IGO) interface.Sunquest has identified a problem with IGO interfaces when performing an Online File Cleanup (OFC) that can lead to the test results of one patient filing to another patient when the order codes are the same. Performing an OFC prior to all results being assigned an accession number may cause the patient results to remain in a temporary file where they can be incorrectly associated to the wrong patient. When this occurs, Sunquest generated reports and inquiry contains the incorrect results.The letter provides the various steps that cause the issue to occur. Sites Potentially Affected by this Problem:Sunquest Laboratory using Instrument Generated Orders (IGO) Interface, Version 5.3 and laterAction Sunquest Information Systems is taking to eliminate this Problem:Change Request 1-62VWLJ has been created to address this issue. This will be available as part of Sunquest Laboratory Instrument Build version 1.6.0023 and later.To request the software correction, customers can create a Service request via the Client support Webpage at www.sunquestinfo.com/supportweb. The notice also provides temporary procedural instructions:Effective immediately, do NOT perform Online File Cleanup (OFC) on any IGO instrument interface methods.-Manual OFC should not be performed using functions OFC, or Instrument Maintenance (IX or IXM), option 11 for IGO interfaces only.-If using function Instrument Maintenance (IX or IXM), option 1, 2, 2, Method, 2 (Automatic File Cleanup). This should be set to "N". Then stop and re-start the Instrument Results Processor.Note: Online file cleanup for non-IGO interfaces needs to be performed as usual.-Pending and Overdue Logs need to be monit |  | 419 | 419 clients had the affected code | Worldwide Distribution --- including USA and countries of Canada, Bermuda, Ireland, United Kingdom, United Arab Emirates, and Saudi Arabia. | Terminated | December 05, 2011 | 1327 | 419 | Software | Output/Calculation | Software update | Software Update
Z-0223-2009 | 49466 | No | CFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Glucose (HEX) Hitachi; 3-HG920-L. (Clinical chemistry) |  |  | September 10, 2008 | 2008 | Pointe Scientific, Inc. | The inability of the product to maintain stated performance specifications through the stated shelf life. | Process control | Consignees are being notified by letter dated 9/10/08 which will be sent out via fax, email or certified mail to all direct account and drop-ship customers. Customers are advised to dispose of any recalled product found and to consider test results obtained with these lots questionable. Contact Pointe Scientific, Inc. at 1-800-757-5313 or 1-734-487-8300 for assistance. |  | 63 | 63 | Nationwide, Canada, China, Romania and Trinidad. | Terminated | November 27, 2009 | 443 | 131 | Not_Computer
Z-0226-2009 | 49470 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BD GeneOhm MRSA 48 ct, Catalog #441244. IDI-MRSA assay is a qualitative in vitro diagnostic test for the direct detection of nasal colonization by methicillin-resistant Staphylococcus aureus (MRSA) to aid in the prevention and control of MRSA infections in healthcare settings. The test performed on the Smart Cycler instrument with a nasal swab specimen from patients at risk for colonization, utilizes polymerase chain reaction (PCR) for the amplification of MRSA DNA and fluorgenic target-specific hybridization probes for the detection of amplified DNA. IDI-MRSA assay is not intended to diagnose MRSA infections nor to guide or monitor treatment for MRSA infections. Concomitant cultures are necessary only to recover organisms for epidemiological typing or for further susceptibility testing. |  |  | July 21, 2008 | 2008 | BD Diagnostics (GeneOhm Sciences, Inc) | BD Diagnostics GeneOhm (BDDG) initiated the recall after they identified that the product has the potential to identify a patient as falsely positive for colonization with methicillin resistant Staphylococcus aureus.The potential health hazard related to a false positive result from this assay would be that a patient may be misclassified as having MRSA colonization when in fact they do not have | Nonconforming Material/Component | BD Diagnostics GeneOhm (BDDG) initial recall notification was made on 07/21/2008 by telephone and followed up with written notification on 07/22/2008. Domestic customers were advised to check their stock to determine if they have inventory remaining of the recalled lot. If so, BD GeneOhm Technical Service will arrange for return of unused product from the customer. The were provided with a Recall Verification form to complete and return. The customers were instructed if they have remaining tests from the affected referenced lot, please discontinue use immediately. They were also told that BD GeneOhm recommend that they review the results of samples tested using this lot of product. Look for any trends outside of the normal expected results, including any changes in MRSA positivity rates compared to your normal prevalence. If they have any questions or concerns following the review, please contact BDDG Technical Service to arrange for consultation with our Medical Director.An telephone number 1-858-334-6356 was given for customers to contact BDDG Technical Service. |  | 60 | 60 kits (48 test/kit) =2880 test | Nationwide to FL, NJ, NY, OH &amp; PA and Internationally to Belgium | Terminated | March 04, 2010 | 591 | 3460 | Not_Computer
Z-0279-2009 | 49392 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Outback LTD Re-Entry Catheter, Sterile, Single Use. The Cordis LuMend OUTBACK LTD Re-Entry Catheter is indicated for placement and positioning of guidewires within the peripheral vasculature. The cannula is a nitinol needle that is used to re-enter the true lumen of the vessel. |  |  | August 20, 2008 | 2008 | Cordis Corporation | The Cordis LuMend Outback LTD Re-Entry Catheter - separation rate of the cannula to the deployment slide of the handle assembly is higher than anticipated. The cannula is unable to be retracted into the device after deployment due to a separation of the inner key from the cannula. | Other | U.S. Customers - A Notification of Recall Letter with an Acknowledgement Form was sent overnight to 5 individuals in each account on 08/18/2008. For non-U.S. customers, email notification with return acknowledgement is used to notify the affiliate distributors, who then notify and reconcile the customers in those countries (08/18/2008). Contact Cordis special Customer Service at (800) 551-7683 for assistance. |  | 11893 | 11,893 | Worldwide | Terminated | December 01, 2008 | 103 | 11893 | Not_Computer
Z-0192-2009 | 49515 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hydrophobic HALO360 Filter that is |  |  | September 18, 2008 | 2008 | Barrx Medical, Inc. | Some units may contain the wrong filter, which does not have the proper lock and may result in a leak. | Mixed-up of materials/components | The recall was initiated on September 18, 2008. All users and international distributors were informed by letter. The replacement filters were sent on 9/19/2008. Contact Barrx Medical, Inc. at 1-408-745-8000 for assistance. |  | 650 | 650 units | Worldwide Distribution, including USA, Austria, Australia, Belgium, Canada, France, Netherlands, Switzerland, Sweden, and UK . | Terminated | April 26, 2011 | 950 | 650 | Not_Computer
Z-0227-2009 | 49484 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Hipstar V40 Femoral Stem |  |  | August 20, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | The warning label "do not use with heads more than +10, offset" does not appear on 135 degree stem boxes. This warning is contained in the instructions for use. | Labeling mix-ups | Recall notification letters were sent by Fed Ex on August 20, 2008 to Stryker branches/agencies and Hospital Risk Management departments. Letters to Branches stated the reason for the recall and the potential hazard. Both letters state that it is recommended that physicians follow patients in a manner that is consistent with a physicians routine standard of care for total hip arthroplasty. Contact Stryker Howmedica Osteonics Corp. at 1-201-831-5000 for assistance. |  | 77 | 77 units | Nationwide | Terminated | June 01, 2009 | 285 | 77 | Not_Computer
Z-0228-2009 | 49492 | No | CAX | Compensated thorpe tube flowmeter. | Anesthesiology | Anesthesiology | 510(K) Exempt | Bag Tee Assembly Model P1407A . An accessory to the dental flowmeter which has a connection function for complete delivery of mixed gases to the breathing circuit portion of the gas scavenging apparatus. The Bag Tee assembly features a non-breathing valve and an Emergency Air Intake located on the bag tee which comply with the American Dental Association guidelines. |  |  | August 26, 2008 | 2008 | Porter Instrument Co Inc | inverted check value on the bag tee assembly which can cause leaking of mixed N2O/O2 gas | Employee error | On 9/2/08, the firm issued an Important Field Correction Notice to their customers informing them of the problem and how to determine if their product was defective and needed to be returned. Contact Porter at 1-800-457-2001 and ask for Dental Customer and refer to the Bag Tee Field Correction Notice. |  | 220 | 220 units | Nationwide | Terminated | September 24, 2010 | 759 | 220 | Not_Computer
Z-0271-2009 | 49525 | No | LGW |  |  | Neurology | PMA | Precision Charger 1.0 (Model Number SC-5300) for the Precision Spinal Cord Stimulator System labeled as BIONICS SCS-CHARGING KIT CONTENTS Charger (1) Base Station (1) Belt Power Supply (1) Adhesive Kit ADVANCED BIONICS CORP 12740 San Fernando Rd. Sylmar, CA 91342 USA. The Precision Spinal Cord Stimulator System (Precision System) is indicated as an aid in the management of chronic intractable pain of the trunk and/or limbs, including unilateral or bilateral pain associated with following: failed back surgery syndrome, intractable low back pain and leg pain. The Precision System includes an implantable 16-output, multi-channel stimulator (IPG - Implantable Pulse Generator) with a rechargeable battery power source. The IPG, commonly implanted in the abdomen or buttock area (less frequently in the subclavicular area), can be configured to accept one or two leads. The leads are implanted in the epidural space, adjacent to the spinal column. The stimulation delivered at the distal end of the lead is intended to mask the pain signals. An externally used Charger is utilized to transcutaneously recharge the IPG (Implantable Pulse Generator) battery. Currently, Boston Scientific distributes Charger 2.0 (Model Number SC-5312). For improved heat management, Charger 2.0 is equipped with thermistor (temperature sensing component), designed to activate/deactivate charging at a set temperature. |  |  | September 22, 2008 | 2008 | Advanced Bionics Corp | Since April 2004, 27 of 8,769 (0.31%) patients have reported receiving second degree burns in the area of charging and 3 of 8,769 (0.03%) patients have reported receiving third degree burns in the area of charging while using Charger 1 .0. Some of these events were the direct result of the patient disregarding the instructions for use, such as sleeping with the Charger on or placing it directly on | Other | In the United States, notification will be sent to patients with the affected devices. Additionally, their physicians on record will be provided with the notification. The Company is currently planning to initiate this field action on September 22, 2008. The U.S. notification will be sent via Fed Ex or U.S. Postal Service certified mail. The letter advises that Boston Scientific Neuromodulation is voluntarily recalling its first generation charging device, Precision Charger 1.0 (Model Number SC-5300) for the Precision Spinal Cord Stimulator System. The Charger 1.0 devices are being recalled and will be replaced with the new second generation Precision Charger 2.0 (SC-5312, currently in distribution) that features a temperature limiter for improved temperature management. The Charger 1.0 can be used until replacement with Charger 2.0 by following the instructions on Page 2 of the letter. Customers are asked to complete the acknowledgement form at the end of the letter and return it to Boston Scientific in the return envelope provided. A representative from Boston Scientific will contact them upon receiving the completed form to schedule an in-person visit at their physician's office to exchange the device. Contact Boston Scientific Technical Serives at (866) 566-8913 for assistance. |  | 6549 | 6549 units | Nationwide and Internationally - Ireland, Korea, Mexico, The Netherlands, United Kingdom, Canada, Australia and Israel | Terminated | January 24, 2012 | 1219 | 6549 | Not_Computer
Z-0212-2009 | 49385 | No | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BioStar OIA FLU AB, Catalog 90007. Packaged 30 devices per tray. BioStar, Louisville, CO 80027 |  |  | July 02, 2008 | 2008 | Inverness Medical BioStar Inc. | Diagnostic kits for flu were distributed with incorrect components. | Mixed-up of materials/components | Consignees were notified by a Important Customer Notification letter on 7/2/2008 and told to re-examine patient records that indicated a negative result and to destroy any remaining recalled product. The accompanying record should be filled out and faxed to Technical Service. If consignees further distributed any kits, they were to forward the notification to these accounts to assure that patient results were reviewed, and destruction of remaining kits could be verified. For assistance, customers can call 800-637-3717. |  | 509 | 509 kits | Nationwide Distribution --- including state of IL. | Terminated | November 04, 2008 | 125 | 509 | Not_Computer
Z-0201-2009 | 49590 | [6 empty] | Vibe Technologies, VIBE machine, Vibrational Integrated Biophotonic Energizer. |  |  | October 01, 2008 | 2008 | VIBE Technologies | Medical device marketed without marketing approval for claims that include cures cancer, infections, depression. | PMA | Vibe Technologies notified their customers by letter on October 1, 2008 and were told to discontinue use as a medical device. This letter notification included: 1. A warning label to be placed on the Vibe machine stating that it is not a medical device and should not be used as such.2. An updated operation manual/users guide containing no medical conditions or human body claims.3. A certification to be signed by the user and returned acknowledging that they: received the letter; affixed the warning label to the Vibe machine; understand that the Vibe machine does not affect the structure/function of the human or animal body; will not promote the Vibe machine as a medical device; will remove any medical claims from their website; destroy any Vibe literature making medical claims. Approximately three weeks after the certified mailing, the firm will begin performing recall effectiveness checks. |  | 851 | 851 units | Worldwide | Terminated | May 21, 2012 | 1328 | 851 | Not_Computer
Z-0131-2009 | 49531 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 18 gage XTW needles, sold as a component in SafeSheath Hemostatic Tear-away Introducer System with Infusion Side Port kits or sold Bulk Non Sterile (BNS). |  |  | September 12, 2008 | 2008 | Thomas Medical Products Inc | Difficulty inserting guidewire through introducer needle. Needle hub is not tapered down to the needle cannula. | Process control | On 9/11/08, Thomas Medical Products issued letters to the distributors informing them of the problem, to cease distribution and await further instructions. Additional instructions to the distributors were to notify all their account via an advisory notice and request that the user not use the needle supplied with the kit and to replace with another commercially available sterile 18 gage introducer needle. |  | 246585 | 246,585 needles | Nationwide distribution via distributors in CA, MA, and MN. | Terminated | May 06, 2009 | 236 | 246585 | Not_Computer
Z-0278-2009 | 48739 | No | MRV | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | S/PÂż Orange 50g Glucose Tolerance Beverage, 10 oz glass bottles, 12 bottles/case, labeled in part ***Cardinal Health McGaw Park, IL 60085-6787 USA*** For prescription use only. Product is used in support of diagnostics testing for hyperglycemia and hypoglycemia. |  |  | April 23, 2008 | 2008 | Nerl Diagnostics LLC | Beverage may contain glass particles. | Other | The recalling firm notified distribution centers by letter dated 04/23/08 and advised of a single report of glass found in a bottle. The notifications further advised distributors to quarantine referenced product, notify their customers and return the enclosed reply form. Distributors were requested to return or destroy product for replacement. Please contact NERL Diagnostics, LLC at 1-401-824-2046 for assistance. |  | 25908 | 25908 bottles | Nationwide. | Terminated | November 26, 2012 | 1678 | 25908 | Not_Computer
Z-0137-2009 | 49426 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Bicor X-Ray System, Model number 9023136. |  |  | August 11, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Unintended movement of the system table and/or C arm. | Software design | Siemens Medical Solutions issued a letter to consignees on 8/11/2008 to inform them of the problem and a workaround should the issue occur prior to software revision. |  | 20 | 20 units | Nationwide. | Terminated | December 21, 2012 | 1593 | 975 | Computer
Z-0193-2009 | 49529 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis. Angiographic X-ray System. |  |  | August 22, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Patients could possibly fall off the table if not properly secured when: 1) moving them onto the patient table, 2) moving them on the patient table or 3) removing them from the patient table. | Labeling design | The recalling firm issued a Customer Safety Advisory Notice dated 7/04/08 (including an Addendum to the Operators Manual) which alerts the operator of possible injury to patients if they are not properly secured to the table during examinations and describes the recommended fastening procedures required to properly secure the patient to the table.For questions or additional information, contact Siemens Medical Solutions, Inc. at 610-448-4634. |  | 1530 | 1,530 units | Nationwide Distribution. | Terminated | September 23, 2009 | 397 | 1638 | Not_Computer
Z-0282-2009 | 50113 | No | JEY | Bone plate. | Dental | Dental | 510(k) | HT X-DRIVE SCREW, Part # 91-2416, W.Lorenz Surgical. Screws for use in the stabilization and fixation of mandibular fractures and mandibular reconstructive surgical procedures. |  |  | July 11, 2008 | 2008 | Biomet Microfixation, Inc. | The packaging for the 91-2416, 2.4X16MM HT X-DRIVE SCREW, Lot 493200 actually contained the 99-6577, 2.0X7MM FOSSA X-DRIVE SCREW. | Labeling mix-ups | An "Urgent Medical Device Recall Notice" was sent to customers on 07/11/2008 by certified letter. The letter described the problem, the risk and asked that customers discontinue use of the product and return them to Biomet. Also, customers were provided with a Inventory Reconciliation sheet and asked to fax it back to Biomet. Please contact Biomet Microfixation at 1-800-874-771 or 904-741-4400, ext. 448 for assistance. |  | 10 | 10 | Nationwide and Canada | Terminated | September 11, 2009 | 427 | 10 | Not_Computer
Z-0213-2009 | 49440 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Mobile Care Server used in conjunction with Web Viewer, Pocket Viewer, and Cellular Viewer WIT with L-WEB05 software. |  |  | July 16, 2008 | 2008 | GE Healthcare | Issue 1 -Possible failure of CIC Pro audible alarms: Requests for telemetry defaults from Mobile Care Server could cause excessive CIC Pro memory consumption. This excessive memory consumption may lead to potential loss of audio at the CIC or a system reset or reboot. Issue 2 - Mobile Care software/Mobile Viewers don not recognize ApexPro FH Telemetry transmitter: Mobile Care server does not ac | Software design | Consignees contacted via a GE Healthcare Urgent Medical device Correction letter dated 8/26/08 and supplemental instruction for installation of software" Mobile care server with software, L-WEB05 and iCentral with software, L-NET05 supplemental Instruction for installation". The letter was addressed to Healthcare Administrator/Risk manger, chief of Nursing and director of Biomedical engineering / IT. The letter also described the problem, affected Products,and Short Term Recommendations . |  | 292 | 292 | Worldwide Distribution --- USA including states of CALIFORNIA, IOWA, INDIANA, VERMONT, and WASHINGTON and countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CROATIA, CYPRUS, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TURKEY, and UNITED KINGDOM. | Terminated | November 24, 2015 | 2687 | 292 | Software | Alarm/Message | Software update | Software Update
Z-0273-2009 | 49554 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Gyrus ACMI, Inc. dissector PlasmaKnife (DPK) |  |  | October 01, 2008 | 2008 | Gyrus ACMI Corporation | Sterility may be compromised | Process control | Gyrus ACMI initiated a voluntary recall by an Urgent: Medical Device Recall letter on 10/01/08 informing accounts of the problem, cease any further use of the affected product, remove any of the affected product service representative to obtain a Return Goods Authorization, return the product, and complete the reply form and fax it to the Regulatory Affairs Dept. so that the response can be tracked. |  | 1340 | 1340 units | Worldwide Distribution --- including USA and countries of Canada, UK, Italy, Spain, Norway, Australia, and Taiwan. | Terminated | May 07, 2012 | 1314 | 1340 | Not_Computer
Z-0284-2009 | 49606 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | ReliOn Insulin syringes - 1cc, 31 G |  |  | October 09, 2008 | 2008 | Covidien LP | Mislabeled: Package labeled as an insulin syringe for use with U-100 insulin contains an insulin syringe for use with U-40 insulin. Risk of overdose of insulin. | Process design | Tyco Healthcare Group LP (Covidien) notified Can-Am Care, Corp. by letter sent by e-mail and by Federal Express on 10/9/08. The letter requests that they immediately inventory, quarantine any remaining stock they may have of the product, return, and notify their customers. Wal-Mart conducted a consumer mailing on October 14, 2008, and is posting placards in their stores. Can-Am Care, whose name appears on the label of the product has concurrently posted a recall notice on their website. |  | 471000 | 471,000 (4710 boxes of 100) | Nationwide | Terminated | July 06, 2010 | 635 | 471000 | Not_Computer
Z-0188-2009 | 49448 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Maquet Alphastar 1132.01AX/BX, 1132.02AX/BX, and 1132.03AX/BX Universal Operating Table |  |  | September 16, 2008 | 2008 | MAQUET Inc. | Table tilts: The Alphastar Surgical Table piston rod head may break and cause a full inclination of the table top during Trendelenberg positioning if there is insufficient lubrication. | Equipment maintenance | Maquet sent Recall Notification letters to all customers on September 16, 2008 by Federal Express. Letters inform customers of the potential problem with the Alphastar tables, a description of the failure and hazards involved. Hospitals are asked to confirm whether yearly maintenance by an authorized technician have been performed. If preventive maintenance has not been performed, customers are requested to contact their Maquet Service rep. A Customer Service Response form is included and customers are to complete if to acknowledge whether of not the facility is still in possession of an affected table. If so, the model number, serial number and facility name &amp; address is to be completed so that Maquet reps can visit the facility and perform the corrective action. |  | 141 | 141 in US | Nationwide | Terminated | January 12, 2010 | 483 | 141 | Not_Computer
Z-0203-2009 | 44851 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Base, 100/120V; Model 801763. For use for up to 6 hours in the extracorporeal circulation of blood for arterial perfusion, regional perfusion and cardiopulmonary bypass procedures. |  |  | July 18, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The pumps may exhibit blank or distorted displays, loss of local control and/or may have local control knobs that are difficult to use. | Other | Consignees were notified via letter dated 7/18/08 and instructed to use the central control module should the problem occur. Contact Terumo Cardiovascular Systems Corp at 1-800-521-2818 for assistance. |  | 626 | 626 | Nationwide, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Viet Nam. | Terminated | July 07, 2009 | 354 | 926 | Other | Device Operation | Notification | Safety Notice/Insructions
Z-0274-2009 | 49600 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | PurePoint System Operator's Manual, Catalog Number 8065751131, Rev. B for use with the Alcon PurePoint Laser, Catalog Number 8065750597 |  |  | September 26, 2008 | 2008 | Alcon Laboratories, Inc. | Alcon had identified that indications for use unapproved by the Food and Drug Administration are included in the PurePoint Laser System Operator's Manual, Catalog Number 8065751131, Revision B. | Employee error | An Alcon representative has been provided with a package of information for each customer site. All domestic customers are being provided an Urgent: Medical Device Correction letter describing the change to the PurePoint Operator's Manual and are being issued a revised PurePoint Operator's Manual (Revision C). The letter advises that Alcon Research, Ltd, has initiated a voluntary correction based upon unapproved indications for use identified within the Alcon PurePoint Laser System Operator's Manual, Revision B. The letter advises that the following indications for use have been updated or removed on page 1.14, Professional Operator's Information, in Revision C of the PurePoint Laser System Operator's Manual:Updated:"AMD; Wet or Dry to include macular degeneration" has been updated to read "Choroidal neovascularization secondary to age-related macular degeneration (AMD)"."Trabeculoplasty for treatment of Chronic/Primary Open Angle Glaucoma (COAG, POAG), Acute Angle ClosureGlaucoma (AACG), and refractory glaucoma" has been updated to read "Trabeculoplasty for treatment of Chronic/Primary Open Angle Glaucoma (COAG, POAG) and refractory glaucoma"Removed:Macular photocoagulation; including grid, focal, laser drusen scatter (panretinal)Transcleral cyclophotocoagulationIntra-ocular tumors; to include choroidal hemangioma, choroidal melanoma, retinoblastomaOtosclerotic hearing loss. No other changes to the manual have been made in this revision and no other Alcon products are affected by this correction. Per the letter, an Alcon Representative will provide customers with a new PurePoint Laser System Operator's Manual, Catalog Number 8065751131, Rev. C. and customers are return their existing PurePoint Laser System Operator's Manual, Catalog Number 8065751131, Rev.B. A signed confirmation of receipt of the letter is required from all customers. |  | 41 | 41 units | Worldwide Distribution --- USA including states of IL, WI, TX, MI, PA, NY, CT, SC, NC, LA, FL, GA, CA, OR, WA, CO, and NV, and countries of India, Bulgaria, Barbados, and Dominican Republic. | Terminated | March 07, 2012 | 1258 | 41 | Not_Computer
Z-0209-2009 | 49092 | No | GIO | Capillary blood collection tube. | Hematology | Hematology | 510(K) Exempt | Iris Sample Processing SafeCrit 40mm Tubes. Catalog number 59-000869-002 H. Iris Sample Processing, 60 Glacier Drive, Westwood, MA 02090. |  |  | August 01, 2008 | 2008 | Drummond Scientific Co | presence of OSCS contaminant | Material/Component Contamination | The recalling firm issued recall letters dated 8/1/08 via certified mail to the label owners informing them of the problem and the need to notified their customers. |  | 4000 | 4000 vials | Nationwide Distribution --- including states of FL and MA. | Terminated | December 18, 2008 | 139 | 19000 | Not_Computer
Z-0281-2009 | 49451 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ASP Automatic Endoscope Reprocessor, AER Plus Product Code 20300 and AER with printer Product Code 20301 |  |  | September 11, 2008 | 2008 | Advanced Sterilization Products | Residual high-level disinfectant solution remaining in endoscopes that have been reprocessed in the ASP Automatic Endoscope Reprocessor. If high-level disinfectant, or other fluids, remains in the endoscope after reprocessing, contact with the mucous membranes may occur and result in chemical burns, irritation (chemical colitis), or other symptoms. | Component design/selection | ASP is mailing customer notification CL-101367 and AD-52976-001 in an Urgent Product Correction envelope starting the week of September 8, 2008. The letter advises that the firm has received customer reports of residual high-level disinfectant solution remaining in endoscopes that have been reprocessed in the ASP Automatic Endoscope Reprocessor (AER). If high-level disinfectant, or other fluids, remain in the endoscope after reprocessing, contact with patient mucous membranes may occur and result in chemical burns, irritation (chemical colitis), or other symptoms. ASP has tested the PENTAX 70 series endoscopes containing FWJ (Forward Water Jet) and has confirmed that residual fluids may remain in the PENTAX 70 series endoscopes with a FWJ after reprocessing. To eliminate any residual fluids, ASP is recommending customers follow the revised instructions when reprocessing PENTAX 70 series endoscopes with a FWJ. A copy of the revised instructions is provided with the notification letter. |  | 1945 | 1,578 units installed in the US and 367 units installed internationally | Worldwide Distribution --- including USA and countries of Italy, UK, Uruguay, Argentina, Chile, Colombia, Ecuador, Puerto Rico, Mexico, Australia, China, Singapore, Taiwan, Middle East, Canada, Czech Republic, Portugal, South Africa, Spain, Hungary, Poland, Brazil, Russia, Egypt, and Slovenia. | Terminated | December 27, 2011 | 1202 | 1945 | Not_Computer
Z-0283-2009 | 49475 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | One Touch Ping Glucose Monitoring System |  |  | August 04, 2008 | 2008 | Animas Corporation | display old, inaccurate values for bolus amount delivered and amount of planned bolus totals | Software design | On 8/6/08, the recalling firm telephoned all users to inform them of problem and the need to return the system. |  | 251 | 251 units | Nationwide Distribution | Terminated | November 12, 2008 | 100 | 251 | Other | Display/Image | Return | Remove or Replace
Z-0215-2009 | 49371 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | VACUETTEÂż, 2 ml Lithium Heparin Venous Blood Collection Tubes. Reference #454237, 13x75 green cap-white ring, 24 racks of 50 pcs (1200 pcs in total, non ridged, Sterile). Manufactured by: Greiner bio-one NA, Inc., 4238 Capital Drive, Monroe, NC 28110. |  |  | July 29, 2008 | 2008 | Greiner Bio-One North America, Inc. | Labeling Error: Incorrect tube label reads "Z Serum Clot Activator" instead of "LH Lithium Heparin" tube.All other information on the tube label including color coding of tube was correct. Case and rack label information showed the correct information with 2 ml Lithium Heparin. | Labeling mix-ups | Consignees were notified by letter on 07/29/2008. Notification letters to distributors instructed them to stop distribution of the product and destroy all affected lots of the product. In addition, notification letters addressed to users instructed them to destroy and discard all products from the item and affected lots. The "Product Disposition Form" must be completed by the consignee and returned to Greiner Bio-One North America, Inc.For questions or additional information, contact Greiner Bio-One North America, Inc. at 704-261-7800. |  | 33 | 33 cases (39,600 pieces) | Worldwide including USA (states of AK, AL, AR, CA, CO, DE, GA, ID, IL, IN, KS, LA, MA, MI, MO, MS, NC, ND, NE, NH, NY, OH, OK, OR, PA, RI, SC, SD, TX, VA, VT, WA, and WI) and country of Canada. | Terminated | August 26, 2009 | 393 | 33 | Not_Computer
Z-0295-2009 | 49604 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Sarns Level Sensor II (yellow) (Part # 195215) alert level sensor transducer for use with Sarns Advanced Perfusion System I, Perfusion System 800 and Perfusion System 900; Terumo Cardiovascular Systems Corp., Ann Arbor, MI.; REF 195215. Level sensors are distributed as a finished device, as part of a kit, with a perfusion system 8000 safety monitor or with an Advanced Perfusion System 1 base. |  |  | May 23, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The level sensor may not properly couple to the reservoir, resulting in a sensor "not attached" message, an alert or alarm condition, or a failure to detect a low level condition. | Nonconforming Material/Component | U.S. consignees were notified by (Urgent Medical Device Removal) letter dated 5/23/08 via FedEx and international affiliates were notified via e-mail on 5/27/08. Consignees were advised of the problem, and to continue using the affected sensors until replacements are available in late 2008.For questions or additional information, contact the Terumo CVS Customer Service at 1-800-521-2818. |  | 574 | 574 | Worldwide including USA and countries of Australia, Belgium, Bolivia, Canada, Chile, Ecuador, Egypt, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Latin America, Malaysia, Panama, Philippines, Saudi Arabia, Singapore, South Africa, South Korea, Sri Lanka, Taiwan, Thailand, United Arab Emirates, Venezuela and Vietnam. | Terminated | August 04, 2009 | 438 | 1121 | Not_Computer
Z-0197-2009 | 49553 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DeRoyal Lap Appy TraceCart, REF53-1776, Rx only, non-sterile. (custom surgical kit) |  |  | October 01, 2008 | 2008 | DeRoyal Industries Inc | Surgical kits contained a recalled Endopath ETX 35 mm Endoscopic Linear Cutter White Reload/Cartridges. | Nonconforming Material/Component | The firm sent out recall notices to their sales representatives, the distributors, and the hospitals addressed to the Recall Coordinator/OR Supervisor on 10/2/08. These notices stated that the recalling firm's sales representatives would perform the field correction in lieu of returning the kits. Contact DeRoyal Industries, Inc. at 1-865-362-6157 for assistance. |  | 4 | 4 kits | LA &amp; KY | Terminated | December 02, 2009 | 427 | 34 | Not_Computer
Z-0280-2009 | 49575 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 7F Bard SSV 13 cm Split Sheath with Valve and sideport |  |  | October 01, 2008 | 2008 | C. R.Bard, Inc./BardElectrophysiology Division | Failure to insert the guidewire through the introducer needle. | Process control | Bard issued a recall notice on 1 October 2008 [note that the letter is dated 29 September 2008] to all recipients of affected product via certified mail. Users are requested to discontinue use and return inventory. Contact Bard Customer Service at 1-800-824-8724 for assistance. |  | 3105 | 3,105 kits | Nationwide | Terminated | March 09, 2010 | 524 | 3105 | Not_Computer
Z-0275-2009 | 49602 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ATEC Breast Biopsy and Excision System 12 gauge, 20 centimeter handpiece, sterile, Hologic, Inc., Indianapolis, IN; Model number ATEC 1212-20. |  |  | September 22, 2008 | 2008 | Hologic, Inc. | The distal tip of the needle may become detached and remain in the patient, requiring surgical removal. | Process control | Consignees were notified via an Urgent Medical Device Recall letter dated 9/17/08 asking them to immediately check their inventories and quarantine all affected product from the listed lots. Also, if the product had been further distributed, users were to be identified and notified of the recall. Products are to be held in quarantine until they are returned to Hologic. The response form must be completed and returned to fulfill regulatory requirements and document actions taken at your facility. Handpieces whose pouches had been opened were to be destroyed, as they could potentially shift during transport and expose dangerous sharp points to personnel handling the boxes. Upon receipt of the attached response form, Hologic will promptly send an equivalent amount of replacement product free of charge. |  | 5120 | 5,120 (1,024 cases of 5) | Worldwide Distribution --- including USA and countries of Australia, Belgium, Brazil and Singapore. | Terminated | November 18, 2009 | 422 | 5120 | Not_Computer
Z-0276-2009 | 50068 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Barco Surgical Display, Model MDSC-2124 |  |  | September 09, 2008 | 2008 | Barcoview | The front protective cover may loosen and completely fall off. | Other | Consignees were notified via a Field Safety Notice letter on 09/09/2008. They were instructed to inspect the units to check if the front panel is tightly secured to the display housing. Any front panel found to be lose or showing a gap between the front plate and the back shell of the monitor is to be reported the BARCO Service &amp; Support organization and is to be secured using adhesive tape. Follow up Field Safety Notice was sent to consignees on 09/26/2008 indicating the root cause and the firm's corrective field actions implemented October 15, 2008. |  | 25 | 25 units | Nationwide Distribution --- including states of AK, GA, IL, and MA. | Terminated | June 02, 2009 | 266 | 25 | Not_Computer
Z-0287-2009 | 49571 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | NAMIC Custom Angiographic Kit, Right Heart Kit, REF/Catalog # 60190254, Sterile, and bulk Non-Sterile Manifolds Assemblies, Product #s 40338200 and 40336200 (foreign product for further processing, e.g., packaging, labeling, sterilization, and distribution under BSC). Manufactured by: Boston Scientific Corporation/NAMIC Technology Center/Glens Falls, New York 12801. |  |  | August 06, 2008 | 2008 | Navilyst Medical | Firm became aware that a specific batch of the product may contain loose plastic particulates in the fluid pathway. | Process design | On 8/12/08, Navilyst Medical sent a fax and phoned the sole domestic hospital account that had received the affected units. Navilyst Medical instructed the consignee to immediately locate, segregate, and quarantine the units. The hospital consignee confirmed that all units were still in their control and that they will maintain them in quarantine awaiting return to Navilyst Medical. The sole international account, Boston Scientific in Ireland was notified of the problem on 8/06/08. This was followed up with an Urgent Medical Device Recall Letter and Verification Form sent via FedEx on 8/20/08. |  | 16220 | Domestic - 20 units; Foreign - 16,200 bulk units | Worldwide Distribution --- USA including state of MO and country of Ireland. | Terminated | June 16, 2009 | 314 | 16220 | Not_Computer
Z-0288-2009 | 49709 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Biomet 2.0 MM Stainless Steel (SS) Crimp Sleeve, sterile, use w/2.0 MM S.S. cable # 350800, Biomet Orthopedics, Inc., Warsaw, IN; REF. 350805. |  |  | September 24, 2008 | 2008 | Biomet, Inc. | The outer package is properly labeled, but the inner package may be labeled as a femoral component. | Labeling mix-ups | The firm's sales force was notified of the problem via an Urgent Medical Device Removal Notice letter dated 9/24/08, and provided instructions to retrieve the product from customer accounts and to provide them with a copy of the removal letter. Distributors were to immediately locate and remove the identified implants from circulation, carefully follow the instructions, and fax a copy of the Response Form prior to the return of the product. If the product was further distributed, hospital personnel MUST be notified via the enclosed "Dear Biomet Customer" notice. The letter MUST be given to hospital personnel responsible for receiving recall notices but distributors are charged with the location and return of those products. |  | 48 | 48 | Worldwide Distribution --- USA including states of Arkansas, Florida, Kansas, New Jersey, and Ohio, and countries of Argentina, Canada, Chile Italy, Mexico and United Kingdom. | Terminated | June 29, 2009 | 278 | 48 | Not_Computer
Z-0218-2009 | 49463 | No | MLY |  |  | Physical Medicine | 510(k) | Gebauer's Fluro-EthylÂż Nonflammable Topical Anesthetic Skin Refrigerant (Aerosol Can) P/N 0386-0020-20. The product consists of a can, a valve and an actuator. |  |  | September 03, 2008 | 2008 | Gebauer Company | Some units of Fluro-Ethyl containing the defective valve which could malfunction and spray refrigerant out from the side of the valve in addition to or instead of spraying in the normal inverted position from the product's actuator. The product can also leak coolant from the side of the valve onto the fingers of the user. | Component design/selection | A recall notification letter (Recall Notice) (dated 9/3/2008) was issued on 9/16/2008 to customers. The notice instructs customers to check their inventory for any remaining recall product and if found, return the product to the Gebauer Company. The letter also advises distributors to notify their customers and follow instructions as noted above, i.e. customers inventory their stock for recalled product and if found, return to the Gauber Company. For questions or additional information, contact Gebauer Company at 216-581-3030 (ext. 120) or 800-321-9348. |  | 12717 | 12,717 Units | Nationwide Distribution. | Terminated | October 13, 2010 | 770 | 12717 | Not_Computer
Z-0214-2009 | 49136 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Sunquest Laboratory System. Software Version: Sunquest Laboratory v5.4.2, v6.1, v6.2 and v6.3. |  |  | August 01, 2008 | 2008 | Sunquest Information Systems, Inc. | While performing in-house testing for the Sunquest Laboratory System and under certain conditions, there have been instances in which an error that occurred at the database level or originated in the operating system (Cache)was not communicated or displayed in the Graphic UserInterface (GUI) application so the activity continued. While this error was registered in the error log in the SystemU | Software design | A notice (PSN-08-07) was issued to all customers on 08/01/2008 via e-mail. The notice advises customers of the problem and actions Sunquest Information Systems has taken to eliminate the Problem. Sunquest Client Support will contact their facility to schedule correcting the error trap setting. The notice also provides temporary procedural instructions to include monitoring error log using System Utilities (Function UTL) and suboption 2, System Error/Message Inquiry. Contact Client Support at 520-570-2252 for questions or additional information. |  | 139 | 139 | Worldwide Distribution including USA and countries of Canada, Denmark, Saudi Arabia and United Arab Emirates. | Terminated | November 28, 2011 | 1214 | 139 | Other | Alarm/Message | Error trap setting correction | Repair
Z-0195-2009 | 49546 | No | EWO | Audiometer. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | GSI Audera (version 2.6 software) Viasys Healthcare, Inc. |  |  | September 16, 2008 | 2008 | Cardinal Health, Inc | GSI Audera systems requires a system software update due to issues with the Split-Screen and the Vestibular Evoked Myogenic Potential (VEMP) functions. A. Split Screen Field Correction:GSI Audera units with version 2.6 software may mislabel Auditory Evoked Potential (AEP) waveform responses with respect to the identification of the stimulus ear. When using 2.6 software in split-screen mod | Software design | Consignees were sent a Cardinal Health "Urgent: Medical Device Correction" letter dated September 16, 2008 on 9/22/08 and 9/26/08. The letter was addressed to the Customer/Distributor. The letter instructed consignees to perform installations if they have the following software: 1) For Version 2.6 software, you must install Version 2.6.5 software to correct a Split- Screen performance issue and to remove the VEMP function. 2) For Version 2.0 or 2.1 software, you must install the Protocol Update Software to remove the VEMP function. Versions 2.0 and 2.1 do not contain Split-Screen capability. "Attachment A-Verification Form" must be completed by the recipient/distributor.If you need assistance in making the software revision determination or have further questions, please contact the Cardinal Health (formerly VIASYS) Neurocare/GSI Technical Support Group at 1-800-356-0007, 1-800-700-2282, 608-273-5000, or by e-mail at gsiservice@viasyshc.com or gsiservice@cardinalhealth.com. |  | 297 | 297 | Nationwide including states of AL AZ CA CO CT DC FL GA HI IA IA IL IN KS LA MA MD ME MN MO MS NC NC NH NJ NM NY OH OK PA RI SC TN TX UT VA WA and WI. | Terminated | May 01, 2012 | 1323 | 297 | Software | Device Operation | Software update | Software Update
Z-0216-2009 | 49427 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Architect LH MasterCheck (List 6C25-05). An in-vitro diagnostic, consisting of 4 - 2 mL vials of LH MasterCheck, Levels 0, 1, 2 and 3. Manufactured for Abbott Laboratories, Abbott Park, IL 60064. |  |  | September 05, 2008 | 2008 | Abbott Laboratories | The values listed in the Architect LH MasterCheck, Lot 82520 data sheet are incorrect. When MasterChecks do not perform as intended, the integrity of the system cannot be verified and patient results would not be generated. Other commercially available materials are available to customers in order to meet CLIA requirements. | Error in labeling | A Product Correction, Immediate Action Required letter and Customer Reply Form dated 9/5/08 were sent to all Architect LH MasterChecks (List Number 6C25-05) customers who received Lot 82520M. Customers were informed that the values listed in the Architect LH MasterCheck Lot 82520 data sheet are incorrect. Customers were requested to check their inventory for the affected lot, and if any kits were found, to substitute the values printed on the data sheet with the values supplied in the letter. In addition, Customers were requested to keep a copy of the letter on file and to complete the Customer Reply Form, indicating that the letter was received, understood and disseminated to the laboratory manager/supervisor responsible for the Architect LH MasterCheck testing, and fax it to Abbott by 9/19/08 at 1-800-777-0051. Follow-up telephone calls will be placed to customers who do not provide a written reply via the Customer Reply Form. U.S. Customers please contact Abbott Laboratories Customer Service at 1-877-422-2688 for questions regarding this information. |  | 30 | 30 kits | Nationwide Distribution including Arizona, California, Florida, Illinois, Kansas, Louisiana, Massachusetts, Missouri, Montana, New York, Ohio, Pennsylvania, South Carolina, Texas, Utah and Virginia. | Terminated | March 18, 2010 | 559 | 30 | Not_Computer
Z-0272-2009 | 49544 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bone Marrow Aspiration Kit, consists of 3 main components packaged in a box. 1. Marrow Loc (Bone Marrow Specimen Acquisition System), 2. "J" Style Bone Marrow Biopsy Needle (11g (3.0mm)x10cm (4.0)) 3. BD 20ml Syringe. |  |  | August 27, 2008 | 2008 | Theken Spine LLC | The kits were manufactured in 2005 with syringes that expire on 10/2009 and labeled with an overall expiration on 10/2010. | Other | On 8/27/2008 the firm sent recall notification letters to each of their customers notifying them of the problem with the product. They were requested to return any remaining product back to the recalling firm and to call Customer Service at (866) 942-8698 for an RGA and to arrange shipment. On 8/28/2008 the firm conducted follow up contact each distributor via telephone. |  | 115 | 115 total | Nationwide Distribution --- including states of CA, NC, NJ, NM, NY, TX, and UT. | Terminated | April 13, 2012 | 1325 | 115 | Not_Computer
Z-0217-2009 | 49441 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Reflex Hybrid 4.0 x 14 mm Variable, Self-Tapping Bone Screw (Non- sterile). |  |  | August 18, 2008 | 2008 | Stryker Spine | The Reflex Hybrid 4.0 x 14 mm Variable, Self Tapping Bone Screw, Lot NSZ, was determined to have been anodized with the incorrect color for that screw. | Nonconforming Material/Component | An URGENT Product Recall Letter and a Stryker Spine Accountability Form were sent via FedEx with return receipt on August 18, 2008 to Stryker Branches and Hospital OR Supervisors. Letters notified branches of the recall and requested that they immediately identify and quarantine affected product and notify hospital locations that received shipment of product within 72 hours prior to identification of the problem. Letters sent to Hospital OR Supervisors described the problem and requested that all inventory locations be checked for affected product. If recalled product is found, Stryker Spine Representatives must be contacted to retrieve the recalled product from the facility. The Stryker Spine Accountability Form must be completed and returned via fax, 201-760-8269 to Stryker Spine. Contact Stryker Spine Clinical and Regulatory Affairs at 201-760-8150 or 201-760-8069 for questions or additional information. |  | 183 | 183 units | Nationwide Distribution. | Terminated | June 19, 2009 | 305 | 183 | Not_Computer
Z-0301-2009 | 48712 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Alumina Ceramic Femoral Head, 12/14 taper, 28 mm diameter, neck length +3.5 mm, sterile, Zimmer, Warsaw, IN; REF 00-6418-028-03 and 00-6428-028-03. |  |  | June 11, 2008 | 2008 | Zimmer Inc. | When used with cobalt/chromium hip stems, the average test values for ceramic head bursting falls below the guidance document limits if the taper is made by casting instead of by forging. | Device Design | The firm issued Product Compatibility Update # 13, dated 6/11/08, to consignees informing them which products are not compatible, and thus should not be used in combination. |  | 49474 | 49,474 of all products | Worldwide Distribution --- including USA and countries of Austria, Australia, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Dubai, England, Finland, France, Germany, Greece, Hungary, India, Israel, Italy, Korea, Lebanon, Mexico, Netherlands, New Zealand, Poland, Romania, Singapore, Slovakia, Spain, Switzerland and Ukraine. | Terminated | November 10, 2009 | 517 | 49474 | Not_Computer
Z-0004-2009 | 49413 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Ultra Test Strips, part number 020-994-02, manufactured by LifeScan Inc, Milpitas, CA, 25 count. |  |  | September 19, 2008 | 2008 | Lifescan Inc | Products exceed inaccuracy threshold , with many complaints resulting from inaccurate low results. | Process control | Urgent: Medical Device Recall letters, dated 09/19/2008, were sent to consumer end users, direct account distributors and wholesalers, health care providers, and pharmacists notifying them of the recall. The letter states that if customers have strips from Lot 2829235, they should call LifeScan at 1-866-247-1281 and tell customer service that they have strips that need to be replaced. Customer service and QA scripts were provided to LifeScan staff to provide to any callers who inquire about the recall. The distributors and wholesalers were to locate and sequester any inventory from the affected Lot 2829235; request all recalled inventory from customers; call Carolina Logistics for an RGA; and when only the affected returned product is received, a credit will be issued. |  | 35308 | 35308 vials | Nationwide Distribution --- including Puerto Rico. | Terminated | December 27, 2010 | 829 | 35308 | Not_Computer
Z-0306-2009 | 49397 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Sapphire Hematology Analyzer, List numbers 08H00-01, 08H00-03 |  |  | August 25, 2008 | 2008 | Abbott Laboratories | Exposed wire resulted in minor electrical shock and burn to the service technician | Device Design | Recall issued August 22, 2008. Technical service bulletin was issued for all field personnel via internal database. No customer communication is planned as customers are instructed *not* to remove the cover, which is necessary to access the area. |  | 958 | 958 units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Brazil, Canada, Chile, China WFOTC, Colombia, Germany, Hong Kong, Japan, Mexico, New Zealand, Panama, Paraguay, Peru, Singapore, South Korea, Taiwan and Venezuela. | Terminated | January 06, 2011 | 864 | 958 | Not_Computer
Z-0297-2009 | 48628 | No | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | AxSYM Digoxin II Reagent Pack, Disk Version 1. (3D53-01 or 3D53-02) |  |  | May 22, 2008 | 2008 | Abbott Diagnostic International, Ltd. | AxSYM TDM/ Transplant Assay Disks 3D53-03 and higher do not contain the AxSYM Digoxin II Assay file.The instrument might require installation of the AxSYM Digoxin II assay file. | Software change control | A Product Information letter, dated May 22, 2008, was issued to customers. It requested that one of the following two actions be taken if the AxSYM Digoxin II assay file is not installed on your AxSYM instrument and you want to run the AxSYM Digoxin II assay. If you have one of the previously manufactured software disks listed, [AxSYM Digoxin II Assay Disk, LN 2C43-01, or higher; or AxSYM TDM/Transplant Assay Disk Version 1.0 (3D53-01) or Version 2.0 (3D53-02)], install the AxSYM Digoxin II assay file. Do not install any other assay file. If you do not have one of the assay disks listed, contact your local Customer Service representative to have this assay file installed or install the file from a previous version if you have it on hand. |  | 9069 | 9069 kits | Worldwide Distribution --- USA including states of AL, AZ, AR, CA, CT, FL, GA, IL, KS, MD, MA, MI, MN, MS, MO, NV, NJ, NY, ND, OK, OR, PA, RI, TX, VA, WA, Hawaii, and Puerto Rico, and countries of Canada, Mexico, Guatemala, Costa Rica, Venezuela, Colombia, Ecuador, Peru, Chile, Uruguay, Argentina, Puerto Rico, Germany, Japan, Singapore, Hong Kong, Taiwan, Thailand, South Korea, Australia, New Zealand, Curacao, and Panama. | Terminated | September 09, 2011 | 1205 | 9069 | Computer
Z-0305-2009 | 49097 | No | FMZ | Neonatal incubator. | General Hospital | General Hospital | 510(k) | Giraffe OmniBedÂż, labeled in part: ***GE Healthcare 8880 Gorman Road Laurel, MD 20723 U.S.A. www.gehealthcare.com***. |  |  | June 20, 2008 | 2008 | Datex Ohmeda, Inc. dba GE Healthcare | Manufacturer of neonatal incubators changed the device labeling to account for new use and care instructions. | Equipment maintenance | A notification letter (Urgent Medical Device Correction) was sent to consignees on 06/20/08. The notification letter addressed healthcare professionals, including neonatal, infection control and nursing departments at medical facilities. In add tion, the letter advised of revisions to the current Giraffe OmniBedÂżÂż cleaning instructions contained in the product Operators Manual. Instructions were enclosed with the notification letter and are intended to supplement those already present in the manual.Contact the GE Healthcare Technical Support center at 1-800-558-7044 for questions or additional information. If outside of the USA, please contact your local GE Healthcare Service Representative. |  | 7878 | 7,878 beds | Worldwide Distribution: The devices were widely distributed to national and international medical facilities. | Terminated | March 24, 2010 | 642 | 7878 | Not_Computer
Z-0289-2009 | 49867 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Vanguard(TM) DCM PS Plus 16mm x 63/67mm Tibial Bearing (REF. 183726), ARCOM UHMWPE, sterile, Biomet Orthopedics, Inc., Warsaw, IN. |  |  | September 22, 2008 | 2008 | Biomet, Inc. | The package is properly labeled, but the laser etch for size incorrectly reads "16 X 63/63" instead of the correct size, "16 X 63/67". | Process control | A notification letter (Urgent Medical Device Removal Notice) dated 9/22/08 was sent to the firm's sales staff with instructions to locate and retrieve the product and to notify customers at affected hospitals of the problem.Questions related to this notice should be directed to the Biomet Orthopedics, LLC Corrective Actions Manager at 800-348-9500 or 574-372-3983, Monday through Friday, 8 A.M. to 5 P.M. (EST). |  | 11 | 11 | Worldwide Distribution: USA including states of Florida, New Jersey, Maryland, Pennsylvania, South Carolina, South Dakota, and Virginia, and country of Finland. | Terminated | September 22, 2009 | 365 | 11 | Not_Computer
Z-0340-2009 | 50073 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100S Sterilizer (Product Code 10101). The sterilizer includes software and hardware components. |  |  | October 17, 2008 | 2008 | Advanced Sterilization Products | 1) Inability of the sterilizer to detect when an injectiontakes place without hydrogen peroxide being transferred to the vaporizer bowl. This situation can result in a cycle completion with the injection of insufficient hydrogen peroxide sterilant and 2) Inability of the sterilizer to detect an obstructionin the door travel path while the door is closing. This situation can result in a hand b | Component design/selection | Customer notifications will be sent the week of November 24, 2008. The notification letter (URGENT: Product Correction) will advise of two situations that could cause the STERRAD 100S Sterilizer to malfunction. The first situation involves the failure of the STERRAD 100S System to detect when an injection takes place without hydrogen peroxide having been transferred to the vaporizer bowl. This can occur under combinations of the following conditions:- the pneumatic pressure is low on the machine, or- the injector valve has a small leak,AND- a cassette fails to advance, or- a used cassette is re-inserted into the sterilizer.This situation can result in a cycle completion with insufficient hydrogen peroxide. The second situation involves the failure of the STERRAD 100S System to detect an obstruction in the door travel path while the door is closing. Under some conditions, the door sensors may become mechanically unsecured/disconnected and will not detect an obstruction. This situationcan result in an operator's hand becoming pinched and bruised while the door is closing. To correct both situations above, Advanced Sterilization Products(ASP) has revised the software and is making physical hardware additions to the STERRAD 100S System.Contact ASP Customer Care Center at 888-783-7723, option 2 for questions or assistance. |  | 6324 | In the U.S., there are 2,623 units installed. There are also 3,701 units internationally. | Worldwide Distribution including USA and countries of France, Germany, Italy, UK, Switzerland, Benelux, Uruguay, Venezuela, Argentina, Chile, Colombia, Ecuador, Belgium, Puerto Rico, Hong Kong, Indonesia, Mexico, Philippines, Australia, Austria, China, Greece, Israel, Japan, Korea, Malaysia, Singapore, Taiwan, Thailand, Turkey, Middle East, Peru, Canada, Czech Republic, Portugal, South Africa, Spain, Sweden, Hungary, Poland, Brazil, Russia, Egypt, India and Slovenia. | Terminated | November 29, 2011 | 1138 | 6324 | Software | Output/Calculation | Software update | Software Update
Z-0323-2009 | 50079 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Arjo Maxi 500 Patient Lift with 4 Point DPS (Dynamic Positioning System) and Scale; a Non-AC-Powered Patient Lift; Models KM560181; BHM Medical Inc., Magog, QC, Canada |  |  | October 20, 2008 | 2008 | Arjo, Inc. | The pivot bolt that attaches the hanger bar to the scale of the patient lift can break, resulting in the hanger bar falling. | Nonconforming Material/Component | Arjo sent an Urgent Device Recall Customer Notification letter dated 10/20/08 to all affected end users to the attention of the Administrator/Risk Manager via UPS 2nd day on the 10/20/08. The letters informed the accounts of the potential for the hanger bar to detach from the scale due to shearing of the pivot bolt. The accounts were advised to review the enclosed Safety Advisory Notice, note the affected model and serial numbers, and avoid using the lift under load on a high friction floor like carpet or soft tile while using the hanger bar to move the lift. The accounts were advised that an Arjo Field Service Technician would contact them within two weeks to schedule a visit to the account to perform repairs on their Maxi 500 lifts. Any questions were directed to Arjo Quality Department at 800-323-1245, ext. 6118. The accounts were also instructed to complete and fax back to Arjo the enclosed Customer Response Form, providing the facility name and address, contact name and phone number, the numbers of affected units, and acknowledgment of receipt and understanding of the Urgent Device Correction. |  | 336 | 336 units | Nationwide Distribution. | Terminated | August 19, 2010 | 668 | 374 | Not_Computer
Z-0310-2009 | 46312 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, 6 inch diameter Roller Pump; Model 801041. |  |  | July 18, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The pumps may fail to power up or experience unplanned pump stops. | Process control | Consignees were advised of the problem via an Urgent Medical Device Correction: Safety Advisory letter dated 7/18/08 and each consignee was contacted by phone to assure receipt. The letter noted that the firm was not recommending that customers stop using the product. If customers found the risk unacceptable they were to contact Technical Service, 1-800-521-2818. Customers were to review the notice, assure all users were aware of the notice, and confirm receipt by completing the attached Customer Response Form and faxing it to the number indicated on the form. |  | 1429 | 1429 | Worldwide Distribution --- including USA and countries of Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Hong Kong, India, Indonesia, Israel, Japan, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Vietnam. | Terminated | July 07, 2009 | 354 | 2548 | Not_Computer
Z-0122-2009 | 49584 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis dTA, Axiom Artis Modular Angiography X-Ray System, Model number 7008605. |  |  | September 02, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Incorrect contour finding: Image may be calibrated to the wrong size catheter. | Software design | Siemens issued a Customer Safety Advisory Notice dated 8/29/08 to inform customers of the potential issue. The recalling firm issued Update Instructions AX040/08/S to the service engineers which provides information to fix the problem. |  | 133 | 133 units | Nationwide. | Terminated | October 05, 2009 | 398 | 577 | Not_Computer
Z-0259-2009 | 49506 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | PREMISE, PART NO. 32813, 10 PACK UNIDOSE PREMISA A3 (INTERNATIONAL ONLY), dental composite |  |  | September 05, 2008 | 2008 | Kerr Corp | Material stiff: Material in several lots of Premise Unidose appears to stiffen and become difficult to extrude over time. | Component design/selection | Kerr sent an Urgent Medical Device Recall letter to consignees via USPS first class mail the week of September 8, 2008, requesting they return the product for replacement or credit.The recall notification requested the return of the affected products because the material appears to stiffen and become difficult to extrude over time. Customers were instructed if they have any of the affected lots of product in their inventory to immediately return it to Kerr. Any product returned will be replaced at no charge or credited to your account. In addition, customers were instructed to fax the attached Recall Return Form to 714-516-7739.If the customers were an authorized Kerr distributor, they were requested to identify and recover the affected product lots listed on the attached Lot Number Table that may have been shipped to their customers.Also, customers were instructed to contact Kerr Customer Care at 1 -800-537-7123 directly to handle the arrangements of a quick return and replacement |  | 22478 | 22,478 packs for all affected lots | Worldwide: USA, Australia, Canada, Hong Kong, Malaysia, Mexico, and Singapore. | Terminated | March 07, 2012 | 1279 | 22478 | Not_Computer
Z-0308-2009 | 49564 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Proven Cemented Semi-Constrained Total Knee. Part number SC1591. The size is 38 and the thickness is 10 mm. |  |  | September 03, 2008 | 2008 | Stelkast Co | Five (5) Three Peg Patella SC1591-29 were packaged in containers labeled for SC1591-38 sharing the same lot number. | Employee error | On 9/4/08, the recalling firm telephoned their customers to inform them of the problem and asked them to immediately recover all SC1591-38 with lot number 19906-052008. Once the product is in their possession, customers should contact Stelkast Customer Service for a Return Authorization (RA) Number prior to shipment to Stelkast. |  | 5 | 5 units | Nationwide Distribution --- including states of IA and WV. | Terminated | December 04, 2008 | 92 | 5 | Not_Computer
Z-0294-2009 | 49476 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Battery Caps used with Animas IR1250 Insulin Pump. |  |  | August 04, 2008 | 2008 | Animas Corporation | Unintentional rebooting: Pump products exhibit an intermittent loss of power due intermittent loss of contact between battery cap and battery canister resulting in the device resetting. The failure of the battery cap may result in failure of the device to administer insulin therapy which may result in hyperglycemia. | Nonconforming Material/Component | The recalling firm issued a notification letter to the patients with insulin pumps on 8/5/08 to inform them of the problem and that they needed to replace the battery. The recalling firm contacted OneTouch Ping users by telephone on 8/5/08 for this issue and another issue. The recalling firm issued letters to distributors informing them of the problem and notified customers to replace the batteries. In addition, a letter was issued to physicians to alert them of the situation with regard to current patients. Consumers can contact 1-800-423-4087 for additional information or questions. |  | 1624 | 1,624 units | Worldwide Distribution including USA (AR, CA, FL, KY, LA, MA, NE, OH, PA, SC, SD, TN, and UT) and countries of Australia, Austria, Canada, Czech Republic, Finland, France, Israel, Italy, New Zealand, Spain, Switzerland, and Sweden. | Terminated | February 04, 2010 | 549 | 7152 | Battery | Treatment/Delivery/Therapy | Replace battery | Remove or Replace
Z-0327-2009 | 50123 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Luminos dRF without ST filter. . Solid state x-ray imager (flat panel/digital imager). Model number 10252048. |  |  | September 19, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Issue may occur where an intended movement of the joystick may unintentionally initiate an x-ray exposure. | Nonconforming Material/Component | The recalling firm issued Update Instruction AX042/08/S dated August 2008, to affected customers. This Update instruction provided Siemens service engineers with instructions for the replacement of the joystick module at each facility. |  | 1 | 1 unit | Nationwide Distribution --- including states of KY, NJ and PA. | Terminated | September 23, 2009 | 369 | 4 | Not_Computer
Z-0415-2009 | 50149 | No | FDG | Gastroenterology-urology fiberoptic retractor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Fredrick's Converse Retractor #5 F/O, Model Number: 01-0485. |  |  | October 24, 2008 | 2008 | Biomet Microfixation, Inc. | Fiber optic cable is missing from the Frederick's Converse Retractor Lot # 042408D08. | Nonconforming Material/Component | Consignee was notified by certified mail via an Urgent Medical Device Recall Notice dated 10/24/08, with return receipt. The letter identifies the product and states the issue with the product. The letter required the customer to discontinue use of the item and return it to Biomet Microfixation so that they may rework the instrument. It also asked the customer to notify any hospital personnel responsible for receiving recall notices of this action if the product had been further distributed. The RGA number should be utilized to return the products. The Inventory Reconciliation sheet should be filled out and faxed to 904-741-9425. |  | 3 | 3 | Nationwide Distribution --- including states of Chicago, IL and La Crosse, WI. | Terminated | September 11, 2009 | 322 | 3 | Not_Computer
Z-0331-2009 | 50096 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Easy Core" Biopsy System, 18 ga. x 15 cm, REF 43-453 , M001434531, Sterilized with ethylene oxide gas. |  |  | October 07, 2008 | 2008 | Boston Scientific Corporation | Boston Scientific Corporation initiated a recall of its easy Core" BIOPSY SYSTEM due to difficulty cocking or arming the cannula latch on the device. This difficulty may result in an inability to use the device. | Device Design | Consignees were sent a "Boston Scientific Urgent Medical Device Recall" letter dated October 21, 2008. The letter was addressed to The Hospital Field Action Contact. The letter described the problem and products being recalled. The letter stated that it was very important that the attached product table be used carefully and consider both inner and outer packaging product codes when searching for affected/recalled product as the UPN numbers on the inner and outer labeling are different. Requested to discontinue use of product, identify, and segregate and return affected products to BSC. Letter also requested to complete and return the Reply Verification Tracking Form. |  | 3808 | 3808 total for all products | Nationwide Distribution --- including states of AL, AZ, AR, CA, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO,MT, NE , NV, NH, NH, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI, and PR. | Terminated | December 10, 2011 | 1159 | 3808 | Not_Computer
Z-0298-2009 | 48645 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | "LOCATOR Abutment for 5.7 Screw-Vent &amp; Compatibles." Zest Anchors, Inc. Product is sold for use with any Zimmer Dental 5.7 mm diameter Tapered Screw-Vent Implant. |  |  | April 09, 2008 | 2008 | Zest Anchors Inc | The recall was initiated after the firm discovered through a customer complaint that the Locator Abutments for 5.7 Screw-vent &amp; Compatibles do not fit properly in all sizes of the Zimmer Dental 5.7mm diameter Tapered Screw-Vent Implants (TSV6B8, TSV6B10, TSV6B11, TSV6B13, TSV6B16, TSV6H8, TSV6H10, TSV6H11, TSV6H13 and TSV6H16). The Locator Abutment bottoms out in the pilot hole of these implants c | Nonconforming Material/Component | The recall notification letter (Recall Notice) was sent to consignees on April 23,2008. The letter includes the following information on how to handle the recall.It is very important that you determine the use of the Locator Abutment(s) as follows, mark the appropriate box, ,and return this letter to Zest Anchors, Inc. in the stamped, self-addressed envelope:1. The Locator Abutment(s) listed above WERE NOT used on any Zimmer Dental 5.7mm diameter Tapered Screw-Vent Implants. THE FIT IS CORRECT AND NO ACTION IS NECESSARY.2. The Locator Abutment(s) listed above WERE used on a Zimmer Dental 5.7mm diameter Tapered Screw-Vent Implant. PLEASE CONTACT OUR DENTAL OFFICE TO ARRANGE THE R PLACEMENT OF THE ABUTMENT WITH A NEWLY DESIGNED LOCATOR ABUTMENT. and3. We have current STOCK of the Locator Abutment(s) listed above with planned use for a Zimmer Dental 5.7mm diameter Tapered Screw-Vent Implant. PLEASE CONTACT OUR DENTAL OFFICE TO ARRANGE THE REPLACEMENT OF THE ABUTMENT WITH A NEWLY DESIGNED LOCATOR ABUTMENT.For questions or additional information, contact Zest Anchors, Inc. at 800-262-2310 or 760-743-7744. |  | 996 | 996 units | Worldwide Distribution including USA and countries of Spain, Canada, Australia and South Africa. | Terminated | December 02, 2008 | 237 | 996 | Not_Computer
Z-0307-2009 | 49561 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Uni-CP Compression Forceps (spreader); |  |  | September 30, 2008 | 2008 | Integra LifeSciences Corp | Reports of the Uni-CP Compression Forceps breaking during the Compression of the Uni-CP plate | Nonconforming Material/Component | Recall Notice letters were sent to all Integra Sales Specialists on September 30, 2008 by first class mail. The letter informed them of the product and the problem concerning the product. The Sales Specialists were asked to return all inventories under their control with part number 339001ND from lots E6WS, E7UY, &amp; E86M. They are required to fill out the attached form identifying the lot number for all Newdeal Compression Forceps (Spreader) in their inventory along with the serial number(s) of their Uni-CP set(s). The form should be filled out even if the sales specialist does not have a Uni-CP. The form should be faxed or mailed back to Walt Ryan in Customer Service, 609-799-3297. A RMA number will be generated for the return of any affected inventory. A replacement will be sent and upon receipt, the forceps (spreader) should be replaced and returned to the address indicated on the letter. |  | 90 | 90 | Nationwide Distribution | Terminated | May 11, 2009 | 223 | 90 | Not_Computer
Z-0358-2009 | 46313 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Terumo Sarns Sternal Saw II System Power Unit, 115 V, Terumo Cardiovascular Systems Corp., Ann Arbor, MI; Catalog # 15670. |  |  | February 14, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The drive power cable cannot be fully inserted into the power unit or can only be inserted with difficulty, resulting in the device not being operational. | Nonconforming Material/Component | Consignees were notified via Urgent Medical Device Removal letter dated 2/14/08 to not use the device if it exhibited the problem, provided a response card, and stated that replacement motors will be shipped to the person listed on the returned response card. |  | 69 | 69 total | Worldwide Distribution --- including USA and countries of Belgium, China, Columbia, Malaysia, Philippines, United Arab Emirates, and Venezuela. | Terminated | July 07, 2009 | 509 | 69 | Not_Computer
Z-0441-2009 | 50078 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 5 F BardSelect 11 cm Percutaneous Catheter Introducer Set |  |  | October 07, 2008 | 2008 | Medtronic, Inc. | Failure to insert the guidewire through the 18 gauge introducer needle | Nonconforming Material/Component | Medtronic notified Medtronic Sales representatives via UPS for hand delivery of the Urgent Notification of Product Recall letter to the affected customers on 10/7/08 advising users to segregate the product so that it would not be used. The letter stated that a Medtronic Sales Rep would visit the site to account for all units that were received and would initiate a credit, as appropriate. Customers were requested to cooperate and assist with completing the necessary documentation. |  | 1830 | 1,830 | Worldwide Distribution --- USA including states of IL and FL, and countries of Belgium, Honduras, and Germany. | Terminated | June 07, 2011 | 973 | 6555 | Not_Computer
Z-0286-2009 | 49498 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | Posey Synthetic Leather and Biothane Waist and Wrist Restraints, Catalog Numbers: 2217SL and 2217B |  |  | July 25, 2008 | 2008 | J T Posey Company | Failure to restrain: This recall was initiated after a complaint was received by the firm regarding a Posey 2217SL unit indicating that a patient was able to pull the rivet or washer through the synthetic leather allowing the patient to have freedom of arm movement. A second complaint was received indicating that a patient was able to tear the strap at the adjustment notch on a Posey 221 7SL unit | Device Design | On July 29, 2008 Posey sent a Recall Letter to all affected customers. The recall letters were sent by certified mail with return receipt signature. The recall letter instructed the customers not to distribute or use the affected products. Any existing inventory subject to this recall was requested to be returned to Posey Company. Also, if they were a distributor, they were asked to please additionally forward a copy of the recall notice and reply form to each of their customers to whom you have further distributed the subject product. A Recall Fax Back form was provided with the recall letter for the affected facilities to fax back to inform Posey of their existing inventory.An RMA number was issured to each facility and a return UPS label was provided for any product that will be returned to Posey. Additionally, the sales representatives for each of the customers will be contacted by Posey Customer Service. |  | 145 | 145 units | Worldwide, USA, Australia, Canada, France &amp; The Netherlands. | Terminated | March 30, 2012 | 1344 | 145 | Not_Computer
Z-0484-2009 | 50071 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX Sterilizer, Product Code 10033 |  |  | October 07, 2008 | 2008 | Advanced Sterilization Products | ASP has discovered a component defect in some of the UV lamp power supplies used in certain STERRAD NX Sterilizers. This defect can potentially cause the Hydrogen Peroxide Monitor to give inaccurate readings. | Nonconforming Material/Component | The Urgent: Product Correction customer notifications, dated October 6, 2008, were hand delivered by field representatives starting October 7, 2008. The letter advises that Advanced Sterilization Products has discovered a component defect in some of the UV lamp power supplies used in certain STERRAD Sterilizers. This defect can potentially cause the Hydrogen Peroxide Monitor to give inaccurate readings. A list of the affected sterilizer serial numbers is included with the letter. The letter advises that the Field Service Engineer has come to their facility to inspect and, if necessary, replace this power supply. Customers are asked to sign the letter acknowledging receipt of this letter upon completion of the work by the Field Service Engineer. |  | 79 | 32 potentially affected systems in the US - 20 installed, 5 uninstalled, 7 potentially affected. There are also 15 units internationally. | Worldwide Distribution --- USA including states of WV, CA, TN, AZ, WI, NC, MO, NJ, DE, GA, SC, WA, FL, MA, NM, MI, and AL and countries of Ecuador, China, Israel, Japan, Korea, Middle East, Australia, Czech Republic, Spain, Brazil, and Russia. | Terminated | November 07, 2011 | 1126 | 79 | Hardware | Output/Calculation | Replace component | Remove or Replace
Z-0456-2009 | 49460 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | GE Healthcare Centricity Perinatal (formerly Quantitative Sentinel) System - Alert and Reminder software; automatic patient data management providing clinical information at the bedside in Labor &amp; Delivery, Mother-Baby and the Neonatal Intensive Care Unit; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | September 25, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software anomalies in the Alert and Reminder feature could result in a delay of treatment. When attempting to select the last visible alert or reminder choice, the next choice on the list below the desired choice is selected, and an inconsistent color may be displayed for the same clinical element across a set of work stations. | Software design | The software was updated to version 6.80.1 on 4/22/08, however, the consignee was notified of the software defects relating to the last visible Alerts and Reminders choice and the inconsistent display for the same clinical element across a set of work stations, via Product Safety Notification letter dated 8/29/08, sent via first class mail on 9/24/08. If there were any questions, the customer was to contact Tech Support at 1-800-433-2009. |  | 1 | 1 unit | Nationwide Distribution --- state of Ohio. | Terminated | December 03, 2008 | 69 | 1 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0457-2009 | 49468 | No | JFP | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL7XX Calibration 1 Solution S1720. P/N S1720, 944-024 |  |  | September 05, 2008 | 2008 | Radiometer America Inc | The firm found that for this particular lot that the barcode, used for entering the characteristics of CAL1 Solution into the analyzer, does not reflect the actual values of the solution. As a consequence, the calibration curves for these parameters and thereby the measured results for patient samples and quality controls will be biased. | Process control | On, 9/2/2008 the firm sent out recall notification letters requesting their customers discard any Calibration 1 Solution (CAL1) S1720 (part number 944-024), lot WA-04, complete the attached return form and fax it to 800-736-0601. The firm will send out replacements for all affected product that the firm shows was shipped to their location. Further questions could be directed to Technical Support at 800-736-0600, option 2. |  | 3200 | 3,200/200ml bottles | Worldwide Distribution --- USA including states of AZ, CA, FL, IA, ID, IL, IN, KY, LA, MD, NC, ND, NE, NY, OH, OK, PA, SD, TX, VA, WA, and WI, and country of Canada. | Terminated | April 25, 2012 | 1328 | 3200 | Not_Computer
Z-0324-2009 | 50085 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo MultiModality WorkPlace (MM WP). Software product(Model Number 10140720). |  |  | September 12, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Values derived from Dynamic CT data sets may be incorrect. The corresponding images are not affected. | Software design | The recalling firm has issued a software update to affected customers via Update Instructions CT034/08/S. This software update corrects the values from the volume perfusion post processing application. A firm representative will visit each consignee to install the software update beginning 9/12/08. |  | 130 | 130 units | Nationwide Distribution. | Terminated | September 23, 2009 | 376 | 130 | Software | Output/Calculation | Software update | Software Update
Z-0408-2009 | 50081 | Yes | NYU | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PALINDROME EMERALD SPORT PACK Cuffed Dual Lumen Catheter with Pre-curved Shaft and Heparin Coating, REF 8888123411, Size 14.5Fr/Ch (4.85mm) x 23cm. TYCO HEALTHCARE GROUP LP, Mansfield, MA 02048. --- NOTE: Both the "sport pack" and "kit" package configurations include the catheter, Venetrac insertion stylets, pull apart sheath, and other accessories necessary for placement. The Venetrac insertion stylet is an optional accessory. The 510k for the Venetrac accessory is K051584. The 510k for the catheter is K060509. |  |  | September 19, 2008 | 2008 | Covidien | Incorrect size of the Venetrac stylets (an optional accessory). They are too short and cannot be used. | Mixed-up of materials/components | Starting on 9/19/08, Covidien Sales Representatives visited each of the hospitals that received the catheters, informed the staff that the Venetrac insertion stylets were too short, and placed pink corrective sticker labels on any unused product. Should you require additional information, please contact jim.welsh@covidien.com or 508-216-8532. |  | 10 | 10 units | Nationwide distribution including states of OH, WI, LA, TX, UT, CA, OR, and WA. | Terminated | December 04, 2008 | 76 | 117 | Not_Computer
Z-0319-2009 | 49713 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | AS3000 Anesthesia Delivery System. Datascope Patient Monitoring; |  |  | October 07, 2008 | 2008 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | Two issues have been identified with the AS3000 unit.1. The threadlocker of the caster (wheels) may not have been utilized, which may allow the caster to loosen and possibly separate from the unit.2. The use of select brands of pre-pack absorber in the absorber canister of the AS3000 has been associated, in some cases, with gas leakage around the pre-pack, rather than through the absorber mate | Component design/selection | Notification (Urgent Product Field Correction) letters were sent to the one consignee on October 7, 2008 by Federal Express overnight delivery. |  | 2 | 2 units | Only 2 units are in distribution, both in the same facility in Redondo Beach, CA. | Terminated | December 08, 2008 | 62 | 2 | Not_Computer
Z-0318-2009 | 49572 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Toddler Chair Hip Strap (Belt). The hip belt is made with nylon webbing and a plastic buckle. 510(k) exempt, Medical Device Listing # R061877. |  |  | September 12, 2008 | 2008 | Community Products, LLC | Some hip straps were assembled incorrectly by threading the buckle on the strap upside down. If a child were left unattended in spite of the warnings on the product and in the manual, a defective hip strap could loosen, allowing the child to slip down in the chair and this could create a choking hazard. | Process control | On 9/12/08, Rifton Equipment began notifying the customers by phone using a phone script. This phone call notified the customers of the recall and also served as an effectiveness check that they were notified. The customers were instructed to remove the product from use until the replacement kit is received and the old strap is returned. The replacement kit contains instructions with ID codes and a place for the customer to verify that they have made the replacement as instructed. It also contains the required number of new hip straps. A verification check was sent with the replacement straps, which the customer will sign and return to the firm verifying that they have replaced the straps for the affected serial numbers. Rifton Equipment is providing the customer with return shipping labels which were sent by UPS 2nd day air. The replacement straps use black thread instead of tan thread for the sewing on the ends of the straps so that the firm can be sure that the returned straps are not the new replacement straps.For questions or additional information, contact Rifton Equipment at 1-800-571-8198. |  | 147 | Potentially 147 units may be affected. | Worldwide Distribution including USA and countries of Spain and Great Britain. | Terminated | March 15, 2010 | 549 | 147 | Not_Computer
Z-0312-2009 | 49569 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis System dFC using software version VB31D with 30x40 cm2 flat detectors. |  |  | August 25, 2008 | 2008 | Siemens Medical Solutions USA, Inc | System may switch to emergency fluoroscopy mode. | Software design | A Customer Safety Advisory Letter dated 8/20/08 was issued with Update Instructions AX045/08/S. The letter informed the customer of the potential issue and provided instruction to avoid its occurrence. A software update (VB31E) was released to correct the issue. |  | 4 | 4 units | Nationwide Distribution including states of AR, CA, CO, FL, IA, IL, IN, KS, KY, MI, NY, TX, VA, and WY. | Terminated | January 26, 2010 | 519 | 21 | Computer
Z-0414-2009 | 50105 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Fisher &amp; Paykel Healthcare RT240 Adult Breathing Circuit Kit. |  |  | October 22, 2008 | 2008 | Fisher &amp; Paykel Healthcare Inc | The recall was initiated because certain lot dates of the RT240 Adult Breathing Circuit Kits manufactured on or before June 27, 2008 include a heated breathing circuit that may be more susceptible to damage when used in excess of the specified seven-day maximum duration of use, which may increase the risk of malfunction or fire. | Process change control | Fisher &amp; Paykel Healthcare, Inc. (FPH) issued an URGENT-PRODUCT RECALL letter (dated October 8, 2008) including a Customer recall letter for consignees/customers to Cardinal Health. Cardinal Health will, in turn, send the Customer recall letter dated October 10, 2008 to its consignees/customers. Fisher &amp; Paykel Healthcare (FPH) also issued a similar set of recall notification letters to Norco on October 15, 2008.Recipients of the recall letter must perform the following:1) Check all Cardinal Health stock of the product for the affected lots, 2) Immediately quarantine the affected product, 3) Complete a Device Recall Response Sheet, and 4) Fax or e-mail the completed Device Recall Response Sheet. Fisher &amp; Paykel Healthcare, Inc.'s "Dear Customer Letter" for Cardinal Health and Norco instructs consignees to complete a customer response (i.e. fax enclosed Customer Acknowledgement Form) to FPH. The customer letter also provides instructions for product return.Fisher &amp; Paykel Healthcare, Inc. recall communication instructions to Cardinal Health and Norco was for all affected product to be returned and quarantined by the distributors. The distributors will return any affected stock to Fisher &amp; Paykel Healthcare, Inc. for documented destruction. |  | 229320 | 229,320 units | Worldwide Distribution including US and country of Canada. | Terminated | March 26, 2012 | 1251 | 229320 | Hardware | Device Operation | Return | Remove or Replace
Z-0343-2009 | 50179 | No | LZP | Intraocular fluid. | Ophthalmic | Ophthalmic | PMA | AMO Healon D Ophthalmic Viscosurgical Device, Product Code: 10204011, Lot Number: UD30654, 30 mg/mL fill size |  |  | October 29, 2008 | 2008 | Advanced Medical Optics, Inc. | Endotoxin levels above specifications have been noted in some syringes tested for lot number UD30654 of the AMO Healon D Ophthalmic Viscosurgical Device. Endotoxin levels above the maximum USP level may be potential causes of an inflammatory response and/or TASS in patients following surgery. | Process design | On Thursday, October 30, 2008, AMO began distribution of an Urgent Device Recall Notice via Federal Express overnight delivery to the twenty-six (26) AMO Healon D OVD, Lot UD30654, customer accounts in the United States. The Urgent Device Recall Notice include&amp; a facsimile reply form.The Recall Notice informed the customers that AMO has received reports that some physicians using lot number UD30654 of the AMO Healon D OVD have observed an inflammatory reaction in their patients at the one-day post-operative visit. The customers were asked to undertake the following actions:(1) STOP using and remove from your inventory all units of AMO Healon D OVD, Lot UD30654. (2) Complete the included Facsimile Form making sure to note on the Facsimile Form the Quantities of AMO Healon D OVD present in your facility to be returned to AMO and fax to AMO Customer Service at 714-247-8722 within the next three (3) business days.The customers were also informed that an AMO Customer Service Representative will be contacting them to assist with the return of product affected by this action. If they have AMO Healon DC OVD units to be returned as part of this recall event and have not been contacted by an AMO Customer Service Representative, please call our toll free number at 1-877-AMO-4LIFE (1-877-266-4543) so that we can arrange the return of the product. |  | 1450 | 1450 | Nationwide to: MA, NJ, NY, WV, NC, SC, KY, OH, IN, MI, MN, OK, TX, WY OR, WA and Puerto Rico | Terminated | February 21, 2012 | 1210 | 1450 | Not_Computer
Z-0311-2009 | 46314 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Sarns Level Sensor II Pads. For use with Sarns Advanced Perfusion System I, Perfusion System 800 and Perfusion System 900. Terumo Cardiovascular Systems Corp., Ann Arbor, MI.; REF 195240. |  |  | January 13, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The level sensor holder may detach from the adhesive pad causing the sensor to lose contact with the venous reservoir, which may result in a detached sensor indication or false alarm. | Nonconforming Material/Component | Consignees were notified via recall letter (Urgent Medical Device Removal letter) dated 2/13/08 and requested to report stocks on hand so that replacements could be sent. |  | 31980 | 31,980 pads (25,620 to U.S. and 6,360 international) | Worldwide Distribution including USA and countries of Australia, Belgium, Canada, China, Dominican Republic, Japan, Saudi Arabia, Singapore and Yemen. | Terminated | July 07, 2009 | 541 | 31980 | Not_Computer
Z-0354-2009 | 44839 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 (4 inch diameter Roller Pump; Model 801040). |  |  | July 18, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | Overspeed and underspeed errors, erratic pump behavior, jerky operation at low RPMs, pump instability, pump slowdowns and pump stops due to overspeed max events may occur due to a grease leak onto the motor speed encoder disk. | Component design/selection | Consignees were notified of the problem via recall letter dated 7/18/08. The firm began conducting field upgrades on the roller pumps in 07/08. For questions or concerns, contact Terumo CVS Customer Service at 1-800-521-2818. |  | 1834 | 1,834 | Worldwide Distribution including USA and countries of Australia, Belgium. Canada, Columbia, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Viet Nam. | Terminated | July 07, 2009 | 354 | 5116 | Not_Computer
Z-0447-2009 | 48470 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Modular Shoulder Body Assembly, Left (SH-1540L-S) and Modular Shoulder Body Assembly, Right (SH-1540R-S). Product labeled in part, "Acumed 5885 NW Cornelius Pass Road Hillsboro, OR 97124-9432 " |  |  | May 13, 2008 | 2008 | Acumed LLC | The packages for Modular Shoulder Body Assembly Left (SH-1540L-S) and Modular Shoulder Body Assembly Right (SH-1540R-S) may contain assemblies for the opposite side as indicated. | Labeling mix-ups | The firm issued a Urgent Notice: Device Recall on 5/13/08. The firm notified consignees of the issue and asked them to return the product to Acumed with the enclosed product recall effectiveness form for customers. For further information contact Acumed customer service at 1-888-627-9957 |  | 20 | 20 units | World-wide distribution including states of GA, MD, CA, MO, and AL, countries of United Kingdom, South Korea, South Africa, and Spain. | Terminated | May 13, 2010 | 730 | 20 | Not_Computer
Z-0407-2009 | 49522 | No | GAZ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Mentor Aseptic Transfer Set, Cat. #350-8400 (contains BD 60mL Luer-Lok Syringe). Aseptic transfer set manufactured by Mentor Texas, LP, Irving, TX 75038 (BD 60 mL Luer-Lok syringe manufactured by Becton Dickinson, Franklin Lakes, NJ 07417). |  |  | August 22, 2008 | 2008 | Mentor Texas, Inc | The Mentor Aseptic Transfer Set contains a component, the BD 60mL Luer-Lok Syringe, which is under recall by Becton Dickinson due to a package integrity issue. | Package design/selection | Firm sent notification, along with a copy of the Becton Dickinson recall letter, to consignees on 08/18/08. Consignees asked to inspect inventory of Mentor Aseptic Transfer Sets for affected BD 60mL syringes and to complete and return the included Device Accountability Worksheet. For additional information, please contact Mentor Customer Service at 800-235-5731. |  | 32480 | 32,480 Aseptic Transfer Sets (4 syringes/set) | Worldwide Distribution; product sent to all states (except Alaska), Washington DC and Puerto Rico. Product also sent to the following countries: Australia, Canada, Costa Rica, Korea, New Zealand and South Africa. | Terminated | August 26, 2010 | 734 | 32480 | Not_Computer
Z-0458-2009 | 50065 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Gravity Compensating Accessory; |  |  | October 03, 2008 | 2008 | Integra LifeSciences Corp. | Integra NeuroSciences has become aware that certain Gravity Compensating Accessory implants have the potential for CSF leakage under certain conditions. | Process control | Urgent: Product Recall Notification letters, dated 10/3/08, were sent out on October 9, 2008. Letters explain the reason for the recall and the catalog and lot numbers affected. If the product has been implanted, it is recommended that the patient be monitored periodically using the standard of care determined appropriate by the patient's attending physician. If the product is found in inventory, customers are asked to quarantine the product and contact their Integra Sales Rep for assistance and for arrangements for product return. A Recall Acknowledgement and Return Form is to be completed and returned to Integra for reconciliation purposes. |  | 49 | 49 units | Worldwide Distribution - MI, IL, NY, IN, CA, IN and NC and Carnegie, Victoria, Melbourne, Australia. | Terminated | October 06, 2009 | 368 | 49 | Not_Computer
Z-0418-2009 | 49523 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Dual Port Wizard Low Profile Replacement Gastrostomy Device, Code 00202W, 16F, 2.4cm, 10cc, Sterile EO, Bard Access Systems, Inc., Salt Lake City, UT 84116. |  |  | September 23, 2008 | 2008 | C R Bard Inc | Gastrostomy device anti-reflux valve may allow leakage from the stomach. | Component design/selection | Customers notified by letter (Product Field Action letter) on September 23, 2008 and instructed to remove and return any affected product to Bard Access Systems, Inc (BAS). Customers were instructed in the recall letter to respond by completing and faxing the enclosed Reply Form and Inventory Reconciliation Form. For assistance, contact BAS customer service at 800-290-1689. |  | 1399 | 1399 units | Worldwide Distribution including US, Latin America and countries of Canada, Japan, Finland, Norway, Sweden, Austria, Belgium, Holland, Great Britain, Italy, France, Switzerland, Spain, Germany. | Terminated | November 08, 2011 | 1141 | 20263 | Not_Computer
Z-0460-2009 | 49458 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis). GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | September 25, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software anomalies result in patient safety issues involving patient jacket content intermittently becoming unintentionally out of synchronization with the images being displayed, and involving Default Display Protocols (DDPs) which are used to layout images when displaying a study. If the user does not check the Study Date Time of the exam on the image title bar, then they may interpret the curr | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation customers via letter (Product Safety Notification) dated 9/25/08 of the software anomalies resulting in patient safety issues involving the patient jacket content intermittently becoming unintentionally out of synchronization with the images being displayed, and involving Default Display Protocols (DDPs) which are used to layout images when displaying a study. If the user does not check the Study Date Time of the exam on the image title bar, then they may interpret the current exam as a historical study and vice versa. The letter provided the users with safety instructions to follow until the software is updated. Direct any questions to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-277-5240. |  | 546 | 546 units | Worldwide Distribution including USA and countries of Australia, Belgium, Brazil, Canada, Cayman Islands, Chile, Denmark, Ecuador, France, Germany, Hong Kong, Italy, Japan, Korea, Malaysia, Malta, Mexico, New Zealand, Singapore, Spain, Switzerland, Taiwan, United Kingdom and Venezuela. | Terminated | September 06, 2010 | 711 | 546 | Software | Display/Image | Software update | Software Update
Z-0485-2009 | 50075 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Enzymatic Creatinine Flex reagent cartridge, DF270. |  |  | September 29, 2008 | 2008 | Dade Behring, Inc. | Reagent may exhibit unflagged inaccurate patient sample results. | Component design/selection | The recalling firm telephoned all customers on 9/29/08 to inform them of the problem. Customers were issued a letter dated 9/29/08 (and a response form) with instructions. Customers were instructed to complete the response form and fax (302-631-7242) it to Siemens Healthcare Diagnostics. For questions or additional information, contact the Siemens Technical Solutions Center at 800-441-9250. |  | 12773 | 12,773 cartons | Nationwide Distribution. | Terminated | March 27, 2009 | 179 | 12773 | Not_Computer
Z-0483-2009 | 49586 | No | LTD |  |  | Toxicology | Contact ODE | RNAgentsÂż Total RNA Isolation System (Cat. # Z5110). Kit Box label/Box drawer label. For Laboratory Use. Promega Corporation, Madison, WI. USA. |  |  | June 04, 2008 | 2008 | Promega Corpopration | Promega Corporation received reports of leaking bottles of Phenol:Chloroform:Isoamyl Alcohol (Part # Z5112), a component of the RNAgentsÂż Total RNA Isolation System (Catalog # Z5110). Promega product Z5112 is a mixture of phenol, chloroform and isoamyl alcohol. If phenol contacts skin, it can cause burns and poisoning. The fumes of this product are also harmful. | Packaging | Consignees were sent a "Promega Urgent Product Alert-Risk of Injury" letter dated June 9, 2008. The letter described the product, problem and provided instructions on handling, inspection and disposition of product. In addition, E-mail notifications were sent to distributors and branch offices on May 29,2008. For questions and additional information, contact your local Promega representative or Promega Technical Service at 800-356-9526 and press "4", or e-mail at techserv@promega.com. |  | 836 | 836 | Worldwide Distribution: USA including states of IN, ME, NE, NY, WI, PA, NC, FL, OR, UT, MI, IL, AL, OR, GA, CA, MD, TX, MA, UT, MO, OH, and LA, and countries of Canada, Mexico, Brasil, China, Japan, Germany, Singapore, Great Britain, Australia and France. | Terminated | June 26, 2012 | 1483 | 836 | Not_Computer
Z-0454-2009 | 48660 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Vidiera NsP Nucleic Sample Preparation, Part Number: A22421, Version 1.0.41 |  |  | August 30, 2007 | 2007 | Beckman Coulter Inc | Incorrect results: under certain circumstances, when transfer volume is not within the set range, the Vidiera NsP software does not report sample exclusion and does not flag the excluded sample in the "Run Results" report. The software error associated with the excluded sample appears as "Liquid was not successfully transferred to well#..........the sample will be excluded from further proce | Software design | The customer notifications were initiated by Beckman Coulter on September 11, 2007 when all affected accounts were contacted via telephone, followed by a Product Corrective Action (PCA) was mailed via US Postal Service to all affected customers. The PCA informed the customers of the issue and instructed them to:(1) Review the "run Summary" event log for your samples. (2) Review the attached instructions describing the steps for identifying the samples affected by this error. (3) Complete and return the enclosed response form within 10 days so that we may maintain our records.The letter informed the customers that Beckman Coulter would provide a software update within 60 days. The letter also asked their customers to share the PCA with their staff and provide a copy of the PCA to any customers that they may have forwarded the affected products to.For questions regarding the PCA, please contact firm at 714-961-3634. |  | 6 | 6 | Nationwide distribution including states of PA, MI, UT, WA, IL, and CA. | Terminated | January 13, 2009 | 502 | 6 | Software | Output/Calculation | Software update | Software Update
Z-0338-2009 | 49420 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Boston Scientific iLab Ultrasound Imaging System, models 120INS and 240INS. Its intended use is for ultrasound examination of intravascular pathology. Manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | September 05, 2008 | 2008 | Boston Scientific Corporation | Computer malware: Products may be infected with a worm (computer malware) which could infect a computer network to which it may be connected. | Process control | Customer letters sent on Friday, Sept 5 via FedEx, 2008 to all customers. This letter will be followed up by a customer service visit. Package delivery will be tracked, as well as Reply Verification tracking forms.For further information, please contact firm at 508-652-5594. |  | 357 | 357 units | World-wide distribution including USA and countries of Europe, Middle East, Africa, and Japan. | Terminated | May 18, 2009 | 255 | 357 | Other | Output/Calculation | Service | Repair
Z-0504-2009 | 50157 | [6 empty] | Single-use sterile puncture attachment UA1256-U used with the BK Medical Ultrasounds: |  |  | September 08, 2008 | 2008 | Bk Medical | No 510k approved | Other | BK Medical notified accounts via a Recall Notice letter dated 9/03/08 to inform customers of the unapproved status and to remove the device from use and it will be replaced with B-K Medical's reusable metal guide UA1256. Customers were to fill out the attached "Reply Form" and return it via fax or the postage paid envelope. Any questions should be directed to B-K Medical Systems, Inc. at 1-800-876-7226. |  | 664 | 664 pieces | Nationwide Distribution --- including states of AR, CA, CO, CT, IL,IN, OK. OH, MA, MS, MD, NC, and NY. | Terminated | January 29, 2016 | 2699 | 664 | Not_Computer
Z-0503-2009 | 50151 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Maxi Move, Model NDA0200-20 Batteries Pack; BHM Medical Inc., 2001 Tanguay, Magog, QC, Canada J1X 5Y5 |  |  | November 05, 2008 | 2008 | Arjo, Inc. | The battery pack NDA0200-20, lot 070826D, sold with the Maxi Move patient lift or as replacement parts, has a defective connector that could lead to the inability of the battery pack to recharge and/or short circuit, resulting in smoke emission. | Component design/selection | Arjo sent the Urgent Device Recall Customer Notification dated 11/5/08 and Field Safety Notice (FSN) dated 10/31/08 to all affected Maxi Move customers on 11/5/08. The accounts were informed of the problem with the connector of the battery pack NDA0200-20, lot 070826D, sold with the Maxi Move patient lift (model numbers starting with KMC)or as replacement parts. They instructed the accounts to discontinue use of the battery pack if it does not recharge or has an intermittent electrical contact. The accounts were requested to complete and return to Arjo Inc., the enclosed customer response form indicating that they have received the notification package. The accounts will be contacted by an Arjo representative within a week to arrange the replacement of the battery pack and corresponding connectors (male part/plug) on the lift and battery charger. |  | 273 | 273 battery packs | Nationwide Distribution | Terminated | January 05, 2010 | 426 | 273 | Hardware | Device Operation | Replace component | Remove or Replace
Z-0482-2009 | 49563 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | ACE and ACE Alera Hemoglobin A1C : EZA1C Reagent, EZA1C Controls, and EZA1C Calibrators. |  |  | August 22, 2008 | 2008 | Alfa Wassermann, Inc. | Internal studies observed occasional unexpected outlier results in whole blood samples tested for Hemoglobin A1c using the EZA1c reagent on the ACE or ACE Alera clinical chemistry systems. | Component design/selection | An initial URGENT: CUSTOMER ADVISORY NOTICE dated August 22, 2008 was sent to all customers and distributors. The letter stated the problem and asked customers to cease using the EZA1C reagent, controls and calibrators. Customers were asked to hold any affected inventory while additional testing was conducted by Alfa Wassermann Diagnostic Technologies, LLC. (AWDT). A second communication, URGENT: CUSTOMER ADVISORY NOTICE (and Product Response Form) dated September 18, 2008 was sent instructing customers to destroy all product and package inserts in stock for the EZA1C products listed in the letter. Information was provided on receiving credit and an alternative product for use. Customers are requested to complete and return the Product Response Form by fax (888-646-2535) to AWDT indicating their receipt of the letter and the amount of product destroyed. Distributors' letters requested their customers be notified of the recall and that the Product Response Form be sent directly to AWDT. Direct questions to AWDT Customer Solution Center at 866-419-2532. |  | 69 | Reagent Lot F2481 = 264 kits and Reagent Lot F2603 = 69 kits. (Control Lot HBC00507-1=322 kits and Calibrator Lot HBS00907-1=262 kits). | Nationwide Distribution. | Terminated | April 15, 2009 | 236 | 69 | Not_Computer
Z-0412-2009 | 50098 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Arjo Maxi 500 Patient Lift with 2-Point Spreader Bar and Scale. A Non-AC-Powered Patient Lift. BHM Medical Inc., Magog, QC, Canada. Model KM560101, and Model KM560001 - Maxi 500 with 2-Point Spreader Bar, if it was equipped with the optional scale kit 700.05505. |  |  | October 30, 2008 | 2008 | Arjo, Inc. | If the spring pin is not properly reinstalled after maintenance, the pivot bolt could unscrew by itself within a limited period of time, resulting in a hanger bar detachment. | Employee error | Arjo, Inc. sent an Urgent Device Recall Customer Notification letter dated 10/29/08 to all affected end users to the attention of the Administrator/Risk Manager via UPS 2nd day on the 10/30/08. The letters informed the accounts of the potential for the hanger bar to detach from the scale if the spring pin is not replaced into the pivot bolt after performing maintenance on the lift. The accounts were advised to review the enclosed Urgent Field Safety Notice, note the affected model and serial numbers, and inspect their lifts for the presence of the spring pin in the pivot bolt. If the user finds that the spring pin is missing, they are instructed to immediately stop using the lift and contact their Arjo representative. The accounts were also instructed to remove previous versions of the Operating and Product Care instructions, and replace them with the current version 001.20815.EN Rev3 dated October 2008 provided with the letter, ensuring that the proper personnel are made aware of the changes outlined in the Field Safety Notice. Direct questions to the Arjo, Inc. Quality Department at 800-323-1245, ext. 6118. The accounts were also instructed to complete and fax (888-594-2756) to Arjo, Inc. the enclosed Customer Response Form, providing the facility name and address, contact name and phone number, the numbers of affected units, and acknowledgment of receipt and understanding of the Field Safety Notice. |  | 46 | 44 lifts, 2 scale kits | Nationwide Distribution. | Terminated | August 19, 2010 | 658 | 427 | Not_Computer
Z-0481-2009 | 49562 | No | LDJ | Cannabinoid test system. | Toxicology | Toxicology | 510(k) | Cannabinoids Intercept Micro Plate EIA 100 plate kit |  |  | September 08, 2008 | 2008 | OraSure Technologies, Inc. | Readings, high: intermittent high absorbance readings. For further information, please contact the firm at 610-882-1820. | Other | On 9/8/08, the recalling firm telephoned the customer to inform them of the problem. As a follow up to the telephone call, the recalling firm issued a recall letter dated 9/11/08 to the customer. |  | 1 | 1 kit | The product was shipped to one customer in New York. | Terminated | December 18, 2008 | 101 | 1 | Not_Computer
Z-0498-2009 | 50103 | No | GZL | Depth electrode. | Neurology | Neurology | 510(k) | FC4000 MicroTargeting Electrode Kit 5x |  |  | October 20, 2008 | 2008 | FHC, Inc. | Mislabeled: May contain an electrode longer than labeled size | Pending | FHC notified customers via an Urgent: Medical Device Recall letter dated 10/20/08 advising users to remove the recalled electrodes from inventory and set the product aside for inspection by your FHC representative. Customers were asked to document their actions by completing the enclosed form and faxing it back to FHC at 207-666-8539. |  | 3 | 3 Kits (30 electrodes-10 electrodes/kit) | Nationwide Distribution --- including states of CA, CO, FL, LA, NM, PA, TN, and PR. | Terminated | May 02, 2012 | 1290 | 62 | Not_Computer
Z-0502-2009 | 50125 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | STA Neoplastine CI Plus 10; Product Catalogue Number: 0667 |  |  | October 20, 2008 | 2008 | Diagnostica Stago, Inc. | Internal investigation has determined the potential lack of homogeneity between product vials. | Process control | Product Recall Notice letters were sent to all customers on October 20, 2008 vis US mail. The letter stated the reason for the notice, the corrective action, and replacement process. Customers are to confirm that they received the Recall Notice by completing the attached "Product Recall Response Form" and return by fax to the Regulatory Affairs Department at 973-867-6028 or 973-695-0095. The firm is prepared to immediately implement the following corrective action: 1) perform QC testing in accordance to directions from the package insert on each vial and destroy vial and replace inventory if determined product was unsatisfactory OR 2) return inventory for replacement with new lots of inventory. If customers have any questions or need any technical assistance, they are to call the DSI Hotline at 1-800-222-2624, option 1. |  | 31207 | 31,207 units | Nationwide Distribution. | Terminated | June 24, 2010 | 612 | 32487 | Not_Computer
Z-0509-2009 | 50172 | No | KKY | Polytetrafluoroethylene with carbon fibers composite implant material. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Custom Cranial Implant Kit, small. Distributed by Stryker Craniomaxillofacial, Kalamazoo, MI; Catalog no. 54-00101. |  |  | October 24, 2008 | 2008 | Stryker Craniomaxillofacial Division | Lack of assurance of sterility. | Process control | The firm's sales representatives and consignee risk managers, as well as implanting surgeons, were notified individually of the problem via recall letter dated 10/24/08, advised of the risk for serious infections, and instructed to follow-up with patients for infections for at least six months after surgery. The recall letter also included details for identifying and returning any remaining implant kits. Direct questions to Stryker's RA/QA Representative 269-324-5346 ext. 4258. The firm issued a press release concerning this recall on 12/23/08. |  | 322 | 322 of all products. | Nationwide Distribution. | Terminated | December 31, 2009 | 433 | 322 | Not_Computer
Z-0496-2009 | 50091 | No | DSQ |  |  | Cardiovascular | PMA | Thoratec Heartmate II Left Ventricular Assist System (LVAS). Model/Catalog Number 1355 and Model/Catalog Number 102139. |  |  | October 24, 2008 | 2008 | Thoratec Corp | The firm has become aware that, over time, wear and fatigue of the percutaneous lead connecting the HeartMate II Left Ventricular Assist System (HM II LVAS) blood pump with external system controller may result in damage that has the potential to interrupt pump function and may require a reoperation to replace the pump. | Device Design | An initial "Urgent: Medical Device Correction" letter dated October 21, 2008 and a follow-up letter dated October 24, 2008 were sent to affected HeartMate II Left Ventricular Assist System (HM II LVAS) customers. The letters identified the problem, immediate action and preventive action to be taken by customers and recommended that the pump be replaced as soon as possible if damage to the percutaneous lead is confirmed. If customers suspect that a HM II LVAS patient may have a damaged percutaneous lead, please contact Thoratec Technical Services (800-456-1477) for assistance. A return response (Acknowledgement Form) was included with the letter for consignees to complete and fax it to Thoratec Regulatory Affairs at 925-847-8628. Direct questions to the Thoratec Regulatory Affairs Department at 925-730-4135 (Monday-Friday, 8 a.m.-5 p.m. PST). Clinicians and patients with questions may contact the firm at 800-528-2577, or if calling outside the USA, 925-847-8600 (7 days a week, 8 a.m.-5 p.m. PST). Copies of the press release may be found on the firm's website, www.thoratec.com, under Investor Relations/Press Releases. |  | 2335 | 2,335 units | Worldwide Distribution. | Terminated | January 26, 2011 | 824 | 2335 | I/O | Treatment/Delivery/Therapy | Replace | Remove or Replace
Z-0500-2009 | 50115 | No | LTJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | FastPack(R) Total PSA Immunoassay; Catalog #: 25000001 |  |  | September 29, 2008 | 2008 | Qualigen Inc | The recall was initiated after the firm investigation of customer complaints of falsely elevated results determined the product was not meeting product performance expectations. | Other | The recall was initiated with Qualigen faxing a copy of the Product Notification letter, dated 9/29/2008, to all customers who have received any sublot of the affected product . Faxes were sent beginning 9/29/08. The recall letter informed the customers that the affected product was not meeting the firm's expectations for product performance. Customers were instructed to please stop using the affected lots immediately as they have been noted to produce falsely elevated results and they should confirm any result that appears to be inconsistent with "a patient's medical history, clinical examination, and other findings". In addition, they may wish to retest any patients that were run with a different FastPack lot.Beginning on 9/30/08, all customers were contacted by phone to confirm that they received the recall letter.The customers were shipped replacement products at no cost and as part of the product replacement process, they were instructed to return any unused Total PSA FastPacks/Kits of lot 0807002 to Qualigen. |  | 720 | 720 kits | Worldwide Distribution --- including USA and countries of Canada, UK , Italy, Switzerland, Sweden, &amp; South Korea. | Terminated | June 02, 2011 | 976 | 720 | Not_Computer
Z-0593-2009 | 50108 | No | MQE | Assisted reproduction needles. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Wallace Oocyte Recovery Needles,Oocyte Recovery sets, 750 mm tube with flushing connector, Product Number: KÂżLN-750, and KÂżLN-750. |  |  | October 23, 2008 | 2008 | Irvine Scientific Sales Co Inc | Smiths Medical International became aware that certain lots of Wallace oocyte Recovery Sets have a damaged needle tip. | Packaging | The recall was initiated by Smiths Medical on October 22, 2008 with the forwarding a recall notification to Irvine Scientific. Smiths Medical International sells the products to Irvine Scientific for distribution in the U.S. Irvine Scientific is the U.S. Agent. Irvine Scientific issued an Urgent Field Safety Notice Letter, dated 10/22/2008, with attached fax-back Customer Response form to their customers via US Postal service beginning on October 24, 2008. Irvine Scientific requested their customers to complete an inventory summary on the Customer Response form and contact the firm to arrange for the return of any remaining product in their inventory to the firm. |  | 71 | 71 units | Nationwide Distribution --- including states of CA, FL, PA, IL, OR &amp; UT. | Terminated | September 02, 2011 | 1044 | 71 | Not_Computer
Z-0486-2009 | 50086 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | FHC microTargeting Platform DBS Measuring Fixture, |  |  | October 08, 2008 | 2008 | FHC, Inc. | Measuring fixture is incorrectly graduated. | Process control | FHC provided letter notification to FHC Field Representatives on 10/08/08 to hand deliver to customers. FHC Field Representatives will remove incorrectly marked product and provide replacement product. Direct questions to an FHC representative or FHC Quality and Regulatory Affairs at 207-666-5651, ext. 2210. |  | 5 | 5 units | Nationwide Distribution including states of CO, FL, NC and NM. | Terminated | May 02, 2012 | 1302 | 5 | Not_Computer
Z-0524-2009 | 49077 | No | JLS | Progesterone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE/IMMULITE 1000 Progesterone (LKPG 1,5) |  |  | July 21, 2008 | 2008 | Siemens Medical Solutions Diagnostics | Siemens Healthcare Diagnostics has confirmed a long-standing high bias with LKPG 1, 5. (The bias is approximately 25% at 10 - 20 ng/mL as opposed to approximately 3% at 1-3 ng/mL - for more detail see the bias plot that is included with this notice.). | Other | Siemens initiated the recall on 07/22/2008 by forwarding A Customer Bulletin by fax or mail to all affected Siemens Healthcare Diagnostic Branches for communication with the affected customers.The Customer Bulletin informed their customers that the firm is conducting a field correction and requested that they discontinue using any IMMULITE/IMMULITE 1000 kit lot prior to LKPG 1,5 330. They were also instructed to complete attached Completion Notification Form and return to their local Siemens Healthcare office or distributor as soon as possible. The were informed that replacement kits will be provided once the completed forms have been received.The Customer Bulletin recommend their customer discuss the content of the letter with their laboratory director regarding the need to review previous test results, conduct patient follow up, and/or repeat testing. The bias should not impact the in-vitro fertilization clinical application because the average change for samples within the range of 1-3 ng/mL is less than 5%. Although there is a greater degree of shift at higher concentrations, the clinical impact is considered to be relatively low when using the test for the diagnosis of ectopic pregnancy and luteal phase defects because in these situations, the progesterone result should always be considered in conjunction with other tests. PROGESTERONE DETERMINATIONS USED IN THE DIAGNOSIS OF PLACENTAL INSUFFICIENCIES OR FOR MONITORING PROGESTERONE THERAPY MAY HAVE BEEN AFFECTED BY THE HIGH BIAS.The letter also indicated SIEMENS IS INTRODUCING A RE-CALIBRATED IMMULITE PROGESTERONE KIT BEGINNING WITH KIT LOT LKPG1, 5 300, WHICH WILL BEGIN ROUTINELY SHIPPING AUGUST 15 TH. AT THAT TIME, KITS WITH THE PREVIOUS CALIBRATION WILL ONLY BE AVAILABLE UPON REQUEST AND AVAILABILITY OF INVENTORY FOR CUSTOMERS WHO STILL NEED TO PERFORM BASELINING STUDIES. An Important Notice, which is reproduced at the end of this notice below, will be included with the recalibrated kits |  | 12908 | 12,908 | Worldwide Distribution --- including USA and countries of Argentina, Aruba, Australia, Austria, Bangladesh, Bolivia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Cyprus, Czech Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Jordan, Korea, Kuwait, Lebanon, Mexico, Netherlands, Nicaragua, Pakistan, Peru, Philippines, Poland, Portugal, Puerto Rico, Republic of Panama, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, United Arab Emirates (UAE), United Kingdom, Uruguay, and Venezuela. | Terminated | July 25, 2012 | 1465 | 12908 | Not_Computer
Z-0479-2009 | 49557 | No | LDJ | Cannabinoid test system. | Toxicology | Toxicology | 510(k) | Micro Plate EIA Oral Fluid Cutoff Calibrator which is a component of the Intercept Micro Plate EIA kit. |  |  | September 08, 2008 | 2008 | OraSure Technologies, Inc. | Results, false-positive test: low % displacement result of 36.0% (spec 39-60%). For further information, please contact the firm at 610-882-1820. | Nonconforming Material/Component | The recalling firm issued a letter entitled "URGENT PRODUCT RECALL NOTICE" dated 9/8/08 to their customers to inform them of the problem and the need to return the product. |  | 2198 | 2198 units | Worldwide Distribution --- USA including states of AZ, CA, CT, DE, FL, GA, ID, KS, KY, MA, MD, MI, NC, NH, NV, NY, PA, VA, and WY, and countries of United Kingdom, Italy, Germany, and Canada. | Terminated | May 05, 2009 | 239 | 2198 | Not_Computer
Z-0476-2009 | 49550 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | VASOVIEW 4 Endoscopic Vessel Harvesting System, with VASOVIEW Flexible Endoscopic Tools in the following configurations: VH-10242: Uniport plus dissection cannula, short port blunt tip trocar, flexible endoscopic scissors; and VH-10243: Uniport plus dissection cannula, short port blunt tip trocar, flexible endoscopic bisector. |  |  | September 09, 2008 | 2008 | CTS, Inc. dba Guidant Cardiac Surgery | Packaging unsealed: Sterility of product may be compromised due to packaging issue. For further information, please contact the firm at 408-635-6874. | Packaging process control | The firm issued a letter entitled "URGENT FIELD ACTION - IMMEDIATE ACTION REQUIRED *** VASOVIEW(TM) 4 Endoscopic Vessel Harvesting System (Kit Only)" via federal express to its consignees on September 9, 2008. The firm has asked its consignees to return the product. Follow-up calls were made to ensure recall effectiveness. |  | 7935 | 7,935 units | Worldwide Distribution --- including USA and country of Australia. | Terminated | December 08, 2010 | 820 | 7935 | Not_Computer
Z-0459-2009 | 50092 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Sun Med GreenLine/D Laryngoscope Blades Compatible with Fiber Optic Green System Handles, Sterile Disposable. Contents 20. |  |  | April 27, 2007 | 2007 | Sun Med, Inc. | A breakage problem of the acrylic bundles at both distal and proximal ends of the light tube for the MD Laryngoscope Blade GreenLine D Mac 3 has been detected. If breakage occurs during use, there is the potential for leaving approximately a one inch piece of acrylic tube in the patient's mouth. | Device Design | SunMed sent initial recall letter on April 27, 2007 to all distributors requesting that kits containing the MacIntosh size 3, consisting of lots number Gj and Ha be returned to facility. Replacement information was included in letter. On December 12, 2008, the firm issued a Device Recall Press Release with additional information on the potential for the tube to break during use leaving an approximate 1 inch piece of the acrylic tube in the patients mouth and subsequently requiring medical or surgical intervention to prevent serous injury to the patient. SunMed has notified its distributors and customers by certified mail and has arranged for return and replacement of all recalled products. Consumers with question may contact their distributor of the GreenLine/D or SunMed at 1-800-433-2797. |  | 2682 | 2,682 boxes (20 per box) | Nationwide Distribution | Terminated | July 06, 2010 | 1166 | 2682 | Not_Computer
Z-0621-2009 | 50487 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Flextome Cutting Balloon Device Over-the-Wire Delivery System, Coronary Atherotomy System, Over-the-Wire Delivery System, Catalog No. CB0340006, 4.00mm, Sterile. Made in Ireland: Interventional Technologies Europe Ltd., A subsidiary of Boston Scientific Letterkenny, County Donegal, Republic of Ireland. US Customer Service 888-272-1001. |  |  | November 21, 2008 | 2008 | Boston Scientific Corporation | For one batch/lot of flextome Cutting Balloon (OTW), an incorrect compliance chart was packaged inside the sterile pouch. For the affected batch/lot, the compliance chart packed inside the pouch of the 4.0mm diameter Flextome Cutting Balloon device is for a 3.5mm diameter Flextome Cutting Balloon device. This labeling discrepancy could create a procedural delay as users seek to explain the discr | Mixed-up of materials/components | An Urgent medical Device Field Correction, dated November 21, 2008, was sent to affected customers/hospitals for delivery on November 24, 2008. The letter described the issue; identified product affected; use or further distribution of the affected product should be ceased immediately; and a correct compliance chart included with the notification is to be secured to the outer sterile pouch. Affected product is not asked to be returned as the correct compliance chart is enclosed and is to be secured to the outer sterile pouch. A Reply Verification Tracking Form is to be completed and faxed to Customer Service Call Center. |  | 75 | 75 | Nationwide Distribution --- including states of AL, AR, AZ, CA, CO, GA, HI, ID, IL KS, MA, MD, MI, MN, MO, MS, NV, NY, OH, PA, TN, TX, UY, VA, and WI. | Terminated | December 10, 2011 | 1114 | 75 | Not_Computer
Z-0622-2009 | 50489 | No | GEA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Genicon 5mm Pyramidal Trocar, Sterile, Genion 6869 Stapoint Court, Suite 114 Winter Park, FL U.S.A. 32792, Rx Only. Catalog Numbers: 100-005-001 &amp; 100-005-002. |  |  | September 25, 2008 | 2008 | Genico, Inc., dba Genicon | The stainless steel tips on the Genicon 5mm Pyramidal Trocar exhibited contamination that resembled rust and there was also pitting on the tip of the trocar. | Manufacturing material removal | Consignees were notified with a letter, dated 9/19/2008, by e-mail and fax on 09/25/2008. The letter stated a description of the issue, analysis and reliability data, indications of Device Malfunction, and recommendations. There was also a request for all defective devices to be returned to the distributor or Genicon for replacement. Genicon Customer service was to be contacted at 407-657-4851. |  | 2965 | 2,965 units (1340 OUS and 1625 US) | Worldwide Distribution --- including USA and countries of Africa, United Kingdom, Denmark, Colombia, Belgium, Italy, Finland, Hong Kong, and Switzerland. | Terminated | January 09, 2009 | 106 | 2965 | Not_Computer
Z-0445-2009 | 44838 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 100V -120V, 15 A (circuit breaker), 50/60 Hz (20A power source required) base; Model 801763. |  |  | April 17, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The power supply may fail to charge the batteries due to various hardware malfunctions. | Nonconforming Material/Component | Terumo notified U.S. consignees via Urgent Medical Device Correction: Safety Advisory recall letter dated 4/17/08 and international consignees were notified via email on 4/24/08 informing them of the potential problem. |  | 256 | 256 | Worldwide distribution: USA, Australia, Belgium, Canada, Colombia, Dominican Republic, Honduras, Hong Kong, Indonesia, Japan, Jordan, Kuwait, Malaysia, Mexico, Philippines, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, United Arab Emirates, and Viet Nam. | Terminated | July 07, 2009 | 446 | 417 | Hardware |  | Notification | Safety Notice/Insructions
Z-0449-2009 | 49712 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | Hologic, Inc., 10 Year Fracture Risk Questionnaire Option for QDR X-Ray Bone Densitometers Software |  |  | September 24, 2008 | 2008 | Hologic, Inc. | Software error may lead to a high estimate of major fracture probability. | Software design | Hologic notified customers by telephone on 9/23 and 24/08 and instructed them to not use the software. A customer notification letter was sent via FedEx to customers on September 25, 2008.The letter requested de-installation of the software and return of the software to Hologic in a pre-paid envelope. |  | 15 | 15 software discs | Nationwide distribution in: CO, MA, MN, NM, NY, NC, OR, OH, PA, and TX. Also distributed in the UK. | Terminated | March 07, 2016 | 2721 | 15 | Software | Output/Calculation | Deinstall and return | Remove or Replace
Z-0467-2009 | 49495 | No | KZD | Pressure infusor for an I.V. bag. | General Hospital | General Hospital | 510(K) Exempt | CASMED 500cc Unifusor Classic with piston gauge and stopcock valve |  |  | September 20, 2008 | 2008 | CAS Medical Systems, Inc. | Leaks: Infusion Cuffs may exhibit air leakage at the union of the tube and bag, and not maintain pressure. For further information, please contact the firm at 203-488-6056. | Equipment maintenance | CAS Medical Systems notified accounts by letter dated 9/19/08 requesting return and sub-recall be conducted to their accounts. The letter was entitled "UNIFUSOR(R) INFUSION CUFFS AND PRIVATE LABEL (CASMED-MANUFACTURED) INFUSION CUFFS *** URGENT PRODUCT CORRECTION NOTICE." For further information, please contact the firm at 203-488-6056. |  | 6384 | 6,384 pc | NationwideForeign: Malaysia, Austria, Belgium, Czech Rep., France, Germany, Greece, Italy, Norway, Spain, Sweden, Switzerland and the UK. | Terminated | August 05, 2010 | 684 | 80864 | Not_Computer
Z-0480-2009 | 49560 | No | KHE | Occult blood test. | Hematology | Hematology | 510(k) | INSURE(R) QUIK FIT(TM) DEVELOPER KIT, Lot number B1478, Product Numbers: 30025 (InSure Quik FIT Developer Kit) and 50025 (INSURE(R) QUIK FIT(TM) containing Developer Kit and Patient Collection Kits. |  |  | October 03, 2008 | 2008 | Enterix, Inc. | False readings: Reduced sensitivity could affect the test line area and/or control line on the test strips, which may lead to inaccurate results. For further information, please contact Enterix, Inc. at 1-800-531--3681. | Other | Medical Device Recall Letters were sent on October 3, 2008 to all customers by certified mail. The letters were entitled "URGENT *** MEDICAL DEVICE RECALL (U.S. ONLY). Questions are to be directed to Enterix at 1-800-531-3681, |  | 1125 | 45 kits (includes 25 patient collection kits and 1 developer kit) | Product was distributed to medical facilities and laboratories throughout the US. | Terminated | March 26, 2009 | 174 | 1125 | Not_Computer
Z-0617-2009 | 50220 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | DeVilbiss Model 515A Oxygen Concentrator |  |  | October 30, 2008 | 2008 | Sunrise Medical | potential for sieve beds to have been incorrectly assembled | Device Design | On 10/30/08, the recalling firm began to telephone customers of the problem. The recalling firm followed up with Urgent: Medical Device Recall letters, dated 11/7/2008. The letters requested that users cease the use of the affected device and immediately contact their provider to make arrangements to have the Concentrator replaced. The firm has set up a toll free line to assist customers with identification of affected units and the return process. |  | 428 | 428 units | Worldwide Distribution --- USA including states of AL, CA, FL ,GA, IN, IL, MA, MD, MI, MS, NC, NE, NH, NM, NY, PA, SD, TN, TX, VA, and WI, and countries of Canada, Columbia, Germany, Mexico, Netherlands, Singapore, and United Kingdom. | Terminated | October 27, 2012 | 1458 | 428 | Not_Computer
Z-0604-2009 | 50154 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | MassTRANSIT Infusion Catheter Kit, 605-731K &amp; 605-780K, Cordis Neurovascular |  |  | June 30, 2008 | 2008 | Cordis Corporation | Two lots of Cordis Neurovascular Masstransit Infusion Catheter Kits has the potential for sterile PTFE contamination within the inner lumen of the catheter. | Pending | A Medical Device Recall letter, dated 6/30/2008, and Acknowledgement Form were sent overnight. The firm asked customers to set aside any affected product in a manner that ensures it will not be used. Customers are asked to review, sign and return the enclosed Acknowledgement Form directly to Cordis Neurovascular. Customers can either return any product with reference to the RA# on the form, or contact their local sales reps to facilitate return of the affected product. Customers will receive a credit for the units they return. If any affected product has been forwarded to another facility, the facility should be contacted to arrange return of the product. If there are any additional questions, Customer Service may be also be contacted at 800-551-7683. |  | 27 | 27 | Nationwide Distribution --- including states of Colorado, Connecticut, Florida, Minnesota, Nebraska, Nevada, Ohio, Oregon, and Virginia. | Terminated | January 02, 2009 | 186 | 27 | Not_Computer
Z-0618-2009 | 50225 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Flexi-Seal Fecal Management System Advanced Odor Control Kit. |  |  | November 20, 2008 | 2008 | ConvaTec | It was discovered that a portion of the inner nylon sleeve of the catheter was visible at the balloon end of a sample device. | Nonconforming Material/Component | Urgent Product Correction Letters were sent by first class mail on November 18, 2008 to Healthcare Professionals and ConvaTec Reps. The reason for the recall was identified along with possible risk to patient. Healthcare professionals were asked to use sufficient lubricating jelly during the insertion process with special care given to those who have a tendency to bleed or are on anti-coagulant medication. Physician should be notified if rectal bleeding or abdominal symptoms develop. ConvaTec reps were asked to contact all health care professionals to whom product or samples were given and reinforce the importance of proper lubrication and following the directions. A return form is to be completed by ConvaTec reps for each visit indicating the name of the Health Care Provider and that the they personally received notice of the product correction. All remaining sample products are to be returned to Besta, Franklin, WI. |  | 795 | 795 | Worldwide Distribution --- including USA and country of Canada. | Terminated | March 09, 2009 | 109 | 795 | Not_Computer
Z-0610-2009 | 50180 | No | JEY | Bone plate. | Dental | Dental | 510(k) | Stryker Leibinger New Generation/Universal Mandibular System, Mand. Recon Plate, Straight, w. Template, Stryker Leibinger GmbH &amp; Co., Freiburg, Germany; REF # 55-28911. |  |  | September 19, 2008 | 2008 | Stryker Leibinger USA | A template is supposed to accompany the product, but it was not included with this lot. | Process control | Consignees were notified via an Urgent Medical Device Recall letter dated September 2008 and requested to examine their inventory for the affected lot and return the recalled product. Customers are to fill out the attached forms to confirm the return-shipment and send it back. The firm will inform customers immediately regarding replacement deliveries. |  | 23 | 23 | Nationwide Distribution | Terminated | December 15, 2009 | 452 | 23 | Not_Computer
Z-0508-2009 | 48177 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Novel PEEK VBR XS 16-22 MM, S-5, 26MM, Part Numbers : 64753-126. |  |  | August 16, 2007 | 2007 | Alphatec Spine, Inc. | Firm confirmed that two lots were mixed during manufacturing and were incorrectly labeled with the wrong part number &amp; lot number. They were incorrectly laser marked with the wrong part numbers and sizes. Two lots were mixed from the following lot numbers: Part Number 64733-120 Lot Number 611884, and Part Number 64753-126 and Lot Number 611881. | Process control | The recall was initiated on August 16, 2007 with Alphatec Spine contacting each of the four direct distributors by telephone. This was followed with a Product Recall Notice, dated August 27, 2007, that was mailed to all affected customers. The distributors were informed that the firm was conducting a voluntary recall because two lots were mixed during manufacturing and were incorrectly labeled with the wrong part number &amp; lot number. The firm asked the customers to return the product, and were informed that replacement product will be sent back. |  | 4 | 4 units | Nationwide Distribution --- including states of AZ, FL, NY &amp; TX. | Terminated | January 12, 2009 | 515 | 7 | Not_Computer
Z-0522-2009 | 49065 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Diagnostic Ultrasound System, Model HD3. Product labeled HD3 Ultrasound System, Distributed by Philips Ultrasound, 22100 Bothell-Everett Highway, Bothell, WA 98021. |  |  | August 12, 2008 | 2008 | Philips Medical Systems | The system has a software defect that affects the display of Estimated Fetal Weight (EFW) Growth Percentile result values. | Software design | On 8/12/08, the firm sent an Urgent - Medical Device Correction letter, dated July 2008, to all affected customers via UPS. The letter describes the reason for the recall, informs them that Philips Ultrasound is correcting the error, states that the software is under development, will be available in September 2008 and will contact the customer to upgrade their system. "In the interim, Philips recommends that you only use the Hadlock or Osaka growth tables to perform your EFW Growth Percentile calculations. You should make the appropriate selections in the Analysis Setup menu on your system." |  | 4279 | 4279 units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Austria, Bahamas, Bahrain, Belgium, Bosnia, Brazil, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Dominican Rep, Ecuador, Egypt, El Salvador, France, Germany, Greece, Guam, Guatemala, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Jordan, Kuwait, Malaysia, Martinique, Mexico, Moldova, Morocco, Netherlands, New Zealand, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Syrian Arab Rep, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Utd Arab Emirr., Venezuela, Viet Nam, Canada, Afghanistan, Albania, Argentina, Australia, Austria, Belgium, Bolivia, Bosnia, Brazil, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Rep, Ecuador, El Salvador, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Italy, Japan, Korea Rep, Latvia, Lebanon, Lithuania, Malaysia, Mauritius, Mexico, Moldova, Morocco, Namibia, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Paraguay, Peru, Poland, Portugal, Puerto Rico, Romania, Serbia, Singapore, SLOVAKIA, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Uganda, Ukraine, United Kingdom, Uzbekistan, Venezuela, and Virgin Isl (Br). | Terminated | July 14, 2011 | 1066 | 6181 | Software | Output/Calculation | Software update | Software Update
Z-0472-2009 | 49517 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Ultra Laboratory System Software; for recording, reporting and distribution of lab results. GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | September 22, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software computer, error: The abnormal result flag is not being sent in the HL7 messages on results that are lower than the normal range. For further information, please telephone GE Healthcare Integrated IT Solutions at 800-437-1171. | Software design | GE Healthcare notified their Centricity Ultra Laboratory Information System customers via letter dated 9/19/08. The letter was entitled "PRODUCT SAFETY NOTIFICATION." The letter states that the software correction was installed or will be installed to correct the abnormal results flag not being sent in the HL7 on results that are lower than the normal range. The letter summarized the safety concern and the action taken. Any questions were directed to GE Technical Support at 888-778-3373 or 905-479-1249. The accounts were requested to complete the enclosed customer reply form, supplying the software version number, and fax it back to 847-277-5240. |  | 3 | 3 units | Globally - United States including the states of New York and Texas; and the country of Canada. | Terminated | May 18, 2010 | 603 | 3 | Software | Alarm/Message | Software update | Software Update
Z-0523-2009 | 49065 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Diagnostic Ultrasound System, Model HDI4000. Product labeled HDI4000 Ultrasound System, Distributed by Philips Ultrasound, 22100 Bothell-Everett Highway, Bothell, WA 98021. Product also marketed as Philips 4000 New Standard, Philips 4000 System and Multibeam System. According to the firm, these units have the same user manual as the HDI4000 system. The different names are used internally as a description for the different configurations of model HDI 4000. Over the product lifecycle, the product configuration changed as follows: MULTIBEAM SYSTEM : This was the original HDI 4000 system branded with the former company name 'ATL' and labeled, "Distributed by ATL Ultrasound Bothell, WA. USA 98041"; PHILIPS 4000 SYSTEM :This configuration changed the legacy 'ATL' brand / colors to the 'Philips' brand / colors; PHILIPS 4000 NEW STANDARD: This Philips brand configuration eliminated hardware that supported the 'CV'. The CV option was initially standard, but later became optional depending on customer preference. The product ID label is the same for both PHILIPS 4000 SYSTEM and PHILIPS 4000 NEW STANDARD. |  |  | August 12, 2008 | 2008 | Philips Medical Systems | The system has a software defect that affects the display of Estimated Fetal Weight (EFW) Growth Percentile result values. | Software design | On 8/12/08, the firm sent an Urgent - Medical Device Correction letter, dated July 2008, to all affected customers via UPS. The letter describes the reason for the recall, informs them that Philips Ultrasound is correcting the error, states that the software is under development, will be available in September 2008 and will contact the customer to upgrade their system. "In the interim, Philips recommends that you only use the Hadlock or Osaka growth tables to perform your EFW Growth Percentile calculations. You should make the appropriate selections in the Analysis Setup menu on your system." |  | 1902 | 1902 Units | Worldwide Distribution --- including USA and countries of Argentina, Australia, Austria, Bahamas, Bahrain, Belgium, Bosnia, Brazil, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Dominican Rep, Ecuador, Egypt, El Salvador, France, Germany, Greece, Guam, Guatemala, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Jordan, Kuwait, Malaysia, Martinique, Mexico, Moldova, Morocco, Netherlands, New Zealand, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Syrian Arab Rep, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Utd Arab Emirr., Venezuela, Viet Nam, Canada, Afghanistan, Albania, Argentina, Australia, Austria, Belgium, Bolivia, Bosnia, Brazil, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Rep, Ecuador, El Salvador, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Italy, Japan, Korea Rep, Latvia, Lebanon, Lithuania, Malaysia, Mauritius, Mexico, Moldova, Morocco, Namibia, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Paraguay, Peru, Poland, Portugal, Puerto Rico, Romania, Serbia, Singapore, SLOVAKIA, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Uganda, Ukraine, United Kingdom, Uzbekistan, Venezuela, and Virgin Isl (Br). | Terminated | July 14, 2011 | 1066 | 6181 | Computer
Z-0544-2009 | 49457 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo Dynamics Workplace, version 5.0, Siemens Medical Solutions, Ann Arbor, MI; 10090310. |  |  | July 23, 2008 | 2008 | Siemens Medical Solutions, USA, Inc | If the text is manually modified and additional phrases added as a result of worksheet interactions, some of the phrases may not populate the report. | Software design | Consignees were notified of the problem via a Customer Safety Advisory Notice dated 7/23/08, which provided guidance on how to minimize the risk of the problem occurring, and told customers that a service patch will be installed when it becomes available. |  | 306 | 306 of all products. | Nationwide Distribution. | Terminated | October 18, 2010 | 817 | 306 | Not_Computer
Z-0595-2009 | 50111 | No | LYT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ossur 3" Ceramic Tip Pin, 4 pack, Model Number: 517DS |  |  | September 26, 2008 | 2008 | Ossur | Ossur has initiated this voluntary recall as a precautionary measure because the firm is unable to guarantee the sterility of the affected products. | Packaging process control | Ossur initiated contact with the customers to locate the pins on Sept. 26, 2008with the Product Manager for this device contacting each consignee by phone.The customers were informed the recall was initiated because the firm was unable to guarantee the sterility of the affected products. Customers were asked to check their inventory for the affected parts. If products were located, they were informed that an Ossur customer service representative will be emailing you with return information any they are to return these pins to the firm as soon as possible. They were also told Ossur does not currently have replacement product available, but they will contact them immediately when they are available.If the products were not located, they were asked to determine where these pins are currently located, if they have been applied to a patient? If so, are they able to track down that patient and bring them in for an evaluation? |  | 12 | 3 pouches (4 pins/pouch) = 12 pins | Nationwide Distribution --- including states of FL, KS, MI, MO, NY, &amp; WA. | Terminated | April 27, 2010 | 578 | 43 | Not_Computer
Z-0285-2009 | 49135 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Solar 8000M System, a multiparameter physiological patient monitoring system intended for use on adult, pediatric and neonatal patients, within a hospital or facility environment |  |  | April 02, 2008 | 2008 | GE Healthcare | A software error has been discovered when using Combination Monitoring with Solar 8000M/i - PDM - Telemetry in combo mode. This error has only been noted with this particular configuration and mode of monitoring. It has not been observed in other monitoring applications. The software error will result in a reset of the Solar 8000M/i, causing a loss of monitoring for approximately 1-2 minutes a | Software design | A GE Healthcare consignee letter was sent out 9/9/08. The letter dated August 26, 2008 was addressed to Healthcare Administrator/Risk Manger. The letter described the problem, product and provided recommendation. |  | 2782 | 2782 | Nationwide to: AL, AR, AZ, CA, FL, GA, IA, ID, IL, IN, MA, MI, MN, MO, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, and WI. OUS to include: Australia, Austria, Belgium, Brazil, Czech Republic, Denmark, Egypt, France, Germany, Indonesia, Ireland, Italy, Jordan, Korea, Lebanon, Lithuania, Malaysia, Netherlands, Norway, Pakistan, Saudi Arabia, Spain, Sweden, United Kingdom, and Venezuela. | Terminated | May 17, 2012 | 1506 | 2782 | Software | Device Operation | Instructions | Safety Notice/Insructions
Z-0651-2009 | 50206 | support, head and trunk, wheelchair | IMS | Wheelchair accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | High-low:x frame for the x:panda mechanical chair. |  |  | November 06, 2008 | 2008 | Snug Seat Inc | Breakage in the welds was discovered in a section of the chair frame. | Process control | Consignees were notified by letter dated 11/6/2008. Snug Seat sent a repair kit to their retail dealers on 11/18/2008. The kit includes a safety plate, written and video replacement instructions as well as the tools (allen wrench, crescent wrench and two nuts) necessary to make the safety update. Customers were identified using purchase order numbers, and product serial numbers to enable the dealers to identify the purchaser of the product. Due to HIPPA regulations, purchaser information was not available to Snug Seat, Inc. to contact directly. Written notification will be followed up with a telephone call to each dealer to ensure they have received the retrofit kit. Snug Seat will also insure they understand the importance of the kit and how to properly install and document receipt of the update for their customers. |  | 67 | 67 units | Nationwide. | Terminated | February 17, 2011 | 833 | 67 | Not_Computer
Z-0439-2009 | 49580 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SteriLyteÂżÂż Liquid Bicarbinate, Formula CL-02; 1 gallon (3.78 L); manufactured by Rockwell Medical Technologies, Inc., Grapevine, TX 76051. This product is used in conjunction with C-Series RenalPureÂżÂż Liquid Acid Concentrate or DC-Series Dri-SateÂżÂż Dry Acid Concentrate in artificial kidney (hemodialysis) machines. |  |  | September 29, 2008 | 2008 | Rockwell Medical Technologies, Inc | Microbial growth was observed in the retain samples of 4 lots of Liquid Bicarbonate after products had been distributed. | Environmental control | Recall notification was faxed and/or e-mailed to consignees on 09/29/08 and then sent via mail on 09/30/08. A follow-up response form was faxed to consignees starting 10/01/08.Consignees asked to return product for a credit. For additional information, contact Rockwell Medical Technologies, Inc., Customer Service at 1-800-449-3353. |  | 2281 | 2,281 cases (4 units per case) | Nationwide. Product distributed to AL, AR, FL, GA, NC, SC, TN, TX, and VA. | Terminated | July 29, 2010 | 668 | 2976 | Not_Computer
Z-0339-2009 | 49607 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy HI-ART Systems with 3.1.2, 3.1.3 or 3.2.1 software. The affected applications include Planning Station, Planned Adaptive, Data Management System, and TomoPortal. The product is intended for use in radiation therapy. |  |  | September 18, 2008 | 2008 | TomoTherapy Incorporated | TomoTherapy Inc. discovered an issue on the Hi-Art system during the course of ongoing testing. It has been identified that treatment plans involving very small structures (volume &lt;2cc), reading in a high dose gradient region, may exhibit inaccuracies with the DVH curve and the Dose Statistics for the small structures. This error can only happen if planning was performed using very small voxels | Software design | Consignees were sent on 9/18/08 a "TomoTherapy Incorporated Urgent Medical Device Correction Field Safety Notice" letter dated September 17, 2008. The letter described the issue, product affected, recommended actions and resolution. Questions may be directed to TomoTherapy Inc. Customer Service Center at 1-868-368-4807. |  | 165 | 165 | AKALARAZCACOCTDCFLHIIAILINKYLAMAMDMIMNMONENJNMNVNYOHOKORPAPRSCSDTNTXVAWAWI andWYBelgiumCanadaSwitzerlandGermanySpainFranceGreat BritainIndiaItalyJapanKoreaNetherlandsTurkey andTaiwan | Terminated | April 19, 2011 | 943 | 165 | Not_Computer
Z-0342-2009 | 50083 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE OEC 9900 Elite GSP fluoroscopic x-ray system GE Healthcare, Surgery, Salt Lake City, UT. The system is intended to provide an x-ray display of human anatomy. |  |  | August 29, 2008 | 2008 | OEC Medical Systems, Inc | Under certain conditions, x-ray system may experience operational impairment. | Device Design | Consignees were notified by letter on 08/29/2008. Interim solutions were provided for proper start-up procedures and for avoidance of mechanical impact to the C-arm. A start Up caution label was included with the notification. Further information available at 800-874-7378. |  | 448 | 448 units | Nationwide, including AK, KY, NC, VA WA. and WI | Terminated | April 24, 2012 | 1334 | 1054 | Not_Computer
Z-0321-2009 | 50077 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Somatom Sensation 64 X-Ray System, Model number 8377520. This is a computed tomography x-ray system which is intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from the same axial plane taken at different angles. |  |  | September 03, 2008 | 2008 | Siemens Medical Solutions USA, Inc | software error - does not allow protocol to properly synchronize with ECG triggers | Software design | The recalling firm issued a Customer Safety Advisory Notice dated 9/3/08 to affected customers via Update Instructions CT045/08/S. This letter informs customers of the potential issued and provides instruction to avoid its occurrence. For additional information, contact Siemens Medical Solutions USA, Inc. at 1-610-448-4500. |  | 9 | 9 units | Nationwide | Terminated | September 23, 2009 | 385 | 177 | Computer
Z-0325-2009 | 50118 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | FS-OMNI, Fusion Omni-Tome, Sterile, Rx only, Wilson Cook Medical, GI Endoscopy, 4900 Bethania Station Road, Winston-Salem, NC 27105, EC Representative: Cook Ireland Ltd, O'Haioran Road, National Technology Park, Limerick, Ireland. The Fusion Omni-Tome is supplied sterile and is a single use device. This device is used in cannulation of the ductal system and for sphincterotomy. |  |  | October 30, 2008 | 2008 | Cook Endoscopy | The wire control port may have been omitted during manufacture. | Process control | Cook Ireland (distributor) was contacted via email on 10/30/08 and was instructed to contact their customers and coordinate the return of all recall product to Cook Endoscopy. A response sheet was attached. Recall status reports will be issued monthly. All devices returned as part of this recall initiative will be dispositioned to scrap and destroyed. Should you have any questions, please contact Cook Endoscopy Customer Quality Assurance Department at 1-336-744-0157. |  | 11 | 11 | Ireland | Terminated | April 07, 2009 | 159 | 11 | Not_Computer
Z-0645-2009 | 50502 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | CardioVations Ethicon, a Johnson and Johnson company, EndoClamp aortic catheter 100 cm, REF EC1001, Sterile R, individually packaged, Heartport, Inc., Somerville, NJ 08876. This product is indicated for use in patients undergoing cardiopulmonary bypass. It occludes the ascending aorta, vents the aortic root and allows delivery of cardioplegia solution to arrest the heart. The pressure lumen allows monitoring of the aortic root pressure |  |  | November 26, 2008 | 2008 | Edwards Lifesciences Research Medical, Inc. | Balloon rupture: Catheter balloon may rupture during cardiopulmonary bypass procedure. For further information, please telephone Edwards Lifesciences Research Medical, Inc. at 801-565-6173. | Nonconforming Material/Component | All consignees were notified by letter on 11/26/2008 and requested to return any unused product to Edwards. The letter was entitled "PRODUCT RECALL." Questions could be referred to the Edwards representative or call 800-424-3278. |  | 963 | 963 units | Globally - USA including the states of AK, AL, CA, CO, CT, DC, DE, FL, GA, IA, IL, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NJ, NV, NY, OH, OK, PA, SC, TN, TX, VA, WA, WI, and WV. VA facilities in VA and WI. No military distribution. Foreign distribution to Australia, Canada, Switzerland and UK. | Terminated | August 10, 2011 | 987 | 2230 | Not_Computer
Z-0448-2009 | 48734 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Homocysteine Reagent Test Kit. The Homocysteine test kits were manufactured under 2 brand names. The first brand name Catch, the label reads "CATCH Catch Inc., Bothell, WA 98011 HOMOCYSTEINE REAGENT KIT, Catalog Number A11167, 100 Test Kit, Manufactured at Intersect Systems, Inc., Longview, WA 98632". The kit is made up of 5 individual portions. These include: Homocysteine Reagent R1 HCY A (28mL), Homocysteine Reagent R2 HCY R2 (13.5 mL), Homocysteine HCY C (2.8 mL), Homocysteine 0.0 umol/L (2mL) and Homocysteine 27.0 umol/L (2mL). The second brand name Equal the label reads "EQual DIAGNOSTICS FOR IN VITRO DIAGNOSTIC USE ONLY HOMOCYSTEINE 3-Part Kit w/calibrators, Catalog Number E02-058-03, Equal Diagnostic, Exton, PA 19341". The kit is made up of 5 individual portions. These include: Homocysteine Reagent R1, Homocysteine Reagent R2, Homocysteine Reagent R3, Homocysteine Calibrator 0.0 umol/L and Homocysteine Calibrator 27.0 umol/L. All individual kits are packaged into individual boxes and then packaged into 1 box. Homocysteine Reagent is intended for in vitro quantitative determination of total homocysteine in serum and plasma. |  |  | January 18, 2007 | 2007 | Catch Incorporated | A high homocysteine level can be reported as falsely low. | Device Design | On 01/18/07 an "URGENT: PRODUCT CORRECTIVE ACTION" letter was sent to the consignees. The letter states the circumstances in which a high homocysteine rate can result in a false low homocysteine rate, a work around for the problem which a flag must be programmed into the user defined homocysteine test parameters and instructions to advise their customers about the situation was addressed in the letter so the consignees can take the recommended corrective action. A confirmation sheet was included with the letter to be signed by a representative of each distributor and returned to Catch Inc. Contact Catch Inc. at 1.425.402.8960 for assistance. |  | 2699 | 2,699 kits | CA and PA. | Terminated | September 28, 2010 | 1349 | 2699 | Not_Computer
Z-0542-2009 | 49382 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche ISE Internal Standard Gen.2 for use on Roche cobas c systems, Roche Diagnostics, Inc., Indpls., IN; Catalog no. 04522320190. Product is an in vitro diagnostic. |  |  | September 15, 2008 | 2008 | Roche Diagnostics Corp. | Test results, low: Sodium, potassium and chloride results on the cobas 6000 ISE unit may exhibit a downward drift in controls and patient samples within the recommended calibration interval of 24 hours. For further information, please contact Roche Diagnostics Corporation by telephone at 800-428-2336. | Package design/selection | Consignees were notified by letter dated 9/15/08, which provided a work-around for them to use until the problem can be permanently resolved. The letter was entitled "URGENT MEDICAL DEVICE CORRECTION." For further information, please contact Roche Diagnostics Corporation by telephone at 800-428-2336. |  | 7909 | 7909 | Nationwide Distribution. | Terminated | December 01, 2009 | 442 | 7909 | Not_Computer
Z-0453-2009 | 50156 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Angiotech, Bone Marrow Harvest Needle, Medical Device Technologies, Inc. The device is intended for use Bone Marrow Biopsy. |  |  | July 16, 2008 | 2008 | Medical Device Technologies, Inc. | Product packaging for Bone Biopsy trays/blisters may have a hole in the formed tray. | Other | Each consignee was notified of recall by letter on 07/16/2008 . The letter described the problem and indicated the corrective action to be taken. A Customer Acknowledgment Form with request to fax form to firm's QA Dept was provided. Contact Medical Device Technologies, Inc., Quality Assurance Department at 1.352.338.0440, ext 350 and 353 for assistance. |  | 130 | 130 units | Nationwide - FL, TX, WI, PA, MN, SC, NC, KY, TN, CA, AR, OR, NJ, and IA. | Terminated | April 21, 2009 | 279 | 430 | Not_Computer
Z-0455-2009 | 49010 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with CXP Software, Part Numbers: 626553, Version 2.2 |  |  | August 01, 2007 | 2007 | Beckman Coulter Inc | Software error: The recall was initiated after Beckman Coulter identified that the CXP User documentation does not adequately characterize the use of the Live Gate feature.The primary use of a Live Gate is for the international removal of unwanted events, such as debris. The data outside of the Live Gate will not be stored to the final listmode file, this reducing the data set size.The softwa | Software design | The customer notifications were initiated by Beckman Coulter on July 23, 2007 when a Product Corrective Action (PCA) was mailed via US Postal Service to all affected customers. The PCA informed the customers that the user documentation does not adequately characterize the use of the Live Gate feature. The letter provides the customer with explanation of Live Gate and actions to be taken. For further information, please contact Beckman Coulter, Inc. by telephone at 714-993-5321. |  | 565 | 565 | Globally to the United States and Canada. | Terminated | January 13, 2009 | 531 | 565 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0607-2009 | 50162 | No | LJB | Rubeola (measles) virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Measles IgG Enzyme Immunoassay Test Kit, DIAMEDIX. The assay is intended for the semi-quantitation of human IgG antibodies to measles virus in human serum by indirect immunoassay to aid in the assessment of the patient's immunological response to measles virus, to determine the immune status of individuals and, when evaluating paired sera, as an aid in the diagnosis of measles infection. |  |  | October 10, 2008 | 2008 | Diamedix Corporation | Possible false-negative test results: The Sample A Diluent Lot # 50768 assigned to Is-Measles IgG Lot #'s 41108 and 41108B has higher conductivity than normal, which can sporadically interfere with the MAGO Plus/Aptus sensing of and picking up of samples. As a result, if the sample is not picked up/processed, it could lead to the possibility of obtaining false negative results in some cases. | Other | Notification of all domestic customers via Certified Mail. Notification of the one international customer via Federal Express. |  | 974 | 974Test Kits | Worldwide Distribution --- including USA and country of England. | Terminated | May 12, 2010 | 579 | 974 | Not_Computer
Z-0601-2009 | 50144 | No | FNM | Alternating pressure air flotation mattress. | General Hospital | General Hospital | 510(K) Exempt | Synergy Dynamic, Powered Patient Rotation Bed, Model number 5607. |  |  | October 09, 2008 | 2008 | Hill-Rom Manufacturing, Inc. | These devices have the potential to overheat. | Device Design | All end users (those who either purchased the device and any customer that has rented) were notified by an Urgent Field Safety Notice on/about 10/31/2008. The letter informed the customers of the safety risk and list steps they can take to mitigate the risk until a Hill-Rom representative is able to modify the device. For the devices that were sold, the consignee of each device was notified by letter for the safety risk and what they can do to mitigate the risk until a Hill-Rom representative is able to modify the device. For the devices that are rented, an internal notification was sent via e-mail to Service Centers (Because the device is owned by the manufacturer, Hill-Rom, each Service Center in possession of the device will be the Consignee). Each Service Center will perform the correction on the units affected. |  | 2083 | 2,083 units (Total all four units) | Nationwide Distribution. | Terminated | December 10, 2009 | 427 | 2083 | Not_Computer
Z-0519-2009 | 49014 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Compress/Finn Instrument Set Compress Anchor Plug Assembly, stainless steel, Biomet Orthopedics, Inc., Warsaw, IN; REF RD425032. |  |  | July 18, 2008 | 2008 | Biomet, Inc. | The set screw may not fully seat, which will prevent the handle from mating to the instrument during use. | Process control | Distributors were notified via an Urgent Medical Device Removal Notice, dated 7/18/08, which instructed them to locate and remove the instruments, carefully follow the instructions on the enclosed "FAX Back Response Form," fax a copy of the Response Form to 574-372-1683 prior to return of product, and to notify the hospital personnel via a copy of a recall letter. The letter also stated that the instruments would be reworked and returned to the distributors. Distributors are to confirm receipt of the notice by calling, 800-348-9500, ext 3983. Questions related to the notice should be directed to Mary Hardesty at 800-348-9500 or 574-372-3983. |  | 41 | 41 of all products. | Worldwide Distribution --- including USA and countries of Australia, Belgium and Finland. | Terminated | July 07, 2009 | 354 | 41 | Not_Computer
Z-0608-2009 | 50167 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi CXR4 Computed Tomography Scanner. This product is an x-ray imaging device that produces cross-sectional images of he body at different angles. The system reconstructs, processes, displays, and stores the collected images. The device output can provide an aid to diagnosis when used by a qualified physician and is intended for general purpose CT applications. |  |  | November 05, 2008 | 2008 | Hitachi Medical Systems America Inc | Software error: The device has a software error that can result in a condition where the displayed image orientation does not match the display orientation markers (right-left reversal). | Software design | On 11/5/2008 Hitachi Medical Systems sent recall notification letters entitled "URGENT - DEVICE RECALL CORRECTION" to its customers informing them of the problem with the software. They stated that a Hitachi service representative will contact each customer to schedule a software upgrade. For further information, please contact Hitachi Medical Systems America Inc. by telephone at 330-425-1313. |  | 42 | 42 | Nationwide Distribution --- including states of AZ, CA, DE, FL, ID, IL, IN, KS, KY, LA, MA, MO, NJ, NY, OH, OK, PA, PR, SC, TX, VA, and WI. | Terminated | September 21, 2012 | 1416 | 42 | Software | Display/Image | Software update | Software Update
Z-0579-2009 | 49573 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Total Knee system Primary Tibial Baseplate #5 Beaded wPA; Catalog Number 5526-B-500, Cementless use only. Sterile. |  |  | September 21, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that one lot of the Triathlon Primary Beaded PA Baseplate Size 5 may not have the Peri-Apatite coating. | Process change control | Important Market Withdrawal notification letters were sent via Fed Ex on 9/20/06 to the direct accounts. Notification letters informed of the reason for the market withdrawal and requested that all branches examine their inventory and hospital consignment locations to identify the product. Product was to be reconciled on the Product Accountability Form and returned to Stryker as per instructions. Contact Stryker Howmedica Osteonics Corp at 1-201-831-5718 for assistance. |  | 10 | 10 | PA, GA, AL, WI, UT, and OR. | Terminated | January 09, 2009 | 841 | 10 | Not_Computer
Z-0647-2009 | 50500 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | CT syngo Dual Energy Viewer on the syngo Multimodality Work Place (MMWP) and on the CT workplace, Model number 10094833. |  |  | November 13, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Indicated orientation does not match actual orientation of the patient. | Software design | The recalling firm issued a Customer Safety Advisory dated 10/8/08 to affected customers via Update Instructions CT052/08/S. The letter informs customers of the potential issue and provides instructions to avoid its occurrence. Customers should consider "no saving of calculated Dual Energy data sets from non-supine and mirror-option scans via "Save Sets CT"-button." The firm also highly recommends for customers "to check the correct display of "left", "anterior" and "posterior"markings, e.g. by checking the position of the spine with respect to the posterior marking in the image." The error should be fixed within certain upgrade and installation activities on the affected software. |  | 374 | 374 units | Nationwide Distribution | Terminated | March 19, 2013 | 1587 | 374 | Software | Output/Calculation | Software update | Software Update
Z-0643-2009 | 50501 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST CVi Angiographic Injection System with Interface Cable for Toshiba Infinix-i cardiovascular X-ray system, Part Number 700318-002. The ACIST injection system is an angiographic injection system used in interventional cardiology, radiology, and vascular surgical procedures. The ACIST injection system supplies radiopaque contrast media to a catheter at a user-determined variable flow rate and volume which can be instantaneously and continuously varied. Manufactured by: ACIST Medical Systems, Inc., 7905 Fuller Road, Eden Prairie, MN 55344, USA. The ACIST CVi injection system is able to synchronize with certain X-ray imaging systems from Siemens, Toshiba, GE and Philips. |  |  | November 10, 2008 | 2008 | Acist Medical Systems | Multiple Injections may result. It has been determined that when using the ACIST CVi Angiographic Injection system to synchronize with the Toshiba Infinix -i cardiovascular x-ray system there is an injection scenario that may result in multiple injections of contract media when one injection was intended. | Software design | ACIST is distributing an Urgent Medical Device Correction letter, dated November 10, 2008, to affected customers. The letter describes the issue and the scenario when the issue to occur, instructs them to discontinue use of the synchronized mode between the ACIST Angiographic Injection System model CVi and the Toshiba Infinix-i X-ray fluoroscopic system by disconnecting the interface cable immediately. The letter states that ACIST is working with Toshiba to resolve this issue. A response form is asked to be returned either by fax, e-mail or mail. If there are any questions, they should be directed to either ACIST at 952-995-9300 or local Sales Reps or Clinical Applications Specialist. |  | 18 | 18 cables | Worldwide Distribution --- USA including states of AR, AZ, FL, and NY and countries of France, Germany, Italy, Japan, and Spain. | Terminated | May 07, 2011 | 908 | 18 | Not_Computer
Z-0505-2009 | 50168 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Sinu-Lift(TM) System, LLC., Model SLS-1, Sterile, Innovative Implant Technology, |  |  | September 30, 2008 | 2008 | Innovative Implant Technology | Possible device failure: The cutter's shaft of the Sinu-Drill of the Sinu-Lift System could potentially get stuck inside the tap, thereby not allowing the cutter to spring out of the tap upon encountering the sinus membrane, increasing the possibility of perforating the sinus membrane. | Nonconforming Material/Component | All consignees will be notified by letter entitled "URGENT: MEDICAL DEVICE RECALL" dated September 30, 2008, using priority mail. Consignees will also be contacted by telephone. Consignees will be asked to return the product using the prepaid FedEx label included with the letter. The returned product will be quarantined. For further information, please telephone Innovative Implant Technology at 305-975-4101 |  | 299 | 299 | Worldwide Distribution --- including USA and countries of Canada, Mexico, Japan, Australia, Japan, Egypt, and Turkey. | Terminated | October 01, 2009 | 366 | 299 | Not_Computer
Z-0871-2009 | 50546 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural-Knee II System Modular Cemented Tibial Baseplate, Right, Size 3, For cemented use only, sterile, Zimmer, Warsaw, IN; REF 6420-01-230. |  |  | November 25, 2008 | 2008 | Zimmer Inc. | Sterility compromised: Lack of assurance of sterility, as a corner flap of the sterile barrier packaging on the inner cavity may be trapped in the seal of the barrier packaging for the outer cavity and the seal integrity of the outer sterile barrier cannot be assured. | Packaging process control | Zimmer distributors were notified by email of letter dated 11/25/08 on 11/25/08. This letter is entitled "URGENT: DEVICE RECALL." Secondary consignees/user accounts were notified by a copy of the 11/25/08 letter, which was sent on or before 12/05/08. Consignees were instructed to stop using the product and to return it to Zimmer. For further information, please contact Zimmer, Inc. by telephone by dialing 800-613-6131. |  | 10 | 10 | Nationwide Distribution --- including states of Arkansas, Florida, Michigan, North Carolina, New York and Wisconsin. | Terminated | November 12, 2009 | 352 | 10 | Not_Computer
Z-0487-2009 | 50087 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Easy Core Biopsy Device, 18 ga. x 21 cm, sterile, latex free. Boston Scientific, Spencer, IN; UPN 500-104, REF M0065001040 and M0065001041. |  |  | October 17, 2008 | 2008 | Boston Scientific Corp | These devices may experience difficulty cocking and arming, resulting in an inability to use the devices. | Nonconforming Material/Component | U.S. consignees were notified via letter dated 10/17/08, which informs them of the problem, instructs them to remove from use and quarantine the recalled lots and requests completion and submission of the Reply Verification Tracking form via fax at 508-683-5578. For questions about this recall, contact your local Boston Scientific representative. |  | 14254 | 14,254 of all devices. | Worldwide Distribution including USA and countries of Algeria, Austria, Australia, Belgium, Canada, Denmark, Egypt, Finland, France, Germany, Great Britain, Greece, Hong Kong, Ireland, Italy, Japan, Mexico, Netherlands, Norway, Philippines, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, Japan and United Arab Emirates. | Terminated | December 15, 2009 | 424 | 15114 | Not_Computer
Z-0870-2009 | 50540 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Centricity Enterprise Web 3.0 software; the software allows physicians convenient, quick access to medical images and related data they need via a web browser, password and a standard PC; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | December 12, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Software error: There are two potential safety situations with Centricity Enterprise Web software - a forced log off may occur while using the Centricity Enterprise Web during an open session, and Centricity Enterprise Web does not utilize calibration performed on image measurements in Centricity RA1000 Workstation. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity Enterprise Web software customers via letter (Urgent Medical Device Correction) dated 12/12/08 of the software anomalies resulting in patient safety issues involving a forced log off during an open session while using the Centricity Enterprise Web and measurements displayed in Centricity Enterprise Web will use the calibration provided by the original scanning modality. Centricity Enterprise Web does not utilize calibration performed on image measurements in Centricity RA1000 Workstation. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653. For further information, please contact GE Customer Care Center by telephone at 800-437-1171. |  | 742 | 742 units | Worldwide Distribution including USA and countries of Australia, Austria, Bahrain, Belgium, Canada, China, Denmark, Egypt, France, Germany, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Malaysia, Malta, Martinique, Netherlands, New Zealand, Nigeria, Norway, Portugal, Reunion, Russia, Saudi Arabia, Slovenia, Singapore, South Africa, Spain, Sudan, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | May 18, 2010 | 522 | 742 | Software | Device Operation | Software update | Software Update
Z-0580-2009 | 49597 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABBOTT Clinical Chemistry Bilirubin Calibrator, List Number 1E66-04. |  |  | September 12, 2008 | 2008 | Abbott Laboratories Inc. | The field action was initiated after an investigation for the Total Bilirubin reagents (List Number 8G62 and List Number 6L45) determined that the matrix of the secondary standard used in the value assignment of the calibrator is sensitive to the Diazo method. The matrix of this secondary standard caused a positive bias. Because of these findings, a new Total Bilirubin calibrator value assignm | Component design/selection | On September 12, 2008, a Product Correction letter with an attached Customer reply form was mailed via US Postal Service to all affected customers that have received the affected lots. The letter instructed customers to take the following actions:1. Identify the lot number(s) of Clinical Chemistry Bilirubin Calibrator currently used in your laboratory;2. Replace the calibrator value sheet with the appropriate attachment provided with this letter. Discard the previous calibrator value sheet;3. Configure the revised Total Bilirubin calibrator values provided in the Attachments, for the appropriate calibrator lot number and reagent list number.ARCHITECT cSystems: Refer to System Configuration in Section 2, Installation Procedures and Special Requirements in the Architect System Operations Manual.AEROSET: Refer to Calibrator/Control Configuration in Section 2, Installation Procedures and Special Requirements of the AEROSET System Operations Manual;4. Calibrate the Total Bilirubin assay. Verify the calibration with at least two levels of controls according to the established quality control requirement for your laboratory;5. Evaluate your quality control. Shifts in quality control and proficiency testing samples may vary among laboratories and should be evaluated according to your laboratory procedures. NOTE: Refer to the Attachments (E and F) for the representative shift with the revised calibrator values (Patient sample and Control/QC recovery); and6. Evaluate your reference range and any shifts in patient results. It is recommended that each laboratory established its own expected ranges, which may be unique to the population it serves.Please retain this communication for future reference. If you have forwarded any Clinical Chemistry Bilirubin Calibrator to another laboratory, please provide a copy of this letter to them.For questions, please contact ABBOTT Customer Service at 877-422-2688 (or 877-4ABBOTT). Customers outside of the |  | 7101 | 7,101 units worldwide (2,267 units in US). | Worldwide Distribution including United States, Mexico, Columbia, Uruguay, Puerto Rico, Canada, Hong Kong, New Zealand, Bahamas, Guatemala, Argentina, Germany, Thailand, Trinidad/Tobago, Venezuela, Chile, Brazil, El Salvador, Singapore, Australia, Curacao and Cayman Island. | Terminated | April 24, 2012 | 1320 | 7101 | Not_Computer
Z-0869-2009 | 50535 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon PKR Insert X3 #1 LM/RL - 8mm, 5630-G-108, Lot code NXDMEE |  |  | November 25, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Label mix-up: Stryker Orthopaedics became aware of a lot for lot mix up between Triathlon PKR Insert X3 #1 LM/RL 8mm, 5630-G-108 lot code NXDMEE and Triathlon PKR Insert X3 #1 RM/LL- 12 mm 5630-G-122 lot code M8JMEE. | Labeling mix-ups | Recall notifications entitled "URGENT PRODUCT RECALL" were sent via Federal Express on 11/25/2008 to Stryker branches, hospital risk management, Hospital Chief of Orthopaedics. For further information, please contact Stryker Howmedica Osteonics Corporation by telephone at 201-831-5718. |  | 29 | 8 mm - 1 unit; 12 mm - 8 units. | Nationwide distribution: Product was distributed to 7 Stryker branches and one hospital. | Terminated | August 30, 2010 | 643 | 29 | Not_Computer
Z-0129-2009 | 48188 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cytomics FC 500 Flow Cytometry System with CXP Software Version 2.2, Part Number: 626553. This product is used for the qualitative and quantitative measurement of biological and physical properties of cells and other particles. These properties are measured when the cells pass through one or two laser beams in single-file. |  |  | October 11, 2007 | 2007 | Beckman Coulter Inc | Software error: This field action was initiated after Beckman Coulter determined that the CXP software does not refresh the dataset automatically under certain circumstances. The User documentation does not clearly state that the user should refresh the dataset after changes are made. | Software design | Software error: A Product corrective Action (PCA) letter was mailed via US Postal Service on October 17, 2007 to all Cytomics FC 500 Flow Cytometry System with CXP Software accounts. This letter was entitled "URGENT: PRODUCT CORRECTIVE ACTION." Customers were informed that the user documentation does not clearly state that the user should refresh the dataset after changes are made. The letters provide the customers with an explanation and action to be taken whenever a protocol containing data is modified by making changes to cytosettings, regions or gates.Customer were instructed to review all data plots and results before reporting results out. any uncharacteristic patterns in the data or unexpected results must be verified by running Quality Control and if necessary, preparing new samples.In addition, customer were instructed to please share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. Also they were asked to complete and return the an attached Customer Response Form within ten days.Customers with questions or concerns regarding the notification, could call Beckman Coulter Customer Service at 1-800-526-7694 in the United States or Canada. Outside the US and Canada, they were to contact their local Beckman Coulter Representative. |  | 1188 | 594 Worldwide: 518 in the United States and 76 in Canada | Globally - United States and Canada | Terminated | January 13, 2009 | 460 | 1188 | Software | Device Operation | Instructions | Safety Notice/Insructions
Z-0578-2009 | 49526 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Custom Waste Management Kit, K10-01756/B, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095 |  |  | September 12, 2008 | 2008 | Merit Medical Systems, Inc. | Leak, disposal bag: A Merit Disposal Depot (MDD) bag contained in certain custom waste management kits may leak. For further information, please contact Merit Medical Systems, Inc. by telephone at 801-208-4748. | Process control | All consignees were notified by sales reps via phone, fax or visit between 09/15/2008 and 10/09/2008. Sales reps were to instruct them to immediately discontinue use and quarantine all affected products. A Product Retrieval Form was to be completed and signed by the customer's site representative. More information available at 801-316-4998. |  | 399 | 399 kits | United States including the states of AR, AZ, CA, CO, FL, KY, IL, IN, MA, MI, MN, MO, NJ, NY, PA, SC, TX, UT, VA, WI, and WV. No military or foreign distribution. VA facilities in CA and FL. | Terminated | January 13, 2009 | 123 | 399 | Not_Computer
Z-0581-2009 | 49608 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension EXL Clinical Chemistry System used with software versions 8.5.1 and 8.5.1SP3. |  |  | September 30, 2008 | 2008 | Dade Behring, Inc. | Error messages, "Wash Aspirator Failure" and "Wash station unable to recover" are not generated when one or both HM wash probes are totally occluded. | Software design | The recalling firm notified consignees by telephone to inform them of the problem. A follow-up letter, "Urgent Field Safety Notice" dated 9/30/08 was issued with information describing the problem, the possible impact of an occluded wash probe on assay results, and emphasizing regular performance of routine wash probe maintenance as detailed in the operator guide to prevent wash probe occlusion. For questions, please contact the Siemens Technical Solution Center at 800-441-9250. |  | 57 | 57 units | Nationwide Distribution including states of AR, CA, CO, IA, ID, IL, IN, LA, MD, MI, MO, MS, NC, NE, OH, OK, PA, SD, TN, TX, VA, and WI. | Terminated | May 05, 2009 | 217 | 57 | Not_Computer
Z-0817-2009 | 50088 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips M8000-64100 Table Mount |  |  | October 22, 2008 | 2008 | Philips Medical Systems | Table Mounts assembled with incorrect length hardware may fall | Process control | Philips issued an "Urgent Medical Device Correction" notification to all potentially affected customers on 10/22/08 via UPS for the US customers. The letter mailed to customers outside the US by Philips' regulatory contacts in each country via a tracked method.The letter informs customers of the problem and gives instructions on how to identify affected table mounts. Field Service Engineers will be dispatched to each customer site to inspect each Table Mount. If customers detect any sign of a loose connection between the monitor and the mounting plate while they are awaiting the Field Service Engineer, they are instructed to make sure that the monitor is not in a position to fall. They are asked to disconnect the monitor from the table mount. For additional information, contact Philips Medical Systems at 1-978-687-1501. |  | 14292 | 14,292 units | NationwideCanadaForeign: Argentina, Aruba, Australia, Austria, Bahain, Bangladesh, Belgium, Bosnia, Brazil, Chile, China, Colombia, Czech Republic, Denmark, Estonia, Egypt, Finland, France, Gabon, Germany, Great Britain, Greece, Hong Kong, Iceland, India, Iraq, Ireland , Italy, Japan, Korea, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Namibia, Netherlands, Norway, Peru, Philippines, Poland,Portugal, Russia, Saudia Arabia, Singapore, SloveniaSouth Africa, Spain.,. Srilinka, Sweden, Switzerland, Thailand, Tunisia, Turkey, Uganda, UAE, Uzbekistan, Venezuela, Vietnam, and Zambia | Terminated | March 24, 2017 | 3075 | 14292 | Not_Computer
Z-0550-2009 | 49510 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Burrs |  |  | August 08, 2008 | 2008 | Sterilmed Inc | SteriMed, Inc is conducting a product removal of a small group of devices individually packaged and sealed within a pouch. The removal is being conducted because the possibility exists that some of the pouches were not properly sealed on one end.A breach in packaging seal or a failure in packaging integrity has the possibility of risk to the patient in terms of transmitting organisms capable | Packaging process control | Consignees were notified of the recall via verbal communication in person by on-site technician and or their local sales representative on 8/8/08. Others were contacted via fax or e-mail followed by on-site visits 8/11/08. SteriMed also provided consignees (Materials Manager/Risk Manager) with a "Urgent Product Removal" recall letter (August 14, 2008) in person informing them of the potential of receiving affected product. Included in the Product Removal Notification was a Customer Reply Form to inventory quantity of product received and quantity returned. Contact SterilMed, Inc. at 1-888-541-0078 for assistance. |  | 125 | 125 | AL, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, MA, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, WA, WI, and WV OUS: CANADA | Terminated | April 26, 2011 | 991 | 5070 | Not_Computer
Z-0619-2009 | 50235 | No | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 1 ml Integra TB Syringe with Retracting Precision Glide Needle; 1 ml 27 g 1/2 (0.4mm x 13 mm); Sterile; REF # 305298 |  |  | November 24, 2008 | 2008 | Becton Dickinson &amp; Company | Expiry date on the shelf carton is incorrect. It reads 2013-13, it should read 2013-03. | Incorrect or no expiration date | Recall letters were sent on November 24, 2008 to all direct customers and potential end users of the product by UPS second day air. BD requests that inventory be examined and affected product be quarantined and returned as per recall letter instructions. Questions should be directed to BD customer service at 888-237-2762, M-F 8.OO AM to 5:00 PM EST. |  | 334400 | 334,400 units | GA, KS, OH, MO, FL, PA, KY and NY. | Terminated | March 27, 2009 | 123 | 334400 | Not_Computer
Z-0646-2009 | 50480 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Drive Belt for Radiation Therapy Simulation systems SIMVIEW 3000 and SIMVIEW NT with CT option or IDI. |  |  | November 06, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Excessive wear on drive belts used for Image Intensifier and/or X-ray head drive mechanisms may result in uncontrolled movement, which may cause injury. | Device Design | Firm has reissued the Urgent Medical Device Customer Letter regarding the belt replacement and a reminder message to replace the belts yearly during the annual Preventative Maintenance. The letter requests that customers follow the guidelines, which consist of making any required SAD adjustments utilizing the Image Intensifier drive, the X-Ray Head Drive, or both; being vigilant of any change in the type or level of noise from either of the drive belt assemblies while they are in use as this may be an indication of wear; discontinuing clinical operation if there is any doubt, and contacting their local Siemens Service organization to inspect the belts. Currently, the firm is distributing the current belt kit to replace the entire installed base of active SIMVIEW Simulators (3000 and NT models). |  | 223 | 223 units | Worldwide Distribution | Terminated | December 09, 2010 | 763 | 223 | Not_Computer
Z-0879-2009 | 50698 | No | HTZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Biomet Microplasty Elite Instruments Vanguard Slidex Femoral Block, 70 mm, Biomet Orthopedics, Inc., Warsaw, IN; REF 32-485156. A manual surgical instrument for general use; a non-powered, hand-held, or hand-manipulated device; instrument guide. |  |  | December 18, 2008 | 2008 | Biomet, Inc. | Falling component: The universal pin, which holds the sliding instrument saw guide in its slot, may loosen during surgical procedures and fall into the wound. | Other | Consignees were notified by letter entitled "URGENT MEDICAL DEVICE REMOVAL NOTICE" dated 12/18/08 advising them of the problem and that the instruments should be removed from use and returned. For further information, please contact Biomet, Inc., by telephone on 800-348-9500. |  | 444 | 444 of all products | Worldwide Distribution --- including USA and countries of Australia, Canada, Japan and the United Kingdom. | Terminated | January 20, 2010 | 398 | 444 | Not_Computer
Z-0838-2009 | 50136 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | NeuroTherm Simplicity III RF Electrode |  |  | July 28, 2008 | 2008 | Neurotherm, Inc. | Distal tip may detach from the probe | Labeling design | NeuroTherrn sales representatives were notifiedtelephonically or through e-mail. End users were notified through visits by NeuroTherm sales representatives on 7/28/08 and 8/15/08. For questions, please contact NeuroTherm, Inc. at 1-978-777-3916. |  | 71 | 71 units | CO, FL, IL, OH, TN, TX , VA, and WYForeign: Australia, China, and UK, | Terminated | September 24, 2009 | 423 | 71 | Not_Computer
Z-0611-2009 | 50193 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ReSolve Hydrophilic Coated Locking Biliary Catheter, REF RBC-8-038, 8.5 F, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095. |  |  | November 07, 2008 | 2008 | Merit Medical Systems, Inc. | Biliary drainage catheter tips may separate from the catheter shaft after placement. | Other | Customers were notified by Merit representatives by phone or visit beginning 11/07/2008. They were to identify and quarantine any unused inventory and complete a Product Retrieval Form. For further information, call 800-356-3748 x 4929. |  | 284 | 284 units | Nationwide. OUS distribution to Canada, Colombia, France, Germany, Mexico, Ireland, Netherlands, South Korea, and UK. | Terminated | August 05, 2009 | 271 | 1164 | Not_Computer
Z-0856-2009 | 50208 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | RoboCouch Patient Support System, a component of the CyberKnife Robotic Radiosurgery System. Model number 025007, manufactured by Accuray Inc., Sunnyvale, CA |  |  | October 30, 2008 | 2008 | Accuray Inc | Product may not be tensioned properly, potentially causing unexpected rotation or descent. | Process control | All affected sites were notified of the potential problem via Customer Advisory Notice dated 30 October 2008. The letter states required actions that the customer should follow at the beginning of each treatment day. The tension of the belt drive and the support bolts in question will be inspected by Accuray Field Service personnel at the affected facilities. In addition, the preventive maintenance procedure will be updated to include regular inspection and replacement of the belt-drive. Any questions or concerns regarding this issue should be directed to Accuray Customer Support at 1-877-668-8667 or 1-408-716-4700 or customersupport@accuray.com. |  | 13 | 13 units | Worldwide Distribution --- including USA and countries of France and Turkey. | Terminated | September 10, 2009 | 315 | 13 | Not_Computer
Z-0582-2009 | 49609 | [6 empty] | QuantumPulse machine. Vibe Technologies, Greeley, CO. |  |  | October 01, 2008 | 2008 | VIBE Technologies | Medical device marketed without marketing approval for claims that include strengthening, recharging and removing toxins from cells. | PMA | Customers were notified by letter on September 16, 2008 and were told to discontinue use as a medical device. This letter notification included: 1. A warning label to be placed on the Vibe machine stating that it is not a medical device and should not be used as such;2. An updated operation manual/users guide containing no medical conditions or human body claims; and3. A certification to be signed by the user and returned acknowledging receipt of the letter; affixed the warning label; will not promote the Quantum Pulse as a medical device; will remove any medical claims from their website; destroy any Quantum Pulse literature making medical claims and destroy any previous operation manual/user guides.Please contact VIBE Technologies at 800-356-9594 for questions or need assistance in complying with these actions. |  | 58 | 58 | Worldwide Distribution. | Terminated | May 03, 2012 | 1310 | 58 | Not_Computer
Z-0858-2009 | 50483 | No | LPB |  |  | Cardiovascular | PMA | RF Ablation System Foot Switch. (Can be used with Boston Scientific Maestro 3000 and EPT-1000XP Cardiac Ablation Cardiac Ablation Systems). Manufactured by Altech Corporation, Flemmington NJ and distributed by Boston Scientific Corporation, San Jose, CA, catalog numbers/UPN numbers 21840/M004 218400. |  |  | November 25, 2008 | 2008 | Boston Scientific Corp | Product malfunction may result in delivery of energy without depressing the foot switch that controls delivery. | Process control | Recall initiated on November 25, 2008. Urgent Company Initiated Medical Device Recall - Immediate Action Required letters, dated 11/26/2008, were sent to consignees via Federal Express and included return and reply forms for consignees to respond. The letters asked that further distribution or use of any foot switch affected by this recall be ceased immediately. Attached to the letter was a document entitled "Instructions for Determining Affected Product," which provided instructions and serial numbers for determining if customers have product affected by this recall. In addition, a document entitled "Recall Instructions" was also attached. Affected product should be segregated and returned to Boston Scientific in accordance to the "Recall Instructions," and the Reply and Verification Tracking Form should be completed and returned. Customers will receive replacement product for all recalled product returned to Boston Scientific. Local Sales Reps should be contacted in regards to any questions customers may have pertaining to this recall. |  | 322 | 322 units | Worldwide Distribution --- including USA and countries of France, Great Britain, Saudi Arabia, Spain, Canada, Portugal, Germany, Italy, Australia, Russia, Netherlands, Ireland, Libya. | Terminated | April 26, 2011 | 882 | 322 | Not_Computer
Z-0857-2009 | 50477 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV Digital Linear Accelerator, when used in combination with Syngo RT Therapist Version 4.1 build 67 and automatic Motion of Flat Panel, part number 8139789, manufactured by Siemens Medical Solutions, Oncology Care System. |  |  | October 23, 2008 | 2008 | Siemens Medical Solutions USA, Inc | Under certain conditions, the flat panel may move unexpectedly and collide with the patient. | Labeling design | The firm issued an Urgent Medical Device Customer Safety Letter, dated 10/23/2008. With the letter included a detailed description of the issue, the condition under which this issue may occur, and recommendation to avoid them. Customers are to read the information carefully and distribute to the responsible persons. The Customer Safety Letter should be included their syngo RT Therapist Express manual and remain there until the next release of software is installed. In the interest of safety, the preventive measures should be performed and all affected personnel should be informed immediately. The firm is tracking responses via return cards and completion protocol. |  | 31 | 31 consignees have the configuration noted above. | Worldwide Distribution --- including USA and countries of Netherlands, Germany, Poland, Norway, Australia, Belgium, France, New Zealand, Malaysia, Spain and Canada. | Terminated | December 09, 2010 | 777 | 31 | Software | Physical Safety Hazards | Software update | Software Update
Z-0396-2009 | 47848 | No | GET | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ConMed Linvatec Hall Surgical Blades, Sterile ST SAG BLDE CRS, Catalog # 00502313300, 00502313700, 00502313000, 00502313100, 00502313500, 00502313200, and 00502314300. |  |  | February 21, 2008 | 2008 | Linvatec Corp. | Sterility compromised. There is a possibility that certain catalog number and lots of ConMed Linvatec Hall Surgical Blades, may have a hole(s) in the corner(s) of the blister package(tray), thereby potentially compromising the sterility of the devices. | Process control | An Urgent Medical Device Recall letter was issued. Customers must check their facility's inventory immediately for the catalog and lot numbers in the attached list. They also need to complete and sign the enclosed Reply Form and mail or fax the Reply Form to ConMed Linvatec using the enclosed self addressed postage paid envelope or fax 727-319-5701. Customers have an option to either re-sterilize the products using the re-sterilization instructions in the letter or they can return the devices to ConMed Linvatec for a credit using the enclosed preprinted FedEx Air Bill with ConMed Linvatec's Bill Number and the return goods number. Recall is being conducted to the distributor and end user level. |  | 13388 | 13388 | Worldwide Distribution | Terminated | January 15, 2009 | 329 | 199977 | Not_Computer
Z-0531-2009 | 49100 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Mainline Confirms hCG urine pregnancy test, Mainline Technologies, Inc., Ann Arbor, MI; Catalog # 6008C. |  |  | August 18, 2008 | 2008 | Mainline Technology, Inc. | Lack of assurance of safety and efficacy: Unapproved for marketing in the U.S., non-compliance with cGMP regulations and no stability data to support labeled expiration dates. | PMA | Consignees were notified via a Medical Device Removal Notice dated 8/18/08 to locate, discontinue use, and return product along with a copy of the letter. The recall was extended to an additional product via letter dated 9/5/08. |  | 33919 | 33919 kits | Worldwide Distribution --- including USA and countries of Belgium and Switzerland. | Terminated | February 07, 2011 | 903 | 45330 | Not_Computer
Z-0538-2009 | 49112 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Depuy DELTA Xtend, Standard Humeral PE Cup, Dia 42/+9 STD, sterile, Depuy France; REF 130742209. |  |  | July 10, 2008 | 2008 | Depuy Orthopaedics, Inc. | Compressed air may become trapped between the cup and the stem, preventing the tapered walls from engaging and causing the cup to feel unstable. | Device Design | A letter, "Urgent Device Correction" dated July 8, 2008 was sent via certified mail to implanting surgeons advising them of the issues and guidance on how to mitigate the problem. A communication was also sent to all sales representatives and territory offices on July 15, 2008. For questions, please contact Depuy Orthopaedics, Inc. Customer Service at 1-800-473-3789. |  | 1521 | 1,521 of all devices. | Nationwide Distribution. | Terminated | October 28, 2009 | 475 | 1521 | Not_Computer
Z-0620-2009 | 50472 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Foot switch, Part Number 1009.81. Maquet GmbH &amp; Co. KG, 76437 Rastatt, Germany. |  |  | December 08, 2008 | 2008 | MAQUET Inc. | A bearing point of the pedal of the foot switch may come loose. With continued use of the foot switch, one or both springs beneath the pedal can fall out of their mounting, and it is possible the pedal comes to rest on the switch causing an unintended motion of the operating table during patient use. | Device Design | An Urgent Device Correction letter (with an enclosed Customer Response form) was sent via Federal Express on Monday, December 8, 2008 to all customers who had received tables. Customers were notified in the letter of the problem and actions to be taken including not to use any affected foot pedal for safety reasons. In addition, customers were instructed to complete and return the Customer Response Form via fax (732-667-1903) to Maquet Inc. Please call USA Maquet, Inc. at toll free number 1-888-627-8383 #1 (then select the option for Surgical Workplace products) for questions. |  | 587 | 587 units total (54 units in US) | Worldwide Distribution. | Terminated | June 18, 2012 | 1288 | 587 | Not_Computer
Z-0855-2009 | 50143 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Eclipse 8.5 (build 8.2.2X) systems with distributed Calculation Framework (DCF) 8.2.22, model number H48, manufactured by Varian Medical Systems, Inc., Palo Alto, CA. |  |  | October 02, 2008 | 2008 | Varian Medical Systems Oncology Systems | When using the multiple static segment option, the device may alter dose delivery. | Software design | An Urgent Medical Device Correction Letter was distributed to all affected users, with a description of the problem and user corrective action steps on October 2, 2008. The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the issue. The firm recommends that customers DO NOT USE the Multiple Static Segment (MSS) option in Varian Leaf Motion Calculator version 8.2.22. The firm also recommends setting the default delivery method for Varian Leaf Motion Calculator to Sliding Window. When using Eclipse 8.5, always follow the Safety Instructions found in chapter 2 of "Eclipse External Beam Planning Instruction for Use", P/N B501357R01A, March 2008. If customers believe a patient treatment has been affected by this issue, they should notify a Varian Service or Support person as soon as possible. |  | 64 | 64 units | Worldwide Distribution --- including USA and countries of Switzerland and the Netherlands. | Terminated | November 27, 2010 | 786 | 64 | Not_Computer
Z-0597-2009 | 50120 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Vertical drive brake (on patient supports). Used on certain Philips Medical Systems MX series and Brilliance series CT scanners. |  |  | September 23, 2008 | 2008 | Philips Medical Systems (Cleveland) Inc | The product is being recalled because the patient support may travel downwards without being commanded to move. | Process control | A On 11/15/2007 &amp; 6/6/2008 the firm sent Urgent Device Notification letters to their customers. They notified their customers that a Field Service Engineer will inspect and repair the brake system. |  | 559 | 559 | Worldwide Distribution: USA including states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI and WV and countries of Argentina, Australia, Austria, Bahamas, Belgium, Brazil, Canada, China, Colombia, Costa Rica, Czech Republic, Denmark, Egypt, England, France, Germany, Greece, China, India, Indonesia, Ireland, Iraq, Italy, Japan, Jordan, Korea, Lebanon, Libya, Lithuania, Malaysia, Malta, Mexico, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Saudi Arabia, South Africa, South Korea, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | March 01, 2012 | 1255 | 559 | Not_Computer
Z-0506-2009 | 50181 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom 70 Semi-Electric Bed, Hill-Rom Company, Batesville, IN.; Model HS-968. The bed is a general purpose bed for use with low to medium acuity patients in the home care environment. |  |  | November 11, 2008 | 2008 | Hill-Rom, Inc. | Failure, mechanical: If the bed mechanism is cranked downward when the bed is already in its lowest position, the springs may cause the bed to rise suddenly if the bed is empty, more slowly if occupied by a patient, or when the patient exits the bed. For further information, please contact Hill-Rom, Inc. at 800-445-3720. | Other | Consignees were notified of the problem by letter dated 11/11/08 and instructed to position the bed at least 3 inches above the lowest bed position in order to prevent the problem, and then to only use the pendant electronic controls to raise and lower the head and foot sections of the bed, until Hill-Rom can repair the bed. The letter was entitled "Urgent Medical Device Correction." |  | 567 | 567 | Nationwide. | Terminated | December 31, 2009 | 415 | 567 | Not_Computer
Z-0625-2009 | 49568 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Vision 3 Flat Panel Shroud, Part #0682-000-414 (Individual Flat Panel Shroud, Ver. 2, Reworked), 0682-000-415 (SV3 Shroud Retro Fit Kit) and 0682-000-416 (Flat Panel Retro Fit Kit); distributed by Stryker Communications, Flower Mound, TX 75028. The 0682-000-414 (FP Shroud version 2 reworked) is a shroud designed to cover the cabling necessary on a Stryker Version 3 Monitor when mounted on the Flat Panel Arm. |  |  | September 02, 2008 | 2008 | Stryker Communications Corp | Paint chips: Shroud used to cover cables on the Stryker Vision 3 flat panel monitor (monitor used in surgery rooms during procedures) was not manufactured/painted to specification, resulting in the possibility that paint chips could fall off and fall into non-sterile and/or sterile fields during surgery. For further information, please contact Stryker Communications Corporation by telephone at 97 | Nonconforming Material/Component | Firm began notifying consignees of recall via letter on 09/02/08. Consignees advised that a representative would shortly visit the facility to replace the shroud with one that complies with specifications. The recall letter is entitled "URGENT: Device Recall." |  | 116 | 116 units. | Worldwide distribution; Product distributed nationwide to consignees in the states of CA, FL, IL, IN, MA, MN, NC, NM, OH and TX and to the following countries: Canada, Spain and the United Kingdom, | Terminated | January 08, 2010 | 493 | 116 | Not_Computer
Z-0616-2009 | 50198 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | Esmark Elastic Bandage, 4" x 3 yards, sterile, for single use only, latex free. Distributed by Cardinal Health, McGaw Park, IL 60085-6787. |  |  | November 18, 2008 | 2008 | Cardinal Health | The natural rubber latex bandages were mislabeled as latex free. | Labeling Change Control | A recall letter dated 11/17/08 was issued to the Directors of Materials Management and sent via DHL 2-day service on 11/18/08. Accounts were advised that the elastic bandages containing latex were mislabeled as latex free. Customers were requested to return their inventories of the affected product to Cardinal Health for credit or replacement by calling the Presource Customer Sales Operations Group at 800-766-0706 for return instructions. Customers were also requested to complete the enclosed Cardinal Customer Acknowledgement form by indicating the amount of affected product being returned and fax it to 847-689-9101. Please direct questions to Quality Systems, Professional Services at 800-292-9332. |  | 9521 | 9,521 units (of all devices) | Worldwide Distribution including USA, Guam and Korea. | Terminated | November 04, 2009 | 351 | 9521 | Not_Computer
Z-0868-2009 | 50525 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Dri-Lok Disposable Cannula, Part numbers 3910-075-500, 3910-075-501, 3910-075-502, 3910-075-650, 3910-075-651, 3910-075-652, 3910-075-800, 3910-075-801, 3910-075-802, 3910-090-500, 3910-090-502, 3910-090-650, 3910-090-652, 3910-090-800, and 3910-090-802. |  |  | December 01, 2008 | 2008 | Stryker Endoscopy | Leak: Leak-resistant barrier may not properly seal and leak when an instrument is inserted, and keep leaking after instrument is removed. | Component change control | Recall initiated 12/01/2008, Sales representatives were notified via telephone and e-mail. Product advisories entitled ""Urgent: Device Removal" and dated December 1, 2008, were sent to all consignees. All unused product is to be returned to the firm. For further information, please contact Stryker Endoscopy by telephone at 408-754-2124. |  | 4194 | 4,194 boxes (24,750 units) | Worldwide Distribution --- including USA and countries of Australia, Brazil, Canada, Chile, China, France, Germany, Greece, India, Italy, New Zealand, Poland, Portugal, Sweden, Spain, Singapore, Switzerland, and UK. | Terminated | January 10, 2011 | 770 | 4194 | Not_Computer
Z-0514-2009 | 48190 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Quantimetrix Dip and Spin Control Urine Dipstick/Microscopics Control. |  |  | June 01, 2007 | 2007 | Quantimetrix Corporation | This recall was initiated due to efficacy concerns with the stabilized erythrocytes, leukocytes and lymphocytes contained in this product group. | Employee error | The recall communication was initiated when customer both direct and distributors who have purchased the lots were mailed a recall notification letter dated June 8, 2008 with instructions which was sent via first class mail.Customers were requested to assist in the voluntary recall of the listed product and if any of the affected lots have been consumed please review their records to ensure that the obtained values conform to the stated specifications found in the lot specific product insert. |  | 24937 | 24,937 sets for all products. | Nationwide Distribution including states of AK, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MN, MO, NC, ND, NH, NM, NY, OR, PA, SC, SD, TN, TX, VA, WA and WY. | Terminated | February 18, 2009 | 628 | 24937 | Not_Computer
Z-0629-2009 | 50207 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Cyberknife Robotic Radiosurgery System. A radiation therapy device, MultiPlan (MP) Treatment Planning Software and Iris Variable Aperture Collimator, Software version 3.0. |  |  | November 14, 2008 | 2008 | Accuray Inc | System may use random incorrect data to calculate dose. Resulting dose calculation error can exceed 100% of correct dose which may lead to serious patient injury. | Software design | An Urgent Advisory Notification letter dated November 14, 2008 was issued to all sites via Fed Ex. The letter identifies the issue and required actions. The user was instructed not to modify current extrapolation depth settings and to contact Accuray Incorporated Customer Support prior to making any changes.Please direct any questions or concerns regarding this issue to Accuray Incorporated Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 (outside of USA) or customersupport@accuray.com. |  | 4 | 4 units | Nationwide Distribution. Product was distributed to 4 medical facilities in the US. | Terminated | December 02, 2010 | 748 | 4 | Software | Treatment/Delivery/Therapy | Instructions to not change settings | Safety Notice/Insructions
Z-0473-2009 | 49539 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 2100IQ Versatile Cardiac and Vascular Imaging System, GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) The product is indicated for use in generating fluoroscopic images of human anatomy for vascular angiography diagnostic and interventional procedures and optionally, rotational imaging procedures. It is also intended for generating fluoroscopic images of human anatomy for cardiology, diagnostic and interventional procedures. |  |  | January 30, 2008 | 2008 | GE Medical Systems, LLC | Software error, computer: GE Healthcare has became aware of an issue with Error Management affecting the Innova 2100, 3100/3100 IQ &amp; 4100/4100IQ. It was reported that when the foot switch was pressed, the live monitor screen was blank and it looked like there were no X-rays being emitted. But the message "Acquisition in progress" was displayed on the DL monitor, even after the foot switch was r | Software design | Consignees sent a "GE Urgent Medical Device Correction" letter dated February 28, 2008. The letter was addressed to Hospital Administrators/Hospital Risk Management Department/Managers of Radiology-Cardiology/ Radiologists-Cardiologists. The letter described the affected products, safety Issue, short term recommendations, long term solution and contact Information. The first batch of letters was received by customers on 5/28/08. For further information, please contact GE Healthcare at 262-544-3894 |  | 427 | 427 | USA including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL IN, KS, KY, LA, ME, MD, MA, MI,MN, MS,MO, MT, NE, NV, NH, NJ, NM, NY,NC, ND, OH, OK OR, PA, RI, SC, SD, TN, TX UT,VT, VA, WA,WV, and WI; the United States territory of Puerto Rico; and countries of ALGERIA, AUSTRALIA, BELARUS, BELGIUM, BOSNIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ,FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN,KOREA, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND,NORWAY, PANAMA, POLAND, PORTUGAL RUSSIAN FEDERATION SAUDI AR ARABIA, SINGAPORE, SLOVAKIA, SPAIN, SUDAN, SWEDEN SWITZERLAND, SYRIA, TAIWAN, TURKEY UNITED KINGDOM, and VENEZUELA. | Terminated | March 07, 2016 | 2959 | 1127 | Software | Display/Image | Notification | Safety Notice/Insructions
Z-0652-2009 | 50141 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Stryker Endoscopy FloControl Arthroscopy Pump Integrated Tubing, Model number 0350600006, for use with the FloControl Arthroscopy Pump. |  |  | October 01, 2008 | 2008 | Stryker Endoscopy | Unable to pump fluids. Silicone tubing is stiffer than usual, resulting in motor defect message, leading to inability to pump saline through tubes for surgery. | Process control | Stryker Endoscopy initiated the recall 10/01/2008. Information gathered and recall packets sent via FedEx to all affected accounts. Packets included advisory notices and acknowledgement card, as well as shipping labels to return product. Each account was instructed to locate and quarantine product for return to Stryker endoscopy. Sales representatives received a voicemail/email notification that included a list of all affected accounts. |  | 1440 | 1440 devices | Nationwide. | Terminated | December 16, 2010 | 806 | 1440 | Not_Computer
Z-0598-2009 | 50127 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Alimed Gait Belt (Transfer Aid), Antimicrobial Plastic buckle, 54" |  |  | October 30, 2008 | 2008 | AliMed Corporation | Misassembled gait (transfer aid) belt. | Employee error | Alimed notified accounts via U.S. Postal Service (Return Receipt Requested) with a Recall Notice dated 10/30/2008 and instructs users to review their records and notify all users to whom the product may have been issued. If any of the product has been issued outside of their facility, customers are to forward the correspondence to them in it's entirety. Customers are to complete the enclosed form for return of their product (new or used) and the issuance of replacement product or credit. |  | 993 | 993 units total | Nationwide Distribution | Terminated | September 09, 2011 | 1044 | 993 | Not_Computer
Z-0628-2009 | 50147 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy HI-ART Systems with versions 2.2.4, 3.1.2, 3.1.3 or 3.2.1 software. The affected applications include Planning Station, Planned Adaptive, Data Management System, and TomoPortal. |  |  | October 14, 2008 | 2008 | TomoTherapy Incorporated | TomoTherapy Incorporated identified a potential issue with the Hi-Art system during the course of ongoing testing. The Operator Station Calibration panel provides access to view and modify machine specific configuration settings. Access to these settings has always been restricted to individuals with appropriate security rights, being limited to only the "Superuser" and "Field Service engine | Software design | A TomoTherapy Incorporated "Urgent Medical Device Correction Field Safety Notice" letter dated October 17, 2008 was sent to all consignees. The letter addressed the Issue, Product Affected, Recommended Actions and Resolution. Please direct questions to the TomoTherapy, Inc. Customer Care Center by email (www.support@tomotherapy.com) or by telephone, North America: 1-866-368- 4807 and all other locations: 1-608-824-2900 or 32-2-400-44-44. |  | 215 | 215 | Worldwide Distribution --- USA including states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, NE, NV, NJ, NM, NY, OH,OK , OR, PR, PA, SC, SD, TN, TX, VA, WA, WI, WY, and DC, and countries of BELGIUM, CANADA, SWITZERLAND, CHINA, GERMANY, SPAIN, FRANCE, GREAT BRITAIN, HONG KONG, ITALY, JAPAN, KOREA, NETHERLANDS, SINGAPORE, TAIWAN and TURKEY. | Terminated | February 10, 2011 | 849 | 215 | Not_Computer
Z-0907-2009 | 50750 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Expressaire, Disposable Tourniquet Cuff 34", 1 line (18), sterile, Manufactured for Sammons Preston, A Patterson Medical Company, Cedarburg, WI.; Cuff type A70107010. Each cuff is marked ExpressAire by Sammons Preston and A70107010. |  |  | November 24, 2008 | 2008 | Virtus, Inc. | The inflation tube may come loose at the connection to the cuff during use. | Employee error | Sammons Preston was notified of the recall by letter dated 11/24/08, which instructed them to examine their inventory immediately for the specific lot numbers. The quantity on hand should be indicated on the table attached tothe letter and the info should be faxed to Virtus at 812-933-0749. Customers that may have tourniquet cuff inventory of identified lot numbers should be contacted. Once recall items have been located in quarantine, Tara Collins, Quality Technician, should be contacted at 812-933-1121 for Return Material Authorization number. |  | 3654 | 3654 | Nationwide Distribution --- including state of Wisconsin. | Terminated | October 22, 2010 | 697 | 5958 | Not_Computer
Z-0626-2009 | 49591 | No | NUC | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Pre Isotonic Personal Lubricant labeled PRE' Patient Lubricant. |  |  | September 22, 2008 | 2008 | INGfertility, LLC | Product contaminated with bacteria (Brachybacterium rhamnosum, a non pathogen gram positive cocci). | Pending | Firm telephoned all wholesale, clinic, and retail customers on 9/22/08 requesting that product be returned. |  |  | Unknown*** | Worldwide Distribution including USA and country of Ireland. | Terminated | May 02, 2012 | 1318 |  | Not_Computer
Z-0885-2009 | 50544 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer NEXGEN Complete Knee Solution Cruciate Retaining (CR) Provisional - Articular surface, Size Blue, 14 MM Height, nonsterile, Zimmer, Warsaw, IN; REF 5971-50-14. |  |  | December 12, 2008 | 2008 | Zimmer Inc. | Some 12 mm thick devices may have been etched and packaged as 14 mm thick, and vice versa. | Employee error | Zimmer sales force and distribution sites were notified of the recall via email on 12/15/08 and instructed to inventory, return the affected products and to provide a copy of the recall letter to those customers to whom they had further distributed the products. The 'Urgent: Device Recall" letter dated 12/12/08, described the problem and the potential risks. For additional information, contact Zimmer, Inc., at 1-800-613-6131. |  | 131 | 131 of both products. | Nationwide, Canada, Denmark, Finland, Germany, Poland, Spain, Sweden and United Kingdom. | Terminated | November 12, 2009 | 335 | 131 | Not_Computer
Z-0854-2009 | 50112 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Integra Ruggles Surgical Instrument ACF Distraction Screws |  |  | October 22, 2008 | 2008 | Integra LifeSciences Corp. | A packaging error had occurred where 16 mm screws were packaged and labeled as 12 mm screws and vice versa. | Packaging | Urgent: Product Recall Notification letters were sent via Fed Ex on October 22, 2008. The letter asked customers to quarantine any of the affected stock on hand and contact their local Integra Sales Representative for further assistance and for arrangement of product returns. Customers were to fill out the attached "Recall Acknowledgement and Return Form" indicating whether or not they possess any of the product identified in the letter. Additional questions can be directed to Christin Roach at 609-936-5413.A Urgent: Product Recall Second Notification dated 1/5/2009 was sent to customers providing them with additional information on the packaging error and included immediate and/or product used during surgery. This notification also requested that all surgeons be notified who were or may have been provided with any of the screws affected by this recall. |  | 351 | 351 units | Nationwide Distribution | Terminated | July 15, 2009 | 266 | 351 | Not_Computer
Z-0818-2009 | 50232 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Tornier AFFINITI Humeral Stem Standard 10mm, Catalog No. 0020010. Sterile. Manufactured for Tornier, 10750 Cash Rd., Stafford, TX 77477. Tornier's Affiniti Shoulder System is a non-constrained glenohumeral prosthesis intended for use as a total or hemi-shoulder replacement system. The total shoulder consists of a metal humeral stem, a metal humeral head and an ultrahigh molecular weight polyethylene glenoid. The hemi-shoulder consists of a metal humeral stem and a metal humeral head. |  |  | October 20, 2008 | 2008 | Tournier, Inc. | Tornier, Medical has received a report of a mislabeled sterile AFFINITI head product. As a result, Tornier has decided to voluntarily recall the potentially affected lots, including two (2) lots of AFFINITI Stems. | Labeling Change Control | An Urgent Product Recall letter, dated 10/17/08, was hand delivered to end users by distributors and sales representatives beginning 10/20/08. The recall letter identified the issue and requested end users to immediately discontinue use of the affected products and return any unused affected products to Tornier. Distributors were to notify any and all accounts that may be end users or to provide this information to Tornier, whom will contact the end users. A questionnaire and return instructions were provided with the letter. The questionnaire was asked to be completed and returned to Tornier as soon as possible. For additional information, contact Tornier Customer Service at 1-888-494-7950. |  | 401 | 401 total | AR, CA, FL, NC, NE, OH, PA, TN, TX, WA, and WI | Terminated | May 07, 2011 | 929 | 401 | Not_Computer
Z-0860-2009 | 50491 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | CombiSet True Flow Series, Hemodialysis Blood Tubing with Priming Set and Transducer Protectors, with Access Flow Reversing Connector (Twister) |  |  | November 25, 2008 | 2008 | Fresenius Medical Care North America | Port may crack and separate resulting in potential blood exposure/blood loss for the patient | Process control | Fresenius Medical contacted customers by e-mail,telephone and hard copy letter advising users to immediately discontinue use and return inventory. All internal customers were emailed on 11/25/08. All external customers were US Certified Mailed on 11/26/08. All internal customers were US Certified Mailed on 12/2/08 &amp; 12/3/08. |  | 351743 | 351,743 units | Worldwide Distribution --- including USA and country of Canada. | Terminated | September 24, 2009 | 303 | 351800 | Not_Computer
Z-0836-2009 | 50538 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | December 12, 2008 | 2008 | GE Healthcare Integrated IT Solutions | There is a potential patient safety issue involving the use of the F12 keyboard Approve &amp; Complete function. The issue may cause erroneous information to be added to a report when trying to approve and complete it. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 12/12/08 of the software anomaly resulting in a patient safety issue involving the use of the F12 keyboard Approve &amp; Complete function. The issue may cause erroneous information to be added to a report when trying to approve and complete it. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653. |  | 138 | 138 units | Worldwide Distribution including USA and countries of Canada, China, Denmark, Germany, Hong Kong, India, Ireland, Italy, Japan, Korea, Lebanon, Mexico, Netherlands, Saudi Arabia, Spain, Taiwan and United Kingdom | Terminated | May 18, 2010 | 522 | 138 | Software | Output/Calculation | Software update | Software Update
Z-0929-2009 | 50770 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Somatom Definition AS, Model number 8098027. |  |  | December 09, 2008 | 2008 | Siemens Medical Solutions USA, Inc | possible improper installation of tilt transport lock/tilt limiter | Labeling Change Control | The recalling firm received a letter dated 11/25/08 from the manufacturer informing them of the problem. As a result, the recalling firm has issued Update Instruction CT065/08/S to have the transport lock/tilt limiter on the device visually inspected. The tilted scan modes have been disabled until the modified lock/tilt limiter has been installed. |  | 42 | 42 units | Nationwide Distribution --- including states of AZ, CA, CO, FL, IL, KS, KY, MA, MI, MN, MO, NE, NY, OH, PA, SC, VA, and WY. | Terminated | September 23, 2009 | 288 | 42 | Not_Computer
Z-0842-2009 | 48760 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Urological Endoureterotomy Stent Set, 10Fr/13 cm, 6.0Fr/24 cm, sterile, Cook Urological, Inc., Spencer, IN; Ref EUSCFS-060024-10-13; GPN REF G18132. Ureteral stent. |  |  | July 31, 2008 | 2008 | Vance Products Incorporated | The bond may break where the two materials meet, resulting in separation of the stent. | Process control | Consignees were notified via "Urgent: Recall Notice" letter dated 7/31/08. The letter described the failure, instructed that the products be isolated and returned to the firm. For additional information, contact Cook Urological at 1-812-339-2235, ext. 7241. |  | 1739 | 1739 of all products. | Nationwide, Australia, Canada, Colombia, Hong Kong, India, Ireland, Israel, Japan, Malaysia, Mexico, South Korea, Taiwan, Venezuela and Viet Nam | Terminated | November 27, 2009 | 484 | 1739 | Not_Computer
Z-0587-2009 | 49603 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable Pump, Model 8627-10. 10 mL Reservoir. The contents [STERILE] of the inner package have been sterilized by ethylene oxide gas. The implantable Medtronic SynchroMed EL programmable Pump is part of the SynchroMed EL Infusion System designed to contain and administer parenteral drugs to a specific site. The implantable components of the SynchroMed EL Infusion System include the pump with or without a side catheter access port, catheters, and catheter accessories. Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604, USA. |  |  | August 25, 2008 | 2008 | Medtronic Neuromodulation | Medtronic issued a letter to healthcare professionals providing important safety information regarding MRI (magnetic resonance imaging) effects on SynchroMed EL and SynchroMed II implantable infusion pumps. As stated in product labeling, the magnetic field of an MRI will temporarily stop the rotor of the pump motor and suspend drug infusion for the duration of MRI exposure for all SynchroMed pump | Labeling Change Control | An Urgent Medical Device Correction, Important Information on Potential MRI Effects letter, dated August 2008, was sent to health care professionals beginning August 25, 2008. The letter described the issue, provided patient management recommendations and included enclosures: Roller Study Procedure to confirm pump function, Post MRI Pump Interrogation Guidelines and New Information Regarding Potential MRI Effects. |  | 76253 | Approx 76,253 of all devices. | Worldwide Distribution --- USA including Washington D.C., and Puerto Rico, and countries of Algeria, Aruba, Austria, Belarus, Belgium, Croatia, Cyprus Czech Republic, Denmark, Egypt, Finland, France, French Polynesia, Germany, Greece, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy , Jordan Kuwait, Lebanon, Luxembourg, Malta, Netherlands, Netherlands Antilles, Norway, Pakistan, Poland, Portugal, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia and Montenegro, Slovakia, Slovenia, Sough Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Vatican City State, New Caledonia, Brazil, Dominican Republic, Mexico, Hong Kong, Singapore, Canada, Australia and Japan. | Terminated | May 17, 2012 | 1361 | 76253 | Not_Computer
Z-0001-2009 | 50138 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | NeuViz Dual Multi-Slice Computed Tomography (CT) Scanner System (software). |  |  | April 10, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | An issue associated with the "Timed Scan" application was found. The system does not wait for the programmed time interval and starts the x-ray after selection of the scan button when certain criteria is met. | Software design | A letter was issued to customers prior to the release of the field correction during June 2008. The system software will be updated to correct the non-compliances. For information or support concerning this issue, please contact the Neusoft Medical Systems Service Support Department (helpdesk@pnms.neusoft.com) or contact Philips Medical Systems North America Co. Phillips at 425-487-7602. |  | 19 | 19 units in the US. | Nationwide Distribution --- USA including states of CO, FL, GA, IL, KY, NC, TN and TX. | Terminated | October 29, 2010 | 932 | 19 | Software | Device Operation | Software update | Software Update
Z-0609-2009 | 50173 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | DLP(R) Pericardial/Intracardiac Sump Cannulae, 20 Fr., P/N: 12112. This product is intended for draining the pericardial sac or the cardiac chambers. |  |  | September 29, 2008 | 2008 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | Separates: Six complaints were received from two locations. In all cases the sump tip assembly separated from the tubing during device removal and had to be retrieved from the patient. There were no adverse events affecting these patients. | Nonconforming Material/Component | Consignees were sent a "Medtronic Urgent Medical Device Recall Notice" dated October 1, 2008. The letter described the product and issue. It also requested its customers to return the product and to complete, sign and fax the Recall Certificate "Medtronic Recall Certificate (FCA 0902)". For further information, please contact Medtronic Cardiovascular Revascularization &amp; Surgical Therap by telephone at 763-391-9000. |  | 560 | 380 US ; 180 OUS | Worldwide Distribution --- including states of NC, PA, IA, WI, IN, VA, IL, KY, and TX, and countries of Australia, France, Hong Kong, Netherlands, and Poland. | Terminated | May 07, 2011 | 950 | 560 | Not_Computer
Z-0887-2009 | 50230 | No | ILX | Shortwave diathermy. | Physical Medicine | Physical Medicine | 510(k) | Regenesis Biomedical PROVANT Wound |  |  | September 29, 2008 | 2008 | Regenesis Biomedical Inc | In the course of device testing, the firm determined that the affected devices may, under certain unusualclinical circumstance, be unable to generate a therapeutic dose of radio frequency energy. As such,for these few devices, the effectiveness of Provant therapy may be impaired. | Finished device change control | Written notification dated September 29, 2008 was sent to consignees by overnight mail notifying them of the nonconformance and providing instructions for return of the affected PROVANT Model 4201A units to Regenesis. Receipt of the notification and return shipping of the units will be verified. Contact Regenesis Biomedical Inc. for additional information at 1-480-970-4970. |  | 7 | 7 units | WY, VA, and NC | Terminated | July 08, 2009 | 282 | 7 | Not_Computer
Z-0888-2009 | 50485 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Solar Shoulder Humeral Stem |  |  | November 13, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware of a visual defect in certain blister packages of certain devices. | Packaging | Recall notification letters were sent via Federal Express on November 21, 2008 to Stryker Branches/Agencies, Hospital Risk Management, Chief of Orthopaedics and Surgeon. The letters explained the problem and indicated that there are no risk factors. Contact Stryker Orthopaedics for additional information at 1-201-831-5718. |  | 322 | 322 total units - all types | Nationwide. OUS: Australia, Canada, Japan, Argentina, Brazil, Germany, Spain, Columbia, Guatemala and Sweden. | Terminated | January 20, 2010 | 433 | 322 | Not_Computer
Z-0884-2009 | 50498 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoration Modular Conical Distal Stem; |  |  | November 21, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker became aware of a lot for lot mix up. | Labeling mix-ups | Recall notification letters sent via Federal Express on Nov. 21, 2008 to Branches/Agencies, Hospital Risk Management, Chief of Orthopaedics and Surgeons. Letters explain the labeling mix up and the potential risks to patients if the wrong size stem is used. Stryker branches are asked to examine inventory and that of hospital locations to identify affected product. All affected product is to be reconciled on the Product Accountability form and the form faxed back to Stryker. Contact Stryker Howmedical Osteonics Corporation for additional information at 1-201-831-5718. |  | 4 | 4 | Class II Recall - Nationwide Distribution | Terminated | June 25, 2012 | 1312 | 4 | Not_Computer
Z-0904-2009 | 50694 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Zimmer TIV Self Tap Screw, 40 mm. Manufactured by Zimmer Manufacturing B.V., A Subsidiary of Zimmer, Inc., Warsaw, IN; Item No. 48231904000. NOTE: This product is not sold in the United States. |  |  | December 12, 2008 | 2008 | Zimmer Inc. | The package may contain a 40 mm long screw instead of the labeled 42 mm long screw, or vice versa. | Process control | The firm's facility in Japan was notified via letter, "Urgent: Device Recall" (dated 12/11/08) and requested to complete an enclosed Inventory Return Certification form. The letter instructed distributors to notify accounts to the user level to stop using the affected implants and return them to the Zimmer Distribution Center in the U.S. and fax (574-372-4265) a copy of the completed Inventory Return Certification form to Zimmer, Inc. For questions about the recall, please contact Zimmer, Inc. at 1-800-613-6131. |  | 100 | 100 | International Distribution including country of Japan. | Terminated | October 30, 2009 | 322 | 508 | Not_Computer
Z-0926-2009 | 50189 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump 2 Blockaid, 400 mL PainPump for Continuous Nerve Block, sterile, 5 per box. Stryker, Kalamazoo, MI; REF 575-100, Product Number 0575-100-000. |  |  | November 12, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | The pump may display an E-4 error message, which causes the pump to stop dispensing the pain medication. | Process design | Consignees were notified via letter dated 11/12/08 and instructed to locate and destroy the affected lots. Customers were also instructed to complete and return the enclosed Business Reply form via fax (336-499-8603) to Stryker Instruments Regulatory Department. For questions regarding the recalled product, please contact Stryker Instruments at 1-800-800-4236, extension 3808 or extension 4056. |  | 2009 | 2,009 boxes | Worldwide Distribution including USA and country of Canada. | Terminated | November 30, 2009 | 383 | 2009 | Other | Treatment/Delivery/Therapy | Remove | Remove or Replace
Z-0930-2009 | 50169 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Invacare Storm TDX SP Power Wheelchair with Stability Lock, |  |  | November 27, 2007 | 2007 | Invacare Corporation | Invacare has become aware that on some chairs, the stability lock feature may not be engaging properly or consistently. If the stability lock feature does not engage properly, the wheelchair may veer to one side or it may rock forward onto its front riggings. | Pending | On 11/27/2007 the firm sent Urgent Field Correction Needed recall notification letters and response forms to their customers informing them of the problem. They instructed their customers to immediately inspect and adjust the Stability Lock feature, distribute the enclosed update to the Owner's Manual to all end-users, and return the enclosed response form notifying them that the inspection and/or adjustments have been completed. |  | 4171 | 4171 Wheelchairs | Worldwide Distribution --- including USA and countries of Canada, Australia, Germany, New Zealand and the United Kingdom. | Terminated | August 01, 2013 | 2074 | 4171 | Not_Computer
Z-0606-2009 | 50155 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Liko AB Uno Lifts, Powered (non-AC) Patient Lift: |  |  | November 11, 2008 | 2008 | Liko North America Inc | Actuator of the lift has the potential to separate when the upper arm is fully extended. | Process control | Hill Rom notified customers with a Field Safety Notice, dated December 5, 2008 to notify customers of the issue, potential risks, and to inspect units per their instructions and remove units from service if the unit fails inspection. Hill-Rom will provide a permanent fix, using an "OuterTube" to be placed on all units. |  | 24070 | 24070 total units | Worldwide Distribution --- including USA and country of Canada. | Terminated | February 04, 2016 | 2641 | 24070 | Not_Computer
Z-0654-2009 | 50503 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | CAD SCIENCES LLC 3TP PrecisionPointTM Software, for biopsy guidance. |  |  | December 09, 2008 | 2008 | iCAD, Inc. | Software modules not approved for this indication. | PMA | iCAD initiated a voluntary field correction by telephone and letter on 12/5/08 by disabling these features at all affected sites. Since Precision Point and TheraMap are controlled by a software license key, iCAD will disable these features by updating the license at the affected sites. |  | 35 | 35 | Nationwide distribution to: CA, FL, HI, KS, NH, NJ. NY, MA, MO, MI ,NC, OH, SC, and WI. | Terminated | January 28, 2016 | 2606 | 35 | Computer
Z-0928-2009 | 50539 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis). GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | November 21, 2008 | 2008 | GE Healthcare Integrated IT Solutions | There is a potential patient safety issue involving patient jacket content, which may result in unintended historical image(s) being displayed. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 11/21/08 of the software anomaly resulting in a patient safety issue involving patient jacket content, which may result in unintended historical image(s) being displayed. The letter provided the users with safety instructions to follow until the software is updated. Please direct any questions to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 414-918-4653. |  | 476 | 476 units | Worldwide Distribution including USA and country of Canada. | Terminated | September 06, 2010 | 654 | 476 | Software | Display/Image | Software update | Software Update
Z-0663-2009 | 50510 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Ethicon Endo-Surgery ENDOPATHÂż Electrosurgery Probe Plus II 29 cm Needle Probe (Product Code EPS09) |  |  | December 09, 2008 | 2008 | Ethicon Endo-Surgery Inc | Nickel exposure: The stainless steel tips on the affected product were degrading, causing a small amount of nickel to be released, which may lead to health risks in some patients, particularly those with a known or suspected nickel allergy or sensitivity. Ethicon Endo-Surgery determined the presence of free nickel in the affected product could elicit a reaction in those patients with nickel sens | Process design | On 12/9/2008 Ethicon Endo-Surgery sent recall notification letters to their customers instructing them to stop using the product, examine all inventory to determine if they have the affected product on hand and remove the affected products from inventory. They requested that each customer complete the Business Reply Form and fax it back to Stericycle and also return any affected product to Stericycle. |  | 2461 | 13 lots, 2,448 pieces | Worldwide distribution: USA, PUERTO RICO, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, DOMINICAN REPUBLIC, EGYPT, ENGLAND, EQUADOR, GERMANY, GREECE, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KOREA, MALAYSIA, MEXICO, PAKISTAN, PANAMA, POLAND, REPUBLIC OF YEMEN, RUSSIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD, TURKEY, UNITED ARAB EMIRATES, URUGUAY, and VENEZUELA. | Terminated | October 20, 2010 | 680 | 48372 | Not_Computer
Z-0932-2009 | 50680 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AMS InteXen LP Collagen Dermal Matrix, REF 72404066. Product Size: 8 x 12 cm. AMS InteXen LP is a sterile non-perforated processed porcine collagen dermal matrix. Sterile EO. American Medical Systems, Inc., 10700 Bren Road West, Minnetonka, MN 55343 USA. |  |  | December 05, 2008 | 2008 | American Medical Systems | The product size shown on the carton labeling does not match the size on the product inside. | Labeling mix-ups | Customers were first notified by phone on 12/05/08. They were told of the situation and were requested to return the devices. Customers were also provided an Important Product Recall letter, dated 12/05/08. This letter identified affected product, told customers to discontinue use of product and to return it. An acknowledgement form was also asked to be returned. AMS Customer Service should be contacted at 1-800-328-3881 to receive a Return Authorization number and arrange for shipment of replacement product. |  | 5 | 5 | Nationwide Distribution --- including states of TN and TX. | Terminated | June 05, 2012 | 1278 | 5 | Not_Computer
Z-0883-2009 | 50119 | Yes | MVK |  |  | Cardiovascular | PMA | Monitor portion of the LifeVest device - specifically the Defibrillator printed circuit assembly (PCA) within the monitor. |  |  | October 21, 2008 | 2008 | Zoll Lifecor Corporation | The charging circuit will not shut down properly. | Component design/selection | Physicians with active patients were notified of the product removal by letter via first-class mail dated 10/21/2008. The letter describes the problem and the potential risk. For additional information, contact ZOLL Lifecor Corporation at 1.800.543.3267. |  | 1685 | 1685 units | Nationwide. OUS: Germany, United Kingdom, Italy, and France. | Terminated | November 04, 2009 | 379 | 1685 | Hardware | Device Operation | Remove | Remove or Replace
Z-0664-2009 | 50128 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stelkast Dual Taper Wedge Pressed Fit Femoral Component. Hip Prosthesis, Part number SC1439 |  |  | October 07, 2008 | 2008 | Stelkast Co | Sterility may be compromised due to breach in packaging. | Package design/selection | On 10/7/08, Stelkast telephoned the two customers of the problem and the need to return the products. |  | 6 | 6 units | USA distribution to dealers in TX and WV. | Terminated | April 01, 2009 | 176 | 6 | Not_Computer
Z-0669-2009 | 50129 | No | JTZ | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | HARDY DIAGNOSTIC McFarland Standard 0.5; in vitro diagnostic, glass vial, one vial per box |  |  | September 22, 2008 | 2008 | Trek Diagnostic Systems | Incorrect result: The absorbance has changed as the standards have aged; the standards' turbidity has become denser. The recalled lot is now reading 0.15 ABS @ 625 nm. TREK'S internal specification and the manufacturer's (Hardy Diagnostics) specification is 0.08 - 0.10 ABS @625 nm. CLSI specifies 0.08 - 0.13 ABS @ 625 nm. | Nonconforming Material/Component | On 9/22/2008 TREK Diagnostic sent a recall notification letter to their customers advising them of the problem with the product. They instructed their customers to inventory their stock for any remaining recall lot and immediately destroy. |  | 4 | 4 vials | USA distribution in OH and IL. | Terminated | July 06, 2010 | 652 | 4 | Not_Computer
Z-0668-2009 | 49711 | No | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus CS Acute Care Troponin I TestPak, Catalog number CCTNI, in vitro diagnostic. |  |  | June 23, 2008 | 2008 | Dade Behring, Inc. | Incorrect result: Potential for false elevation of CTnI results | Nonconforming Material/Component | Siemens issued an Urgent Field Safety Notice dated 6/23/08 with instructions to immediately discontinue use and discard the affected lots. Customers were also instructed to contact Siemens Customer Service Center at 800-241-0420 to request replacement product at no charge. |  | 5868 | 5868 cartons | Worldwide distribution: USA, Singapore, Saudi Arabia, Japan, Germany, Canada, Bahrain, Australia, Korea, and India. | Terminated | January 29, 2009 | 220 | 5868 | Not_Computer
Z-0837-2009 | 50140 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Atlantis PV Peripheral Imaging Catheter, model number H749364560, catalog number 36456, manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | October 09, 2008 | 2008 | Boston Scientific Corporation | Product is mislabeled as having a shelf life of 3 years, where the actual shelf life is 1 year. | Labeling mix-ups | Recall initiated on October 6, 2008, and notification letters sent on October 8, 2008. Letters were sent via FedEx. Responses are being tracked along with delivery verification. For additional information contact Boston Scientific Corporation at 1-508-652-5594. |  | 1033 | 1033 units | Worldwide Distribution --- USA including Puerto Rico and country of Panama | Terminated | July 27, 2009 | 291 | 1033 | Not_Computer
Z-0666-2009 | 50137 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic Maximo II VR D284VRC, for ventricular pacing and defibrillation. |  |  | July 25, 2008 | 2008 | Medtronic Inc. Cardiac Rhythm Managment | Maximo II VR ICD, DR ICD and CRT-D inadvertently include the Capture Management algorithm (Atrial/Right Ventricular/Left Ventricular Capture management). The Capture management algorithm monitors daily pacing thresholds and may increase or decrease pacing amplitudes in response to patients' needs. Although the algorithm is operating correctly, it is not accessible to the clinician for reprogramm | Process design | A Medtronic Inc. letter dated July 30, 2008 was sent to Physicians. The letter described the product and problem. The letter advised to contact Technical Services for specific questions regarding the care of patients or how to program the Maximo II devices. The letter instructed to collect and hold all non-implanted Maximo II devices and that a Medtronic Representative will contact the doctor to make arrangements to update the programmer 2090. |  | 376 | 376 | No US distribution. OUS to include: AUSTRIA, DENMARK, FINLAND, GERMANY, ITALY, SPAIN, SWEDEN, UK, GREECE, HONG KONG, INDIA, JORDAN, and SAUDI ARABIA. | Terminated | May 07, 2011 | 1016 | 1108 | Not_Computer
Z-0935-2009 | 50479 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bard 3D Max Mesh, Left large , Sterile |  |  | December 08, 2008 | 2008 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: Product labeled as Left large, may contain a Right Large | Mixed-up of materials/components | Davol issued a letter notification dated 12/25/08 via Federal Express requesting accounts discontinue use, remove product from inventory and return product.For additional information, contact Davol at 1-401-463-7000. |  | 232 | 232 units | Nationwide | Terminated | May 02, 2012 | 1241 | 232 | Not_Computer
Z-0913-2009 | 50749 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific EndoVive 20Fr (6.7 mm) Safety PEG Kit, Push Method, latex free, sterile, Boston Scientific Corp., Natick, MA; Cat. no./REF 6653, inner pouch UPN M00566530 and outer carton UPN M00566531. |  |  | December 30, 2008 | 2008 | Boston Scientific Corp | The user may be unable to pass the device over the guidewire for placement in the patient, resulting in an inability to use the device. | Other | Consignees were notified by "Urgent Medical Device Recall" letter dated 12/30/08. The letter described the problem with the product, instructed customers to determine whether they store these products in the outer carton or the inner pouch at their facility, to locate all affected products, and to return them to the firm. For additional information, contact Boston Scientific Corporation at 1-508-683-4678. |  | 2460 | 2460 of all products | Nationwide. | Terminated | October 30, 2009 | 304 | 2460 | Not_Computer
Z-0670-2009 | 44850 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Air Bubble Detection Module, Terumo Cardiovascular Systems, Ann Arbor, MI; Catalog number 802110. |  |  | November 20, 2007 | 2007 | Terumo Cardiovascular Systems Corporation | The system may malfunction due to t-filter corrosion on an internal circuit board. | Nonconforming Material/Component | U.S. consignees were notified of the problem by letter on 11/16/07. International affiliates were informed via email on 11/21/07. Consignees were informed that Terumo would schedule appointments to install new pressure modules and reconfigure the systems beginning in Feb. 2008 or that parts would be shipped to consignees wishing to perform the work themselves. For additional information, contact Terumo CVS Customer Service at 1-800-521-2818. |  | 1219 | 1219 (total) | Nationwide, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Viet Nam. | Terminated | August 06, 2009 | 625 | 1219 | Hardware | Device Operation | Replace component | Remove or Replace
Z-0867-2009 | 50516 | [6 empty] | REF 01-6851 Office System 1.5/IMF Small Office Set, 156MMX125MM, (6.4"X4.9") Polypropylene, Ultem, Radel, AL, SS. Non Sterile Product. W.LORENZ SURGICAL, Walter Lorenz Surgical, Inc. 1520 Tradeport Drive |  |  | August 06, 2008 | 2008 | Biomet Microfixation, Inc. | Biomet Microfixation is recalling the Lactosorb Endobrow Lift Complete, Office Fixation Kit, 1.5/IMF Small Office Set and IMF Set trays containing instruments associated with the Lactosorb and IMF systems. The finished trays were packaged and shipped to facilities without the instructions for use and without the single use symbol on the packaged tray label. | Other | An IFU and correction letter will be issued to the field. Each distributor will receive instructions for delivering the correction letter to all customers. The customer letter will contain a reconciliation form to be returned once the customer receives the IFU for the twist drill. Contact Biomet Microfixation for additional information at 1-800-874-7711. |  | 33 | 33 | Distributed domestically throughout the U.S. Distributed internationally to Argentina, Australia, Canada, Columbia, Hong Kong, Japan, Mexico, New Zealand, Puerto Rico, Taiwan, Thailand, Denmark, France, Netherlands, Greece, Italy, Norway, Finland, Spain, Czech Republic, and Venezuela. | Terminated | September 11, 2009 | 401 | 420 | Not_Computer
Z-0927-2009 | 50515 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV SA Linear Accelerator with syngo RT Therapist Version 4.1 build 67 and build 122. Manufactured by Siemens Medical Solutions Oncology Care Systems, Concord, CA. |  |  | October 23, 2008 | 2008 | Siemens Medical Solutions USA, Inc | If a plan contains a sequence of beams, including at least 1 ARC treatment beam and if during the setup stage of a beam following an ARC beam in the delivery sequence a termination occurs, and the user selects the immediate resumption feature, syngo RT Therapist Express systems will display and record a second delivery of a prior beam and not the delivery of the current beam in certain circumstanc | Software design | An Urgent Medical Device Customer Safety Letter dated October 23, 2008 was issued to consignees via certified mail. The letter identified the issue and preventive measures for consignees. Monitoring will be done via return cards. Follow-up to be conducted via telephone. |  | 36 | 36 units | Worldwide Distribution including USA and countries of Netherlands, Germany, Poland, Norway, Australia, Hungary, Belgium, New Zealand, France, Malaysia, Spain, and Canada. | Terminated | December 09, 2010 | 777 | 36 | Other | Display/Image | Notification | Safety Notice/Insructions
Z-0903-2009 | 50512 | No | OEX | Catheter introducer. | Cardiovascular | Cardiovascular | Enforcement Discretion | Spectranetics SLS II 14Fr. Laser Sheath Kit, REF 500-012, Sterile EO, individually packaged, Lot number 080314B, Spectranetics, 96 Talamine Court, Colorado Springs, Colorado 80907. |  |  | October 15, 2008 | 2008 | Spectranetics Corporation | Catheter pouch was mislabeled with the wrong size. | Error in labeling | Consignees were notified by sales representatives who visually inspected their inventory. The sales representatives were given a form to fill out to confirm whether the consignee had the catheters with the serial number still in inventory. The forms were faxed back to Spectranetics indicating if the catheters were used or not and if they were not used, were they labeled correctly. For assistance, contact Spectranetics Corporation at 1-719-447-2319. |  | 24 | 24 units | CA, FL, MI, NC, NY, OH, PA, TX, and VA. | Terminated | February 26, 2009 | 134 | 24 | Not_Computer
Z-0946-2009 | 50521 | No | OCL | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | Epicor LP Connecting Cable, CC-6-LP-CE, Catalog Number 12263, an accessory to the Epicor Cardiac Ablation Control System, manufactured by St.Jude Medical, Sunnyvale, CA |  |  | December 09, 2008 | 2008 | St. Jude Medical Inc | Firm has received multiple reports of difficulties connecting the product and/or the presence of error code messages occurring with the use of the connecting cable and the Ablation Control System unit. | Other | The firm issued notification by "Urgent Product Recall" letter dated 12/4/2008, to all consignees. The letter describes the issue, potential risk and states that the product should be returned. For additional information, contact St Jude Medical at 800-722-3423. |  | 158 | 158 units distributed. | Worldwide - NY, IN, NC, UT, OR, CA, IL, MI, TX, TN, and FL and country of Belgium. | Terminated | December 28, 2010 | 749 | 158 | I/O | Device Operation | Remove | Remove or Replace
Z-0931-2009 | 50523 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Low Profile U-Joint Driver for the SynFix Mini-Open System, Catalog number 03.802.231. |  |  | November 20, 2008 | 2008 | Synthes Spine | screwdriver tip may break off and lodge in the screw head | Device Design | The recalling firm emailed all Spine Sales Consultants on 11/20/08. They were instructed to immediately examine their stock and discuss appropriate technique with surgeon users. In addition, a Medical Device Correction letter dated 11/19/08 was issued via USPS certified mail to the medical facilities. This initial letter stated the issue and requested the users take the following actions: discard all SynFix Technique Guides and replace with the updated literature contained in the mailing packet with the letter, and is identified by code J7022-E (Rev. E); to ensure receipt and understanding of the information, the Verification Section of the letter needs to be completed and returned to Synthes by either fax: 610-719-5120 or scan/e-mail: lewis.lynne@synthes.com. An Urgent: Medical Device Recall letter, dated 1/16/2009, was also sent out. This latter letter requested that users examine their stock immediately and remove any of the affected devices from their inventory. Customers are also to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number; complete the Verification Section; and return the letter with the recalled product to the address listed. There are also actions for those who do not have the affected product. Questions should be directed to 800-620-7025 x 5453 or 610-719-5453. |  | 381 | 381 units | Class II Recall - Worldwide Distribution --- USA including states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, PA ,SC, TN, TX, UT, VA, WA, and WI, and countries of Canada and Switzerland. | Terminated | October 27, 2009 | 341 | 381 | Not_Computer
Z-0949-2009 | 50787 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Artiste MV Digital Linear Accelerator System Equipped with 160 MCL Accessory Holder, Manufactured by Siemens Medical Solutions, Oncology Care System, Concord, CA |  |  | December 20, 2008 | 2008 | Siemens Medical Solutions USA, Inc | The range of tolerance is to high for the accessory holder, therefore, Wedges and other accessories do not latch to the locking mechanism correctly. | Device Design | On 12/20/08, the firm issued letters via certified mail to direct consignees, informing them of the affected product and providing instructions on the recall.Contact Siemens Medical Solutions USA, Inc. at 1-925-246-8407 for additional information. |  | 43 | 43 units | Nationwide and Worldwide | Terminated | December 09, 2010 | 719 | 43 | Not_Computer
Z-0933-2009 | 50756 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Ergolift 400 Lbs. Floor Lift, Manufactured by BHM Medical Inc., 2001 Tanguay, Magog, QC, Canada, J1X 5Y5; model/part numbers ERGOLIFT and ERGOLIFT-2 |  |  | December 08, 2008 | 2008 | Arjo, Inc. | The manufacturer, BHM Medical, has received reports of structural failure of the mast during use. | Process control | Arjo sent the Urgent Device Recall letter dated 12/22/08 and Field Safety Notice (FSN) dated 12/8/08 to their Ergolift customers on 12/22/08. The accounts were informed of the potential structural failure of the mast on the BHM Medical Ergolift. The letter instructed the accounts to discontinue use of the lifts immediately. Those lifts with serial numbers ERLI-0001 and ERLI-0899 will be taken out of service permanently, and will be replaced with a new device at a significant discount. Those lifts with serial numbers ERLI-0900 and ERLI-1717 will be corrected by a reinforcement kit installed by a qualified technician. The accounts were requested to complete and return to Arjo Inc. within 10 days the enclosed customer response form indicating that they have received the notification, indicating the serial numbers on hand needing correction, and whether they need any replacement lifts. The accounts will be contacted by an Arjo technician within weeks to arrange for the correction of the lifts and/or an Arjo representative for the replacement lifts. |  | 220 | 220 lifts | Nationwide Distribution. | Terminated | January 21, 2011 | 774 | 220 | Not_Computer
Z-0948-2009 | 50751 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | iSite PACS (Picture Archiving and Communications System) Versions 4.1.x up to and including 4.1.51.3, manufactured by Philips Healthcare Informatics, Foster City, CA. |  |  | December 29, 2008 | 2008 | Philips Healthcare Informatics, Inc. | Product may display clinical reports with the study date and time in GMT format, which may not be obvious to the end user, who may expect to see the local time in the report. | Software design | Product notification was issued on 12/29/2008, instructing users to identify affected products, and instructing users to refer to image overlays for local study date and time information. The firm is planning a field correction expected to be available in "Q1 2009". For additional information, contact Philips at 650-293-2624. |  | 40 | 40 units | International - Italy, Denmark, Germany, Netherlands, Sweden, Spain, Portugal, Switzerland, Spain, Turkey, France, Saudi Arabia, Australia, Thailand, New Zealand, Japan, and Brazil. | Terminated | November 29, 2010 | 700 | 40 | Software | Display/Image | Instructions | Safety Notice/Insructions
Z-0906-2009 | 50737 | No | EMG | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Obwegeser Ramus Clamp w/ wide shield, 20 cm, 8", Right. ORAL-MAXILLO-CRANIO-FACIAL SURGERY, Original OBWEGESER Instruments. Walter Lorenz Surgical, Inc., Subsidiary of Biomet Microfixation, Inc. NOTE: None distributed domestically. |  |  | May 21, 2008 | 2008 | Biomet Microfixation, Inc. | The Biomet Microfixation Obwegeser Ramus Clamp, Wide Right 01-7996 were incorrectly etched as the left part number 01-7997. | Labeling mix-ups | Customers were contacted by e-mail notification with read receipt. Customers were notified in the e-mail to immediately discontinue use of the recalled product and return the affected items to Biomet Microfixation, Inc. The e-mail also included an Inventory Reconciliation form for customers to complete and return via fax (904-741-1801) to Biomet Microfixation, Inc. For questions related to this notice, please call 1-800-874-7711 or 904-741-4400 extension 448, Monday through Friday, 8 a.m.-4 p.m. Eastern Standard Time (EST). |  | 4 | 4 | International Distribution including countries of Ireland and Japan. | Terminated | September 11, 2009 | 478 | 4 | Not_Computer
Z-0923-2009 | 50184 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump, Two Site Infusion Kit with ExFen 2.5", 120 mL Infusion Kit, sterile, 5 per box. Manufactured for Stryker Instruments, Kalamazoo, MI. Product Number 0500-166-025, REF 500-166-25. |  |  | November 07, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | There is a potential for the pump to overinfuse. | Process control | A "Medical Device Recall Notification" letter dated November 11, 2008 was issued to consignees. The letter identified the problem and instructed consignees to locate and destroy all affected products. Customers were also instructed to complete the enclosed Business Reply form and fax it to Stryker Instruments Regulatory Department at 866-521-2760. For questions regarding the recalled product, contact Stryker Instruments at 1-800-800-4236, extension 4056 or extension 3808. |  | 118 | 118 boxes of all products. | Nationwide Distribution. | Terminated | December 01, 2009 | 389 | 118 | Not_Computer
Z-0794-2009 | 48991 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | King Systems Universal Flex2 Breathing Circuit, Latex Free, 50 per case, King Systems Corporation, Noblesville, IN; REF # DF477S-6121Z. |  |  | August 01, 2008 | 2008 | King Systems Corp. | The inner line may be severed or broken into two pieces, thus hindering the delivery of oxygen and anesthetics to the patient and possibly elevated patient carbon dioxide levels. | Pending | Consignees were notified by an Urgent Product Recall letter on or about 7/29/2008 to cease sales, notify their customers to the hospital level, and to return product on hand. |  | 500000 | Approx. 500,000 units of all products. | Worldwide Distribution - including USA and countries of Australia, Colombia, Germany, Netherlands, and Peru. | Terminated | November 12, 2009 | 468 | 500000 | Not_Computer
Z-0939-2009 | 50192 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | VOYAGER NC Coronary Dilatation Catheter, 4.5 X 12.0 mm. |  |  | November 05, 2008 | 2008 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | The recall was initiated after Abbott Vascular received reports that, in a small percentage of cases, the protective, yellow balloon sheath of the 4.5mm diameter VOYAGER NC Coronary Dilatation Catheter has been tight. If unnoticed by the user, there is a potential to cause damage to the balloon dilatation catheter during removal of the sheath. The balloon sheath covers the balloon, maintaining th | Package design/selection | An Urgent Device Recall letter (dated November 6, 2008) with an Efficacy Check Reconciliation form was issued and hand-carried to affected customers by their respective Sales Representative who assisted in locating and returning the recalled units. The letter described the problem and informed customers that Abbott Vascular will replace returned recalled product with like product or with VOYAGER NC sizes when they become available. Additionally, customers were instructed to contact their Abbott Vascular local account representative to review their inventory, complete and return the attached Efficacy Check Reconciliation Form and return affected products to Abbott Vascular. |  | 126 | 126 units | Worldwide Distribution. | Terminated | August 13, 2012 | 1377 | 224 | Not_Computer
Z-0956-2009 | 50204 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | HR90K Platinum Series Headpiece Model Number AB-7300-001 ADVANCED BIONICS LLC 12740 San Fernando Rd. Sylmar, CA 91342 USA. The HiRes90K headpiece is used in conjunction with speech processors manufactured by Advanced Bionics. |  |  | July 24, 2008 | 2008 | Advanced Bionics LLC | Mistuned headpiece - Advanced Bionics recently became aware of a tuning issue with some Platinum Series and S-Seriesheadpieces shipped from May 2007 through April 2008. Tuning refers to the process through which the headpiece's communicationfrequency is adjusted to match that of the internal device. Potential clinicalsymptoms include:-No lock-Intermittent lock-Implant locks, but cann | Employee error | Advanced Bionics is notifying clinicians of this situation so that they can provideappropriate clinical management. The notification letter sent to clinicians in theUnited States was sent via FedEx so that delivery information could be confirmed. The letter advises that Advanced Bionics recently became aware of a tuning issue with some Platinum Series and S-Series headpieces shipped from May 2007 through April 2008. Auria and Harmony headpieces are notaffected by this issue. Tuning refers to the process through which the headpiece's communication frequency is adjusted to match that of the internal device.Since discovery of this issue firm has conducted an extensive technical analysis showing that the use of these headpieces does not pose any safety risk to the patient. While there has been no increase in thenumber of headpiece-related complaints, firm advises of the potential clinicalsymptoms:-No lock-Intermittent lock-Implant locks, but cannot measure impedances-Reduction in battery life, primarily for C1 patients with thick flaps.Per the letter, as long as the system maintains lock, there is no degradation in sound quality orstimulation. Further, the impact on battery life for CII and HiRes 90K patients is minimal.Medical professionals are advised that If they have any patients who display any of these clinical symptoms that they attribute to an affectedheadpiece, to advise the patient to contact Advanced Bionics Customer Service at 877-244-9541 and they will send them a replacement promptly. |  | 2049 | 2,049 units | Nationwide | Terminated | March 23, 2012 | 1338 | 4617 | Not_Computer
Z-0969-2009 | 50758 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | January 12, 2009 | 2009 | GE Healthcare Integrated IT Solutions | If a user site created a custom default display protocol (DDP) to display the current exam versus the historical exam and the user-defined DDP fails, the system displays the exams using a GE default hanging protocol. This may create a potential for misdiagnosis if the user does not review the date and time stamps in the title bars. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation software customers via letter (Urgent Medical Device Correction) dated 1/12/09 of the software anomaly resulting in a patient safety issue involving Default Display Protocols. The issue may create a potential for misdiagnosis if the user does not review the date and time stamps in the title bars. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479. |  | 778 | 778 units | Nationwide, Canada, Bermuda, Mexico, Ecuador, Chile, and Brazil | Terminated | September 06, 2010 | 602 | 778 | Software | Display/Image | Software update | Software Update
Z-0934-2009 | 50134 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | Posey Connecting Straps &amp; Belts/Locking Cuffs. |  |  | September 08, 2008 | 2008 | J T Posey Company | The recall was initiated after an investigation conducted by the Posey Company determined that a patient may be able to catch the belt on a stationary object (wall, pipe, etc.) and use their body force to tear or break the strap, creating freedom of arm movement. In some patient populations, such freedom of movement could pose an unwanted hazard of serious injury to a patient or caregiver. | Other | An "Urgent Medical Device Recall (Correction)" letter dated August 15, 2008 was mailed via US Postal Service to affected customers. The letter described the issue, instructed customers to secure a replacement label with the correct application technique over the picture in the 2008 Posey Catalog and provided a corrected Posey Locking Cuffs Instruction Sheet (19273). The Posey Company also provided updated literature for the Posey Locking and Non-Locking Connecting Straps 19270 and 19254.The Posey Company asked customers to complete and return the "Reply Fax-Back Acknowledgment" Form via fax (626-443-5418) or e-mail (vcimmarusti@posey.com) within 7 days. Contact your Sales Representative with any questions or concerns. |  | 51200 | 51,200 (Approximately) | Worldwide Distribution including USA and countries of Australia, Belgium, Canada, France, Germany, Ireland, Italy, Korea, Lithuania and Netherlands. | Terminated | April 16, 2012 | 1316 | 51200 | Not_Computer
Z-0936-2009 | 50774 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker AVS AL Spacer Trial Inserter; |  |  | September 08, 2006 | 2006 | Stryker Spine | DESIGN FLAW: Issues were reported concerning fracture of handle, difficulty in removing inserter and inserter difficult to lock. | Device Design | Market Withdrawal letters were sent out on September 8, 2006 to all branches who received device. The letter asked customers to examine their inventory and hospital consignment locations to identify the product. Also, customers needed to reconcile all products on the attached Product Accountability Form and fax a copy back. The affected products were to be retrieved and returned using the "Product Returns Procedure". |  | 37 | 37 | Nationwide Distribution --- including state of NJ. | Terminated | February 05, 2009 | 881 | 37 | Not_Computer
Z-0961-2009 | 50545 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PROMOS Reverse, Glenosphere Loosening Adapter, REF 24 000 047, S&N; 75005271, Rx only, QTY(1), non-sterile |  |  | November 21, 2008 | 2008 | Smith &amp; Nephew Inc | The product was breaking during use. | Other | All affected Smith &amp; Nephew International Distributors were notified via email on 11/21/2008. They were instructed to immediately notify their accounts concerning the recall. Contact Smith &amp; Nephew Inc at 901-399-6771 for assistance. |  | 51 | 51 units | Germany, France, Italy Austria, Switzerland, Belgium, United Kingdom, Greece, Australia, South Africa, Spain, and Cypress | Terminated | November 27, 2012 | 1467 | 210 | Not_Computer
Z-0962-2009 | 50699 | No | NRA |  |  | Orthopedic | PMA | Biomet Oxford Unicompartmental Knee, Phase 3, Right Medical Tibial Tray, Standard Size B, CoCr, sterile, Biomet UK Ltd, Bridgend, UK; . |  |  | December 19, 2008 | 2008 | Biomet, Inc. | Packages labeled as "right" actually contain a "left" component. | Packaging process control | Biomet salesmen were notified by phone and instructed to retrieve the products. One customer was notified by letter dated 12/19/08 that use of the product should cease and that a salesman would pick up the product. A second customer had implanted the product and was hand delivered a letter dated 12/19/08 instructing that the patient should be monitored and symptoms treated as necessary. Contact Biomet at 574-362-3983 for assistance. |  | 2 | 2 | Massachusetts and West Virginia. | Terminated | July 07, 2009 | 200 | 2 | Not_Computer
Z-0968-2009 | 50677 | No | NPQ | Factor V Leiden DNA mutation detection systems. | Hematology | Hematology | 510(k) | NFlex Stabilization System 150 mm, Catalog number NF60150-1. |  |  | November 07, 2008 | 2008 | Synthes Spine | the dimensional specifications are incorrect | Nonconforming Material/Component | The recalling firm issued an e-mail on 11/18/08 to their sales force to inform them of the problem and the need to remove the product.The recalling firm issued Medical Device recall letters dated 11/19/08 to their customers to inform them of the problem and the need to return the product. Contact Synthes at 800-620-7025 ext. 7025 or 610-719-5750 for additional information. |  | 1508 | 1508 (total) | Worldwide: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI, and WV. OUS: Germany, Denmark, Switzerland, and Korea. | Terminated | October 05, 2009 | 332 | 1508 | Not_Computer
Z-0872-2009 | 50676 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | SERVO Ventilator 300A (SV300A); |  |  | December 19, 2008 | 2008 | MAQUET Inc. | SERVO Ventilator 300A could cause hypoventilation if the use of the PC or PRVC mode with the Automode function set to "On" in Neonatal and Pediatric patient ranges with small tidal volumes and short set inspiratory rise time. | Device Design | URGENT: Device Correction letters were sent out on December 22, 2008. Letters alert users to the potential problem and the situation in which it may occur. Letter indicate the set inspiratory rise time should be set in the range of 7-10% and that the humidifier chamber water level be maintained according to manufacturers instructions. The letter is enclosed as a supplement to the SV300A Operating Manual. Consignees were directed to complete a "Device Correction-Customer Fax Response" to indicate number of units of the affected product and that they followed the instructions for including the attached supplemental instructions to the operating manual. For questions about the recall, contact your local Maquet, Inc. representative. |  | 3688 | 3,688 units in US | Nationwide | Terminated | July 08, 2009 | 201 | 3688 | Not_Computer
Z-0945-2009 | 50755 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX Cardiovascular Polling Procedure version 1.00.09, used in IMPAX 6.2.1 and 6.3.1.; Catalog numbers: E18R4000, L5NQJ000, L57OP000, LBCDB000, and E2ZCK000. |  |  | December 24, 2008 | 2008 | AGFA Corp. | Polling procedure may fail when multiple merge/split is carried out. | Software design | All sites potentially affected by this product were notified of the issue directly via telephone. Each site was directed to observe a ten minute polling window as an interim mitigation until the normal mitigation by software Service Update is performed at the site by Agfa Support. An Urgent Safety Notice was sent via FED-EX on/about 12/24/2008 to all sites. The letter describes the potential issue and both interim and formal mitigation. Acknowledgment, via FAX-Back, that the information was received and understood was requested from the sites. |  | 5 | 5 units | Nationwide Distribution --- including states of FL, IL, OR, SC, and SD. | Terminated | April 01, 2010 | 463 | 5 | Software |  | Software update | Software Update
Z-0970-2009 | 50759 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | November 20, 2008 | 2008 | GE Healthcare Integrated IT Solutions | Unread exams may inadvertently get marked as dictated in PACS workstation when the user signs off the report in Talkstation and opens the next exam. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation with Talk Extend software customers via letter (Urgent Medical Device Correction) dated 11/10/08 of the software anomaly resulting in a patient safety issue involving Talk Extend, which could contribute to a delay of care or misdiagnosis due to delay in communicating pertinent exam results or failure to treat due to lack of an attached dictated report by the radiologist for physician review. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479. |  | 22 | 22 units | California, Florida, Illinois, Iowa, Louisiana, Oregon, Pennsylvania, Texas, Utah, Virginia and Wisconsin. | Terminated | February 03, 2010 | 440 | 22 | Software | Output/Calculation | Software update | Software Update
Z-0624-2009 | 45967 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Medtronic O-Arm Intraoperative Imaging System |  |  | November 27, 2007 | 2007 | Medtronic Navigation, Inc | Navigation accuracy problem: Accuracy may fall outside of specification when used in combination with the Medtronic Synergy Experience StealthStation System Spine Software with the O-Arm Intraoperative Imaging System and tracker. | Software design | Medtronic Navigation notified Field representatives on 11/27/07 to perform site visits to implement FCA 2007-005 - O-Arm Intraoperative Imaging System Accuracy, and deliver the Customer Letter notification. Confirmation of the visit was documented. |  | 54 | 54 units | Worldwide Distribution: USA, Canada, Czech Republic, Denmark, Germany, Netherlands, S. Korea, Poland, South Africa, and Singapore | Terminated | January 22, 2016 | 2978 | 54 | Software | Output/Calculation | Software update | Software Update
Z-0540-2009 | 49114 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | MicroScan Synergies Plus Dried Overnight Gram Negative MIC/Combo Panels, 20 panels per box. Medical device designed for use in determining quantitative and/or qualitative antimicrobial agent susceptibility and/or identification to the species level of colonies, grown on solid media, of rapidly growing aerobic and facultative anaerobic gram-negative bacteria. |  |  | July 31, 2008 | 2008 | Siemens Microscan | Falsely lowered imipenem MICs may result, leading to major errors in results. | Other | Consignees were notified on July 31, 2008 by letter (Urgent Device Correction) sent via FedEx overnight. International consignees were notified by Siemens Healthcare Diagnostics associates in affected countries. The letter described the problem and provided users with safety instructions to follow until an update is provided. Contact Siemens Healthcare Diagnostics Inc. at 916-374-3075 for additional assistance. |  | 569375 | Total of both Negative and Synergy: 569,375 boxes | Worldwide | Terminated | November 29, 2010 | 851 | 569375 | Not_Computer
Z-0650-2009 | 50178 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Medical Stair-PRO Evacuation Chair civilian use stair chair, patient transport device, Model 6253. |  |  | November 18, 2008 | 2008 | Stryker Medical Div. of Stryker Corporation | Handle could come off. The bolts may back out over time, resulting in a risk of dropping the patient and/or injury to transport personnel, a delay in medical treatment, or a delay in evacuation. | Process control | Consignees were notified by Urgent Medical Device Correction letter dated 11/18/08 and given the serial numbers of their units, an upgrade kit and instructions for performing replacement of the bolts. Consignees were informed Stryker will perform the upgrade if requested. |  | 102 | 102 | Worldwide Distribution: USA, India, Netherlands and United Kingdom. | Terminated | October 27, 2010 | 708 | 213 | Not_Computer
Z-0951-2009 | 49474 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ACCELERATOR APS Centrifuge Module (Hettich Robotic centrifuge); List Number: 07L02-01; manufactured by Inpecco SpA, Segrate (MI), Italy. |  |  | September 09, 2008 | 2008 | Abbott Laboratories, Inc | ACCELERATOR APS Centrifuge Module being recalled due to potential for micro-cracks to form over time in the metal centrifuge rotor buckets that may cause the centrifuge buckets to disconnect from the rotor in the centrifuge during centrifugation. | Device Design | Notification letters sent to consignees on 09/09/08. Consignees provided instructions for inspecting the buckets for cracks and were asked to contact customer service reps for replacement buckets. Consignees also asked to respond via fax using Customer Reply form acknowledging receipt, understanding and completion of correction letter. |  | 6 | 6 | Nationwide Distribution --- including states of TX, UT, and VA. | Terminated | October 04, 2010 | 755 | 6 | Not_Computer
Z-0950-2009 | 48172 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Architect / Aeroset Clinical Chemistry Iron/Magnesium Calibrator, List Number: 1E69-03, Lot Numbers: 54187M200 |  |  | April 24, 2008 | 2008 | Abbott Laboratories Inc. | An incorrect Magnesium calibrator 2 value (CAL 2) was provided for Clinical Chemistry Iron/Magnesium Calibrator lot number 54187M200. Use of the incorrect Magnesium CAL 2 value causes quality control and patient magnesium results to be reduced by as much as 17.24% above 0.8mEq/L (1.0 mg/dL or 0.4 mmol/L). | Error in labeling | On April 24, 2008, a Product Correction letter with attached Customer reply form was mailed via US Postal Service to all affected customers that have received the affected lots. Customer reply form were included in all US customers letter. The letter instructed customers to:1. Identify the lot number of Clinical Chemistry Iron/Magnesium Calibrator currently used in your laboratory2. If you have lot number 54187M200, cross out the incorrect Magnesium CAL 2 value in the value sheet, commodity 30-3535/R1, June 2005. Place a copy of this letter in all calibrator kits with this lot number.3. Enter the revised Magnesium CAL 2 value provided in this letter in your system. NOTE: Ensure that the value used corresponds to the units of concentration used in your laboratory. *ARCHITECT cSystems - Refer to Enter a calibrator concentration (cSystem) in Section 2, Installation procedures and special requirements of the ARCHITECT System Operations Manual. *AEROSET - Refer to Calibrator/Control Configuration in Section 2, Installation Procedures and Special Requirements of the AEROSET System Operations Manual.4. Calibrate the Magnesium assay. Verify the calibration with at least two levels of controls according to the established quality control requirement for your laboratory.Customers were also instructed to please retain this communication for future reference and if they have forwarded any Clinical Chemistry Iron/Magnesium Calibrator to another laboratory, please provide a copy of this letter to them. An 800 number was provided in the recall letter for any questions the affected customers or any of the health care providers they serve may have. |  | 1287 | 1287 | Worldwide Distribution --- USA including states of AL, AZ, CA, CT, FL, GA, IA, IL, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, and WI, and countries of Mexico, Chile, Puerto Rico, Canada, Thailand, Curacao, Guatemala, Argentina, Dominican Republic, Germany, New Zealand, Jamaica, Ecuador, Brazil, El Salvador, Hong Kong, and Bermuda. | Terminated | July 09, 2009 | 441 | 1287 | Not_Computer
Z-0971-2009 | 50790 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | 3.7 V Li-Polymer Battery, used in Precess MRI Compatible Patient Monitoring System and Achieva, Intera and Panorama 1.0T. REF 9065 or REF 9067. Invivo Corporation, Orlando, FL 32826 |  |  | August 13, 2008 | 2008 | Invivo Corporation | Battery Failure: The Invivo Lithium Polymer Wireless Module Battery Part Number 9065 and the Lithium Polymer Fiber Optic Module Battery Part Number 9067 is being recalled because the battery, itself, when dropped from 3 feet/1 meter onto a hard surface can produce a flame and then emit smoke. The firms investigation shows that the battery failure does not occur when a module is dropped with a ba | Process design | Urgent Medical Device Recall Notice letters, dated August 2008, and response forms were mailed to each consignee. Distributors were instructed to forward copies of the recall letter and response forms to their customers. The letters identified the affected products, gave the reason for the recall, the hazard involved, how to identify the product, actions to prevent the failure, and procedures to mitigate risk if failure occurs. The letter stated that Invivo would provide replacement batteries but until replacements were provided, customers should handle batteries over an elevated area and to make sure batteries are handled in the absence of paper or other flammable materials. If customers wish to remove the battery from the module in order to conserve battery life, they are to disengage the battery from the module and place the battery on a table or in the battery charger. Do not allow the battery to be disengaged and remain partially inserted in the module as this may cause the battery to slip from the module and fall to the floor. |  | 8476 | 8,476 | Worldwide Distribution | Terminated | November 19, 2010 | 828 | 8476 | Not_Computer
Z-0952-2009 | 50209 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | EPOC BGE Test Card |  |  | October 16, 2008 | 2008 | Epocal | elevated range of potassium levels | Other | The recalling firm visited the hospital on 10/18/08 to conduct an investigation. At the end of the visit, the firm representatives removed the remaining cards from the hospital and returned to the manufacturing location. |  | 1000 | 1000 test cards | Nationwide Distribution --- including state of CA. | Terminated | February 06, 2009 | 113 | 1000 | Not_Computer
Z-0973-2009 | 50848 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Sterile Field Saline Flush; |  |  | January 13, 2009 | 2009 | Excelsior Medical Corp | Sterility Issue: There is the possibility that the affected lots have pouches that were not sealed properly and could potentially compromise the external sterility of the syringes. | Packaging process control | An Urgent Medical Device Recall letter (dated January 13, 2009) was sent to consignees via certified mail, with return receipt requested. The letter provided the reason for the recall and instructions to discontinue use and return (or destroy) the recalled product. The letter also instructed customers to complete the enclosed Customer Reply Form and fax (866-360-6211) to Excelsior Medical Corporation. If the affected product is destroyed on site, fax a copy of the certificate of destruction with the completed Customer Reply Form. Arrangements will be made by Excelsior Medical Corporation for the return of recalled product. Please direct questions to Excelsior Medical Customer Service at 800-487-4276. |  | 341580 | 341,580 units | Nationwide Distribution. | Terminated | August 06, 2009 | 205 | 341580 | Not_Computer
Z-0635-2009 | 50227 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Terumo TenderFlow Pediatric Arterial Cannulae, 3.3mm (10 Fr) OD, with 1/4 inch non-vented connector, 9 inch (22.9 cm) long; Catalog no. 813572. |  |  | October 16, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The introducer or stylet is difficult to remove from the pediatric arterial cannula body, which may result in the need to remove the cannula and replace it. | Device Design | An Urgent Medical Device Recall letter dated 10/16/08 was sent to consignees advising them of the problem; requesting that they immediately verify whether or not they had unused cannulae with the affected lot numbers; requesting that they isolate any affected product to prevent use; complete the Customer Response Form, indicate on form if it is medically necessary to use product without introducer; fax Customer Response Form to Terumo CVS at the number listed; and Terumo CVS would contact customers to arrange for the return of unused product. The letter stated that credit would be issued for any unused TenderFlow cannulae that were returned. The quantities to be returned should be indicated on the Customer Response Form. Consignees were advised that, if it was medically necessary to use the affected cannulae, they should be used without the introducer and in accordance with instructions for such use. |  | 288 | 288 (170 plus 78 samples to US and 40 international) | Worldwide Distribution --- including USA and countries of Australia, Belgium and Canada. | Terminated | August 04, 2009 | 292 | 2892 | Not_Computer
Z-0974-2009 | 50887 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Zimmer Dental Calcitite 4060-2 Dense Nonresorbable Hydroxylapatite Bone Grafting Material, 4 ampules, 0.5g each, 40-60 mesh size, sterile. Zimmer Dental, Carlsbad, CA; Item # 0003. |  |  | January 22, 2009 | 2009 | Zimmer Inc. | Fragments of the glass vial may be present in some of the packages from these lots. | Packaging process control | The consignee in Japan was notified by e-mail and "Urgent: Device Recall" letter dated January 22, 2009 to stop using the product and to return it to the firm. |  | 225 | 225 kits | International Distribution including the country of Japan. | Terminated | October 23, 2009 | 274 | 225 | Not_Computer
Z-0972-2009 | 50825 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Endoscopy brand 3.5mm Super 90-S SERFAS Energy Suction Probe; |  |  | June 20, 2008 | 2008 | Stryker Endoscopy | The probes may have inadequate suction due to obstruction, which can result in unintended ablation due to reduced visibility within the surgical area. | Process control | On 6/20/08, the firm initiated the recall and its notification was via Urgent: Device Removal letters explaining the reason for the recall. The letter recommended that recipients of the letter inspect their inventory for the part number/lot numbers listed and quarantine any of the affected product. Affected products should be returned to Stryker Endoscopy using the Pre-paid shipping label provided. Customers should call Stryker Endoscopy's Customer Service Department at 1-800-624-4422 for replacement product. Importantly, customers should sign and complete the enclosed "acknowledgment of receipt" card and return it to Stryker Endoscopy or fax it to 408-754-2521 with the customer's account info included. If there are any questions regarding the notice, a Regulatory Affairs Representative should be contacted at 408-754-2161 or 408-754-2000. |  | 2463 | 2463 units | Worldwide Distribution. | Terminated | June 22, 2009 | 367 | 2463 | Not_Computer
Z-0975-2009 | 50161 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Surgical Membrane Oxygenator. |  |  | October 17, 2008 | 2008 | Medtronic Inc | Medtronic has identified an issue with the labeling of the I-4500 Silicone Oxygenators (P/N 61399402655). Lot numbers of this oxygenator manufactured between August 2007 and September 2008 have an incorrect "Recommended maximum blood flow range" on the label. The label on the device states a range of 5.0 - 8.5 L/min but should state a range of 5.0 - 6.5 L/min. These oxygenators are acceptable t | Error in labeling | A consignee letter " Urgent Medical Deive Notice" dated October 17, 2008 was sent to the impacted customers on October 17, 2008. The letter described the issue and product and confirmed the correct maximum internal flow of 6.5 L/min for the I-4500 Silicone Oxygenators. The letter states the product will not be removed from the field. Customers were instructed to complete the" Medtronic Product Notification Certificate (FCA 0903)" acknowledging receipt of the information and fax the completed certificate to 763-391-9603. Please direct questions about the recall to your Medtronic Sales Representative or call 763-391-9124. |  | 408 | 408 | Worldwide Distribution USA including states of CA, CO, FL, IL, IN, KY, LA, MD, MI, MN, MO, NV, NJ, NY, OH, OK PA, TN, TX, UT, VA, and WI, and countries of LATVIA, ITALY, ARMENIA, RUSSIAN FEDERATION, POLAND, CHILE and CHINA. | Terminated | April 19, 2011 | 914 | 408 | Not_Computer
Z-0980-2009 | 50223 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Medical Universa Soft Ureteral Stent and Positioner with Hydrophilic Coating, 5.0Fr/26cm, sterile, Cook Urological Inc., Spencer, IN; REF USH-526-R, GPN REF G49966. The product was placed in a shipping box which is preprinted with the word "firm" instead of "soft", but a product specific label stating "soft", as above, was also placed on the box. |  |  | November 10, 2008 | 2008 | Vance Products Incorporated | "Soft" stents were placed in "firm" stent outer boxes. The soft stent is only to be left in the patient for 6 months, but the firm stent can be left in the patient for 12 months, so this error could result in the stent being left in the patient for longer than indicated for the device. | Labeling mix-ups | Consignees were notified via recall letter, "Urgent: Recall Notice" dated November 10, 2008, and by follow-up phone call. Customers were instructed to isolate the recalled product, complete the enclosed form and fax to Cook Urological at 812-829-2353. After receipt of the completed form, Cook Urological will issue a Return Goods Authorization number for each customer to return the affected product to Cook Urological via FedEx or UPS. |  | 23 | 23 | Nationwide Distribution including states of Georgia, Oregon and Texas. | Terminated | November 25, 2009 | 380 | 23 | Not_Computer
Z-0863-2009 | 50522 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DeVilbiss Model 7310PR-D VacuAide Compact Suction Unit. A portable, AC-powered or compressed air-powered suction pump. The device may include a microbial filter. |  |  | October 27, 2008 | 2008 | Sunrise Medical | Improper suction: Regulator knob does not turn properly which can affect the suction level. | Device Design | DeVilbiss notified dealers and users (through the dealers) to recall the units, by a telephone call October 27, 2008 followed by an Urgent Medical Device Recall mailing November 3, 2008. The letter advised customers what models are affected and how to contact the firm of dealer to have the affected unit replaced. Corrective and preventative action on current manufacturing has been implemented. |  | 78 | 78 units | Worldwide Distribution: USA, Mexico, Guatemala, Colombia, and Costa Rica. | Terminated | September 25, 2012 | 1429 | 78 | Not_Computer
Z-0915-2009 | 48598 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Tibial Array and Femoral Array. |  |  | May 02, 2008 | 2008 | Mako Surgical Corporation | The product in its current state may fail during use, which would cause the instrument to be in-accurate. | Component design/selection | A MAKO employee will visually verify product at each account for any product affected by this recall. In addition, a recall letter dated July 1, 2008 was mailed to users (surgeon and owner of system, e.g. hospital). Customers were instructed to complete the enclosed Acknowledgement form and return it to your Clinical Technical Specialist (CTS) or in the enclosed pre-paid addressed envelope to the Mako Surgical Corporation. Please direct questions to your CTS or call 954-927-2044, extension 604. |  | 44 | 44 | Nationwide Distribution USA including states of FL, NY, CA, TX, GA, TN, NC and PA. | Terminated | August 21, 2009 | 476 | 44 | Other |  |  | Other
Z-0976-2009 | 50548 | [6 empty] | Turbo Elite 0.9mm Turbo Elite (TE) RX, Model #410-154, tip diameter 0.9mm/.038 in, 4 Fr, sterile EO. Manufactured by Spectranetics, 96 Talamine Court, Colorado Springs, CO 80907. |  |  | July 04, 2008 | 2008 | Spectranetics Corporation | The catheter pouch label was incorrectly labeled with the wrong size. | Error in labeling | The mislabeled catheter was returned to Spectranetics. The second possible mislabeled catheter was visually inspected by a Spectranetics representative and confirmed the catheter was labeled correctly. |  | 2 | 2 units | Nationwide Distribution including states of TX and FL. | Terminated | February 13, 2009 | 224 | 2 | Not_Computer
Z-0982-2009 | 50700 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | DePuy DNP Anatomic Fracture Repair System, JIG DNP Anatomic, Depuy, Warsaw, IN; REF DNPAJIG. |  |  | December 23, 2008 | 2008 | Depuy Orthopaedics, Inc. | The DNP nail plate may not attach to the alignment guide (jig) instrument, making it difficult to insert the nail plate into position or to drill needed holes accurately. | Device Design | The firm's sales force was instructed via "Urgent Device Correction" letter issued December 29, 2008 to inspect all of these instruments in their possession, and those instruments at their hospital customers, and to leave a copy of the letter with each hospital customer who has the instruments on site.The letter describes the issue and corrective actions for customers. Please direct questions about this recall to Depuy Orthopaedics, Inc. at 888-783-7156. |  | 3250 | Approx. 3,250 of all products | Nationwide Distribution. | Terminated | October 18, 2010 | 664 | 3250 | Not_Computer
Z-0979-2009 | 50697 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gold Probe 7FR Working Length: 350 cm, Probe Tip O.D. 2.3 mm (7Fr), Rx Only. Electrical Connector: Standard. Single Use Only. Legal Manufacturer, Boston Scientific Corporation, One Boston Scientific Place Natick, MA 01760-1537 USA. Made in USA: 8600 N.W. 41st Street Miami, FL 33166-6202 USA. |  |  | December 24, 2008 | 2008 | Boston Scientific Corporation | Packaging tray for the Gold Probe Single-Use BiPolar Electrohemostasis Catheters may have a crack. Cracks in the tray or damage to the tray of the Gold Probe device may lead to a compromised sterile barrier. | Packaging process control | A notification letter (dated December 24, 2008) was issued to customers. The letter identified the issue. Consignees are instructed to: 1) Discontinue use and segregate affected product, 2) Complete a Reply Verification Tracking Form and fax to the Boston Scientific Corporation (BSC) at 508-683-5578 and 3) Package and ship recalled product as instructed in the letter. If affected product is found, customers are instructed to call BSC Customer Service at 1-866-868-4004 between 8:30 a.m. through 6:00 p.m. (Eastern Standard Time).Please direct questions about the recall to your local Sales Representative. |  | 10841 | 10,841 | Worldwide Distribution. | Terminated | May 11, 2010 | 503 | 109610 | Not_Computer
Z-0862-2009 | 50499 | No | GES | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stainless Steel Surgical Blades Part no 4-312B. The product is supplied sterile individually packaged in boxes containing 100 blades. |  |  | November 17, 2008 | 2008 | Miltex, Inc. | Blade may break: Over milling of the blade resulted in tip more narrow than specified. | Process control | An Urgent Medical Device Recall letter dated 11/17/08 with an accompanying fax back form was sent to each customer. Each distributor was instructed to review their inventory, complete the form, and return any affected product. The distributors were also instructed to conduct a sub recall notification. |  | 200 | 200 boxes | Nationwide Distribution and Canada. | Terminated | February 23, 2009 | 98 | 200 | Not_Computer
Z-0861-2009 | 50496 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Eclipse treatment planning system, Model Number H48, versions 6.5/7.x and Inspiration 8.0. |  |  | November 21, 2008 | 2008 | Varian Medical Systems Oncology Systems | Possible incorrect wedge orientation in treatment: Software labeling was issued without warning message relating to preservation of user selected values for Siemens wedge orientation, wedge limits and SAD values for Elekta machines, and collimator operating limits for Saturne Machines when the software version is upgraded. | Software design | An Urgent Device Correction product notification letter/field safetly letter dated November 21, 2008, was distributed to all affected users, with a description of the problem and user corrective action steps. The letter was also distributed to Varian Sales, Marketing and Service organizations, informing them of the issue. |  | 516 | 516 units | Worldwide Distribution: USA, Austria, Belgium, Canada, Czech Republic, Finland, France, Germany, Greece, Hungary, Italy, Ireland, Lithuania, Norway, Portugal, Romania, Slovakia, Spain, Sweden, Switzerland, Netherlands, Turkey and the United Kingdom. | Terminated | March 10, 2011 | 839 | 516 | Software | Treatment/Delivery/Therapy | Instructions for correction | Repair
Z-0910-2009 | 50554 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1; 220V -240V, 7A (circuit breaker), 50/60 Hz (10A power source required) base, Terumo Cardiovascular Systems, Ann Arbor, MI. Model 801764. (Not distributed within the United States). |  |  | November 07, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | The power manager board may cause a short circuit, resulting in loss of system power, loss of central control monitor or electrical arcing. | Nonconforming Material/Component | An "Urgent Medical Device Correction" letter was sent to consignees on November 7, 2008. The letter described the reason for correction and provided instructions to customers. Terumo Cardiovascular Systems Corporation will replace the power manager boards on the affected devices. Customers were instructed to complete and fax the Customer Response Form to CVS Customer Service at 1-800-292-6551 (US only) or respond by telephone at 1-800-521-2818. Please direct questions to CVS Customers Service at 1-800-521-2818. |  | 515 | 515 | Worldwide Distribution including USA and countries of Australia, Belgium, Canada, Columbia, Costa Rica, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates and Viet Nam. | Terminated | July 07, 2009 | 242 | 1080 | Not_Computer
Z-0940-2009 | 49610 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker InTouch 1.0 Critical Care Bed; Model 2130. |  |  | January 05, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The touch screen on the footboard may display a "tilt over range" error, resulting in a precautionary stop of the bed motions which result in the bed appearing to have locked out some of its bed positioning features. | Device Design | Consignees were notified of the problems via letter dated January 5, 2009, advised to check their beds and to take them out of service if the foot board touchscreen displayed a "tilt over range" error, and that a Stryker Medical field service representative would contact them to perform software upgrades on the affected beds. Please direct questions to Stryker Medical at 1-269-324-6644. |  | 1288 | 1,288 of all models | Worldwide Distribution including USA and countries of Canada, Japan and Switzerland. | Terminated | November 05, 2010 | 669 | 1288 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0919-2009 | 50064 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Cardiac Science Powerheart Automated External Defibrillator G3 9390A |  |  | October 13, 2008 | 2008 | Cardiac Science Corporation | Potential for AED to not deliver defibrillation therapy. | Software design | On October 13, 2008 Cardiac Science began sending via Fed-Ex letters and software upgrade kits. Mailing of letters and software upgrade kits is anticipated to be completed by November 21, 2008. The letter instructs consignees to install the software update, assure that the AED information test file generated by the update software is returned to the firm, and complete the postage paid response card and return it to the firm. If the affected devices are no longer in the possession of the consignee, the consignee is to contact the firm and forward the update kit and letter to the person who has the AED in their possession. A separate letter is being sent to distributors. The letter advised to install the software upgrades for any units in stock or to forward the software upgrade along with the end user letter enclosed and that the AED information text file generated by the update software must be returned to the firm. The firm will visit the locations of individual national account customers and install the update. OEM partners will be notified and provided the information and update kits to provide to their customers. |  | 2572 | 2,572 devices | Worldwide Distribution: USA, Puerto Rico, Virgin Islands, Australia, Belgium, Bermuda, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czechoslovakia, Denmark, Dominican Republic, France, Germany, Greece, Guam, Hong Kong, Hungary, Ireland, Israel, Italy, Jamaica, Japan, Kenya, Kuwait, Libya, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, UK, and Ukraine. | Terminated | December 08, 2011 | 1151 | 24183 | Computer
Z-0944-2009 | 50541 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Symbiq One-Channel list number 16026-04 and Two-Channel list number 16027-04 Infusion System, a general purpose infuser, Distributed and Manufactured by Hospira Inc, Morgan Hill, CA. |  |  | December 03, 2008 | 2008 | Hospira Inc | Hypersensitivity to normal motor friction/drag may cause malfunction errors and alarms that are designed to stop pump operations. | Software design | Recall initiated 12/03/2008. Consignees were notified of the issue via an Urgent: Device Field Correction letter, dated 12/3/2008, and were requested to respond appropriately to any alarm or warnings displayed by the device. The firm is making arrangements to screen each Symbiq pump in customer facilities and implement the appropriate enhancements. |  | 11840 | combined 11,840 units | Worldwide Distribution --- including USA and countries of Canada and Australia. | Terminated | April 03, 2012 | 1217 | 11840 | Not_Computer
Z-0990-2009 | 50552 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Minimally Invasive Solutions Technology Offset Rasp Handle, Left, 30 degrees, Zimmer, Warsaw, IN; REF 00-7806-035-01. |  |  | December 19, 2008 | 2008 | Zimmer Inc. | The c-clip located near the connection end may fracture or fall off during use in surgical procedures. | Labeling design | U.S. consignees were notified by "Urgent: Device Recall" letter issued via e-mail on December 19, 2008. The letter described the problem and instructed customers to immediately quarantine and return the recalled product along with the completed Inventory Return Certification form to the Zimmer, Inc. Distribution Center. Consignees were also instructed to fax a copy of the completed Inventory Return Certification Form at 1-574-371-8603. International consignees were also notified, but Zimmer, Inc. will gather the instruments from those customers. Direct questions about the recall to Zimmer, Inc. at 1-800-613-6131. |  | 1461 | 1,461 of both products. | Worldwide Distribution --- including USA, Belgium, Czech Republic, Finland, Germany, Greece, Italy, Poland, Romania, Spain and United Kingdom. | Terminated | December 22, 2009 | 368 | 1461 | Not_Computer
Z-0627-2009 | 50139 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Excelsior SL-10 Microcatheter, placed in sterile pouch. Manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | October 07, 2008 | 2008 | Boston Scientific Corporation | Packaging issue may result in failure to maintain sterile barrier. | Package design/selection | Consignees were notified by letter dated October 7, 2008. The letter identified the issue and affected product. Customers were instructed to: 1.) Discontinue use of and segregate recalled product, 2.) Complete and return the enclosed Reply Verification Tracking Form, and 3.) Package/ship the recalled product and send to the Boston Scientific Corporation (BSC).Return the completed Reply Verification Tracking Form via fax (800-836-4804) to the BSC Customer Service Call Center.A follow-up letter dated October 8, 2008 was also issued to customers. The letter notified customers of the correct date range of expiration for product within the scope of the recall which is September 2008 through October 2010. There were no other changes identified in the follow-up letter. |  | 21423 | 21,423 units | Worldwide Distribution. | Terminated | March 02, 2010 | 511 | 21423 | Not_Computer
Z-0985-2009 | 50746 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Orthosoft Navitrack, OS Unicondylar Knee 1.0 - Universal; software application for orthopedic surgical stereotaxic instrument. |  |  | August 19, 2008 | 2008 | Zimmer Inc. | Incorrect parameters: If the user switches from one surgeon profile to another surgeon profile after completing anatomical landmark acquisition tasks, the application will continue to use the initial surgeon profile parameter values in the calculation of the predicted hip-knee-ankle alignment estimates, which may have different resection or component thickness parameters than are appropriate for | Software design | Zimmer notified consignees via letter dated 8/19/08, which advised them not to change the selected surgeon profile after the acquisition of landmarks has begun, and the issued will be fixed with the next software release. |  | 75 | 75 (34 U.S. and 41 international) | Worldwide distribution: USA, Australia, Austria, Belgium, Canada, France, Germany, Italy, Singapore, South Africa, Switzerland and United Kingdom. | Terminated | April 28, 2010 | 617 | 75 | Software | Output/Calculation | Software update | Software Update
Z-1038-2009 | 50839 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Triple Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division, Deerfield, IL 60015 U.S.A., Made in Singapore. Product codes: 2M8153, 2M8163, 2M9163, DNM8153, DNM9163, BRM8153, BRM8153T, BRM8153RT, PNM8153, PNM8153T, PNM8153RT, CNM8153, GNM8153, HNM8153, TRM8153 and WNM8153. |  |  | January 23, 2009 | 2009 | Baxter Healthcare Corp. | Baxter has identified failure codes that lead to a delay or an interruption of therapy, hazards associated with improper cleaning, and damaged battery messages for the Colleague Infusion Pumps. | Device Design | Baxter Healthcare Corporation sent an "Urgent Device Correction" letter dated January 23, 2009 to all Colleague Infusion Pump customers, to the attention of the Vice President of Nursing with copies to the Chief Operating Officer, Director of Biomedical Engineering and Director of Risk Management to advise them of failure codes that lead to an interruption of therapy, hazards associated with improper cleaning, and damaged battery messages associated with the Colleague pumps. The letters listed the failure codes involved that lead to an interruption of therapy which could lead to serious injury and/or death; advised institutions to have contingency plans to verify that back-up pumps are available; and provided steps for addressing an interruption of therapy with any failure code. The letters also provided information about the hazards associated with improper cleaning of the pumps and instructions for addressing damaged battery messages. Customers were instructed to complete the attached Customer Reply Form acknowledging receipt of the letter and fax it to 888-569-1585. Direct questions to your local Baxter Representative or to Baxter Global Technical Services by calling 1-800-843-7867, Prompt 2, between 7AM and 5 PM Central Standard Time (CST). |  | 81775 | 81,775 pumps | Worldwide including USA, Puerto Rico, Aruba, Australia, Austria, Bahamas, Bahrain, Belgium, Bermuda, Brazil, Canada, Cayman, Chile, China, Colombia, Curacao, Denmark, Dominican Republic, El Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Honduras, India, Ireland, Israel, Italy, Jamaica, Korea, Kuwait, Lebanon, Macedonia, Malaysia, Malta, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Peru, Philippines, Qatar, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | October 04, 2011 | 984 | 302872 | Battery | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-0955-2009 | 50796 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Meter Remote for OneTouch Ping Glucose Management System |  |  | October 27, 2008 | 2008 | Animas Corporation | incorrect dosing recommendation | Software design | The recalling firm issued Urgent: Medical Device Correction letters dated 10/27/08 to their customers/user to inform them of the problem. The letter stated that the firm was in the process of correcting the issue and that a new OneTouch Ping Meter-Remote would be provided as soon as one became available. In the interim, customers could still safely use their OneTouch Ping System provided that they use the ezCarb calculator on the meter-remote when it is paired and communicating with the pump. The letter listed guidelines to ensure that the device is properly communicating with the pump. The letter stated that customers should follow the directions in the Owner's Booklet to troubleshoot any communication issues. Questions regarding this letter were to be directed to Animas at 866-423-4087. A patient acknowledgment letter dated November 18, 2008 was also issued. Another Urgent: Medical Device Correction letter, dated November 2008, was issued along with the replacement device. This letter asked customers to download their current meter-remote data, set up their replacement meter-remote, and then return their current meter-remote. |  | 5093 | 5093 units | Nationwide Distribution --- including states of KY, UT, FL, OH, MA, and TX. | Terminated | December 15, 2009 | 414 | 5093 | Not_Computer
Z-0947-2009 | 50736 | No | ETN | Surgical nerve stimulator/locator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | VARI-STIM III Nerve Locator/Stimulator, Sterile, Medtronic Xomed, Inc. |  |  | November 21, 2008 | 2008 | Medtronic Xomed, Inc. | Device may fail to operate or operate intermittently during use and such operation could result in a lack of nerve stimulation when nerve tissue is actually contacted by the stimulator. | Employee error | All consignees were contacted by telephone (initiated on November 10, 2008). After telephone contact, each consignee was sent a follow-up fax/letter, "Urgent Product Recall Notification". The letter described the issue and actions for customers to take in response to the recall. Customers were instructed to check their inventories and return affected product to Medtronic Xomed, Inc. Customers were also instructed to fax their completed "Customer Contact and Product Return Checklist" to Medtronic Xomed, Inc. at 1-904-296-2386. Direct questions to Medtronic ENT Customer Service at 1-800-874-5797 (select Option 1 for Customer Service) and ask for the Vari-Stim Field Action Contact Person. |  | 22 | 22 boxes (220 units) | Worldwide Distribution --- including USA and country of Canada. (22 boxes were distributed within U.S. Three boxes were distributed to Medtronic Canada but returned) | Terminated | October 01, 2009 | 314 | 22 | Not_Computer
Z-1013-2009 | 50738 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon X3 Posterior Stabilized Tibial Insert |  |  | December 18, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware of a lot for lot mix-up between part numbers 5532-G-713, lot code LBD528 (Triathlon X3 Posterior Stabilized Tibial Insert, size 7-13 mm) and part 5532-G-813, lot code LBD512 (Triathlon X3 Posterior Stabilized Tibial Insert Size 8-13 mm) | Labeling mix-ups | Urgent Product Recall letters were sent via Federal Express on December 18, 2008 with return receipt to Stryker branches, Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. The letter stated the issue, the potential hazards, and the risk mitigation factors. The firm asked customers to examine their inventory and hospital locations to identify affected product; retrieve all affected product and return it to their branch or agency warehouse for reconciliation; reconcile all products on the attached Product Accountability Form; and return affected product. Questions should be directed to Karen Ariemma at 201-831-5718. |  | 126 | 9 of lot code LBD528; 14 of lot code LBD512 | Nationwide Distribution: FL and IL. | Terminated | September 11, 2012 | 1363 | 126 | Not_Computer
Z-0992-2009 | 50132 | No | MTJ |  |  | General &amp; Plastic Surgery | 510(k) | Bone Wax. |  |  | October 28, 2008 | 2008 | Ethicon, Inc. | A hospital reported to their Ethicon sales representative that they had an "unmarked" box of bone was on their shipping dock that they could not receive due to lack of appropriate documentation. Upon investigation, this box was determined to be missing from a shipment between two steps of the manufacturing process. Further checks determined that two other boxes were missing. | Process design | Letters were sent via UPS next day air on October 28, 2008 to all customers who purchased finished product codes W31G and DW31B (on inner box labels) from January 1, 2008 through October 22, 2008. |  | 2200 | 2,200 foil pouches | Nationwide Distribution. | Terminated | April 29, 2009 | 183 | 2200 | Not_Computer
Z-0997-2009 | 51160 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Shiley PED3.0 Cuffless Pediatric Tracheostomy Tube, REF 3.0 PED, Sterile. Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Assembled in Mexico. The product is provided as one each shipped in an inner carton to end customers. |  |  | January 14, 2009 | 2009 | Covidien Limited | Difficulty inserting an obturator and/or suction catheter into the tracheostomy tube. | Finished device change control | Consignees were sent an "Urgent Device Alert" letter dated 01/14/2009, by mail via Federal Express overnight delivery and instructed to evaluate any of the affected lot numbers of tracheostomy tubes prior to placement into a patient's stoma. If either the obturator or the desired sterile suction catheter does not fit properly into the tube, the tube must not be inserted into the patient. Further instructions are available at Nellcor Technical Services by calling 1-800-635-5267, option 3, then option 1. New product was expected by mid February, 2009, to address the issue. |  | 8931 | 8,931 units | Nationwide Distribution, including Puerto Rico. | Terminated | May 21, 2012 | 1223 | 8931 | Not_Computer
Z-1031-2009 | 50176 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Silhouette Axial Torque Handle, Zimmer Spine, Inc., Minneapolis, MN; REF 7010-0151-01. |  |  | November 07, 2008 | 2008 | Zimmer Inc. | Some of the torque wrenches apply less than the minimum torque to the screw locking cap when the torque indicators are aligned as described in the surgical technique. | Process control | The firm's sales force was notified by an Urgent: Device Recall letter dated 10/7/08, which instructed them to stop using the instrument and quarantine it immediately; physically count all affected product in their territory and record data on the Inventory Return Certification Form; fax a copy of the completed form to Zimmer Spine; and return the affected devices with the completed Inventory Return Certification Form. User accounts/implanting physicians were informed by an Urgent: Device Recall letter dated 12/10/08 and instructed to monitor their patients. |  | 76 | 76 | Worldwide Distribution --- USA including states of California, Minnesota, Ohio, Pennsylvania, and Tennessee, and countries of Denmark, Germany, Iran, Italy, Russia, Singapore, Spain, Switzerland, and United Kingdom. | Terminated | December 22, 2009 | 410 | 76 | Not_Computer
Z-1033-2009 | 49372 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | InteliFUSE, Drill Bit, 2.00 mm, J-Latch, QTY: 1, REF: DB.200.JL, non-sterile, 1515 Poydras st., Ste 1490, New Orleans, LA 70112, MDCI Arundel House, 1 Liverpool Gardens, Worthing, West Sussex BN11 1SL, United Kingdom |  |  | July 30, 2008 | 2008 | Intelifuse Inc | Manufacturing specification for drill bit was incorrect which, when used in conjunction with the drill guide, could cause the production of metal shavings. | Process control | The firm initiated their recall on July 31, 2008 via telephone and Urgent Recall Notice to their current distributor, Integra. The letter asked them to complete, sign, and return the enclosed Recall Acknowledgment and Return Form. The letter also stated that InteliFUSE would travel to their designated Distribution Center(s) to ensure testing and replacement of all defective units once the form has been received. The firm sent follow-up Urgent Recall Notice letters on August 14, 2008 directing the distributor to recall from the hospitals to whom they distributed the product and to two hospitals which were shipped the products directly. The letter asked that use of the affected products be immediately ceased and that the any stock on hand be quarantined. The Recall Acknowledgment and Return Form should be completed, signed, and returned and the products should be returned to InteliFUSE for immediate replacement. |  | 2149 | 2149 units | Nationwide Distribution - NY and NJ. | Terminated | June 22, 2011 | 1057 | 4282 | Not_Computer
Z-1034-2009 | 50899 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Integra Mosaik Putty, 15cc, Catalog # PTY10256 |  |  | January 16, 2009 | 2009 | Integra LifeSciences Corp | Potential for open seals in the inner foil pouch of certain products. The outer seal remains intact and therefore the sterility of the inner pouch is maintained. | Pending | Urgent: Product Recall Notification letters were sent on January 21, 1009 via Federal Express. The letter asked customers to immediately quarantine any stock on hand and contact their local Integra Sales Representative for further assistance and arrangement of product returns. The enclosed "Recall Acknowledgement and Return Form" indicating whether or not customers possess any of the affected product should also be completed and returned. Company is conducting recall to the hospital level. Questions should be directed to Judy O'Grady at 609-936-2317. |  | 88 | 88 units | Nationwide Distribution -- including states of FL, CA, GA, FL, CT, MT, IL, TX, WA, TN, NC, IN, AZ, MD, CO, MO, and TN. | Terminated | June 25, 2009 | 160 | 119 | Not_Computer
Z-0953-2009 | 50511 | No | GNW | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | LIGHT DIAGNOSTICS SimulFluor Flu A/Flu B DFA Kit |  |  | November 24, 2008 | 2008 | Millipore Corporation | Certain lots of Simulfluor Flu A / Flu B Reagent may cross-react with Respiratory Synctial Virus. The continued use of Simulfluor Flu A / Flu B Reagent Catalog #5250 may lead to a potential misdiagnosis of influenza B. | Pending | Millipore is performing a voluntary recall by notifying customers and distributors by telephone on November 24, 2008. In addition, an Urgent: Medical Device Recall notification letter with the same date was sent via certified mail.Customers were advised to examine their stock immediately and discontinue the use of the specified lots. Customers were instructed to complete the enclosed Effective Check Survey confirming disposal of defective lots and return this form to Millipore's Regulatory Department. Customer service is to be contacted at 1-800-437-7500 for replacement product. Millipore Corporation will issue credit or replace the recalled lots that customers currently have in inventory. |  | 404 | 404 units | Worldwide Distribution -- USA including states of AL, CA, CO, CT, FL, ID, IL, KS, MA, MI, MS, NM, NY, OH, OK, OR, PA, SD, TX, UT, VA, VT, WA, WV, and WI and countries of Canada, Mexico, Italy, Belgium,Germany, France, Spain, andUnited Kingdom. | Terminated | August 10, 2012 | 1355 | 662 | Not_Computer
Z-1006-2009 | 50146 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 automated external defibrillator and MRL Jumpstart (collectively "AED 10") |  |  | December 15, 2008 | 2008 | Welch Allyn Protocol, Inc | Reliability issues - potential to shock a non shockable rhythm or not shocking a shockable rhythm; | Component design/selection | Welch Allyn sent a Customer Recall Notice dated February 26, 2009, to all customers (customer list supplied by distributors). Welch Allyn will upgrade all units with software version 2.02 and lower to version 2.06.02 by Proc. AED10 Prod Service Min HW/SW Update Rev. E 820-0178-00. |  | 10735 | 10, 735 | Worldwide Distribution. | Terminated | March 19, 2012 | 1190 | 10735 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1018-2009 | 50711 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical Medfusion Model 3500 Syringe Infusion Pump, Software version V3.0.6. A small, lightweight and portable syringe infusion pump. |  |  | December 18, 2008 | 2008 | Smiths Medical MD, Inc. | Delay in therapy. Keypad issues with Medfusion Syringe Infusion Pump, model numbers 3500-306 and 3500-415 ("Pump") have occurred during implementation or shortly after first being placed into use. These issues can be intermittent, and may render the Pump inoperable or the Pump may not respond to certain keypad button presses. | Nonconforming Material/Component | Consignees were initially notified of the recall via a telephone conversation beginning December 18, 2008 and continued through January 2009. After the initial telephone contact, an "Urgent Medical Device Recall" letter dated December 22, 2008 was sent via electronic mail. The letter was addressed to "Affiliates, Distributors, Clinicians, and other Customers". The letter gave a Description of the Problem and Advise on Action to be Taken by the User. Direct questions to Smiths Medical Customer Service department at 1-800-426-2448 (option 1, then option 3).A Smiths Medical Customer Service Representative will have contacted customers. These pumps will be fitted with new keypads with gold connectors. A second Smiths Medical "Urgent Medical Device Recall Additional Affected Units"recall letter was sent after the recall was expanded. The second letter dated March 6, 2009 was addressed to Safety/Risk Mangers, Healthcare Providers, Biomedical Engineers, Distributors, Clinicians and other Customers. The letter listed the Details of Affected Devices, Description of the Problem, Advise on Action to be Taken by the User, Results of Investigation and Contact Details. |  | 957 | 957 | Worldwide, including USA, AMERICAN SAMOA, AUSTRALIA, CANADA, CHILE, GERMANY, UNITED KINGDOM, HONG KONG, IRELAND, ISRAEL, SOUTH KOREA, NETHERLANDS, SAUDI ARABIA, SWEDEN | Terminated | July 06, 2012 | 1296 | 3509 | Not_Computer
Z-0996-2009 | 51039 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Most Hinged Tibia Baseplate Spacer, size 1, 6 mm, sterile, Zimmer, Warsaw, IN; REF 5020-06-001. |  |  | January 29, 2009 | 2009 | Zimmer Inc. | The package contains 11 mm screws instead of the specified 6 mm screws. | Packaging process control | The consignee in Germany was notified by an Urgent: Device Recall letter dated 1/29/09, and delivered by email on 2/19/09. The letter stated the reason for the recall and clinical implications. The firm asked the consignee to stop using the provisional and quarantine it immediately; physically count all affected product and record data on the Inventory Return Certification Form included with letter; fax a copy of the completed form to Zimmer, Inc at 574-371-8603; request an IRA from Lindsey Arnett; and return recalled product along with completed Inventory Return Certification Form and IRA to Simmer Distribution Center. The letter states that credit will be issued upon receipt of the returned product. Also, if product has been further distributed, a copy of the recall letter should be provided to the customers. Questions or assistance should be directed to Zimmer, Inc at 1-800-613-6131. |  | 3 | 3 | International Distribution only -- Country of Germany. | Terminated | November 12, 2009 | 287 | 3 | Not_Computer
Z-1005-2009 | 50142 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 Automated External Defibrillator |  |  | December 15, 2008 | 2008 | Welch Allyn Protocol, Inc | Potential for device to shut down prematurely under certain circumstances. | Device Design | On 4/13/07 the firm released Service Bulletin part no. 810-2614-00, SB - AED 10 S/W V020600 DOWNGRADE to V020500. That mandatory bulletin, for units with software version 2.06.00 and up indicates that all returning AED 10 units shall have their software version downgraded to 2.05.00. |  | 89 | 89 units | Worldwide distribution: USA, Belgium, Canada, France, Ireland, Israel, Mexico, and United Kingdom. | Terminated | February 15, 2012 | 1157 | 89 | Software | Device Operation | Software update | Software Update
Z-1022-2009 | 50084 | No | KIJ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Stryker Rotary Handpiece, Single Trigger, For use with Stryker System 6. Stryker Instruments, Kalamazoo, MI; REF 6203. |  |  | October 10, 2008 | 2008 | Stryker Instruments Div. of Stryker Corporation | The handpieces may continue to run even when the trigger is not depressed, which could result in the unintended cutting of bone or soft tissue damage. | Component design/selection | Consignees were notified of the problem via letter dated October 10, 2008, provided interim instructions to reduce the risk, and informed that Stryker instruments will replace the trigger on all handpieces. Consignees were also instructed on the following: 1) Review the letter and distribute to all affected departments, 2) Complete and sign the enclosed Business Reply Form and fax to 336-499-8603 and 3) Call Stryker Instruments Service Department at 888-308-1983 to schedule service for the handpieces. |  | 25312 | 25,312 of all units | Worldwide Distribution including USA and countries of Australia, Brazil, Canada, Chile, Colombia, Costa Rica, England, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Malaysia, Netherlands, New Zealand, Pakistan, Panama, Peru, Philippines, Poland, Romania, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan and Turkey. | Terminated | October 27, 2010 | 747 | 25312 | Not_Computer
Z-1007-2009 | 50532 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 automated external defibrillator and MRL Jumpstart (collectively "AED 10") |  |  | December 15, 2008 | 2008 | Welch Allyn Protocol, Inc | Potential for device to deliver less than the programmed energy. | Process control | Welch Allyn notified distributors by letter on 12/15/2008 requesting that they notify their customers. Welch Allyn sent an Urgent: Medical Device Recall letter to distributors and customers on February 26, 2009. |  | 14012 | 14,012 devices | Worldwide distribution. | Terminated | March 19, 2012 | 1190 | 14012 | Other | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-1060-2009 | 50856 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PROSTIVA RF Model 8929 Hand Piece is the delivery system component of the PROSTIVA RF Therapy System. |  |  | June 05, 2008 | 2008 | Medtronic Neuromodulation | Reported incidents of needle retraction failure when using the PROSTIVA RF Model 8929 Hand Piece. The Medtronic Prostiva RF Model 8929 Hand Piece is the delivery system component of the PROSTIVA RF Therapy System. | Device Design | A Medtronic "Urgent Medical Device Correction" letter addressed to "Healthcare Provider" was sent August 6, 2008. The letter described the issue, product involved and the Prostiva hand piece Disassembly Procedure that should be followed in the event the user is unable to retract the needles during a procedure. The letter also requested the return of the "Medtronic Reply Form/receipt Confirmation" form. |  | 40168 | 40,168 | Worldwide Distribution -- AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NV, NJ, NM, NY, NC, ND, OH, OK, OR, PA, PR, RI, SC, SD, IN, TX, UT, VA, VT, WA, WI, WY, and WV and countries of Australia, New Zealand, China, Austria, Denmark, Finland, France, Germany, Italy, Kuwait, Lebanon, Norway, Qatar, Russian Federation, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | May 17, 2012 | 1442 | 40168 | Not_Computer
Z-1059-2009 | 51176 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AcuSnare Polypectomy Soft Snare, Order number SASM-1, Wilson Cook Medical, GI Endoscopy, 4900 Bethania Station Road, Winston-Salem, NC 27105, packed 5 boxes, 10 units per box. |  |  | January 21, 2009 | 2009 | Cook Endoscopy | The snare head is hexagonal but the handle and packaging indicate an oval reorder number. | Process control | Consignees were notified via USPC, Certified Mail and email on/about 1/21/2009 with an Urgent - Product Recall Notice. The distributor (Cook Australia) was instructed to return quarantined product to Cook Endoscopy. All medical users were instructed to review the attached list and quarantine any affected product remaining at their facility. They were to complete the form at the bottom of the page and send/fax it. Users are to return any affected product to Cook Endoscopy. All devices returned as part of this recall initiative will be dispositioned to scrap and destroyed. |  | 10 | 10 boxes (100 snares) | Worldwide Distribution -- USA including states of MI, NY, and WI and country of Australia. | Terminated | April 06, 2010 | 440 | 10 | Not_Computer
Z-1010-2009 | 50861 | No | LQR | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Stonetome Stone Removal Device. Triple Lumen Sphincterotome, 5.5 Fr, 20 mm, 11.5 mm, sterile, Boston Scientific Corp., Natick, MA; REF 3511. |  |  | January 08, 2009 | 2009 | Boston Scientific Corp | The ultraviolet light barrier on the pouch may be discolored, which will allow ultraviolet light into the package and may cause the latex retrieval balloon to become brittle and to leak or burst when inflated. | Packaging | Consignees were notified via letter dated January 8, 2009 to discontinue use of the affected product and to return it to Boston Scientific. Customers were instructed to complete and fax the Reply Verification Tracking Form to Boston Scientific at 508-683-5578. Direct questions about this recall to your local Sales Representative. |  | 294 | 294 of all products. | Nationwide. | Terminated | October 22, 2009 | 287 | 294 | Not_Computer
Z-1091-2009 | 50131 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Glucose Oxidase Reagent as follows: Pointe Scientific Glucose Oxidase Reagent Set, Pointe Scientific, Inc, Canton, MI; Catalog # G7519-1L, G7519-500 and G7519-SAM. 2MM Diagnostics Glucose Oxidase Reagent Set, 2MM Corporation, Miami, FL; Reorder # RT-160-500. Pointe Scientific Glucose Oxidase Reagent, Pointe Scientific, Inc., Canton, MI; Catalog, MI; Catalog # 2-G7519-L, 7-G7519-100, 7-G7519-200, 7-G7519-1000, 8-G7519-100, and 12-RT-160-500. |  |  | November 03, 2008 | 2008 | Pointe Scientific, Inc. | The product is unable to maintain the specification for linearity through the shelf life of the product. | Process control | Consignees were notified via an Urgent: Medical Device Recall letter dated 11/3/08 that all test results obtained with these lots are questionable and that attending physicians should be informed to determine whether additional testing is required. Also, consignees should examine their inventories for any of the lot numbers listed and dispose of them. If the product has been further distributed, identify customers and notify them of the product recall by including a copy of the recall letter. The Recall Return Response Form should be completed and returned and upon receipt Pointe Scientific will replace any disposed of reagent. Questions should be directed to the Technical Support Department at 1-800-757-5313. |  | 1500 | 1500 liters | Worldwide Distribution -- USA, Hong Kong, Indonesia, Israel, Jordan, Mexico, Nigeria, Poland, Russia, Serbia, Sri Lanka, Trinidad, and Venezuela. | Terminated | December 03, 2009 | 395 | 1500 | Not_Computer
Z-1014-2009 | 50185 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Medtronic TSRH 3DX Implant Set; SPS002058. Sets are unlabeled, but contain individually labeled 3Dx Splined Connectors in sizes small, medium and large. |  |  | October 31, 2008 | 2008 | Warsaw Orthopedic, Inc. dba Medtronic SOFAMOR DANEK | The thickness of the connector may be under specification, which could cause the implant to break during surgery or after implantation. | Process control | Hospital risk managers and implanting physicians were notified separately by letters dated 1/9/09. The letters advised of the recall, reported that a firm representative would visit the facility to pick up the recalled products. Physicians were instructed to closely monitor their patients during postoperative visits, to notify any other physician who may be monitoring the patient in lieu of the implanting physician, and to advise their patients to contact them immediately if they have any sudden and significant symptoms (increased discomfort, sudden change in their clinical symptoms, or new symptoms such as leg pain, back pain or irritation of local tissue). Customers were also instructed to complete and return the attached questionnaire to Medtronic SOFAMOR DANEK via mail (enclosed self-addressed stamped envelope) or fax at 901-399-2047. Direct questions to your Medtronic SOFAMOR DANEK Sales Representative or the firm's Global Quality Department at 800-876-3133, extension 6333. |  | 43 | 43 sets | Nationwide Distribution including states of Arizona, Connecticut, Georgia, Indiana, Minnesota, New Jersey, New York, South Carolina, Texas and Washington. | Terminated | November 18, 2009 | 383 | 4153 | Not_Computer
Z-1092-2009 | 51159 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 3100 / 3100 IQ, Cardiovascular Imaging System, GE Healthcare, Waukesha, WI 53188 (The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | March 06, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare reported an issue with the Exposure Switch Redundancy system with Innova Systems 3100 &amp; 4100, in which this error stays on, even after the pedal is released, inhibiting further x-rays, and a system reset is required to recover. | Software design | On site visits were performed from 3/6/08 - 12/19/08 to update software on all 5 units. |  | 4 | 4 | Worldwide Distribution -- USA including states of MD, TX, and VA and country of China. | Terminated | September 04, 2014 | 2373 | 5 | Software | Device Operation | Software update | Software Update
Z-1074-2009 | 51205 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Magic3 (Intermittent Catheter), Sterile, Uncoated 6 French Pediatric Length REF # 52306 |  |  | January 14, 2009 | 2009 | Rochester Medical Corp | Rochester Medical observed some Magic3 and HydroSil(R) catheter packages in which the catheter was caught in the package seal during the packaging process. | Package design/selection | Consignees were contacted by telephone or in-person visit. The communication provided the reason for the recall and a list of affected products and their lot numbers, and instruction for the removal. Distributors were given instructions to notify their customers in turn. For the customers that were not reached via telephone call or in-person, a letter, "Voluntary recall of magic product samples," was used for the recipients of product samples and sent via e-mail. The letter was dated January 14, 2009 and asked customers to destroy all of their remaining sample inventory of the affected lot numbers. |  | 12 | 12 | Worldwide Distribution -- USA including states of CA, FL, GA, ID, IL, LA, MD, MA, MI, MN, NJ, NM, NY, OH, PA, and WA and countries of United Kingdom. | Terminated | April 20, 2011 | 826 | 1960 | Not_Computer
Z-1070-2009 | 50471 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OneTouch SureStep Test Strips, Part numbers 020-052-01 and 010-359-04, for use with SureStep brand Blood Glucose Meters, manufactured by LifeScan Inc., Milpitas, CA |  |  | November 25, 2008 | 2008 | Lifescan Inc | Deformity in the test strip may result in insufficient blood transfer to reaction area, resulting in inaccurate test results. | Process design | Urgent: Medical Device Recall letters, dated 11/24/2008 were issued to all consignees on November 25, 2008, requesting return of affected lots to LifeScan and instructing how to obtain replacement strips free of charge. A request for sub-recall communications were made as well, and the firm posted notification on its website. Customer service and Sales account reps were provided information to address questions regarding the recall. |  | 23288 | 23,288 vials | Worldwide Distribution -- USA and the Caribbean. | Terminated | December 06, 2010 | 741 | 23288 | Not_Computer
Z-1088-2009 | 50689 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Triage Drugs of Abuse Panel, Model Number: 95000. |  |  | November 06, 2008 | 2008 | Biosite Inc | The recall was initiated after the firm's investigations confirmed that, while lot #186625 passed all internal specifications and quality requirements, the incidence of false positive PCP results has exceeded their post market surveillance thresholds for customer complaints in the field. In addition, the firm has been able to reproduce false positive results reported by customers and as a result | Other | Consignee notification was initiated on 11/6/2008 with Urgent Medical Device Recall notifications, dated 11/5/2008 and mailed to domestic customers sent via US Postal Service Priority Mail and to foreign customers via FedEx. The notice informed customers of the reason for the recall and requested the customers immediately discontinue all use of lot #186625 and discard all affected product in accordance with your local regulations. The customers were also instructed to complete and fax an attached Customer/Distributor Verification form within 10 days. |  | 2634 | 2634 kits | Worldwide Distribution -- USA and country of Canada. | Terminated | October 13, 2011 | 1071 | 2634 | Not_Computer
Z-0642-2009 | 50476 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medtronic COMPLETE SE Self Expanding Biliary Stent System, Model number SC760L, includes a pre-loaded self-expanding Nickel-Titanium (Nitinol) stent with 8 tantalum radiopaque markers, 4 on each end, and an over the wire retractable sheath delivery system. |  |  | November 24, 2008 | 2008 | Medtronic Cardiovascular | Mislabeled: Product may contain wrong instructions for use-- may contain International IFU instead of the correct US (for Biliary indication use only) IFU. | Process control | Medtronic sent a notification entitled "URGENT - VOLUNTARY MEDICAL DEVICE RECALL" to consignees on 11/21/2008. All products are to be returned to the firm's distribution center. For further information, please contact Medtronic Cardiovascular by telephone at 707-566-0111. |  | 23 | 23 units | Nationally in the states of NY, TX, CA, IA, AZ, PA, IN, and KS. | Terminated | March 27, 2009 | 123 | 23 | Not_Computer
Z-1090-2009 | 49548 | No | MKQ |  |  | Pathology | PMA | BD PrepStain System, Model Number: 799-14000-00; Catalog Number: 05CR000021, 05CR00021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 799-14000-00R; Catalog Number: 05CR000021, 05CR000021R, 799-13000-00, 799-13000-00R, 799-14000-00 and 7914000-00R. |  |  | September 26, 2008 | 2008 | Bd Diagnostic Systems Tripath | Software Error in the non-gyn program of some devices may cause non-gyn samples to be incorrectly transferred from centrifuge tubes to slides potentially result in mismatched patient results. | Software change control | BD Diagnostics began contacting known non-gyn customers by phone on 09/26/2008. Customers were notified by Urgent - Important Product Information letter on 09/29/2008. They were advised to stop using the system until they complete a list of actions provided in the letter. A telephone contact number for Technical Support was provided in the letter. A permanent software based solution will be provided to customers when completed. They anticipate the software will be available in October 2008. |  | 662 | 662 units | Worldwide Distribution | Terminated | January 15, 2013 | 1572 | 662 | Software | Output/Calculation | Software update | Software Update
Z-1008-2009 | 50224 | No | JWX | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Dryspot (R) Staphytect Plus, Catalog #DR0100M, 120 tests/kit. |  |  | October 07, 2008 | 2008 | Remel, Inc | Control and test wells are reversed on some of the reaction cards. | Pending | The recalling firm issued Important Medical Device Customer Advisory Notice letters dated 10/7/08 via regular mail to their customers explaining the reason for recall. The letter requested that they review their inventory for the product lots indicated in the letter; discard all of the lot remaining in their inventory; complete and return the Product Inventory Checklist. |  | 32 | 32 kits dist. by recalling firm | Nationwide Distribution -- including MN, MI, IL, and NC. | Terminated | July 30, 2009 | 296 | 32 | Not_Computer
Z-1009-2009 | 50170 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche COBAS c 111 Analyzer with ISE module option, GMMI No. 04777433001; Roche Diagnostics, Indianapolis, IN. |  |  | November 17, 2008 | 2008 | Roche Diagnostics Corp. | Due to a manufacturing error at the supplier, not all cobas c 111 pinch valve caps are properly glued to the correct position on the pinch valve shaft. Without the glue, the cap can become loose or misaligned. The effect of these faulty valves is that they may not open or close properly which can disturb the fluid transportation within the ISE module resulting in problems with ISE calibration, con | Nonconforming Material/Component | Consignees were notified via an Urgent Medical Device Correction letter dated 11/17/08 and instructed Roche will contact them to make an appointment to inspect their analyzers. A work around was provided, for interim use, in the letter. Customers were instructed to complete and fax back the "ISE Controls Request - Purchased from a Distributor." A no charge replacement product will be shipped to customers. |  | 66 | 66 | Nationwide Distribution | Terminated | December 08, 2009 | 386 | 66 | Not_Computer
Z-1041-2009 | 50174 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio-Rad Variant II TURBO Hemoglobin Testing System with CDM v 4.0, model number 270-2600, Manufactured by Bio-Rad Laboratories, Hercules, CA |  |  | October 17, 2008 | 2008 | Bio-Rad Laboratories Inc | Sample results may be assigned to the wrong barcode number and tube position numbers, leading to discrepant sample results. | Pending | An Urgent Medical Device Correction letter was sent out on October 17, 2008 via fax and phone call. The letter recommended that customers review the Summary Report for duplicate tube position numbers or barcode IDs within the same rack. If there are questions or a need for assistance, customers are to contact their regional Bio-Rad office. Follow-ups will be made once root cause is identified. |  | 92 | 92 units | Worldwide Distribution | Terminated | November 12, 2010 | 756 | 92 | Not_Computer
Z-1040-2009 | 50163 | No | NQN | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | QCA and iQCA software, a video microscopy software system for quantitative estrogen receptor immunohistochemistry; Cell Analysis Inc., Evanston, IL 60201 or Highland Park, IL 60035. |  |  | November 06, 2008 | 2008 | Leica Microsystems, Inc. | A change in the software to open up the software architecture to include other analytes other than Estrogen Receptor was made without FDA review and approval. | Software change control | Leica Microsystems sent Medical Device Correction letters dated 11/6/08 to the direct accounts, informing them that all versions 3.7 and above of QCA and iQCA software are not in compliance with the original 510(k) cleared software for scoring of immunohistochemical stains in clinical applications. The QCA/iQCA software is only cleared for Estrogen Receptor staining under clinical applications. The accounts were provided with red warning stickers that state: "WARNING The clinical application of QCA software has been cleared only for the measurement of images from microscopic slides of breast cancer specimens stained for the presence of estrogen (ER) nuclear receptor protein only. Performance characteristics of the product, when used outside of this intended use, have not been established." They were instructed to apply the stickers to any existing manuals and or at each point of use, and to instruct all users about the notification. The accounts were also requested to complete and return the enclosed reply form confirming that the changes have been made. |  | 5 | 5 units | Nationwide Distribution -- including states of CA, IL, KS, and NY. | Terminated | March 16, 2011 | 860 | 5 | Software |  | Instructions for attahing stickers | Safety Notice/Insructions
Z-1160-2009 | 50867 | No | LRS | Intravascular catheter. | General Hospital | General &amp; Plastic Surgery | Enforcement Discretion | Tri-State Centurion PICC Insertion Tray, sterile, 15 per case, Tri-State Hospital Supply Corp., Howell, MI.; Reorder CVI2415A. |  |  | January 14, 2009 | 2009 | Tri-State Hospital Supply Corporation | Lack of assurance of sterility, as the package seals may be inadequate. | Packaging process control | Tri-State Hospital Supply Corporation notified all consignees via Urgent Product Recall Notice letter dated January 14, 2009. Consignees were instructed to remove the product from inventory, to inform the firm of amount on hand, and to hold the product for retrieval. |  | 60 | 60 | Nationwide. | Terminated | December 31, 2009 | 351 | 1186 | Not_Computer
Z-1133-2009 | 51287 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Medtronic Innervision Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27782B. Sizes: 5 CM in length, REF (240) 27782B. Each catheter is provided in a sealed Tyvek plastic tray placed within a sealed Tyvek/Mylar pouch. The sealed pouch is then placed within a chipboard carton along with product IFU and patient wallet cards. The product is ETO sterilized. |  |  | February 12, 2009 | 2009 | Medtronic Neurosurgery | The ventricular catheters may become detached from the Snap Base assembly after implantation and this condition may increase the need for emergency revision surgery. | Device Design | An Urgent - Medical Device Recall notification was sent on February 13,2009 to the attention of risk managers, advising them to cease use of the Innervision Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27782 ( all versions and all lot numbers), the Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27802 ( all versions and all lot numbers), and the Snap Shunt Ventricular Catheter, BioGlide, Catalog Number 27708 ( all version and all lot numbers). Each customer is to cease the use of the listed product and any unused unit needs to be returned to Medtronic. Further, each account is required to account for product used/implanted and to advise surgeons of the recall information. Each account will recieved an individualized Customer Product Reconciliation Form provided by their Medtronic Sales Representative for determining the status of any product shipped to them. |  | 3048 | 3048 of all units | Worldwide Distribution -- USA, Ireland, and Australia. | Terminated | November 08, 2010 | 634 | 3048 | Not_Computer
Z-1052-2009 | 50896 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Stryker Impaction Bur Guard REF 2296-301, 2296-300, 0260-901-300, &amp; 5400-300; Stryker Instruments, Kalamazoo, MI. |  |  | January 27, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | Bur guards were distributed without an expiration date, and if they are used after they have expired it could result in the patient or operating room staff experiencing a burn or laceration. | Process control | Customers were notified via a Medical Device Recall Notification letter dated 1/27/09 and instructed to immediately check all stock areas or operating room storage and inspect their bur guards for an expiration date. They were also instructed to quarantine any affected product found, to indicate on the Business Reply Form the quantity of affected product on hand, to destroy any affected product found without an expiration date by breaking off at least 3 of the tabs on the bur guard, then dispose of the affected product in normal trash, and complete and return the business reply form. Customers are to fax the Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2760. Upon receipt of the Business Reply Form, a replacement bur guard(s) will be sent for any affected product destroyed. |  | 70 | 70 | Nationwide Distribution | Terminated | December 17, 2009 | 324 | 70 | Not_Computer
Z-1042-2009 | 50704 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Diasorin LIAISON (R) Chemiluminescence Analyzer Model # 15970. |  |  | December 05, 2008 | 2008 | Diasorin Inc. | DiaSorin determined that there was a potential for a short to occur in a a residual starter pump electrical cable in the LIAISON (R) Analyzer. | Device Design | DiaSorin Inc authorized Field Service Engineer has visited the sites and performed the necessary field correction. Upon completion of their work, the field service engineer generated a Field Service Report and a copy was left with the customer. |  | 9 | 9 | Worldwide Distribution -- USA including states of NC, NJ, PA, CA, WA, MO, UT, and OR and country of Canada. | Terminated | May 22, 2012 | 1264 | 9 | Not_Computer
Z-1089-2009 | 49438 | No | LJM | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Sm ELISA Test System |  |  | August 19, 2008 | 2008 | Zeus Scientific Inc | The Calibrator Value is printed incorrectly on the Sm ELISA Test System, lot # 08041819. The Calibrator Value (CV) printed on the label is 102; the correct CV should read 266. | Employee error | Urgent Recall Notice letters were sent out by Federal Express night mail on August 19, 2008 to all direct account. Wampole (Inverness) will notify their customers in a separate letter. Recall letters explained the reason for the recall and that the product does meet the company's criteria when it is used qualitatively to interpret index values or OD rations as outlined in the kit package insert. Customers are advised in the letter that the Calibrator Value of 102 is incorrect and that it should be 266. Instructions to correctly express the specimen results using specimen and calibrator OD values semi-qualitatively as Unit Values and the Calibrator Value (CV) or 266 are included. Letters also state that results for all patients using the incorrect CV of 102 should be recalculated using the correct CV of 266. All letters included a fax back form indicating that they have received the recall notification, have notified customers to correct the CV value for the Semi-Quantitative Conversion and have notified customers that have received the kit to correct the Semi-Quantitative Conversion of Optical Density to AAU/mL for the affected lot number of the Sm ELISA Test System. |  | 265380027 | 33 of lot 08041819, 27 of lot 08041811. | Nationwide Distribution | Terminated | April 30, 2009 | 254 | 265380027 | Not_Computer
Z-1073-2009 | 50520 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Dual Taper Wedge Pressed Fit Femoral Component. |  |  | October 31, 2008 | 2008 | Stelkast Co | A possible breach of the sterile barrier packaging (Tyvek pouch). | Packaging | The recalling firm telephoned the distributors to inform them of the problem and the need to return the product. |  | 187 | 187 units | Nationwide Distribution -- states of NC, TX, and WV. | Terminated | April 14, 2009 | 165 | 187 | Not_Computer
Z-1020-2009 | 51177 | No | NQO |  |  | Orthopedic | PMA | Spacers of varying lengths used with the InSTANT Stranding System, Part numbers: TG3138, TG3139, TG3140, TG3141 and TG3142. Theragenics Corporation, 5203 Bristol Industrial Way, Buford GA 30518. |  |  | December 08, 2008 | 2008 | Theragenics Corporation | Some pouches were found to have weak seals potentially compromising sterility. | Packaging process control | Consignees were notified first by telephone on December 8, 2008 and sent a follow up letter dated January 13, 2009. They were instructed to return any of the affected lots remaining in their inventory. |  | 62 | 62 units | Nationwide Distribution -- states of VA, WI, and CO. | Terminated | April 06, 2009 | 119 | 62 | Not_Computer
Z-1043-2009 | 50821 | No | LRF | Cryptococcus neoformans serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BEC Growth Chek Microbial Suspension - Candida albicans CA1 (100 dose), Product #10231-100 |  |  | January 13, 2009 | 2009 | Lexamed | Candida albicans microbial suspension, certified to deliver less than 100 organisms per dose, was found to have a population count which exceeded 100 organisms per dose, following distribution. | Process control | The recalling firm contacted the customer and notified them of this recall via telephone on 1/13/2009, followed by an Urgent Product Recall Notice which was sent to the customer via Certified mail on 1/14/2009. The letter stated the problem with the affected lot and asked that the customer examine their inventory, remove the affected lot, and return the used and unused vial to LexaMed. The customer is to complete and return the enclosed acknowledgement card by mail or fax. Questions should be directed towards Anne Schuler, QA Manager at 419-693-5307 or by e-mail aschuler@lexamed.net. |  | 1 | 1 vial | USA: Utah | Terminated | December 09, 2009 | 330 | 1 | Not_Computer
Z-1204-2009 | 50482 | No | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | diaDexus Inc. PLAC Test Reagent Kit, Turbidimetric Immunoassay for the Quantitative Determination of Lp-PLA2 in Human Plasma or Serum, liquid ready-to-use two-reagent kit assembled in a kit box, manufactured by diaDexus Inc, South San Francisco, CA 94080, Catalog number 90110 |  |  | November 04, 2008 | 2008 | DiaDexus, Inc | Storage equipment failure at distributor site due to Hurricane Ike caused product to experience a temperature variation and go out of range. The integrity of the product could not be assured. | Other | The firm and its distributor contacted the consignees via phone on 11/04/08. The firm informed them of the issue and inquired about the products' status. An Urgent: Device Recall letter was issued on 11/04/08 via FedEx. The letter requested that use of the kit be immediately discontinued and returned promptly. A phone call was made to follow up requesting return of the product. |  | 27 | 27 units | Nationwide Distribution | Terminated | December 21, 2009 | 412 | 27 | Not_Computer
Z-1122-2009 | 51236 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ILED Surgical Light, Model numbers: ILED3, ILED5 and ILED5K |  |  | February 27, 2009 | 2009 | Trumpf Medical Systems, Inc. | Cracks may form around the screw connections of the cover plates causing brittleness, which may lead to fragments breaking off and falling into the operating area. | Other | Consignees were notified via an Urgent Safety Information letter starting 02/27/2009. The Recall Notice asks customers to visually inspect the ILED covers for cracks and in particular to focus on the corners. If cracking is visible they are asked to remedy by taping with a clear tape over the cracks and report it to TRUMPF Medical Systems, Inc. immediately. TRUMPF will be replacing all covers with a newly designed and tested cover. Original priority in the United States will be set to start with the oldest units affected. If TRUMPF receives notification that cracking does exist then that facility will rise to the highest priority. Additionally, TRUMPF will issue a new approved cleaning and disinfection instruction, since investigations have shown that the cracking is significantly influenced by the use of certain disinfectants containing alcohol. The operator manual previously stated to avoid using alcohol-based cleaners on some parts of the ILED, but after reviewing this it was determined that it was not clearly stated to avoid use on the covers. |  | 4480 | 4480 units (540 in U.S.) | Worldwide Distribution. | Terminated | October 22, 2009 | 237 | 4480 | Not_Computer
Z-1206-2009 | 51053 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll AED Plus Defibrillator |  |  | February 12, 2009 | 2009 | ZOLL Medical Corporation, World Wide Headquarters | Device fails to discharge the defibrillation energy. | Process design | On 3/31/09, the firm revised their recall strategy and required all users to a perform a "Mandatory Upgrade" software The software upgrade allows all potentially affected devices to monitor battery charging performance through periodic self-testing. If defective batteries are detected at any time prior to the recommended maximum of five years, users are prompted by the device to install fresh batteries. Frim issued Press Release on April 2, 2009.In the first notification Zoll notified customers by letter on 2/12/09 via certified mail. Customers instructed to replace device batteries every three years. In lieu of replacing batteries every three years, customers will also have the choice to update theirdevices with software that will monitor batteries for the identified defect. ZOLL will be send email notification to all consignees who have provided an email address as part of their contact information. ZOLL will also be publishing information of the notification on industry magazines. http://www.zollaedplusbatteryhelp.com/ |  | 183535 | 183,535 units | Worldwide distribution | Open, Classified |  |  | 183535 | Battery | Device Operation | Software update | Software Update
Z-1053-2009 | 50914 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Coonrad/Morrey Total Elbow, Interchangeable Humeral Assembly, extra small, 4 inch length, sterile, Zimmer Inc., Warsaw, IN; REF 32-8105-027-04. |  |  | January 30, 2009 | 2009 | Zimmer Inc. | The surgical technique is inadequate for extra-small sized implants, resulting in the possibility of the provisional becoming stuck in the bone. | Labeling design | Consignees were notified by "Urgent Device Correction" letter dated January 30, 2009 and provided revised surgical techniques for these extra-small implants. Direct questions about this recall to Zimmer, Inc. at 574-267-6131. |  | 14613 | 14,613 of all products | Worldwide Distribution, including USA, Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Czech Republic, Denmark, England, Finland, France, Germany, Hong Kong, Hungary, India, Italy, Japan, Jordan, Latin America, Netherlands, Russia, Singapore, South Africa, South Korea, Spain, Switzerland, Taiwan, Thailand, Turkey and Venezuela. | Terminated | December 22, 2009 | 326 | 14613 | Not_Computer
Z-1231-2009 | 51223 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Electric Stretcher with the Quick Drop CPR option, Model 1550. |  |  | February 27, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The head section backrest (fowler) may not raise up, or, if up, may drop suddenly due to wearing out of the actuator threads on the fowler. | Device Design | Consignees were notified by an Urgent - Medical Device Correction letter dated 2/27/09, advised of the problem and instructed to locate stretchers and verify that there is no grinding noise when activating the fowler electrically. Consignees were also instructed to remove any stretchers with a fowler that was non-operational, or which makes a grinding sound, from service until the firm makes repairs and contact the firm. Stretchers will be corrected within the next four months by a Stryker field service representative. The enclosed post card should be returned to confirm receipt of the notification. The notice should be forwarded to any users that have borrowed or purchased the stretchers and the firm should be notified of their new location. Also, if any of the affected stretchers have been disposed of or are no longer in use, the firm needs to be provided with their serial numbers. Questions or concerns should be directed to Andrea Kuthe at 269-324-6581. |  | 3607 | 3607 | Worldwide Distribution | Terminated | April 25, 2011 | 787 | 3645 | Not_Computer
Z-1044-2009 | 50518 | No | DHN | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | ANA Fluorescent Test System and ANA Colorzyme Test System, Sub assembly catalog number SA 2013, Kit Catalog Numbers, 2000, 2002, 2003, 2004, 2005, 2100, 2200, 4000, 4200, 8024.1, and 8060, manufactured by Immuno Concepts N.A. Ltd, Sacramento, CA. |  |  | December 04, 2008 | 2008 | Immuno Concepts Inc | Samples can occasionally run out of the wells and cross contaminate adjacent wells. | Nonconforming Material/Component | The firm notified customers via a "Notice of Market Withdrawal" letter, e-mail and telephone. The letter described the problem and instructed customers to removed affected product from their inventory, complete the attached form and return it via fax to Immuno Concepts, Inc. at 916-363-2843. The firm will issue replacements as requested. Direct questions to Immuno Concepts, Inc. at 916-363-2649. |  | 378 | 378 kits | Worldwide Distribution, USA (NC, WI, MS, MD, IL, GA, CA, MI and Puerto Rico) and countries of Sweden, Mexico, Spain and Denmark. | Terminated | November 10, 2009 | 341 | 378 | Not_Computer
Z-1039-2009 | 51065 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, S/5 iCentral and iCentral Client, |  |  | January 26, 2009 | 2009 | GE Healthcare | GE Healthcare is providing information about two issues associated with iCentral L-NETO5 and L-NETC05 v4.3, 4.4, 4.5 and 5.0 when used in combination with Dash 2000, 3000, 4000 and 5000; Solar 7000, 8000, 8000M and 8000i; Eagle 3000 and 4000 and ApexPro. The iCentral does not provide audio or visual alarming for the following: A FIB, VT&gt; and PVC(1) alarms when used with ApexPro and VT&gt;2 and PVC | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated January 26, 2009 addressed to Healthcare Administrator/Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the safety issue, affected product details, safety instructions, product correction and contact information. Direct questions to GE Healthcare at 1-800-558-7044, Option 2 and ask for Technical Support or contact your local GE Healthcare Distributor. |  | 2146 | 2,146 | Worldwide Distribution, USA (and PR) including states of AZ, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY,LA, ME, MI, MN, MS, MO, NH, NJ, NY, NC, OH, PA, SC, TX, VT, VA, WA, and WI and countries of AUSTRALIA, AUSTRIA, BARBADOS, BELGIUM, BOLIVIA, BRASIL, BULGARIA, CANADA, CHILE, CHINA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, ECUADOR, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, GAYANA, HONG KONG, HUNGARY, ICELAND, INDIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LIBYAN ARAB JAMAHIRIYA, MACEDONIA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NORWAY, PAKISTAN, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SOUTH AFRICA, SLOVEN IA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, SYRIA ARAB REPUBLIC, THAILAND, TUNAS IA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA and YEMEN. | Terminated | August 13, 2013 | 1660 | 2146 | Software | Alarm/Message | Instructions | Safety Notice/Insructions
Z-1045-2009 | 50519 | No | LLL | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | RELISA ENA Multiparameter Antibody Screening test, enzyme immunoassay test system, manufactured by Immuno Concepts N.A. Ltd., Sacramento, CA. |  |  | November 25, 2008 | 2008 | Immuno Concepts Inc | In certain patient samples, it is difficult to properly discriminate between anti-Sm/RMP positive samples and samples that are only anti-Sm positive. | Process control | The firm has issued telephone calls, e-mails and notification letters to its consignees. |  | 25 | 25 kits | Worldwide Distribution including USA, Australia, England, Sweden and Austria. | Terminated | November 17, 2009 | 357 | 25 | Not_Computer
Z-0999-2009 | 50875 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Viasys Healthcare, Sonora Transcranial Doppler (TCD) System. Model Sonara. Part Number of Units: 982A0470 Sonara Bilateral. |  |  | January 06, 2009 | 2009 | Cardinal Health, Inc | Internal testing has identified several several scenarios where unexpected System lockups, shutdowns or system errors would render the Sonara unavailable for use until the Sonara can be restarted. These issues can be corrected by installing the Version 6.0 b045 software updating the Audio Driver and updating to XP Service Pack 3. Additionally, improper wiring could cause the cooling fan to not op | Device Design | Consignees were sent a Cardinal Health "Urgent: Medical Device Correction" letter dated January 12, 2009 beginning the week of February 2, 2009 and continuing into week of February 9, 2009. The letter was addressed to customers describing the problems and systems affected and solutions. Cardinal Health requested customers to complete the actions listed in the letter to install the Field Correction and return the completed "Verification Form" via mail, e-mail or fax to Cardinal Health-NeuroCare. Further information can be obtained by calling 1-800-356-0007 ext 2191 or fax to 1-877-244-3805. |  | 64 | 64 systems | Worldwide Distribution -- USA including states of FL, DC, LA, NJ, NY, PA, TN, TX, WI, CA, MO, and LA and countries including THAILAND, SINGAPORE, SERBIA-MONTENEGRO, PORTUGAL, PHILIPPINES, POLAND, NORWAY, NETHERLANDS, KOREA, ITALY, ISRAEL, HUNGARY, HONG KONG, GREAT BRITAIN, SPAIN, GERMANY, CZECH REPUBLIC, SWITZERLAND, CANADA, BULGARIA, and BOSNIA-HERZEGOVINA. | Terminated | July 03, 2012 | 1274 | 100 | Software | Device Operation | Software update/Audio Driver update/OS update | Software Update
Z-1061-2009 | 51136 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet Offset Tibial Tray Adaptor, 5.0 MM, sterile, Biomet Orthopedics, Inc., Warsaw, IN, REF 141491. |  |  | February 13, 2009 | 2009 | Biomet, Inc. | A 2.5mm locking insert was packaged with the 5.0mm Offset Tibial Tray Adaptor instead of the 5.0mm locking insert. | Packaging process control | Consignees were notified by "Field Safety Corrective Action" letter dated February 13, 2009 to conduct a recall to the hospital level and to locate and return the recalled product to Biomet Orthopedics. Customers were also instructed to complete the enclosed "Fax Back Response Form" and return it via fax to Biomet Orthopedics at 1-574-372-1683. Direct questions about this recall to 1-800-348-9500 or 1-574-372-3983, Monday through Friday, 8 a.m. to 5 p.m. |  | 17 | 17 | Worldwide Distribution including Germany, Italy, Netherlands, Spain and the United Kingdom. | Terminated | August 11, 2009 | 179 | 17 | Not_Computer
Z-1171-2009 | 51173 | Yes | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | Internal Sterilizable Paddles (electrodes) and Handles without discharge control for use with LIFEPAK 9 and LIFEPAK 9P defibrillators/monitors. |  |  | February 16, 2009 | 2009 | Physio Control, Inc. | Sterilization efficacy can not be confirmed for the current methods of sterilization listed in the product labeling. | Other | Recalling firm sent Urgent Medical Device Recall letters, dated 2/2009, advising customers to discontinue use of the listed sterilization methods and to remove them from service since there is no sterilization method that can be supported by current data. They are to use the enclosed form to identify and confirm if they have the listed products and follow the instructions on the form for returning information to the firm. |  | 3218 | TOTAL: 3218 units; 2052 in the US, 1166 internationally | Worldwide Distribution | Terminated | August 17, 2011 | 912 | 3218 | Not_Computer
Z-1174-2009 | 51402 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | Integra Bilayer Matrix Wound Dressing, 8 in X 10 in (20 cm x 25 cm); Non-pyrogenic, Do not Re sterilize, Store flat at room temperature, Avoid Excessive heat, Avoid freezing. Sterile. Catalog number BMW810. |  |  | March 10, 2009 | 2009 | Integra LifeSciences Corp. | Package Seal Integrity: Integra became aware of the potential of open seals in the foil pouch of the Integra Bilayer Matrix Wound Dressings. | Nonconforming Material/Component | Urgent: Product Recall Notification letters, dated March 10, 2009, were sent out by first class mail to all consignees. The letters asked consignees to quarantine any stock on hand and contact their local Integra Sales Representative for further assistance and for arrangement of product returns. Consignees also need to complete and return the "Recall Acknowledgement and Return Form" indicating whether or not they possess any affected product. Additional questions should be directed to Judith O'Grady at 609-936-2311. |  | 125 | 125 units | Nationwide Distribution -- including states of CA, FL, MN, TX, NC, GA, NY, OH, MI, NE, and PA. | Terminated | June 26, 2009 | 108 | 125 | Not_Computer
Z-1170-2009 | 50795 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | ReadyMed Elastomeric Ambulatory Pump, Model Numbers: R050050, R100100, R100200, R250050, R250100, and R250167. |  |  | December 19, 2008 | 2008 | Cardinal Health 303 dba Cardinal Health | The recall was initiated as a result of a potential for out of specification flow rates. The product may not consistently meet the labeled flow rate range of +/- 15%. | Process design | Cardinal sent an Urgent: Medical Device Recall Notification letter, dated 12/19/2008, to consignees on December 22, 2008 stating the Public Reason for Recall, and informing them of the Short Term and Long Term Actions. In addition a modification to the "Dear Pharmacist" letter, FAQ and Recall Response Card was mailed beginning January 26, 2009 to clarify that the pharmacist's ReadyMed products should be sent directly to their distributor and the credit for their product will also be provided through their distributor. Cardinal Health will work directly with the distributor and not the pharmacist as they are not Cardinal Health's direct consignee.Short Term: All ReadyMed elastomeric ambulatory infusion system models are subject to this recall. Cardinal Health is requesting that the use of the ReadyMed elastomeric ambulatory infusion system be discontinued and that the product be returned to Cardinal Health for credit.-Immediately complete and return the enclosed pre-addressed and postage paid Recall Response Card. Note on the card the ReadyMed elastomeric ambulatory infusion system models, lot numbers and quantities that will be returned. Cardinal will provide a courier mailing package for that return and will process consignees' credit.-Customers who purchased through a distributor will receive their credit through that distributor and not directly from Cardinal Health.-Consignee distributor has already been notified of this recall. If you have any questions, please contact your distributor directly, or call the Cardinal Health ReadyMed Recall Call Center at the number below.-Enclosed is a Frequently Asked Questions (FAQ) Letter for Pharmacists to address any questions or concerns related to the use of the ReadyMed elastomeric ambulatory infusion system.Long Term Actions: cardinal Health is continuing to investigate the out of specification condition of the ReadyMed elastomeric ambulatory infusion system and reviewing possible corrective measu |  | 2600000 | Approximately 2.6 million units | Nationwide Distribution | Terminated | September 13, 2011 | 998 | 2600000 | Not_Computer
Z-1096-2009 | 50696 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Dual Port Wizard Low Profile Replacement Gastrostomy Device, REF 00200W, 16F x 1.2cm, 10cc, Sterile EO, Bard Access Systems, Inc., Salt Lake City, UT 84116. |  |  | December 22, 2008 | 2008 | C R Bard Inc | Out of specification replacement gastrostomy device was distributed. | Process control | Consignees were notified by "Product Recall" letter distributed on January 7, 2009 via Federal Express overnight mail and asked to check all inventory locations, remove unused affected product and return to Bard Access Systems, Inc (BAS). Customers were also instructed to complete the enclosed Reply Form and Inventory Reconciliation Form and fax it to BAS at 1-801-522-5674. Additional information is available by calling BAS Customer Service at 1-800-290-1689. |  | 450 | 450 units | Worldwide Distribution, USA (states of AL, AZ, CA, CO, CT, FL, GA, IA, IL, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, NC, NH, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV) and countries of Canada, Japan, Finland, Norway, Sweden, Austria, Belgium, Holland, United Kingdom, Italy, France, Switzerland, Spain and Germany, Australia, and Latin America. | Terminated | May 18, 2010 | 512 | 29939 | Not_Computer
Z-1172-2009 | 51135 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic RV Lead Integrity Alert , Cat. # SW012, 1.0 software in EnTrust(R) (D153ATG, D153VRC, D154ATG, D154VRC) and EnTrust(R) Escudo (D144DRG, D144VRC) defibrillators. |  |  | November 03, 2008 | 2008 | Medtronic Inc. Cardiac Rhythm Managment | Medtronic has identified an issue related to installation or removal of the Lead Integrity Alert software in EnTrust(R) and EnTrust(R) Escudo" defibrillators. In those devices only, installation or removal of LIA will inadvertently turn off two audible patient alerts. | Software design | Between November 24 and December 9, 2008 US CRDM are responsible to hand deliver the consignee letter, a Medtronic "Urgent Patient Management Information," dated November 2008, to impacted physicians to make them aware of the issue and continue to promote and install LIA (by using the new LIA tipcard). For any physician listed that a confirmation sheet was not received by end of day, December 9, 2008, the letter was mailed to the physician on Dec 10, 2008. The letter described the issue and the product involved. The letter also gave recommendations and informed consignees in regard to a future update to the LIA software for EnTrust devices. |  | 55755 | 55,755 | Worldwide Distribution -- USA including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, DC, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MI, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and countries of Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Bolivia, Brazil, Canada, Cayman Islands, Chile, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jamaica, Jordan, Kuwait, Lebanon, Macedonia, Martinique, Netherlands, Netherlands Antilles, Norway, Oman, Peru, Poland, Portugal, Puerto Rico, Qatar, Reunion, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom Vatican City State, Venezuela, and Virgin Islands. | Terminated | May 10, 2012 | 1284 | 55755 | Software | Alarm/Message | Software update | Software Update
Z-1072-2009 | 50771 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Soft Tip Sheath Multipurpose Introducing Sheath Kit, Model Numbers: |  |  | December 24, 2008 | 2008 | Boston Scientific Corp | The marker band may detach from the sheath and a clinical failure could result in an embolism. | Process design | The firm initiated the recall on December 24, 2009 and notified consignees via "Urgent Company Initiated Medical Device Recall - Immediate Action Required" letters which explain the reason for the recall and request the affected product be returned to the Boston Scientific Corporation Distribution Center. An amendment letter with two additional Convoy Advanced Delivery Sheath Kit models (M0045748ST0 and M0045749ST0) was sent to affected customers on January 27, 2009. Direct questions about this recall to your local Sales Representative. |  | 31340 | 31,340 of all units | Worldwide Distribution. | Terminated | April 06, 2010 | 468 | 31340 | Not_Computer
Z-1181-2009 | 51300 | No | LRW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Pilling(R) Coronary Scissors, 120Deg w/bar, Catalog numbers: and 352178, Teleflex Medical, Research Triangle Park, NC 27709 |  |  | March 03, 2009 | 2009 | Teleflex Medical | The product has a rough square finish instead of a smooth rounded finish which can cause damage to arteries. | Process control | Consignees were notified by an Urgent Medical Device Recall letter, dated 03/03/2009, via UPS Next Day. Hospitals were instructed to immediately discontinue use of the products, inspect the scissors for the date codes and to return the product to Teleflex Medical for replacement. Distributors were instructed to forward the letter to their consignees to retrieve relevant product from those locations. A Recall Acknowledgement Form was attached to be completed and faxed back to Teleflex Medical. A second mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the second mailing or they may be contacted by phone, e-mail or fax. A record of these contacts will be maintained. Affected product retrieved by Teleflex Medical will be destroyed and replacement product provided to the consignee. Records of destruction will be maintained. |  | 497 | 497 of all units | Worldwide Distribution -- USA including states of AL, AZ, CA, CO, DC, DE, FL, IA, IN, KS, MA, MD, MI, MO, MS, NH, NJ, NY, NV, OH, OK, OR, PA, RI, TN, TX, UT, and WA, and countries of Australia, Canada, France, Germany, India, Japan, and Singapore. | Terminated | March 31, 2011 | 758 | 497 | Not_Computer
Z-1065-2009 | 50171 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche/Hitachi Modular Analytics System, clinical chemistry analyzer Modular D2400 Module, Roche Diagnostics Corp., Indpls., IN; GMMI No. 04998618001. |  |  | November 12, 2008 | 2008 | Roche Diagnostics Corp. | If a masked module is unmasked during the reagent registration process, the calibration parameters may be assigned to the wrong reagents, resulting in the reporting of incorrect control and sample results. | Software change control | Consignees were notified via an Urgent Medical Device Correction letter dated 11/12/08 and instructed to only execute "module unmasking" when the unit is in "standby" mode and to consult with their physicians at the facility to determine the clinical implications and whether any clinical intervention is needed for patients treated based upon results obtained with this software. Questions should be directed to Roche Diagnostics Technical Support at 1-800-428-2336. |  | 22 | 22 systems in all | Nationwide Distribution -- including states of California, Florida, Georgia, Hawaii, Indiana, Maryland, New York, North Carolina, Pennsylvania, and Virginia. | Terminated | September 22, 2009 | 314 | 22 | Computer
Z-1064-2009 | 50478 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-Dyn Ruby Hematology Analyzer, List Number 08H67-01 |  |  | November 17, 2008 | 2008 | Abbott Laboratories | Carryover failures in software revisions 1.0ML and 2.0ML result in elevated PLT background count. | Other | Recall initiated on 11/17/2008. A Product Correction letter and customer reply form is sent to all currently active CELL-DYN Ruby customers. The customers are asked to keep the communication with their CELL-DYN Ruby System Operator's Manual and to note that the information listed in the table replaces the carryover specification provided in all applications of the manual. Questions from customers in the US should be directed to Customer Support at 1-877-4ABBOTT. Customers outside the US should contact their local hematology customer support representative. |  | 899 | 899 units | Worldwide Distribution | Terminated | November 27, 2010 | 740 | 899 | Software | Output/Calculation | Notification | Safety Notice/Insructions
Z-0986-2009 | 50753 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Trio Small Offset Connector |  |  | September 05, 2006 | 2006 | Stryker Spine | Reports were received of the ball ring ejecting from the connector's body when other connectors have already been locked down. | Device Design | A notification letter dated September 12, 2006 and a Product Accountability Form were sent via Federal Express (with return receipt) to all affected branches/agencies. Further information can be found by calling 1-201-760-8256. |  | 2417 | 2,417 units | Nationwide Distribution including OR, CA, MA, VA, GA, FL, NY, NM, NJ, UT, KS, OH, MI, MN, IN, MT, TX, WI, AL, CO, and TN. | Terminated | April 14, 2009 | 952 | 2417 | Not_Computer
Z-1121-2009 | 50927 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare, Infinia HawkeyeÂż 4, Model Numbers: 5178728, ASM500060, |  |  | June 02, 2008 | 2008 | GE Medical Systems, LLC | The metal pin in the adjustable head holder, which is used to acquire Brain SPECT data on Infinia systems (Infinia, Infinia HawkeyeÂż, and Infinia HawkeyeÂż 4), creates an artifact on the patient's cerebellum, which may lead to the potential for misdiagnosis. | Other | Consignees were sent a GE "Product Safety Notification" letter dated April 2008. The letter was addressed to Hospital Administrators, and Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information. Contact your local sales or service representative for further information regarding this notification by calling 1-800-437-1171 (within the US or Canada) and 1-262-896-2890 (outside the US). |  | 288 | 122 US and 166 OUS | Worldwide Distribution -- USA (DC and Puerto Rico) including states of AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, and WI and countries ofAustralia,Austria,Belgium,Brazil,Canada,China,Colombia,Czech,Denmark,Ecuador,Egypt,Finland,Finland,Egypt,Finland,France,Germany,Greece,Hungry,India,Indonesia,Ireland,Israel,Italy,Japan,Korea,Malaysia,Martinique,Mexico,Netherlands,New Zealand,Norway,Pakistan,Panama,Philippines,Poland,Portugal,Republic of Georgia,Russia,South Africa,Spain,Sweden,Switzerland,Taiwan,Thailand,Turkey,Ukraine, and United Kingdom. | Terminated | May 24, 2012 | 1452 | 1256 | Not_Computer
Z-1094-2009 | 50990 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Infinia Hawkeye(R), Model numbers: |  |  | September 08, 2008 | 2008 | GE Medical Systems, LLC | There is a possibility for misregistration when brain scans are performed with Infinia with the Hawkeye option. This misregistration could present the potential for misdiagnosis when the patient headrest sags under the weight of a head. | Software design | Consignees were sent from September 2008 - Feb 2009 a GE "Product Safety Notification" letter dated 4/2008. The letter was addressed to Hospital Administrators, and Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information. It instructed consignees to continue to use the Infinia Hawkeye system normally for all types of Hybrid scans and stated that as a general practice tomograms should be reviewed carefully using the Xeleris Volumetrix processing protocol to ensure proper registration. Local sales or service reps should be contacted for further information regarding the notification. |  | 417 | 153 US, 264 OUS | Worldwide Distribution -- USA including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, DC, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK ,OR, PA, PR, SD, TN, TX, UT, VT, VA, WA, WV, and WI and countries of AUSTRALIA, BELGIUM , BRASIL, CANADA, CHINA, COLUMBIA, CZECH REPUBLIC, DENMARK, ECUADOR, FINLAND, FRANCE, GEORGIA, GERMANY, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KOREA, MEXICO, NETHERLANDS, NEW Zealand, NORWAY, POLAND, PUERTO RICO, RUSSIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, UNITED KINGDOM. | Terminated | May 24, 2012 | 1354 | 639 | Other | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-1156-2009 | 51195 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | NurseMate with Physio Module; |  |  | February 13, 2009 | 2009 | St. Jude Medical | Installation problem: there is a risk that the monitor may not be adequately secured to the stand, resulting in the monitor falling off which may cause injury. | Nonconforming Material/Component | St. Jude Medical received a recall notice dated January 16, 2009 from its supplier, Smiths Medical PM, Inc. St. Jude Medical prepared and sent their own recall letter dated February 13, 2009 to affected customers. The recall letter described the issue and possible associated risks and provided instructions to customers. A St. Jude Medical representative will contact customers to facilitate replacement of the fasteners. Questions should be directed to your St. Jude Medical sales representative. |  | 4 | 4 units | Nationwide Distribution -- CA, NV, and FL. | Terminated | May 04, 2009 | 80 | 4 | Not_Computer
Z-1154-2009 | 51297 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Alm Prismatic Surgical Light 5004 SF; |  |  | February 23, 2009 | 2009 | MAQUET Inc. | The screws holding the Prismatic triple surgical lights may break due to dynamic multidirectional bending stress. The firm has received one complaint report. | Device Design | An "Urgent Device Correction" letter dated February 23, 2009 was mailed by FedEx Tracking to all affected consignees. The letter described the issue, potential hazard and recommended precaution. Customers were instructed to complete the enclosed "Customer Response Fax Form" and return it via fax (1-732-667-1903) to Maquet, Inc. Maquet, Inc. will schedule repair service of the affected units once the "Customer Response Fax Form" is received. Direct questions to your local Maquet, Inc. Representative. |  | 54 | 54 of all units, both article numbers | Nationwide Distribution. | Terminated | September 29, 2009 | 218 | 54 | Not_Computer
Z-1063-2009 | 50850 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | AxSOS Torque Limiting Screwdriver, T8 1.5 Nm 3.0mm Locking Set. REF 702749. This product is non sterile. |  |  | November 24, 2008 | 2008 | Stryker Howmedica Osteonics Corp. | Potential Breakage: Stryker Orthopaedics became aware that there is the potential for the tip of the AxSOS Screwdriver and the AxSOS Torque Limiter to break. | Other | "Urgent Product Recall" letters were sent to Stryker Orthopaedics Branches/Agencies, Chief of Orthopaedics, Hospital Risk Management and Surgeons on November 24, 2008 via Federal Express. The letters described the issue and provided the following instructions:1. Examine your inventory and determine if you have any of the affected product,2. Remove and quarantine affected product from your inventory,3. Complete the attached Product Recall Acknowledgement Form (and fax to Stryker Orthopaedics at 201-831-6069), and4. Return the affected product to your Stryker Orthopaedics Sales representative. Direct questions to your Stryker Orthopaedics Sales Representative or call Stryker Orthopaedics at 201-831-5000. |  | 415 | 415 units | Nationwide Distribution. | Terminated | April 06, 2010 | 498 | 882 | Not_Computer
Z-1173-2009 | 51292 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Medtronic STREAMLINE Bipolar Temporary Myocardial Pacing Lead, 6495, Sterile EO, REF 6495F, Medtronic, INC., Minneapolis, MN 55432 |  |  | February 06, 2009 | 2009 | Medtronic Inc | Medtronic is recalling one lot of STREAMLINE Bipolar Myocardial Temporary Pacing Leads, Model 6495, Lot JGD880205F. This lot of product contains incorrect inner pouch labeling, instead of 6495 it is labeled 6500. The outer carton, shelf box labeling, and instructions for use are correct, and the actual product shipped was consistent with the product ordered. No adverse patient effects have b | Process control | Consignees were sent on 2/6/09 a Medtronic " Urgent Medical Device Recall Notice" dated February 6, 2009. The letter was sent to Risk Mangers. The letter described the product, problem, and requested the removal and quarantine of the affected, unused product. They asked consignees to sign and return the attached recall certificate regarding the status of affected product remaining at their facility. Their Medtronic Sales Representative will contact them to arrange the return of any affected inventory or customers may follow the instructions provided on the attached certificate. |  | 612 | 612 | Nationwide Distribution -- including states of AK, AZ, CA, CO, CT, ID, IL, IN, KS, MI, MS, NE, NV, NH, NM, NY, NC, OH, OR, PA, TN, TX, VA, WA, and WI. | Terminated | May 07, 2011 | 820 | 612 | Not_Computer
Z-1046-2009 | 50816 | No | LOM |  |  | Microbiology | PMA | AxSYM HBsAg Reagent Pack; Abbott Laboratories, Abbott Park, IL 60064; AxSYM HBsAg is a microparticle enzyme immunoassay (MEIA) intended for the qualitative detection of hepatitis B surface antigen (HBsAg) in neonatal serum, and adult and pediatric serum (including serum collected in serum separator tubes) or plasma (collected in potassium EDTA, sodium citrate, sodium heparin, lithium heparin, or plasma separator tubes); list 9B01-20; |  |  | January 19, 2009 | 2009 | Abbott Laboratories | Defective caps: Customers have observed reagent bottles not opening appropriately, which could lead to a probe crash on the instrument. | Employee error | Consignees were sent a Product Correction-Immediate Action Required letter on 1/19/09. The letter informed users of the reported increased frequency of probe crashes when using the affected reagent pack lots and provided instructions on what actions the laboratory can take if they have a defective reagent pack or experience a probe crash with their system. Users need to inspect each reagent pack for broken or improperly sealed caps; manually open each bottle to ensure the flipper bar raises the reagent cap correctly; close each flipper bar prior to loading reagent pack on instrument; and discard any reagent packs that do not open properly. If a probe crash is observed prior to testing samples again, a probe crash recovery must be performed following the procedure in the Operations Manual. A copy of the communication needs to be provided to any laboratory that you may have forwarded affected lots to. Users were requested to return a customer reply letter acknowledging receipt and understanding of the letter. For additional information, contact Abbott Labs at 1-877-4ABBOTT. |  | 1540 | 1,540 kits | Worldwide Distribution -- USA including Puerto Rico and countries of Germany, Mexico, Hong Kong, and Japan. | Terminated | September 11, 2010 | 600 | 10123 | Not_Computer
Z-1198-2009 | 51178 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Low-Level Output Cable. |  |  | February 04, 2009 | 2009 | Datascope Corporation | Non-functioning cable: the Low-Level Output Cable when connected to a Datascope Intra-aortic Ballon Pump (IABP) and patient monitor does not provide a signal to the patient monitor. | Process design | An "Urgent Medical Device Recall" letter dated February 4, 2009 was delivered to affected consignees by Datascope Field Representatives. Direct questions to your local Sales/Service Representative or call the Datascope Corporation at 1-973-244-6314. |  | 20 | 20 | Nationwide Distribution -- states of IL and FL. | Terminated | August 11, 2009 | 188 | 20 | Not_Computer
Z-0989-2009 | 50872 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Deltec CozmoÂż Insulin Pump, Model 1700 containing 1800 software. |  |  | November 20, 2008 | 2008 | Smiths Medical MD, Inc. | Smiths Medical has become aware of a display irregularity with the Deltec CozmoÂż Insulin Pump, Model 1800 and Model 21-1700 containing Model 1800 software. There have been adverse events reported that the amount of Extended Bolus delivered was not accurately displayed by the Pump. The amount of extended Bolus displayed on the Pump Home Screen 2 and in the Bolus Summary Report is less than what w | Software design | Consignees were sent a Smiths Medical "Urgent Device Recall" letter dated November 20, 2008. The letter was addressed to Distributors, Clinicians, Patients and other Customers. The letter included a Description of the Problem, Advice on Action to be Taken by the User and requested the return of the Confirmation Form.In particular, users were given the option to check the "Replacement Pump" option on the Confirmation Form initiating the process for pump replacement with a recertified pump followed by a return of the defective product. The Confirmation Form can be faxed to 1-800-628-6322 in the U.S. or 1-651-628-7485 outside the U.S., completed at www.cozmore.com/notice or calling 1-800-501-5748 in the U.S. |  | 586 | 586 | Worldwide Distribution --- Including states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MI, MA, MD, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and DC and countries of AUSTRALIA, NEW ZEALAND, UNITED KINGDOM, IRELAND, CANADA, CHINA, GREECE, POLAND, SLOVENIA, SWEDEN, ISRAEL and BRAZIL. | Terminated | June 05, 2012 | 1293 | 17047 | Computer
Z-1212-2009 | 51080 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | 8 Fr 30cc Ultra 8 IAB, |  |  | February 02, 2009 | 2009 | Arrow International, Inc. | A fault in the connector of the pump tubing assembly may result in the volume setting on the pump defaulting to 2.5 ccs or 5 ccs (depending on IABP model) rather than the appropriate 30, 40, or 50 cc volume. | Process design | An "Urgent Medical Device Recall" letter dated February 2, 2009 was issued to consignees via US Certified Mail to notify users of the problem. Users were instructed to check their inventory for the affected products, quarantine, and return the them to Arrow International, Inc. Customers were also instructed to follow the "Mitigation Directions" in circumstances where alternate Intra-Aortic Balloon (IAB) catheters/adaptors or other forms of therapy are not available (until receipt of available product). Canadian accounts were notified on February 9, 2009 by courier and International accounts notified on 6 February 2009 and 10 February 2009 via Registered Post/UPS.Direct questions to your local Sales Representative or the Arrow International Customer Service at 1-800-523-8446 (US) or 1-919-361-4062 (outside US).Press Release issued April 13, 2009 is posted to firm's webiste:http://www.arrowintl.com/iab_recall.asp. |  | 4258 | 4,258 units | Worldwide Distribution: USA, Argentina, Australia, Taiwan, Singapore, New Zealand, India, Brazil, Belgium, Korea, Peru, Finland, and Denmark. | Terminated | July 06, 2010 | 519 | 61778 | Not_Computer
Z-1203-2009 | 50529 | No | JTC | Microtiter diluting and dispensing device. | Microbiology | Microbiology | 510(K) Exempt | 300 microliter CO-RE tips for use with Microlab STAR series instruments, Part Number 235902, 5,760 tips per case, manufactured by Hamilton Company, Reno, NV. |  |  | December 01, 2008 | 2008 | Hamilton Co | Some tips may be occluded and are unable to aspirate or dispense liquid. | Process control | An "Urgent: Product Recall Letter" dated December 1, 2008 was issued to consignees via certified mail and e-mail. The letter described the problem and potential effects of use. Consignees were instructed to complete the enclosed "Hamilton Company Product Recall Questionnaire" and return it via e-mail or fax (775-856-7259) to the Hamilton Company. After receipt of the completed questionnaire, the firm's Customer Service Department will contact consignees to provide Returned Material Authorization information for returning affected product. Direct questions to the Hamilton Company at 775-858-3000, ext. 236. |  | 28 | 28 cases | Nationwide Distribution -- MA, CA, MN, CT, and NC. | Terminated | July 21, 2009 | 232 | 28 | Not_Computer
Z-1251-2009 | 51212 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Modular Replacement System (MRS) Cemented Stems, Sterile, Catalog Numbers: 6485-3-000, 6485-3-008, 6485-3-009, 6485-3-010, 6485-3-018, 6485-3-019, 6485-3-300, 6485-3-308, 6485-3-309, 6485-3-310, 6485-3-318, and 6485-3-319. |  |  | January 30, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Instruction for use incorrect: There are no statements or warning on labels or Instructions for Use that 8, 9 and 10 mm diameter MRS Cemented Stems are indicated only for use in the GMRS Proximal Tibial and Distal Femoral Components. | Labeling False and Misleading | Urgent Product Recall letters, dated 1/30/2009, were sent to all Branches/agencies, Hospital Risk Management, Chief of Orthopaedics and Surgeons on January 30, 2009 via Federal Express. The letter asked consignees to examine their inventory and determine if they had any affected products, remove and quarantine affected product from their inventory, and complete the attached Product Recall Acknowledgment Form. The letter also stated that a Stryker Orthopaedics Sales representative would contact the consignee to reconcile and return the product . Consignees should contact their Stryker sales rep if they weren't not contacted by them. Questions should be directed to Colleen O'Meara at 201-831-5970. |  | 1161 | 1161 units | Nationwide Distribution | Terminated | April 08, 2010 | 433 | 1161 | Not_Computer
Z-1202-2009 | 51042 | No | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pointe Scientific Direct Bilirubin Reagent Set, Pointe Scientific Inc., Canton, MI. |  |  | February 02, 2009 | 2009 | Pointe Scientific, Inc. | The product performance fails due to precipitation of the R2 component. | Process control | Consignees were notified by "Urgent: Medical Device Recall" letter dated February 2, 2009 and mailed on February 6, 2009, and instructed to conduct the recall to the retail level if they had further distributed the product, to destroy any stocks on hand, and to complete and return the "Recall Return Response Form" via mail, e-mail, or fax (734-483-1592) to Point Scientific, Inc.Direct questions to the firm's Technical Support Department at 1-800-757-5313 or 1-734-487-8300. |  | 24 | 24 kits | Worldwide Distribution -- USA, Chile and Mexico. | Terminated | October 13, 2010 | 618 | 24 | Not_Computer
Z-1239-2009 | 51492 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Swan-Ganz Intellicath CCO/VIP Thermodilution Catheter with AMC Thromboshield (an Antimicrobial *Heparin Coating) For use with Vigilance Monitor. REF: 139HF75 |  |  | February 09, 2009 | 2009 | Edwards Lifesciences, Llc | Cardiac output measurements were unavailable and the following message is displayed in the Vigilance monitor: "Fault. CCO Catheter Verification. Use Bolus Mode". | Software design | Edwards initiated the worldwide recall via Urgent Customer Notification letters dated February 9, 2009 and are being sent by Federal Express starting on February 10, 2009. These letters are requesting that customers cease using Swan-Ganz" Continuous Cardiac Output (CCO) Thermodilution Catheter products from the indicated lots and that they return all unused product. Edwards Lifesciences requested that the customer provide written verification of stock and identify any unused product to be returned. Questions should be directed to Technical Service at 800-822-9837. |  | 107 | 107 units | Worldwide Distribution -- AL, AR, AZ, CA, CT, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, MT, NC, NE, NJ, NM, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, Hawaii, and Puerto Rico and countries of Canada, Austria, Belgium, Switzerland, Germany, Denmark, Spain, Finland, France, UK, Greece, Ireland, Israel, Italy, Kuwait, Netherlands, Norway, Portugal, Sweden, Colombia, Australia, and Singapore. | Terminated | April 11, 2011 | 791 | 10754 | Not_Computer
Z-1250-2009 | 51499 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Diamondback 360, CSI ViperWire Orbital Atherectomy Guide Wire, REF: GW14335FRM. Sterilized with Ethylene Oxide. Manufactured for Cardiovascular Systems, Inc., St. Paul, MN 55112, USA. The ViperWire atherectomy guide wire has a shaft diameter of 0.014" and is constructed of stainless steel with a smooth finish and is silicone coated. The spring tip configuration is atraumatic, radiopaque, and can be bent to form a steerable system. |  |  | March 16, 2009 | 2009 | Cardiovascular Systems, Inc. | CSI is recalling 10 devices due to a mislabeling of the shelf life as shown on the device labeling. The labeling shown on the affected devices represents a 24 month shelf life rather than a 6 month shelf life that should have been indicated. | Incorrect or no expiration date | Customer letters were FedEx'd to CSI Sales Professionals on 03/16/2009, who delivered the letters and helped coordinate the removals of affected product. The letter identified the product, provided a description of the issue and provided instructions to immediately remove from use and return the affected guidewires to CSI. A Customer Acknowledgement Form was asked to be completed and returned with the devices. |  | 10 | 10 devices | Nationwide Distribution | Terminated | April 20, 2012 | 1131 | 10 | Not_Computer
Z-1207-2009 | 51316 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Biomet Marrowstim Concentration Kit w/single Disposable Separator, Bone Marrow Needle, and 30 ml ACD-A. For use with BBI centrifuges only, sterile, Biomet Biologics, Warsaw, IN; REF 800-0611A and 800-0613A. |  |  | March 05, 2009 | 2009 | Biomet, Inc. | A 30 ml tube is packaged with the kit instead of the labeled 60 ml tube. | Process control | Consignees were notified by "URGENT MEDICAL DEVICE REMOVAL NOTICE " dated March 5, 2009 and instructed to locate and return the recalled products. Direct questions about this recall to Biomet Orthopedics at 1-800-348-9500 or 574-372-3983, Monday through Friday, 8AM to 5PM. |  | 27 | 27 total kits of both products | Worldwide Distribution -- USA including states of Massachusetts and Minnesota and countries of Germany and Italy. | Terminated | July 07, 2009 | 124 | 27 | Not_Computer
Z-0987-2009 | 50829 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CADD Medication Cassette Reservoir 100 ml with Clamp and Female Leur, Nonvented stopper included REF 21-7002-24, Sterile EO. |  |  | December 09, 2008 | 2008 | Smiths Medical MD, Inc. | Smiths Medical noticed an increased trend in leakage associated with 100 ml CADD Medication Cassette Reservoirs Reorder Number 21-7002-24 Lot Numbers 213X18 and 214X18. If a leak occurs in a CADD medication Cassette Reservoir, it may allow spillage of medication, contamination of the fluid path, and/or air to enter the Reservoir. Smiths Medical has not received any report of patient injuries due | Process control | An "Urgent Medical Device Recall" letter was sent via electronic mail to consignees on December 8, 2008. The letter was addressed to Distributors, Clinicians and other Customers. The letter described the problem and affected product and listed the "Advise on Action to be Taken by the User". Smiths Medical MD, Inc. requested the user to complete and return the "Urgent Medical Device Recall Confirmation Fax Return Form" and fax to the firm at 1-651-628-7485 or e-mail to info.md@smiths-medical.com (in US) or international.customer.service@smiths-medical.com (outside US). Direct questions to the Smiths Medical MD, Inc. Customer Service Department at 1-800-426-2448, Option 1. |  | 214416 | 17,868 cassettes (1,489 boxes 12 cassettes/box) | Worldwide Distribution USA (state of PA) and countries of El Salvador, Australia, France, Finland, Sweden, and Germany. | Terminated | June 07, 2012 | 1276 | 214416 | Not_Computer
Z-1004-2009 | 50543 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Reflection, Microstable Acetabular Liner, REF 71740250, 32mm ID, 50-52 OD, 1 each, sterile EO. Component is a liner used to hold the femoral head used in total hip replacement arthroplasty procedures. |  |  | November 21, 2008 | 2008 | Smith &amp; Nephew Inc | Liners measuring 28 mm ID and were mismarked as 32 mm. | Process control | All affected Smith &amp; Nephew International Distributors were notified via e-mail dated November 24, 2008. In a letter titled "Urgent: Product Recall Notice - Immediate Action Required", distributors were instructed to immediately notify their accounts concerning the recall. All affected product is to be located and returned to the firm immediately.Direct questions to Smith &amp; Nephew, Inc. by calling 1-901-399-5441. |  | 19 | 19 units | International Distribution to Canada, France, Germany and Switzerland. | Terminated | November 21, 2011 | 1095 | 19 | Not_Computer
Z-1003-2009 | 50918 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Infinia HawkeyeÂż 4, Model Numbers: 5178728, ASM500060, H3000WC and H3000WD. |  |  | June 02, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has recently became aware of excessive pressure applied by the collimator pressure sensitive cover of Infinia (Infinia, Infinia HawkeyeÂż, Infinia HawkeyeÂż 4) that may impact patient safety. It was reported at a customer site that the pressure used to activate the Pressure Sensitive Device (PSD) should be lowered and that the body contouring software requires correction to prevent po | Device Design | Consignees were sent from June 2008 to February 2009 a GE "Product Safety Notification" letter dated April 2008. The letter was addressed to Hospital Administrators and Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information.A GE Healthcare service representative will contact each consignee to schedule the installation of the Pressure Sensitive Device (PSD). Please contact your local sales or service representative for further information regarding this recall. Further information can be found by calling 1-800-437-1171 in the US or Canada or 1-262-896-2890 if outside the US. |  | 288 | 122 US, 166 OUS | Worldwide Distribution US (and Puerto Rico) including states of AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV and WI and countries of Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Czech, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Hungry, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Malaysia, Martinique, Mexico, Netherlands, New Zealand, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Republic of Georgia, Russia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine and the United Kingdom. | Terminated | May 24, 2012 | 1452 | 1256 | Computer
Z-0995-2009 | 47442 | No | LYJ |  |  | Neurology | PMA | Cyberonics VNS Therapy Demipulse Duo Generator, Model 104; an implantable, multiprogrammable, vagus nerve stimulator. |  |  | February 07, 2008 | 2008 | Cyberonics, Inc | Reset/disabling of the VNS Therapy Demipulse Generator and Demipulse Duo Generator due to magnet interference, resulting in the loss of stimulation. | Software design | Cyberonics notified consignees of recall via letter on 02/07/08. Consignees were notified of reset issue in which during a strong magnetic field, the software may report an error even though no error occurred. This error causes the generator to reset, which disables the generator (i.e., no stimulation delivered).Physicians should inform any potential and currently implanted patients of the possibility of this event. Patients should continue with their regularly scheduled visits as determined by their physician.If this event is suspected, physicians should interrogate the device as described in the Physicians Manual to confirm the generator has been reset and Cyberonics should be contacted for any additional instruction. |  | 108 | 108 units. | Worldwide Distribution - USA, Belgium, Germany, Norway, Spain, Sweden, the Netherlands, and United Kingdom. | Terminated | February 28, 2011 | 1117 | 795 | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-1205-2009 | 51215 | No | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | PLAC Test Calibrator Kit, Five-level liquid calibrator set for use with the diaDexus PLAC Test Reagent Kit, assembled in a kit box with an outer label, Catalog numbers 90108 and 10-0108, manufactured by diaDexus inc., South San Francisco, CA. |  |  | February 09, 2009 | 2009 | DiaDexus, Inc | Premature Expiration-- the product is projected to fall short of expected shelf life, and may result in inaccurate results. | Process control | Distributors and Direct accounts were notified by phone followed by Notification letter and certificate of destruction via FedEx. International customers were notified via e-mail with PDF copies of notices, to be followed by hard copies via FedEx. Customers are requested to record the number of kits (full and partial ) they had on the Certificate of Destruction and to dispose of that number of kits with two people signing the form, and then to fax that form to diaDexus. |  | 105 | 105 kits | Worldwide Distribution | Terminated | January 05, 2010 | 330 | 105 | Not_Computer
Z-1069-2009 | 50530 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Hematology Reagent Syringe, List number 08H49-02, used on the CELL-DYN Sapphire Automated Hematology Analyzer, List Number 08H00-01, manufactured by Abbott Diagnostics Division, Santa Clara, CA. |  |  | December 10, 2008 | 2008 | Abbott Laboratories | Some syringes were improperly assembled and are therefore not usable, due to an incorrect nut on the syringe shell. | Process control | Recall initiated on December 10, 2008. A product recall letter and reply form were cent to CELL-DYN Sapphire customers who received syringes between September 16, 2008 and October 28, 2008 to US customers, and to November 5, 2008 for Outside-of-US customers. |  | 30 | 30 syringes | Worldwide Distribution -- USA, Australia, Canada, and Germany. | Terminated | December 27, 2010 | 747 | 30 | Not_Computer
Z-1200-2009 | 51044 | No | LBR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific Liquid Auto HDL Cholesterol Reagent Set, Pointe Scientific, Inc., Canton, MI, Catalog Numbers: 10-H7545-80, 12-H7545-162, 3-H7545-L, 7-H7545-1000, 7-H7545-320, 7-H7545-40, 7-H7545-80, H7545-1000, H7545-320, H7545-40, H7545-80, HH945-240, and HH945-480. |  |  | January 29, 2009 | 2009 | Pointe Scientific, Inc. | The R1 reagent has microbial contamination. | Nonconforming Material/Component | Consignees were notified by "Urgent: Medical Device Recall" letter dated January 29, 2009, to conduct the recall to the retail level, to examine inventory and dispose of affected stocks on hand, and to complete the "Recall Return Response Form". Consignees were instructed to return their completed forms via mail, fax (1-734-483-1592) or e-mail to Pointe Scientific, Inc. (PS). Direct questions about the recall to PS Technical Support Department at 1-800-757-5313 or 1-734-487-8300. |  | 590 | 497 kits, 93L | Worldwide Distribution -- US, Chile, Dominican Republic, Egypt, Ghana, Greece, India, Israel, Italy, Jamaica-West Indies, Korea, Malaysia, Mexico, Philippines, Poland, Romania, Serbia, Trinidad-West Indies, Turkey and Venezuela. | Terminated | October 20, 2010 | 629 | 1471 | Not_Computer
Z-1123-2009 | 51276 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Operon Electrical Operating Room Table. |  |  | February 02, 2009 | 2009 | Berchtold Corp. | Difficulty in pressing the button needed to release a bed part resulted in an injury to a caregiver's finger which became stuck in metal latch under the gas cylinder of the bed piece. | Device Design | An official notification letter dated February 3, 2009 was hand-carried by a local Berchtold Corporation Technical Service representative to each account. The letter indicated that the Berchtold Corporation Service Technician will complete recommended upgrades. Direct questions concerning this recall to your BERCHTOLD Technician Service at 1-800-243-5135 and press the Number 2. |  | 453 | 453 units | Worldwide Distribution -- US and Canada. | Terminated | June 04, 2009 | 122 | 453 | Not_Computer
Z-1197-2009 | 49019 | No | LTJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Beckman Coulter Access Immunoassay System Hybritech PSA Calibrators, Product Number: 37205. |  |  | January 08, 2008 | 2008 | Beckman Coulter Inc | Beckman Coulter incorrectly shipped Access Hybritech PSA calibrators containing dual calibration cards (Hybritech and WHO) to some Access Hybritech PSA Customers. The WHO calibration for PSA is not currently approved in the United States. | Packaging process control | An "Urgent: Product Corrective Action" letter dated January 10, 2008 was issued via fax or mail to customers. The letter described the issue and instructed consignees to discontinue use and discard the affected product. Consignees were also asked to complete and return a "Fax-Back Form" via fax (1-952-368-7610) or mail to Beckman Coulter, Inc. within 10 days of receiving the notification letter.Further questions should be addressed to Beckman Coulter Customer Technical Support at 1-800-854-3633. |  | 123 | 123 Units | Nationwide Distribution -- Including states of AL, AZ, CA, FL, GA, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NY, OK, OR, PA, SC, TN, TX, VA, WI and WV. | Terminated | March 15, 2012 | 1528 | 123 | Not_Computer
Z-1266-2009 | 51563 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging model number 10014063. Affected systems are V30A or earlier with an NFS connected archive and V30A systems where image data was stored to LTS using software version older than V30A. |  |  | February 09, 2009 | 2009 | Siemens Medical Solutions USA, Inc | An issue concerning archived image data that is sporadically saved with corrupted data may occur. This issue may occur when syngo Imaging is configured to archive image data on a network file system (NFS) connected via a connected archive (Mount Point Solution). | Software design | A Customer Safety Advisory Notice, dated 02/09/2009 was sent to affected customers via Update Instruction IM013/09S. The letter informs customers of when the malfunction occurs and the potential issue, provides instructions to avoid its occurrence, and how the issue will be resolved. Customers are to ensure that the Safety Advisory is placed in the system's instructions for use and contact their SIEMENS Uptime Service Senter if they have any questions. |  | 20 | 20 units | Nationwide Distribution -- including states of AL, CA, CT, FL, LA, MI, MO, NC, NJ, OH, PA, TN, and WI. | Terminated | September 23, 2009 | 226 | 20 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-1234-2009 | 51469 | No | ELR | Root canal post. | Dental | Dental | 510(K) Exempt | Dentsply, MAILLEFER, CYTCO-K TITANIUM POST SYSTEM, dental implant components, Model numbers (SKU): C106KL0050A20 (12 posts) and C117K00000020 (4 posts), Swiss Made, Manufactured for DENTSPLY Maillefer, Johnson City, TN 37604 |  |  | March 20, 2009 | 2009 | Dentsply Worldwide Headqrt | The product is manufactured with a left-hand thread instead of a right hand thread. This error could cause tooth breakage if the posts were removed without knowing that the threads were reversed. | Process control | Dentsply notified their distributors and dentists of the recall by email sent on 03/20/2009 by Urgent Recall Notice letter explaining the problem and requesting return of the product from their inventory and from their consignees. The letter requests faxback response (enclosed with the recall notification) within five (5) working days. The letter included a response page to ensure the receipt of the notice. A second follow up letter will be sent to all that do not respond to the first letter.The firm identified another distributor within the company, DENTSPLY Tulsa Specialty Products, who also sent out identical recall letters to those they distributed the product on 03/25/09. |  | 1588 | 1588 posts | Worldwide distribution: USA, Canada, Germany, France, Holland, Switzerland, The Netherlands, Italy, Kuwait, Poland, Argentina, Greece, Russia, Czech Republic, Belgium, Turkey, Hungary, India, and Colombia. | Terminated | February 12, 2010 | 329 | 1588 | Not_Computer
Z-1208-2009 | 51431 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Screwdriver for 3.5mm Screws, DARCO system, instrument for use in orthopaedic implant surgery. REF: DC4261, Non-sterile, Rx only, 1 each, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002 |  |  | February 19, 2009 | 2009 | Wright Medical Technology Inc | The package insert lists 3 sterilization procedures, 2 of which are not validated. Pre-vacuum steam sterilization is validated. | Process design | Wright Medical notified their distributors of the recall on 03/11/2009 by letter explaining the problem and requesting return of the product from their inventory and from their consignees. The letter included a "fax back" response page to ensure the receipt of the notice. The hospitals and surgeons were notified of the recall also on 03/11/2009. All letters were delivered by Fed Ex. |  | 682 | 682 devices | Worldwide distribution: USA, Canada, Germany, France, Belgium, the Netherlands, Italy, Finland, Austria, Denmark, Norway, Turkey, South Africa, United Kingdom, and Australia. | Terminated | February 03, 2010 | 349 | 682 | Not_Computer
Z-1235-2009 | 51423 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging, picture archiving and communications system (PACS), Model number 10 014 063 |  |  | January 12, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Incorrect and/or incomplete data: Several potential safety issues while working with the syngo Image product including, but not limited to, data loss, errors in patient data due to different patient IDs, delays in image availability, inaccurate results during image measurements, incorrect review or diagnosis of displayed or printed images, measurement inaccuracies from the host modality and image | Software design | syngo Imaging has issued an Addendum to the User Manual, which was distributed to affected customers via Update Instructions IM118/08/S in November, 2008. Two safety related hotfixes (VB20G_HF4 and VB30A_HF26) will be released to those affected only after Update Instructions IM118/08/S has been performed. The addendum informs the user about safety relevant notes which pertain to syngo Imaging versions VB20G and VB30A. Update is a prerequisite for the following safety related hotfixes: VB20G_HF4 and VB30A_HF26. |  | 55 | 55 units | Nationwide Distribution. | Terminated | September 23, 2009 | 254 | 55 | Software | Output/Calculation | Software update | Software Update
Z-1187-2009 | 51239 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump2, 400 mL PainPump with 5.0 inch ExFen 2 Site Catheter Set, 5 per box, sterile, Stryker, Kalamazoo, MI; REF 540-50. |  |  | February 11, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | An electronic failure during preparation of the unit, or during patient use, may cause the pump to shut down. | Process design | A "Medical Device Recall Notification" letter dated February 11, 2009 was issued to consignees via phone, mail and/or email. The letter described the issue and instructed customers to discontinue use of the recalled product, destroy it according to instructions provided with the notification and return the "Business Reply Form" by fax to Stryker Instruments Regulatory Department, 1-866-521-2762.Further questions should be addressed to Stryker Instruments at 1-800-800-4236 extension 4086. |  | 24675 | 24,675 pumps | Worldwide Distribution -- USA including states of AK, AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, WA, WI, WV and WY and countries of Argentina, Australia, Brazil, Canada, England, Singapore and South Africa. | Terminated | December 01, 2009 | 293 | 24675 | Not_Computer
Z-1319-2009 | 51439 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Tri-State Centurion Endoscopy Tray, sterile, 10 per case, Tri-State Hospital Supply Corp., Howell, MI; Reorder MNS4635. |  |  | January 14, 2009 | 2009 | Tri-State Hospital Supply Corporation | Compromised sterility, as the package seals may be inadequate. | Packaging process control | A letter titled "Urgent Product Recall Notice" was sent by Tri-State Hospital Supply Corporation to all consignees via certified mail titled dated January 14, 2009. Consignees were instructed to remove the product from inventory, to inform the firm of amount on hand by sending the Accountability Information Form to the firm via fax (1-517-546-3356) or email and to hold the product for retrieval. Further questions should be addressed to Tri-State Hospital Supply Corp. at 1-517-546-5400 extension 1135. |  | 10 | 10 | Nationwide including states of NC, FL, TX, NJ, IL, MA, GA and IA. | Terminated | December 24, 2009 | 344 | 550 | Not_Computer
Z-1308-2009 | 51708 | No | OHS | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Lumi 8 Intense LED Light Facial System, Sybaritic, Inc.,9220 James Ave South Minneapolis, Minnesota 55431., Mode of Operation: Operation with Intermittent/continuous load, Visible and Invisible LED Radiation, Class 3R LED Product, Refer to IEC 60825-2001 for further information., Input: 100 -230 V AC 50/60 HZ Power Input: 1509 V A....Protective Eyewear required Wavelength 590 &amp; 630 nm. |  |  | June 18, 2008 | 2008 | Sybaritic, Inc | Sybaritic stated that no 510(k) has been filed for the device (Lumi 8 Intense LED Light Facial System).No injuries or illness reported. | PMA | Letter sent to consignee on June 18, 2008 confirming a phone conversation on June 18 regarding the required Field Correction of the device. The letter also described the changes to be performed by the local technician to be in compliance with the FDA. The customer was asked to allow the technician access to their LUMI 8 system; allow technician to replace the computer chips and owner's manual; give current LUMI 8 owner's manual to technician; and sign the work order provided by the technician. Questions should be directed to Fred Berg at 800-445-8418 or fberg@sybaritic.com. |  | 1 | 1 | USA: MI | Terminated | July 05, 2012 | 1478 | 1 | Not_Computer
Z-1236-2009 | 51319 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cerner Bridge Medical Transfusion Administration and Specimen Collections, Versions 3.4.0, 3.4.1 and 3.4.2. |  |  | January 19, 2009 | 2009 | Cerner Corp | Specimen labels printing incorrect patient demographics. | Software design | Cerner Corporation issued a "Priority Review Flash" dated January 15, 2009 via e-mail to consignees using the software. The letter described the violative nature of the software and explained the proper navigation controls needed to prevent further issues.Further questions should be addressed to Bridge Client Support at 1-800-786-0239. |  | 4 | 4 | Nationwide distribution to states of OH, NC, MO and MI. | Terminated | March 22, 2012 | 1158 | 4 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-1237-2009 | 51432 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Patient-Controlled Analgesia (PCA) Module (formerly Medley PCA Module) Model 8120 used with the Alaris PC Unit with software versions 8 through 9.1. |  |  | March 04, 2009 | 2009 | Cardinal Health 303 dba Cardinal Health | The Alaris PCA module can potentially infuse above or below the intended infusion dose. | Software design | The field action was initiated on March 11, 2009 with Cardinal Health sending by registered return receipt mail a Safety Alert to all customers using the Alaris PC Unit with software versions 8 through 9.1, which encompasses the subset of those customers that use the Alaris PCA Module. The Safety Alert was addressed to the Directors of Nursing, Pharmacy, Risk Management, and Biomedical Engineering at each hospital facility. If a customer did not respond to the Safety Alert, then Cardinal Heath notified that customer by telephone to verify receipt of the Safety Alert.This Safety Alert Informed the customers about a potential risk, and required them to: (1) Determine whether you use version 8 through 9.1 of the Alaris PC Unit. (2)Forward this Safety Alert to all Data Set owners (3) To immediately identify the potential risks review enclosed FAQs and Instructions for Data Set Review. (4)Continue to follow the Institute for Safe Medication Practices (ISMP) and the Joint Commission recommendation for safe medication administration - two nurses checking to ensure that the correct prescription is being entered into the PCA pump when first ordered and when there are any changes in the prescription (paying close attention to dosing units and infusion modes). The customers were also instructed to call our Support Center at 1-888-562-6018 with any questions. An updated Customer Recall Notification was forwarded by registered return receipt mail by Cardinal Health on April 15,2009. This notice reinforced the information in the first notice in addition to informing the customers of the hazard of an over-infusion or under-infusion of medication. Also, customers were given instruction for reporting any adverse reactions experienced with the use of the affected product and/or quality problems via the FDA's MedWatch Program. |  | 164000 | Approximately 164,000 units | Worldwide Distribution -- USA, Canada, Germany, Malaysia, and Singapore. | Terminated | June 26, 2012 | 1210 | 164000 | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-1317-2009 | 51562 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom TotalCare Bed, Model P1900. |  |  | March 09, 2009 | 2009 | Hill-Rom, Inc. | An e-ring may not have been installed during production, which could result in an inability to latch the siderail. | Process control | Consignees were notified by an Urgent Medical Device Correction letter dated 3/9/09 and were instructed that a Hill-Rom technician will visit each consignee to inspect and to correct the beds, if necessary. The firm asks that consignees locate and test their beds to make sure the siderails are latching correctly; forward a copy of the letter to users that were loaned or purchased the device and notify Hill-Rom of their new locations or owners; and if affected beds were disposed of or are no longer in use, to inform the firm of their serial numbers. If there are questions customers should contact their Hill-Rom Technical Support for their geographic region. |  | 53 | 53 beds | Worldwide Distribution -- USA and Canada. | Terminated | October 13, 2010 | 583 | 53 | Not_Computer
Z-1232-2009 | 51253 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CIC Pro" Software Versions 5.0.3, 5.0.6 ,5.0.7, 5.0.8 and 4.1. |  |  | February 27, 2009 | 2009 | GE Medical Systems Information Tecnology | GE Healthcare is aware of potential network time issues associated with the time change aspects of certain CARESCAPE CIC Pro" software versions. These issues may impact patient safety. No reported patient injuries or illnesses for the use of this device in conjunction with this device correction. | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated February 27, 2009. The letter was addressed to the Chief of Nursing/Chief of Biomedical Engineering and Risk Manager/Healthcare Administrator. The letter describes the Safety Issue, Affected Product Details, Safety instructions, Product Correction, and Contact Information.Direct questions to your local GE Healthcare Field Service Representative or the GE Healthcare Technologies Support Center by calling 1-800-558-7044 (Option 2, Technical Support). |  | 12737 | 12,737 | Worldwide Distribution including US (states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, DC, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, PR, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WY, WI, and WV) and countries of UNITED ARAB EMIRATES, AUSTRIA, AUSTRALIA, BOSNIA AND HERZEGOVINA, BANGLADESH, BELGIUM, BRAZIL, BAHAMAS, CANADA, SWITZERLAND, CHILE, CHINA, COLOMBIA, COSTA RICA, CUBA, CYPRUS, CZECH REPUBLIC, GERMANY, ESTONIA, EGYPT, SPAIN, FINLAND, FRANCE, GHANA, GREECE, HONG KONG, HUNGARY, HONDURAS, INDONESIA, IRELAND, ISRAEL, INDIA, IRAQ, BRITISH INDIAN OCEAN TERRITORY, ITALY, JORDAN, JAPAN, SOUTH KOREA, KUWAIT, LEBANON, LITHUANIA, MALTA, MEXICO, MALAYSIA, NIGERIA, NETHERLANDS, NORWAY, NEPAL, NEW ZEALAND, PANAMA, PERU, PHILIPPINES, PAKISTAN, POLAND, PORTUGAL, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SUDAN, SW DEN, SINGAPORE, SLOVENIA, SLOVAK REPUBLIC, SYRIA, THAILAND, TURKEY, TRINIAD, TOBAGO, TAIWAN, UKRAINE, UNITED KINGDOM, VENEZUELA, YEMEN, and SOUTH AFRICA. | Terminated | September 15, 2014 | 2026 | 12737 | Software | Device Operation | Correction | Repair
Z-1244-2009 | 51421 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | TEVADAPTOR Closed Drug Reconstitution and Transfer System, a closed system used to prepare drugs, including cytotoxic drugs, for intravenous infusion or injection. Model number is MG412110. The device is a single use device. |  |  | December 18, 2008 | 2008 | Migada Inc. | Syringe adaptor is disintegrating when attached to the end of the syringe. | Component design/selection | A Field Safety Notice, dated 12/8/2008, was sent to all TEVADAPTOR customers advising them not to use drugs containing N,N-Dimethylacetamide. Teva Medical is presently revising the instruction leaflet supplied with each TEVADAPTOR to include a warning against the use of drugs containing N,N-Dimethylacetamide. Teva Medical will introduce this revised leaflet to current production as soon as practicable. In the meantime, Teva Medical will advise all users of the potential problem by means of the Urgent Field Safety Notice. Customers are to ensure that the safety notice is passed to all those who need to be aware within their organizations. Questions should be directed to their local contacts at B Braun, USA. |  | 900000 | 900,000 units | Nationwide Distribution | Terminated | July 27, 2009 | 221 | 900000 | Not_Computer
Z-1254-2009 | 51560 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM Sireskop SD Model numbers 8890415 and 8890407. The product is intended for use as radiographic and fluoroscopic imaging device. |  |  | January 08, 2009 | 2009 | Siemens Medical Solutions USA, Inc | During interventional use and applications, the possibility exists for liquids to enter the systems. | Equipment maintenance | A Customer Safety Advisory Notice dated January 8, 2009 was issued to affected customers via Update Instructions AX069/08/S. This letter informed customers of the potential issues and provided instructions to avoid its occurrence. The firm has released an Addendum to the User Manual warning of the potential hazards if liquids penetrate the systems during interventional use and applications.Direct questions about this recall to Siemens Medical Solutions USA, Inc. by calling 1-610-219-6300. |  | 240 | 240 units | Nationwide Distribution. | Terminated | September 01, 2009 | 236 | 909 | Not_Computer
Z-1328-2009 | 49127 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stat Profile Critical Care Xpress (CCX and CCX+), CCX Creatinine Membrane Kit for CCX analyzer |  |  | August 11, 2008 | 2008 | Nova Biomedical Corporation | Erroneous, low creatinine results on CCX analyzer | Process control | Initial:USA and Canada - Customer Advisory Notice notification by phone contact with key customer contact facilitated using a Field Correction Script. This contact will be documented within the Nova Biomedical Technical Support database (Pivotal) and on a Field Correction Checklist. All affected inventory will be replaced at customer sites.International - Customer Advisory Notice and Faxback Form being communicated and sent to Nova Biomedical Subsidiaries and Distributors by three methods of delivery (Email, Fax and Phone Call).Follow-up:USA and Canada - Sending the Customer Advisory Notice by email and/or fax to the key customer contact from the initial phone contact.International - Sending the Customer Advisory Notice by email and/or fax to the key contact. |  | 1230 | 1230 | Worldwide Distribution -- USA, Canada, Germany, France, Czech Republic, Taiwan, Japan, Thailand, Italy, Hungary, India, UK, Ireland, Venezuela, Brazil, and Hong Kong. | Terminated | July 22, 2011 | 1075 | 1230 | Not_Computer
Z-1233-2009 | 51494 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Siemens Mammomat Novation DR, full field digital mammography system. Model Number 06646900. |  |  | January 23, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Image may appear dark, requiring repeat image acquisition. | Software design | A Customer Safety Advisory Notice dated December 16, 2008 was issued to affected customers via Update Instructions SP040/08/S. The letter described the potential issue and provided instructions to avoid its occurrence. A software update to correct the issue is currently being developed and will be installed on affected systems when it becomes available. |  | 497 | 497 units | Nationwide Distribution. | Terminated | March 19, 2013 | 1516 | 497 | Software | Display/Image | Software update | Software Update
Z-1391-2009 | 51925 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Triage Cardiac Panel, Biosite Catalog Number 97000HS, 25 individually pouched devices in a labeled kit box. |  |  | April 27, 2009 | 2009 | Biosite Inc | The recall was initiated because the use of the affected lot may lead to reporting falsely low values for troponin I and myoglobin, possibly resulting in missed or incorrect diagnosis. | Process control | The recall was initiated with "Urgent Medical Device Recall" notices, dated 4/28/2009, with attached "Customer/Distributor Verification Form" being sent to the affected customers via fax, e-mail or direct mail on 04/28/2009. Customers were informed of the reason for the recall and instructed to immediately discontinue all use of the listed products and discard all affected product in accordance with your their regulations. The customers were give two options; "Option 1: Complete and FAX the enclosed Verification Form within 10 days to confirm your receipt of this notice, to indicate the number of kits discarded from your inventory and product replacement will be sent to you. Option 2: Contact inverness Medical Technical Service at 1-877-441-7440 and select option 2 from the phone prompt. A representative will confirm the number of kits discarded from your inventory and arrange for replacement product to be sent to you."Customers were also provided telephone, fax and e-mail contact information for Biosite should they have any questions about the information contained in the recall notification. |  | 401 | 401 kits | Nationwide Distribution | Terminated | July 06, 2010 | 435 | 401 | Not_Computer
Z-1209-2009 | 51206 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | KyphX Xpander Inflatable Bone Tamp, size 15/3, Catalog Number K09A, Manufactured by Kyphon Sarl, Switzerland. For use in spinal surgery. |  |  | February 04, 2009 | 2009 | Medtronic Spine LLC, formerly Kyphon Inc | Product not sterile: 10 units were released for distribution without undergoing sterilization process. | Process change control | Sales representatives were contacted via e-mail on February 4, 2009, and instructed to contact hospitals and inform them of the issue and sequester affected product. They were further instructed to immediately travel to the hospitals and obtain affected product for return to Medtronic LLC. A Customer Communication Notice URGENT NOTICE: DEVICE REMOVAL dated February 9, 2009, was sent to the two consignees. |  | 10 | 10 units affected. | Product was distributed to two hospitals in OH and FL. | Terminated | July 14, 2009 | 160 | 10 | Not_Computer
Z-1252-2009 | 51473 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Servo Ventilator 300 (Model SV300) and |  |  | March 23, 2009 | 2009 | MAQUET Inc. | The expiratory valve in the SV 300 ventilator may stay in the open position if the printed circuit board fails. | Device Design | An "Urgent: Device Correction" letter dated March 23, 2009 was sent by Federal Express to customers. The letter described the affected product, problem, indications, potential hazard and actions for users. Customers were instructed to complete the "Device Correction -- Customer Fax Response" form and return via fax (732-667-1903) to Maquet Inc. RA/QA. Direct questions to your local Maquet Inc. representative. |  | 7847 | 7,847 units | Nationwide Distribution. | Terminated | November 16, 2009 | 238 | 7847 | Hardware | Device Operation | Instructions | Safety Notice/Insructions
Z-1309-2009 | 51049 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Coated VICRYL (Polyglactin 910) Suture, Ethicon Inc., A Johnson &amp; Johnson Company, Somerville, NJ 08876 |  |  | February 02, 2009 | 2009 | Ethicon, Inc. | Sterility Issue: The firm decided to remove the product because it cannot assure the sterility of the products due to a failure in the sterilization process. | Process control | Urgent Voluntary Product Recall notification letters were sent via UPS Next Day Air with receipt on February 2, 2009. Letters were sent to customers, Directors of Surgical Services, Director of Risk Management, Materials Manager and Distributors. The letters requested that customers immediately discontinue use of the products with lot numbers listed. Customers are to identify and return only the boxes identified in the attached tables. The affected product should be returned to Ethicon via the pre-paid authorized shipping label. Customers that are not returning affected product should mail the enclosed business reply card to Ethicon. Customers should direct questions regarding this recall to 1-888-202-3694 or their Ethicon sales representative.5/13/09: It was discovered that a lot of PDS II sutures that had been previously identified as affected but not distributed, were mistakenly distributed. Revised letters including this number were sent to 25 customers on 5/13/09. |  | 1758360 | 1,758,360 total | Worldwide Distribution -- USA, Argentina, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Greece, Hong Kong, Israel, Japan, Korea, Luxemburg, Mexico, New Zealand, Norway, Panama, Puerto Rico, Sweden, Taiwan, United Kingdom, Uruguay, and Venezuela. | Terminated | October 21, 2009 | 261 | 1758360 | Not_Computer
Z-1262-2009 | 51320 | No | HTQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Nexgen Complete Knee Solution Fluted Stemmed Tibial Broach, Size 7, 8, nonsterile; REF 5997-61-07. |  |  | March 13, 2009 | 2009 | Zimmer Inc. | The surface finish inside of the broach is not to specification and the instrument may corrode if cleaned and sterilized according to instructions. | Finished device change control | Distributors were notified by an "Urgent: Device Recall" letter dated March 13, 2009. The letter described the affected product, reason for recall, health risks and actions for distributors and user facilities. Customers were instructed to quarantine affected instruments and return them to the Zimmer Product Service Department. Distributors were also instructed to provide the name and address of hospitals where the instruments were used so that a copy of the recall letter can be forwarded to the user facilities.Direct questions about this recall to Zimmer, Inc. by calling 800-613-6131. |  | 37 | 37 | Worldwide Distribution -- US, Australia, and Germany. | Terminated | December 22, 2009 | 284 | 37 | Not_Computer
Z-1327-2009 | 47475 | No | JMO | Iron-binding capacity test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Diagnostics Chemical Limited Iron Assay |  |  | March 14, 2008 | 2008 | Diagnostic Chemicals, Ltd. | Precipitate in product may cause elevated background absorbance interference. | Process control | Genzyme Diagnostics P.E.I. Inc. notified the 1 account by a Removal Notification letter on 3/14/08 requesting that they stop using and discard any remaining stock. The customer is to request replacement through their normal channels. Questions or concerns should be directed to Penny White of Genzyme. |  | 153 | 153 kits | Nationwide Distribution | Terminated | May 16, 2012 | 1524 | 153 | Not_Computer
Z-1263-2009 | 51447 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Carina Home Ventilator. |  |  | January 21, 2009 | 2009 | Draeger Medical, Inc. | The device may not generate a pilot line disconnect alarm if the pilot line becomes disconnected, which may potentially result in a substantial reduction in tidal volume delivery. | Software design | An "Urgent - Medical Device Recall" Letter dated January 2009 was issued to consignees via certified mail. The letter described the issue, specific conditions, background information, and precautions for customers. An additional letter dated January 22, 2009 was included with the "Urgent - Medical Device Recall" letter which instructed consignees to inform an enduser of the contents of the notification letter because some of the consignees are home care device/service providers. When the new hardware/software solution is available, affected devices will be updated free of charge. Direct questions about the notification letter to Drager Medical, Inc. by calling 1-800-543-5047 (press "1" at the prompt and then press "2349#"). Direct questions regarding updating your device when the new software is available, contact Drager Service Technical Support at 1-800-543-5047 (press "4" at the prompt). |  | 40 | 40 units | Nationwide Distribution -- states of CA, FL, GA, HI, IL, LA, MA, MN, OH, and TX. | Terminated | January 05, 2010 | 349 | 40 | Software | Alarm/Message | Software update | Software Update
Z-1326-2009 | 50740 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Smiths Medical BCI(R) Advisor(R) Vital Signs Monitor, model 9200, catalog numbers: 920654225, 920674225, 920674235, 925454220, 925454235, 925454320, 925454325, 925454330, 925454335, 925457325, 925474325, 925654225, 925654325, 925754325, 925754335, and 92M774325. |  |  | December 22, 2008 | 2008 | Smiths Medical PM, Inc. | Recent improvements were made to the Advisor case plastic resulting in the inserts being deeply embedded into the plastic. This affects the ability of the monitor to be adequately secured on the wall mount bracket or roll stand bracket, resulting in the monitor falling off. There have been no adverse events or injuries reported. | Device Design | An Urgent Product Safety Information, Voluntary Recall Notification, dated 12/22/08, was sent to affected consignees. The notification contained a Safety Action Bulletin which provided actions to be taken. The Safety Action Bulletin described the issue, requested customer to contact Smiths Medical to ensure proper mounting hardware is utilized for those consignees using a roll stand or wall mount. If the consignee further distributed the product, they were asked to forward on the Safety Action Bulletin to the end user. Page 3 of the Safety Action Bulletin was asked to be completed and returned to Smiths Medical PM, Inc.A second "Urgent Product safety Information Voluntary Recall Notification" letter dated January 16, 2009 was sent to consignees. The second letter included a risk statement associated with the recalled product. |  | 143 | 143 monitors (17 US, 126 OUS) | Worldwide Distribution -- USA including states of FL, GA, NJ, NY, and TN and countries of Argentina, Brazil, Ecuador, Greece, Peru, Romania, Russia, Singapore, Slovenia, Trinidad and Tobago, United Kingdom, and Uruguay. | Terminated | December 17, 2011 | 1090 | 143 | Not_Computer
Z-1397-2009 | 50201 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Centricity TriWin Laboratory Information System; GE Healthcare Integrated IT Solutions, Barrington, IL 60010 |  |  | November 17, 2008 | 2008 | GE Healthcare Integrated IT Solutions | A software defect in the Centricity Lab (TriWin) product may incorrectly link patient information to a wrong report. Also, when viewing results in Patient Viewer, and a coagulation result is higher or lower than the linearity limit, then all of the results following the higher or lower result will show the same result as the first flagged result. | Software design | GE Healthcare notified their Centricity TriWin Laboratory customer via letter dated 11/17/08 of the software anomalies resulting in patient safety issues involving the patient information incorrectly being linked to a wrong report, and when viewing results in Patient Viewer, and a coagulation result is outside of a linearity limit, then all of the results following the out of linearity result will show the same result as the first flagged result. The user was informed that a software patch has been prepared to resolve the issues and will be installed on their system. Any questions were directed to GE Technical Support at 888-778-3375 or 905-479-1249, select option 1. The accounts was requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax it back to 905-305-0046 to schedule the software upgrade. |  | 26 | 26 units | Worldwide Distribution -- USA and country of Canada. | Terminated | September 11, 2010 | 663 | 26 | Software | Output/Calculation | Software update | Software Update
Z-1255-2009 | 51071 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche MagNA Pure LC Workstation; Product Numbers: 12236931001, 12236931444, 12236931692 and 03670325001. Robotic workstation for automated isolation of nucleic acids from different kinds of crude sample material such as whole blood, serum, blood cells, culture cells, tissue, bacteria and fungi. |  |  | February 09, 2009 | 2009 | Roche Diagnostics Corp. | Protruding screw threads may cut the user when cleaning the interior of the instrument. | Device Design | In a letter titled "Urgent Medical Device Correction" dated February 9, 2009, consignees were notified of the protruding screws and instructed to use caution when cleaning the instrument until a protective plate is installed by the firm's field representative.Further questions should be addressed to Roche Molecular Diagnostics Technical Support at 1-800-526-1247. |  | 371 | 371 | Nationwide Distribution including Puerto Rico and states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI and WV. | Terminated | December 07, 2009 | 301 | 371 | Not_Computer
Z-1261-2009 | 51687 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer Sirus Intramedullary nail for tibia, cannulated, diameter 12 L, 330 mm, sterile, Zimmer GmbH, Winterthur, Switzerland; REF 02.02631.233. |  |  | March 23, 2009 | 2009 | Zimmer Inc. | The package is labeled as containing 330 mm length nails, but actually contains 300 mm length nails. | Other | Distributors were initially notified by phone on March 23, 2009. Distributors and user accounts were notified by an "Urgent: Device Recall" letter dated March 31, 2009. The letter described the affected product, reason for recall, risks to health, other information and actions for customers. Customers were instructed to quarantine recalled product and to return it to the Zimmer Distribution Center. For questions about the recall, contact Zimmer, Inc. at 1-800-613-6131. |  | 32 | 32 | Worldwide Distribution -- US (states of Florida, Missouri, New Jersey, New York, Ohio, Oklahoma and Texas) and countries of Austria, Japan, Korea and Germany. | Terminated | October 23, 2009 | 214 | 32 | Not_Computer
Z-1195-2009 | 51193 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 2121-IQ/3131-IQ Biplane Cardiovascular Imaging System. The Innova Biplane systems are the first and only digital flat panel biplane systems with a full sized lateral plane to cover lateral anatomy without requiring multiple contrast injections and radiation exposures. The product is indicated for use in generating fluoroscopic images of human anatomy for diagnostic and interventional cardiac angiography procedures. These systems can be operated in a mobile or fixed site environment. |  |  | November 21, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has become aware of a potential image storage failure on Innova 2121 IQ/ 3131 IQ cardiovascular imaging systems that may impact patient safety. This intermittent communication failure on the image storage chain may result in the following situations: Record abort, Record not available, Fluorostore function not available, Review not available, Boot/Reset failure. Fluoroscopy imag | Device Design | A GE Healthcare "Urgent Medical Device Correction" letter dated February 2, 2009 was sent to consignees via mail. The letter was directed to Hospital Administrator/Risk Manger, Managers of Radiology/Cardiology and Radiologists/Cardiologists addressing the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.Further questions should be addressed to GE Healthcare at 1-800-437-1171. |  | 85 | 85 | Worldwide Distribution -- USA including states of CA, CO, FL, HI, IA, IL, IN, KS, MD, MA, MN, MO, NY, OH, OR, PA, TN, TX, WA, WV, WI and VA and countries of the United Kingdom, Romania, Japan, Taiwan, Bahrain, France, Malaysia, India, Canada, Chile, Indonesia, Ireland and Italy. | Terminated | November 21, 2013 | 1826 | 85 | Other | Device Operation | Correction | Repair
Z-1395-2009 | 50531 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio-Rad Variant II Hemoglobin Testing System with CDM v 4.0, model number 270-2000, Manufactured by Bio-Rad Laboratories, Hercules, CA |  |  | December 09, 2008 | 2008 | Bio-Rad Laboratories Inc | The wrong barcode number and tube position number may be assigned to samples leading to discrepant sample results. | Pending | Recall initiated on December 10, 2008. Notification conducted by phone call and fax for domestic consignees. Each subsidiary will receive the subsidiary Urgent: Medical Device Correction notice, customer medical device correction response form and number of customers distributed to by e-mail. Each subsidiary must contact customers and provide the customer notice and response form as needed. The notice stated that Bio-Rad Laboratories would release a follow up communication providing a permanent solution upon completion of an investigation. In the interim, it was recommended that customers carefully review the Summary Report for duplicate barcode IDs or or tube position numbers within the same rack. If there are any questions or assistance needed, the local regional Bio-Rad office should be contacted. |  | 42 | 42 units | Worldwide Distribution -- USA, Argentina, Australia, Brazil, Canada, China, Czech Republic, France, Germany, Hungary, India, Italy, Mexico, New Zealand, Singapore, South Korea, Sweden, Thailand, and the UK. | Terminated | November 12, 2010 | 703 | 42 | Not_Computer
Z-1199-2009 | 51455 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart XL Model M4735A Defibrillator, Automatic, External |  |  | March 30, 2009 | 2009 | Philips Healthcare Inc. | Rotary energy select switch in affected M4735A devices may fail and prevent the user from turningthe device on, rendering the device unusable for monitoring and defibrillation therapy. | Nonconforming Material/Component | An Urgent - Medical Device Recall Letter and Notice dated March 2009 were issued to affected customers. Mailings of the Urgent - Medical Device Recall Letter and Notice were handled by Philips Healthcare representatives in each affected geography. Customers were asked to follow the "Actions to be taken by customer/user" section of the Urgent - Medical Device Recall Notice. In this section, customers were instructed to remove the device from service if possible, while they await the correction of their device. If the customer is unable to remove the device from service, they are asked to identify a readily available backup device to use in the event that their device becomes unusable.All customers will be contacted by a Philips Healthcare representative to arrange for replacement of the energy select switch in the affected devices (Mandatory Field Change Order FCO86100076). Contact your local Philips Healthcare representative for additional information or support concerning this issue. |  | 26992 | 26,992 | Worldwide Distribution -- UNITED STATES, CANADA, ALBANIA, ARGENTINA, AUSTRIA,BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BOSNIA &amp; HERZEGOWINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, DOMINICAN REPUBLIC, EGYPT,EL SALVADOR, FINLAND, FRANCE, GERMANY,GREAT BRITAIN, GREECE, GUADELOUPE, HONG KONG,HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND,ISRAEL, ITALY, JAPAN, JORDAN,KAZAKHSTAN, KENYA,KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACAU, MACEDONIA, MALAYSIA,MALI, MARTINIQUE, MAURITIUS, MEXICO, MONGOLIA, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, OMAN, PALESTINE, PANAMA, PARAGUAY, PERU,PHILIPPINES, POLAND,PORTUGAL, PUERTO RICO,ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA,SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN,TANZANIA,THAILAND, TRINIDAD AND TOBAGO, TURKEY,TUNESIA, TURKMENISTAN, UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA,VIETNAM, VIRGIN ISLANDS and YEMEN. | Terminated | September 04, 2012 | 1254 | 26992 | Not_Computer
Z-1400-2009 | 51131 | No | GTQ | Streptococcus spp. exoenzyme reagents. | Microbiology | Microbiology | 510(K) Exempt | Remel Staphaurex Plus Kit, Catalog No. R30950102 (150 tests) and Catalog No. R30950201 (450 tests), rapid latex agglutination test for the identification of Staphylococcus aureus The responsible firm listed as Remel Europe Ltd, Dartford, England. |  |  | November 18, 2008 | 2008 | Remel, Inc | These lots may exhibit false positive reactions with strains of Staphylococcus epidermidis that may result in the mis-identification of Staphylococcus aureus. | Nonconforming Material/Component | Consignees were notified of the recall via letter dated 11/20/08 explaining the reason for recall and requesting any remaining units from the lots be discarded. A Product Inventory Checklist was enclosed to report the amount of product remaining in their inventory. Any non-responding consignees were contacted via telephone during the period of 12/30/08 to 1/11/09 to ensure the effectiveness of the recall. When the recall was expanded to include additional lots, the recalling firm issued an Important Medical Device Customer Recall Notice letters dated 3/16/09 again explaining the reason for recall. The letter also asked consignees to review their inventories for the product lots indicated and requested that remaining units from the lots be discarded. This letter also had a Product Inventory Checklist enclosed for completion. |  | 446 | 446 kits | Worldwide Distribution -- USA and Canada. | Terminated | December 03, 2009 | 380 | 446 | Not_Computer
Z-1411-2009 | 51886 | Yes | FLS | Apnea monitor. | Anesthesiology | Anesthesiology | 510(k) | SmartMonitor 2 Infant Apnea Monitor. |  |  | April 23, 2009 | 2009 | Respironics, Inc. | Audible alarm failure. | Process design | A notification letter dated April 23, 2009 was issued to consignees informing them of the issue. Consignees were instructed to locate the affected units in their inventory and return them to Philips Children's Medical Ventures-Youngwood Service. Any units that were shipped by consignees to their customers are to be retrieved and returned to Respironics. Consignees were also instructed to complete the enclosed Business Reply Form and Serial Number Reconciliation List and fax to Stericycle at 888-345-5369 or email to respironics1985@stericycle.com. For questions and additional information about the recall, contact Stericycle Customer Service by calling at 888-345-4630. |  | 4992 | 4,992 units | Nationwide and Canada. | Terminated | October 19, 2011 | 909 | 4992 | Not_Computer
Z-1196-2009 | 50202 | No | MDS |  |  | Clinical Chemistry | PMA | Abbott brand FreeStyle Navigator, Receiver Warranty Kit, Blood Glucose Monitoring System, Part Number: 70789-01. |  |  | November 01, 2008 | 2008 | Abbott Diabetes Care, Inc. | The product has the Units of Measure set to mmol/L rather than mg/dL. The blood glucose results could be misinterpreted. This may lead to under treatment and potential for hyperglycemia. | Packaging process control | The Consignees of the affected FreeStyle Navigator Receiver Kit was contacted via telephone, Federal Express or an Abbott representative beginning on November 1, 2008. Contact Customer Service at 1-866-597-5520 with further questions. |  | 2 | 2 units | US Distribution to states of NY and AZ. | Terminated | June 03, 2009 | 214 | 2 | Not_Computer
Z-1394-2009 | 51561 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche cobas 6000 analyzer series (including cobas c 501 and cobas e 601 modules) Middleware Solutions Instrument Manager Server as follows: v.8.04.02.10, catalog no. 04785568001; v.8.04.03.01, catalog no. 04785568002; and v.8.06.02, catalog no. 05109019001. |  |  | April 01, 2009 | 2009 | Roche Diagnostics Corp. | Some data alarm flags do not transfer from the analyzer to the laboratory information system, since not all data flags are recognized by the instrument manager/middleware. | Component design/selection | Consignees were notified by "Urgent Medical Device Correction" letter dated April 1, 2009 and instructed to add error codes to the Instrument Manager driver and to contact the manufacturer of their LIS for instructions on adding data alarm codes to the LIS. Consignees were also instructed to complete the attached fax form after adding the error codes to their system and fax to 1-888-345-4690. Contact Roche Diagnostics Technical Support at 1-800-428-2336 for questions about information contained in the notification letter. |  | 498 | 498 of all products. | Nationwide Distribution. | Terminated | December 07, 2009 | 250 | 498 | Not_Computer
Z-1344-2009 | 51334 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific, FleximaÂż and PercuflexÂż Drainage Catheter Family (Pigtail Style) |  |  | March 18, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific Corporation is initiating an urgent Field Correction related to its FleximaÂż and PercuflexÂż Drainage Catheter Family (Pigtail Style). During the attempted removal of the device, sutures were separated from the catheter and remained in the patient. | Labeling False and Misleading | Consignees were sent a Boston Scientific " Urgent Medical Device Field Correction Notification" letter dated March 18, 2009. The letter was addressed to "The Hospital Field Action Contact". The letter described the problem, product affected and the revised catheter removal instructions. They also requested the return of the Acknowledgement Form. |  | 586276 | 586,276 US; 276,013 OUS total for all products | Worldwide Distribution -- USA including DC, Puerto Rico and Guam and countries of United Arab Emirates Austria, Bosnia and Herzegovina, Bahrain, Switzerland, Germany, Egypt, Spain, Finland, France, Great Britain, Ireland, Italy, Kuwait, Netherland, Pakistan, Portugal, Saudi Arabia, Sweden, Slovak Republic, Tunisia, Turkey, South Africa, Netherlands, Costa Rica, Japan, Argentina, Australia, Brazil, Canada, Chile, China, Columbia, Ecuador, El Salvador, Guatemala, Honduras, Hong Kong, Macau, Malaysia, Mexico New Zealand, Panama, Peru, Philippines, Puerto Rico, Singapore, South Korea, Sri Lanka, Taiwan, Thailand, Uruguay, and Venezuela. | Terminated | February 23, 2012 | 1072 | 586276 | Not_Computer
Z-1640-2009 | 51702 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | In Situ Bender Cutter Kit Wide Sterile, Catalog number 530.521S |  |  | April 03, 2009 | 2009 | Synthes Spine | Sterility may be compromised due to a ineffective package design. | Package design/selection | All affected Synthes DMF Sales Consultants were notified of the Medical Device Recall by an Urgent: Medical Device Recall letter notification which was sent via USPS certified mail on 4/6/09. The letter stated the problem and risk. The firm recommended that patients who have had a surgical procedure in which these products were utilized be monitored for signs/symptoms of infection for a period of six months following surgery. If customers have any of the identified devices, they are to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number; complete the form; and return the letter with the recalled product to the address provided. Customers that do not have the identified product, are to complete the attached form and return the document to the address provided acknowledging their receipt of the recall information. Questions should be directed to their Synthes CMF Sales Consultant. |  | 2706 | 2706 units total | Worldwide Distribution -- USA including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and the country of Canada. | Terminated | September 15, 2010 | 530 | 2706 | Not_Computer
Z-1642-2009 | 51868 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoration Modular Cone Bodies; titanium Alloy and Hydroxyapatite, Sterile. Catalog number 6276-1-021, Catalog number 6276-1-025, Catalog number 6276-1-219, and Catalog number 6276-1-325. |  |  | April 02, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Packaging burst test values for the Restoration Modular Cone Body family were found to be incorrectly transferred from the finalized validation to the work specification. Testing showed that burst test results for recalled products fell below the validated parameters. | Process control | Urgent Product Recall letters were sent via Federal Express on April 2, 2009 to Stryker branches, Chief of Orthopaedics, Hospital Rick Management and Surgeons. The letter stated the issue, the potential hazards, and the risk mitigation. Customers were asked to examine their inventory and hospital locations to identify the affected products; retrieve all effected products and return it to their branch or agency warehouse for reconciliation; reconcile all products on the attached Product Accountability Form and fax a copy to the number listed; return the affected product via the Inter-Org Transfer and ship it to the attention of the individual listed; attach the fluorescent orange PRODUCT REMEDIATION sticker to their return, indicating the Product Remediation #, and mark the box with the words "PRODUCT RECALL". Questions should be directed to Colleen O'Meara at 201-831-5970. |  | 17 | 17 units | Nationwide Distribution | Terminated | August 27, 2012 | 1243 | 17 | Not_Computer
Z-1396-2009 | 48692 | No | DDR | Myoglobin immunological test system. | Immunology | Immunology | 510(k) | Access Myoglobin Reagents Kits, Part Number: 973243. |  |  | September 27, 2007 | 2007 | Beckman Coulter Inc | Beckman Coulter initiated the field action as a result of recent internal testing following the CLSI EP-5A - Evaluation of Precision Performance of Clinical Chemistry Devices protocol, it was determined that the total imprecision for the Access Myoglobin Reagent is greater than or equal to 10% across the expected physiological range rather than the stated claim less than 8%. Based on internal tes | Labeling False and Misleading | An Urgent: Product Corrective Action (PCA) letter was mailed via US Postal Service on October 08, 2007 to the customers that received Access Myoglobin Reagent kits informing them about the issue. The letter also informed customers that the product insert in the Access Assay Manual will be revised with the new range to reflect the new imprecision claim. Customers were instructed to:(1) Share this information with their laboratory staff and retain the letter for theirQuality System documentation,(2) Complete and return the enclosed fax-back response form within 10 days so that Beckman Coulter, Inc. can be assured that the customer received the notification, and(3) Contact Customer Technical Support at 1-800-854-3633 or your local Beckman Coulter, Inc. representative for questions or additional information regarding the notification. |  | 20727 | 20,727 | Nationwide and Canada. | Terminated | May 27, 2009 | 608 | 20727 | Not_Computer
Z-1398-2009 | 51207 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Cell-Dyn Ruby Hematology Analyzer with System Software Version 2.0ML, List Number 08H67-01, Manufactured by Abbott Diagnostics Division, Santa Clara, CA. |  |  | February 12, 2009 | 2009 | Abbott Laboratories | Burn Risk-- when power is cycled to the instrument, there is a possibility that the Mixer Motor Assembly may reach 100 degrees Celsius (212 degrees Fahrenheit ). If the customer touches the motor, burns may occur. | Device Design | A "Product Correction" letter dated February 12, 2009 and "Customer Reply" form were sent to all currently active CELL-DYN Ruby customers using Software Version 2.0 ML. The letter described the affected product, necessary actions and contact information for customers. Customers were instructed to acknowledge receipt of the "Product Correction" letter by completing the enclosed Customer Reply form and return via fax to 1-800-777-0051. A product information letter is being added to the accessory kits shipped with new instruments and version 2.0 ML System Software upgrade kits. For questions regarding this information, contact Abbott Laboratories Customer Support by calling 1-877-422-2688. Customers outside the US, please contact your local hematology customer support representative. |  | 324 | 324 units distributed | Worldwide Distribution -- US, Argentina, Canada, Columbia, Germany, Jamaica, Japan, Mexico, Puerto Rico, Singapore, South Korea and Venezuela. | Terminated | December 27, 2010 | 683 | 324 | Software | Physical Safety Hazards | Software update | Software Update
Z-1256-2009 | 50695 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | HemoSense INRatio2 Prothrombin Time Monitor, manufactured by HemoSense Incorporated. |  |  | December 23, 2008 | 2008 | HemoSense Inc | INRatio Monitors that were manufactured with a specific EEPROM memory chip fail during use. The monitors report an error code and do not report a test result. | Nonconforming Material/Component | Letters for each product titled "Urgent: Medical Device Recall Notice" dated December 19, 2008 were issued explaining the problem and instructing users to discontinue use, complete the provided reply form and return the device by mail to Inverness Medical. Distributors of the violative device were instructed to contact their customers and provide the same recall information.Further questions should be addressed to HemoSense Technical Service at 1-877-441-7440. |  | 921 | 921 units | Worldwide Distribution in the US, Germany, Spain and Australia. | Terminated | March 25, 2011 | 822 | 921 | Hardware | Display/Image | Discontinue | Other
Z-1265-2009 | 51424 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Gendex CB-500, 3D Dental Imaging Systems, Model Number G1-15-1-0. This product is manufactured by Imaging Sciences International for Gendex Corporation. |  |  | January 08, 2009 | 2009 | Imaging Sciences International, LLC | faulty calibration | Nonconforming Material/Component | Imaging Sciences International issued a letter to Customers dated January 7, 2009 informing them of the affected device. Users were advised to discontinue use of the device until they are contacted by a service representative and a certified re-calibration of the device is performed. Users were also asked to return a completed "Acknowledgement Form" by fax to the firm at 1-215-997-5665.Additional questions can be made to Imaging Sciences International by calling 1-800-205-3570 Extension 299. |  | 24 | 24 units | Worldwide Distribution -- US to states of FL, OH and PA and countries of Canada, Germany, England, Italy and Switzerland. | Terminated | June 04, 2009 | 147 | 24 | Not_Computer
Z-1638-2009 | 51898 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgidyne S-Vac round fluted channel silicone drain, 1/8" trocar, 10 FR, round, full fluted, 28 cm length, 4 channel, contains 10, sterile, Sterion Inc., Ham Lake, MN; REF 332187. |  |  | April 13, 2009 | 2009 | Aspen Surgical Products, Inc. | An adaptor, which may be required when connecting the drain to an external device, was not included in the package. | Process change control | An "Urgent Medical Device recall" letter dated April 13, 2009 was issued to consignees via certified mail. The letter described the affected product. and actions for customers. Consignees were instructed to carry out the recall to the retail level and to quarantine and return the product to Aspen Surgical Products, Inc. Consignees were also instructed to complete the enclosed response form and return it via mail, fax or e-mail. Direct questions about the recall by calling Aspen Surgical Products, Inc. at 1-616-698-7100, extension 142 or extension 141. |  | 9 | 9 cases of both products | Nationwide Distribution -- California, Florida, Georgia, Massachusetts and Missouri. | Terminated | November 23, 2009 | 224 | 9 | Not_Computer
Z-1399-2009 | 50849 | No | DGC | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | Immulite 2000 and Immulite 2500 digital Fluidics PC Board, |  |  | January 05, 2009 | 2009 | Siemens Medical Solutions Diagnostics | Design Flaw: The redesign of the Digital Fluidics PCB used in IMMULITE 2000 and IMMULITE 2500 is missing a jumper. | Process design | Siemens Healthcare Diagnostics Inc. issued an "Urgent: Device Recall" notice dated January 2009 mailed to all direct customers of the affected product. The letter described the problem, provided "work-around" instructions and also provided a Recall Response Form to be returned to the firm via fax. For further information, please contact your local Siemens Healthcare support representative or Siemens Healthcare Diagnostics Inc. directly at 1-973-927-2828. |  | 295 | 295 Digital Fludics PCB (88 IMMULITE 2000 and 2500 and 207 as spare parts). | Worldwide Distribution to states including MA, NY, PA, MD, GA, FL, OH, IN, MI, IL, MO, LA, UT, AZ and CA and countries of Canada, Czech Republic, Germany, Poland, Hungary, Belgium, Russia, Republic of South Africa, Greece, Israel, Turkey, Spain, Portugal, Denmark, France, Great Britain, India, China, South Korea and Mexico. | Terminated | April 19, 2010 | 469 | 295 | Hardware |  | Remove | Remove or Replace
Z-1267-2009 | 51284 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Argon Medical Devices, ArgoGuide Hydrophilic NiTi Guidewire, Catalog Number: AG087301, .035/150cm, regular/straight/standard, sold as a single hydrophilic guidewire in bulk non-sterile packaging. Argon Medical Devices, Inc. Athens, TX 75751. |  |  | February 26, 2009 | 2009 | Argon Medical Devices, Inc | Hydrophilic guidewires have the potential to demonstrate degradation of the Pebax cladding. | Other | Argon Medical Devices Inc. issued an "Urgent Field Safety Notice" letter dated February 26, 2009 informing all consignees of the affected product. Consumers were asked to immediately quarantine and return any product remaining followed by completion and return of the provided response form via email or fax (1-903-677-9396).For further questions or additional information call Argon Medical Devices Inc. at 1-903-677-9319. |  | 3300 | 3,300 units | Worldwide Distribution -- US including Puerto Rico and states of AL, CA, CO, FL, GA, IL, KS, LA, MA, MI, MN, MO, MT, NC, NJ, NV, NY, PA, SC, TN, TX, UT, VA and WI and countries of Germany, Thailand, Turkey and the Netherlands. | Terminated | January 07, 2010 | 315 | 43940 | Not_Computer
Z-1643-2009 | 51964 | No | NGY | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Alcon 25 GA Total Plus Pak, Catalog/Item #8065750220; for use with ACCURUS Vitrectomy System; distributed by Alcon Laboratories, Inc., Fort Worth, TX 76134. |  |  | April 29, 2009 | 2009 | Alcon Laboratories, Inc | Specific lots of Alcon 25 GA Trocar Blades do not conform to finished product specifications in that oxidation can occur on some of the blades. | Process control | Firm began notifying consignees of recall by Urgent: Medical Device Recall letter on 04/29/09 (letters sent via FedEx). Consignees instructed to check inventory for affected lots; to complete and return the Medical Device Recall Response Form; to contact Alcon to arrange for return of the devices; that credit will be issued for the recalled product when it is returned; and in the case that there is no inventory remaining, to indicate that on the Response Form and return it as well. Consignees should contact Alcon Technical Consumer Affairs at 1-800-561-6466, Option 8 for technical assistance. |  | 3523 | 3,523 surgical paks. | Worldwide Distribution -- USA and Canada. | Terminated | October 14, 2009 | 168 | 3582 | Not_Computer
Z-1646-2009 | 51695 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Roche ACCU-CHEK Spirit insulin pump. Part numbers 4540492001, 4759702001 and 4759729001. |  |  | April 30, 2009 | 2009 | Disetronic Medical Systems, Inc. | The "up" and "down" buttons, which are used for changing the program in the menu or to administer additional insulin through a bolus delivery, may experience intermittent failure or completely cease to operate. | Device Design | The firm issued a press release on 4/30/09 and sent letters with attached Urgent Medical Device Recall (Field Corrective Action) notification dated 5/4/09 to distributors, health care providers and users. Disetronic will replace pumps within the affected serial numbers (SN02119552 to SN10006093) that have intermittent or complete loss of function of the up and/or down button promptly upon contact from the customer with a new pump that is not affected by the potential button failure. For pumps within the affected serial numbers (SN02119552 to SN10006093) which are functioning properly, no replacement is available/possible at this time. Disetronic will gradually phase in replacement pumps for the remaining pumps within the affected range, as they become available. Customers are to complete the enclosed Customer Reply Card to acknowledge receipt of the notification.On 5/17/11, after being notified of the recall's ineffectiveness (in that only 40% of the pumps had been replaced), the firm re-issued recall notification letters including the aforementioned information to their consignees. |  | 14929 | 14,929 | Worldwide Distribution -- USA and Canada. | Terminated | June 05, 2012 | 1132 | 14929 | I/O | Device Operation | Replace | Remove or Replace
Z-1418-2009 | 51994 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical Table Headrest, Catalog/Part #0788-300-012, component of VERTIER Surgical Table. Distributed as part of the VERTIER Surgical Tables or individually by Stryker Communications, Flower Mound, TX 75028. |  |  | May 08, 2009 | 2009 | Stryker Communications Corp | Potential for health care provider/hospital staff member to sustain injury to fingers or hands when removing/handling the headrest component of the VERTIER Surgical Table. | Device Design | An "URGENT: Device Recall" letter dated May 11, 2009 was issued to consignees. Consignees were informed that Stryker service representatives will visit each location to affix a caution label identifying the pinch point on each headrest. Additionally, consignees will be provided with an updated surgical table user manual. Direct questions to Stryker Communications Corporation by calling 972-410-7100. |  | 230 | 230 headrests. | Nationwide Distribution and Canada. | Terminated | April 26, 2010 | 353 | 230 | Not_Computer
Z-1645-2009 | 51959 | No | LYZ | Non-powdered patient examination glove. | General Hospital | General Hospital | 510(k) | Kimberly-Clark, Synthetic Powder Free Vinyl Exam Gloves, Small, Medium and Large, 100 gloves per box, Non-Sterile, For Medical Use, Catalogue Numbers: 55031, 55032, and 55033. Distributed by Kimberly Clark LLC, Roswell, GA |  |  | April 30, 2009 | 2009 | Kimberly-Clark Corporation | The product may contain defects/holes reducing the product's effectiveness as a barrier. | Release of Material/Component prior to receiving test results | Consignees were contacted by an Urgent: Medical Recall letter on/about 4/30/2009. The letter instructs the Distributor to extend the recall to the end-user level by notifying each of their customers of the recall via a Customer Recall Letter. Both the Distributor Letter and Customer Letter include Response Sheets that allows for both Distributor and Customer (end-user) to indicate the quantity of affected product they have in their inventory or indicate whether they do not have any affected product. Distributors were instructed to cease further shipment of the product and to physically quarantine the product and fax the completed Response Sheet within 5 business days of receipt of the letter. The customer was instructed to immediately cease further use of the gloves and physically quarantine the product and fax the completed Response Sheet within 5 business days of receipt of the letter. Upon receipt of the Response Sheets, KC will arrange pick-up of the affected products and deliver it to a quarantine location at the contracted warehouse. |  | 9200 | 920 cases/10 boxes per case | Nationwide Distribution -- USA including states of CT, MA, MD, NJ, and NY. | Terminated | January 07, 2010 | 252 | 9200 | Not_Computer
Z-1417-2009 | 51962 | No | GDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CHIBA NEEDLE MARKED, Model CNM-1806, Size: 18 GA x 15 CM, Qty: 10 each, Remington Medical, Inc., Alpharetta, GA 30005. |  |  | April 06, 2009 | 2009 | Remington Medical Inc. | Incorrect needle (Franseen Needle, FNM-1806) was packaged as a Chiba Needle Marked. | Other | The distributor was notified on Monday, April 6, 2009 via telephone by the QA Manager and a return authorization was issued. A notification letter dated April 13, 2009 was issued via certified mail to the affected customers. The letter described the affected product, reason for recall and actions for customers. Customers were also instructed to conduct recall to user level. For questions or additional information, contact Remington Medical, Inc. by calling 800-989-0057, extension 123. |  | 60 | 60 units | Nationwide Distribution -- KY and NC. | Terminated | May 27, 2009 | 51 | 60 | Not_Computer
Z-1392-2009 | 51254 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Reaction Vessel, 2000 units per carton; for in vitro diagnostic use with the ARCHITECT i System; Abbott Laboratories, Abbott Park, IL 60064; list 7C15-01. |  |  | February 26, 2009 | 2009 | Abbott Laboratories | Abbott has had an increase in complaints due to the receipt of error code 1006 (unable to process test, background read failure), error code 1007 (unable to process test, activated read failure) and incorrect results when using these lots of ARCHITECT Reaction Vessels. | Nonconforming Material/Component | Consignees were contacted by phone and sent an "Urgent Field Safety Notice Product Recall" dated February 26, 2009. The letter informed users of the problem and detailed how to identify and discard the affected product. In addition, Abbott Laboratories asked the consignees to send a copy of the letter to any laboratory to which they may have forwarded any of affected lots. Abbott sent a second letter dated May 12, 2009 informing consignees of a recall expansion of additional affected lots. The recall was again expanded on 12/17/09 for an addtional 40 lots of reaction vessels.For additional information, contact Abbott Labs at 1-877-4ABBOTT. |  | 27163 | 27,163 cartons | Nationwide including Puerto Rico, and internationally to Canada, Mexico, Guatemala, Costa Rica, Venezuela, Colombia, Chile, Argentina, Dominican Republic, Honduras, Nicaragua, Germany, Japan, Hong Kong, Taiwan, Thailand, South Korea, Uruguay, Austria, Bahrain, Belarus, Belgium, Bosnia &amp; Herzegovia, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Greece, Hungary, India, Indonesia, Ireland, Italy, Jamaica, Kuwait, Luxembourg, Malaysia, Netherlands, Norway,Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom, Zimbabwe and Australia. | Terminated | November 04, 2010 | 616 | 27163 | Not_Computer
Z-1323-2009 | 51668 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Lactosorb Pectus Stabilizer, model number: 01-3805. Distributed and Manufactured by : Biomet Microfixation, Jacksonville, FL. |  |  | March 20, 2009 | 2009 | Biomet Microfixation, Inc. | This recall is based on the occurrence of known adverse effects identified in the device instructions for use. Bending, fracture, loosening, rubbing and migration of the devices can occur as a result of excessive activity, trauma or load bearing.Implantation of foreign materials can result in an inflammatory response or allergic reaction. | Device Design | An "Urgent Medical Device Recall Notice" was issued via certified letter or email on March 20, 2009 to consignees. Customers were instructed to immediately remove and return the affected in-stock product to Biomet Microfixation, Inc. Customers were also asked to complete and return the enclosed Inventory Reconciliation sheet via fax (904-741-9425) to the firm, and to notify recipients of any recalled products that had been further distributed.For questions about the recall, contact Biomet Microfixation at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm EST. |  | 2117 | 2,117 | Worldwide Distribution -- US (including DC and states of AL, AZ, CA, CO, FL, GA, HI, IL, IN, KS, KY, LA, MI, MO, MN, NC, NE, NH, NJ, NM, NY, OH, OK, PA, TX, VA and WI) and countries of Argentina, Austria, Belgium, Chile, Colombia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Guatemala, Israel, Italy, Lativa, Netherlands, South Africa, Spain, Sweden, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | September 11, 2009 | 175 | 2117 | Not_Computer
Z-1421-2009 | 51703 | No | OBP | Repetitive transcranial magnetic stimulation system. | Neurology | Physical Medicine | 510(k) | NeuroStar TMS Therapy System/SenStar Treatment Link. |  |  | March 16, 2009 | 2009 | Neuronetics, Inc. | SenStar Treatment Link may have an open circuit that during use with the NeuroStar TMS Therapy system may cause arcing and heat damage to the SenStar and/or treatment coil. Heating of the SenStar during use has a potential risk of minor scalp burn to patients. | Process control | Neuronetics Inc. notified consignees by telephone and sent a follow-up fax/letter dated March 19, 2009 titled "Urgent Product Recall Notification." The letter describes the issue and actions for consumers to take in response to the recall. Neuronetics will visit each customer in person to complete a "Returned Goods: SenStar Recall Inventory Form." Customer questions should be directed to Neuronetics Customer Service at 1-877-600-7555. |  | 3123 | 3123 Units | Nationwide Distribution to AZ, CA, DC, GA, FL, IL, MD, MI, MN, NH, NY, PA, RI, SC, and WA. | Terminated | May 29, 2009 | 74 | 3123 | Hardware | Physical Safety Hazards | Return | Remove or Replace
Z-1401-2009 | 51863 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panorama Patient Monitoring Network with ambulatory telepacks, and software version 8.5.5, 8.5.6, 8.6.1 or 8.7. |  |  | April 10, 2009 | 2009 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | Software Anomaly: Software anomaly can occur when the Panorama is attempting to play an audio file, either a physiological or technical alarm sound, at which time the system software may cause the system to freeze due to an interaction in the alarm system. | Software design | Urgent Product Field Correction letters, dated March 27, 2009, were sent by certified mail, return requested, and signature required. The letters identified the affected product and the problem. It also stated the adverse effect on patients and that customers' Panorama software will be upgraded to a current software revision. A Datascope Service Representative will contact customers to arrange for the software upgrade of the Panorama(s) in their facility. This will be done at no cost to the customer. |  | 294 | 294 sites | Worldwide Distribution -- USA and Australia. | Terminated | August 03, 2010 | 480 | 294 | Software | Device Operation | Software update | Software Update
Z-1436-2009 | 51982 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | CareVue Chart Release C.O |  |  | May 05, 2009 | 2009 | Philips Healthcare Inc. | Software: A software defect under certain circumstances could lead to the unintended administration of a duplicate dose of a scheduled medication within a single 24 hour period | Software design | Philips Medical notified accounts by letter on 5/6/09 the Urgent - Medical Device Correction notification. While customers await the correction of their devices, customers are given the following instructions:Clinicians should not rely solely on the worklist within CareVue Chart to determine actionable orders for the patient until their software is upgraded. Clinical users should always use the medication administration record (MAR) and the administration record in determining the care their patients receive in addition to the critical care worklist. Philips is currently developing a correction that will be available at no additional charge. A Philips representative will notify customers when the software correction becomes available. If further assistance is needed, customers are to contact the Philips Customer Care Center at 1-800-722-9377. |  | 103 | 103 units | Worldwide Distribution -- USA including states of CA, FL, MA, OH, OK, RI, VT, and WV and countries of Canada, Australia, Austria, Belgium, Denmark, Finland, France, Germany, Italy, Malaysia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, and the United Kingdom. | Terminated | March 06, 2012 | 1036 | 103 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1224-2009 | 44848 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1 Central Control Monitor. Terumo Cardiovascular Systems, Ann Arbor, MI; Catalog number 802100. |  |  | January 12, 2009 | 2009 | Terumo Cardiovascular Systems Corporation | The monitor may exhibit start-up failure, video/display issues, calibration issues, and partial or complete loss of functionality during use. | Device Design | Consignees were notified via an Urgent Medical Device Correction: Safety Advisory letter dated 1/12/09 advised of the reason for the correction, potential problems, and provided clinical information should the problem(s) arise. Customers are to review the notice, assure that all users are aware of the notice, and confirm receipt of the communication by completing the attached Customer Response Form and faxing it to the number indicated on the form. Questions or concerns should be directed to Terumo CVS Customer Service at 1-800-521-2818. |  | 951 | 951 | Worldwide Distribution -- USA, Australia, Belgium, Canada, Dominican Republic, Egypt, Germany, Honduras, Hong Kong, India, Indonesia, Japan, Jordan, Kuwait, Malaysia, Mexico, Pakistan, Philippines, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, Turkey, United Arab Emirates, and Viet Nam. | Terminated | July 14, 2009 | 183 | 951 | Software | Device Operation | Instructions | Safety Notice/Insructions
Z-1225-2009 | 51220 | No | DYE | Replacement heart valve. | Cardiovascular | Cardiovascular | PMA | Medtronic Mosaic Porcine Bioprostheses with Cinch, Heart Valve; Model 305; Catalog #s 305C21AA, 305C2701, 305C2301, 305C2501, and 305C2901; Sterile LC; Rx only; Medtronic Inc 710 Medtronic Parkway, Minneapolis, MN 55432-5604, Manufactured at: Santa Ana, Ca. |  |  | January 26, 2009 | 2009 | Medtronic Inc | Medtronic identified several heart valves (Mosaic Porcine Bioprosthesis with Cinch &amp; Hancock II Bioprosthesis with Cinch) that were inadvertently distributed after exposure to temperatures exceeding the 77 F (25 C) maximum storage temperature published on the labeling. The product is being retrieved as they do not meet the handling requirements. | Storage | A Medtronic "Urgent Medical Device Recall Notice" dated January 26, 2009, was sent to Risk Managers of each affected account and to each Implanting Physician. The letter described the problem, product and informed consignees that they will be retrieving the non-implanted vales. If additional information is needed, customers are to contact their local Medtronic sales representative or technical services at 1-877-526-7890. |  | 5 | 5 | Nationwide Distribution -- Including states of CT, FL, MD, MA, MI, NC, OH, OR, PA, SD, TN, TX, WI, and WA. | Terminated | May 22, 2012 | 1212 | 15 | Not_Computer
Z-1324-2009 | 51777 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Square Nut MR Safe, Catalog Number 03.311.060. |  |  | February 09, 2009 | 2009 | Synthes Spine | Numbers are etched incorrectly on the product. | Process control | An "Urgent: Medical Device Recall" letter dated February 6, 2009 was sent to all affected Synthes Trauma Sales Consultants via USPS certified mail on February 09, 2009 and affected In House employees were notified via return receipt email on February 09, 2009. All affected hospitals were notified of the recall through a notification letter mailed via USPS certified mail on February 09, 2009. The notification letter described the affected product, issue and actions for recipients. Direct questions about the recall to Synthes Spine by calling 800-620-7025, Extension 5455 or 610-719-5455. |  | 1137 | 1,137 units | Worldwide Distribution -- US (states of AR, CA, CO, CT, FL, IA, IL, LA, MD, ME, MI, MN, MS, NJ, NY, OK, PA, SC, TX, and UT) and country of Switzerland. | Terminated | December 15, 2009 | 309 | 1137 | Not_Computer
Z-1249-2009 | 51422 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes InSitu Bender Cutter Kit Narrow Sterile Catalog number 530.522S, used with Synthes Rapid Resorbable Fixation System for craniofacial skeleton fracture repair and reconstructive procedures. |  |  | December 18, 2008 | 2008 | Synthes Spine | The bender is heating prior to activation of the power button, requiring disassembly to stop heating. | Process control | All affected Synthes CMF Sales Consultants were notified of the Medical Device Recall by phone on 12/19/08 and by letter sent via return receipt email on 12/22/08. All affected hospitals were notified by letter mailed on 12/22/08 via USPS certified mail. |  | 40 | 40 units | Nationwide distribution: AL, CA, KS, LA, MD, NJ, NV, NY, OR, PA, TX, and UT | Terminated | December 15, 2009 | 362 | 132 | Not_Computer
Z-1227-2009 | 51166 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare ApexPro Telemetry System for mobile patient physiological monitoring. |  |  | February 24, 2009 | 2009 | GE Healthcare | Intermittent or continuous loss of patient monitoring. | Nonconforming Material/Component | Consignees were sent on 2/20/09 a GE Healthcare "Urgent Medical Device Correction" letter dated 2/12/09. The letter was addressed to Risk Manger/Chief of Nursing/Director of Biomedical Engineering. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. |  | 291 | 291 | Nationwide distribution and Canada. | Terminated | May 29, 2012 | 1190 | 291 | Not_Computer
Z-1318-2009 | 51449 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Mammo Test Model number 10144185, x-ray guided stereotactic biopsy system |  |  | January 21, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Table may unintentionally lift during procedure | Nonconforming Material/Component | Siemens Medical Solutions issued a Safety Advisory Notice dated 1/22/09 which informs the customers of the potential issues and recommends that they stop using their MammoTest System until the prone table has been repaired. Linear Medical visited the customers to repair the weld joint on the prone table by January 23, 2009. |  | 4 | 4 units | USA distribution to IL, NY, SC, and WA | Terminated | June 01, 2009 | 131 | 4 | Not_Computer
Z-1352-2009 | 51693 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Dual Offset Broach Handle, Clamping Style, Left, REF 31-555403. For use in primary hip arthroplasty. |  |  | March 31, 2009 | 2009 | Biomet, Inc. | The handle body may fracture or crack at the slot where the locking lever is inserted to clamp down onto the broach. | Device Design | Biomet notified their distributors by recall letter dated 3/31/09, instructed them to pick up the insturments at the hospitals and to notify the hospital personnel of this recall action in writing. |  | 178 | 178 of all products. | Nationwide, Canada and Finland. | Terminated | December 23, 2009 | 267 | 178 | Not_Computer
Z-1228-2009 | 51495 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Sterile Disposable Syringe Kit, catalog number DSK 130-Q, Angiographic Syringe for delivery of contrast media. |  |  | January 29, 2009 | 2009 | Medrad Inc | Possible compromised sterility due to holes in styrene tray | Packaging process control | On 1/29/09, Medrad issued an Urgent Medical Device Recall letter. Customers were requested to check their inventory for any of the affected batch. If the customer has affected product, they are instructed to contact the firm to make arrangement for the return of the product. Non responders were sent additional notifications via FedEx on 3/2/09 and 4/1/09. |  | 2958 | 2958 kits | Nationwide: AL, AZ, CA, FL, IL, LA, MA, MT, NC, NY, OH, TX, and WI. | Terminated | June 04, 2009 | 126 | 2958 | Not_Computer
Z-1264-2009 | 51438 | No | OEX | Catheter introducer. | Cardiovascular | Cardiovascular | Enforcement Discretion | Spectranetics SLS II Laser Sheath Kit, 14F, REF 500-012, Individually packaged catheter introducer. |  |  | March 16, 2009 | 2009 | Spectranetics Corporation | Catheter outer jacket may split during use. | Process control | Spectranetics notified the 8 recipients of the catheters by letter on 3/17/2009. They were told that a Spectranetics sales rep would contact them to coordinate the return of the suspect devices. Questions were referred to 719-447-2462 or 719-447-2539. |  | 17 | 17 units | Nationwide distribution: AL, KY, LA, MI, NC, NY, and PA. | Terminated | November 28, 2011 | 987 | 17 | Not_Computer
Z-1351-2009 | 51409 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Stryker, Gamma3, The Compact Version of the Gamma Nail System, Operative Technique for Hip Fracture System, Trochanteric and Long Nails. |  |  | March 02, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Operative Technique: further clarification and guidance is needed to help reduce the potential for distal mis-drilling of the short Gamma3 Trochanteric Nails. | Labeling design | Stryker sent Urgent Product Recall letters by Federal Express on March 2, 2009 to Stryker Branches, Chief of Orthopaedics, Hospital Risk Management and Surgeons. Letters were sent to all who have used the Gamma3 products in the past. |  | 18232 | 18232 | Nationwide. | Terminated | August 17, 2012 | 1264 | 38964 | Not_Computer
Z-1390-2009 | 51673 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator/Monitor, M3535A/M3536A |  |  | April 10, 2009 | 2009 | Philips Healthcare Inc. | Spontaneous turn-on which could deplete the battery, rendering the device unusable until power is restored. There is also the possibility of a failure mode in which the device fails to respond to user initiated turn-on, rendering it unusable for monitoring and therapy. | Nonconforming Material/Component | Philips notified customers via Urgent Medical Device Recall Letter &amp; Notice dated March 2009. Customers are asked to follow the "Actions to be taken by customer/user section of theUrgent Medical Device Recall Notice. In this section, customers are given the following instructions:"During the interim period, as you await the upgrade for your device, you may continue to use the Heartstart MRx. However, to reduce the risk that a spontaneous turn-on could deplete the battery, it is advised that you keep:a charged spare battery with the Heartstart MRx AC power applied, if possible, when the device is not in useYou should also consider identifying a readily available backup defibrillator to use in the unlikely event that the device fails to turn on. Devices that exhibit a spontaneous turn-on or failure to respond to user initiated turn on should beremoved from service as soon as possible." Philips will replace the therapy switch in affected units free of charge per Mandatory Field Change OrderFC086100077. All customers will be contacted by a Philips Healthcare representative to arrange for replacement of the therapy switch in the affected devices. |  | 22176 | 12,385 units US; 9,791 units OUS | Worldwide Distribution -- USA, Canada, Afghanistan, Argentina, Aruba, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bosnia &amp; Herzegovina, Brazil, Bulgaria, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece, Guadeloupe, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Mexico, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Norway, Oman, Panama, Papua, New Guinea, Peru, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudia Arabia, Serbia &amp; Montenegro, Singapore, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, and Venezuela. | Open, Classified |  |  | 22176 | Other | Device Operation | Replace | Remove or Replace
Z-1435-2009 | 52049 | No | GZI | External functional neuromuscular stimulator. | Neurology | Neurology | 510(k) | WalkAide Stimulator System Control Module, Part Number 20-0100; component of the WalkAide Stimulator System; distributed by Innovative Neurotronics, Austin, TX 78746. |  |  | August 06, 2008 | 2008 | Innovative Neurotronics, Inc. | Knob responsible for regulating the amount of stimulation administered by the WalkAide System control unit may fall off while device is in use. | Nonconforming Material/Component | Innovative Neurotronics Inc notified the consignee via e-mail on August 6, 2008 of the affected product. The Consignee was asked to place a stop shipment on the WalkAides and to return inventory for replacement.Further questions can be directed to Innovative Neurotronics Inc at 1-512-721-1906. |  | 137 | 137 units | Nationwide Distribution to the state of GA. | Terminated | July 28, 2009 | 356 | 137 | Not_Computer
Z-1434-2009 | 51869 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ADDitIV Primary IV Set with Universal Spike, Backcheck Valve, two Injection Sites and Spin-Lock Connector, Catalog # V1443. |  |  | March 27, 2009 | 2009 | B. Braun Medical, Inc. | Device defect is occlusion/no flow. | Mixed-up of materials/components | An Urgent - Medical Device Recall notice was issued via US Postal Service Certified Mail with registered return receipt mail or FedEx to all customer/distributors in receipt of affected product on 3/27/09. Direct customers received letters addressed to Materials Management that informed them of this recall and provided instructions for disposition of affected product in their inventory. Distributors received the direct customer recall letter along with a cover letter outlining their responsibilities to execute the recall and notify their customers as requested. Distributors are to immediately discontinue further distribution of the affected product, determine their customers' inventory, and have their customers return the product to them. The attached "Product Removal Acknowledgement" should be used to indicate the specifics of the combined distributor and customer inventories. The Customer Support Department should be contacted to arrange for return. A Customer Support Representative will provide instructions for handling the affected product and assist with replacement product. |  | 15450 | 15,450 units | Worldwide Distribution -- USA and Canada. | Terminated | December 15, 2009 | 263 | 397191 | Not_Computer
Z-1325-2009 | 51701 | No | NAW |  |  | Radiology | PMA | TheraSphere Administration Set (Part Number K125914-034). |  |  | February 24, 2009 | 2009 | MDS Canada Inc. DBA MDS Nordion | Instructions for the use of the pinch clamp that is an integral part of the TheraSphere Administration Set were not included with the distribution of the device. | Error in labeling | A User Bulletin (TS-2009-1) dated February 23, 2009 was issued to all TheraSphere user sites. The bulletin described the issue, affected units, and recommended actions for users. The bulletin instructed users on the proper use of the pinch clamp. MDS Nordion also updated the package insert with these instructions.For questions about the recall, customers in North America contact MDS Nordion Customer Service at 1-866-363-3330 and customers outside North America call 1-877-474-5999 and request Customer Service. Customers may also e-mail MDS Nordion at TheraSphere@mdsinc.com. |  | 26 | 26 | Worldwide Distribution -- US (states of AZ, CA, CO, FL, IL, IN, KS, MD, MI, NC, RI, OH, OR, PA, WI) and country of Canada. | Terminated | June 08, 2009 | 104 | 26 | Not_Computer
Z-1348-2009 | 51781 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Zero-P StarDrive Screwdriver Shaft, T8, self-retaining, quick coupling, for use in spinal fusion surgery. Catalog # 03.617.902 Technique Guide. |  |  | February 10, 2009 | 2009 | Synthes Spine | Tip breakage, due to inadequate instructions for use. The Technique Guide provided incorrect information related to use of the Torque Limiting Attachment during screw insertion and final tightening. | Labeling design | All affected Synthes Spine Sales Consultants were notified of the Labeling Correction via email on February 10, 2009. They were instructed to review changes with their surgeons. They were supplied with an electronic version of the corrected Technique Guide on March 11, 2009. All affected Synthes Sales Consultants and Customers were notified of this labeling change via distribution of the updated Technique Guide and a notification letter. This packet was sent via USPS certified mail on March 26, 2009. To ensure a broad and accurate distribution of notification, affected consultants and customers are defined as having been shipped Zero-P 3.0 mm Titanium Cervical Spine Locking Screws. These devices are used only with the Zero-P system. |  | 1766 | 1766 units | Distributed Nationwide and Canada. | Terminated | October 05, 2009 | 237 | 1766 | Not_Computer
Z-1414-2009 | 51877 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | 8F Prelude Short Sheath Introducer, 4CM in Length with Marker Tip and 0.038" X 50CM, REF No.: PSS-8F-4-038MT, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095. |  |  | April 14, 2009 | 2009 | Merit Medical Systems, Inc. | External tubing may separate during use, allowing blood loss. | Process control | All of Merit's US sales representatives and distributors were notified by Prelude Short Sheath Introducer - Product Recall letter on 4/14/2009. They were instructed to promptly contact their customers. Customers were instructed to isolate, remove from inventory, and return all affected inventory. A Product Retrieval Form was to be returned to Merit. Further information is available at 801-316-4998 or 801-208-4344. |  | 188 | 188 units | Worldwide Distribution -- USA, including states of CA, CO, GA, IL, MO, OK, and OR and countries of Austria, Belgium, Canada, Denmark, France, Germany, Hong Kong, Japan, Latvia, Lithuania, Netherlands, New Zealand, Spain, Sweden, and UK. | Terminated | February 09, 2010 | 301 | 18972 | Not_Computer
Z-1510-2009 | 52076 | No | DXY | Implantable pacemaker pulse generator. | Cardiovascular | Cardiovascular | PMA | Medtronic SIGMA 100/200/300 Series IPGs, Single Chamber Pacemaker; models SS103, SS106, SS203, and SS303. |  |  | May 18, 2009 | 2009 | Medtronic Inc. Cardiac Rhythm Disease Managment | One or more bond wire pairs will lift or separate from the bonding terminals on the device hybrid. This may present clinically as loss of rate response, premature battery depletion, loss of telemetry, or no output. | Device Design | Medtronic Sales representative hand delivered an Important Patient Safety Information letter, dated May 2009, to physicians beginning 05/18/09. The letter describes the issues, provides a predicted failure rate for the 3 populations of devices, and provides Patient Management Recommendations. The letter recommended that physicians should advise their patients to seek medical attention immediately if they experience symptoms (e.g., fainting or lightheadedness). It was also recommended that physicians should consider device replacement for patients who are both pacemaker dependent and who have been implanted with a device in the affected subsets. Medtronic will offer supplemental device warranty if the device is not already at elective replacement time. The last recommendation was that physicians should continue routine follow up in accordance with standard practice for those patients who are not pacemaker dependent. The letter also provides Physician and Patient Support. Letters will be mailed to patients with registered devices beginning 05/27/09. Questions should be directed to a local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 9434 | 9434 | Worldwide Distribution -- USA, including Washington D.C., Guam, Puerto Rico, and Virgin Islands; Albania, Algeria, Argentina, Armenia, Aruba, Australia, Austria, Bahamas, Bangladesh, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, Georgia, Germany, Greece, Guadeloupe, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Latin America, Latvia, Lebanon, Luxembourg, Martinique, Mauritius, Mexico, Morocco, Nepal, Netherlands, Netherlands Antilles, Norway, Oman, Pakistan, Palestinian Territory Occupied, Peru, Poland, Portugal, Qatar, Reunion, Romania, Saudi Arabia, Senegal, Slovakia (Slovak Republic), Slovenia, South Africa, Spain, Sri Lanka, Sudan, Suriname, Sweden, Switzerland, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Vatican City State (Holy See), and Venezuela. | Terminated | March 28, 2012 | 1045 | 85378 | Not_Computer
Z-1422-2009 | 51681 | Yes | JOQ | External programmable pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | EP-4 Computerized Cardiac Stimulator; SJM Part Numbers: 09-1527-0000 / 09-1527-0002; |  |  | March 23, 2009 | 2009 | St. Jude Medical | Non-compliance with IEC Standard 60601-1 Clause 57.6 relative to fusing of AC input line. The neutral line is not fused. The hot line in fused. Clause 57.6 calls for both lines to be fused. | Device Design | All foreign customers were sent Urgent Product Safety Information letters on March 23, 2009. Letters described units affected, the reason for action, and the actions to take to eliminate risk. It is recommended that the information in the letter be circulated to all users, risk managers and the biomedical engineering department responsible for maintaining and inspecting electrical equipment located within their institution. The letter also stated that a SJM representative will be in touch to perform the EP-4 Stimulator fuse modification. Customers are to complete and return the attachment. |  | 694 | 694 total (US, Canada, and OUS) | Worldwide distribution. | Terminated | January 20, 2010 | 303 | 694 | Not_Computer
Z-1419-2009 | 51413 | No | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 50 Sterilization System Product Code 10050 |  |  | March 09, 2009 | 2009 | Advanced Sterilization Products | ASP has identified a secondary, lower-occurring cause of oil mist within certain STERRAD NX Systems and STERRAD 50 Systems due to potentially defective oil fill plugs. | Component design/selection | Firm mailed out Urgent Device Correction letters dated March 9, 2009, advising customers/users that on April 23, 2008, Advanced Sterilization Products (ASP) initiated a device correction notification related to certain models of STERRAD Sterilization Systems that have the potential to emit oil mist due to a premature mechanical failure in the oil mist filter and/or filter assemblies. All affected STERRAD Systems have been corrected since the notification was issued.Recently, a secondary, lower-occurring cause of oil mist has been identified within the STERRAD NX System and STERRAD 50 System oil fill plugs. STERRAD NX System and/or STERRAD 50 System may potentially have defective oil fill plugs. Less than 1% of the installed STERRAD NX Systems have been affected. ASP has received 18 complaints in 10 months related to this failure for the STERRAD NX System. No complaints related to this issue for the STERRAD 50 System have been received. As of December 22, 2008, ASP has been actively replacing the plugs in these systems with stainless steel plugs during service calls and will continue to replace the oil fill plugs untilall affected units are corrected. The letter advises that there is a risk of symptoms, such as nausea, shortness of breath, dry throat, burning, watery and blurry eyes, and headache occurring if customers inhale oil mist. If you see mist, haze or smoke in the room or detect an oil odor, personnel should complete the following: - Cancel the cycle if the sterilizer is running and as always, do not use instruments from an incomplete cycle- Leave the room- Avoid working in the room until the mist has cleared- Discontinue use of the STERRAD System until the system is repaired.Any additional questions can be directed to the Clinical and Technical Support staff at 1-888-4613. |  | 4222 | 2,209 units installed in US; 2013 units OUS | Worldwide Distribution -- USA, Germany, Italy, France, UK, Switzerland, Benelex, Uruguay, Venezuela, Argentina, Chile, Columbia, Ecuador, Belgium, Puerto Rico, Hong Kong, Indonesia, Mexico, Philippines, Austria, China, Greece, Ireland, Israel, Japan, Korea, Malaysia, Singapore, Thailand, Taiwan, Turkey, Middle East, Peru, Canada, Czech Republic, Portugal, South Africa, Spain, Sweden, Hungary, Poland, Brazil, Russia, Egypt, India, and Slovenia. | Terminated | March 14, 2012 | 1101 | 4222 | Not_Computer
Z-1441-2009 | 51947 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Mavig Portegra 2 which is an accessory to the Siemens angiographic, fluoroscopic and radiographic imaging systems. Component material number 10281151 Surgery lamp, 115V. |  |  | May 01, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Improper installation of accessory support arm, which could potentially result in the loosening of the arm over time causing it to fall. | Other | The recalling firm has decided to inspect all of the installed units for this condition and where necessary replace the missing clip. The recalling firm has issued an Update Instructions AX011/09/S dated 03.09, which provides instructions on how to perform a visual inspection of the Mavig Portegra 2 arm for the presence of the safety clip and the installation instructions on how to properly install the clip if it is missing. |  | 427 | 427 units | Nationwide Distribution. | Terminated | September 01, 2009 | 123 | 427 | Not_Computer
Z-1416-2009 | 52008 | No | HFE | Endometrial brush. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cook Medical Tao Brush I.U.M.C. Endometrial Cytology/Histology Sampler (Tao Brush), 9.0Fr/26 cm, 3.5 cm, sterile, Cook Ob/Gyn, Spencer, IN; REF J-ES-090500-BL. |  |  | April 02, 2009 | 2009 | Cook OB/Gyn, Inc. | There is an ink on the brush which has not been tested for use with the Tao Brush. | Process control | The consignee was notified by an Urgent: Recall Notice letter dated 4/2/09. The customer is to forward the notice to all personnel that need to be aware of it; pull from their stock and isolate any affected product; and complete the form. Upon receipt of the form, Cook Women's Health will issue a Return Goods Authorization number and the customer should return the affected products via FedEx using the Return Goods Authorization number. |  | 15396 | 15,396 | Nationwide Distribution. | Terminated | December 03, 2009 | 245 | 15396 | Not_Computer
Z-1469-2009 | 50745 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard E LUMINEXX Biliary Stent, Size: 14mm x 80mm, Product Number: ZBM14080. |  |  | December 23, 2008 | 2008 | Bard Peripheral Vascular Inc | This action is being taken because a label for an indicated use was applied to product sizes that are not approved for that use. The incorrect indication for use label stated "Now Approved for Vascular Use". | Employee error | Bard Peripheral Vascular, Inc. (BVP) sent an "Urgent - Product Recall Notice" letter to consignees dated December 22, 2008 outlining the reason for the recall. In the letter, customers were informed to check all inventory locations, stop any further distribution of and return affected product to BPV. Consignees were to complete the enclosed Recall and Effectiveness Check Form and fax to 1-800-440-5376 or send information by phone at 1-800-321-4254 extension 2727 or e-mail at silvia.carillo@crbard.com. |  | 94 | 94 units total for all product codes | Nationwide Distribution including states of AL, AR, CA, CT, DE, FL, GA, IL, IN, KY, MA, MD, MO, MS, NC, ND, NJ, NM, NY, OH, PA, TN, TX, VA and WI. | Terminated | April 02, 2012 | 1196 | 94 | Not_Computer
Z-1260-2009 | 51699 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Remstar Pro M-Series Heated Humidifier System. |  |  | February 05, 2009 | 2009 | Respironics | Malfunction of the J3 connection between the Heater Plate Main printed circuit assembly (PCA) and the heater plate may have an intermittent connection which may lead to excessive heat build-up on the humidifier PCA. | Nonconforming Material/Component | A Service Communication dated February 5, 2009 was provided to all Respironics customers that are approved to service Respironics devices. The communication described the affected device and actions required.Further questions may be addressed by phone at 1-800-345-6443, and by email at service@respironics.com or clinical@respironics.com. |  | 449585 | 449,585 units | Worldwide Distribution -- US including states of OR, OH, KY, TX, MO, AZ, IN and NV and countries of United Arab Emirates, Bosnia, Belgium, Bulgaria, Bahrain, Canada, Czech Republic, Denmark, Egypt, Spain, Finland, Greece, Croatia, Hungary, Ireland, Israel, Jordan, Kuwait, Lebanon, Latvia, Macedonia, Netherlands, Norway, Oman, Poland, Portugal, Romania, Qatar, Russia, Slovenia, Turkey, Serbia and South Africa. | Terminated | July 05, 2012 | 1246 | 449585 | Hardware | Physical Safety Hazards | Notification | Safety Notice/Insructions
Z-1540-2009 | 51211 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT EG7+ Cartridge; pH, PCO, PO, Na , K, iCa, Hct. |  |  | February 23, 2009 | 2009 | Abbott Point of Care Inc. | Equipment problem: i-STAT cartridges were not properly sealed as the packaging system jammed, causing compromised seals. | Packaging process control | Urgent Recall Notice letters pertaining to the EG7+ Cartridges were sent on February 23, 2009 by first class mail to all customers. The firm found one other affected lot of G3+ cartridges and expanded the recall to include one other lot. Letters for the expansion are dated May 2009. Customers were asked to use the attached form to record all of the box numbers if they have any cartridges from the affected lot. Cartridges from the box numbers listed will be replaced an any remaining unused cartridges from those specific boxes are to be returned as instructed on the form. To receive replacement cartridges, customers are to follow the instructions on the enclosed customer reply form. If any cartridges from the affected lot were forwarded to another facility, a copy of the recall letter needs to be provided to them and the information at the bottom of the customer reply form should be completed. Questions should be directed to 1-800-366-8020, Option 1. |  | 26200 | 1048 boxes of 25 cartridges each. | Worldwide Distribution -- USA, Austria, China, India, Saudi Arabia, South Africa, and Taiwan. | Terminated | July 29, 2009 | 156 | 54600 | Not_Computer
Z-1452-2009 | 52027 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Esteem Blue with Neu Thera Surgeons Glove; a sterile, single use powder-free synthetic polyisoprene surgical glove, latex-free, used as an underglove when double-gloving; 1 pair of gloves per sterile pouch, 200 pairs per case; REF2D73PB90 - size 9; Cardinal Health, McGaw Park, IL 60085 USA, Made in Thailand. |  |  | April 24, 2009 | 2009 | Cardinal Health | Increase in the number of customer complaints due to cuff tears. | Device Design | Urgent: Product Recall letters dated 4/24/09 were sent to the direct accounts on 4/24/09 via UPS 2-day mail, to the attention of the Director of Materials Management. The accounts were informed of the increased complaints of cuff tears, and were instructed to stop using them and destroy them. The letter included a list of gloves that could be substituted for the recalled gloves. The accounts were requested to call the appropriate customer service group to arrange for credit for the destroyed product, and to complete and return by fax the enclosed acknowledgement form, indicating the number of gloves destroyed. |  | 29697 | 29,697 cases total | Worldwide Distribution -- USA, Australia, Austria, Belgium, France, Germany, Japan, New Zealand, Switzerland, Spain, UK, Netherlands, Norway, Sweden, and Turkey. | Terminated | May 11, 2010 | 382 | 29697 | Not_Computer
Z-1470-2009 | 51865 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Wiseguide Guide Catheter, Sterilized. |  |  | March 11, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific has initiated a recall of the Wiseguide Guide Catheter. They have received reports of difficulties in connecting the product hubs with the Y-Adaptors and toughy borst connectors. If there is difficulty making such connection, there is a potential for prolongation or delay of the procedure in order to exchange the catheter or connector. There have been 41 complaints in the US as | Process change control | Consignees were sent a "Boston Scientific Urgent Medical Device Recall" letter dated March 11, 2009. The letter was addressed to Risk Manager/Field Action Contacts describing the problem including instructions for affected device handling and a request to return the enclosed Reply Verification Tracking Form.Further questions can be addressed to Boston Scientific at 763-494-1634. |  | 11308 | U.S. and U.S. territories: 22,728; OUS: 11,308 | Worldwide Distribution -- USA, United Arab Emirates, Dutch Antilles, Argentina, Austria, Belgium, Brazil, Canada, Switzerland, Chile, Colombia, Czech Republic, Germany, Algeria, Egypt, Spain, Finland, France, Great Britain, India, Indonesia, Ireland, Iraq, Italy, Jamaica, Japan, Kuwait, Sri Lanka, Libya, Netherlands, Norway, Poland, Puerto Rico, Qatar, Saudi Arabia, Sweden, Singapore, Syria, Turkey and South Africa. | Terminated | January 07, 2012 | 1032 | 11308 | Not_Computer
Z-1474-2009 | 49524 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 2100 IQ Versatile Cardiac and Vascular Imaging System, GE Healthcare, Waukesha, WI 53188. |  |  | June 17, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has become aware that a few customers have experienced an unusual occurrence of system lockups on Innova single plane cardiovascular X-ray imaging systems while using the image auto-archiving function. when occurring during X-Ray imaging, the acquisition fails and the user has to rest the system. To date, one patient death has been reported related to this issue. | Process control | GE Healthcare sent a "GE Urgent Medical Device Correction" letter dated May 2008. The letter was addressed to Health Care Professionals describing the problem with the affected devices including instructions to avoid further problems until a GE Healthcare service representative is able to make contact and upgrade the system. Further questions may be addressed by calling 1-800-437-1171. |  | 514 | 514 | Worldwide Distribution - US including Puerto Rico and states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY,NC, ND, OH, OK OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA,WV and WI and countries of ALGERIA, AUSTRALIA, BELARUS, BELGIUM, BOSNIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, FINLAND, FRANCE, GERMANY, GEORGIA, GREECE, GUATEMALA, HONG KONG, HUNGARY, INDIA, INDONESIA, IRAQ, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL RUSSIAN FEDERATION SAUDI ARABIA, SINGAPORE, SLOVAK REPUBLIC, SLOVENIA, SPAIN, SUDAN, SOUTH AFRICA, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, TURKEY UNITED KINGDOM, VENEZUELA and SOUTH AFRICA. | Terminated | October 22, 2014 | 2318 | 15816 | Not_Computer
Z-1507-2009 | 50706 | No | HAX | Skull tongs for traction. | Neurology | Neurology | 510(k) | ReSolve Halo - Open Back Halo Ring, Model Number: 505300D |  |  | November 09, 2008 | 2008 | Ossur Engineering, Inc | Ossur initiated the recall after their investigation of some complaints found that in some instances the ceramic pins are difficult to screw through the halo ring. | Process change control | Ossur contacted direct consignees and sub-accounts beginning on November 10 , 2008 via a phone call informing them of the recall, and the reason as stated in the public reason for recall. Consignees were asked to check their inventory for any of the identified products and contact them with this information. The firm developed a telephone script for the customer service representatives to use.Ossur then forwarded a notification letter to the affected customers beginning on November 12, 2008 via US Postal Service. In the event that any of these parts have already been applied to a patient, Ossur advised customers to follow their normal follow-up protocols, ensuring that the ceramic pins are properly fixed in the skull.At the time of the phone calls Ossur did not have sterilized open back replacement rings available; however, they offered their ReSolve Closed Back Halo ring as an alternative. The firm estimated that the normal product offering should be available once again in approximately four to six weeks from the date of the phone call. |  | 1 | 1 piece | Worldwide Distribution -- USA including states of CA, GA, FL, MD, MI, MO, &amp; WA and countries of Australia, Cyprus, Germany, Italy, Kuwait, Saudi Arabia, Singapore, South Africa, Spain, United Arab Emirates, and the United Kingdom. | Terminated | April 22, 2010 | 529 | 92 | Not_Computer
Z-1425-2009 | 51202 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Datex-Ohmeda Aisys, Anesthesia Gas-Machine. |  |  | February 05, 2009 | 2009 | GE Healthcare | Unintended shut-down: The anesthesia machine may start shutting down without human intervention due to a faulty On/Standby Switch. Prior to shutdown, the system will issue an alarm and play a message that shutdown will occur in 8 seconds. The failure only affects the electrical portion of the switch. If the clinician does not intervene, the patient could experience hypoventilation leading to h | Component design/selection | GE Healthcare contacted consignees via two "Urgent Medical device Correction" letters dated February 5, 2009 and March 10, 2009. The letters are addressed to 3 titles within the affected accounts; Healthcare Administrator/Risk Manger, Chief of Intensive Care and Director of Biomedical Engineering. The letters describes the Safety Issue, Affected Product details, Safety Instructions, Product Correction and Contact Information. The second letter dated March 10, 2009 includes 700 serial numbers that were missing from the first letter. The second letter was only mailed to the consignees of the additional 700 units. |  | 3082 | 3082 | Worldwide distribution: USA, Canada, Mexico, YEMEN, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, UKRAINE, TUrKEY, TUNISIA, THAILAND, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIA, QATAR, PORTUGAL, POLAND, PERU, PANAMA, PAKISTAN, NORWAY, NEW ZEALAND, NETHERLANDS, MOROCCO, MALAYSIA, MACEDONIA, LITHUANIA, LIBYA, LEBANON, LATVIA, KUWAIT, KOREA, KENYA, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, INDIA, HUNGARY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CZECH REPUBLIC, CROATIA, COLoMBIA, CHINA, CHILE, BRAZIL, BELGIUM, BAHRAIN, AUSTRIA, AUSTRALIA, ARGENTINA. | Terminated | December 10, 2011 | 1038 | 9325 | Not_Computer
Z-1389-2009 | 51496 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 automated external defibrillator and MRL Jumpstart (collectively "AED 10") |  |  | February 26, 2009 | 2009 | Welch Allyn Protocol, Inc | Issue #1: AED 10 displayed a "Do Not Use" symbol as a result of a blown F1 fuse.Issue #2: AED 10 models that incorporate a spring connector with the speaker experienced intermittent electrical connection to the circuit board causing a loss of voice prompts. | Component design/selection | Welch Allyn sent an Urgent Medical Device Recall letter dated February 26, 2009 to all distributors and customers, describing 5 failure modes. Customers were instructed to respond to the letter by calling 888-345-5356, and decide on replacement or discounted new purchase. |  | 9754 | Issue #1: 4568 units; Issue#2: 5186 units | Worldwide distribution. | Terminated | March 07, 2012 | 1105 | 9754 | Hardware | Display/Image | Replace | Remove or Replace
Z-1423-2009 | 51716 | No | ISA | Therapeutic massager. | Physical Medicine | Physical Medicine | 510(K) Exempt | Sybaritic, Inc., SlimLine Health Environment Capsule, |  |  | June 18, 2008 | 2008 | Sybaritic, Inc | Misbranded and sale without pre-market notification. | PMA | Sybaritic sent a letter to the consignee on June 18, 2008. The letter confirmed their phone conversation on June 16 regarding the required Field Correction of the device. The letter also described the changes to be performed by the local technician to be in compliance with the FDA. |  | 1 | 1 | UT | Terminated | July 06, 2012 | 1479 | 1 | Not_Computer
Z-1424-2009 | 51952 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, Transfer Spike, Blood Processing Accessory, REF BT945, Sequestration Kits |  |  | February 25, 2009 | 2009 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | Sterility may be compromised. Non-vented caps that connect to the 3-way stopcock attached to the transfer spike could prevent Ethylene Oxide from contacting the inside surfaces of the stopcock during sterilization, leading to a potentially non-sterile product. Use of the product may cause serious health consequences from potential sepsis/infection and the associated remote risk of death. | Device Design | Consignees were sent a Medtronic "Urgent Medical Device Recall Notice" letter dated March 5, 2009. The letter was addressed to Risk Manager. The letter described the product and problem and requested the consignee to complete, sign and return the certificate acknowledgement. |  | 20 | 20 units | Distribution to one medical facility in MN. | Terminated | May 07, 2011 | 801 | 20 | Not_Computer
Z-1500-2009 | 51887 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Arrow Balloon and stylet temporary pacing catheter |  |  | February 11, 2009 | 2009 | Arrow International Inc | Leakage: Material not lipid resistant, may crack and/or leak. Use of this device with lipid containing solutions or 70% isopropyl alcohol may cause leakage or air embolism. | Device Design | Arrow International issued a letter dated 2/18/09 to all customers informing them of the problem and if a lipid resistant stopcock is required to contact the firm for replacements. New labels state: Warning - do not use with lipid containing solutions or 70% isopropyl alcohol, as these solutions may cause leakage or air embolism. |  | 900000 | 900,000 total, all products | Worldwide distribution: USA, Argentina, Australia, Bahamas, Belarus, Belgium, Brazil, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Guatemala, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Peru, Philippines, Poland, Singapore, Slovakia, South Africa, South Korea, Spain, Sri lanka, Sweden, Taiwan, Thailand, Tunisia, Turkey, UK, Ukraine, and Venezuela. | Terminated | July 22, 2010 | 526 | 900000 | Not_Computer
Z-1486-2009 | 52098 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific FortÂżÂż Guide Wire, Floppy Marker Wire, 300cm x 2 cm, REF 34949-02, for use in intravascular interventional procedures. |  |  | May 12, 2009 | 2009 | Boston Scientific Corporation | Product packaged incorrectly. One batch/lot of wires, that should contain radiopaque marker bands, is missing the marker bands; and a second batch/lot of wires, which should not contain radiopaque marker bands, has the marker bands. Since the marker bands issue would not be visually apparent to the physician there is a potential for prolongation or delay of the procedure in order to exchange the | Process control | Consignees were sent a "Boston Scientific Urgent Medical Device Recall" letter dated may 19, 2009. The letter was addressed to "Dear Risk Manger/ Field Action Contact". The letter described the problem, customer steps for recall and listing of product batches / lots. The letter requested consignees to return the Reply Verification Tracking Form. |  | 5 | Total of 5 boxes (5 units /box) , | USA distribution: AR, MO, CA, FL, and NV. | Terminated | December 11, 2011 | 943 | 9 | Not_Computer
Z-1504-2009 | 51817 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | AccuSure, U-100 31 Gauge, 1cc Insulin Syringe, 5/16" (8mm) Needles, Short Needle, packaged 10 syringes per package, 10 packages per carton, NDC 0603-7002-2 |  |  | April 09, 2009 | 2009 | Vintage Pharmaceuticals LLC DBA Qualitest Pharmaceuticals | Needles separate from the barrel of the syringes. | Nonconforming Material/Component | On 04/09/2009, Qualitest issued a press release and sent out recall notifications to their consignees by US Mail. The recall was extended to the User level. Qualitest will send a second letter to all non-responders after thirty days. Thirty days after the second letter, Qualitest will perform a recall effectiveness check by telephone on a percentage of non-responders (in accordance with FDA request). |  | 559500 | 559,500 syringes | Nationwide. | Terminated | November 01, 2010 | 571 | 559500 | Not_Computer
Z-1505-2009 | 50773 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions E.CAM Patient Handling System (PHS) Field of View (FOV) indicator bars; a component of the E.CAM Emission Computed Tomography System; part number 43 66 618 (right) and 43 81 088 (left). |  |  | December 18, 2008 | 2008 | Siemens Medical Solutions USA, Inc. | Patient fingers can get pinched if they attempt to pull themselves up from the patient handling system without any assistance and grabbing onto the flip-up FOV indicator bars. | Device Design | Siemens sent Urgent Field Correction Recall letters dated 12/18/08 to all users of the E.cam Systems. The accounts were informed of the potential risk to patients when they attempt to pull themselves up from the patient handling system (PHS) without any assistance and grabbing the field of view (FOV) indicator bar (aka flip-up armrests). The customers were advised to always assist the patient when getting on/off of the PHS. They also were advised that a Siemens Customer Service Representative would contact them within 90 days to schedule a corrective upgrade to the PHS with a modified FOV guide to mitigate the risk for patient injury free of charge. Any questions were directed to contact Siemens Medical Solutions USA, Inc. at 1-800-767-2313 (USA) for assistance. |  | 1107 | 1,107 units | Worldwide distribution: USA, Guam, Argentina, Australia, Austria, Bangladesh, Belgium, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Czech Republic, Ecuador, Egypt, Finland, France, Germany, Great Britain, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Latvia, Lebanon, Mauritius, Mexico, Myanmar, Netherlands, New Zealand, Norway, Pakistan, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, Turkey and United Arab Emirates. | Terminated | May 03, 2011 | 866 | 1107 | Not_Computer
Z-1506-2009 | 51871 | No | GTZ | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Inverness Medical BioStar Strep A OIA MAX, 100 test kit size, |  |  | April 06, 2009 | 2009 | Binax, Inc. dba IMPD | Potential for false positive Strep A results. | Other | Inverness Medical notified Customers by letter sent via Fedex April 6, 2009 . The letter included a request for a response from the customer. The product is discontinued. |  | 1729 | 1,729 kits | Nationwide. | Terminated | January 06, 2010 | 275 | 1729 | Not_Computer
Z-1488-2009 | 52023 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 20 ml Syringe |  |  | April 30, 2009 | 2009 | Becton Dickinson &amp; Company | Potential for compromised sterility: Insufficient packaging seals. | Nonconforming Material/Component | Becton Dickinson sent Urgent Product Recall letters on April 30, 2009 to all direct customers. Questions are to be directed to BD Customer Service at 1-888-237-2762.Letters to customers of distributors went out on May 12, 2009. |  | 201600 | 201,600 units | Nationwide Distribution. | Terminated | September 28, 2009 | 151 | 201600 | Not_Computer
Z-1523-2009 | 52003 | No | NJL |  |  | Orthopedic | PMA | Zimmer NexGen Complete Knee Solution Legacy Knee- Posterior Stabilized LPS Femoral Component Option Size D Left, sterile, Zimmer, Warsaw, IN; REF 00-5996-014-51. |  |  | April 17, 2009 | 2009 | Zimmer Inc. | The implant surface may not have been polished adequately, which could result in increased wear and polyethylene debris generation. | Employee error | Consignees were notified by phone beginning on 4/17/09 and email and letter notifications were sent on 4/22/09. A dear surgeon letter entitled Urgent: Medical Device Recall, dated April 28, 2009, is being sent to implanting surgeons recommending patient monitoring. Consignees are to return the response form indicating their receipt of the notification. Questions or concerns should be directed to (800) 846-4637 or www.zimmer.com. |  | 10 | 10 | Worldwide Distribution -- USA including states of Arizona, California, Colorado, Hawaii, Illinois, Indiana, Michigan, Minnesota, Mississippi, New Jersey, New Mexico, New York, Ohio, Oregon, and Pennsylvania and country of Canada. | Terminated | November 12, 2009 | 209 | 37 | Not_Computer
Z-1473-2009 | 51960 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow Embolectomy Catheters: 4 French, 2 Lumen x 40cm, Product Number: W01798, Single use, Latex Free, Rx only, Arrow International, 2400 Bernville Road, Reading, PA 19605. |  |  | March 25, 2009 | 2009 | Arrow International Inc | The inflation arm tubing may separate from the stopcock adapter. | Process control | Arrow International issued an "Urgent Medical Device Recall" letter dated March 25, 2009 describing proper handling of the affected device. Consignees were also instructed to complete a Recall Acknowledgement and Stock Status Form and fax to Arrow International at 1-800-343-2935.Further questions should be directed to Arrow International, Customer Service Department at 1-800-523-8446. |  | 6910 | 6910 units (total all three products) | Worldwide Distribution - USA and Panama. | Terminated | February 01, 2010 | 313 | 6910 | Not_Computer
Z-1477-2009 | 52004 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Gender Solutions Natural Knee Flex System, Gender Solutions Male, Femoral Component, nonporous, size 4, left, sterile, Zimmer, Warsaw, IN; Product Number: 00-5410-017-01. |  |  | April 29, 2009 | 2009 | Zimmer Inc. | The tyvek lid, which is intended to assure sterility until the device is transported to the operating room, may be missing from the inner tray. | Employee error | Zimmer, Inc. issued an "Urgent: Device Recall" letter dated April 29, 2009 informing all consignees of the affected product. Users were instructed to locate and quarantine all unused devices and ship them to the firm with a completed Inventory Return Certification form. Return Certification forms are also to be sent via fax to Zimmer, Inc. at 1-574-372-4265. For further questions, contact Zimmer, Inc. at 1-800-613-6131. |  | 14 | 14 | Worldwide Distribution -- US including states of FL, MN, MS, NY, OH, TN, TX and WI and Switzerland. | Terminated | October 21, 2009 | 175 | 14 | Not_Computer
Z-1524-2009 | 52031 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Normed TX 8 Driver, Non-Cannulated (used with the Ascension Total Foot System); Catalog #503004268; distributed by Ascension Orthopedics, Inc., Austin, TX 78754. |  |  | April 14, 2009 | 2009 | Ascension Orthopedics, Inc | Screwdriver used with the Ascension Total Foot System not manufactured to specification, resulting in the driver not properly engaging and/or stripping the screws making it difficult to properly seat the screws during surgery. | Pending | Firm notified their distributors of recall via phone starting 04/14/09. Distributors replaced the recalled screwdrivers with ones that were manufactured to specification. |  | 9 | 9 units. | Nationwide Distribution -- including states of CO, OR, and TX. | Terminated | September 24, 2010 | 528 | 9 | Not_Computer
Z-1478-2009 | 52026 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Philips Avalon Fetal Monitor FM20; Model Number: M2702A. |  |  | May 12, 2009 | 2009 | Philips Healthcare Inc. | Interruption of power-up/boot sequence may cause the paper scale and recorder speed settings of the monitor to revert to factory default may lead to misinterpretation of the fetal trace by the user and to potential incorrect clinical decisions | Software change control | Philips Healthcare issued an "Urgent-Medical Device Correction" notification dated May 15, 2009 to Customers in the United States via UPS and notifications outside the US will be managed by Philips representatives in each affected geography. Customers were asked to follow the guidelines in the Action to be Taken by Customer/User section of the Correction Notice until they receive their software upgrade.Questions can be made to Philips response center at 1-800-722-9377. |  | 5663 | 5,663 units | Worldwide Distribution | Open, Classified |  |  | 10451 | Software | Output/Calculation | Software update | Software Update
Z-1525-2009 | 52093 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Xia Dual Anterior Staple; |  |  | April 12, 2007 | 2007 | Stryker Spine | Labeling Issue: The large Xia Anterior Staple was laser marked with an "S", indicating small. | Error in labeling | Market Withdrawal letters were sent to Stryker branches who received product on April 13, 2007 by Federal Express. The letters asked that they examine their inventory and hospital locations to identify the implant; reconcile these implants on the attached Product Accountability Form and fax a copy to (201) 760-8456; retrieve and return the affected implants; and contact their Stryker Customer Service rep to re-order the product that they are returning as part of the regulatory action. |  | 38 | 38 in US | Nationwide Distribution. | Terminated | June 29, 2009 | 809 | 38 | Not_Computer
Z-1481-2009 | 52005 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Zimmer Steinmann Pins, 7/64 in. (2.8 mm) diameter, 9 in. (229mm) length, style 6, nonsterile, qty 6, Zimmer, Warsaw, IN; Model Number: 00-0187-03-69. |  |  | May 01, 2009 | 2009 | Zimmer Inc. | The package contains pins with an incorrect diameter. | Packaging process control | Zimmer, Inc. issued an "Urgent: Device Recall" letter informing users of the affected product. Users were instructed to locate and quarantine all unused devices and ship them to the firm with a completed Inventory Return Certification form. Return Certification forms are also to be sent via fax to Zimmer, Inc. at 1-574-372-4265.For further questions, contact Zimmer, Inc. at 1-800-613-6131. |  | 37 | 37 | Worldwide Distribution -- US including states of CA, GA, IL, MA, MO, OR, TX and WI and Japan. | Terminated | November 12, 2009 | 195 | 37 | Not_Computer
Z-1559-2009 | 52041 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Dixtal Novametrix Neonatal/Pediatric SpO2 Sensor, Foam Wrap Style |  |  | May 13, 2009 | 2009 | Dixtal Medical, Inc. | Neonatal/Pediatric SPU Sensors may be incorrectly packaged in pouches labeled for Adult/Pediatric SPU Sensors. | Labeling mix-ups | Urgent Product Recall Notification was issued to affected accounts by email/facsmile on May 8, 2009 requesting to cease distribution, quarantine and return recalled product. Contact Dixtal Medical with questions; to inquire about your on-hand inventory; or to arrange to receive replacement product at Customer Service: sales@dixtal.com, Phone Number 1-203-269-1112, Fax Number 1-203-269-1760. |  | 675 | 27 Boxes (Packaged 25 sensors/box) | Worldwide Distribution -- US (TX and CT) and countries of Italy, Japan, France and Turkey. | Terminated | March 23, 2012 | 1045 | 675 | Not_Computer
Z-1539-2009 | 52046 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Bard StatLock Intravenous (IV) Catheter Stabilization Kit, Catalog number: IVK0124, C.R. Bard, Inc., Covington, GA 30014, USA. |  |  | April 02, 2009 | 2009 | C.R. Bard, Inc., Urological Division | The affected units may contain extension sets with the microbore tubing instead of extension sets with the macrobore tubing. | Process control | All consignees were notified by an Urgent: Product Recall Notification letter via Fed-Ex carrier on 04/02/2009. They were instructed not to use or further distribute any of the affected products. They were advised to check their inventory and to remove and return affected product to Bard Medical Division. A Recall &amp; Effectiveness Check Form was included to be filled out and returned via fax, to Bard. |  | 16100 | 16,100 units | Nationwide Distribution - including states of AZ, CA, CO, GA, IA, MN, NE, and TX. | Terminated | December 04, 2009 | 246 | 16100 | Not_Computer
Z-1582-2009 | 52214 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT software v2.5.0, Model 728306. The Brilliance iCT is packaged INA carton and the software is contained within the Brilliance CT scanner. |  |  | December 22, 2008 | 2008 | Philips Medical Systems (Cleveland) Inc | Software anomaly. An artifact may appear in brain scans and brain perfusion studies, when using 80kV and a UA, UB, or UC filter with this CT scanning device. | Software design | An "URGENT - Field Safety Notice" dated December 22, 2009 was issued via certified mail to affected customers. The letter described the affected product, problem, hazards involved, and actions for customers to take. Philips will provide customers with a software update for all affected systems within next 6 months. The firm also sent a second notification in the form of an "URGENT -Medical Device Correction" report dated January 29, 2009 which contained information noted in the "URGENT - Field Safety Notice". Direct questions to your local Philips representative or Philips Customer Care Center (1-800-722-9377), option 5: Other Imaging Systems, option 1: CT or to your local Philips Healthcare office. |  | 44 | 44 | Worldwide Distribution -- US including DC (states of CA, GA, IL, IN, KY, NY, MD, MN, OH, OR, and MD), Germany, Greece, Israel, Italy, Japan, Netherlands, Norway, Russia, South Korea, Switzerland, Taiwan, Canada, and Thailand. | Terminated | September 24, 2010 | 641 | 44 | Software | Display/Image | Software update | Software Update
Z-1531-2009 | 51980 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet OSS Elliptical Proximal Femoral Trial 7 cm left provisional, stainless steel, Biomet Orthopedics, Inc., Warsaw, IN; REF 32-472125. |  |  | May 19, 2009 | 2009 | Biomet, Inc. | The retaining ring may fall out of the instrument during surgery. | Process control | Distributors and consignees were notified by an Urgent Medical Device Recall Notice dated 5/19/09. Distributors and consignees were instructed to locate and remove the devices from circulation, carefully follow the instructions on the "FAX Back Response Form," and fax a copy of the Response Form to 574-372-1683 prior to return of the product. If the product has been further distributed, consignees must notify hospital personnel responsible for receiving recall notices of the action, via the enclosed "Dear Biomet Customer" notice. Consignees are responsible for the location and return of products to the firm. The letter stated that the instruments would be reworked and returned. |  | 387 | 387 of all products | Worldwide Distribution -- USA, Australia, Belgium, Canada, Costa Rica, Finland, France, Germany, Mexico, New Zealand, Poland, South Wales, and Spain. | Terminated | October 08, 2010 | 507 | 387 | Not_Computer
Z-1247-2009 | 50841 | No | IYB | Diagnostic x-ray tube mount. | Radiology | Radiology | 510(K) Exempt | Ceiling Tube Mount Ceiling Tube Mount (CTM)/Digital Ceiling Tube Mount (DCTM) with 5 inch or 10 inch Column Extensions; a radiographic x-ray tube suspension system designed to facilitate a wide range of procedures; Del Medical Systems Group, 11550 West King Street, Franklin Park, IL 60131-1330. |  |  | January 12, 2009 | 2009 | Del Medical Systems Group | There is the potential for the supporting cable to become detached from the column causing the column to fall and possibly contact the patient or operator. | Component design/selection | Del Medical sent a "Field Correction Recall" letter with an acknowledgement form to the direct account dealers/assemblers on January 12, 2009 - February 4, 2009, informing them that required maintenance inspection of a retaining ring inside the column assembly may not be performed as required. If the retaining ring is not regularly inspected and replaced when worn or loose, there is the potential for the supporting cable to become detached from the column causing the column to fall and possibly contact the patient or operator. The accounts were requested to locate the affected systems and immediately take them out of service until correction by installing a Cable Mount Retention Retrofit Kit to eliminate the potential for the cable to become detached. Affected accounts were instructed to complete the enclosed acknowledgement form and return it via fax (847-288-7911) or mail to Del Medical. If there are any questions, please contact Del Medical Technical Support at 800-800-6006 or 847-288-7000. |  | 213 | 213 devices | Worldwide Distribution -- US, Canada, Scotland and Pakistan. | Terminated | June 17, 2010 | 521 | 213 | Not_Computer
Z-1485-2009 | 51406 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural-Knee II System Modular Cemented Tibial Baseplate, Right, Size 0, TIB B/PL, for cemented use only, sterile, Zimmer, Warsaw, IN; Model Number: 6420-01-200. |  |  | March 18, 2009 | 2009 | Zimmer Inc. | The tibial spacer will not affix to the distal surface of the baseplate, as intended. | Nonconforming Material/Component | Zimmer, Inc. issued an "Urgent-Device Recall" letter dated March 18, 2009 instructing users to cease using and return the affected product accompanied with a completed Inventory Return Certification form. The inventory Return Certification form should also be faxed to Zimmer, Inc. at 1-574-372-4265.Further questions should be addressed to Zimmer, Inc. at 1-800-613-6131. |  | 834 | 834 of all products | Worldwide Distribution - USA including states of CA, NY, MI, OH, WI, IA, MN, NJ, IL, MO, TX, LA, NC, SC, PA, AR, MS, FL, GA, AL, OK, KS, TN, CT, NM, UT, NV, ID and CO and Switzerland. | Terminated | October 23, 2009 | 219 | 834 | Not_Computer
Z-1533-2009 | 51924 | No | LXT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TransFx External Fixation System Drill Bit, Quick Connect, 2.5 mm diameter, 180 mm length, nonsterile, Zimmer, Warsaw, IN; REF 4450-56. TransFx drill bits can be found in the following kits: TransFx Intermediate External Fixation System, REF 4450-05-10; TransFx Large External Fixation System, REF 4450-010-10 and TransFx Small External Fixation System, REF 4450-00050-15. |  |  | May 01, 2009 | 2009 | Zimmer Inc. | Drill tip is out of specifications, reducing the efficiency of drilling action, requiring more force and potentially heating the bone surface during operation. | Process control | Distributors and consignees were notified by an Urgent: Device Recall letter dated 5/1/09. Consignees were advised to stop using the device and quarantine it immediately. Distributors were instructed to carry out a physical count of all affected product and record the data on the Inventory Return Certification Form; fax a copy of the completed form to Zimmer, Inc.; and return recalled product along with the Inventory Return Certification Form to the firm. |  | 1391 | 1391 | Worldwide Distribution -- USA, Australia, Canada, El Salvador, England, Germany, Latin America, and Singapore. | Terminated | November 03, 2010 | 551 | 6085 | Not_Computer
Z-1556-2009 | 52015 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Signa Ovation 0.35 T MR System. |  |  | April 16, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has identified a potential safety issue with the Signa OpenSpeed and Signa Ovation MR Scanner tables that may impact patient safety. The table may begin lowering on its own while a patient is on the table. If this occurs while the table is in the magnet bore, the table may tilt and could result in patient injury. The patient's head or feet depending on patient orientation, may hit | Device Design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated April 16, 2009 addressed to Hospital Administrators, Mangers of Radiology/Cardiology and Radiologists/Cardiologists. The letter describes the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. A GE Healthcare Field Engineer will schedule a field visit to correct the device.Further questions or concerns may be addressed to GE Healthcare at 1-262-521-6681. |  | 203 | 203 | Worldwide Distribution -- US including PR and states of AL, AZ, AR, CA, CO,CT, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MN, MS, MO,MT, NV, NH, NE, NJ, NY, NC, OH, OK,OR, PA, SC, SD, TN, TX, VA, WA, WV, WI and countries of BRAZIL CHINA, DENMARK, EGYPT, ENGLAND, GERMANY, GUATEMALA, HUNGARY, ITALY, IVORY COAST, JAPAN, KOREA, KUWAIT, OMAN, MEXICO, OMAN, PHILIPPINES, RUSSIAN FEDERATION, SPAIN, SWITZERLAND, TURKEY, UNITED ARAB EMIRATES and VENEZUELA. | Terminated | April 23, 2013 | 1468 | 374 | Not_Computer
Z-1480-2009 | 52009 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Optivac Total Hip Kit, Maximum Capacity Cartridge, 1 kit, sterile, latex free, Biomet Orthopedics, Inc., Warsaw, IN; Product Number: 418000. |  |  | April 09, 2009 | 2009 | Biomet, Inc. | The package may contain two short mixing rods instead of one long and one short rod. | Packaging process control | Biomet Orthopedics issued an "Urgent Medical Device Recall Notice" dated April 9, 2009 instructing distributors to visit each hospital consignee, deliver a copy of the recall notification and replacement product and to pick up, and return the affected product.Further questions may be addressed to the Biomet Orthopedics at 1-800-348-9500. |  | 250 | 250 | Nationwide Distribution - Including AL, AR, FL, IL, LA, MN, NE, NC, NV, OH, OK, SC, SD, WA and WI. | Terminated | October 06, 2010 | 545 | 250 | Not_Computer
Z-1570-2009 | 52124 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | Sterrad 50 Sterilizer, Product Code 10050. |  |  | January 12, 2009 | 2009 | Advanced Sterilization Products | The capacitor in the vacuum pump component within the Sterrad 50 System could overheat and potentially create a hazardous condition inside the system, resulting in smoke and/or fire. | Other | Advanced Sterilization Products (ASP) sent an "Urgent: Product Correction" letter dated January 12, 2009 to each consignee informing them of the issue and instructed them to report smoke or a burning odors in or around their STERRAD 50 System. ASP strongly recommended for consignees to follow their facility's safety protocols and if possible, immediately disconnect the STERRAD 50 System from its electrical source. Customers were instructed to contact ASP to schedule an onsite visit to investigate the issue. If the issue is caused by a malfunctioning capacitor in the vacuum pump, the firm will replace the malfunctioning component right away. To correct this situation, an ASP Field Service Engineer (FSE) will visit all customers facilities within thenext sixty days to replace the capacitor in the vacuum pump on the affected STERRAD 50 System. |  | 1151 | 1,151 units. | Worldwide Distribution -- US, Germany, Italy, France, United Kingdom, Switzerland, Benelex, Uruguay, Venezuela, Argentina, Chile, Columbia, Belgium, Puerto Rico, Mexico, Philipines, Australia, Japan, Korea, Malaysia, Taiwan, Middle East, Peru, Canada, Czech Republic, Portugal, Spain, Sweden, Hungary, Poland, Brazil, Russia, Finland, Norway, Panama and Switzerland. | Terminated | December 28, 2011 | 1080 | 1151 | Hardware | Physical Safety Hazards | Replace component | Remove or Replace
Z-1549-2009 | 51501 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS (Picture Archiving and Communication System) software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | April 03, 2009 | 2009 | GE Healthcare Integrated IT Solutions | There is a potential safety issue associated with the use of GE Centricity PACS software related to configuration of DICOM query retrieval with rejected images. | Software design | GE Healthcare Integrated IT Solutions sent an "Urgent Medical Device Correction" letter dated April 2, 2009 to their Centricity PACS software versions 2.1.x and 3.0.x customers with attention to the Director of Radiology, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator. The letters advised users of the patient safety issue associated with the affected device and provided safety instructions to follow until the software is updated. User were asked to complete the enclosed Customer Reply form and fax the completed form to 1-847-939-1479.Centricity PACS Software Versions 1.0.x and 2.0.x customers were sent separate letters explaining the issue and the safety instructions to follow, but there will be no software updates since those software versions are no longer being supported.Any questions should be directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171. |  | 1188 | 1,188 units | Worldwide Distribution - United States including Puerto Rico and countries of Australia, Austria, Belgium, Brazil, Canada, Chile, China, Denmark, Ecuador, Egypt, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia, Malta, Mexico, Netherlands, New Zealand, Portugal, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Terminated | May 25, 2012 | 1148 | 1188 | Software | Display/Image | Software update | Software Update
Z-1956-2009 | 50739 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Genicon, RX Only, Sterile EO. Genicon 6869 Stapoint Court, Suite 114 Winter Park, FL U.S.A. 32792. EC REP Genicon Europe, Ltd. 44 Airedale Avenue, Chiswick London U.K. W4 2NW. Catalog Number 700-005-010, 5MM. |  |  | May 08, 2008 | 2008 | Genico, Inc., dba Genicon | Genicon is recalling the Genicon Disposable Suction Irrigation 5mm 700-005-010 and 10mm 700-005-005 Tube Sets. The caps on the trumpet valves of the devices could fall off, resulting in the malfunction of the Suction Irrigation Device. | Nonconforming Material/Component | All four consignees were contacted by email and fax on May 8, 2008. Wholesalers were requested to forward copies of Genicon's recall letter dated May 8, 2009 and response form to their customers. Instructions for product return was provided. Direct questions about the recall to Genicon by calling 407-657-4851 (in United States) or 0118 930 5333 (in United Kingdom). |  | 2790 | 2,790 | Worldwide Distribution -- US (MO and IA), United Kingdom, Netherlands and United Arab Republic. | Terminated | October 01, 2009 | 511 | 3260 | Not_Computer
Z-1544-2009 | 51484 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BrainLab Micro-Multileaf-Collimator, model 41600; BrainLab AG, Kapellenstrasse 12, 85622 Feldkirchen, Germany. |  |  | March 19, 2009 | 2009 | Brainlab AG | The Integrated Multileaf-Collimator (MLC) of the linear accelerator could interfere with the intended conformal fields of the BrainLab m3 micro-Multileaf-Collimator (mMLC) during patient treatment. | Other | BrainLAB AG issued an "Urgent Field Safety Notice/Product Notification" letter dated March 19, 2009 informing all customers of the affected device. Consignees were instructed on proper steps to take to avoid further issues. In addition, the most current m3 mMLC user guide, version 7.4, was scheduled to be sent to users in April 2009.For further questions, contact BrainLAB AG by phone at 1-800-597-5911). |  | 13 | 13 units | Nationwide Distribution - AL, CA, CO, FL, GA, IL, KS, KY, LA, MD, MI, MN, MS, ME, NJ, NY, OR, TN, TX, UT, VA, WV, WA and WI. | Terminated | August 19, 2010 | 518 | 50 | Not_Computer
Z-1580-2009 | 52176 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | TumorLoc Software Application for the EBW (Extended Brilliance Workstation, Model # 455011002031, sold as an accessory for the GEMINI PET/CT Systems running CT software versions 3.5.x. , CT host 2.2. |  |  | February 19, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | A software anomaly in the TumorLoc software for the firm's Brilliance CT scanners was discovered. The same software anomaly was subsequently determined to be present in the firm's GEMINI PET/CT units. The software defect results in a the potential for data to be lost during reconstruction of images in which slices can either be discarded, or interleaved based upon their series ID. | Software design | An "URGENT - Field Safety Notice" dated February 12, 2009 was sent via FedEx to domestic customers. The notice described the affected products, issue, involved hazard should problem occur and corrective actions for the customer/user. The letter also notified users/customers that a Philips Healthcare Field Service Engineer will contact and visit their site within 6 months to install a TumorLoc upgrade. Notification will also be sent to the international consignees by Philips personnel. For further information or support concerning this issue, please contact your local Philips representative: Customer Care Center (1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: Oncology) or to your local Philips Healthcare office. |  | 11 | 11 units | Worldwide Distribution -- US, Columbia, Germany, France, Netherlands, Canada, Belgium and Slovenia. | Terminated | September 21, 2012 | 1310 | 27 | Computer
Z-1568-2009 | 52047 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Dental Application on the Extended Brilliance Workspace v4.0 and v4.0.1. Model #455011202691 for Computed Tomography (CT); Model #455011002031 for Nuclear Medicine (NM). |  |  | April 27, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare has decided to recall the Dental Application on the Extended Brilliance Workspace v4.0 and v4.0.1 due to the fact that when images are saved or filmed that contain the Curved Multi-planar Reformation Section (cMPRS) views that were created with a short curve the Right/Left annotations will be reversed. | Software design | An "URGENT - Field Safety Notice" dated April 27, 2009 was sent to all consignees via certified mail. The letter described the affected products, issue, and actions for the customer/user. The notification letter provides a work around until the new software is available. For further information or support concerning this issue, please contact your local Philips representative: Customer Care Center (1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: Oncology) or to your local Philips Healthcare office. |  | 417 | 417 units | Worldwide Distribution -- US including DC (states of AL, AR, AZ, CO, DC, FL, IL, KY, LA, MA. ME, MI, MN, MO, NJ, NY, OH, OR, PA, SC, TN, TX, VT, and WI) and countries of Africa, Albania, Austria, Australia, Bangladesh, Belgium, Brazil, Canada, China, Columbia, Cyprus, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Holland, India, Iraq, Israel, Italy, Japan, Korea, Lithuania, Mexico, Netherlands, New Zealand, Oman, Pakistan, Peru, Philippines, Poland, Puerto Rico, Russia, Singapore, Spain, South Africa, Sweden, Switzerland, Syria, Tahiti, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uzbekistan, and Venezuela. | Terminated | September 21, 2012 | 1243 | 417 | Software | Display/Image | Software update | Software Update
Z-1615-2009 | 52128 | No | LWJ | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer femoral stem, revision taper, 15 mm diameter, 185 mm stem length, w/compression nut, sterile, Zimmer, Warsaw, IN; Model Number: 9982-15-18. |  |  | May 15, 2009 | 2009 | Zimmer Inc. | The threads on the screw are incomplete/non-functional. | Process control | Zimmer, Inc. began Consignee notification of the affected device via telephone on May 15, 2009. A certified "Urgent: Device Recall" notification dated May 21, 2009 was sent to all consignees instructing them of the affected device, to stop use of and return it with a Return Certification form. The Return Verification form should also be faxed to Zimmer, Inc. at 1-574-372-4265.For shipping assistance, questions or other concerns, contact Zimmer, Inc. at 1-800-613-6131. |  | 10 | 10 | Worldwide Distribution -- US states of IA, NE and CA and countries of Canada, Switzerland, United Kingdom, Hungary, Netherlands and Slovakia. | Terminated | November 12, 2009 | 181 | 10 | Not_Computer
Z-1657-2009 | 52077 | No | MGB |  |  | Cardiovascular | PMA | Abbott Vascular, StarClose Vascular Closure System Part Number 14677-01, for closure of femoral artery access sites after cardiovascular catheterization. |  |  | May 06, 2009 | 2009 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Potential sterility issue. Internal component quarantined for possible contamination, was released for sale after going through sterilization procedure. | Process design | Between May 6 and May 8, 2009, Sales Representatives visited accounts to recover the product. Recall notifications (Urgent Product Recall letters, dated May 6, 2009) were hand delivered to customers advising them Abbott Vascular is voluntarily recalling one lot of the StarClose Vascular Closure System for a potential stenlity issue discovered internally. Customers are asked to cease use of devices from this lot and work with their local account representative to review their inventory. complete the attached Recall Effectiveness Check Form and rerturn Identified products to Abbott Vascular. |  | 240 | 240 units | Nationwide distribution. | Terminated | December 21, 2011 | 959 | 240 | Not_Computer
Z-1656-2009 | 51923 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Welch Allyn AED 10 automatic external defibrillator and MRL Jumpstart (collectively "AED 10") |  |  | April 24, 2009 | 2009 | Welch Allyn Protocol, Inc | Possible for users to misunderstand signals. The original product labeling told users that a flashing low battery indicator signals that the battery is low and requires changing soon. If a user experiences a flashing low battery status indicator and does not understand that the device can continue to be used, the user might choose not to continue operating the device, thereby causing therapy to | Labeling design | The firm has developed expanded labeling for users and will revise the user manuals for software versins 2.03, 2.04, 2.05, and 2.06. Customers receiving the updated manuals will be sent a cover letter directing them to discard the original user manual and replace it with the updated manual. The April 24, 2009, cover letter will highlight the clarification of directions for use regarding what to do when a low battery indicator is displayed. |  | 46740 | 46740 units | Worldwide distribution: USA, Guam, Puerto Rico, Canada, Mexico, Andorra, UAE, Afghanistan, Australia, Azerbaijan, Belgium, Belarus, Bermuda, Bahrain, Brazil, British West Indies, Switzerland, China, Cyprus, Czech Republic, Germany, Denmark, Algeria, Egypt, Spain, France, French Guiana, French Lesser Antilles, Greece, Haiti, Hong Kong, Hungary, Ireland, Israel, India, Iraq, Italy, Japan, Kenya, Kyrgyz, Lebanon, Lithuania, Luxembourg, Latvia, Monaco, Malaysia, Netherlands, New Zealand, Phillipines, Pakistan, Poland, Portugal, Reunion Island, Russia, Saudi Arabia, Singapore, South Korea, Taiwan, United Kingdom, Uraguay, Yemen, and South America. | Terminated | August 11, 2011 | 839 | 46740 | Not Computer
Z-1655-2009 | 52051 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Blackstone Medical: |  |  | May 21, 2009 | 2009 | Blackstone Medical, Inc. | Product marketed without an approved 510(k) | Finished device change control | Blackstone issued the recall notice to distributors by FedEx and email on May 21, 2009. This notification provides instructions to the field representatives how to return the product to Blackstone Medical. The Dear Doctor (Surgeon) letters are being sent Tues, May 26, 2009. |  | 143 | 63 implant kits, 80 instrument kits | Nationwide | Terminated | April 07, 2015 | 2147 | 143 | Not_Computer
Z-1601-2009 | 50093 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Duracon Total Knee, A-P Lipped Tibial Insert; 11MM |  |  | May 05, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | A box packaged and labeled as a Duracon Total Knee Duration A-P Lipped Tibial Insert 11MM may actually contain a 9MM insert. | Labeling mix-ups | A notification letter and product accountability form were sent via Federal Express on May 5, 2006 with return receipt to branch accounts that had received the affected product. For questions about the recall contact Stryker Howmedica Osteonics Corp. by calling 1-201-831-5000. |  | 8 | 8 | Nationwide Distribution (AL, OH, MN, NM, and CA). | Terminated | July 07, 2009 | 1159 | 8 | Not_Computer
Z-1649-2009 | 51483 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions USA, Inc., Symbia T Series SPECT-CT (Single-Photon Emission Computed Tomography and Computed Tomography) System; |  |  | March 11, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Patient injury can occur when the patient is lying on the patient bed and the automatic collimator exchanger is undergoing a collimator change operation at the launch of an acquisition workflow. Detector motions required for this change can cause collision with the patient or pallet if it is in brain scan position, i.e. the patient or pallet is into the field of view. | Software design | Siemens sent Urgent Field Correction Recall letters dated 1/20/09 to all users of the Symbia S and Symbia T Systems. The accounts were informed of the potential injury to patients while lying on the patient bed and the automatic collimator exchanger is undergoing a collimator change operation at the launch of an acquisition workflow. The customers were advised to perform all necessary collimator changes before the patient is positioned on the patient bed, and to examine each of their workflows that include a Tomo Acquisition activity and make sure the Auto Collimator Change check box on the Camera Parameters Tab is not checked. They also were advised that a Siemens Customer Service Representative would contact them within 60 days to schedule a corrective software upgrade free of charge. Any questions were directed to contact Siemens Medical Solutions USA, Inc. at 1-800-767-2313 (USA) for assistance. The accounts were requested to complete the Mandatory Fax-Back Form acknowledging receipt and understanding of the letter. |  | 148 | 148 units | Worldwide distribution, including USA, Australia, Austria, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Ireland, Italy, Japan, Korea, Kuwait, Netherlands, Norway, Poland, Portugal, Qatar, Saudi Arabia, South Africa, Spain, Sweden and Switzerland. | Terminated | September 11, 2010 | 549 | 300 | Computer
Z-1597-2009 | 51953 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ASC2000 Ambulatory Surgical Table. Model #MTS110003. |  |  | April 17, 2009 | 2009 | Steris Corp | Upon investigation, the firm determined that repetitive rubbing of the hydraulic hoses against the interior of the stainless steel shroud may result in the rupture of the hydraulic hoses causing loss of control of the table surface. | Device Design | A Field Correction Notice dated April 17, 2009 was issued to customers. The letter described the potential problem of the unit. The letter also notified customers of the corrective actions to be taken by the firm. A Steris technician will visit affected accounts to inspect the hydraulic hoses and replace any that show abrasion and will make other preventive corrections. For further information or if you have questions regarding the servicing of your unit, please contact the STERIS Field Service Dispatch at 1-800-288-2805. |  | 412 | 412 units | Worldwide Distribution -- US (states of AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KY, LA, MD, MI, MN, MO, NC, NH, NJ, NY, OH, OK, OR, PA, PR, SC, TN, TX, VA and WV), AUSTRALIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, ECUADOR, GERMANY, HONG KONG, INDIA, INDONESIA, ITALY, KOREA, MEXICO, PHILIPPINES, PORTUGAL, SINGAPORE, SPAIN, SRI LANKA and THAILAND. | Terminated | March 01, 2012 | 1049 | 412 | Not_Computer
Z-1246-2009 | 51675 | No | HTQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer TM-400 Broach Handle; Part No. 96-210-10001. The product is a surgical instrument for use in conjunction with Zimmer TM-400 Trabecular Metal implants during spinal surgery. |  |  | April 29, 2009 | 2009 | Zimmer Inc. | The broach handle may fracture at the threaded tip during use, which would result in surgical intervention to retrieve the broach or the trial as the threaded tip fragment prevents attachment of another handle to retrieve the device. | Device Design | Distributors and user accounts were notified by "Urgent: Medical Device Recall" letter dated April 29, 2009. The letter described the affected product, reason for recall, health risks, and action for consignees. Distributors were instructed to retrieve the devices and to return them to Zimmer no later than May 13, 2009. This instrument recall will result in an inability to implant the associated devices until a new instrument becomes available.For questions or assistance about the recall please contact Zimmer Spine by calling 800-777-7505. |  | 435 | 435 | Nationwide, Austria, Australia, Belgium, Canada, Finland, Germany, Greece, Korea, South Africa, Sweden, Switzerland, Taiwan and the United Kingdom. | Terminated | October 08, 2010 | 527 | 435 | Not_Computer
Z-1611-2009 | 52090 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Mantis Rod Inserter; Non Sterile, Manufactured by Stryker Spine SAS, Cestas, France. Distributed by Howmedica Osteonics Corp; 325 Corporate Drive, Mahwah, NJ 07430. |  |  | April 30, 2009 | 2009 | Stryker Spine | The Mantis Rod Inserter Shaft was found to have deformation of the stainless stell split sphere when subjected to high forces, prohibiting the user from being able to grip the rod implantation and the split sphere may separate and fall into the patient. | Device Design | Stryker Spine issued an "Urgent Product Recall" notification dated April 29, 2009 via Federal Express to all Consignees. Users were instructed to quarantine all affected product and fax a copy of the Product Accountability Form to 201-760-8406. A Stryker sales representative will retrieve the quarantined product. For further information, contact Stryker Spine at 201-760-8206. |  | 175 | 175 in US | Nationwide Distribution. | Terminated | January 23, 2013 | 1364 | 562 | Not_Computer
Z-1654-2009 | 51979 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Clinical Information Portfolio |  |  | May 05, 2009 | 2009 | Philips Healthcare Inc. | Error in therapy scheduling: If scheduled orders are documented and edited prior to the orders' original schedule time, ICIP can create a second instance of the same order marked as pending. A second instance of an order may result in a patient receiving duplicate therapy or patients may not receive their scheduled therapies. | Software design | Philips notified customers by Urgent Medical Device Correction Notification letter on May 6, 2009. The Customer notifications in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography.Customers will be asked to follow the guidelines in the Action to be taken by Customer/User section of the Urgent Medical Device Correction Notice until they receive their software corrections. The corrections will be in the form of software upgrades and will be performed by Philips service personnel.On May 27, 2009, Philips issued the second Urgent Medical Device Correction notification. While customers await the correction of their devices, customers are given the following instructions:Clinicians should not rely solely on the worklist within lCIP to determine actionable orders for the patient until their software is upgraded. Clinical users should always use the medication administration record (MAR) and the administration record in determining the care their patients receive in addition to the critical care worklist. |  | 105 | 105 units | Worldwide distribution, including USA, Australia, Belgium, China, Denmark, France, Germany, Ireland, Italy, Kenya, Malaysia, Netherlands, Singapore, Spain, Switzerland, Taiwan, UAE, and UK. | Completed |  |  | 105 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1582-2009 | 52223 | No | DRN | External counter-pulsating device. | Cardiovascular | Cardiovascular | 510(k) | Cardiomedics CardiAssist External Counterpulsation System, circulatory assist device, Series 4000. |  |  | April 09, 2009 | 2009 | Cardiomedics Incorporated | Cuff may underinflate, or device may stop working during use. Cardioassist software version 4.3.1 utilized a signal from the plethysmograph that was found not to be sensitive enough due to the shape of the waveform in certain circumstances where there was excessive motion artifact. | Software design | Service visit to customer in May 2009 to upgrade and validate CardioAssist 4000 software to version 4.3.5. |  | 1 | 1 | Florida | Terminated | August 27, 2012 | 1236 | 1 | Software | Display/Image | Software update | Software Update
Z-1553-2009 | 51951 | No | GZL | Depth electrode. | Neurology | Neurology | 510(k) | Cyberkinetics Inserter Wand &amp; V2 Trigger Assembly, Expiration: July 14, 2009; Catalog Number: 4468. |  |  | February 04, 2009 | 2009 | Cyberkinetics Neurotechnology Systems, Inc. | Sterility of product cannot be assured | Packaging process control | Cyberkinetics Neurology Systems (CKI) notified the Distributor, Blackrock Microsystem, of the affected device. Blackrock notified the final user by email on February 4, 2009. |  | 2 | 2 units | Nationwide Distribution - State of MA. | Terminated | August 20, 2009 | 197 | 2 | Not_Computer
Z-1606-2009 | 52016 | No | MNS | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics BiPAP Focus Noninvasive Ventilator System. |  |  | September 22, 2008 | 2008 | Respironics California Inc | Power Supply failures have occurred on some ventilators. A discrepancy between the power supply strain relief clip and the power cord connector may allow unwanted movement of the cord and the unit's power supply. A disconnection or intermittent electrical connection between the power cord and the unit's power supply during operation or charging can result in transient power surges, which may cau | Component design/selection | Respironics California, Inc. issued an "Attention - Medical Device Field Correction" letter addressed to the Customer/Distributor dated September 22, 2008. The letter informed consignees of the affected device and instructions for removal, device replacement and return of the Business Response form by fax to 1-760-918-0169.Further question may be addressed to Respironics US Customer Service at 1-9800-345-6443 extension 5 or Respironics International Customer Service at 1-724-387-4000. |  | 2270 | 2,270 Power Cords | Worldwide to Argentina, Australia, Belgium, Brazil, Cambodia, Canada, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Libya, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Portugal, Puerto Rico, Russian Fed., Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, United States and Venezuela. | Terminated | July 06, 2010 | 652 | 2270 | Not_Computer
Z-1571-2009 | 52130 | No | OKE | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | Enforcement Discretion | Tri-State Hospital Supply Corp., Tri-State Centurion Venous Access Port Kits, Qty/case 20, sterile, Code No./Reorder KNV80. Each kit contains 1 sterile heparin flush injection solution as follows: Tyco Healthcare Monoject Prefill 100U/mL Heparin Lock Flush Syringe, 10mL, REF # 8881590121. |  |  | April 08, 2008 | 2008 | Tri-State Hospital Supply Corporation | The kits contain Monoject prefill heparin lock flush syringes, which were recalled by the manufacturer due to contamination. | Material/Component Contamination | Tri State Hospital notified their consignees by letter dated 4/8/08 and instructed to discontinue use of the product and to either discard or return it. |  | 100 | 100 kits | Indiana, Mississippi, Ohio and Tennessee. | Terminated | November 19, 2009 | 590 | 810 | Not_Computer
Z-1653-2009 | 52388 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow Pediatric Two-Lumen Central Venous Catheterization Set with Blue FlexTipÂż Catheter, 5 Fr., 2 Lumen, 13cm catheter length, Catalog number: CS-14502. |  |  | June 17, 2009 | 2009 | Arrow International Inc | The spring wire guide is too large in diameter to pass through the needle, dilator and catheter. | Process design | Arrow notified consignees by letter on/about 06/17/2009. They were instructed to cease use and distribution and quarantine all affected product. They were to contact Arrow's Customer Service Department for a Return Authorization Number. Enclosed was a Recall Acknowledgement &amp; Stock Status Form to be completed and faxed to Arrow International to verify receipt of notification and acknowledge the amount of return. |  | 3019 | 3019 units | Worldwide distribution, including USA, Argentina, Brazil, India, Taiwan and Thailand. | Terminated | December 03, 2009 | 169 | 3019 | Not_Computer
Z-1652-2009 | 52430 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Hitachi ECLOS Computed Tomography (CT) Scanners, Software Version: 2.11 or less. |  |  | April 14, 2009 | 2009 | Hitachi Medical Systems America Inc | Incorrect scale on image: A software error can occur if two (2) different Field-of-View (FOV) settings are registered in a protocol using Multi-Reconstruction. This could cause the image generated to be displayed with the wrong scale factor rendering anatomical measurements incorrect. | Software design | HMSA issued an URGENT - DEVICE CORRECTION notification to their customers on 4/14/2009. The notification describes the potential software error and the condition under which these image distortions can occur. The notification goes on to provide the customer with a work-around procedure which they can use to avoid encountering this issue with their current software configuration(s). The letter states that the firm will be installing updated software on all affected systems beginning in June, 2009, and that a Hitachi Service Rep. will be contacting each customer and making arrangements to come out and install a software upgrade on the affected system(s) at no cost to the customer. Lastly, the letter provides a telephone # to the firm's Applications Helpline: 1-877-228-2777, which the customer can use to obtain additional information, if desired. |  | 4 | 4 units | Nationwide distribution, including Puerto Rico (PR). | Terminated | September 21, 2012 | 1256 | 51 | Computer
Z-1650-2009 | 52436 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging Version V31 Picture Archiving and Communication System, model number 10014063. |  |  | June 01, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Risk of misinterpretation of patient response to treatment: Date displayed may be incorrect. Filtering based on examination date may lead to an incorrect list of examinations. | Software design | syngo Imaging has issued a Customer Safety Advisory Notice, dated May 28, 2009, to affected customers via Update Instruction IM019/09/S. The notice informs the customers of the potential issue and provides instructions to avoid its occurrence. A software update to correct this issue is currently being developed and will be installed on affected systems by sygo's service representatives when it becomes available. |  | 16 | 16 units | Nationwide. | Terminated | March 19, 2013 | 1387 | 16 | Software | Output/Calculation | Software update | Software Update
Z-1577-2009 | 52171 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | XoranConnect, software version 4.1.0.6, Xoran Technologies, Ann Arbor, MI. |  |  | February 03, 2009 | 2009 | Xoran Technologies, Inc. | An incorrect patient name label may appear on the screen displaying the selected patient's scan. | Software design | The firm released software version 4.1.0.7 by uploading it through a remote software upgrade on February 3, 2009. User accounts were notified by letter dated February 9, 2009. Direct questions to Xoran Technologies at 1-800-709-6726. |  | 14 | 14 | Nationwide Distribution (Alabama, Arizona, Colorado, Georgia, Illinois, Louisiana, New York, North Carolina, Oklahoma, Rhode Island, Tennessee, Texas and Utah). | Terminated | November 19, 2009 | 289 | 14 | Software | Display/Image | Software update | Software Update
Z-1542-2009 | 50824 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TumorLoc software for the Brilliance Big Bore (BB) and the CT Extended Brilliance Workstation (EBW), Model #4550 116 02431 (TumorLoc). |  |  | January 28, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | A software error in the TumorLoc software was discovered upon receipt of a customer complaint. There is the potential for data to be lost during reconstruction of images in which slices can either be discarded, or interleaved based upon their series ID. | Software design | Philips Medical issued an 'URGENT -Medical Device Correction' document dated January 28, 2009. This notification informs consignees of the TumorLoc imaging issue, the inherent hazards involved should this problem occur, the circumstances under which this problem may occur and corrective actions which the user/customer should perform in the event that this problem is encountered. It also informs that user/customer that Philips will have their Service Engineers contact the customers within 6 months to install a TumorLoc upgrade. For further questions, contact the Phillips Customer Care Center at 1-800-722-9377. |  | 356 | 356 units | Worldwide Distribution - US, Australia, Austria, Bahrain, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Israel, Italy, Jordan, Libyan Arab Jamahiriya, Lithuania, Malaysia, Netherlands Antilles, Netherlands, New Zealand, Norway, Philippines, Portugal, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Tahiti, Taiwan, Thailand, United Kingdom and Venezuela. | Terminated | September 19, 2012 | 1330 | 356 | Software | Device Operation | Software update | Software Update
Z-1543-2009 | 51252 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS RA1000 Workstation software (for diagnostic image analysis); GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | March 24, 2009 | 2009 | GE Healthcare Integrated IT Solutions | Software anomalies may allow the image orientation to be misinterpreted, posing a potential patient safety issue. | Software design | GE Healthcare Integrated IT Solutions notified their Centricity PACS RA1000 Workstation Software Versions 2.1.x and 3.0.x customers via an Urgent Medical Device Correction notice dated March 2, 2009 of the affected device resulting in patient safety issues. The letter provided users with safety instructions until the software is updated. Consignees are requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 1-847-939-1479.Centricity PACS RA1000 Workstation Software Versions PathSpeed 7.x, 8.x, Centricity 1.X and 2.0.X customers were sent separate letters explaining the recall including instructions. There will be no software updates since those software versions are no longer being supported.For questions, call the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171. |  | 984 | 984 units | Worldwide Distribution - US including Puerto Rico and countries of Brazil, Canada, Chile, Ecuador, Mexico, Australia, China, Hong Kong, Japan, Korea, Malaysia, New Zealand, Singapore, Taiwan, Austria, Belgium, Switzerland, Germany, Denmark, Egypt, Spain, France, Great Britain, Ireland, Israel, Italy, Kuwait, Lebanon, Malta, Netherlands, Portugal, Saudi Arabia, Sweden and Turkey. | Terminated | September 06, 2010 | 531 | 984 | Software | Display/Image | Software update | Software Update
Z-1576-2009 | 52170 | No | MSH | Surgical apparel. | General &amp; Plastic Surgery | General Hospital | 510(k) | Inovel N95 HealthCare Particulate Respirator, Model/Catalog Number: FRN95-SEZ, Expiration Date: 07/13/2010. |  |  | July 23, 2008 | 2008 | Inovel LLC | Display boxes labeled as FRN95-MLEZ, Medium/Large Size Respirators were inadvertently mixed with Display Boxes labeled for FRN95-SEZ Small Size Respirators. | Packaging process control | The recalling process was initiated on July 23, 2007. The affected distributor was notified of the recall through a certified recall letter with an attached Customer Response card. The distributor was instructed to notify and obtain recall product from their customers and return all affected product to Inovel. Direct questions about the recall to Inovel at 1-866-546-6835. |  | 80 | 80 cases | US (Illinois only). | Terminated | July 17, 2009 | 359 | 80 | Not_Computer
Z-1552-2009 | 51785 | No | OEU | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | Enforcement Discretion | Medtronic Ideal MAXPAC Total System, Catalog Numbers: TS1262R6, TS1262R7, TS1420R, TS1420R1, TS1442R, TS1446R, TS1464T, TS1470R, TS1470T1, TS1480R3, TS1480R5, TS1480R6, TS1487R, TS1487R2, TS8040R28, TS8108R6, TS8163R10, TS8387R5, TS9412R3, TS9453R28, TS9764R9, TS9800R28, TS9800R29, TS9800R31, TS9987R33 and TS9987R34. |  |  | March 23, 2009 | 2009 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | Medtronic has identified a small number of One Source Kits and MAXPAC/Total System lots with incorrect Use Before Dates on the outer kit package label. The Use Before Date on the outer label of the kit/pack does not accurately reflect the Use Before Date of the components within the kit/pack. For affected product, the outer label lists a date which exceeds the earliest Use Before Date of the com | Packaging process control | An "Urgent Medical Device Recall Notice" dated March 23, 2009 was sent to customers describing the affected device. Consignees were asked to quarantine product, complete and sign a certificate indicating acknowledgement of the notification and the amount of affected product at the facility, and to return the affected product.For further questions, contact Medtronic at 1-763-514-9605. |  | 1363 | 1363 kits total (1215 US, 148 OUS) | Worldwide Distribution -- US including states of AZ, CA, CO, CT, DE, FL, HI, IL, IN, LA, MI, MN, NC, NJ, NM, NV, NY,OH, OK, PA, TN, TX, and WI and countries of Australia, Japan and Venezuela. | Terminated | May 22, 2012 | 1156 | 1363 | Not_Computer
Z-1605-2009 | 52036 | No | IMB | Moist steam cabinet. | Physical Medicine | Physical Medicine | 510(K) Exempt | DERMALIFE Takara. Part #E30006-02 - Takara; Renew 2G, w/Pedestal. |  |  | June 22, 2007 | 2007 | Sybaritic, Inc | A correction or removal, conducted to reduce a risk to health posed by a device, was not reported in writing to FDA.Sybaritic, Inc. conducted a recall on DERMALIFE 2G equipment (SpaJet, Spa Oceana, Spa Feng Shui) because of IR Fan failure which resulted in the overheating and failure of the IR emitter and surrounding assembly. | Device Design | Sybaritic, Inc. issued a "Safety Alert" dated April 15, 2008 informing "All Owners/Users of DermaLife 2G Equipment" of the affected device. The letter advised consignees on device use until the mandatory Filed Correction Kit is delivered. For further questions, contact Sybaritic, Inc. at 1-800-445-8418. |  | 10 | 10 | Worldwide Distribution -- US including states of CA, IA, MN, NE, NJ, NY, TX and WA and countries of Macedonia, Hungary, Greece, South Africa, Italy, Hong Kong, United Arab Emirates, Greece, Croatia, Australia, Kuwait, Turkey, Mexico, Belgium, England, Barbados, France, Japan, Lebanon, Moldova, Poland, Portugal, Romania, Russia, Spain and Thailand. | Terminated | July 05, 2012 | 1840 | 363 | Not_Computer
Z-1569-2009 | 52099 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Amico E-Cylinder Canister Holder, distributed by Stryker Communications, Flower Mound, TX 75028. |  |  | May 19, 2009 | 2009 | Stryker Communications Corp | Design flaw on the E-Cylinder canister holder may result in early fatigue and failure of the bracket screws, which could cause the cage to become detached from its mount and fall. | Device Design | An "Urgent: Device Recall" letter dated May 19, 2009 was issued to customers. Customers were advised to discontinue use of all affected E-cylinder holders. Consignees were asked to return the "Product Accountability Form" included with the recall letter, after which consignees will be provided with an improved cylinder holder at no cost. Direct questions to Stryker Communications at 1-972-410-7100. |  | 66 | 66 units. | Nationwide and Canada. | Terminated | April 26, 2010 | 342 | 66 | Not_Computer
Z-1567-2009 | 52042 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Smith &amp; Nephew Endoscopic Disposable Blades:TURBOWHISKER BLD,EP-1,4.5MM(6) |  |  | May 07, 2009 | 2009 | Smith &amp; Nephew, Inc. Endoscopy Division | Product sterility is compromised due to breach in the sterile packaging. | Pending | Smith &amp; Nephew, Inc. issued a notification letter dated May 5, 2009 via Federal Express to User Facility and/or Sales Representatives on May 7, 2009 and an Email was sent to Sales Representatives. Accounts were requested to 1)complete the form and fax to Smith &amp; Nephew, Inc. at 1-508-261-3636 and 2) return (using the firm's return authorization number) affected product to Smith &amp; Nephew, Inc. Direct questions about this recall to Smith &amp; Nephew, Inc. by calling 1-508-261-3655. |  | 948 | 948 units | Worldwide Distribution -- US, Switzerland, Canada, United Arab Emerites, Belgium, Mexico, South Africa, Great Britain, Australia, Korea, Japan,Germany, Israel, India, Taiwan, Greece, Italy, Turkey and Chile. | Terminated | May 08, 2012 | 1097 | 13638 | Not_Computer
Z-1613-2009 | 52092 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Biolox-Delta Modular Ceramic Head, 28 mm head diameter, plus 5 neck, type 1, taper, sterile, Biomet Orthopedics, Inc., Warsaw, IN; Catalog Number 12-115112. |  |  | March 16, 2009 | 2009 | Biomet, Inc. | The firm has received reports of fractured heads, requiring revision surgery. | Process control | Biomet Orthopedics issued an "Urgent Medical Device Recall Notice" dated March 16, 2009 to users describing the affected device. Consignees were instructed to locate and discontinue use of the product and quarantine until a sales representative can arrange for removal. Consignees were also asked to complete and return the attached "Fax-Back Response Form" to Biomet at 1-574-372-1683.Further questions can be directed to Biomet Orthopedics at 1-800-348-9500. |  | 174 | 174 | Worldwide Distribution -- US including PR and states of AZ, AL, AR, CA, CO, FL, IL, IN, KY, KS, MA, MI, LA, NY, SD, GA, MN, PA, NJ, MD, MI, OH, ME, TX, WA, WI and VA and countries of Australia, Belgium, Canada, Costa Rica, Finland, Netherlands and South Wales. | Terminated | October 06, 2010 | 569 | 335 | Not_Computer
Z-1679-2009 | 51713 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific TELIGEN, Implantable Cardiac Defibrillator models E102, E110, the following models are not available in the US: models F102, F110. |  |  | March 23, 2009 | 2009 | Boston Scientific CRM Corp | Lead connection issues or fracture increases probability of inappropriate therapy: Inappropriate shocks, inappropriate anti-tachycardia pacing therapy, or inappropriate inhibition of needed bradycardia pacing therapy. | Device Design | Boston Scientific sent an Urgent Medical Device Information letter, dated March 23, 2009, to physicians following implanted devices. The letter described the issue, the rate of occurance and provided Patient Management Recommendations. |  | 49148 | 49,148 total | Worldwide distribution: USA, Guam, Puerto Rico, Aruba, Austria, Belgium, Brazil, Canada, Cayman Islands, Colombia, Cyprus, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Ireland, Israel, Italy, Kuwait, Latvia, Lebanon, Luxembourg, Mexico, Netherlands, New Caledonia, Norway, Philippines, Poland, Polynesia, Portugal, Reunion, Russia, Saudi Arabia, Slovakia, South Africa, Spain, Sweden, Switzerland, United Kingdom, and Venezuela. | Terminated | October 18, 2012 | 1305 | 49148 | Not_Computer
Z-1673-2009 | 52461 | No | KYB | Intraocular lens guide. | Ophthalmic | Ophthalmic | 510(K) Exempt | Optiflex Lens Positioning Unit (LPU) Product Code: 10970, for use in vitreoretinal ophthalmic surgical procedures. |  |  | January 23, 2009 | 2009 | Volk Optical Inc | Lens may fall onto eye surface during procedure. Ophthalmologic device had a loose set screw on the Ball-Seal bushing assembly. As a result, the set screw becomes disengaged from the set screw groove in the bushing assembly and the Lens Positioning Unit slides downward. | Process design | Volk Optical sent a 'MEDICAL DEVICE RECALL' letter to their sole customer on 1/23/2009. The letter requests the return of the remaining two (2) units in the customer's possession (this firm experienced the device failure and reported the consumer complaint which prompted the recall; and had already returned the affected unit to the Volk for refund or replacement). The letter provides Volk's Fedex and UPS numbers for the customer to use to return the recalled units and also provides the E-mail address for the firm's contact person concerning this recall should additional assistance be required. |  | 3 | 3 units | USA distribution to a single customer located in MO. | Terminated | April 12, 2012 | 1175 | 3 | Not_Computer
Z-1618-2009 | 52191 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Revision Instruments; 1,2,7,8 TGC Upper Tray; Non Sterile, |  |  | May 15, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | The screen print for the Valgus Adapter catalog numbers on the Triathlon Trial Cutting Guide Instrument Trays is incorrect. | Process change control | Stryker Orthopaedics issued an "Urgent Product Correction" notice dated May 15, 2009 via Federal Express addressed to Branch/Agency Manager/Quality Contact informing them of the affected device. The letter outlined instructions for correction and requested consignees to complete and return a Product Correction Acknowledgment Form by fax to 1-201-831-6069.Questions can be directed to Colleen O'Meara, Manager, Divisional Regulatory Reporting at 201-831-5970. |  | 27 | 27 | Worldwide Distribution -- US and Canada. | Terminated | April 05, 2010 | 325 | 53 | Not_Computer
Z-1245-2009 | 51934 | No | IMB | Moist steam cabinet. | Physical Medicine | Physical Medicine | 510(K) Exempt | Hydration Station with Pedestal (Part # E30003-01), Hydration Station with Pedestal &amp; DVD (Part # E30003-01D), Sybaritic, Inc., 9229 James Avenue South, Minneapolis, MN 55431. |  |  | May 15, 2007 | 2007 | Sybaritic, Inc | Sybaritic conducted a recall on the Hydration Station 806 because of the potential for hot water to spurt through steam ports into unit and cause burns. Also, potential for customers to inadvertently place their legs over the side wall of the unit directly onto the steam ports and cause burns.There were 18 injuries and 8 malfunctions reported. | Device Design | Sybaritic, Inc. issued a letter dated May 15, 2007 addressed to "Hydration Station Customer". The letter informed the customer of the new 2007 Hydration Station Upgrade Improvement Kits including instructions and advised the customer to coordinate the installation of the components with their Technical Service Department.Further questions can be addressed to Sybaritic, Inc. Technical Administration at 1-800-445-8418 extension 373. |  | 370 | 370 | Worldwide Distribution - US including states of AL, AZ, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MI, NE, NV, NH, NJ, NM, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, WA, WV and WI and CANADA. | Terminated | July 05, 2012 | 1878 | 370 | Not_Computer
Z-1624-2009 | 51903 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ConforMIS iUni ( Unicondylar Knee System iUni Right Lateral Implant Kit Product Code: M57220600040. |  |  | April 23, 2009 | 2009 | Conformis Inc | Microscopic cracks potential for pre-mature revisions in patients implanted with the system. | Process control | ConforMIS contacted accounts and presented the surgeon with the customer communication letter dated April 23, 2009, customized for the surgeon by serial number and patient identification. Product not yet implanted will be recovered. A patient evaluation and monitoring protocol will be recommended for potentially affected product already implanted.Contact ConforMIS for further questions at 1-781-345-9164. |  | 1 | 1 unit | Worldwide Distribution - US including states of CA, LA, MI and TX and Germany. | Terminated | October 14, 2009 | 174 | 17 | Not_Computer
Z-1554-2009 | 51986 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Waste bottles Part Numbers/Article Numbers: 905-590/D512 and 905-802/D513. |  |  | April 16, 2009 | 2009 | Radiometer America Inc | When removing the waste bottle from the analyzer waste material may leak both from the analyzer and from the waste bottle. | Process design | Radiometer America Inc. issued a recall notification letter dated May 5, 2009 to all customers containing information on replacement of all potentially defective waste containers and instruction on how to handle any mounted waste containers.For further questions, contact Radiometer America Inc. at 1-800-736-0600. |  | 905590 | 905-590 (D512): 35,552 bottles; 905-802 (D513): 45,204 bottles | Worldwide Distribution to US and Canada. | Terminated | May 09, 2012 | 1119 | 905590 | Not_Computer
Z-1705-2009 | 52521 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Paradigm Quick-set Infusion Sets, Models MMT-396, MMT-397, MMT-398 and MMT-399. |  |  | June 29, 2009 | 2009 | Medtronic MiniMed | Approximately 2% of the infusion sets in the affected lots may not allow the insulin pump to vent properly. Venting is necessary to equalize the pressure in the syringe compartment with the surrounding atmosphere. If the vent does not work properly, this could potentially result in too much or too little insulin being delivered. | Process control | Medtronic issued an "URGENT MEDICAL DEVICE RECALL" notification letter dated July 6, 2009 to health care professionals, followed by notification letters dated July 7, 2009 to distributors, inactive and active customers. Letters to healthcare professionals and distributors were sent via UPS 2-Day Air with signature confirmation of delivery required. On July 8, 2009 and July 9, 2009, a notification letter, along with replacements sets, were mailed via UPS to active patients or customers (who have ordered since January 2009). The notification letter to patients (inactive customers) who last ordered prior to January 1, 2009 were issued via First Class U.S. Mail. All letters to patients include a preaddressed, postage paid card to confirm that the patient has received the recall notice and to identify whether or not they have unused sets to return. As described in the patient letters, patients have the option to confirm receipt of the notification and to order replacement infusion sets online or via an automated telephone voice response system.Distributors who purchased affected Paradigm Quick-set infusion sets must either notify all patients to whom they have provided infusion sets or provide patient names and addresses to Medtronic MiniMed so that the firm can carry out the notification.The firm's website, www.medtronicdiabetes.com/lot8, is available for customer concerns and questions not addressed in the recall letter. The website is also the most efficient way to exchange Lot 8 Quick-set infusion sets and/or place new supply orders.In the event the above website is not accessible or for additional questions, please call 1-800-345-8139 at Medtronic Diabetes 24- hours a day. Doctors who would like to speak with a Medtronic Diabetes medical officer may call 1-818-576-4211, 24-hours a day. Report product problems by calling 1-800-345-8139 at any time.Medtronic issued a press release on July 10, 2009. |  | 28007256 | 28,007,256 units | Worldwide Distribution -- US, Canada, Mexico, Bermuda, Brazil, United Kingdom, Ecuador, Paraguay, Kuwait, Germany, El Salvador and Turkey. | Terminated | November 08, 2010 | 497 | 28007256 | Not_Computer
Z-1620-2009 | 51345 | Yes | LWP |  |  | Cardiovascular | PMA | Guidant ContakÂż RenewalÂż 3 CRT-D Model H170; 3CRT-D, Model H175; 3HE CRT-D, Model H177; 3HE CRT-D, Model H179; 4CRT-D REF H190; 4CRT-D REF H195; 4HE CRT-D REF H197; 4HE CRT-D REF H199; 4AVT CRT D, REF M177; 4AVT HE CRT-D REF 179. Cardiac Resynchronization Therapy Defibrillator, Sterile EO, Guidant Corporation, 4100 Hamline Avenue North, St Paul, MN 55112-5798. |  |  | March 04, 2009 | 2009 | Boston Scientific CRM Corp | In April 2007, Boston Scientific CRM communicated with physicians regarding the potential for reduced Elective Replacement Indicator (ERI) to Battery End of Life (EOL) time in a subset of implantable defibrillators due to degradation of a low-voltage capacitor. Since that time, the April 2007 advisory population has not experienced any clinically significant changes to either the rate of occurrenc | Nonconforming Material/Component | Boston Scientific CRM notified Doctors via sales representatives of an "Urgent Medical Device Information" letter dated March 4, 2009. The letter described the problem and included Description of Issue, Rate of Occurrence, Patient Management Recommendations, Devices Affected, Warranty Program &amp; Further Information. A Boston Scientific " Product Advisory" letter dated March 4, 2009, was also made available to physicians to distribute to patients. The letter informed the patient that their defibrillator was included in a group of devices that might experience a performance issue. For further questions, Contact Boston Scientific CRM at 1-800-227-3422. |  | 112 | 112 | Worldwide Distribution - US including PR and states of AL, AZ, AR, CA, CO, CT, DC, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI and WY and countries of ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, IRELAND, ITALY, JAPAN, LEBANON, NETHERLANDS, NEW ZEALAND, POLYNESIA, PORTUGAL, SPAIN, SWITZERLAND, TURKEY and UNITED KINGDOM. | Terminated | October 04, 2012 | 1310 | 1383 | Not_Computer
Z-1548-2009 | 51485 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Gendex VixWin Platinum Imaging Software, Versions 1.0, 1.1 and 1.2; Gendex Dental Systems, 901 West Oakton Street, Des Plaines, IL 60018. |  |  | March 06, 2009 | 2009 | Gendex Dental Systems | There is a compatibility issue when the VixWin Platinum versions 1.0/1.1/1.2 software is used in conjunction with Easy Dental, PracticeWorks, EagleSoft and SoftDent practice management software, resulting in patient files that may contain images that are mixed between patients. | Software design | Gendex issued an "Urgent Medical Device Report" dated March 6, 2009 to the direct account dealers of the affected device via first class mail. The noticed described the issue and provided a software update directed to fixing the problem. In addition, customers were asked to complete and return an Acknowledgement Form by fax to 1-847-718-0601.Further questions should be directed to the Gendex Technical Support team at 1-888-275-5286. |  | 1256 | 1,256 units | Nationwide Distribution. | Terminated | January 07, 2011 | 672 | 1256 | Software | Output/Calculation | Software update | Software Update
Z-1674-2009 | 52129 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Essential Medical Supply brand Essential Rollator with Loop Hand Brakes, models W1700B (Blue) and W1700R (Red), Four Wheel Aluminum Walker/Rollator |  |  | May 08, 2009 | 2009 | Essential Medical Supply, Inc. | The caster may separate from the frame completely causing injury to the user. | Device Design | Essential Medical Supply Inc issued a Notice of Quarantine Letter to all customers on 04/30/09, identifying affected products as W1700B &amp; W1700R Essential Rollators, Serial #'s KLD-E07090001 through KLD-E08020570, received after 11/05/2007. Firm issued a Device Recall Letter to all direct customers on 05/08/09 instructing them to do the following: (1) segregate and quarantine any affected product; (2) identify and notify any customers that may received the affected products, instructing them to return the product to their dealer; and (3) either return the affected products to Essential Medical Supply or destroy the products and send proof of destruction. |  | 2265 | 2265 | Nationwide distribution. | Terminated | September 15, 2010 | 495 | 2265 | Not_Computer
Z-1599-2009 | 52172 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Cmax 110v Surgical Table Hand Control. |  |  | May 12, 2009 | 2009 | Steris Corp | A limited quantity of the Cmax Hand Controls, P150832-500, were manufactured without Loctite thread locker on the screws that attach the mounting clip to the hand control case. This clip provides a means of storing the hand control on the table side rails. It is possible that screws without Loctite can work loose and the clip can become separated from the hand control. If both screws back out c | Process control | Steris Corp. issued an "Urgent Product Correction" notice dated May 12, 2009 informing consignees of the affected device and that a Steris Service Representative with be in contact to make the necessary inspection and repairs to the device. Steris Field Service Dispatch can be reached at 1-800-288-2805.For further questions, contact Steris at 1-800-548-4873 extension 27318. |  | 217 | 217 Units | Worldwide Distribution -- U.S. states of AL, AR, CA, CO, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NE, NH, NM, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, WI and WV and countries of Canada and Taiwan. | Terminated | February 26, 2013 | 1386 | 217 | Not_Computer
Z-1711-2009 | 52431 | No | EXQ | Urodynamics measurement system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Laborie Medical Technologies Urodynamic Catheters, Twin Lumen 6 Fr , 40cm For Cystometry. Catalog Number: CAT-206. |  |  | April 09, 2009 | 2009 | Laborie Medical Technologies Cor | Laborie Medical Technologies reports that there was a mix up in the labeling and packaging of their Urodynamic Catheters. There were 8Fr catheters mixed inside the 6Fr catheter packages. | Packaging process control | Laborie Medical Technologies notified customers by telephone/email beginning April 10, 2009 advising them of the mislabelled product.For further questions, contact Laborie Medical Technologies at 1-800-522-6743. |  | 28 | 28 boxes (10's) | Nationwide Distribution -- CA, GA, IL, KS, MD, PA, OH, OK, UT and WA. | Terminated | August 20, 2009 | 133 | 28 | Not_Computer
Z-1710-2009 | 52453 | No | HKS | Gonioscopic prism. | Ophthalmic | Ophthalmic | 510(K) Exempt | G-4 HighMag Flange, Model #VG4HM. Standard Gonio Lens for High Magnification Static and Dynamic Gonioscopy. |  |  | June 30, 2008 | 2008 | Volk Optical Inc | The ring of the Gonio lenses distributed was engraved as, 'G4 High Mag Gonio LNF'; however, the engraving should have read simply: 'G4 High Mag Gonio.' | Process design | Volk Optical Inc. contacted their customers and notified them of affected device via E-mail, letter, and Fax transmittals starting June 30, 2008. The recall notification(s) describes the engraving error and instructs the users on steps to replace the affected product.For further questions, contact Volk Optical Inc. at 1-800-345-8655 extension 443 . |  | 15 | 15 | Worldwide Distribution -- Great Britain, China and the U.S. | Terminated | March 09, 2012 | 1348 | 15 | Not_Computer
Z-1598-2009 | 52283 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fusion Omni-Tome, Rx Only, Wilson Cook Medical, GI Endoscopy, 4900 Bethania Station Road, Winston-Salem, NC 27105. |  |  | June 12, 2009 | 2009 | Cook Endoscopy | The wire control port of these devices may be missing due to omission of a manufacturing activity. | Process control | The distributor (Cook Ireland) was instructed to contact their customers and coordinate the return of all recall product to Cook Endoscopy. All devices returned as part of this recall initiative will be dispositioned to scrap and destroyed.Contact the Customer Quality Assurance Department at Cook Endoscopy for more information at cws.complaints@cookmedical.com or 1-336-744-0157. |  | 20 | 20 units | International Distribution - Italy, Netherlands, Germany and Spain. | Terminated | April 06, 2010 | 298 | 20 | Not_Computer
Z-1680-2009 | 50901 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Super Secur Fit Plus. |  |  | January 06, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Product deviation : Specific lots of sprayed HA hip stems did not meet Stryker's Internal Material Specification for tensile bond strength and crystallinity. | Process control | An "Urgent Product Recall" letter was sent via Federal Express on January 6, 2009 with return receipt to Stryker Branches/Agencies, Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. The letter described the affected product, issue, potential hazards and risks. Stryker Branches and Hospital Risk Management were asked to examine their inventory for affected product. Stryker Representatives were instructed to contact customers to reconcile and return recalled product to Stryker Orthopaedics. All letters include a response form to be completed and returned indicating that they have been informed of the recall and the amount of product on hand, if any. Direct questions about the recall to Stryker Orthopaedics at 1-201-831-5718. |  | 1628 | 1,628 total stems, all varieties. | Nationwide Distribution. | Terminated | June 25, 2012 | 1266 | 1628 | Not_Computer
Z-1607-2009 | 52072 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Offset Adaptor Trial - 2MM, Triathlon Revision Instruments. Non-Sterile; Howmedica Osteonics Corp: 325 Corporate Drive, Mahwah, NJ 07430. |  |  | April 16, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Triathlon Offset Adaptors may seize during surgery and the OR staff may be unable to disassemble the instrument using the removal tool. | Process control | "Urgent Product Correction" letters dated April 16, 2009 were sent to all Stryker Representatives, Hospital Rick Management Departments, Hospital Chief of Orthopaedics, Surgeon and countries by Federal Express. The notice explains the Issue, Potential Hazards and Risk Mitigation Factors of the affected product(s) including a request to complete and return the Product Correction Form.For further questions, contact Stryker Orthopedics, Division of Regulatory Reporting at 201-831-5970. |  | 2498 | 2498 units total, all sizes. | Worldwide Distribution. | Terminated | April 05, 2010 | 354 | 2498 | Not_Computer
Z-1699-2009 | 52227 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Allen medullary cement plugs, 1 - 20 dia. flange/10 mm dia. core, 1 - 24 mm dia. flange/12 mm dia. core (polyethylene with barium sulfate) with inserter, sterile, Zimmer, Warsaw, Indiana; Catatog Number: 00801102001. |  |  | June 08, 2009 | 2009 | Zimmer Inc. | The seal on the sterile barrier pouch may be inadequate, resulting in lack of assurance of sterility. | Packaging process control | Zimmer Orthopedic Surgical Products issued an "Urgent: Medical Device Recall" notice dated June 8, 2009 informing users of the affected devices and actions to be taken including the return of product and acknowledgement of receipt of notice by returning the Recall Certification Form by fax to 1-330-364-0974. |  | 133 | 133 | Worldwide Distribution -- United States, Canada and Japan. | Terminated | December 15, 2009 | 190 | 258 | Not_Computer
Z-1596-2009 | 52159 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Stat (TM) Super-Fine (TM) Pen Needles and Allison Medical, Inc. Sure Comfort Pen Needles with Silicone Tip, for human use with insulin pen injector device to administer insulin subcutaneously. |  |  | May 29, 2009 | 2009 | Stat Medical Devices, Inc. | Stat Pen Needles with silicone coating have not been cleared by FDA for use with silicone coating which could significantly affect the safety and effectiveness of the device. | PMA | An "Urgent: Medical Device Recall Notification" letter dated June 18, 2009 was issued to direct customers requesting return of affected product and requesting they send notification to any customers which may have received the affected product from them. Direct questions about the recall to Stat Medical Devices by calling 1-978-774-1469. |  | 6424100 | 6,424,100 total pen needles | Worldwide Distribution -- US (CO and Puerto Rico), Canada, and Saudi Arabia. | Terminated | September 07, 2010 | 466 | 6424100 | Not_Computer
Z-1636-2009 | 51273 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Passive Clip Slings; washable sling; Manufactured in Belgium for Medibo NV, Heikant 5, 3930 Hamont Achel, Belgium; these slings are intended for use with ARJO and BHM Passive Patient Lifts and Ceiling Lifts with four-point hanger bars intended for use with clip slings. |  |  | February 26, 2009 | 2009 | ArjoHuntleigh | If the sling clips are not applied or are applied incorrectly, there is the potential for the patient to fall, resulting in injury to both the patient and/or the caregiver. | Other | An "Urgent Field Safety Notification" letter dated January 30, 2009 to all customers with ARJO/BHM Passive Patient Lifts and Ceiling Lifts with four-point hanger bars via UPS 2nd day on the February 26, 2009. The notification letter informed the accounts of the potential for clip detachments if the sling clips are not applied to the hanger bar attachment lugs. The accounts were instructed to do the following: 1) Apply the stickers provided to the passive patient lift and/or ceiling lift hanger bars as shown in the enclosed Sticker Application Instructions sheet; 2) Retrain lift operators in accordance with the enclosed instructional DVD, the revised Passive Clip Sling Operating and Product Care Instructions manual and the revised Flites Clip Operating and Product Care Instructions manual; and 3) Make lift operators aware of the Clip Sling/ Flites Clip Application Instructional poster and its location within their facility. The accounts were also instructed to complete and fax back to Arjo the enclosed Customer Response Form, providing the facility name and address, contact name and phone number, acknowledging receipt and understanding of the Field Safety Notice.If you have any questions or require any additional information, contact the ARJO, Inc. Quality, Customer Support, or Technical Support Departments via e-mail at quality@arjousa.com, via phone at 1-800-323-1245, or via fax at 1-888-594-2756. |  | 103108 | 103,108 total slings | Nationwide. | Terminated | February 25, 2011 | 729 | 103108 | Not_Computer
Z-1802-2009 | 52125 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Medtronic INTREPID Spinal System Coverplate, size: Small, Part Number: REF 7961330; medium, REF 7962330; large, REF 7963330; STERILE R. |  |  | February 06, 2009 | 2009 | Medtronic Sofamor Danek USA Inc | The firm received complaints regarding the use of the product in patients with poor bone quality, translational instability, difficulty implanting the cover plate and case of cover plate detachment. | Device Design | Medtronic Sofamor Danek USA, Inc. instructed Sales Representative on starting February 6, 2009 by phone to immediately return product from the field.The firm further issued a "Subject: Medtronic Intrepid Intervertebral Body Fusion Device Recall" notice dated April 2009. The recall notices were mailed directly to Risk Managers and Surgeons, with Recall Questionnaires and Surgeon Confirmation forms included respectively.For further questions, contact your Medtronic Sales Representative or the Medtronic Global Quality Department at 1-800-876-3133 extension 6333. |  | 1059 | 1059 units | Worldwide Distribution -- United States, Australia, Germany, The Netherlands, and South Africa. | Terminated | April 19, 2010 | 437 | 9087 | Not_Computer
Z-1588-2009 | 52038 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ansell, Derma Prene Ultra, Powder-Free, Smooth Surgical Gloves, REF 8517, Size 8 1/2, 4 boxes of 50, Sterile, Distributed by Ansell Healthcare, 1635 Industrial Road, Dothan, AL, 36303, USA, Made in Sri Lanka. |  |  | May 01, 2009 | 2009 | Ansell Healthcare Products LLC | Premature degradation which could result in tearing during use. | Process control | The firm initiated the recall on May 1, 2009 by letter to Health Canada and by letter dated May 4, 2009 to its US and Canadian consignees. The letters described the affected product, issue and actions for customers. The recall was extended to the hospital and clinic level requesting that the product be returned. Direct questions to your Ansell Healthcare Products LLC representative. |  | 560000 | 560,000 total units--all lots. | Nationwide Distribution and Canada. | Terminated | April 12, 2010 | 346 | 560000 | Not_Computer
Z-1795-2009 | 52575 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx/MRxE Defibrillator/Monitor Models: M3536A (HeartStart MRx) and M3536J (HeartStart MRxE). Note: M3536JIMRxE was sold to Japan only. |  |  | July 07, 2009 | 2009 | Philips Healthcare Inc. | External transport use environments, the mechanical/electrical connection between the padstherapy cable including pads/CPR cable and the MRx/MRxE may wear at an increased rate ultimately prevent the device from sensing that the padstherapy cable is connected, or, cause the device to inappropriately identify external paddles or internalpaddles when the pads therapy cable is connected | Device Design | Philips Healthcare issued an "Urgent - Voluntary Medical Device Correction" notice dated July 2009 providing information on the affected device, namely, how to identify the affected devices and instructions on actions to be taken. Consignees are asked to contact their local Philips representative to confirm receipt of the recall letter and to arrange for service.Contact Philips Healthcare for further questions at 1-800-722-9377. |  | 11579 | 11,579 units | Worldwide Distribution -- United States, Afghanistan, Albania, Australia, Austria, Bahrain, Brazil, China, Columbia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kenya, Latvia, Lebanon, Libyan Arab Jamaica, Lithuania, Netherlands, New Caldonia, New Zealand, Norway, Panama, Peru, Poland, Portugal, Reunion, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, United Arab Emirates, United Kingdom and Venezuela. | Terminated | February 19, 2013 | 1323 | 11579 | Not_Computer
Z-1616-2009 | 52169 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Alumina Ceramic Femoral Head, 12/14 taper, 32 mm diameter, neck length 0 mm, sterile, Zimmer, Warsaw, IN; Catalog Number: 6428-32-02. |  |  | May 29, 2009 | 2009 | Zimmer Inc. | The corner tab of the inner compartment lid was not folded under before the outer lid was sealed, resulting in lack of assurance of sterility. | Packaging process control | Zimmer, Inc. issued "Urgent: Device Recall" notices dated May 29, 2009 informing Consignees of the affected devices. Users were instructed to locate and return affected product and complete and return an Inventory Return Certification form via fax to the firm at 1-547-372-4265.For further questions, contact Zimmer, Inc. at 1-800-613-6131. |  | 147 | 147 of both products | Worldwide Distribution -- United States, Brazil, Canada, Chile, Czech Republic, Finland, France, Germany, Italy, Korea, Lebanon, Switzerland and United Kingdom. | Terminated | December 22, 2009 | 207 | 147 | Not_Computer
Z-1692-2009 | 52715 | No | MEQ |  |  | Gastroenterology/Urology | PMA | Prolieve Thermodilatation Catheter Kit, a Transurethral Microwave Therapy (TUMT) System, for BPH (Benign Prostatic Hyperplasia) Thermotherapy, Material Numbers M0068808022 and M0068808170. |  |  | July 20, 2009 | 2009 | Boston Scientific Corporation | Balloon leaks: Boston Scientific has received complaints from the field involving product leaks associated with the anchor balloons and compression balloons. Correspondence between BSC and FDA has led to BSC's decision to remove the Prolieve Catheter Kits from the field. | Process design | Boston Scientific sent Customer notification letters via Federal Express Priority mail on July 20, 2009. A total of 812 US consignees were notified and a total of 26 consignees in Puerto Rico. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product in their inventory. Follow-up communications will be made to the accounts that do not respond to this inquiry. |  | 26935 | 26,935 kits | Nationwide distribution, including Puerto Rico. | Terminated | March 04, 2010 | 227 | 26935 | Not_Computer
Z-1806-2009 | 52705 | No | MAF |  |  | Cardiovascular | PMA | POWERSAIL Coronary Dilatation Catheter: 2.75x18mm (US) |  |  | June 17, 2009 | 2009 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Distal shaft of the POWERSAIL coronary Dilatation Catheters exhibits damage that may result in a leak of contrast material and functional failures. | Process control | Abbott Vascular sent an Urgent Device Recall letter (w/return form) to customers on June 17, 2009 requesting return of identified products to Abbott Vascular Territory Manager. Abbott Vascular issued a press release dated July 28, 2009. |  | 557 | 557 total devices | Worldwide: United States, Australia, Belguim, Canada, China, Czech Republic, France, Germany, Greece, India, Italy, Japan, Jordan, Mexico, Netherlands, Norway, Poland, Puerto Rico, Spain, Sri Lanka, Switzerland, Thailand, and United Kingdom. | Terminated | July 06, 2010 | 384 | 557 | Not_Computer
Z-1635-2009 | 50800 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Datex-Ohmeda S/5 Anesthesia Delivery Unit |  |  | October 30, 2008 | 2008 | GE Healthcare | GE Healthcare has received reports of breathing circuit or tubing misconnections of the Anesthesia Delivery Unit (ADU), that may impact patient safety. | Other | A GE Healthcare "Urgent Medical Device Correction" letter was sent to consignees on 1/16/09. The letter was addressed to Health Care Administrator, Chief of Anesthesia and Director of Clinical Engineering. The letter described the Safety Issue; Affected Product Details; Safety Instructions, which include clinical personnel being properly trained on the correct connections of breathing circuit tubing and proper execution of System Checkout according to the ADU URMs; Product Correction and Contact Information. |  | 6621 | 6,621 | USA including states of AL, AK, AZ, AR, CA, CO, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NV, NH, NM, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, VT, VA, WA, and WI and countries of Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Bolivia, Bosnia &amp; Herzegovina, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Egypt, England, Estonia, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Iraq, Ireland, Isle of Man (UK), Israel, Italy, Japan, Jordan, Kazahkstan, Kuwait, Latin America, Latvia, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Paraguay, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Yemen. | Terminated | May 17, 2012 | 1295 | 9633 | Not_Computer
Z-1873-2009 | 52721 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Welch Allyn Vital Sign Monitor 5300 Series/Class II. (VSM 300). |  |  | July 28, 2009 | 2009 | Welch Allyn Protocol, Inc | The Welch Allyn VSM300 Vital Signs Monitor may experience a loss of audio which includes a loss of audio alarms. | Device Design | An "Urgent: Correction" letter dated July 28, 2009 was sent to customers describing the risk associated with loss of audible alarms and recommend that the users mitigate the risk by checking for proper audible alarm function each time they power on the device. (Current labeling describes the variety of beeps and tones that should be present). Users were instructed: 1) if the audio fails, do not use the device, 2) Remove it from service and 3) contact Welch Allyn Protocol, Inc. at the following telephone number (1-800-289-2501) to arrange the return and repair of your device to Technical Services. Direct questions about this recall to Welch Allyn Protocol, Inc. by calling 1-503-530-7935. |  | 68123 | 68,123 | Worldwide Distribution -- US, United Arab Emirates, Afghanistan, Austria, Australia, Belgium, Bulgaria, Bahrain, Botswana, Canada, Switzerland, China, Cyprus, Germany, Denmark, Egypt, Spain, Figi, France, Great Britain, Greece, Hong Kong, Croatia, Indonesia, Ireland, Israel, India, Iceland, Italy, Jordan, Kenya, Kyrgyzstan, South Korea, Kuwait, Lebanon, Sri Lanka, Malaysia, Nigeria, The Netherlands, Norway, New Zealand, Oman, Phillipines, Pakistan, Poland, Puerto Rico, Panama, Portugal, Qatar, Russia, Saudi Arabia, Sweden, Singapore, Thailand, Taiwan and South Africa. | Terminated | August 10, 2011 | 743 | 68123 | Not_Computer
Z-1675-2009 | 52435 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Somatom Definition Model AS/AS+, Model number 8098027. |  |  | May 29, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Support mechanism may not engage properly due to improper stowage after opening which may result in internal component damage or potential injury. | Equipment maintenance | The firm has issued a modification of the telescope support arm parking bracket via Update Instruction CT028/09/S. A firm service representative will install the modified support arm parking brackets on the affected systems. |  | 93 | 93 units | Nationwide Distribution -- including states of AR, AZ, CA, CO, CT, FL, IA, IL, IN, KS, KY, MA, MD, MN, MO, MT, NC, ND, NE, NH, NY, OH, PA, SC, SD, TN, VA, WA, WI, and WY. | Terminated | September 01, 2009 | 95 | 93 | Not_Computer
Z-1555-2009 | 52007 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Stratus CS Acute Care Troponin I CalPak. |  |  | April 28, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc. | The firm received and confirmed that some lots of Cardiac Troponin I (cTnl) CalPaks (CCTNI-CR) may exhibit calibration issues that cause lower than typical signal during calibration with the current population of cTnl TestPaks (CCTNI). This issue may result in failed calibrations due to "Slope Error" messages or falsely elevated patient sample and quality control (QC) results on calibrations that | Labeling False and Misleading | Siemens Healthcare Diagnostics, Inc. issued an Urgent Field Safety Notice dated April 29, 2008 to all affected customers instructing them to discontinue use of 3 CalPak lots and reduce the shelf-life for the remaining lots.For further information, contact Siemens Healthcare Diagnostics, Inc. at 1-302-631-7672. |  | 4415 | 4415 cartons | Worldwide Distribution - including United States, United Arab Emirates, Austria, Bermuda, Belize, Canada, Germany, Egypt, India, Japan, Saudi Arabia and Bengal. | Terminated | July 14, 2009 | 442 | 4415 | Not_Computer
Z-1626-2009 | 52189 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific EndoVive 24F (8 mm) Safety PEG Kit, Push Method for use with sterile Xylocaine ampule, latex free, sterile, Boston Scientific Corp., Natick, MA; Cat. no./REF 6649, inner pouch UPN M00566490 and outer carton UPN M00566491. |  |  | May 05, 2009 | 2009 | Boston Scientific Corp | The user may be unable to pass the device over the guidewire for placement in the patient, resulting in an inability to use the device. | Process control | Boston Scientific Corporation issued an "Urgent Medical Device Recall - Immediate Action Required" notice dated May 5, 2009 as follow-up to their "Recall of the Safety and Standard Push PEG Kits" dated December 30, 2008. The notice outlined the issue and Consignees were instructed to locate and return all affected product to the firm. In addition, the firm requests that users complete and return a Reply Verification Tracking Form via fax to 1-508-683-5578.For further questions, contact the Boston Scientific Corporation Field Action Team by phone at 1-508-683-4678. |  | 140 | 140 | Worldwide Distribution -- United States, India, Japan and Slovenia. | Terminated | October 13, 2010 | 526 | 10259 | Not_Computer
Z-1790-2009 | 52119 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Alaris PC Unit, Model Numbers 8000 and 8015 (formerly Medley PC Unit) with software version 4 and above. |  |  | June 12, 2009 | 2009 | Cardinal Health 303 dba Cardinal Health | There is a potential risk for medication not being delivered when IV tubing sets are occluded in the Alaris Pump module. This can occur if IV sets are set up in the infusion pumps but not used for an extended period of time. If users fail to clear the occlusion, then the occluded tubing may cause medication not to be delivered. A delay of medication may result in serious injury and/or death, an | Device Design | The recall was initiated on June 12, 2009 with the firm forwarding via UPS an Urgent: Medical Device Recall Notification with attached Customer Response Card to the Chief Administrative Officer, Director of Pharmacy, Director of Biomedical Engineering, Director of Nursing &amp; Director of Risk Management of all of the affected consignees. Customers were also asked to please promptly complete and return the enclosed Customer Response Card to expedite the correction process.The Recall Notification informed the consignees of:(A) PROBLEM #1: Occlusion Warning Message for Alaris Pump Module; (B) PROBLEM #2 Syringe volume Warning Message for Alaris PCA module; (C) PROBLEM #3 Electrostatic Discharge (ESD) protection circuitry; (D) PROBLEM #4 Fluid Ingress and (E) PROBLEM #5 Alaris System Inter-Unit Interface (IUI) Connectors Labeling. The recall notice included "Required Action for Users &amp; Follow-up actions by Cardinal Health for all five problems. Customers were also asked to report any adverse reactions experienced with the use of this product, and/or quality problems to the FDA's MedWatch Program. In addition, customer were provided the following contact information for the recall: (1) Cardinal Health Recall Center for general questions: By phone at 1-888-562-6018 7am to 5pm (Pacific) ; (2) Customer Advocacy for adverse reports: By phone at 1-800-854-7120, Option 1, Option 1, Option 3 or by email at customerfeedback@cardinalhealth.com 24 hours a day, Sunday-Saturday; (3) Technical Support for Technical Questions regarding the Alaris System: By phone at 1-888-562-6018, 6am-5pm (Pacific).The firm followed-up with the release of a nationwide press statement regarding this issue on July 29, 2009. The press release indicated serial numbers of affected devices, as well as CareFusions short term instructions to customers, and the firms strategy to fix the affected devices can be found at: www.cardinalhealth.com/alaris/medical-device-recall/ (This informatio |  | 186000 | 186,000 units | Worldwide Distribution -- US (including Puerto Rico and Guam), Germany, England, Canada, Netherlands, Malaysia, Singapore, Australia, and New Zealand. | Terminated | March 19, 2012 | 1011 | 1016000 | Not_Computer
Z-1875-2009 | 52032 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | FlexicairÂż MC3 (model M4000) Low Airloss Therapy Units, for the prevention of pressure sores. |  |  | May 12, 2009 | 2009 | Hill-Rom Manufacturing, Inc. | The devices may emit smoke from the blower box assembly. | Equipment maintenance | Urgent Field Safety Notice letter dated May 11, 2009. Hill-Rom service centers were notified by letter on 05/13/2009. Action to be taken are as follows: Rental Customers; In the next 60 days, Hill Rom will be removing the FlexicairÂż II and FlexicairÂż MC3 low Airloss Therapy Units from service to correct the devices. During that time, Hill-Rom may provide an equivalent substitute therapy to accommodate the patients who continue to need a pressure reducing therapy device. If t his malfunction occurs before a substitution or a corrected device is provided, unplug the device, remove the patient from the surface and notify Hill-Rom immediately. Facility-Owned Customers: In the next 60 days, for all facility owned Flexicair II and FlexicairÂż MC3 Low Airloss Therapy Units, a Hill-Rom service technician will either replace the affected components in the blower assembly or deliver an equivalent substitute unit from Hill-Rom's rental fleet. If this malfunction occurs before the device is corrected or a substitute provided, do as follows: unplug the device, remove the patient from the surface, and immediately notify Hill-Rom Technical Support using the attached listing their geographical region. |  | 889 | 889 units | Nationwide and Canada | Terminated | April 03, 2012 | 1057 | 946 | Not_Computer
Z-2046-2009 | 46722 | No | DTS | Cardiopulmonary bypass cardiotomy return sucker. | Cardiovascular | Cardiovascular | 510(K) Exempt | Terumo Sarns cardiovascular disposable products, Weighted Flexible Sucker, sterile, Terumo Cardiovascular Systems Corp., Ann Arbor, MI.; Catalog No. 804113. |  |  | February 05, 2008 | 2008 | Terumo Cardiovascular Systems Corporation | A loose strand of stainless steel left on the device during manufacturing may result in an unretrieved metal fragment remaining in the patient. | Process control | Consignees were notified via Urgent - Medical Device Removal letters dated 2/5/08 to cease use of the product. Terumo CVS will contact customers to arrange replacement for the affected product upon receipt of the attached response form. Questions are to be directed to Terumo CVS Customer Service at 1-800-521-2818. |  | 7079 | 7079 | Worldwide Distribution --- including USA and countries of Belgium, Brazil, Dominican Republic, Germany, Jordan and Latin America. | Terminated | November 01, 2009 | 635 | 7079 | Not_Computer
Z-1676-2009 | 52091 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia S Series SPECT System; Single-Photon Emission Computed Tomography; Siemens Medical Solutions USA, Inc., Hoffman Estates, IL 60195-2061 |  |  | April 22, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | The pinion gear could crack and cause the gantry to spin freely thereby damaging the Symbia S or T System and potentially injuring a patient during a scan. | Finished device change control | Siemens sent Urgent Field Correction Recall letters dated 4/22/09 to the affected users of the Symbia S and Symbia T Systems. The accounts were informed of the potential faulty pinion gear which could crack and cause the gantry to spin freely, damaging the system and potentially injuring a patient during a scan. The customers were advised to stop using the system if they hear any abnormal grinding noise from the gantry. They also were advised that a Siemens Customer Service Representative would contact them within 60 days to schedule a replacement of the pinion gear free of charge. Any questions were directed to contact Siemens Medical Solutions USA, Inc. at 1-800-767-2313 (USA) for assistance. The accounts were requested to complete the Mandatory Fax-Back Form acknowledging receipt and understanding of the letter. |  | 37 | 37 units | Worldwide Distribution -- USA, Australia, Austria, Brazil, Bulgaria, Canada, China, Finland, France, Germany, India, Italy, Japan, Mexico, Netherlands, Norway, Philippines, Russia, Spain, Taiwan, and Turkey. | Terminated | June 21, 2010 | 425 | 75 | Not_Computer
Z-1702-2009 | 52226 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer ITST Intertrochanteric/subtrochanteric fixation taper reamer - Asia, nonsterile; REF 2258-50-01. (Note: This product is not distributed in the United States). |  |  | June 05, 2009 | 2009 | Zimmer Inc. | Reamers labeled as having a standard diameter may actually be the dimensions of an Asia reamer, and vice versa. | Process control | Zimmer sent an Urgent Device Recall letter dated 6/5/09 to all consignees and instructed to locate and return the products. |  | 120 | 120 of both products | Worldwide distribution: USA, China, Germany, Japan, Korea and Thailand. | Terminated | December 22, 2009 | 200 | 120 | Not_Computer
Z-1704-2009 | 52177 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | IMHS CP, 3.2mm, Tip Threaded Guide Pin, REF71687000, for use with IntraMedullary Hip Screw in orthopedic surgery. |  |  | January 22, 2009 | 2009 | Smith &amp; Nephew Inc | Incorrect depth gauge mark etched onto guide wire pin. Guide pin is too long. | Process design | Urgent Medical Device Recall letter dated January 23, 2009. All affected Smith &amp; Nephew Sales Representatives, Distributors and Hospitals were notified of problem and the recall via Fed Ex and e-mail on 01/22/2009. They were instructed to immediately check their inventory and return the product. |  | 99 | 99 units | Distribution to USA, Switzerland and Australia. | Terminated | August 18, 2011 | 938 | 99 | Not_Computer
Z-1695-2009 | 52228 | No | LPQ |  |  | Orthopedic | PMA | Exogen 4000+ Ultrasound Bone Healing System, SAFHS, Part Numbers 71034100 and 71034101, Rx only, Smith &amp; Nephew, 1450 Brooks Road, Memphis TX 38116. |  |  | June 05, 2009 | 2009 | Smith &amp; Nephew Inc | Problems with transducer may result in a reduced ultrasound output. | Other | All affected Smith &amp; Nephew consignees and distributors were notified of problem and the recall by letters (dated June 5, 2009 and June 8, 2009) via US mail. Patients were directed to return the instruments; physicians were informed that their patients were being contacted; and distributors were requested to return their stock and to notify anyone to whom the product was distributed to also return the product. Direct questions to Smith &amp; Nephew, Inc. by calling at 1-800-836-4080. |  | 4139 | 4139 units | Worldwide Distribution -- US, Switzerland, Canada, Malaysia, Japan, and UK. | Terminated | November 18, 2010 | 531 | 5030 | Not_Computer
Z-1703-2009 | 52289 | No | MRA |  |  | Orthopedic | PMA | Trident Acetabular System Surgical Protocol. Instructions for use of hip prosthesis. Literature Number LSP55. |  |  | June 09, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | The Surgical Protocol has been modified to create separate and distinct surgical protocols, one for the TridentÂż PSL Shell (LSP68) and one for the TridentÂż Hemispherical Shell (LSP69) in order to clarify and highlight the differences in reaming technique. | Labeling design | Stryker sent Urgent - Product Correction - Surgical Protocol letters by Federal Express on June 9, 2009 to all Stryker branches/agencies, hospital Risk Management, hospital Chief of Orthopaedics, surgeons and all foreign consignees. |  | 3702 | 2376 in USA, 1326 internationally | Worldwide distribution. | Terminated | May 08, 2012 | 1064 | 3702 | Not_Computer
Z-1697-2009 | 49148 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | Foot Choice Therapeutic Foot Massager, Model Number YS-322. |  |  | August 07, 2006 | 2006 | Suarez Corporation Industries | Complaints of product causing burns to consumers. | Device Design | On August 9- August 14, 2006, the firm sent a letter to all customers which included an "Important Customer Reminder." The Reminder now listed additional warnings about serious burns occurring regardless of the control settings. Direct questions to Suarez Corporation Industries by calling 1-330-494-5504. |  | 71512 | 71,512 units | Nationwide and Canada. | Terminated | October 13, 2010 | 1528 | 71512 | Not_Computer
Z-1671-2009 | 52376 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 12 French 3 Lumen 16 cm Large-Bore Central Venous Catheters. Product numbers AK-12123-H, CS-12123-E, MA-12123, and MC-12123. |  |  | May 08, 2009 | 2009 | Arrow International Inc | Wire guide can have difficulty passing through catheter juncture hub or being removed from the catheter. | Nonconforming Material/Component | The recalling firm issued an Urgent Medical Device Recall letter dated 5/11/09 to their customers informing them of the problem and the potential consequences. The letter asked customers to check their stock, cease use and distribution, and quarantine the affected products. They are to contact Arrow's Customer Service Department at 1-800-523-8446 for a Return Authorization Number; complete the Recall Acknowledgement &amp; Stock Status Form and fax to 1-800-343-2935; and return any affected product. The letter states that Arrow can provide either replacement with a similar product at no charge or credit customers' accounts. Additional information or clarification can be obtained by calling Customer Service at 1-800-523-8446. |  | 46610 | 46,610 total | Worldwide Distribution -- USA, Belgium, Canada, Cyprus, Czech Republic, Estonia, Finland, France, Germany, Hungary, Indonesia, Ireland, Israel, Italy, Lithuania, Japan, Republic of Korea, Netherlands, New Zealand, Poland, Singapore, Slovakia, South Africa, Spain, Taiwan, Ukraine, and United Kingdom. | Terminated | July 29, 2010 | 447 | 46610 | Not_Computer
Z-1593-2009 | 51963 | No | LFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche ACCU-Chek Aviva Meter Kit, Roche Diagnostics, Indianapolis, IN; REF 04893247001 includes the Roche ACCU-CHEK Softclix Lancing Device, Distributed by Roche Diagnostics, Indianapolis, IN. Model 3144844001 (10 lancets). |  |  | May 13, 2009 | 2009 | Roche Diagnostics Corp. | The lancet cap may be missing, which could result in an unintended lancet stick to the user. | Other | An "Urgent Medical Device Correction" letter dated May 13, 2009 was sent to all affected customers, distributors and consumer consignees who have either registered their blood glucose meter, contacted the firm about the lancets since 10/1/08, or have been shipped lancets by Roche since 10/1/08 were notified of the problem. Distributors were requested to notify each patient self-tester to whom a meter kit was distributed since October 1, 2008. The firm issued a press release on May 18, 2009. Diabetes Care customers: If you have questions regarding the recall, please contact the ACCU-CHEKÂż Softclix Lancet Hotline at 1-800-778-7057.CoaguChek meter customers: If you have questions regarding the recall, please contact the ACCU-CHEKÂż Softclix Lancet Hotline at 1-800-778-7505. |  | 449533 | 449,533 packages of lancets | Nationwide. | Terminated | January 05, 2011 | 602 | 449533 | Not_Computer
Z-1813-2009 | 52439 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics Esprit Ventilator System, Model Number V1000 with Optional External Battery; Product Number: 1001456. |  |  | April 22, 2009 | 2009 | Respironics California Inc | The recall was initiated because through routine product monitoring Respironics discovered that under certain conditions, when operating on AC line power, a portion of the current from the External Battery charger is erroneously diverted to the ventilator. When this occurs, the External Battery charger may not transition correctly. Prolonged operation in this condition results in overcharging and | Device Design | Respironics California, Inc. issued an "Urgent - Medical Device Field Correction" notice dated April 22, 2009. Customers were informed of the affected device including corrective actions taken by the firm. A Respironics representative will contact customers to schedule service of the device.Further questions can be addressed by contacting Respironics US Customer Service at 1-800-345-6443. |  | 508 | 508 | Worldwide Distribution: United States, Hong Kong, India, Mexico, South Africa and Vietnam. | Terminated | July 13, 2012 | 1178 | 508 | Battery | Device Operation | Repair | Repair
Z-1815-2009 | 52523 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Phoenix Infusion System with MedNet Software, Symbiq One-Channel Infusion System, Software versions v2.1 &amp; v3.0, Manufactured by Hospira Inc., Morgan Hill, CA. |  |  | March 31, 2009 | 2009 | Hospira Inc | Potential delay/underinfusion of critical therapy-- Devices experience increased frequency of alarms that result in temporary pump stoppage. | Software design | Hospira, Inc. issued an "Urgent: Device Field Correction" dated March 31, 2009 via Federal Express to Healthcare Professionals and Valued Hospira Customers. Consignees were informed of the affected devices and will be contacted by a Hospira representative to arrange necessary device upgrades.For further questions, contact Hospira Global Product Safety and Complaints at 1-800-441-4100. |  |  |  | Product was distributed to 59 consignees to all 50 States. | Terminated | December 18, 2010 | 627 |  | Other | Treatment/Delivery/Therapy | Upgrade | Repair
Z-1731-2009 | 51998 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Cardiac Troponin I Flex reagent cartridge (CTNI). |  |  | April 08, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc. | possibility of falsely elevated results | Process control | Siemens Healthcare Diagnostics issued a letter to their customers starting April 8, 2008 informing them of the problem and to discontinue use of the product. Further questions should be addressed to Siemens Healthcare Diagnostics, Inc. at 1-800-441-9250. |  | 4574 | 4,574 cartons | Worldwide Distribution -- Including countries of the United States, Bahamas, Belgium, Czech Republic, France, Germany, Italy, Poland, Portugal, Saudi Arabia, Spain, South Korea, and Slovakia. | Terminated | August 13, 2009 | 492 | 4574 | Not_Computer
Z-1811-2009 | 52601 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 60ml Syringe Luer-Lok tip, Sterile, BD, Franklin Lakes, NJ 07417, Made in USA. |  |  | July 07, 2009 | 2009 | Becton Dickinson &amp; Company | Insufficient Seal: Becton Dickinson received complaints of insufficient/open seals on one lot of 60 ml Luer-Lok Syringes. | Packaging process control | Becton Dickinson issued an "Urgent: Product Recall" notice via UPS tracking dated July 7, 2009 to all distributors informing them of the affected product. The firm asked all distributors to identify end customers in an attempt to notify all customers. The firms asks that all affected product be returned.Further questions should be directed to BD Customer Service representatives at 1-888-237-2762. |  | 133600 | 133,600 units | Nationwide Distribution -- Including states of IL, CA, VA, PA, and FL. | Terminated | October 02, 2009 | 87 | 133600 | Not_Computer
Z-1878-2009 | 52708 | No | KKO | Teething ring. | Dental | Dental | 510(K) Exempt | Gel Filled Teethers, various shapes and sizes, marketed under the names: Nuby, Cottontails, and Playskool. |  |  | July 17, 2009 | 2009 | Luv N' Care Ltd | Gel filled teethers were found to contain to contain Bacillus circulans and Bacillus subtilis. | PMA | The firm initiated their recall by telephone and email, and followed with letters to their consignees directing a recall to the consumer level. Consumers are urged to stop using the affected product and discard or return to the place of purchase. Consumers with questions about the recall may contact Luv N' Care, LTD by calling 1-800-256-2399. The firm issued a Press Release on July 17, 2009.Press Release about this recall is located on the website address: http://www.fda.gov/Safety/Recalls/ucm172689.htm. |  | 546320 | 546,320 units | Worldwide Distribution -- US (states of NY, IA, WI, IL, FL, NE, OH, MI, TX, NY, RI, and VA), Canada, Haiti, West Indies, Surinam, Antilles, Uruguay, Mexico, Bermuda, Canada, Dominican Republic, Dutch West Indies, Bahamas, North Antilles and Paraguay. | Terminated | February 28, 2013 | 1322 | 546320 | Not_Computer
Z-1685-2009 | 50097 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident Hemispherical Acetabular Shell; 72mm, Arc Deposited, Hydroxylapatite Coated, Multi holed; Sterile; Stryker, France. |  |  | May 29, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that Trident Hemispherical Shells, Catalog Number 508-11-74J, are actually marked and labeled as Catalog Number 598-11-72J, Lot Code 20282301 | Labeling mix-ups | Stryker Orthopaedics issued an "Important Market Withdrawal" notice via Federal Express dated May 29, 2007. Consignees were asked to locate and return affected product and fax the attached Product Accountability Form to 1-201-831-6069.For further questions, contact Stryker Orthopaedics at 1-201-831-5825. |  | 2 | 2 units | Distributed to 2 Stryker branches. | Terminated | August 17, 2009 | 811 | 2 | Not_Computer
Z-1734-2009 | 52079 | No | GKH | Automated blood cell diluting apparatus. | Hematology | Hematology | 510(K) Exempt | CellTracks AutoPrep System for in vitro diagnostic use, software version 2.2 or greater. Distributed by Veridex, 33 Technology Drive, Warren, NJ 07059. Manufactured by Immunicon, Huntingdon Valley, PA. Immunicon was acquired by Veridex in August of 2008. |  |  | May 06, 2009 | 2009 | Veridex, LLC | Bottle degradation: Investigation has determined that degradation may occur in the plastic waste bottle of the CellTracks AutoPrep System. | Device Design | Veridex, LLC issued a recall notice to consignees dated March 12, 2009 detailing the affected product. The firm will provide replacement bottles to all customers in the near future.For further questions, contact Veridex Customer Technical Services at 1-877-837-4339. |  | 196 | 196 instruments | Worldwide Distribution | Terminated | October 05, 2010 | 517 | 196 | Not_Computer
Z-1720-2009 | 52022 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, CT Perfusion 4 on Advantage Workstation, CT and PET/CT consoles. |  |  | May 04, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware of the need to adjust default settings associated with the use of CT Perfusion 4 on the Advantage Workstation, ST and PET/CT consoles that may impact patient safety.The default parameter settings in CT Perfusion 4 may not be optimal in the case of data acquired with longer time sampling intervals such as axial scans with inter scan delay (ISD) greater th | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated April 23, 2009. The letter was addressed to Hospital Administrators / Risk managers and Radiology department mangers. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information instructing the user to follow the enclosed User Guide Addendum to mitigate further problems.For further questions, contact GE Healthcare at 1-800-437-1171 (US), 0120-055-919 (Japan); for other countries, contact your local GE Healthcare field service representative. |  | 444 | 444 | Worldwide Distribution -- Including states of AZ, CA, CT, DC, FL, HI, IL, IA, MA, MI, NJ, NY, NC, OK, PA, PR, TX, WA, WV and WI and countries of ARGENTINA, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, PAKISTAN, POLAND, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM and VENEZUELA. | Terminated | May 16, 2012 | 1108 | 444 | Other |  |  | Other
Z-1735-2009 | 52602 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb, BostonÂż Conditioning Solution, Original Formula, For Rigid Gas Permeable Contact Lenses, 1 fl. oz. (30 ml), Bausch &amp; Lomb Incorporated, Rochester, NY 14609. |  |  | June 30, 2009 | 2009 | Bausch &amp; Lomb Inc | The product failed to meet one of the shelf life specifications at the end of the expiry period. | Other | Bausch &amp; Lomb issued an "Urgent Voluntary Medical Device Recall" notice dated June 26, 2009. All consignees were requested to review their inventory, and return any of the affected product to Stericycle Pharmaceutical Services, Indianapolis, IN. A Business Reply Card was included to be returned to Bausch &amp; Lomb via US mail. A prepaid UPS Packing Slip was also provided to return the affected product.For further questions, contact Bausch &amp; Lomb Customer Resource Center at 1-800-828-9030. |  | 56000 | 56,000 bottles | Nationwide Distribution. | Terminated | May 10, 2012 | 1045 | 56000 | Not_Computer
Z-1719-2009 | 51120 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare MUSE Cardiology Information Systems vs. 7.0, 7.0.1, 7.0.2, 7.1, 7.1.1. |  |  | February 02, 2009 | 2009 | GE Healthcare, LLC | GE Healthcare has become aware of an issue resulting from a workflow situation related to the automated labeling of patient data in the MUSE Cardiology Information System. When a prior record has been confirmed with an invalid (default) patient identifier (PID), MUSE will accept future records with the same default PID as correct. If subsequent records don not contain a name, MUSE will pre-popu | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated October 13, 2008, addressed to Healthcare Administrator/Risk Manager, Cardiology Administrators and Head of Cardiology. The letter listed the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, Background Information and Contact Information.For further questions, contact GE Healthcare at 1-800-558-7044. Outside the US, please contact your local GE Healthcare Service Representative. |  | 327 | 327 | Worldwide Distribution -- Including states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VT, VA, WA, WI, WY and DC and countries of UNITED KINGDOM , TAIWAN, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, QATAR, NEW ZEALAND, NETHERLANDS, KUWAIT, REPUBLIC OF KOREA, JORDAN, JAPAN, ITALY, IRELAND, GERMANY, FRANCE, FINLAND, ESTONIA, DENMARK, CANADA, BELGIUM, AUSTRIA and AUSTRALIA. | Terminated | January 16, 2014 | 1809 | 327 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-1670-2009 | 51138 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | DePuy Spine Expedium SFX Cross Connector System Torque Driver Shaft, Non-Sterile, REF: 2894-10-300. |  |  | February 11, 2009 | 2009 | DePuy Spine, Inc. | Tip of Driver may break during use | Nonconforming Material/Component | DePuy Spine notified direct accounts by visit and distributors by letter titled "Urgent - Voluntary Product Recall Notification". Accounts are requested to return the notification fax form and all affected product.For further questions, contact Depuy Spine at 1-800-227-6633. |  | 117 | 117 units | Nationwide Distribution -- Including states of MD, PA, NC, MO, NV, AZ, WI and CT. | Terminated | July 09, 2015 | 2339 | 117 | Not_Computer
Z-1730-2009 | 51722 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | GE Healthcare, Drapes and Sterile Covers |  |  | August 29, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has found that the sterile products manufactured by Contour Fabricators, Inc. (CFI) for GE healthcare (GEHC) may have sterile seals that could be compromised during the shelf life of the products. This issue could result in the product not maintaining its sterile state. | Release of Material/Component prior to receiving test results | Consignees were sent 8/29/08 a GE Healthcare "Urgent Recall Notice" dated August 15, 2008. The letter was addressed to Director / Manager of Radiology and Director of Surgery. The letter described the problem and provided a list of products, solutions and requested the return of the attached fax forms. In the event of any questions or concerns, a service team should be contacted at 800-874-7378 option #8. |  | 137689 | 137,689 (123,373 US; 14,316 OUS) | Worldwide Distribution -- USA, Canada, Australia, Bahamas, China, Hong Kong, Italy, Japan, Korea, Mexico, Malaysia, Singapore, and El Salvador. | Terminated | May 15, 2012 | 1355 | 137689 | Not_Computer
Z-1733-2009 | 52111 | No | JSY | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Remel Middlebrook 7H10 Agar, Ref R01600, packaged 10 units/package. The firm name on the label is Remel, Lenexa, KS. |  |  | May 06, 2009 | 2009 | Remel, Inc | The plate labeling is incorrect, but the outer box labeling is correct | Error in labeling | The recalling firm issued Important Medical Device Customer Recall Notice letters dated 5/6/08 via regular mail explaining the reason for recall and requesting all remaining inventory of the recalled lot be discarded. A Product Inventory Checklist was enclosed for completion to report the amount of inventory destroyed. A self-addressed postage-paid envelope was enclosed for return of the Product Inventory Checklist. Customers are to contact the firm's Technical Services Department at 800-447-3641 if they have inquiries concerning the letter. |  | 118 | 118 packages | Nationwide Distribution -- including states of FL, NJ, IN, FL, OH, KY, TX, and WI. | Terminated | October 15, 2009 | 162 | 118 | Not_Computer
Z-1718-2009 | 51322 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | AdvantageSim MD versions 7.4 through 7.6, Model 5160092. |  |  | February 02, 2009 | 2009 | GE Medical Systems, LLC | It was reported by a customer site that when using Advantage Sim MD on Advantage Workstation, the series could be incorrectly labeled in image view when multiple series of an exam are loaded simultaneously in Advantage Sim MD and if their series dates are different. A mismatch of series label for structure sets may lead to under-treatment of a tumor due to too small coverage of volume treated. | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated March 11, 2009. The letter was addressed to AdvantageSim Customers, Hospital Administrators/ Risk Mangers and Oncology Department Mangers. The letter addressed the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information.For further questions, contact GE Healthcare at 1-800-437-1171 (US), 0120-055-919 (Japan) and for other countries, contact your local GE Healthcare sales or field service representative. |  | 230832 | 687 units (351 US / 336 OUS) | Worldwide Distribution -- Including states of AL, AZ, CA, CO, CT, FL, GA, DC, ID, IL, IN, IA, KS, LA, ME, MD, MA, MI, MN, MS,MO, MT,NE, NH, NJ, NM, NY, NC, ND, OH, OK PR, PA, SC, SD, TN, TX, UT, VA, WA and WI and AUSTRIA, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, MALAYSIA, NETHERLANDS, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TOKELAU, TURKEY, UNITED ARAB EMIRATES and the UNITED KINGDOM. | Terminated | December 10, 2011 | 1041 | 230832 | Not_Computer
Z-1714-2009 | 52457 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | SJM Peel Away Introducer, REF 405153, Lot 2694252, Length 14 cm, 7F (2.33 mm), Max Guidewire o.d .038" (.965 mm), contents 5, 2011-08, Manufacturer St Jude Medical 14901 DeVeau Place, Minnetonka, MN 55345-2126 USA, Sterile EO, Rx only. |  |  | June 24, 2009 | 2009 | St. Jude Medical | St. Jude Medical has initiated a voluntary recall of two lots of Peel Away Introducers . Isolated lots could contain some items that have been mis-labelled with the incorrect French size of the outer box label. Some 9F products have been labelled as 7F and vice versa. Only the outer box label is affected by this error, the labels on the individual product pouches correctly identify the product. | Labeling mix-ups | The Consignee was provided with a St. Jude Medical "Urgent Product Information" letter and a doctor Acknowledgement form, via e-mail on June 2009. The letter described the problem and the product involved and advised to isolate and discontinue use of the product involved.The Consignee will notify the end user and remove any remaining affected units and will notify St Jude Medical of the results. |  | 200 | 200 | OUS only: Belgium. | Terminated | March 29, 2012 | 1009 | 200 | Not_Computer
Z-1812-2009 | 52160 | No | LIO | Microbiological specimen collection and transport device. | Microbiology | Microbiology | 510(k) | Remel BactiSwab NPG Collection and Transport System, Sterile, 100 swab units/box, Catalog 12300. The firm name on the label is Remel, Lenexa, KS. |  |  | May 07, 2009 | 2009 | Remel, Inc | Incorrect expiration date on the product | Incorrect or no expiration date | The recalling firm issued Important Medical Device Product Recall Notice letters dated 5/7/09 via regular mail to their direct accounts explaining the reason for recall and requesting they discard all remaining units of the affected lot number in their inventory. A Product Inventory Checklist was enclosed to be returned in the enclosed self-addressed, stamped envelope to report the amount of product that was discarded. The Technical Services Department should be contacted at 800-447-3641 if there are any inquiries concerning the letter. |  | 103 | 103 packages | Nationwide Distribution | Terminated | January 15, 2010 | 253 | 103 | Not_Computer
Z-1821-2009 | 52525 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Dall Miles Cable Grip System; Catalog number 6704-0-210; Stryker Ireland. |  |  | June 18, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Sterilization issue: During post-sterilization inspection of the outer blister packaging of the affected lots, it was determined that the outer blister did not meet the required minimum seal width. | Process control | Urgent Product Recall letters were sent via Federal Express on June 18, 2009 to all Stryker Branches/Agencies, hospital Risk Management, hospital Chief of Orthopaedics and surgeons who may have used the recalled products. The letters stated the issue, potential hazards, and the risk mitigation. Customers may contact Colleen O'Meara at 201-831-5970 with any questions. |  | 13 | 13 total, (both products) | Nationwide Distribution | Terminated | August 27, 2012 | 1166 | 13 | Not_Computer
Z-1818-2009 | 52379 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | T2 Ankle Arthrodesis Nail Surgical Protocol, Catalog Number LT2AA-OT, Stryker Orthopaedics |  |  | November 17, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Operative Surgical Technique was found to have a typographical error where hyphens were inadvertently omitted in incision lengths (i.e. 1520 cm should be 15-20 cm). | Labeling Change Control | Important Product Correction Letters, dated November 17, 2006, were sent to all Stryker Branches/agencies via Federal Express. The letter stated the issue and asked consignees to examine their inventory and contact their sales reps. It also told them to destroy the affected Operative Techniques by cutting them in half and discarding them, complete the attached Acknowledgment Form, and fax a signed copy of that form to 201-831-6069. The letter stated that Stryker Orthopaedics Customer Service reps would place consignees' re-orders for the product being destroyed. Also, the letter stated that in the interim a correct revision of the Operative Technique could be found online. Customers are to contact Dave O'Dell at 201-831-5277 or Rita Intorrella at 201-831-5825 with any questions. |  | 1270 | 1270 | Nationwide Distribution | Terminated | August 20, 2009 | 1007 | 1270 | Not_Computer
Z-1933-2009 | 52809 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific TELIGEN, Implantable Cardiac Defibrillator, Models E102 and E110. Sterilized using ethylene oxide. Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 66112-5798, USA. |  |  | July 20, 2009 | 2009 | Boston Scientific CRM Corp | Boston Scientific Cardiac Rhythm Management retrieved devices that could be subject to a potential for acute non-secure lead connections when implanted. Nonsecure lead connections can manifest as high impedance, or lead electrogram artifacts resulting in oversensing and/or inappropriate delivery of therapy. | Device Design | Boston Scientific Corporation field personnel began retrieving devices containing the original set screw system on Friday, July 17, 2009. If needed a letter, dated 07/17/2009 was left with the hospital administrator letting them know that product has been collected for return to Boston Scientific Corporation. Replacement devices will be provided at no charge. If the field personnel was unable to contact the hospital in the time provided, a notification letter dated July 22, 2009 was sent via Federal Express to the hospital. This letter was directed to the EP Catheter Lab Manager with a cc to the Hospital Risk Manager, Hospital Administrator and Purchasing and Central Supply. The letter asked that the affected devices not be implanted and to contact their local representative for pick-up and return. The firm will replace the affected devices. Please direct questions to your local Sales representative or Boston Scientific Corporation CRM Technical Services by calling 1-800-227-3422. |  | 104 | 104 | U.S. Distribution only (states of AZ, CA, CT, FL, GA, IA, IL, KY, MD, MO, NH, NJ, NY, OH, PA, SC, TN, TX, WI, and Washington D.C.) and Puerto Rico. | Terminated | June 26, 2012 | 1072 | 149 | Hardware | Output/Calculation | Return | Remove or Replace
Z-1954-2009 | 52853 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic Virtuoso VR, Model D154VWC, Single Chamber implantable cardioverter defibrillator with atrial therapies. Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. |  |  | July 27, 2009 | 2009 | Medtronic Inc. Cardiac Rhythm Managment | Medtronic has detected a specific pattern of MOSFET IC malfunctions in its Concerto, Virtuoso and EnRhythm family of devices. The probability of occurrence decreases with time and, to date over 90% of the malfunctions related to the pattern have occurred within the first twelve months after implant. | Process control | Medtronic, Inc. personnel retrieved 2 non-implanted devices remaining at 2 hospitals. A Performance Notes letter was left at each facility. The Performance Notes described the issue, actions to address issue and recommendations to physicians. The Performance Notes will also be posted on the firm's website. Direct questions about this recall to Medtronic, Inc. by calling 1-763-514-4000. |  | 1 | 1 | Nationwide Distribution-- (states of IA and NJ). | Terminated | May 31, 2012 | 1039 | 1 | Not_Computer
Z-1694-2009 | 51956 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Capless Li Pedicle Screw System, Diameters - 4.75mm, 5.5mm, 6.5mm, 7.5mm, 8.25mm, Lengths 30mm-55mm, labeled non-sterile. |  |  | March 24, 2009 | 2009 | X Spine Systems Inc | A defect cause was discovered following a customer complaint pertaining to a post operative dissociation of the Capless Li Screw construct. The screw assembly is made up of three parts; a cup which locks onto a 5.5mm rod, a yoke that cradles the rod during fusion, and a screw that attached to the vertebra. The screw ball on the screw was manufactured out of tolerance (undersized). | Process control | X Spine Systems Inc. issued advisory notices in March 2009 to Consignees notifying them of the problem and requesting all affected product be returned to the firm. The firm also conducted telephone calls and emails for the return of the product.For further information, contact X Spine Systems Inc. at 1-937-847-8400 extension 115. |  | 9613 | 9,613 units | Worldwide Distribution -- Including states of AZ, CA, CO, FL, ID, IN, MO, MS, NV, OK, TN, TX, UT and WY and countries of Spain, South Africa and the United Kingdom. | Terminated | April 25, 2012 | 1128 | 9613 | Not_Computer
Z-1736-2009 | 52452 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump catheter and split introducer needle, sterile, Stryker Instruments, Kalamazoo, MI; REF 500-23. |  |  | June 16, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | The catheter may break when removed from the body, or the catheter may have kinks or be occluded such that the patient will not receive the intended quantity of medication. | Nonconforming Material/Component | Stryker Instruments issued a recall letter dated June 16, 2009 and instructed Consignees to cease using and destroy the product and to return a business reply form for credit via fax to Stryker Instruments Regulatory Department, 1-866-521-2762.For further questions, contact Stryker Instruments at 1-800-800-4236 extension 3808. |  | 268 | 268 | Worldwide Distribution -- Including the United States, Argentina, Australia, Brazil, Canada, England, France, Germany, Hong Kong, Italy, Japan, Latin America, Malaysia, Mexico, Netherlands, New Zealand, Singapore, South Africa, Spain, Sweden, Switzerland and Taiwan. | Terminated | October 27, 2010 | 498 | 47091 | Not_Computer
Z-1958-2009 | 50783 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Stryker Reflex Hybrid Level Anterior Cervical Plate; |  |  | June 14, 2006 | 2006 | Stryker Spine | Design Flaw: The design did not prevent over angulation of the screws when the recommended instrument, All in One Guide, was not used. | Device Design | Stryker Spine issued an "Important Market Withdrawal" notice dated June 14, 2006. Consignees were advised to complete and return a Product Accountability Form via fax to 1-201-831-6069 prior to returning the affected product. Product return was to be performed during certain phase periods when upgrades were being shipped.For further questions, contact Stryker Spine at 1-201-831-5825. |  | 38987 | 38987 plates | Nationwide Distribution -- Including states of NC, OR, AZ, TN, CA, MA, MD, IL, GA, FL, NJ, NM, KY, LA, PA, MI, UT, MO, OK, KS, OH, MN, NE, IN, MT, TX, AL, CO and WI. | Terminated | August 20, 2009 | 1163 | 38987 | Not_Computer
Z-1721-2009 | 51317 | No | IXQ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Foot Lok Cusion, Part Number MTAFS01, reusable nonsterile (20.3 x 42 x 25.4cm) cushion. |  |  | March 06, 2009 | 2009 | Med Tec Inc | The silkscreened indexing scale was reversed. | Error in labeling | Med Tec inc. (CIVCO Medical Solutions, Radiation Oncology) issued a "Recall Notification" notice dated March 6, 2009 via certified mail explaining the reason for recall and providing the customer a choice of returning the product for replacement or not take any action. A response form was enclosed to be returned via FAX to indicate the number of units in their possession and whether or not they are being returned. Telephone calls were initiated April 20, 2009 to nonresponders or for any consignee whose letter was returned as undeliverable.Foreign customers were notified by e-mail or phone beginning March 16, 2009.For further information, contact CIVCO Medical Solutions, Radiation Oncology at 1-800-842-8688. |  | 231 | 231 units | Worldwide Distribution -- Including the United States, Czech Republic, Colombia, Thailand, Morocco, Brazil, the United Kingdom, India, Costa Rica, The Netherlands, Russia, Spain, Israel, Taiwan, Mexico, Turkey, Korea, Australia, Malaysia, Italy and Japan. | Terminated | November 05, 2009 | 244 | 231 | Not_Computer
Z-1732-2009 | 52018 | No | JBP | Activated whole blood clotting time tests. | Hematology | Hematology | 510(k) | Hemochron Jr. ACT- LR Assay (JACT- LR); |  |  | May 05, 2009 | 2009 | International Technidyne Corp. | One lot of ACT-LR was released when one finished good testing criteria acceptance limit did not meet acceptance criteria. | Employee error | An "Urgent - Medical Device Recall" letter dated May 4, 2009 was sent to healthcare professionals. The recall letter described the reason for recall and the possible health risks involved. Consignees were asked to check their inventory for the affected lot number and return the completed Customer Account Tracking form via mail, e-mail or fax to ITC. ITC will make arrangements for the return of affected product. If there is no affected product in the inventory and no product or safety concerns, the letter is for information purposed only. Note: the Customer Account Tracking form should be completed in all cases.If you have questions regarding the recall information, please call ITC Technical Support at 1-800-579-2255, or email ITC at JACTLR@itcmed.com. |  | 7785 | 173 boxes of 45 cuvettes | Nationwide. | Terminated | September 30, 2009 | 148 | 7785 | Not_Computer
Z-1686-2009 | 52088 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | DePuy P.F.C. Sigma Knee System, Non-Porous Cruciate Retaining Femoral Component, 73 mm M/L 69 mm A/P, 5 left, sterile, DePuy Orthopaedics, Inc., Warsaw, IN; REF 96-0005. |  |  | May 07, 2009 | 2009 | Depuy Orthopaedics, Inc. | There may be a crack on the lateral side of the condyle in the posterior chamfer region.The device is used as a Orthopedic knee implant. | Process control | The firm's sales staff was notified by email on May 7, 2009, and the recall was expanded to include additional lots on May 29, 2009, instructing them to remove the affected lots from their inventory and the inventory of any consignees, to provide the consignees with a copy of the recall letter, and to return the product to the firm. Hospital consignees were notified by letter dated May 7, 2009, as well as by updated letter on May 29, 2009. A dear doctor letter dated May 29, 2009 was issued to implanting surgeons advising them of the issue. Direct questions to Depuy Customer Service by calling 1-800-366-8143. |  | 260 | 260 | Nationwide. | Terminated | October 25, 2010 | 536 | 260 | Not_Computer
Z-1724-2009 | 52057 | No | CCN | Bourdon gauge flowmeter. | Anesthesiology | Anesthesiology | 510(K) Exempt | Amvex Integrated Flowmeter; a nonback-pressure compensated Bourdon gauge flowmeter; Amvex Corporation, 25B East Pearce St, Richmond Hill, Ontario L4B 2M9, Canada; Model No. FI. |  |  | May 29, 2009 | 2009 | Amvex Corporation | The affected flowmeters may have a damaged o-ring which may cause the flowmeter to deliver higher amounts of gas and or not be able to shut off fully. | Employee error | An "Urgent: Medical Device Recall" letter dated May 29, 2009 was issued to the affected direct accounts on the same date, informing them that the Integrated Flowmeters shipped between April 12, 2007 and August 31, 2007 may experience damage to the o-ring which helps to control the flow of gas from the flowmeter portion of the outlet. Units with a damaged o-ring might deliver higher amounts of gas and/or not be able to shut off fully. The letter provided a list of the affected serial numbers, and asked the accounts to confirm that they have flowmeters from the recall time period. The accounts were provided with interim instructions to use the flowmeters during the replacement period, and were instructed to call an Amvex Corporation Representative at 1-866-462-6839 with at least one of the affected serial numbers and provide the number of units that need replacement. The Amvex Corporation will send new flowmeters for replacement, with instructions on how to replace, record the serial numbers and return the recalled units. If the accounts distributed the units, they were requested to forward the recall letter to the end users. The accounts were also requested to complete the enclosed reply sheet acknowledging receipt of the letter and fax it back to the Amvex Corporation at 1-905-764-7743. Direct questions to the Amvex Corporation by calling 1-866-462-6839. |  | 1424 | 1,424 units | Worldwide Distribution -- US (including Puerto Rico), Canada, Iran, Peru, Taiwan and Thailand. | Terminated | August 19, 2011 | 812 | 1424 | Not_Computer
Z-1712-2009 | 52633 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Twin-PassÂż Dual Access Catheter , REF 5200, sterilized with Ethylene Oxide, RX Only, Non-Pyrogenic, Manufactured by Vascular Solutions, Inc., 6464 Sycamore Court, Minneapolis, MN 55369. |  |  | July 02, 2009 | 2009 | Vascular Solutions, Inc. | Three occurrences of Twin-PassÂż catheters containing a piece of PTFE beading material in the OTW lumen have been reported to Vascular Solutions. The PTFE beading material is a processing aide used during hydrophilic coating. The beading is to be removed followed by a visual inspection prior to packaging | Process control | Vascular Solutions, Inc. issued an "Urgent: Medical Device Recall - Lot Specific" notice dated July 6, 2009. The letter described the problem and product involved and asked Consignees to immediately check their Twin-PassÂż catheter inventory, remove any affected units and complete the Product Recall Inventory Form and return to Customer Service Department.For further questions, contact Vascular Solutions, Inc. at 1-763-656-4300. |  | 5885 | 5885 | Worldwide Distribution -- Including Puerto Rico and states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY and DC and countries of CANADA, SPAIN, GREECE, IRELAND, NORWAY, SWITZERLAND, ENGLAND, TURKEY, ISRAEL, GERMANY, AUSTRIA, TAIWAN, FINLAND, SWEDEN, FRANCE, IRELAND and PORTUGAL. | Terminated | April 03, 2012 | 1006 | 5885 | Not_Computer
Z-1027-2011 | 52903 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SEQUENCER component of MOSAIQ versions 1.30 (all builds), 1.40 (all builds), 1.50 (all builds), 1.60 (all builds), and 2.0 (all builds). Manufactured by IMPAC Medical Systems, Inc. Sunnyvale, CA. |  |  | July 17, 2009 | 2009 | Impac Medical Systems Inc | Potential for improper treatment. A software error may result in potential major underdosing to targeted areas from jaw blockage of the MLC aperture (jaw intrusion), and dosing to non-targeted areas from interstitial leaf leakage (jaw extrusion). | Software design | Impac Medical Systems Inc sent an Important Safety Notice letter to all affected consginees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact Impact Software support as soon as possible if they have the affected configurations identified in the letter. Customers were asked to complete the form included with the letter and and fax to:European customers: +44 1293 654401 American &amp; rest of world customers 702-992-5002 or e-mail to support@impac.com Customers were asked to distribute the Safety Notice to any and all users of IMPAC software at their organization who are potentially affected by this issue. For questions e-mail support@impac.com or callEuropean customers: 44 1293 654320American &amp; rest of world customers: 800-488-4672. |  |  |  | Worldwide Distribution - USA including AL, CA, CT, FL, GA, IA, IL, KY, MI, MN, MS, ND, NJ, OH, OR, PA,TN, TX, VA, WA, WI, and WV and the country of Canada. | Terminated | November 25, 2011 | 861 |  | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-1367-2011 | 53012 | Yes | NIQ |  |  | Cardiovascular | PMA | Endeavor Resolute Zotarolimus-Eluting Coronary Stent System with Rapid Exchange Delivery System, UPN 00613994400505, model #IDERES25024X. Manufacturer: Medtronic, Inc. Minneapolis, MN. Manufactured in: Medtronic Ireland, Galway, Ireland. Product is packaged inside a sterile, inner pouch, labeled in part "Use immediately". Inner pouch is supposed to be within a foil pouch containing desiccant, and moisture absorber. Foil pouch label reads in part "RX Endeavor Resolute Zotarolimus-Eluting Coronary Stent System...2.5 mm X 24 mm...REF IDERES25024X...Contents: One (1) Endeavor Resolute zotarolimus-eluting stent mounted on a rapid exchange stent delivery system...Guide Catheter/Minimum Inner Diameter &gt;=5F/0.056" (1.42 mm)...Maximum Guidewire Diameter 0.014" (0.36 mm)...Maximum Stent I.D. 3.00 mm (0.12"...Nominal Pressure 9 atm (912 kPa)...Rated Burst Pressure 15 atm (1520 kPa)...MR Conditional...Sterilized using ethylene oxide Consult instructions for use...Manufacturer: Medtronic, Inc. 710 Medtronic Pkwy NE Minneapolis, MN 55432 USA...Manufactured In: Medtronic Ireland Parkmore Business Park West, Galway, Ireland Tel:+353-91-708000 Fax:+353-91-757524 Investigational Devices / Returned Goods: Tel: (800) 556-4247 / Fax: (877) 518-2469...BioLinx Medtronic Polymer System". Foil pouch contents and Instructions for Use are placed within an outer box, which is labeled identically to the foil pouch. |  |  | August 06, 2009 | 2009 | Medtronic Vascular Galway Limited | The product was removed from investigational sites due to a packaging issue, in that the device units were packed without the foil pouch, oxygen absorber and Instructions For Use. | Packaging | The firm, Medtronic, sent their initial "Urgent Medical Device Recall Notice" to customers with affected product via e-mail on August 7, 2009. An additional "Urgent Medical Device Recall Notice" was sent to all customers on August 12-13, 2009 via FedEX and UPS. The notice described the product, problem and action to be taken by the customers. The customers were instructed to remove the affected units from their inventory and isolate the units in a secure location to prevent additional use. A Medtronic CardioVascular Clinical Customer Service (CCS) representative will contact their site to arrange for return of the units. This will be followed by a shipment of replacement units to their sites. If there were any questions regarding the return process, customers were to contact CCS at 800-556-4247.If you have any further questions or concerns regarding this notice, please contact the Sr. Clinical Research Manager at 707-543-2284 or ann.depaoli@medtronic.com. |  | 49 | 49 units | Nationwide distribution: AL, AR, CA, FL, IN, MD, MI, NY, OH, PA, RI, TX, VA, and VT. | Terminated | February 17, 2011 | 560 | 49 | Not_Computer
Z-1820-2009 | 52582 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LEGION Pressfit Stem, Straight, 20 MM X 160 MM, REF 71424053, Sterile, Smith &amp; Nephew, Memphis TX 38116. |  |  | January 22, 2009 | 2009 | Smith &amp; Nephew Inc | Products were mispackaged. | Process control | All affected Smith &amp; Nephew Sales Representatives were notified of the problem and the recall via email and overnight mail on January 22, 2009. All affected Smith &amp; Nephew International Distributors were notified via email and telephone on January 22, 2009. They were instructed to immediately quarantine the product for return. |  | 19 | 19 units | Worldwide Distribution -- US, Canada, Australia, Germany, GB, Italy, Switzerland, Spain, and France. | Terminated | February 07, 2012 | 1111 | 35 | Not_Computer
Z-1722-2009 | 51314 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare Signa Ovation 0.35T, Model Numbers: |  |  | November 03, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has identified a potential pinch point hazard on Ovation MR Scanner when using CTL Array XL Body Flex and Opened Body coil. A patient's hand may be pinched between the magnet bore ceiling cover and the CTL Array, XL Body Flex, or Open Body coil if the patient puts his/her hand on the top of the coil. 2/20/09: Recall being expanded to include the OpenSpeed Systems. | Device Design | GE Healthcare issued a "Product Safety Notification" letter dated October 23, 2008. The letter was addressed to Hospital Administrators, Mangers of Radiology/Cardiology &amp; Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information for the Signa Ovation 0.35T Systems.A second GE Healthcare "Urgent Medical Device Correction" letter dated March 5, 2009 was sent 3/9/09 to consignees which addressed the expanded recall to the Signa OpenSpeed0.7T MR Systems.For further questions, contact your local GE Healthcare sales or service representative. |  | 43736 | 308 (143 US / 142 OUS / 23 TBD) | Worldwide Distribution-- US (including states of AL, ALZ, AR, CA, CO,CT, DE, FL, FA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MT, MO, NE,NH, NV, NJ, NY, NC, OH, OK, OR, PA, PR, SC SD, TN, TX, VA, WA, WV and WI), CHINA, CYPRUS, ARGENTINA, BRAZIL EGYPT, DENMARK, FEDERAL REPUBLIC OF GERMANY, FEDERAL REPUBLIC OF FRANCE, GET 7 D CHINA, GREECE, HUNGARY, ITALY, IVORY COAST, JAPAN, KOREA, KUWAIT, MALAWI, MEXICO, OMAN, PANAMA, PHILIPPINES, RUSSIAN FEDERATION, SAUDI ARABIA, SPAIN , SWITZERLAND, TURKEY, UKRAINE, UNITED ARAB EMERITUS, UNITED KINGDOM and VENEZUELA. | Terminated | May 24, 2012 | 1298 | 43750 | Not_Computer
Z-1728-2009 | 51692 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | Hologic Delphi series Bone Densitometer |  |  | April 07, 2009 | 2009 | Hologic, Inc. | System may not turn off at the end of the x-ray and cause additional x-ray exposure | Component design/selection | Hologic issued an Urgent: Medical Device Recall customer notification letter, dated April 17, 2009, by certified mail to inform customers of the problem.With the notification letter, customers will receive the following:-Software version Apex 12.6.2, 2.3.2, or 2.4.2-Instructions to upgrade their systems with the provided software and to verify the new version on the system as a confirmation that the software was successfully loaded.-New Operator's Manual for installations where the operatingcharacteristics of the system have changed due to the software upgrade.For Customers with systems with software versions other than 12.6, 2.3, or 2.4:Customers were asked to upgrade their QDR software, return their old QDR software disk, and complete and return to Hologic the enclosed Confirmation of Notification. |  | 1914 | 1914 units total | Worldwide Distribution -- USA, Canada, Europe, Austria, Belgium, Bulgaria, Czech Republic, Denmark, France, Germany, Great Britain, Greece, Hungary, Iceland, Italy, Latvia, Lithuania, Macedonia, Netherlands, Norway, Poland, Portugal, Romania, Russia, Serbia, Slovak Republic, Slovenia, Spain, Sweden, Switzerland, Turkey, Ukraine, Uzbekistan, Algeria, Arab Emirates, Argentina, Australia, Bahrain, Chile, China, Colombia, Costa Rica, Dominican Republic, Ecuador, Egypt, Hong Kong, Indonesia, India, Iran, Iraq, Israel, Japan, Jordan, Kenya, Korea, Kuwait, Lebanon, Malaysia, Morocco, Mexico, Pakistan, Peru, Puerto Rico, South Africa, Saudia Arabia, Singapore, Syria, Taiwan, Thailand, Tunisia, Venezuela, and Vietnam. | Terminated | November 29, 2010 | 601 | 1914 | Computer
Z-1902-2009 | 52524 | No | INX | Air-fluidized bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | CLINITRONÂż CII Air Fluidized Therapy Units. |  |  | June 24, 2009 | 2009 | Hill-Rom Manufacturing, Inc. | Incorrectly repaired power cord wires can potentially overheat resulting in smoldering sound foam in the base of the therapy bed. | Equipment maintenance | Hill-Rom issued an "Urgent Field Safety Notice" letter dated February 2, 2009 addressed to Facility Risk Managers and Facility Administrators. The notice detailed the issue and instructed users on actions to be taken.For further information, please contact Hill-Rom Customer Service at 1-800-638-2546. |  | 625 | 625 units (total both units) | Worldwide Distribution -- United States, Abu Dhabi, Brazil, Canada, Chile, Columbia, Mexico, Newfoundland, Puerto Rico, Saudi Arabia, Singapore, Taiwan, Turkey and Venezuela. | Terminated | July 11, 2012 | 1113 | 625 | Battery | Physical Safety Hazards | Notification | Safety Notice/Insructions
Z-1906-2009 | 51336 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Specimen Gate Lab - MSMS Data Suite, Product Code: 5002-0310. |  |  | January 15, 2009 | 2009 | Perkin Elmer | If Result Codes are used and if they are manually edited, they are correct unless the laboratory specifically uses the View Assay button in theSpecimen History screen to open another assay, and then returns back to the original assay | Software design | Perkin Elmer began notifying US customers January 15, 2009. Notification included instructions for how to avoid the situation and instructions for how to obtain a correction. The firm requested consignees to send in a Response Form by fax (358-2-2678-357) or by email (TurkuQMresponse@perkinelmer.com).For further information, please contact your closest Specimen Gate support center at specimen.gate.support@perkinelmer.com for North America and specimengatesupportfi@perkinelmer.com for outside of North America. |  | 13 | 13 units | Worldwide Distribution -- United States (AL, AZ, CA, CO, CT, FL, GA, IN, KS, KY, MI, MO, OH, OK, OR, TN and TX) and Canada. | Terminated | September 09, 2011 | 967 | 39 | Not_Computer
Z-1913-2009 | 52629 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Edwards LifeSciences, Percutaneous Sheath Introducer Set, Rx only, For single use only. |  |  | May 11, 2009 | 2009 | Edwards Lifesciences, Llc | Edwards LifeSciences LLC received several complaints, in which the blue connector of the introducer assembly has broken (fracture/separation) after several hours of use. | Nonconforming Material/Component | Edwards LifeSciences LLC initiated the worldwide recall via customer letters dated May 11, 2009 and are being sent by Federal Express stating that customers cease using affected product from the indicated lots and return all unused products. The firm requested that the customer provide written verification of stock and identify any unused product to be returned. For further information, contact Edwards LifeSciences LLC Customer Service at 1-800-424-3278. |  | 67357 | 67,357 Units | Worldwide Distribution: United States (AL, AR, CA, CO, CT, DC, FL, GA, IA, IL, IN,KS, KY, MA, MD, ME, MI, MO, MS, ND, NH, NJ, NM, NV, NY, OH, PA, RI, SC, TN, TX, VA, WV, HI and PR), Austria, Canada, Switzerland, Czech,republic, Germany, Denmark, Algeria, Estonia, Egypt, Spain, Finland, France, Great Britain, Hungary, Iceland, Ireland, Italy, Kuwait, Lithuania, Luxembourg, Latvia, Netherlands, Norway, Poland, Portugal, Romania, Russia, Sweden, Slovak republic, Tunisia, Turkey, Aruba, Hong Kong, Malaysia and Singapore. | Terminated | February 23, 2012 | 1018 | 67357 | Not_Computer
Z-1955-2009 | 52718 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Stabilet Infant Warming System, model numbers 200, 300, 1250, 1500, 200/3000, 2000, 2200/3200, 3000, and 3200. |  |  | July 17, 2009 | 2009 | Draeger Medical Systems, Inc. | Hot particles fell into the infant bassinet beneath a Radiant Infant Warmer. The resulting fire injured an infant. These devices are obsolete. | Equipment maintenance | On 7/20/09, Draeger Medical, the recalling firm, issued a press release to medical media outreach organizations and an Urgent Medical Device letter to all known customers. This information explained the incident and the need to remove and replace the units in question as soon as feasibly possible. |  | 6489 | 6489 units | Worldwide distribution: USA, Brazil, Canada, China, Dominican Republic, France, Korea, Lebanon, Mexico, Nigeria, Panama, Poland, Saudi Arabia, Taiwan, Turkey, United Kingdom, Uruguay, and Venezuela. | Terminated | June 01, 2011 | 684 | 6489 | Not_Computer
Z-1934-2009 | 52802 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences "VAMP Jr" (Venous Arterial blood Management Protection system), Lot numbers (all product manufactured between 8/01/06 and 2/29/08.) |  |  | March 14, 2008 | 2008 | Edwards Lifesciences AG | Edwards Lifescience initiated a recall of on Edwards "VAMP Jr" Pressure Monitoring Products manufactured between 8/01/06 and 2/29/08 due to the potential for the cap on the reservoir of the "VAMP Jr" to become disconnected. | Equipment maintenance | Edwards initiated the recall in the U.S. via a Urgent Customer lNotification letter with attached response confirmation fax-back attachment dated March 10, 2008 and was sent by Federal Express on March 12, 2008 requesting that customers cease using all "VAMP Jr" products and that they return all unused product.Customers were instructed to fax-back the confirmation attachment to 800-422-9329 and contact Edwards Customer service at 888-570-4012 to arrange for return of any unused product and obtain information about replacement product.Edwards Lifesciences requested that the customer provide written verification of stock and identify and unused product to be returned. This will enable a 100% effectiveness check for this recall. For product that was shipped to distributors, contact was made directly from Edwards Lifesciences Corporate Office in Irvine, CA. All product that was distributed internationally will be handled regionally. (i.e. US, Canada, Europe/EMEA, Japan, Intercontinental-including Latin America and Asia Pacific) |  | 61199 | 61,199 units | Worldwide Distribution including USA and the countries of Australia, Austria, Belgium, Brazil, Canada, Canaric Island (Spain), China, Columbia, Czech Republic, Denmark, France, Germany, Hungary, Italy, Japan, Kuwait, Netherlands, New Zealand, Norway, Puerto Rico, Spain, Sweden, Turkey, and United Kingdom. | Terminated | July 03, 2012 | 1572 | 61199 | Not_Computer
Z-1899-2009 | 52712 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Vertier Surgical Table, an AC powered, general purpose, mobile surgical table with electro-hydraulically controlled movements which are part of a control feedback system. Part Numbers: 0788100000 Standard table, 0788100001 Standard table with 5th wheel, 0788100000S SIDNE-enabled standard table, and 0788100001S SIDNE-enabled 5th wheel table |  |  | November 13, 2008 | 2008 | Stryker Communications Corp | Surgical tables were produced using an unapproved tilt frame bushing. | Nonconforming Material/Component | Stryker Communications issued an "Urgent: Medical Device Recall" notice dated November 13, 2008 informing Consignees of the affected product. The firm intends to contact each facility to coordinate service by a trained representative.For further information, contact Stryker at 1-972-410-7100. |  | 31 | 31 | Worldwide Distribution -- United States (OR, IL, WA, CA, AR, TX, UT, SC and GA) and Canada. | Terminated | August 28, 2009 | 288 | 31 | Not_Computer
Z-1915-2009 | 52676 | device, prosthesis alignment | IQO | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | Stryker Orthopaedics PAL Pelvic Alignment Level, Ref: PAL 400; Howmedica Osteonics Corp, 325 Corporate Drive, Mahwah, NJ 07430. |  |  | September 11, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The bubble in a vial of the Pelvic Alignment Level (PAL) was wider than the level marking. Upon opening the PAL and removing the vial, it was discovered that the vial was leaking | Device Design | Stryker Orthopaedics issued an "Important Product Removal" letter dated September 11, 2007 via Federal Express informing consignees of the affected devices. The firm requested that users identify and return any affected product and account for product by returning a Product Accountability Form via fax.For further information, contact Stryker Orthopaedics at 1-201-831-5825. |  | 615 | 615 | Nationwide Distribution | Terminated | October 15, 2009 | 765 | 615 | Not_Computer
Z-1916-2009 | 52641 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Triathlon AR 4:1 Cutting Blocks - Size 1-8; |  |  | February 12, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Triathlon Cutting Blocks did not assemble to the Triathlon Impactor/Extractor Handle. | Process control | Important Product Correction letters were sent to all branches/agencies on February 12, 2008 by Federal Express. The letter asked consignees to examine their inventory and hospital locations to identify the affected product. Customers were asked to perform a functional inspection by assembling the Triathlon Impactor/Extractor Handle to all of their MIS Cutting Blocks. If the cutting block does not assemble, customers are to list the affected lot codes on the Attached Product Accountability Form and fax a copy to (201) 831-6069. The letter states that product should not be returned until further instructions are received. Questions should be directed to (201) 831-5825. |  | 690 | 690 total units - both varieties | Worldwide Distribution | Terminated | October 15, 2009 | 976 | 690 | Not_Computer
Z-1882-2009 | 52455 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | MR Coaxial 10g x 118mm (VC10118MR). |  |  | March 26, 2009 | 2009 | Bard Peripheral Vas | Some Vacora 10G Plastic Coaxial Cannulas may leave a small metal filament in the breast or the sample obtained post procedure. | Process control | Bard Peripheral Vascular, Inc. issued an "Urgent Medical Device Recall Notification" dated March 27, 2009 informing Customers of the affected devices. The firm requested that all affected product be removed and returned along with returning a Recall and Effectiveness Check Form via fax.For further information, contact Bard Peripheral Vascular, Inc. at 1-800-321-4254. |  | 1440 | 1,440 | Worldwide Distribution. | Terminated | April 02, 2012 | 1103 | 10425 | Not_Computer
Z-1890-2009 | 52782 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Sheridan Preformed Cuffed Tracheal Tube, Nasal 6.5, Catalog Number: 5-22313, Teleflex Medical Ltd, Research Triangle Park, NC 27709 |  |  | July 06, 2009 | 2009 | Teleflex Medical | The Nasal Preformed Tracheal Tubes were found to have incorrect insertion depth mark printings. | Process control | Consignees were notified by an Urgent Medical Device Recall letter on/about 07/06/2009. They were instructed to discontinue use of and return any of the recalled product to Teleflex Medical. Distributors were advised to forward the letter to their consignees to retrieve relevant product from those locations. A Recall Acknowledgement Form was attached to be returned to Teleflex Medical. Replacement product or credit will be provided to the consignee. |  | 1160 | 1160 total | Worldwide Distribution -- USA, including states of CA, FL, IL, KY, MA, NC, NM, NJ, OH, PA, RI, TN, TX, UT, VA, WA, and WI and countries of France and Japan. | Terminated | August 24, 2010 | 414 | 1160 | Not_Computer
Z-1693-2009 | 52835 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Heater Elements for Resuscitaire Infant Radiant Warmers model numbers RW82, RW82 VHA, WBR82, and WMRW82, 120 volt configuration only. |  |  | July 29, 2009 | 2009 | Draeger Medical Systems, Inc. | Spark, flame and smoke: A mis-positioned crimp on heater element lead wire resulted in flame and smoke aerating from the top of the warmer assembly. | Process control | Draeger Medical issued an Urgent Medical Device Recall letter via certified mail on 8/7/09. The letter explained the problem, the need to conduct an inspection of the infant warmer (instructions included), and request replacement elements. |  | 2915 | 2915 elements | Worldwide distribution: USA, Canada, Saudi Arabia, Bermuda, Virgin Islands, Mexico, Panama, Colombia, and Costa Rica. | Terminated | October 13, 2010 | 441 | 2915 | Not_Computer
Z-1809-2009 | 52654 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varis, Aria Radiation Oncology - Version 8.1.15, 8.5.11, and 8.6.07, Model number HIT-INFSYS, Manufactured by Varian Medical Systems Inc., Palo Alto, CA |  |  | May 08, 2009 | 2009 | Varian Medical Systems Oncology Systems | Potential mistreatment-- due to software issues, retired plans may be used and treated repeatedly. | Software design | An Urgent Medical Device Correction Letter/ Field Safety Letter, dated May 8, 2009, were distributed to all affected users, with a description of the problem and user corrective action steps. The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the issue. The firm intends to produce a software fix for the problem. If further clarification is needed, users should contact their Varian Customer Support District or Regional Manager. |  | 1587 | 1,587 units | Worldwide Distribution | Terminated | December 23, 2011 | 959 | 1587 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1928-2009 | 52118 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AMPLILINK Software, version 3.1.0 to 3.2.2 with COBAS AmpliPrep. Catalog number 03563383001, catalog number 04855094001, catalog number 04807197001, catalog number 04862392001, catalog number 05093236001, and catalog number 05201357001. |  |  | June 02, 2009 | 2009 | Roche Molecular Systems, Inc. | Software bug: A unique series of events involving the COBAS AmpliPrep instrument running AMPLILINK software version 3.2.2 that results in specimens being amplified and detected with the wrong parameters. The error can occur on either COBAS AmpliPrep/COBAS TaqMan 48, COBAS AmpliPrep/COBAS TaqMan (docked) or cobas s 201 system (docked) running Amplilink Software versions 3.1 or 3.2 series up to and | Software design | Urgent Medical Device Correction Letters were sent on June 2, 2009 by first class mail. The letter described the issue; impact on results; clinical implications; the workaround; and the actions required, which include following the workaround until a new version of AMPLILINK software is available and filing the letter for future reference. Questions or concerns should be directed to Roche Diagnostics Technical Support at 1-800-526-1247. |  | 282 | 282 | Nationwide Distribution | Terminated | August 13, 2010 | 437 | 282 | Software | Output/Calculation | Software update | Software Update
Z-1901-2009 | 52717 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific (1) Flexima Regular Biliary Catheter System with RO Marker, 8 FR (2.7 mm) with Glidex hydrogel Coating Locking Pigtail and Radiopaque Marker, REF /Catalog No. 27-260 UPN; M001272600, Use before 2011-11, Sterilized with Ethylene Oxide Gas, Single Use Only, Made in USA 780 Brookside Drive, Spencer, IN 47460-1080. |  |  | July 01, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific is initiating a recall of a single batch/lot of this product. We received a report that the labeling on the device does not match the labeling on the product packaging. Specifically, the packaging pouch for the product and outer carton labeling reflect that the device is a Flexima Biliary Catheter, which is correct, but the heat shrink inscription directly below the hub on the | Labeling mix-ups | Boston Scientific issued an "Urgent Medical Device Recall -Immediate Action Required" letter dated July 1, 2009. The letter was addressed to Risk Manger / Field Action Contact describing the problem and units affected and that further distribution or use of affected product should cease immediately. The Reply Verification Tracking Form enclosed with the letter was requested to be completed and returned.For further information, contact Boston Scientific at 1-508-683-6036. |  | 1 | 79 units (1 box = 1 unit) | Worldwide Distribution -- United States (AZ, AR, CO, DE, FL, GA, MI, MT, NE, NC, OH and WI), Spain, Finland and Netherlands. | Terminated | December 02, 2011 | 884 | 1 | Not_Computer
Z-1885-2009 | 52761 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IntelePACS Picture Archiving Communications System InteleViewer module, Model #3.X.X. |  |  | October 27, 2008 | 2008 | Intelerad Medical Systems Inc. | Product software versions 3.5.2.P1 to 3.5.2.p56 inclusive and 3.6.1.P1 to 3.6.1.P82 inclusive were defective. | Software design | Intelerad Medical Systems issued an Urgent Medical Device Filed Safety Notice" dated October 24, 2008 describing the affected product and steps required for system upgrade. Users are requested to complete and return an acknowledgement form via email to quality@intelerad.com.For further information, contact Intelerad Medical Systems at 1-514-931-6222. |  | 17 | 17 (including 11 free trials) | Worldwide Distribution -- United States (NM, AZ, CT, MN, TX, NC, TN, FL and CO), Canada, Australia, and New Zealand. | Terminated | September 24, 2010 | 697 | 18 | Software | Device Operation | Software Update | Software Update
Z-1817-2009 | 52375 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Lateral Decubitus Alignment Guide, Non Sterile; manual surgical instrument. The MIS Hip Acetabular instrument set includes the curved positioner/impactor, supine alignment guide, and decubitus alignment guide. |  |  | April 27, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | The Lateral Decubitus Alignment Guide either did not assemble to its mating part, the Curved Impaction Handle, or partially engages the handle, but does not fully lock. | Nonconforming Material/Component | Important Market Withdrawal Letters were sent to Stryker Branches on April 27, 2006 vis Federal Express. The letter asked consignees to examine their inventory and identify the product. They are to retrieve all affected product lots and return to their branch or agency warehouse for reconciliation. Also, they are to reconcile all product on the attached Product Accountability Form and fax a copy. Affected should be returned and their Stryker Orthopaedics Customer Service Rep should be contacted to re-order the product that is being returned as part of this recall. Questions should be referred to Rita Intorella at 201-831-5825. |  | 61 | 61 units | Nationwide Distribution | Terminated | October 15, 2009 | 1267 | 61 | Not_Computer
Z-1904-2009 | 52725 | No | NUC | Condom. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | ConceiveEase (TM), Fertility Friendly Lubricant, Reproductive Lab, Memphis, TN |  |  | July 10, 2009 | 2009 | Memphis Fertility Laboratory Inc | Product was marketed without a 510(k). | PMA | The firm notified its sole consignee on 07/10/2009 by telephone and followed by an Urgent Medical Device Recall letter. The letter requested that they examine their inventory and quarantine all products subject to recall and immediately contact Dr. William Kutteh for return instructions. The enclosed response form should be completed and returned. Questions should be directed to Dr. Kutteh at 901-261-2401. |  | 800 | 800 units | Nationwide Distribution -- USA, including the state of MA. | Terminated | October 05, 2009 | 87 | 800 | Not_Computer
Z-1729-2009 | 52636 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Clinac with Version 7.x Software, manufactured by Varian Medical Systems, Palo Alto, CA. |  |  | May 29, 2009 | 2009 | Varian Medical Systems Oncology Systems | Unexpected Movement: if the stereotactic motion disable function is turned on, the couch can be moved via float mode unexpectedly. | Device Design | Varian Medical Systems issued an "Urgent Medical Device Correction" dated May 29, 2009 distributed to all affected users, as well as Varian Sales, Marketing and Service organizations, informing them of the issue. Consignees were informed that a Varian Service representative will be in contact to schedule the proper correction.For further questions, contact your local Varian Customer Support District Manager. |  | 1933 | 1,933 units | Worldwide Distribution. | Terminated | May 18, 2011 | 719 | 1933 | Other | Physical Safety Hazards | Correction | Repair
Z-1945-2009 | 52126 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Maxilift Patient Lift, Non-AC Powered; a patient lift with a four point tilting frame, adjustable chassis, and optional electronic scale, to move patients in a hospital/nursing home situation; Arjo Inc., Roselle, IL 60172; Model numbers 21102406, 21200001, 21200006, 21000-01, 210024-06, 211010-06, 211024-06, 211026, 212000-06, 212000-06R, 212000-6, 212024-06, 212026, 212624, 218100-01, KGA0200, KMA0520, KMA1003-04, KMA1004, KMA1005, KMA1006, KMA1040, KMA1501-06, KMA1513-06, KMA211024, KMA82124-24, KMB2351, KMB4800-06, KMB4900-06, KMB5501-12US, KRA0300, KRA0310, KRA1000, KRA1500-12, KRA2100, KRA3000-12, MA0500, MA0510, MA1010, MA1020, MA1100, MA1500, Maxi Pillar, Maxilift, Maxipillar, MAXIPILLER, MB0500R, MB0510, and MB0610. |  |  | April 27, 2009 | 2009 | ArjoHuntleigh | The roll pin on certain configurations of the Maxilift hanger bar could break due to a combination of severe corrosion and fatigue. Breakage of the roll pin would lead to the detachment of the hanger bar from the jib of the lifter. | Equipment maintenance | ArjoHuntleigh sent a Field Safety Notice DCHU-UK/0209/001FSN to all affected end users via UPS Ground on 4/27/09. The notice informed the accounts of the potential for the roll pin to break due to a combination of severe corrosion and fatigue, leading to the detachment of the hanger bar from the jib of the lift. The accounts were advised to identify if they have any Arjo Maxilift patient lifters in their facility, identify which version of Arjo Maxilift hanger bar they have using the photographs in Figures 1,2 and 3 in the letter. Any lifters found with a hanger bar attachment as shown in Figure 1 must be immediately and permanently taken out of use, as the replacement hanger bar assembly is no longer available. Any lifters with a hanger bar attachment as shown in Figure 2 must be taken out of use until a correction has been performed by ArjoHuntleigh. Any lifters with a hanger bar attachment design as shown in Figure 3 are not affected and can continue being used. The accounts were provided with a product upgrade offer and a MaxiLift Preventative Maintenance Schedule. The accounts were requested to complete and fax, e-mail or mail the enclosed Customer Response Form to Arjo at 630-925-7974 or quality@arjousa.com, indicating the number of Maxilifts in each category that they have and that the products have been taken out of service for replacement or correction, listing the model numbers and serial numbers of their lifts. Questions were directed to the ArjoHuntleigh Customer Support or Quality Department at 1-800-323-1245. |  | 8015 | 8,015 lifts | Worldwide Distribution -- Countries of USA and Mexico. | Terminated | February 25, 2011 | 669 | 8015 | Not_Computer
Z-1897-2009 | 52640 | Yes | FIR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Owner Manual Modifications for PlasmaPro Plasma Pump Model PP-05. Apheresis Technologies, Inc. Palm Harbor, FL. - No labeling or Packaging. |  |  | April 01, 2009 | 2009 | Apheresis Technologies, Inc. | Apheresis Technologies, Inc. is taking the following Corrective Action: Modification of the Owners Manual Only for the PlasmaPro Plasma Pump Model PP-05. Operator's Manual will modify tubeset loading instructions so as to reduce possible loading error which might result in the tube being pulled into the pump head. | Labeling design | PlasmaPro Plasma Pump Operators Manual was modified and distributed to firms consignees. Firms received written notification of re-issuance of Operator's Manual. |  | 7 | 7 Owner's Manuals | Nationwide Distribution -- Florida, Colorado and Utah. | Terminated | July 06, 2010 | 461 | 7 | Not_Computer
Z-1872-2009 | 52199 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST Automated Manifold Kit , REF Model BT2000, SKU # 014613, Sterile R, Made in Mexico, Manufacturer; ACIST LMedical Systems Inc, 7605 Fuller Road, Eden Prairie, MN .55344. |  |  | May 11, 2009 | 2009 | Acist Medical Systems | ACIST Medical Systems initiated a recall due to an increase in the incidence of field reports related to a bulge on the side and/or leaks in the bond of the 2.5 high-pressure tubing of BT2000 kits during a procedure which can result in unsatisfactory performance of the ACIST BT2000 Automated Manifold Kit. | Nonconforming Material/Component | Consignees were sent on 5/11/09 a ACIST Medical Systems "Urgent: Medical Device Recall" letter dated May 11, 2009. The letter described the problem and the product. It advised consignees to discontinue the use or further distribution of the recall product and to return the inventory and the enclosed "Medical Device Recall Form". Please use the enclosed form with firms address: 7905 Fuller Rd, Eden Prairie, MD 55344. Additional questions please contact the firm at 1-877-272-2269, option 1. |  | 62 | 62, 280 (57,110 US, 5130 Canada, 40 Hong Kong) | Worldwide Distribution including the states of AL, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, ID, IN, IA, KS, KY, LA, MD, MI, MN, MS, MO, NE,NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT VA, WA, and WI.The countries ofCANADA and HONG KONG | Terminated | February 04, 2012 | 999 | 62 | Not_Computer
Z-1929-2009 | 52055 | No | KEY | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | VP 2000 Processor; an automated tissue stainer |  |  | September 15, 2008 | 2008 | Abbott Molecular | Thermal runaway conditions, in that the target temperature was exceeded and the temperature display could not track the actual temperature within the heated slide basins of the VP 2000 Slide Processor. | Device Design | Abbott sent Device Recall/Urgent Field Safety Notice letters dated 9/15/08 to the VP2000 customers on 9/12/08, informing them that thermal runaway conditions may occur within the heated slide basins of the VP2000 Slide Processor. The customers were given actions to take to prevent this condition: a) VP2000 should be powered down when not in use and when the controlling computer is not in use; b) Liquid level in all heated basins should be checked periodically to ensure that there is adequate reagent present for slide processing and to keep the sensor probe immersed for accurate temperature sensing; and c) If large, rapid and random fluctuations in indicated basin temperature occurs the unit should be shut down and a Field Service Representative be contacted. Any questions were directed to their local Abbott Technical Service Representative. |  | 261 | 261 units | Worldwide Distribution -- USA, Australia, Austria, Belgium, Canada, Columbia, Czech Republic, Egypt, France, Germany, Hong Kong, India, Ireland, Italy, Japan, Korea, Malaysia, Norway, Russia, Saudi Arabia, Singapore, Switzerland and the United Kingdom. | Terminated | January 07, 2010 | 479 | 261 | Not_Computer
Z-1900-2009 | 52722 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ArthroWand Saber 30 with Integrated Cable wand, Catalog number AC4330-01, Manufactured by ArthroCare Corporation, Sunnyvale, CA |  |  | August 05, 2008 | 2008 | ArthroCare Corporation | Potential Sterility Loss-- Due to wearing or puncture in the Tyvek lid, the product may lose sterility. | Package design/selection | Firm initiated this recall on August 5, 2008. Consignees were notified via letter asking them to contact consignees and delivered via courier. Distribution discontinued on July 9, 2008. |  | 23443 | 23,443 units (includes distribution of previously returned product) | Worldwide Distribution including USA and countries of Netherlands, Slovak Republick, Ireland, Spain, Switzerland, Singapore, Russian Federatin, Greece, Japan, Saudi Arabia, India, Slovenia, Israel, South Africa, South Korea, Poland, Thailand, Jordan, Belgium, Finland, Portugal, Cyprus, Italy, UAE, Kuwait, China, Bahrain, Denmark, Norway, Sweden, Australia, Germany, Austria, UK, New Zealand, Mexico and Canada. | Terminated | June 14, 2011 | 1043 | 23443 | Not_Computer
Z-1990-2009 | 51999 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BHM Combi Sling, 100% Polyester Shell/100% Polyester Strap; soft polyester net fabric sling for ceiling and floor patient lifts; BHM, Magog, Canada, made in China. |  |  | April 07, 2009 | 2009 | B.H.M. Medical, Inc. | Premature failure of the stitching at the junction of the shoulder strap and sling body of the BHM/Medi-Man Combi Sling used with patient lifts. | Device Design | Arjo sent "Urgent Device Recall Customer Notification" cover letters dated April 15, 2009 to the end user accounts who received the affected BHM/Medi-Man Combi Slings, to the attention of the Administrator/Risk Manager, advising them of the potential for possible stitching failure that could lead to a hazardous situation. Enclosed with the letter was a copy of the BHM Medical Field Safety Notice, FSN 10/2009/ca, dated April 3, 2009 which was sent to their BHM customers in the U.S. on April 7, 2009. The FSN identified the part numbers of the BHM/Medi-Man Combi Slings, described the potential health hazard, and listed the actions to take. The accounts were instructed to inspect and assess their slings' suitability for use per the Slings Maintenance Instructions. If the sling is showing this kind of breakage, the account was instructed to remove it from use and discard it immediately. If the inspection reveals no deterioration of the stitching, the sling is safe to use. The accounts were requested to complete and return the enclosed FSN receipt form within 10 days, acknowledging the receipt and understanding of the letter, and indicating the number of slings in use prior to inspection, the number of slings in use after the inspection, and the number of slings rejected. The accounts will be contacted for the replacement of their eligible old slings (in use up to two years). |  | 1396 | 1,396 slings | Nationwide | Terminated | January 07, 2011 | 640 | 1396 | Not_Computer
Z-1894-2009 | 52560 | Yes | KDJ | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter Minicap Extended Life PD Transfer Set Extra Short (4") with Twist Clamp; six sets per carton; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 5C4483 The set is used in disconnect applications and cycler applications where aseptic connections and disconnections are performed at the transfer set/cycler juncture. |  |  | July 02, 2009 | 2009 | Baxter Healthcare Corp. | The peritoneal dialysis transfer set has a defective white sleeve twist clamp. The twist clamp will not open or close completely, which could cause leakage and lead to peritonitis. | Nonconforming Material/Component | Baxter Healthcare Corporation telephoned the accounts receiving the affected lots and then faxed or e-mailed "Urgent Product Recall" letters dated July 2, 2009, to them, to the attention of the Peritoneal Dialysis Nurse, to inform them of the defective white sleeve twist clamp, which may result in leakage and possible peritonitis. The accounts were asked to immediately discontinue the use of the affected lot numbers and examine their inventory for the affected lots. If any affected product was found, they were requested to quarantine it and contact Baxter Center for Service at 1-888-229-0001 to arrange for its return to Baxter. The accounts were also requested to review their patient records to determine if any of their patients had received the affected product. If so, they were advised to contact the patients to see if they had experienced any leaking. If so, the patients should return to the facility to have the set changed per the facility's transfer set change procedure. Any questions were directed to the Center for One Baxter at 1-800-422-9837, Monday through Friday, 8:00 a.m. to 5:00 p.m. Central Standard Time (CST). The accounts were requested to complete the enclosed reply form acknowledging receipt of the letter and that they followed the instructions and disseminated the information to their staff. |  | 2436 | 2,436 sets | US (states of Arkansas, Illinois, Mississippi, North Carolina, Tennessee and Virginia) and El Salvador. | Terminated | July 08, 2010 | 371 | 3150 | Not_Computer
Z-1869-2009 | 52663 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet Coronary Perfusion Cannulae with Balloon Tip, with either self-inflating balloon tip or cone distal tips, straight or right angle, available with proximal end fittings for either 1/4 inch connector or male luer fitting, Model numbers 210800, 210801, 210802, 210803, 210804, 210805, 210806, 210807, 210810, 210811, 210812, 210813, 210814, 210815, 210816, 210817, 215800, 215802, 215803, 215804, 215805, 215806, 215807, 215810, 215811, 215812, 215813, 215814, 215815, 215816, and 215817. |  |  | May 19, 2009 | 2009 | Maquet Cardiovascular | Undeclared allergen: Products are missing the warning label declaring that the products contain natural rubber latex. | Process control | Maquet Cardiovascular issued a notification letter to consignees on May 19, 2009. |  | 2226 | 2,226 devices | Nationwide distribution. | Terminated | December 21, 2010 | 581 | 2226 | Not_Computer
Z-1995-2009 | 51278 | No | MLD | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Datex-Ohmeda S/5 Compact Critical Care Monitor with L-CICU01, L-CICU02A, L-ICU05 &amp; L-ICUO5A software. |  |  | February 06, 2006 | 2006 | GE Medical Systems Information Tecnology | GE Healthcare (Datex-Ohmeda) S/5 Critical Care Monitor and S/5 Compact Critical Care Monitor with L-CICU01, L-CICU02A, L-ICU05 &amp; L-ICUO5A software had an issue where the Heart Rate (HR) was displaying as ---(3 dashes) when the monitor was used continuously for 49 days without a cold start. A missed alarm or missed HR value could cause a delay in recognition of a cardiac event that may result i | Software design | GE Healthcare Technologies issued a Marketing Bulletin "Ref: MB MS HKI 06007" dated February 6, 2006 to explain a software update. |  | 507 | 507 | ITALY, AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, BRASIL, BULGARIA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, IRAN, ISRAEL JORDAN, LITHUANIA, ITALY, MACEDONIA, MALAYSIA, MEXICO, NETHERLAND, NICARAGUA, NORWAY, POLAND, PORTUGAL PUERTO RICO, RUSSIA, SINGAPORE, SLOVAK IA, SOUTH AFRICA, SPAIN, SWEDEN, SYRIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA. | Terminated | May 17, 2012 | 2292 | 2760 | Computer
Z-1936-2009 | 52427 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle Lite Blood Glucose Test Strips, CSP, 50's US, Checkpoint, FS Lite, for use with FreeStyle Lite Blood Glucose Meters and FreeStyle Freedom Lite Blood Glucose Meters, Part Number 70822-02, manufactured by Abbott Diabetes Care, Inc., Alameda, CA. |  |  | June 10, 2009 | 2009 | Abbott Diabetes Care, Inc. | Low results: Freestyle Lite Blood Glucose Test strip may generate low blood glucose readings when tested with the control solution from lot 0822524 with Freestyle Freedom Lite and Freestyle Lite Blood Glucose Meters. | Process control | An "Important: Medical Device Notification" letter dated June 10, 2009 was issued to consignees. First consignees received notifications via letters to both direct consignees and registered customers of FreeStyle Lite and FreeStyle Freedom Lite blood glucose meter users in the states where the affected test strip lot was distributed. Affected accounts and customers were instructed to discontinue distribution and use of the product. The notification will be posted to the Abbott Diabetes Care website. Direct questions about the recall to Abbott Diabetes Care, Inc. by calling 510-749-5400. |  | 9252 | 9,252 vials | US only (states AL, AZ, CA, FL, GA, IL, IN, MD, NJ, NY, OH, PA, RI, SC, TN, TX and WI). | Terminated | January 10, 2011 | 579 | 9252 | Not_Computer
Z-1914-2009 | 52168 | No | DSQ |  |  | Cardiovascular | PMA | Abiomed AB5000 Portable Driver |  |  | May 08, 2009 | 2009 | Abiomed, Inc. | The Emergency System Operation (ESO) mode may not provide adequate patient support under a specific operating condition. | Finished device change control | Abiomed notified customers via an Urgent - Voluntary Medical Device Field Removal Notice on May 8, 2009 advising users that the Emergency System Operation (ESO) mode may not provide adequate patient support under a specific operating condition. The letter stated that customers may continue to support patients with the AB Portable Driver during the process since the failure has not occurred during patient support. An Abiomed Field Service Engineer will contact customers and arrange for a temporary exchange for the affected driver with a corrected driver. The firm recommends that customers switch the patient to hand pump support and switch patient to backup driver until they can confirm that the Driver(s) has been corrected by an Abiomed Field Service. Questions are to be directed to Michael O'Hara at 978-646-1543 or a Field Service at 800-554-8666. |  | 18 | 18 units | Worldwide Distribution -- USA, including states of AL, IN, FL, NJ, NY, and OH and countries of Germany, Greece, and Sweden. | Terminated | January 26, 2012 | 993 | 18 | Not_Computer
Z-1898-2009 | 52762 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | SmartPort CT Port Access System, Catalog number CT80STPD, AngioDynamics Incorporated |  |  | July 16, 2009 | 2009 | Angiodynamics, INC | The product may contain an incorrect size catheter and introducer. | Packaging process control | Notification of the AngioDynamics, Inc. sales force began on July 16, 2009 and they were asked to assist in the return of all affected devices to the Manchester, GA division of AngioDynamics, Inc. Urgent Recall letters were sent to all customers on July 16, 2009, requesting the completion of a FAX BACK form. Those currently in possession of any remaining devices will be provided a Return Goods Authorization (RGA) number and pre-paid shipping boxes for return of the product to AngioDynamics, Inc. All returned devices will be physically and electronically quarantined until final disposition of the devices. All devices will be inspected for damage to the devices and any damaged devices destroyed (scrapped) via the AngioDynamics procedure for Disposition of Nonconformances, procedure. Non-damaged product may be reworked to remove the incorrect catheter and introducer and replaced with the correct catheter and introducer, repackaged and re-sterilized in accordance with the AngioDynamics procedure for Rework. If necessary, follow-up letters, phone calls or visits by the AngioDynamics sales representatives may be used to complete the effective checks. Level A (100%) of consignees to be contacted).Please contact your local representative or Recall Coordinator at 1-706-846-3126, for assistance. You may e-mail your questions to customerservice@angiodynamics.com. |  | 150 | 150 kits | AL, CA, CO, FL, GA, ID, IL, IN, KS, LA, MI, MO, MN, MS, NC, NY, OH, OR, PA, TN and TX. | Terminated | October 27, 2010 | 468 | 150 | Not_Computer
Z-1910-2009 | 52758 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Jejunal Feeding Set, Product Order number FLOWJ-9-20. |  |  | July 24, 2009 | 2009 | Cook Endoscopy | The English only version of the Patient Care Manuals that accompany the device were shipped to a medical facility in Italy. | Process control | The consignee was notified by an Urgent: Field Safety Notice, dated 07/24/2009. The distributor (Cook Ireland) was instructed to contact the Italian Company Representative. The appropriate Patient Care Manual with the Italian translation was provided to the medical facility with the assistance of the Area Representative. The Area Representative instructed the medical facility staff to provide the translated Patient Care Manual to the caregiver(s) of the patient that received the Jejunal Feeding Sets. The distributor is already aware of the situation. The lower portion of the letter was to be filled out and sent to Cook Endoscopy. |  | 4 | 4 units | International Only Distribution -- Country of Italy. | Terminated | April 06, 2010 | 256 | 4 | Not Computer
Z-1947-2009 | 52604 | No | DZA | Intraoral dental drill. | Dental | Dental | 510(K) Exempt | ACT3020 Advanced Cutting technology (ACT) Twist Drill 3.0mm(D) X 20mm(L). BIOMET 3i 455 Riverside Drive, Palm Beach Gardens, FL 33410 USA. Made in Switzerland. |  |  | June 29, 2009 | 2009 | Biomet 3i, LLC | Some of the packages of ACT3020, drills may actually contain an ACT3220 drill. | Mixed-up of materials/components | Consignees were notified of recall by telephone and by faxed an Urgent: Medical Device Recall letter, dated June 30, 2009. The letter stated the problem. The letter also asked customers to check their surgical kits and confirm the drill is placed in the correct space within the kit. Customers are to verify that the shank of the drill is marked as ACT3020 on one side and as 3.0 x 20mm on the opposite side. The letter also stated the clinical implications. Customers are to respond with the attached response fax and then return the drills to BIOMET 3i. Once the returned product is received, replacement drills will be sent to the customers. Questions or concerns should be directed to Kelly Taylor at 561-776-6906. |  | 29 | 29 units | Worldwide Distribution -- Korea, Netherlands, Taiwan, United Kingdom, and USA. | Terminated | January 05, 2010 | 190 | 29 | Not_Computer
Z-1706-2009 | 52561 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | GE Medical Systems, OEC, Vascular Drape, 00-902776-01, Sterile. 150" long x 87" wide. Distributed by GE OEC Medical Systems, Salt Lake City, UT 84116 USA. |  |  | August 15, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has found that the sterile products manufactured by Contour Fabricators, Inc. (CFI) for GE healthcare (GEHC) may have seals that could be compromised during the shelf life of the products. This issue could result in the product not maintaining its sterile state. Eleven complaints and 3 MDR's have been filed. No illness or injuries reported. | Packaging process control | GE Healthcare issued an Urgent Recall Notice dated August 15, 2008. The letter described the problem, provided a list of products, instructed customers to stop using the products and to return or dispose. To receive credit, customers must complete the attached product return form and include it in the package with returned items. All customers were asked to complete and fax back the attached response form. |  | 2760 | 2760 | Nationwide Distribution -- AZ, CA, FL, MD, NY, OR, PA, SC and TX. | Terminated | May 15, 2012 | 1369 | 5820 | Not_Computer
Z-1951-2009 | 52803 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | PowerProMax PRO5100M Battery Single Trigger, |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | There is a potential the devices may self activate for their PowerPro, PowerProMax, and MPower I Handpieces manufactured prior to June 1, 2008. | Device Design | Consignees were notified via letter by FedEx delivery or equivalent method. If the product has been further distributed the consignee is requested to notify their consignees. The letter indicated the following risk mitigation actions until affected handpieces are returned to the manufacturer for preventive maintenance: 1) Apply supplied label to the identified handpieces by specific serial number, 2) If self-activation is experienced, discontinue use immediately and return to ConMed Linvtec and 3) Advise/remind the consignee to follow the manufacturers preventive maintenance intervals specified in the instructions for use manuals. Direct questions to ConMed Linvatec at 1-800-237-0169.Recall Expansion: New amended customer and distributor letters will be sent to those identified customer/distributor accounts. The letter will address the additional units. |  | 5553 | 5,553 (4,235 INITIAL AND 1,318 EXPANDED) | Worldwide Distribution including United States. | Terminated | December 19, 2011 | 871 | 23426 | Not_Computer
Z-2007-2009 | 52812 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | MC5057 12ft/3.6 meters Universal Cable. Made in USA. NON STERILE Rx Only. |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | Sufficiently worn/damaged Universal Cable with product number MC5057, when used with various ConMed Linvatec electric Handpieces, may cause the handpiece to self activate. | Component design/selection | ConMed Linvatec notified customers via letter, cautioning the users to inspect cords for signs of excessive wear or damage and urging them to discontinue use and return cables when appropriate. Direct questions to Linvatec Corporation by calling 1-800-237-0169. |  | 12205 | 12,205 units | Worldwide Distribution including United States. | Terminated | July 11, 2011 | 710 | 12205 | Not_Computer
Z-1715-2009 | 52217 | No | JDL | Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | CONSERVE, TOTAL A-CLASS HEAD W/BFH TECH., REF 38AM-3604, O.D. Size: 36mm, Neck Length: -3.5mm, Taper: SLT, Sterile R, 1 each, Rx only, Wright Medical Technology, Inc., 5677 Airline, Road, Arlington, TN 38002. |  |  | May 14, 2009 | 2009 | Wright Medical Technology Inc | Incorrect packaging, swapped sizes | Packaging process control | Wright Medical Technology, Inc. notified hospitals, surgeons and their distributors of the recall situation via an "Urgent: Medical Device Voluntary Recall" notice dated May 14, 2009 explaining the problem, the hazard involved and requesting confirmation that the product was removed from their inventory. The letter included a "Fax Back" response page to ensure the receipt of the notice.For further questions, contact Wright Medical Technologies, Inc. at 1-800-874-5630. |  | 14 | 14 units | Worldwide Distribution -- Including states of NC, TN, IA, IN, MD, WI, WV, OK, WA, PA, TN, TX, FL and CA and country of Canada. | Terminated | December 09, 2009 | 209 | 29 | Not_Computer
Z-1717-2009 | 52428 | No | KXK | Radionuclide brachytherapy source. | Radiology | Radiology | 510(k) | M-19 (Iridium 192): Source Production and Equipment Company, St. Rose, LA. |  |  | March 03, 2008 | 2008 | Source Production &amp; Equipment Inc | Radioactive source broke during use. | Other | Source Production and Equipment, Inc. notified the Distributor by phone and letter March 2008 and requested the return of all affected units. |  | 3 | 3 units | Nationwide Distribution -- GA, NV and NY. | Terminated | August 19, 2009 | 534 | 3 | Not_Computer
Z-1895-2009 | 52565 | No | FHN | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Boston Scientific Clipping Device, Resolution Clip, working length 155 cm, working channel 2.8 mm, Manufactured for Boston Scientific Corporation, Natick, MA; Product codes M00522600 (1 clip) and M00522602 (20 clips). |  |  | June 17, 2009 | 2009 | Boston Scientific Corporation | The clip may not deploy or, if deployed, the clip may not release from the tissue. | Process control | Consignees were notified by an "Urgent Recall Letter " dated June 17, 2009 and instructed to cease using the product because the calibration of equipment used to manufacture these products was not in control. This may potentially result in an inability to deploy or release the clip from the catheter. Consignees were instructed to return a business reply form for return authorization. Please contact Cindy Finney, Field Action Team Lead at 1-508-683-4678 for assistance. |  | 217 | 217 | Nationwide, France, Germany, Great Britain, Italy, Netherlands, Spain and Sweden. | Terminated | October 08, 2010 | 478 | 4486 | Not_Computer
Z-1889-2009 | 52679 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | NX 2008 Central Monitoring System (CMS). Product code: E2FB6 |  |  | June 22, 2009 | 2009 | AGFA Corp. | After opening and closing a study on a Central Monitoring System while a study is still open on an in-room NX system, a problem of image mix-up can occur on the in-room NX system. | Software design | An "Urgent Safety Notice" was sent to all potentially affected sites on/about June 22, 2009 to make them aware of the issue. A FAX Back form is requested with each letter to ensure someone at the account acknowledges the information was received and understood. If a FAX Back form is not received, the account will be contacted by telephone. A software service update to solve this issue is expected to be released by August 15, 2009. A Mandatory Service Bulletin will be deployed to correct the problem in the software. If you have questions, please contact Agfa Healthcare at 1-800-581-2432 prompts 1, 3 and then 2 for CR Support. |  | 97 | 97 units | Worldwide - IL, PA, TX, Canada, Europe, Asia Pacific and Latin America | Terminated | July 22, 2010 | 395 | 97 | Software | Display/Image | Software update | Software Update
Z-1444-2009 | 51704 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sybaritic, NannoLight MP50. |  |  | April 12, 2008 | 2008 | Sybaritic, Inc | Sybaritic Inc. had an issue with the 585nm filter used with the NannoLight MP50. The filter may cause unwanted burning and should not be used for vascular lesion treatments. | Device Design | A "Sybaritic" notification letter dated April 15, 2008 was issued to consignees. The letter described the modifications to the protocols for the NannoLight MP50 and the temporary replacement of 585 Filter for the 560 Filter for Vascular Lesions until further notice. They requested the immediate use of the protocols replacing all other versions.A follow-up letter dated June 3, 2008 was issued to customers. The letter included an attachment " Product Notice Nannolight MP50 585 nm "special" filter from Sybaritic, Inc. The letter described the new 585 nm "bandwidth filter". Also the notification advised consignees to discontinue the use of the old 585 filter and to use 560nm filter until they receive the new 585 nm bandwidth filter.Please direct any questions to your Account Manager or to Sybaritic, Inc. by calling 1-800-445-8418 or 1-952-888-8282. |  | 59 | 59 | Worldwide Distribution -- US (states of WA, CA, MN, FL and AR), POLAND, THAILAND, CHILE, CANADA, KOREA, AZERBAIJAN, INDIA, SPAIN, ARGENTINA, RUSSIA, GREECE, PHILIPPINES, POLAND, INDIA, JORDAN, TURKEY, MEXICO, URUGUAY,BELGIUM, LEBANON, BULGARIA and SINGAPORE. | Terminated | July 05, 2012 | 1545 | 59 | Not_Computer
Z-1923-2009 | 52115 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Malibu Polyaxial Screw, 7.5 x 35mm Part Number 12-7535 |  |  | March 04, 2009 | 2009 | SeaSpine, Inc | SeaSpine has received reports relating to rust on the internal washer contained within the titanium seat assembly of certain Malibu Polyaxial and Malibu Reduction Screws. | Mixed-up of materials/components | In an Urgent Voluntary Product Recall Notification letter dated March 5th, 2009, SeaSpine informed consignees about the reason for recall and asked them to immediately check the products in their possession and hospital inventories to determine if they have any of the affected product. Alternatively, there were told if they would prefer to return to SeaSpine all Malibu Screw inventory that is in their possession, SeaSpine will sort out the affected product upon receipt and will replace their inventory with Malibu Screws that conform to all SeaSpine specifications. The part number and lot are laser marked on the outside of the seat assembly of each screw.SeaSpine asked their consignees to return affected product to them immediately. Attached was a verification form, that they were told to complete and return even if they did not have any product to return. Furthermore their local Sales Manager could assist them in completing the form. They were also encouraged to contact their Sales Manager or Steve Rybka at (760) 727-8399 x202 if they had any questions regarding the recall, any of their products, or would like assistance with the recall. |  | 1116 | 1,116 Screws in total | Worldwide Distribution -- USA, including states of AL, AZ, CA, CO, FL, GA, IN, MD, MI, MO, MS, NM, OK, TN, TX, VA, and WI and country of Greece. | Terminated | April 28, 2011 | 785 | 1116 | Not_Computer
Z-2050-2009 | 52695 | No | BSO | Anesthesia conduction catheter. | Anesthesiology | Anesthesiology | 510(k) | Portex Continuous Epidural Tray, Reference 4037-20 |  |  | July 15, 2009 | 2009 | Smiths Medical ASD, Inc. | Portex Continuous Epidural Tray was distributed with a 20G closed end, three eyed, nylon catheter without a stylet instead of the 20 G open end Teflon catheter with stylet specified on the label. | Mixed-up of materials/components | Smiths Medical notified consignees by an Urgent Device Recall letter dated July 15, 2009, to locate and quarantine the affected product, stop distribution, and complete and fax back the attached Distributor Reply Letter to Smiths Medical. Credit or replacement will be arranged as requested for the return of any affected product. All distributors will be required to return any affected product in stock. All of the distributors will be required to send their customer list to Smiths Medical who will notify their customers of this recall. |  | 190 | 190 kits (19 cases) | Nationwide Distribution -- including states of New York, Texas, Ohio, Utah, Alabama, Arizona, Kentucky, &amp; Wyoming. | Terminated | September 14, 2012 | 1157 | 190 | Not_Computer
Z-2004-2009 | 52724 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Roche CoaguChek XS System, CoaguChek XS Plus Care Kit, Roche Diagnostics, Indianapolis, IN; Model Number 05021537001. |  |  | July 22, 2009 | 2009 | Roche Diagnostics Operations, inc. | The recommended cleaning procedures for the meter may cause a malfunction and falsely elevated test result reports. | Device Design | Roche Diagnostics Operations, Inc. notified Consignees by letter dated July 22, 2009 informing them of the problem and instructing them to use revised cleaning procedures.For further questions, contact Roche Diagnostics Operations, Inc. at 1-317-521-3194. |  | 1327 | 1,327 in US | Worldwide Distribution | Terminated | April 27, 2011 | 644 | 66326 | Not_Computer
Z-1998-2009 | 52817 | No | FDG | Gastroenterology-urology fiberoptic retractor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Ultra-Thin LightMat Surgical Illuminator, Disposable one time use. Model Number: UA2550, Exp Date: 2012-05. Product is packaged in individual sterile pouch w/5 pouches per box. |  |  | July 01, 2009 | 2009 | Lumitex Inc | Adhesive which is used to attach LightMat to retractor may not have been tested to Medical standards. The vendor provided an adhesive strip that did not conform to Lumitex specifications. The adhesive strip provided may not have met Lumitex requirements to be medical grade, EtO compatible, and biocompatible. | Nonconforming Material/Component | Lumitex, Inc. issued a recall letter dated July 2, 2009 to their Customers describing the affected device and requesting its return.For further information, contact Lumitex, Inc. at 1-440-243-8401. |  | 295 | 59 packs/5 individual units | Product was distributed to the following states: CA, FL, KS, MA, ME, MI , NC, NM, NV, NY, OH, and TX. | Terminated | March 09, 2012 | 982 | 295 | Not_Computer
Z-1961-2009 | 52611 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Celsite Access Port Model ST305P. |  |  | June 22, 2009 | 2009 | B. Braun Interventional Systems | B. Braun Interventional Systems was notified by B. Braun Medical France, the manufacturer of the Celsite Access port that they had received 11 reports of polyurethane (type P) catheter disconnections. B. Braun Medical France has decided to recall all CelsiteÂż ports supplied with the affected catheter lot. To date, B. Braun Interventional Systems (US) has not received any reports of catheter dis | Nonconforming Material/Component | B. Braun Interventional Systems, Inc. issued an "Urgent: Medical Device Recall" notice dated June 22, 2009. The letter identified the problem and product. The letter reminded the consignee of the general instruction for use of access ports and instructed them to contact patients who had the Celsite port system implanted.For further questions, contact B. Braun Interventional Systems, Inc. at 1-800-523-9676 extension 4818. |  | 40 | 40 in US | Nationwide Distribution -- NY. | Terminated | June 05, 2012 | 1079 | 40 | Not_Computer
Z-2010-2009 | 52898 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Transtracheal (Jet Ventilation) Catheter ACU 1060.1, labeled sterile if unopened. |  |  | July 09, 2009 | 2009 | Instrumentation Industries Inc | user instructions not updated to correspond with product change | Component change control | Instrumentation Industries, Inc. issued a notification letter dated July 9, 2009 to their customers explaining the problem and the need to replace current inventory user instructions with the revised user instructions accompanying the letter.For further questions, contact Instrumentation Industries, Inc. at 1-800-633-8577. |  | 396 | 396 units | Nationwide Distribution -- AK, AL, AR, CA, CO, FL, GA, IA, ID, IL,IN, MA, MD, ME, MI, MN, NC, ND, NH, NJ, OH, PA, SC, SD, TX, VA and WA. | Terminated | October 15, 2009 | 98 | 396 | Not_Computer
Z-1965-2009 | 52798 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | iLink Bone Void Filler. Packaged in a double Tyvek pouch, then placed in a box. Model Number: 50-05-0010, Exp. 2010-04. |  |  | August 26, 2008 | 2008 | Theken Spine LLC | Failure to label the product properly. Additionally, 15 pieces were reworked and the expiration date was misprinted as 2011-04 instead of the actual date of 2010-04. | Error in labeling | Theken Spine, LLC contacted the Consignee August 26, 2008 via telephone and instructed them to inventory their stock and return affected product to the firm.For further questions, contact Theken Spine, LLC at 1-330-475-8662. |  | 22 | 22 units | Nationwide Distribution -- TX. | Terminated | March 09, 2012 | 1291 | 22 | Not_Computer
Z-2048-2009 | 52993 | No | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(K) Exempt | Stratus OCT Model 3000 instruments distributed with version 6.0.0 software, product number 2660021119710 (new instrument) or 2660022119710 (refurbished instrument), and Stratus 6.0.0 Software Kit, product number 2660021133164. |  |  | July 31, 2009 | 2009 | Carl Zeiss Meditec Inc | Mis-labeled image. A bug in version 6.0.0 software and 6.0.0 software kit used with Stratus OCT Model 3000 can produce a report that incorrectly labels OD or OS. | Software design | Firm initiated notification to customers on July 2, 2009, by sending a letter via FedEx to US customers, and via e-mail and FedEx to outside US consignees. Customers were informed of the issue and steps needed to correct it. A revised software version will be provided to the customers. Any questions should be directed to Carl Zeiss Meditec, Inc. Customer Care at 1-800-341-6968. |  | 244 | 244 units | Worldwide - US, Austria, UK, Japan, Singapore, Argentina, France, Hong Kong, New Zealand, Canada, Colombia, Germany, Korea, Venezuela, India and China. | Terminated | April 19, 2010 | 262 | 244 | Software | Display/Image | Software update | Software Update
Z-2029-2009 | 50785 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport e360 Ventilator, Rx only. |  |  | December 15, 2008 | 2008 | Newport Medical Instruments Inc | In December 2008, Newport initiated a field correction to upgrade e360 ventilators with software versions 3.0 or higher to software version 3.3. As a result of this upgrade, some customers experienced unexpected device alert alarms. Newport's investigation into these complaints confirmed that the 3.3 software responds too quickly to certain situations, giving a false device alert. Newport condu | Software design | Newport Medical initally sent out a Field Correction Notice to a limited number consignees dated December 17, 2008 and then the notice was expanded to include all e360 consignees dated June 15, 2009. The firm notified consignees that they have revised the e360 software to version 3.8. A customer verification form was attached to verify upgrade completion. For further questions, contact Newport Customer Technical Support Department at 1-800-451-3111 or +1-714-427-5811 Extension 500. |  | 4117 | 4,117 units (214 Domestically; 3,903 Internationally) | Worldwide Distribution -- US (including AL, AZ, CA, CO, FL, GA, IL, IN, KY, MA, MD, ME, MI, NJ, NM, NY, OH, OK, OR, PA, UT and VA), Albania, Algeria, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Brunei, Bulgaria, Canada, Chile, Colombia, Czech Republic, Dominican Republic, Ecuador, Egypt, England, Finland, Germany, Greece, Guatemala, Hong Kong, India, Indonesia, Japan, Jordan, Kazahkstan, Kenya, Korea, Kuwait, Lebanon, Liechtenstein, Lithuania, Malaysia, Mauritius, Mexico, Morocco, Nambia, Nepal, New Zealand, Nicaragua, Nigeria, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Sri Lanka, Taiwan, Thailand, Tunisia, Turkey, AUE, Ukraine, United Kingdom, Yemen Republic and Zimbabwe. | Terminated | November 04, 2011 | 1054 | 4117 | Software | Alarm/Message | Software update | Software Update
Z-1558-2009 | 52019 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Datex-Ohmeda S/5 Device Interfacing Solution, N-DISVENT 02. |  |  | May 08, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware of an issue in the calculations associated with Datex-Ohmeda S/5 Device Interfacing Solution N-DISVENT-02 interfaced between certain Datex-Ohmeda patient monitors (Datex-Ohmeda S/5 Anesthesia Monitor (AM), Datex-Ohmeda S/5 Critical Care Monitor (CCM), Datex-Ohmeda S/5 Compact Anesthesia Monitor (CAM), Datex-Ohmeda S/5 Compact Critical Care Monitor (CCCM)) an | Software design | An "Urgent Medical Device Correction" letter was issued to consignees on or about April 9, 2009. The letter was addressed to Healthcare Administrator/Risk Manager, Chief of Nursing, Director of Biomedical Engineering. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. A representative from GE Healthcare will contact customers to schedule implementation of the correction when it is available for installation at the customer's facility. Direct questions regarding this notice to GE Healthcare Technical Support or your local GE Distributor and refer to FMI 36042. |  | 2638 | 2,638 | Worldwide Distribution -- US (including DC and states of AL, AK, AZ, AR, CA, CO, FL, IL, IN, IA, KS, KY, LA, ME, MD, MI,MN, MO,MT, NE, NH, NJ, NY, NC, OH, OK OR, PA, SC, TX, UT,VT, VA, WA, WV and WI) and countries of AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, CANADA, CHINA, CYPRUS, DENMARK, ESTONIA, FRANCE, GERMANY, INDIA, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PHILIPPINES, PORTUGAL SAUDI ARABIA, SERBIA, SOUTH AFRICA, SWEDEN SWITZERLAND, THAILAND, TURKEY, UKRAINE, UNITED ARAB EMIRATES and UNITED KINGDOM. | Terminated | February 04, 2012 | 1002 | 2638 | Not_Computer
Z-2016-2009 | 51966 | No | LIF | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HACH SteriChek Residual Peroxide Reagent Strips, 5/100-strip bottles per kit, Kit Part #811905, Bottle part #1905LB, Hach Co., Ames, IA. |  |  | April 21, 2009 | 2009 | Hach Co | Some of the reagent strip bottles contain the incorrect strips | Labeling mix-ups | Hach Company issued an "Urgent Medical Device Recall" letter dated April 21, 2009 explaining the reason for recall and requested product remaining in inventory be returned. Subrecall to the end-user level was requested. A Return Form was also enclosed to be completed to indicate the amount of product being returned.Between April 22-23, 2009, the firm telephoned their customers to assure they received the recall letter.For further information, contact Hach Customer Service at 1-800-548-4381 extension 3560. |  | 275 | 275 kits | Nationwide Distribution -- NJ, CA, IN, PA, NV, CA, and ME. | Terminated | October 07, 2009 | 169 | 275 | Not_Computer
Z-2008-2009 | 52209 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Model 100 enFlow Fluid/Blood Warmer |  |  | March 23, 2009 | 2009 | Enginivity LLC | 1. Fluid contamination of the electronic circuitry can cause short circuits resulting in permanent fuse failure or localized circuit board overheating with possible smoldering and discoloration. 2. Electromagnetic (EMI) Interference associated with the Warmer may produce artifact in ECG, EEG, or EMG records (cardiac or neuro monitoring). | Device Design | Vital Signs, Inc. issued Product Safety Notification on March 23, 2009 informing users of potential EMI and fluid ingress issues. All warmers in the field will be replaced beginning in May 2009. |  | 3052 | 3052 units | Worldwide Distribution -- USA, Philippines, Saudi Arabia, India, Israel, Australia, Mexico, Germany, Republic of Singapore, China, Guatemala, Switzerland, Dubai United Arab Emirates, New Zealand, Malaysia, India, Ecuador, &amp; United Kingdom. | Terminated | May 02, 2012 | 1136 | 3052 | Hardware | Device Operation | Replace component | Remove or Replace
Z-1862-2009 | 52163 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker 550 mL AutoFuser PainPump, single site continuous with bolus, 5 mL/hr, 1 x 5.0 in. catheter, sterile, REF 0535-500-050. |  |  | May 14, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | Failure to infuse and potential compromised sterility: The seal on the tyvek tray cover may fail, compromising sterility, and the device may fail to infuse the drug, resulting in a loss of pain mitigation. | Process control | Stryker sent all consignees a Medical Device Recall Notification letter dated 5/14/09 and instructed to locate and destroy the recalled products and to complete and return the business reply card. |  | 28755 | 28,755 of all pumps | Nationwide and Canada. | Terminated | January 20, 2010 | 251 | 28755 | Not_Computer
Z-1808-2009 | 52697 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | American Optisurgical, Horizon Phacoemulsification System, Model: HZN-2030, for ophthalmic surgery. |  |  | March 24, 2009 | 2009 | American Optisurgical Inc | Device operation interrupted: A software bug causes Phaco to stop when vacuum goes above 300 mmHg while in Multi Burst mode. Use of this product at the parameters specified above could lead to a delay in surgery that might present a health risk. | Software change control | American Optisurgery sent an E-mail to customers on April 2, 2009 stating the reason for recall and informed them to immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, please identify customers and identify them at once of this product recall. Their notification to customers may be enhanced by including a copy of this recall notification letter. They were then told to refer to the enclosed Recall Questionnaire for instructions on what to do with the recalled product: The recall should be carried out to the user level and return the enclosed response form as soon as possible. If there are any questions, they are advised to contact David Salzberg at (800) 576-1266. |  | 12 | 12 units | Worldwide distribution: USA (California), Brazil, Algeria, France, South Korea, India, and Peru. | Terminated | September 04, 2009 | 164 | 12 | Software | Device Operation | Remove | Remove or Replace
Z-1600-2009 | 52002 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CIC Pro Clinical Information Center Central Station (CIC Pro version 5.1, MP100 Hardware). |  |  | March 10, 2009 | 2009 | GE Medical Systems Information Tecnology | GE Healthcare has become aware of a potential issue associated with the failure to restart CARESCAPE CIC Pro v5.1 on the MP100 hardware. This issue may impact patient safety. | Component change control | A GE Healthcare "Urgent Medical Device Correction" letter dated February 24, 2009 was sent to consignees. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, and Contact Information. Affected units will be replaced with new units byGE Field Service team members.For questions regarding this notice or experience a failure to restart the device, please contact GE Healthcare Technologies Support Center at 1-800-558-7044, enter Option 2 and ask for Technical Support or your local GE Field Service Representative. |  | 352 | 352 | Worldwide Distribution -- US (states of AL, AZ, CA, CT, FL, GA, IL, IN, KY, LA, MD,KY, MI, MN, MS, MO, NV, NH, NJ, NY, NC, OH, OK, PA, RI, TN, TX, VA, WA, WV and WI), AUSTRIA, AUSTRALIA, CANADA, CZECH REPUBLIC, GERMANY, EGYPT, SPAIN, FRANCE, GREAT BRITAIN, ITALY, POLAND and SWEDEN. | Terminated | October 16, 2012 | 1316 | 352 | Hardware | Physical Safety Hazards | Replace | Remove or Replace
Z-1992-2009 | 52701 | No | MBF | Shoulder joint metal/polymer/metal nonconstrained or semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Turon Revision Glenoid Poly Liner. |  |  | July 20, 2009 | 2009 | Encore Medical, Lp | Product devices used for testing are not identical to what was manufactured. Under certain circumstances snap feature could allow poly liner to move on baseplate reducing tightness of fit. | Release of Material/Component prior to receiving test results | Firm notified consignees by an Urgent: Medical Device Recall letter, dated July 20, 2009. The letter stated the reason for the recall, asked customers to quarantine, and return the affected product. The enclosed response form needed to be completed and returned as soon as possible. Questions were to be directed to David Harris, Manager of Customer Service at 512-834-6330. |  | 60 | 60 units | Nationwide Distribution -- Including states of TX, FL, and GA. | Terminated | January 07, 2010 | 171 | 60 | Not_Computer
Z-2026-2009 | 52773 | valve, prosthesis | ISP | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Icelock Expulsion Valve 551 Model L-551000 |  |  | June 23, 2009 | 2009 | Ossur Engineering, Inc | The plastic caps have come off or become loose when the user stepped into their prosthesis. | Process change control | Customers were sent a Safety Notice Voluntary Product Recall, dated June 23, 2009, that advised them to return any unused valves to Ossur for a free replacement. Those using valves in combination with a cushion liner and suspension sleeve should replace immediately if problems are experienced, otherwise replace at next routine service. Customers using valve in combination with an Iceross Seal-In liner should replace ASAP and return used valves to Ossur for free replacement. Customers with questions can call Darlyne Roa 1-800-233-6263 ext 1023743. |  | 4647 | 4,647 | Worldwide Distribution | Terminated | March 06, 2012 | 987 | 4647 | Not_Computer
Z-1989-2009 | 52775 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Power Supply Accessory (model AD-0660G) |  |  | June 18, 2009 | 2009 | Breg Inc, An Orthofix Company | Power supplies short-circuited at the electrical outlet upon initial use of the power supply, causing the circuit breaker for the electrical outlet to trip. | Employee error | An Urgent: Medical Device Recall Notification was sent to customers on June 18, 2009. FOR USERS: Customers are advised to immediately open all unused Kodiak cold therapy units and inspect the lot code of the Power Supply included with the Kodiak unit. If the lot code is 3508, 3908, 4308, or 4808, then remove the power supply, cut the cord using scissors and discard. Call Breg Customer Service for a replacement: 1-800-321-0607. Complete and return the response form ASAP.FOR DISTRIBUTORS: Send Breg your distribution information, including contact and quantities. Breg will notify your customers of the recall directly. |  | 20439 | 20,439 | Worldwide Distribution | Terminated | March 28, 2011 | 648 | 20439 | Battery | Device Operation | Remove | Remove or Replace
Z-1996-2009 | 52801 | fiber, medical, absorbent | FRL | Medical absorbent fiber. | General Hospital | General Hospital | 510(K) Exempt | Coloplast, InterDry Ag 10" x 144" Roll P/N 6791001400, Manufactured for Coloplast A/S DK-3050 Humlebaek, Denmark, Distributed by Coloplast Corp Minneapolis MN 55411. |  |  | July 21, 2009 | 2009 | Coloplast Corp Skin Care Div | Coloplast determined that a portion of one lot of InterDry" Ag Lot number 1789435, did not meet the product specification for silver content. In this portion of the lot antimicrobial properties of the product could be compromised although the absorption and transportation of moisture would be as expected. | Process control | Consignees were faxed on 7/27/09 a Coloplast "Product Recall" letter dated July 2009. The letter addressed the problem and requested consignees to contact Coloplast Customer Service at 1-800-533-0464 to communicate the status of the product. Replacement product will be provided upon receipt of returned product. It also advised distributors to contact their customers immediately and advise them of the recall and to return the product them. Questions or concerns should be directed to the customer's customer care rep at 1-800-533-0464. |  | 47216.4 | USA, 876 cases (10 rolls/case) ; Canada 53.9 cases (10 rolls/case) | Worldwide Distribution -- USA, including states of AL, AZ, CA, CT, FL, GA, IL, IN, IA, KY. LA, MI, MA,MO, MN, MS, MD, NE, NV, NH, NJ, NM, NY, NC, OH, OK OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI and country of CANADA. | Terminated | January 11, 2011 | 539 | 47216.4 | Not_Computer
Z-2015-2009 | 52826 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Expo, Angiographic Catheter, Sterilized with ethylene oxide gas, Made in Mexico, Parque Industrial La Mesa, Edificion G1, Suites B. H, Tijuana, Baja California 22650 MEX. |  |  | August 05, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific is initiating a recall removal of numerous batches/lots of Wiseguide Guide Catheter, Impulse Angiographic (Diagnostic) Catheter, Expo Angiographic (Diagnostic) Catheter. Boston Scientific has determined that the sterile barrier in the packaging may be compromised. During internal inspections, they discovered gaps between the side seals and the top seal approximately 3/8 in len | Packaging | Boston Scientific issued an "Urgent Medical Device Recall Removal" letter dated August 5, 2009. The letter was addressed to Risk Manger / Field Action Contact. The letter described the problem and product involved. The Firm recommended to segregate affected product immediately and return to Boston Scientific and requested return of the "The Reply Verification Tracking Form" enclosed with the letter.For further questions, contact your local Boston Scientific Sales Representative or call 1-763-494-1634. |  | 367898 | 367,898 (US 167,496; OUS 224,292) | Worldwide Distribution -- US (including DC and Puerto Rico), Japan, United Arab Emirates, Albania, Austria, Belgium, Bulgaria, Bahrain, Switzerland, Czech Republic, Germany, Denmark, Algeria, Egypt, Spain, Finland, France, Great Britain, Greece, Hungary, Ireland, Israel, Iraq, Iceland, Italy, Jordan, Kuwait, Lebanon, Lithuania, Libya, Morocco, Netherlands, Norway, Pakistan, Poland, West Bank Gaza Strip/ Palestine, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Sweden, Slovakia, Syria, Turkey, Ukraine, South Africa, Argentina, Australia, Brazil, Canada, Chile, Colombia, Dominican Republic, Ecuador, Guatemala, Indonesia, India, South Korea, Mexico, Malaysia, Nicaragua, New Zealand, Panama, Philippines, Puerto Rico, Paraguay, Singapore, El Salvador, Thailand, Taiwan, Uruguay, American Virgin Islands and Vietnam. | Terminated | January 07, 2012 | 885 | 787578 | Not_Computer
Z-1948-2009 | 52768 | No | FHQ | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BARDÂż Simpulse" SOLO System with Soft Splash Shield Tip |  |  | July 30, 2009 | 2009 | Davol, Inc., Sub. C. R. Bard, Inc. | Product sterility may be compromised due to lack of package integrity. | Process control | Davol issued a Recall Letter dated July 30, 2009 requesting user facilities examine inventory and immediately discontinue use of any product from the affected lot and return it to Davol. Direct questions about the recall to Davol, Inc., Sub. C. R. Bard, Inc. by calling 1-401-463-7000. |  | 1140 | 1,140 units | Nationwide. | Open, Classified |  |  | 1140 | Not_Computer
Z-1883-2009 | 52767 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Pyxis Anesthesia System 3500, Cardinal Health, Clinical Technologies and Services, San Diego, CA. |  |  | December 24, 2008 | 2008 | Cardinal Health Manufacturing LLC | Cardinal Health has sent out a Safety Alert Notification to inform its customers that the SR2b Software for the Pyxis Anesthesia System 3500 (PAS 3500) has received several customer reports of system "non-responsiveness" (lock ups) during the use of the product since its release in September 2008. | Software design | The "Safety Alert Notification" letter was mailed to all Pyxis Anesthesia System 3500 customers on December 17, 2008 by registered return receipt mail. The letter was addressed to the Director of Pharmacy at each facility. The mailing packages contain the "Safety Alert Notification" letter and the revised pages of the Pyxis Anesthesia System 3500 User Guide and Safety Manual. The "Safety Alert Notification" letter reinforced to customers how to maintain back-up processes to allow quick access if a system lock-up condition is experienced.The letter informed customers the situation will be corrected and a solution will be implemented. If you experience any system unresponsiveness during the use of the product, please report the issue to the Cardinal Health Technical Support Center at 1-800-727-4102. If you have any questions regarding this issue, please contact your local field support representative or the Customer Advocacy Department at Customer.Advocacy@cardinalhealth.com. |  | 250 | 250 | Nationwide and Germany. | Terminated | March 28, 2012 | 1190 | 250 | Software | Device Operation | Software update | Software Update
Z-2053-2009 | 51674 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Integra NeuroSciences Ventrix Ventricular Tunneling Pressure Monitoring Kit NL950-V and Integra NeuroSciences Ventrix Ventricular Tunneling Pressure Monitoring Kit with Ventricular Drainage System NL950-VC. |  |  | March 16, 2009 | 2009 | Integra Neuro Sciences | Due to a dimensional error, the diameter of the stylet provided with the device may be too large to insert into the lumen of the catheter tip. | Process design | An "Urgent-Product Recall Notification" letter dated March 23, 2009 was issued to consignees. The letter described the issue and instructed customers to examine their inventory immediately to determine whether they have any NL950V - Ventrix Ventricular Tunneling Pressure Monitoring Kits, or NL950VC - Ventrix Ventricular Tunneling Pressure Monitoring Kits with Ventricular Drainage System. Customers were further instructed not to use the above recalled product.Call Integra Customer Service at 1-800-654-2873, and then select 5 for 'Returns and Repairs' to obtain a Returned Material Authorization number, and return the product as directed.Direct questions about the recall to Integra NeuroSciences by calling 1-858-455-1115, extension185. |  | 274 | 274 kits | Worldwide Distribution -- US (states of AR, CA, OH, TX, CT, NY, LA, FL, NY, IL, and VA), Belgium and Malaysia. | Terminated | July 19, 2012 | 1221 | 274 | Not_Computer
Z-1946-2009 | 52232 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | ARJO Carendo Multi-Purpose Hygiene Chair; Model BIB2003-01, a non-AC powered patient lift; Manufactured by Arjo Hospital Equipment AB, Eslov, Sweden. |  |  | June 11, 2009 | 2009 | ArjoHuntleigh | Male patient genitals may be pinched between the seat and the frame of the patient lift chair when the Care Raiser seat is returned to its lowest position. | Device Design | ArjoHuntleigh sent Urgent Device Field Correction letters dated June 11, 2009, along with the Urgent Field Safety Notice ID 090526AHE dated May 26, 2009, via UPS Ground delivery. Customers were informed of the affected device and were requested to complete the response form with new contact information and fax it back to ArjoHuntleigh at 630-567-5012 or 630-925-7969. For further questions, contact ArjoHuntleigh at 1-800-323-1245. |  | 1987 | 1,987 hygiene chairs | Nationwide Distribution | Terminated | August 19, 2010 | 434 | 1987 | Not_Computer
Z-1930-2009 | 52000 | No | DCN | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | Dimension Flex Reagent Cartridge C-Reactive Protein Extended Range (RCRP). |  |  | April 15, 2008 | 2008 | Siemens Healthcare Diagnostics, Inc. | falsely elevated results in EDTA plasma samples | Labeling False and Misleading | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice on April 15, 2008 informing their customers that the EDTA plasma specimen claim had been rescinded. Current lots in the firm's inventory have been put on hold and reworked with an Alert Card rescinding the EDTA plasma claim. Subsequent lots use a revised instruction for use stating "Do not use EDTA plasma" in the Specimen Collection and Handling Section.For further questions, contact Siemens Healthcare Diagnostics at 1-302-631-6564. |  | 25760 | 25,760 cartons | Worldwide Distribution -- US, Austria, Bosnia, Czech Republic, Hungary, Slovakia, Serbia, Denmark, Estonia, Latvia, Lithuania, Norway, Sweden, Belgium, Cyprus, Canary Islands, France, Great Britain, Germany, Ireland, Italy, Netherlands, Poland, Russia, Portugal, Saudi Arabia, Spain, South Africa, Switzerland, Turkey, Yemen, Bahamas, Brazil, Canada, Mexico, Australia, Bangladesh, China, Hong Kong, Japan, Malaysia, Pakistan, and Taiwan. | Terminated | September 09, 2009 | 512 | 25760 | Not_Computer
Z-1949-2009 | 52672 | No | EJL | Dental bur. | Dental | Dental | 510(K) Exempt | BIOMET MICROFIXATION Trauma One SYSTEM 1.8X115 MM STOP DRILL, 1.8X115 MM WITH 26 MM STOP STAINLESS STEEL, NON-STERILE, BIOMET MICROFIXATION, INC. 1520 TRADEPORT DRIVE JACKSONVILLE, FL 32218 USA. |  |  | February 16, 2009 | 2009 | Biomet Microfixation, Inc. | Biomet Microfixation is recalling the 46-1004, Trauma One Drill, 18.x115 mm, 26 mm Stop Lot Numbers 619450 and JC1002. The drills may be color-banded blue instead of brown. The color is a visual aid in determining what system the drill goes with. The blue indicates the drill is for the 2.0mm system. The brown indicates the drill is to be used for our 2.3mm system. The drill is etched and labe | Process control | Biomet Microfixation, Inc. issued an Urgent Medical Device Recall Notice dated February 19, 2009 via certified mail with return receipt. The notice informed users of the affected device and instructed them to immediately discontinue use and return to the firm. Consignees were also instructed to complete and return an Inventory Reconciliation sheet via fax to 1-904-741-9425.For further information, contact Biomet Microfixation, Inc. at 1-800-874-7711 or 1-904-741-4400. |  | 79 | 79 | Nationwide Distribution -- Texas, Maryland, Florida, Washington DC, Michigan, Kansas Colorado and Missouri. | Terminated | November 05, 2009 | 262 | 79 | Not_Computer
Z-1887-2009 | 52757 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Agfa Impax R6.3 DB Server. |  |  | July 16, 2009 | 2009 | AGFA Corp. | Measurements are inaccurate on multiple workstations. | Software design | Agfa Healthcare issued an "URGENT FIELD SAFETY NOTICE" dated July 16, 2009. The notice informed of the problem and provided details to identify potentially compromised data that could result in measurement inaccuracies. A feedback form was provided to be completed and returned to the firm.For further information, contact Agfa Healthcare at 1-773-592-4769. |  | 1 | One unit | Nationwide Distribution -- IL | Terminated | April 28, 2011 | 651 | 1 | Not_Computer
Z-1912-2009 | 52694 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Broselow/Hinkle Pediatric Emergency System; |  |  | June 30, 2009 | 2009 | Vital Signs Inc | Not compatible: The luer connector on the Intraosseus needle provided does not fit securely on the luer connection of the IV Extension Set in the module. | Device Design | Vital Signs, Inc. issued an "Urgent Medical Device Correction" letter dated June 30, 2009 to all direct and indirect accounts by first class mail. Consignees were informed of the affected device, asked to complete a Confirmation form for return and to return unused product for replacement.For further questions, contact Vital Signs, Inc. at 1-973-790-1330 extension 404. |  | 23429 | 23,429 kits | Worldwide Distribution -- US, Canada, United Arab Emirate, Venezuela, Guam, Korea and United Kingdom. | Terminated | August 27, 2012 | 1154 | 23429 | Not_Computer
Z-2024-2009 | 52642 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 The Endovascular Company, Protege GPS 6mm x 80mm Self Expanding Peripheral Stent System , Rx Only, and SERP65-06-80-120. |  |  | June 10, 2009 | 2009 | Ev3, Inc. | ev3 Inc. is conducting a voluntary recall of specific lots of the Protege GPS Self-Expanding Stent System because labeling did not match the actual stent size contained in the package. The affected product was labeled as a 6mm x 80mm Protege GPS stent but actually contained an 8mm x 40mm stent. This situation can lead to vessel damage or inadequate coverage of the target lesion. | Labeling mix-ups | Consignees were emailed a letter with a subject heading of "Medical Device Recall". The e-mail was dated June 10, 2009. The letter informed the consignee of the problem and product involved. They requested that any affected devices remaining in inventory be identified and returned and to fill out and return the "Device Recall Field Action Reconciliation Form". |  | 2 | 2 | International Distribution Only -- Brazil, Uruguay, &amp; Malaysia. | Terminated | December 11, 2011 | 914 | 13 | Not_Computer
Z-2047-2009 | 52927 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Pedi-Cap CO2 Detector, 1 kg - 15 kg, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. Assembled in Mexico. Sold in packages of 6 (PediCap6) or 24 (PediCap). |  |  | August 14, 2009 | 2009 | Covidien Limited | CO2 detector may increase resistance to air flow in an intubated patient. | Component change control | Covidien issued an Urgent Medical Device Recall letter dated August 14, 2009 informing Consignees of the affected devices. Users were told to stop use and arrange for return of all affected product. For further information, contact Covidien at 1-800-635-5267. |  | 276390 | 276,390 units | Worldwide Distribution -- US (FL, ID, MD, NC, VA and PA including Puerto Rico), Australia, Belgium, Canada, Japan, Singapore. | Terminated | June 01, 2012 | 1022 | 276390 | Not_Computer
Z-2042-2009 | 52885 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Biomet 8.0MM cannulated reamer for the Aimer system, Biomet Sports Medicine, Ontario, CA; REF 909618. |  |  | August 06, 2009 | 2009 | Biomet, Inc. | These surgical instruments are etched with incorrect measurement markings. | Nonconforming Material/Component | Consignees were notified by letter dated 8/6/09 to remove the instrument from use and to return the recalled instrument to Biomet. Implanting surgeons were notified by letter dated 8/6/09, advised of the problem and it was suggested that the patients be monitored for instability if deemed necessary by the surgeon. |  | 2 | 2 | Nationwide Distribution -- Including the state of Texas. | Terminated | October 06, 2010 | 426 | 2 | Not_Computer
Z-2041-2009 | 52804 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ST ULTRAPOWER DIAMOND WHEEL (RED) 25.4mm, Catalog Number 00702109200, STERILE, Single-use. Made is USA. |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec is recalling Ultrapower Bur, Diamond Wheel for possibility of a packaging breach in the sterile barrier. | Package design/selection | Firm notified international and domestic consignees by an Urgent Medical Device Recall letter on July 31, 2009. Customers were asked to check their inventory; complete, sign, and return the enclosed Reply Form; and return the devices for credit. Questions should be directed to Customer Service at 1-800-237-0169. |  | 2502 | 2502 | Worldwide Distribution | Terminated | May 11, 2010 | 284 | 2502 | Not_Computer
Z-2049-2009 | 52963 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Penumbra Neuron 5F Select Catheter, Neuro Deliver/Intracranial Access System Catheter, manufactured by Penumbra, Inc., Alameda, CA. |  |  | July 02, 2009 | 2009 | Penumbra Inc. | Leakage-- Due to a manufacturing error, some products may have a fold in the innermost layer, causing disruption in the layering process and may result in blow holds or exposed wire braids, which may result in further damage. | Process control | Penumbra, Inc. initiated Recall activities on May 29, 2009. The Firm notified consignees via telephone and email of the problem and asked them to return the product to Penumbra.For further information, contact Penumbra, Inc. at 1-510-748-3223. |  | 9 | 9 units | Worldwide Distribution -- US (IA, NJ, NY and VA), and Canada. | Terminated | January 05, 2011 | 552 | 9 | Not_Computer
Z-1877-2009 | 52659 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Boston Scientific, LeVeen Needle Electrode, 25 cm, 5.0 cm, sterile; REF M001262170. For tissue coagulation. |  |  | July 06, 2009 | 2009 | Boston Scientific Corp | The tines may detach from the core wire during use, which may result in bleeding or soft tissue trauma requiring medical intervention. | Process control | Boston Scientific notified consignees by letter dated 7/6/09 and instructed to cease using and to return the products, and to complete and return a reply form. |  | 347 | 347 | Nationwide. | Terminated | October 08, 2010 | 459 | 3121 | Not_Computer
Z-1911-2009 | 52652 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Adapt AO, Akreos Advanced Optics Aspheric Lens. Bausch &amp; Lomb Incorporated, Rochester, NY. Manufactured by Bausch &amp; Lomb Incorporated, Clearwater, FL. |  |  | December 08, 2008 | 2008 | Bausch &amp; Lomb Inc | Affected lenses were inadvertently shipped to Bausch &amp; Lomb's European Logistics Center (ELC) prior to release authorization. Initial testing indicated the product did not meet release criteria for endotoxin limits. | Release of Material/Component prior to receiving test results | All consignees were notified via phone, telefax and letter. Products are to be returned to Bausch &amp; Lomb's European Logistics Center (ELC) for destruction. Direct questions to Bausch and Lomb, Inc. by calling 1-585-338-6612. |  | 158 | 158 | Foreign Distribution only -- Algeria, Belgium, Bulgaria, Croatia, Germany, Hungary, Ireland, Israel,Italy, Mexico, Portugal,Spain, Thailand andUnited Kingdom. | Terminated | November 05, 2009 | 332 | 158 | Not_Computer
Z-1999-2009 | 52667 | No | JTC | Microtiter diluting and dispensing device. | Microbiology | Microbiology | 510(K) Exempt | 300 Microliter CO-RE Tips for use with Microlab STAR, manufactured by Hamilton Company, Reno, NV. |  |  | April 09, 2009 | 2009 | Hamilton Co | Product Defect-- some tips from the mold cavity 28 are too short and do not always aspirate or dispense liquid correctly. | Process control | The firm initiated the recall on April 10, 2009, and issued notification to consignees via certified mail and e-mail. Customers were instructed to return the affected product to the Hamilton Company. A fax-back form was supplied with each notification. Direct questions about the recall to the Hamilton Company by calling 1-775-858-3000, extension 236. |  | 4719000 | 20 cases of Part 235950 and 27 cases of Part 235902. | Worldwide Distribution -- US (states of DE, NJ, PA, MA and CA), New Zealand and Japan. | Terminated | January 04, 2010 | 270 | 4719000 | Not_Computer
Z-2052-2009 | 52668 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Neuroform 3 Microdelivery Stent System, Model Number M003E3400300, Catalog model number and product upn number M00490310, manufactured by Boston Scientific Corporation, Fremont, CA. |  |  | May 08, 2009 | 2009 | Boston Scientific Corporation | Mislabeling-- Product was placed in a carton, which has printed on it the name of another product. A correct product label was placed on the carton. | Employee error | The firm issued an Urgent Product Field Action Notice to all listed customers, along with a reply verification form, requesting the return of affected products. |  | 8 | 8 units | International Distribution Only -- Countries of Germany, Great Britain, France, Ireland, and Italy. | Terminated | October 21, 2009 | 166 | 8 | Not_Computer
Z-1993-2009 | 52423 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Tibial Tray, component of Ortho Development Balanced Knee System, Nonporous, Sterile R, Size 3, REF 162-1300, Ortho Development, Draper, Utah 84020. |  |  | May 29, 2009 | 2009 | Ortho Development Corporation | Erroneous lot number on 3 of 28 units of Balanced Knee System. Three units from Lot Number 0055245 were mixed in with the recalled lot (0054891). | Mixed-up of materials/components | Consignees were notified via e-mail on May 29, 2009 and June 5, 2009 requesting return of units. Direct questions about the recall to the Ortho Development Corporation by calling 1-801-619-3450. |  | 13 | 13 units | Worldwide -- US (states of FL and TX) and Japan. | Terminated | May 11, 2010 | 347 | 13 | Not_Computer
Z-1870-2009 | 52605 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | GE Healthcare, Uni-Circuit Kit, Adult, Disposable, 1.8 m/72 in, REF M1089028, M1087972, M1087975, M1126445, M1127508, M1127509, M1129780, |  |  | June 05, 2009 | 2009 | GE Medical Systems, LLC | Breathing circuit inner tube may separate from the rigid connectors under certain conditions. A separation of the inner tube from the connector may prevent adequate gas flow (O2, N2O and anesthetic agents) to the patient. A separation can also cause CO2 re-breathing. This could lead to hypoxia and/or hypercarbia and inadequate anesthesia delivery. | Process control | GE Healthcare sent consignees 6/5/09 an "Urgent Medical Device Correction" letter dated June 5, 2009. The letter was addressed to Health Care Administrator/Risk Manager; Chief of Anesthesia; Director of Clinical Engineering; Risk Manager. The letter described the Safety Issue, Affected Product Details, Safety Instructions (included "Recall Confirmation" form fax to 973-790-4150), Product Correction and Contact Information. |  | 85097 | 85097 | Worldwide distribution: USA, AUSTRALIA, DENMARK, FRANCE, ESTONIA, GERMANY, ITALY, JORDAN, NETHERLANDS, NORWAY, SWEDEN, and UNITED ARAB EMIRATES. | Terminated | January 16, 2014 | 1686 | 85097 | Hardware | Treatment/Delivery/Therapy | Correction | Repair
Z-2054-2009 | 52434 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | EUB-6500/HI VISION 6500 Diagnostic Ultrasound Scanners, DICOM Option EZU-FC5, Software Versions: V01-03, V01-05, V02-04, V03-02, V04-02, V04-03, V04-04, V05-03, V05-05A, V05-07A and V05-30A. |  |  | May 12, 2008 | 2008 | Hitachi Medical Systems America Inc | The firm's parent company discovered a DICOM UID (Universal Identifier) code mix-up which can occur due to a software error. This error can cause the images from two (2) different patient scans to become intermingled and consequently, labeled incorrectly. | Software design | HMSA issued a 'DEVICE CORRECTION' notification to their customers on 5/12/2008. The notification describes the potential data transmission problem (potential to archive studies done on different patients into the same patient folder) and the conditions under which this problem can occur. The notification cautions the user/customer to verify all images stored on the PACS (Picture Archiving and Communications System) in order to ensure that they do not contain images from a different patient. The DEVICE CORRECTION letter also states that the firm will be installing updated software on all affected systems beginning in June, 2008, in order to address this issue. It also states that a Hitachi Service Rep. will be contacting each customer and making arrangements to set up a convenient time to come out and perform a software upgrade. Lastly, the letter provides a toll-free telephone # to the firm's 'Customer Service Center: 1-800-800-4925, which the customer can use to obtain additional information, if desired. |  | 10 | 10 units | The recalled device system units were distributed to customers located in the following states within the U.S ony.: AR, CA, and TX. | Terminated | July 20, 2011 | 1164 | 16 | Software | Output/Calculation | Software update | Software Update
Z-2044-2009 | 52814 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Engenex Cover Strip; Product Code 446471; |  |  | July 29, 2009 | 2009 | ConvaTec | Labeling: The expiration date does not appear on the product label. | Labeling design | Urgent Product Recall letters dated 7/29/09, were hand delivered to all direct customers by ConvaTec Sales Reps. Customers were asked to return any unused products to their supplier. The letter also stated that a credit would be issued for the product. Customers are to contact Karen Krenta at the firm's Customer Interaction Center, 1-800-422-8811. |  | 128 | 128 units | Nationwide Distribution | Terminated | October 02, 2009 | 65 | 456 | Not_Computer
Z-2051-2009 | 52811 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | REF IA-2000-S, LightWave Suction Ablator, STERILE, Rx ONLY, MADE IN USA. |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec has determined there is a possibility that the device may not shut off if the ablate button is released, which may result in users/patients suffering burns. | Device Design | ConMed Linvatec is notifying their direct consignees by an Urgent Medical Device Recall letter, requesting that customers check their inventory for the affected products; complete, sign, and return the Reply Form; and return of any affected lot numbers of the product. The product is to be returned to ConMed Linvatec. Questions should be directed to Customer Service at 1-800-237-0169. |  | 28074 | 28,074 units | Worldwide Distribution | Terminated | February 25, 2011 | 574 | 28074 | Not_Computer
Z-2000-2009 | 51896 | No | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | Roche/Hitachi Preciset TDM I, Roche Diagnostics, Indianapolis, IN; Catalog No. 03375790 190. |  |  | April 22, 2009 | 2009 | Roche Diagnostics Corp. | The bottle caps may be more difficult to remove than usual, and minor injuries may be caused by the application of the force required to open the bottles. | Packaging process control | Roche Diagnostics Corp. notified Consignees via letter dated April 22, 2009 and instructed to not use excessive force when opening bottles and contact the firm for product replacement instead. A Fax Back Form was included for customers to return via fax to 1-888-345-0480.For further information, contact your Roche Diagnostics representative or call 1-800-428-5074. |  | 3970 | 3,970 bottles of both products | Nationwide Distribution | Terminated | December 07, 2009 | 229 | 3970 | Not_Computer
Z-1963-2009 | 52107 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BD Diagnostics, BD GeneOhm MRSA Assay, REF: 441242 (200 Reaction Kit) and 441244 (48 Reaction Kit). |  |  | November 11, 2008 | 2008 | BD Diagnostics (GeneOhm Sciences, Inc) | BD Diagnostics GeneOhm (BDDG) has determined that the Cepheid SmartCycler reaction tubes packaged in the BD GeneOhm MRSA test kits identified above may be defective. | Other | BD Diagnostics GeneOhm began contacting consignees by phone November 17, 2008 followed by a letter dated November 21, 2008 instructing discontinued use of the affected lots. Consignees were also asked to complete and return the enclosed Field Action Verification form.To arrange for product replacement or for further information, contact BD Diagnostics Technical Service at 1-888-436-3646 extension 2. |  | 114 | 114 kits | Worldwide Distribution -- US (CA, IL, IN, KT, MD, ME, MN, MO, MT, NC, NJ, NY, OK, PA, TN, TX, UT, VA, and WI), Belgium and Japan. | Terminated | November 29, 2011 | 1113 | 114 | Not_Computer
Z-1987-2009 | 52662 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Apex Hip System Bipolar Head Bipolar Head, 28 mm ID, 59 mm OD |  |  | July 10, 2009 | 2009 | Omni Life Science | The APEX Hip System Bipolar Head implant may not properly capture the femoral head assembly of the bipolar head onto the femoral head. | Device Design | Omni Life Science notified distributors and implanting surgeons by an Urgent, APEX Hip System Bipolar Head Voluntary Recall letter dated June 29, 2009. The method of notification is both e-mail and FedEx overnight. The letter requires a signature acknowledging receiving the letter and the letter is to be faxed back to Omni Life Science. The Distributors/Sale Representatives are instructed to fill in the "Quantity Returned" on the recall letter and fax back to Omni. They are to ship all existing Apex Hip System Bipolar Head inventory by 07/10/2009 to the directed address. |  | 17 | 17 units | Nationwide Distribution -- Including states of AZ, CA, CO, FL, IL, MI, NJ, OK, UT, VA, and WI. | Terminated | June 28, 2012 | 1084 | 295 | Not_Computer
Z-2040-2009 | 52114 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Leckey 4-Point Pelvic Harness Belts, catalog 081326990; James Leckey Design Ltd., Kilwee Business Park, Upper Dunmurry Lane, Belfast, United Kingdom; Distributed by Patterson Medical; sold as replacement belts and as a component of the following adjustable mechanical chairs: |  |  | April 17, 2009 | 2009 | Patterson Medical Holdings, Inc. | The plastic buckles on the hip belts of the 4-Point Pelvic Harness may break. | Nonconforming Material/Component | Sammons Preston, the U.S. agent for James Leckey Design Ltd., issued recall letters dated 4/17/09 to all of their customers who purchased the Leckey 4-Point Pelvic Harness, used in the Leckey Contoured Advance Seat, Leckey Early Sitting System and the Leckey Squiggles Saddle Seat, informing them that the manufacturer, James Leckey Design Ltd., is recalling the harness because the pelvic belt buckle may not lock securely and/or break. A copy of the Urgent Field Safety Notice dated 1/28/09 issued by Leckey in the United Kingdom was included with the Sammons letter. The customers were advised to be vigilant and check the hip belts daily to ensure that the user is secure in their seating system until the replacement hip belt is received. The accounts were requested to contact Carla Dreher at 1-800-288-3693, ext. 8724 for instructions on how to receive a replacement hip belt. |  | 226 | 226 belts | Nationwide and Canada | Terminated | May 13, 2011 | 756 | 226 | Not_Computer
Z-2033-2009 | 52383 | No | ECX | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Empty D aluminum oxygen cylinders packaged in shipping containers labeled as: Product #0022-58 P, Desc CYL 02 ALUM 400 LITER, Catalog #31-10-2014. The brand and firm name on the label affixed to the cylinder is Chemetron, Allied Healthcare Products Inc., St. Louis, MO. |  |  | May 15, 2009 | 2009 | Allied Healthcare Products Inc | Medical valves do not meet quality specifications which may cause the valve to fail and result in uncontrolled release of oxygen when the valve is in the open position. | Nonconforming Material/Component | The recalling firm first notified their customers via a letter dated 5/15/09 issued regular mail explaining the reason for recall and requesting the cylinders with specific post valves be returned. A second letter dated 6/1/09 was issued via regular mail which provided specific shipping dates that each customer received the suspect product. Neither of these two letters requested subrecall. A third letter dated 6/17/09, which included copies of the first two letters, via certified mail. This letter informed the customer the recall was to be conducted to the user level and requested their customer conduct a subrecall. Requested to remove the cylinder from service and call Allied Healthcare Products for return and a replacement at 1-800-268-1661. |  | 177 | 177 cylinders | Class II Recall Worldwide Distribution including states of NY, NC, KY, PA, WI, NH, NJ, CA, CT, FL and Singapore. | Terminated | April 23, 2012 | 1074 | 288 | Not_Computer
Z-1937-2009 | 52352 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | EASYMax N. Self-Monitoring Blood Glucose System, distributed in cardboard carton kits. Each kit contains the following components: Blood Glucose Meter; Test Strip x 10 ct.; Normal Control Solution; Carrying Case; Alkaline Battery x 2 ct.; Lancets x 5 ct.; Lancing Device; AST Lancing Device Cap; User's Manual; Test Strip Instructions and Control Solution Instructions. |  |  | May 26, 2009 | 2009 | Oak Tree International Holdings, LLC, Inc. | Blood glucose meters, which were set to the mmol/L. unit of measure, rather than the accepted/labeled mg./dL unit of measure, were distributed. | Process control | EPS Technology issued a certified notification letter to the sole distributor/customer: GEMCO, Coconut Creek, FL dated May 26, 2009. The letter describes the potential preset error and its possible consequences to the user/customer. The letter goes on to inform the distributor that EPS Technology Corp. will be sending out letters to all affected user/customers during a direct mailing. If the error is found, the user/customer can either: 1) have the firm's Customer Service Dept. instruct them as to how to fix the unit, or 2) request that a replacement meter be sent to them.For further information, contact EPS Technology Customer Service at 1-866-994-3345. |  | 5396 | 5,396 units | Nationwide Distribution -- AL, AR, AZ, CA, CT, FL, GA, IL, IN, KS, KY, LA, MD, ME, MI, MS, NC, NE, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, VA, WA, WY and HI. | Terminated | September 27, 2012 | 1220 | 5396 | Not_Computer
Z-1997-2009 | 52614 | No | KKO | Teething ring. | Dental | Dental | 510(K) Exempt | MDI Hybrid Implant, O-Ball abutment, sizes 2.9mm x 13mm and 2.9mm x 15mm, endosseous dental implant. |  |  | July 07, 2009 | 2009 | IMTEC Corporation | Size printed on product packaging does not match size of enclosed product. | Other | A notification letter dated July 7, 2009 was issued to customers. The letter described the issue and instructed customers do not use the affected product and call IMTEC, a 3M Company, at 1-800-879-9799, extension 240 for assistance. Customers were further instructed to complete the enclosed Recall Response Fax-Back Form with information on their inventory (whether or not the recalled product is in their stock). The firm's representative will follow-up with customers if no response within three business days of receiving the notification letter. |  | 605 | 605 units | Worldwide -- US, Azerbaijan, Guatemala, Turkey, and Germany. | Terminated | December 02, 2009 | 148 | 605 | Not_Computer
Z-2073-2009 | 53129 | No | LKK |  |  | General Hospital | PMA | Medtronic Indura 1P Intrathecal Catheter, model 8709SC, Method of Sterilization: Ethylene Oxide. Medtronic, Inc., Minneapolis, MN 55432-5604 USA. The implanted infusion system components consist of a Medtronic pump and a model 8709SC Catheter. The catheter connects to the pump at the catheter port. Contents: 81.4-cm Catheter with length markers and guide wire, 15 T-guage introducer needle, Sutureless pump connector with attached 7.6-cm catheter and connector pin, Transparent strain-relief sleeves, Anchors. |  |  | August 27, 2009 | 2009 | Medtronic Neuromodulation | Current labeling for Medtronic Sutureless Connector (SC) catheters and revision kits (hereafter referred to as SC catheters ) incorrectly states that SC catheters are compatible with Medtronic IsoMedÂż constant-flow infusion pumps. SC catheters are compatible with Medtronic SynchroMedÂż II and SynchroMed EL pumps. Medtronic has determined that SC catheters are not compatible with IsoMed pumps. | Device Design | A "Medical Device Correction Notification" letter was issued on August 14, 2009 to implanting and managing physicians. The letter described the nature of the issue, scope of affected product, potential severity of the issue and provided recommendations. Physicians were instructed that Sutureless Connector (SC) catheters must not be used with IsoMed pumps for future revisions and implants. A copy of the letter was also sent to the Risk Manager at each hospital associated with physicians using SC catheters. Healthcare professionals were asked to complete and return the enclosed reply card.Direct questions to your Medtronic field representative or contact Medtronic Neuromodulation Technical Services at 1-800-707-0933. Patient Management information is also available at http://www.professional.medtronic.com under the Heading, "Advisories". |  | 20000 | 20,000 of all products | Worldwide Distribution -- US (including Washington D.C. and Puerto Rico), Algeria, Aruba, Australia, Austria, Belgium, Canada, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Luxembourg, Malta, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, Norway, Poland, Portugal, Reunion, Romania, San Marino, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom and Vatican City. | Terminated | December 12, 2011 | 837 | 20000 | Not_Computer
Z-2028-2009 | 52810 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Diamondback 360 Orbital Atherectomy Device, 1.50 LT Classic, Model #DB-150L, CSI Part # 70025-02 |  |  | April 16, 2009 | 2009 | Cardiovascular Systems, Inc. | CSI is correcting (8) packages of Diamondback 360 due to a missing "Use By Date" on the device labeling. The labelling should have been a 12 month shelf life (2010-02) rather than no information for shelf life. The units affected by this correction are sterilized products, and CSI believes that there is no immediate risk to health associated with these distributed products recognizing the valid s | Incorrect or no expiration date | Cardiovascular System, Inc. issued an "URGENT: Medical Device Correction" letter dated April 16, 2009. The letter described the product and the problem. The letter informed the consignees that the "Device" re-labeling would be performed by company representatives. Devices will not be returned.For further information, contact Cardiovascular System, Inc. 1-877-274-0901. |  | 8 | 8 | Nationwide Distribution -- CA, GA, MI, NY, NY and TN. | Terminated | December 10, 2011 | 968 | 8 | Not_Computer
Z-1938-2009 | 51281 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Mirror assembly to magnetic resonance device (1.5T Atlas SPEEDER Head-Neck Coil) Coil Part #Q7000021, Quality Electrodynamics, Mayfield Village, OH 44143. |  |  | December 29, 2008 | 2008 | Quality Electrodynamics Llc | The firm received a complaint indicating that the mirror is high risk and could drop on a patient. | Device Design | Quality Electrodynamics, LLC notified Customers via email beginning December 29, 2008 requesting the affected device be returned to the firm for improvement.For further information, contact Quality Electrodynamics, LLC at 1-440-484-2228. |  | 5 | 5 units | Nationwide Distribution: IL and OH. | Terminated | April 25, 2012 | 1213 | 5 | Not_Computer
Z-2011-2009 | 51726 | No | ECX | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Post Medical Valves VO870A-Y and VO870-A-Y-T PSI 2015, PSI 2015, CGA 870, Inlet .750 STDS, Gas services (Oxygen), PSI 2015, Aluminum Cylinders, InterMed Gas Products LLC Industrial Division, 2056 Vista Parkway West Palm Beach, FL 33411. |  |  | March 27, 2009 | 2009 | Intermed Gas Products Corporation | Firm received reports from customers that valves malfunctioned when installation and/or filling process. Customers forwarded pictures of the defect where the top portion of the packing nut broke off from the threaded portion of the net; which may cause the valve to fail and result in uncontrolled release of oxygen when the valve is in the open position. | Component design/selection | Intermed Gas Products Corporation notified customers of the affected product by phone beginning March 27, 2009 and followed up with a formal "Product Recall Notice" letter dated March 31, 2009. All affected product should be located and returned to the firm for replacement.For further information, contact Intermed Gas Products at 1-561-483-1364. |  | 8767 | 8,767 | Nationwide Distribution -- FL, SC, TX, MO, TN, NC, NY, GA, MN, AZ. | Terminated | May 17, 2011 | 781 | 8767 | Not_Computer
Z-2039-2009 | 52700 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Advanced Bionics HiRes 90K Cochlear Implant, Model: CI-1400-01 and CI-1400-02H, Implant Magnet Case. |  |  | June 10, 2009 | 2009 | Advanced Bionics LLC | The recall was initiated after Advanced Bionics became aware of fourteen (14) devices manufactured between May 27, 2005 and June 30, 2005 that are confirmed to have or potentially have an issue related to the internal magnet. These devices are were made with an internal magnet where the metal case surrounding the magnet was not properly welded. Patients with these devices may experience a decreas | Process control | Advanced Bionics, LLC issued an "Important Notification" dated June 10, 2009 via Fed Ex to Consignees in the US and Canada. Outside of the US, the notification letters will be distributed in such a manner that delivery confirmation can be obtained. There were three type of notification letters distributed: one for Patients, one for Surgeons and one for Clinicians. Letters described the affected device prompting users on what to do next. An acknowledgement form was enclosed and should be completed and sent via fax to 1-661-362-7621.For further information, contact Advanced Bionics at 1-877-829-0026. |  | 18 | 18 | Worldwide Distribution -- US, Canada, France and South Korea. | Terminated | March 26, 2012 | 1020 | 18 | Not_Computer
Z-2211-2009 | 53124 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Integrated ECG Cable are supplied as part of the start-up kits provided with some Spacelabs' monitors. The kit PNs are: 718-0017-10, 718-0017-11, 718-0017-12, and 718-0017-13. The cables are also offered as replacement parts under Spacelabs PNs: 700-0008-22, 700-0008-23, 700-0008-24, and 700-0008-25. |  |  | July 23, 2009 | 2009 | Spacelabs Healthcare, Incorporated | The Spacelabs Medical Integrated ECG Cables, PN: 700-0008-22, 700-0008-23, 700-0008-24, and 700-0008-25 will not withstand the high voltage of a defibrillation shock. | Process control | Spacelabs Healthcare issued an "Urgent: Medical Device Recall" dated July 20, 2009. Consignees are being sent replacement ECG cables for each item ordered and are requested to replace the affected cables. Also, they are instructed to destroy and dispose of the recalled cables according to their own protocol or local regulations.For further information, contact Spacelabs Healthcare Technical Support at 1-800-522-7025, option 2 (US) or +86-512-87171888, extension 304 (outside US). |  | 1395 | 1,395 | International Distribution -- Bangladesh, Turkey, Trinidad and Tobago, Romania, Guatemala, Algeria, Netherlands, Egypt, Chile, Venezuela, Spain, Colombia, India, the United Kingdom, Singapore, France and Ecuador. | Terminated | July 21, 2011 | 728 | 1395 | Not_Computer
Z-2117-2009 | 52713 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Cancellous Bone Screw; |  |  | June 26, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Raw material supplied by a third party used to manufacture various medical devices does not comply with metallurgical requirements outlines in ASTM standard for titanium surgical implants. | Nonconforming Material/Component | Stryker Orthopaedics issued "Urgent Product Recall" letters via Federal Express priority overnight beginning June 26, 2009 to Branches, Hospital Risk Management and Chief of Orthopaedics and Surgeons. All users are requested to examine their inventory, retrieve all affected product and return for reconciliation to Stryker Orthopaedics. Users are also requested to complete the enclosed Product Accountability Form and fax with necessary information to the firm at 1-201-831-6069.For further information, contact Stryker Orthopaedics at 1-201-831-5970. |  | 51113 | 51,113 total - all sizes. | Worldwide Distribution | Terminated | July 27, 2010 | 396 | 51137 | Not_Computer
Z-1810-2009 | 52643 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Treatment Delivery System, a subsystem of the CyberKnife Robotic Radiosurgery System, manufactured by Accuray Inc., Sunnyvale, CA. |  |  | June 22, 2009 | 2009 | Accuray Inc | Targeting accuracy out of specification, Error alert does not render the system down, which may result in mistreatment in the wrong area. | Software design | Accuray, Inc. issued an "Urgent Device Correction" notice dated June 22, 2009. The consignees were informed of the defect including methods for detecting the error. The firm is in the process of developing a software patch to address the issue.For further information, contact Accuray, Inc. Customer Support at +1-877-668-8667 (US) or +1-408-716-4700 (outside US) or customersupport@accuray.com. |  | 61 | 61 units | Worldwide Distribution -- US including states of MA, FL, AL, TN, DE, CA, NJ, OK, CO, NC, IL, PA, MD, DC, MS, CN, NV, RI, SC, UT, VA, MI, NE, MO, TX and AL and countries of Canada, Russia, UK, Ukraine, Saudi Arabia, Switzerland, Turkey, Taiwan, India, China, Korea and Thailand. | Terminated | January 05, 2011 | 562 | 61 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2111-2009 | 50802 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | I-Flow Corporation Easy Pump LT 60-24 (60 ml x 2 ml/hr), REF: 04434366, assembled in Mexico. |  |  | February 11, 2008 | 2008 | I-Flow Corporation | I-Flow Corporation received two complaints involving Easypump LT 60-24 (Lot 612996) containing a 100 ml x 2 ml/hr pump in the sterile package instead of a 60 ml x 2 ml/hr pump (Illustration was included to show consignee what both the correct and incorrect part looks like). Lot 612996 was manufactured on January 6, 2006. As of February 7, 2008, there have been no further complaints received on t | Labeling mix-ups | I-Flow Corp. issued an "Important Recall Information" letter dated February 7, 2008 informing Consignees to quarantine and return any remaining inventory and notify any end users of the affected lot.For further information, contact I-Flow Corp. at 1-949-206-2653. |  | 4450 | 4,450 Distributed | Nationwide Distribution to B. Braun, Spain, who then further distributed to Italy, Germany, Netherland, Latvia, Spain, Portugal and Singapore. | Terminated | February 04, 2010 | 724 | 4450 | Not_Computer
Z-2030-2009 | 53035 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Medtronic Physio-Control LIFEPAK CR Plus Defibrillator/Monitor. |  |  | July 31, 2009 | 2009 | Physio Control, Inc. | The K1 relay on the Analog printed circuit board assembly may short due to moisture ingress in a highly humid environment. K1 shorting may affect the ECG amplitude causing the Shock Advisory System(SAS) to not analyze the ECG rhythm correctly thus preventing the device from delivering therapy. | Nonconforming Material/Component | Physio-Control, Inc. notified Customers by phone beginning August 18, 2009. In addition, the recalling firm Issued an "Urgent Medical Device Recall" letter dated August 2009 informing Customers of the affected product and informing them that a Physio-Control representative will call and arrange for immediate device replacement.For further information, contact Physio-Control Customer Care at 1-800-442-1142. |  | 34 | 34 devices (2 domestically &amp; 32 internationally) | Worldwide Distribution -- US (DC and MN) and Japan. | Terminated | March 15, 2013 | 1323 | 34 | Hardware | Output/Calculation | Replace device | Remove or Replace
Z-2179-2009 | 50678 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Access Systems, Inc., Gauderer Genie* II PEG (Percutaneous Endoscopic Gastrostomy) Safety System Ponsky "Pull" Placement Technique REF 000373, 20F, STERILE. |  |  | December 04, 2008 | 2008 | C.R. Bard Access Systems, Inc | Snare can detach in patient. The Polypectomy Snares and the Entake Peg System (which contain these snares) are being recalled because the looped portion of the snare can become detached and drop into the stomach or other internal canal during use, necessitating retrieval of the remaining piece. | Nonconforming Material/Component | Bard contacted all 1,883 worldwide customers who received the distributed suspect 467 product code lots, comprising 177, 614 units that were either not expired or were still available in the field. These will be returned by the customer to Bard Access Systems. 1-800-290-1689. |  | 177614 | 177,614 units | Worldwide, including USA, Belgium, Canada, Australia, and Japan. | Terminated | November 09, 2011 | 1070 | 177614 | Not_Computer
Z-2027-2009 | 52617 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Smith &amp; Nephew Orthopaedics PIP (proximal Interphalangeal) FIXATION HINGE KIT, REF 101638, for external fixation of finger joint. |  |  | January 23, 2009 | 2009 | Smith &amp; Nephew Inc | Hinge failure: The product is cracking or breaking during use. | Nonconforming Material/Component | The Smith &amp; Nephew Sales Representatives were notified of the problem and the recall via certified letter and e-mail on 01/23/2009. A second notice was sent to the hospitals on 06/06/2009. All affected Smith &amp; Nephew International Distributors were notified via email on 01/23/2009. They were instructed to immediately notify their accounts concerning the recall. Smith &amp; Nephew c/o Return Goods, 4287 Delp St, Memphis, TN 38118 |  | 7089 | 7,089 units | Worldwide Distribution: USA, Turkey, Denmark, Germany, Malaysia, Hong Kong, Korea, Switzerland, Japan, Singapore, Canada, New Zealand, Australia, and Italy. | Terminated | July 29, 2013 | 1648 | 7089 | Not_Computer
Z-2109-2009 | 53179 | No | NAJ | Vascular embolization device. | Cardiovascular | Obstetrics/Gynecology | 510(k) | Boston Scientific, Contour SE Microspheres, for tumor treatment. |  |  | August 17, 2009 | 2009 | Boston Scientific Corporation | Compromised sterility: Boston Scientific is initiating a Recall Removal for twenty (20) UPN codes and 638 associated lots/batches of its Contour SE Microspheres product. All affected product is packaged within a vial. Boston Scientific has determined that, for affected products, the sterile barrier in the packaging (pouch) that contains the vial may be breached. The breach of sterility could l | Process design | Boston Scientific sent consignees an "Urgent Medical Device Recall Removal" letter dated August 28, 2009. The letter was addressed to Risk Manger / Field Action Contact. The letter described the problem and the product and advised the consignee to cease immediately the distribution and use of the product. Requested consignees to return affected product to Boston Scientific. The letter enclosed the Recall Instructions, Reply Verification Tracking Form and Field Action Return Shipping Label. |  | 45724 | 45,724 (13,048 US, 45,724 OUS) | Worldwide distribution: USA, United Arab Emir., Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, China, Colombia, Costa Rica, Czech Republic, Germany, Denmark, Algeria, Ecuador, Egypt, Spain, Finland, France, Great Britain, Greece, Hong Kong, Hungary, Ireland, Israel, Italy, Jordan, South Korea, Kuwait, Lebanon, Latvia, Libya, Mexico, Namibia, Nicaragua, Netherlands, Norway, New Zealand, Oman, Panama, Poland, Puerto Rico, Portugal, Russian Fed., Saudi Arabia, Sweden, Slovakia, Tunisia, Turkey, Taiwan, Ukraine, and South Africa. | Terminated | December 18, 2011 | 853 | 45724 | Not_Computer
Z-2220-2009 | 48960 | No | MTG |  |  | Immunology | PMA | Beckman Coulter Access Hybritech free PSA Calibrators, REF 37215. |  |  | January 08, 2008 | 2008 | Beckman Coulter Inc | Beckman Coulter incorrectly shipped Access Hybritech free PSA calibrators containing dual calibration cards (Hybritech and WHO) to your lab. The WHO calibration for free PSA is marketed without a PMA. | Other | Beckman Coulter contacted customers by phone beginning January 11, 2009 and sending a Product Corrective Action letter informing them to discontinue use and discard the referenced product immediately.For further information, contact Customer Technical Support at 1-800-854-3633 or contact your local Beckman Coulter Representative. |  | 22 | 22 Kits | Nationwide Distribution -- OH, CA, TX, OR, MD, MO, FL and PA. | Terminated | November 29, 2011 | 1421 | 22 | Not_Computer
Z-1931-2009 | 52200 | No | DSQ |  |  | Cardiovascular | PMA | iPulse Circulatory Support System |  |  | April 23, 2009 | 2009 | Abiomed, Inc. | iPulse Console SC1035 displayed a "Low Pressure / Low Flow" alarm and stopped pumping during patient transport in battery operation. | Device Design | An "Urgent Voluntary Medical Device Field Correction Notice" was issued April 8, 2009 to affected customers. The customers were notified that the iPulse Console SC1032 displayed a "Low Pressure/ Low Flow" alarm and stopped pumping during patient transport in battery operation. Instructions were provided to users in case they encounter this failure mode. ABIOMED Field Service Engineers will contact customers to schedule a time to visit site and correct affected consoles. Please contact ABIOMED's Field Service at 1.800.554.8666, if you have any questions. |  | 79 | 79 | Domestic: MA, MI, SC, NY, NJ, OH, CA, PA, AR, MI, MD, CT, IL, WI, TX, NJ, TN, IN, MI, DE, FL, IA, WA, KY, and WVForeign: Germany, Greece, Italy, Latvia, Spain, and Turkey | Terminated | June 05, 2014 | 1869 | 79 | Other | Device Operation | Instructions | Safety Notice/Insructions
Z-1886-2009 | 52658 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker Trio Connector Hex Socket for Torque Wrench 3.5MM |  |  | July 01, 2009 | 2009 | Stryker Spine | The Trio 3.5MM Hex socket has been reported to break under high stresses. | Other | Product Correction letters were sent on July 1, 2009 to all Branches/Agencies, Surgeons and Hospital Risk Management OR Supervisor by Federal Express. The letter indicates the potential hazards associated with the breakage and it reinforces the appropriate surgical technique in using the hex socket. Customers are requested to complete an acknowledgement form and return it via FAX to Aminah Crawford at 201.760.8370 within five (5) days of receipt of the notice. If you have questions, please contact Tiffani Rogers, Regulatory Compliance Manager at 201.760.8206. |  | 143 | 143 | Nationwide - OR, AR, CA, MA, KY, MD, GA, FL, NY, NM, UT, FL, OK, KS, MI, MN, IN, MT, TX, WI, and TN. | Terminated | October 20, 2009 | 111 | 143 | Not_Computer
Z-2005-2009 | 52449 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Roche ACCU-CHEK Performa blood glucose test strips, Roche Diagnostics, Indianapolis, IN. |  |  | June 10, 2009 | 2009 | Roche Diagnostics Corp. | There may be a significant negative bias when testing neonate samples. | Process control | Roche Diagnostics notified consignees in the affected countries of the problem by "Urgent Field Safety Notice" dated June 10, 2009 and instructed to cease using the Accu-Chek Performa and Accu-Chek Inform II systems on neonates, except for systems using test strips with advanced chemistry. For questions and additional information, contact your local Roche Diagnostics sales affiliate or sales representative. |  | 1518 | 1,518 of total products | International Distribution -- Argentina, Australia, Brazil, Chili, China, Colombia, Ecuador, France, Germany, Great Britain, Hong Kong, India, Korea, Malaysia, Mexico, New Zealand, Pakistan, Panama, Peru, Philippines, Singapore, South Africa, Taiwan, Thailand, Uruguay and Venezuela. | Terminated | October 29, 2010 | 506 | 1518 | Not_Computer
Z-1892-2009 | 52704 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView Gamma Camera Radius Drive Assembly, manufactured by Philips Medical Systems (Cleveland) Inc., Cleveland, OH, 3/8" tilt, model number 4535604 34811, 3/4"tilt, model number 4535674 34801, and 3/8" fixed, model number 4535604 34781, all catalog number 882480 |  |  | July 02, 2009 | 2009 | Phillips Nuclear Medicine | Unexpected movement-- if the drive assembly is severely compromised, the detector may slide to its hardware limit. | Process control | Firm initiated its recall notification on July 6, 2009 issuing a field safety notice to its consignees via FedEx and internationally through its personnel. Responses will be tracked for effectiveness. Questions may be directed to Philips Customer Care Center at 1.800.722.9377. |  | 288 | 288 units | Worldwide | Terminated | April 12, 2010 | 284 | 288 | Not_Computer
Z-2009-2009 | 52766 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | Cobham AMOGS Advanced Medical Oxygen Generating System, 70309-03900-101. Concentrator, Oxygen, Molecular Sieve, S/A P/N: 70309-03900-101, Mfg P/N 3261165-0101. The firm name on the label is Carleton Life Support Systems Inc., Davenport, IA. |  |  | July 24, 2009 | 2009 | Carleton Life Support Systems Inc | The backup oxygen supply control valve on the oxygen concentrator could fail resulting in it sticking in the open or closed positions. | Device Design | Cobham Life Support Systems, Inc. issued letters to their two customers via e-mail. One customer letter was dated July 24, 2009 and the other was dated July 28, 2009. Both letters notified the customer of the reason for the recall and that the usage of the concentrator must be restricted to exclude actual use with a patient until corrected. The letter informs customers they are redesigning the backup supply control valve and will coordinate the retrofit in the future. For further information, contact Cobham Life Support at 1-563-383-6421. |  | 23 | 23 units | Nationwide Distribution -- CT and AL. | Terminated | March 22, 2012 | 972 | 23 | Not_Computer
Z-1940-2009 | 52823 | No | BZA | Airway connector. | Anesthesiology | Anesthesiology | 510(K) Exempt | Hudson RICAÂż Sher-I-Swiv &amp; Sher-I-Swiv F/O, Teleflex Medical, Research Triangle Park, NC 27709. |  |  | July 31, 2009 | 2009 | Teleflex Medical | The tether attaching to the double swivel cap may partially or completely break at the attachment points, resulting in the potential of the accidental aspiration of the tether into the patient's lungs. | Process control | Teleflex Medical, Inc. notified consignees via overnight mail starting July 31, 2009 requesting acknowledgement of the recall and a request to return all of the affected product. Distributors were instructed to forward the letter to their consignees and retrieve relevant product from those locations. Affected product received by Teleflex Medical will be destroyed or reworked and replacement product or credit will be provided to the consignee. For further information, contact Teleflex Medical at 1-919-433-4986. |  | 569465 | 569,465 (total for all types) | U.S. Nationwide, Canada, Mexico, Europe, India, Asia Pacific Rim, South and Latin America. (Including: Malaysia, Hong Kong, Vietnam, Singapore, Thailand, Korea, China, Japan, Philippines, Australia, New Zealand, France, Germany, Ireland, Sweden, Costa Rica, Puerto Rico, Jamaica, Brazil, Argentina, Uruguay and Colombia). | Terminated | September 15, 2011 | 776 | 569465 | Not_Computer
Z-2212-2009 | 53135 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Drager Surgical Lights Sola 500 and Sola 700. |  |  | August 26, 2009 | 2009 | Draeger Medical, Inc. | welded seam may fail | Nonconforming Material/Component | Draeger Medical, Inc. issued a "Medical Device Recall" letter dated August 2009 via certified mail with return receipt. Consignees were informed on the affected product and that a service representative will be in contact shortly to schedule a replacement.For further information, contact Draeger Medical at 1-800-543-5047, press 1 at the prompt and then 2349. |  | 20 | 20 units | Nationwide Distribution -- IL, KY, and PA. | Terminated | March 08, 2010 | 194 | 20 | Not_Computer
Z-2219-2009 | 51309 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Ventri, GE Healthcare, Waukesha, WI. |  |  | November 03, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare become aware of an incident which caused a slight patient injury during his unload following a prone-oriented scan on Ventri. Although clear safety instructions were not followed in this case, additional measures will be implemented in an effort to reduce the probability or repeat occurrence. | Process control | GE Healthcare issued a "Product Safety Notification" dated January 2009 to consignees. The letter was addressed to Hospital Administrators, Managers of Radiology/Cardiology &amp; Radiologists/Cardiologists. The letter describes the Safety Issue, Affected Product, Safety Instructions, Product Correction and Contact Information. A GE Field Engineer will visit each site and implement the correction.For further information, contact your local GE Healthcare service representative at 1-800-437-1171 (US or Canada) or 1-262-896-2890 (outside the US). |  | 308 | 308 (252 US, 57 OUS) | Worldwide Distribution -- US (AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, IL, IN, KY, LA, ME, MD, MA, MI, MS, MO, NE, NV, NH, NJ, NY, NC, OH,OK PA, RI,SC, TN, TX, UT, VA, WA, WI and WV), AUSTRALIA, BRAZIL CANADA, DENMARK, FRANCE, GERMANY, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, MALAYSIA, MEXICO, NETHERLANDS, NORWAY,SWEDEN, SWITZERLAND, TURKEY and UNITED KINGDOM. | Terminated | July 12, 2012 | 1347 | 308 | Not_Computer
Z-1939-2009 | 52769 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | The BladderScan BVI 9400 is a portable ultrasound instrument that provides a noninvasive measurement of urinary bladder volume. The device consists of an ultrasound probe that scans the patient's bladder, and a battery operated console that provides measurement related information. The BVI 9400 measures ultrasonic reflections on multiple planes inside the body and produces a three dimensional image. Based on the image, the BVI 9400 calculates and displays bladder volume. |  |  | May 18, 2009 | 2009 | Verathon, Inc. | Three specific issues are being addressed:1. Calculation and use of an incorrect year in dates after 12/31/2009.2. Double scans, continuous scans, and double printing.3. Calibration sensitivity. | Software design | U.S. customers (including Government) were sent a post card on or about 5/18/2009 and were called about the upgrade. Units are to be upgraded by the customers using an online tool called ScanPoint. Customers may elect to return their units to the firm for upgrading.Canadian customers were phoned by Verathon.For foreign units sold and distributed through one of the recalling firm's related firms, Service Bulletin SB-0014 was sent to the distributors located in The Netherlands, United Kingdom, and Australia.Contact Verathon Customer Care at 1-800-331-2313 if you have any questions. |  | 1978 | 1,978 | Worldwide - US, Canada, Malaysia, Hong Kong, Singapore, Argentina, New Zealand, Barbados, Dominican Republic, and Costa Rica. | Terminated | October 07, 2010 | 507 | 1978 | Software | Output/Calculation | Software update | Software Update
Z-2019-2009 | 52862 | No | KGC | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard * Button Device Continuous Feeding Tube with 90 degree Adaptor, 18F x 1.7cm, NON STERILE, REF 000282, Bard Access Systems, Inc., Salt Lake City, UT 84116 USA. |  |  | January 20, 2009 | 2009 | C R Bard Inc | Possible restriction/occlusion in adaptors/tube accessories causing difficulty to infuse during routine feeding treatment. | Equipment maintenance | Customers were notified by letter and instructed to not use or further distribute any of the affected lot numbers. Customers were further instructed to return the recalled products to Bard Access Systems (BAS) through a return goods authorization number. Questions and requests for assistance should be directed to BAS Customer Service at 800-290-1689. |  | 4381 | 13, 337 total of all products | Worldwide Distribution -- US, Canada, Austria, Australia, Belgium, Finland, France, Germany, Great Britain, Holland, Italy, Norway, Spain, Sweden and Switzerland, Japan, and Latin America. | Terminated | December 09, 2009 | 323 | 4381 | Not_Computer
Z-2225-2009 | 50528 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Sytec 6000, Sytec 8000, and Lemage/ProSeed/ProSpeed families. |  |  | December 01, 2008 | 2008 | GE Medical Systems, LLC | During internal testing, GE Healthcare has recently determined that the collimator used on our CT/e and HiSpeed X/i families of Computed Tomography systems does not comply with the specific requirements in United States Title 21 of the Code of Federal Regulations, and IEC 60601-1-3, related to leakage radiation from the diagnostic source assembly, which includes the collimator. The excess leak | Radiation Control for Health and Safety Act | GE Healthcare issued an "Important - Electronic Product Radiation Warning" to affected customers. GE Healthcare field service representative will perform a field corrective action on all affected units. All potential users of the device should be made aware of the notification and the information contained.For further information, contact GE Healthcare at 1-800-437-1171 (US) or by contacting your local GE Healthcare service manager. |  | 643 | 643 units ( 136 US; 507 OUS) | Worldwide Distribution -- UNITED STATES (AL, AZ, AR, CA, CT, DC, FL, GA, IL, IN, KS, KY, LA, MD, MA, MI, LMN, MS, MO, NE, NJ, NM, NY, NC, OH, OK, PA, PR,SC, SD, TN, TX, VA, WA, WV and WI), ALGERIA, ANTIGUA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BELGIUM, BENIN, BOLIVIA, BOSNIA &amp; HERZEGOVINIA, BRAZIL BULGARIA, BYELORUSSIAN, CAMERON, CANADA, CHILE, CHINA, COLOMBIA, CONGO, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK , DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, IVORY COST, JAPAN, JORDAN, KAZAKHSTAN, KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, , MACEDONIA, MALAYSIA, MALTA, MARTINIQUE, MAURITANIA, MAURITIUS, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGER, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL QATAR, ROMANIA, RUSSIAN F=FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND , SYRIA ARAB REPUBLIC, TAIWAN, THAILAND, TRINIDAD &amp; TOBAGO TUNISIA, TURKEY,, UGANDA, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM, YEMEN and YUGOSLAVIA. | Terminated | January 27, 2016 | 2613 | 7648 | Not_Computer
Z-2223-2009 | 52218 | No | JJQ | Colorimeter, photometer, or spectrophotometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BioTek Gen5CL Microplate Data Collection and Analysis Software. |  |  | March 12, 2009 | 2009 | BioTek Instruments, Inc. | Software design allowed false negatives to be read when wells are dark. | Software design | BioTek Instruments, Inc. issued a "Product Recall Notification" dated March 13, 2009 information customers of the affected product. Follow-up letters with software installation instructions were sent March 25, 2009.For further information, contact BioTek Instruments, Inc. at 1-888-451-5171 extension 405. |  | 13 | 13 | Worldwide Distribution -- United States (MN, CA, IL, MA, MI, CA, GA, FL, AZ and TX) and Peru. | Terminated | May 18, 2015 | 2258 | 13 | Software | Output/Calculation | Software update | Software Update
Z-2222-2009 | 51275 | No | LOJ | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | Abbott AxSYM Cancer Assay Disk Version 5.0, for in vitro diagnostic use with the AxSYM System; Abbott Laboratories, Abbott Park, IL 60064; List Number 3D50-05. |  |  | February 20, 2009 | 2009 | Abbott Laboratories | The AxSYM Cancer Assay Disk (LN 3D50-05 version 5.0) contains an AxSYM AFP assay file where the Low Concentration flag for the AFP 1:101 automated dilution protocol the instrument to report results between 0 ng/mL and 0.4 ng/mL (40.4 ng/mL accounting for the 1:101 dilution). The sensitivity of the assay is 0.4 ng/mL. | Device Design | Abbott sent "Product Correction Immediate Action Required" letters dated February 2, 2009 to all AxSYM AFP customers. Customers were requested to install the updated version of the AxSYM AFP assay using version 6.0 of the AxSYM Cancer Assay Disk provided with the letter and to retain a copy of this communication for their records. Any questions regarding the letter were directed to Customer Service at 1-877-4ABBOTT. |  | 8205 | 8,205 disks | Worldwide Distribution -- United States (including Puerto Rico), Australia, Bahamas, Barbados, Bermuda, Brazil, Canada, Cayman Islands, Chile, Columbia, Costa Rica, Curacao (Neth Ant), Dominican Republic, Germany, Guatemala, Hong Kong, Jamaica, Japan, Mexico, New Zealand, Peru, Singapore, South Korea, Taiwan, Thailand, Trinidad and Tobago, Turks and Caicos and Venezuela. | Terminated | August 23, 2010 | 549 | 8205 | Computer
Z-2221-2009 | 50488 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ACCELERATOR APS System Input/Output Module (IOM); List Number: 7L01-01; manufactured by Inpecco SpA, Segrate (MI), Italy. |  |  | October 27, 2008 | 2008 | Abbott Laboratories, Inc | When an ARCHITECT c16000 System is connected to an ACCELERATOR APS System, the software on the ACCELERATOR APS System may not generate an error warning when a Sample Presentation/Sample Queue error occurs, which could result in contamination of subsequent samples. | Software design | Abbott Laboratories notified Customers of the affected device via a "Product Correction" letter dated October 27, 2008. Consignees provided with a temporary work around until a service representative upgrades the system software.For further information, contact your local Abbott Laboratories, Inc. service representative. |  | 15 | 15 units | International Distribution -- Spain, Italy, the United Kingdom, Sweden, Switzerland, the Netherlands, France, Canada and Scotland. | Terminated | December 02, 2009 | 401 | 15 | Software | Alarm/Message | Software update | Software Update
Z-2226-2009 | 52774 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical 15x100mm Kii Optical Access System, Non-Threaded 6/Box, Model Number: C0R36, Labeled Sterile. |  |  | June 04, 2009 | 2009 | Applied Medical Resources Corp | Applied Medical is conducting a coluntary recall of two lot numbers, 1084426 and 1085506, of the 15mm Kii non-threaded optical access system, model C0R36, due to a potential weakness in the cannula tip, possibly resulting in tip breakage. | Component design/selection | Applied Medical Resources Corp. notified Customers via an overnight "Voluntary Recall Notification" Letter dated June 9, 2009. The letter informed customers of the reason for recall and asked that the listed product be returned. The recall letter provided customers with a contact and telephone number for product return questions (1-949-713-8652) and a separate contact for regulatory questions (1-949-713-8369). |  | 934 | 934 units | Class II Recall - Worldwide -- United States, Australia, New Zealand, France, Canada, Norway and Great Britain. | Terminated | July 14, 2011 | 770 | 934 | Not_Computer
Z-2213-2009 | 52230 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker Epic II critical care bed, 115v, Stryker Medical, Portage, MI; Model 2030. |  |  | August 10, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The backrest of the bed may not stay in an upright position (may drift down when left raised). | Device Design | Stryker Medical issued an "Urgent - Medical Device Correction" letter dated August 10, 2009 informing customers to locate the affected product listed and verify whether it is functioning properly. In the case the device is not properly functioning, remove from service and contact the firm as soon as possible at 1-800-STRYKER, option 3. Customers are also asked to return an enclosed return receipt post card and also forward the notification to users who have loaned or purchased the device. A representative from Stryker field service will be contacting you to perform device corrections shortly.For further information, contact Stryker Medical at 1-269-324-6527. |  | 12599 | 12,599 total products | Worldwide Distribution -- United States, Brazil, Canada, Chile, China, India, Korea, Latin America, Mexico, Singapore and Switzerland. | Terminated | April 25, 2011 | 623 | 12599 | Not_Computer
Z-2228-2009 | 53121 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical FAST-CATH Hemostasis Introducer 7F, REF 406702, Length 12 CM, Lot 6372 Max. Guidewire O.D. .038", 2010-01, Manufacturer, St. Jude Medical, 14901 DeVeau place, Minnetonka, MN 55345-2126, Sterile EO. |  |  | August 17, 2009 | 2009 | St. Jude Medical | St. Jude Medical has initiated a voluntary recall of two lots of Fast-Cath Hemostasis Introducers-reorder number 406709, lot number 4219 and reorder number 406702, lot number 6372. They discovered that the product lots could contain some items which have been mis-labeled with the incorrect French size. In one case, the pouch label does not correctly identify the content of the pouch. In another | Error in labeling | Field Representatives were sent a "St Jude Medical Urgent Product Information" letter to be hand delivered to their customers. The letter listed the products and problem and advised consignees to discontinue use of the product. A form was attached for the return of products to the firm.Contact your local St Jude Medical representative for further information. |  | 2000 | 2,000 | Worldwide Distribution -- US (CA, MI, PA and TN), China and Belgium. | Terminated | December 17, 2011 | 852 | 2490 | Not_Computer
Z-2110-2009 | 53087 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | Biphasic LIFEPAK 15 Monitor/Defibrillator. |  |  | August 12, 2009 | 2009 | Physio Control, Inc. | Some LIFEPAK 15 Monitor/Defibrillators have a Non-Invasive Blood Pressure (NIBP) module that was supplied by CAS Medical Systems, Branford, Connecticut who is recalling the modules. The NIBP modules may be susceptible to sudden changes in cuff pressure. Sudden changes in cuff pressure may possibly be caused by bumping or squeezing thecuff. In some instances this may result in a screen display | Software design | The firm notified customers by letter dated August 2009. Customers will have the Noninvasive Blood Pressure (NIBP) module replaced. If additional questions about the notification, contact Technical Support by calling 1-800-442-1142, option 5, Monday-Friday 6:00 a.m. to 4:00 p.m. (Pacific Standard Time).. |  | 445 | 445 devices worldwide | Worldwide Distribution -- US, AUSTRALIA, GERMANY and NETHERLANDS. | Terminated | August 10, 2012 | 1094 | 445 | Not_Computer
Z-2070-2009 | 53075 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | PortexÂż Tracheal Tube, Oral/Nasal, Murphy Eye, Clear, 3.5 mm |  |  | August 25, 2009 | 2009 | Smiths Medical | Pediatric Trachel Tube manufactured with internal diameter smaller than indicated in the label may result in inability to remove secretions and cause partial or complete obstruction of the airway | Process control | Smiths Medical issued an "Urgent Field Safety Notice" dated August 24, 2009 via UPS Second Day Mail including a Guidance Chart. Return of the enclosed Confirmation Form was also requested by fax to 1-603-355-8157. Smiths Medical issued a revised recall letter dated September 9, 2009. Customers are requested to return unused affected product to the firm.For further information, contact Smiths Medical ASD, Inc. by phone at 1-800-348-6064, prompt 4 extension 2273. |  | 24582 | 24,582 (US) 15,620 (International) | Worldwide Distribution -- United States, Andorra, Argentina, Armenia, Australia, Austria, Bahrain, Bangladesh, Barbados, Belgium, Bermuda, Borneo, Brazil, Brunei, Bulgaria, Canada, Chile, China, CÂżÂżte d'Ivoire, Czech Republic, Denmark, East Timor, Egypt, Estonia, Finland, France, Georgia, Germany, Gibraltar, Greece, Guatemala, Vatican City, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Libya, Liechtenstein, Lithuania, Luxembourg, Malaysia, Maldives, Malta, Mexico, New Zealand, Netherlands, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Senegal, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Suriname, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom and Uruguay. | Terminated | November 16, 2012 | 1179 | 655164 | Not_Computer
Z-2012-2009 | 52828 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | Acuvue Advance for Astigmatism Diagnostic Product. |  |  | April 20, 2009 | 2009 | Vistakon | Vistakon is recalling Acuvue Advance for Astigmatism Diagnostic product for mislabeling. | Other | Firm notified consignees by phone from April 7, 2009 to April 27, 2009. Firm requested product disposition. Direct questions about the recall to Vistakon by calling 1-904-443-1763. |  | 73 | 73 | Nationwide. | Terminated | October 01, 2009 | 164 | 73 | Not_Computer
Z-2251-2009 | 52988 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BrainLab Radiotherapy Treatment Planning Software; |  |  | August 12, 2009 | 2009 | Brainlab AG | Failure to conduct important safety checks when using the BrainLab radiotherapy treatment planning software in combination with BrainLab conical collimators could result in unintended radiation outside the conical shaped field, which may lead to serious injury of the patient. | Other | BrainLAB AG issued an "Urgent Field Safety Notice" letter dated August 12, 2009 informing all customers of the BrainLab radiotherapy planning software using conical collimators of the important safety checks to avoid the failure to apply and verify an adequate jaw-shaped field size and field position at the LINAC for the used conical collimator(s) or the failure to mount and verify the correct, intended conical collimator. Consignees were instructed on proper steps to take to avoid further issues. For further questions, contact BrainLAB AG by phone at 1-800-597-5911.*** On February 1, 2010, BrainLab sent supplemental letters to 50 of the accounts who received the August 12, 2009 letter and have Novalis Dose Delivery Systems (excluding all Novalis Tx systems). The accounts were informed that in addition to the safety checks listed in the August letter, BrainLab will be installing a mandatory software modification on their systems that limits the field size to a maximum of 4x4 cm2 when the conical collimators are inserted. The accounts would be contacted by a service representative within the next six weeks to schedule the update. Any questions were directed to BrainLab at 1-800-597-5911. *** |  | 230 | 230 units | Nationwide. | Terminated | March 16, 2011 | 581 | 230 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1935-2009 | 52781 | [6 empty] | Chiba Needle Marked Model: CNM-2308, Size: 23 GA x 20 CM, Remington Medical, Inc., Alpharetta, GA 30005. |  |  | June 19, 2009 | 2009 | Remington Medical Inc. | Two unsealed pouches were identified after product release, compromising product sterility. | Process control | Initial notification to consignees was issued via telephone on June 19, 2009 and by certified mail on June 30, 2009. They were instructed to examine their inventory and quarantine product subject to the recall. If the product was further distributed, they were advised to identify their customers and notify them at once of the recall via the recall letter. A Return Authorization number was provided. Forty of 50 units in distribution were returned from the two customers. |  | 50 | 50 units | US - CA and WI. | Terminated | October 23, 2009 | 126 | 50 | Not_Computer
Z-2112-2009 | 50687 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | EnteraLite Infinity Enteral Feeding Pump, REF INFP01, Zevex Enteral Nutrition Delivery Systems, Salt Lake City, UT 84123. |  |  | December 18, 2008 | 2008 | Zevex International, Inc. | Enteral feeding pumps may not accurately deliver fluid, potentially under-infusing or over-infusing. | Software design (manufacturing process) | All of Zevex' customers and distributors were notified by an Urgent Medical Device Correction letter, dated 12/18/2008. The letters explained the issue, the patient impact, and what necessary actions should be taken. Customers were instructed to contact Zevex to schedule the upgrade of their installed base of pumps at 800-970-2337 or sales@zevex.com. Distributors were to contact their customers and notify them of the letter. |  | 32545 | 32,545 | Nationwide, including 1 VA facility in CA. Foreign distribution to 1 consignee in Canada. | Terminated | October 27, 2011 | 1043 | 32863 | Not_Computer
Z-2078-2009 | 52211 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | DYONICS 25 Fluid Management System Disposable Inflow/Outflow Tubeset With Forked Suction |  |  | May 27, 2009 | 2009 | Smith &amp; Nephew, Inc. Endoscopy Division | The sterility of the device cannot be assured | Package design/selection | Smith &amp; Nephew notified direct accounts by letter dated May 27, 2009 . International accounts notified by Smith &amp; Nephew Inc. Endoscopy Division to each affected Outside United States (OUS) country or distributor by electronic communication, with telephone follow-up as required. The distributor is expected to notify each of their customers that has been confirmed to have product included in this recall. |  | 67697 | 67,697 units | Worldwide Distribution -- USA, Korea, Australia, Hong Kong, India, Mexico, New Zealand, United Kingdom, Spain, Singapore, Austria, Poland, Russia, Finland, Denmark, Japan, Switzerland, China, Italy, Puerto Rico, France, Asia Pacific, Belgium, Canada, Norway, Thailand, The Netherlands, Malaysia, Brazil, Argentina, Chile, Colombia, Costa Rica, Venezuela, Turkey, Israel, Germany, UAE, Greece, and South Africa. | Terminated | May 08, 2012 | 1077 | 67697 | Not_Computer
Z-2116-2009 | 51725 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Uniflex Ti Buttress Thread Screw; |  |  | March 27, 2009 | 2009 | EBI, L.P. | Certain lots of raw material contain inclusions that could potentially reduce the strength of some devices. | Nonconforming Material/Component | Letters to Biomet/EBI reps were sent on March 29, 2009 by US mail to alert them of the product recall and the affected parts that may be in their possession. They were also asked to notify each customer and locate all affected inventory. All letters advised of the recall and the specific lot numbers involved. A letter was sent to the hospital and physician that received affected parts and asked that implanting physicians make any medical decisions. |  | 5 | 5 | Worldwide Distribution -- USA and Spain. | Terminated | September 23, 2009 | 180 | 56 | Not_Computer
Z-2085-2009 | 52683 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Single Shot Epidural Tray, Disposable, Convenience Kit, Reorder Number CUS511-08, Integra LifeSciences Corporation, Salt Lake City, Utah 84104. |  |  | July 13, 2009 | 2009 | Integra Life Sciences Corporation | Lack of packaging seal integrity on various pain management kits may compromise sterility of contents. | Packaging process control | Consignees were notified by telephone on July 09, 2009 and by "Urgent Medical Device Recall" letter on July 13, 2009. Customers were instructed to immediately examine inventory, cease use and distribution of any affected product and to contact Integra Pain Management Customer Service by calling 1-800-241-2210 to arrange for return and replacement. Refer questions to Integra Pain Management by calling 1-800-241-2210. |  | 350 | 350 units | Worldwide Distribution -- US (states of AK, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IN, KY, ME, MO, NC, NJ, NV, NY, OH, OK, PA, RI, SD, ST, TX and UT), Belgium and Japan. | Terminated | October 23, 2009 | 102 | 3594 | Not_Computer
Z-2232-2009 | 52959 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Oridion Medical Capnostream20 capnograph/pulse oximeter monitor and monitor with printer. Part numbers CS 08651, CS 08653, CS08657, and CS08659. |  |  | June 30, 2009 | 2009 | Oridion Medical 1987 Ltd. | 1) The monitor's ability to withstand electrostatic discharge (ESD) to exposed metal connectors on the back of the monitor which may cause the monitor's display to freeze.2) The unit generates electromagnetic radiation above the limits set by applicable standards for medical devices | Device Design | Oridian sent a letter to all its distributors June 25, 2009, instructing them to gather all the units to be exchanged for factory reconditioned units that have the hardware and software enhancements and meet all performance requirements in full. |  | 326 | 326 units | Nationwide. | Terminated | October 01, 2010 | 458 | 326 | Other | Display/Image | Upgrade | Repair
Z-2252-2009 | 53207 | No | LZH | Infusion pump. | General Hospital | General Hospital | 510(k) | Flexiflo Quantum Enteral Pump Model # 50596, 50598 and 59322. |  |  | September 04, 2009 | 2009 | Abbott Nutrition, Abbott Laboratories | The firm initiated this voluntary recall because some of its Flexiflo Quantum Enteral Pumps, have a 110-volt AC power cord plug that may crack or fail. This potential power cord plug failure may cause electrical shock and/or a delay in feeding that could represent a potential health and safety risk. | Device Design | An "Urgent Device Recall" letter dated September 4, 2009 was issued via FedEx overnight delivery to all direct consignees. The letter described the affected product, issue and immediate actions for customers. Direct consignees were further instructed to immediately contact affected customers and notify them about the Urgent Device Recall. Abbott Nutrition (AN) Device Call Center (DCC) will telephone those customers who do not responded promptly to the "Urgent Device Recall" letter with the completed AN Account Summary of the Flexiflo Quantum Enteral Pump Shipments. AN Device Call Center will make arrangements for the service or replacement of the recalled pumps. Please contact the AN Device Call Center at 1-877-457-0249, Monday through Friday between the hours of 8:30 am and 5:00 pm Eastern Standard Time (EST). An AN representative will work with you to arrange service or replacement of the recalled enteral pumps. |  | 5173 | 5,173 pumps | Nationwide Distribution (states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY). | Terminated | September 19, 2012 | 1111 | 5173 | Battery | Physical Safety Hazards | Replace | Remove or Replace
Z-2243-2009 | 52729 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art SystemÂż, Version 3.X |  |  | August 10, 2009 | 2009 | TomoTherapy Incorporated | TomoTherapy received one customer report indicating that after starting a patient procedure from the StatusConsole, the radiation on light and the audible indicator on the Status Console came on as expected,however the Operator Station (OS) did not display the dose rate and the procedure time, or the couch position updates. Within a few minutes, the customer pressed the Stop button on the Sta | Software design | Consignees were sent a TomoTherapy "Urgent Medical Device Correction Field Safety Notice" letter dated August 10, 2009. The letter addressed the Issue, Product Affected, Investigation, Recommended Actions and Resolution.Customers were recommended to perform the following actions (in the event that the dose rate, procedure timer, and couch position updates are not present on the OS, after pressing Start on the status console and the radiation on light and audible indicators are active):1) Immediately press the Stop button on the Status Console and 2) contact your authorized TomoTherapy Service Representative for further direction. The affected issue will be corrected in the next software version, which is scheduled for release in early 2010. The TomoTherapy Customer Contact Center can be reached 24 hours a day, 7 days a week, by e-mail atsupport@tomotherapy.com or by phone at 1-866-368-4807 (North America). |  | 259 | 259 (169 US; 91 OUS) | Worldwide Distribution -- US (states of AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL IN, IA, KS, KY, LA, MD, MA, MI, MN,MO, NE,NV,NJ, NM, NY, OH, OK OR, PA, PR, SC, SC, TN, TX, UT, VA, WA, WI and WY) andBelgiumCanadaSwitzerlandChinaGermany Spain FranceEnglandHong KongIndiaItaly Japan KoreaNetherlandsPolandSaudi ArabiaSwedenSingaporeTurkey andTaiwan. | Terminated | December 17, 2011 | 859 | 259 | Software | Display/Image | Software update | Software Update
Z-2312-2009 | 51730 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Multi-Axial Screw (TI); 8.00MM (D) x 70MM(L) W/Collar; |  |  | April 02, 2009 | 2009 | EBI, L.P. | Some devices could have been made with raw material lots that may have inclusions, and could result in decreased strength of the device. | Nonconforming Material/Component | The firm sent Urgent Recall Notices to all direct accounts, hospital Risk Management and Physicians by first class mail.Letters are dated March 29, 2009, April 2, 2009 and April 16, 2009 respectively. The letter describes the problem and that that there is a possible risk of a device failure.An additional e-mail was sent on May 22, 2009 to 5 customers due to an expansion of the recall to include 4 additional items.On June 19, 2009, the firm again expanded the recall to include 2 additional products and notified customers.Questions should be directed to William Hennig, at 1.973.299.9300. ext. 1557. |  | 1 | 1 | Worldwide - US -NY, NJ, MI, MD, FL, GA, PR, MN, SD, OK, TX,, LA, CA, PA, OH, NV, HI, ME, NC, KS, UT, CO, TN, WV, VA, and WA. Foreign: Brazil, Colombia, Denmark, Germany, Finland, Greece, Italy, Japan, The Netherlands, Switzerland, Sweden, Portugal, Spain, United Kingdom, Saudi Arabia, and Trinidad. | Terminated | September 23, 2009 | 174 | 708 | Not_Computer
Z-2209-2009 | 52780 | No | MNB |  |  | Obstetrics/Gynecology | PMA | Boston Scientific HydroThermAblator Endometrial Ablation System- |  |  | July 31, 2009 | 2009 | Boston Scientific Corporation | Cracked procedure sheaths, incorrect care/use of device, and console malfunctions may result in fluid leaks and loss of a cervical seal that are responsible for burns. | Process design | Boston Scientific issued an Urgent Medical Device Recall - Immediate Action Required notification to hospitals and distributors dated July 31, 2009 via Federal Express and requested the immediate discontinued use and segregation of Recall product for return and Complete and return the Reply Verification Tracking Form. An email was sent to Physicians to provide the recall information. Boston Scientific issued a Clarification letter dated August 6, 2009 via Federal Express as a follow-up to the July 31, 2009 Recall letter for the following reason: customers may remove product from the outer box and store it on the inventory shelf by the inner package only. If this is a practice at the facility, a customer must consider both the inner and outer packaging product codes when searching for affected recalled product as the UPN numbers on the inner and outer labeling are different. A reply verification tracking form for customers was provided in case they have located any additional product to return. |  | 22744 | 22,744 eaches | Worldwide Distribution -- USA, Austria, Denmark, Finland, France, Great Britain, Italy, Netherlands, and Sweden. | Terminated | May 03, 2011 | 641 | 24299 | Not_Computer
Z-2227-2009 | 52631 | helmet, cranial, for protective use | NDA | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | 3i Custom # 5753, OSCM2913, 2.9MMX 3.4MM X 13MM, OSSEOTITE CUSTOM, MICROMINI IMPLANT, LOT # 470258, Sterile. Implant Innovations, Inc. 4555 Riverside Drive, Palm Beach Gardens, FL 33410. |  |  | December 11, 2006 | 2006 | Biomet 3i, LLC | Healing Abutment was not made within required manufacturing specification. The product did not have the relief behind the thread, which does not allow proper seating. | Nonconforming Material/Component | Each of the US customers was contacted by phone. Following the phone contacts, Biomet 3i performed a follow- up of each call with a faxed letter. Each customer was instructed to check their respective inventory for the affected product and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory. They were also instructed to return the affected product to Biomet 3i. Contact Biomet 3i, LLC at 1.561.776.6706, if there are any questions. |  | 24 | 24 | USA and Italy | Terminated | October 01, 2009 | 1025 | 24 | Not_Computer
Z-0003-2009 | 53158 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Revolution XR/d systems with control room PC model 5117866-2 with Software collector 5194381. GE Healthcare, 3000 North Grandview, Waukesha, WI 53188. |  |  | February 14, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has recently become aware of an error in the generator software on the Revolution XR/dconfigured with HP 8200 PC consoles that reports a mAs reading that is higher than is actually exposed to patients. | Software design | Consignees were sent a GE Healthcare "Product Safety Notification" letter dated March 7, 2008. The letter was addressed to Hospital Administrators, Mangers of Radiology, and Radiologists. The letter listed the Safety Issue, Affected Product, Safety Instructions, and Contact Information. |  | 83 | 83 (63 USA, 20 OUS) | Worldwide Distribution -- USA, including states of AL, AZ, CA, CO, FL, ID, IL, IN, LA, MA, MI, MO, NY, NC, OH, OR, PA, SC, SD, TN, TX, WA, and WI and countries of Canada, China, Japan, Korea, and Taiwan. | Terminated | December 18, 2011 | 1403 | 83 | Software | Output/Calculation | Notification | Safety Notice/Insructions
Z-2256-2009 | 52711 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | Posey Connecting Strap/Belt, Cotton, 36" (232036). |  |  | April 20, 2009 | 2009 | J T Posey Company | Posey has initiated a recall on certain Posey product units which were manufactured using an incorrect key lock buckle. The intended use of these products is to limit a patient's movement to the extent necessary for treatment, examination or protection of the patient or others. | Employee error | JT Posey Company issued an "Urgent Medical Device Recall (Removal)" letter dated May 14, 2009 informing customers of the affected products. Any used and/or existing inventory subject to this recall is requested to be returned to Posey including the Recall Reply Fax-Back Form.For further questions, contact your local sales representative or JT Posey Company at 1-800-447-6739 extension 111. |  | 47 | 47 | Class II Recall - Worldwide Distribution -- United States, Canada, France, Belgium and the Netherlands. | Terminated | April 11, 2012 | 1087 | 3136 | Not_Computer
Z-2230-2009 | 52824 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Maxi Move Scoop Stretcher; Note: The stretcher is an accessory to the Maxi Move Patient Lift; Made by Arjo Hospital Equipment AB, Eslov, Sweden; product number KRA2210-20. |  |  | August 18, 2009 | 2009 | ArjoHuntleigh | A Scoop Stretcher is made up of two frames that are held together with two locking devices, one on each end. Each locking device contains a safety latch. The safety latch could unintentionally release when the side frames are pulled for positioning. | Manufacturing material removal | ArjoHuntleigh sent their customers a cover letter and the Field Safety Notification, informing them of the potential problem with the safety latches on the stretcher and a pending visit from an ArjoHuntleigh Service Technician to correct the problem. These letters were sent via UPS 2nd day delivery on 8/18/09. The customers were requested to test their stretchers upon receipt of the letter following the instructions for checking the safety latches on the stretcher:* Place the stretcher on its side on a non-scratch surface.* Try to pull the two sides apart with the safety latch engaged.* Test both the head and foot ends of the stretcher.* If the stretcher pulls apart, remove it from use immediately.It is important to note that all affected stretchers will be upgraded, whether they fail the test or not. |  | 214 | 214 units | Worldwide Distribution | Terminated | January 07, 2010 | 142 | 214 | Not_Computer
Z-2239-2009 | 53132 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | ILUMAVision Software, version 2.1.0 Picture Archiving and Communication system |  |  | August 03, 2009 | 2009 | IMTEC IMAGING L.L.C. | Under certain circumstances, software Version 2.1.0 has potential for incorrect data imaging and patient identification. | Software design | Firm sent notifications to all ILUMAVision customers on 8/3/2009 and offered upgraded software. The notification letter indicated the the reason for the recall and identified the potential issues associated with the device. Contact 3M IMTEC Technical Support at 1.800.879.9799, if there are any questions. |  | 223 | 223 | Worldwide - US - Foreign: Australia, Austria, Belgium, Bulgaria, Canada, Croatia, Egypt, France, Germany, Hong Kong, Italy, Korea, Malaysia, Mexico, The Netherlands, Poland, Saudi Arabia, Spain, Turkey, and United Kingdom | Terminated | January 27, 2010 | 177 | 223 | Software | Output/Calculation | Software update | Software Update
Z-2231-2009 | 53045 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Eclipse Proton Convolution Superposition DCF, version 8.6.15, Model H48, radiation therapy treatment planning system. |  |  | July 17, 2009 | 2009 | Varian Medical Systems Oncology Systems | Software Issue/Mistreatment Potential-- Due to software issues, dose distribution may be incorrect and unintended radiation doses may be delivered. | Software design | Varian Medical Systems Recall initiated on July 17, 2009. A product notification letter was distributed to the affected user, with a description of the problem and user corrective action steps. The letter will be distributed to the sales, marketing and service organizations via certified mail. A software correction is planned. |  | 1 | 1 | Product was distributed to one user in Texas | Terminated | November 27, 2010 | 498 | 1 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2236-2009 | 53041 | No | JAD | X-ray radiation therapy system. | Radiology | Radiology | 510(k) | Action Products, Inc., Bol-X-II, 30x30x0.5cm, item number 32052, packaged 10/carton. Polymer pad used during radiation therapy. |  |  | June 26, 2009 | 2009 | Action Products Incorporated | Medical device intended for use by healthcare practitioners was incorrectly labeled. 1.0 cm thick pad was labeled as 0.5 cm thickness, and vice versa. | Process control | Action Products Inc. notified one affected distributor by phone and email on 06/26/09 and issued a Return Merchandise Authorization. |  | 1 | 1 carton | Nationwide distribution. | Terminated | September 29, 2009 | 95 | 1 | Not_Computer
Z-2235-2009 | 50473 | Yes | NVZ |  |  | Cardiovascular | PMA | Sorin Group Reply DR Rate responsive dual-chamber pacemaker. Made in Italy by: Sorin BioMedica CRM S.r.l., Via Crescentino s.n., 13040 Saluggia (VC) Italy. ELA Medical, Inc., 14401 West 65th Way, Arvada, CO 80004, USA. REPLY DR is a DDDR-type dual-chamber rate-responsive pacemaker. It is equipped with a physiological sensor and an accelerometer to allow adaptation of pacing to suit the patient's activity. The REPLY device can be programmed to separate pacing and sensing polarities, independent of chamber. |  |  | September 23, 2008 | 2008 | ela Medical Llc | Analysis showed that the screwdriver was not fully inserted in the setscrew's hex cavity, which damaged the setscrew. | Employee error | A Reply Technical Note, dated September 2008, titled Additional Instructions for use at Implant for REPLY pacemakers, was issued to physicians on September 24, 2008. The Technical Note describes the background for an increased incidence of complaints relating to the screwdriver and setscrew. The letter provided a reminder for inserting the screwdriver, and additional instructions for insertion of the screwdriver. |  | 233 | 233 | Nationwide Distribution -- USA, including states of AL, AZ, CA, CO, CT, FL, GA, KS, KY, LA, MA, MI, NC, NJ, OH, OK, SC, TN, and TX. | Terminated | December 17, 2011 | 1180 | 233 | Not_Computer
Z-0345-2009 | 50742 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Integris H5000F |  |  | December 10, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Tube failure occurs involving fluoroscopy of large patients and a partial failure of the grid switch. | Radiation Control for Health and Safety Act | An Electronic Product Radiation Warning, dated October 2008, was sent to customers. The letter identified the problem and under what circumstances that it could occur; the actions that should be taken by the customer/user to prevent risks for patients or users; and the actions planned by Philips to correct the problem. Further information or support concerning the issue could be obtained by contacting Bob Pettis at 425-487-7469 or the Philips Call Center at 800-722-9377.Firm FCO tracking numbers 72200107, 72200108, 72200111, 72200112, 72200117, 72200118, 72200119, 72200122, 72200123, 72200124, 72200125, 72200129, 72200142, 7989980, 7080085, 70600034, 70600035, 72200132, 72200133, 72200137, 72200138, 72200140 and 72200141. C &amp; R # 1217116-10/20/08-016C |  | 929 | 929 units | Worldwide Distribution | Terminated | August 17, 2012 | 1346 | 3361 | Not_Computer
Z-1155-2009 | 51856 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Fluorostar 7900 Mobile Fluoroscopy System |  |  | February 19, 2009 | 2009 | Ge Healthcare | The alarm timer is reset after each exposure and therefore never reaches the five minute alarm state unless one exposure exceeds 5 minutes in duration. | Radiation Control for Health and Safety Act | GE Healthcare sent a customer notification out to all sites that have potentially affected systems. GEHC Surgery will also install a software correction into the affected systems in order to resolve this issue at no charge to the customer. Until the GEHC OEC Fluorostar 7900 has received the corrective software installed, they are encouraging facilities to exercise caution in monitoring the amount of cumulative fluoroscopy time when the affected system is in use. |  | 554 | 554 units | Worldwide Distribution | Terminated | July 22, 2011 | 883 | 554 | Software | Alarm/Message | Software update | Software Update
Z-2241-2009 | 52703 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | InnovaÂż 3100/3100IQ GE Innova 3100 / 3100 IQ, Cardiovascular Imaging System, GE Healthcare, Waukesha, WI 53188 (The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector). |  |  | July 31, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently became aware of a potential power supply failure with the InnovaÂż 2000, 3100 and 4100 systems that may impact patient safety. A potential power supply failure inside the GE equipment cabinet could occur with no advance warning and thereby prevent x-ray generation and image processing. Should this occur, the system cannot be restarted. | Component design/selection | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated June 25, 2009. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology or Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.GE Healthcare will correct all affected systems by replacing the power supply located in the GE equipment cabinet. A GE Healthcare service representative will contact you to arrange for this correction. Direct questions about the recall to the firm's Call Center by calling: 1-800-437 -1171, option 4 (United States) or 0120 - 055 - 919 (Japan). For other countries, please contact your local GE Healthcare field service Representative if you have any questions or concerns regarding the notification. |  | 6 | 6 | Worldwide Distribution -- US (states of AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA., MI, MN, MO, NE, NV, NH, NJ, NY, NC, OH, OK , OR, PA, PR, RI, SC, TN, TX, UT VA, WA, WI and WV),TUNISIA, THAILAND , TAIWAN, SYRIAN ARAB REPUBLIC, SWEDEN SPAIN, SOUTH AFRICA SAUDI ARABIA, RUSSIAN FEDERATION , POLAND, NORWAY, NEW ZEALAND, NETHERLANDS, MOROCCO, MEXICO, MARTINIQUE, MACAU, LEBANON, KUWAIT, REPUBLIC OF KOREA, JAPAN, ITALY, ISRAEL, IRAN, INDONESIA, INDIA, GREECE, GERMANY, FRANCE , CHINA, CANADA, BULGARIA, BRAZIL BELGIUM, BELARUS and AUSTRALIA. | Terminated | May 15, 2012 | 1019 | 508 | Not_Computer
Z-1888-2009 | 52723 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE MV Systems (Part Number 8139789), equipped with a syngo RT Therapist (Part Number 8162815) or RT Therapist Assist system (Part Number 8162807), with Software version NET VA10A for RTT4.1.122 or RTT 4.1.125, manufactured by Siemens Medical Solutions, Concord, CA. |  |  | April 09, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Unexpected treatment plan alteration-- Software issue may result in unintended modifications to treatment | Software design | Siemens Medical Solutions USA, Inc. notified consignees of the affected product by advisory letter via Certified Mail on April 09, 2009.For further information, contact Siemens Medical Solutions at 1-925-602-8157. |  | 53 | 53 units | Worldwide Distribution -- US (MA, UT, OH, NJ, NE, FL, OH, MI and PA), Germany, Poland, the UK, Ireland, Italy, France, South Korea, and Belgium. | Terminated | April 12, 2011 | 733 | 53 | Software | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-2250-2009 | 53080 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | B. Braun Blistered Omnifix L.L. Syringe packaged in the Easypump LT 60-24, 60 ml x 2 ml/hr |  |  | August 01, 2008 | 2008 | I-Flow Corporation | I-Flow identified a B. Braun syringe package that had a tear/hole in the paper portion of the primary package.Used to provide an optional filling device for the Easypump. | Packaging process control | I-Flow sent a recall letter dated August 1, 2008 to consignees telling them of the above mentioned recall and asking them to Inspect their Easypump inventory to determine if they have any of the affected lots identified on this Recall Notification Letter. They were told not to distribute affected devices and quarantined remaining inventory immediately. They were asked to check the appropriate response on the attached I-Flow Fax sheet and return the form back to I-Flow within 48 hours of receipt.The I-Flow Customer Service number (949) 206-2700, and hours of operation (7:00 am - 5:00 pm PST) were provided to answer questions. |  | 1608 | 1,608 units | France | Terminated | October 07, 2009 | 432 | 2177 | Not_Computer
Z-2314-2009 | 52616 | No | LZY | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Tandem Unipolar, 46 MM CO-CR, REF 126646, Sterile, Smith &amp; Nephew, Memphis TX 38116 |  |  | January 12, 2009 | 2009 | Smith &amp; Nephew Inc | 46 mm Femoral head was mismarked and packaged as being 45 mm | Process control | All affected Smith &amp; Nephew Sales Representatives were notified of problem and sent recall notice in overnight mail on 01/12/2009. All affected Smith &amp; Nephew International Distributors were notified via e-mail and telephone on 01/12/2009. They were instructed to immediately quarantine the product for return. The firm re-issued the letter on 04/06/2009 to address hazard to health.Contact Smith &amp; Nephew, Inc at 1.901.399.6771 if there are any questions. |  | 10 | 10 units | CA, FL, MI, and Canada | Terminated | March 25, 2011 | 802 | 10 | Not_Computer
Z-1036-2009 | 53176 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 4100/4100 IQ. , GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector). The product is indicated for use in generating fluoroscopic images of human anatomy for vascular angiography diagnostic and interventional procedures and optionally, rotational imaging procedures. It is also intended for generating fluoroscopic images of human anatomy for cardiology, diagnostic and interventional procedures. |  |  | November 04, 2008 | 2008 | GE Medical Systems, LLC | One of three copper spectral filters could remain within the x-ray while the system displays no spectral filter. The displayed Air Kerma rate and thus cumulative Air Kerma are overestimated. The Air Kerma rate may be over estimated by up to 60%. | Radiation Control for Health and Safety Act | Affected customers were notified by means of a Important Electric Product Radiation Warning letter. GE field personnel will be replacing the collimators. |  | 166 | 166 | Nationwide Distribution -- AL, AZ, CA, CO, FL, GA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MS, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and Puerto Rico. | Terminated | May 16, 2012 | 1289 | 166 | Not_Computer
Z-2244-2009 | 53143 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Renegade" Hi-Flo Microcatheter, Peripheral Vascular Use. Sterilized with Ethylene Oxide. Legal Manufacturer: Boston Scientific Cork Ltd., Business &amp; Technology Park, Model Farm Road, Cork, Ireland. |  |  | August 26, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific has determined that the sterile barrier in the packaging of the affected product may be compromised. It has been reported that the pouch seal (top seal) may be breached. | Packaging | Consignees were sent a Boston Scientific Urgent Medical Device Recall Removal - Immediate Action Required letter Dated August 26,2009. The letter was addressed to Risk Manager / Field Action Contact. The letter described the problem and product involved. Advised consignees to segregate the affected product and return it to Boston Scientific. Requested the return and completion of The Account Reply Verification Tracking form enclosed. Replacements will be issued for all recalled product that is properly returned to Boston Scientific. |  | 80885 | 80,885 (51,491 US; 29,394 OUS) | Worldwide Distribution -- USA, excluding New Mexico and Wyoming, and countries of Europe, the Middle East and Africa, Japan, American Virgin Islands, Argentina, Aruba, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dutch Antilles, Egypt, Finland, France, Georgia, Germany, Great Britain, Greece, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Jordan, Kuwait, Latvia, Libya, Luxembourg, Macau, Malaysia, Mexico, Namibia, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Poland, Portugal, Puerto Rico, Romania, Saudia Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, and Vietnam. | Terminated | August 24, 2012 | 1094 | 118920 | Not_Computer
Z-2316-2009 | 52858 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Centricity Enterprise Archive (EA) versions 3.0 to 3.0.7 software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | August 14, 2009 | 2009 | GE Healthcare Integrated IT Solutions | Software error: There is potential safety issue with Centricity Enterprise Archive (EA) 3.0.x software where study split operations are not correctly replicated to a secondary 'shadow' archive. | Software design | GE Healthcare Integrated IT Solutions sent Urgent Medical Device Correction letters dated August 7, 2009 to the attention to the Director of Radiology, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator. The letters advised users of the potential safety issue and provided safety instructions to follow until the software is updated. Users were asked to complete the enclosed Customer Reply form and fax the completed form to 1-847-939-1479. Any questions should be directed to the Remote Online Center or the GE Customer Care Center at 1-800-437-1171. |  | 212 | 212 units | Worldwide Distribution -- United States, Belgium, Canada, Cayman Islands, China, Denmark, France, Germany, Great Britain, Hong Kong, India, Israel, Italy, Korea, Malaysia, Malta, Mexico, New Zealand, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan and Turkey. | Terminated | February 25, 2011 | 560 | 212 | Software | Device Operation | Software update | Software Update
Z-1690-2009 | 53260 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Innova 4100, GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector). |  |  | October 09, 2008 | 2008 | Ge Healthcare | GE Healthcare has become aware of an Air Kerma Rate issue associated with a large Equivalent Patient Thickness (EPT) on the Innova system that my impact patient safety. During a Maximum Air Kerma Rate (AKR) calibration procedure check, GE Healthcare found the AKR calibration procedure does not exactly reflect the worst-case clinical usage of the product. This may result, when using large Equival | Radiation Control for Health and Safety Act | Consignees were sent a GE Healthcare "Important electronic Product Radiation Warning" letter dated November 3, 2008. The letter was addressed to Healthcare Administrator/Risk Managers, Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Defect Description and Related Hazards, Affected Product Details, Safety Instructions, Product Correction and Contact Information.For further information, contact the GE Healthcare Technical Support Center at 1-800-437-1171. |  | 929 | 929 units total | Nationwide Distribution -- United States including Puerto Rico and Guam. | Terminated | January 07, 2012 | 1185 | 929 | Not_Computer
Z-0002-2009 | 53157 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Revolution XR/d system control room PC model 5117866-2 with Software collector 5115434 |  |  | February 14, 2008 | 2008 | GE Medical Systems, LLC | GE Healthcare has recently become aware of a potential risk of observing a blank image after exposure onRevolution XR/d systems equipped with HP 8200 PC user consoles that may impact patient safety.It was reported that images acquired with off-center collimation for extremities with theposition selected as Digits to front or Digits to back are displayed as blank. Investigation intoth | Software design | Consignees were sent a GE Healthcare "Product Safety Notification" letter dated March 7, 2008. The letter was addressed to Hospital Administrators, Mangers of Radiology and Radiologists. The letter listed the Safety Issue, Affected Product, Safety Instructions, Corrective Action and Contact Information. |  | 129 | 129 | Nationwide Distribution -- AL, AZ, CA, CO, FL, GA, HI, IL , IN, KY, LA, MD, MA, MI,MN,MS, MO, MT, NE, NH, NJ, NM, NY, NC, OH, OK OR, PA, SC, TN, TX, VA, WA, and WV. | Terminated | January 07, 2012 | 1423 | 129 | Software | Display/Image | Notification | Safety Notice/Insructions
Z-2313-2009 | 52989 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Penumbra Neuron Delivery Catheter 053, 115 cm/12 cm-- MP, manufactured by Penumbra Inc., Alameda, CA |  |  | July 01, 2009 | 2009 | Penumbra Inc. | Misbranding--Some products may be identified with the wrong ID band content 'Neuron 6F 070' instead of 'Neuron 6F delivery' | Mixed-up of materials/components | Firm initiated its recall on June 30, 2009. Consignees notified on July 6 and 7, with follow up contact on July 7, 2009. Contact Penumbra Inc. at 1.510.618.3223, if there are any questions. |  | 37 | 37 units | CA, FL, NC, MI, MO, VA, AZ, CT, IA, NY, and CO, Foreign: Germany and Australia. | Terminated | November 30, 2010 | 517 | 37 | Not_Computer
Z-2234-2009 | 52794 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | S&W; Electrodes for Defibrillation, 10 Defibrillation Electrodes, Reorder #90101, S &amp; W Healthcare Corp, 1525 Flight Path Drive Brooksville, Florida 34604 |  |  | July 10, 2009 | 2009 | Leonhard Lang Gmbh | A potential risk was identified associated with the connectors of the defibrillation electrodes DF20. It was found that some of the connectors do not slip easily into the sockets of the defibrillator or an extension cord of the defibrillator. In such cases increased force or some wiggling is required to insert the connector fully. A potential risk for patients exist, if (a) such a connector (b) | Nonconforming Material/Component | A Leonhard Lang letter dated June 9, 2009 was sent June 6, 2009 to its customer. The letter identified the problem and the product. The letter instructed the consignee to stop using or distributing the product and to quarantine them. It also requested them to report all quantities of product that they had in stock using the attached form. |  | 3044 | 3044 | Worldwide Distribution - US - FL Foreign: Austria, Germany, Romania, France, Dubai, Greece, Italy, South Korea, United Kingdom, Hungary, Latvia, Saudi Arabia and Lebanon | Terminated | December 17, 2011 | 890 | 4477 | Not_Computer
Z-2248-2009 | 52849 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | AccuSure, U-100, 31Gauge, 1cc Insulin Syringe, 5/16" (8mm), Short Needle, NDC 0603-7002-21, packaged 10 syringes per package, 10 packages per carton |  |  | August 21, 2009 | 2009 | Vintage Pharmaceuticals LLC DBA Qualitest Pharmaceuticals | Needles separate from the barrels of the syringes | Nonconforming Material/Component | The firm initiated the recall via a press release on 08/21/2009. They began notifying their consignees of the problem and the recall by letter on 08/21/2009. email: recallresponse@qualitestrx.com |  | 1652000 | 16,520 boxes of 100 units | Nationwide | Terminated | November 24, 2010 | 460 | 5386000 | Not_Computer
Z-1666-2009 | 53034 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Diagnost 96/97 DSI |  |  | July 06, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | The x-ray system fail to meet the requirements of 21 CFR 1020.32(d)(2)(ii) when a specific tube failure occurs involving fluoroscopy of large patients and a partial failure of the grid switch | Radiation Control for Health and Safety Act | A Customer Notification letter describing the potential grid switch failure and actions planned by Philips to correct the problem will be sent to all users of the affected systems during the week of July 6, 2009. |  | 17 | 17 units distributed in US. | Worldwide Distribution | Terminated | December 16, 2011 | 893 | 401 | Not_Computer
Z-0623-2009 | 50748 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems MultiDiagnost Eleva Flat Detector System with software version 4.3.1 or higher |  |  | December 30, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | System fails to meet requirements when it is used at kV levels greater than 115 kV at certain frame rates and in certain acquisition modes, as grid switch leakage may occur. | Radiation Control for Health and Safety Act | A Customer Notification letter describing the failure, the means, by which a customer can prevent the failure mode from occurring, and actions planned by Philips to correct the problem were sent to all users of the affected systems. Philips will modify the system software and will be installed at no cost to the customer. |  | 30 | 30 systems | Nationwide distribution including states CA, CO, IA, ID, KS, MN, NC, NJ, NY, OH, SC, TX, UT, VA, and WI. | Terminated | September 28, 2010 | 637 | 30 | Other |  | Software update | Software Update
Z-1223-2009 | 51867 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | MilCAM Recon III, Family of IR Imagers that incorporates a Class IIIb infared laser pointer |  |  | March 10, 2009 | 2009 | Flir Systems, Inc. | MilCAM Recon III models were sold to non-DoD customers prior to certification and reporting as required in Title 21 of the Code of Federal Regulations. Conditions of sale were not stated in variance approval letters issued for similar laser products. | Radiation Control for Health and Safety Act | FLIR Systems notified direct account by letter of the Corrective Action Plan and updated certification label that references CDRH issued variance and instructions on how to install it. |  | 58 | 58 | Nationwide Distribution -- Washington, D.C. | Terminated | May 28, 2015 | 2270 | 58 | Not_Computer
Z-1118-2009 | 51490 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Definium 5000 X-ray System, model number 5220493 and 5220393-2 |  |  | March 24, 2009 | 2009 | GE Healthcare, LLC | A defect in the AEC board in the system's control causes failure to terminate additional exposures at the exposure time selected for the initial exposure. | Radiation Control for Health and Safety Act | A notification was sent to all affected customer sites alerting them to this issue, with instructions for safe use of the affected systems. |  | 85 | 85 | Nationwide Distribution | Terminated | May 16, 2012 | 1149 | 85 | Hardware | Device Operation | Instructions | Safety Notice/Insructions
Z-0937-2009 | 51488 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | MilCAM Recon III Model Family of IR Imagers that incorporate a Class I infared laser for Laser Range Finding (LRF). |  |  | January 30, 2009 | 2009 | Flir Systems, Inc. | LRF models were sold to non-DoD customers prior to certification and reporting as required in Title 21 of the Code of Federal Regulations, Subchapter J, Parts 1000 - 1050 (21 CFR 1000 - 1050). A filter defect allows transmission of laser energy such that the affected LRF would be Class 3B laser product, however the product is reported to CDRH as a Class I product. | Radiation Control for Health and Safety Act | Customers will be notified by a letter instructing them to stop using the LRF, contact a Customer Service Representative, and to detach and return the Laser Range Finder. For each returned LRF, FLIR will provide a certified Class I LRF. |  | 120 | 120 | Nationwide Distribution | Terminated | June 02, 2015 | 2314 | 120 | Not_Computer
Z-1259-2009 | 51879 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare CT Systems Table Model Number 2269165 used on LightSpeed 1X through 4X and HiSpeed Qxi CT Systems. |  |  | March 12, 2009 | 2009 | GE Medical Systems, LLC | Failure to provide the certification label required by 21 CFR 1010.2. | Radiation Control for Health and Safety Act | A GE Field Service Representative will visit each affected customer site and affix the certification label required. A Field Modification Instruction (FMI Number 25412) has been developed for release to all facilities in possession of the affected systems. |  | 1179 | 1179 total of all units | Nationwide Distribution | Terminated | December 10, 2011 | 1003 | 1179 | Not_Computer
Z-0005-2009 | 53451 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix fluoroscopy DFP-8000D Digital Radiography System with version 3.40 or higher software. Controller for Infinix interventional angiography systems. |  |  | September 29, 2008 | 2008 | Toshiba American Medical Systems Inc | Audible signal is not available at all times when using the HLC fluoroscopy mode. Special means of activation are not designed in for selection of High Level Controls, as required by regulation. | Software design | Toshiba issued a letter entitled Important Electronic Product Radiation Warning, September 30, 2008, explaining the issue. The customers were requested to reply to the letter via fax-back form, and a Toshiba service representative would install the corrective action. |  | 129 | 129 | Nationwide | Terminated | October 04, 2012 | 1466 | 129 | Not_Computer
Z-1658-2009 | 52173 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Celex/GE Precision MPi, REF no.: 02980000 |  |  | May 01, 2009 | 2009 | NRT - Nordisk Rontgen Teknik A/S | The GE Precision MPi is in violation with 21 CFR Part 1020.32 requiring a signal audible to the fluoroscopist to sound for each passage of 5 minutes of fluoroscopic irradiation time during an examination or procedure. | Radiation Control for Health and Safety Act | An Advisory Notice concerning GE Precision MPi, dated May 1, 2009, was issued to Healthcare Professionals. The notice discussed the problem, risk, recommended immediate action, future actio, and affected system and serial numbers. A GE Field Service Representative will visit each affected U.S. customer site and introduce a speaker connected to the generator touch screen (to which the audible tone is delivered) in the operator's room, to be placed on the wall in the x-ray room. For further information, customers were to contact their local GE field engineer. |  | 38 | 38 units (US) | Nationwide Distribution, including CA, CO, FL, IN, MA, MI, MN, NV, NY, OH, OR, PA, TN, VT, WA, WI. | Terminated | October 20, 2009 | 172 | 38 | Not_Computer
Z-2276-2009 | 52730 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMedÂż II Programmable Pump, 8637, Rx Only, Sterilized by ethylene oxide gas, Medtronic inc., Minneapolis, MN 55432-5604. (The SynchroMedÂż II Pump is supplied in 20 ml or 40 ml reservoir size.) |  |  | July 01, 2009 | 2009 | Medtronic Neuromodulation | This recall is related to the potential for reduced battery performance in a small percentage of Medtronic Model 8637 SynchroMedÂż II pumps with batteries manufactured during two distinct time periods prior to April 2005.Nature of the Device Issue: As part of ongoing analysis of returned explanted product, Medtronic has confirmed that reduced battery performance resulted in eight (8) occurrence | Nonconforming Material/Component | Consignees were sent on July 7, 2009 a Medtronic " Urgent: Medical Device Correction" letter dated July 2009. The letter was addressed to Healthcare Professionals. The letter described the nature of the Device Issue, Scope, Potential Severity of the Issue, Recommendations, Ongoing patient Management Recommendations and Additional Information. Important patient management information is also available at http://www.professional.medtronic.com under the heading Product Advisories.Direct questions to your Medtronic Field Representative or Medtronic Neuromodulation Technical Services by calling at 1-800-707-0933. |  | 18572 | 18,572 | US (including DC and Puerto Rico), Australia, India, canada, Hong Kong, Japan, Austria, Belgium Cyprus, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, Iceland, Ireland, Italy, Luxembourg, Netherlands, Norway, Poland, Portugal, San Marino, Saudi Arabi, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, United Kingdom and Vatican City. | Terminated | May 10, 2012 | 1044 | 18572 | Battery | Device Operation | Notification | Safety Notice/Insructions
Z-2315-2009 | 53162 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ZOLL R Series Defibrillator/Pacemaker/Monitor |  |  | August 31, 2009 | 2009 | ZOLL Medical Corporation, World Wide Headquarters | Failure of Device to display ECG could delay delivery of defibrillation or pacing therapy. | Nonconforming Material/Component | Zoll notified domestic consignees on August 31, 2009 and international consignees on September 1, 2009 by an Urgent Medical Device Correction letter. Customers were instructed to continue to perform daily tests per the recommendations provided in the R Series Operator's Guide. In the event of a failure, customers are to remove the device from service and contact biomedical or clinical engineering department per their facility's protocol. Units that operate as expected can remain in service and back-up monitor should be available in the event of a failure. Zoll's technical support department should be contacted at 1-800-348-9011 or +1-978-421-9460 to schedule the return of affected devices with the defect for replacement of the suspect component. |  | 247 | 247 units | Worldwide Distribution -- USA, UK, Austria, Australia,, Canada, Switzerland, Germany, France, French Guiana, Jordan, Macedonia, Netherlands, Panama, Poland, Romania, and Saudi Arabia. | Terminated | May 30, 2012 | 1003 | 247 | Other | Display/Image | Replace component | Remove or Replace
Z-1991-2009 | 52861 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Informatics System included with the Dimension Vista System. |  |  | July 14, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | incorrect results may be transmitted to LIS | Software design | Siemens issued an "Urgent Device Correction" letter dated July 2009. The letter informed the user of the affected product including instructions for avoiding the problem.For further information, contact Siemens Healthcare Diagnostics Technical Solution Center at 1-800-441-9250. |  | 359 | 359 units | Worldwide Distribution -- United States, Australia, Austria, Belgium, Canada, France, Germany, Italy, Japan, Malaysia, Netherlands, Portugal, Slovenia, Spain and Switzerland. | Terminated | January 05, 2010 | 175 | 359 | Not_Computer
Z-1814-2009 | 52680 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ConvaTec, Engenex Tunnel Dressing, for use with Engenex Negative Pressure Wound Therapy, (1 x 15 cm) |  |  | July 13, 2009 | 2009 | ConvaTec | Labeling: the actual manufacturing date (2/2009) and actual expiration date (2/2011) are switched. | Labeling Change Control | Recall notification letters were hand delivered on July 13, 2009 to all direct customers (tier 1) by ConvaTec reps. Letters were provided to all direct customers so that they could notify all customers that received the product. |  | 764 | 764 | Nationwide distribution and ConvaTec sales rep in the UK and France. | Terminated | October 02, 2009 | 81 | 764 | Not_Computer
Z-2238-2009 | 48191 | No | DKJ | Salicylate test system. | Toxicology | Toxicology | 510(k) | SYNCHRON Systems Salicylate (SALY) Reagent; Part Number: 378194. |  |  | October 22, 2007 | 2007 | Beckman Coulter Inc | The Salicylate Reagent lots listed here may fail calibration and/or fail to maintain linearity at the upper end of analytical range. | Nonconforming Material/Component | An "Urgent: Product Corrective Action" letter dated October 22, 2007 was issued to customers. Customers were instructed to do the following: 1) Discontinue use and discard appropriately the affected lots of Salicylate Reagent., 2) Contact your Customer Service Representatives at 1-800-526-3821 in the United States or local Beckman Coulter Representative to discuss replacement options and 3) Complete and return the enclosed response form within 10 days. Contact your Customer Service Representative at 1-800-526-3821 in the United States or 1-800-463-7828 in Canada or, if you are outside the United States or Canada, please contact your local Beckman Coulter Representative to discuss replacement options.Direct any questions regarding this Product Corrective Action letter to the Customer Support Center (Hotline) at 1-800-854-3633 in the United States and Canada. Outside the United States and Canada, please contact your local Beckman Coulter Representative. |  | 7842 | 7,842 total | Nationwide and Canada. | Terminated | July 16, 2012 | 1729 | 7842 | Not_Computer
Z-2326-2009 | 51156 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | DCL (Diagnostic Chemicals Limited) Creatine Enzyme Reagent (R2) Catalog Number: 265-30. |  |  | February 12, 2009 | 2009 | Diagnostic Chemicals, Ltd. | Labeled with extended expiration date (24 months instead of 18 months) | Labeling Change Control | Genzyme Corp. notified consignees by Letter dated February 12, 2009 via Federal Express. Accounts are requested to discontinue use and discard units. Customers were also asked to complete and sign the Confirmation of Notification the fax to 1-902-628-6504.For further information, contact Genzyme Technical Support at 1-800-565-0265. |  | 82 | 82 kits | Worldwide Distribution -- United States (AL, CA, NJ, MO, PA, KY, FL, LA, MN, NY, NC, OH, TX and UT), Canada and Korea. | Terminated | May 23, 2012 | 1196 | 107 | Not_Computer
Z-2317-2009 | 51258 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AxSYM System Tubing Decontamination Solution; 500 mL bottle; Abbott Laboratories, Diagnostics Division, Abbott Park, IL 60064; List Number 07B05-09. |  |  | February 18, 2009 | 2009 | Abbott Laboratories | Bottle labels from lot 211974, AxSYM System Tubing Decontamination Solution (List Number 7B05-09), were found to be missing the following information: lot number, date of manufacture, and expiration date. | Manufacturing material removal | Abbott sent Product Recall Immediate Action Required letters dated February 18, 2009 to customers informing them of the affected product. Customers were instructed to examine their inventory and dispose all affected product. Accounts were requested to complete and fax back customer reply sheet acknowledging receipt of the letter, as well as the number of bottles destroyed, and to retain a copy of the letter for their records. For further information, contact Abbott Customer Service at 1-877-4ABBOTT. |  | 1965 | 1,965 kits | Worldwide Distribution -- United States including Puerto Rico and Argentina, Brazil, Canada, Chile, Columbia, Germany, Hong Kong, Japan, Mexico, Peru, South Korea, Thailand and Trinidad &amp; Tobago. | Terminated | July 08, 2010 | 505 | 1965 | Not_Computer
Z-2328-2009 | 51872 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Martel Printer Accessory to i-STAT Portable Clinical Analyzer, rechargeable battery pack. |  |  | April 22, 2009 | 2009 | Abbott Point of Care Inc. | It appears that the rechargeable battery pack is overheating with or without the presence of smoke. | Nonconforming Material/Component | Urgent Recall Notice letters dated April 2009 were sent out the third week of April by Federal Express. The letter discussed the background, required actions, and additional information. Customers are to identify all affected Martel printer rechargeable battery packs, record the number of affected packs on the attached business reply cards, and return the card along with the affected batteries as instructed. Questions should be directed to Abbott Technical Support at 800-366-8020. |  | 2552 | 2552 | Worldwide Distribution | Terminated | February 25, 2010 | 309 | 2552 | Battery | Physical Safety Hazards | Return | Remove or Replace
Z-2325-2009 | 51928 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i System Assay CD ROM - US, list 6E58-22, Version 22.01, and ARCHITECT i System Assay CD-ROM - WW (excluding US), list 6E59-23, Version 23.0, containing the ARCHITECT LH assay file; Abbott Laboratories, Diagnostic Division, Abbott Park, IL 60064 USA; Product List Number: 6E58 (US) and 6E59 (WW) Versions 23 and lower. |  |  | April 24, 2009 | 2009 | Abbott Laboratories | There continues to be an elevated level of complaint activity for calibration error code 1227, when using ARCHITECT LH reagent lots (6C25-22 or 6C25-27) due to the use of assay disk versions that do not contain the updated ARCHITECT LH assay file contained on ARCHITECT i System Assay CD ROM versions 24 and higher (US LN 6E58, WW LN 6E59). | Software change control | Abbott issued a "Product Correction-Immediate Action Required" letter dated April 24, 2009. The letter provided users with the steps to take to determine what version of the ARCHITECT LH Assay is on their instrument and instructions for installing the assay file provided with the letter if needed. Users were requested to return a customer reply letter acknowledging receipt of the letter. For additional information, contact 1-877-4ABBOTT. |  | 5713 | 1,037CD-ROMs - US, 4,676 CD-ROMs - WW | Worldwide Distribution -- United States including Puerto Rico and Argentina, Australia, Brazil, Canada, Chile, Colombia, Costa Rica, Curacao, Germany, Guatemala, Hong Kong, Japan, Mexico, New Zealand, Peru, Singapore, South Korea, Taiwan, Thailand, Uruguay and Venezuela. | Terminated | November 30, 2009 | 220 | 5713 | Other | Device Operation | Update file | Software Update
Z-1871-2009 | 52089 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Whatman Body Fluid Collection Paper 903 |  |  | May 06, 2009 | 2009 | Whatman, Inc. | The two portions of the card are coming apart prematurely. | Process control | Whatman issued an Urgent Medical Device Correction letter to to the direct account on 5/6/09 requesting this notice is distributed to users. The notice states the safety issue; requests users to inspect 903 cards prior to use, if the 903 Paper Portion of the card partially or completely detaches from the Demographic portion of the card, discard the card obtain a new card and inspect. On May 26, 2009 the firm expanded the notification to additional users of the 903 cards to inspect cards prior to use. |  | 2853410 | 2,853,410 cards | Nationwide Distribution -- AZ, FL, GA, NC, IL, IN, LA, KS, KY, MI, MO, NE, NY, OH, OK, PA, SC, OR, TX, TN,UT VA, WI, WI, &amp; PR. | Terminated | October 25, 2011 | 902 | 2853410 | Not_Computer
Z-2324-2009 | 52813 | No | JTZ | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | SensititreÂż Mueller-Hinton Broth w/Lysed Horse Blood (MHB w/LHB), Model Number: CP112-10; Lot number 150444SA. |  |  | June 04, 2009 | 2009 | Trek Diagnostic Systems | Thirty-one boxes of unapproved quarantine product, CP112-10 MHB w/LHB, Lot # 150444SA, were inadvertently shipped to customers on 6/2/09 instead of shipping approved Lot # 150344SA. | Employee error | On 6/4/2009 each customer was telephoned and instructed to immediately return the recall lot back to the firm. Replacement product was shipped. |  | 310 | 31 boxes/10/11 ml tubes | Nationwide Distribution -- PA, VA &amp; WA. | Terminated | February 26, 2010 | 267 | 310 | Not_Computer
Z-2253-2009 | 52961 | No | LXV |  |  | Ear Nose &amp; Throat | 510(k) | VEST 6.6 or earlier software for use with Neuro Kinetics products. |  |  | July 24, 2009 | 2009 | Neuro Kinetics | Marketed without a 510k or PMA submission to include the normative data display. | PMA | A notification letter was emailed to customers starting on August 10, 2009. The letter stated that the normative display needed to be removed pending FDA clearance.Direct questions to Neuro Kinetics, Inc. by calling 1-412-963-6649. |  | 72 | 72 units | Worldwide Distribution -- US (states of AL, AZ, CA, DC, FL, GA, HI, IN, MA, MD, MI, MN, NC, NY, OR, PA, SC, TN, UT, VA, WA, and WV), Korea, United Kingdom, and Spain. | Terminated | January 13, 2010 | 173 | 72 | Software |  | Notification | Safety Notice/Insructions
Z-0001-2010 | 53426 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Penumbra Neuron Delivery Catheter 070, percutaneous catheter, catalog numbers PND6F0701058, PND6F070956, PND6F0701058M, and PND6F070956M; manufactured by Penumbra, Inc., Alameda, CA. |  |  | September 22, 2009 | 2009 | Penumbra Inc. | Failure-- product can kink or ovalize, which may result in difficulty in catheter advancement or delivery of other devices and damage to catheter, which could lead to patient injury or death as a result of portions being retained in the patient. | Device Design | The firm issued an Urgent Voluntary Field Removal Notice letter on 09/22/2009 to its consignees, asking them to inspect their stocks; remove affected product and place in quarantine for return to Penumbra Inc.; and complete and return the attached product identification / return form. |  | 2246 | 2246 units | Worldwide Distribution -- USA, Canada, Hong Kong, Chile, the UK, Denmark, Norway, Sweden, Finland, Netherlands, Germany, Switzerland, Austria, and Spain. | Terminated | November 24, 2010 | 428 | 2246 | Not_Computer
Z-0010-2010 | 52202 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Multi-Drug 8 Panel Integrated iCup Drug Screen Test Cup. Part Number DUE-187-071-08A. |  |  | April 27, 2009 | 2009 | Biosite Inc | The cup associated with these lots of product may have a pin-sized hole in the floor of the cup. As a result, there is a potential for affected lots of cups to leak urine after an extended period of time. | Process control | Biosite, Inc. issued a "Notice of Correction" letter dated May 1, 2009 to consignees describing the problem and instructed users of proper use in order to avoid any cross contamination of samples until all affected product is completely used.For further information, contact Biosite, Inc. using the following information:Biosite Incorporated, an Inverness Medical Company9975 Summers Ridge RoadSan Diego, CA 92121Phone: 1-877-441-7440, option 2Fax: 1-858-695-7100Email: productadvisory@invmed.com |  | 113575 | 113,575 cups distributed in total | Nationwide Distribution | Terminated | April 02, 2012 | 1071 | 113575 | Not_Computer
Z-0018-2010 | 52866 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | LifeCare PCA with Hospira Mednet, list numbers 20709-04-01, 20709-04-03, 20709-04-51, 20709-04-53, 20709-04-77; AC cord manufactured by Electri-Cord, Westfield PA, Infusion pumps manufactured by Hospira, Morgan Hill, CA. |  |  | August 11, 2009 | 2009 | Hospira Inc | Fire/Shock hazard-- A component used in the device may cause sparks/flashes and poses a fire and shock hazard. | Component design/selection | Firm initiated its recall on 08/14/2009. A nationwide press release was issued to AP, and an Urgent: Device Field Correction letter to consignees was sent via UPS. The letter recommended that customers check their infusion pumps and spare parts in their inventory to determine if the specified AC Power Cords are in use at their facility; and, to inspect the affected cords for bent or cracked prongs and bridge connection, burnt plastic, or excessive wear and tear. Also, it was recommended that customers immediately discontinue use of the affected cords exhibiting the described characteristics and contact their local Hospira representative for assistance. The letter stated that customers may continue to use and monitor cords that are not exhibiting the described characteristics until replacement cords are sent to their facility. Hospira will contact customers to schedule replacement. Customers are to complete and return the attached Reply Form. Question should be directed to Hospira Customer Service at 1-800-241-4002. |  | 204999 | 204,999 total | Worldwide Distribution | Terminated | November 30, 2010 | 476 | 204999 | Not_Computer
Z-0036-2010 | 52760 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ROC Lumbar bolt caddy, Part Number: 81997-01-05. |  |  | December 16, 2008 | 2008 | Alphatec Spine, Inc. | Alphatec discovered that the lot of bolt caddies had the incorrect bolt length laser etched on one side of the caddy. The actual ROC bolts themselves are correctly laser etched. | Employee error | Alphatec contacted consignees via phone on December 16, 2008 and informed consignees that they will arrange to have Federal Express contact them for pick-up.Consignees were asked to check their inventory to see where this product is located, and ensure that it is put on hold for return; and what days and hours they are in the office, (Confirm address: attach copy of the pick-up request form.)If they have any follow up questions, they were told to call 1-800-922-1356. |  | 7 | 7 Screws | Nationwide Distribution -- Including states of TX, FL, &amp; NJ. | Terminated | November 13, 2009 | 332 | 7 | Not_Computer
Z-0038-2010 | 53081 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | AGI-2238 Portable 4 Cylinder Manifold Cart, for connecting medical gas cylinders. |  |  | August 06, 2009 | 2009 | Airgas Puritan Medical | Yoke may become displaced. Yoke index safety system could become compromised, allowing a user to attach the wrong gas type to the yoke connection. | Component design/selection | Airgas issued a Medical Device Removal letter dated 8/6/09 to its customers informing them of the problem and the need to have a yoke back-up plate installed on the Manifold Block Assembly. |  | 208 | 208 | Nationwide distribution. | Terminated | January 26, 2010 | 173 | 208 | Not_Computer
Z-0059-2010 | 52847 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Medical Gas Yoke Assembly. Generic 4-Cylinder Manifold Cart Model #8070. |  |  | August 10, 2009 | 2009 | Western - a Scott Fetzer Company | The yoke pins of the attached medical gas yoke assemblies could become displaced to such a degree that the yoke pin index safety system could become compromised and allow a post valve of one gas type to become interchangeable with a yoke connection of another gas type. | Component design/selection | Western Enterprises issued an "Urgent-Device Recall" notification dated August 20, 2009 informing Consignees of the affected product. Consignees are to contact the firm to request backup plates and installation instructions or request an RGA and shipping instructions for return of the manifold block assembly for modification.For further information, contact Western at 1-800-783-7890. |  | 18 | 18 cart assemblies containing 72 yoke assemblies (36 Oxygen yokes and 36 Nitrous Oxide yokes) | Nationwide Distribution -- AL, AR, CA, CO, FL, GA, HI, IL, IN, MA, OH, OR, PA, TX, WA and WI. | Terminated | April 25, 2012 | 989 | 363 | Not_Computer
Z-0064-2010 | 53033 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | 200996-Cable Assembly, Hybrid Extended Length, |  |  | July 17, 2009 | 2009 | Mako Surgical Corporation | Cables fail to provide adequate transmission. MAKO Surgical Corporation is recalling 200996 Cable Assembly, Hybrid Extended Length; 200997 Cable Assembly, Hybrid, Guidance Module; and 200998 Cable Assembly, Hybrid, UI Panel. These parts are sub components of the 20300 Rio Standard System. | Device Design | A Recall Notice, dated July 15, 2009, was sent to owners of the RIO system (hospital OR staff) with a signed acknowledgement returned to MAKO. Surgeons using the RIO system during the recall time period, regardless of whether their system has already been upgraded, are given the notification letter and asked to sign and return the acknowledgement memo. The letter stated that customers' RIO will be upgraded to replace the potentially faulty cable with a new cable. Their MAKOplasty Specialist (MPS) will notify them when their system was scheduled to have their fiber optic cable upgrade. The upgrade will be installed by a Field Service Engineer (FSE) free of charge and the RIO will be completely tested prior to being returned to service. Customers should direct questions or concerns to their MPS, FSE, or Jim Pomeroy, who can be reached at 954-927-2044. |  | 23 | 23 | Nationwide Distribution | Terminated | May 12, 2010 | 299 | 23 | Not_Computer
Z-0065-2010 | 52964 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Emerald Hematology Analyzer, model number 09H39-01. Distributed by Abbott Diagnostics Division, Santa Clara, CA. |  |  | July 16, 2009 | 2009 | Abbott Laboratories | Software bug. The CELL-DYN Emerald WBC L1 flag is configured incorrectly. The flag is not generated according to system requirements. | Software design | Abbott Diagnostics issued a "Product Correction" notice dated July 16, 2009 describing the affected product. The notice was sent to all currently active users and included a customer reply form, a software update and installation instructions.For further information, contact Abbott Customer Support at 1-877-4ABBOTT. Customers outside the U.S. should contact your local hematology customer support representative. |  | 411 | 411 units | Worldwide Distribution -- United States (AL, AR, AZ, CA, CT, FL, GA, IL, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI and WV), Canada, Germany, Puerto Rico, Singapore, and South Korea. | Terminated | December 06, 2010 | 508 | 411 | Software | Alarm/Message | Software update | Software Update
Z-0066-2010 | 52632 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Sterrad CycleSure Biological Indicator, Product Code 14324. |  |  | June 23, 2009 | 2009 | Advanced Sterilization Products | Biological Indicator indicating a positive result does not maintain yellow color for the entire incubation period and some indicators experience media evaporation prior to the end of the labeled maximum incubation time. | Device Design | An Important Product Labeling Correction - New Instructions for Use dated, June 26, 2009, has been proposed to be sent by the firm for Cyclosure Biological Indicators to consignees. The revised instructions reduce the maximum read time for the indicators to 3 days (72 hours). The letter also states why ASP changed the Instructions. Questions are to be directed to ASP Customer Care Center. |  |  |  | Worldwide Distribution. | Terminated | November 29, 2011 | 889 |  | Not_Computer
Z-0067-2010 | 52808 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Arthro-Knife, ConMed, Linvatec 11311 Concept Blvd, Largo FL 33773, Made in USA. RX only. U.S. PAT NO. 5.141.517. |  |  | July 31, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed has determined there is a possibility that the device may not fully retract once it is deployed. | Process control | ConvaMed Linvatec USA issued an "Urgent Medical Device Recall" notification via Fed Ex to all direct consignees of the affected product. Users were asked to inventory and return all affected product to the firm and to return a reply form by mail or fax.For further information, contact customer service at 1-800-237-0169 for US or your local ConMed Linvatec office for outside US or email at Custserv1@linvatec.com. |  | 66144 | Approximately 66,144 products. | Worldwide Distribution | Terminated | February 25, 2011 | 574 | 66144 | Not_Computer
Z-0069-2010 | 53498 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Extended PTFE Coated Blade electrosurgical electrodes, 6.5" (16.51CM) Length, 12 each &amp; 20 each per box, Labeled MEDLINE brand. |  |  | August 28, 2009 | 2009 | Unimed Surgical Products, Inc. | Unimed Surgical Products, Inc., Largo, FL initiated a voluntary recall of the Extended PTFE Coated Blade electrosurgical electrodes, product number ES0014, due to concerns regarding product sterility for Lot #042209-03. | Process control | Unimed Surgical Products, Inc. issued a Medical Device Removal letter dated August 28, 2009. The customer, Medline Industries, Inc., was instructed to quarantine any undistributed product and to notify its customers of the recall.For further information, contact Unimed Surgical Products, Inc. at 1-727-546-1900 or 1-800-886-1225. |  | 1920 | 1920 units | Nationwide Distribution -- MI | Terminated | May 12, 2010 | 257 | 1920 | Not_Computer
Z-0068-2010 | 53435 | No | MAF |  |  | Cardiovascular | PMA | Boston Scientific LiberteÂżÂż MonorailÂż and Over-the-Wire Coronary Stent Systems to be relabeled as VeriFLEX". |  |  | September 11, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific initiated a field correction for the LiberteÂż Bare-Metal coronary stent products. They have received reports from cardiac cath labs that TAXUS Liberte Paclitaxel-Eluting (TAXUSÂż LibertÂżÂż Drug-Eluting) coronary stents have been inadvertently selected when the interventional cardiologist intended to implant a LiberteÂż Bare-Metal stent, and LibertÂżÂż Bare-Metal stents have been ina | Labeling design | Boston Scientific issued an "Urgent Medical Device Field Correction" letter dated September 18, 2009. The letter was addressed to Risk Manger/Field Action Contact. The letter described the product, problem and actions which included renaming the LiberteÂżÂż Bare-Metal stent as VeriFlex" Bare-Metal stent. Enclosed with the letter included the Direction for Field Correction and a Reply Verification Tracking Form.For further information, contact you local Boston Scientific Sales Representative or call 1-800-811-3211. |  | 391769 | 391,769 | Nationwide Distribution -- United States including Puerto Rico. | Terminated | June 28, 2012 | 1021 | 391769 | Not_Computer
Z-0073-2010 | 53163 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biotrak PLLA Pin, Resorbable Fixation System, Biotrak 20 mm Pin; sterile pin packaged in sterile bag, enclosed within a box labeled Biotrak Resorbable Fixation System Biotrak 20 mm Pin REF 30170320-S Manufactured or Distributed by: ACUMED LLC 5885 NW Cornelius Pass Rd Hillsboro, OR 97124. |  |  | August 21, 2009 | 2009 | Acumed LLC | Labeling may contain 2 lot numbers and 2 model numbers.It was observed that the box top label (30170060-S, W67391) does not match the two side labels (3010320-S, W67385). The part and patient labeling inside the box match the side labels (30170320-S, W67385). | Pending | Urgent Notice: Device Recall Letters dated August 21, 2009 were sent to all consignees advising of the issue and and required action. Customers are to identify and quarantine the product having the affected lot code(s). Customers are to also return all product from the specified lot(s) to Acumed immediately. If product was further distributed, then those consignees should be notified of the recall and return of the product(s) should be requested. Questions should be directed to Acumed Customer Service at 1-888-627-9957. |  | 125 | 125 | Worldwide Distribution -- USA and South Africa. | Terminated | November 06, 2010 | 442 | 125 | Not_Computer
Z-0074-2010 | 53296 | No | EMA | Dental cement. | Dental | Dental | 510(k) | Kerr TempBond Clear, Part # 33351, 1717 West Collins Avenue, Orange, CA 92867. |  |  | September 09, 2009 | 2009 | Kerr Corp | The recall was initiated because the catalyst portion of the syringe may have polymerized, therefore making the product unusable. | Nonconforming Material/Component | Kerr Corporation issued an "Urgent Medical Device Recall" dated September 8, 2009. The letter informed customers that the product recalled was shipped to their establishment between June and July 2009 and if they have any of the affected products listed to contact Kerr Customer Care at 1-800-537-7123 to receive an RMA number. The RMA will allow for a quick return and replacement or credit. |  | 13518 | 13,518 units | Worldwide Distribution -- United States, Canada, Mexico, Hong Kong, Israel, Singapore, Colombia, Dominican Republic and Europe. | Terminated | December 12, 2011 | 824 | 13518 | Not_Computer
Z-0080-2010 | 53327 | No | JEY | Bone plate. | Dental | Dental | 510(k) | SYNTHES TI MATRIXMANDIBLE SAGITTAL SPLIT PL 1.5MM THK/4H/45MM, Catalog # SD448.005. |  |  | September 03, 2009 | 2009 | Synthes USA (HQ), Inc. | The recalling firm has determined that the customized plates instruments do not have a cleared premarket notification. | PMA | The customer account was notified by an Urgent: Medical Device Removal letter via USPS certified mail on 9/4/2009. The letter stated that the products did not have clearance for the hospital's intended use and that if the customer has any of the identified devices, they should call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Also, the customer should complete the Verification Section, and return the letter along with any remaining product. For questions the hospital could call 610-719-5865. |  | 15 | 15 | Nationwide Distribution -- State of CA. | Terminated | March 08, 2010 | 186 | 136 | Not_Computer
Z-0086-2010 | 52048 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Ortho Development Corp., Pivot BiPolar Head, 28mm, Hip Prosthesis, Femoral Head. |  |  | April 29, 2009 | 2009 | Ortho Development Corporation | A Pivot BiPolar Head, the femoral head in a hip prosthesis, might not move freely after assembly due a stacking up of manufacturing tolerances. | Device Design | Ortho sent a Field notification Memo dated 4/2/2009 to agents and distributors, giving them pre-implant examination guidelines. Decision to recall non-implanted units was made on 4/29/2009 and this was communicated by telephone to distributors. |  | 23 | 23 | AZ, CA, FL, MA, MI. No military, government or foreign distribution. | Terminated | May 25, 2010 | 391 | 23 | Not_Computer
Z-0083-2010 | 53073 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Foot Pedal, Foot Control for SC2000 and SC2100 Powered, simple cranial drills, burrs, trephines, and their accessories. Catalog Numbers: EMAX2-FP and FPLUS-FP. |  |  | August 13, 2009 | 2009 | The Anspach Effort, Inc. | The Anspach Effort, Inc. is recalling the Foot Control/ Foot Pedal for SC2000 and SC2100 Powered, simple cranial drills, burrs, trephines, and accessories. The Foot Control Cables lacks the shield connection from the cable to the connector body. The lack of shielding may result in radiated emissions which could possibly cause interference on other electrical equipment adjacent to the foot pedal | Device Design | Anspach issued an "Urgent Product Removal" Letter dated August 17, 2009 informing customers of the affected product. Consignees are asked to screen their inventory and return any affected product including a completed Product Replacement Form to the firm.For further questions, contact Anspach at 1-800-327-6887 or 1-561-627-1080. |  | 14 | 14 | Worldwide Distribution -- Canada, Indonesia, United Kingdom, Korea and United States (NC, OK, TN, CA, KS, MI and WA). | Terminated | January 05, 2010 | 145 | 14 | Hardware | Device Operation | Return | Remove or Replace
Z-0035-2010 | 51228 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | WarmAir 135 Thermal Regulating System. Catalog Numbers: 86186-115 V, 06186-115 V (sold in the U.S.) and #86185-100 V ( not sold in the U.S.). Intended for warming |  |  | February 16, 2009 | 2009 | Cincinnati Sub-Zero Products Inc | The blower on the WarmAir 135 Warming Unit may overheat. Resistance in a connector between circuit board and wiring harness results in excessive heat that may melt connector or deform the outer casing of the blower unit. | Process control | The firm sent an "URGENT Medical Device Recall' notice via First Class Mail on 2/16/2009 (Note: the letter is dated 2/13/2009). The notice informs the customers of the potential for the device to overheat and of the 'small number of customer complaints received by CSZ concerning this issue. In addition, the notice asks that the customers inspect their units for signs of discoloration indicative of a faulty connector. In the notice, the customers were instructed to return the product to CSZ for repair by contacting their Customer Service division and obtaining a Return Authorization number and specific instructions concerning packaging and returning of the unit(s) for repair. The customer is also instructed to complete the attached Recall Response Form which accompanies the recall notification. The notification provides that customer with a toll-free number at (800) 989-7373ext.3234 or call collect at (513) 719-3234 to reach CSZ representative between the hours of 8:30 a.m. until 4:30 P.M. Monday through Friday EST. Once the customer ships the recalled unit to CSZ, the firm will repair the unit(s) using new connectors manufactured, after which time, the repaired unit is returned to distribution for customer use. |  | 2674 | 2,674 units | Nationwide Distribution -- DE, TX, PA, NJ, LA, MI, GA, MA, OH, KY, NY, FL, NV, NC, MT, RI, WA, MN, AL, CT, MD, TN, UT, VA, CO, MO, IA, AZ, AR, IN, NE, KS, WI, ME, MS, and Puerto Rico. | Terminated | September 19, 2012 | 1311 | 2674 | Hardware | Physical Safety Hazards | Repair | Repair
Z-0090-2010 | 53192 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Vacuum Collection Bottle, 1 gallon (3200 mL); Ohio Medical Company, Gurnee, IL 60031; part number 0212-0302-600 (6700-209-500); the bottles are sold alone and with various cap &amp; float assemblies: |  |  | September 21, 2009 | 2009 | Ohio Medical Corporation | The re-useable collection bottles were made with polycarbonate instead of the specified polysulfone. The recommended cleaning of the bottles include use of ammonia and autoclaving, which may have an adverse impact on polycarbonate. | Process control | Ohio Medical Corporation issued recall letters beginning September 17, 2009. Consignees were informed of the affected product and requested to return affected collection bottles to the firm. The accounts were requested to complete and return the enclosed acknowledgement form, indicating the number of affected units on hand, to obtain a returned material authorization for the return of the recalled bottles and their replacement. For further information, contact Ohio Medical Quality Assurance Manager, Anita Christ, at 847-855-6270. |  | 823 | 823 bottles | Worldwide Distribution -- United States, Canada, Chile, Costa Rica, Cyprus, Egypt, El Salvador, Greece, Hong Kong, India, Israel, Japan, Jordan, Korea, Kuwait, Malaysia, Netherlands, Oman, Paraguay, Philippines, Qatar, Saudi Arabia, Singapore, Thailand, Turkey, United Kingdom and Venezuela. | Terminated | June 22, 2011 | 639 | 6715 | Not_Computer
Z-0088-2010 | 53086 | No | LNQ |  |  | Gastroenterology/Urology | HDE - Humanitarian Device Exemption | Medtronic Enterra Therapy, Gastric Electrical Stimulation System, consisting of model 3116 Enterra Therapy Neurostimulator and model 4351 Unipolar Intramuscular Lead. Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. |  |  | August 25, 2009 | 2009 | Medtronic Neuromodulation | Medtronic is in the process of updating the Enterra Therapy labeling to include information regarding the risks of bowel obstruction and/or perforation. Medtronic has received 15 reports since November 2002 of bowel obstruction and/or perforation of the bowel associated with Enterra Therapy System leads. Of the 15 reported events, 11 required surgical intervention (all patients recovered withou | Device Design | Medtronic Neuromodulation issued a "Medical Device Correction" letter beginning August 25, 2009 addressed to Healthcare Professionals. The letter described the Nature of the Issue, Scope, Potential Severity of the Issue, Enterra Therapy System labeling updating information warnings, Recommendations and other Additional Information.For further information, contact Medronic Neuromodulation Technical Services at 1-800-707-0933. |  | 1909 | 1909 (1095 US, 814 OUS) | Worldwide Distribution. | Terminated | May 22, 2012 | 1001 | 1909 | Not_Computer
Z-0094-2010 | 53436 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Wallach LL50x Cryosurgical Freezer CO2 Unpinned Yoke System |  |  | September 29, 2009 | 2009 | Cooper Surgical, Inc. | CO2 Cryosystem thermal load may exceed the exhaust gas phase, resulting in ice formation in the exhaust valve. Should this occur, the Cryo tip may cool at a less than optimal rate, resulting in ice formation in the exhaust valve. | Device Design | Cooper Surgical notified customers by letter titled "Notice of Recall - Urgent" dated September 29, 2009 advising them of the problem and requesting return of the device. International accounts were notified by letter Beginning October 1, 2009.For further information, contact Cooper Surgical by fax at 1-203-799-2002 or phone at 1-203-799-2000. |  |  |  | Worldwide Distribution -- United States, Canada, Australia, Belgium, England, France, Germany, Greece, Hungary, Iceland, Italy, Ireland, Latvia, Lithuania, Netherlands, Portugal, Spain, Solvenia, Sweden, Hong Kong, Republic of China, Taiwan, Angola, Cameroon, Congo, Croatia, Egypt, Ghana, Guam, Guineae Bissau, Indonesia, Kenya, Kuwait, Lebanon, Madagascar, Mali, Mauritania, Mozambique, New Zealand, Niger, Pakistan,Qatar, Russia, Sao Tome, Saudi Arabia, Senegal, Serbia, Singapore, South Africa, St Vincents &amp; Grenadines, Oman, Thailand, Turkey, UAE and Zimbabwe. | Terminated | April 17, 2012 | 931 |  | Not_Computer
Z-0119-2010 | 53492 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Nucleus Non-Magnetic Plug for removable magnet implants, REF Z50100, STERILE EO, Cochlear Limited. |  |  | September 30, 2009 | 2009 | Cochlear Americas Inc. | Cochlear implant device component was mis-labeled. | Labeling mix-ups | On September 28, 2009, each surgeon who received a non-magnetic plug after December 5, 2008 was sent a notification letter via courier. The letter described the mislabeling issue and actions for consignees. Customers were requested to return affected product (and the quantity discarded) to the firm and review patient records to verify the implantation of the non-magnetic plug and not the mislabled product (sterile magnetic). Direct questions about the recall by calling at 1-800-523-5798. |  | 10 | 10 units | Worldwide Distribution -- United States (AL, AZ, CA, GA, IA, NC, MA, MI, MN, MO, NY, OH, OR, PA, TN, TX and WA), Brazil, and Argentina. | Terminated | December 22, 2009 | 83 | 10 | Not_Computer
Z-0039-2010 | 53489 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Infant Manual Pulmonary Resuscitator with reservoir bag, face mask &amp; Gas sampling elbow. Product ID # 5801, 12-8713, 732-E, and 2206. |  |  | October 06, 2009 | 2009 | Unomedical | In some Manual Pulmonary Resuscitator units the duckbill valve and retention ring inside the patient valve housing may become dislodged and fall into the MPR bag/body. If this occurs, the MPR may lose its ability to generate the positive pressure necessary necessary to ventilate the patient. | Device Design | Unomedical issued Press on October 6, 2009. Letters were sent to customers on October 7, 2009 via UPS - return receipt. Letters explained the reason for recall and the potential hazard to patients. The letter contains a list of all affected lot numbers as well as a picture which shows that units with a clearly visible blue ring are NOT affected. Distributors are asked to examine inventory and quarantine all affected product and return the response form, even if no recalled product is in inventory. If the product was further distributed, the letter requests that a copy of the letter be sent to all customers and for returns, follow Distributor actions listed in the letter. Retailers and end users/Hospitals/EMT's are asked to perform a count of recalled product, segregate it and return the response form. All customers are asked to call 1-800-634-6003 where they will be assigned a RGA number and given instructions on how to return product.Contact information: 8:00AM - 5:30 PM Central Time Monday through Thursday and 8:00AM to 1:00 PM on Fridays, Central Time. |  | 200 | over 200 total lots | Worldwide distribution: USA, Argentina, Canada, Denmark, France, Greece, India, Italy, Japan, Mexico, Paraguay, Peru, Poland, Russia and Vietnam. | Terminated | August 10, 2010 | 308 | 200 | Not_Computer
Z-0104-2010 | 53583 | No | CFR | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pointe Scientific Liquid Glucose (HEXO) Reagent Set, Pointe Scientific, Inc, Canton, MI; Catalog #'s G7517-120,G7517-120, G7517-500, and G7517-1L; 8-G7517-120, 8-G7517-500, 3-G7517-L, 7-G7517-500 and 7-G7517-1000. (Clinical chemistry) |  |  | October 19, 2009 | 2009 | Pointe Scientific, Inc. | There is a loss of linearity. The product fails to maintain linearity specification of 500 mg/dL. | Process design | The firm issued a press release, and an amended press release, on 10/30/09. Consignees were notified by an Urgent: Medical Device Recall letter dated 10/19/09, which identified the problem; and, instructed testing laboratories to inform patients' attending physicians, and determine, with their input, whether confirmation of the previous test results will be required. Customers are to examine their inventory for any of the lot numbers listed on the letter; and, for any product that has been further distributed, those customers need to be identified and notified of this product recall by including a copy of the recall letter. Customers are to complete and return the Recall Return Response form and upon receipt, Pointe Scientific will replace any disposed of reagent. The recall was expanded to include two additional lots by letter dated 11/2/09. |  | 474 | 400 kits plus 59 liters and 15 vials. | Worldwide Distribution -- USA, Israel, Mexico, South Africa, and West Indies. | Terminated | September 27, 2010 | 343 | 474 | Not_Computer
Z-0121-2010 | 53329 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens ARTISTE syngo RT Therapist, Therapist Assist, Therapist Express Basic, and Therapist Express Assist Systems, manufactured by Siemens Medical Solutions, Concord, CA |  |  | September 10, 2009 | 2009 | Siemens Medical Solutions USA, Inc | A failed data transfer is possible that may result in data not being recorded in the interfaced system. Siemens Medical Solutions, Inc. USA became aware of this potential issue with its ARTISTE Linear Accelerator's Syngo RT Therapist v4.1 Workspace (containing the Siemens brand of Verify and Record application called PRIMEVIEW) and the LANTIS Oncology Information System product. | Software design | The firm issued a TH015/09/S Safety Modification syngo RT Therapist 4.1 Customer Advisory letter, dated September 10. 2009. the letter stated the reason for modification, the affected systems, working time, deadline, reimbursement, and reporting. The firm is preparing a software upgrade for the affected products to correct the problem. The letter was sent via certified mail. |  | 97 | 97 units | Worldwide Distribution -- USA, Spain, Germany, Poland, Australia, Malaysia, New Zealand, the UK, India, Ireland, Egypt, Italy, Belgium, Republic of Korea, and Canada. | Terminated | December 06, 2010 | 452 | 97 | Software | Device Operation | Software update | Software Update
Z-0087-2010 | 53574 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Cordis CROSSOVER. Cordis CROSSOVER Sheath Introducer. Thomas Medical Products, Inc. 65 Great Valley Parkway, Malvern, PA 19355 USA. |  |  | October 15, 2009 | 2009 | Cordis Corporation | Cordis has received complaints of stretching or fracture of the sheath shaft during use. | Component design/selection | A Medical Device Recall letter dated October 19, 2009 was issued to consignees by sending an overnight letter to each account. The letter described the affected device, issue and actions for customers.A Cordis Corporation representative will follow-up and contact consignees as necessary to facilitate obtaining signature, faxing the acknowledgement form to the Cordis Corporation, collecting and returning units. A press release from the Cordis Corporation announced a nationwide recall of the CROSSOVER Sheath Introducer available on (http://www.webwire.com/ViewPressRel.asp?aId=106460).In addition to the local sales representative, customers may contact the special Customer Service number at 1-866-854-1672, Monday through Friday from 8:30 a.m. to 8:00 p.m. Eastern Standard Time, if you have any additional questions relative to product return or alternate product. |  | 2362 | 2,362 units | Worldwide Distribution -- United States, Belgium, Switzerland, Germany, Finland, France, Italy, Netherlands, Sweden and United Kingdom. | Terminated | August 12, 2013 | 1397 | 2362 | Not_Computer
Z-0124-2010 | 53142 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CIVCO MTAPSID2932 Type-S Thermoplastic, Disposable nonsterile (48.3 cm) IMRT Style 29 thermoplastic (3.2 mm), Ref. MTAPSID2932, disposable frame mask, 5 per box. The firm name on the label is CIVCO, Orange City, IA. |  |  | August 31, 2009 | 2009 | Med Tec Inc | Some of the packages contain the incorrect mask style. | Process control | The recalling firm initiated their recall via letter dated 8/27/09 issued certified mail for the domestic customers and e-mail for the foreign customer. The communication informs the customer of the reason for recall and offers to replace the product. For questions or for return of the mask, a phone number is provided. A response form was enclosed with the letter for completion to indicate the amount of affected product in the possession of the customer and whether or not they plan to return it. |  | 12 | 12 boxes | Worldwide Distribution -- USA, including states of FL, TX, and OH and country of Australia. | Terminated | March 07, 2012 | 919 | 12 | Not_Computer
Z-0085-2010 | 53576 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | EndoClamp aortic catheter label reading in part: "CARDIOVATIONS EndoClamp aortic catheter 65 cm,STERILE R, REF EC65, Edwards |  |  | September 24, 2009 | 2009 | Edwards Lifesciences Research Medical, Inc. | Spontaneous balloon rupture during use of the product. | Process control | Customers were notified by an Urgent - Product Recall letter, dated 9/29/09, and instructed to check their inventory, identify any unused product and return it to the company. Customers are to contact Customer Service Organization at 1-800-424-3278 to obtain a Returned Goods Authorization number, information concerning replacement product, and if they have any questions. |  | 2737 | 2,737 total | Worldwide Distribution -- US, Europe, Canada, Australia, and Thailand. | Terminated | April 25, 2012 | 944 | 2737 | Not_Computer
Z-0151-2010 | 50763 | No | EKC | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Neos Vanadium Excavator Part Number 591/0 |  |  | November 14, 2008 | 2008 | Sybron Dental Specialties | The tips of the Neos Vanadium Excavators may prematurely release from the handle during use, and if so, the tip may be accidentally swallowed. | Component change control | Sybron Dental wrote to the distributor by letter dated November 20, 2008 and posted Nov 24, 2008 titled "URGENTMEDICAL DEVICE RECALL," informing them of the recall, and the reason for recall. Sybron informed the distributor to return affected batches of product. Any product in stock with the mentioned batch number was replaced at no charge or credited to their account.The consignee was also asked to contact KerrHawe Customer Care at +41 91 610 05 05 to receive an RGA number. The affected product should be returned to KerrHawe at the following address:KerrHawe SAVia Strecce 4CH-6934 BioggioSwitzerlandPLEASE LABEL RETURNED PRODUCT "RECALLED PRODUCT RETURN". In addition, they asked for the completed RETURN FORM to be returned by fax to +41 91 610 05 14 |  | 135 | 135 | Worldwide distribution - Austria, Belgium, Denmark, France, Germany, Italy, Netherlands, Portugal, Spain, Switzerland, China, and the United States. One US consignee. | Terminated | August 09, 2012 | 1364 | 355 | Not_Computer
Z-0145-2010 | 53257 | No | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Power cord for Q2 Continuous Cardiac Output (CCO)/SO2 Computer, list #s 52235-04-01, 52235-04-03, 52236-04-05, 52235-04-07, 52235-04-11, AC cord manufactured by Electri-Cord, Westfield PA, Infusion pumps manufactured by Hospira, Morgan Hill, CA. |  |  | August 11, 2009 | 2009 | Hospira Inc | Fire/Shock hazard-- The power cord used in the device may cause sparks/flashes and poses a fire and shock hazard. | Component design/selection | Hospira initiated its recall on 08/14/2009. A nationwide press release was issued to AP, and a letter to consignees was sent via UPS. The firm will replace affected cords and effectiveness checks will be conducted through tracking of servicing. |  | 204999 | 204,999 power cords distributed for all devices | Worldwide distribution. | Terminated | April 04, 2012 | 967 | 204999 | Not_Computer
Z-0149-2010 | 53251 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Medical Stair-PRO Stair Chair with wheels and a track system, patient transport device, Model 6252. |  |  | September 02, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The front of the handle may become detached from the chair, which could result in a patient fall or an operator injury. | Device Design | Consignees were sent a letter titled URGENT MEDICAL DEVICE CORRECTION dated 9/9/09, advising them of the problem and to remove the devices from service until they have been repaired. U.S. customers were notified that Stryker will complete the correction while international customers were sent a kit with instructions for completion of the correction. International consignees were told that Stryker will complete the correction if the consignee does not feel comfortable in doing so. Questions or comments regarding this issue should be addressed to Sarah Marie Puterbaugh of Stryker Medical at (800) 869-0770 ext. 6902. |  | 165 | 165 | Worldwide distribution - Colorado, Ohio, Wisconsin, and Germany, Italy, Netherlands, Poland and Switzerland. | Terminated | February 16, 2011 | 532 | 230 | Not_Computer
Z-0139-2010 | 53301 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | John Bunn 1392 Series Jackson Tracheostomy Stainless Steel Tube. |  |  | September 14, 2009 | 2009 | Gf Health Products, Inc. | Incorrect material. Manufacturer changed the material of the trach tube from Stainless Steel to German silver/nickel, contrary to product specifications. | Nonconforming Material/Component | GF Health Products (Graham Field) notified consignees by certified US Mail on/about 09/14/2009. The distribution to the dealers includes the announcement of the recall to the dealer the letter to send to the end users and a Proof of Compliance card for the dealer to submit to Graham Field following completion of the recall instructions. Graham-Field instructed the dealers to return any impacted products on site, as well as any delivered by end-users, customers or other impacted third parties. Dealer will be compensated for executing the recall. Graham Field does not have direct access to end user names; however it plans to reach end-users by financially supporting a mailing from dealers to their individual customers notifying them of this recall. A follow up 60 day certified mailing will be performed for non-responders. The dealers will receive financial support for this second mailing. A secure third company will be contracted to hold for destruction all products affected by this recall. The method of destruction will be to cut the products into pieces, make them not usable. The FDA will be notified prior to the destruction of product and detailed logs and records will be maintained. Periodic effectiveness status report will be maintained for this recall. |  | 396 | 396 units | Nationwide distribution. | Terminated | December 30, 2010 | 472 | 396 | Not_Computer
Z-0134-2010 | 53177 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Stryker Imaging, Guardian Services Software, Part Number 0240-009-430 with Part Numbers 0240-020-013 and 0240-020-022 (Office PACS Power). |  |  | August 31, 2009 | 2009 | Stryker Imaging | Firm has determined Guardian Services does not contain 100% data backup for 12 OfficePACS Power customers on a daily basis. | Other | Stryker Imaging issued an "Urgent: Device Correction" notice dated August 18, 2009. Customers were notified of the defective product and informed that Stryker will be in contact shortly to coordinate the product correction. Actions to be taken during the interim were also detailed.For further information, contact Stryker Imaging at 1-972-410-7100. |  | 12 | 12 devices affected | Nationwide Distribution -- AR, TN, IL, GA, TX, WA, MI, PA, and MS. | Terminated | April 14, 2010 | 226 | 12 | Software | Device Operation | Prodcut correction | Repair
Z-0132-2010 | 52833 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT i2000 System and ARCHITECT i2000 SR wash buffer and wash zone mechanism grounding kit. |  |  | July 17, 2009 | 2009 | Abbott Laboratories, Inc | Under certain conditions, the product's wash buffer containing sodium azide can come in contact with the wash zone mechanism ground cables on either of the two wash zones. The copper in the ground strap may corrode and form an unstable chemical substance (metal azide) that may be sensitive to direct pressure and impact. | Other | A product correction notice was issued to consignees on July 17, 2009. Field representatives were instructed to visit consignees, inspect cables and replace them |  | 4526 | 4,526 | Worldwide Distribution | Terminated | January 26, 2010 | 193 | 4526 | Not_Computer
Z-0138-2010 | 52056 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Innova 2121-IQ/3131-IQ Biplane Cardiovascular Imaging System. |  |  | April 16, 2009 | 2009 | GE Medical Systems, LLC | Potential pump failure associated with the Coolix 4000 chiller of the Innova single plane and biplane systems that may stop the chiller from working. If a chiller failure occurs, the user will be informed with a message on the n-room monitor stating that three minutes are left before the X-ray is inhibited. This could lead to a delay in treatment and possible additional X-ray exposure and contra | Device Design | Consignees were sent on May 1, 2009 a GE Healthcare " Urgent Medical Device Correction" letter dated April 16, 2009. The letter was addressed to Hospital Administrator/ Risk Manger, Mangers of Radiology/ Cardiology, and Radiologists/ Cardiologists. The letter described the problem, Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. |  | 61 | 15 AND 46 UNLOCATED | Worldwide distribution: USA, ARGENTINA, BELGIUM, BRAZIL, BULGARIA, CHINA, CROATIA, EGYPT, FRANCE, GERMANY, HONG KONG, HUNGARY, INDIA, ITALY, JAPAN, REPUBLIC OF KOREA, LATVIA, LIBYA, MEXICO, NORWAY, POLAND, RUSSIA, SINGAPORE, SPAIN, SWEDEN, TAIWAN, TURKEY, UNITED KINGDOM, ROMANIA, ISRAEL, UKRAINE, YEMEN, BELARUS, and AUSTRALIA. | Terminated | December 03, 2015 | 2422 | 234 | Not_Computer
Z-0003-2010 | 52990 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx (M3535A/M3536A) Defibrillator/Monitor. |  |  | August 21, 2009 | 2009 | Philips Healthcare Inc. | incorrect default configuration settings, insufficientbattery power as the user may not be alerted that no battery is installed or that a low battery condition exists prior to disconnecting from AC or DC power | Process change control | Philips Healthcare issued a Urgent Voluntary Medical Device Correction Notification on August 21, 2009 advising users of the problem and Customers are asked to follow the "Action to be taken by customer/user" section of the UrgentVoluntary Medical Device Correction Notification. In this section, customers are given the instructions to correct their devices. For further information or support concerning this issue, please contact Philips representative at 1-800-722-9377. |  | 1523 | 1,523 units | NationwideCanadaAustralia, Austria, Beharin, BRUNEI DARUSSALAM,Chian,. Finland, France, Germany,Hong Kong, India, Indonesia, Ireland, Israel,Italy, Japan, Libya, Lithuania. Mexico, Nambia, Netherlands, Palestine, Philippines, Poland, Portugal, Qatar,Russia, Saudi Arabia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, UAE, UK | Terminated | April 25, 2017 | 2804 | 1523 | Other | Alarm/Message | Instructions | Safety Notice/Insructions
Z-0141-2010 | 53137 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cardinal Health, Craig Calibrated Cannula, overall length 5-3/8" (13.7 cm), catalog #OS5000-005, sold separately, or as a component of the Craig Vertebral Biopsy Set, catalog #OS5000 |  |  | August 28, 2009 | 2009 | Cardinal Health | The Craig Calibrated Cannula does not interface with the toothed cutter properly, which may prevent the toothed cutter from properly advancing through the cannula to the biopsy site. | Nonconforming Material/Component | Cardinal Health sent Urgent Product Recall letters dated 8/28/09 to all customers who received the product during the recalled period on the same day via UPS Next Day Air. The accounts were informed of the problem with the cannula, and were instructed to functionally test the cannula by passing the toothed cutter within it to ensure the cutter advances freely into the cannula. If any restriction is noticed, the product should be segregated and returned to Cardinal Health. The accounts were requested to complete and return by fax the enclosed acknowledgement form, indicating the number of affected cannulas being returned. Any questions were directed to V. Mueller Customer Service at 800-323-9088. |  | 19 | 19 units | Nationwide distribution. | Terminated | May 12, 2010 | 257 | 19 | Not_Computer
Z-0130-2010 | 53294 | Yes | LWS |  |  | Cardiovascular | PMA | This device is not approved in the United States. |  |  | September 09, 2009 | 2009 | Medtronic Inc. Cardiac Rhythm Disease Managment | A subset of active implanted Concerto CRT-D and Virtuoso ICD devices worldwide (estimated at 6,300) may not meet expected device longevity due to gradually increasing current drain caused by low voltage capacitor degradation. This issue may present in the affected devices as reaching the Recommended Replacement Time (RRT) earlier than projected. This issue does not compromise device functionalit | Nonconforming Material/Component | An "Important Medical Device Information" letter, dated September 2009 was hand delivered by Medtronic Representatives beginning 09/09/09. The letter describes the issue, identifies affected devices and provides recommendations for patients with devices in the affected subset.. A Patient letter was mailed to patients beginning 09/28/09.For additional information: Doctors are advised to contact your local Medtronic Representative or Medtronic Technical Services at 800-723-4636;Device tracking managers are advised to contact your Medtronic representative, or the Patient Registration Department at 1-800-328-2518 x41832; andPatients are advised to contact Medtronic Patient Services number at 800-551-5544 ext. 41835. |  | 439 | 439 | Worldwide distribution (except Idaho and Wyoming) including countries of Australia, Austria, Belgium, Canada, Czech Republic, Denmark, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Kuwait, Netherlands, Norway, Portugal, Saudi Arabi, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | May 29, 2012 | 993 | 6294 | Not_Computer
Z-0133-2010 | 53461 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | VARiS 6.2, 6.1 and 6.0 when used with LINK, Model numbers H88 and H85. Manufactured by Varian Medical Systems, Palo Alto, CA. |  |  | September 15, 2009 | 2009 | Varian Medical Systems Oncology Systems | Software issue: Under certain circumstances, when a therapy treatment field is transferred to VARiS Vision 6.2, 6.1 and 6.0 via LINK procedures, the rotation direction will be reversed. | Software design | Recall initiated September 15, 2009. An "Urgent - Field Safety Notice" letter was distributed to all affected users, with a description of the problem and corrective action steps. The letter was also distributed to Varian sales, marketing, and service organizations, informing them of the issue. For further information in the USA and Canada, contact Varian at 1-650-424-5731 or 1-888-VARIAN5 (888- 827-4265). In Europe, contact +41 41 749 8844. |  | 100 | 100 units | Worldwide distribution to 100 consignees throughout the US in IL, NY, MI, NJ, AL, TN, MD, OH, FL, and TX, and to Brazil, Chile, China, Czech Republic, Egypt, Germany, Hungary, Israel, Italy, Japan, Malaysia, the Netherlands, Romania, Slovakia, Slovenia, South Korea, Spain, Taiwan, Thailand and the United Kingdom. | Terminated | July 29, 2011 | 682 | 100 | Software | Device Operation | Instructions | Safety Notice/Insructions
Z-0108-2010 | 53509 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Synthes USA, TI Synex (TM) II Central Body 14mm - 19mm, Catalog number 04.808.001. |  |  | September 14, 2009 | 2009 | Synthes USA (HQ), Inc. | loss of device height | Device Design | Synthes USA issued an "Urgent: Medical Device Recall" letter to hospital accounts, Synthes personal and sales consultants dated September 23, 2009. Users were asked to take inventory, complete and return a Verification form and return all affected product to the firm.For further information regarding the recall, contact Synthes at 1-800-479-6329 or 1-800-620-7025 extension 5453. |  | 267 | 267 units | Worldwide Distribution -- Including the United States and Canada. | Terminated | July 09, 2015 | 2124 | 2275 | Not_Computer
Z-0125-2010 | 53532 | No | MVN | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ClearSight(TM) 1-Day (ocufilcon D) Toric 30 daily disposable contact lenses for astigmatism --- Name and Address on Label: CooperVision Hamble, SO31 4RF, UK, and Scottsville, NY 14546 USA --- Each lens is supplied sterile in a blister containing buffered saline solution. The blister is labeled with the base curve, diameter, dioptric power, manufacturing lot number, and expiration date of the lens. ClearSight contact lenses are made from a material containing 55% water and 45% ocufilcon D, a random copolymer of 2-hydroxyethylmethacrylate and methacrylic acid. The lenses have a tint (VAT Blue 6) which is added to make the lens more visible for handling. |  |  | August 28, 2009 | 2009 | CooperVision Inc. | Thirty-eight lots of Ocufilcon D 55% 1-Day Toric Contact Lens with a modified packaging solution sold under the brand name ClearSight 1-Day Toric were inadvertently shipped into the US market prior to complete evaluation of the 510(k) by the FDA. | Other | Urgent: Medical Device Recall - Contact Lens letters, dated 8/28/09, Response Forms, along with reorder forms and postage paid return labels were sent to the consignees via first class mail with instructions for returning the affected product. Customers are to examine their inventory and customer records to determine if they have any of the affected product. The affected products are to be returned and CooperVision will reimburse for all postage associated with the return and replace the lenses free of charge. Questions should be directed to Bonnie Tsymbal at (800) 341-2020 ext. 3210. |  | 1440 | 1440 cartons | Nationwide Distribution | Terminated | June 12, 2012 | 1019 | 1440 | Not_Computer
Z-0140-2010 | 53482 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Olsen Medical, MIDAS TOUCH Âż NEEDLE ELECTRODE, WITH MODIFIED TIP. Model Numbers: 30-1011, 30-1012, 30-1013, 30-1014 and 30-6011. |  |  | September 24, 2009 | 2009 | Kentucky Packaging Service LP | THERE WERE INCONSISTENCIES NOTICED ON THE INSULATION OVERLAP. THE ELECTRODE HAS A LARGE DIAMETER SHAFT THAT TRANSITIONS TO A SMALLER DIAMETER SHAFT. THE 1ST LAYER OF INSULATION WHICH COVERS THE SMALL DIAMETER SHAFT OVERLAPS WITH THE 2ND LAYER OF INSULATION THAT COVERS THE LARGER DIAMETER SHAFT. ON SOME ELECTRODES IT APPEARS THAT THE OVERLAP IS TOO SHORT TO COVER THE TRANSITION WHICH MAKES THIS | Process control | Olsen Medical issued an "Urgent: Device Recall" letter to Consignees beginning September 24, 2009 . Customers were instructed to inventory their stock and for all remaining product immediately discontinue use and return. Consignees were also instructed to contact their customers and notify them of the recall.For further information, contact Olsen Medical at 1-502-772-4272. |  | 3617 | 3,617 pieces | Nationwide Distribution -- AZ, FL, IA, MD, MO, NM, OR, NV, TN, TX and VA. | Terminated | October 05, 2010 | 376 | 3617 | Not_Computer
Z-0169-2010 | 52639 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Byrd Screw, 2.0X6MM SD, Model Number: 01-0264 |  |  | February 19, 2009 | 2009 | Biomet Microfixation, Inc. | Byrd Screw, Self Drilling, may break during usage. BIOMET has determined that the parts were not correctly processed by the vendor , i.e. that the parts were not hardened in the final processing and that lead to weaker than acceptable material. | Nonconforming Material/Component | Biomet plans to send letters to customers on February 19, 2009. Customers are told to discontinue use of the item and return it to Biomet Microfixation. The firm plans to follow up with an e-mail to their customers. Questions pertaining to this issue, should be directed to Christy Cain at 1.800.874.7711 or 904.741.4400. |  | 22 | 22 | AZ, CA, GA, IL, IN, NM, NY, TX, VA, Puerto Rico and Canada. | Terminated | May 03, 2012 | 1169 | 22 | Not_Computer
Z-0170-2010 | 53128 | No | LGW |  |  | Neurology | PMA | Eon Rechargeable Implantable Pulse Generator (IPG) System |  |  | December 22, 2006 | 2006 | Advanced Neuromodulation Systems, Inc | Product's external charging system on occasion was failing to communicate with the Implantable Pulse Generator preventing re-charging of the pain relief system. | Device Design | 12/22/2006, Marketing Update 5015 was issued to field sales representatives providing information on improving communication between the Internal Pulse Generator and its charger. May 15, 2007, Marketing Update 5022 was issued to field sales representatives to reiterate avoiding contact between the charger antenna and conductive materials and place the preferred side of the charging antenna face-down over the implant. Troubleshooting Guide was created in January 2008 for field sales representatives and technical services departments regarding isolation of the communication issue to the external devices instead of the implanted component. March 25, 2008, the Clinician's Manual was updated adding a recommendation for a minimum depth of pocket for device implantation. |  | 24750 | 24,750 units total | Worldwide | Terminated | July 23, 2010 | 1309 | 24750 | Not_Computer
Z-0166-2010 | 53353 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Koodinat M/MP Patient Table used with Angiostar Model 5294009. |  |  | July 01, 2009 | 2009 | Siemens Medical Solutions USA, Inc | The table swivel locking mechanism may not completely engage adn could result in uncontrolled patient movement during transfer resulting in a patient fall. | Labeling False and Misleading | The recalling firm issued a Customer Safety Advisory Notice dated July 1, 2009 to affected customers via Update Instructions AX025/09/S to inform customer of the potential issues and provide instructions to prevent occurrence. The firm also provided an addendum to User Manual. Direct questions about the recall to Siemens Medical Solutions USA, Inc. by calling 1-610-219-6300. |  | 107 | 107 units | Nationwide Distribution. | Terminated | January 12, 2010 | 195 | 469 | Not_Computer
Z-0168-2010 | 53133 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Guardian Services, Part Number 0240-009-430 with Part Numbers: 0240-030-004 and/or 0240-030-002, 0240-030-008 (Net Practice). |  |  | July 31, 2009 | 2009 | Stryker Imaging | Product serving as data backup protection does not contain 100% data backup for all NetPractice customers on a daily basis. | Other | Firm notified consignees by an Urgent: Device Correction letter, dated July 31, 2009. The letter states the problem and that Stryker is upgrading its internal infrastructure to remediate the problem. The letter also provides Stryker's recommendations for reducing risk to patients. Technical instructions on the backup storage for the NetPractice are included with the letter. Stryker will not bill customers for the Guardian Service until the problem is remediated and Stryker notifies customers. If there are questions, customers should contact Stryker at 972-410-7100. |  | 55 | 55 devices affected | Nationwide Distribution | Terminated | January 08, 2010 | 161 | 55 | Software | Device Operation | Backup instructions/Upgrade infrastructure | Repair
Z-0159-2010 | 52987 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb ReNuÂż Multi-Purpose Solution, 12 fl.oz. (355 ml), Bausch &amp; Lomb, Incorporated, Rochester, NY 14609. |  |  | August 17, 2009 | 2009 | Bausch &amp; Lomb Inc | During routine testing the product failed to meet shelf life specifications. | Other | Consignees were notified by letter on/about August 17, 2009. All consignees were asked to review their inventory and return any of the affected product to Stericycle Pharmaceutical Services.For questions about the recall process or need additional shipping labels, please contact Stericycle Inc. at 1-888-345-83 16. For questions regarding this recall, please contact Bausch &amp; Lomb's Customer Resource Center at 1-800-828-6974. |  | 5000 | 5,000 bottles | Distribution: Nationwide, Trinidad, St. Croix, St. Thomas, Aruba and Bermuda. | Terminated | May 10, 2012 | 997 | 5000 | Not_Computer
Z-0002-2010 | 52834 | [6 empty] | Edwards Lifesciences Fogaty Occlusion Catheter Ref # 62080822F. |  |  | June 22, 2009 | 2009 | Edwards Lifesciences, Llc | Several complaints were the two parts of the tubular package became loose under the shrink wrap that holds the DFU on the packaging tube. | Packaging change control | Edwards initiated a worldwide recall via customer letter dated June 29, 2009, sent via Federal express on same day, requesting that customers cease using Fogarty Occlusion Catheter, Model No. 62080822F products from the indicated lots and that they return all unused products. Edwards requested that the customers provide written verification of stock and identify any unused products to be returned to Edwards Lifesciences LLC, Att: Field Action Coordinator, ODW Logistics, Inc., 13725 Pipeline Ave., Chino, CA 91710. Additionally provided Customer Service Tel # 800.424.3278. |  | 9548 | 9548 | AL AR AZ CA CO CT DC DE FL GA HI IA ID IL IN KS KY LA MA MD MI MN MO MS MO NC ND NE NH NJ NM NV NK OH OR PA RI SC SD TN TX UT VA VT WA WI WV WY, Puerto Rico Australia, Belgium, Brazil, Bulgaria, Canada, Canary Island, China, Czech Republic, Egypt, Finland, France, Germany, Great Britain, Greece, Hungary, Iceland, India, Ireland, Italy, Japan, Luxembourg, Netherland, New Zealand, Poland, Portugal, Rumania, Russia, Singapore, Slovak Republic, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, and Turkey. | Terminated | February 23, 2012 | 976 | 9548 | Not_Computer
Z-0150-2010 | 52791 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL 80 CO-OX SYSTEM, software versions 1.30 and 1.31 |  |  | February 20, 2009 | 2009 | Sendx Medical Inc | Inverness have recently become aware of a potential error contained in V1.30 and V1.31 software for the ABL 80 CO-OX platform. The error can occur if the analyzer is set to automatically upload patient records to Radiance or an HIS/LIS system. This can result in the reporting of incorrect results for one or more CO-OX derived parameters. Incorrect results can be displayed on the screen and the | Other | Radiometer sent out a letter to consignees informing them of public reason for recall, and telling them that they have provided a copy of the upgrade software and instructions to correct this potential problem immediately. Software upgrade instructions were included with the recall letter. Consignees were informed that if they preferred that their local service representative perform the upgrade, to follow the Instructions for disabling auto upload of patients records.SOFTWARE UPGRADE INSTRUCTIONS:The upgrade software must be installed via the Software Upgrade option under Menu&gt;Utilities&gt;Software Upgrade. It is not a "boo table" CD.~The Upgrade process will take approximately 10 minutes. The analyzer should be allowed to complete the entire process without user interaction. Do not attempt to turn the analyzer on or off manually at any time during the process.~The upgrade process will include the following general steps:1. Automatic power down and reboot of the analyzer2. File transfer process which includes a variety of progress bars to communicate the status3. Automatic power down and reboot of the analyzer a second time4. A dialog box on the main menu screen that states, "ABL80 FLEX software has been upgraded to version 1.32." This signifies a successful upgrade process.~If a solution pack and a sensor cassette are already installed on the analyzer, their installation status will not be affected by the software upgrade.~All stored data will be retained in all logsThey were also asked to complete and return the attached fax form, with their signature, to 1-800-736-0601.If their were any questions regarding the letter, or upgrade, to contact Techline at 1-800-736-0600 opt. |  | 317 | 317 | Domestically and Worldwide: Germany, France, Spain, United Kingdom, Poland, Austria, Italy, Tokyo, Australia, and DenmarkDomestically to OH | Terminated | July 16, 2012 | 1242 | 317 | Software | Output/Calculation | Software update | Software Update
Z-0161-2010 | 53136 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medcomp, 10F Bioflex Tesio Kit with cutting trocar. Catalog number BFR-6CTT 5 per box. |  |  | September 02, 2009 | 2009 | Medical Components, Inc dba MedComp | mislabeled as right are actually left | Employee error | Medcomp issued a "Product Alert" letter dated September 2, 2009 to all customers informing them of the problem and the need to return the product. For further information, contact Medcomp at 215-256-4201. |  | 100 | 100 trays | Worldwide Distribution -- United States (AR, FL, KS, LA, MO, PA, TN, and TX) and Austria. | Terminated | April 16, 2010 | 226 | 100 | Not_Computer
Z-0160-2010 | 53134 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | Guardian Services, Guardian Software, Part Number 0240-009-430 (Guardian Services) with Part Numbers 0250-003-016, 0250-003-017 and 0250-003-018 (OrthoPad) |  |  | August 03, 2009 | 2009 | Stryker Imaging | Guardian Services data backup does not contain 100% data for all OrthoPad customers on a daily basis. | Other | Firm notified consignees by an Urgent: Device Correction letter, dated 7/31/2009. The letter states the problem and that Stryker is upgrading its internal infrastructure to remediate the problem. The letter also states the firm's recommendations to reduce risk to patients. Stryker will not bill customers for the Guardian service until the problem is remediated and the firm notifies customers. If there are any questions, customers are to contact the firm at 972-410-7100. |  | 85 | 85 | Nationwide Distribution | Terminated | January 26, 2010 | 176 | 85 | Software | Device Operation | Upgrade infrastructure | Repair
Z-0164-2010 | 53453 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Siemens Biograph mCT-S 64, model number 10248672. |  |  | August 11, 2009 | 2009 | Siemens Medical Solutions USA Inc. | In the unlikely event that the device is modified to operate in the tilted position, the telescoping support mechanism will not engage properly. | Device Design | Siemens Medical Solutions USA, Inc. issued an "Urgent Field Correction - Recall" notice to consignees by letter dated September 3, 2009 informing them of the problem and that a Siemens service representative will be in contact shortly to service the affected product.For further information, contact Siemens at 1-800-888-7436. |  | 2 | 2 units | Worldwide Distribution -- United States (IL and OH), Germany, France and Hong Kong. | Terminated | December 22, 2009 | 133 | 3 | Not_Computer
Z-0171-2010 | 53531 | No | FHW | Penile inflatable implant. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | Coloplast Titan OTR Inflatable Penile Prosthesis,18cm scrotal Bioflex cylinder set w/OTR Pump, P/N QSR9181400. Sterile EO. Manufacturer: Coloplast A/S, 3050 Denmark. Distributed by: Coloplast Corp., Minneapolis, MN 55411 USA. |  |  | October 06, 2009 | 2009 | Coloplast Manufacturing US, LLC | Coloplast determined that a portion of Titan OTR pumps were manufactured using a process that where validation activities were not completed. The Titan OTR pumps were made to our standard product specifications and have passed all in-process testing to ensure mechanical reliability of these pumps. Coloplast is not recommending product explant because there are no identified risks since the prod | Process control | Coloplast issued a customer notification letter dated October 2009 describing the affected product.For further information, contact Coloplast at 1-800-533-0464. |  | 3 | 3 | Nationwide Distribution -- FL, GA, NC. | Terminated | December 12, 2011 | 797 | 3 | Not_Computer
Z-0173-2010 | 53535 | No | HET | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Gynecare Morcellex Tissue Morcellator; |  |  | September 28, 2009 | 2009 | Ethicon, Inc. | There is the potential for the duckbill seal to become dislodged. | Device Design | Recall notification letters titled "URGENT VOLUNTARY PRODUCT RECALL" were sent on 10/8/2009 via UPS next day. The letters describe the lots within expiration that are recalled. The motor drive unit is not included in the recall. Letters have photos and illustrations of the product which clearly show the location of the product code and expiration date. Recalled product is to be returned to Stericyle using an enclosed pre-paid authorized shipping label. Ethicon will notify customers when the product is available.Questions are to be directed to Ethicon at 1-888-276-5901. |  | 60930 | 60,930 in US. | Worldwide distribution - US and United Arab Emirates, Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, China, Czech Republic, Germany, Denmark, Egypt, Spain, Finland, France, United Kingdom, Greece, Israel, Italy, Jordan, Republic of Korea, Mexico, Malaysia, Netherlands, Norway, New Zealand, Puerto Rico, Reunion, Russian Federation, Sweden, Slovina, Turkey, Saudi Arabia, Venezuela and S. Africa. | Terminated | August 02, 2010 | 308 | 60930 | Not_Computer
Z-0174-2010 | 53564 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging XS VA60 and VA70 in combination with a NAS or HSM license |  |  | September 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Incomplete amount of image data; specifically, The image count in the series table of the online and archive database will display the full amount of images received at the Syngo Imaging XS as archived. However, the zip file on the NAS or HSM will contain fewer images-those which have been archived until the time of the error/interruption. | Software design | The recalling firm issued a "Customer Safety Advisory Notice" dated 9/14/09 to all affected customers via Update Instructions IM101/09/S. The notice informs the customer of the potential issue and provides instructions to avoid its occurrence. Questions regarding this issue are directed to the local SIEMENS Uptime Service Center. |  | 2 | 2 units | The product was shipped to medical facilities in NY and OH | Terminated | July 05, 2012 | 1009 | 2 | Software | Output/Calculation | Update | Other
Z-0165-2010 | 53545 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | * (Box label) SUCTION COAGULATOR 10FR QTY 25 ***** |  |  | September 30, 2009 | 2009 | ConMed Electrosurgery | One production lot of Suction Coagulators received ETO sterilization instead of gamma radiation sterilization. | Process control | Recalling firm notified all distributors by phone from 9/30 - 10/1/09. All customers were notified by letter titled "URGENT: DEVICE RECALL" dated 09/30/2009 on 10/2/2009. Customers were instructed to check their inventory for the affected product, complete the reply form and return affected product. For questions, please contact Damaris Velez or Jeff Dickinson of ConMed Electrosurgery at 800-552-0138, ext 5225 or email succoagrecall AT conmed.com. |  | 2700 | 2700 units | US only. CA, FL, GA, ID, IL, IN, KY, LA, MD, MI, MO, MS, NC, NY, OH, OR, PA, SC, SD, TX, VA | Terminated | March 16, 2010 | 167 | 2700 | Not_Computer
Z-0176-2010 | 53676 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | RESPIRONICS V60 VENTILATOR |  |  | October 26, 2009 | 2009 | Respironics California Inc | The recall was initiated to correct an issue with a component on the V60 Ventilator Power Management (PM) Printed Circuit Board (PCB). A component on the PCB has the potential to break off or become damaged if the PCB is removed from the ventilator. If this occurs, the ventilator may not properly switch between back up battery power and AC power. This can result in power loss during ventilation fo | Device Design | The recall notification was initiated October 26, 2009 with the firm mailing a Urgent Medical Device Field Correction letter to the affected customers. The letter informed the customers of the reason for the notice and how Respironics Field Service Technicians were going to correct this issue. The letter also gave the customers instructions to follow pending the completion of the field inspection/replacement.For questions or additional information contact Respironics US Customer Service at 1.800.345.6443, Option #5 then Option #1. |  | 524 | 524 | Nationwide and Internationally to :Libya, Hong Kong, France, Italy, India, Spain, United Kingdom,Germany, Netherlands and Saudi Arabia | Terminated | October 13, 2011 | 717 | 524 | Hardware | Device Operation | Replace | Remove or Replace
Z-0175-2010 | 53582 | No | FFK | Electrohydraulic lithotriptor. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific 1.0mm x 570 mm rigid pneumatic probe for use only with Swiss LithoClast Select System model no. 840-300K, Swiss LithoClast Ultra System model nol 840-700K, or Swiss LithoClast Ultra Upgraded System model no. 840-701U, compatible with 330mm Ultrasound Probes (REF 840-714 and 840-715), Rx, sterile, no detectable latex, Packaged and Distributed by Boston Scientific Corporation, Natick, MA; REF/catalog no 840-733. |  |  | September 17, 2009 | 2009 | Boston Scientific Corp | The label incorrectly reads "compatible with 330mm ultrasound probes (Ref 840-714 and 840-715)" instead of bearing the correct statement "compatible with 403mm ultrasound probes (Ref 840-716 and 840-7170." | Other | Consignees were notified by letter titled "Urgent Medical Device Recall-Immediate Action Required" dated 9/17/09 and instructed to cease using the product and to call the firm's customer service for return authorization. For additional information, call the firm at 866-868-4004. |  | 173 | 173 | Nationwide. | Terminated | October 08, 2010 | 386 | 173 | Not_Computer
Z-0177-2010 | 53536 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Trabecular Metal Shoulder, humeral stem, 12 mm stem diameter, 130 mm stem length, 42 degrees, sterile, Zimmer, Warsaw, IN, REF 4342-112-13. |  |  | September 17, 2009 | 2009 | Zimmer Inc. | Some of the devices may possess a taper which will not mate with the humeral head, as intended. | Other | Zimmer distributors were notified by email on 9/19/09 to locate and return the products. Letter dated 9/29/09 were sent to the Zimmer salesforce and distributors with a letter to be given to each affected hospital account to inform them of the problem. Zimmer accounts were again asked to locate and return the devices. Hospital accounts were requested to consider the available facts when monitoring potentially affected patients. |  | 12 | 12 | Nationwide. | Terminated | December 22, 2009 | 96 | 12 | Not_Computer
Z-0190-2010 | 52759 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Alphatec Spine, Inc., NOVEL SPINAL SPACER SYSTEM, Part Number(s): 64817-010, 64817-012, 64817-013, 64817-014, 64817-107, 64113-012, 64113-010, 64765-109, 64733-124. |  |  | May 13, 2009 | 2009 | Alphatec Spine, Inc. | Due to a deficiency in the process validation related to the machine that manufacturedthese parts, Alphatec would like to have the parts returned to them so that they may perform anadditional inspection of each implant. | Process control | Alphatec Spine notified consignees of the affected product via telephone. The firm requested that consignees immediately remove the affected devices in their possession and return to Alphatec Spine. For further information, contact Alphatec Spine at 1-800-922-1356. |  | 37 | 37 devices | Nationwide Distribution -- CA, NV, FL, CT, TX, IA and WI. | Terminated | November 13, 2009 | 184 | 37 | Not_Computer
Z-0179-2010 | 53139 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | System 1000 family of Hemodialysis Instruments, including the System 1000, TINA, AURORA and AltraTouch; Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A. |  |  | August 31, 2009 | 2009 | Baxter Healthcare Renal Div | When the main power is lost temporarily, a software anomaly prevents the instrument from automatically changing the ultrafiltration rate (UFR) to the minimum value as described in the Operator's manual. | Software design | Baxter sent an Urgent Device Correction letter dated August 31, 2009 to all System 1000 Series of Hemodialysis Instruments customers via first class mail on the same date, to the attention of the Hemodialysis Administrator. The letters informed the accounts that a software anomaly in software versions 4.B.0.9, 4.B.1.1 and 4.B.1.3 was discovered when the main power was lost temporarily during a simulated treatment. The anomaly prevents the instrument from automatically changing the ultrafiltration rate (UFR) to the minimum value as described in the Operator's Manual. The accounts were informed that instruments with software versions 4.B.0.9, 4.B.1.1 and 4.B.1.3 will be upgraded to version 4.B.1.4 software, provided the instruments have a 486 Motherboard or Single Board Computer, and that instruments containing versions of 4.A software are not affected. The accounts were requested to complete the enclosed Customer Reply form, indicating the instrument serial number and software version for each machine at the site, and if they would be requesting Baxter Technical Services to upgrade their instruments or have the software sent to them to make the upgrades, and fax the sheet to 1-847-270-5457. Any questions were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 4, option 1. |  | 8258 | 8,258 units | Worldwide Distribution -- United States (including Puerto Rico and the Virgin Islands), Algeria, Argentina, Austria, Bahrain, Bangladesh, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Honduras, India, Indonesia, Ireland, Italy, Jordan, Korea, Latvia, Lebanon, Lithuania, Malaysia, Mauritania, Mexico, Morocco, Nigeria, Oman, Panama, Paraguay, Peru, Philippines, Poland, Romania, Russia, Saipan, Singapore, Slovakia, South Africa, Spain, Sweden, Taiwan, Thailand, Trinidad, Tunisia, Turkey, Ukraine, United Kingdom, Uruguay and Venezuela. | Terminated | November 15, 2010 | 441 | 8258 | Software | Output/Calculation | Software update | Software Update
Z-0180-2010 | 53083 | Yes | LWS |  |  | Cardiovascular | PMA | Guidant LATITUDEÂż Communicator, model 6481, Boston Scientific 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. The LATITUDEÂż Patient Management system enables physicians to periodically monitor both patient and device status remotely. |  |  | August 07, 2009 | 2009 | Boston Scientific CRM Corp | The L ATITUDEÂż communicator (Model 6481) for a patient may not have been able to receive measurements from their weight scale or blood pressure monitor afer initial setup. The issue has been resolved. If any patient weight or blood pressure readings are stored in the patient's sensor, they will be uploaded to LATITUDE after the patient's communicator calls in to Boston Scientific's servers. It | Software design | Boston Scientific used phone calls as the delivery method for these communications. The communication began August 6, 2009 for the original 6 patients affected. Notifications via phone calls to additional patients affected began on August 17, 2009. Direct questions about the recall to the LATITUDE Customer Service by calling1-651-582-4000. |  | 97 | 97 | Nationwide Distribution -- AZ, AR, CA, CT, FL, GA, IL, IN, MA, MO, NV, NJ, NY, OH, OR, PA, SC, TX, UT, VA and WA. | Terminated | December 04, 2012 | 1215 | 97 | I/O |  |  | Other
Z-0189-2010 | 53494 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions USA, Inc., Siregraph CF X-Ray System, Model Number: 4466033. |  |  | September 23, 2009 | 2009 | Siemens Medical Solutions USA, Inc | possible injury to hand/fingers | Device Design | Siemens Medical Solutions USA, Inc. issued a "Customer Safety Advisory Notice" dated September 11, 2009 to inform customers of the potential hazard to patients when using the Siregraph CF system.For further information, contact Siemens at 1-610-448-4500. |  | 14 | 14 units | Nationwide Distribution -- FL, IL, MI, PA and TX. | Terminated | January 12, 2010 | 111 | 14 | Not_Computer
Z-0181-2010 | 53458 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Organogenesis, Inc., Apligraf (draftskin) packaging, Canton, MA 02021. |  |  | September 27, 2009 | 2009 | Organogenesis, Inc. | Unit contaminated with Staphylococcus epidermidis | Packaging process control | Organogenesis, Inc. initiated a recall notification dated September 28, 2009 via fax to Physicians asking for return of unused affected product. If the unit was used, Physicians are advised to use standard wound care precautions to assure the safety of the patient. Users were asked to acknowledge receipt by completing and returning the notification to the firm.For further information, contact Organogenesis, Inc. at 1-888-432-5232. |  | 60 | 60 units | Nationwide Distribution | Terminated | March 31, 2010 | 185 | 60 | Not_Computer
Z-0188-2010 | 53499 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | XTEN LED DF cupola, P/N: 567815999. |  |  | October 07, 2009 | 2009 | MAQUET Inc. | Several reports have been received regarding one model of spring arm (Acrobat 200 series) in combination with surgical light from various Manufacturers. All incidents pertain to a crack on the front joint of the Acrobat arm. | Device Design | Urgent Device Correction letters, dated October 7, 2009, were sent out by FedEx. The letter identified the affected products and asked customers to call 1-888-627-8383 #1, if they have any questions. The letter also provided a Technical Description along with a Customer Fax Response form. The letter stated the potential hazard and recommended precaution. If customers need additional info or have any questions they should contact their local MAQUET representative. |  | 346 | 346 | Nationwide Distribution. | Terminated | June 04, 2012 | 971 | 940 | Not_Computer
Z-0191-2010 | 53581 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet 5.5 mm peek Allthread PEEK Knotless Anchor, knotless, sterile, Biomet sports Medicine, Inc., Ontario, CA; REF 904842P. |  |  | September 09, 2009 | 2009 | Biomet, Inc. | These anchors may be difficult to remove, or may not detach from the inserter shaft. | Device Design | Biomet issued an "Urgent Medical Device Recall Notice" dated September 9, 2009 informing consignees of the affected product and requesting return of the device. Distributors were instructed to notify their customers copy of the recall letter and to locate and return the products from user facilities.For further information, contact Biomet at 1-800-348-9500 or 1-574-372-3983. |  | 894 | 894 | Worldwide Distribution -- United States and Belgium. | Terminated | January 19, 2011 | 497 | 894 | Not_Computer
Z-0105-2010 | 53550 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Tri-State Hospital Supply Corp., Centurion Healthcare Products Premie Packs, 50, sterile, Tri-State Hospital Supply Corp., Howell, MI; code no. LM110. |  |  | September 22, 2009 | 2009 | Centurion Medical Products | The pediatric tracheal tubes used in these kits were manufactured with an internal diameter smaller than indicated in the label, which could result in an inability to remove secretions and cause partial or complete obstruction of the airway. | Process control | The firm's sales representative was notified of the recall on September 22, 2009 via email and the sole customer was notified by letter dated September 23, 2009. All product was removed from the customer account on September 28, 2009 and returned to the recalling firm. |  | 650 | 650 | Illinois. | Terminated | August 03, 2010 | 315 | 1400 | Not_Computer
Z-0198-2010 | 52836 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer VerSys Hip System Femoral Head, 12/14 taper, 28 mm dia., neck length +0 mm, sterile, Zimmer, Warsaw, IN; REF 8018-28-02. |  |  | August 03, 2009 | 2009 | Zimmer Inc. | Manufacturing material may be present on the surface of the implant. | Manufacturing material removal | An Urgent: Device Recall letter, dated 8/3/2009, was issued to consignees. The letter identified the affected product, stated the reason for the recall, the clinical implications (risks to health), and actions to be taken. Customers are to locate any unused devices and quarantine them immediately, carry out a physical count of affected product and record the data on the Return Certification Form included with the letter, and fax a copy of the completed form back to Zimmer, Inc. at 574-372-4265. Customers are to return the recalled product along with the completed Inventory Return Certification Form. A "Dear Surgeon" letter will be issued to any hospitals where the devices have been implanted. |  | 104 | 104 | Worldwide Distribution -- USA, Brazil, Canada, and Switzerland. | Terminated | December 15, 2009 | 134 | 104 | Not_Computer
Z-0197-2010 | 53471 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Equalizer Balloon Catheter, Order No./REF 17-112 (M001171120). EQL/40/7/2/100. Sterilized using Ethylene Oxide. Single Use Only. Legal Manufacturer Boston Scientific Cork Ltd., Business and Technology Park, Model Farm Rd, Cork, Ireland. |  |  | September 29, 2009 | 2009 | Boston Scientific Corporation | Boston Scientific is initiating a Recall Removal of numerous batches/lots of Equalizer Occlusion Balloon Catheter because the sterile barrier in the packaging of the affected product may be compromised. The outer pouch seal may be breached. If a compromised seal is not detected and a sterility-compromised catheter is used clinically, there is a risk of adverse health consequences. | Other | An Urgent Medical Device Recall Removal letter dated 09/29/2009 was sent to affected facilities and addressed to Risk Manager / Field Action Contact, beginning 09/29/2009. The letter explained the issue and identified product. Customers were asked to identify affected product within their inventory, segregate it and immediately return it to Boston Scientific. Customers will receive replacements for all recalled product that is returned. Customers were also to fax back to Boston Scientific the Reply Verification Tracking Form. Questions regarding this Recall Removal are directed to the local Sales Representative. |  | 9800 | 9800 total (2373 US and 7427 OUS) | Worldwide. US distribution includes: AL, AR, AZ, CA, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT. NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and Washington D.C. and Puerto Rico. Worldwide distribution includes: Algeria, Argentina, Austria, Belgium, Brazil, Bulgaria, Canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Lithuania, Luxembourg, Malaysia, Morocco, Netherlands, Norway, Poland, Portugal, Qatar, Saudi Arabia, Serbia, Slovenia, South Korea, Spain, Sweden, Switzerland, Thailand, Tunisia, Turkey, United Arab Emirates. | Terminated | December 17, 2011 | 809 | 9800 | Not_Computer
Z-0260-2010 | 53517 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GEMINI PET/CT Systems Upper Patient Pallet for the GEMINI GXL, 16 Slice and 6 Slice scanning device configurations. Model Numbers: 4535-679-71891 and 4535-6797591, 510(k) #K051170. |  |  | January 26, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | Upper patient pallet component of the recalling firm's Gemini PET/CT system may be thinner than normally required, which may reduce the product's load bearing capacity, resulting in cracking under less than the designed full-load weight. | Process change control | The firm notified customers with an "URGENT - Field Safety Notice" which was sent to the user/customers via FedEx on 1/26/2009. The letter informs customers of this potential safety hazard; the actions to be taken by the user/customer in order to mitigate risk to patients while undergoing scans; and the actions planned by Philips Medical's Service Engineers in order to correct the problem. The notification refers the user/customers to the applicable sections of their Systems Operations and Systems Information Manuals for a review on procedures for safely loading; unloading; and monitoring patients before, during, and after scanning studies. The notice goes on to inform the customers that replacement parts for the upper patient pallet(s) are on order and that a Philips Healthcare Field Service Engineer will be visiting each customer site within two (2) months in order to replace each affected pallet. Lastly, the notification provides a toll-free telephone number and instructions which the customer can use to contact a local Philips representative if additional information, or assistance is required. The firm's Customer Care Center contact is 1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: Oncology. |  | 2 | 2 units | The recalled products were distributed to customers in the following states within the U. S.: OH, GA, MI, and CA and to customers overseas in: Germany, France Japan, and South Korea. | Terminated | November 24, 2009 | 302 | 11 | Not_Computer
Z-0262-2010 | 53578 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Intramedullary Skeletal Kinetic Distractor (ISKD) is an internal limb lengthening device consisting of a telescoping Internal Limb Lengthener, Locking Screws, Instrumentation and an external hand-held Monitor; catalog numbers: T10-215-265, T10-215-265NS, T10-255-305, T10-255-305NS, T10-255-335, T10-300-350, T10-300-350NS, T12-215-265, T12-245-295, T12-255-305, T12-255-305NS, T12-255-335, T12-300-350, T12-380-380, F12-255-305, F12-255-305NS, F12-255-335, F12-255-335NS, F12-300-350, F12-300-350NS, F12-300-380, F12-300-380NS, F12-345-395, F12-345-395NS, F12-345-425, and F12-345-425NS. |  |  | October 09, 2009 | 2009 | Orthofix, Inc | ISKD lengthening device has the potential for not pre-distracting. This may cause or contribute to device malfunction, and if not discovered prior to use, may require surgical re-intervention or removal. | Other | Firm notified consignees by letter sent via certified mail on 10/9/2009. Follow up is planned for nonresponders. Questions are directed to the firm at (800) 527-0404. |  | 539 | 539 units | Worldwide distribution - In the US: FL, MD, TN, OH, NY, CA, GA. Also Mexico, Italy, Canada. | Terminated | September 24, 2010 | 350 | 539 | Not_Computer
Z-0199-2010 | 53399 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Gender Solutions Natural-Knee Flex System, Prolong Articular Surface, congruent, Size 00,0, left, 9 MM height, use with baseplate size 00 or 0, sterile, Zimmer, Warsaw, IN; REF 00-5424-010-09 |  |  | October 08, 2009 | 2009 | Zimmer Inc. | The surgeon may have difficulty inserting the polyethylene into the tibial baseplate, which may result in intraoperative damage to the device. | Other | An "Urgent: Field Correction" dated October 8, 2009 was issued to Consignees. The notification letter described the affected product, issue, health risks and actions for customers to take. The letter notified consignees of a changed to the surgical technique, to include the use of a new peripheral rib impactor tool.Notifications of this field notification were sent to all affected direct accounts of Zimmer, Inc. Contact your Zimmer Sales Representative ifyou have any questions regarding this action. For shipping assistance, questions or assistance in notifying your accounts about the field correction please contact Zimmer, Inc. at 1-800-613-6131. |  | 26404 | 26,404 of all products | Worldwide Distribution -- United States, Austria, Belgium, France, Germany, Japan, Spain and Switzerland. | Terminated | May 16, 2011 | 585 | 26404 | Not_Computer
Z-0192-2010 | 53543 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Revolution Blood Pump, "revOlution Centrifugal pump Catalog# 050300000***Rx ONLY***SORIN GROUP ITALIA via Statale 12 Nord, 86 41037 MIRANDOLA (MO) (Italy) Distributed in the U.S. by: Sorin Group USA, Inc. 14401 W. 65th way Arvada (CO) 80004-3599" |  |  | September 25, 2009 | 2009 | Sorin Group USA, Inc. | A weakness in the plastic packaging tray could cause the tray to crack; compromising product sterility. | Packaging change control | Customers were notified by an Urgent: Field Safety Notice, dated 9/25/09, and instructed to: check their inventory for the affected product, complete the customer response form and return or visually inspect all affected product before use. For additional information contact Sorin Group Customer Service at 1-800-650-2623. |  | 808 | 808 units | Worldwide Distribution -- US, Canada, Japan, and Europe. | Terminated | December 14, 2009 | 80 | 1064 | Not_Computer
Z-0259-2010 | 53603 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Medtronic Integrated Power Console System, IP X1 |  |  | July 17, 2009 | 2009 | Medtronic Xomed, Inc. | An error code (Code 13) is displayed on some of the consoles when a Medtronic Stylus or EHS handpiece is connected to the IPC indicating that the handpiece has not been recognized by the console. | Software design | A Product Correction Notification was issued to each consignee on June 15, 2009. Medtronic sales representatives will visit each facility to advise consignees of the issue and to install a software upgrade that will correct the handpiece recognition issue on the console(s). |  | 493 | 493 | Worldwide Distribution | Terminated | January 20, 2010 | 187 | 493 | Software | Alarm/Message | Software update | Software Update
Z-0268-2010 | 52546 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GMD UNIVERSAL URINARY INCONTINENCE SLING, Rx ONLY Product Code 1020-18. |  |  | June 18, 2009 | 2009 | Generic Medical Devices | Product that is approved only for use outside the U.S. was shipped to two customers in the U.S. | Employee error | On June 17, 2009 the recalling firm called the consignees and arranged for the unused product to be returned. On July 13, 2009 the firm sent a notification letter to the consignee that had implanted 1 of the devices. |  | 15 | 15 devices | USA - NJ, and PA | Terminated | March 14, 2012 | 1000 | 15 | Not_Computer
Z-0264-2010 | 53131 | No | LYJ |  |  | Neurology | PMA | Cyberonics VNS Therapy Programming M250 Programming System, model number 250, version 4.6, 6.1, 7.0 and 7.1. The Model 250 Handheld software is provided on flashcards packaged in software envelopes or pre-installed on handheld computers. |  |  | July 24, 2009 | 2009 | Cyberonics, Inc | Some VNS Therapy System replacement Demipulse generators reporting low lead impedance readings. In rare instances, a system diagnostic test using Model 250 Programming Software (versions 7.1 and earlier) may report "Lead Impedance: OK" when a short-circuit condition exists. | Device Design | Cyberonics sent an Important Safety Notice on July 24, 2009, to all physicians explaining the event, risks presented and recommended actions to take. Labeling updates are being developed concurrently with the notification. The firm will mail or hand-deliver the notifications via firm representatives. |  | 9500 | 9,500 programming devices | Worldwide: USA, Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Luxembourg, Malaysia, Mexico, Morocco, Netherlands, Norway, Poland, Portugal, Puerto Rico (US Territory), Russia, San Marino, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | September 26, 2012 | 1160 | 9500 | Software | Output/Calculation | Labeling update | Other
Z-0266-2010 | 53497 | No | CCT | Laryngotracheal topical anesthesia applicator. | Anesthesiology | Anesthesiology | 510(k) | Nasal Mucosal Atomizer Device with syringe, LATEX FREE For Single Patient Use Only Packaged Clean -Non-Sterile- MADE IN CANADA. |  |  | March 05, 2009 | 2009 | Sharn, Inc. | Sharn Anesthesia Inc., Tampa, FL received reports of leaking containers following use of their Sharn Anesthesia Microblast nasal and pediatric laryngeal atomizer device. | Device Design | Recalling firm is contacting all direct accounts (distributors &amp; user facilities) by telephone and letter. They request completion and return of the of the Return Goods Form and of any affected product by UPS using the firm's account number. |  | 550 | 550 | Nationwide Distribution | Terminated | July 13, 2012 | 1226 | 3225 | Not_Computer
Z-0269-2010 | 53496 | No | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Norian CRS Rotary Mixer Cement 10cc sterile |  |  | September 14, 2009 | 2009 | Synthes USA (HQ), Inc. | Non-conforming material - increased set time. Potential clinical ramifications associated with increased set time and decreased strength include extrusion of unset material and fragmentation of cured material. The risk of Norian fragmentation post cure increases with a larger defect and a thicker layer of product.Customers were advised to call Synthes at 1-800-479-6329 to obtain a Return Autho | Nonconforming Material/Component | The firm issued an Urgent Medical Device Recall letter dated 9/17/09 via USPS certified mail. The letter informed customers of the problem and the need to return the product. Sales Consultants that were shipped the device were notified by teleconference on 9/14/09. Consultants were instructed to go to their accounts to remove the product from inventory. Customers were advised to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Questions are directed to Synthes at 800-620-7025 x 5453 or 610-719-5453 or to your Synthes CMF Sales Consultant. |  | 15 | 15 units | The product was shipped to medical facilities in CO, KS, MA, MO, NY, SC, TX, UT, and WA. | Terminated | March 08, 2010 | 175 | 15 | Not_Computer
Z-0263-2010 | 53549 | No | DZI | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | Stryker Instruments 2.1 mm Cross-Cut Fissure Bur, carbide extended, sterile; REF 2296-101-621. |  |  | October 05, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | The cross-cuts are missing on the device, which will result in poor cutting and overheating and/or smoking of the device. | Vendor change control | Consignees were notified by letter dated 10/5/09 and requested to return the devices to Stryker. The international account was notified by email. |  | 93 | 93 | USA (California, Georgia, Louisiana, Ohio, Wisconsin) and Brazil. | Terminated | November 30, 2009 | 56 | 93 | Not_Computer
Z-0270-2010 | 53563 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging XS VA60 and VA70 in combination with JPEG Compression. |  |  | September 29, 2009 | 2009 | Siemens Medical Solutions USA, Inc | Distorted image. The "Distorted" image cannot be used for diagnostic findings, and, depending on the customer scenario, a repeated exam could be necessary if the image cannot be reported on any other workplace or recovered from another source. | Software design | The recalling firm issued a Customer Safety Advisory notice dated 9/14/09 to all affected customers via Update Instructions IM100/09/S. The notice informs the customer of the potential issue and provides instructions to avoid its occurrence. To avoid the potential risk of the issue, the local compression function can be disabled as described in the letter. Depending on the image data, the disabling of compression will result in increased storage usage.The issue will be resolved in the release of version VA70B.For all potentially affected customers with "JPEG Compression" license an update instruction will be released which defines the necessary corrective actions in case of "Distorted" images. The update instruction will be performed on request by the customer service group. Further questions regarding this issue are directed to the local SIEMENS Uptime Service Center or +49 (9131) 84-2056. |  | 28 | 28 units | The product was shipped to medical facilities in CA, FL, GA, IL, MD, MI, MO, NE, NC, NY, OH, PA, TN, TX, and WV. | Terminated | July 05, 2012 | 1010 | 58 | Not_Computer
Z-0286-2010 | 50135 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Respiratory Gating System, Catalog #NPTA540, Model #4535 674 26701, which includes: Pulmonary digital assembly kit and Pulmonary toolkit oncology Pro available on the following GEMINI PET/CT systems: GEMINI GXL 16 slice configuration, Model Number: 4535 679 75691. |  |  | November 24, 2008 | 2008 | Philips Medical Systems (Cleveland) Inc | A leak between the Tube Interface and the Outlet Tube of the transducer for the Respiratory Gating System component may result in a failure of the device to produce respiratory correlated images. | Process design | Philips Medical Systems Nuclear Medicine division initiated this recall by sending a 'Customer Advisory Notification to their customer accounts on 11/24/2008. This notification informs the customer of the potential for the device equipped with the Pulmonary Gating Option to develop an air leak which could result in the device being unable to produce the desired respiratory correlated images. The notification provides field test instructions and recommends that the customer test the bellows on each day of use in order to assess that the unit is functioning properly. If the bellows device fails the test the recalling firm recommends that the unit not be used and that the recalling firm be contacted via telephone or E-mail. The firm intends to conduct a field correction of the affected units in which their Field Service Engineers will install upgrade kits on the affected devices tentatively scheduled to begin in September, 2009. |  | 20 | 20 units | Worldwide distribution: USA, Puerto Rico, Canada, Belgium, China, Denmark, Egypt, France, Germany, Hong Kong, India, Japan, Netherlands, Pakistan, Philippines, Poland, Saudi Arabia, Singapore, South Korea, Spain, and Sweden. | Terminated | September 24, 2010 | 669 | 71 | Not_Computer
Z-0277-2010 | 52229 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker Medical transport stretcher, Stryker Medical, Portage, MI; Model 734. |  |  | August 10, 2009 | 2009 | Stryker Medical Div. of Stryker Corporation | The brake/steer pedal may seize, resulting in the brakes not working. | Nonconforming Material/Component | Consignees were notified by letter dated August 10, 2009 that Stryker will contact them to upgrade the stretchers within 6 months. The recall was expanded in scope to include additional serial numbers by a second letter dated September 17, 2009, which was issued to additional consignees who had not been previously notified of the recall, as well as the consignees who had been issued the August 10, 2009 letter. |  | 241 | 241 | Nationwide, Argentina, Australia, Brazil, Canada, Chile, China, France, Germany, Greece, India, Italy, Japan, Korea, Latin America, Malaysia, Mexico, Netherlands, Philippines, Poland, Portugal, Singapore, South Africa, Spain, Switzerland, Taiwan, and United Kingdom, | Terminated | October 29, 2010 | 445 | 17807 | Not_Computer
Z-0332-2010 | 53605 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 32CM ALTA LR GOLD STD KIT |  |  | October 23, 2009 | 2009 | Spire Biomedical, Inc. | There is a possibility that the catheter tubing could separate from the hub/bifurcation of the device. If this occurs, bleeding is the primary clinical complication likely to result from catheter tubing/hub separation. A second potential clinical complication is air embolism, particularly for the segment of the patient population that develops negative central venous pressure during normal respira | Process control | Spire Medical issued an "Urgent device Recall" letter dated 10/22/09 by FEDEX to end users. Users are requested to return product. For questions, contact Spire Biomedical Customer Service at 1-877-876-8635. |  | 666 | 666 units | Worldwide distribution - Nationwide, Government. Foreign: Canada, France. | Terminated | June 09, 2010 | 229 | 55183 | Not_Computer
Z-0273-2010 | 53615 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Magnum Disposable Biopsy Core Needle 16 gauge, Catalog Number: MN1610. |  |  | October 08, 2009 | 2009 | Bard Peripheral Vascular Inc | The recall was initiated after the firm determined that some Magnum Disposable Biopsy Core Needle 16 gauge lots may exhibit an exposed sample chamber, due to extended stylet length resulting from a loose hub connection. The exposed sample chamber will typically result in an inability to obtain the biopsy sample. In the majority of reported cases, there was no injury to the patient or user. When us | Process design | The recall notification was initiated by the firm issuing an urgent medical device notification to it's consignees on October 08, 2009 via USPS requesting them to check all inventory locations for the affected products, remove and return affected products to Bard. The consignees were instructed to complete and return a recall effectiveness check form. Bard provided this 800 number 1-800-321-4254 for any questions consignees may have. |  | 127169 | 127,169 units | Nationwide and internationally to CanadaMexicoAlgeria Andorra Austria Bahrain Belarus Belgium Croatia Cyprus Czech Republic Denmark Finland France French Guiana Germany Gibraltar Greece Hungary Ireland Israel Italy Ivory Coast Kenya Kuwait Latvia Lebanon Luxembourg Malta Martinique Mauritius Morocco Netherlands New Caledonia Nigeria Norway Oman Poland Portugal Republic of Kosovo Republic of Macedonia Reunion Romania Russia Saudi Arabia Senegal Serbia Seychelles Slovenia South Africa Spain Sweden Switzerland Syria Tunisia Turkey Ukraine United Arab Emirates United Kingdom AustraliaNew ZealandCanadaJapanBangladeshChinaHong KongIndiaIndonesiaKoreaMalaysiaPakistanSingaporeSri LankaTaiwanThailandVietnamArgentinaBahamasBarbadosBrazilChileColombiaDominican RepublicMexicoPeru andSuriname | Terminated | April 02, 2012 | 907 | 334701 | Not_Computer
Z-0272-2010 | 53586 | No | HGB | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Richard Wolf Hulka Clip; for single patient use only; a sterile tubal occluding device; 2 clips per procedure tray/pouch, 10 pouches per shipping carton; Richard Wolf Medical Instruments Corp., Vernon Hills, IL 60061; item 4986.90. |  |  | October 16, 2009 | 2009 | Richard Wolf Medical Instruments Corp. | The sterile pouch seal may fail to remain intact during handling rendering the Hulka clip non-sterile. | Packaging | Richard Wolfe sent recall letters dated October 16, 2009 to the direct accounts via certified mail on the same date. The accounts were informed that the lot of Hulka clips did not meet their sterile seal burst test requirements, and that the sterile seal may fail to remain intact during handling, resulting in the clips becoming contaminated or non-sterile. They were instructed to inspect their inventory for the affected lot, quarantine any product found and call Richard Wolf at 800-323-9653, ext 225 or 380 to arrange for the return of the affected product for replacement. The accounts were also requested to complete and return the enclosed response form to Richard Wolf, even if they do not have any of the affected lot in stock. |  | 150 | 150 cartons | Nationwide | Terminated | March 16, 2011 | 516 | 150 | Not_Computer
Z-0338-2010 | 51389 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom 100 Low Bed, Model P3930. AC-powered adjustable hospital bed. |  |  | April 20, 2009 | 2009 | Hill-Rom, Inc. | The siderails may exhibit false latching, or unintentional lowering, or may become inoperable. | Process control | Hill Rom notified consignees by recall letter dated 4/20/09 and advised them of the problem and requested them to locate the beds. Consignees with patients in need of siderails were directed to be moved to a different model bed with siderails and, if the patient did not need siderails; consignees were requested to remove the siderails, move the siderails to a stowed position, to tie the siderail to the bed frame or to tag the siderail as out of service. The letter stated the consignee would be contacted within 90 days by a firm representative to do a correction on all affected units. |  | 4017 | 4017 beds | Worldwide distribution: USA, Canada, France, India, Singapore and Spain. | Terminated | October 15, 2010 | 543 | 4017 | Not_Computer
Z-0365-2010 | 51876 | No | NSX |  |  | General Hospital | Contact ODE | PICIS, CareSuite Family- OR Manager (4235), Preop Manager (4250), Anesthesia Manager (4220), PACU Manager (4245) &amp; Critical Care Manager (1110), patient information systems. |  |  | April 08, 2009 | 2009 | Picis Inc. | Demographic data, most notably allergy and precaution data, can be overwritten with incomplete data or blanks by the interface between non-Picis clinical systems (your "HIS" or "CIS") and certain Picis applications, if that interface is not configured properly. | Software design | On April 8th 2009, Picis initiated a customer notification letter to the installed base of potentially affected users, explaining the potential source of patient data conflict and to provide support in choosing the appropriate configuration option and implementation. The notification letters included with this report package contain additional details regarding configuration options. |  | 325 | 325 | Worldwide distribution: USA, Australia, Austria, Belgium, Denmark, Finland, France, Germany, Italy, Malaysia, Netherlands, Norway, Portugal, Singapore, Spain, Sweden, and United Kingdom | Terminated | December 22, 2014 | 2084 | 325 | Other |  |  | Other
Z-0366-2010 | 53604 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | AGFA IMPAX HeartStation ECG Management System. |  |  | May 07, 2009 | 2009 | AGFA Corp. | HeartStation does not refresh content changes to the PDF reports that are printed emailed or faxed from the Index page. This may lead to a discrepancy in patient demographics between what is displayed in the HeartStation Client interface and the report delivered to the consumer. Questions are being directed to the local Agfa representative; the contact information was included in the letter. | Other | An Urgent Safety Notice was sent via FED-EX to all the sites with the affected units (a total of 6) on May 7, 2009. The three foreign accounts were notified by letter on May 14, 2009. The letter describes the potential issue and mitigation. Acknowledgment via FAX-Back that the information was received and understood was requested from the sites. Customers who use the print, email, fax functionality of this product were verbally notified of the defect and they have been advised of a workaround. In the notice letter, Agfa HealthCare is also providing customers with the recommended workaround. The workaround is to only print from the Viewer screen or to print the ECG once confirmed. The Viewer screen, however, does not allow the user to print batches of reports as does the Index screen. Questions about this matter are being directed to the local Agfa HealthCare representative. |  | 9 | 9 units | CA, KY, MA, NY, OH, SC, Belgium, Italy and Singapore | Terminated | July 27, 2010 | 446 | 9 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0370-2010 | 51933 | No | NQP | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Zimmer Dynesys top-loading cutter, spacer, Zimmer Spine, Minneapolis, MN; REF 07.01284.001. |  |  | May 28, 2009 | 2009 | Zimmer Inc. | The spacer cutter may jam during surgery, preventing use of the instrument and causing surgical delay. | Other | Consignees were notified by letter dated 5/20/09 titled "Urgent Medical Device Correction" that instructed customers that they may continue to use the recalled product until adequate replacements become available, but that they should always ensure they have a backup spacer cutter (of a different model which does not have this issue) available in case this model should jam during use. Users were also advised to ensure the spacer cutter is properly lubricated and is sharp. For questions, please contact Zimmer Spine at 1-800-777-7505. |  | 177 | 177 | Nationwide, Austria, Belgium, Canada, Italy, Netherlands, Singapore, Spain, Switzerland and United Kingdom. | Terminated | October 05, 2011 | 860 | 177 | Not_Computer
Z-0369-2010 | 50784 | No | LPB |  |  | Cardiovascular | PMA | Biosense Webster NaviStar RMT Diagnostic/Ablation Steerable Tip 8mm DS Catheters; Part Numbers: D-1259-01 &amp; D-1259-02 |  |  | October 02, 2008 | 2008 | BioSense Webster, Inc. | Biosense Webster initiated the recall because of the potential concern with one of the Printed Circuit Board (PCB) components used in the NaviStar RMT Catheters, which may cause the ablation catheter icon to shift on the map viewer when RF energy is being applied. Therefore the user, attempting a linear ablation in the location displayed on the CARTO RMT screen, may ultimately apply RF energy to u | Process control | The firm sent an "Urgent Medical Device Correction" letter to consignees on October 6, 2008 informing consignees. Users were advised to pay "close attention to the location of the ablation catheter icon on the map viewer when applying RF energy."If consignees suspect that they have a defective NAVISTAR RMT Catheter, they were instructed to notify Biosense Webster and to return the defective catheter for evaluation and refund/replacement.Customer were instructed to contact their local Biosense Webster representative for any questions or further information. No return reply was requested. |  | 2180 | 2180 Units | CA, CT, FL, GA, IA, IL, IN, KS, KY, MA, MI, MN, MO, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI | Terminated | November 23, 2009 | 417 | 4739 | Not_Computer
Z-0367-2010 | 53580 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Biomet K-Wire/Trocar Point, both ends/Dia. 045, length: 9 in., certified stainless steel, sterile, Biomet Orthopedics, Inc., Warsaw, IN; REF 35-351534. |  |  | August 06, 2009 | 2009 | Biomet, Inc. | Packages labeled as having a .045 inch diameter trocar point actually contain a diamond point with a diameter of .062 inches. | Process control | An "Urgent Medical Device Recall Notice" dated August 6, 2009 was issued to consignees. The letter described the affected product, issue and actions for consignees. Customers were instructed to complete and return the "FAX back Response Form" (fax number 1-574-372-1683) and all affected product in inventory to Biomet Orthopedics.Direct questions related to the recall notification by calling at 1-800-348-9500 or 1-574-372-3983, Monday through Friday, 8AM to 5PM. |  | 19 | 19 | United States (Florida and Pennsylvania). | Terminated | October 06, 2010 | 426 | 19 | Not_Computer
Z-0389-2010 | 52615 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Oxinium Femoral Head, 36 MM OD-3, 12/14 Taper, REF 71343603, Sterile, Smith &amp; Nephew, Memphis TN 38116 |  |  | January 23, 2009 | 2009 | Smith &amp; Nephew Inc | Radius at the end of the taper of the femoral implant is smaller than specification. This error could result in the tapers not locking fully or properly on femoral stems or in causing impingement and/or increased wear debris. The deficiency may not be noticed by the user. Impingement and/or wear debris may lead to revision. | Other | All affected Smith &amp; Nephew Sales Representatives were notified of problem and the recall overnight mail on 01/23/2009. All affected Smith &amp; Nephew International Distributors were notified via email and telephone on 01/23/2009. They were instructed to immediately quarantine the product for return. Further questions are directed to the contacts at Smith and Nephew listed in the letter. |  | 925 | 925 units | Nationwide; 17 international: Switzerland, Dubai, Brazil, Canada, South Africa, Australia, Belgium, Spain, Thailand Turkey, Costa Rica, Denmark, India, Italy, and France. | Terminated | February 07, 2012 | 1110 | 925 | Not_Computer
Z-0371-2010 | 53772 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stelkast Unicondylar Knee Femoral Component SC2462-1. |  |  | May 05, 2009 | 2009 | Stelkast Co | incorrect size markings on component | Nonconforming Material/Component | The recalling firm telephoned the medical supply companies on 5/5-19/09 to inform them of the problem and the need to return the products. If there are questions regarding this issue, please call 1.888.273.1583. |  | 20 | 20 units | Nationwide | Terminated | November 20, 2009 | 199 | 20 | Not_Computer
Z-0107-2010 | 53579 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Guider XF Softip Guide Catheter, manufactured by Boston Scientific Corporation, Maple Grove, MN, distributed by Boston Scientific Corporation, Fremont, CA. |  |  | October 06, 2009 | 2009 | Boston Scientific Corporation | Product degradation-- a polymer portion of the distal section of the catheter may degrade prematurely, leading to potential systemic embolization of particulates with ischemic complications. | Environmental control | The firm initiated its 'Safety Alert' on October 6, 2009. A notification letter was sent to each customer reminding them to follow storage and handling instructions. The firm has not ceased distribution. Direct questions about this notification to the Boston Scientific Corporation at 1-510-624-1734. |  | 94296 | 94,296 units | Worldwide Distribution -- United States, Argentina, Australia, Austria, Azerbaijan, Bahrain, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Jordan, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macau, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nicaragua, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad, Tobago, Tunisia, Turkey, Ukraine, UAE, Uruguay, Venezuela, Vietnam and Yemen. | Terminated | December 19, 2011 | 804 | 94296 | Not_Computer
Z-0122-2010 | 53614 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stryker Instruments Navigation System II- CART, Stryker Navigation, Kalamazoo, MI; Stryker REF 7700-100-000. |  |  | October 26, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | Multiple software-related malfunctions may occur. The device may suddenly stop working, the screen may freeze, the screen may only be updated very slowly, the unit may not respond to user input, or the unit may only respond to user very slowly. | Software design | The firm issued a press release on 11/4/09. Consignees were notified by an Urgent: Medical Device Recall Notification letter dated 10/23/09, which advised of the affected products, the reason for the recall, the recall plan, and the actions to be taken. Customers were to immediately review this notification and distribute to all departments affected. Also, customers were to immediately quarantine any affected product and label as recalled product; complete and sign the enclosed Business Reply Form and fax a copy to 866-521-2762. Upon receipt of the form, a Stryker Navigation Field Service Rep would contact them to schedule the repair of their equipment. All U.S. units were corrected by 10/30/09. |  | 21 | 21 | Worldwide Distribution -- USA, including states of Florida, Kansas, Kentucky, Massachusetts, Michigan, New Jersey, Pennsylvania, South Carolina, Texas, and Washington and countries of Australia, China, Germany, Japan, Korea, Switzerland, and the United Kingdom. | Terminated | December 31, 2009 | 66 | 23 | Software | Device Operation | Repair | Repair
Z-0413-2010 | 53619 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Merit Medical Systems, Inc., Convenience Trays, custom packaged under the following Merit MedicalÂż labels: (1) Prentice PICC Kit, Catalog number K12T-03698; (2) Insertion Kit, catalog number K12T-03050C; and (3) Angio Procedure Pack, catalog number K09T-09530A. |  |  | May 19, 2009 | 2009 | Merit Medical Systems, Inc. | Surgical gloves in hospital convenience kits may be degraded and tear during use. | Nonconforming Material/Component | Merit Medical Systems notified customers by phone, fax, or visit by sales representatives on 05/19/09. The notification advised of the convenience kit correction due to the recall of the Derma Prene Ultra Surgical Gloves, and that use of the product should cease immediately. Sales representatives were instructed to provide replacement gloves and "warning" stickers to be placed on identified kits in inventory. The warning stickers advise the customer to discard the surgical gloves in the kit and replace. Customers were requested to verify product disposition by signing a response form. |  | 378 | 378 units | USA distribution: nine hospitals in Texas and Colorado. | Terminated | December 03, 2009 | 198 | 378 | Not_Computer
Z-0414-2010 | 53475 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STERIS Harmony LL 500/700 Surgical Light, Vertical Spring-Loaded Arm. |  |  | September 11, 2009 | 2009 | Steris Corp | DrÂżger Medical AG &amp; Co., KGaA informed STERIS Corporation that they received (4) reported international cases where a welded seam at the joint of the spring-loaded arm of the DrÂżger SOLA 500/700 surgical light broke as a result of a crack that had slowly developed over time. The STERIS Harmony LL 500/700 surgical light is a Private Label version of the DrÂżger SOLA 500/700 surgical light. | Device Design | The firm sent Urgent Product Correction letters, dated Sept. 11, 2009, to their customers. The letter identified the problem and stated that Drager had supplied STERIS the replacement spring arms and that a STERIS Service Technician would perform the replacement. The letter also stated that the correction would be performed at no charge to the customers. According to the letter, customers would be contacted by their STERIS Service Representative to arrange a convenient time for replacement of potentially affected spring arms. In the meantime, the letter states that customers should use their surgical lights in accordance with the guidelines presented in the Operator's Manual. Questions or concerns regarding the letter should be directed to Barbara Rich at 440-392-7318 or your local STERIS account manager. |  | 388 | 388 units | Worldwide Distribution -- USA, including the states of AL, AR, CA, CO, CT, FL, GA, HI, IL, IN, KS, LA, MA, MN, MO, MS, NC, NH, NJ, NV, NY, OH, PA, RI, SC, TN, TX, UT, VA, WA, WI &amp; WY and the country of Canada. | Terminated | March 01, 2012 | 902 | 388 | Not Computer
Z-0374-2010 | 53688 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | SureSigns Vital Signs Viewer (VSV):Spot Check Vital Signs Monitor with NBP (options to include Predictive Temperature only; SPo2 only; SPo2 and Recorder only; SPo2 and Predictive Temperature only; or SPo2, Predictive Temperature, and Recorder |  |  | October 30, 2009 | 2009 | Philips Healthcare Inc. | Use of these devices at their maximum volume setting may result in the premature failure of the internal speaker | Software design | Customer notifications dated October 2009, in the United States will be delivered via UPS. Customer notifications outside the US will be managed by Philips representatives in each affected geography. Customers will be instructed to lower the volume setting while using existing units during the first notification. As part of the second notification to the customers, the field action correction will be a customer installable software upgrade. Customers may contact Philips Response Center with questions or concerns at 1.800.722.9377.On December 14, the firm issued the follow-up notification letter implementing the action plan to provide the software upgrade kit for installation and instructions. |  | 221 | 221 units | Nationwide Foreign: CanadaAfghanistan, Albania, Algeria, Angola, Argentina, Austria, Australia, Bahrain, Bosnia &amp; Herzegowina, Bangladesh, Belgium, Bulgaria, Bahrain, Bolivia, Brazil, Switzerland, Chile, China, Colombia, Costa Rica, Croatia/Hrvatska, Cyprus, Czech Republic, Germany, Denmark, Dominican Republic, Algeria, Ecuador, Estonia, Egypt, Spain, Finland, France, Gabon, Ghana, Great Britain, French Guiana, Guadeloupe, Greece, Guatemala, Hong Kong, Hungary, Indonesia, Ireland, Israel, India, Iceland, Italy, Iraq, Jordan, Japan, Jamaica Republic of Korea, Kenya, Kosovo, Kuwait, Lebanon, Sri Lanka, Lithuania, Luxembourg, Latvia, Libyan Arab Jamahiriya, Morocco, Norway, New Zealand, Republic of Moldova, Macau, Malawi, Mali, Malta, Martinique, Mauritius, Mexico, Malaysia, Namibia, Nepal, Netherlands,Nicaragua, Nigeria, Norway, Philippines, Pakistan, Oman, Panama, Peru, Poland, Puerto Rico, Palestinian Territories, Portugal, Qatar, Romania, Republic of Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, Spain, Sri Lanka, Switzerland, El Salvador, Tanzania, Thailand, Tunisia, Turkey, Taiwan, Uganda, Venezuela, Vietnam, South Africa, and Zimbabwe United Arab Emirates | Terminated | February 20, 2013 | 1209 | 49648 | Computer
Z-0397-2010 | 53694 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FlexShaft, product code FS14; FlexShaft II, product code FS214; FlexShaft III, product code FS312; FlexShaft III Sterilization Connector Guard, product code FS3SCG; FlexShaft II Reconditioned, product code FS214RC; Manual Release for FlexShaft II &amp; III, product code MR14. |  |  | September 23, 2009 | 2009 | Surgical Devices/Covidien | Lack of assurance that products were manufactured under GMPs. This action is being undertaken due to certain products potentially exhibiting compromised staple formation or cutting which in some instances may result in extended surgical time for the patient. | Nonconforming Material/Component | The recalling firm issued an Urgent Medical Device Recall letter dated 9/22/09 to all customers. The letter informed customers that Covidien acquired PMI, about the problem with the products, and about the need to return any products remaining inventory. Customers were asked to return the entire inventory of Power Medical Interventions product, regardless of its expiration date. This process can be initiated by contacting Power Medical Innovations customer service at 1-866-769-3763. |  | 1100 | 1100 units | The products were shipped medical facilities nationwide. The product was also shipped to distributors in Canada, Japan, Germany, Ireland, Italy, Spain, Taiwan, Belgium, Sweden, Netherlands, Israel, Portugal, India, Austria, Turkey, and Greece. | Terminated | March 08, 2010 | 166 | 240432 | Not_Computer
Z-0391-2010 | 53434 | No | JDG | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Howmedica Osteonics, Skirtless 5 40' LFIT V40 Femoral Head (32 mm +0 mm offset); hip prosthesis component. |  |  | September 04, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Packaging mix up between the Taper V40 Femoral Head and the LFIT V40 Femoral Head. | Process control | Stryker sent Notification letters and Product Accountability Forms via FedEx on September 4, 2009 with return receipt. Corrected letters were sent on September 9, 2009. |  | 10 | 10 units | Nationwide distribution. | Terminated | June 26, 2013 | 1391 | 17 | Not_Computer
Z-0383-2010 | 52806 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural Nail System, Antegrade Femoral Recon Module, standard, nonsterile, Zimmer, Warsaw, IN; REF 00-2490-001-08. |  |  | July 01, 2009 | 2009 | Zimmer Inc. | The guides may come apart during use, some guides may experience surface cracking, guide handles and modules may not target the femoral nail holes accurately and the instruments may not be able to be assemble together due to a change in the threading method used. | Other | On July 1, 2009, U.S. distributors were notified by e-mail to return two instruments. The recall was expanded with international accounts notified by e-mail on August 5, 2009, and U.S. accounts being notified by e-mail on August 6, 2009, to return 7 instruments. The scope of the recall was then expanded again to include 14 instruments and another letter dated September 14, 2009 to distributors and user accounts.The notifications described the affected products, problem, health risks and actions for consignees. Direct questions about the recall to Zimmer, Inc. by calling 1-800-613-6131. |  | 327 | 327 of total product | Worldwide Distribution -- United States, France, Germany, Italy, Switzerland and United Kingdom. | Terminated | October 05, 2011 | 826 | 327 | Not_Computer
Z-0415-2010 | 53299 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | AlMER ENDOFEMORAL 4MM OFFSET (PURPLE) |  |  | September 03, 2009 | 2009 | Smith &amp; Nephew, Inc., Endoscopy Div. | Device may have improper Laser Marking, 4mm device may be marked as 6mm and 6mm device may be marked as 4mm. | Labeling design | A notification letter was issued via Fedex on September 3, 2009 to domestic consignees and international consignees were notified via e-mail. The letter described the issue as the affected device may have improper laser marking. The 4mm devices are marked as 6mm and the 6mm devices are marked at 4mm. Customers were instructed to return the completed response form and affected product in inventory to Smith &amp; Nephew. Direct questions about the recall to your local Smith &amp; Nephew Sales Representatives or call 1-508-261-3655. |  | 134 | 134 individual units | Worldwide Distribution -- United States including Puerto Rico, Australia, Austria, Belgium, Canada, Costa Rica, Dubai, Finland, France, Germany, Great Britain, Netherlands, Norway, Portugal, Spain, Sweden, and Switzerland. | Terminated | May 08, 2012 | 978 | 134 | Not_Computer
Z-0455-2010 | 52886 | No | KWH | Toe joint polymer constrained prosthesis. | Orthopedic | Orthopedic | 510(k) | Tornier, LMP-50 (Lesser Metatarsal Phalangeal Implant), Size 50, Sterile EO, , Manufactured by Tornier, Inc, Edina, MN 55435. |  |  | April 06, 2009 | 2009 | Tornier, Inc. | Tornier, Inc. has received a report of a mislabeled product due to a product mix-up. The package label indicated that the product was a model LMP-50 however the product inside the package was Tornier FGT-50 implant (Primus Flexible Great Toe Implant With Grommetts). | Labeling mix-ups | Consignees were sent on April 6, 2009 a Tornier "Urgent-Product Recall" letter dated April 3, 2009. The letter was addressed to Healthcare professionals and distributors. The letter described the problem and product involved. Advised consignees to immediately discontinue use of the affected product and to return any unused affected product to their facility. Advised their distributors to notify all accounts. They also attached a questionnaire needed for inventory status. |  | 77 | 77 | Worldwide Distribution -- USA, including states of CA, FL, MD, MI, NJ, OH, and PA, and country of France. | Terminated | December 17, 2011 | 985 | 77 | Not_Computer
Z-0456-2010 | 53462 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity PACS (Picture Archiving and Communication System) RA1000 Workstation software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | September 25, 2009 | 2009 | GE Healthcare Integrated IT Solutions | There are potential safety issues, including misdiagnosis or delay in treatment, associated with the use of GE Centricity PACS RA1000 Workstation software that may impact patient safety. | Software design | GE Healthcare Integrated IT Solutions sent Urgent Medical Device Correction letters dated September 23, 2009 to their Centricity PACS software versions 2.1.x and 3.0.x customers, to the attention of the Hospital Risk Manager, the Hospital Administrator, the Head of the Radiology Department and the PACS Administrator, on 9/25/09. The letters advised the users of the patient safety issues associated with the use of GE Centricity PACS related to unintentionally adding new images to an approved exam; deletion of images from PACS if Cancel Checked or Cancel Selected is chosen in the CA Tool - Archive Queue; inadvertently getting two or more exams out of sync with the corresponding reports resulting in the RIS-IC reports being associated with the wrong patients exam; and when rejecting a duplicate image(s) from an unspecified exam that is still in the Arrived or Verified status, the duplicate image(s) is also rejected from the Dictated, Transcribed, or Completed status exam. The letter provided the users with safety instructions to follow until the software is updated. Any questions were directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. The accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479.Centricity PACS Software Versions 1.0.x and 2.0.x customers were sent separate letters explaining the issues and the safety instructions to follow, but there will be no software updates since those software versions are no longer being supported. Any questions should be directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 1-800-437-1171. |  | 1188 | 1,188 units | Worldwide Distribution -- United States including Puerto Rico and countries of Australia, Austria, Belgium, Brazil, Canada, Chile, China, Denmark, Ecuador, Egypt, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Korea, Kuwait, Lebanon, Malaysia, Malta, Mexico, Netherlands, New Zealand, Portugal, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Terminated | May 25, 2012 | 973 | 1188 | Software | Output/Calculation | Software update | Software Update
Z-0120-2010 | 53672 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Reinforced Catheter Introducer System, 5 per carton. The product is sold under the Cordis Crossover and Cardiovascular Systems Inc (CSI) ViperSheath labels. |  |  | October 21, 2009 | 2009 | Thomas Medical Products Inc | Sheath may break or fracture which may cause vessel damage or require surgical intervention to retrieve catheter pieces. | Component design/selection | An "Urgent Medical Device Correction" letter dated October 22, 2009 was issued to the direct accounts. The letter described the problem and actions for consignees. Consignees were instructed to cease distribution of the recalled product and to notify their subaccounts of the problem. In addition, customers were instructed to return the affected, unused product to Thomas Medical Product, Inc.'s contract warehouse, Inner-Space Warehouse, Inc.Contact your local sales representative for questions about the recall, or Customer Service by calling 1-866-446-3003, Monday through Friday from 8:30AM to 5:OOPM Eastern Standard Time, for questions about product return. |  | 17236 | 17,236 units | United States (FL and MN). | Terminated | November 08, 2010 | 383 | 17236 | Not_Computer
Z-0420-2010 | 53681 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | The Natus Blue Light Phototherapy is also called the "neoBLUE LED Phototherapy System". The "neoBLUE mini LED Phototherapy System" |  |  | October 01, 2009 | 2009 | Natus Medical Incorporated | The clamp connecting the neoBLUE mini phototherapy light to the accessory pole stand can sustain stress fractures or break when excessive force is used when tightening the clamp. No injuries have been reported. | Device Design | On 10/1/09 Natus notified domestic consignees by Urgent Product Advisory Notice letter that advised them of the issue, instructed them to inspect the clamps, and included pictures of the affected clamp. Consignees are instructed to return the recalled clamps for a new replacement to be shipped by Natus.International distribution partners were notified by letter on 10/1/2009 and instructed to identify and contact their consignees. Replacements will be sent to the distribution partner for their consignees. |  | 2127 | 2127 | Worldwide distribution: USA, Argentina, Australia, Austria, Belgium, Bermuda, Canada, Chile, Denmark, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iraq, Israel, Italy, Japan, Kenya, Kuwait, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Norway, Philippines, Poland, Portugal, PR, Qatar, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UAE, and UK | Terminated | September 22, 2011 | 721 | 2127 | Not_Computer
Z-0422-2010 | 53561 | No | DSQ |  |  | Cardiovascular | PMA | HeartMate 12 volt sealed lead acid Battery Clips that are components of the HeartMate II LVAS (PMA P060040) and the HeartMate XVE LVAS (PMA P920014) Systems. |  |  | October 14, 2009 | 2009 | Thoratec Corp | Patient reports alerted Thoratec that during the process of disconnecting power leads from the battery clips, the threaded connector of the battery clip housing was observed to be loose, and in some cases dislodged from the battery clip housing. The problem affects HeartMate 12V SLA Battery Clips including those used for the HeartMate II LVAS and HeartMate XVE systems. If not detected and handled | Other | Firm has issued a notification letter to consignees, dated October 14, 2009 asking them to return proof of receipt, and to remove all affected unused devices from stock, and asking ongoing patients to return to hospitals for inspection, replacement if needed, and retraining on the process of switching power sources and safely switching to a back-up system controller. Contact Thoratec Customer Services at 1-800-456-1477 for any questions or concerns about this issue. |  | 8194 | 8,194 battery clips | Worldwide - USA- AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI and the District Columbia. Internationally to Austria, Belgium, Bulgaria, Canada, Czech Republic, Denmark, France, Germany, Greece, Ireland, Israel, Italy, Lebanon, Lithuania, Luxembourg, Monaco, Netherlands, Singapore, South Africa, Sweden, Switzerland, Taiwan and the UK. | Terminated | April 18, 2011 | 551 | 8194 | Not_Computer
Z-0418-2010 | 53593 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | IntelliVue Clinical Information Portfolio (ICIP) Critical Care Release D, Version D.02. |  |  | October 21, 2009 | 2009 | Philips Healthcare Inc. | A software defect exists in the ICIP Release 0.02 scheduling function for order management such that if a user edits either the Schedule Time or the Stop Time for an existing order, the corresponding intervention(s) for the order may not be created. | Software design | Philips Medical Systems notified accounts by letter on October 21, 2009, of the release ICIP version D.03, which contains a software correction . A Philips representative will schedule the implementation of this upgrade. Please contact the Philips Customer Care Center with questions or concerns about this correction at 1-800-722-9377. |  | 53 | 53 units | MD, RI, WV, and TXAustraliaBelgiumDenmarkFranceItalyMalaysiaSingaporeSpainSwitzerlandTaiwanUnited Arab EmiratesUnited KingdomandIreland | Completed |  |  | 53 | Software | Device Operation | Software update | Software Update
Z-0443-2010 | 53606 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Stryker Thor Anterior Plating System; |  |  | October 02, 2009 | 2009 | Stryker Spine | Surgeons may apply an accidental axial torque to and/or over-angulate the All-In One Guide while implanting the screw. This may cause the screw not to seat properly in the locking ring and may cause the Thor screwdrivers to twist or break at the tip. | Device Design | Recall notification letters were sent to all Stryker branches, Hospital Risk Management and Surgeons on October 19, 2009 by Federal Express. Questions are to be directed to Tiffani Rogers, Regulatory Compliance Manager at 201-760-8206. |  | 262 | 262 | Nationwide | Terminated | January 23, 2013 | 1209 | 9550 | Not_Computer
Z-0421-2010 | 52445 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Kuschall K Junior Manual wheelchair |  |  | June 15, 2009 | 2009 | Invacare Corporation | The handle has been breaking at a calculated weight of 65 kg which may have serious consequences (broken bones) depending on the circumstances of the break (e.g., user on ramps, hills, etc.). | Device Design | The firm contacted their customers by telephone and then sent a recall notification letter to each customer on 6/15/2009. Questions are directed to email: Recalls@Invacare.com or phone: 800-333-6900. |  | 20 | 20 Wheelchairs | Product was distributed to the following states: CA, CO, MN, NJ, OR, PA, TX &amp; VA and one to Canada. | Terminated | November 21, 2011 | 889 | 20 | Not_Computer
Z-0457-2010 | 52720 | No | IZZ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | GE Medical Systems, COMPAX 40E X-ray Systems, Models; 2208203, 2208205, 2244282, 2244284, 2225863, 2225864 , 2225865, 2225870, 2225881, and 2225882. |  |  | July 30, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware that table lock safety features, which are provided with Compax40E X-ray systems, are not always being utilized and has resulted in unintended lateral tabletop movementthat may impact patient safety. | Other | Consignees in US and Canada were sent on July 31, 2009 a GE Healthcare "Urgent Medical device Correction" letter dated July 30, 2009. The letter was addressed to Hospital Administrators/ Risk mangers; Radiology Department Mangers; Radiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product correction, and Contact Information. The letter instructs customers that the table manual mechanical lock is provided with the system and should be engaged when loading or unloading patients to ensure the tabletop is stabilized and not moving. For the rest of the world, the customer letter was dated October 22, 2009 and was mailed October 30, 2009. |  | 1046 | 1046 | Worldwide Distribution -- USA, including states of AL, AK, AZ, AR, CA, CO, CT, DC,FL, GA, HI, ID, IL , IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH,NJ, NM, NY, NC, OH, OK OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WY and countries of UNITED KINGDOM, THAILAND, TAIWAN, SWITZERLAND, SPAIN, SAUDI ARABIA, PUERTO RICO, PORTUGAL NEW CALEDONIA, MARTINIQUE, KOREA, JAPAN, JAMAICA, ITALY, ISRAEL, HONG KONG, HONDURAS, GREECE, GERMANY, FRANCE, EGYPT, DOMINICAN, CHINA, CANDA, BRAZIL AUSTRIA, and AUSTRALIA. | Open, Classified |  |  | 1046 | Not_Computer
Z-0419-2010 | 53261 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Philips Lifeline Personal Help Button Pendant Classic and Slimline Style. |  |  | September 11, 2009 | 2009 | Philips Lifeline Inc | Pendant Personal Help Button neck cord may not break away and if not will present a potential choking risk. | Other | Philips Lifeline initiated Safety Notification letters on 9/21/09 to users and advised of the potential choking risk of the Pendant Personal Help Button The firm recommends that users consult their health care providers to determine which style of emergency button, including those that are worn on the wrist, is most beneficial for them. A postage-paid reply envelope had been provided. Questions should be addressed to Philips at 1-877-221-8756 (Monday  Friday, 8 a.m.  8 p.m. EST). Firm issued Press on 9/11/09. FDA Press issued 9/22/09 |  | 715000 | 715,000 | NationwideCanada | Terminated | October 02, 2012 | 1117 | 715000 | Not_Computer
Z-0458-2010 | 53801 | No | MUD | Oximeter. | Cardiovascular | Cardiovascular | 510(k) | Nonin Model 7600B cerebral oximeter pod. Used with the Nonin Model 7600 Regional Oximeter with Equanox Technology and Bluetooth Wireless Technology. Nonin Medical, Inc., 13700 First Avenue North, Plymouth, Minnesota 55441-5443 USA. |  |  | October 23, 2009 | 2009 | Nonin Medical, Inc | A malfunction was detected with one of the components that, under rare circumstances, may result in the overheating of the POD. This overheating, while unlikely to occur, has the potential to cause a burn if the POD is in direct contact with the patient. | Device Design | An Urgent: Medical Device Recall letter was emailed to customers on 10/23/2009. The letter described the issue and hazard and asked that the use of the device be discontinued and to immediately contact Nonin for return and repair/replacement instructions. Questions or concerns should be directed to John Dalpee at 763-577-3166. |  | 62 | 62 | Worldwide Distribution -- USA, including states of LA, MN, NC, and NY and countries of Canada, Germany, Italy, Netherlands, and United Kingdom. | Terminated | March 29, 2012 | 888 | 62 | Not_Computer
Z-0479-2010 | 52121 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cargille Modified Kingsbury-Clark Albumin Standards For in vitro Diagnostics (Set of 8 Standards: 5, 10, 20, 30, 40, 50, 75, 100 mg/ml); Catalog number: 52064; |  |  | July 07, 2006 | 2006 | Cargille-Sacher Laboratories Inc. | The 10 ml size standard of Cargille modified Kingsbury Clark Albumin Standards of lot Jan-12-05 appear shrunken or liquefied. | Storage | Cargill sent recall letters mailed July 7, 2006 by first class mail. |  | 5 | 5 kits | The product was distributed to 5 customers in OH, Washington, DC, PA, WI, and VA. | Terminated | December 02, 2009 | 1244 | 5 | Not_Computer
Z-0459-2010 | 53161 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | Bath Safe Adjustable Shower Seat, Without Back and Arms, Catalog number 727142001 |  |  | July 09, 2009 | 2009 | Nutech Manufacturing Corp | Safety chair used in shower/bath may not meet specifications for weight load. | Device Design | The subsidiary distributor, Maddak, Inc. notified customers by A "Voluntary Product Recall Notice", phone and e-mail on 07/09/2009. The own label distributor instructed customers to return their inventory of affected products received between May 1, 2009 and July 7, 2009. Distributors were instructed to immediately quarantine then return all affected inventory to Maddak and notify their accounts. Distributors were requested to instruct dealers to display a "storefront" notice and provide consumers with a customer response form. Distributors were also requested to forward a request to dealers to post a website notice if they have one. Recovered products will be replaced. |  | 3627 | 3,627 total units | Worldwide Distribution -- USA and Canada. | Terminated | May 29, 2013 | 1420 | 3627 | Not_Computer
Z-0467-2010 | 53588 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Scorpio CR Waffle Femur Lfit with posts; |  |  | October 02, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | It was determined that a change made in the dimensions of the collar component of the packaging resulted in the inner blister material being overstretched causing a potential hole in the sterile barrier system. | Other | An "Urgent Product Recall" letter was sent on October 2, 2009 by Federal Express to Stryker branches and Hospital Risk Management Departments in the United States as well as foreign consignees. The letter described the issue, potential hazard and actions for affected accounts. Consignees were instructed to examine their inventory and return recalled product as instructed in the notification letter to Stryker Orthopaedics. In addition, consignees were instructed to reconcile affected product by completing the attached Product Accountability form (with spreadsheet) and faxing to 1-201-831-6069. Direct questions to Stryker Orthopaedics by calling 1-201-972-2100. |  | 77 | 77 of total product | Worldwide Distribution -- United States (MA, NJ, IN, MI, CA and PA), Hong Kong, Germany, Netherlands, Japan, Korea, Switzerland, Malaysia, and the United Kingdom. | Terminated | April 05, 2010 | 185 | 77 | Not_Computer
Z-0478-2010 | 53867 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ProtÂżgÂż EverFlex" Self-Expanding Biliary Stent System (for United States distribution). 6 x 150mm, Model # PRB35-06-150-080. Sterile EO, ev3 Inc, 4600 Nathan lane North, Plymouth, MN 55442-2920, Rx only. |  |  | November 06, 2009 | 2009 | Ev3, Inc. | ev3 has determined that a lot of 150mm ProtÂżgÂż EverFlex Biliary Stent System contains a 100mm length Self-Expanding stent. The implantation of a stent 50mm shorter than expected may not cover the full length of the target lesion and may require subsequent intervention. | Labeling mix-ups | Two separate letters were sent to United States and international consignees on November 9, 2009. The letter to US consignees was addressed to Cath Lab Managers while the letter to foreign consignees was addressed to the Risk Manager or Cath Lab Manager. The letters described the affected product, gave an Issue Summary and Required Action, which advised consignees to locate and remove the recalled product. An Ev3, Inc. sales representative will contact consignees to arrange return of all unused product to Ev3, Inc. Direct questions regarding the recall to Ev3, Inc. by calling Customer Service at 1-800-716-6700. |  | 6 | 6 (United States) and 6 (Outside United States) | Worldwide Distribution -- United States (TX and FL), GERMANY, FRANCE and UNITED KINGDOM. | Terminated | March 08, 2012 | 853 | 6 | Not_Computer
Z-0480-2010 | 53612 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Maquet Accessories, Connection Bracket. Catalogue # 1180.36A0. |  |  | October 22, 2009 | 2009 | Maquet Inc. | Product monitoring has determined that the connection bracket shows potential mechanical play at the OR table interface for some procedures. This is due to the hole in the lock washer, which is responsible for locking the accessory to the OR table, being too large. | Device Design | Urgent Device Correction Letters were sent by Federal Express to affected US hospitals on October 22, 2009. The letter stated the technical description, potential hazard, and the recommended precaution. The letter asked customers to immediately verify whether their Connection Bracket is affected thru the FaxBack provided. Questions or need for additional information can be directed to customers' local MAQUET representative. |  | 9 | 9 units in US. | Nationwide Distribution -- USA, including states of SC and PA. | Terminated | May 18, 2010 | 208 | 9 | Not_Computer
Z-0486-2010 | 53555 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | E-Z CLEAN Non Stick Cautery Tip. Label reads in part: "MEGADYNE 11506 SOUTH STATE STREET DRAPER, UTAH 84020 USA***CAT/REF 0029M E-Z CLEAN NON STICK CAUTERY TIP***STERILE *** Expiration Date: 2104-07***" |  |  | October 13, 2009 | 2009 | Megadyne Medical Products, Inc. | Product labels reference an incorrect expiration date. Expiration date referenced on labels is 2104-07 instead of 2014-07. | Error in labeling | A notification letter dated October 13, 2009 was sent to the first level consignees. Consignees were instructed to examine their stock, remove affected product and contact Megadyne Customer Service at 1-800-747-6110 (United States) or 1-801-576-9669 for product return instructions. |  | 324 | 324 units | Worldwide Distribution -- United States (CA, FL, NE, LA, MT, OK, OR, PA, SC and UT) and New Zealand. | Terminated | February 17, 2010 | 127 | 324 | Not_Computer
Z-0482-2010 | 53673 | No | EBG | Temporary crown and bridge resin. | Dental | Dental | 510(k) | TempSpan Dual-Cure Temporary Crown and Bridge Material, PART NUMBER N69AD, |  |  | October 22, 2009 | 2009 | Jeneric/Pentron | Pentron Clinical Technologies LLC received two (2) complaints alleging that the TempSpan Dual-Cure Temporary Crown and Bridge Material was not setting.Investigation of product in inventory was conducted and it was confirmed that some of the bottles had been mispackaged with an incorrect catalyst. | Employee error | The recall customer notification was initiated on October 22, 2009 with the firm forwarding an urgent medical device letter with recall return form via USPS first class mail to the affected consignees. The letter informed the consignees of the reason for recall and instructed them to return any used or unused product of the affected lot. The firm provided an 800 number for the consignees to call to arrange for product return and replacement. |  | 58 | 58 units | Nationwide | Terminated | July 19, 2012 | 1001 | 58 | Not_Computer
Z-0481-2010 | 53748 | No | MRZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas ezManager Max Diabetes Management System version 2.0.9 |  |  | November 02, 2009 | 2009 | Animas Corporation | Printed hard copy of results is not the same as the system screens- carbohydrate to insulin ratio incorrect. | Software design | Animas issued an Urgent Medical Device Correction letter dated 11/2/09 to the pump user and another letter to the health care professional that prescribed the pump. The letters explain the issue. The patient letter is accompanied with a CD that contains the upgraded software and instructions to follow. |  | 20648 | 20,648 units | Nationwide distribution, and Canada, Cayman Islands, and Denmark. | Terminated | October 27, 2010 | 359 | 20648 | Software | Output/Calculation | Software update | Software Update
Z-0483-2010 | 53844 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis MP, Model number 5904466 |  |  | October 26, 2009 | 2009 | Siemens Medical Solutions USA, Inc | c-arm may tilt | Other | The firm has issued a Customer Safety Advisory Notice via Update Instructions AX051/09/S to affected customers. This letter informs customers of the potential malfunction and of the possible danger to patients, users or other persons. |  | 70 | 70 units | Nationwide | Terminated | January 26, 2010 | 92 | 113 | Not_Computer
Z-0491-2010 | 50685 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Zodiac Polyaxial Spinal Fixation System. A spinal interlaminal fixation orthosis is a device intended to be implanted made of an alloy, such as stainless steel, that consists of various hooks and a posteriorly placed compression or distraction rod. The device is implanted, usually across three adjacent vertebrae, to straighten and immobilize the spine to allow bone grafts to unite and fuse the vertebrae together. The device is used primarily in the treatment of scoliosis (a lateral curvature of the spine), but it also may be used in the treatment of fracture or dislocation of the spine, grades 3 and 4 of spondylolisthesis (a dislocation of the spinal column), and lower back syndrome. |  |  | November 03, 2008 | 2008 | Alphatec Spine, Inc. | Alphatec Spine recently discovered that this lot of screws has the incorrect length laser etched on one side of the screw body. As a result, they are conducting a voluntary recall of the Zodiac Titanium cannulated polyaxial screw 6.5 mm x 50 mm L. | Error in labeling | Alphatec Spine contacted consignees via phone on October 16, 2008 and requested that consignees immediately remove the Zodiac Titanium cannulated polyaxial screws (6.5 mm x 50 mm L) from the surgical set to help ensure that they will not be used in surgery. There is a script the company developed to assure consistent responses to customers, and a phone number for more information.The strategy calls for products to be returned to Alphatec Spine. The company will arrange to have Federal Express contact the consignees for pick-up of the screws. Follow up questions are directed to the company at 1-800-922-1356. |  | 56 | 56 | AZ, CA, FL, TX | Terminated | December 08, 2009 | 400 | 56 | Not_Computer
Z-0063-2010 | 53383 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Heartstart FR2+ automated external defibrillator Models M3860A and M3840A are equipped with an ECG Module. M3860A is under the Philips label and M3840A is under the Laerdal label. |  |  | October 01, 2009 | 2009 | Philips Medical Systems | A potential exists for a component failure, which, if it occurs, could result in the inability to deliver therapy. | Nonconforming Material/Component | On September 28, 2009 Philips issued a nationwide press release. On October 1, 2009 the firm issued notification letters to consignees arranging for the return and replacement of the affected units. The firm also set up a page on their website with a serial number look-up tool.For further information or support concerning this issue, customers are asked to contact Philips at 1.800.263.3342 or visit www.philips.com/FR2PlusAction for the serial number look-up list, answers to frequently asked questions, directions for the battery insertion test, and the replacement AED selection form. |  | 5808 | 5,808 | Domestic and Worldwide distribution to:United States, Canada, Australia, Austria, Bulgaria, China, Finland, France, Germany, Gibraltar, Great Britain, Hong Kong, India, Indonesia, Ireland, Italy, Japan, Jordan, Lebanon, Malaysia, Mexico, Netherlands, Norway, Portugal, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Arab Emirates. | Terminated | September 27, 2013 | 1457 | 5808 | Not_Computer
Z-0492-2010 | 53538 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ADVANCE (R) CANAL FILLING STEM EXTENSION, REF: KSP1-0140, DIAMETER 11mm, LENGTH 140mm, 1 EACH, STERILE R, (FOR USE WITH CEMENT ONLY), SINGLE USE ONLY, (Ti6A14V), Wright Medical Technology, Inc., 5677 Airline Road, Arlington, TN 38002 |  |  | September 22, 2009 | 2009 | Wright Medical Technology Inc | Titanium used in the recalled devices was discovered to have internal microstructure conditions of alloy segregation, primary alpha and/or a transformed beta matrix in some of the heat lots. The absence of alloying elements results in an area with less strength. For those instances where the product has been implanted, there is the potential increased risk of failure for these patients. | Process control | Wright Medical initiated the recall of the product on September 22, 2009 by contacting the distributors by telephone followed with letters addressed to the hospitals and surgeons dated October 9, 2009. At the firm, Customer Service is the contact for return instructions and replacement inventory. This phone number is 800-238-7117. Other questions regarding this matter are directed to 800-874-5630. |  | 37 | 37 units | Nationwide, Columbia, Brazil, Portugal, UK, Belgium, Austria, and Greece | Terminated | December 16, 2010 | 450 | 62 | Not_Computer
Z-0488-2010 | 53812 | No | EKM | Gutta percha. | Dental | Dental | 510(K) Exempt | Elements Gutta Percha and Elements RealSeal Cartridges; Part Number: 972-1002, 972-1003, 972-2001, 972-2002 |  |  | October 27, 2009 | 2009 | Sybronendo | The recall was initiated because during the manufacturing process of the affected units, the molding process resulted in lock nuts that did not have the level of acceptability required, and therefore, may not withstand the pressures exhibited by the Elements Obturation Unit with which the cartridges are used. More Specifically, there exists a possibility that if the lock nut were to fail, the cart | Component design/selection | The recall communications was initiated on November 06, 2009 with the firm forwarding a Recall Letter to the affected consignees via US Postal Service 1st class mail. The recall letter informed the consignees of the reason for recall, the hazard and the affected products. Consignees were instructed to call SybronEndo Customer Care at 1-800-346-3636 directly to handle returns &amp; replacements. In addition, consignees were asked to complete the enclosed Recall Return Form and return it by fax to 909-962-5605.The recall notice also instructed consignees that are an authorized SybronEndo distributor to identify and recover the affected product lots listed above that may have been shipped to their customers. |  | 5046 | 5046 units (3720 units in US) | Nationwide and Worldwide to: USA, Canada, Chile, China, Ecuador, Guatemala, Hong Kong, Japan, Malaysia, Singapore, Taiwan and Venezuela | Terminated | March 22, 2012 | 877 | 5046 | Not_Computer
Z-0487-2010 | 53751 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Buretrol Solution Sets, a sterile fluid pathway of 105 (2.7 m), 60 drops/mL, with 150 mL Burette Drip Chamber Filter Valve, 2 injection Y-sites, 3 Luer Activated Valves and Male Luer Lock Adapter with Retractable Collar; 48 individually pouched sterile sets per case; Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A. |  |  | November 03, 2009 | 2009 | Baxter Healthcare Corp. | The tubing may separate from the distal Clearlink Y-site of the set, which could result in patient injury if the separation occurs during patient use. | Employee error | An "Urgent Product Recall" letter dated November 3, 2009 was sent to the direct accounts via first class mail on the same date, to the attention of the Director of Nursing and the Director of Materials Management. The letter informed the accounts that there have been reports of a separation of the tubing from the distal Clearlink Y-site of the set, with one instance resulting in a serious injury. The accounts were instructed to examine their inventory and remove any sets of Product code 2C8864, Lot number UR09E21238 from inventory and return the affected product for credit or replacement product by calling Baxter Healthcare Center for Service at 1-888-229-0001, Prompt 2, during hours of operation 7:00 a.m. to 6:00 p.m. Central Standard Time. Questions regarding the communication were directed to the Center for One Baxter at 1-800-422-9837 during the hours of 8:00 a.m. to 5:00 p.m. Central Standard Time. If the accounts further distributed the Buretrol Sets, they were requested to forward the communication to those accounts. The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter. |  | 173 | 173 cases | Worldwide Distribution -- United States (Alabama, California, Colorado, Florida, Georgia, Illinois, Kentucky, Louisiana, Michigan, Mississippi, Missouri, Nevada, New Mexico, Ohio, Tennessee, Texas, Utah and Wisconsin), Canada and New Zealand. | Terminated | July 08, 2010 | 247 | 173 | Not_Computer
Z-0499-2010 | 53813 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Zimmer Dental AdVent Implant System, Abutment, Tapered, Rx, sterile, Zimmer Dental, Inc., Carlsbad, CA; REF AVACT. |  |  | November 06, 2009 | 2009 | Zimmer Inc. | The devices may not meet specifications, which may result in fracturing during tightening of the component. | Process control | User accounts were notified of the issue by phone on or about 11/6/09. Clinicians were requested to schedule a follow-up appointment and to remove and replace the recalled components for any patients who had this device placed without fracturing during the procedure. Questions are directed to the company at 1-800-854-7019. |  | 738 | 738 | Ohio, Canada and France. | Terminated | October 08, 2010 | 336 | 738 | Not_Computer
Z-0498-2010 | 51734 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ILLICO Posterior Thoracolumbar Access Retractor, Part Number: 73520 |  |  | January 28, 2009 | 2009 | Alphatec Spine, Inc. | The recall was initiated after Alphatec Spine received a complaint on the ILLICO Posterior Thoracolumbar access retractor. It was reported that one of the retractor arms in the retractor broke during surgery. | Device Design | On January 28, 2009 Alphatec Spine contacted their domestic consignees via telephone and the international consignee was contacted via telephone and electronic mail. Alphatec Spine's telephone call consisted of informing the consignees of the product affected, the problem, what actions needed to be taken and what action Alphatec Spine would take to arrange for pick-up of the affected products. |  | 14 | 14 | AZ, CA, OR, TX, and Tokyo, Japan. | Terminated | December 22, 2009 | 328 | 14 | Not_Computer
Z-0500-2010 | 53689 | No | LGW |  |  | Neurology | PMA | Medtronic Accessory Kit , TitanÂż Anchor Accessory Kit, 3550-39, Single Use, RX Only, Contents of the inner package are STERILE., Medtronic, Inc, Minneapolis, MN 55432-5604 |  |  | October 27, 2009 | 2009 | Medtronic Neuromodulation | Medtronic is issuing a medical device correction for the Model 3550-39 TitanÂż anchor due to the potential for lead migration as a result of insert separation within the anchor. The Medtronic Model 3550-39 TitanÂż anchor accessory kit is designed for use with Medtronic percutaneous in-lineconnector spinal cord stimulation leads. Medtronic has received reports of separation of the titaniuminsert | Nonconforming Material/Component | Consignees were sent a Medtronic "Urgent: Medical Device Correction" letter on 10/27/09. The letter was addressed to "Dear Healthcare Professional" and was dated October 2009. The letter described the problem and the product as well as the Scope and Potential Severity and gave Recommendations. Enclosures with the letter included: Use By Date (UBD) location on the Titan anchor accessory kit product packaging, Titan anchor patient letter, Reply card and Return envelope. Questions are direct to Medtronic via email at rs.neuroqa1a@medtronic.com. |  | 55600 | 55,600 | All states in the United States including Puerto Rico.Australia, New Zealand, Canada, Andurra, Belgium, Czech Republic, Denmark, Finland, France, Germany, Greece, Netherlands, Ireland, Israel, Italy, Kuwait, Lebanon, Luxembourg, Norway, Portugal, South Africa, Spain, Sweden, Switzerland, Tukey, United Kingdom. | Terminated | December 24, 2011 | 788 | 55600 | Not_Computer
Z-0502-2010 | 53144 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Alphatec Spine Solanas Titanium Pedicle Screw, Part Number 63035-12 |  |  | August 05, 2009 | 2009 | Alphatec Spine, Inc. | Alphatec Spine discovered that the affected lots were not manufactured correctly in that the screw body portion of the assembly has only 22.5 degrees of angulation in the east-west plane instead of the designed 38-40 degrees. | Process control | Alphatec Spine telephoned consignees informing them of the above mentioned Public Reason for Recall on August 5, 2009. They were asked to check their inventory to see where this product is located.If the product had not already been used Alphatec arranged to have Federal Express contact consignee for pick-up of the product(s). They were told to expect to receive the replacement(s) within 2-3 days. |  | 163 | 163 Units | Nationwide distribution. | Terminated | June 07, 2010 | 306 | 321 | Not_Computer
Z-0511-2010 | 53557 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Merit Medical Systems, Inc. Connection tubes label reading in part: CONNECTING TUBE W/ STOPCOCK 12" CATALOG NUMBER K10-04131 LOT F720653 Connection tubing can be used with any drainage catheter and drainage bag |  |  | September 29, 2009 | 2009 | Merit Medical Systems, Inc. | Potential sterility breach caused by excessive shipment damage. | Other | Merit Medical Customers were contacted directly by phone and letter on 9/29/09. Customers were instructed to identify and quarantine all unused affected product. Further information was available at 801-208-4365. |  | 1 | 1 connection tube | US (AL, LA, MO, KY, TN, GA). No military, government or foreign distribution. | Terminated | December 15, 2009 | 77 | 352 | Not_Computer
Z-0525-2010 | 53454 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Abbott Spine Thoracolumbar Common, SpdLnk 5.5 Trans Cntr Med 43-51 mm, non-sterile, Manufactured by Abbott Spine, Inc., Austin, TX; Ref. 1302-55M. Note: This lot was not distributed in the United States. |  |  | September 21, 2009 | 2009 | Zimmer Inc. | The titanium stock used to manufacturer the product is reported to exhibit areas of alloy segregation, which could result in a weakening of the strength of the metal. | Nonconforming Material/Component | Consignees were notified by letter dated September 21, 2009 and instructed to return the affected product to Zimmer, Inc. |  | 13 | 13 | Worldwide Distribution -- Canada, Mexico and South Africa. (Product not distributed in the United States) | Terminated | October 08, 2010 | 382 | 13 | Not_Computer
Z-0501-2010 | 53860 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Gender Solutions Nexgen Complete Knee Solution, Legacy Knee- Posterior Stabilized, Gender Solutions Female, Femoral Component, option, LPS-flex, size C, right, for cemented use only; REF 00-5764-013-52. |  |  | November 13, 2009 | 2009 | Zimmer Inc. | The inner package may interfere with the outer seal, resulting in lack of assurance of sterility. | Packaging process control | Zimmer sent a recall letter, dated 11/12/09, to each initial consignee on 11/13/09. Zimmer field staff is to remove the product from affected hospitals, and to provide them with a copy of the recall letter. |  | 22 | 22 | Nationwide, Korea and Switzerland. | Terminated | October 15, 2010 | 336 | 22 | Not_Computer
Z-0535-2010 | 52837 | No | LRK | Intraoral devices for snoring and intraoral devices for snoring and obstructive sleep apnea. | Dental | Dental | 510(k) | Repose G2 Bone Screw System, Sterile, REF 76353200; Lot # 62450300. |  |  | July 27, 2009 | 2009 | Medtronic Xomed, Inc. | Medtronic Xomed, Inc. is recalling the Repose Bone Screw System product and Bone Screw product due to improper labeling. The labeling for the Repose System was inadvertently placed on the Repose Screw and vice versa. | Labeling mix-ups | Firm notified customers by phone on 7/27/2009. Firm asked customers to check inventories for the affected products listed on the enclosed Checklist.1. Fill-in the "quantity on-hand "column on the checklist.2. If you have any this product lot, contact Medtronic ENT Customer Service at 800-874-5797 (select Option 1 for Customer Service) and ask for the REPOSEField Action Contact Person for help to arrange for returns and replacements or credit.3. Fax your Checklist to Medtronic ENT at the number provided.4. Return products clearly marked. |  | 7 | 7 | 9 consignees in the continental United States. | Terminated | March 15, 2010 | 231 | 20 | Not_Computer
Z-0509-2010 | 53495 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | IMPACT Catheters, Balloon Dilatation and Esophageal Dilatation Catheter, Model Number 307, STERILE, Rx only |  |  | September 23, 2009 | 2009 | Numed Inc | The firm received a complaint from the sole direct account that while performing an inventory review it was observed that the packaging of the IMPACT catheter was cracked and deteriorating. The product was still in its packaging and never used. The customer referred to it as looking like "paint crackle." The risk is that for continued sterility to be assured, the packaging must remain intact. | Labeling design | An informal e-mail was sent to B. Braun to give them notice not to distribute any Impact devices until further notice. This was followed up by a formal URGENT MEDICAL DEVICE RECALL letter (dated 9/24/09) to the sole consignee. The products are to be returned to the manufacturer, including any products further distributed. Customer questions should be directed to the recall contact (replace " AT " with @) nlaflesh ATnumedusa.com |  | 7367 | 7,367 units domestically | One distributor in Pennsylvania | Terminated | May 23, 2012 | 973 | 7367 | Not_Computer
Z-0530-2010 | 53584 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BD Recalled component is labeled: |  |  | January 09, 2009 | 2009 | Biomet Microfixation, Inc. | Biomet Microfixation, Inc. is recalling their OnPoint TMJ Kit (Part Numbers 24-3050, 24-3055) and Arthrosimplicity TMJ Arthocentesis Kit (Part Numbers 922070, 922090, 922075). These two kits contain a BD sterile disposable 60 ml Luer-Lok syringe, Part number 309653 which is under recall by Becton Dickinson (Recall Number Z-2347-2008). Becton Dickinson has determined that unit package seal integ | Device Design | Consignees were notified by a combination of telephone, fax, e-mail and certified mail. Questions related to this notice should be directed to Christy Cain at 1-800-874-7711. |  | 121 | 121 Kits | Product was distributed Internationally to Argentina, Denmark, Norway, and UK. Product was distributed domestically to AL, CA, CT, FL, GA, HI, IA, IL, IN, KS, LA, MA, MD, MN, MO, MS, NC, OH, NY, PA, SC, TN, TX, UT, and VA. | Terminated | March 17, 2010 | 432 | 235 | Not_Computer
Z-0529-2010 | 53858 | No | HTS | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Sterile Knee Instrument, Patella Cutter; |  |  | November 11, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Disposable Patella Cutters - 30 mm (catalog number 8000-0030, lot V31257001) were packaged and labeled as 33 mm Disposable Patella Cutters (catalog number 8000-0033, lot V31257001 and vice versa. | Labeling mix-ups | An "Urgent Product Recall" letter dated November 11, 2009 was sent to all affected Stryker branches, hospital risk management, hospital Chief or Orthopaedics and surgeons by Federal Express. The letter described the issue, potential hazards and risk mitigation factors and requested that the Recall Acknowledgement form be completed and returned. The Recall Acknowledgement form has information regarding the return of affected product to the firm. Direct questions to Stryker Orthopaedics by calling 1-201-831-5832. |  | 12 | 12 | Nationwide Distribution (MA, NJ, FL, TN, OR and NY). | Terminated | August 30, 2010 | 292 | 12 | Not_Computer
Z-0537-2010 | 50486 | joint, knee, external limb component | ISY | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Otto Bock Modular Knee Joint, 3R66, Otto Bock HealthCare GmbH, Max-Nader-StraBe 15 . 37115 Duderstadt /Germany. |  |  | October 28, 2008 | 2008 | Bock,Otto,Orthopedic Ind,Inc | Otto Bock Health Care LP has discovered that in some cases, the securing screws of the joint axes may not have been properly mounted. This can cause the axes to gradually move out of the joint and result in a risk of falling for the user. | Employee error | Consignees were sent on 10/28/08 a Otto Bock "Urgent: Correction Notice" letter dated October 27, 2008. The letter described the product and problem. It provided instructions if the product was in stock and instructed the consignee to contact the patients (a patient letter was provided) and to complete and return the reply card. Consignees were to contact Otto Bock at their Customer Service to request a return authorization. |  | 9 | 9 | The product was only distributed within the US by the US company. Any OUS distribution was handled by the manufacturer in Germany. | Terminated | December 10, 2011 | 1138 | 9 | Not_Computer
Z-0526-2010 | 52010 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Proteus XR/a. Model # 2259988. |  |  | August 13, 2009 | 2009 | GE Medical Systems, LLC | Certain GE Radiographic fixed X-ray tables (Proteus XR/a, Revolution XR/d, and Definium 8000) were found to have tabletops that move unexpectedly in a bi-directional mode. In rare cases, a malfunction may occur, which could result in the tabletop brakes falling to engage when the control foot pedal is released to lock the floating tabletop in place. | Nonconforming Material/Component | A GE Healthcare "Urgent Medical Device Correction" letter dated April 9, 2009 was issued to customers. The letter was addressed to Hospital Administrators/Managers of Radiology and Radiologists. The letter described the Product Issue, Affected Product, Safety Instructions, Product Correction and Contact Information. Operators should verify that the tabletop is locked before the patient is loaded on or off the table. A Field Modification will be performed on the affected products. A GE Healthcare Field Engineer willschedule a field visit to make this modification. Direct questions or concerns regarding the recall to your local GE Healthcare service representative or by calling 1-262-853-7685. |  | 5943 | 5,943 of total product | Worldwide Distribution -- UNITED STATES (including Puerto Rico), VIETNAM, VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, UKRAINE, TURKS AND CAICOS ISLANDS, TURKEY, TUNISIA, TRINIDAD AND TOBAGO, THAILAND, TAIWAN, SWITZERLAND, SWEDEN SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIA, PORTUGAL POLAND, PANAMA, PAKISTAN, NORWAY NEW ZEALAND, NETHERLANDS , MEXICO MARTINIQUE, MALTA, MALAYSIA, LITHUANIA, LABA NON, LATVIA, KUWAIT, REPUBLIC OF KOREA, KAZAKSTAN, JORDAN, JAPAN, JAMAICA, ITALY, ISRAEL, IRELAND, INDONESIA, INDIA, ICELAND, HUNGARY, HONG KONG, GUAM, GREECE, GERMANY, FRENCH POLYNESIA, FRANCE FINLAND , ECUADOR, DENMARK CZECH REPUBLIC, CROATIA, COLUMBIAN, CHINA, CHILE , CANDA, BULGARIA, BRAZIL, BOSNIA, BELGIUM AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | May 24, 2012 | 1015 | 5943 | Not_Computer
Z-0172-2010 | 53809 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | MX Series Products, Model numbers MX 20, MX 20 High Definition (HD), MX 15, MX 15i (integrated) and MX 15 HDi. |  |  | August 01, 2009 | 2009 | Wescam Inc | Products do not comply with 21 CFR 1040.11. | Radiation Control for Health and Safety Act | L-3 Wescam has issued notification of the MX Series Products describing a non-compliance with 1040.11(b). The notice instructs customers to refrain from using the infrared laser illuminator, based on a serious eye hazard from viewing the direct or reflected laser emission. Products will be repaired or removed if they cannot be repaired and a refund will be implemented. |  | 306 | 306 total | Nationwide Distribution | Terminated | January 15, 2013 | 1263 | 306 | Not_Computer
Z-0541-2010 | 53519 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | GE Healthcare Voluson E6, GE Voluson E8, and GE Voluson E8 Expert Diagnostic Ultrasound Systems |  |  | October 22, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware of an incorrect operating procedure associated with software versions 9.0.0, 9.0.1, and 9.0.2 of the Voluson E6 or Voluson E8 Diagnostic Ultrasound Systems that may impact patient safety.When using the combination of 2D-CRI and PW with a steered Doppler angle, the displayed velocity scale is incorrect. The consequence can be an underestimation of flow v | Software design | Consignees were sent on 10/23/09 a GE healthcare "Urgent Medical Device Correction" letter dated 10/22/09. The letter was addressed to Hospital Administration/Risk manger, Primary Physician and Biomedical Dept. Manger. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Questions and concerns regarding this issue are directed to the local GE Healthcare service representative or to 1-800-437-1171.GE also plans to update all of the affected GE Voluson E6, GE Voluson E8, and GE Voluson E8 Expert Diagnostic Ultrasound Systems that were distributed. |  | 715 | 715 | AL, AR, CA, CT, FL, GA, IL, KS, KY, LA ME, MD, MA, MI, MN,MS, MO, NE NJ, NY, NC, OH, OK PA, SC, TN,TX, VA,WI.AUSTRALIA, AUSTRIA , BELGIUM, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, HUNGARY, INDIA, INDONESIA, IRAN, IRELAND, ITALY, JAMAICA, JAPAN, KOREA, KUWAIT, LITHUANIA, MALAYSIA. NETHERLANDS, NE ZEALAND, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, UNITED ARAB, EMIRATES, UNITED KINGDOM, MEXICO, VENEZUELA. | Terminated | December 17, 2011 | 786 | 715 | Software | Display/Image | Software update | Software Update
Z-0540-2010 | 53596 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2121-1Q/3131-1Q Biplane Cardiovascular Imaging Systems. |  |  | October 14, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has become aware of a potential limit switch failure with Innova 2121IQ and 3131IQ (Biplane) systems that may impact patient safety.A potential limit switch failure could occur due to collision between the lateral plane of Innova Positioner and other equipment in the room. This collision may break the x-ray and positioning switches, and may not allow exam completion. To date, no in | Employee error | GE Healthcare issued an "Urgent Medical Device Correction" notice dated October 14, 2009. The letter was addressed to Hospital Administrators/Risk mangers, Managers of Radiology/Cardiology, Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.For further information, contact GE Healthcare Systems at 1-800-437-1171, option 4 (U.S.) or contact your local GE Healthcare Service Representative. |  | 114 | 114 | Worldwide Distribution -- Including the UNITED STATES, YEMEN, UNITED KINGDOM, TAIWAN, SINGAPORE, RUSSIA, ROMANIA, MALAYSIA, LYB IA, JAPAN, ITALY, IRELAND, INDONESIA, INDIA, GERMANY, FRANCE, EGYPT, CHILE , CANADA and BAHRAIN. | Terminated | February 04, 2012 | 843 | 114 | Not_Computer
Z-0538-2010 | 53943 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Biomet 1.1 X 50MM, 3.5MM STOP STAINLESS STEEL, 1.5MM SYSTEM TWIST DRILL WITH J NOTCH, REF 01-7141. |  |  | November 06, 2009 | 2009 | Biomet Microfixation, Inc. | The Twist Drill 1.1x50mm, 3.5mm stop, with J notch Lots #'s 537407, 536306, and 535933 which were packaged and shipped without the instructions for use (IFU). | Process change control | Consignees were notified of the recall by a combination of telephone calls, letters, and emails during the week of November 16, 2009. Customers were instructed to complete the Inventory Reconciliation Sheet and fax (1-904-741-9425) to the firm. Customers who further distributed the affected product was instructed to notify recipients of the recall. Direct questions about the recall to Biomet Microfixation, Inc. by calling 1-800-874-7711 or 1-904-741-4400, extension 468, Monday through Friday, 8 a.m. to 4 p.m. Eastern Standard Time. |  | 283 | 283 pieces | Worldwide Distribution -- United States (including Puerto Rico), Australia, Brazil, Canada, Chile, Europe, France, Germany, Greece, Japan, Lebanon, Spain, and the United Arab Emirates. | Terminated | March 02, 2010 | 116 | 283 | Not_Computer
Z-0539-2010 | 52384 | No | HTD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Holding Forceps for Trident Acetabular System; non Sterile; |  |  | January 24, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | The tips of the forceps were larger than the positioning holes located on the trials and the implant liners. | Pending | Important Product Correction letters, dated January 24, 2007, were sent to all consignees on via Federal Express. The letter stated the problem and asked customers to examine their inventory and hospital locations to identify the affected product. All affected forceps product should be retrieved from hospitals and returned to their branch or agency warehouse for reconciliation. Customers are also to reconcile all affected product on the attached Product Accountability Form and fax a copy of that form to 201-831-6069 within 5 days of receipt of the letter. All affected forceps product is to be destroyed and discarded at their branch location. Questions should be directed to their Stryker Orthopaedics Customer Service rep at 1-866-OR-ASSIST. |  | 351 | 351 units | Worldwide Distribution -- Including countries of Hong Kong, Australia, Japan, New Zealand, Hong Kong, Singapore, Germany, Sweden, The Netherlands, France, Spain, Portugal, and the United Kingdom. | Terminated | December 21, 2009 | 1062 | 351 | Not_Computer
Z-0543-2010 | 53771 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | i-Cat Classic, 3D Dental Imaging System manufactured by Imaging Sciences International, Hatfield, PA. |  |  | October 30, 2009 | 2009 | Imaging Sciences International, LLC | overhead carriage drops quickly | Device Design | Imaging Sciences International issued an "Urgent Medical Device Recall" notification dated October 30, 2009 to all customers informing them of the problem and that a field service representative would schedule a visit to install an upgrade kit.For further information, contact Imaging Sciences International at 1-800-205-3570, option 5. |  | 98 | 98 units | Worldwide Distribution -- Including the United States, Austria, Australia, Belgium, Brazil, Canada, France, Germany, India, Israel, Italy, Jamaica, Spain, South Korea and Taiwan. | Terminated | September 30, 2010 | 335 | 98 | Not_Computer
Z-0544-2010 | 53577 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Centrifugal Pump (with X-coating), catalog 164275X, packaged in 8 piece shelf boxes, and convenience kits with single pump units. |  |  | July 06, 2009 | 2009 | Terumo Cardiovascular Systems Corporation | The tubing- to- pump connection in a medical device used in surgical procedures may disengage resulting in blood loss. | Device Design | Terumo Cardiovascular Systems Corporation notified consignees of an "Urgent Medical Device Correction" by letter on July 06, 2009. The letter was addressed to hospital Perfusion, Cardiovascular Surgery, and Operating Room departments and contained specific measures to be taken to ensure proper use of the affected device. The firm recommended that users continue to use the devices and follow the corrective measures enclosed with the notification. Consignees were requested to route the notification to all users and confirm receipt of the notification by returning a customer response form.For further information, contact Terumo CVS Customer Service at 1-800-521-2818. |  | 20702 | 20,702 units | Nationwide Distribution | Terminated | November 09, 2011 | 856 | 34455 | Not_Computer
Z-0546-2010 | 53996 | No | HBB | Pneumatic cranial drill motor. | Neurology | Neurology | 510(k) | Medtronic Midas Rex Legend Footed Attachment, AF01. |  |  | November 30, 2009 | 2009 | Medtronic Sofamor Danek Usa, Inc - Dallas Distribution | Component of product attachment may be outside established tolerance with potential for footed portion to break during normal use. | Nonconforming Material/Component | Medtronic issued an "Urgent: Medical Device Recall" letter dated November 30, 2009. Consignees were notified of the issued and asked to quarantine any affected product and return affected product to the firm for replacement.For further information, contact Medtronic Powered Surgical Solutions Customer Service at 1-800-433-7080 or 1-817-788-6400. |  | 64 | 64 | Worldwide Distribution -- Including the United States (DC, TN, NC, KY, AL, FL, NJ, OH, MI, IN, NY, WA, CA, NV, NM, IA, MO, NE and UT), Great Britain, Brazil, Turkey, Singapore and the Netherlands. | Terminated | May 12, 2010 | 163 | 64 | Not_Computer
Z-0548-2010 | 53987 | No | HDD | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Hill-Rom Affinity Four Birthing Bed, Hill-Rom Company, Inc., Batesville, Indiana; Model P3700B. |  |  | November 16, 2009 | 2009 | Hill-Rom, Inc. | The brakes may not hold or lock. | Device Design | Hill-Rom issued an "Urgent Safety Notice" notification dated November 16, 2009 to customers. The letter advised users of the problem and asking them to locate the beds and to remove them from use until corrected if possible and, if not possible to inspect them for the problem and remove them from use if it is present. The firm will visit each account to make corrections.For further information, contact Hill-Rom Technical Support at 1-800-445-3720. |  | 2551 | 2551 (US), 4171 (OUS) | Worldwide Distribution -- United States, Canada, Argentina, Australia, Bahamas, Brazil, Chile, China, Denmark, Dominican Republic, France, Great Britain, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Libiya, Netherlands, Poland, Puerto Rico, Russia, Saudi Arabia, Serbia, Singapore, Spain, Sweden, Syria, Thailand, Turkey and United Arab Emirates. | Terminated | July 02, 2012 | 959 | 2551 | Not_Computer
Z-0549-2010 | 53889 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Philips Avalon Fetal Monitor FM 20 |  |  | November 23, 2009 | 2009 | Philips Healthcare Inc. | Potential for inaccurate ultra-sound derived fetal heart rate readings in Philips Avalon Fetal Monitors. Although the types of inaccuracies described in these complaints reflect known limitations inherent to ultrasound fetal heart rate monitoring, the frequency of such complaints is greater for these devices than for the previous generation Philips Series 50 fetal monitors. Clinical decisions bas | Process design | Philips notified customers with an Urgent Medical Device Recall notice and Instructions for Use Addendum on 11/23/09 via UPS. Philips is providing the Addendum to reinforce the Avalon Instructions for Use regarding:"Recognizing these phenomena as they relate to the Avalon Fetal monitor and fetal monitors in general; and"Understanding the differences between the Philips Avalon Fetal Monitors and their predecessors, the Philips Series 50 Fetal Monitors, as they relate to the display of ultrasound-derived FHR.Philips is providing access to application specialists who are on call 24 hours a day to assist users with questions related to device use by calling 1-800-722-9377, press &lt;3&gt; then &lt;1&gt; for patient monitoring application issues and ask to speak with an application specialist. Customers will be requested to incorporate the information contained in the Urgent Medical Device Recall notice and the Addendum to the Instructions for Use with the existing Avalon labeling (Instructions for Use) provided with the product. Customers will be asked to complete an enclosed Confirmation Form indicating that they have received and disseminated this information, and then return the completed form to Philips. Questions about this field action or reports of product problems should be directed to the Philips Response Center at 1-800-722-9377. |  | 10128 | 10,128 units | Nationwide (US and Puerto Rico).Foreign: AfghanistanAlgeriaArgentinaAustraliaAustriaBahrainBelgiumBoliviaBosnia and HerzegovinaBotswanaBrazilBulgariaCanadaChileChinaColombiaCosta RicaCroatiaCzech RepublicDenmarkDominican RepublicEcuadorEgyptEl SalvadorEstoniaFinlandFranceFrench GuianaGabonGermanyGreeceGuadeloupeGuatemalaHong KongHungaryIcelandIndiaIndonesiaIrelandIsraelItalyJapanJordanKazakhstanKenyaKorea, Republic ofKosovoKuwaitLatviaLebanonLibyan Arab JamahiriyaLithuaniaLuxembourgMacauMalaysiaMartiniqueMauritiusMexicoMoldovaMoroccoNamibiaNetherlandsNetherlands AntillesNew CaledoniaNew ZealandNorwayOmanPakistanPanamaPeruPhilippinesPolandPortugalPuerto RicoQatarReunionRomaniaRussiaSaudi ArabiaSingaporeSlovakiaSloveniaSouth AfricaSpainSri LankaSwedenSwitzerlandSyrian Arab RepublicTaiwanThailandTunisiaTurkeyTurks and Caicos IslandsUkraineUnited Arab EmiratesUnited KingdomUzbekistanVenezuelaViet NamZambia | Completed |  |  | 18940 | Not_Computer
Z-0547-2010 | 53938 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Vascular Solutions Pronto-Short Extraction Catheter. Model 5030. The PRONTO-Short Extraction catheter is a dual lumen, over-the-wire (OTW) catheter with related accessories. The catheter is designed to be delivered through a 6F or larger introducer sheath over the included 0.018" (0.45mm) guidewire. Sterilized with ethylene oxide. Manufactured by: Vascular Solutions, Inc., 6464 Sycamore Court, Minneapolis, MN 55369 USA. |  |  | November 20, 2009 | 2009 | Vascular Solutions, Inc. | It was found that three Pronto Short (model 5030) device lots were manufactured and labeled with an expiration date which is inconsistent than the expiration of the included guidewire component. The guidewire is sterilized separately and was realized to have an expiration date of October 31, 2009. | Process control | Vascular Solutions issued an "Urgent: Medical Device Recall - Lot Specific" notification dated November 20, 2009, via USPS. The letter described the issue, identified affected product, asked for affected devices to be removed from inventory, quarantine, and call customer service for a Return Authorization number. Units will be replaced upon Vascular Solutions (VSI) receiving the returned devices. A Product Recall Inventory form was asked to be completed and faxed back to VSI.For further information, contact your local sales representative or Vascular Solutions Director of Marketing at 1-763-656-4300. |  | 274 | 274 (150 US, 124 OUS) | Worldwide Distribution -- United States (CA, CO, FL, GA, IA, IL, IN, KS, LA, MI, MO, NC, NE, NH, NV, NY, OH, OR, PA, SC, TX, VA, VT, WA and WI), Canada, France, Ireland, Israel, Italy, Netherlands, Spain, Sweden, Taiwan and Turkey. | Terminated | April 05, 2012 | 867 | 274 | Not_Computer
Z-0557-2010 | 51180 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Howmedica Osteonics, Primary Super SecurFit Stems, hip prosthesis, Sterile; |  |  | January 14, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that there may be difficulty in mating stem to implantation/extraction instrument. | Device Design | Stryker sent Urgent product Correction Letters on January 14, 2009 by Federal Express to Stryker branches/agencies, hospital risk management, chief of orthopaedics and surgeon. |  | 786 | 786 total stems | Nationwide. Product was distributed to 3 Stryker branches/agencies and 2 hospitals in PA and CA. | Terminated | July 18, 2012 | 1281 | 786 | Not_Computer
Z-0555-2010 | 51219 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer, Insall/Burstein II Modular Knee System, Tibial Wedge and 2 Screws, Size 59, 7 Deg. x 9.5 mm Height, Tivanium TI-6AL-4V Alloy, sterile; REF 5220-62-01. |  |  | February 26, 2009 | 2009 | Zimmer Inc. | Smaller diameter screws than required were included in the package. | Process control | Zimmer notified consignees by recall letter dated 2/25/09 instructing them to cease use of the device and to return it. Zimmer distributors and international accounts were notified by email on 2/26/09. |  | 13 | 13 | USA (Connecticut, Indiana, Pennsylvania), Australia, Serbia, Spain, Switzerland and the United Kingdom. | Terminated | February 28, 2011 | 732 | 13 | Not_Computer
Z-0554-2010 | 52805 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Zimmer / Abbott Spine Sequoia dorsal height &amp; revision tool, non-sterile; REF 3367-1. Surgical instrument for pedicle screw system. |  |  | August 20, 2009 | 2009 | Zimmer Inc. | The tip of the instrument may fracture during use. | Device Design | Zimmer sent a letter dated 8/17/09 to distributors and consignees on or about 8/20/09. Users were notified of the potential problem via the letter and advised that the instruments will be removed and replaced when a new version of the instrument becomes available (around 10/09). |  | 215 | 215 | Nationwide, Belgium and New Zealand. | Terminated | April 13, 2011 | 601 | 215 | Not_Computer
Z-0558-2010 | 53687 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Zimmer/CAS Power Cord, Sesamoid Plasty, NA, Zimmer/CAS, Montreal (Quebec), Canada; REF 20-8000-070-12. The device is the power cord component for the Sesamoid Plasty CAS workstation which connects the workstation to the power mains. |  |  | November 02, 2009 | 2009 | Zimmer Inc. | The power cord female receptacle is not recessed sufficiently and may present a shock hazard. | Device Design | Zimmer sales staff were notified by letter dated 11/2/09 and instructed to locate the units and to upgrade the cords and to notify consignees by copy of a letter addressed to risk managers and dated 11/2/09. |  | 71 | 71 | Nationwide. | Terminated | October 15, 2010 | 347 | 71 | Not_Computer
Z-0560-2010 | 53878 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Toggleloc System, Biomet Ziptight Ankle Syndesmosis Fixation System - SS with Ziploop Technology, polyethylene/polypropylene/ polyester/ stainless steel, sterile Biomet Sports Medicine, Warsaw, IN.; REF 909857. |  |  | November 10, 2009 | 2009 | Biomet, Inc. | Packages labeled as containing a stainless steel device actually contain a titanium device, and vice versa. | Labeling mix-ups | Distributors, hospital consignees and an implanting surgeon were notified of the recall by an Urgent Medical Device Recall Notice, dated 11/10/09. Distributors were instructed to locate and remove the identified products from circulation. The instructions on the enclosed "FAX Back Response Form" should be followed and product from their accounts should be returned. Hospital consignees and the implanting surgeon were informed of the recall and advised of the potential health hazards presented by the mix-up. |  | 24 | 24 of both products. | Nationwide Distribution -- Including states of California, Florida, South Carolina, and Virginia. | Terminated | November 03, 2010 | 358 | 24 | Not_Computer
Z-0542-2010 | 53946 | No | HQY | Sunglasses (nonprescription). | Ophthalmic | Ophthalmic | 510(K) Exempt | Polycarbonate prescription lenses (single vision, progressive, and trifocal) SKU #'s: 3703154, 4286102, 4286103, 4286104 and 4286106 |  |  | November 24, 2009 | 2009 | Luxottica | The lenses made for sunglass frames (Liberty Sport Slam or MaxMorpheus ) do not meet the lens thickness specification established by the frame manufacturer. The lens thickness range from 1.0 to 1.9 millimeters instead of the required 2.0 millimeter minimum. There is a potential risk that the lenses could break or disengage from the frame in the event of an impact. | Employee error | On 11/24/2009 the firm sent recall notification letters to their customers indicating the problem with the lenses. they instructed their customers return the product to the place of purchase for replacement or refund. Questions are directed to the firm at 1-877-488-2407 from 8:00 a.m. to 5:00 p.m. ET. |  | 276 | 276 polycarbonate prescription lenses (138 pairs of lenses) | Product was distributed to the following states: AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MO, MT, NC, NH, NJ, NM, NV, NY, OH, OR, PA, TN, TX, UT, VA, WA, WI, &amp; WV. | Terminated | October 27, 2010 | 337 | 276 | Not_Computer
Z-0553-2010 | 53873 | No | MRC | Rigid gas permeable contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb BostonÂż Conditioning Solution, Original Formula, Bausch &amp; Lomb Incorporated, Rochester, NY 14609. The product is packed in 1 fl oz (30 mL), 3 fl oz (90 mL), 3.5 fl oz (105 mL) and 4 fl oz (120 ml) bottles. |  |  | October 30, 2009 | 2009 | Bausch &amp; Lomb Inc | The product failed to meet Out of Specification shelf life at the 9 months time point. This could result in increased risk of an ocular infection, especially in the presence of other contributing or risk factors. | Other | Consignees were notified by mail on/about 11/06/2009. All consignees were asked to review their inventory, and return any of the affected products to Stericycle Pharmaceutical Services. A Business Reply Card was included to be returned via US mail within 5 days. A packing slip was provided to be affixed to a pre-paid USP shipping label for return. |  | 159672 | 159,672 bottles | U.S. Nationwide, Canada and Hong Kong | Terminated | August 29, 2012 | 1034 | 159672 | Not_Computer
Z-0561-2010 | 47370 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT System for ENT (server and work station), Xoran Technologies, Ann Arbor, MI. |  |  | September 17, 2008 | 2008 | Xoran Technologies, Inc. | The system may overheat, causing poor image quality, interruption in image acquisition and/or damage to the system. Also, if the system is not allowed to cycle down prior to being restarted after an emergency stop, the x-ray and gantry motions will both resume. | Device Design | A letter dated 9/17/08 was sent to each account telling them that the PLC would be upgraded. The firm visited each account and made corrections. |  | 82 | 82 | Nationwide Distribution | Terminated | December 23, 2009 | 462 | 82 | Hardware | Device Operation | Repair | Repair
Z-0566-2010 | 53817 | No | MMZ | Urological table and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Stille Sonesta Gynecology and Urology Table 6303 ; an electric cystometric table; Stille AB, Solna, Sweden. |  |  | November 03, 2009 | 2009 | Stille AB | There is a potential hazard to the patient if the bolts holding the side rail break, causing the accessories attached at the side rails of the tables and chairs to fall off during a medical procedure. | Nonconforming Material/Component | Stille sent Field Safety Notices dated 10/30/09 to the direct accounts, informing them that there is a potential hazard with the Sonesta tables if the bolts holding the side rail break, causing the accessories attached at the side rails of the tables and chairs to fall off during a medical procedure. Stille developed a retrofit kit consisting of six side rail bolts and instructions for installing the bolts, which were sent to the end users with the safety notice. The accounts were requested to replace the side rail bolts with the new bolts provided, and to complete and return the enclosed acknowledgement card to Stille by fax or e-mail as soon as possible and no later than 12/11/09. Contact Stille Surgical Inc. at 224-612-5404 for assistance. |  | 68 | 68 units total | Worldwide Distribution -- Including countries of USA and Canada. | Terminated | December 10, 2010 | 402 | 68 | Not_Computer
Z-0562-2010 | 53820 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart Vehicle Wall Mount, Model M5528A |  |  | November 12, 2009 | 2009 | Philips Healthcare Inc. | HeartStart Vehicle Wall Mount (Model M5528A) with DC Power Module (Model M5529A)--Incorrect wiring during installation by user may result in a smoke or fire hazard. | Device Design | Philips Healthcare, Inc. notified customers by an "URGENT - Voluntary Medical Device Correction" letter, plus revised Instructions For Use (IFU) for the Vehicle Wall Mount accessory and the DC Power Module accessory of the HeartStart MRx which includes a revised description of correct installation and verification of performance. Customers were asked to follow the Actions in the customer/user section of the Urgent - Voluntary Medical Device Correction Notification. In this section, customers were given the instructions to correct their devices. Distribution of the Urgent - Voluntary Medical Device Correction Notification along with the IFU will be managed by Philips representatives in each affected geography. Direct questions or concerns about the recall to your local Philips representative. |  | 3265 | 3,265 units | Class II Recall -- Worldwide Distribution -- United States, Canada, Australia, Brazil, China, France, Germany, Hong Kong, Iceland, India, Indonesia, Italy, Japan, Jordan, Kenya, Libyan Arab Jamahiriya, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Romania, Saudi Arabia, Spain, Sri Lanka, Switzerland, Trinidad, Tobago, United Arab Emirates, and United Kingdom. | Terminated | January 25, 2012 | 804 | 3265 | Not_Computer
Z-0563-2010 | 53837 | No | ETP | Air or water caloric stimulator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | AquaStar Water Caloric Irrigator; water caloric stimulator; Manufactured for/Distributed by Micromedical Technologies, Inc., 10 Kemp Drive, Chatham, IL 62629 USA, Instrumentation Difra, Made in Belgium. |  |  | November 16, 2009 | 2009 | Micromedical Technologies, Inc | Some of the filters on both the display board and the main electronic board (PCB) become oxidized causing the LCD display to be garbled and the timer will not operate. | Nonconforming Material/Component | Micromedical Technologies, Inc. sent letters dated November 16, 2009 via e-mail (with read receipt) to each customer and their representative. If they did not have the customer's e-mail address, the sales representative was requested to forward the letter to the customer and copy Micromedical. The letters advised the customers that one or more electronic components in the AquaStar caloric irrigator may fail due to contact corrosion. In the event of this failure, unheated tap water will continue to flow through the device back to the water drain, but the LCD display will be garbled and the timer will not operate. If this problem develops, the account should stop using the irrigator and contact Micromedical Technologies, Inc. at 1-217-483-2122 or via e-mail at mmtinfo@micromedical.com to report the problem for repair of the unit. |  | 10 | 10 units | Worldwide Distribution -- United States (New York and Pennsylvania), Canada, Egypt, Korea, Poland and South Africa. | Terminated | April 18, 2011 | 518 | 10 | Hardware | Display/Image | Repair | Repair
Z-0567-2010 | 53952 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | SeQual Technologies Inc. Eclipse Oxygen System, Model 1000 |  |  | October 01, 2009 | 2009 | Sequal Technologies Inc | Unit may shut down when operating from the Power Cartridge. A temporary power interruption would cause the unit to sound an alarm and shut down. If the patient does not have access to an alternate supply of oxygen or access to AC or DC power, this may lead to shortness of breath, hypoxia, or in some patients may cause serious health consequences. | Pending | The recall notification was initiated on November 12, 2009 with the firm forwarding a "Medical Device Correction" letter to all of the affected consignees via US Postal Service for US Consignees and FedEx for Foreign Consignees. The Medical Device Correction Letter contained the following:(1) Medical Device Correction Letter; (2) List of Serial Number affected that were distributed to the Provider; (3) Reply Card that the Provider will be requested to send back regarding receipt notification of the Medical Device Correction Letter and confirming that they have the affected product; (4) Additional Medical Device Correction Letters for Providers to distribute to their patients.The Medical Device Correction Letter informed the consignees of the reason for recall, the hazard involved and the actions that were required. For further details or information, consignees were instructed to contact SeQual Technical Support at 1-800-826-4610 from 8:00 AM to 5:00 PM PST. |  | 6588 | 6,588 units | Worldwide Distribution -- Including countries of USA, ARGENTINA, AUSTRIA, BAHAMAS, BERMUDA, CANADA, CHILE, COLUMBIA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, ITALY, ISRAEL,JAPAN, JORDAN, KUWAIT, MALTA, NETHERLANDS, POLAND, PUERTO RICO, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | June 18, 2012 | 991 | 6588 | Battery | Device Operation | Notification | Safety Notice/Insructions
Z-0497-2010 | 53879 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | TrailBlazer Support Catheter, ev3 inc, 4600 Nathan Lane North, Plymouth, MN 55442-2920, Rx Only, Sterile EO. |  |  | November 06, 2009 | 2009 | Ev3, Inc. | Catheter may crack near the radiopaque marker band at the distal end of the device during use. Cracking of the catheter shaft in this location can result in material separation, and potential embolization. | Component design/selection | Consignees were sent a Medical Device Recall letter, dated 11/6/09. The letter was addressed to Risk Manger or Cath Lab Manage. The letter provided the Issue Summary, identified the Affected Product, and stated the Required Action, which asked consignees to locate and remove the affected product. Questions should be directed to ev3 Customer Service at 800-716-6700. |  | 70 | 70 | Nationwide Distribution -- USA, including states of CA, FL, GA, IL, IN, IA, KS, KY, MI, MT, NJ, PA, SD, TN, TX, and WI. | Terminated | January 05, 2011 | 425 | 70 | Not_Computer
Z-0568-2010 | 53339 | stretcher, hand-carried | FPP | Hand-carried stretcher. | General Hospital | General Hospital | 510(K) Exempt | Paramed Polypropylene evacuation sled for non-ambulatory patients. |  |  | October 29, 2008 | 2008 | Paramed Systems | Possible partial board disassembly due to inadequate webbing attachment. | Process change control | Customers were notified by letter and phone during 10/20/2008 and 10/29/2008 to schedule visits by Paramed to correct the devices. |  | 722 | 722 Units | Nationwide Distribution -- Including states of AZ, CA, FL, GA, IL, IN, KS, KY, MA, MI, MT, NY, PA, SC, TN and UT. | Terminated | January 28, 2010 | 456 | 722 | Not_Computer
Z-0574-2010 | 53709 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | AccuSure, U-100 30 Gauge, 1cc Insulin Syringe, 5/16" (8mm) Needles, Short Needle, packaged 10 syringes per package, 10 packages per carton, NDC 0603-7000-21 |  |  | October 27, 2009 | 2009 | Vintage Pharmaceuticals LLC DBA Qualitest Pharmaceuticals | Needles are becoming dislodged from the barrels. | Other | The recall is being conducted to the Consumer level. A Press Release was issued on October 27, 2009 stating that Qualitest Pharmaceuticals was issuing a nationwide recall of all Accusure ÂżÂż Insulin Syringes. The consignees were asked to post a copy of the press release in their retail stores and to alert patients of this action. Also, Recall Notifications were sent to all of the firm's consignees by First Class Mail on October 30, 2009. The consignees were instructed to notify their customers or sub accounts who may have been sold or dispensed the affected syringes starting in 2002, and to return the affected product. Questions are directed to the firm at telephone number 800-444-4011. |  | 17142100 | 171,421 cartons of 100 | Nationwide | Terminated | November 24, 2010 | 393 | 237867600 | Not_Computer
Z-0580-2010 | 53777 | No | DTB | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | PMA | Enpath Lead Adapter, Ref: 501214. The contents of the inner package have been sterilized by ethylene-oxide gas. Enpath Medical, Inc., 2300 Berkshire Lane North, Minneapolis, MN 55441 USA. |  |  | September 01, 2009 | 2009 | Greatbatch Medical | The silicone tubing of affected leads was processed outside of validated parameters. Although the material remains biocompatible, test data regarding the performance characteristics are not available. | Component change control | Customers were contacted by phone beginning 09/01/09. They were asked to hold product and determine if they had distributed any product. A written notification followed which identified the product, described the issue and asked for product to be returned. Product was distributed to two consignees and contained the following day. Questions are directed to the company at 612-559-2613. |  | 45 | 45 units | CA and MN | Terminated | December 17, 2011 | 837 | 80 | Not_Computer
Z-0586-2010 | 53901 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 30 ml Syringe; |  |  | November 23, 2009 | 2009 | Becton Dickinson &amp; Company | Reports received of open seals on lot # 9175046, BD 30 ml Luer-Lok Syringes. This indicates product sterility may be compromised. | Other | Recall notification letters were sent to all distributors on 11/23/2009 by Federal Express. BD called all distributors to get customer lists and will start sending letter to customers commencing on December 8, 2009. Questions are directed to BD Customer Service at 1-888-237-2762, Monday  Friday, 8:00 AM  5:00 PM Eastern Time. |  | 166240 | 166240 units | Distributors nationwide. | Terminated | April 06, 2010 | 134 | 166240 | Not_Computer
Z-0582-2010 | 53875 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Delta/Delta XL Monitors with Omega S Package |  |  | November 09, 2009 | 2009 | Draeger Medical, Inc. | monitor is rebooting/resetting | Software design | Draeger Medical Systems, Inc. issued and Urgent Medical Device Recall letter dated November 2009 to their affected customers informing them of the problem and a work around until software can be validated. A Technical Service Bulletin has been generated and distributed globally that instructs the affected International Drager subsidiaries to inform customers in their countries.For further information, contact Draeger Medical Systems, Inc. at 1-800-543-5047, extension 1-2349. |  | 4188 | 4,188 units | The product was shipped to medical facilities in CA, CO, CT, ID, IN, KY, MS, NY, OK, and WI. The product was also shipped to United Arab Emirates, Argentina, Austria, Australia, Bosnia Herz, Belgium, Benin, Brunei Darussalam, Bolivia, Brazil, Belarus, Canada, Switzerland, Chile, China, Columbia, Costa Rica, Cyprus, Czech Republic, Germany, Denmark, Egypt, Spain, Finland, France, Great Britain, Greece, Hong Kong, Croatia, Hungary, Ireland, India, Iran, Italy, Japan, Kenya, Korea, Kuwait, Kazakhstan, Lebanon, Lithuania, Latvia, Libya, Morocco, Mexico, Malaysia, Netherlands, Norway, Nepal, Oman, Poland, Portugal, Qatar, Reunion, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovak Republic, Thailand, Turkey, Taiwan, Ukraine, Venezuela, Vietnam, and South Africa. | Terminated | August 23, 2011 | 652 | 4188 | Software | Device Operation | Sofrware validation | Software Update
Z-0579-2010 | 53871 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Invacare Storm TDX SR Power Wheelchair with Stability Lock Model numbers: TDXSR, TDXSR-CG, TDXSR-CG-HD, TDXSR-HD, TDXSR-MCG, TDXSR-MCG-HD, TDXSRV-HD |  |  | September 15, 2009 | 2009 | Invacare Corporation | For some Invacare chairs, the stability lock feature may not be engaging properly or consistently. If the stability lock feature does not engage properly, the wheelchair may do one or more of the following: (1) veer to one side; (2) rock forward onto its front riggings; (3) drive in an unintended circular motion; and/or (4) fall forward or to one side. In these cases there is a risk to injury for | Device Design | On 9/15/2009 the firm sent recall notification letter to their customers instructing them to immediately contact the dealer to have the stability lock inspected and/or adjusted. If the dealer is unavailable or the repair will take longer than 60 days, they are to contact the firm and arrange for an inspection and/or adjustment. Invacare requested contact names to whom it can send additional information. Invacare plans to mail, via FEDEX, information regarding the recall to each consignee. Initial consignees received a response card and instructions to forward cards to any customer to whom they have distributed the product. The products are to be inspected at the cylinder label and if the label does not have a long alpha-numeric string demonstrated in the label, the cylinders on the chair have to be replaced. Questions are directed to Recalls@Invacare.com or to 800-333-6900. |  | 36 | 36 | Product was distributed to the following states: AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, MA, ME, MI, MN, MO, MS, NC, ND, NM, NY, OH, SC &amp; TX. Product was also distributed in Canada. | Terminated | August 25, 2014 | 1805 | 36 | Not_Computer
Z-0585-2010 | 53937 | No | CCL | Oxygen gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Oxygen Sensor #5804, Distributed by: Teleflex Medical, Research Triangle Park, NC 27709 |  |  | November 25, 2009 | 2009 | Teleflex Medical | Sensor causes the oxygen monitor to display prematurely a Low Sensor warning which indicates an early maintenance requirement for the instrument. This may result in interrupted treatment. | Vendor change control | Consignees were notified via next day delivery on or about 11/25/2009. Each customer received a letter requesting acknowledgement of the recall and a request to return the affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. A second and third mailing to non-responding customers on the consignee list will be made, or as an alternative, the third contact may be by phone, e-mail, or fax. Affected products received by Teleflex Medical will be destroyed or reworked and replacement product or credit will be provided to the consignee. Records of contacts and of destruction or rework will be maintained. Questions are directed to the company at 1-866-246-6990. |  | 51 | 51 eaches | FL, KY, MI, OH, PA, SD, TN, Australia and Canada. | Terminated | April 07, 2011 | 498 | 51 | I/O | Alarm/Message | Replace | Remove or Replace
Z-0583-2010 | 53919 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia S Series SPECT System; Single-Photon Emission Computed Tomography; Siemens Medical Solutions USA, Inc., Hoffman Estates, IL 60195-2061. |  |  | December 11, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Some of the hand controllers of the Symbia S &amp; T systems are missing a resistor switch, and the absence of the switch may lead to unintended system motion during camera set up activities. | Process control | Siemens issued an Urgent Field Correction Recall letters dated November 18, 2009 and December 11, 2009 to the affected users of the Symbia S and Symbia T Systems, respectivley. The customers were advised of precautionary measures and that Siemens will replace the affected device. They also were advised that a Siemens Customer Service Representative would contact them within a few days to schedule a replacement of the device. For further information, contact Siemens Medical Solutions USA, Inc. at 1-800-888-7436 (USA). |  | 135 | 135 units | Worldwide Distribution -- United States, Algeria, Australia, Austria, Belgium, Canada, Chile, China, Finland, France, Germany, Hong Kong, India, Indonesia, Italy, Korea, Mexico, Netherlands, Norway, Philippines, Portugal, Qatar, Russia, Saudi Arabia, Spain, Sweden, Taiwan, Turkey, Venezuela and Vietnam. | Terminated | August 27, 2012 | 990 | 232 | Hardware | Physical Safety Hazards | Replace | Remove or Replace
Z-0490-2010 | 53845 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | LACTOSORB SYSTEM 1.5MM ADJUSTABLE SELF-DRILLING TAP FOR POWER DRIVER, REF: 915-2185, NON-STERILE PRODUCT, |  |  | September 15, 2009 | 2009 | Biomet Microfixation, Inc. | Biomet Microfixation is recalling LactoSorb Self Drilling Adjustable Tap Model Numbers 915-2075 Lot # 064350 and 915-2185 Lot # 885580. The two lots identified have a new style stop and an old style tap; they are not compatible, as the depth etchings do not line up correctly. Depending on the tap (1.5mm VS 2.0mm) you may drill too shallow or too deep. | Component change control | Biomet Microfixation notified consignees of the recalled product beginning September 2009 via telephone and e-mail correspondence. Users were asked to return the affected product to the firm.For further information, contact Biomet Microfixation at 1-904-741-9468. |  | 7 | 7 pieces | USA - FL, MO, and KS; Foreign: Brazil and China. | Terminated | January 23, 2010 | 130 | 22 | Not_Computer
Z-0589-2010 | 53843 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Maquet Servo Ventilator SV900C/D/E (SV900) |  |  | November 17, 2009 | 2009 | MAQUET Inc. | Some Servo Ventilators 300/300A and Servo Ventilator 900C/D/E should not be used with a system that may generate negative pressure below -100 cm H2O (Closed system suctioning) due to the risk of damaging the ventilator's pressure transducers. | Device Design | Maquet sent Device Correction letters commencing 11/17/09 and ending 11/23/09 by Federal Express. |  | 29364 | 29364 (8438 of those in US) | Worldwide distribution | Terminated | June 04, 2012 | 930 | 50201 | Not_Computer
Z-0590-2010 | 52207 | No | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Hocoma Lokomat-System with Pediatric Module |  |  | May 27, 2009 | 2009 | Hocoma AG | Changes to the stored "sensitivity" values may occur and leads to higher limits in the force and trajectory deviation | Software design | Hocoma issued a letter dated May 27, 2009 to Lokomat users informing them of the affected device. Instructions were provided to avoid further problems with the product and users were asked to complete and return the provided Declaration to Field Safety Notice.For further information, contact Hocoma at info@hocoma.com. |  | 7 | 7 | Nationwide Distribution -- including IL, MA, MI, NY and OH. | Terminated | February 22, 2010 | 271 | 7 | Software
Z-0504-2010 | 53386 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Accolade System Surgical Technique; |  |  | August 31, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Certain surgical protocols do not provide specific instructions for initiating power to the calcar planer prior to engagement with bone. | Error in labeling | "Urgent Product Correction - Surgical Protocol" letters and Product Accountability Forms were sent to consignees via Fed Ex on August 31, 2009 and September 1, 2009 with return receipt. Notification letters described the issue, protocols affected, potential hazard and harm and risk mitigation. Direct questions about the recall by calling 1-201-972-2100 or Customer Service 1 -866-OR-ASSIST (1 -866-672-7747). |  | 5670 | 5,670 | Worldwide Distribution. | Terminated | July 17, 2013 | 1416 | 14396 | Not_Computer
Z-0592-2010 | 53957 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Signa Excite 1.5T, Signa HD 1.5T, Signa HDx 1.5T, Signa HDxT 1.5T, Signa HDe, GE DISCOVERY MR450 and the Signa 3.0T MR750 MRI systems when used in conjunction with Confirma's CADstream (These systems contain the PURE software option.) |  |  | October 06, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has recently become aware that the use of Confirma, Inc.'s CADstream in conjunction with the GE Healthcare's PURE (Phased Array Uniformity Enhancement) software for dynamic MRI imaging (Signa Excite 1.5T, Signa HD 1.5T, Signa HDx 1.5T, Signa HDxT 1.5T, Signa HDe, GE DISCOVERY MR450 and the Signa 3.0T MR750 MRI systems ) may require a modification to the study preference settings to p | Software design | GE Healthcare issues an "Important Product Information" letter dated October 23, 2009. The letter was addressed to Hospital Administrators, Managers of Radiology and Radiologists. The letter listed the Issue, Affected Product Details, Instructions, Product correction and Contact Information.For further information, contact GE Healthcare at 1-262-548-2731. |  | 5460 | 5460 | Worldwide Distribution -- All States in the US including DC and Puerto Rico. OUS: ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BELGIUM, BOLIVIA, BOSNIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, ISLAMIC IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KOREA REPUBLIC, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB, LITHUANIA, LUXEMBOURG, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB, UNITED KINGDOM, UZBEKISTAN, VENEZUELA and VIETNAM. | Terminated | August 03, 2012 | 1032 | 5460 | Software
Z-0594-2010 | 53983 | Yes | OAD |  |  | Cardiovascular | PMA | Powerlink System with IntuiTrak Delivery System (Endovascular Graft): Model Number 34-34-80LE, LOTS W09-2967 AND W09-3000. |  |  | November 25, 2009 | 2009 | Endologix Inc | This action is being taken because the product has the potential for the polyimide tubing to detach from the inner core due to the bond failure. If this occurs, treatment may be delayed or interrupted. The firm's initial health hazard analysis indicated the event may cause temporary or medically reversible adverse health consequences or an outcome where the probability of serious adverse health | Process design | The recall communication was initiated on 11/25/09 by phone call and followed by recall letter via Fed-Ex Overnight. The recall letter informed the consignees of the reason for recall and products affected. Users were instructed to share this information with physicians that performed the affected procedures at their facility. Endologix informed their consignees that their personnel would retrieve the affected devices and return it to Endologix at no cost and would replace the affected units. Consignees with questions or comments about the notice are directed to (800) 983-2284. |  | 28 | 14 UNITS DISTRIBUTED: 3 LOT W09-2967 AND 1 LOT W09-3000 STILL AT CONSIGNEES. 10 UNITS USED. | Recalled devices were distributed to medical centers and hospitals in OR, PA, FL, TX, MI and LA. | Terminated | May 07, 2010 | 163 | 34 | Not_Computer
Z-0591-2010 | 53949 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Bausch &amp; Lomb SoFlex SE Foldable Intraocular Lens. |  |  | November 19, 2009 | 2009 | Bausch &amp; Lomb Inc | Bausch &amp; Lomb Incorporated, Rochester, NY is recalling their Silicone L161SE Intraocular Lens Lot 4916928 (23.5D) after receiving three customer complaints for what appears to be opacification of the intraocular lens. The opacification occurred within one week of implantation. In each reported case the surgeon explanted the lens. | Other | Bausch &amp; Lomb, Inc. notified consignees via phone beginning November 19, 2009 and by follow-up letter dated November 23, 2009. Titled "Urgent - Medical Device Recall", the letter instructed users to discontinue implantation of the affected product and return any unused lenses along with the enclosed acknowledgement form to the firm.For further information, contact Bausch &amp; Lomb at 1-585-338-6612. |  | 85 | 85 lenses | Lenses were distributed to one distributor and 35 retailers. Worldwide Distribution -- Distributed to Canada, Taiwan and throughout the United States. | Terminated | July 22, 2010 | 245 | 85 | Not_Computer
Z-0595-2010 | 53982 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Stryker MIS Bur, 2.0mm Neuro Diamond, sterile, Stryker Instruments, Kalamazoo, MI; REF 8420-107-120. |  |  | July 06, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | The bur may fracture, resulting in fragments within the surgical site, injury to the patient or surgical team, or both. | Device Design | Consignees were sent a letter dated 7/10/09 instructing them to locate and return product on hand for replacement. Surgeons who use the product were notified via a separate, untitled letter, dated July 13, 2009. A speed warning will be added to the product labels. Questions are directed to the firm at: 1-800-800-4236 ext. 3808. |  | 5438 | 5438 of all products | Nationwide distribution: Australia, Canada, England, France, Hong Kong, Japan, Netherlands, South Africa, Sweden, Switzerland and Taiwan. | Terminated | October 25, 2010 | 476 | 5438 | Not_Computer
Z-0605-2010 | 53781 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | AMO Tecnis 1-Piece Intraocular Lens (Model ZCB00) |  |  | November 11, 2009 | 2009 | Abbott Medical Optics Inc (AMO) | Abbott Medical Optics (AMO) has received complaints of some Tecnis 1-Piece IOLs (model ZCB00) have adhered to the lens optic after insertion into the eye. This can lead to a need for additional manipulation to separate the haptics from the optic during surgery which could lead to an injury. | Process design | Beginning on Friday, November 6, 2009, AMO began distribution of the Recall Notification Letter via Federal Express to the 874 AMO Tecnis 1-piece IOL customer accounts in the US. Included with the Recall Notification letter is a facsimile delivery confirmation that each customer was instructed to fax back to AMO to confirm receipt of the Recall Notification Letter. --- Beginning on Monday, November 9, 2009, AMO began distribution of the Recall Notification letter to the AMO Tecnis 1-piece IOL customer/distributors accounts outside the USA."THIS RECALL NOTIFICATION DOES NOT INVOLVE TECNIS 1-PIECE IOL THAT HAVE BEEN IMPLANTED. ONCE THE HAPTICS HAVE BEEN SUCCESSFULLY DEPLOYED, THESE LENSES FUNCTION PROPERLY."Questions should be directed to an AMO Customer Service Representative at 1-877-AMO-4LIFE (1-877-266-4543). |  | 246000 | 246000 Lenses | AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY, Hawaii, and Puerto Rico.INTERNATIONAL: Austria, Azerbaijan, Belgium, Canada, Switzerland, Cyprus, Czech Republic, Germany, Estonia, Spain, Finland, France, United Kingdom, Hungary, Ireland, Israel, Italy, Lebanon, Lichtenstein, Libya, Martinique, Netherlands, Poland, Palestine, Portugal, Saudi Arabic, Sweden, Slovenia, Turkey, and South Africa. | Terminated | July 25, 2012 | 987 | 246000 | Not_Computer
Z-0607-2010 | 54031 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Medtronic Sofamor Danek Extended Length Trocar, Size : L=50MM, Item No.: 7578014, Mat'l: Stainless Steel, non-sterile, Rx only, Trocars used with the CD HORIZONÂż LEGACY" Percutaneous Peek Spinal System |  |  | December 07, 2009 | 2009 | Medtronic Sofamor Danek USA Inc | The Trocar may not fully engage with the Rod Inserter or may prematurely disengage. | Device Design | On 12/7/2009, the Sales Representatives were notified by conference call and email regarding the recall. The recall notices were mailed directly to Risk Managers and Sales Representatives, with Recall Questionnaires and Confirmation forms included respectively, on 12/09/2009 with a delivery date of 12/10/09. |  | 266 | 266 | Nationwide distribution, and the Netherlands | Terminated | January 14, 2011 | 403 | 465 | Not_Computer
Z-0616-2010 | 53788 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Omega IV and Omega V tables on GE Advantx-E imaging system, used for general purpose diagnostic angiographic fluoroscopy and radiographic studies. |  |  | December 22, 2009 | 2009 | GE Medical Systems, LLC | Potential table rotation of Omega IV and Omega V tables affecting Advantx-E and Innova 2000, 2100IQ, 3100, 3100IQ, 4100, 4100IQ, 2121IQ and 3131IQ cardiovascular X-ray imaging systems.Due to potential weakness of some table rotational brakes, an unintended table rotation may occur while the table is locked. This issue could lead to a patient fall if the unintended motion occurs during patient | Device Design | Consignees were sent on 12/22/09 a GE Healthcare " Urgent Medical Device Correction" letter dated 12/1/09. The letter was addressed to Hospital Administrators / Risk Mangers, Mangers of Radiology / Cardiology, Radiologists / Cardiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. |  | 1502 | 1502 | Worldwide distribution: USA, SRI LANKA, ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, BELARUS, BELGIUM, BOLIVIA, BOSNIA &amp; Herzegovina, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, GERMANY, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and YUGOSLAVIA. | Terminated | November 13, 2015 | 2152 | 3635 | Not_Computer
Z-0611-2010 | 54114 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet 2.9 Micromax L-15 suture anchor, flexible long shaft, #2 blue/white maxbraid, PLLA/PGA L-15, sterile, Biomet Sports Medicine, Warsaw, IN; REF 905456. |  |  | December 11, 2009 | 2009 | Biomet, Inc. | The anchor label states it contains a flexible shaft, but this lot actually contians a rigid shaft. | Labeling mix-ups | Biomet Sports Medicine notified the single consignee by letter dated December 11, 2009. The firm requested that users locate and return the affected product and return a business reply card.For further information, contact Biomet at 1-800-348-9500 or 1-574-372-3983. |  | 6 | 6 | Nationwide Distribution -- Alabama. | Terminated | October 18, 2010 | 311 | 6 | Not_Computer
Z-0623-2010 | 53953 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Cordis Aquatrack Hydrophilic Nitinol Guidewire, for use in angiographic procedures, REF C3526-SSA, Guidewire Diameter .035" Tip Length Flexible 3 cm, Sterile. |  |  | August 25, 2009 | 2009 | Cordis Corporation | Some of the guidewires in these two lots have a straight tip (incorrect) rather than an angled tip (correct configuration for that catalog number). | Process control | Recall for Lot 80000521 was initiated on August 25, 2009 with Cordis Medical Device Recall letters being mailed for delivery on August 25, 2009. Response form was also provided to obtain product disposition.Recall for Lot 80000584 was initiated on October 19, 2009 with Cordis Medical Device Recall letters being mailed for delivery on October 20, 2009. Response form was also provided to obtain product disposition. |  | 145 | 145 units | Nationwide distribution. | Terminated | August 11, 2010 | 351 | 145 | Not_Computer
Z-0631-2010 | 53980 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | GE ApexPro Oximeter systems (ApexPro CH or ApexPro FH with ApexPro Telemetry version 3.9 software) |  |  | September 25, 2009 | 2009 | General Electric Medical Systems Information Technology | Potential misuse of the ApexPro Telemetry System and incorrect message and alarm information in the ApexPro Operator's Manual may impact patient safety. When ApexPro Telemetry systems are used with SpO2 oximeters:1. ApexPro FH Transceiver can operate with the ApexPro FH Xpod connected, but without ECG lead wires attached to the transceiver. This is unintended use of the device that is not do | Labeling design | Consignees were sent on 10/27/09 a GE Healthcare "Urgent Medical Device Correction" letter dated September 25, 2009. The letter was addressed to Risk Manger, Chief of Nursing, director of Biomedical Engineering. The letter described the Safety Issues, Affected Product Details, Product Correction, Safety Instructions and Contact Information. "ApexPro Telemetry System Operator's Manual Supplement for SpO2 was enclosed with the letter. |  | 8305 | 8305 | Worldwide distribution: USA, AUSTRALIA, AUSTRIA, BELGIUM, BERMUDA, CANADA, CHILE, CHINA, COSTA RICA, CROATIA, CUBA, CYPRUS, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUAM, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LEBANON, MALAYSIA, MALTA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PHILIPPINES, POLAND, QATAR, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED KINGDOM, and VENEZUELA. | Terminated | May 24, 2012 | 972 | 8305 | Other | Physical Safety Hazards |  | Other
Z-0622-2010 | 53981 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Foundation Tibia Baseplate, Catalog number 324-01-108. The tibial is modular in design, such that plastic inserts of various thicknesses and intramedullary stem are attached to the metal baseplate intraoperatively. |  |  | December 03, 2009 | 2009 | Encore Medical, Lp | Product does not meet specifications for a size 8 Right Foundation Tibia Baseplate. | Mixed-up of materials/components | Firm notified consignees by an Urgent Field Safety Notice dated 12/3/2009. The letter identified the product, stated the problem, and the risk associated with the problem. Customers are to immediately return the affected product and notify all of their customers if product was further distributed. Customers are to complete and return the enclosed response form. Questions should be directed to David Harris, Manager of Customer Service at 512-834-6330 or Valerie Moulton, International Customer Service at 512-834-6275. |  | 17 | 17 | Worldwide Distribution -- USA, including states of GA, OK, FL, CA, and SD, and countries of Germany and Turkey. | Terminated | May 05, 2010 | 153 | 17 | Not_Computer
Z-0609-2010 | 53994 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific COGNIS CRT-D, Cardiac Resynchronization Therapy Defibrillator, models N118, N119. the following models are not available in the US: models N106, N107, P106, P107, P108. Sterilized using ethylene oxide. Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. |  |  | December 01, 2009 | 2009 | Boston Scientific CRM Corp | BSC CRM provided important patient management information regarding Boston Scientific COGNIS CRT-Ds and TELIGEN ICDs. Engineering simulations and field reports indicate that in some cases, implanted devices may encounter sufficient mechanical stress to weaken the bond between the header and case when positioned subpectorally. A weakened header bond may alter lead impedance and introduce noise | Device Design | Boston Scientific CRM issued an "Urgent Medical Device Correction" letter dated December 01, 2009. The letter described the issue and potential device behavior. The letter provided recommendations to avoid issues relating to the affected device until improved devices become available.For further information, contact your local Boston Scientific CRM representative, United States Technical Services at 1-800-227-3422, or European Technical Services at +32-2-416-7222. |  | 39873 | 39,873 | Worldwide Distribution -- United States including Guam, Puerto Rico, and U.S. - Virgin Islands. OUS to include: Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Belgium, Bermuda, Brazil, Canada, Cayman Islands, Chile, China, Cocos (Keeling) Islands, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Finland, Gabon, Germany, Greece, Hong Kong, Hungary, India, Iran, Ireland, Israel, Jordan, Kuwait, Latvia, Lebanon, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Trinidad &amp; Tobago, Tunisia, United Arab Emirates, United Kingdom and Venezuela. | Terminated | November 06, 2012 | 1071 | 100084 | Not_Computer
Z-0637-2010 | 53890 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Gyrus ACMI PKS Seal Open Forceps, model 915010PK, 9-3/4" Angled Jaw. Sterile EO. The device is intended to be used with the G400 electrosurgical generator. Gyrus ACMI, Inc., 136 Turnpike Road, Southborough, MA 01772-2104 USA. |  |  | August 18, 2009 | 2009 | Gyrus ACMI Corporation | Gyrus ACMI received complaints of a problem with 3103PK, 3104PK, 3105PK, 915000PK, 915005PK, 915010PK - Gyrus ACMI PKS Seal Open Forceps. The metal shim between the jaws may detach during the procedure and fall into the surgical field. The detached shim may be undetected and be left behind in the patient. This circumstance could subsequently cause complications. Further use of this product sh | Environmental control | An Urgent Medical Device Recall letter, dated 08/15/09 was sent to customers notifying them of the issue. The letter identified affected product, asked that they cease further use, remove affected devices from stockrooms, contact Gyrus ACMI customer service for a return goods authorization and return product to Gyrus ACMI at no charge. A reply form is also asked to be returned to Gyrus ACMI via fax.Questions and requests for a Return Goods Authorization are directed to Gyrus ACMI customer service at 1-888-524-7266. |  | 93665 | 93,665 total (63,418 US, 30,247 OUS) | US and International distribution. OUS to Include: Australia, Belgium, Bolivia, Bulgaria, Canada, Chile, China, Columbia, Croatia, Czech, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece , Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Panama, Puerto Rico, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, United Kingdom. | Terminated | March 06, 2012 | 931 | 93665 | Not_Computer
Z-0642-2010 | 54020 | No | LYJ |  |  | Neurology | PMA | VNS Therapy Demipulse Generator, Model 103 and VNS Therapy Demipulse Duo Generator, Model 104. |  |  | November 20, 2009 | 2009 | Cyberonics, Inc | Under certain conditions, product's battery life can be reduced. | Other | Firm sent "An Important Safety Notice" letter to all physicians explaining events, risks presented and recommended actions. Letter was mailed November 20, 2009. Physicians should complete and return the effectiveness card as soon as possible to acknowledge that you have read and understood the Safety Alert Letter. If further information is needed, customers in the US can contact Clinical Technical Support at 866-882-8804 or via email at clinicaltechnicalservices@cyberonics.com. |  | 8667 | 8,667 units | Worldwide Distribution -- USA, Argentina, Austria, Belgium, Canada, Columbia, Costa Rica, Croatia, Czech Republic, Denmark, Finland, France, Greece, Hungary, Ireland, Iceland, Israel, Italy, Kuwait, Lebanon, Luxembourg, Netherlands, Norway, Panama, Portugal, Republic of Germany, Saudi Arabia, Sweden, Slovakia, Spain, Switzerland, Taiwan, UAE, United Kingdom, and Venezuela. | Terminated | July 24, 2012 | 977 | 8667 | Battery |  | Notification | Safety Notice/Insructions
Z-0630-2010 | 54000 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sorin / Clearglide Precision Bipolar, REF PBD01, Sterile R, Manufactured for: Sorin Group USA, Inc., 14401 W.65th Way, Arvada, CO 80004. |  |  | November 25, 2009 | 2009 | Sorin Group USA, Inc. | Vessel clamping device may break during use. | Component change control | Consignees were notified by "Urgent Field Safety Notice" dated November 24, 2009 with instructions to examine the device before, during and after use. Customers were instructed to not use the device, if it is damaged. Customers were also instructed to contact the Sorin Group Customer Service at 1-800-650-2623 to arrange shipment of replacement product and requested to complete and return the response form. Direct questions to the Sorin Group Customer Service at 1-800-650-2623. |  | 756 | 756 units | Worldwide Distribution -- United States, Australia, Canada and Italy. | Terminated | October 04, 2010 | 313 | 5817 | Not_Computer
Z-0643-2010 | 53755 | No | KZI | Enriched culture medium. | Microbiology | Microbiology | 510(K) Exempt | BBL (TM) Brain Heart Infusion, Catalog number 221778, package of 100 tubes, 0.5 mL tubes. Product is labeled in part *** Becton Dickinson and Company, 7 Loveton Circle, Sparks, Maryland 21152 USA*** 800-638-8663 |  |  | February 19, 2009 | 2009 | Becton Dickinson &amp; Co. | In vitro diagnostic test reagent may contain residual material that inhibits the growth of gram positive/gram negative organisms. | Equipment maintenance | The recalling firm notified direct accounts by an Urgent Product Recall letter on 02/19/09. The notification identified the affected product and stated the problem. The letter also advised that users discontinue use of the product and to discard any remaining packages. The attached response form should be completed and returned for replacement of product. If further assistance is needed regarding replacements, customers can contact BD Customer Service Department at 1-800-675-0908. |  | 144 | 144 boxes | Nationwide Distribution -- Including states of CA, CO, FL, GA, IL, MI, MN, MT, NY, OH, PA, TN, TX, UT, and VA. | Terminated | February 12, 2010 | 358 | 144 | Not_Computer
Z-0646-2010 | 53945 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Vanguard Lock-On Femoral Impactor, Biomet Orthopedics. Warsaw. IN; REF 32-486206. |  |  | December 01, 2009 | 2009 | Biomet, Inc. | The locking hook tabs may fracture and fragments remain in the patient post-operatively. | Device Design | Urgent Medical Device Recall Notices were sent to distributors on 12/01/2009 with a notice for them to provide to medical facilities where distributor-owned inventory is stored. Distributors were instructed to immediately locate and remove the instruments from circulation; follow the instructions on the "FAX Back Response Form" and fax a copy back prior to returning the product; and to return the product to the firm. Questions should be directed to 800-348-9500. |  | 92 | 92 | Worldwide Distribution -- USA, Canada, China, Costa Rica, Denmark, Germany, Netherlands, New Zealand, and United Kingdom. | Terminated | October 18, 2010 | 321 | 92 | Not_Computer
Z-0652-2010 | 53960 | No | BZA | Airway connector. | Anesthesiology | Anesthesiology | 510(K) Exempt | Hudson RCI One Way Valve, 22mm E.D. to 22mm O.D., REF 1665, Rx Only, Teleflex Medical, Research Triangle park, NC 27709 |  |  | December 02, 2009 | 2009 | Telefelx Medical | Products 1664 and 1665 are each one-way valves that offer low flow resistance, each in one, opposite direction. Each of these valves demonstrate the appropriate direction with a red arrow.The reason for this recall is incorrect packaging; product code 1664 was incorrectly packaged as product code 1665.If the red arrows are not followed, the valve may be placed in the circuit in the wrong | Packaging | Consignees were notified by letter via next day delivery. Each consignee was sent a letter requesting acknowledgement of the recall and a request to return all of the affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. A second mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the second mailing or as an alternative, they may be contacted by phone, e-mail, or fax. A record of these contacts will be maintained. Affected products received by Teleflex will be destroyed or reworked and replacement product or credit will be provided to the consignees. Records of destruction or rework will be maintained.Questions are directed to the firm's sales representatives or to Teleflex Customer Service at 1-866-246-6990 |  | 2450 | 2450 eaches | CA, IL, NE, NJ, TX, VA, WI and Canada. | Terminated | August 11, 2010 | 252 | 2450 | Not_Computer
Z-0644-2010 | 54042 | No | JTZ | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | remel Mueller Hinton Agar w/4% NaCl w/6 mcg/ml Oxacillin, Ref R01626, 10/box. The firm name on the label is Remel, Lenexa, KS. |  |  | December 08, 2009 | 2009 | Remel, Inc | The product may fail to adequately grow methicillin resistant Staphylococcus aureus ATCC 43300 and Staphylococcus aureus ATCC 33591. | Nonconforming Material/Component | The recalling firm issued two different recall letters dated December 8, 2009 via regular mail. One letter was addressed to Microbiology Laboratory Supervisor and the other to Distributors. Both letters explained the reason for recall, requested the customer discard all remaining units in their possession, and complete the enclosed Product Inventory Checklist to indicate the amount of product that was discarded. The letter issued to distributors also requests they notify their customers to the laboratory level and the Product Inventory Checklist has a block to be marked indicating they have notified their customers. Direct questions to the Technical Services Department at 1-800-255-6730 (United States) or 1-913-888-0939 (International). |  | 267 | 267 units | Nationwide. | Terminated | July 12, 2011 | 581 | 267 | Not_Computer
Z-0641-2010 | 53968 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Orthofix Spinal Implants; Pillar SA PEEK Spacer System Instrument Case. |  |  | November 19, 2009 | 2009 | Orthofix, Inc | The Pillar SA Implant Insertion Instrument 49-1100 may experience resistance in turning the activation knob while releasing an implant after intra-operative positioning. | Nonconforming Material/Component | A notification letter dated November 19, 2009 was sent to customers via Federal Express. The letter described the recalled product (PILLAR SA Implant Insertion Instrument, 49-1100) and problem and actions for customers to take. The affected product, PILLAR SA lmplant lnsertion lnstrument, 49-1100 is in the PILLAR SA set (set part number 49-0021) which was shipped to customers. An enclosed attachment provide instructions on product removal and return. Customers were instructed to complete the enclosed Tracking and Verification Form and fax to Orthofix Spinal Implants at 1-973-406-2912. A replacement instrument will be provided as available. In the interim, the Implant Inserter (Part Number 49-1000) included in the PILLAR SA UPE set (set number 49-0020) that was provided earlier this year may be used as an alternative to the affected product. For questions regarding the removal and return of the recalled product to Orthofix Spinal Implants, please contact your Account Services Representative.Direct other questions to the Regulatory Affairs Department at 1-973-406-2812. Direct product questions to Marketing at 1-862-221-1335 or Product Development at 1-973-406-2824. |  | 29 | 29 instruments | United States (IA, NV, CO, AZ, NY, MD, IL, VA, MN, MI, NC, CT, LA, AR, GA, CA, MO, WA, OK and AL). | Terminated | February 02, 2010 | 75 | 29 | Not_Computer
Z-0651-2010 | 53940 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical Medfusion model 3010a syringe Infusion pump. REF 3010A. Rx Only. Smiths Medical MD, Inc., St. Paul, Minnesota 55112 (USA) |  |  | November 16, 2009 | 2009 | Smiths Medical ASD, Inc. | Smiths Medical has become aware of an increased trend in reports of Motor Not Running (MNR) and Motor Rate Error (MRE) alarm message events with certain MedfusionÂż Syringe Infusion Pumps, Models 3010A and 3500 (Pump). MNR/ MRE are high priority alarm messages which can occur for a variety of reasons as part of the usual self-test functionality of the Pump. The Pump is designed so that when its so | Device Design | An Urgent Medical Device Recall letter, dated November 16, 2009, was sent to customers beginning November 16, 2009. The letter described the issue, identified affected devices, and asked that affected devices be returned to Smiths Medical to be inspected, repaired, if necessary, and returned back to medical facility. A Confirmation Form is also asked to be completed and returned to Smiths Medical.To return Pumps or for questions regarding this Urgent Recall Notice, please contact Smiths Medical Customer Service Department at 1~866-270-7060.Please report any issues with these Pumps please contact Smiths Medical Technical Customer Service Department at 1866-216-8806. |  | 43 | 43 | AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY,OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV,including Guam, Puerto Rico. OUS to include: Canada, Chile, Czech Republic,, England, France, Germany, Hong Kong, Ireland, Israel, Lebanon, Malaysia, Poland, Saudi Arabia, Scotland, Singapore, Sweden, Switzerland, United Arab Emirates. | Terminated | December 18, 2011 | 762 | 5436 | Not_Computer
Z-0647-2010 | 53779 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Model 91367 SL 2200 Compact Monitor |  |  | November 06, 2009 | 2009 | Spacelabs Healthcare, Incorporated | Potential for monitor to shut down and not being able to power on again. | Component design/selection | On 11/6/09 the firm sent Urgent - Device Correction notifications to customers by certified mail informing them of the monitor failure and that a field service engineer will contact the customer to fix the affected device at no cost. On 11/13/2009 the firm's Global Tech Support sent notification via email to all foreign subsidiaries and distributors informing them of the monitor failure and that a field service engineer will contact the customer to fix the affected device at no cost. Questions are to be directed to Spacelabs Medical's Technical Support at 1-800-522-7025, and 1-425-657-7200 x5089 for those customers outside of the US. |  | 422 | 422 units | Worldwide Distribution | Terminated | July 21, 2011 | 622 | 3413 | Not_Computer
Z-0656-2010 | 54036 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Tornier Arthro Tunneler. For transossous fixation of tendons to bone in rotator cuff repair. |  |  | December 04, 2009 | 2009 | TAG Medical Corp. | Device failed to deploy the distal loop | Process change control | Medicsense, the US Agent notified the sole distributor Tornier Inc., Edina, MN of the recall and requested notification to users requesting to discontinue and return product. Letter to customers dated December 4, 2009. |  | 60 | 60 UNITS | Nationwide | Terminated | February 29, 2016 | 2278 | 60 | Not_Computer
Z-0655-2010 | 53199 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Complete SE Biliary Stent System, SC680L (6 mm x 120 mm), lot number 0000983806, UPC code 006133994186669, use by date 03/17/2011, manufactured by Medtronic CardioVascular, Santa Rosa, CA. |  |  | July 17, 2009 | 2009 | Medtronic CardioVascular | Mislabeling-- A lot of 5 mm x 12 mm stents were labeled as 6 mm x 80 mm stents and vice versa. | Labeling mix-ups | Medtronic, Inc. issued an "Urgent - Voluntary Medical Device Recall" communication dated July 17, 2009 asking consignees to segregate labeled lot numbers until a representative from the firm arrived to retrieve the products.For further information, contact your local Medtronic representative or call Medtronic, Inc. at 1-707-525-0111. |  | 20 | 20 units | Product was distributed to 24 consignees in TX, MO, PA, MN, NY, IL, CA, IN, AZ, TN, OH, MI, and CT. | Terminated | June 21, 2010 | 339 | 20 | Not_Computer
Z-0004-2010 | 53491 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE OEC 9900 Elite ESP |  |  | August 12, 2009 | 2009 | Ge Healthcare | A software defect - may result in a false indication and warning on the OEC 9900 Workstation and C-arm display that the X-Ray tube Anode is over-heated. This will also cause the displayed air kerma rate and cumulative air kerma to be inaccurate. | Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter will be sent to customers identifying the issues, interim solution, and permanent solution. Questions or concerns are to be directed to the GEHC OEC service team at 1-800-874-7378. |  | 1411 | 1411 units total | Nationwide and Canadian Distribution | Terminated | April 23, 2012 | 985 | 1411 | Software | Alarm/Message | Notification | Safety Notice/Insructions
Z-0577-2010 | 54019 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Ellex Motorised Safety Filter, Zeiss; Model number 6339585. |  |  | February 11, 2009 | 2009 | Ellex Medical Pty Ltd | Potential for two screws important in the design of the safety filter to come loose. | Radiation Control for Health and Safety Act | Ellex issued a bulletin stating the problem and identifying the affected products. Customers are to remove the Motorised Safety Filter from clinical use if the green laser light is visible while viewing through the microscope at any time during use of the Solitaire laser. The Bulletin also states that customers should contact an Ellex approved Service Engineer to investigate and perform the Field Corrective Action listed. The manufacturer will remedy the defect or bring the product into compliance free of charge. Ellex will inspect and rework all affected safety filters sold prior to February 2009. |  | 200 | 200 units total | Worldwide Distribution. | Terminated | September 07, 2012 | 1304 | 200 | Not_Computer
Z-0645-2010 | 53164 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare MAC 800 RESTING ECG ANALYSIS SYSTEM. |  |  | December 10, 2009 | 2009 | GE Medical Systems, LLC | GE Healthcare has become aware of an issue with MAC800 laser printouts. When more than one ECG report in the file manager is batch printed in laser print, the potential exists for an incorrect patient identification or name to be displayed on the print out. Misidentification of patient data potentially contributing to delayed or incorrect treatment is then possible. | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated December 10, 2009. The letter was addressed to Healthcare Administrator/Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.GE Healthcare will provide a software upgrade to resolve the issue. A GE Healthcare representative will contact customers to schedule implementation of the correction when it is available for installation at the customer facility.Direct questions to the GE Healthcare Technologies Support Center at 1-800-558-7044, enter Option 2 and ask for Technical Support or your local GE Field Service Representative. |  | 567 | 567 | Worldwide Distribution -- United States (CA, CO, CT, FL, IL,MD, NY, NC, OH, TN, TX, WA and WI), UNITED KINGDOM, UNITED ARAB EMIRATES, THAILAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVAKIA, SINGAPORE, ROMANIA, QATAR, PORTUGAL NETHERLANDS, LITHUANIA, KOREA, JAPAN, ITALY, ISRAEL, INDIA, HUNGARY, HONG KONG, GREECE, GERMANY, FRANCE, CZECH REPUBLIC, CROATIA, CHINA, BAHRAIN, AUSTRIA and AUSTRALIA. | Terminated | August 06, 2013 | 1335 | 567 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-0336-2010 | 53704 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Imager" II Angiographic Catheters, Sterilized with ethylene oxide gas, Manufactured for Boston Scientific Corporation, One Boston scientific Place, Natick, MA 01760-1537, USA, Made in Ireland, Unit 7,8 &amp; 9 Annacotty Business Park, Annocotty, Co. Limerick, Ireland: |  |  | October 22, 2009 | 2009 | Boston Scientific Corporation | The sterile barrier in the packaging (pouch) that contains the catheter may be breached. The breach of sterility could lead to a contamination of the device with subsequent transfer of infectious agents to the patient. | Packaging process control | Two Boston Scientific "Urgent Voluntary Medical Device Recall" letters were sent to consignees on October 29, 2009. The letters described the problem and the affected product. Consignees were instructed to cease use and distribution of the product and to complete and return the Account Reply Verification Tracking Form. The first letter with the heading, "Boston Scientific IMAGER II Angiographic Catheter (Peripheral Intervention)", was addressed to the Risk Manager/Field Action Contact. The second letter with the heading, "Imager II Urology Torque Catheters" was addressed to the Risk/Materials Manager. Direct questions about the recall to your local Sales Representative or Boston Scientific Corporation by calling at 1-763-494-1700. |  | 28321690 | 8845 boxes (5/box) US ; 3202 boxes (5/box) OUS | Worldwide Distribution -- United States (AL, AZ, AR, CA, CO,CT, DE, DC, FL, GA, HI, IL IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO,MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SC, TN, TX, UT, VT, VA, WA, WV and WI), United Arab Emir., Dutch Antilles, Austria, Australia, Belgium, Brazil, Canada, Switzerland, Chile, Colombia, Cyprus, Czech Republic, Germany, Denmark, Egypt, Spain, Finland, France, Great Britain, Greece, Hong Kong, Hungary, Israel, Italy, South Korea, Kuwait, Lebanon, Lithuania, Mexico, Malaysia, Netherlands, Norway, New Zealand, Poland, Puerto Rico, Portugal, Qatar, Russian Fed., Saudi Arabia, Sweden, Singapore, Slovakia, Syria, Thailand, Tunisia, Turkey, Trinidad, Tobago, Taiwan, and South Africa. | Terminated | December 12, 2011 | 781 | 28323022 | Not_Computer
Z-0587-2010 | 54153 | No | NVK | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | Dental | 510(k) | SoftLase Pro Dental Lasers: Zap SoftLase, SoftLase Pro, OrthoLase, and HygieneLase models, Model numbers ZA-SLP, ZA-ORT, ZA-HYG, G22002SLP, ZA-SOFTLASE2WSII, ZA-SOFTLASER2W, Manufactured by Zap Lasers LLC, Pleasant Hill, CA. |  |  | August 08, 2008 | 2008 | Zap Lasers LLC | The SoftLase Lasers are in need of compliance upgrades due to lack of a remote interlock connector, an emission delay, user guide labels and locations, and calibration procedures and schedule. | Radiation Control for Health and Safety Act | Zap Lasers sent a letter to affectedcustomers stating that their SoftLase diode laser is eligible for a compliance maintenance upgrade. The upgrades will be free of charge and will enhance the safety of the laser. Customers are encouraged to call the firm's customer service line at 888-876-4547 to schedule their RMA number and shipping instructions. The letter also lists the upgrades that are needed. |  | 1375 | 1375 units | Worldwide Distribution to 1397 consignees | Terminated | May 25, 2010 | 655 | 1375 | Not_Computer
Z-0495-2010 | 53883 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Infinia with Hawkeye 4, model numbers: H3000WC and H3000WD. GE Healthcare, Waukesha, Wisconsin 53118. |  |  | August 01, 2009 | 2009 | GE Medical Systems, LLC | Failure to provide the complete User Manualinformation. | Radiation Control for Health and Safety Act | A Field Modification Instruction will be released in February 2010. GE Healthcare will correct the Operator Manuals in forward production and the installed base units. |  | 343 | 343 total | Nationwide Distribution | Terminated | February 04, 2012 | 917 | 343 | Not_Computer
Z-0613-2010 | 53941 | No | GKG | Hematocrit measuring device. | Hematology | Hematology | 510(K) Exempt | Haematokrit 24 Centrifuge with 2050 Rotor |  |  | November 21, 2009 | 2009 | Hettich Instruments Management I | Centrifuge Rotor Failure: may crack, break, and explode outside the centrifuge. | Device Design | Hettich Instruments notified Distributors by email on 11/21/09, to the Request For Urgent Action notifications, dated 11/23/2009. There is a description of how to identify a recalled rotor as well as a description of the replacement metal rotor that is to be used in place of the list of centrifuges affected by this recall. Customers are to cease using the setup described and contact their distributor or the firm. The firm will replace the plastic housing with a metal housing. Questions should be directed to 978-232-3957. |  | 46 | 46 units total | Worldwide Distribution -- USA, including states of CT and IN, and country of Canada. | Terminated | October 25, 2011 | 703 | 46 | Not_Computer
Z-0638-2010 | 54203 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Exel Huber Needles (Right angle and Straight Models with Catalog Numbers: 26901CE, 26902CE, 26903CE, 26904CE, 26905CE, 26906CE, 26907CE, 26908CE, 26909CE, 26910CE, 26911CE, 26921CE, 26922CE, 26923CE, 26924CE, 26925CE, 26926CE, 26927CE, 26928CE, and 26929CE). |  |  | December 23, 2009 | 2009 | Exelint International Co. | Non-coring needles that, when inserted into an implanted port, were shown to core the silicone port's septum. Coring of port's septum may lead to: decreased effectiveness of port, infection, necrosis, swelling, and embolization, which may result in pulmonary embolism, limb ischemia, stroke, local extravasation of medication such as chemotherapy or other caustic material; may lead to local tissue | Device Design | Distributors have been advised to cease distribution of any recalled needles and infusion sets; to quarantine any remaining; to document amount on Customer Response Form; to return form to Exel (firm); to arrange for return of product to Exel for credit. Distributors also are advised to share recall notice with all affected customers. Contact phone: 727-827-1922x0 (M-F 9a-5p EST). Contact email: info@exelint.com |  | 849000 | 849,000 | US (VA TX NV KY OH NH MS NY OH NM DE WA AZ IL PA OR CO FL GA LA UT MI WI NV CA NJ MO IA) and Canada. | Terminated | June 06, 2011 | 530 | 2010360 | Not_Computer
Z-0662-2010 | 53802 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | AC Power Cords, 110V Manufactured by Electri-Cord Manufacturing Company in Westfield, PA. Power Cord #4100079 used with Kinair IV Model 22600, Kinair IV 60601 Model 216050, Triadyne Proventa Model 406000, Therapulse ATP/II Models 227800 and 227080.S, Kinair Medsurg Model 201001W, and Kinair Medsurg Pulse Model 40900. |  |  | November 04, 2009 | 2009 | KCI USA, Inc. | Power cords may crack and fail inside plug with potential for fire hazard. | Nonconforming Material/Component | Kinetics Concepts, Inc. (KCI) issued an "Urgent Medical Device Correction" letter dated November 4, 2009. Consignees were informed of the affected product and instructed to take necessary action to avoid any further issues associated with the product. For further information, contact KCI's technical support team at 1-800-275-4524 option #3. |  | 4637 | 4,637 units | Nationwide and Canada | Terminated | January 21, 2011 | 443 | 79267 | Not_Computer
Z-0716-2010 | 53955 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Outlook Pump Nitroglycerin Set w/Universal Spike &amp; Spin Lock Connector, 135 in |  |  | November 24, 2009 | 2009 | B. Braun Medical, Inc. | System error 9 alarms. A System Error 9 alarm indicates a potential pressure loss in the pump cassette delivery chamber (chamber closest to the patient). This alarm is generated during the pumps initiation sequence and will not allow the set to be used. This requires a new set to be obtained, thus creating a potential for delay in therapy. | Nonconforming Material/Component | The recalling firm issued a Product Removal Notice via US Postal Service Certified Mail with registered return receipt or FedEx to all customers and distributors in receipt of the affected product. The notices were sent 11/30/09 and 12/1/09. Direct customers also received letters addressed to Material Management that informed them of the recall and provided instruction for disposition of affected product in their inventory. Distributors received the direct customer recall letter along with a cover letter outlining their responsibilities to execute the recall and notify their customers.Questions are directed to the company's Customer Support Departments: US customers please call (800) 227-2862. Canadian customers, please call (800) 624-2920.International customers, please call (610) 691-5400. |  | 57480 | 57480 units | The products were shipped to distributors and medical facilities nationwide. The product was also shipped to Afghanistan, Canada and LU. | Terminated | February 14, 2011 | 447 | 5435249 | Not_Computer
Z-0653-2010 | 53554 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Edwards Lifesciences Aquarius System Model GEF09800. |  |  | March 16, 2009 | 2009 | Edwards Lifesciences, Llc | There have been reports of clinically significant fluid imbalance when a certain level of fluid imbalance is detected the Aquarius will trigger an alarm, however, users are able to override this alarm and continue therapy. It is possible to remove too much fluid from or replace too much fluid to the patient. In extreme cases, this may result in serious injuries or death. | Employee error | Edwards is continuing a field corrective action related to the possibility of fluid imbalance when using the Aquarius. In order to mitigate some of the risk associated with this problem, Edwards sent notifications to consignees advising of mitigation steps. Consignees were provided: a description of the affected product, problem, and potential hazard; advice on action to be taken; and revised device labeling. Additionally, consignees were notified of a planned software upgrade to prevent users from bypassing the fluid balance alarm more than five times in a 20-minute period. Although Edwards does not believe there is a need to stop using the Aquarius system at this time, clinicians should weigh the risks/benefits to the patients before continuing the use of the device Edwards first provided some of this information to consignees in a Field Safety Notice on March 16, 2009. On January 11, 2010, Edwards provided an Urgent Product Recall letter to consignees which updated and expanded the earlier communication. The FDA classified this as a Class I recall. |  | 135 | 135 units | Nationwide Distribution. | Terminated | January 18, 2012 | 1038 | 135 | Software | Physical Safety Hazards | Software update | Software Update
Z-0602-2010 | 53692 | No | KGI | Bone densitometer. | Radiology | Radiology | 510(k) | DPX-NT series, GE Lunar DEXA bone densitometers operating with enCORE software versions 12.0, 12.1, 12.2,13.0, &amp; 13.1. |  |  | February 01, 2009 | 2009 | Ge Healthcare | Failure to properly pre-indicate the technique factors to be used during a patient scan. | Radiation Control for Health and Safety Act | An Important Electronic Product Radiation Warning letter will be sent to customers. The letter will discuss the defect description and related hazards, the affected product details, safety instructions, and product correction. Customers are to contact their local GE Healthcare service representative with questions. |  | 1764 | 1,764 total | Nationwide Distribution | Terminated | December 31, 2011 | 1063 | 1764 | Not_Computer
Z-0748-2010 | 51104 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Jeanes Hospital Custom Cath Procedure Tray, K09T-08945A, Sterile EO, Merit Medical Systems, Inc., 1600 West Merit Parkway, South Jordan, Utah 84095. |  |  | January 29, 2009 | 2009 | Merit Medical Systems, Inc. | A manufacturing defect may prevent adequate sterilization of the entire device. | Process control | Applicable Merit sales representatives and distributors were notified by phone and letter beginning January 29, 2009 and instructed to immediately contact all affected consignees. Consignees were to immediately quarantine any remaining product and arrange for return to the firm.For further information, contact Merit Medical Systems, Inc. at 1-800-356-3748 extensions 4929 or 4998. |  | 157 | 157 units | Worldwide Distribution -- United States ( PA and TX) and Japan. | Terminated | June 16, 2011 | 868 | 910 | Not_Computer
Z-0749-2010 | 53607 | Yes | NIM |  |  | Cardiovascular | PMA | RX ACCULINK Carotid Stent System, 10x40 mm, Product Number: 1011344-40, Lot Number: 9081851. |  |  | October 01, 2009 | 2009 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | The recall was initiated after Abbott Vascular discovered, through internal testing, that the affected lot may not meet our quality specifications for catheter shaft tensile strength. | Process control | Abbott Vascular's Sales Representatives initiated customer contact with the identified consignees beginning October 1, 2009 . They provided the consignees with a recall letter with attached fax-back effectiveness check form. The recall letter instructed the consignees to work with their local account representative to review their inventory, complete the attached recall effectiveness check form and return identified products to Abbott Vascular.For further information, contact Abbott Vascular Customer Service at 1-800-227-9902. |  | 22 | 22 units | Nationwide Distribution. | Terminated | September 27, 2011 | 726 | 22 | Not_Computer
Z-0732-2010 | 54149 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Stryker CBCII ConstaVac Blood Conservagtion System w/ Quick Disconnect and 1/8in. (3.2mm) PVC Drain, sterile, Stryker Instruments, Kalamazoo, MI REF 225-28-914. |  |  | November 12, 2009 | 2009 | Stryker Instruments Div. of Stryker Corporation | The product may not be sterile due to package not being intact. | Packaging process control | Stryker issued a "Medical Device Recall Notification" via email to foreign consignees dated November 12, 2009. Stryker further issued recall letters via overnight mail to domestic consginees beginning November 20, 2009. These notifications informed consignees of the affected devices instructing them to destroy all affected product and return a completed Business Reply Form to the firm.For further information, contact Stryker Instruments at 1-800-800-4236 extension 3808. |  | 79617 | 79,617 | Worldwide Distribution. | Terminated | October 25, 2010 | 347 | 79617 | Not_Computer
Z-0757-2010 | 54138 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Total Stabilizer Femoral Component; Size # 2 |  |  | December 10, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Triathlon TS femoral components may have been damaged during manufacturing, possibly preventing the assembly of a stem extension. | Nonconforming Material/Component | An "Urgent Product Recall" Letter dated December 10, 2009 was sent by Federal Express to all Stryker Branches, Chief of Orthopaedics, Hospital Risk Management and Surgeons. The letter described the affected product, issue and potential hazards. Consignees were instructed to complete Product Recall Acknowledgement Form and return to Stryker Orthopaedics via fax at 1-201-831-6069. All questions regarding the recall should be directed to Stryker Orthopaedics at 1-201-972-2100. |  | 5 | 5 | Worldwide Distribution -- United States (DE, NC, OH. MI, TX, CA, MA, NJ, TN, KS, GA, FL, AL, IN, MN, MD, NE, AR, OK, CO, AZ, OR, and MA), Australia, United Kingdom and Canada. | Terminated | August 28, 2012 | 992 | 90 | Not_Computer
Z-0755-2010 | 53749 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | MPS Acacia Extension Set, 15cm Small Bore, Spin Nut w/BD Q-Syte, REF 385151 |  |  | October 19, 2009 | 2009 | Becton Dickinson Medical Systems | Manufacturing defect may allow air entry into a central venous catheter, possibly resulting in an air embolism. | Process control | Consignees were notified by B-D Urgent Medical Device Recall letter on 10/28/2009 and asked to return all unused affected lots. For information, call the company at 800-453-4538, option 2, extension 2341. Another letter was sent to all affected customers on 02/08/2010 notifying them of the extension of the recall to include Nexiva products. A response card was included to be returned whether or not product was on hand. B-D issued a press release on Febriuary 8, 2010. It can be found on www.bd.com at http://www.bd.com/contentmanager/b_article.asp?Item_ID=24486&ContentType;_ID=1&BusinessCode;=20001&d;=BD+Worldwide&s;=&dTitle;=&dc;=&dcTitle;= |  | 45350 | 45,350 units | Nationwide distribution. | Terminated | October 19, 2010 | 365 | 7895560 | Not_Computer
Z-0779-2010 | 54141 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Biomet bone mulch screw, 10.5mm x 20 mm, TI-6AL-4V alloy, sterile, Biomet Sports Medicine, Warsaw, IN; REF 907320. |  |  | December 22, 2009 | 2009 | Biomet, Inc. | The screw may not have the hex dimension inside of the head of the screw, which will prevent the device from being implantable. This may cause a delay in the procedure while another screw is obtained. | Employee error | Hospitals, distributors, and other customers were notified by letter dated 12/22/09 and instructed to locate and return these lots of product. Questions related to this notice should be directed to the firm at 800-348-9500 or 574-372-3983, Monday through Friday, 8AM to 5PM. |  | 169 | 169 | Arkansas, California, Delaware, Florida, Kansas, Kentucky, Michigan, Missouri, New York, Pennsylvania, South Dakota, Texas, Washington, Wisconsin, Chile, Germany, Italy, Japan and Norway. | Terminated | October 22, 2010 | 304 | 169 | Not_Computer
Z-0764-2010 | 54172 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | Percutaneous Catheter Introducer Set, VNUS Medical Technologies Cath Shth Intro 7F x 7cm, catalog number IS-7F07, Disposable, Single Use, Do Not Re-sterilize, Manufactured in USA for VNUS, 5799 Fontanus Way, San Jose CA 95138 USA; General: 408-360-7200. |  |  | January 06, 2009 | 2009 | Argon Medical Devices, Inc | One lot of Percutaneous Catheter Introducer Set contained incorrect size needle. | Other | Firm notified sole consignee by an Urgent Medical Device Recall certified letter dated 1/6/2010. The letter identified the product being recalled and asked the consignee to immediately quarantine any product remaining at their facility. The recall should be conducted to the user level. Affected products at the consignee's facility and with their customers, should be returned using the instructions on the response form. The response forms are to completed and returned by fax or e-mail as indicated. Questions or the need for additional information concerning this recall should be directed to Sarah Larson at 903-677-9351 or Sarah.Larson@argonmedical.com. |  | 500 | 500 units | Nationwide Distribution -- State of CA. | Terminated | July 14, 2010 | 554 | 500 | Not_Computer
Z-0661-2010 | 54193 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Thomas Medical Products Safesheath CSG (Coronary Sinus Guide) Hemostatic Tear away Introducer System with Infusion Side Port, for the introduction of various types of pacing or defibrillator leads and catheters. The product is shipped 5 per carton. Product is distributed by Pressure Products. Catalog numbers: CSG/MSP-00-09, CSG/90-09, CSG-KR-90-09, CSG/WORLEY-1-09, CSG-WORLEY-Bcor-1-09, CSG/WORLEY/L/Bcor-1-09, CSG/WORLEY-2-09, CSG/WORLEY/Bcor-2-09, CSGMSP0009M, CSGWORLEY109M, CSGWORL19M, CSGWORBC19M, CSGWORLBC19M, CSGWORBC29M, CSG/Worley/BCor-1-07, CSG/Worley/BCor-2-07, CSG/WORLEY/BCOR-M-09, and CSG/Worley/R-1-09. |  |  | December 17, 2009 | 2009 | Thomas Medical Products Inc | Radiopaque sheath tip may fracture | Component design/selection | GE Healthcare Thomas Medical Products firm issued an Urgent Medical Device Correction email on 12/23/09 to the distributor to discontinue distribution and return any affected products. The distributor was instructed to conduct a sub recall regarding the issue. Press issued 1/29/2010. |  | 60982 | 60,982 units | Nationwide. | Terminated | November 01, 2010 | 319 | 60982 | Not_Computer
Z-0668-2010 | 54178 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Discovery CT 750 HD model numbers: 5232083, 5232083-2, 5232083-3, and 5232083-4. |  |  | October 01, 2009 | 2009 | Ge Healthcare | Failure to properly document the CTDI in the Technical Reference or User Manual. | Radiation Control for Health and Safety Act | GE is sending an Important Electronic Product Radiation Warning letter to customers. The letter will describe the defect, identify the affected product details, provide instructions, state the product correction, and the necessary contact information. |  | 111 | 111 units total | Nationwide Distribution -- Including states of CT, SD, VA, WA, MA, PA, GA, IL, CA, MD, NY, NC, SC, AZ, TX, MT, FL, OH, OK, MI, CO, ME, WV, TN, KS, and WI. | Terminated | December 10, 2011 | 800 | 111 | Not_Computer
Z-0782-2010 | 53947 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | TempPadÂż Cold Therapy Pad. |  |  | November 14, 2009 | 2009 | Cincinnati Sub-Zero Products Inc | Product is labeled as "Sterile". It was observed that incomplete seals were created in the packaging materials during production. When the seal is compromised, the sterility of the product inside the package cannot be assured. | Other | On November 14, 2009, Cincinnati Sub-zero Products, Inc. (CSZ) sent a recall notification letter and response letter to customers. A second notice dated December 10, 2009 was also issued to customers. Customers were instructed in the notifications to assume the recalled product is not sterile and do not use. CSZ requested all affected product be returned to their firm for replacement. Customers are to complete and return the response letter acknowledging receipt of the recall letter. Direct questions to CSZ by calling Customer Service at 1-800- 989-7373. Customers may also call 1-800-989-7373, extension 3234 or 1-51 3-71 9-3234 for recall instructions. Representatives will be available to take your call from 8:30 a.m. until 4:30 p.m. eastern standard time, Monday through Friday. |  | 6006344 | 8,872 cartons of ten + 677 each (89,397 total) | Worldwide Distribution -- United States (AL, AR, AZ, CA, CO, CT, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MT, NC, ND, NE, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI &amp; WV), Canada, France, Germany, Ireland, Israel, New Zealand, Singapore, and The Netherlands. | Terminated | January 25, 2012 | 802 | 6007439 | Not_Computer
Z-0608-2010 | 54226 | No | REG | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Model ST-400 Solafeet Foot portable tanners; Solafeet, Inc. |  |  | September 08, 2009 | 2009 | Solafeet, Inc. | Product failed to have a label that contained a recommended exposure schedule; the operator's manual provided inadequate instructions for use; and label implication may result in overexposure. | Radiation Control for Health and Safety Act | Solafeet sent letters to customers informing them that the labeling and instructions included with the product were not in accordance with the FDA requirements. Attached with the letter was a replacement label and a replacement set of instructions. Customers are to attach the new label on top of the old label and use the new instructions, noting the change in exposure limits. Questions should be directed to 941-359-2608 or info@solafeet.com. |  | 72 | 72 units | Nationwide Distribution | Terminated | August 21, 2012 | 1078 | 72 | Not_Computer
Z-0789-2010 | 54274 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Xia Precision System; Jam Shidi 10 Gauge 9 inch; Sterile; Manufactured for: Stryker Spine SAS, Cestas, France; Distributed in the US by Howmedica Osteonics Corp, 325 Corporate Drive, Mahwah, NJ. |  |  | January 18, 2010 | 2010 | Stryker Spine | Stryker Spine received reports of packaging failures during shipping for the Jam Shidi instruments, which may compromise the sterility of the product. | Package design/selection | Stryker Spine issued an "Urgent Product Recall" notification dated January 15, 2010 via Federal Express to all consignees. The letter informed users of the issue and potential hazards of the affected product.For further information, contact your local Stryker Spine Sales Representative or call 201-760-8206. |  | 2727 | 2727 total - all sizes. | Nationwide distribution. | Terminated | September 06, 2012 | 962 | 2727 | Not_Computer
Z-0794-2010 | 50099 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Navigation Reamer T-handle; Non Sterile; Howmedica Osteonics, 325 Corporate Drive, Mahwah, NJ 07430. |  |  | August 08, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics learned that the T-handles have the potential to not engage and not attach to the reamer properly. | Device Design | Stryker Orthopaedics issued notification dated August 8, 2005 via Federal Express. Consignees were informed of the affected product and necessary steps to take for product replacement.For further information, contact Stryker at (201) 831-5825. |  | 22 | 22 | Worldwide Distribution -- United States, Germany, The Netherlands, Spain, Italy, Switzerland, Poland, Romania, South Africa, England, Columbia, Venezuela, Japan and Australia. | Terminated | February 18, 2010 | 1655 | 295 | Not_Computer
Z-0795-2010 | 50101 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Scorpio PS Femoral Waffle (No post) and LFIT; Use only with Scorpio PS Tibial Bearing inserts and Scorpio Patellas, Stryker Orthopaedics. |  |  | July 30, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that a box packaged and labeled as a Scorpio PS Femoral Waffle (No Post) with LFIT, may actually contain a Scorpio Femoral Waffle (With Post) and LFIT. | Labeling mix-ups | Stryker Orthopaedics issued an "Important Product Removal" letter via Federal Express dated July 30, 2007 to all direct accounts. Users were asked to identify and return all affected product to the firm.For further information, contact Stryker Orthopaedics at (201) 831-5825. |  | 7 | 7 | Worldwide Distribution -- United States and France. | Terminated | February 19, 2010 | 935 | 7 | Not_Computer
Z-0796-2010 | 54198 | No | HTJ | Depth gauge for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Depth Gauge; Catalog Number: 703707; |  |  | December 22, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | The Depth Measuring Gauges have 5 scale marking instead of the required 4 scale markings. | Process control | Urgent Product Recall letters, dated December 22, 2009, were sent to all Stryker Branches/Agencies. Letters were also sent to hospitals and surgeons that use the VariaAx Elbow System were also notified. All letters were sent by Federal Express. The letters identified the affected product, stated the issue, and the hazards involved. It also asked customers to examine their inventory and hospital locations to identify affected products. All affected products are to be retrieved and returned to their branch or agency warehouse for reconciliation. All products should be reconciled on the attached Product Accountability Form and a copy should be faxed to 201-831-6069 within 5 days of receipt of the notice. Follow the instructions for shipping the affected product back to the firm. Questions should be directed to 201-972-2100. |  | 49 | 49 | Nationwide Distribution | Terminated | March 16, 2010 | 84 | 49 | Not_Computer
Z-0780-2010 | 53989 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MRI Eclipse 1.5T System Model #781230. |  |  | December 07, 2009 | 2009 | Philips Medical Systems (Cleveland) Inc | Polymeric RF connector block on the tabletop overheated and eventually ignited during scanning due to combination of arcing of electrical components associated with the RF trap assembly located below the connector and flame rating of the connector block material. | Other | An "Urgent-Field Safety Notice" dated December 4, 2009 was issued via certified mail to customers. The notification letter described the affected product, problem, hazard involved and action to be taken by customer. The customer will ensure to maintain frequent contact with patients during scanning and that the patient comfort kit is installed for all patient studies. The Customer will be visited by a Philips Service Representative to inspect their MRI System and install the newly remanufactured RF Connector Block. Additionally, step-by-step trap/cable installation instructions including pictures are being drafted to ensure correct installation of cabling. If you need any further information or support concerning this issue, please contact your local Philips representative or 1-800-722-9377. |  | 290 | 290 units | Worldwide distribution- United states: (AL, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, LA, MA, MD, ME, MN, MO, NE, NH, NJ, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA &amp; WI), ARGENTINA, BRAZIL, CANADA, CHINA, CYPRUS, FRANCE, GERMANY, GREECE, ITALY, JAPAN, KOREA, NEW ZEALAND, PAKISTAN, POLAND, RUSSIA, SWITZERLAND, TAIWAN &amp; UNITED KINGDOM. | Terminated | September 19, 2012 | 1017 | 290 | Not_Computer
Z-0798-2010 | 54233 | No | ETN | Surgical nerve stimulator/locator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | 2-Channel Laryngeal Surface Electrode (Model LSE500DCS and LSE500DCL) |  |  | January 11, 2010 | 2010 | Neurovision Medical Products Inc | This recall has been initiated due to a conductive glue used to attach the lead wire to electrode, which has been found in a small number of electrodes, to cause a short circuit in one of the channels. This is due to migration of the adhesive from one pad to the other creating an electrical short. Use of one of these defective electrodes could lead to a failure to monitor. Our investigation has c | Other | URGENT RECALL: Medical Device Letter dated 1/7/10.The letter identified the affected product and provided a list of necessary actions for customers to take.The customers should immediately examine their inventory and quarantine product subject to recall. In addition, if they have further distributed this product, please identify their customers and notify them at once of this product recall. Their notification to their customers may be enhanced by including a copy of this recall notification letter.This recall should be carried out to the user level. Your assistance is appreciated and necessary to prevent any failure to monitor.Please complete and return the enclosed response form as soon as possible. If you have any questions, call Ryan Rea or Kevin Rea, at 866-815-6999. |  | 4143 | 4143 | Nationwide: CA CO FL GA IL LA MI ND NY OH PA TX UT WA WI) | Terminated | April 21, 2011 | 465 | 4143 | Hardware | Device Operation | Instructions | Safety Notice/Insructions
Z-0811-2010 | 54005 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Biphasic LIFEPAK 12 defibrillator / monitor. |  |  | March 31, 2009 | 2009 | Physio Control, Inc. | Affected devices contain a Biphasic PCBA that may have a solder defect on the H Bridge component that could prevent the device from providing a defibrillation shock. There could be a delay in therapy or prevention of defibrillation therapy if the defibrillator were needed in a cardiac arrest situation. Delay or prevention of defibrillation therapy may contribute to the death of a viable patient. | Nonconforming Material/Component | Consignees were sent a notification letter dated March 2009 that explained firm representatives would follow up with each customer within 30 days to arrange an update to affected devices.Affected Biphasic PCBAs will be replaced. Customers advised to keep the devices in use and to continue to perform daily User Test. |  | 1718 | 1718 | Distributed domestically and internationally to hospitals, clinics, etc.Foreign distribution to the following countries:Australia, Brazil, Canada, Colombia, Costa Rica, Germany, Hong Kong, Mexico, Netherlands, and Puerto Rico. | Terminated | August 08, 2013 | 1591 | 1718 | Hardware | Treatment/Delivery/Therapy | Repair | Repair
Z-0808-2010 | 54195 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LifeShield Latex-Free Primary SYMBIQ Set, Piggyback with Backcheck Valve, 3 Prepierced Y-Sites, 106 inch, Non-DEHP, 15 drops/mL; an Rx sterile single use fluid path way for administration of I.V. fluids using the Symbiq infusion pump; 48 sets per case; Hospira, Inc., Lake Forest, IL 60045 USA, Made in Costa Rica; list number 16096-28. |  |  | January 08, 2010 | 2010 | Hospira Inc. | There is a potential for fluid to leak from the backcheck valve which could result in inaccurate delivery and occupational exposure. | Process control | Hospira Medical issued an "Urgent Device Recall" notification dated January 8, 2010 informing consignees of the affected product. The accounts were instructed to examine their inventory and quarantine affected product immediately and arrange for its return. Users were also requested to complete the enclosed reply form and return it to the firm. For further information, contact Hospira Medical Communications at 1-800-615-0187. |  | 47904 | 47,904 sets | Minnesota and Utah | Terminated | July 20, 2011 | 558 | 49440 | Not_Computer
Z-0810-2010 | 54074 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | LOCATOR Restorative Abutment, ILOA002, LOT 877423, Certain LOCATOR ABUTMENT, 4.1 mm(D) X 2mm(H), STERILE Made in USA, Sterile. |  |  | November 23, 2009 | 2009 | Biomet 3i, LLC | Package labeled as sterile may have an incomplete or partial seal thereby possibly compromising sterility. | Process control | Each US customer will be contacted by telephone using a prepared phone script. Following telephone contact, Biomet 3i will perform a follow-up of each call with a faxed letter. The Urgent: Medical Device Recall letter, dated 11/23/2009, identified the affected product and asked customers to return unused product for replacement. The letter also explained the clinical implications and asked customers to respond with the attached response fax and then return the abutment(s) to BIOMET 3i. Questions or concerns can be directed to Kelly Taylor at 561-776-6906. |  | 97 | 97 | Worldwide Distribution -- USA, Middle East, Italy, Europe, and Greece. | Terminated | December 15, 2010 | 387 | 97 | Not_Computer
Z-0806-2010 | 54485 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Endotek Alimaxx-B Uncovered Biliary Stent for Endoscopic Delivery, REF 70129-210, 10 x 80, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095. |  |  | January 28, 2010 | 2010 | Merit Medical Systems, Inc. | Degradation of the outer portion of the delivery catheter may prevent proper deployment of the biliary stent. | Component change control | Merit Medical Systems, Inc initiated customer notification by phone and letter beginning January 28, 2010. Consignees were instructed to immediately contact their affected customers, advise them of the recall activity and to identify and quarantine any unused inventory. A Recall Notification Form is to be completed and signed by the sales rep and by the site representative. Affected product is to returned to the firm. For further information, contact Merit Medical Systems, Inc. at 1-801-208-4228. |  | 51 | 51 units | Nationwide Distribution -- AL, CA, CO, FL, IL, KY, NY, MA, MD, ME, MN, MO, PA, TX, UT, VA, WI, VA, TX and WI. | Terminated | April 27, 2010 | 89 | 165 | Not_Computer
Z-0813-2010 | 52754 | No | JBP | Activated whole blood clotting time tests. | Hematology | Hematology | 510(k) | Abbott i-STAT ACT Celite Cartridges. |  |  | August 18, 2009 | 2009 | Abbott Point of Care Inc. | Labeling discrepancy: The manual instruction state that the line should be flushed with 5 mL of saline and the first 5 mL of blood or six dead space volumes should be discarded. The Cartridge Test Information sheets that it should be 2 mL of blood - which is incorrect. | Labeling False and Misleading | Abbott Point of Care issued an "Important Product Information" letter beginning August 18, 2009. Consignees were instructed on how to reconcile all affected product.For further information, contact Abbott Point of Care Technical Support at 1-800-366-8020, Option 1. |  | 696675 | 696,675 Cartridges | Worldwide distribution. | Terminated | May 18, 2010 | 273 | 1807450 | Not_Computer
Z-0819-2010 | 53025 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CellSearch Circulating Tumor Cell Kit (Epithelial) IVD |  |  | August 12, 2009 | 2009 | Veridex, LLC | Veridex has received 4 complaints regarding reagent bottles having been incorrectly placed or duplicate bottles places within the reagent tray of CellSearch Circulating Tumor Cell Kit. Delayed detection of disease relapse or progression may occur due to false low CTC results. | Other | Urgent Product Correction Notification letters were sent to all customers on August 12, 2009 by Federal Express. Customers were advised to inspect the kits to assure the proper reagents are in the proper position. If not, do not use the kits and contact the company at 877-837-4339 (chose option 1). |  | 2021 | 1641 in US, 380 internationally. | Product was distributed worldwide, US (one government account: National Cancer Institute, Frederick, MD). Also 8 international accounts in Belgium, France, Germany, Italy, Spain, United Kingdom, Netherlands, Norway and Israel. | Terminated | July 18, 2012 | 1071 | 2021 | Not_Computer
Z-0817-2010 | 53553 | No | JWX | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Oxoid dry Spot Staphytect Plus, Catalog #DR0100M, 120 tests/kit. Oxoid Ltd., Wade road, Basingstoke, Hants, UK. |  |  | July 20, 2009 | 2009 | Remel, Inc | The products may give false positive results with some non-Staphylococcus aureus strains | Process control | Thermo Fischer Scientific, Remel Inc. issued letters dated July 20, 2009 via regular mail explaining the reason for recall, risks involved, and requested they discard all remaining product. A Product Inventory Checklist was enclosed to report the amount of product their customer discarded. A second letter dated October 19, 2009 was issued via certified mail.For further information, contact Thermo Fischer Scientific, Remel Inc. at 1-800-255-6730, Extension 3304185. |  | 27 | 27 kits | Nationwide Distribution -- MN, CT, FL, GA, CA, NY, MI, MO, PA, and IL. | Terminated | April 25, 2012 | 1010 | 77 | Not_Computer
Z-0812-2010 | 53335 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT 1 Analyzer, Abbott Point of Care Inc., Abbott Park, IL 60064 USA, Product of USA. |  |  | September 09, 2009 | 2009 | Abbott Point of Care Inc. | The storage temperature on the label on the side of the i-Stat Portable Clinical Analyzer shipping box does not match product specification. | Labeling mix-ups | An "Urgent Recall Notice" was issued to customers during September 2009. The letter described the affected product, issue and recommended actions. Customers were asked to verify storage conditions on the box are correct and if not, discard the outside box. If the customer requires a replacement box, the customer must contact Abbott Point of Care Technical Support. Direct questions and requests for additional information to the Abbott Point of Care Technical Support by calling 1-800-366-8020, Option 1 or contact your Abbott Point of Care representative. |  | 3918 | 3,918 analyzers | Worldwide Distribution (including the United States). | Terminated | December 08, 2014 | 1916 | 3918 | Not_Computer
Z-0827-2010 | 54184 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | RambaChrom KPC M/N 11500. RambaCHROM KPC is for detection of gram negative bacteria with a reduced susceptibility to most of the carbapenem agents. |  |  | December 17, 2009 | 2009 | Gibson Laboratories Inc | Discrepancies between the product package label and technical insert are conflicting and may be misleading to user. The technical insert included the statement for In Vitro Diagnostic Use. The product insert should have indicated for Research Use Only instead of for In Vitro Diagnostic Use. | Employee error | On 12/10/2009 the firm sent letters to their customers notifying them of the problem and instructing them to replace insert with a new one. They were also instructed to complete and return a "Recall Response Form." |  | 3440 | 3,440 units | Distribution of the product was limited to AL, FL, IL, KY, MA, MI, MO, NY, OH, PA, &amp; VA and one customer in France. | Terminated | September 21, 2012 | 1009 | 18770 | Not_Computer
Z-0823-2010 | 54271 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Stryker XIA 3 Polyaxial Screw, 8.5mm x 65 mm. Catalog Number: 482318565. |  |  | December 30, 2009 | 2009 | Stryker Spine | The Xia 3 8.5mm Polyaxial screws, lot A91109 were anodized turquoise instead of Fuchsia. The turquoise color is reserved for the 9.5 mm Xia 3 screws. | Process control | Stryker Spine branches that received the recalled product were notified of the recall by an Urgent Product Recall letter, dated January 15, 2010, sent via Federal Express. The letter identified the affected product and the issue involved. It also discussed the potential hazards, risk mitigation, and the return process. Customers are to examine their inventory and hospital locations to identify the product. Customers are also to reconcile the product by utilizing the Customer Response Form and distribution sheet. In addition, they are to retrieve and return the affected product. Questions should be directed to Tiffani Rogers at 201-760-8206. |  | 21 | 21 | Nationwide Distribution -- IL, MO, DA, MA and CO. | Terminated | September 06, 2012 | 981 | 21 | Not_Computer
Z-0826-2010 | 54197 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Encore Medical Foundation Knee Revision Baseplate, Part Number 321-01-106, Size 6, Left |  |  | December 11, 2009 | 2009 | Encore Medical, Lp | Partial lot of product was not manufactured to specification. Central locking screw was missing threads. There is a potential for dissociation because the screw is not able to be used as a secondary fixation method. | Process control | Firm notified consignees by letter on 12/11/2009. The letter informs physicians that patients with implanted devices should counseled about the issue and should report any changes to their treating physician. |  | 17 | 17 | TX, FL, GA, CA, AZ, Japan | Terminated | September 24, 2010 | 287 | 17 | Not_Computer
Z-0825-2010 | 54238 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes In Situ Bender/Cutter Kit-Wide-Sterile. Catalog Number 530.521S. |  |  | January 06, 2010 | 2010 | Synthes USA (HQ), Inc. | There is the potential for the bender/cutter attachment to continue heating after release of the powerbutton. There is potential for patient and/or staff injury if continuous heating is unrecognized. | Other | An "Urgent: Medical Device Recall" letter dated January 8, 2010 was issued to customers via USPS Certified mail. The letter described the issue and affected product. Customers were notified to cease use of the recalled product, examine inventory and follow instructions for return to the firm as indicated in the recall letter. For questions, please call 1-800-620-7025 x 5452 or 1-610-719-5452 or contact your Synthes CMF Sales Consultant. |  | 187 | 187 units | Nationwide Distribution. | Terminated | October 06, 2010 | 273 | 187 | Not_Computer
Z-0822-2010 | 53702 | No | KIR | Cultured animal and human cells. | Hematology | Pathology | 510(K) Exempt | Rabbit Kidney Culture, brand name: 'FreshCells'. The product is labeled as 'Rabbit Kidney' also known as 'Rbk' , 2-square cm area of a cell monolayer covered with 2-mL culture medium. The product is packed in glass culture tubes 16x125mm, with screw caps. The product is distributed in EPS shipping racks which can hold up to 50 tubes each. |  |  | January 21, 2009 | 2009 | Diagnostic Hybrids Inc | During QC post release monitoring, the firm determined that the monolayer culture of rabbit kidney cells previously distributed by the firm, appeared to contain another type of cells (monkey kidney cells). | Process control | Diagnostic Hybrids issued a "Product Notification" letter dated January 23, 2009 in addition to verbal communication of the affected product beginning January 21, 2009. The firm requested that all affected product be discarded /destroyed and not use for diagnostic testing. Consignees were asked to complete and return a Fax-Back form.For further information, contact Diagnostic Hybrids Technical Services at 1-866-344-3477 or 1-740-589-3300. |  | 1127 | 1127 units | Worldwide Distribution -- United States (CA, FL, GA, MA, MO, NY, NJ, TN, TX and VA), Ontario Canada and Manitoba Canada. | Terminated | September 21, 2010 | 608 | 1127 | Not_Computer
Z-0846-2010 | 54249 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker T2 Ankle Arthrodesis Nail, right; 012 x 300 mm; Catalog number: 18191230S; |  |  | December 29, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that there is the potential for damage to the sterile packaging of the product during transport. | Packaging | Urgent Product Recall notification letters were sent to branches on 12/29, 2009, with a corrected version that included a list of catalog numbers and code numbers on January 11, 2009. Letters were also sent by Fed Ex on January 11, 2009 to Hospital Risk Management, Chief of Orthopaedics and surgeons who may have used the device. The letters identified the affected product, described the issue, and the hazards involved. The letter also asked customers to examine their inventory and hospital locations to identify the affected product. They are to retrieve all affected product and return it to their branch or agency warehouse for reconciliation. Customers are to reconcile all products utilizing the Product Recall Accountability Form. Questions should be directed to 201-972-2100. |  | 702 | 702 total, all sizes. | Nationwide Distribution. | Terminated | August 06, 2012 | 951 | 702 | Not_Computer
Z-0853-2010 | 53785 | No | DXH | Telephone electrocardiograph transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Cybernet Medical MedStar Telemedicine System, Manufactured by Cybernet Medical, Ann Arbor, MI; Model 533-380. |  |  | January 23, 2008 | 2008 | Cybernet Systems Corp. | A software problem renders these devices incapable of capturing lung function, oxygen saturation and pulse data collected from multiple tests in accordance with the instructions for use. | Component design/selection | The firm sent out a recall letter dated 10/27/09. Consignees were asked to return the devices in prepaid shipping containers. The recalling firm would do a software upgrade and return the device to the consignee. The firm expects it to take 3-6 months to upgrade all devices.Contact Cybernet Medically immediately, at 1-800-CYBERNET, with the serial number and code versions of the devices in question and Cybernet will start the RMA process for returning the devices for the updates, free of cost. |  | 24 | 24 | Nationwide distribution: California, Florida, Georgia, Illinois, Louisiana, Michigan, Mississippi, New York, Oklahoma and West Virginia. | Terminated | October 22, 2010 | 1003 | 88 | Computer
Z-0850-2010 | 48978 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | COULTER LH SlideMaker, LH Barcode Specification, LH SM Intended Use, LH Slide and Sample Flow, Part Number 6605633, P/N 4277248, 4277299, and 277299. |  |  | May 29, 2008 | 2008 | Beckman Coulter Inc | Beckman Coulter has become aware of a potential issue in which the Sample ID information embedded in the barcode affixed to a slide made by the Coulter GEN*S and LH SlideMaker does not match the text. | Pending | An Urgent: Product Corrective Action letter, dated June 9, 2008, was sent to customers. The letter identified the affected products, stated the background, discussed the issue, and stated the actions to be taken. Customers are to: 1. Only use upper case alpha characters in the Sample Identifier2. If you must use lower case alpha characters in the Sample Identifier, configure the SlideMaker Slide labels without a bar-code. Questions regarding the Product Corrective Action should be directed to Beckman Coulter Customer Service at 1-800-526-7694 or your local Beckman Coulter Rep. |  | 843 | 843 | Worldwide Distribution -- USA and Canada. | Terminated | November 29, 2011 | 1279 | 847 | Not_Computer
Z-0858-2010 | 53026 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Flex Reagent Cartridge Cardiac Tropinin-I, Catalog Number: RF421C and RF521. |  |  | March 02, 2007 | 2007 | Siemens Healthcare Diagnostics, Inc. | Tests exhibit falsely elevated results. | Material/Component Contamination | The recalling firm issued a Field Correction letter dated 3/2/07 informing user of the problem and to contact the firm at 800-241-0420 for replacements. Customers are to follow the precautions provided in the letter. Questions should be directed to Dade Behring Technical Assistance Center at 800-441-9250. The notification should be forwarded to anyone to whom the product may have been further distributed to. |  | 9767 | 9767 cartons | Worldwide Distribution -- USA, Canada, Germany, Egypt, New Zealand, Saudi Arabia, Singapore, and BS. | Terminated | February 25, 2010 | 1091 | 9767 | Not_Computer
Z-0856-2010 | 53679 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxI 800 and 600 Access Immunoassay Systems, UniCel DxC 880i, 860i, 660i, and 680i SYNCHRON Access Clinical Systems with Dxl System Software versions 4.2 or 4.2.1. |  |  | October 28, 2009 | 2009 | Beckman Coulter Inc. | This recall was initiated because of software issue was identified in System Software 4.2 and 4.2.1, which may result in the incorrect handling of a wash arm dispense plate motion failure. Instead of generating a red warning event and stopping sample processing, the system will post a yellow caution event and continue to operate with the dispense plate in an elevated position until the instrument | Software design | Beckman Coulter issued an "Urgent: Product Corrective Action" letter dated October 28, 2009. Customers were notified of the affected product and how the issue will be resolved.For further information, contact Beckman Coulter Inc. Technical Support at 1-800-854-3633. |  | 377 | 377 Instruments | Worldwide Distribution -- United States and Canada. | Terminated | May 07, 2012 | 922 | 377 | Software | Alarm/Message | Instructions | Safety Notice/Insructions
Z-0859-2010 | 53810 | No | LBR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid HDL (PEG) Cholesterol Reagent Set. Pointe Scientific HDL Cholesterol Reagent Set, Pointe Scientific, Inc, Canton, MI; Catalog # H7511-60 and Catalog # H7511-SAM. HDL Cholesterol Reagent Set, Pointe Scientific, Inc, Canton, MI; . (Clinical chemistry). |  |  | November 10, 2009 | 2009 | Pointe Scientific, Inc. | The firm initiated the recall due to the presence of visible contamination of the product. | Process control | Consignees were notified by an Urgent: Medical Device Recall letter (sent by email, fax or through the mails) dated 11/10/09. The letter identified the affected product and instructed customers to immediately examine their inventory. If the product was further distributed their customers should be identified and notified of the recall. Customers are instructed to dispose of any product in stock. Questions should be directed to the Technical Support Department at 800-757-5313 or 734-487-8300. |  | 379 | 379 sets | Worldwide Distribution -- USA, Bolivia, Ghana, Jamaica, Kenya, Korea, Malaysia, Mexico, Nicaragua, Nigeria, Philippines, Poland, Sri Lanka and Trinidad. | Terminated | February 11, 2011 | 458 | 625 | Not_Computer
Z-0857-2010 | 54364 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Cardinal Health EnVe Ventilator Cat # 19250-001, Manufactured by Cardinal Health, Minneapolis, MN 55447-1341. |  |  | January 08, 2010 | 2010 | CareFusion 203, Inc. | CareFusion is recalling the EnVe Ventilator because the EnVe Ventilator does not fully meet applicable standards for immunity to electro-Magnetic Interference (EMI) when tested in the laboratory environment. CareFusion has identified certain device malfunctions that may affect device performance and patient safety in the intended use environment. | Component design/selection | CareFusion issued an "Urgent Medical Device Recall notification" letter dated January 07, 2010. The letter was addressed to the Director of Respiratory Care describing the problem and the product. Consignees were told that a Service Representative will contact them to arrange for the prompt return of the affected product for correction.For further information, contact CareFusion Technical Support at 1-800-231-2446, Option 1. |  | 9 | 9 | MN | Terminated | December 17, 2011 | 708 | 9 | Not_Computer
Z-0870-2010 | 53800 | No | JXB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | ERIC (Electronic RapID Compendium) Version 1.0.77, Catalog Number R8323600, used in conjunction with: remel RapID Yeast Plus System is a qualitative micromethod employing conventional and chromogenic substrates for the identification of medically important yeast, yeast-like, and related organisms isolated from human clinical specimens. A complete listing of the organisms addressed by the RapID Yeast Plus System is given in the RapID Yeast Plus Differential Chart. |  |  | October 28, 2009 | 2009 | Remel, Inc | There is an error in Version 1.0.77 of the electronic code compendium (ERIC) software. The error is limited to the database corresponding to the R8311007 RapID Yeast Plus System. The database erroneously includes Candida dubliniensis as well as displays an incorrect temperature range as an accessory test for codes with an overlap between Candida albicans and Candida dubliniensis. When using the | Software design | The firm issued an "IMPORTANT MEDICAL DEVICE PRODUCT RECALL NOTICE" dated October 29, 2009. After completion of the firm's investigation and risk assessment, Remel decided to notify customers of this situation and quarantine any remaining units on hand. The notification provided the consignees with instructions for obtaining and downloading the software patch. A response form was included with the recall letter and the consignee was asked to complete the form and return it to Remel.The consignees should contact the Technical Services Department at 800-255-6730 (USA) or 913-888-0939 (International), if they have inquiries concerning this advisory. |  | 143 | 143 units of software | Worldwide Distribution: USA, Europe and Canada | Terminated | March 03, 2010 | 126 | 143 | Software | Output/Calculation | Software update | Software Update
Z-0864-2010 | 53323 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista System CHEM 1 CAL, Catalog Number KC110, Lot Number 9DM001 |  |  | August 31, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics, Inc. has confirmed high accuracy shifts with Dimension Vista URCA Flex reagent cartridge (various lots) when calibrating with some vials from CHEM 1 calibrator lot 9DM001. | Other | An Urgent Field Safety Notice dated August 31, 2009, was issued to all affected customers. Customers were instructed to discontinue the use of Dimension Vista System CHEM 1 Cal lot 9DM001 and discard any remaining product. If customers have any question they can contact Technical Solutions at 1-800-241-0420. |  | 1474 | 1474 Cartons | Nationwide and Foreign: Australia, Canada, Denmark, and Malaysia. | Terminated | December 13, 2010 | 469 | 1474 | Not_Computer
Z-0861-2010 | 53341 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Huber Plus Non-Coring Needle Safety Infusion Set, 19g, 1 inch, Needless Y-site, REF 011901NY, Sterile EO, Assembled In Mexico, Bard Access Systems, Inc., SLC, UT 84116. |  |  | January 14, 2010 | 2010 | Bard Access Systems | After activation of the safety mechanism, the tip of the needle protrudes from the safety sheath creating the potential for needlestick injury. | Component change control | Bard Access Systems, Inc. issued and "Urgent: Huber Plus Non-Coring Needle Safety Infusion Set Recall" notification beginning January 2010. Users were instructed to quarantine all affected product and to notify any of their customers that had received the product. Quarantined product is to be returned to the firm. For further information, contact Bard Access Systems, Inc. at 1-800-290-1689. |  | 4975 | 4,975 units | Worldwide Distribution -- United States (VA, KY, MA, NJ and TN) and Belgium. | Terminated | May 14, 2010 | 120 | 5750 | Not_Computer
Z-0865-2010 | 53483 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | i-STAT HL7 Interface |  |  | September 29, 2009 | 2009 | Abbott Point of Care Inc. | Abbott Point of Care has determined that it is possible under certain circumstances that some test records could have their status update messages mis-matched with another record, resulting in the incorrect patient name, order number of comment being included in the CDS test record only. | Software design | Urgent Recall Notification letters were sent by Fed Ex to all US customers on September 29, 2009 and worldwide on October 1, 2009. Questions should be directed to Abbott Point of Care Technical Support at 1-800-366-8010, Option 1. |  | 36 | 36 | Nationwide and Foreign: Germany, the United Kingdom, Switzerland and the Netherlands. | Terminated | July 01, 2010 | 275 | 36 | Other | Output/Calculation | Notification | Safety Notice/Insructions
Z-0868-2010 | 54237 | No | KRI | Cardiopulmonary bypass accessory equipment. | Cardiovascular | Cardiovascular | 510(K) Exempt | IDEAL MAXPAC TOTAL SYSTEM (USA) , Sterile and Nonpyrogenic, ETO or Gamma Sterilized, Medtronic Perfusion Systems, Model TS9453R29, TS9453R30, TS1470R1, TS1442R1, TS9800R32, TS8299R4. |  |  | December 14, 2009 | 2009 | Medtronic Inc | Medtronic has identified a small number of Autotransfusion One Source Kits and MAXPAC/Total System lots with incorrect Use Before Dates (UBD) on the outer kit package label. The UBD on the outer label of the kit/pack does not accurately reflect the UBD of the components within the kit/Pack. For affected product, the outer label lists a date which may be later than the USB of the components inclu | Incorrect or no expiration date | Consignees were sent a Medtronic "Urgent Medical Device Recall Notice" dated December 14, 2009. The letter was addressed to "Valued Customer." The letter described the product and the problem and requested them to quarantine the product and to complete and sign the attached certificate, "Medtronic Recall Certificate (FCA 1002)." For additional information, contact the firm at 763-514-9605. |  | 24 | 24 | CA, CO, JAPAN, CHINA. | Terminated | April 20, 2011 | 492 | 36 | Not_Computer
Z-0867-2010 | 53649 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Variant II Turbo Hemoglobin A1c Program, model number 270 2417, manufactured by Bio-Rad Laboratories, Inc, Hercules, CA. |  |  | October 12, 2009 | 2009 | Bio-Rad Laboratories Inc | Product update kit CD rom will not upload. As a result, the device cannot be run. | Other | The firm initiated this action on 10/12/2009. Consignees were contacted by phone with a fax follow up. Subsidiaries, as well as consignees that received the product, were provided with a Medical Device Correction notice, dated October 13, 2009, and response form. Each subsidiary will contact customers. The letter asks customers to destroy the affected lot included with their reorder packs and replace them with the lot number that's enclosed. Customers are to complete the Customer Medical Device Correction Response Form and fax it to Bio-Rad CSD Regulatory Affairs Department at the number provided. Questions should be directed to customers' regional Bio-Rad office. |  | 21 | 21 kits | Worldwide Distribution -- USA, Hong Kong, and New Zealand. | Terminated | December 17, 2010 | 431 | 21 | Not_Computer
Z-0866-2010 | 54355 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | EXXCEL Soft Standard Wall EPTFE Vascular Graft; |  |  | January 20, 2010 | 2010 | Maquet Cardiovascular, LLC | Labeling mix up: Shelf Carton label for Exxcel Soft Thin Wall Vascular Grafts, lot # 12633379 were mislabeled as Exxcel Soft Standard Wall Grafts. Bleeding could occur if wrong product is used. | Labeling mix-ups | Recall letters were sent to all US customers on January 20, 2010 by certified mail. Letters for foreign consignees were to be sent as soon as translation was completed. Customer are directed to call the company at 973-709-7515 for more information. |  | 18 | 18 units | Product was distributed to hospitals in TX, TN, CT, AR, AZ as well as to Austria and France. | Terminated | July 06, 2010 | 167 | 18 | Not_Computer
Z-0895-2010 | 54487 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Electri-Cord Manufacturing Corporation (312 E. Main St., Westfield, PA 16950) AC power cords used with either Bair PawsÂż Temperature Management Units: Models 850 and 875 |  |  | December 17, 2009 | 2009 | Arizant Inc | Arizant Healthcare Inc. is initiating a nationwide voluntary recall of 100/110-volt power cords (Arizant part number 502221 Model 90024) attached to its temperature management units and pressure infusion power pack (Bair PawsÂż Temperature Management Units: Models 850 and 875Bair HuggerÂż Temperature Management Units: Models 500/OR, 505 (Human &amp; Vet.), 750 (Human &amp; Vet) and 775RangerÂż Blood/Flui | Nonconforming Material/Component | An "Urgent Medical Device Power Cord Recall" letter dated December 17, 2009 was sent to customers. The letter was addressed to Recall Coordinator/Biomed Dept. The letter described the affected product, problem, hazard involved, and action to be taken by customer. The customers are being ask to inspect their temperature management units (as described in FDA's recommendation dated October 19, 2009) and contact the firm with the results of their inspection and provide a contact name via website www.arizant.com/powercords or phone (877) 947-1487. Replacement cords and instructions will be provided to customers based on priority listed in the letter and availability of replacement cords. If the customer does not have the means to make the changes, alternate means will be taken (i.e., an Arizant employee or outside service organizations may be brought in to do the replacements).If you have any questions, please contact the Aziant at (877) 47-1487. |  | 84537 | Total distributed for all devices: 84,537 US, 7,380 OUS. | Worldwide Distribution: USA and ARGENTINA, THE BAHAMAS, CANADA, COLOMBIA, COSTA RICA, CUBA, DOMINICAN REPUBLIC, ECUADOR, GERMANY, GUAM, JAPAN, LEBANON, MEXICO, NICARAGUA, PANAMA, PUERTO RICO, QATAR, SAUDI ARABIA, SOUTH KOREA, SURINAME , TAIWAN, TRINIDAD AND TOBAGO, QATAR. | Terminated | April 19, 2012 | 854 | 84537 | Not_Computer
Z-0871-2010 | 53768 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PTS Panels Cholesterol test strips for self-testing or professional use with CardioChek analyzers, Manufactured by Polymer Technology Systems, Inc., Indianapolis, IN.Catalog # 1711. |  |  | April 02, 2009 | 2009 | Polymer Technology Systems, Inc. | The memory chip was coded with incorrect product information. | Process control | Polymer Technology Systems, Inc. issued an "Urgent - Field Correction" letter dated April 2, 2009. Consignees were advised of the problem and asked to identify how many kits were in inventory so that the firm could replace product.For further information, contact Polymer Technology Systems, Inc. Customer Service at 1-877-870-5610. |  | 486 | 486 vials | Nationwide Distribution | Terminated | October 20, 2010 | 566 | 486 | Other |  | Replace | Remove or Replace
Z-0800-2010 | 54285 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HomeChoice PRO Automated Peritoneal Dialysis Systems; Catalog Numbers; 5C8310, 5C8310R, R5C8320, R5C8320R, T5C8300, and T5C8300R. |  |  | January 08, 2010 | 2010 | Baxter Healthcare Renal Div | Potential overfill of the peritoneal cavity, also referred to as Increased Intraperitoneal Volume (IIPV). IIPV can result in serious injury or death. | Labeling design | Urgent Product Recall letters dated January 8, 2010 were mailed to all Peritoneal Dialysis Clinicians on January 8, 2010 and letters dated January 12, 2010 to all Home Care Coordinators and Home Dialysis Patients on January 12, 2010, informing them of steps to take to help reduce or eliminate overfill, also referred to as Increased Intraperitoneal Volume, associated with HomeChoice/Homechoice Pro cyclers. Baxter Healthcare Corporation explained that they had received complaints of IIPV. The accounts were requested to complete and return the enclosed response form, indicating that they received and understood the notification, and if they still had a HomeChoice System or not. Any questions from the Home Care Coordinators are to be directed to the Baxter Global Technical Services at 1-800-553-6898. |  | 32678 | 32,678 units | Worldwide Distribution -- United States, Austria, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Ecuador, Finland, France, Germany, Greece, Hong Kong, India, Iraq, Ireland, Israel, Italy, South Korea, Kuwait, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Panama, Peru, Philippines, Poland, Portugal, Qatar, Russia, South Africa, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, Thailand, Taiwan, Tunisia, Turkey, United Arab Emirates, United Kingdom and Venezuela. | Terminated | June 21, 2017 | 2721 | 253508 | Not_Computer
Z-0908-2010 | 49449 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | Rubella IgG EIA test kit, containing 1 Rubella IgG Microplate, 1 Negative Control, 1 High Positive Control, 1 Low Positive Control, 1 Substrate, 1 Stop Reagent. Labeled: Distributed by: Bio-Rad Laboratories-Redmond, WA 98052 |  |  | August 20, 2008 | 2008 | Bio-Rad Laboratories | Elevated optical density (OD) values with the kit Negative Control. | Employee error | An "Urgent Product Correction" letter dated August 26, 2008 was sent via overnight carrier to all customers. The letter described the affected product, problem, and action to be taken by customer. The customers are to review their current inventory of the product for any remaining kits of the specified lot number. The customer is to complete the form and fax to the Technical Support department if replacement kits were needed. The remaining kits will be replaced once the form is completed and faxed.If you have any questions, please contact Technical Support at 1-800-224-6723, options 2 then 4. |  | 82 | 82 kits | Nationwide distribution: NE, NY, FL, GA, MA, OR, WI, MN, MI, MS, NC, CA, MD, MO, IA, DC | Terminated | April 08, 2010 | 596 | 82 | Not_Computer
Z-0921-2010 | 48648 | No | DCK | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | IMMAGE Immunochemistry Systems C-Reactive Protein (CRP) Reagent, Part Number 447280. |  |  | January 10, 2008 | 2008 | Beckman Coulter Inc | IMMAGE CRP Reagent lots M705441, M709279, &amp; M709319, may demonstrate a prozone or hook effect on samples with extremely high CRP analyte concentration. | Other | Customers were sent Urgent: Product Corrective Action letters dates, 1/22/2008. The letter identified the issue, discussed the impact, and listed the following actions/resolution.* To eliminate the hook effect when using these lots, run CRP samples beginning with the 1:216 dilution. ~ A limited correlation of patient samples run in the 1:36 dilution compared to the same samples run in the 1:216 dilution verifies acceptable performance (there may be a shift of up to 0.8 mg/dL for samples between 1.0 and 8.0 mg/dL).* Controls run in the 1:216 dilution may recover differently than those run in the default 1:36 dilution. A limited number of control samples were also run in the 1:216 dilution.* The IMMAGE 800 system can be set to default to a non-standard dilution for each chemistry. Change CRP to a Non-Standard Dilution of 1:216 in the Setup section of the software. Detailed instructions for configuring a Non-Standard dilution as default are located in chapter 5 of the IMMAGE 800 Immunochemistry System Operations Manual.* The IMMAGE system does not allow resetting the default dilution of a chemistry. Instructions for running CRP samples beginning in a Non-Standard Dilution are enclosed.* Complete and return the enclosed response form within 10 days so we may maintain our records.Customers are to share this information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. If you have forwarded any of the affected product listed above to another laboratory, please provide a copy of this letter to them. Questions regarding this letter should be directed to 800-854-3633 or customers can contact their local Beckman Coulter Representative |  | 878 | 878 kits | Nationwide Distribution -- Including states of AL, CA, CT, DC, FL, FL ,GA, HI, IA, IL ,IN, KY ,MA, MD, ME, MI, MO, NE, NJ, NY, OH, OR, PA, RI, TX, VA, WI, and WV. | Terminated | July 02, 2012 | 1635 | 878 | Not Computer
Z-0910-2010 | 50150 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon X3 UHMWPE Tibial Inserts and Patellar Components, multiple sizes, Stryker Orthopaedics, Mahwah, NJ. |  |  | February 17, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware of a potential anomaly related to the packaging process of referenced lots of X# polyethylene used on Scorpio, Triathlon and Trident inserts. | Process design | Stryker Orthopaedics issued an "Important Market Withdrawal" notification dated February 17, 2007 via federal express. Consignees were informed of the affected product and asked to retrieve, reconcile and return product via Return Material Authorization to the firm.For further information, contact Stryker Orthopaedics Quality Assurance at 1-201-831-5825. |  | 594 | 594 units total - all products. | Worldwide Distribution -- United States and Canada. | Terminated | March 03, 2010 | 1475 | 594 | Not_Computer
Z-0914-2010 | 50703 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Peristaltic Pump component of: UniCel DxI 800 Access Immunoassay Systems, UniCel DxI 600 Access Immunoassay Systems, UniCel DxC 880i SYNCHRON Access Clinical Systems. |  |  | October 09, 2008 | 2008 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports of a new premature failure of the upper aspirate peristaltic pump in the systems listed above. Premature failure of the pump may negatively affect precision. The failure is limited to those UniCel DXI 800, DXI 600, and DxC 880i systems which have a pump identified with one of the new pump manufacturing numbers (Mfg. Nos.) listed ab | Other | Beckman Coulter contacted consignees via "Urgent Product Corrective Action" letter dated October 16, 2008 to inform them of the recall and to arrange a date to service the affected product as soon as possible. For further information, contact Beckman Coulter: (a)In the U.S.A. or Canada, contact Technical Support at 800-854-3633 or online at www.beckman.com/customersupport. (b) Outside the U.S.A. and Canada, contact your technical support representative. |  | 107 | 107 (US), 7 (OUS) | Worldwide Distribution -- United States and Canada. | Terminated | May 04, 2012 | 1303 | 107 | Not_Computer
Z-0885-2010 | 54228 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Mistique Infusion Catheter, MIC5-135-20-038, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095 |  |  | January 06, 2010 | 2010 | Merit Medical Systems, Inc. | Infusion catheters may become brittle and break during use. | Storage | All affected sales reps were notified by phone and a Mistique Infusion Catheter - Product Recall letter on 01/06/2010 and instructed to contact their accounts, inform them of the recall situation, and complete the Recall Notification form with a site representative. Accounts were to immediately cease use of any affected units and to arrange for return to Merit. Recall actions are to be completed by Jan 22, 2010. Questions should be directed to Greg Turner at 801-316-4998 or Liz Pratt at 801-208-4828. |  | 18 | 18 uinits | Worldwide Distribution -- USA, including states of AL, CA, IN, LA, MO, MS, NC, NY, OH, PA, SC, UT, and VA and countries of India, Japan, Poland, San Marino, and UK. | Terminated | September 13, 2010 | 250 | 1389 | Not_Computer
Z-0907-2010 | 54258 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion One Whole Body X-Ray CT Scanner |  |  | January 07, 2010 | 2010 | Toshiba American Medical Systems Inc | When W-Volume is performed in Prospective CTA or Target CTA studies with ECG gating, unnecessary intervals may be set before starting the volume scan. As a result, scanning cannot be performed at the appropriate timing of constant enhancement and clinically useful images may not be obtained. | Device Design | A corrective measure will be installed to prevent measures to avoid this phenomenon until corrective actions have been taken.Customers are advised: (1) When performing scanning in Prospective CTA or Target CTA studies with ECG gating, do not use W-Volume scanning but use Helical Scanning. (2) Please share this information with all users and reviewing radiologists as well as clinical engineering or Biomedical group at facility (3) Please complete and return the attached form and fax it to the toll-free number at the top or by email to raffairs@tams.com. Contact: Paul Biggins |  | 28 | 28 units | Nationwide: AL AR AZ CA CO FL GA IA MA MD ND NJ NV NY OH PA PR andVA | Terminated | September 12, 2012 | 979 | 28 | Not_Computer
Z-0899-2010 | 53872 | No | JIY | Iron (non-heme) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific Iron/TIBC Reagent Set, for the quantitative determination of serum iron and iron binding capacity, Pointe Scientific, Inc, Canton, MI, Catalog #s: I7504-60, I7504-400, I7504-SAM, 10-I7504-60. |  |  | November 16, 2009 | 2009 | Pointe Scientific, Inc. | Turbidity formation could occur, resulting in invalid test results being reported out. | Device Design | Consignees were notified by recall letter dated 11/16/09 and instructed to discard any remaining product. Please call 1-800-757-5313 if you have questions. |  | 417 | 417 sets | Nationwide, Australia, Austria, Bolivia, Canada, Cyprus, Egypt, Ghana, Greece, Guyana, India, Indonesia, Israel, Jordan, Kenya, Korea, Lebanon, Macedonia, Malaysia, Mexico, Nicaragua, Nigeria, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Trinidad, Turkey, Ukraine and Venezuela. | Terminated | June 06, 2011 | 567 | 1480 | Not_Computer
Z-0915-2010 | 54359 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NextGen Knee Gender Solutions CR-Flex Femoral Component, Precoat , Size D, Left. 00-5750-014-01 Zimmer, Inc. Warsaw, IN. |  |  | January 27, 2010 | 2010 | Zimmer Inc. | The package should contain the left knee component but actually contains the right knee component. Delay in completing surgery may result, and a new or different implant may need to be identified. | Labeling mix-ups | The firm sent an immediate notification by e-mail to all first consignees on 1/12/2010. They sent a formal recall notification on 1/27/2010. Customers should contact their Zimmer Sales representative with additional questions. |  | 14 | 14 units | US and Korea. | Terminated | October 17, 2010 | 263 | 14 | Not_Computer
Z-0913-2010 | 54112 | No | KNK | Uric acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Pointe Scientific Liquid Uric Acid (R1) Reagent Set, Pointe Scientific, Inc., Canton, MI; Catalog # HU482-R1. |  |  | December 01, 2009 | 2009 | Pointe Scientific, Inc. | A process deviation occurred during the production of the R1 component which may affect its performance. | Device Design | Pointe Scientific, Inc. issued an "Urgent: Medical Device Recall" notification dated December 1, 2009. Consignees were instructed to examine their inventory for the affected product and discard the affected product immediately.For further information, contact Pointe Scientific, Inc. Technical Support at 1-800-757-5313 or 1-734-487-8300. |  | 1 | 1 kit | Worldwide Distribution -- United States (California, Florida, Kentucky, Michigan, Ohio, Oklahoma and Virginia), Chile and Malaysia. | Terminated | October 08, 2010 | 311 | 13 | Not_Computer
Z-0922-2010 | 54124 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | DiaSorin Rubella IgG ELISA Kit, Cat. # 7545, 2 x 6 Wells, DiaSorin Inc., 1951 northwestern Avenue, Stillwater, MN 55082-0285, Lot 123127A, Exp. 2010//07/28. |  |  | December 16, 2009 | 2009 | Diasorin Inc. | DiaSorin Rubella IgG kits of lot 123127A were packed with one coated microtiter plate instead of two microtiter plates as indicated in the product labeling. The risk is the possibility of delayed patient test results. | Employee error | Consignees were sent a DiaSorin Customer Notification letter dated December 16, 2009. The Letter was addressed to Customer. The letter described the product and the problem and indicated a replacement product will be provided. Requested consignees to complete the acknowledgement receipt form. For questions or concerns please contact DiaSorin Inc., Technical Services at 1-800-328-1482 or 651-439-9710. |  | 10 | 10 kits | US distribution: CA, CT, MA, MD, IL, OR | Terminated | March 29, 2012 | 834 | 10 | Not_Computer
Z-0916-2010 | 53431 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | Herculite Ultra (Part Number 34333), dental composite resorative material. Manufactured by Kerr Corporation, Orange, CA. |  |  | February 23, 2009 | 2009 | Sybron Dental Specialties | Sybron Dental Specialties is initiating a field correction on the Herculite Ultra, a dental composite restorative material, due to a mistake in the directions for use located in the Herculite technique card. | Labeling mix-ups | Kerr Corporation issued a letter dated February 23, 2009. Users were informed of the affected product and instructed on proper return.For further information, contact Kerr Customer Care at 1-800-537-7123. |  | 442 | 442 | Nationwide distribution | Terminated | March 19, 2012 | 1120 | 442 | Not_Computer
Z-0919-2010 | 54360 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | BCI 3180 Pulse Oximeter uses power cord WW3005 (Electri-Cord Manufacturing Co., 312 E. Main St., Westfield, PA 16950). |  |  | January 15, 2010 | 2010 | Smiths Medical PM, Inc. | Smiths Medical is conducting a voluntary recall of a limited number of medical device AC power cords, manufactured by Electri-Cord Manufacturing Co., for use with BCI 3180 Pulse Oximeter(Human), 9200 Advisor Vital Signs Monitor(Human or Veterinary), 8400 Capnocheck II Capnometer(Human or Veterinary) and V6400 Invasive Pressure Monitor (Veterinary).The affected Electri-Cord AC power cords are eq | Nonconforming Material/Component | Consignees were sent a Smiths Medical PM, Inc letter dated January 12, 2010 starting on 1/15/10. The letter was addressed to Risk/Safety Managers, Distributors, Clinicians, Critical Care Physicians, Pediatricians, Pediatric Intensivists, Neonatologists, and Nursing, Clinical/Biomedical Engineering, and Anesthesia Professionals. The letter described the problem and products involved. It provided "Advice on Action to be Taken by the User" and requested the return of the "Urgent Recall Notice Confirmation Form." Customers are advised to monitor affected equipment. To return Power Cords, for questions regarding this Urgent Recall Notice, or to report any issues with these Power Cords, please contact the Smiths Medical Technical Service Department at 1-800-558-2345. |  | 1358 | 1358 cords for all devices (1032 US, 326 OUS) (871 Human, 487 Vet), (Cord # 8404 5 Vet &amp; 325 Human), (Cord # WW3005 482 Vet &amp; 546 Human) | US distribution: AL, AZ, AR, CA, CO, CT, DE, FL, GA, HI, IL, IN, IA, KS, KY, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, ND, OH, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, WI, WY.Worldwide distribution: Trinidad and Tobago, Mexico, Taiwan, Ecuador, Norway, Colombia, Brazil, France, Japan, Canada. | Terminated | March 14, 2012 | 789 | 1358 | Not_Computer
Z-0917-2010 | 53480 | No | KIF | Root canal filling resin. | Dental | Dental | 510(k) | Sealapex Express (Part Number: 33639), root canal filling material. Manufactured by SybronEndo, Glendora, CA. |  |  | February 23, 2009 | 2009 | Sybron Dental Specialties | Sybron Dental Specialties is initiating a recall on the Sealapex Xpress due to the presence of crystals in the catalyst. This crystallization may affect proper canal filling by preventing Gutta Percha Points from reaching the bottom of the root canal which may lead to an inadequate seal. | Process control | SybronEndo initiated an "Urgent Medical Device Recall" notification dated February 19 , 2009. Consignees were instructed to complete a "return form" and return any affected product in their inventory to the firm. Any product returned will be replaced at no charge.For further information, contact SybronEndo Customer Care at 1-800-346-3636. |  | 1309 | 1309 units | Worldwide Distribution -- United States, Australia, Canada, United Arab Emirates, Belgium, Bulgaria, Germany, Spain, France, UK, Greece, Italy, Kuwait, Lebanon, Portugal, Russia, Sweden, Hong Kong, India, Malaysia and Taiwan. | Terminated | March 21, 2012 | 1122 | 1309 | Not_Computer
Z-0968-2010 | 52901 | No | NGE | Luteinizing hormone test system. | Clinical Chemistry | Toxicology | 510(K) Exempt | OvuChek Saliva Ovulation Disposable Test Strips; 5 disposable test strips per pouch, packaged in a box with a specimen cup and instructions for use, 12 boxes per case; a rapid colorimetric indicator for the qualitative changes (ferning) of salt content in saliva to aid in the prediction of ovulation; Distributed by Quest Products, Libertyville, IL |  |  | November 07, 2008 | 2008 | Quest Products Inc. | The ovulation test strips were shipped without 510(k) premarket notification clearance from the FDA. | Other | Quest Products notified their major customer by telephone and e-mail on 11/7/08, requesting the return of all inventory of the OvuChek Saliva Ovulation Test Strips due to a lack of 510(k) clearance for marketing of the device. |  | 3930 | 3,930 cases | Nationwide | Terminated | March 03, 2010 | 481 | 3930 | Not_Computer
Z-0969-2010 | 54152 | No | MKZ | Chlamydia serological reagents. | Microbiology | Microbiology | 510(k) | BD Probetec" ET CT/GC Endocervical Specimen Collection and Dry Transport System-Pink Cap, Catalog number 440476, packed in shelfpack unit of 100 swabs. |  |  | September 18, 2009 | 2009 | Becton Dickinson &amp; Co. | Product was shipped beyond its expiration date. | Employee error | The recalling firm notified end users by letter flagged as "Urgent Product Recall" on 09/18/09. The notification advised that the referenced lot number had been correctly labeled with expiration date but had been shipped after the labeled date. End users were instructed to discontinue use discard product for replacement and return enclosed response form. If further assistance is needed regarding replacements, please contact SD Customer ServiceDepartment at 1-800-675-0908. For all other inquiries, please contact SD Technical Services Department at 1-800-638-8663. |  | 242 | 242 packs | The product was sold for distribution to clinical laboratories nationwide and to the firm's international affiliate in Australia. | Terminated | March 24, 2010 | 187 | 242 | Not_Computer
Z-0954-2010 | 54269 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | JOURNEY " UNI TIBIAL INSERT RIGHT MEDIAL/ LEFT LATERAL, A/P 52 MM, M/L 30 MM, REF 71422265, QTY: (1), STERILE EO, SMITH &amp; NEPHEW, INC. MEMPHIS, TN 38116 |  |  | January 23, 2009 | 2009 | Smith &amp; Nephew Inc | The firm has received complaints of the base plates breaking in situ. If a breakage occurs, the broken baseplate must be revised and/or changed. If left unrevised a broken base could lead to instability and premature wear of the construct. | Device Design | All affected Smith &amp; Nephew Sales Representatives were notified of problem and the recall overnight mail on 12/23/2009. All affected Smith &amp; Nephew International Distributors were notified via e-mail and telephone on 01/23/2009. They were instructed to immediately quarantine the product for return. For questions, please contact the firm, Leah Easley at 901-399-6137. |  | 1344 | 1344 units | US and Argentina, Switzerland, Australia, Canada, Belgium, Italy, Spain, Thailand, Austria; PROC, Finland, Germany, Malaysia, Portugal, Netherlands, Singapore, and England | Terminated | March 25, 2011 | 791 | 36181 | Not_Computer
Z-0923-2010 | 54381 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic Orthopedic Dovetail Clamp, REF 9730289, Medtronic Navigation, Louisville, CO 80027. |  |  | November 03, 2009 | 2009 | Medtronic Navigation, Inc | Orthopedic clamp may slip if used with non-indicated system, resulting in possible inaccurate navigation information being displayed. | Error in labeling | All affected users were notified by a Product Correction Notification letter on 11/03/2009 providing a description of the issue and informing them of the updated instructions for use for the Orthopedic Dovetail Clamp. Customers are to ensure that all users read the letter, place a copy of the letter with the IFUs for the related software applications listed in the letter, and fill out the attached response form indicating the receipt of the information. Further information is available at 720-890-3309. |  | 244 | 244 units | Worldwide Distribution -- USA, including VA, IA, MN and OK and countries of Australia, Austria, Belarus, Belgium, Brazil, Canada, China, Croatia, Czech Republic, Denmark, France, Germany, Hong Kong, Iceland, India, Israel, Italy, Japan, Korea, Lebanon, Malta, Mexico, Netherlands, Poland, Russian Federation, Saudi Arabia, South Africa, Switzerland, Turkey, United Arab Emirates, and UK. | Terminated | August 13, 2010 | 283 | 244 | Computer
Z-0978-2010 | 54068 | No | JTO | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BBL(tm) Lactophenol Cotton Blue Stain Droppers, Catalog number 261188, shelf package of 50 droppers, incorrectly labeled in part ***Methylene Blue Loeffler***Becton Dickinson and Company***Packaged in Mexico*** |  |  | September 28, 2009 | 2009 | Becton Dickinson &amp; Co. | An in-vitro diagnostic reagent used by clinical laboratories was mislabeled. | Labeling mix-ups | An "URGENT PRODUCT RECALL" notice dated September 2009 was issued via UPS letter to customers. The notification letter described the product, problem and the action to be taken by customer. The recalling firm notified distributors and end users to discontinue use of the identified product and discard for replacement. The customers were requested to return an enclosed response form. The customers were requested to provide, within one week, an excel file or a list with name, address and phone number of end-users who were shipped the product from their distribution branches so that the BD may inform them of the issue.For further assistance regarding product replacement contact our Customer Service Department at 1-800-675-0908. For all other inquiries call our Technical Services Department at 1-800-638-8663. |  | 100 | 100 shelf packs | Worldwide distribution: USA and Puerto Rico, Korea, and Singapore. | Terminated | March 24, 2010 | 177 | 100 | Not_Computer
Z-0974-2010 | 53560 | No | JCF | Lymphocyte separation medium. | Hematology | Hematology | 510(K) Exempt | Sigma-Aldrich Histopaque-1077, Ref. 10771, packaged in 500-ml. btls. The firm name on the label is Sigma-Aldrich, Inc., St. Louis, MO. |  |  | September 04, 2009 | 2009 | Sigma-Aldrich Mfg LLC | Some of the product was contaminated with Pseudomonas mendocina | Other | An "IMPORTANT PRODUCT NOTIFICATION" dated September 4, 2009 was sent to customers via regular mail. The notification letter described product, problem and action to be taken. The customers were requested to immediately discontinue use of the product and discard all remaining material. Subrecall was requested if the product was further distributed. The customer is to complete and fax the notification form to (314) 286-7829 to assure proper replacement for material they have not used and which is to be discarded.Should you have any questions, please contact Signma-Aldrich Technical Service at (800) 325-5832. |  | 6808 | 6,808/500-ml. bottles | Worldwide distribution: USA and Great Britain, France, China, Germany, Australia, Canada, Mexico, Korea, Singapore, India, Japan, Colombia, Hong Kong, Peru, Taiwan, and Venezuela. | Terminated | April 10, 2012 | 949 | 6808 | Not_Computer
Z-0775-2010 | 53827 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | GE Responder 2019198 automated external defibrillator. This semi-automatic model requires the user to press a shock button in order for the device to deliver defibrillation energy (when appropriate). |  |  | November 13, 2009 | 2009 | Cardiac Science Corporation | Potential for devices not to deliver therapy | Process control | Cardiac Science issued an "Urgent - Voluntary Medical Device Correction AED Resistor Issue" letter dated November 13, 2009. Consignees were informed of the affected product and instructed on preventive measures.For further information, contact Cardiac Science at 1-877-901-1788 in the US and +44-161-926-0011 for outside the US. |  | 4518 | 4518 | Worldwide Distribution -- UNITED STATES, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BELARUS, BELGIUM, BERMUDA, BOLIVIA, BRAZIL, BULGARIA, CANADA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, CZECHOSLOVAKIA, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, GRECE, GUAM, GUATEMALA, HOLLAND, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KENYA, KOREA, KUWAIT, LEBANON, LIBYA, LITHUANIA, MALAYSIA, MALTA, MEXICO, MONACO, MORROCCO, NAMIBIA, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SCOTLAND, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD &amp; TOBAGO, TURKEY, UAE, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM, VIRGIN ISLANDS AND YEMEN. | Terminated | October 24, 2014 | 1806 | 288468 | Not_Computer
Z-0973-2010 | 53750 | No | DCK | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | "K-ASSAY CRP (3) Calibrator E". |  |  | May 01, 2009 | 2009 | Kamiya Biomedical Company, LLC | The long-term stability of level 2 (1.0 mg/dL) may not be within the specifications due to a decrease in CRP concentration. | Other | On May 1, 2009 the recalling firm, Kamiya Biomedical Company contacted the 3 consignees by telephone and notified them of the issue. A follow-up e-mail notification on May 4, 2009 was sent to consignees requesting to return the affected product to Kamiya Biomedical Company for a free replacement kit. After the e-mail notifications, an Urgent Medical Device Recall Notice dated May 13, 2009 was also sent to consignees describing the issue and actions to take for returning opened and unopened affected product to the recalling firm for replacement. Direct questions to the Kamiya Biomedical Company by calling 1- 206-575-8068. |  | 15 | 3 kits each containing 5 x 2 ml vials | United States (NY, NJ, and NC). | Terminated | January 21, 2011 | 630 | 15 | Not_Computer
Z-0824-2010 | 54167 | No | DSQ |  |  | Cardiovascular | PMA | Abiomed AB5000 Circulatory Support System, Catalog Number: 0015-0000. Manufactured by Abiomed, Inc. Danvers, MA. |  |  | December 16, 2009 | 2009 | Abiomed, Inc. | Console shut down with no audible alarm | Process control | Abiomed initiated the recall by telephone contact and followed with a Recall Letter dated January 27, 2010. Consignees were informed that a Field Service Engineer will be in contact and schedule a time to visit and correct the affected device.For further information, contact Abiomed Quality Assurance at 1-978-656-1543 or Abiomed Field Service at 1-800-554-8666. |  | 496 | 496 units | Worldwide Distribution -- United States, Australia, Brazil, China, Hong Kong, Japan, Mexico and Germany. | Terminated | July 02, 2014 | 1659 | 496 | Other | Device Operation | Correction | Repair
Z-0976-2010 | 53956 | No | JCF | Lymphocyte separation medium. | Hematology | Hematology | 510(K) Exempt | Sigma-Aldrich Histopaque-1077, Ref. 10771, packaged in 100- and 500-ml. btls. and a 6x100-ml. bottle package. The firm name on the label is Sigma-Aldrich, Inc., St. Louis, MO. |  |  | December 08, 2009 | 2009 | Sigma-Aldrich Mfg LLC | Instructions for Use (IFU) was revised to recommend sterile filtering of the product under certain circumstances. | Other | The recalling firm issued an "Urgent Field Correction" notice dated December 7, 2009 via regular mail on December 8, 2009 which explains the reason for the field correction and instructs the user to review the enclosed, revised Instructions for Use (IFU) prior to further use of the product. The customer is requested to complete the acknowledgement information and return it to the recalling firm via FAX or e-mail. Direct questions to Sigma-Aldrich by calling 1-800-505-6275, extension 1232. |  | 63281175 | 11,015/100-ml. btls; 5,745/6/100-ml btl packages; and 14,872/500-ml. btls. | Worldwide Distribution -- United States, Canada, Mexico, United Arab Emirate, Australia, Bolivia, Brazil, Chile, China, Colombia, Costa Rica, Germany, Dominican Republic, Ecuador, France, Great Britain, Hong Kong, Israel, India, Italy, Jamaica, Japan, Korea, Panama, Peru, Paraguay, Saudi Arabia, Singapore, Thailand, West Indies, Taiwan, Venezuela, and South Africa. | Terminated | April 10, 2012 | 854 | 90393495 | Not_Computer
Z-0807-2010 | 54365 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | StatSpin Express 4 Horizontal Centrifuge |  |  | January 27, 2010 | 2010 | StatSpin, Inc., dba Iris Sample Processing | Centrifuge shield micro-switch failed due to damage, and the unit opened, ejecting pieces. | Device Design | Iris Sample Processing notified distributors by an Urgent Device Recall letter advising of the recall and the defect. Users are to immediately quarantine the affected products and choose between the two options listed in regards to correcting the affected centrifuge. Customers may return the unit or an upgrade kit with instructions will be sent to upgrade the unit. Questions should be directed to Gail Magrath at 800-782-8774 ext. 6123. |  | 1510 | 1510 units | Worldwide Distribution -- USA, CANADA, AUSTRALIA, JORDAN, UK, BELGIUM, FRANCE, CZECH, RUSSIA, KUWAIT', NORWAY, GREECE, TURKEY, DENMARK, FINLAND, SEOUL, KOREA, and HONG KONG. | Open, Classified |  |  | 1510 | Not_Computer
Z-0797-2010 | 54356 | No | ETD | Tympanostomy tube. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | MICRONÂż Âż BOBBIN VENT TUBE, 1.27 MM I.D., TITANIUM, REF 145281-ENT, QTY 1, STERILE EO, GYRUS ACMI INC., 136 TURNPIKE ROAD, SOUTHBOROUGH, MA 01772 USA. |  |  | January 08, 2010 | 2010 | Gyrus ACMI Corporation | Units of the product in lot MH136952 may have been shipped without being sterilized. | Pending | The firm initiated their recall on 01/04/2010 by telephone or visit to their consignees. A follow-up Urgent: Medical Device Recall correspondence, dated 1/14/2010, was sent on 01/27/2010, return receipt requested, to both the consignees and the implanting surgeons. The follow-up letter identified the affected product and described the issue. Customers are to immediately cease any further use of any affected product and quarantine it. Also, they are to obtain a Return Goods Authorization and return the affected product to the firm. The reply form should be completed and also returned. Questions should be directed to customer service at 1-800-773-4301. |  | 114 | 19 boxes of 6 units | Nationwide Distribution -- Including states of FL, GA, OH, TX, RI, IL, MI, and KS. | Terminated | July 27, 2010 | 200 | 114 | Not_Computer
Z-0863-2010 | 54344 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter UniCel DxC Synchron Clinical System ISE Flow Cell. |  |  | January 04, 2010 | 2010 | Beckman Coulter Inc. | Excessive Buildup of protein, bacteria and sample tube additives in the ISE flow cell may cause erroneous NA (sodium) results. | Device Design | An Urgent: Product Corrective Action letter dated January 4, 2010, was mailed to customers advising of erroneous sodium results derived from build up of protein, bacteria and additives in the ISE flow cell. The letter discussed the impact and actions to be taken along with some additional recommendations. Customers are advised to follow the new maintenance instructions included with the letter. The information in the letter should be shared with laboratory staff and the notification should be retained as part of the Quality System documentation. A copy of the letter should be provided to any other laboratory that affected product was forwarded to. The enclosed response form should be completed and returned. Questions should be directed to the Customer Support Center at 1-800-854-3633. |  | 2630 | 2630 | Worldwide Distribution -- USA and Canada. | Terminated | April 17, 2012 | 834 | 2630 | Not_Computer
Z-0815-2010 | 54525 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Transseptal Sheath Introducer Kits (under trade names: HeartSpan, Channel FX, Torflex, Braided Guiding Introducer Kit). The product is shipped 5 per carton. |  |  | February 01, 2010 | 2010 | Thomas Medical Products Inc | Radiopaque tip may fracture | Component design/selection | Thomas Medical Products, A GE Healthcare Company, issued an "Urgent Medical Device Correction" letter dated February 1, 2010. Consignees were informed of the affected product then asked to cease distribution, return any product to firm and contact all sub-accounts.For further information, contact Thomas Medical Products Customer Service at 1-866-446-3003. |  | 13973 | 13,973 units | Worldwide Distribution -- United States (CA, MA and NJ), Canada, Germany, and the United Kingdom. | Terminated | January 05, 2011 | 338 | 13973 | Not_Computer
Z-0788-2010 | 53954 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Multi-Lumen Central Venous Catheterization Kit with Blue Flex-Tip(R), ARROWg+ard Blue PLUS (R) Catheter, Sharps Safety Features and Maximal Barrier Precautions. (contains Q-Syte) |  |  | November 24, 2009 | 2009 | Arrow International Inc | Q-Syte Component contained in kits was recalled by manufacturer due to the potential for embolism if used with a central venous catheter. | Nonconforming Material/Component | The recalling firm issued an Urgent Medical Device Recall letter dated 11/24/09 to their customers informing them of the problem the need to return any kits in inventory and notify their customers of the recall. For additional information customers can contact the Customer Service Department at 800-523-8446. |  | 180 | 180 kits | The products were shipped to distributors in GA, MD, and VA. | Terminated | January 18, 2012 | 785 | 608 | Not_Computer
Z-1001-2010 | 50691 | No | LJB | Rubeola (measles) virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Measles IgM ELISA 96 Well Kit, Catalog Number: MS019M |  |  | August 27, 2008 | 2008 | Calbiotech Inc | The recall has been initiated due to the "For Export Only" labeling found on some Calbiotech, Inc. ELISA kits. Although there is no potential health hazard related to this recall, Calbiotech is fully committed to compliance with all State and Federal regulations. | Labeling False and Misleading | Immediately examine your inventory and quarantine product subject to recall. In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.Please return any product you currently have in stock which is subject to this recall and noted on the attached list. Any product you return will be replaced free of charge. Please use our FedEx account: 2284-6203-9 and send any packages via "FedEx Saver" service.This recall should be carried out to the wholesale level in order to comply with Federal regulations.Please complete and return the enclosed response form as soon as possible. If you have any questions, please contact David Barka at 619-660-6162. |  | 3 | 3 Kits | FL, and VA | Terminated | March 05, 2010 | 555 | 651 | Not_Computer
Z-1015-2010 | 51983 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Reference Electrode Membranes for use with blood gas analyzer ABL 800 FLEX, when used on ABL 827 and ABL 837 with Creatinine modules (8x7). The device is packaged four (4) units per box. |  |  | March 26, 2009 | 2009 | Radiometer America Inc | Analysis of customer data, (for ABL 8X7 customers) has shown that the reference membranes are out of analytical specifications when replaced as specified in the device labeling. These errors in analytical specifications could lead to inadequate or clinical interventions (therapy) or misleading interpretation of patient status in clinical settings. | Other | The recalling firm contacted their customers and notified them of this recall with a letter on 3/26/2009. The letter informs the customer that Radiometer has determined that the Reference Membrane D711 for use in the ABL8x7 Flex Creatinine Analyzer has a shorter in-use lifetime than the previously recommended 1-month interval listed in the Operators Manual. It states that consequently, the firm has changed the recommended replacement intervals as they correspond to the number of samples per day analyzed by the customer and a revised insert is being supplied to the customer which should be placed in the Operators Manual to reflect this new information. The recall notification is also accompanied by an attached response FAX FORM which the customer is asked to complete and to fax back to Radiometer so as to acknowledge receipt of the recall notification letter and the revised insert.Please direct questions to the company TechLine at 1-800Âż736Âż0600 opt 2. |  | 50 | 50 membranes in U. S. distribution | The recalled product was distributed customers located in the following states within the U.S.: CT, FL, GA, IL, MA, MD, MN, MO, NC, NM, OH, PA, VA, WA, WI, and WV. Product was also sold in the following countries: Algeria, Australia, Austria, Bahrain, Brazil, Czech Rep., Denmark, England, Estonia, France, Germany, Greece, Hungary, India, Italy, Japan, Netherlands, Norway, Poland, Russia, Spain, Sweden, Switzerland &amp; Turkey. | Terminated | May 09, 2012 | 1140 | 50 | Not_Computer
Z-1005-2010 | 52101 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Whatman Body Fluid Collection Paper 903, Ref: 10534790, for IVD use. Whatman Inc., Florham Park, NJ. |  |  | May 19, 2009 | 2009 | Whatman, Inc. | '903' newborn Neonate screening cards have duplicate numbers assigned on cards and may result in cross-over of patient test results | Process control | Whatman issued an "Urgent Medical Device Correction " letter on May 19, 2009. The consignee was instructed to remove cards with duplicate numbers. For further information, contact Whatman Inc. Quality Assurance at 1-207-459-7557. |  | 11300 | 11,300 Cards | Nationwide Distribution -- CA. | Terminated | September 28, 2015 | 2323 | 11300 | Not_Computer
Z-1014-2010 | 52851 | No | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pointe Scientific Direct Bilirubin Reagent Set, Pointe Scientific Inc., Canton, MI., Catalog # HB936-1250. |  |  | July 29, 2009 | 2009 | Pointe Scientific, Inc. | R2 reagent component may fail because it may contain R1 reagent. The reagent was rendered inoperative and would not pass the required quality control (calibration and quality control testing) before beginning running the bilirubin assay. | Error in labeling | Consignee was notified by letter dated 7/29/2009 and instructed to identify and notify them at once of this product recall. Product is to be returned or destroyed. Customers with questions should contact the firm directly. |  | 2 | 2 | California. | Terminated | December 06, 2010 | 495 | 2 | Not_Computer
Z-1012-2010 | 53246 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS Panels CHOL+HDL+GLU Panel test strips for professional use with CardioChek P-A analyzers, Manufactured by Polymer Technology Systems, Inc., Indianapolis, IN. |  |  | August 04, 2009 | 2009 | Polymer Technology Systems, Inc. | Possible under recovery of cholesterol, HDL and glucose. A loss of recoverable cholesterol and HDL cholesterol is a manageable to moderate health risk. The clinical applications analysis shows that lot 1891 shows such a loss of activity. A loss of recoverable Glucose measurement is a moderate to high risk, based upon the degree of under-recovery, and the frequency of occurrence. | Other | Consignees were notified by letter dated 7/28/09 to discontinue use of, and to destroy, the reagent and to contact the firm for replacement product. For questions, please contact the company Customer Service toll-free at 877-870-5610 or + 1-317-870-5610 (outside the US). |  | 962 | 962 | Nationwide. | Terminated | October 20, 2010 | 442 | 962 | Not_Computer
Z1006-2010 | 53708 | No | NPR | Factor V Leiden DNA mutation detection systems. | Hematology | Hematology | 510(k) | FACTOR II (PROTHROMBIN) G20210A Kit, LightCycler 2.0 Instrument. Material Number 03610195 001 |  |  | March 12, 2009 | 2009 | Roche Molecular Systems, Inc. | Within the Precautions and Warning Section, Polymorphism heading, of the Product's Package Insert, it stated that mutations at positions 20207, 20209, 20218 and 20221 exist and these mutations are spanned by the mutation probe. These rare mutations will lead to an unknown result after performing genotyping. | Labeling design | Product Advisory Notices were sent via e-mail on 3/12/09 to Roche Local Safety Officers and Roche county General Manager. Local Safety Officers and General Managers were to communicate to the end users with the Important Product Notice. In addition, Reagent Bulletin 09-052, dated 3/27/09, was issued to Roche U.S. Affiliate of U.S. customers. |  | 9624 | 9624 kits | Nationwide Distribution | Terminated | June 09, 2010 | 454 | 9624 | Not_Computer
Z-1013-2010 | 52719 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bio-Rad VARIANT II Hemoglobin A1c Program Recorder Pack with Calibrator/Diluent Set, Model Number 270-2101NU Manufactured by Bio-Rad Laboratories, Hercules, CA |  |  | June 15, 2009 | 2009 | Bio-Rad Laboratories Inc | Calibration Failure-- due to high intercept values, the device fails calibration and users are unable to run samples. | Other | Recall initiated June 15, 2009, with notification initially conducted by phone with Fax Notification follow up domestically. Each subsidiary will be given an urgent notice, subsidiary response form, customer notification, customer acknowledgement form. Each subsidiary is to contact customers and translate notification and response form as needed. Contact the recalling firm for additional information. |  | 613 | 613 units | Product was distributed to 127 customers throughout the US and internationally to Argentina, Australia, Canada, Chile, Czech Republic, France, Germany, Hong Kong, Hungary, India, Italy, New Zealand, Poland, Singapore, Sweden, Taiwan, and the United Kingdom | Terminated | February 02, 2011 | 597 | 613 | Not_Computer
Z-0980-2010 | 54039 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista LOCI Reaction Vessel |  |  | November 24, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Some vessels may leak. If leakage occurs, patient results will be flagged and are not reportable. | Nonconforming Material/Component | The recalling firm issued an Urgent Field Safety Notice dated 11/2009 to all affected customers. Customers have been told to discard any bags of the affected lot. If the affected lot is the only lot the customer has, customers can examine a vessel from the bag. If the vessel mold cavity identity is B8, customers should discard the bag. Thus customers can use the lot if the cavity mold is anything other than B8. In addition, customers have been told to empty the LOCI vessel holder on all systems and to discard vessels in case any affected vessels have been loaded onto the system.For technical questions or concerns, please contact the Siemens Technical Solutions Center at 800-441-9250 for further assistance. |  | 408 | 408 units | The device was shipped to medical facilities in CA, CO, FL, GA, IL, LA, MD, MI, MS, NE, NC, NJ, NY, OH, SC, TN, TX, VA, WA, and WI. The product was also shipped to Belgium, France, Germany, Italy, Netherlands, Portugal, Spain, and Switzerland. | Terminated | September 08, 2011 | 653 | 408 | Not_Computer
Z-1008-2010 | 53961 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BD BACTEC(TM) Plus + Aerobic/F Medium, catalog 442192, packed in tubes, 50 tubes/case and labeled in part ***Becton Dickinson and Company, Sparks, MD 21152*** |  |  | November 11, 2009 | 2009 | Becton Dickinson &amp; Co. | The draw volume on test tubes of blood culture medium is marked with graduations that may be confusing for healthcare practitioners using the tubes for appropriate collection of patient samples for microbiological laboratory testing. | Component change control | An "URGENT CUSTOMER NOTIFICATION" was sent via UPS to customers. The notification letter described the product, problem and action to be taken. The letter alerted customers that the firm had recently manufactured identified lots of BACTEC(tm) Blood Culture Media with a new label containing a red line which should not be used as an accurate indication for blood draw volume. The firm also advised that the manufacturer is working to build stock with a revised label, but all of the identified lots distributed had passed testing and release criteria and will continue to perform as expected when used as instructions indicate. The customers are being urged to pay close attention to their BACTEC blood collection procedure until they receive bottles with the revised labels. Drawing too much blood may contribute to false negative or false positive results. The customer is being ask to communicate this information to all specimen collection sites within their facility.If further assistance is needed, please contact BD Technical Services Department at 1-800-638-8663. |  | 2940473 | 2,940,473 bottles | Worldwide Distribution: USA and Canada, Japan, Australia, Singapore, Korea, Brazil, Taiwan, China, Mexico, Germany, New Zealand, Hong Kong, Uruguay | Terminated | May 13, 2010 | 183 | 5025698 | Not_Computer
Z-1020-2010 | 53541 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | RapidPointÂż 405 System Analyzers |  |  | October 30, 2009 | 2009 | Siemens Healthcare Diagnostics Inc | Software: Reference sensor failure may result in inaccurate pH reuslts | Component design/selection | An "Urgent Field Safety Notice" dated September 2009 was sent via overnight mail to customers and included software Version 3.5.2 with instructions for install that will detect and prevent the reporting of potentially errant test results. The customers are requested to please install the included software update immediately and return the attached Fax back form. Siemens will provide the mandatory update to all customers with the Rapidpoint 400 /405 test system.If you have any questions, please contact your local Technical Support at 508-359-3877. |  | 4095 | 4095 units | Worldwide distribution: USA and Australia, Austria, Bahamas, Barbados, Belgium, Bosnia Herzeg, Brazil, Canada, Canary Islands, Chile, Croatia, Czech Republic, Egypt, France, Germany, Holy See (Vatican City State)Hong Kong, Ireland, Italy, Japan, Macedonia, Malaysia, MÂżxico, New Zealand Not assigned P.R. China, Poland, Portugal, Puerto Rico, Refurb.Systems U.S.ARepublic Korea, Saudi Arabia, Singapore, South Africa, Spain, Switzerland, U.A.E., United Kingdom | Terminated | September 04, 2015 | 2135 | 5093 | Software | Output/Calculation | Software update | Software Update
Z-1017-2010 | 52958 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista System BUN Flex Reagent Cartridge, Catalog number K1021, in vitro diagnostic. |  |  | June 16, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Incorrect result: patient sample BUN results reported may be higher or lower than actual level. | Nonconforming Material/Component | Siemens issued an Urgent Field Safety Notice dated 6/16/09 to all affected customers instructing them to discontinue use of the Dimension Vista BUN Flex reagent cartridge lot 09061AA. |  | 1254 | 1254 cartons | Worldwide distribution: USA, Australia, Canada, Germany, Japan, and Korea. | Terminated | November 30, 2010 | 532 | 1254 | Not_Computer
Z-0972-2010 | 53546 | No | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | Roche/Hitachi GENT2 ONLINE TDM Gentamicin for Roche cobas c 501 analyzer systems; Catalog no. 04490843190. |  |  | October 07, 2009 | 2009 | Roche Diagnostics Operations, Inc. | A high recovery of the assay may be generated, which would lead to an underdosage of gentamicin being administered to the patient. | Other | Consignees were notified by recall letter dated October 7, 2009. The letter described the affected products, issue and instructed consignees to immediately discontinue use of the products. Please contact Roche Diagnostics Technical Support at 1-800-428-2336 if you have questions aboutthe information contained in the recall letter. |  | 13187 | 13,187 of all products. | Nationwide Distribution. | Terminated | December 10, 2010 | 429 | 13187 | Not_Computer
Z-1024-2010 | 53650 | No | JSM | Transport culture medium. | Microbiology | Microbiology | 510(k) | BD Vacutainer Urine Complete Kit with UA Tube and S&C; Preservative Tube for Midstream specimens; |  |  | October 12, 2009 | 2009 | Becton Dickinson &amp; Company | A small amount of discoloration was noted on Towelettes included in BD Vacutainer Urine collection kits. The entire kitcomponents retain sterile interiors and no viable growth was detected on the towelettes. | Other | Recall notification letters were sent on 10/12/09 with return receipt requested. Customers were advised to contact your SO Sales Consultant for additional information, via e-mail at Vacutainer_TechServices@SD.com or via telephone at 1-800-6310174. |  | 588900 | 588900 | Nationwide distribution to hospitals and medical device distributors. | Terminated | March 10, 2010 | 149 | 1217900 | Not_Computer
Z-1018-2010 | 54131 | No | LCX | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Mainline Confirms hCG Urine Pregnancy Test Catalog # 6008C. |  |  | October 07, 2009 | 2009 | Mainline Technology, Inc. | Lack of assurance of safety and efficacy: unapproved for marketing in the U.S., non-compliance with cGMP regulations and no stability data to support labeled expirations dates. | PMA | Consignees were notified via recall letters dated October 7, 2009 to cease using and to return the recalled products to Mainline Technology. Direct questions about the recall to Mainline Technology by calling 1-800-541-4467 or 1-734-930-2043. |  |  | Information not provided. | Worldwide Distribution -- United States, Belgium and Switzerland. | Terminated | February 04, 2011 | 485 |  | Not_Computer
Z-1021-2010 | 54089 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | SmartInfuser PainPump, part numbers P49230, P49220, P49224-D, P49224N, P49230-D and P49234. |  |  | January 08, 2009 | 2009 | Hsmg, Inc. D/b/a Smartinfuser Usa | Product leaks associated with a new pump reservoir, and incident related to regulator fast priming feature. | Process control | SmartInfuser USA sent letters to consignees on 1/8/2009 and 1/22/2009. |  | 6393 | 6,393 units | Nationwide. | Terminated | April 25, 2011 | 837 | 6837 | Not_Computer
Z-1016-2010 | 51296 | No | KEY | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Xmatrx FISH (fluorescent in situ hybridization) Automated Staining System; Model Number/Catalog Number: AS4010B; |  |  | January 23, 2009 | 2009 | Biogenex Laboratories | Under certain conditions, product system failure has the potential to cause operator injury due to splash of Hemo-De solution. | Device Design | On 1/23/09, the firm issued an Advisory Notice to all its direct consignees, informing them of the affected product and providing instructions on the recall.Please contact BioGenex Technical Support at 1-800-421-4149 for any assistance you may need. |  | 12 | 12 units | The firm distributed the product to one (1) distributor as a consignee of our product. The firm indicated that no customer has received the product directly from BioGenex, due to the fact that the product is exclusive distributed by Abbott Molecular. | Terminated | December 02, 2010 | 678 | 12 | Not_Computer
Z-1040-2010 | 54232 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Teleflex Medical, Hudson RCI, Concha Therm Neptune, Heated Humidifier. Product numbers: 425-00, 425-10 and 525-30. |  |  | January 12, 2010 | 2010 | Telefelx Medical | When turned on, the Neptune heater may cause an artifact on the patient's heart monitor. | Device Design | Teleflex consignees were notified via Urgent Medical Device Recall letter via next day delivery on/about 01/12/2010. They were asked to return all of the affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. A second mailing to non-responders on the consignee list will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the second mailing or as an alternative, they may be contacted by phone, e-mail or fax. A record of these contacts will be maintained. Affected products received by Teleflex Medical will be reworked with software upgrade and placed back into distribution. Consignees will either receive a software upgrade on site by a third party or replacement product with the software upgrade. Records will be maintained accordingly. |  | 15134 | 15,134 units | U.S. Canada, India, Japan, Puerto Rico, Germany, France, Italy and Ireland. | Terminated | April 05, 2011 | 448 | 15134 | Software | Output/Calculation | Software update | Software Update
Z-0979-2010 | 54194 | No | JSC | Microbiological assay culture medium. | Microbiology | Microbiology | 510(K) Exempt | Oxoid CM0225 Brain Heart Infusion IVD. The firm name on the label is Oxoid Ltd., Basingstoke, Hampshire, England. |  |  | December 21, 2009 | 2009 | Remel, Inc | Product failure - Failure to grow quality control strains of Streptococcus pneumoniae. | Other | The recalling firm issued separate recall letters dated December 21, 2009 addressed to Microbiology Laboratory Supervisors or Distributors explaining the reason for recall and requesting all remaining product be discarded. A Product Inventory Checklist was enclosed to report the amount of product that was discarded by the customer. The Distributor letter requests the consignee contact their customers to notify them of the recall. Direct inquiries concerning this recall to the Technical Services Department at 1-800-255-6730 (United States) or 1-913-888-0939 (International). |  | 105 | 105 bottles | United States (FL, NY, MO, NY, DE, and PA). | Terminated | July 12, 2011 | 568 | 105 | Not_Computer
Z-1031-2010 | 53893 | No | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Exacta Mix TPN (total parenteral nutrition) Bag, Model 739, 1000mL, Baxa Corporation |  |  | November 12, 2009 | 2009 | Baxa Corporation | TPN bags may leak fluid due to inadequate sealing. | Process control | Consignees were notified by Baxa Urgent Exactamix Bag Recall letter on 11/12/2009 and then again on 11/17/2009 when more lots were found to be affected. Instructions were to stop use and return all affected product to Baxa. More information is available at 800-678-2292. |  | 105500 | 2110/50 bag cases | Worldwide distribution, including USA, Argentina, Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, Hong Kong, Japan, Malaysia, Mexico, New Zealand, Panama, Peru, Singapore, South Korea, Taiwan, and Uruguay. | Terminated | February 24, 2011 | 469 | 293300 | Not_Computer
Z-1036-2010 | 54094 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Pointe Scientific Creatinine Reagent R2, Pointe Scientific, Inc., Canton, MI; Catalog 3-C7539-L, 7-C7539-R2-30, 7-C7539-R2-125, 7-C7539-R2-250, 8-C7539-R2-30, 8-C7539-R2-125 and 8-7539-R2-250. |  |  | December 01, 2009 | 2009 | Pointe Scientific, Inc. | An additional 10% picric acid was added to the R2 component during production. | Process change control | Consignees were notified by An Urgent: Medical Device Recall letter, dated 12/1/09, that identified the affected product. Customers were asked to immediately examine their inventory for the affected product and to notify any customers that the product may have been further distributed to by giving them a copy of the recall notice. Customers are to dispose of the affected product immediately. Questions should be directed to the Technical Support Department at (800)757-5313 or (734)487-8300. There is also a Recall Response form attached with this notice. |  | 625 | 625 vials | Worldwide Distribution -- USA, Bolivia, Greece, Guyana, India, Indonesia, Israel, Kenya, Korea, Malaysia, Mexico, Pakistan, Philippines, Russia, Serbia, Sri Lanka, Tanzania, Trinidad and Ukraine. | Terminated | October 15, 2010 | 318 | 1517 | Not_Computer
Z-1053-2010 | 54080 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Latex-Free Symbiq Pump Set, convertible pin, 104 inch piggyback with backcheck valve, 2 prepierced injection sites, 0.2 micron filter and OPTION-LOK, non-DEHP, 15 drops/mL; a sterile Rx single use device; Hospira Inc, Lake Forest, IL USA. Made in Costa Rica; product 160410428, list number 16041-28. |  |  | November 24, 2009 | 2009 | Hospira Inc. | Company has confirmed reports of backflow of fluid past the backcheck valve on Symbiq administration sets. | Process control | Hospira, Inc. issued an "Important Administration Set Information" letter dated November 24, 2009 to their direct accounts, informing them of the affected product. Consignees were instructed to return the defective product to the firm.For further information, contact Hospira Medical Communications at 1-800-615-0187. |  | 21312 | 21,312 sets | Worldwide Distribution -- United States, Australia, Canada, Costa Rica, India, Mexico, Netherlands, Philippines, Singapore and United Arab Emirates. | Terminated | January 29, 2012 | 796 | 3558848 | Not_Computer
Z-1027-2010 | 54187 | eosin y | HYB | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | SS-035C, Hematology Reagent C Eosin Stain, 500 mL, IVD, |  |  | January 06, 2010 | 2010 | Wescor, Inc | Stains used for diagnostic use may be contaminated with bacteria. | Component design/selection | Customers were notified by letter on Jan. 6, 2010 and asked to destroy any remaining product. A response form was included to note whether or not customer had any remaining stock for replacement purposes. Additional information is available at 800-453-2725 or regaff@wescor.com. |  | 24466 | 24,466 bottles | Nationwide, Foreign distribution to Algeria, Australia, Austria, Canada, Cyprus, Denmark, France, Germany, Hungary, Ireland, Italy, Japan, Kuwait, Malta, Netherlands, Oman, Panama, Poland, Qatar, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, United Arab Emirates, UK, and Venezuela. | Terminated | April 20, 2010 | 104 | 26962 | Not_Computer
Z-1039-2010 | 53686 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista V-LYTE Fluids: Diluent catalog number K835 |  |  | October 22, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Results can elevate after sensor replacement. Sodium urineresults can be falsely elevated 15 mmol/L within 24 hours after installation of a new sensor. | Other | The recalling firm issued an Urgent Field Safety Notice dated October 2009 to all affected customers. Customers were given instruction which minimize urine specimen use and report urine specimen results when QC is acceptable. Sensors are to be replaced when urine QC is out. Replacement sensors are available by contacting the firm at 800-241-0420. Please contact the Siemens Technical Solutions Center at 800-441-9250 for further assistance. |  | 150 | 150 units | The product was shipped to medical facilities nationwide. The product was also shipped to Australia, Canada, Belgium, France, Germany, Italy, Japan, Malaysia, New Zealand, Netherlands, Portugal, Spain, South Korea, and Switzerland. | Terminated | December 16, 2010 | 420 | 150 | Not_Computer
Z-1043-2010 | 53784 | No | CIC | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Calcium (CPC) R2 Reagent Set. Pointe Scientific Liquid Calcium (CPC) R2 Reagent Set, Pointe Scientific, Inc, Canton, MI; Catalog #'s HC402-R2. Calcium CPC Hitachi Reagent Set for repackaging, Catalog # 3-HC902-R2. |  |  | November 10, 2009 | 2009 | Pointe Scientific, Inc. | Non-compliance with the Good Manufacturing Practice regulations, as the manufacturing process for this product had not been validated. | Process control | Consignees were notified by "Urgent: Medical Device Recall" letter dated November 11, 2009 and instructed to conduct a recall to the retail level, that users should dispose of the product immediately, and that a business reply card should be returned to account for products. If you have any questions, please call Pointe Scientific, Inc. Technical Support Department at 1-800-757-5313 or 1-734-487-8300. |  | 13 | 4 sets and 9 L | Worldwide Distribution (United States, Chile, India, Republic of Yemen and Trinidad). | Terminated | December 10, 2010 | 395 | 84 | Not_Computer
Z-1033-2010 | 53675 | No | LFX | Rubella virus serological reagents. | Microbiology | Microbiology | 510(k) | IMMULITE 2000 and IMMULITE 2500 |  |  | October 21, 2009 | 2009 | Siemens Healthcare Diagnostics Inc. | The recall was initiated because Siemens Healthcare Diagnostics has identified there is the potential for patient results to be biased low (-0.3 Sample to Cutoff [S/CO] ratio units) due to incompatibility of the Toxoplasma IgM kits with this particular lot of IgG/IgM Sample Diluent. All available lots of the IMMULITE 2000/2500 Toxoplasma IgM kits (L2KTM/L5KTM) are incompatible with Diluent lot 150 | Nonconforming Material/Component | The customer notification was initiated on 10/22/2009 with the firm forwarding an Urgent Field Safety notice with attached Field Correction Effectiveness Check form to the affected consignees informing them of the issue and the products affected. The consignees were informed that Siemens is initiating the Field Correction (FC) to prevent IMMULITE 2000/2500 Toxoplasma IgM from being run with Diluent lot 150.The consignees were instructed in the interim, until replacement kit lots of RUB IgG and CMV IgG are received, your laboratory can continue using Rubella IgG, CMV IgG, or Toxoplasma IgM as long as Diluent Lot 150 is not loaded on the analyzer when running Toxoplasma IgM. Diluent Lot150 can be used with RUB IgG and CMV IgG.Once the replacements lots are received, any unopened kit lots listed in the above table MUST BE RETURNED to Siemens per instructions that will accompany a separate Fed Ex Call Tag that will explain the proper return of the affected product.The consignees were also instructed to complete the attached Field Correction Effectiveness Check form and fax to the firm's technical solutions center. Additional questions should be directed to the firm at 973-927-2828. |  | 557 | 557 units | Nationwide | Terminated | April 22, 2010 | 183 | 567 | Not_Computer
Z-1064-2010 | 53942 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer NexGen Complete Knee Solution Minimally Invasive Solutions Procedures Trabecular Metal Technology Tibial Tray, fixed bearing, size 8, sterile, Zimmer, Warsaw, Indiana; REF 00-5954-057-02. |  |  | November 16, 2009 | 2009 | Zimmer Inc. | The titanium portion of the implant may separate from the trabecular metal material. | Other | On December 1, 2009, an e-mail was sent to consignees for lot 60866613 to quarantine and return the product. On December 2, 2009 an e-mail was sent to U.S. distributors, and on December 3, 2009 and e-mail was sent to international distributors, instructing them to suspend usage and return all lots of size 8 trays. On December 9, 2009, a recall letter was sent to U.S. user accounts who received the products directly from Zimmer, U.S. distributors and Zimmer's U.S. sales force informing them of the recall. Direct questions about this recall to Zimmer, Inc. by calling 1-800-613-6131. |  | 846 | 846 | Worldwide Distribution -- United States, Australia, Austria, Belgium, Canada, Denmark, France, Germany, Italy, Saudi Arabia, Spain, Sweden, Switzerland and the United Kingdom. | Terminated | March 23, 2011 | 492 | 846 | Not_Computer
Z-0816-2010 | 54526 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Torflex Transseptal Guiding Sheath, Baylis Medical Company Inc., 5959 Trans-Canada Highway, Montreal, QC, Canada H4T1A1, (514) 488-9802; INCLUDES: One (1) Radiopaque Sheath, One (1) Radiopaque Dilator; And One (1) B* x 135 cm Guidewire. |  |  | February 08, 2010 | 2010 | Baylis Medical Corp * | Radiopaque soft tip may fracture. | Component design/selection | Baylis Medical Issued an "Urgent Medical Device Correction" letter via overnight courier beginning February 8, 2010. Consignees were instructed to segregate and return all affected product to the firm.For further information, contact Baylis Medical Regulatory Scientific Affairs at 1-905-602-4875 Extension 252. |  | 1085 | 1,085 units (508 in US, 577 Foreign) | Worldwide Distribution -- United States, Canada, Chile, China, Colombia, Denmark, Egypt, France, Germany, Hong Kong, Italy, Mexico, Netherlands, New Zealand, Portugal, Qatar, Russia, Singapore, South Africa, Sweden, Switzerland, United Arab Emirates and the United Kingdom. | Terminated | September 13, 2010 | 217 | 1085 | Not_Computer
Z-1044-2010 | 53896 | No | BSN | Anesthesia conduction filter. | Anesthesiology | Anesthesiology | 510(k) | Millipore Millex Syringe Filter Unit, 33 mm Durapore PVDF 0.22um, Sterile, 50 Filters per box. Catalog Number: SLGVM33RS. |  |  | November 25, 2009 | 2009 | Millipore Corporation | Membrane integrity defects may allow the passage of viable or non-viable contaminants to pass through the filter. | Process control | Millipore Corporation issued an "Urgent: Medical Device Recall" letter dated November 24, 2009 requesting users to examine inventory and immediately discontinue use and return all affected product to the firm.For further information, contact Millipore Technical Services at 1-800-645-5476. |  | 26950 | 539 Boxes - 50 per box | Worldwide Distribution -- United States, Japan and Hungary. | Terminated | May 23, 2011 | 544 | 26950 | Not_Computer
Z-1067-2010 | 53841 | No | JSO | Culture medium for antimicrobial susceptibility tests. | Microbiology | Microbiology | 510(k) | CHROMagar VRE Base; |  |  | November 09, 2009 | 2009 | DRG International Inc | CHROMagar Products are labeled incorrectly. Label currently states that the product is for IVD use, when it should state that the product is intended For Research Use, Not for use in diagnostic procedures. | Other | Two recall letters were sent via certified mail to customers. One letter was dated November 9, 2009 referencing the CHROMagar VRE product and the other letter dated November 12, 2009 referencing the CHROMagar 0157 product. The letters describe the affected product, problem, and action to be taken by the customer. The customers are to please STOP the use or distribution of the listed products and IMMEDIATELY examine your stock and return any un-expired CHROMagar product slated for recall for replacement the DRG International, Inc. The customers are being requested to please fill out and return the disposition sheet even if you have none of these items on hand since the information is important.If you have any questions, please contact customer service at (908) 233-0758. |  | 325 | 325 (5 x 5000ml), 65 5 L packs | Nationwide distribution | Terminated | May 14, 2012 | 917 | 2561 | Not_Computer
Z-1081-2010 | 54253 | No | GXQ | Dura substitute. | Neurology | Neurology | 510(k) | Suturable DuraGen Dural Regeneration Matrix; |  |  | January 13, 2010 | 2010 | Integra LifeSciences Corp | One lot of Suturable DuraGen Dural Regeneration Matrix (# 1091575) did not pass stability testing at certain time points for Suture Retention Strength and Shrink Temperature specifications. | Other | An "URGENT PRODUCT RECALL NOTIFICATION" dated January 13, 2010, was sent via Federal Express to customers. The notification described the product, problem, hazard involved and action to be taken by customers. In the event that the customers have in their possession any products from the specific lot number 1091576, the customers are being ask to please immediately quarantine all stock on hand and contact their local Integra Sales representative for further assistance and for arrangements of product return. The customers are being ask to return the enclosed "Recall Acknowledge and Return Form" indicating whether or not they possess any of the affected product identified. Integra will replace any product affected by this Recall that is returned.Please feel free to contact me at 609-936-5407 should you have any additional questions. |  | 94 | 94 | Nationwide distribution:Ca, NY, TX, CO, VA, SC, IA, MO, IL, LA, OR, OH, WI, FL, DC, DE, MD, MN. | Terminated | March 18, 2010 | 64 | 94 | Not_Computer
Z-1069-2010 | 54076 | centrifuges (micro, ultra, refrigerated) for clinical use | JQC | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CLAY ADAMSÂż DynacÂż III Centrifuge, 120V, Model Number 420104 |  |  | August 17, 2009 | 2009 | Becton Dickinson &amp; Co. | Medical laboratory centrifuge component may have been assembled incorrectly and cause disabling of the safety feature which prevents customers from opening the centrifuge while it is spinning. | Nonconforming Material/Component | An "Urgent Product Recall" letter dated August 2009 was sent to the customers. The letter described the affected product, problem and the action to be taken by customers. The customers were advised to discontinue use of the centrifuge for repair or replacement with refurbished unit , provide an Excel file or a list with name, address and telephone number of end-users who were shipped this product from their distribution branches so that BD may inform them of the issue,and return the enclosed response form.For further assistance regarding product replacement and returns contact our Technical Services Department at 1-800-638-8663. |  | 411 | 411 units | Worldwide distribution: USA and Mexico, Phillipines, Columbia, Chile, Ecuador, Venezuela, Puerto Rico, Canada and Singapore. | Terminated | April 22, 2010 | 248 | 411 | Not_Computer
Z-1100-2010 | 53892 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | PET-511BTM Multi-plane transesophageal transducer (TEE) used with Toshiba Diagnostic Ultrasound System |  |  | May 11, 2009 | 2009 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems Inc is removing the PET-511BTM Multi-plane Transesophageal Transducer from the market after receiving several reports of injury regarding the transducer. Bending or buckling of the flexible portion of the PET-511 BTM multi-plane transesophageal transducer (TEE) may result in injury including perforation. | Device Design | Toshiba America Medical Systems sent a Field Safety Notice to all customers on 5 MAY 2009 indicating its intent to remove the units from the market. The letter requested immediate cessation of use and return to the manufacturer site. It announced plans to replace unit with a new multi-plane transducer at no addle charge.The notice also included a description of the problem that occurred, the proper insertion technique into the esophageal passage, how to avoid and detect bending or buckling of the transducer, and what to do if the transducer does buckle. A revised safety manual and poster reflecting these precautions for the PET-51 IBTM was sent to each customer using this transducer. Shipments of the transducer from December 2008 also contained the Field Safety Notification materials. For additional information, contact the firm at (714) 669 2458. |  | 70 | 70 | Nationwide | Terminated | March 22, 2012 | 1046 | 70 | Not_Computer
Z-1102-2010 | 54230 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Howmedica Osteonics Corp. Restoration Modular Revision Hip System, Plasma Distal Stem; Titanium, Alloy and Hydroxyapatite. Catalog number: 6276-5-014. |  |  | December 30, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | The Restoration Modular Hip Stem may have two different package labels. The 14 mm diameter stem may be labeled with both a size 14 mm diameter and a 26 mm diameter. | Process control | Stryker Branches were notified by e-mail on December 30, 2009. This was followed by a formal recall letter by Fed Ex sent on 1/11/10 to branches as well as the hospital and surgeon that used the device. |  | 1 | 1 | 1 device distributed to a hospital in GA. | Terminated | September 14, 2012 | 989 | 1 | Not_Computer
Z-1071-2010 | 54113 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | DeRoyal TOTAL KNEE TRACECART, REF 53-1818.13, Rx ONLY, NON-STERILE. Custom surgical kit containing Stryker CBCII Blood Conservation Kit. |  |  | December 09, 2009 | 2009 | DeRoyal Industries Inc | Custom surgical kits contain Stryker CBCII Blood Conservation Kit which may lack sterility | Nonconforming Material/Component | DeRoyal initiated its recall by visit on 12/09/2009 and followed by letter dated 12/18/2009 |  | 194 | 194 units | LA &amp; MS | Terminated | May 20, 2010 | 162 | 313 | Not_Computer
Z-1065-2010 | 54117 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom TotalCare Bed System, Hill-Rom, Batesville, IN; Model P1900 and P1840. |  |  | January 01, 2010 | 2010 | Hill-Rom, Inc. | The controls for nurse call, bed movement and entertainment functions on the intermediate rail may fail to function. | Nonconforming Material/Component | The firm sent a letter to each consignee dated January 4, 2010 advising them that a firm representative will visit the facility to make corrections. Hospitals were requested to take all of the affected beds out of service, if possible, until they are corrected, or to inspect and remove those with a problem out of service until they are corrected if all cannot be removed from service. If you have any questions concerning the recall, please contact Hill-Rom Technical Support at 1-800-445-3720 (U.S.). |  | 185 | 185 | Worldwide Distribution -- United States, Australia, Canada, France, Germany, Mexico, United Arab Emirates and United Kingdom. | Terminated | February 17, 2012 | 777 | 185 | Not_Computer
Z-1103-2010 | 54358 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | NIPRO, GLUCOPRO INSULIN SYRINGES, SINGLE USE INSULIN SYRINGE with needle, 1CC 30g 5/16", U-100 INSULIN, LATEX FREE. Distributed by: NIPRO MEDICAL CORPORATION 3150 N.W. 107th Ave., Miami, Florida 33172. |  |  | January 13, 2010 | 2010 | Nipro Medical Corporation | Nipro Medical Corporation, Miami FL is initiating a nationwide recall of all GlucoPro Insulin syringes (This does not include the GlucoPro syringe specific for use with the amigo Insulin pump). These syringes may have needles that detach from the syringe. If the needle becomes detached from the syringe during use, it can become stuck in the insulin vial, push back into the syringe, or remain in | Nonconforming Material/Component | Each consignee will receive a Recall Notification Letter and response form by either fax or e-mail.Product should be returned to point of sale.Product not distributed should be collected and returned to Nipro Medical Warehouse in Memphis TN and Carolina Puerto Rico.Any questions or concerns can be addressed to the company at 787.769.2522. |  | 2268600 | 2,268,600 pieces | US (including Puerto Rico) Foreign Distribution: Canada, Barbados, Honduras, Jamaica, Panama, Trinidad and Tobago, and Costa Rica. | Terminated | September 15, 2010 | 245 | 2268600 | Not_Computer
Z-1083-2010 | 54512 | No | CGZ | Chloride test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OPTIÂżÂż LION Cassettes, E-Plus, Model BP7507 (25 per box), OSME TECH critical care. |  |  | April 20, 2009 | 2009 | OPTI Medical Systems, Inc | Outer box labeling shows an incorrect storage temperature range. | Process control | OPTI Medical Systems, Inc. issued a "Medical Device Correction" letter dated April 20, 2009. Consignees were instructed on how to corrected the affected product and were asked to return a completed Correction Response Card.For further information, contact OPTI Medical Systems, Inc. at +01 770 510 4444 (outside US) or 1-800-490-6784 (US). |  | 3175 | 127 boxes of 25 cassettes each for a total of 3175 cassettes. | Worldwide Distribution -- United States (CA, KY, MA and OK), Djouti (Africa), Egypt, Guatemala, India, Iraq, Pakistan, United Arab Emirates and Vietnam. | Terminated | April 17, 2012 | 1093 | 3175 | Not_Computer
Z-1082-2010 | 54482 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | Electri-Cord AC Power Cords, IEC 60320 Coupler, Product Reorder Number 21-2145-01 only for use with Smiths Medical CADDÂżÂż-Solis ambulatory infusion pumps, Model 2100. These cords were sold as an accessory for these pumps under Manufacturing Number A23.164.098.0A, Electri-Cord part Number PC2000403. |  |  | January 27, 2010 | 2010 | Smiths Medical ASD, Inc. | Potential for the power cord's prongs to crack and fail at/or inside the plug. Other medical device manufacturers have reported incidents of sparking, charring, and fires from the affected power cords used with their devices.Smiths Medical has received no reports of incidents regarding the use of the affected Electri-Cord powercords with their devices. | Nonconforming Material/Component | Consignees were sent a Smiths Medical, Inc. "Urgent Medical Device Recall" letter dated January 27, 2010. The letter was addressed to Risk/Safety Managers, Clinicians, Critical Care Physicians, Pediatricians, Pediatric Intensivists, Nenatologists, and Nursing, Pharmacy, Clinical/Biomedical Engineering, and Anesthesia Professionals. The letter described the problem, products involved and provided "Advice on Action to be Taken by the User". Consignees were requested to complete and return the "Urgent Recall Notice Confirmation Form". To return Power Cords, for questions regarding the recall notice or to report any issues with the affected product, please contact the Smiths Medical Technical Service Department at 1-800-558-2345. |  | 62 | 62 | United States (KY and OH) and Canada. | Terminated | December 10, 2011 | 682 | 62 | Battery | Physical Safety Hazards | Return | Remove or Replace
Z-1070-2010 | 54251 | No | JOX | Automated heparin analyzer. | Hematology | Hematology | 510(k) | Medtronic HMS PLUS, Hemostasis Management System, Model 30514. |  |  | December 14, 2009 | 2009 | Medtronic Inc | Recently the US Food and Drug Administration (FDA) issued an alert, communicating a change in the United States Pharmacopeia (USP) monograph for heparin, introducing a new USP reference standard that will in effect, change potency, and harmonize the USP unit dose with the WHO International Standard (IU) unit dose. Users of the Medtronic Hemostasis Management System (HMS Plus) must be aware o | Other | Consignees were sent a Medtronic "Urgent Medical Device Notice' dated December 14, 2009. The letter was addressed to "Valued Customer". The letter described the product and the problem. Advised customers to configure the HMS Plus appropriately.The Field Notification was sent to all current HMS Plus users. The current HMS Plus users were identified by the sales of HMS Plus disposable product -Heparin Assay Cartridges from the period of August 1, 2008 to date. This time was considered appropriate, being more than twice the longest expiration date of any of the HMS disposable. |  | 580 | 580 | All states in USA (Including DC) except RI and VT. | Terminated | December 24, 2011 | 740 | 580 | Not_Computer
Z-1101-2010 | 54369 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | AGA Medical Corporation, Amplatzer TorqVue Delivery System 45-degree. REF: 9-1TV10F45/80. Designed to facilitate the introduction of transvenous devices to chambers and coronary vasculature of the heart and for introducing therapeutic devices to a location within the peripheral vasculature. |  |  | December 18, 2009 | 2009 | AGA Medical Corporation | In this lot of TorqVue Delivery Systems (PN: 9-ITV10F45/80, LN: 0906151781), the dilator cannot be fully secured to the sheath as directed by the instructions for use. | Process design | An AGA Medical Product Recall Notice, dated December 18, 2009 was sent to customers on 12/18/2009. The letter identified the affected product, described the issue and asked for devices to be returned. Customers are asked to complete a form attached to the product recall notice and return it to AGA via fax or email. |  | 57 | 57 | Worldwide distribution: USA, Colombia, Greece, Italy, Japan, Mexico, Slovakia, and Taiwan. | Terminated | March 29, 2012 | 832 | 57 | Not_Computer
Z-1097-2010 | 54250 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Smart Perfusion Pack, REF: 088501900, Sterile EO, Sorin Group USA, Inc., 14401 W. 65th Way, Arvada, CO 80004 |  |  | December 30, 2009 | 2009 | Sorin Group USA, Inc. | Tubing set packaging seal between lid and tray may not be adequately sealed, compromosing sterility. | Packaging process control | Customers were notified by letter on December 30, 2009 and asked to examine all inventory for seal voids and arrange for return if any were found defective. Examination instructions were included. The two distributors were given the same instructions and were told to share all the information with any of their customers. |  | 9 | 9 units | AZ, CA, FL, MO, NJ, SC, SD, TN, TX, UT, VA and WI. | Terminated | March 15, 2010 | 75 | 280 | Not_Computer
Z-1111-2010 | 54267 | No | JWX | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Wampole Staph Latex Test, 6.0 ml -120 Test Kit |  |  | January 15, 2010 | 2010 | NERL Diagnostics Corporation | Unreactive Staph latex reagent | Process control | NERL Diagnostics notified distributors on January 15, 2010, by letter to advise of the recall and request they examine inventory, discontinue use of affected lot numbers and discard any remaining inventory. In addition, if product was further distributed, notify the customers of the product recall. If you have any questions, please telephone 1-401-824-2046. |  | 146 | 146 kits | NJ | Terminated | October 05, 2011 | 628 | 271 | Not_Computer
Z-1133-2010 | 53674 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Martel Printer |  |  | November 06, 2009 | 2009 | Abbott Point of Care Inc. | Some Martel Printers used with the i-STAT 1 system have the potential of being damaged due to the use or rechargeable battery packs that do not contain a fuse. | Component design/selection | An "Urgent Recall Notice" dated October 2009 was sent to customers. The notice described the product, problem and action to be taken by the customers. The customers are requested to remove the printers from use, if they have a Martel printer with the specified serial numbers, that do not have a battery door for access to the rechargeable battery pack. The customers are ask to turn in to Abbott Point of Care or a designated repair facility, all affected Martel printers for battery pack replacement. Abbott will return the Martel printers to your facility at no charge. The customers are ask to please follow the instructions on the enclosed customer reply cared. If in doubt regarding identification of the affected Martel printers, contact Abbott Point of Care Technical Support. If you have any questions regarding this information or if removing the affected Martel printers from use will negatively impact patient care, please contact Abbott Point of Care Technical Support at 800-366-8020, option 1 or your Abbott Point of Care representative. |  | 6432 | 6432 printers (1208 have already been returned and repaired or scrapped) | Worldwide distribution | Terminated | September 14, 2011 | 677 | 6432 | Battery
Z-1132-2010 | 53759 | No | MRR | Lipoprotein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PTS Panels LDL Cholesterol test strips for professional use with CardioChek P-A test systems, Manufactured by Polymer Technology Systems, Inc., Indianapolis, IN; Catalog No. 1753. |  |  | May 29, 2008 | 2008 | Polymer Technology Systems, Inc. | The expiration date was extended by 20 weeks without justification. | Process change control | An "URGENT-FIELD CORRECTION" letter dated May 29, 2008, was sent to customers. The letter described the product, problem and action to be taken by the customers. The customers are to please immediately identify the number of remaining LDL kits they currently have remaining in inventory of the specified lots and indicate the amount on the included Customer Acknowledgement and Confirmation letter and return via fax to 217-870-5608. The firm will ship a new lot of the LDL Cholesterol strips with new expiration dates to the customers. The customers are to please save the updated information and make the appropriate changes to their records. If you have any questions, please contact Customer Service at 317-870-5610 or toll-free at 877-870-5610. |  |  | not available. | Worldwide distribution | Terminated | October 20, 2010 | 874 |  | Not_Computer
Z-1137-2010 | 54231 | No | MOV | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Modified Novy Cornual Cannulation Set with clear tip, G17558, REF J-NCS-504070, Cook OB/GYN. |  |  | January 08, 2010 | 2010 | Cook Women'S Health | The distal tip may separate from the catheter while the catheter is in the patient. | Process control | Cook Women's Health sent Urgent Recall Notice letters dated 1/8/2010. Consignees were instructed to return product to the manufacturer for credit. |  | 721 | 721 both products | USA, Ireland, China, Canada, India, Australia, Hong Kong. | Terminated | December 10, 2010 | 336 | 3596 | Not_Computer
Z-1135-2010 | 54282 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips Switched Internal Defibrillator Paddles (used with Philips HeartStart XL and HeartStart MRx monitor/defibrillators) used during open heart surgery. |  |  | January 19, 2010 | 2010 | Philips Healthcare Inc. | Paddle shock switch may fail to actuate and delay or prevent delivery of defibrillation therapy | Nonconforming Material/Component | Philips issued Urgent Medical Device Recall notification to customers by letter on 1/25/10 via UPS advising of this defect and providing instructions on how to determine if they have a defective paddle. The customers will be assisted by field service engineers to remove and replace these defective units (those within the date range of September 1, 2008 to December 30, 2008). Customers are asked to follow the Actions to be taken by customer/user section of the Urgent Medical Device Recall Notification.Removed units will be replaced free of charge and will be destroyed by Philips. Distribution of the Urgent Medical Device Recall Notification will be managed by Philips representatives in each affected geography.Customers are asked to follow the Actions to be taken by customer/user section of the Urgent Medical Device Recall Notification. |  | 1592 | 1592 units | Worldwide distribution: USA, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Dominican Republic, Egypt, Estonia, Finland, France, Georgia, Germany, HONG KONG, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Republic of Korea, Lebanon, Lithuania, Malaysia, Mauritius, Netherlands, NICARAGUA, Norway, Peru, Philippines, POLAND, Portugal, Puerto Rico, Russian Federation, Saudi Arabia, Serbia &amp; Montenegro, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad/Tobago, Turkey, UKRAINE, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, andVietnam. | Open, Classified |  |  | 1592 | I/O | Treatment/Delivery/Therapy | Replace device | Remove or Replace
Z-1136-2010 | 53773 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PTS Panels Lipids Panel test strips for self-testing or professional use with CardioChek P-A analyzers, Manufactured by Polymer Technology Systems, Inc. Catalog # 1710. |  |  | June 05, 2009 | 2009 | Polymer Technology Systems, Inc. | Exhibits under-recovery of all analytes. | Nonconforming Material/Component | Consignees were notified by letter dated 6/5/09, advised of the problem, and that the firm would replace product. |  |  | Not available. | Nationwide and international. | Terminated | October 20, 2010 | 502 |  | Not_Computer
Z-1139-2010 | 54281 | No | JSM | Transport culture medium. | Microbiology | Microbiology | 510(k) | remel A.C.T. II Sterile Pack Tube, Collection and Transport System, Ref 12402, 10/pk, each Tyvek pouch contains 1 tube of medium and 1 rayon-tipped sterile swab. The firm name on the label is remel, Lenexa, KS. |  |  | January 18, 2010 | 2010 | Remel, Inc | Pouch may be inadequately heat sealed compromising sterility | Packaging | An IMPORTANT MEDICAL DEVICE PRODUCT RECALL NOTICE" dated January 18, 2010, was sent to customers via regular mail. The notice described the product, problem, and action to be taken by the customers. The customers are requested to: review their inventory for the indicated product, discard all remaining units of the product lot, complete the "Product Inventory Checklist" enclosed to report the amount of product they wanted replaced or credited, and return by February 1, 2010 regardless if they have affected product and/or have verbally contacted the firm.If you have any inquiries concerning this advisory, please contact our Technical Services Department at 800-255-6730 (USA) or 913-888-0930 (International). |  | 154 | 154/10-unit packages | Nationwide distribution: MI, MA, WV, CO, ME, FL, NJ, MD, WA, AZ, VA, CA, KS, IL, MO, AL, TN | Terminated | April 25, 2012 | 828 | 154 | Not_Computer
Z-1131-2010 | 54191 | No | OBO | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(k) | Optovue RTVue Optical Coherence Tomography (OCT), Model-RT100 with software versions 3.0.x.x and higher, with Normative Database, manufactured by Optovue, Fremont, CA |  |  | January 08, 2010 | 2010 | Optovue Inc. | Marketed without proper 510(k)-- No 510(k) clearance for Optovue RTVue Model-RT100 with Software Versions 3.0.x.x and higher which includes the Normative Database. | Other | The firm is preparing a software "Roll Back" version of software to disable the NDB and other features, ensuring that this version will work properly. A customer notification letter is provided as well as a correction verification form, installation instructions and user manual.If there are any questions, please call 1-866-344-8948. |  | 1800 | 1800 units | Nationwide and Worldwide | Terminated | March 24, 2011 | 440 | 1800 | Software |  | Software rollback | Software Update
Z-1140-2010 | 54623 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Centurion Healthcare Products Pump Fill Kit DT10175. CONTENTS: 1 CSR WRAP, 2 GLOVES, 1 FENESTRATED DRAPE, 1 EXTENSION LINE Approx. Priming Vol.: 0.3mL, 1 SYRINGE, 1 Exel Huber NEEDLE, 1 ChloraPrep(R), 1 TRAY |  |  | February 01, 2010 | 2010 | Centurion Medical Products | A supplier notified this firm of a recall of their Huber needles because the needles that are labeled as non-coring could core 60-72% of the time. Centurion is doing a sub recall of their kits that contain the referenced needles. | Other | An "URGENT MEDICAL DEVICE WITHDRAWAL" notice dated January 26, 2010, was sent to customers via certified mail. The notice described the product, problem and action to be taken by the customers. The customers were instructed: DO NOT USE THESE AFFECTED PRODUCTS. PLEASE RETURN ALL AFFECTED PRODUCTS IN ACCORDANCE WITH THE INSTRUCTIONS PROVIDED. The customers were also instructed to report any adverse reactions experienced with the use of the product and/or quality problems to MedWatch Program by phone at 1-800-FDA-1088, by fax at 1-800-FDA-1078 by mail at MedWatch, HF-2, FDA, 5600 Fishers Lane, Rockville, MD 20852-9787 or on the MedWatch website at www.fda.gov/medwatch. The customers were also instructed to check the lot numbers of products in inventory to determine if they have products affected by this withdrawal and complete and return the enclosed Customer Response Form by fax or email.Direct all inquiries to: EXEL International Customer Service 727-827-1922 ext. 0 Monday through Friday 9am-5pm EST or E-mail info@exelint.com |  | 880 | 880 kits | USA only | Terminated | November 08, 2010 | 280 | 880 | Not_Computer
Z-1108-2010 | 54622 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ACM Bio-Prep with Merocel pkg/6, Rx, Sterile, REF 0306-708-000S4, Stryker Instruments, Kalamazoo, MI. |  |  | January 25, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | The firm did not have packaging validation to support the labeled shelf life. There is potential for a breach in sterility specific to the sterile packaged femoral sponge which may result in a non-sterile product. | Other | A "Medical Device Recall Notification" dated January 21, 2010 was issued to customers. The notification described the affected product, issue and actions to take. Consignees were asked to forward the notification to all their affected locations. Customers were instructed to immediately check inventory and quarantine any recalled product found. Consignees were instructed to destroy all product and report how much product was destroyed. For questions regarding this recall, please contact Stryker Instruments by calling 1-800-800-4236, ext. 4354 |  | 9274 | 9,274 all products | Worldwide Distribution -- United States, Canada, Mexico, Australia, Brazil, Switzerland, Japan, Netherlands, England, UK, Sweden, Hong Kong, Korea, Malaysia, Singapore, Argentina, Spain, and Italy. | Terminated | October 25, 2010 | 273 | 9274 | Not_Computer
Z-1154-2010 | 48976 | No | JEC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stableloc (TM) Wrist External Fixation System, Quantity 1. Packaged labeled, "Stableloc Wrist Fixator System... Desc:stabeloc (TM) External Fixator Kit... Ref FX-4001-S... Manufactured or Distributed by: Acumed (R) 5885 NW Cornelius Pass Road, Hillsboro, OR 97124...STERILE" |  |  | July 09, 2008 | 2008 | Acumed LLC | Parts are non-functional. The user is unable to fully pass the fixator pins through the body of the Stableloc assembly as required. The kit may also be missing the set screw. | Process control | The firm sent "Urgent Field Safety Notice: Medical Device Recall" letters to customers 7/9-11/08. In addition, the firm sent the letter via e-mail to some of the customers. The notification advised the customer of the recall, identified the product, explained the reason, and asked the customer to locate and segregate any affected product. The customer was asked to notify their customers who they had distributed the affected product to and asked them to complete and return an enclosed response form. The firm provided a model recall letter for its agents/distributors to issue to users. Upon receipt of the response form, the firm will contact customers to arrange for replacement of any affected product. Questions should be directed to the "Distributor phone number" or Richard Snyder, Quality Manager at 1-888-627-9957. |  | 80 | 80 kits | Worldwide Distribution -- USA, including states of VA, NC, FL, ME, WA, GA, TX and PA and countries of Spain, Japan and Greece. | Terminated | May 21, 2010 | 681 | 80 | Not_Computer
Z-1141-2010 | 50102 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Howmedica Osteonics Corp., Scorpio NRG Femoral Size 7 Right; Knee prosthesis (Stryker France). Product Code: 814407R. |  |  | August 20, 2007 | 2007 | Stryker Howmedica Osteonics Corp. | A box labeled as a Scorpio NRG Femoral Size 7 Right, may actually contain a Scorpio PS Femoral Size 5 Right. | Process design | Stryker Important Product Removal Letters were sent via Federal Express on 8/20/2007. |  | 4 | 4 units | Product went to 2 US Stryker branches. | Terminated | March 16, 2010 | 939 | 4 | Not_Computer
Z-1109-2010 | 52618 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PrecisePLAN 2.15 Treatment Planning System. |  |  | September 30, 2008 | 2008 | Elekta, Inc. | The new "DICOM CT with Body Outline" method to create a patient case in PrecisePLAN 2.15 may produce incorrectresults for some customers, given several input conditions encountered in the field. Discontinue use of this method ofcreating a patient case until FCO 772 00 540 064, (entitled "Update for Patient creation method using DICOM CT with Body Outline"), is received from Elekta to update th | Other | Important Notice A297, "PrecisePLAN release 2.15 with DICOM CT with Body Outline" method of patient creation (dated May 8, 2007) was distributed to all customers. An additional Field Change Order 540 073, dated September 30, 2008 was distributed to customers. Important Notice A316 "Cancellation of Important Notice for PrecisePLAN" was issued. The Important Notice instructed affected users to dispose of Important Notice A297 and replace it with Important Notice A316.If you have any questions regarding the notification, please contact your local Elekta representative directly. |  | 280 | 280 units | United States -- (AL, AK, AZ, AR, CA, FL, GA, HI, IL, IN, IA, KY, LA, MD, MA, MI, MS, MO, NE, NV, NJ, NM, NY, , NC, OH, OK,OR, PA, SC, TN, TX, VA, WA and WI) &amp; France. | Terminated | December 28, 2010 | 819 | 280 | Not_Computer
Z-1146-2010 | 53597 | No | MVC |  |  | Pathology | PMA | VENTANA PATHWAY anti-HER-2/neu (4B5) RABBIT MONOCLONAL PRIMARY ANTIBODY |  |  | September 17, 2009 | 2009 | Ventana Medical Systems Inc | Incorrect result: Light staining inconsistencies with lot number 680957 of the Ventana PATHWAY HER-2/neu (4B5) Rabbit Monoclonal Primary Antibody. | Nonconforming Material/Component | The recall was initiated on 09/17/2009 with Ventana Medical Systems forwarding Urgent Voluntary Medical Device Recall letters to all the affected customers advising them to discontinue the use of lot number 680957 and have their laboratory retest with product from a different lot of PATHWAY HER2 (4B5) any breast cases scored as 0, 1+, or 2+ that were not reflexed to an approved in situ hybridization (ISH) method.Customers were also instructed to contact Ventana Customer Support at 1-800-227-2155 and select 1 for Technical Support to speak to a Customer Support agent who will work with then to replace or credit any product from this specific lot number. |  | 433 | 433 units | Worldwide distribution: USA, Austria, Denmark, Germany, Hungary, Netherlands, Sweden, United Kingdom, Brazil, Canada,India, Korea, Morocco &amp; Thailand | Terminated | April 22, 2010 | 217 | 433 | Not_Computer
Z-1148-2010 | 54190 | formulations, paraffin, all | KEO | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Paraffin Pretreatment Reagent Kit III; each kit contains the following: |  |  | January 22, 2010 | 2010 | Abbott Molecular | The Paraffin Pretreatment Reagent kits do not have any hazardous or MSDS information included in their labeling. | Package design/selection | Abbott Molecular sent Field Correction Letter/Urgent Field Safety Notice letters dated 1/21/10 to the Paraffin Pretreatment Regent Kit customers on 1/22/10, informing them that the MSDA symbology for hazardous and corrosive material is missing from the Paraffin Pretreatment Regent Kit II and Paraffin Pretreatment Regent Kit III labels. The information is missing from the vial label, kit label and package insert. The letter listed the risk and safety statements that should be included in the package insert, and advised the customers should use the kits in accordance with the listed risk and safety statements. The accounts were instructed to review this information with laboratory personnel and retain the letter for future reference, making copies of the letter and placing them in each kit in stock. Any questions were directed to their local Abbott Molecular Representative. Abbott Molecular plans to telephone each customer to ensure that they have received the letter and have disseminated the information to their laboratory personnel. |  | 17 | 17 kits | Worldwide distribution: USA, Argentina, Austria, Belarus, Belgium, Canada, Colombia, France, Germany, Greece, Hong Kong, Hungary, India, Italy, Japan, Kuwait, Lebanon, Lithuania, the Netherlands, Panama, Poland, Singapore, Slovenia, South Africa, Sweden, Switzerland, Taiwan, Turkey, United Kingdom and Vietnam. | Terminated | January 29, 2012 | 737 | 322 | Not_Computer
Z-1155-2010 | 53848 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Piccolo Comprehensive Metabolic Panel Reagent Disc, 10-pack box, for use with the Piccolo Blood Chemistry Analyzer or Piccoloxpress Chemistry Analyzer, Catalog No. 400-0028, manufactured by Abaxis Inc., Union City, CA |  |  | November 16, 2009 | 2009 | Abaxis Inc | Wrong barcode applied to a lot of items, which may result in incorrect calibration factors and results. | Labeling mix-ups | An "URGENT: PRODUCT RECALL NOTICE" dated November 16, 2009, was sent to customers. The notification described the product, issue/impact, and the action/resolution to be taken by the customers. The customers are to discontinue use of the specified lot, contact Abaxis for immediate replacement and complete and return the enclosed Response form within 5 working days. They may fax the form to Abaxis Technical Support at 1-510-315-3140. The customer are requested to ensure all laboratory staff are informed about this notice and retain this notification as a part of their laboratory Quality System documentation. The firm is currently contacting all direct shipment customers.If you have any questions regarding this Product Recall Notice, please contact Abaxis Medical Technical Support at 800-822-2947, extension 6505, or at medtechsupport@abaxis.com. |  | 1010 | 1010 boxes | Worldwide distribution: USA and Canada, EU and Asia | Terminated | December 14, 2010 | 393 | 1010 | Not_Computer
Z-1150-2010 | 54490 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Bond-1 Primer/Adhesive. |  |  | August 24, 2009 | 2009 | Sybron Dental Specialties | Some of the material has been found to be gelled and therefore becomes unusable. | Nonconforming Material/Component | An "Urgent Medical Device Recall" letter and "Recall Return Form" were sent by first class mail on August 24, 2009 to consignees. The letter acknowledged the "gelling" issue of identified lots shipped between July 2009 &amp; August 2009. Customers were requested to return any of the remaining affected lots in their inventory. Contact Customer Service by calling 1-800-551-0283, (option 1) directly to handle the arrangements of a quick return and replacement. |  | 1400 | 1,400 | Worldwide Distribution -- United States, Canada, Netherlands, Sweden, Taiwan and Thailand. | Terminated | February 24, 2012 | 914 | 1400 | Not_Computer
Z-1157-2010 | 54537 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Auto Suture ProTack Fixation 5 mm |  |  | February 08, 2010 | 2010 | Covidien LP | Fixation device may fail to fire and staple. | Process control | Covidien notified accounts by "Urgent Medical Device Recall" letters dated 2/8/10. The letter identified the affected product and the reason for recall. Customers are to examine their inventory, and remove and return recalled lots. Customers are to follow the instructions provided in the letter to return the affected product. Questions should be directed to a Covidien Representative. On June 23, 2010, the firm issued a letter to expand the recall to include 5 additional lots. |  | 5638 | 5,638 units | Nationwide Distribution | Terminated | February 28, 2013 | 1116 | 5638 | Not_Computer
Z-1153-2010 | 53477 | Yes | NIM |  |  | Cardiovascular | PMA | Xact 10 mm x 20 mm x 136 cm (Part Number 82099-01) |  |  | March 19, 2009 | 2009 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Abbott Vascular is initiating a recall on the XACT Carotid Stent System because the product was distributed prior to approval of a PMA supplement for a manufacturing line move between sites. There are no product quality issues associated with this action. | PMA | Notification was sent out to all consignees about the product recall. Representatives from Abbott Vascular personally visited consignees and removed product from site. |  | 49 | 49 | Nationwide: GA, NY, IN, LA, MA and PA | Terminated | March 30, 2012 | 1107 | 49 | Not_Computer
Z-1142-2010 | 54363 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Neutralect Diathermy Cable, |  |  | January 22, 2010 | 2010 | ConvaTec | Mislabeling: Boxes labeled with the order code 3508M may contain a cable for the order code 3405M, and boxes labeled with order code 3405M may contain a cable for order code 3508M. | Other | Unomedical Urgent Medical Device Correction letters were sent to distributors on January 22, 2010 by certified mail. Letters state that the only way to determine if the incorrect product is contained in the carton is to open the box. Distributors are asked to inspect all Diathermy Cables in their possession with order code 3508M or 3405M to ensure that the correct cable is inside. The number of mislabeled product is to be recorded on the Questionnaire for Distributors, and then mislabeled product is to be destroyed by cutting the product in half. The questionnaires include an area for names of customers stating that the distributor will contact them, or requesting that Unomedical contact them directly. Copies of the correction notice and questionnaire for end users are to be forwarded to all customers, asking them to complete the questionnaire and return it to the distributors. Once received by the distributors, end users can discard the cables by cutting them in half. All completed forms should be returned to Unomedical and local reps are to be contacted to arrange for replacement of mislabeled stock and/or credits. |  | 6400 | 5130 cables worldwide, 1270 in US. | Worldwide distribution. | Terminated | June 10, 2010 | 139 | 6400 | Not_Computer
Z-1144-2010 | 54507 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Ascent Posterior Occipital Cervical Thoracic System, Single Looped Cable with Two Crimps, One Each. Distributed by Blackstone Medical, Springfield, MA. |  |  | February 02, 2010 | 2010 | Blackstone Medical, Inc. | The Ascent Posterior Occipital Cervical Thoracic System Single Looped Cable with two crimps may be made of stainless steel instead of Titanium. | Labeling mix-ups | Orthofix Spinal Implants contacted customers by phone beginning December 22, 2009 and issued a Notification of Product Removal dated February 2, 2010. Consignees were instructed to Identify, Remove and Return all affected product to the firm.For further information, contact Orthofix Spinal Implants, Regulatory Affairs Department at 1-973-406-2812. |  | 25 | 25 | Worldwide Distribution -- United States (Texas and California), Australia, Colombia and Greece. | Terminated | April 09, 2010 | 66 | 25 | Not_Computer
Z-1149-2010 | 54374 | No | LIA | Shigella spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BD Shigella poly Group D Antiserum S. sonnei; types I,II, packaged in 3 mL vials, Catalog #24103, and labeled in part. Becton, Dickinson and Company, 7 Loveton Circle, Sparks, MD 21152 USA. |  |  | January 25, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic test reagent may exhibit decreased reactivity. | Nonconforming Material/Component | BD Diagnostic Systems issued an "Urgent Product Recall" notification dated January 2010. Consignees were advised to discontinue use of the affected product and discard for replacement or credit.For further information, contact BD Diagnostic Systems, Customer Service at 1-800-675-0908. |  | 1096 | 1096 units | Nationwide Distribution | Terminated | May 21, 2010 | 116 | 1096 | Not_Computer
Z-1143-2010 | 54278 | No | GZL | Depth electrode. | Neurology | Neurology | 510(k) | FHC microTargeting Guideline 4000 ver 1.4 (GL4K) Software, |  |  | December 28, 2009 | 2009 | FHC, Inc. | The raster tab for the GL4K version 1.4 software may assign incorrect depths for waveform recordings and if used to identify the implantation target for the DBS lead, the implant may be placed in the wrong location. | Software change control | FHC notified customers by telephone and by an Urgent: Medical Device Correction email, dated December 2009, beginning December 28th. Customers have been instructed to keep the instruction with the GL4K until FHC can provide a revision to the software; and, to use only "distance from zero" or "distance traveled" modes (not "distance from target" mode) during procedures using the raster tab. Customers are to notify their FHC representative to confirm receipt of the notification. A software patch has been verified and validated and will be installed on all affected units. The new version of the corrected software is Guideline 4000 version 1.41. Questions should be directed to Jason Moyer, product manager or Keri Seitz, recall coordinator at 207-666-8190. |  | 10 | 10 | Worldwide Distribution -- USA, including states of GA, MA, NC, NJ, OH, RI, TN, TX, and WV and Bucharest, Romania. | Terminated | March 18, 2010 | 80 | 10 | Software | Output/Calculation | Software update | Software Update
Z-1145-2010 | 52889 | No | DHC | Systemic lupus erythematosus immunological test system. | Immunology | Immunology | 510(k) | BIO-RAD 425-2020 Autoimmune EIA Anti-Cardiolipin IgM, 96 Tests...2010 02 16...Lot CL-863. Positive Control included in lot is identified as "REF 425-2025" and "LOT CPM-591". in vitro diagnostic. |  |  | July 14, 2009 | 2009 | Bio-Rad Laboratories Inc | The firm received customer complaints of microbial contamination in Positive Controls in its Anti-Cardiolipin IgM Test Kit Lot CL-863. | Nonconforming Material/Component | Bio-Rad initiated its field correction on July 14, 2009. The firm notified domestic consignees by phone followed by fax, and each subsidiary will be given an urgent notice, subsidiary response form, customer medical device correction notification, acknowledgement form and a listing of consignees. The subsidiaries must contact customers as instructed. |  | 213 | 213 kits | Worldwide distribution: USA, Australia, Canada, France, Italy, Korea, Mexico, Panama, Singapore, and Trinidad. | Terminated | March 18, 2011 | 612 | 213 | Not_Computer
Z-1156-2010 | 53463 | No | LPB |  |  | Cardiovascular | PMA | CircuCool Fluid Circulation Pump used in the Chilli Cooled Ablation System, Model Number M00480050, manufactured by Barnant ThermoFisher Scientific, Barrington, IL. Distributed by Boston Scientific Corporation, San Jose, CA. |  |  | September 22, 2009 | 2009 | Boston Scientific Corp | Product failure-- Product may cease operation as a result of potential speed control board failure when the device is shut off, and may not turn back on. A component on the circuit board fails unexpectedly because of a power surge, causing a diode to burn out. | Nonconforming Material/Component | The firm has issued notification letters via Federal express with reply verification tracking forms. A follow up letter and telephone survey as needed will be sent to follow up with non-responders. The firm is asking its consignees to return any recalled product.If there are any questions, please call 1.408.935.4952. |  | 75 | 75 units | US and to Spain, France, Germany, Italy, Russia, Poland, the Netherlands, Israel and Belgium | Terminated | December 27, 2010 | 461 | 75 | Hardware | Output/Calculation | Return | Remove or Replace
Z-1151-2010 | 54372 | No | JEY | Bone plate. | Dental | Dental | 510(k) | BIOMET MICROFIXATION, "1.5MM" SYSTEM BONE SCREW LP, SELF DRILL, X-DRIVE, SGL.LEAD AUGER POINT, 1.5 X 3.5MM, TITANIUM, NON STERILE PRODUCT. Model 01-7293 and 99-7293-01. BIOMET MICROFIXATION, 1520 TRADEPORT DRIVE, JACKSONVILLE, FL 32218 USA. |  |  | October 06, 2009 | 2009 | Biomet Microfixation, Inc. | Biomet Microfixation is recalling Model 01-7293 an Model 99-7293-01, 1.5X3.5mm Self Drilling, Low Profile, Cross Drive screw, 5/PK was found to contain a 91-6104 (1.5X4mm High Torque, Self Drilling, Cross Drive screw. | Process control | Biomet Microfixation Issued an "Urgent Medical Device Recall Notice". Consignees were notified of the recall instructed to indentify and return all affected product to the firm. An Inventory Reconciliation sheet should be completed and faxed to 1-904-741-9425. For further information, contact Biomet Microfixation, Global Complaint Coordinator at 1-800-874-7711 or 904-741-4400, extension 468, Monday through Friday, 8am-4pm ET. |  | 831 | 831 screws | Worldwide Distribution -- United States (WI), Spain, Italy, Japan, Brazil and Korea. | Terminated | April 07, 2010 | 183 | 831 | Not_Computer
Z-1158-2010 | 54554 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Perfexion, Article #715000. |  |  | July 02, 2007 | 2007 | Elekta, Inc. | Need to modify the closing speed of the shielding doors in the event of an emergency exit | Other | A field action (FCO TN-07-0007) has been performed to modify the shielding door speed on all sites in clinical use. A second Field Change Order 200 069 "Replacement of Door Servo Board" dated 9/16/08 was released. |  | 19 | 19 | Worldwide Distribution -- USA, including states of AR, AZ, CA, FL, IL, MI, MN, MS, NJ, OH, OR, PA, UT, VA and countries of France and United Kingdom. | Terminated | March 25, 2010 | 997 | 19 | Not_Computer
Z-1163-2010 | 54634 | [6 empty] | Olympus Trocar Spikes, Model numbers: A5821, A5948, A5949, WA58341L, and WA58343L |  |  | February 08, 2010 | 2010 | Olympus America Inc. | weak weld seams | Nonconforming Material/Component | The recalling firm issued a Medical Device Recall letter dated February 8, 2010, informing their customers of the problem and the need to return the product. |  | 45 | 45 | Nationwide: AL, AR, CA, FL, MI, MN, MO, NY, PA, OH, TX, and WY. Foreign: Canada | Terminated | May 25, 2010 | 106 | 45 | Not_Computer
Z-1159-2010 | 54559 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta SL Series/Precise Digital Accelerator. |  |  | October 01, 2005 | 2005 | Elekta, Inc. | Customer reported larger than expected variations in wedge factory with gantry rotation. | Device Design | Important Notice A280 "Wedge Matching Procedure" (dated 11/18/05) was sent to all effected customers. The Important Notice described the issue, warnings and cautions and corrective instructions/Maintenance to take to eliminate occurrence. Direct questions to your local Elekta Limited representative. |  | 315 | 315 units | United States, Canada and Mexico. | Terminated | March 25, 2010 | 1636 | 315 | Not_Computer
Z-1160-2010 | 54619 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ATEC 0909-12 Breast Biopsy and Excision System , 9 gauge, 9 centimeter handpiece (part number 0909-12), Rx, Sterile, Suros Surgical Systems, Inc. 6100 Technology Center DR. Indianapolis, IN 46278. |  |  | February 01, 2010 | 2010 | Hologic, Inc. | The product is labeled as ATEC0909-12 but it actually contains 0909-20. | Employee error | Formal recall notification letters and response forms were sent to customers via FEDEX on February 1, 2010. If there are any questions, please call 317.344.7630. |  | 190 | 190 | US and Puerto Rico. | Terminated | October 08, 2010 | 249 | 190 | Not_Computer
Z-1162-2010 | 52299 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | SilverHawkÂż Peripheral Cutter Driver Plaque Excision System |  |  | January 21, 2010 | 2010 | Ev3, Inc. | ev3 Inc. is conducting a voluntary recall of specific lots of the SilverHawkÂż Cutter Drivers (Model Number FG 02550) because of damage to the packaging. | Packaging change control | Consignees were sent an ev3 "Medical Device Recall" letter dated January 21, 2010. The letter was addressed to Cath Lab Manger/Risk Manager. The letter described the Issue Summary, Affected Product, and Required Action. The letter instructed customers to locate and remove from use the specified Cutter Driver(s). Customers are to complete and return the "Device Recall Field Action Reconciliation Form." An ev3 rep will assist in arranging for the return of all unused product and a replacement will be provided. Questions regarding the recall should be directed to ev3 Customer Service at 1-800-716-6700. |  | 8330 | 8,330 | Worldwide Distribution -- USA, with the exception of AK and VT, and countries of AUSTRIA, BELGIUM, BRAZIL, CANADA, CHINA, CZECH REPUBLIC, FRANCE, GERMANY, INDIA, IRELAND, ISRAEL, ITALY, JORDAN, LEBANON, POLAND, SAUDI ARABIA, SOUTH KOREA, SPAIN, SWEDEN, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | April 26, 2012 | 826 | 8330 | Not_Computer
Z-1164-2010 | 52825 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Aquilion 32 CT System; Model number TSX-101 A/E |  |  | January 08, 2009 | 2009 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems Inc (TAMS) has initiated a field correction on the Aquilion 32 CT Scanners due to the potential for the bolus tracking software (SureStart) to cause delay of the helical scan, up to 10 seconds, when large volumes of data are being reconstructed on the secondary console. | Component design/selection | Customer notification was sent and a software upgrade will be issued to all customers affected by the field correction.If you have questions, please call (800) 521-1968. |  | 65 | 65 | Nationwide | Terminated | June 01, 2012 | 1240 | 692 | Computer
Z-1171-2010 | 52832 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Aplio 50; SSA-700A; software version 5.5r002 |  |  | November 12, 2007 | 2007 | Toshiba American Medical Systems Inc | Toshiba America Medical System Inc initiated a field corrective action on Aplio 50; SSA-700A; software version 5.5r002 and later, Aplio 80; SSA-700A; software version 5.5r002 and later, and Xario; SSA-660A; software version 1.0 and later, because the product does not meet the Safety Standard requirements. | Component design/selection | Toshiba America Medical Systems (TAMS) has issued a Field Modification Instruction notification to all of its customers about the affected devices. The correct software will be administered through TAMS' representatives. Reimbursement for the modification will be given upon receipt of the reimbursement claim invoice. |  | 8 | 8 units | Nationwide Distribution | Terminated | April 18, 2012 | 1619 | 200 | Computer
Z-1174-2010 | 54703 | No | MEQ |  |  | Gastroenterology/Urology | PMA | UROLOGIXÂż CTC Advance" Standard Microwave Catheter (3.0-5.0 cm), CTC Advance" Long Microwave Catheter (4.5+ cm), CTC Advance" Short Microwave Catheter (2.5 -3.5 cm), RTU Plus Balloon, and RTU Plus Handle, Sterile EO, Manufactured by UROLOGIXÂż, INC, 14405 Twenty-First Avenue North, Minneapolis, MN 55447. |  |  | February 18, 2010 | 2010 | Urologix, Inc. | UrolgixÂż is recalling certain lots of CTC Advance" catheters and RTU Plus devices which are integral components of the TargisÂż System. The reason for this Recall is a labeling error only; the issue involves a discrepancy in the"Use Before" date printed on the product label. In the affected lots, the "Use Before" date , commonly known as the expiration date, is printed as 2012-12 when it should r | Incorrect or no expiration date | Consignees were sent a UROLOGIX " Important Recall Information" letter dated February 18, 2010. The letter was addressed to the client's name. The letter described the problem and product involved. It provided Instructions for "Returning and Replacing Affected Product" and requested the return of the "Customer Acknowledgement Form".If you have questions, please call (763)-475-1400. |  | 154 | 154 | AL, AZ, CA, FL, IN, MI, MN, MS, NY, OK PA, TN, TX, and WI. | Terminated | March 08, 2012 | 749 | 154 | Not_Computer
Z-1177-2010 | 52816 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba America Medical Systems Inc (TAMS) Digital Radiography System Radrex-i DRAD-3000E; Electrostatic x-ray imaging system |  |  | May 22, 2009 | 2009 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems Inc (TAMS) has initiated a field correction on the DRAD-3000E; Radrex-I, Electrostatic x-ray imaging system due to a software bug that results in a loss of image. | Component design/selection | An Urgent - Medical Device Correction letter was sent to customers. The letter described the issue, the corrective action, and the avoidance. Toshiba America Medical Systems informed customers that software was being prepared to address the issue and planned to send representatives to each customer affected by the field correction to provide a software upgrade to the DRAD-3000E; Radrex-I, Electrostatic x-ray imaging system. Customers were to complete the attached customer reply form. Questions were to be directed to Aaron Ybarra, X-ray Product Manager at 800-421-1968, ext. 2492. |  | 20 | 20 | Nationwide Distribution -- Including states of PA, VA, OH, TN, LA, IL, CA, MT, AZ, and TX. | Terminated | March 21, 2012 | 1034 | 20 | Software | Display/Image | Software update | Software Update
Z-1178-2010 | 53474 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Peripheral Dilatation Catheter, Viatrac 14 Plus 5.0mm x 15mm x 80cm, Part Number 1008194-15. Manufactured by Abbott Vascular, Temecula, CA. |  |  | May 06, 2009 | 2009 | Abbott Vascular-Cardiac Therapies dba Guidant Corporation | Abbott Vascular is initiating a recall on the Viatrac 14 Plus 5.0x15 mm Peripheral Dilation Catheter due to a mistake in labeling/packaging. | Labeling mix-ups | Abbott Vascular issued an "Urgent Device Recall" notification dated May 7, 2009. Consignees were informed of the affected product and instructed to work with local sales representatives in identifying any unused devices and arrange for return to the firm.For further information, contact Abbott Vascular Customer Service at 1-800-227-9902. |  | 26 | 26 | Nationwide Distribution -- IL, FL, CA, WA, AL, OH, NY, MA, TX and IN. | Terminated | March 30, 2012 | 1059 | 26 | Not_Computer
Z-1175-2010 | 54134 | gown, examination | FME | Examination gown. | General Hospital | General Hospital | 510(K) Exempt | Thermoflect Product Line including items for adults and pediatrics. |  |  | December 23, 2009 | 2009 | Encompass Group LLC | The products are not compatible for use in the Magnetic Resonance Imaging (MRI) environment. | Other | A press release distributed through specialized news media, and trade magazines, on/about 12/26/2009. The press release was also posted on the thermoflect.com website. Letters were sent via first class mail to customers, hospitals and doctors in radiology on/about 12/23/2009. In addition to the written notification to the user institutions, Encompass provided adhesive labels sent along with the letter notification. The notice instructed that the labels should be affixed to remaining inventories of the products(s), specifically to the individual package, which holds 5 blankets and to the outermost packaging or carton. The label reads Not for use in MRI, and a symbol showing a circle with a diagonal line through the letters MR. Thermoflect product literature has been quarantine at this time. Instructions were provided showing where the labels where the labels were to be affixed. Encompass also provided a poster indicating the products are for use in the MR environment which is to be displayed in the MR area. All new product literature will indicate that the product is not for ruse in the MR environment. Thermoflect sales representatives will visit with each user institution to confirm their receipt of the recall notice and to be available to answer any question concerning the recall. The sales representatives all of whom have been briefed on the recall strategy will be available to provide training for Radiation Safety Officers and radiology staff, if requested. Following the recall notice, Encompass will conduct Effectiveness Checks to assure that all purchasers/users of the Thermoflect product have been notified and have taken appropriate action to notify their radiology centers and staff. The Encompass recall team will maintain records of the contacts for review by the FDA. A Status Report on the effectiveness of the recall effort will be filed with the Atlanta District Office on a monthly basis until the recall is terminated. |  | 700000 | 700,000 units (of various product lines) | AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY and Switzerland. | Terminated | October 27, 2010 | 308 | 700000 | Not_Computer
Z-118-2010 | 54538 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | AOS Trochanteric Nail, 13mm x 20cm x 130Âż, Product Part Number: 1037-200. |  |  | January 14, 2010 | 2010 | Advanced Orthopaedic Solutions Inc | The recall was initiated after Advanced Orthopaedic Solutions (AOS) became aware of a manufacturing defect involving the short Trochanteric Nail, where one piece was found to be incorrectly bent. If this issue were to recur, surgery time could be extended, but only slightly as nail exchanges take place commonly in a surgery of this type. | Process control | The recall was initiated with the firm forwarding a recall notice with attached customer response form on January 14, 2010 to the affected consignees. The recall notice informed the consignees of the reason for recall and gave them three steps to follow to inspect the affected Short Trochanteric Nail products. Consignees were provided with product return instruction for any incorrectly bent products found during their inspection. Customers with questions are instructed to contact the firm at (310) 533-9966. |  | 65 | 65 | Nationwide and Japan | Terminated | April 26, 2012 | 833 | 1981 | Not_Computer
Z-1176-2010 | 54633 | No | NFF | Medical image storage device. | Radiology | Ophthalmic | 510(K) Exempt | Image Consultant (IC) Software |  |  | February 05, 2010 | 2010 | R.O. Gulden &amp; Co., Inc. | Software not validated; inaccurate results may occur. | Other | On 2/5/10, the recalling firm issued letters via certified mail to its consignees informing them of the problem and to cease use of the software for diagnosis. |  | 20 | 20 | The product was shipped to eye doctors/clinics in AL, FL, LA, MA, NC, OK, OR, PA, TX, WI, and WA. | Terminated | June 14, 2010 | 129 | 20 | Software | Output/Calculation | Remove | Remove or Replace
Z-1198-2010 | 50777 | No | HIF | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Pnemosure Insufflators, model numbers 620-040-600 and 620-040-610, manufactured by Stryker Endoscopy, San Jose, CA. |  |  | December 29, 2008 | 2008 | Stryker Endoscopy | Unexpected powering down/failure to boot up: A problem with a transistor in the heater/temperature sensor component in the device may cause the device to power off during surgery or to fail to boot up at all. The risk includes delay and interruption of ongoing surgical procedures. | Component design/selection | Recall initiated 12/29/2009. Sales representatives were notified via phone and e-mail, while product advisory notices were sent to consignees via FedEx. Consignees are provided with a response form, and are instructed to run a software correction. |  | 884 | 884 devices | Product was distributed throughout the US and to Argentina, Brazil, Canada, Chile and Latin America, as well as to 16 government facilities. | Terminated | March 24, 2011 | 815 | 884 | Hardware | Device Operation | Software update | Software Update
Z-1197-2010 | 54773 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Exacta-Mix 2400 Valve Set, 24 - Port Valve Assembly, REF: 724, STERILE, 10 per box. Baxa Corporation, Englewood CO. |  |  | February 10, 2010 | 2010 | Baxa Corporation | Molding defect in the valve body causes leaking. | Nonconforming Material/Component | Baxa Corporation notified customers of the affected product by letter beginning February 10, 2010. Users were instructed to return all affected product to the firm.For further information, contact Baxa Corporation at 1-303-617-2242. |  | 104329000 | 3230 cases; 32300 units (10 units/case) | Nationwide Distribution -- Including AL, AR, AZ, CA, CO, DC, MN, IL, NY, DE, FL, PA, IN, TN, MS, MO, TX, MI, MA, MT, MO, SC, VA, WV, NJ, OH, CT, MO, NV, LA, WA, WI, GA, OK, NE, NH, NC, MD, KS, UT, IA, KY, HI, OR, SC, ME and SD. | Terminated | April 09, 2012 | 789 | 104329000 | Not_Computer
Z-0847-2010 | 54641 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DigitalDiagnost. General radiographic examinations and applications wherever a solid state X-ray imaging device can be used. |  |  | January 21, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Device failed to display appropriate labels indicating Radiation Symbol along with the required wording. | Radiation Control for Health and Safety Act | Urgent - Field Safety Notice letters, dated January 20 &amp; 21, 2010, were sent to customers. The letter identified the affected products, gave a description of the problem &amp; the hazard involved, stated how to identify affected products, actions to be taken by the customer or user, and actions planned by the firm. Customers are to make sure that the relevant safety precautions are observed and check their operators console (EEC) if the labeling is present as shown in the drawing. If the labeling is missing, a Philips rep should be contacted. For further information or support concerning the issue, a Philips rep should be contacted at 800-722-9377. |  | 117 | 117 units | Devices were distributed through out the US in the following states:AL, AR, AZ, CA, CO, FL, GA, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TX, UT, VA, WA, WI, and WY. | Terminated | January 25, 2011 | 369 | 144 | Other | Display/Image | Instructions | Safety Notice/Insructions
Z-1196-2010 | 53816 | Yes | NSA | Automated external defibrillator system. | Cardiovascular | Cardiovascular | 510(k) | Heartstart HS1 Defibrillator Model M5068A |  |  | November 02, 2009 | 2009 | Philips Medical Systems | Potential for failure of a capacitor during use which would prevent the AED from delivering effective defibrillation therapy when indicated. | Nonconforming Material/Component | The firm notified their US distributors of the recall and instructed them to contact customers and provide information how to return the product to Philips for a free replacement. Firm contacted distributors via phone and email on 11/2/09 and by letter on 11/9/09. A web site providing information related to the recall was activated on 11/3/09. |  | 34 | 34 units distributed domestically | Worldwide Distribution. | Terminated | May 23, 2013 | 1298 | 2588 | Not_Computer
Z-1202-2010 | 54663 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Tessier Osseous Microtome TOM(R) Bone Mill, 01-15402, Stryker Portage, MI. |  |  | February 18, 2010 | 2010 | Stryker Craniomaxillofacial Division | The firm has received 328 customer complaints for this device. Most of the complaints were involving non-functioning of the ratcheting handle or complete device, or the presence of metal shavings or discoloration. | Equipment maintenance | An "URGENT MEDICAL DEVICE CORRECTION" letter dated February 18, 2010, was sent to the customers. The letter describes the product, problem, scope of affected products, potential patient risks and actions to be taken by the customers. The customers should review the letter with all users of the TOM(r) Bone Mill and follow the instructions on the proper inspection and maintenance of the device enclosed with the letter. The customers should fax the accountability form within 5 days to help Stryker ensure their receipt of the information. If the accountability form is not received, addition communication will be sent.If you have any urgent questions or concerns, please contact us at (269) 323-4258. Please ref: PFA #2009-128. Our normal business hours are Monday-Friday 8am-5pm (GMT+1:00)On 7/8/2010 the firm sent an Urgent Medical Device Recall letter to its customers. The letter stated thatt upon consultation with regulatory agencies outside of the United States, the firm had intiated a devicd removal. The letter instructed the customers to return any devices that they had to the recalling firm. |  | 596 | 596 units | Worldwide distribution: USA and Belarus, Bulgaria, Denmark, Finland, Hungary, Jordan, Norway, Russia, Saudi Arabia and the UAE. | Terminated | November 24, 2010 | 279 | 596 | Not_Computer
Z-1200-2010 | 54845 | No | DTB | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | PMA | Medtronic, Sutureless, unipolar, myocardial, screw-in pacing lead, 5071-15, 5071-25, 5071-35 &amp; 5071-53, Medtronic, Inc., Minneapolis, MN 55432, Manufactured in: Villalba, Puerto Rico. |  |  | February 25, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic has identified a packaging issue for a subset of Epicardial Leads Model 5071. They determined that specific package seals could be compromised. Medtronic is not aware of any patient injury due to this packaging issue and their evaluation indicated the risk of patient injury is remote. However, a compromised package seal could potentially affect product sterility. | Packaging | Consignees were sent a Medtronic "Important Medical Device Recall" Letter dated February 2010 on 2/25/10. The letter described the problem and the product. Medtronic recommended to immediately remove the affected leads from inventory. Customers with questions, please contact your Medtronic Representative or Medtronic Technical Services at 1-800-505-4636. |  | 4863 | 4,863 (US); 6,328 (OUS) | All States in continental USA including DC except Alaska.OUS: Algeria, Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Belgium, Bolivia, Bosnia And Herzegovina, Brazil, Canada, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, El Salvador, Finland, France, Germany, Greece, Guam, Guatemala, Honduras, Hong Kong, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Latvia, Lithuania, Luxembourg, Malaysia, Malta, Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Romania, Saudi Arabia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Vatican City, Venezuela, and Yemen. | Terminated | January 07, 2012 | 681 | 4863 | Not_Computer
Z-1199-2010 | 54453 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Avanta Fluid Management Injection System, Avanta Multi-Patient Disposable Set (MPAT) High Pressure Check Valve, Catalog number AVA 500 MPAT, Material number 3018231. |  |  | January 06, 2010 | 2010 | Medrad Inc | The product may be defective resulting in a reduction of the saline delivery rate and inadequate air purging. | Nonconforming Material/Component | The recalling firm issued an Urgent Medical Device Recall letter dated 1/6/10 to all customers that received the affected lots. The customers were instructed to determine whether or not they had the affected product lots. Customers are to discontinue use and contact MEDRAD Customer Support or their MIP Rep to make arrangements for return of the recalled product. Any affected product returned will be replaced at no charge to the customer. Customers are also to complete and sign the attached form and fax it back to the firm. Questions concerning this recall are to be directed to Deb Easler, Product Manager, at 412-767-2400. |  | 2625 | 2625 pieces | Worldwide Distribution -- USA, including stated of AL, AZ, CA, DC, FL, GA, IL, KS, MD, MO, NM, NY, OH,OK, PA, TN, TX, and WI. and countries of Belgium, Germany, and Italy. | Terminated | April 08, 2010 | 92 | 2625 | Not_Computer
Z-1205-2010 | 53321 | No | CDT | Triglyceride test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista System StreamLAB Analytical Workcell. Dimension Vista System Part Numbers 1000035896, 1000035899, 1000036563, and 1000036567. StreamLAB Analytical Workcell Part Number 746002.901, for in vitro diagnostic use. |  |  | September 01, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Incorrect result: An intermittent low result may be seen with the last test run from an aliquot on a Dimension Vista System when connected to a StreamLAB Analytical Workcell. The sample delivery sequence may deliver insufficient surplus volume to the low sensitivity aliquot (second aliquot). The surplus sample volume is occasionally insufficient to and can cause short sampling on the last test. | Software design | Siemens healthcare Diagnostics issued an Urgent Field Safety Notice letter to each customer regarding the issue dated September 2009. The letter informed the user of the problem and volume change instructions. Customers can call Technical Solutions Center at 800-441-9250 |  | 81 | 81 units | Worldwide distribution: USA, Australia, Canada, France, Germany, Malaysia, Netherlands, Spain, and Switzerland. | Terminated | September 15, 2010 | 379 | 81 | Software
Z-1203-2010 | 54481 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Vygon Nutrisafe 2 syringe 5 ml, For Oral use only with Nutrisafe 2 |  |  | January 29, 2010 | 2010 | Churchill Medical Systems, Inc. | Mislabeled feeding syringes : 6 ml syringe mislabeled as 5 ml. | Labeling Change Control | Churchill Medical issued an Urgent: Product Advisory Customer Notification on January 29, 2010 to customers and distributors via overnight delivery. The letter identified the affected product along with the corresponding lot numbers and explained the problem. Accounts are requested to complete the enclosed form and fax it to 215-672-6740, attention Marie Benson. The affected products should be quarantined and customer service should be contacted to arrange return of product. Questions should be directed to Marie Benson at 800-473-5414. |  | 29700 | 29,700 eaches | Nationwide Distribution -- CA, NM, NJ, NY, PA, VA, and TX. | Terminated | September 23, 2010 | 237 | 129600 | Not_Computer
Z-1206-2010 | 53501 | No | EXX | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | VIPA Veronikis Intra-Fiberoptic Pelvic Assistant, Set Catalog No 350-370. |  |  | April 09, 2009 | 2009 | Marina Medical Instruments Inc. | Marina Medical, Sunrise, FL initiated recall of 46 different 350 series devices and 47 different 270 series products because the material used to manufacture the products was not proven to be biocompatible, resulting in risk of temporary skin or tissue irritation. | Nonconforming Material/Component | A "PRODUCT RECALL NOTICE" letter dated April 8, 2009, was sent to the customers. The letter describes the product, problem and the action to be taken by the customers. The customers are asked to identify the subject part in their instrument inventory and quarantine it. The Quality Control team at Marina Medical will contact the customer with instructions to return the part. When contacted, the customer should report any incidents of adverse reactions that have been observed at their institution. If you have any questions, please call 1-800-697-1119 ore (954) 924-4418. |  | 47 | 47 | Nationwide distribution. | Terminated | September 03, 2010 | 512 | 96 | Not_Computer
Z-1210-2010 | 54669 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific, Ultra-thin" Diamond" Balloon Dilatation Catheters, M001164040 (DT/3-4/5/40, 16-404), M001164360 (DT/5-2/5/40, 16-436),M001164620 (DT/6-4/5/40, 16-462), Made in Ireland, Ballybrit Business Park, Galway, Ireland |  |  | February 18, 2010 | 2010 | Boston Scientific Corporation | Boston Scientific Corporation is conducting a Medical Device Recall Removal of specific lots/batches of the Blue Max and Ultrathin Diamond Balloon Dilatation Catheters. Boston Scientific has determined that, the sterile barrier in the catheter packaging (i.e. outer pouch side seal) of the affected products may be compromised. An undetected breach in sterility of the outer pouch may lead to contami | Packaging | Two "URGENT MEDICAL DEVICE RECALL REMOVAL -IMMEDIATE ACTION REQUIRED" letters dated February 17, 2010 and February 18, 2010 were sent to the customers. The letters described the Action Description, Affected Product and Recall Instructions. The customers should immediately discontinue use of and remove all affected recalled product from your inventory, segregate the product in a secure location for return to Boston Scientific, complete and return the enclosed Account Reply Verification Tracking Form (even if you do not have any product to return) -fax to: Customer Service Call Center (866) 213-1806, and package/ship the recalled product per instructions. Your local sales Representative can answer any questions that you may have regarding this Recall Removal.If you have any question, please contact (763) 494-7971. |  | 99 | 99 | Worldwide distribution: USA-AL, CA, FL, IL, MI, MS, NJ, NY, NC, PA, RI, SC, TN, VA, WI, and CANADA, JAPAN, GREAT BRITAIN, NETHERLANDS. | Terminated | December 10, 2011 | 660 | 172 | Not_Computer
Z-1209-2010 | 55008 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO Radiation Treatment Planning System, Release 4.51.00. |  |  | February 22, 2010 | 2010 | Computerized Medical Systems Inc | Inconsistent MLC positions, dose, and monitor units | Software design | Elekta contacted customers by phone and mail beginning February 22, 2010, explaining the affected product and providing software update information.For further information, contact Elekta at 1-408-830-8023. |  | 25 | Planning system is loaded 25 in locations | Worldwide Distribution -- United States (PA, AZ, CA, MN, OH, WI, NC, FL, TX, NY and WI), Australia, The Netherlands and India. | Terminated | August 24, 2011 | 548 | 25 | Software | Output/Calculation | Software update | Software Update
Z-1212-2010 | 53430 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Zimmer PTW4321 20 Ncm Dental Torque Wrench, QTY 1 Pro-Dex Inc, Irvine, CA 92614 |  |  | January 19, 2009 | 2009 | Pro-Dex Inc dba Dyna-Dent | Pro-Dex has initiated a recall on their Dental Torque Wrench. The recall was initiated due to an inaccurate laser etch identification of the torque value. | Error in labeling | A "Voluntary Firm Initiated Recall" letter dated January 22, 2009, was sent to the customer. The letter describes the product, problem and action to be taken by the customer. The customer should immediately cease distribution or use of the noted lot; evaluate inventory and distribution records to ensure that this lot is removed from any potential point of use, and return the product to Returns Department at Pro-Dex, Inc . and include the RMA number on the shipping box. If you have any questions or concerns regarding this matter, please call (949) 769-3200. |  | 70 | 70 | Nationwide distribution: CA | Terminated | July 06, 2011 | 898 | 70 | Not_Computer
Z-1213-2010 | 54616 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Axiom Iconos R200 X-Ray |  |  | January 29, 2010 | 2010 | Siemens Medical Solutions USA, Inc | table can tilt/move when it is not intended | Software design | The recalling firm has issued a letter to affected customers and will send technicians to apply the fix to each defective device. |  | 169 | 169 units | Nationwide Distribution. | Terminated | August 11, 2010 | 194 | 169 | Software
Z-1214-2010 | 54477 | No | LKK |  |  | General Hospital | PMA | Medtronic N'Vision 8870 Application Card for use in the Medtronic model 8840 Physician Programmer. Manufactured for: Medtronic, Inc., Minneapolis, MN 55432-5604 USA. |  |  | March 23, 2009 | 2009 | Medtronic Neuromodulation | Routine monitoring of Medtronic's vigilance program identified user related programming errors. Medtronic has received reports of both overdose and underdose that were the result of programming use of SynchroMed implantable infusion pumps. These errors were most commonly associated with: --Priming Bolus Programming; and --Bridge Bolus Programming. Pump programming errors that result in overdos | Other | Medtronic is exchanging 8870 versions AAG, AAJ and BBG and replacing it with the updated versions AAH01 and BBH01. The updated versions contain user interface improvements to reduce the chance of user related programming errors when programming SynchroMed II and SynchroMed EL Infusion Pump system.A Medtronic "Medical Device Correction" letter was left behind for consignees. The letter was addressed to "Healthcare Provider" and was dated February 2009. The letter provided information regarding changes to the programming software used in SynchorMed Infusion System. For additional information, please contact your Medtronic Representative or Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 8499 | 8499 US | All States in the US including DC.OUS: AustriaBelgiumCroatiaCyprusCzech RepublicDenmarkEstoniaFaroe IslandsFinlandFranceFrench PolynesiaGermanyGreeceHungaryIndiaIrelandIsraelItalyKuwaitLebanonMaltaMartiniqueNetherlandsNew CaledoniaNorwayPolandPortugalReunionRussian FederationSan MarinoSaudi ArabiaSouth AfricaSpainSwedenSwitzerlandTurkeyUnited KingdomVatican City StateCanadaMexicoBrasil | Terminated | July 10, 2012 | 1205 | 8499 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1079-2010 | 54181 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Acacia Catalog # 385165 "Y" IV Extension Set 20cm Macro Bore, Spin Nut w/BD Q-Syte |  |  | December 18, 2009 | 2009 | Medical Product Specialist Inc dba MPS Acacia | Some lots of IV Extension Sets with BD Q-Syte Luer Access Split Septum have a manufacturing deviation which could cause the product to not function properly, possibly resulting in an air embolism. | Process design | A letter dated December 18, 2009 was sent to Acacia customers with the following instructions. 1. Immediately review your inventory of Acacia IV Extension Sets with BD Q-Syte for REF numbers and lot numbers listed and remove them from all inventory printed on the unit packaging , shelf box and shelf shipper containers.2. Complete the Distributor Recall response Card (Attachment 1) and return by fax immediately to the number on the Distributor Recall Response Card for product in your facility at the time of receipt of this letter. This form must be completed and f axed to Acacia, Inc even if there is no product in inventory.3. Acacia, Inc is currently working to ensure that replacement products are readily available. Once you have identified how much product you will be returning, please call your Customer Representative at 1-800-453-4538, opt. 2, ext. 2534, to arrange for product replacement.Questions are to be directed to the Customer Representative at 1-800-453-4538, opt. 2, ext. 2534, between 8:00AM and 5:00PM MST, Monday-Friday. |  | 69726 | 69,726 | US Only | Terminated | January 05, 2011 | 383 | 136441 | Not_Computer
Z-1124-2010 | 54723 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 10 Drop IV Sets, with 3 Pre-Pierced Y-Injection Sites and 3-Gang Pressure Activated Manifold Valve. (Approx. 10 Drops/ML); with Two Y-Injection Sites and Rotating Luer Lock (Approx. 10 Drop/ML); with Positive Pressure Pump with 4-Way "Hi-Flo" Stopcock and 4 Pre-Pierced "Y" Injection Sites (approx. 10 Drop/mL). |  |  | February 01, 2010 | 2010 | Arrow International Inc | Sterility may be compromised. | Component design/selection | The recalling firm issued an "Urgent Medical Device Recall" letter dated February 19, 2010 to their consignees to inform them of the issue. Customers were instructed to check their inventory for the affected products, cease use and distribution, and quarantine all affected product immediately. Consignees were further instructed the following: 1) Complete the enclosed Recall Acknowledgement &amp; Stock Status Form and immediately fax the completed form to Arrow International at 1-800-343-2935 and 2) Return all affected product to Arrow International, Inc. 312 Commerce Place, Asheboro, NC 27203 Attn: Asheboro Quality Assurance Manager. If additional information or clarification regarding the recall is needed, please contact Arrow International, Customer Service Department at 1-800-523-8446. Details about the recall can be accessed on the firm's website at http://www.teleflexmedical.com/spotlight/ivTubingRecall/index.html. |  | 70959969 | 70,959,969 of total units | Worldwide Distribution -- United States, Canada, Columbia, South Africa, Turkey, Kazakhstan, Venezuela, Germany, Cuba, Panama, El Salvador, Costa Rica, Singapore, and United Kingdom. | Terminated | November 08, 2012 | 1011 | 70959969 | Not_Computer
Z-1073-2010 | 53880 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medical Action Industries, IV Start Bundle Kit, using MPS Acacia with BD Q-Syte (BD #385151), Part number: 69946B. |  |  | November 06, 2009 | 2009 | Medical Action Industries Inc | Manufacturing defect may allow air entry into a central venous catheter, possibly resulting in an air embolism. | Nonconforming Material/Component | On 11/6/2009 a letter was drafted and sent by email to three (3) Medical Action Industries customers. They were instructed to return the recalled products in inventory to MAI. They further advised, if the product was distributed to notify their customer, inform them of the recall and instruct them to return the affected lots to MAI. |  | 12000 | 12,000 kits | Nationwide distribution | Terminated | June 01, 2011 | 572 | 18800 | Not_Computer
Z-1215-2010 | 54677 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ADVANCE TIBIAL WEDGE AUGMENT, REF KTAG-W315, SIZE 3/2+, Thickness 15 DEG, SURFACE NON-POROUS, 1 EACH, Wright Medical Technology, 4677 Airline Road, Arlington, TN 38002. |  |  | February 19, 2010 | 2010 | Wright Medical Technology Inc | The product contained screws with the incorrect length (incorrectly contained 15 mm length screws instead of the correct 5 mm length screws). | Employee error | The firm initiated their recall of the product by letter on February 19, 2010. The notification letters were addressed to distributors, hospital administrators, and surgeons with a fax back response form. The letter described the affected product and reason for the recall. Customers were instructed to examine their inventory and return the recalled product to Wright Medical Technology, Inc. Customers may contact Customer Service by calling 1-800-238-7117 for return instructions and replacement inventory. Direct questions regarding this matter to Wright Medical Technology, Inc. by calling 1-800-874-5630. |  | 47 | 47 units | Worldwide Distribution -- United States, Canada, Netherlands, Germany, Japan, S. Korea, UK, Czech Republic, Turkey and Brazil. | Terminated | July 08, 2011 | 504 | 47 | Not_Computer
Z-1217-2010 | 54655 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Live Better 1cc, 31 gauge x 5/16 inch U-100 Insulin Syringe short needle. 100 syringes, Manufactured by UltiCare. |  |  | February 05, 2010 | 2010 | Ulti Med, Inc. | UltiMed, Inc. is recalling one lot of 1cc, 31 gauge x 5/16 inch U-100 Insulin Syringe Live Better because it is mislabeled and actually contains 1/2 cc, 30 gauge x 1/2 inch U-100 insulin Syringe. | Employee error | An " IMPORTANT - DEVICE PRODUCT RECALL"letter dated February 3, 2010, was sent to customers. The letter describes the product, problem and action to be taken by customers. The customers should check their inventory for the affected lot, complete the recall tracking form and return by mail or fax and contact UltiMed for return instructions.If you have any questions concerning this recall, please call UltiMed, Inc. at 877-854-3434 and ask to speak to Jill. |  | 2500 | 25 units (each dispense and dispose sharps unit contains 100 syringes) | Nationwide distribution: CT and NJ | Terminated | May 07, 2011 | 456 | 2500 | Not_Computer
Z-1216-2010 | 55074 | No | FNM | Alternating pressure air flotation mattress. | General Hospital | General Hospital | 510(K) Exempt | AccuMax Quantum Complete (Model CU3/CU4), an alternating pressure with turn assist &amp; on demand low air loss system; Encompass Therapuetic Support Systems, Glendale, CA. |  |  | March 05, 2010 | 2010 | Encompass Therapeutic Support Systems dba BG North America | Potential out gasing of the battery and short term ignition of the gas within the control unit enclosure. This failure mode includes a risk for fire. | Nonconforming Material/Component | An "URGENT RECALL NOTIFICATION" letter dated March 5, 2010, was sent to the customers. The letter describes the product, problem and action to be taken by the customers. The customers should immediately discontinue use of the control unit, turn off the unit and disconnect from the wall socket and quarantine the unit away from any patient's room. The customers should immediately examine your inventory and quarantine the products listed; identify and notify any facility where they may have further distributed or transferred the listed products; and complete and return the REPLY/FAX-BACK FORM as soon as possible to Fax: (626) 796-7856 or mail to EncompassTSS, 100 E. Corson Street, Ste 310, Pasadena, CA 91103.Please contact EncompassTSS Customer Service at 800-822-8288 if you have any questions regarding this recall notice. |  | 349 | 349 | Nationwide distribution: CA, TX, IN, OR, and MO. | Terminated | March 21, 2012 | 747 | 349 | Not_Computer
Z-1218-2010 | 54742 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Harmony Lux Advantage LA Surgical Lighting and Visualization System, also known as the Harmony LA 500, B129382391. |  |  | February 23, 2010 | 2010 | Steris Corporation | Premature Bulb Failure | Device Design | STERIS issued an "Urgent Field Correction Notice" dated February 23, 2010 via Fed Ex. Consignees were informed of the affected product and were instructed on proper handeling of the device until a STERIS Service Technician contacts the consignee to schedule a service visit.For further information, contact STERIS Field Service Dispatch at 1-800-8828. |  | 10433 | 10,433 units | Worldwide Distribution -- United States, Canada, Algeria, Australia, Bahrain, Bangladesh, Brazil, Chile, China, Costa Rica, Ecuador, El Salvador, France, Germany, Guam, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Malaysia, Mexico, New Zealand, Nigeria, Pakistan, Philippines, Poland, Qatar, Russia, Saudi Arabia, South Africa, South Korea, Spain, Taiwan, Thailand, Turkey, United Arab Emirates, United Arab Emirates, United Kingdom, Italy, Yemen and United Arab Emirates. | Terminated | June 04, 2012 | 832 | 29460 | Not_Computer
Z-1179-2010 | 54719 | No | KLC | Dental chair and accessories. | Dental | Dental | 510(k) | Pelton &amp; Crane Dental Chairs, Chairman 5000 with magnetic headrest, Pelton &amp; Crane, Charlotte, NC 28273. |  |  | February 24, 2010 | 2010 | Pelton &amp; Crane Company | Magnetic headrests may affect patients with implantable pacemakers and defibrillators. | Device Design | Pelton &amp; Crane Issued an "Urgent Field Safety Alert" notification dated February 24, 2010. Consignees were instructed of the affected product and informed of proper actions to take. For further information, contact Pelton &amp; Crane at 1-704-588-2126. |  | 498 | Total: 498 chairs and 619 replacement upholstery spare parts | Worldwide Distribution -- United States, Australia, the Netherlands, Canada and Germany. | Terminated | October 28, 2010 | 246 | 498 | Not_Computer
Z-0975-2010 | 53761 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PTS Panels CHOL+HDL Panel test strips for professional use with CardioChek P-A analyzer, Manufactured by Polymer Technology Systems, Inc., Indianapolis, IN.Catalog # 1795. |  |  | February 20, 2009 | 2009 | Polymer Technology Systems, Inc. | Extended expiration date; the product expired on February 12, 2009 but was labeled with a April 4, 2009 expiration date. | Process control | Consignees were notified by letter dated Febuary 20, 2009, advised of the problem, and that the firm would replace product not consumed by the true expiration date. Direct questions to Polymer Technology Systems, Inc. |  | 3748 | 3,748 vials | Worldwide Distribution. | Terminated | October 20, 2010 | 607 | 3748 | Not_Computer
Z-1222-2010 | 54656 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Publix Sterile Preserved Saline Solution 12oz, Product SKU#041415-003767. |  |  | February 03, 2010 | 2010 | K C Pharmaceuticals Inc | The recall was initiated because the affected lots were manufactured between two failed media fills; therefore the lots produced cannot be assured as being free of contamination and may not be sterile. | Other | An "Urgent: Medical Device Recall Notification" letter dated February 17, 2010 was issued to customers. Consignees were instructed to immediately hold and discontinue selling the existing, affected stock.Consignees were also instructed to check inventories at all distribution locations for affected products shipped between June 8, 2OO9 and August 7, 2009 and perform a SUBRECALL to the retail level and report back to KC Pharmaceuticals. If the retailers indicate they have any of the affected lots of product, please have them contact KC Pharmaceuticals at 1-888-527-4276 for a shipping authorization number.Consignees were asked to complete the enclosed Recall Effectiveness Check Form, even if they did not have the products in question and return the form to KC Pharmaceuticals via mail within five (5) business days. The recalled lot and completed packing slips were asked to be shipped to:KC Pharmaceuticals-SALP Recall,3201 Producer Way,Pomona, CA 91768.Direct questions to KC Pharmaceuticals by calling 1-888-527-4276, extension 27. |  | 2004 | 2,004 units | Nationwide Distribution. | Terminated | January 05, 2011 | 336 | 46428 | Not_Computer
Z-1228-2010 | 51926 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 10 degree Tibial Augment Wedge, a component of Ortho Balanced Knee Modular Tibial System. |  |  | March 19, 2009 | 2009 | Ortho Development Corporation | A component of knee prosthesis was assembled incorrectly. | Device Design | E-mail was sent to all agents and distributors on March 19, 2009 with illustrations and instructions on how to detect the upside down assembly and how to correct the problem if detected during surgery. Direct questions to the Ortho Development Corporation by calling 1-801-619-3444. |  | 422 | 422 units | United States (AZ, FL, OR, TX and UT) and Japan. | Terminated | June 30, 2010 | 468 | 790 | Not_Computer
Z-1161-2010 | 53214 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System IS2000 Endoscopic Instrument Control System, rev. A51_P7, manufactured by Intuitive Surgical, Sunnyvale, CA. |  |  | July 15, 2009 | 2009 | Intuitive Surgical, Inc. | Gripper or Scissor jaws may close inadvertently, and will not open on command and various other reported modes of failure. Control by surgeon may fail and this failure may be difficult to detect. | Software design | Recall initiated the recall on July 1, 2009. US consignee recall was initiated on July 7, 2009. All affected consignees have been notified via e-mail communication followed by a field service engineer visit. Customers with questions or with need for additional assistance regarding the notification, please contact the firm's Intuitive Surgical Customer Service Department at 1-800-876-1310 toll free. |  | 782 | 782 units | Product was distributed to 551 consignees throughout the US to AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, WV AND WY, as well as to Puerto Rico, and to Austria, Belgium, Bulgaria, Czech Republic, Denmark, Finland, France, Germany, Greece, Ireland, Italy, Netherlands, Norway,Qatar, Romania, Spain, Sweden, Switzerland, Turkey, the United Kingdom, Argentina, Australia, Brazil, Canada, China, India, Israel, Japan, Korea, Mexico, Russia, Saudi Arabia, Singapore and South Korea. | Terminated | February 11, 2011 | 576 | 782 | Software
Z-1235-2010 | 54284 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD-LEGACYÂż 1 Model 6400, Ambulatory Infusion Pump, |  |  | December 23, 2009 | 2009 | Smiths Medical ASD, Inc. | The CADD-LegacyÂż 1 Model 6400 Ambulatory Infusion Pump (Serial Number 370969) was shipped without documentation that the required accuracy testing had been performed during the manufacturing process. If the subject infusion pump did not go through the accuracy testing, and if it does not meet accuracy specifications, there is a potential for over or under-delivery. | Employee error | Smiths Medical telephoned the customer on December 23, 2009, to return the CADD-LegacyÂż 1 Ambulatory Infusion Pump, serial number 370969, since there was no record of the completed accuracy testing for the infusion pump. |  | 1 | 1 | MI | Terminated | December 10, 2011 | 717 | 1 | Not_Computer
Z-1232-2010 | 55063 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Perfexion, Product Number: 715000. Radionuclide radiation therapy system. Elekta, Inc. Norcross, GA. |  |  | November 06, 2009 | 2009 | Elekta, Inc. | There may be a situation where the Frame Adapter might lock the stereotactic Frame in the wrong position. | Software design | Elekta, Inc. issued a Field Change Order 200 091 titled "Update Frame Adapter with New Frame Fixation Axle" dated November 6, 2009. Consignees were given instructions to prevent further issues related to the affected device.For further information, contact Elekta, Inc. at 1-770-300-9725. |  | 32 | 32 units | Nationwide Distribution -- AZ, CA, FL, IL, KY, MI, MN, MS, MO, NY, NC, OH, OR, PA, SC, TX, UT and VA. | Terminated | December 14, 2010 | 403 | 32 | Software
Z-1231-2010 | 54886 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Perfexion. Radionuclide radiation therapy system. Elekta, Inc. Norcross, GA. |  |  | May 30, 2007 | 2007 | Elekta, Inc. | There has been an issue with "Image Fushing" where low precision calculation caused images to become inaccurate. | Software design | Elekta, Inc. issued a Field Change Order, Reference Number 0000946 dated March 14, 2008. Consignees were informed of the affected product and instructed on necessary steps to take.For further information, contact Elekta, Inc. at 1-770-300-9725. |  | 18 | 18 units | Nationwide Distribution -- AR, AZ, CA, FLA, IL, MI, MN, MO, MS, NJ, OH, OR, PA, UT and VA. | Terminated | December 28, 2010 | 1308 | 18 | Software | Display/Image | Instructions | Safety Notice/Insructions
Z-1237-2010 | 54999 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis Systems operating with A100 Generator |  |  | March 04, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Due to a contact problem at the low voltage power supply of the generator controller, in rare cases where a user has invoked radiation, it may occur that no radiation is released. Under certain circumstances, radiation is possible by repeated release of the x-ray footswitch or button.To remedy the issue described above, a revised cabling has been installed in all potentially affected syste | Nonconforming Material/Component | The recalling firm issued a Customer Safety Advisory to affected customers via Update Instructions AX006/10/S. The letter informs customers of the potential issues and provides additional information on the implemented resolutions. For additional information, please contact the firm at 888-826-9702. |  | 1373 | 1373 units | The product was shipped to medical facilities nationwide. | Terminated | July 05, 2012 | 854 | 1373 | Hardware | Treatment/Delivery/Therapy | Repair | Repair
Z-1230-2010 | 54564 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Perfexion, Radionuclide radiation therapy system. Article Number 715000, Elekta, Inc. Norcross, GA 30092. |  |  | November 09, 2008 | 2008 | Elekta, Inc. | Radiation unit doors could close too fast on emergency exit. | Software design | Field Change Order 200 070 "MCU CS SW 8.3" dated September 11, 2008 was released. Consignees were informed of the necessary steps to avoid further issues relating to the affected product.For further information, contact Elekta, Inc. at 1-770-300-9725. |  | 19 | 19 units | Nationwide Distribution -- AR, AZ, CA, FL, IL, MI, MN, MS, MO, NJ, OH, OR, PA, UT and VA. | Terminated | December 28, 2010 | 779 | 19 | Software
Z-1238-2010 | 53993 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Vertier Surgical Table, Catalog numbers 0788100000, 0788100000S, 0788100001, and 0788100001S |  |  | November 13, 2009 | 2009 | Stryker Communications Corp | In certain positions, product is susceptible to fluid instrusion which could possibly result in short circuit or unsolicited movement. | Other | Firm notified all consignees by letter sent 11/13/2009. Contact the firm at 972-410-7100 with questions relating to this recall. |  | 206 | 206 | United States (GA, AR, ND, KY, CO, CA, SD, PA, OH, TX, SC, NY, MS, IL, KS, NJ, OR, WA, MN, VA, IL, OK, WY, IN, ID, UT, MO, FL, and DC) and Canada. | Terminated | August 11, 2010 | 271 | 206 | Hardware | Treatment/Delivery/Therapy | Notification | Safety Notice/Insructions
Z-1234-2010 | 54279 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ATAC Serum Control Kit. |  |  | December 04, 2009 | 2009 | Vital Diagnostics, Inc. | Vital Diagnostics observed a bias in the Quality Control Consensus Program between published values for ALP on the ATAC 8000 and reported customer means.The ranges were adjusted to reflect the analyzer population. | Other | Customers were mailed copies of Technical Bulletin 8100 and 5024 dated September 30, 2009 which advised of the Alkaline Phosphatase QC value revisions for the Envoy 500 and the ATAC 8000. If you have any questions, please contact Vital Diagnostics Technical Support at 1-800-345-2822. |  | 802 | 802 Kits | Nationwide. | Terminated | July 25, 2012 | 964 | 1673 | Not_Computer
Z-1251-2010 | 54144 | No | DXX | Steerable catheter control system. | Cardiovascular | Cardiovascular | 510(k) | Artisan brand Extend Control Catheter; An Steerable Guide Catheter; Component of Sensei X Robotic Catheter System; Model Number: 07546; Distributed and Manufactured by Hansen Medical, Inc, Mountain View, CA |  |  | October 19, 2009 | 2009 | Hansen Medical Inc | Certain Artisan Extend Catheters may develop a leak in the flush assembly of the catheter which has the potential to cause flush fluid leakage and subsequent decrease in distal flush; inability to aspirate air from flush lines, and/or introduction of air into the catheter. | Other | Recall Initiation Date: October 19,2009 (FedEx shipments on October 19, for October 20 delivery) via letters explaining the reason for the recall and requesting the product be returned. Customers with questions are instructed to call the company at 650-404-5940. |  |  |  | Nationwide to 12 states. | Terminated | February 09, 2011 | 478 |  | Not_Computer
Z-1250-2010 | 54540 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | MaxTorque Cannulated Screw Model # MSD-010-55-060S |  |  | April 07, 2009 | 2009 | Orthohelix Surgical Designs Inc | A pinhead sized disc was found in the cannulation of a MSD-010-55-60S screw. | Other | On April 7, 2009 the firm sent letters to their customers with detailed instructions on how to correct the problem. Pictures were included for clarification.For questions please contact either Jim Bragg via phone at 330-869-9562 ext 211 or Kristin Wolff at ext 209. |  | 45 | 45 pieces | Product was delivered to: CA, CO, GA, NM, NY, OH, VA &amp; WA. | Terminated | October 20, 2010 | 561 | 45 | Not_Computer
Z-1244-2010 | 54517 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Powerheart 9300E automated external defibrillator. |  |  | February 03, 2010 | 2010 | Cardiac Science Corporation | A component of the AED device may fail during a rescue attempt and the AED will not be able to deliver therapy. | Nonconforming Material/Component | The firm began initial notification on 2/3/2010 by sending letters to all affected consignees via USPS certified mail (domestic) and FedEx (international). Also on 2/3/2010 the firm issued a global press release and posted recall information on their website.Customers are instructed to remove their affected units from service, and have been given instructions for returning their affected devices. Replacement AED(s) will be sent as they are available. For questions, please contact the firm's Cardiac Science Support at 1.888.402.2484 or. AED195@cardiacscience.com to arrange for replacement(s). |  | 2640 | 2640 | Distributed to health care and emergency medical personnel and US Gov't/Military in the US and the following 44 countries:AUSTRALIA, BELGIUM, BELIZE, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, DENMARK, EGYPT, FRANCE, GERMANY, GRENADA, GUATEMALA, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, MALAYSIA, MALTA, MEXICO, NETHERLANDS, NORWAY, POLAND, PORTUGAL, PUERTO RICO, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, TAIWAN, THAILAND, TRINIDAD, AND GREAT BRITAIN. | Terminated | December 08, 2011 | 673 | 10340 | Not_Computer
Z-1252-2010 | 54934 | No | NSX |  |  | General Hospital | Contact ODE | CareSuite" - Critical Care Manager, PACU Manager and Anesthesia Manager |  |  | March 08, 2010 | 2010 | Picis Inc. | Under specific timing of conditions and in configuration with 3rd party infusion pumps, an error within the clinical application causes the manual documentation of clinical orders to be changed to an automatic update status. At this point, the order documentation becomes frozen at the last dose and/or rate and the user is unable to modify, stop or delete it. This error could impact the fluid bala | Software design | Picis issued letter notification beginning on March 8, 2010 with software correction information. The letter communicates details of the problem, along with recommendations for mitigating the potential to experience this error condition. Use of the product should be discontinued until the update described in the letter is performed. Customers with questions should call the company at 781-557-3000. |  | 23 | 23 | NJCanada, Australia, Austria,Denmark, Finland, Germany Norway,Portugal, Sweden, France,Netherlands, Spain. | Terminated | July 01, 2013 | 1211 | 23 | Software | Device Operation | Software update | Software Update
Z-1254-2010 | 54618 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Qwix 4.3mm Stabilization Screws - 34mm |  |  | February 08, 2010 | 2010 | Integra LifeSciences Corporation | The screws were etched incorrectly as 32mm instead of the correct 34mm length. | Labeling mix-ups | Consignees were notified via conference calls and/or FedEx mailing. Recall aknowledgement forms were provided with written notification. This requires consignees to review their inventory and to indicate any affected lots . This form requires a signature and must be returned. |  |  |  | Nationwide distribution - Alabama, Arizona, Colorado, Connecticut, Maryland, Michigan, New England, New Jersey, North Carolina, Ohio, South Carolina, Nebraska, and Utah. | Terminated | October 27, 2011 | 626 |  | Not_Computer
Z-1255-2010 | 54702 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Foundation Knee System, Patella Trial, 26 mm, part number 802-01-019, Lot number 53944551 |  |  | February 22, 2010 | 2010 | Encore Medical, Lp | One lot of product was incorrectly manufactured. Pegs are larger than specification. | Process control | Firm notified consignees by e-mail on 2/22/2010. |  | 2 | 2 units | United States (FL and ID). | Terminated | July 22, 2010 | 150 | 2 | Not_Computer
Z-1242-2010 | 54891 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PROMOS Inclination set, SIZE 26mm, REF 42 186, S&N; 75006725, QTY: (1), STERILE R, Smith &amp; Nephew Orthopaedics AG, CH-Rotkreuz-Switzerland. |  |  | January 14, 2010 | 2010 | Smith &amp; Nephew Inc | Firm received complaints of Inclination Set Screws fracturing. | Device Design | All affected Smith &amp; Nephew Sales Representatives were notified of problem and the recall overnight mail on January 23, 2009. All affected Smith &amp; Nephew International Distributors were notified via email and telephone on January 23, 2009. They were instructed to immediately quarantine the product for return to Stericycle, Inc., 2670 Executive Dr. Suite AIndianapolis, IN 46241. Direct questions to Smith &amp; Nephew, Inc. by calling 1-901-399-6153. |  | 362 | 362 units | Worldwide Distribution: United States, Australia, Austria, Belgium, China, Denmark, France, Germany, Greece, Great Britain, Italy, Netherlands, Poland, Portugal, Spain, Sweden and Switzerland. | Terminated | April 10, 2012 | 817 | 3025 | Not_Computer
Z-1256-2010 | 54982 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | NicoletOne vEEG System, VIASYS Healthcare, Inc., Madison, Wisconsin 53711. |  |  | March 08, 2010 | 2010 | CareFusion | CareFusion NeuroCare is voluntarily implementing a field correction to delete certain NicoletOne software protocols, which users may interpret incorrectly when conducting patient evaluations. This may result in the misinterpretation of the left versus right side of the brain.Users are being provided with instructions for removingthese protocols and verifying that the deletion has beenprop | Software design | Consignees were sent a CareFusion "Urgent Medical Device Correction" letter dated March 3, 2010. The letter described the problem and product involved. The letter included "Action to be taken by Users," "Instructions for deleting protocols," and "Photos-Amplifier label and Headbox insert to be discarded". It also included a "Verification Form" to be completed by recipient.For questions, notify CareFusion-NeuroCare, Attention: Sue Niesen, at (800) 356-0007 ext. 5129, or via email at susan.niesen@carefusion.com |  | 6131 | 2,984 US; 3,147 OUS | All States in the United States including DC.UNITED ARAB EMIRATES, ALBANIA, ARGENTINA, AUSTRIA, AUSTRALIA, AZERBAIJAN, BANGLADESH, BELGIUM, BULGARIA, BAHRAIN, CANADA, SWITZERLAND, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECHOSLOVAKIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, ESTONIA, EGYPT, SPAIN, ETHIOPIA, FINLAND, FRANCE, GREAT BRITAIN, GREECE, GUAM, GERMANY, HONG KONG, HUNGARY, CRO ATIA, INDONESIA, IRELAND, ISRAEL, INDIA, ICELAND, ITALY, JORDAN, JAPAN, KOREA SOUTH, KWAI T, KAZAKHSTAN, LEBANON, LIECHTENSTEIN, LATVIA, MOROCCO, MALDOVA, MACEDONIA, MALTA, MAURITIUS, MEXICO, MALAYSIA, NIGERIA, NETHERLANDS, NORWAY, PANAMA, PERU, PHILIPPINES, PAKISTAN, POLAND, PUERTO RICO, PALESTINIAN TERRITORY, QATAR, ROMANIA, SERBIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVEN IA, SLOVAK REPUBLIC, THAILAND, TURKEY, TAIWAN, URUGUAY, UKRAINE, VENEZUELA, YEMEN, VI ET NAAM, YUGOSLAVIA, SOUTH AFRICA, | Terminated | July 05, 2012 | 850 | 6131 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-1257-2010 | 53869 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | AC power cords manufactured by Electri-Cord with tall bridge design. Used across all Philips product lines, as well as Respironics/Philips products, Dixtel owned products and Dixtal (formerly Novametrix or Respironics/Novametrix products. Philips HeartStart MRx Defibrillator/Monitor model number M3535A, ID number 861288. |  |  | November 19, 2009 | 2009 | Philips Healthcare Inc. | Power cords' prongs may crack or fail -- Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside the plug. Potential for burns or electrical shock. | Nonconforming Material/Component | Philips was notified that FDA is investigating whether certain types of Electri-Cord Manufacturing Company power cords used with medical devices may be defective in that the power cord's prongs may crack and fail at/or inside plug. The issue affects medical device power cords equipped with a plug that has a prong and ground-pin insert design and a black plastic bridge connecting the terminal prongs on the plug ("taller bridge" design). Medical device power cords that do not have a black bridge connecting the terminal prongs on the plug are not affected. Urgent Medical Device Correction letters were sent to Distributors describing the problem and giving them instructions for notifying customers and actions to be taken. Customers are to visually inspect the affected cords, complete the Customer Reply Power Cord Order Form, and follow the instructions to dispose of affected cords. Further information or support concerning this issue can be obtained by contacting a Philips representative at 1-800-722-9377. |  | 16705 | 16,705 | Worldwide Distribution -- UNITED STATES, ALGERIA, ARGENTINA, Aruba, AUSTRALIA, BAHRAIN, BANGLADESH, BELGIUM, BRAZIL, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, Denmark, ECUADOR, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAMAICA, JAPAN, KUWAIT, LEBANON, LIBYAN ARAB JAMAHIRIYA, MALAYSIA, MAURITIUS, MEXICO, NETHERLANDS, NETHERLANDS ANTILLES, NEW ZEALAND, NICARAGUA, NORWAY, PAKISTAN, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, RUSSIA, SAINT LUCIA, SAIPAN, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIET NAM, and VIRGIN ISLANDS. | Open, Classified |  |  | 38473 | Battery
Z-1310-2010 | 54823 | No | NAD | Occlusive wound dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Aplicare Advantage Dressing Change System, Central Line Dressing Change, Sterile |  |  | March 03, 2010 | 2010 | Aplicare, Inc. | Incorrect expiration date on outer kit label-extended. | Error in labeling | Aplicare notified customers by telephone on March 3, 2010 followed with "Product Recall Notice" letter issued by e-mail to the distributor level. Product will be labeled with correct expiration date. Direct questions to your sales representative or customer service by calling 1-800-760-3236. |  | 103 | 103 cases (20 each) | United States (CT, RI, NJ and OH). | Terminated | March 16, 2012 | 744 | 103 | Not_Computer
Z-1313-2010 | 54851 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Datascope AS3000 Anesthesia System; |  |  | March 05, 2010 | 2010 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | An issue affecting the caster mount on the AS3000 System have been identified, specifically reports of the caster mount breaking, causing the wheel to fall off and the unit to tip. | Other | An "URGENT PRODUCT FIELD CORRECTIVE ACTION" letter dated March 5, 2010 was sent certified mail to all facilities/customers that have the AS3000 units. The letter described the product, problem and corrective Action/remediation by the firm. A Mindray Service Representative will contact the customers to arrange for the caster mounts to be replaced on the AS3000 units in their facilities at no cost.If you have any questions, please contact Karen Maine of the Mindray DS USA, Inc. Marketing Department at 201-995-8612 or k.maine@mindray.com. |  | 479 | 372 units in US, 107 internationally. | Worldwide Distribution: USA and Mexico, Columbia, Turks &amp; Caicos Islands, BWI; Ecuador, Bolivia, Peru, Dominican Republic, Netherlands Antilles, Chile.SA | Terminated | April 25, 2012 | 782 | 479 | Not_Computer
Z-1315-2010 | 54981 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | KODAK T-MAT G/RA Film, Catalog/REF # 151 0023, 24x30 cm, 5-100 Sheet Boxes FOLHAS/Hojas, Made in U.S.A. by Carestream Health, Inc. 150 Verona Street, Rochester, NY 14608 --- Device Listing # E420904 |  |  | March 12, 2010 | 2010 | Carestream Health, Inc. | Product mix-up. The firm received a customer complaint indicating that the customer had opened a box of film and although it was supposed to be T-MAT L/RA, it was performing like T-MAT G/RA. | Mixed-up of materials/components | A "MEDICAL DEVICE RECALL (VOLUNTARY)" letters and Response Forms (dated 3/12/10) were sent to the customers on 3/12/10 via Certified Mail, Return Receipt Requested. The foreign sales offices are being notified via e-mail (started on 3/12/10) so that they in turn can notify the foreign consignees of the recall action and report back to Carestream Health. The letter describes the products, problem and action to be taken by the customers. The customers should identify and remove from inventory shelves any product that is listed in the recall, complete and return the attached Product Information Form whether they have affected product or not by fax to Carestream Health Customer Service 1-800-445-9967. Carestream Health will organize the return through designated carriers and will pay for transportation; and supply replacement stock in a timely manner.If you have any questions, concerns, or product technical questions, please call or fax Customer Service or Technical support at 1-800-328-2910. |  | 875 | 875 boxes | Worldwide distribution: USA and Puerto Rico, Australia, Canada, Chile, El Salvador, Honduras, and New Zealand | Terminated | July 07, 2010 | 117 | 1435 | Not_Computer
Z-1312-2010 | 54846 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis Systems Operating with VB22, VB23, VB30, and VB31. Model number 5917054, 7008605, 7412807, 7413078, 7555357, 7555365, 7555373, and 7728392. |  |  | February 16, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Communication problems can occur which may result in unavailability of Bypass Fluoro. | Device Design | The recalling firm issued a Customer Safety Advisory Notice dated 2/16/10 to inform their customers of the problem. The letter identified the affected product and stated that the problem had been solved by the service technician by the time that customers would receive the letter. The letter also explained when the problem occurred and the possible risks associated with it. In addition, the firm stated what will be done to avoid the problem. |  | 640 | 640 units | Nationwide Distribution | Terminated | July 05, 2012 | 870 | 640 | Not_Computer
Z-1316-2010 | 53479 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT System Software v5.00 and v5.10, List Number 5F48-22 and 5F48-23; Part Number 7-204717-01 and 7-204717-02 when used with PM Algorithm LLS Brd, part number 7-94255-01 and PM LLS kit, U10 Chip and Ejectors, part number 7-200788-01. |  |  | September 11, 2009 | 2009 | Abbott Laboratories, Inc | ARCHITECT System Software v5.00 and v5.10 are incompatible with ARCHITECT LLS Board. Software patch re-establishes correct parameters. | Software design | Technical Service Bulletins dated September 11, 2009 were issued to affected consignees. The correction was issued via the Technical Service Bulletins (TSB) to provide a software patch to re-establish the correct parameters on the LLS board. Direct questions to Abbott Laboratories, Inc. by calling 1-972-518-6000. |  | 115 | 11 in the USA; 104 worldwide | Worldwide. | Terminated | January 25, 2011 | 501 | 115 | Software |  | Software update | Software Update
Z-1311-2010 | 54809 | No | KLC | Dental chair and accessories. | Dental | Dental | 510(k) | Pelton &amp; Crane Spirit 3000 Dental Chair. |  |  | February 09, 2010 | 2010 | Pelton &amp; Crane Company | The Dental Chair lacks a valid Device History Record. | Other | An "URGENT MEDICAL DEVICE RECALL" letter dated February 9, 2010, was delivered in person to the customer. The letter describe the product, problem and action to be taken by the customer. The customer was requested to provide the name and location of the consignee who received the chair so that Pelton &amp; Crane can contact the end user to arrange to replace the affected chair with a new dental chair, and to complete the enclosed Recall Acknowledgement/Return Form and fax to (704) 587-7204. A monthly status report will be submitted to FDA.Please contact Pelton &amp; Crane Customer Care at 1-800-659-6560 or call Mr. Frank Ray directly at 704-587-7227 to let us know where the device is currently stored. |  | 1 | 1 unit | Nationwide distribution: NC | Terminated | August 11, 2010 | 183 | 1 | Not_Computer
Z-1317-2010 | 54892 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator Monitor (Models M3535A, M3536A and M3536J) used with Philips 14.8-Volt Lithium Ion Battery (ModeI M3538A). |  |  | March 04, 2010 | 2010 | Philips Healthcare Inc. | Batteries older than 2 years have a higher potential for intermittent electrical contact, may contribute to an interruption in monitoring or therapy if the monitor/defibrillator is operating on battery. | Component design/selection | An "Urgent Medical Device Correction" letter with a Field Safety Notice was issued on March 4, 2010 to affected customers. The Instructions for Use Addendum/Application Note is included with the communication describing when to replace the unit's 14.8-Volt Lithium Ion Battery (Product #M3538A) Number M3538A in the the Philips HeartStart MRx Monitor/Defibrillator. Should you have any questions or concerns about the Device Correction, please contact your local Philips representative. |  | 37674 | 37,674 units | Worldwide: UNITED STATES, CANADA, AFGHANISTAN, ALBANIA, ,ARGENTINA, ARUBA,AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM BOSNIA&HERZEGOVINA;,BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CHILE, CHINA, COLOMBIA, CROATIA,CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, FRANCE,GERMANY, GREECE, GUADELOUPE, HONG KONG, ICELAND, INDIA, INDONESIA,IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KUWAIT, LATVIA, LEBANON,LIBYAN ARAB JAMAHIRI, LITHUANIA, LUXEMBOURG, MALAYSIA, MEXICO, NAMIBIA,NETHERLANDS, NETHERLANDS ANTILLES, NETHERLANDS ANTILLES, NEWCALEDONIA, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PAKISTAN, PALAU,PANAMA, PAPUA NEW GUINEA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTORICO, QATAR, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE,SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN,SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY,UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIETNAM and YEMEN. | Completed |  |  | 37674 | Battery | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-1319-2010 | 54375 | No | HWR | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | BIOMET MICROFIXATION, PATIENT DISTRACTOR DRIVER 10.87cm, (4.28")(L), RADEL, STAINLESS STEEL, NON-STERILE PRODUCT. |  |  | January 15, 2010 | 2010 | Biomet Microfixation, Inc. | The patient Distractor Driver, used in Osteotomy procedures, may not ratchet and advance the distraction device as intended. If the distraction device does not advance, bone consolidation may occur and a second surgery may be necessary. | Device Design | Consignees will be notified by letter on 01/15/2010 be either fax, email or mail. |  | 73 | 73 | US ( AZ,CA,DC,FL,GA,ID,IL,IN,MO,NC,TX,UT), Japan, Finland, Africa | Terminated | May 14, 2010 | 119 | 73 | Not_Computer
Z-1320-2010 | 52175 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | ThermoFisher Scientific Microm STP 420D automated tissue processor. |  |  | June 03, 2009 | 2009 | Thermo Fisher Scientific | The firm has received complaints regarding specimen integrity (higher rate of specimen fragmentation) as well as complaints that the chamber may stop rotating prematurely, resulting in sample deterioration. | Other | Consignees were notified of the problem by letter, on or about 6/3/09, updating the instructions for use. The units were subsequently upgraded on or before 9/8/09. For additional questions and information, please contact ThermoFisher's Technical Support department , Subject: PT Module Advisory Notice. |  | 16 | 16 | Hawaii, Kentucky, Louisiana, Michigan, Missouri, Nebraska, New York, North Carolina, Ohio, Pennsylvania, Tennessee, Texas and West Virginia. | Terminated | June 29, 2011 | 756 | 16 | Not_Computer
Z-0762-2010 | 54169 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Stryker Reflex Hybrid Screw Extractor; |  |  | October 12, 2009 | 2009 | Stryker Spine | The threaded tip of the Reflex Hybrid Screw Extractor has been found to fracture the head of the screw, causing problems in removing the screw and possible damage to the screw and/or plate while attempting to remove the screw. | Employee error | An "Urgent Product Recall" letter was sent to Stryker Branches/Agencies on October 12, 2009 and to surgeons who have used the device on October 26, 2009. The recall letter described the affected products, issue and potential hazards. Consignees were instructed to examine their inventory and hospital locations for the affected products and return to Stryker Spine as indicated in the recall letter. Customers were also instructed to complete the enclosed Customer Response form and fax to 1-201-760-8370.Direct questions about the recall to Stryker Spine by calling 1-201-760-8206. |  | 1791 | 1,791 | Worldwide Distribution. | Terminated | September 06, 2012 | 1060 | 2774 | Not_Computer
Z-1170-2010 | 54087 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | LM Aculight Corporation Perseus Pulsed Fiber Laser, Model number: ML36L-1541. Additional Model numbers with non-defective type"A" fiber: M22L-1540, M27L-1541, M281-1540, M30PL-1541, &amp; M33L-1541. Lockheed Martin Aculight, Bothell, WA, 98021. |  |  | May 01, 2009 | 2009 | Lockheed-Martin Aculight | Defect of optical fiber - The Perseus device with type "B" fiber has been found to emit potentially hazardous laser radiation. | Radiation Control for Health and Safety Act | A Radiation Warning Notification was sent to all customers. The notification stated the purpose and problem. In addition, a Notification of Potential Perseus Laser Safety Issue letter was sent to the only customer who received the defective type "B" fiber product. The letter identified the affected product and stated that the product was not considered eye-safe. The firm recommended that the customer cease using the laser with the Part Number: 35-001466-032 and Corresponding Unit Serial Numbers: 00001,00002, and 00003. The customer is to return the affected product for additional testing and warranty repair. Questions should be directed to Rick Whitty on 425-482-1100 or at Richard.j.whitty@lmco.com. |  | 3 | 3 | Worldwide Distribution -- USA, including states of PA, FL, MD, &amp; CO and the countries of Austria and Canada. | Terminated | April 30, 2012 | 1095 | 3 | Not_Computer
Z-1318-2010 | 54729 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems Accu Tnl Reagent Kit when used in conjunction with UniCel Dxl Access Immunoassay Systems; Reagent Part Numbers: A78803 &amp; 33340; Calibrator Part Number: 33345 |  |  | February 06, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed customer reports that different results have been obtained usingthe same patient samples on Access/Access 2 and UniCel Dxl platforms. Values obtained withUniCel Dxl systems have been demonstrated to have a positive bias compared to valuesobtained with Access or Access 2 systems. | Other | An "URGENT: PRODUCT CORRECTIVE ACTION (PCA)" letter dated February 5, 2010, was sent to the customers. The letter describes the product, problem and action to be taken by the customers. The customers are recommended to immediately discontinue using AccuTnI Reagents and Calibrators on the UniCel DxI system until further notice. If this is not possible, the customers are directed to notify their stakeholders of the bias until further notice. The customers should complete, fax (786-639-4000) or mail the enclosed RESPONSE FORM.If consignees need assistance or have any questions regarding the notification, they were instructed to contact TechnicalSupport at 1-800-854-3633 in the United States and Canada. |  | 196071 | 196,071 kits | Worldwide distribution: USA and Canada | Terminated | May 09, 2012 | 823 | 196071 | Not_Computer
Z-1321-2010 | 54930 | No | NYI | Gene expression profiling test system for breast cancer prognosis. | Immunology | Immunology | 510(k) | Agendia MammaPrint, Catalog Numbers MP01, MP US version 03 |  |  | March 03, 2010 | 2010 | Agendia Inc | Over a period of about 6 months in about 15% of results provided Agendia has been "over reporting" the chance of metastasis risk being 29% risk of recurrence instead of 10% risk.This over reporting of risk did not cause direct injury to patients. However, if used by physicians as the sole basis for decision-making, affected patients may have received unnecessary therapy. | Other | Communication with physicians was initiated by the firm's Chief Medical Officer on March 3, 2010. The firm's representatives will meet with treating physicians in person while having the corrected results in hand to inform each in person of the action, to allow a decision of the best course of action. Physicians will be provided with background information on the issue at hand and corrected results.As follow up, the physician would receive a corrected report per patient by FedEx in order to register the reception of corrected information. Because each user received a personal visit, no other contact information was planned or provided by Agendia. |  | 107 | 107 | Nationwide. | Terminated | June 18, 2010 | 107 | 107 | Not_Computer
Z-1322-2010 | 54373 | Yes | FKP | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter 1550 Single Patient System Hemodialysis Machines; Baxter Healthcare Corporation, Deerfield, IL 60015; product codes 5M5538, 5M5538R, 5M5575, 5M5551 and 5M5551R. |  |  | January 29, 2010 | 2010 | Baxter Healthcare Renal Div | Some of the 1550 Hemodialysis Instruments may have replacement circuit boards with incorrect versions of software. | Software change control | An "Urgent Device Correction" letter dated January 29, 2010 was sent to the direct accounts via first class mail on the same date, to the attention of the Hemodialysis Administrator. The letter informed the accounts that subsequent to a prior device field correction in 2003-2004, some 1550 hemodialysis instruments may have replacement circuit boards with incorrect versions of software. The accounts were instructed to confirm the software version in their systems and replacement boards. Instructions for checking the software version were attached to the letter. A complementary software upgrade will be made available to the account if their 1550 instrument and/or spare circuit board inventory has a software version other than 3.0A or 4.1B. The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter and listing the serial numbers and software versions of those instruments that do nit have software 3.0A or 4.1B.Home patients were not notified directly, the information was sent to their Dialysis Center. The 1550 operator's manual was revised during the 2003 device field correction. The manual correctly warns the operator to discontinue treatment if the 1550 experiences a FL04, FL07, or FLOg fault code.To make arrangements for the software upgrade or questions regarding the communication were directed to Baxter Global Technical Services at 1-800-553-6898, prompt 3.For additional questions, contact the company Monday through Friday from 8:00 AM to 5:00 PM CST at (Toll Free): 800-422-9837. |  | 443 | 443 boards | Worldwide Distribution -- United States - nationwide and Puerto Rico, Argentina, Brazil, Canada, Chile, China, Colombia, Ecuador, Guatemala, Honduras, Mexico, Panama, Philippines, Romania, Turkey and Venezuela. | Terminated | February 03, 2011 | 370 | 443 | Software |  | Software update | Software Update
Z-1323-2010 | 55137 | No | DTN | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | The Bigger Better-Bladder, Catalog Number BBB38 (with 3/8 inch ID tubing), Sterile, Single use only --- Manufactured for: Circulatory Technology Inc., 21 Singworth Street, Oyster Bay, 11771 (516) 624-2424 --- 510 (k) 981284. |  |  | June 03, 2009 | 2009 | Circulatory Technology Inc | During an FDA inspection, it was discovered that the firm had conducted a silent recall of its Bigger Better Bladder device after receiving a complaint from a hospital that the device collapsed during a procedure thereby obstructing flow. | Process control | The firm, Circulatory Technology, Inc., sent an email to Bigger Better-Bladder end user accounts on 6/03/09 and again on 6/13/09. The firm followed up with phone calls to the user customers. The emails and calls discussed the product, problem and the actions to be taken by the customers. The customers are ask to examine their stock of Bigger Better-Bladder for the affected units and then visually examine each unit to assure that adhesive has been applied and when the tube is bent away from cap, no separation is visible along the joint between the two according to the picture provided. The customers should return for full credit or a replacement of any product that has a separation of the seal.Please feel free to contact Mr. Yehuda Tamari at 516-624-2424 with any questions. |  | 144 | 144 units | Arizona, Arkansas, California, Florida, Indiana, Iowa, Maine, Michigan, Nebraska, New York, North Carolina, Pennsylvania, Ohio, Texas, Utah, Washington DC, and Germany. | Terminated | May 14, 2010 | 345 | 144 | Not_Computer
Z-1236-2010 | 55001 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife. |  |  | October 01, 2007 | 2007 | Elekta, Inc. | After updating the LGK actuator in spare part 810361, the old sleigh became obsolete due to causing insufficient locking of the helmet in combination with the new actuator. | Other | Field Change Order 0000203 "Replacement of obsolete actuator sleigh on Helmet Changer", was sent to all customers that were affected. Direct questions to Elekta, Inc. by calling 1-770-670-2548. |  | 5 | 5 units | United States (CA, TN, TX and WA) and Taiwan. | Terminated | December 28, 2010 | 1184 | 5 | Not_Computer
Z-1391-2010 | 54079 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Lifeshield, latex-free Secondary I.V. Set, convertible pin, 34 inch, piggyback with OPTION-LOK, non-DEHP, 15 drops/mL, a sterile Rx single use device, Hospira Inc, Lake Forest, IL USA. Made in Costa Rica, list 12667-28. |  |  | November 24, 2009 | 2009 | Hospira Inc. | Reports of difficulty in adjusting the flow rate, which could result in flow rate variances. | Component design/selection | Hospira sent the Urgent Device Recall letters dated 11/25/09 to their direct accounts of the affected products and requested these consignees to notify their customers of the recall. The accounts were informed that Hospira ceased manufacturing the I.V. gravity and pump administration sets with the CAIR clamp manufactured by Borla S.P.A. of Italy in March 2008, due to customer reports related to a difference in the zone of control used to adjust the flow compared to the original CAIR clamp design. To avoid the possibility of flow rate variances, Hospira is recalling all remaining sets with the Borla designed CAIR clamp. The accounts were requested to quarantine all inventory of the affected lots, complete the enclosed reply form indicating the amount of product being returned, fax a copy of the completed form to Stericycle at 1-877-377-5114, and include a copy with the returned goods to Stericycle at 2670 Executive Drive, Suite A, Indianapolis, IN 46241. |  | 308064 | 308,064 sets | Worldwide Distribution -- U.S., Puerto Rico, Dominican Republic, Bahamas, Virgin Islands, Argentina, Australia, Bahrain, Belgium, Canada, China, Denmark, Ecuador, EI Salvador, Finland, France, Germany, Greece, Guatemala, Hong Kong, Ireland, Italy, Jordan, Kuwait, Lebanon, Libya, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Peru, Philippines, Poland, Portugal, San Marino, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Uruguay, and United Arab Emirates. | Terminated | January 29, 2012 | 796 | 35723417 | Not_Computer
Z-1399-2010 | 54542 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Quick Connect Adapter Model #CSS-QA-Adapter MaxTorque(R) Maxtorque Quick Connect Lot #s: 1014001 &amp; 1101011, QTY 1 NON STERILE MAT:17-4PH |  |  | October 16, 2009 | 2009 | Orthohelix Surgical Designs Inc | The adapters are able to come apart when a retaining ring in the proximal end of the adapter can dislodge from the retaining ring groove and the two ends of the adapter fall apart and are unable to be reassembled without special tools and knowledge of the instrument.Also the adapters have a wobble between the two ends. This is caused when the distal and proximal ends of the adapter assembly were | Other | Two "IMPORTANT RECALL NOTICE" letters, one dated October 16, 2009, and the second letter dated November 19, 2009, were sent to the customers. The letters described the product, problem and action to be taken by the customers. The customer are requested to discontinue use of the adapters and return them to the firm, Orthohelix Surgical Designs. The firm will replace the adapters as soon as possible. Should you have any questions about the systems please contact Kristin Wolff at 330-247-1445 or Jim Bragg at 330-247-1444. |  | 127 | 127 pieces | Nationwide distribution: CA, CO, FL, GA, IL, KY, MD, MI, MO, MS, NC, NJ, NM, NY, OH, OR, PA, SD, TN, TX, VA, WA , and WI. | Terminated | April 11, 2012 | 908 | 127 | Not_Computer
Z-1397-2010 | 54175 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | SmarTemp Probe Covers for use with the DPM3 Monitor; Rx only, Mindray DS USA, Inc. Mahway, NJ 07430. Packaged in boxes of 20, Model ST010 also distributed in boxes of 2000 (10 individual boxes of 200); SmarTemp Module Kit (contains one box of 20 probes) and SmarTemp Temperature Probe Kits (contains one box of 20 probes). |  |  | December 31, 2009 | 2009 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | SmarTemp Disposable Probe covers may contain microscopic holes in the molded seam that would allow liquids or body fluids inside the cover and potentially contaminate the temperature probe. | Process control | Mindray DS USA, Inc. issued an "Urgent Product Recall" notification dated December 31, 2009 by certified mail. Consignees were informed of the affected product and requested to discard all remaining stock. Replacement product will be provided by the firm at no cost.For further information, contact Mindray DS USA, Inc. at 1-201-995-8000. |  | 679520 | Total: 33,976 boxes of 20 (boxes of 20 packaged into boxes of 200) in US, 801 boxes OUS | Worldwide Distribution -- United States, Canada and Australia.Nationwide distribution to hospitals, medical facilities and distributors as well as to Canada and Australia. | Terminated | December 09, 2011 | 708 | 679520 | Not_Computer
Z-1395-2010 | 55040 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity Web Diagnostic 1.0(WebDX) software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | February 18, 2010 | 2010 | Ge Healthcare It | There is a potential safety issue associated with the use of GE Centricity PACS-IW and PACS Web Diagnostic software related to use of built-in reconstruction functions which may result in incorrect anatomic orientation markers. | Software design | An "Urgent Medical Device Correction" letters dated February 18, 2010 to their Centricity PACS-IW software version 3.x customers and Centricity PACS Web Diagnostic software version 3.7.4.x customers, The letters advised the users of the patient safety issues associated with the use of the software when using built-in reconstruction functions, which may result in incorrect anatomic orientation markers. The separate letters provided the users with specific safety instructions for that software version to follow until the software is updated. The customers were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479.Any questions should be directed to the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. |  | 36 | 36 units | Worldwide distribution: USA and Australia, Austria, Belgium, Canada, Finland, France, Germany, Hungary, India, Ireland, Israel, Italy, Japan, Korean, Kuwait, Latvia, Lebanon, Libya, Lithuania, Malaysia, Netherlands, Philippines, Poland, Qatar, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Switzerland, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 670 | Software
Z-1007-2010 | 51256 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACSCount Controls Kit, An Automated Differential Cell Counter, Model/Catalog Number: P/N: 340166, |  |  | January 29, 2009 | 2009 | BD Biosciences, Systems &amp; Reagents | The product is mislabeled, in that, zero/low controls were mislabeled as medium/high, and the medium/high controls were labeled zero/low. | Package design/selection | On 1/29/09, the firm initiated the recall and its notification was via Important Product Information letters identifying the affected product and explaining the reason for the recall. Customers who experienced error codes with the affected product were to contact their local BD Biosciences representative or distributor to receive replacement BD FACSCount controls. |  | 17 | 17 units | Nationwide Distribution -- Including states of WA, MO, NV, CA, NY, PA, AL, and OR. | Terminated | March 08, 2011 | 768 | 17 | Not_Computer
Z-1404-2010 | 53031 | No | NVJ | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Renessa RF System, model PR0918, manufactured by Novasys Medical, Newark, CA |  |  | July 27, 2009 | 2009 | Novasys Medical Inc | Instructions for Use (IFU) revision-- Because of complaints received by the firm of worsened stress incontinence symptoms 3-4 weeks after treatment, the instructions for use were revised to emphasize potential side effects. Although Novasys has not, to date, identified a device malfunction associated with the symptoms discussed above, the investigation over time shows that there is a correlation | Labeling design | A Field Action notification letter was sent to all current customers (physicians) along with the revised instructions for use, and a small supply of the revised patient information brochures. The company also updated physician information brochures and patient/physician DVDs to include the information. This was shipped to physicians in August/September of 2009. Additional copies of the Patient Brochures can be ordered by contacting Novasys Customer Service at 866-784-4777. Questions should be addressed to the firm at the same telephone number. |  | 72 | As of 6/12/09 there were 436 Probes at customer accounts. | US nationwide only. (TX, FL, CA, LA, KY, MO, GA,OR WA, HI, AZ, OK AL, TN) . | Terminated | January 08, 2011 | 530 | 72 | Not_Computer
Z-1400-2010 | 54004 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | V-Probe, Endocare Cryo V Probe model # Cryo 206V. |  |  | December 05, 2008 | 2008 | Endocare, a Wholly Owned Subsidiary of HealthTronics, Inc. | Endocare received a report of a patient death that occurred during a prostate cryoablation procedure. An autopsy attributed the cause of death to an air embolism. The safety notice for the cryoprobes was issued to refocus attention on the Indications and Safety Review included with the Physician Training Materials and to reinforce additional safety information. | Other | The recall was initiated with Endocare sending a notification letter to identified customers on December 5, 2008 and offering an optional replacement unit. The initial correspondence was sent out via an e-mail if the correspondence was returned undeliverable or an e-mail address was not available it was then sent Federal Express to those recipients.The notification letter explained the reason for recall and requested providers refocus their attention on the Indications &amp; Safety Review information contained in the Endocare's Physician Training Materials. Physicians were instructed to contact the firm at 1-949-450-5412 to discuss any questions they may have about the information in the notification letter. |  | 22741 | 22,741 units total for all products | Nationwide. | Terminated | October 21, 2010 | 685 | 22741 | Not_Computer
Z-1407-2010 | 54783 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Product is a pull cord type monitor, bed patient alarm used to detect patient/resident movement. An adjustable cord is attached to both the alarm and the patient. When the patient moves beyond the length of the cord the pin is pulled and the alarm activated. |  |  | February 12, 2010 | 2010 | AliMed Corporation | Alimed became aware of issue via internal receiving inspection activities on 1/20/09. Firm states that subsequent to this observation, Customer Dissatisfaction inquiry #100192 was generated on 1/22/10. A single unit was received back by AliMed and 0 n/2/1/20 from a rehabilitation center that indicated that it had received a unit and it did not function upon its initial use. Failure of the alarm | Device Design | Letter sent to customers stating that some units may have been subjected to excessive heat during the manufacturing process causing internal distortion to a key alarm component. The letter also request that the Health Care professional review records and notify all users to whom the product may have been issued and if any product has been issued outside of their facility, to forward this correspondence to them. Firm has enclosed a form for return of the product and issuance of a new product. For questions, please contact the firm using the information provided in the letter. |  | 30 | 30 units | All product was distributed domestically with the exception of one unit sold in Canada. AliMed sells and distributes this product to medical health professionals and distributors as well as directly to end users | Terminated | July 25, 2011 | 528 | 30 | Not_Computer
Z-1403-2010 | 55037 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Centricity Web Diagnostic (WebDX) software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. |  |  | February 11, 2010 | 2010 | Ge Healthcare It | There is a potential safety issue associated with the use of GE Centricity PACS-IW and PACS Web Diagnostic software when dragging-and-dropping images from the Navigator causing flipping of the images that may impact patient safety. | Software design (manufacturing process) | GE Healthcare issued "Urgent Medical Device Correction" notifications dated February 11, 2010. The initial notification advised users of the patient safety issues associated with the affected product. Separate letters provided the users with specific safety instructions. Accounts were requested to complete the enclosed customer reply form, supplying the workstation software or build number, and fax the completed form to 847-939-1479.For additional information, contact the Remote Online Center, Field Engineer or the GE Customer Care Center at 800-437-1171. |  | 13 | 13 units | Worldwide Distribution -- United States (California, Illinois, Indiana, Minnesota, Nevada, New York and Wisconsin), Austria, Canada, Finland, France, Germany, Ireland, Lebanon, Japan, Philippines, Saudi Arabia, Singapore, Switzerland, Turkey, United Arab Emirates and the United Kingdom | Terminated | November 15, 2010 | 277 | 42 | Software
Z-1409-2010 | 55190 | No | LGW |  |  | Neurology | PMA | Bionic NAVIGATOR Clinician Programmer Model Number SC-7150-4 |  |  | March 19, 2010 | 2010 | Boston Scientific Neuromodulation Corporation | Toshiba has determined that certain Sony battery packs installed in Toshiba portable computers can overheat, posing fire hazard to customers. Based on Toshiba's investigation with Sony Corporation, they concluded that the battery cells in specific manufacturing lots could be affected. | Other | In a letter dated March 19, 2010 Boston Scientific informed consignees that it determined that certain Sony battery packs installed in Toshiba portable computers can overheat, posing fire hazard to customers, and telling them that they may be in possession of an affected product. They were advised that a Boston Scientific representative will contact them to arrange for an inspection of the Clinician Programmer (SC-7150-4).Until they have determined if the battery pack is affected, consignees were told they can only use their Toshiba portable computer with the original AC adapter power cord supplied with that computer, and were asked to remove the battery pack until its status is determined.Included with the letter was a Reply Verification Tracking Form which consignees were asked to complete the form following the instructions on the form and return to Boston Scientific in the enclosed postage paid return envelope or fax it to (877) 856-3546 by March 31, 2010.If they experience any issues they were told to cease using the Clinician Programmer and contact Boston Scientific's 24-hour Patient Care Hotline at (866) 360-4747.Customers can also contact the Field Action team by phone at (866) 684-7504 or by fax at (877) 856-3546 if they had any questions. |  | 62 | 62 units | Domestically to AR, CA, CO, DC, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NY, OH, OR, PA, PR, SC, TX, VA, WIInternationally: Italy | Terminated | February 29, 2012 | 712 | 62 | Battery | Physical Safety Hazards | Remove | Remove or Replace
Z-1408-2010 | 55177 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | The BleaseSirius and BleaseFocus Anaesthesia Systems. |  |  | March 26, 2010 | 2010 | Blease Medical Equipment, Ltd. | Potential for delivery of a lower than set concentration of anaesthetic to a patient due to a sticking valve. The lower shuttle may become jammed, not permitting flow of anaesthesia gas through the vapouriser. The vapouriser back bar mounting valves will require exchange only on the systems identified in this matter, no other systems are affected. | Other | On 2/26/2010 notification letters were sent to all international consignees. The 2 domestic consignees were sent notification letters via FEDEX on 3/26/2010. The domestic consignees were also contacted by phone on 3/26/2010.The notification letter is titled URGENT MEDICAL DEVICE CORRECTION. It advises users of the issue and provides additional information which would permit the continued use of the Blease system until replacement back bar valves can be fitted by the recalling firm's service representatives.For users who choose to continue to use the device, the manufacturer recommends the use of an Agent Monitor and user education; also checks that vapourisers have been mounted correctly and the system is leak free, and testing in advance of a procedure. For additional questions, US Customers are directed to call Global Technical Support - +1 425-657-7200 ext: 5089.Non US Customers, call Global Technical Support - +44 1494 784422. |  | 40 | 40 units | Distributed to 2 U.S. hospitals and to customers in the following countries:Panama, Poland, Australia, Singapore, United Kingdom, Bangladesh, Nicaragua, Vietnam, Argentina, Finland, and Columbia. | Terminated | August 23, 2011 | 515 | 40 | Not_Computer
Z-1405-2010 | 55101 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI Medical Linear Accelerator. Model Numbers R3.5, R4.0 and R4.2. Elekta, Inc. Norcross, GA 30092. |  |  | December 12, 2008 | 2008 | Elekta, Inc. | It is possible that two (or more) CT reference data sets are necessary for one patient, for example, if you have to scan a patient during breath hold and then again during free breathing. | Software design | Elekta, Inc. issued an "Important Notice" dated December 12, 2008. Customers were informed of the affected device and instructed on proper steps until an update is provided.For further information, contact Elekta, Inc. at 1-770-300-9725. |  | 76 | 76 units | Nationwide Distribution -- AL, AZ, CO, CT, FL, GA, ID, KS, KY, MA, MI, MN, MS, NE, NY, OH, OK, OR, PA, TN, TX, VA, WA and WV. | Terminated | December 28, 2010 | 746 | 76 | Other |  |  | Other
Z-1406-2010 | 55100 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PrecisePLAN 2.10 Treatment Planning System |  |  | March 15, 2006 | 2006 | Elekta, Inc. | Error in dose rates which were calculated at exactly twice the expected values. | Software design | As described in a prior related submission, Phase I - "Important Notice A285, Scrolling Selected Frame Dose Calculator Error" dated March 15, 2006 was distributed to affected customers to inform them of the potential and what steps should be taken.Phase 2 -- Software release PrecisePLAN 2.11 was released July 11, 2006 to correct this problem. All customers with PrecisePLAN 2.10 will be upgraded to the revised software PrecisePLAN 2.11 through Field Change Order #772 00 540 052 dated June 5, 2006. Phase 3 -- All customers with PrecisePLAN 2.10 will be upgraded to the revised software PrecisePLAN 2.11 through Field Change Order #772 00 540 052 dated June 5, 2006. The company is directing questions to the local Electa representative or to their offices in the U.K. at telephone 44 (0)-1293-654200. |  | 40 | 40 units | Worldwide: IL, NY, PA, TN, Canada, China, Czech Republic, Germany, India, Israel, Italy, Mexico, Poland, Russia, Spain, Switzerland | Terminated | December 30, 2010 | 1751 | 40 | Software | Output/Calculation | Software update | Software Update
Z-1201-2010 | 54478 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Howmedica Restoration ADM Trial Cup Holder; |  |  | January 21, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The Restoration ADM Trial Cup Holder may not have been assembled correctly. | Other | An "URGENT PRODUCTION OPERATION" letter dated January 21, 2010, was sent to customers via FedEx. The letter described the product, problem and actions to be taken by customer. The customers should follow the instructions on the enclosed PRODUCT BULLETIN. The customer should check their internal inventory immediately and quarantine all affected devices, identify if the device is correctly assembled and complete the Product Correction Acknowledgement Form and fax it to (201) 831-6069. If you have any questions, feel free to contact Rich Wolyn, Manager, Divisional Regulatory Reporting at (201) 831-5158. |  | 5 | 5 units | Nationwide distribution: AZ, CA and OR | Terminated | June 20, 2012 | 881 | 5 | Not_Computer
Z-1421-2010 | 54547 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Cordis, AVANTI+ Introducer, STERILE EO, For one use only. |  |  | January 28, 2010 | 2010 | Cordis Corporation | Sterility compromised: Cordis found a pin hole in one of the packaging tray cavities.On January 28, 2010 Cordis Corporation initiated a recall on their AVANTI + Introducer. During routine inspection, a visual irregularity was detected in packing trays. | Package design/selection | Cordis initiated the recall January 28, 2010. (The packets were sent overnight on January 28, 2010, with first receipt by customers January 29, 2010.) Immediate corrective action was taken through an inspection and removal of all affected products. |  | 1850 | 1850 units | Nationwide. | Terminated | July 27, 2012 | 911 | 1850 | Not_Computer
Z-1413-2010 | 54640 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | 3M Micropore Single Use Rolls, Surgical Tape, REF 1530S-1, 1 in x 1-1/2 yd. |  |  | January 29, 2010 | 2010 | 3M Company / Medical Division | Reduced adhesion: Micropore Surgical Tape for use as the primary securement device for 1) dialysis needles used to access A-V fistulas or grafts; or 2) pressure dressings used after dialysis treatments. If the tape does not perform, the dialysis needle may loosen and could then dislodge. In the interim, please offer 3M Micropore Surgical Tape, 1 inch x 10 yard rolls (CAT #1530-1) as a substitutio | Nonconforming Material/Component | An Urgent Medical Device Recall letter was sent to 3M customers beginning 01/29/2010. The letter described the issue, asked distributors to check existing inventory, quarantine affected lots. Distributors were asked to notify their customers and end-users (3M has provided example letters and return forms to facilitate). End users can return product either to the distributor or contact 3M directly. A Product Recall Form is asked to be faxed back to 3M.A second 3M "Urgent Medical Device Recall" letter dated February 22, 2010 was sent to customers. This letter informed the cusotmers that 3M was expanding the recall to include bulk-packaged rolls of Micropore Surgical Tape. The letter addressed to "Valued Kit manufacturer" described the problem and the product, provided instrucitions for "Action Requried by the Kit Manufacturer", "Action Requried by the Kit Manufacturer to the End-User", and "Action Required for Future Tape Orders from 3M". A Product Recall Form is aked to be faxed back to 3M. |  | 16582600 | 165,826 boxes, 100 rolls per box (16,582,600 rolls) | Nationwide distribution, and Puerto Rico. | Terminated | February 04, 2012 | 736 | 20445231 | Not_Computer
Z-1422-2010 | 54769 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra, Licox Ref 1P2.P; |  |  | February 18, 2010 | 2010 | Integra LifeSciences Corp. | One probe that had failed a product release test was inadvertently packaged and released into inventory in Integra's distribution center. | Process control | An Integra NeuroSpecialist was sent to retrieve the probe on February 18, 2010, followed by a formal letter. |  | 1 | 1 | One hospital in Omaha, NE. | Terminated | April 20, 2010 | 61 | 1 | Not_Computer
Z-1423-2010 | 55108 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging XS model number 10496279 |  |  | February 16, 2010 | 2010 | Siemens Medical Solutions USA, Inc | images can be overwritten | Software design | A "Customer (Safety) Advisory notice" dated February 16, 2010, was sent all affected customers via Update Instructions IM019/10/S. This notice describes the product, problem and the action to be taken by the customers. The customers should immediately instruct their employees of the advisor notice and maintain awareness. If the customer sold the device/equipment and is no longer in possession of it, they should forward the notice and inform Siemens of the new owner. A software update to correct this issue will be distributed and installed by a Siemens Service Engineer on all affected systems under Update Instructions IM020/10/S.If you have any further questions regarding this issue, do not hesitate to contact your local SIEMENS Uptime Service Center (michael.mib.braun@siemens.com). |  | 6 | 6 units | Nationwide distribution: CA, NC, NY, PA, and TX | Terminated | July 05, 2012 | 870 | 6 | Software | Display/Image | Software update | Software Update
Z-1412-2010 | 54648 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Eclipse Treatment Planning System, with v client versions 7.3 (build 7.3.10), 8.0 (builds 7.5.x), 8.1 (builds 8.1.x) and 8.5( builds 8.2.x). For radiation therapy. |  |  | December 09, 2009 | 2009 | Varian Medical Systems Oncology Systems | Potential for misdelivery of treatment-- When preparing a treatment plan involving a split, the device may calculate excessive or insufficient monitor units for the treatment plan. | Software design | Recall initiated on 12/09/2009. Varian issued a notification letter to its consignees, including a description and corrective action step. The firm is in the process of preparing a software fix for the issue. |  | 3177 | 3177 unit | Nationwide distribution and Mexico. | Terminated | February 24, 2012 | 807 | 3177 | Software | Output/Calculation | Software fix | Software Update
Z-1426-2010 | 54858 | No | GGP | Factor deficiency test. | Hematology | Hematology | 510(k) | REF 822 IMUBIND(R) Plasma PAI-1 ELISA |  |  | March 05, 2010 | 2010 | American Diagnostica, Inc. | Unacceptable microwell to microwell variation in the section of the plate that is typically used for the calibrators | Process control | Three emails "REF 822 lot 102401 needs to be recalled" dated March 1, 2010 was sent to three customers. The emails described the product, problem and action to be taken by the customers. The customers should send the kit to the firm (SEKISUI american diagnostica inc.) attention of Leigh Ayres, Director, Regulatoy Affairs and Quality Assurance.If you have any question, please call 1-203-602-7777. |  | 25 | 25 kits | Worldwide distribution: USA and UK, Taiwan and Korea. | Open, Classified |  |  | 25 | Not_Computer
Z-1410-2010 | 54894 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Biphasic LIFEPAK 15 Monitor/Defibrillator. Manufactured by Physio-Control Inc., a division of Medtronic, Inc. Redmond, Washington USA. |  |  | March 04, 2010 | 2010 | Physio Control, Inc. | Potential for the device to power off then on by itself, or to power off by itself and requiring the operator to turn it back on, or the device doesn't turn off. | Device Design | Physio-Control Inc. issued an "Urgent Medical Device Correction" notification dated March 2010 to domestic consignees. Foreign consignees were notified by overseas representatives. Consignees were advised to keep the affected device in service and to test the units in accordance with operating instructions. They were further advised that a local service representative will call and schedule a service visit within 60 days. Consignees were asked to forward the notification to all of their sites and to notify the recalling firm if they no longer own the device.For further information, contact Physio-Control Inc. Technical Support at 1-800-442-1142, Option 5, 6:00 am to 4:00 pm (Pacific), Monday-Friday. |  | 3609 | 3,609 | Worldwide Distribution -- United States, Netherlands, Germany, Canada, Australia, and Hong Kong. | Terminated | January 03, 2012 | 670 | 3609 | Other | Device Operation | Repair | Repair
Z-1425-2010 | 55142 | No | GZB | Implanted spinal cord stimulator for pain relief. | Neurology | Neurology | 510(k) | Swift-Lock (TM) Anchor Accessory Kit, Directions For Use Model 1192, ANS, A St. Jude Medical Company. |  |  | February 22, 2010 | 2010 | Advanced Neuromodulation Systems, Inc | Directions For Use of product incorrectly stated lead compatibility information. | Other | An "Urgent Medical Device Correction" letter dated March 8, 2010, was sent to the customers (note: Firm verbally informed consignees on February 22, 2010.). The letter describes the product, problem and action to be taken by the firm -St. Jude Medical. The firm's Sales Representatives will contact the customers to update any DFUs in their possession.If you need additional copies of the insert or more information and if you have any questions or concerns, please contact your St. Jude Medical Sales Representative at 1-651-490-4470. |  | 1120 | 1,120 units (218 went to doctors, remaining 902 were distributed to sales representatives) | Nationwide distribution: AZ, CA, CO, DC, FL, GA, IL, IN, MI, MO, MS, NC, NM, NJ, NV, NY, OH, OK, OR, PA, SC, TN, WA AND WV. | Terminated | July 22, 2010 | 150 | 1120 | Not_Computer
Z-1428-2010 | 55144 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Leksell SurgiPlan |  |  | May 23, 2005 | 2005 | Elekta, Inc. | Images from a different patient can be imported without a warning for the user. | Other | Phase I - A technical note for LGP and SPS describing the risk for importing wrong image files and including recommendations on how to reduce this risk has been written and distributed to users of both Leksell GammaPlan and Leksell SurgiPlan on May 23, 2005 via FCO DOC 05-0001. Phase 2 - A maintenance release was created of the Scanner Interface (SIF) to correct the problem. This was distributed to Leksell SurgiPlan users on November 30, 2005 via FCO0000153-1. Additional questions are directed to the firm to Product Manager lrobert.aslund AT elekta.com, or After Sales gustaf.piehl AT elekta.com.. |  | 57 | 57 units | Nationwide, Canada and Mexico. | Terminated | December 14, 2010 | 2031 | 57 | Software | Output/Calculation | Software update | Software Update
Z-1430-2010 | 54541 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Mini Max Torque 2.6 mm diameter Drill Bit, Model #HCS-072-26 &amp; 3.8 mm Countersink, Model # HCS-070-38. |  |  | September 03, 2009 | 2009 | Orthohelix Surgical Designs Inc | The epoxy is flaking off the drill bits and countersinks. | Nonconforming Material/Component | On 9/3/2009 the firm sent letters to their customers requesting return of all recalled product in stock and for those products to be replaced with the ones listed in the letter. Questions should be directed to 330-869-0625. |  | 74 | 74 pieces | Nationwide Distribution -- CO, OH, PA, SD, TX &amp; VA. | Terminated | February 28, 2012 | 908 | 74 | Not_Computer
Z-1429-2010 | 55148 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife C |  |  | August 01, 2006 | 2006 | Elekta, Inc. | Due to a material change in a critical component of the linear actuators the quality of said part was found to be inadequate. | Other | Phase 1 - Technical Notes Leksell Gamma Knife C version 1.2 (dated August 2006) "Potential Problem with Error Handling" was sent to all users. Phase 2 - Mandatory Field Action 0000748, "Helmet Changer Actuator, MTB and MCU SW Upgrade Kit was released December 3, 2007. This Field Change Order includes two mod kits and a software upgrade to provide a permanent solution to the issues. Phase 3 - Technical Note 200 068, "Safety issue with helmet host actuator", released on May 15, 2008. This Technical Note includes a warning to customers of a potential quality problem with the helmet hoist actuator that is installed in the Leksell Gamma Knife. Phase 4 - Mandatory Field Action 200 067, "Investigation and Correction of Helmet Changer Actuators", released on May 15, 2008. This FCO includes a list of user machines which must be monitored for quality issues and provides a solution for replacement of parts should a deficiency be detected. Phase 5 - Mandatory Field Action 200 075, "Helmet Changer Actuator, MTB and MCU SW upgrade, release two", released on September 16, 2008. This FCO applies to machines within specific serial number ranges. Due to a material change in FCO-0000748 and baselines (conducted without notice by the manufacturer) the quality of said part was found to be inadequate. After reversing the material change, FCO 200 075 is released to update the installed base with the new type of actuator and to finish the remaining sites never or partially updated by the now superseded FCO-0000748.For further information, contact Elekta R&D; Jonas Hallstensson onphone no. +46 709 843 553 or via e-mail jonas.hallstensson AT elekta.com. |  | 51 | 51 units | AL, CA, CO, CT, FL, GA, IL, MD, MA, MI, MN, MO, NE, NY, NC, OH, OK, OR, PA, TN, TX, VA, WA | Terminated | December 14, 2010 | 1596 | 51 | Not Computer
Z-1431-2010 | 54975 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Unilect ECG Electrodes |  |  | March 11, 2010 | 2010 | ConvaTec | In an emergency situation where defibrillation is necessary and electrodes are already in use, the electrodes may malfunction and incorrectly report lack of cardiac activity. | Process design | The one distributor was notified personally on March 9, 2010. All remaining stock was put on hold, and a "URGENT : MEDICAL DEVICE RECALL" letter dated March 11, 2010, was sent to customers. The letter describes the product, problem and action to be taken by customers. All customers were asked to immediately stop distributing and quarantine all recalled lots. The customers were ask to perform a count of recalled product currently in inventory, complete and return the Recall Response Form even if no recalled product is in inventory, send a copy of response to distributor in order to receive reimbursement, and to call 1-800-582-6514 for instructions on how to return the recalled product.If you have any questions, please contact Convatec at 908-904-2425. |  | 567000 | 567,000 units | Nationwide distribution: CA | Terminated | June 21, 2011 | 467 | 567000 | Not_Computer
Z-1411-2010 | 55099 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI |  |  | September 19, 2007 | 2007 | Elekta, Inc. | Firm became aware of a failure in the position detection system which could result in incorrect positioning of the detector in the longitudinal or lateral direction. | Device Design | The following Important Notices were sent to customers: 1) Important Notice A305-Possible Synergy XVI panel position error (dated September 19, 2007) was sent to customers with versions R3.5, R4.0 and R4.1., 2) Important Notice A306-Possible Synergy XVI panel positioning error (dated November 8, 2007) was released to encompass customers with version R4.2., and 3) Important Notice A328 dated March 10, 2009, and sent again on April 29, 2009, included a modification kit to address the original issue and an expanded range of serial numbers affected. This Important notice A328 entitled, Elekta Synergy kV Detector Position Readback Modification Kit supersedes A305 and A306 for customers who have serial numbers within the range specified. The letter describes the product and the problem and advises the customer to follow any instructions or recommendations covered in this Notice.If you have any questions, please contact Elekta at 770-670-2548. |  | 166 | 166 units | Worldwide distribution: USA and Puerto Rico, Canada and Mexico. | Terminated | August 09, 2012 | 1786 | 166 | Not_Computer
Z-1432-2010 | 55103 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Precise Treatment Table |  |  | January 11, 2007 | 2007 | Elekta, Inc. | In some instances, the bolts which are supplied with the table have sheared off, resulting in inability to operate the powered table. | Other | Mandatory Field Change Order (FCO) 528 053 "Check for zinc plated Precise Table 'Z' mechanism fixation bolts", dated 1/11/2007 was distributed to all affected customers. The Change Order identified the product, stated the reason for the FCO, and the procedure to check and resolve the problem identified. |  | 185 | 185 units | Worldwide Distribution -- USA, Puerto Rico, Argentina, Australia, Austria, Bangladesh, Belgium, Canada, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hong Kong, India, Indonesia, Iraq, Israel, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Pakistan, Peru, Poland, Russia, Serbia, Spain, Sudan, Sweden, Taiwan, Thailand, Turkey &amp; United Kingdom. | Terminated | December 28, 2010 | 1447 | 185 | Not_Computer
Z-1427-2010 | 54562 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | InterGard Knitted Collagen Coated Vascular prostheses, Catalog number: IGK1206. InterVascular SAS, France. |  |  | December 09, 2009 | 2009 | Intervascular S.A.S. | Packaging was mislabeled. | Labeling mix-ups | InterVascular SAS issued a "Recall Notification - Product Recall RC 020" notification dated January 29, 2009 to the sole consignee. The consignee had previously been notified by phone and all affected product had been segregated.For further information, contact InterVascular SAS at +33 442 08 77 84. |  | 9 | 9 | Nationwide Distribution - Arizona. | Terminated | April 27, 2010 | 139 | 9 | Not_Computer
Z-1437-2010 | 55252 | No | GWE | Evoked response photic stimulator. | Neurology | Neurology | 510(k) | Diopsys Inc., Enfant Pediatric VEP Vision Testing System, |  |  | April 02, 2007 | 2007 | Diopsys, Inc. | The computer based system may lose test counts when the time in the device was changed during a changeover from standard time to daylight savings time in March 2007. | Software design | Field Service personnel were sent an e-mail notification on 4/2/2007 which included a TZedit file to be downloaded. E-mails instructed reps to install the patch on their customers devices immediately. Field correction was completed 7/31/2007. |  | 175 | 175 units | Nationwide distribution. | Terminated | April 27, 2010 | 1121 | 175 | Software | Device Operation | Software update | Software Update
Z-1436-2010 | 55064 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iView GT Electronic Portal Imaging Device, used with radiation therapy treatments. |  |  | July 31, 2009 | 2009 | Elekta, Inc. | Image problems that have an effect on patient position corrections and on the accuracy of the radiation treatment. | Software design | Elekta sent IMPORTANT NOTICE number A329, dated 31 July, 2009 to all customers. This corresponds with Field Change Order - Documentation 764 00 551 060 "An Important Notice (A329) has been issued for replacement of iViewGT R3.4 Service Pack 2 (SP2) installation CD". Mandatory Field Change Order 764 00 551 059 "Replacement of iViewGT R3.4 installer disk"supersedes FCO 764 00 551 053. |  | 121 | 121 lots | Nationwide &amp; Puerto Rico | Terminated | December 14, 2010 | 501 | 121 | Software | Display/Image | Instructions | Safety Notice/Insructions
Z-1433-2010 | 55171 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKNEE SYSTEM, 3D Knee Tibial Insert-DCM, Sterile, Right, size 4 x 11mm. Encore Medical, LP. Austin, TX 78758-5445. |  |  | March 19, 2010 | 2010 | Encore Medical, Lp | Affected product had incorrect color label which is a secondary visual indicator for sizing. | Packaging process control | Encore Medical, LP notified consignees by e-mail dated March 19, 2010. Consignees were requested to return the affected product to the firm.For further information, contact Encore Medical, LP at 1-888-838-3668. |  | 7 | 7 units | Worldwide Distribution -- United States (UT, IA) and Germany. | Terminated | July 22, 2010 | 125 | 7 | Not_Computer
Z-1438-2010 | 55180 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 2.0/2.4 Cannulated Screw Driver Bit, part #210-24-003 |  |  | April 14, 2009 | 2009 | Trilliant Surgical LTD | Firm received several complaints of field failures, e.g., "broke during surgery" and the firm initiated a design change. A communication to sales staff 04/14/2009 instructed them to recover and reconcile these products from user physicians. | Other | Firm notified sales force by e-mail on 4/14/2009 of the need to recover and reconcile this product . Additional documentation is not available from the firm. Questions may be directed directly to the firm at 800-495-2919. |  | 32 | 32 units | WI, TX, MO, NM, MN, OK, MA, CA, GA, ND, NE, IA, NY, LA, and MD | Terminated | May 03, 2010 | 384 | 32 | Not_Computer
Z-1440-2010 | 54686 | No | DXY | Implantable pacemaker pulse generator. | Cardiovascular | Cardiovascular | PMA | Medtronic EnRhythm, P1501DR, Dual Chamber Rate Responsive Pacemaker with RapidRead Telemetry. The EnRhythm pacing system is an implantable medical device that monitors, detects, and treats atrial tachyarrhythmia episodes. It also provides bradycardia pacing and monitoring of ventricular tachycardia (VT) episodes. Manufactured at: Juncos, Puerto Rico, USA. Medtronic, Inc., Minneapolis, MN 55432 USA. |  |  | February 11, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic is informing doctors of two specific battery issues with EnRhythm pacemakers that will be addressed by a Medtronic software update available mid-2010. If the software update referenced above is not implemented, there will be a potential risk of loss of device functionality or risk for loss of therapy at or near ERI in a small number of devices. The software will eliminate this issue, fo | Nonconforming Material/Component | Consignees were given a Medtronic "Important : Medical Device Correction" letter addressed to "Dear Doctor" and dated February 2010. In the letter, Medtronic is informing doctors of two specific battery issues with EnRhythm pacemakers that will be addressed by a Medtronic software update available mid-2010. The letter included recommendations including that Medtronic recommends physicians continue to use the ERI notification to determine time for device replacement. At this time, no other action, reprogramming or change in the frequency of follow-up is recommended. For additional information or assistance, please contact your local Medtronic Representative or Medtronic Technical Services at 800-505-4636. |  | 120061 | approximately 120,061 worldwide | All States in the USA including Guam, Puerto Rico and DC.OUS: Customers in the following countries: Andorra, Argentina, Aruba, Australia, Austria, Bangladesh, Belarus, Belgium, Brazil, Brunei, Darussalam, Canada, Cayman Islands, Chile, China, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guam, Hong Kong, Hungary, India, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kenya, Kuwait, Latvia, Lebanon, Luxembourg, Malaysia, Malta, Mexico, Morocco, Netherlands, New Caledonia, New Zealand, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Vatican City, Venezuela, Viet Nam, Virgin Islands | Terminated | January 08, 2013 | 1062 | 120061 | Battery | Device Operation | Software update | Software Update
Z-1424-2010 | 53608 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Cardiac Science, Automated External Defibrillator (AED) under the following brand names (with respective model numbers): Powerheart AED G3: 9300A, 9300C, 9300D, 9300E, 9300P, 9390A/E; CardioVive: 92532 and 92533; Nihon Kohden: 9200G, and 9231; and Responder: 2019198 and 2023440. Cardiac Science, Bothell, WA 98021. |  |  | October 12, 2009 | 2009 | Cardiac Science Corporation | Certain models of Powerheart G3 AEDs may not have received adequate electrical safety testing prior to being shipped from the factory. Affected models are: Powerheart 9300A, 9300C, 9300D, 9300E, 9300P, 9390A, 9390E, CardioVive 92532, 92533, NK 9200G, 9231, and Responder 2019198 and 2023440. The affected devices were manufactured between October 29, 2003 and February 10, 2009. | Process design | Cardiac Science issued "Urgent - Medical Device Corrective Action" notices dated October 2009. Customers were instructed to identify the affected product and arrange for product replacement with the firm. Where possible, end user information was requested for further follow-up by the firm.For further information, contact Cardiac Science at 1-877-402-4599 or by email at AED155@cardiacscience. com. |  | 1171 | 1171 (US: 569, OUS: 602) | Worldwide Distribution -- Distributed to the United States, Australia, Belgium, Canada, Chile, China, Denmark, Egypt, France, Germany, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Korea, Malaysia, Marshall Islands, Nigeria, Norway, Philippines, Poland, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Taiwan, and the United Kingdom. | Terminated | February 13, 2012 | 854 | 1171 | Hardware |  | Replace | Remove or Replace
Z-1441-2010 | 55098 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy |  |  | March 18, 2008 | 2008 | Elekta, Inc. | Installation of a touchguard to prevent injury to patients or staff by stopping machine movements in the case of accidental collision. | Finished device change control | An Important Notice A319 "Instructions for Use for Elekta Synergy kV Source Touchguard" dated 3/18/08 is in distribution to all affected customers. This letter identified the product and asked customers to file the Notice in the User Notice section of the appropriate User Manual. It also asked them to follow the instructions and the reason for the Notice. Questions should be directed to a local Elekta representative. Field Change Order (FCO) 514 259 "kV Source touchguard for systems with touchguard ready source covers" dated 10/17/08 was sent to all affected customers. FCO 514 260 "kV Source touchguard for systems with the original source covers" dated 10/17/08 was sent to all affected customers. |  | 184 | 184 units | Nationwide Distribution -- Including Puerto Rico. | Terminated | December 14, 2010 | 1001 | 184 | Not_Computer
Z-1443-2010 | 55049 | No | HWB | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Literature for Implant Extraction Set - Implant Extraction Guide Module One &amp; Two. |  |  | March 02, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Literature did not include the design change to a smaller cannulation diameter. | Labeling Change Control | An Urgent Correction - Operative Technique letter was sent to direct accounts, Hospital Risk Management, Chief of Orthopaedics and Surgeons that use the Gamma3 system. Notification letters were sent via Federal Express on March 2, 2010. The letter identified the product and described the issue. It also discussed the potential hazards and risk mitigation. Customers are to contact Customer Service at 1-866-OR-ASSIST or any of the other firm representatives listed in the letter. |  | 790 | 790 unit | Nationwide Distribution | Terminated | August 06, 2012 | 888 | 790 | Not_Computer
Z-1134-2010 | 54631 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell GammaPlan. |  |  | October 23, 2008 | 2008 | Elekta, Inc. | Although the co-registration looks good during the verification step in the co-registration dialog, the obtain transformation may include an error that depends on the voxel sizes and acquisition parameters of the co-registered image studies. | Other | Important Notice A325 "Error in ImageMerge" dated November 21, 2008 was sent to affected customers and field service engineer as notification of the problem. The notice advises users not to use ImageMerge except if it is judged that the transformation error has no clinical significance. Information is provided in the release to help the user to make an informed decision on this issue. Field Change Order 200 084, "NTPS 8.3 SP1 for NTPS 8.0, 8.2 and 8.3" released January 28, 2009 was sent to affected customers as notification of the fix to this issue. This notice includes instructions to assist Elekta field service personnel in successfully upgrading the affected systems. Once the upgrade is in place, the issue will be corrected. |  | 72 | 72 units | Nationwide Distribution -- (AL, AZ, AR, CA, CO, FL, GA, HI, IL,KY, MA, MI, MS, NE, NV, NV, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN TX and UT). | Terminated | December 14, 2010 | 782 | 72 | Not_Computer
Z-1442-2010 | 54995 | No | HGX | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Instruction for Use for Symphony Double Pumping Kit; Medela, Inc., P.O. Box 660, 1101 Corporate Drive, McHenry, IL 60051-0660; the instructions were packaged with the following non-sterile single user accessory kits for use with Symphony breast pumps: |  |  | March 15, 2010 | 2010 | Medela Inc | The non-sterile devices were erroneously packaged with instructions stating the devices are sterile and do not require cleaning before first use. | Labeling Change Control | An "IMPORTANT:DEVICE RECALL" letter dated March 12, 2010 to the the customers on March 15, 2010, via certified mail, return receipt requested. The letters were tailored to each customer listing the affected shipments sent to them, and informing them of the labeling insert problem. The customers were requested to cease distribution of the affected kits and return them to the firm (Medela) for replacement, and if applicable, to notify their retail customers. The customers were instructed to call Medela Customer Service at 877-320-2301 for assistance in processing returns or for further clarification on the recall. |  | 13834 | 13,834 kits | Nationwide distribution | Terminated | December 10, 2010 | 270 | 13834 | Not_Computer
Z-1434-2010 | 55182 | No | JDW | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Orthofix US XCaliber Meta-Diaphyseal Complete Kit, Sterile. 99-91600, manufactured by Orthofix SRL, Via delle Nazioni, 9, 27012 Bussolengo (VR) Italy. |  |  | February 22, 2010 | 2010 | Orthofix Srl | During standard quality controls, firm identified a potential for breakage of the XCaliber clamp when locking the cam. | Nonconforming Material/Component | Orthofix issued an "Urgent: Field Safety Notice" dated February 22, 2010. Consignees were informed of the affected product and instructed to cease use and to return the product to the firm. |  | 9 | 9 | Nationwide Distribution -- TX, ND, NY, FL and CA. | Terminated | May 05, 2010 | 72 | 34 | Not_Computer
Z-1439-2010 | 54889 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | InterGard Knitted Collagen Coated Vascular prostheses marketed under 510(k) # K964625; |  |  | December 09, 2009 | 2009 | Intervascular S.A.S. | The external lid of Intervascular's prosthesis has a cut in it which may compromise package integrity and therefore sterility. | Other | The firm sent recall letters to all customers on December 9, 12009. Calls had been made prior to the letter being sent. Because all product was recovered, no additional questions are anticipated. |  | 5 | 5 | US Distribution went to one distributor in Arizona. Also Chile, France, Germany, Hong Kong, Poland and Turkey. | Terminated | April 29, 2010 | 141 | 5 | Not_Computer
Z-1531-2010 | 52777 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Violet Ink sterilization indicators Violet Ink (Class 6) and Violet Ink (Bowie Dick). Product numbers: CI 144, BD 115, PL 315, PL319, PL232, PL236, PL237, PL 238. |  |  | July 15, 2009 | 2009 | Steritec Products Mfc. Co., Inc. | This action notifies affected SteriTec customers as well as private label distributors of the necessity to modify the shelf life of their sterilization indicator products. There is insufficient data on file to support the shelf life dating on the product labeling. | Other | Notify accounts by e-mail and letter, offering to exchange product for product with revised expiration date. Products exhausted by proper expiration date need not be returned or relabeled.Questions are directed to the firm at (303) 660-4201. |  | 2656 | 2656 cases | Nationwide distribution: IL, MI, NY, OH, VA | Terminated | June 28, 2010 | 348 | 2719 | Not_Computer
Z-1528-2010 | 55016 | No | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Baxter RenAmin (Amino Acid) Injection; Pharmacy Bulk Package Not for Direct Infusion, Rx Only, 250 mL type II glass bottles; Baxter Healthcare Corporation, Clintec Nutrition Division, Deerfield, IL 60015 USA, Made in USA; product code 2A6222. |  |  | February 15, 2010 | 2010 | Baxter Healthcare Corp. | Routine stability testing of the evacuated container product, along with trending of results, indicates that the pH will exceed the upper specification limit prior to the expiration date of this product. Baxter decided to remove all potentially affected products even though there have been no out-of-specification pH results at the 0, 6 and 12 month intervals. The recall was initiated on 2/15/10. | Nonconforming Material/Component | Urgent Product Recall letters dated 2/15/10 were mailed on the same date via first class mail to the affected direct accounts to the attention of the Director of Pharmacy, the Director of Materials Management, and the Director of Risk Management. The accounts were informed of the potential for the pH of the products to exceed the upper specification limits for pH prior to the product expiration date. The accounts were requested discontinue use and segregate the affected lots of products from inventory and call Baxter Healthcare Center for Service at 1-888-229-0001, Monday through Friday, 7:00 AM to 6:00 PM CT, to arrange for the return of the affected product for credit.The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter and the amount of each product being returned. If the accounts further distributed the products, they were requested to forward the communication to those accounts. Questions regarding the communication are directed to the Medical Information Services at Baxter at 1-800-933-0303. |  | 4164 | 4,164 units | United States -- Nationwide and Puerto Rico, Canada and Saudi Arabia | Terminated | January 29, 2012 | 713 | 148032 | Not_Computer
Z-1525-2010 | 55175 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | PowerPort isp M.R.I Implanted Port without Suture Plugs with attachable 6F Chronoflex Polyurethane Open-ended Single -Lumen Venous Catheter, BARD Access Systems Assembled in Mexico |  |  | March 18, 2010 | 2010 | Bard Access Systems | 8F PowerPort stems were mixed with 6F PowerPorts and packaged. | Mixed-up of materials/components | The firm, Bard Access Systems, Inc. (BAS) sent a "URGENT: Powerport isp M.R.I with 6F Chronoflex Catheter RECALL" letter dated March 18, 2010, to the customers. The letter describes the product, problem and actions to be taken by customers. The customers were instructed to immediately examine their inventory and quarantine all products subject to this recall and notify any customers they may have distributed the product to; DO NOT USE THIS PRODUCT OR FURTHER DISTRIBUTE ANY OF THE IMPLICATED LOTS; remove and return the affected lots to BAS, and complete and Fax the Reply Form and Inventory Reconciliation Form, if unable to Fax call and report information. Only unused products, in its original packaging, should be returned to BAS.Should you have any questions or require assistance in this matter, please contact BAS Customer Service by calling our toll-free number 1-800-290-1689. |  | 200 | 200 units total (100 units of each lot, see code information) | Nationwide distribution: AZ, CA, DC, FL, MA, MD, NC, NE, NV, NY, PA and SC. | Terminated | October 01, 2010 | 197 | 200 | Not_Computer
Z-1533-2010 | 54850 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife C |  |  | September 01, 2006 | 2006 | Elekta, Inc. | Leksell Gamma Knife C 1.2 with MCU version 4.0.0.6 improperly adjusted trolley for 4mm helmet caused jamming when helmet changer attempted to lock helmet. Lock did not fully engage and could still be raised. When the lock does not engage, then when the helmet changer rises, the helmet could disengage from the helmet changer and fall back onto the helmet trolley. | Other | Phase 1 - Field Change Order FCO-DOC-06-0003 "Technical Note - Potential risk of dropping the collimator helmet" (dated Aug. 28, 2006) was sent to affected customers as notification of the problem. Phase 2 - Technical Notes Leksell Gamma Knife C version 1.2 (dated July 2006) "Important Notice - Potential risk of dropping the collimator helmet" was sent to all Elekta Service Engineers as notification of the problem.The firm planned to create an upgrade kit for the effected Leksell Gamma Knife C1.2 systems. A service representative from Elekta will contact each site to decide on a suitable implementation schedule. |  | 88 | 88 units | Nationwide, Austria, Belgium, Canada, China, Croatia, Italy, Japan, Korea, Norway, Spain, Sweden, Taiwan and Turkey. | Terminated | December 14, 2010 | 1565 | 88 | Computer
Z-0820-2010 | 52635 | No | MDS |  |  | Clinical Chemistry | PMA | Freestyle Navigator Continuous Glucose Monitor System, ART 16467 Rev A, ABB SYSTEM KIT(ORANGE)PACK NAV, Abbott Item Number 70684-01, distributed by Abbott Diabetes Care Inc, Alameda, CA. |  |  | April 14, 2009 | 2009 | Abbott Diabetes Care, Inc. | Potential failure or inaccurate readings--the device's plastic housing near the battery compartment may crack, and allow moisture to enter. | Process control | Notification letters were issued to FreeStyle Navigator customers via certified mail and notification was posted on the Abbott Diabetes Care's Website. Customers were notified of the potential issue and actions to take: 1) if moisture is observed in the FreeStyle Navigator transmitter now or at any time, discontinue use of the transmitter and call Abbott Diabetes Care Customer Care at 1-800-418-6427 and 2) if a loss of connection is encountered or if your results are not reliable or are inconsistent with how you feel, consult your User's Guide for troubleshooting guidelines. If the issue is not resolved, call Abbott Diabetes Care Customer Care at 1-800-418-6421. For outside-of-US accounts, the firm's regional offices will distribute notices to identified customers in the appropriate language. |  | 5449 | 5,449 of total units | Worldwide Distribution -- United States (OH, FL, NJ, CA, MS, MN, MA, and NY), France, Australia, the Netherlands, Sweden, the UK, Israel, Ireland and Germany. | Terminated | March 31, 2011 | 716 | 5449 | Not_Computer
Z-1535-2010 | 53836 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista CHEM 3 calibrator |  |  | October 23, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | patient results are depressed | Other | An "Urgent Field Safety Notice" dated October 2009 was issued by Siemens to all customers. The notice described the product, problem and actions to be taken by the customers. The customers were instructed to recalibrate the AMON method (AMON Flex reagent cartridge, K3019) with Vista Chem 3 Claibrator lot 9ED058 using the revised bottle values provided in the notice and to please consult with their laboratory director regarding re-evaluation of previous results. The customers should forward this notification to anyone to whom they may have distributed this product.If you have any technical questions regarding this information, please contact Siemens Technical Solutions Center at 1-800-441-9250. |  | 1753 | 1753 cartons | Worldwide distribution: USA and Australia, Canada, Japan, New Zealand, South Korea, and Switzerland. | Terminated | June 22, 2010 | 242 | 1753 | Not_Computer
Z-1534-2010 | 54264 | No | IWE | Light beam patient position indicator. | Radiology | Radiology | 510(K) Exempt | Hitachi Oasis Magnetic Resonance Scanners Software V1.0C |  |  | September 01, 2009 | 2009 | Hitachi Medical Systems America Inc | The device has a software error that could potentially lead to a misdiagnosis. | Software design (manufacturing process) | Sites have already been corrected. Software updated from V1.0C to V2.0B |  | 43 | 43 | Nationwide distribution: AZ, CA, CT, DE, FL, GA, IA, IL, IN, LA, MD, MI, MO, NC, NY, OH, OK, OR, PA, RI, SC, TX, VA, WA, WI, and WY. | Terminated | February 24, 2012 | 906 | 43 | Software | Output/Calculation | Software update | Software Update
Z-1536-2010 | 55251 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bard Ventrio Small Oval Hernia Patch 8cm X 12cm (3.1" x 4.7"), Warwick, RI |  |  | April 08, 2010 | 2010 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: Product labeled as Small Oval (3.1" x 4.7") may contain Small Circle (3.0") | Labeling mix-ups | An "URGENT RECALL: BARD VENTRIO SMALL OVAL HERNIA PATCH (LOT DATB0015" letter dated April 8, 2010, was issued by the firm (Davol) to all customers via FedEx overnight with signature required. The letter describes the product, problem and action to be taken by the customers. The customers were instructed to: please examine their inventory and immediately discontinue use of this lot; contact the Davol Customer Service Department at 1-800-556-6275 for instructions on how to return product to firm; complete and fax the EFFECTIVENESS CHECK sheet to 1-401-825-8753, and forward a copy of this letter, if they are a distributor, to their customers who have received the product and lot number. All products returned will be held at the Davol, Warwick location.Please contact your Davol Territory Manager or C.R. Bard's Medical Service Support Department at 1-800-562-0027 if you have any other questions. |  | 198 | 198 units | Nationwide distribution | Terminated | October 27, 2011 | 567 | 198 | Not_Computer
Z-1524-2010 | 55191 | No | NEG |  |  | Orthopedic | PMA | Ascension PyroCarbon PIP Total Joint - Trials; Ascension PIP, sizes 10P, 20P, 30P and 40P. Ascension Orthopedics, 8700 Cameron Road, Suite 100, Austin, Texas 78754; TEL: 512 836-5001. |  |  | June 17, 2009 | 2009 | Ascension Orthopedics, Inc | Trial product was larger than actual implant. Trial product was re-designed to more closely match the implant. | Device Design | Ascension Orthopedics, Inc. issued e-mail notification to distributors and sales representatives beginning June 2009. Consignees were instructed return affected product to the firm for replacement.For further information, contact Ascension Orthopedics, Inc. at 1-512-836-5001 extension 1549. |  | 129 | 129 units (64 US, 65 OUS) | Worldwide Distribution - Worldwide Distribution - United States, United Kingdom, France, Australia, Belgium, Czech Republic, Canada, Germany, The Netherlands and South Africa. | Terminated | September 27, 2010 | 467 | 129 | Not_Computer
Z-1537-2010 | 55169 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Medtronic HMS PLUS, Hemostasis Management System. Medtronic, Inc., Cardiovascular Division, Moundsview, MN 55112. |  |  | March 08, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | The heparin assay controls may run longer than normal and in some cases, fail to give an acceptable result. | Process change control | Medtronic issued an "Urgent Medical Device Notification" letter dated March 8, 2010 addressed to "Dear Valued Customer" and was sent to three locations at each customer site: Risk Management, Cardiovascular Surgery and Clinical Laboratory-Point of Care. The letter described the product and the issue, provided instructions for control and patient testing. The customer was requested to return an enclosed certificate as proof of notification.For further information, contact your Medtronic Sales Representative or call 1-763-391-9915. |  | 544 | 508 US, 36 OUS | Worldwide Distribution -- United States, International West (countries included in Western Europe and the Middle East), Australia and Canada. | Terminated | December 31, 2011 | 663 | 544 | Not_Computer
Z-1538-2010 | 54091 | No | KGZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Percutaneous Sheath Introducer (PSI) Kit with Integral Hemostasis Valve/Side Port for use with 7.5 - 8 Fr. Catheters, Arrow International, Inc., 2400 Bernville Road, Reading, PA 19605 |  |  | December 15, 2009 | 2009 | Arrow International Inc | Some central venous catheter sets had the incorrect lid stock. | Packaging | Consignees were notified by an Urgent Medical Device Recall letter on/about 12/15/2009. The letter identified the affected product and stated the reason for the recall. Customers were instructed to immediately check their stock and cease use, distribution, and quarantine affected product. They are also to contact Arrow's Customer Service Department at 800-523-8446 for a Return Authorization Number. Customers are to return any affected product along with the Advisory Acknowledgement and Stock Status Form. Arrow will either provide replacement with a similar product or credit customers' accounts. Customer Service can be contacted with questions or for clarification. |  | 295 | 295 kits | Nationwide Distribution -- MA. | Terminated | July 29, 2011 | 591 | 295 | Not_Computer
Z-1221-2010 | 54165 | No | KGZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Shutter HX SP5, a component of the Leica Point Scanning Confocal Microscope Models TCS SP5 and TCS SP5 X. |  |  | November 12, 2009 | 2009 | Leica Microsystems, Inc. | A manufacturing defect was identified that could potentially result in the laser shutter not closing as intended which could result in the shutter not completely closing disrupting the laser beam safely. | Radiation Control for Health and Safety Act | A Service Note was sent to service organizations on November 12, 2009. The communication asked customers to check if they must exchange their shutters and send their planned exchange date by 2009-11-20 to clsm.qc@leica-microsystems.com. Customers were also asked to check the serial numbers of the one-wire chips of the shutters. Then lastly, the were asked to exchange the shutters and confirm the exchange. |  | 23 | 15 microscopes and 8 spare part shutters | Worldwide Distribution -- USA, Sweden, Japan, Hong Kong, Italy, the Netherlands, Singapore, UK, and Spain. | Terminated | May 11, 2010 | 180 | 23 | Not_Computer
Z-1540-2010 | 55286 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AEM Disposable Electrode, Spatula, 45cm, REF ES0301-45, Qty 10 each Sterile R, Manufactured for: Encision Inc., Boulder, CO 80301, made in Taiwan. |  |  | April 05, 2010 | 2010 | Encision, Inc. | Electrosurgical electrode tip insulation may crack and break off during use. | Nonconforming Material/Component | A "Confirmation of Product Recall" letter dated April 5, 2010, was faxed by the firm (Encision, Inc.) to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to either return or destroy the affected electrodes and to contact any customers or sales representatives to whom they distributed the product and confirm they have been removed from use and/or they will pick them up for return. The customers is to complete form and fax to the attention of RA/QA at 1-303-444-2693 or return with the items. The number to call to arrange return and/or if you have any questions is 1-800-998-0986. |  | 20 | 20 units | Worldwide distribution: USA including states of CA, CT, IA, IN, KS, KY, MI, OK, TX, WV, and countries of New Zealand and The Netherlands. | Terminated | May 21, 2010 | 46 | 1140 | Not_Computer
Z-1517-2010 | 55354 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Shiley Specialized Tracheostomy Tube, M8FEN, Sterile EO, Made in Mexico, Tyco Healthcare Group LP, Nellcor Puritan Bennett Division, Pleasanton, CA. |  |  | April 13, 2010 | 2010 | Nellcor Puritan Bennett Inc. (dba Covidien Ltd) | Inflatable cuff Tracheostomy tubes may leak air, resulting in inadequate ventilation. | Process design | Covidien issued an "Urgent Medical Device Recall" notification dated April 13, 2010. Consignees were alerted to the affected device. If an affected product is already in use, the firm recommends that the tube be replaced as soon as the patient's physician determines it is clinically appropriate. All affected product is to be returned to the firm. For further Information, contact Covidien at 1-800-635-5267. |  | 20 | 20 units | Worldwide Distribution -- United States (including Guam and Puerto Rico), Australia, Belgium, Brazil, Canada, Chile, Colombia, Honduras, Japan, Mexico, Panama, Singapore and Uruguay. | Terminated | January 18, 2012 | 645 | 559545 | Not_Computer
Z-1550-2010 | 55179 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Inion 1.5 x 6mm CPS(R) Baby Screw, 1.5mm Inion CPS(R) Baby 5+E, REF SCR 1221, 5/pack, STERILE/R Stryker Craniomaxillofacial, Portage, MI. |  |  | March 05, 2010 | 2010 | Stryker Craniomaxillofacial Division | During a sampling inspection, 2 lots of screws did not meet shear and torsion testing specifications. | Nonconforming Material/Component | The firm, Stryker, issued two "URGENT PRODUCT RECALL NOTIFICATION" letters, dated March 5, 2010 and March 19, 2010, to customers. The letters describe the product, problem and action to be taken by the customers. The customers were instructed to examine their inventory for the catalogue and lot numbers, return the enclosed confirmation indicating the location of the affected product, return any remaining product to Stryker at the noted address, and provide a copy of this notification to any site that they may have distributed the product to.If you have any further questions, contact Rob Yamashita, Sr. Regulatory Affairs Rep at (269) 323-4258 or email: rob.yamashita@stryker.com |  | 184 | 184 packs (5/pack) | Nationwide distribution: CA, CO, CT, FL, GA, IL, LA, MA, MD, ME, MO, MT, NC, NH, NJ, NM, NY, OH, PA, RI, SC, TX and WI. | Terminated | June 06, 2012 | 824 | 184 | Not_Computer
Z-1547-2010 | 55115 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ZOLL E Series Defibrillator/Pacemaker/Monitors, BLS Model, with software versions 4.XX, 6.XX and 7.0X. Zoll Medical Corporation, Chelmsford, MA. |  |  | March 24, 2010 | 2010 | ZOLL Medical Corporation, World Wide Headquarters | Device issued Shock Advised message but failed to auto-charge the defibrillator. | Software design | Zoll Medical Corporation issued an "Urgent Medical Device Correction" notification dated March 26, 2010 via certified mail to consignees who have received affected product. Customers will be contacted and given the option to have a software upgrade kit to update their devices or return devices to the firm to have the device upgraded.For further information, contact Zoll Medical Corporation 24/7 Technical Support at 1-800-348-9011 or 1-978-421-9460. |  | 1383 | 1383 units | Worldwide Distribution -- United States, Australia, Germany, Libyan Arab Jamahiriya, Malaysia, Singapore, United Arab Emirates and the United Kingdom. | Terminated | January 23, 2014 | 1401 | 1383 | Software | Device Operation | Software update/Replace | Software Update
Z-1548-2010 | 55147 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 7.5mm Ti Solid Humeral Nail 230 mm Sterile Catalog number 462.723S |  |  | February 25, 2010 | 2010 | Synthes USA (HQ), Inc. | potential compatibility issue | Device Design | The recalling firm, Synthes, issued an "URGENT: MEDICAL DEVICE RECALL" letter dated February 25, 2010, return receipt requested, to all customers. The letter describes the product, problem and action to be taken by customers. The customers were instructed to immediately cease use of the products, examine their inventory for the part/lot combinations and remove the recalled devices from their shelves, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number and return the recalled product with the letter; however, if the customers do not have the identified product , they should complete the attached verification section and return document to Synthes by Fax at 610-719-5120 or scan/email at lewis.lynne@synthes.com. If you should have any questions, please call 800-620-7025 x 5452 or 610-719-5452 or contact your Synthes Trauma Sales Consultant. |  | 1 | 1 unit | Nationwide distribution: LA, MA, ME, NC, PA and VA. | Terminated | October 20, 2014 | 1698 | 6 | Not_Computer
Z-1551-2010 | 54566 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | Peloris Rapid Tissue Processor; Leica Biosystems Melbourne Pty Ltd, Mount Waverly, 3149 VIC, Australia; catalog number 26.0001 and 26.0005 |  |  | February 09, 2010 | 2010 | Leica Microsystems, Inc. | Confusion during manually changing reagents, general users running unvalidated protocols, and wax valve leakage have contributed to tissue damage during processing, leading to rebiopsies. | Other | Leica Microsystems sent Medical Device Correction letters dated 2/9/10 to the affected customers via first class mail on the same date, informing them of the potential for tissue damage during processing and outlining the software upgrade being made to minimize the risk of tissue damage. Customers were informed that their local Leica Microsystems representative will contact them to arrange the software upgrade and to provide training on the use of the new software. The letters also provided the necessary steps required to avoid tissue damage during processing while they are awaiting the software upgrade. The accounts were asked to complete the attached Acknowledgement Form, listing the serial numbers of the units at their facility and fax it back to Leica at 1-847-607-3147.Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance. |  | 574 | 574 units | Worldwide distribution -- U.S. (nationwide), Canada, Australia, Europe and Asia. | Terminated | June 30, 2011 | 506 | 574 | Other | Physical Safety Hazards | Software update | Software Update
Z-1560-2010 | 54239 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | One Touch Ping Insulin Pump using ezManager Max diabetes management software |  |  | December 18, 2009 | 2009 | Animas Corporation | Software malfunction which disallows users to download, view and print information from their pump (including blood glucose, insulin delivery, carbohydrate and pump events). Additionally, users are unable to customize the built-in food database and pump settings using the software. | Software design | The recalling firm telephoned the patients/users to inform them of the problem. A letter dated 12/22/09 was also sent to users. Users will have a replacement unit sent to them so that they will be able to fully utilize the pump and the software. As a courtesy a letter dated 12/22/09 was issued to the health care provider/prescribing physician for the affected users. |  | 21 | 21 pumps | The pumps were provided to patients/users in CA, CT, FL, GA, ID, MA, MO, NJ, OH, OR, PA, TX, and WA | Terminated | June 22, 2010 | 186 | 21 | Software | Display/Image | Replace unit | Remove or Replace
Z-1553-2010 | 55216 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific TELIGEN, Implantable Cardiac Defibrillator, Models E102 and E110. Sterilized using ethylene oxide. Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 66112-5798, USA. The product is indicated for patients who are at risk for sudden cardiac death due to ventricular arrhythmias and may require pacing support. COGNIS devices also provide bi-ventricular pacing. |  |  | March 15, 2010 | 2010 | Boston Scientific CRM Corp | Boston Scientific Corporation has stopped shipment and is retrieving unimplanted devices from US hospital shelves of all its implantable cardioverter defibrillators (ICDs) and cardiac resynchronization therapy defibrillators (CRT-Ds) after determining that some manufacturing process changes were not submitted for approval to the US Food &amp; Drug Administration. | PMA | BSC issues press release dated 15 March 2010. FDA press release posted 18 March 2010. On 15 March 2010, BSC Sales Representatives begin removal of devices from health care facilities stock. Starting 22 March 2010 BSC sent a physician letter dated 22 March 2010 to implanting and following physicians via FedEx. The Urgent Medical Device Recall letter described the issue of why devices were being removed from hospital shelves. The letter did NOT recommend any changes to normal follow-up procedures for patients implanted with an ICD or CRT-D device. The letter included ALL potential US models. To be complete and mitigate the possibility of retrieval errors, the ICD and CRT-D US models that may be on customer's shelves were included in the letter. The devices subject to retrieval, however, were those non-implanted devices with domestic customers.For questions, please contact the local representative (1-800-227-3422) for pick up and return or contact Boston Scientific Technical Services at 1.800.CARDIAC (227.3422). |  | 1330 | 1330 | nationwide within US only (OUS distribution not affected). | Terminated | April 05, 2012 | 752 | 2490 | Not_Computer
Z-1561-2010 | 55117 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | PM1000 Flow Selector - An accessory to an emergency ventilator (resuscitator). The selector is attached to a flowmeter and is used to direct the outlet flow from any standard flowmeter to any one of the three outlets it provides. |  |  | March 05, 2010 | 2010 | Precision Medical, Inc. | The on/off outlet is mislabeled so that when a user/operator turns selector to "on", it actually may be turned "off" | Labeling mix-ups | The recalling firm issued Customer Alert Product Recall letters on 3/5/10 to inform users of the problem and how to identify affected product. |  | 19518 | 19,518 | The products were shipped to distributor and medical facilities nationwide. The products were also shipped to Taiwan, Germany, Hong Kong, Canada | Terminated | December 13, 2010 | 283 | 19518 | Not_Computer
Z-1574-2010 | 54083 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics, Scorpio Ceramic 4:1 Guide Size 3; |  |  | December 03, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | The ceramic guide rails, within the XCelerate 4:1 Ceramic Cutting Blocks, may fracture and displace from the block. | Device Design | Urgent Product Recall Letters were sent on December 3, 2009 via Federal Express to Stryker Branches/Agencies, Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. Foreign accounts were notitied on December 7, 2009. Questions or comments should be addressed to Colleen O'Meara at 201-972-2100. |  | 9833 | 9833 Scorpio Cutting Blocks, total. | Worldwide Distribution: USA, Sweden, Ireland, Canada, Hong Kong, Germany, Spain, India, Australia, Italy, New Zealand, and United Kingdom. | Terminated | September 14, 2012 | 1016 | 9833 | Not_Computer
Z-1590-2010 | 54683 | No | LMB | Medical image storage device. | Radiology | Radiology | 510(K) Exempt | An Imaging device for medical procedures, to provide surgeons or operating room personnel the ability to capture still images and to record video during the surgical procedures in various file formats. Images can be transferred to external media. |  |  | October 23, 2009 | 2009 | Stryker Endoscopy | Video device is not able to assign the correct time/date stamp. | Device Design | Stryker Endoscopy has notified the 2865 domestic consignees affected by the device correction. The notification was mailed to affected consignees and their corresponding sales representatives. It included two mailings of the Urgent: Device Correction notice, SDC HD 3.0z correction software, and a SDC HD upgrade dongle. 100% of consignees have been contacted and notified of the device correction needed. The effectiveness of the correction has been tracked via acknowledgement cards mailed, faxed, or emailed. The local Stryker Sales Representative will install a software upgrade to fix this issue. Additional questions should be directed to the firm's Customer Service at (800)624-4422 or (408)754-2039. |  | 12896 | 12896 devices | US government facilities and other US domestic consignees(nationwide), and to foreign countries | Terminated | February 17, 2011 | 482 | 12896 | Software | Output/Calculation | Software update | Software Update
Z-1552-2010 | 49383 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Intramedullary Fixation Rod, common name Fibula Rod, labeled IM Rod System Manufactured or Distributed by: ACUMED Hillsboro, OR |  |  | August 21, 2008 | 2008 | Acumed LLC | Products improperly handled after cleaning, prior to sanitation and may be contaminated with inorganic material. | Other | On 8/21/2008 the firm initiated telephone contact with consignees, followed by an e-mail, and a letter. Consignees were instructed to return product to the firm. Consignees were provided with a response form to complete and return to the firm. For additional information, please contact the firm directly at telephone 888-627-9957. |  | 36 | 36 devices | Product was distributed to 6 consignees, 1 consignee located each in MA, CA, NM, TX, and the United Kingdom and South Africa. | Terminated | May 14, 2010 | 631 | 36 | Not_Computer
Z-1562-2010 | 55189 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ConMed Goldline Hand Controlled With Needle Electrode, Rocker Switch, Safety Holster and 10 FT Cable, REF 130301A, Sterile R, Assembled in Mexico, ConMed Corporation, 525 French Rd., Utica, NY 13502. |  |  | March 29, 2010 | 2010 | ConMed Electrosurgery | Electrosurgical pencil may remain activated or self activate after power switch is released. | Device Design | Customers were notified by letter on March 25, 2010, and were instructed to return all affected devices and to forward the recall information to anyone that may have received the devices to the end user. More information is available by contacting the firm directly at 800-552-0138. |  | 1400 | 1400 units (US) | Nationwide, including 2 VA facilities in GA and 1 military facility in UT. Foreign distribution to Australia, Belgium, Canada, China, Dominican Republic, El Salvador, Japan, Korea, Mexico, Puerto Rico, Saudi Arabia, Spain, Taiwan. | Terminated | September 27, 2010 | 182 | 683440 | Not_Computer
Z-1591-2010 | 55446 | No | LYJ |  |  | Neurology | PMA | VNS TherapyÂż DemipulseÂż Generator, Model 103 and |  |  | April 06, 2010 | 2010 | Cyberonics, Inc | Battery life projection is inaccurate. | Software design | Safety Alert sent April 6, 2010 to all physicians of record for Model 103 or Model 104 generators and to all hospitals with non-implanted affected product, that explains the event, risks presented and recommended actions to take. The generator can be reprogrammed by the physician the next time the patient is seen in the clinic.The issue can also be corrected during the initial implant procedure with the instructions provided.Additional questions should be directed to the firm; United States customers can contact Clinical Technical Support at 866-882-8804. |  | 18722 | Model103: 8109 units; Model 104: 1722 units; Programmers: 8891 | Worldwide distribution. | Terminated | August 20, 2012 | 867 | 18722 | Other | Alarm/Message | Reprogram | Other
Z-1592-2010 | 54680 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | Micro Introducer Kit 5F, Model number: KIT-014-29, &amp; Customer Model number: 7204. Vascular Solutions, 6464 Sycamore Court, Minneapolis, MN 53369. |  |  | February 08, 2010 | 2010 | Galt Medical Corp | Potential problem with pouch seal thus affecting sterile product. | Packaging process control | Firm notified distributor by telephone 2/8/2010. Firm and distributor partnered to notify customers by an Urgent: Medical Device Recall - Lot Specific letter, dated 2/10/2010, beginning 2/16/2010. The letter identified the affected product and the reason for the recall. The letter asked customers to immediately check their inventory remove any affected kits and place them in a secure location. Customers are to complete the Product Recall Inventory Form and fax it to the Customer Service Department. Customers are also to contact Customer Service for a Return Authorization Number at 1-866-240-6001. Affected products will be replaced upon receipt. Questions or concerns should be directed to a local Sales Representative or Jill Eisenzimmer, Product Manager, at 763-656-4300. |  | 400 | 400 units | Nationwide Distribution -- MI, PA, IL, &amp; Washington, D.C. | Terminated | March 21, 2011 | 406 | 400 | Not_Computer
Z-1527-2010 | 55355 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE, 1009-9002-000 Avance Anesthesia Machine and Monitor |  |  | March 12, 2010 | 2010 | Ge Healthcare, Llc | Unanticipated shut-down: A specific lot of GE Healthcare Aisys and Avance machines control board wiring harness have a defect, which can cause the machine to unexpectedly shut down, terminating ventilation, anesthetic delivery, and potentially patient monitoring. | Process design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated March 12, 2010. The letter was addressed to Chief of Anesthesia, Health Care Administrator/Risk Manager, Director of Biomedical/Clinical Engineering.The letter listed the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. |  | 18 | 18 (4 USA, 14 OUS) | Worldwide distribution: USA (CA, FL, TN), CANADA, DENMARK, FINLAND, FRANCE, ITALY, SAUDI ARABIA, SOUTH AFRICA, SPAIN, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | January 30, 2012 | 689 | 33 | Hardware | Device Operation | Correction | Repair
Z-0350-2010 | 52852 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Orthopediatrics IM Femoral Nail - 8mm x 34 cm, Rt, non-sterile, Orthopediatrics, Warsaw, IN; Part # 10-1500-036. |  |  | July 21, 2009 | 2009 | Orthopediatrics Corp | The potential exists for the nails to fracture during insertion of the nail. | Device Design | Distributor consignees were notified by email dated July 28, 2009 and requested to return the recalled products and instrument sets. If there are questions regarding this issue, please contact Gary Barnett at 574-268-6379. |  | 23 | 23 of all products | Nationwide. | Terminated | October 25, 2010 | 461 | 23 | Not_Computer
Z-1595-2010 | 55305 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Basic Core Pack # OR1805, Sterile, Tri State Hospital Supply Corporation Manufacturers of Centurion Healthcare Products. |  |  | April 09, 2010 | 2010 | Centurion Medical Products | The firm is conducting a sub recall of kits containing ConMed Goldline Rocker Switch Pencils. The pencils activation switch may remain in the "on" position after the switch has been depressed with excessive force, thereby continuously activating the pencil. | Nonconforming Material/Component | Centurion sent out an "Urgent Product Recall Notice" to one customer on 4/9/2010. The customer was instructed to discontinue use of the product. The customer was also instructed to return any instock product . |  | 74 | 74 kits | TX | Terminated | October 25, 2010 | 199 | 74 | Not_Computer
Z-1594-2010 | 55213 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics, Accolade Distal Centralizer 13 mm Medium; Catalog No.: 1059-4513. For use with Accolade hip prosthesis. |  |  | March 31, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | One lot of 13 mm medium distal ring centralizers and 14 mm large distal ring centralizer have been labeled incorrectly. The size 14 Large centralizers have been labeled as size 13 Medium centralizers and size 13 Medium centralizers have been labeled as size 14 Large centraliizers. | Process control | Notification letters and Product Accountability Forms have been sent to Stryker Orthopaedics Branch and Agencies via Fed Ex on March 31, 2010 with return receipt.. Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons were sent notificatio letters and Product Recall Acknowledgment Forms on April 1, 2010. with return reciept. |  | 15 | 15 | Nationwide distribution. | Terminated | July 09, 2010 | 100 | 32 | Not_Computer
Z-1598-2010 | 55303 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 ProtÂżgÂż EverFlex Self-Expanding Biliary Stent System, Sterile EO, (REF: PRB35-06-150-080, REF PRB35-06-120-120 and PRB35-06-150-120), 4.5 - 5.5mm, Expiration: 2011-10-01, Rx only. ev3 Inc., 4500 Nathan Lane North, Plymouth, MN 55442-2929. |  |  | March 31, 2010 | 2010 | Ev3, Inc. | ev3 is conducting a voluntary recall on three lots of the ProtÂżgÂż EverFlex" Biliary Stent System. They determined that a lot of 120mm ProtÂżgÂż EverFlex" Biliary Stent system may contain a 150mm length Self-Expanding stent and two lots o 150mm ProtÂżgÂż EverFlex" Biliary Stent System may contain a 120 mm length Self-Expanding stent. | Mixed-up of materials/components | ev3 issued "Medical Device Recall" notifications dated March 31, 2010 addressed to "Dear Risk Manager or Cath Lab Manager". Each letter addressed the affected product. Users were instructed to locate and remove from use the affected product and a sales representative will be in contact shortly to arrange for return.For further information, contact ev3 Customer Service at 1-800-716-6700. |  | 19 | 19 | Nationwide Distribution -- CA, FL, IL, LA, MA, MI, NH, PA, TN, TX, UT, VA and WA. | Terminated | March 08, 2012 | 708 | 19 | Not_Computer
Z-1601-2010 | 55199 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Distal Centralizer, Exeter II Hip System. Catalog number - 0920-2-920. |  |  | March 29, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker became aware that device labels have incorrectly printed manufacture dates and thus incorrect expiration dates on the product label. | Error in labeling | Urgent Product Recall letters and Product Accountability Forms have been sent via Fed Ex on March 29, 2010 with return receipt. Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeon were also sent Notification letters and Product Recall Acknowledgment Form on March 29, 2010 with return receipt. The letters explained the issue, the potential hazards, and risk mitigation. The firm asked customers to examine their inventory and hospital locations to identify the affected products. All affected product should be retrieved and returned to their branch or agency warehouse for reconciliation. The attached Product Accountability form should be used to reconcile all product(s) and it should be faxed back to 201-831-6069. |  | 24 | 24 units | Nationwide Distribution -- FL, MN, &amp; OR. | Terminated | September 10, 2012 | 896 | 24 | Not_Computer
Z-1600-2010 | 55222 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea 5.1. Vitrea fX 3.0, Vitrea Enterprise Suite 1.2 with the licensing of CT lung or CT Colon. |  |  | February 23, 2010 | 2010 | Vital Images, Inc. | Vital images, Inc is issuing a recall notice on their medical device software in the use of CT Lung and CT Colon in Vitrea 5.1, Vitrea fX 3.0 or Vitre Enterprise Suite 1.2. because incorrect measurements may result. Questions or comments are directed to the firm's Customer Support at support@vitalimages.com or at 1-800-208-3005. | Software design | Consignees were sent a Vital "Urgent Medical Device Software Recall Notice" letter dated 23 February 2010. The letter was addressed to "Dear Customer." The letter described the problem, product affected and provided recommendation to their customers. Questions or comments are directed to the firm's Customer Support at support@vitalimages.com or at 1-800-208-3005. |  | 95 | 95 | AZ, AR, CA, COT, CT, FL, GA, IN, IA, KY, LA, MD, MA, MI, MN, NV, NJ, NY, ND, OH, PA, SC, TX, VA, WA, WV, WI.AUSTRALIA, SPAIN, FRANCE, BELGIUM, NETHERLANDS, PERU, NIGERIA, SWEDEN, CANADA, SWITZERLAND, EGYPT, CHINA, GERMANY, PANAMA, TAIWAN, HONG KONG, VENEZUELA, SOUTH AFRICA, CHILE, MEXICO, KOREA, COLOMBIA. | Terminated | June 07, 2012 | 835 | 95 | Software | Output/Calculation | Notification | Safety Notice/Insructions
Z-1597-2010 | 55254 | No | HNN | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | BD Visitec EdgeAhead MVR Knife, BD Opthalmic Systems, Waltham, MA. |  |  | April 06, 2010 | 2010 | Becton Dickinson and Company | A field action was initiated due to a recent customer complaint from Japan there is a labeling mixup and that some of the BD Visitec EdgeAhead MVR Knife 0.90mm (20G), REF 585230, shelf packs may contain BD Visitec EdgeAhead Slit Knife 2.3mm (40 degree), REF 8009918. | Packaging process control | BD Medical issued a "Voluntary Advisory Notice" dated April 6, 2010. Consignees were informed of the affected product and asked to complete the enclosed Customer Response Form indicating whether the product will be returned to the firm.For further information, contact BD Medical Opthalmic Systems at 1-781-906-7952 or via email at ahusebo@bd.com. |  | 1170 | 1170 (US 810, OUS 360) | Worldwide Distribution -- United States, Canada and Japan. | Terminated | August 06, 2012 | 853 | 1170 | Not_Computer
Z-1599-2010 | 55503 | No | LGW |  |  | Neurology | PMA | Eon Mini Implantable Pulse Generator (IPG). Advanced Neuromodulation Systems, Inc. Plano, TX. |  |  | April 23, 2010 | 2010 | Advanced Neuromodulation Systems, Inc | Due to a missed inspection step during manufacturing, a small number of devices may experience a gradual or sudden loss of charging capability. | Manufacturing material removal | St. Jude Medical issued an "Important Medical Device Information" notification dated April 2010. Consignees were informed of the affected product and instructed on necessary steps. Any returned product should be sent to the firm.For further information, contact your St. Jude Medical Neuromodulation Division Sales Representative or call 1-972-309-2154. |  | 7 | 7 devices | Nationwide Distribution -- GA, NC, CA, OK, SC and NJ. | Terminated | July 27, 2015 | 1921 | 7 | Not_Computer
Z-1603-2010 | 53197 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Stryker Endoscopy PrepStrap Shoulder Sleeve, Model Number 3910900200, Manufactured by Stryker Endoscopy, San Jose, CA |  |  | July 27, 2009 | 2009 | Stryker Endoscopy | Latex-- Certain lots of Product contain undeclared latex. | Nonconforming Material/Component | An Urgent: Device Removal notice, dated 7/27/2009 was sent to the sole distributor, Stryker Orthopaedics in New Jersey. The distributor will execute the sub-recall. The firm intends to scrap all returned product at that facility, who will provide a record of disposition. The distributor will notify consignees via a mail packet that will be sent via FedEx overnight mail. Consignees will be instructed to return all product to the distributor, with credit provided per policy. |  | 842 | 842 units | Nationwide Distribution | Terminated | December 08, 2010 | 499 | 842 | Not_Computer
Z-1602-2010 | 55184 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Polarus(R) Humeral Fixation System, size: 3.5mm x 25.0mm Cortical Screw, QTY 1, Ref: HCO3250-S Acumed, LLC Hillsboro, OR |  |  | May 19, 2009 | 2009 | Acumed LLC | Mix-up of parts - Product labeled cortical bone screw actually contains Acutract fixation screw | Employee error | On May 18, 2009, the firm, ACUMED, began calling customers. On May 19, 2009, the ACUMED sent e-mails and a "URGENT NOTICE: DEVICE RECALL" Initial Notice letter to all customers. The letter describes the product, problem and actions to be taken by customers. The customers were instructed to identify the product having the lot code(s) and quarantine them, return all products from the lot code(s) to Acumed immediately, to please notify the customers of this recall if you further distributed this product, and complete and return the Acumed Product Recall Effectiveness Form.Please contact Acumed Customer Service at 1-888-627-9957 for any questions regarding this notification. |  | 83 | 83 | Worldwide distribution: USA including FL, GA, MI, MN, NY, TX, VA and countries of United Kingdom, Herzeles, South Korea, South Africa, China, and Milan | Terminated | May 17, 2010 | 363 | 83 | Not_Computer
Z-1606-2010 | 55317 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | EXXCEL Soft Standard Wall EPTFE Vascular Graft; D:8 mm; L: 50 cm straight. |  |  | March 26, 2010 | 2010 | Maquet Cardiovascular, LLC | This is an expansion of Recall # Z-866-2010, adding an additional two lots to the recall of EXXCEL soft Stand Wall EPTFE Vascular Grafts due to a labeling mix up. | Labeling mix-ups | The firm, MAQUET Cardiovascular, LLC, sent a "DEVICE REMOVAL IMMEDIATE ACTION REQUIRED" letter dated March 23, 2010, to all customers via Federal Express. The letter describes the product, problem and action to be taken by customers. The customers were instructed to examine their stock immediately to determine if they have any of the affected product from the lot, discontinue dispensing the lot and complete and return by fax (973-709-6511) the enclosed Field Action Response Form whether or not they have any affected products.Please contact your local MAQUET cardiovascular sales representative or Customer Service at 1-888-880-2874 for the exchange. |  | 22 | 22 units | Worldwide distribution: USA including states of TX, NJ, SC, MO, ID, TN, FL, CA, PA and Canada and Germany. | Terminated | April 23, 2012 | 759 | 23 | Not_Computer
Z-1605-2010 | 55498 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Universal Flat Panel Yoke, a component of the Stryker Flat Panel Arm. The yoke is intended to hold a flat panel monitor. |  |  | April 07, 2010 | 2010 | Stryker Communications Corp | Some of the pitch shafts were not properly welded which results in the monitor becoming unlevel. | Nonconforming Material/Component | An Urgent: Medical Device Recall letter dated April 19, 2010 was sent to customers. It was recommended that customers continue use of the product unless the monitor becomes unlevel. The customer was notified that a Stryker Representative would contact their facility to coordinate a service call to inspect the affected units. Questions are directed to the firm via phone at 972-410-7617 or via email at |  | 1498 | 1498 | Worldwide --US and Canada, Australia, China (Hong Kong), Latin America (Mexico, Chile, Ecuador, Costa Rica) . | Terminated | March 30, 2011 | 357 | 1498 | Not_Computer
Z-1604-2010 | 54890 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell GammaPlan 8.0 image fusion |  |  | January 07, 2008 | 2008 | Elekta, Inc. | The precision of the calculation used to create the fused study in LGP 8.0 is too low and should not be used until the system is upgraded to LGP 8.2. | Other | Phase 1 - Important Notice and Technical Note #200 058 - "Accuracy problems in Leksell GammaPlan 8.0 image fusion", dated December 10.2007 and Field Change Order #200 058, "Technical Note: Accuracy problems in Leksell Gamma Plan 8.0 image fusion", dated January 7, 2008 were sent to affected customers and all field service engineers as notification of the problem. The notice describes the product, problem and actions to be taken by customers. The firm recommended that the customers not use the fusion function until their Leksell GammaPlan(R) is upgraded to version 8.2.For further question, do not hesitate to contact your Elekta support respresentative at |  | 66 | 66 units | Nationwide distribution: AL, AR, CA, CO, CT, FL, GA, IL MD, MA, MI, MN, MS, MO, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, and WA. | Terminated | December 28, 2010 | 1086 | 66 | Software | Output/Calculation | Software update | Software Update
Z-1608-2010 | 55428 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Coated VICRYL RAPIDE (polyglactin 910) Braided Coated Synthetic Absorbable Suture, Undyed. |  |  | April 20, 2010 | 2010 | Ethicon, Inc. | Channels in the packaging seal that can lead to suture degradation and loss of packaging integrity. Recall Z-0021-2009 was initiated in 2008. Ethicon had decided to recall all sutures product during the may to June 2007 timeframe. | Other | "Expansion of 2008 Voluntary Recall" notification letters dated April 20, 2010 were sent by second day UPS to customers, distributors, Materials Managers, Director of Surgical Services, and Directors of Risk Management. The letters state that the expansion includes 37 lots of products. The letters describe the product, problem and action t to be taken by customers. The customers were instructed to immediately discontinue use of any of the affected product and to remove from all inventory and returned to Stericycle. If they are not returning product, they should mail the enclosed business reply card to Stericycle. The customers should notify their end customers of this recall. If their end customers have the affected product they should first call 1-888-345-0483 to get appropriate shipping labels. THIS RECALL EXPANSION ONLY IMPACTS THE VICRYL RAPID SUTURE. THIS DOES NOT INCLUDE THE VICRYL* ADN VICRYL* PLUS PRODUCTS.Please don't hesitate to contact Christie Bielinski directly at 908-218-2553 or via email at cbielinsi@its.jnj.com should you have any questions or require additional information. |  | 8275666752 | 118332 eaches were sent to the US Distribution Center, of which 69936 eaches were shipped to primary consignees. | Worldwide distribution: USA and Austria, Belgium, Canada, Chile, China, Czech Republic, Egypt, Germany, Greece, Hong Kong, Hungary, Israel, Italy, Japan, Korea, Malaysia, Middle East, Netherlands Antilles, New Zealand, Poland, Portugal, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom | Terminated | July 19, 2011 | 455 | 8275666752 | Not_Computer
Z-1609-2010 | 55280 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Odyssey |  |  | March 23, 2010 | 2010 | PerMedics, Inc. | Anomaly 1, Invalid Characters in Patient ID Error:Currently, Odyssey considers a valid ID to be comprised of numbers and/or letters. If a patient ID in a study contains a character other than a letter or number (including spaces). Odyssey removes the invalid character when moving it to a patient folder. Due to this, it is possible that two different patient studies could be listed within Odyssey | Other | A letter dated March 23, 2010 was mailed to customers acknowledging the anomalies and informed the reader that the mandatory version correcting the software will be provided by May 19, 2010. Customers with version 4.4 will receive 4.4.04:t hose with version 4.5 will receive 4.5.04; and those with version 4.6 will receive 4.6.01 containing the corrections to these issues. In the meantime, the following recommendations are provided. For Anomaly I . Invalid Characters in Patient ID Error, it 1s recommended that the users do not create patient IDS that contain characters other than letters or numbers. For Anomaly 2.0dysse.v Muchitle Library Collimator Type Error, it is recommended that the customer contact PerMedics ifany modifications to the library have been performed. The primary view should also be examined and used to detect and avoid the conditions causing the anomaly. Finally, independent field shape and monitor unit checks should be performed as part of a routine quality assurance program. For Anomaly 3, OOdyssey MLC Block Error, it is recommended to avoid using the MLC Block feature until the mandatory version correcting this anomaly is installed. If you need to use this feature, then it is recommended that a physical block tray is used during treatment for each beam that contains an Odyssey MLC Block. If customers believe a patient treatment has been affected by any of these issues, they should notify Customer Support as soon as possible at 877-4PerMed (877-473-7633) or support@permedics.com |  | 16 | 16 | Worldwide distribution: USA including NY, GA, CA, MA, and countries of Brazil, Hungary, China, India, Italy, Mexico. | Terminated | July 30, 2012 | 860 | 16 | Software | Output/Calculation | Software update | Software Update
Z-1596-2010 | 55165 | No | LOL | Hepatitis A virus (HAV) serological assays. | Microbiology | Microbiology | 510(k) | ARCHITECT HAVAB-M Reagent |  |  | December 30, 2009 | 2009 | Abbott Diagnostics International, Ltd. | There is a potential to generate a false grayzone or reactive HAVAB-M result when the ARCHITECT AUSAB assay precedes the ARCHITECT HAVAB-M assay. | Pending | Direct accounts (hospitals, private laboratories, and other institutions) were sent a Product Correction letter, dated 12/30/2009. The letter identified the affected product, explained the issue and patient impact, and listed several necessary actions. Customers were instructed to follow one of 4 different alternatives to protect the integrity of test results. The letter is to be retained for laboratory records and a copy should be sent to other laboratories that kits were forwarded to. Questions regarding this information should be directed to Customer Service at 1-877-4ABBOTT. |  | 1326 | 1,326 | Nationwide Distribution -- Including Puerto Rico. | Terminated | August 31, 2011 | 609 | 1326 | Not_Computer
Z-1610-2010 | 55215 | No | GAG | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Disposable Skin Stapler - 35 Wide; Surgical Skin Stapler; Rx, sterile, single use skin stapler, individually packaged in a thermo-formed tray with a Tyvek lid, 5 trays per box; Distributed by Cardinal Health, Dublin, OH 43017, Made in China; catalog number FS-35W. |  |  | April 01, 2010 | 2010 | Cardinal Health | The sterility of the staplers is questionable. | Packaging process control | Cardinal Health issued an "Urgent: Product Recall" letter dated April 1, 2010 to all customers who may have received affected product. Customers were instructed to immediately quarantine any affected inventory, return the notification acknowledgement form and return the product to the firm.For further information, contact Cardinal Health Customer Service at 1-888-444-5440. |  | 1638 | 1,638 units | Nationwide Distribution | Terminated | August 17, 2011 | 503 | 1638 | Not_Computer
Z-1612-2010 | 55508 | No | HWX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 5.5MM Bone Tap, Modular; Manufactured by: Blackstone Medical, Inc. 1720 Bray Central Drive, McKinney, TX 75069. |  |  | April 21, 2010 | 2010 | Blackstone Medical, Inc. | Certain lots of the non-cannulated Modular Bone Taps used with the Firebird Spinal Fixation Systems may contain a manufacturing non conformance resulting in the taps creating a single lead screw thread rather than the required dual lead screw thread. | Nonconforming Material/Component | Orthofix issued an "Urgent Product Recall" notification dated April 21, 2010 by USPS certified mail to hospitals and distributors. Consignees were asked to return all affected product to the firm.For additional information, contact the Orthofix Regulatory Affairs Department at 1-973-406-2847. |  | 94 | 94 | Nationwide and to one location each in Australia and Spain. | Terminated | August 10, 2010 | 111 | 465 | Not_Computer
Z-1616-2010 | 55555 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi AIRIS, AIRIS II, AIRIS Elite MRI Scanner Compensation Coil. Part numbers: MR-FC-2EO, MR-FC-3EO |  |  | April 19, 2010 | 2010 | Hitachi Medical Systems America Inc | Hitachi determined that the mounting method for a Compensation Coil did not include an appropriate factor of safety, so it is possible that the coil could fall, contacting the patient table and potentially putting a patient at risk. | Other | On 4/15/2010 the firm sent Device Correction Notice letters to their customers identifying the affected product and informing them of the potential problem. The letter stated that the firm will change the coil mounting method used at the customer's facility, if necessary, at no cost to the customer. A service rep will contact them to schedule a visit to correct the problem. |  | 56 | 56 devices | Nationwide Distribution -- AL, AR, AZ, CA, CT, DC, DE, FL, ID, IL, IN, KY, LA, MD, MO, NJ, NY, OH, OK, OR, RI, TN, TX, VA &amp; WA. | Terminated | September 21, 2012 | 886 | 56 | Not_Computer
Z-1617-2010 | 55504 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport e360 Ventilator |  |  | April 14, 2010 | 2010 | Newport Medical Instruments Inc | The recall was initiated due to a recent increase in complaints related to the early failure of a low pressure transducer which is on Printed Circuit Board (PCB) 2104A in the Newport e360ventilator. Ventilators manufactured between September and December 2009 areaffected by this field correction. All affected ventilators require the replacement of the PCB2104A board to avoid early failure of t | Other | The firm, Newport Medical, sent a "Field Correction Notice" dated April 14, 2010, to all customers. The notice describes the product, problem and actions to be taken by the customers. The customers were instructed to: 1) provide your affected customers with a copy of this letter and advise them to pay close attention to pressure errors which can be an early warning sign of the pressure transducer failure; 2) Schedule their ventilators for replacement of the PCB2104A board as quickly as possible to avoid any unanticipated disruption in their use of their product;3) Contact our Technical Support Department (800.451.3111 for US customers, 714.427.5811x500 for international customers) to request the replacement kits you will need; and4) Complete the attached form and return it to Newport Medical as you complete your PCB2104A replacements. This will allow us to ensure that all affected ventilators have been addressed and close this field action.For questions, customers were instructed to contact Newport Medical Technical Support Department (800.451.3111 for US customers, 714.427.5811x500 for international customers). |  | 990 | 990 | Worldwide distribution: USA including states of CA, KY, NM, NY, UT and internationally to: Armenia, Bolivia, Brazil, Bulgaria, China, Colombia, Ecuador, Egypt, Greece, Hong Kong, India, Malaysia, Mauritius, Mexico, Nicaragua, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Russia, Saudi Arabia, Serbia, South Africa, Sri Lanka, Syria, Taiwan, Thailand, Tunisia, Turkey, UAE, Ukraine, Vietnam. | Terminated | August 18, 2011 | 491 | 990 | Hardware |  | Replace board | Remove or Replace
Z-1628-2010 | 53212 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN 1700, list numbers 03H5301 and 03H53-03, manufactured by Abbott Hematology, Santa Clara, CA |  |  | June 08, 2009 | 2009 | Abbott Laboratories | Use of Syringe LN 04H36-01 with the Die-cast syringe drive may result in leakage, leading to delay in reporting patient results. | Nonconforming Material/Component | A Product Correction Letter and Customer Rely Form were sent to all currently active Cell -Dyn Sapphire customer. New customers are being notified with a Product Information Letter.For questions regarding this information, U,S, customers should call Customer Support at 1-877-4ABBOTT. Customers outside the U.S., please contact your local hematology customer support representative. |  | 274 | 274 | Worldwide distribution: United States, Canada, Mexico, Anguilla, Australia, Argentina, Belize, Bermuda, Brazil, China, Chile, Columbia, Costa Rica, Ecuador, El Salvador, Germany, Guatemala, Haiti, Honduras, Hong Kong, Japan, Paraguay, Peru, Puerto Rico, Singapore, South Korea, Uruguay, and Venezuela. | Terminated | December 03, 2010 | 543 | 2090 | Not_Computer
Z-1627-2010 | 52527 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Endotec Prothesis Ankle, Titanium Alloy, Sterilized, BP Ankle System. |  |  | June 05, 2009 | 2009 | Endotec, Inc. | Endotec is recalling B-P Hip, B-P Shoulder, and B-P Knee Replacement System devices due to sterility issues involving validation, process variations and test results. | Other | Recall Letters were sent to each distributor on 06/07/2009. Questions are directed to Endotec at (407) 822-0021. |  | 156 | 156 Pieces | Worldwide distribution -- US in OH, TX, CA, FL, NJ, TN; and International Distribution: Korea, Australia, Switzerland, The Netherlands, Spain.Domestic: OH, TX, CA, FL, NJ, TN. | Terminated | October 22, 2013 | 1600 | 3477 | Not_Computer
Z-1620-2010 | 49126 | No | LOM |  |  | Microbiology | PMA | Bio-Rad's Monolisa (TM) Anti-HBs EIA test kit, Labeled as, "MONOLISA TM Anti-HBs EIA 25220; |  |  | August 11, 2008 | 2008 | Bio-Rad Laboratories | The package insert provided with this kit contains a typographical error, giving incorrect directions to prepare a Working Conjugate Solution. If these directions are followed, it would result in a 1:6 dilution and invalidate the assay. | Packaging process control | All domestic customers were notified via an Urgent Product Correction customer letter sent by DHL or FedEx on 8/11/08. Enclosed with the letter was a corrected package insert. The Canadian subsidiary was notified via e-mail on 8/11/08. The letter advised customers of the correct wording, and advised that a new corrected package insert was enclosed and to make note of the correction in the laboratory procedures. Questions should be directed to Bio-Rad Laboratories Technical Support at 1-800-224-6723. |  | 795 | 795 kits | Worldwide Distribution -- USA and Canada. | Terminated | May 20, 2010 | 647 | 795 | Not_Computer
Z-1619-2010 | 55304 | No | GWL | Physiological signal amplifier. | Neurology | Neurology | 510(k) | Inserter Wand Holder, Non-Sterile, Serial # 4273-0000, Blackrock Microsystems, Salt Lake City, UT / or / Cyberkinetics Neurotechnology Systems, Inc., Salt Lake City, UT. |  |  | March 26, 2010 | 2010 | I2s Micro Implantable Systems Llc | Component of cranial array may not allow sufficient tightening to maintain inserter position. | Device Design | The firm, Blackrock Microsystems, sent a "RECALL" letter dated March 26, 2010, to all customers. The letter describes the product, problem and actions to be taken by the customers. The customers were instructed to discontinue use of all affected devices and promptly return any inventory, perform a physical account of your inventory for the recalled product and email a list of all serial numbers to Andrew Gotshalk at agotshalk@blackrockmicro.com, and to please email Andrew Gotshalk even if you do not have any inserter wand holders or have destroyed or removed from use the inserter wand holders. Blackrock will rework or replace all Inserter Wand Holders that is returned under this recall. If you have any questions regarding this recall please call 801-582-5533 x 275. |  | 22 | 22 units | Nationwide distribution: CA, MA, NY, RI, and UT. | Terminated | May 25, 2010 | 60 | 22 | Not_Computer
Z-1633-2010 | 55241 | No | FSO | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Leica Wireless Footswitch, Type B used in conjunction with Leica Surgical Microscopes - Models M620, M844/820, M525 and M720; There are two wireless models: |  |  | March 29, 2010 | 2010 | Leica Microsystems, Inc. | The pedals of the Footswitch Type B stick and the Zoom/Focus moves to the end position. | Other | Leica Microsystems sent Medical Device Correction letters dated 3/29/10 to the affected customers via first class mail on the same date, informing them of the potential for injury to the patient's eye if the cross pedal of the footswitch sticks. They were informed that their local Leica Microsystems representative will contact them to arrange for the replacement of the cross pedal. The letters also provided the interim necessary steps required to avoid the risk of patient injury while they are awaiting the cross pedal replacement, i.e. do not use the affected footswitches with a surgical microscope that has a BIOM accessory. The accounts were requested to complete the attached Acknowledgement Form, listing the serial numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. A company contact was provided in the letter for questions. |  | 3 | 3 units | Worldwide distribution -- U.S. (nationwide), Canada, Australia, Europe and Asia | Terminated | February 15, 2011 | 323 | 57 | Not_Computer
Z-1631-2010 | 53322 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific, LATITUDEÂż Communicator, Distributed by Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 55112. Models 6476 (USA) and 6468 (OUS) |  |  | September 10, 2009 | 2009 | Boston Scientific CRM Corp | Boston Scientific has determined that alert conditions for a limited subset of out-of-range Daily Measurements related to leads may not be recognized by Model 6476 (United States) or Model 6468 (Europe) LATITUDE in-home communicators. If a communicator fails to recognize a valid alert situation, notification for this or subsequent alerts for the same test will not be displayed for physician review | Software design | Consignees were sent on 9/10/09 a Boston Scientific "Urgent Medical Device Information" letter dated September 9, 2009. The letter was addressed to "Dear Doctor". The letter provided the information related to the LATITUDE Patient Management system recall as to Background, Description, Rate of Occurrence, Patient management, Devices Affected and Further Information. Patients were also sent a Boston Scientific letter dated September 9, 2009. The letter addressed Background information, Why are we communicating, What You Should Do and Questions. |  | 25656 | 25,656 (USA), 390 (OUS) | All States in the US including Puerto RicoUnited Kingdom, Switzerland, Netherlands, Italy, Germany, France, Finland, Denmark, Belgium, Austria | Terminated | April 05, 2012 | 938 | 25656 | Other | Alarm/Message | Notification | Safety Notice/Insructions
Z-1639-2010 | 54174 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Churchill Medical Systems Bag Access Device w/Smartsite Needleless Injection Site, Sterile, Non pyrogenic, Latex Free Product Code: KSA-140. Churchill Medical Systems, Inc. Dover, NH 03820. |  |  | October 20, 2009 | 2009 | Churchill Medical Systems, Inc. | Bond between the spike and needle may leak, break or detach resulting in leakage or spillage of chemotherapeutic agents | Process change control | Churchill Medical Systems issued an "Urgent: Product Advisory Customer Notification" dated October 20, 2009 requesting return of affected product and notification to end users. For further information, contact Churchill Medical Systems at 1-603-743-5988 or via fax at 1-603-743-6328. |  | 283900 | 283,900 each | Distributed to Florida. | Terminated | March 28, 2013 | 1255 | 283900 | Not_Computer
Z-1636-2010 | 55436 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ADVANCE DURAMER TIBIAL INSERT, REF: KIMP-312L, SIZE 3 LEFT MEDIAL PIVOT, THICKNESS 12MM, 1 EACH, STERILE EO, Wright Medical Technology, 4677 Airline Road, Arlington, TN 38002. |  |  | November 24, 2009 | 2009 | Wright Medical Technology Inc | The left medial pivot and right medial pivot inserts were co-mingled | Packaging process control | Wright Medical Technology, Inc. issued an "Urgent: Medical Device Voluntary Recall" notification dated November 24, 2009. The letters were addressed to distributors, hospital administrators, and surgeons. Consignees were asked to return all affected product in addition to a completed Response Form to the firm.For further information, contact Wright Medical Technology, Inc. at 1-800-874-5630. |  | 23 | 23 units | Worldwide Distribution -- United States, The Netherlands, Italy, Republic of Korea and Japan. | Terminated | October 21, 2010 | 331 | 70 | Not_Computer
Z-1635-2010 | 55407 | No | MOZ |  |  | Orthopedic | PMA | Euflexxa (1% sodium hyaluronate). Pre-filled syringe, Rx only, 3 syringes per carton; Manufactured by Ferring Pharmaceuticals In, Parsippany, NJ 07054; |  |  | April 16, 2010 | 2010 | Ferring Pharmaceuticals Inc | Out of Specification results for molecular weight or viscosity through the established expiration date. | Process control | Ferring Pharmaceuticals issued a "Medical Device Recall" notification dated April 16, 2010 by Federal Express to all distributors, retail pharmacies and physicians who received the product. Consignees were asked to stop distribution, return affected product to the firm and fax a completed Product Inventory Form to the firm.For further information, contact Ferring Pharmaceuticals at 1-973-796-1600 or via fax at 1-973-796-1660. |  | 88128 | 29,376 cartons with 3 syringes per carton. | Nationwide distribution. | Terminated | July 27, 2010 | 102 | 88128 | Not_Computer
Z-1645-2010 | 55477 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Ceiling Suspensions (Surgical Light and Flat Panel Suspensions); Catalog Numbers: 0682000863, 06820001009, 0682001010, 0682001076, 0682000864, 0682000857, 0682000858, 0682001077, 0682000865, 0682001011, 0682001078, 0682000866, 0682001012, 0682001080, 0682000867, 0682000959, 0682001079, 0682000850, 0682000851, 0682000852, 0682001269, 0682001567, 0682001569, 0682001570, 06820012691, 06820015671, 06820015691, &amp; 0680015701. |  |  | April 15, 2010 | 2010 | Stryker Communications Corp | Installation records are incomplete. | Employee error | An Urgent: Medical Device Product Correction customer notification letter, dated 4/15/2010, has been sent out to all affected accounts via regular mail in the form of a certified, return request receipt letter. Accounts will receive a follow-up phone call from a Stryker Representative to schedule a re-inspection of equipment to see if it mets established critical specifications or not. The firm recommends that customers continue using all Stryker Communications products listed in the notice. Questions can be directed to Katie Blackwell at 972-410-7660 or via email at StrykerCommunicationsRA@stryker.com. |  |  |  | Nationwide Distribution | Terminated | January 24, 2011 | 284 |  | Not_Computer
Z-1638-2010 | 52291 | No | HTS | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | BD Beaver Arthro-Lok Pointed Tip 4mm Blade, REF 376780. Becton Dickinson and Company, Waltham, MA 02454. |  |  | June 12, 2009 | 2009 | Becton Dickinson and Company | Mislabeled: Arthro-Lok Pointed Tip 4mm Blade curving right instead of left. | Packaging process control | Becton Dickinson issued an "Urgent: Field Safety Notice" beginning June 25, 2009 by telephone and letter fax to remove and discard affected product from their possession. Customer verification responses were requested.For further information, contact Becton Dickinson and Company at 1-781-906-7900. |  | 354 | 354 units | Worldwide Distribution -- United States, Japan and Belgium. | Terminated | January 26, 2012 | 958 | 354 | Not_Computer
Z-1618-2010 | 55235 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator Monitor, Model M3535A |  |  | April 07, 2010 | 2010 | Philips Healthcare Inc. | ECG signal processing chip may fail causing failure to deliver therapy in AED mode or the failure to indicate poor electrical contact (very high impedance) with the patient in Manual Defib mode. | Component design/selection | The firm, Philips Healthcare, sent an "URGENT voluntary Medical Device Correction" Field Safety Notice, letter dated April 7, 2009, to all customers via UPS. The letter describes the product, problem and the actions to be taken by customers, with precautions to follow if the device is used before upgrade is installed. The customers were instructed to check their serial numbers for the affected product, if the device is affected, it may remain in service prior to correction using the following precautions 1) if using AED, have a back up AED; 2) if using external defibrillator paddles, disregard paddles and manually ensure good paddlet-to-patient contact, and 3) if monitoring, be prepared to switch back to the primary ECG that was initially obtained for a given patient if other leads are lost. Customers will be contacted and assisted by Philips field service engineers to install the software upgrade which corrects the problem.Should you have any questions or concerns about this Device Correction, please contact your local Philips representative at 1-800-722-9377. |  | 1379 | 1379 units | Worldwide distribution:USA and Canada, Argentina, Australia, Austria, Belgium, Brazil, China, Finland, France, Germany, Iceland, India, Ireland, Italy, Japan, Korea, Republic of Lithuania, Mexico, Netherlands, New Zealand, Norway, Poland, South Africa, Spain, Switzerland, Thailand, Trinidad, and Tobago, Ukraine, United Arab Emirates, United Kingdom. | Terminated | February 20, 2013 | 1050 | 1379 | Hardware | Treatment/Delivery/Therapy | Software update | Software Update
Z-1651-2010 | 55079 | No | OPU |  |  | Microbiology | EUA - Emergency Use Authorization | Cepheid Xpert Flu A Panel (GXFLUPANEL-10), manufactured by Cepheid, Sunnyvale, CA. |  |  | February 19, 2010 | 2010 | Cepheid | Firm has had an increasing number of customer complaints regarding higher than expected Indeterminate Test Results due to "Invalid" test reports from our system when they have been conducting their own internal validation studies using banked frozen repository specimens. | Other | Cepheid issued a notification to consignees regarding the affected product. Additional instructions were provided to consignees in order to avoid further "Indeterminate Results".For further information, contact Cepheid at 1-888-336-2743. |  | 1803 | 1803 kits | Worldwide Distribution -- United States, France, Belgium, Germany, Norway Austria, Denmark, Spain, Taiwan and Hong Kong. | Terminated | March 23, 2011 | 397 | 2747 | Not_Computer
Z-1663-2010 | 54268 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | Bausch &amp; Lomb Millennium Microsurgical System. |  |  | January 15, 2010 | 2010 | Bausch &amp; Lomb Inc | The AC power cords used on these products are subject to an FDA Alert -- Safety Investigation of Certain Medical Device Power Cords: Initial Communication, Date Issued: October 19, 2009. | Nonconforming Material/Component | A "URGENT: POWER CORD RECALL" letter and copies of the FDA Alert dated Oct. 19, 2009 were sent to all customers on or about Jan. 15, 2010, by Stericycle, a recall notification service. The letter describes the product, problem and actions to be taken by the customers as well as FDA's recommendations and actions. The customers were made aware of the FDA Alert concerning the AC power cords and notified that these power cords were used in the devices. They were requested to inspect the power cords for damage and notify the recalling firm if damage is noted or if they have experienced sparking, charring, or firing from the power cords and/or cracking or failure of the prongs. In the absence of those conditions, customers were given the option of returning the power cord for replacement and/or during the regularly scheduled servicing of the device. If you would like a new power cord, please contact Customer Service at 1-800-338-2020 for a free replacement. |  |  | Quantity never provided by firm | Worldwide distribution: United States &amp; Puerto Rico; Venezuela, India, Trinidad and Tobago, Colombia, Sri Lanka, Hong Kong, Vietnam, and Peru. | Terminated | May 17, 2013 | 1218 |  | Battery |  | Replace component | Remove or Replace
Z-1681-2010 | 54473 | No | ESD | Hearing Aid. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | InSound Medical, Inc, Lyric in-canal hearing aid. |  |  | January 19, 2010 | 2010 | InSound Medical, Inc. | Manufacturing error could result in electrolyte leakage from the product's battery. | Process design | InSound sent to its consignees a "Dear Lyric Provider" letter of explanation on 2/1/2010. |  | 2019 | 2019 units | Distribution only in US ( nationwide). No foreign consignees involved. | Terminated | August 02, 2011 | 560 | 2019 | Not_Computer
Z-1657-2010 | 55176 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CoulterÂż LH 500 Series System. Beckman Coulter, Inc. Brea, CA. |  |  | February 04, 2010 | 2010 | Beckman Coulter Inc. | Issue 1: The system will substitute or omit the characters # @ [\] ' {I) when the system is configured in a language other than English or Chinese. Example: Using the character @, a substitution in the string '123@456' would result in '123#456' and an omission would result in '123456'Issue 2: The system will omit the character ~ regardless of language.Impact: There is a potential for Specimen | Software design | Beckman Coulter, Inc. issued "Urgent: Product Corrective Action" notifications dated February 12, 2010. Consignees were informed of the affected product and instructed on proper actions. The firm also asked that the enclosed response forms be returned.For further information, contact your local Beckman Coulter Representative or call 1-800-526-7694. |  | 1520 | 1520 units | Distributed Nationwide and Canada. | Terminated | September 02, 2014 | 1671 | 1576 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-1654-2010 | 55300 | No | LNR |  |  | Gastroenterology/Urology | PMA | Therakos CellEx Photopheresis System; |  |  | March 29, 2010 | 2010 | Therakos Inc | Higher than normal blood leaks on THERAKOS CELLEX Procedural Kits that are intended for use with the THERAKOS CELLEX Photopheresis System | Nonconforming Material/Component | On March 29, 2010, Therakos Photopheresis sent Important Product Correction Notification letters by FedEx to all direct end use and distributor customers who received lot X120 in addition to phone contact. On March 31, 2010 those who received lot X123 were sent letters by Fed Ex as well as phone contact.. The one Canadian customer was notified by phone on March 31, 2010 followed by a Fed Ex letter. On April 12, consignees who were shipped Lot X119 were contacted by phone, followed by a hard copy letter.On May 7, 2010 all direct end use and distributor consignees were notified of the additional lot being added to the recall. Questions may be directed to Therakos Customer Technical Services at 1-877-865-6850. |  | 840 | 280 units in US (1 sales unit contains 3 kits each) | Nationwide Distribution, and Canada and Germany | Terminated | July 18, 2013 | 1207 | 840 | Not_Computer
Z-1653-2010 | 55002 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Blackstone Medical Inc./ Orthofix Spinal Implants, Hallmark Anterior Cervical Plate System; Plate Bender Part number 62-0070. |  |  | March 15, 2010 | 2010 | Blackstone Medical, Inc. | The Hallmark Anterior Cervical Plate Bender may have a non conformity causing difficulty accepting 3-4 level plates or to not accept 3-4 level plates that are intended to be additionally contoured. | Nonconforming Material/Component | Orthofix send letters of notification of planned product removal to distributors who keep the instrumentation set on 3/29/2010 by Fed Ex. Letters of planned product removal explain the reason for the planned removal and the specific product involved. Replacement product will not be available until July 2010, however, the firm elected to notify customers of the potential problem and the date that formal notification and replacement Plate Benders will be available. Letters state that if there is any problem with the Plate Benders, questions can be directed to David Thompson in Marketing at 972-529-3024 or Jose Charvet in Product Development at 973-406-2824. |  | 112 | 98 units domestically, 14 units internationally. | Nationwide distribution, including Puerto Rico, and to Australia. | Terminated | September 09, 2011 | 543 | 112 | Not_Computer
Z-1659-2010 | 55445 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 4100 IQ: GE Innova 4100 / 4100 IQ., Cardiovascular Imaging System (The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | March 24, 2010 | 2010 | Ge Healthcare, Llc | A potential software failure related to the Worklist functionality has been identified on a subset of Innova 2100IQ, 3100IQ and 4100IQ systems. When using the Worklist function to import patient data to Innova, there are two scenarios thatwill cause the next exam to fail and the inability to recall acquired sequences. | Software design | Consignees were sent on 3/24/10, a GE Healthcare " Urgent Medical Device Correction" letter dated March 23, 2010. The letter included the Safety Issue, Affected Product Details, Product Correction and Contact Information. |  | 28 | 28 | Worldwide distribution: USA, Puerto Rico, VENEZUELA, UNITED KINGDOM, TURKEY, THAILAND, TAIWAN, PROVINCE OF CHINA, SWEDEN, SPAIN, RUSSIAN FEDERATION, ROMANIA, POLAND, PHILIPPINES, NORWAY, NETHERLANDS, MOROCCO, REPUBLIC OF KOREA, KAZAKSTAN, JAPAN, ITALY, ISLAMIC REPUBLIC OF IRAN, INDONESIA, INDIA, GERMANY, FRANCE, FINLAND, EGYPT, CHINA, CHILE, CANADA, BRAZIL, BELGIUM, BELARUS, AUSTRALIA, and ARGENTINA. | Terminated | May 21, 2013 | 1154 | 195 | Computer
Z-1656-2010 | 55194 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number 629029. Beckman Coulter, Inc. Brea, CA. |  |  | March 02, 2010 | 2010 | Beckman Coulter Inc. | Issue 1: The Diagnostic procedure, Simulate-CDNR Cycle, is not functioning properly. Impact: There is a potential for inadequate sample preparation of samples following the use of this Diagnostic function impacting WBC, Differential, Reticulocyte and/or NRBC results.Issue 2: Use of the Auto Prune function causes the deletion of Quality Control files.Impact: If the last run for a control lot | Software design | Beckman Coulter, Inc. issued an "Urgent: Product Corrective Action" notification dated March 10, 2010. Consignees were informed of the affected product and instructed on proper steps to take and requested to return the enclosed acknowledgement form. For further information, contact your local Beckman Coulter Representative or call 1-800-526-7694. |  | 63 | 63 units | Nationwide Distribution (AZ, CA, CO, FL, GA, IA, IL, IN, KS, LA, MO, NC, NJ, NV, NY, OH, PA, SC, TN, TX, UT, WA, and WI) and Canada. | Terminated | May 23, 2012 | 813 | 63 | Software
Z-1655-2010 | 55361 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta XiO Radiation Treatment Planning System |  |  | April 15, 2010 | 2010 | Computerized Medical Systems Inc | Under certain circumstances, XiO will over-estimate dose in the region where the backup jaw is inside the MLC opening | Software design | Elekta issued an undated "Important Safety Notice" to all affected sites via e-mail on 4/20/2010. The notice was also sent to the sites on 4/6/10 and 4/12/10 (for 4 sites) before the firm made the decision to recall. The letter explains the correct functioning of the system, the problem, when and why it occurs, the clinical impact, and the work around. Although not stated in the Safety Notice, a self-addressed postcard was enclosed for the customer to return as evidence they received and understood the notice. Sites that do not return the card will be contacted by phone. |  | 34 | The release is being used by 34 customers that also have an Elekta linear accelerator | Worldwide Distribution: USA, Canada, Australia, Spain, Italy, the Netherlands, United Kingdom, Taiwan, Malaysia, France, Ireland, and Israel. | Terminated | June 07, 2012 | 784 | 34 | Software | Treatment/Delivery/Therapy | Instructions/Workarounds | Safety Notice/Insructions
Z-1680-2010 | 55462 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 Coulter Cellular Analysis System, Part Number: 629029. Beckman Coulter, Inc. Brea, CA. |  |  | December 08, 2009 | 2009 | Beckman Coulter Inc. | The Corrective Action was initiated after Beckman Coulter confirmed three issues related to host transmissions with the UniCel DxH 800. This information only affects instruments interfaced to a host system. | Software design | Beckman Coulter, Inc. issued an "Urgent: Product Corrective Action" notification dated December 8, 2009. Consignees were informed of the affected product and provided temporary solutions until further support will be made available.For additional information, contact your local Beckman Coulter Representative or call 800-526-7694. |  | 42 | 42 units (40 in US) | Distributed Nationwide and Canada. | Terminated | June 04, 2012 | 909 | 42 | Not_Computer
Z-1683-2010 | 53668 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Responder 2000 Defibrillator/Monitor |  |  | October 15, 2009 | 2009 | Cardiac Science Corporation | The Operators Manual may not have adequate information to clearly explain some device behaviors. Pressing and holding down the Charge or Shock button on the Responder 2000 for more than 5 seconds causes a "Button Stuck" error. With the "Button Stuck" error, the device will automatically exit Manual mode and enter Monitor mode. The user cannot apply therapy while the device is in Monitor mode. | Other | The firm, Cardiac Science, sent an "Urgent-Medical Device Corrective Action" letter dated October 15, 2009, was sent to customers. The letter describes the product, problem and action to be taken by the customers. Cardiac Science requested GE to provide the end users with the Operator's Manual addendum. The customers are instructed to place the Operator's Manual addendum with their Responder 2000 Operator's Manual for future reference.If you have any questions about your Responder 2000, please contact Cardiac Science technical support at +44.161.926.0011 or our EU Representative, Ludger Moller, MDSS GmbH at +49.511.6262.8630. |  | 1838 | 1838 devices | Worldwide distribution: USA and Australia, Brazil, Columbia, Costa Rica, Denmark, Finland, France, Germany, Italy, Korea, Mexico, New Zealand, Panama, Paraguay, Peru, Singapore, South Africa, Spain, and Uruguay. | Terminated | December 07, 2011 | 783 | 1838 | Not_Computer
Z-1682-2010 | 55385 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow 4FR Peripherally Inserted Central Catheter (PICCs) Kits. Product numbers: ASK-01451-BMH; ASK-05041-CHC; ASK-05041-CHC1; ASK-05041-CHM; ASK-05041-HMC; ASK-05041-QV; ASK-07041-BMH; CN-05041-LW; MSO-01451-UCL; MTO-01451-RH; PL-05041; PR-04041-HPX; PR-05041; PR-05041-HP; PR- 05041-HPX; PR-05041-LW; PR-05041-MW; PR-05041-T; PR-05042; PR-05541-HPX; PR-07041-PTSP; and RJ-01451-W. |  |  | March 29, 2010 | 2010 | Arrow International Inc | Tight fit in sheath. Arrow has received complaints of the PICC catheter fitting too tightly in the "kit supplied" peelable sheath. In some cases, clinicians have been unable to insert the catheter through the peelable sheath during the insertion procedure. If the catheter will not pass through the sheath, the clinician is required to place a guide wire back into the sheath, remove the sheath, a | Nonconforming Material/Component | Arrow International issued an Urgent Medical Device Recall letter dated 3/29/10 to inform their customers of the problem and the need to return the product. |  | 6719 | 6719 | Worldwide distribution: USA, Australia, Brazil, Canada, New Zealand, Hong Kong, Malaysia, and Chile. | Terminated | September 27, 2012 | 913 | 6719 | Not_Computer
Z-1668-2010 | 55365 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Mallory Head Calcar Femoral, 34 mm resection/porous coated, 13.0 x 170 mm,, Sterile, REF 11-104904, Biomet Orthopaedics, Warsaw, IN. |  |  | March 10, 2010 | 2010 | Biomet, Inc. | The firm has found that some lots of Mallory Head Revision Stems had flute geometry that was out of specification. The flute geometry was too deep. | Process change control | Biomet Orthopedics issued an "Urgent Medical Device Recall Notice" dated March 10, 2010. The letter went to distributors and users of the recalled device. The letter instructed the consignees to locate the recalled product and return the product to the firm.For further information, contact Biomet at 1-800-348-9500 or 1-574-372-3983, Monday through Friday, 8:00 AM to 5:00 PM. |  | 6 | 6 | Worldwide Distribution -- United States (OH, IN, FL, KY, NV and CA), Mexico, Chile, China, Japan and The Netherlands. | Terminated | October 18, 2010 | 222 | 81 | Not_Computer
Z-1704-2010 | 54168 | No | GWJ | Evoked response auditory stimulator. | Neurology | Neurology | 510(k) | Intelligent Hearing Systems Corporation, SmartEP with Vestibular Evoked Myogenic Potentials Version 3.96. Model M811016 |  |  | October 21, 2008 | 2008 | Intelligent Hearing Systems Corp. | This was a marketing correction mandated by the United States Food and Drug Administration (USFDA). USFDA notified Intelligent Hearing Systems that 501(k) premarket notification application must be submitted and cleared in order to market and sell devices with Vestibular Evoked Myogenic Potential (VEMP) specific functions. The Food and Drug Administration considers that the VEMP testing functions | Device Design | Intelligent Hearing Systems sent a letter "Information Regarding the Discontinued Use of the SmartEP Evoked Response System for Vestibular Evoked Myogenic Potentials (VEMP) Testing" dated September 30, 2008 to customers and distributors. Each customer will receive an upgrade and VEMP removal kit which will remove the VEMP feature. |  | 50 | 50 | Nationwide distribution. | Terminated | September 21, 2011 | 1065 | 50 | Not_Computer
Z-1686-2010 | 55529 | binder, medical, therapeutic | MDR | Therapeutic medical binder. | General Hospital | General Hospital | 510(K) Exempt | Juxta-Lite compression wrap XL Short 28cm Length with Anklet Catalog 23026017 |  |  | April 20, 2010 | 2010 | Circaid Medical Products Inc | Some systems in these lots have Built-In Pressure System lines with the incorrect amount of spacing. Some systems in these lots had I inch space between each line with no tension applied. The system should have 1.25 inches of space. The correct spacing (1.25") is necessary to achieve the measured tension ranges. The incorrect spacing may result in the inability to achieve the correct tension level | Process control | CircAid Medical Products Important Device Product Recall letter dated Apr 20, 2010 was mailed to consignees describing problem, listing affected product and distribution dates. Customers were instructed to locate and return these Juxta-Lite Systems from stock and notify their customers A Retum Authorization number and contact name was provided.this listing for product returns. |  | 372 | 372 total, all sizes | Worldwide distribution: USA, Japan, Canada, Netherlands | Terminated | January 06, 2011 | 261 | 372 | Not_Computer
Z-1705-2010 | 55538 | No | HQD | Rigid gas permeable contact lens. | Ophthalmic | Ophthalmic | 510(k) | Boston ES Rigid Gas Permeable Contact lens manufactured with Fluoro Silicone Acrylate Rigid Gas Permeable Contact Lens Material. Bausch &amp; Lomb, Inc. Wilmington, MA 01887. |  |  | April 20, 2010 | 2010 | Bausch &amp; Lomb Inc | The amount of D&C; #6 Dye added to the formulation exceeded specified amounts. | Process control | Bausch &amp; Lomb issued an "Urgent Voluntary Medical Device Recall" notification dated April 20, 2010. Consignees were informed of the issue and asked to identify and return all affected product to the firm.For further information, contact Bausch &amp; Lomb, Inc. at 1-585-338-5477. |  | 2575 | 2575 buttons | Nationwide Distribution -- TX, IN and WI. | Terminated |  |  | 2575 | Not_Computer
Z-1706-2010 | 55570 | No | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | AIR OPTIX (lotrafilcon B) for Astigmatism Soft Contact Lenses. CIBA Vision, Corp. Duluth, GA 30097. |  |  | April 27, 2010 | 2010 | Ciba Vision Corporation | The lenses inside the package do not match the prescription information for power labeled on the primary package. | Packaging process control | CIBA Vision, Corp. notified consignees of the affected product via letter beginning April 15, 2010. Users were asked to return the product to the firm.For further information, contact your CIBA Vision sales representative or call 1-877-542-5928. |  | 5040 | Lot #8717082, total 5040 lenses and Lot #8717066 total 5394 lenses | Nationwide Distribution -- AZ, AL, CA, FL, GA, IA, IL, MA, MO, NC, NJ, NY, OH, OR, PA TN, VA and WI. | Terminated | October 22, 2010 | 178 | 5040 | Not_Computer
Z-1707-2010 | 55715 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | PULSESPRAY INFUSION SYSTEM, 90 cm x 20 cm infusion length, Catalog No./Ref 12500402, STERILE, EO. AngioDynamics, 603 Queensbury Avenue, Queensbury, NY 12804. |  |  | May 17, 2010 | 2010 | Angiodynamics, Inc. | This lot was packaged with the wrong occluding ball guidewire. The PulseSpray Catheter is 90 cm in length and requires a 109 cm occluding ball guidewire, however this lot was packaged with a 60 cm occluding ball guidewire. | Process control | Angiodynamics issued an "Urgent Medical Device Recall" notification (dated 5/11/10) to user accounts via Certified Mail. Consignees were asked to identify, segregate and return all affected product to the firm in addition to sending in a completed Reply Form via fax.For further information, contact Angiodynamics Customer Service at 1-800-772-6446. |  | 20 | 20 units | Worldwide Distribution -- United States (AR, NY, PA, MI, MN, LA and WA), Spain, Austria and Denmark. | Terminated | July 25, 2012 | 800 | 20 | Not_Computer
Z-1703-2010 | 55744 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | S1813OB 1.8X13mm Sendax MDI Mini Dental Implant O-Ball Prosthetic Head |  |  | May 14, 2010 | 2010 | 3m Imtec Corporation | Product has Incorrect size Labeling. Actual length is 13mm. One label is correct, and one states 10mm. | Packaging process control | IMTEC notified consignees by phone on May 14, 2010 and by Certified Letter dated May 14, 2010. |  | 189 | 189 Implants | Nationwide. | Terminated | July 14, 2010 | 61 | 189 | Not_Computer
Z-1732-2010 | 55647 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis systems with Motor Controller Unit. Siemens Medical Solutions USA, Inc. Malvern, PA 19355. |  |  | April 12, 2010 | 2010 | Siemens Medical Solutions USA, Inc | sporadic malfunctions of the Artis Systems with Motor Controller caused by tolerances on electronic parts that might increase during the lifetime of the system. | Device Design | Siemens issued an "Urgent Field Safety Notice" to consignees describing the affected device instructing users to contact the recalling firm's service organization to schedule maintenance.For further information, contact Siemens Medical Solutions USA, Inc. at 1-610-448-3237. |  | 83 | 83 units | Nationwide Distribution | Terminated | April 11, 2013 | 1095 | 83 | Hardware | Device Operation | Repair | Repair
Z-1714-2010 | 54200 | No | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb Stellaris Optimized Stability Vacuum Pack, Model BL5112, 6 packs/box, RX, Sterile. The firm name on the label is Bausch &amp; Lomb Incorporated, Rochester, NY. |  |  | January 08, 2010 | 2010 | Bausch &amp; Lomb Inc | Users may incorrectly assemble the plastic needle wrench to the phacoemulsification needle resulting in the generation of plastic particulate | Component design/selection | The recalling firm, Bausch &amp; Lomb, issued an "URGENT: PHACO NEEDLE WRENCH CORRECTION" letter dated January 7, 2010, via first class mail to all customers. The letter describes the product, problem and actions to be taken by the customers. The letter also provides instructions and pictorial representations demonstrating proper seating of the needle to the wrench as well as a description of the appropriate force to tighten the needle to the handpiece. The customers were instructed to please read and follow the attached instructions entitled (Potential for Particulate Matter in the eye and proper assembly and use of needle wrench) and fill-out and return/fax at 949-716-8362, the Bausch &amp; Lomb Phacoemulsification Needle Wrench Particulate correction form. If you have any questions or concerns please contact, Chris Willis or Stephen Haight at 1-800-387-3284. |  | 14685 | 14,685/6-pack boxes | Worldwide distribution: USA and Germany, France, Italy, Norway, Sweden, Finland, The Netherlands, Switzerland, U.K., Austria, Luxembourg, Ireland, Hong Kong, New Zealand, India, Australia, and Japan. | Terminated | April 02, 2012 | 815 | 72120 | Not_Computer
Z-1735-2010 | 55723 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Voiceware clip software update to certain Horizon Medical Imaging Systems. McKesson Medical Imaging Company. Richmond, Canada V6X 3G5. |  |  | May 07, 2010 | 2010 | Mckesson Medical Immaging | Voice Clip issue: When the user is recording voice clips in Horizon Rad Station, there is a potential for voice clips to be saved to the incorrect study of the same patient. | Software design | McKesson Medical Imaging Company issued an "Advisory Notice" dated May 7, 2010. Consignees were notified of the affected product and advised to contact McKesson Customer Support department to review their sites current configuration and install the appropriate product software update(s). For further information, contact McKesson Medical Imaging Customer Support at 1-800-663-2533. On June 19, 2010 the recall was amended to include the following actions to follow: Instead of highlighting the anchor study in the Patient Portfolio Sidebar before saving t he voice clip in the Patient Portfolio Sidebar, McKesson recommends that you save the voice clip to the desired study first, before browsing through other documents. A software update that prevents the problem from occurring is currently under development for all affected releases and will be available in approximately three months. Consignees will be notified when the software update is ready to be installed on their system. |  | 100 | 100 units | Distributed Nationwide (AK, AL, AR, AZ, CA, CO, FL, GA, IN, IA, IL, KY, LA, MA, MD, MI, NC, ND, NE, NM, NY, NV, OH, OR, PA, RI, SC, TN, TX, WY, VA and WV) and Canada | Terminated | August 15, 2012 | 831 | 100 | Software | Output/Calculation | Software update | Software Update
Z-1734-2010 | 55713 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | ANGIODYNAMICS BENEPHIT(TM) CV Infusion System, 35 cm, Targeted Renal Therapy Infusion system, Catalog/Model/REF Number 60035, STERILE, EO, Intended for Single Use Only. AngioDynamics, Incorporated 603 Queensbury, NY 12804. |  |  | April 30, 2010 | 2010 | Angiodynamics, Inc. | Benephit Infusion Catheters may exhibit a hole or tear in the primary sterile barrier packaging. The product is intended to be sterile. Breaches in the sterile packaging may result in non-sterile product, which could potentially cause a patient infection which is a serious medical condition. | Packaging change control | Angiodynamics, Inc. issued an "Urgent Medical Device Recall" notification (dated 4/28/10). Customers were advised of the affected product and instructed to return all recalled product to the firm.For further information, contact Angiodynamics, Inc. Customer Service at 1-800-772-6446. |  | 22 | 22 | Distributed Nationwide and to Italy | Terminated | July 25, 2012 | 817 | 763 | Not_Computer
Z-1731-2010 | 55661 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Vessel Sizing Angiographic Catheter, REF 7602-2M65, 5F PIG-VSC-2B, UHF Pigtail, Performa, Sterile EO, Merit Medical Systems, Inc., South Jordan, Utah 84095. |  |  | May 05, 2010 | 2010 | Merit Medical Systems, Inc. | Marker bands on angiographic catheters may not be properly attached, allowing movement of the bands. | Process control | Merit Medical Systems, Inc. began notifying sales representatives beginning May 5, 2010 via phone and instructed to contact all of their affected accounts. All affected product was to be returned to the firm.For further information, contact Merit Medical at 1-713-757-5067 or 1-801-208-4468. |  | 96 | 96 units | Nationwide Distribution -- GA, IL, MO, NY, PA and WI. | Terminated | June 01, 2010 | 27 | 96 | Not_Computer
Z-1736-2010 | 55659 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | EVA 450 Patient Lift Battery Charger. |  |  | April 28, 2010 | 2010 | Romedic, Inc. | Incorrect battery chargers were supplied with the EVA 450 Patient Lifts. Charger unit stops working and will not charge lift. New chargers were supplied to customers. | Other | The recalling firm sent out a letter with instructions as follows: 1) Locate EVA 450 Patient Lift(s), 2) Fill out Customer Response Form and fax to 888-691-1091, 3) Replace battery charger(s) with new battery charger(s) supplied with this notification; NEW battery charger to be used according to EVA 450 Instructions, 4) Repackage replaced battery charger(s) in original packing in which shipment was recieved, 5) Apply supplied sticker "EVA 450 Battery Charger RECALL" on outside of box, 6) Apply supplied UPS return label on outside of box and contact UPS for pickup and mail to the following Address: RoMedic, Inc., 300 West Chestnut St., Ephrata, PA 17522; Attn: EVA 450 Battery Charger Recall.Additional questions are directed to the company at phone: 888-625-4343. |  | 76 | 76 units | Nationwide distribution. | Terminated | September 30, 2010 | 155 | 76 | Not_Computer
Z-1737-2010 | 55234 | No | GDY | Nonabsorbable gauze for internal use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ARTERIAL LINE KIT Arterial Line Kit #2 20ga X 3" REF/REORDER NO. 194284 Labeled Sterile, Rx. Argon Medical Devices, Inc. 1445 Flat Creek Road Athens, TX 75751 800-927-4669. |  |  | March 29, 2010 | 2010 | Argon Medical Devices, Inc | There are 1-2 places on some of the multi-purpose kit blisters/trays in which there are holes and/or the material is so thin that holes are easily created. This would potentially cause the sterility of the products to be compromised. | Package design/selection | Argon Medical Devices issued an "Urgent Medical Device Recall" notification via certified mail dated March 29, 2010 to distributors and end users. A Response Form was attached. Consignees were informed of the affected product and asked to quarantine and return all unused product to the firm.For further information, contact Argon Medical Devices at 1-903-677-9375. |  | 10 | 10 kits | Distributed to United States (IL, AL, NY, TX, OK, MI, NC, CA, OH, VA, KY, CO, LA, WA, MO, AND IA) and Canada. | Terminated | July 29, 2010 | 122 | 199455425 | Not_Computer
Z-1747-2010 | 55387 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage 2.2 workstations with software version 22_01 and 22_02 |  |  | February 09, 2010 | 2010 | Ge Healthcare, Llc | GE Healthcare has recently become aware of inaccurate Scale factor annotation provided by SenoAdvantage 2.2 workstations that may impact patient safety.When using the Print Image or Print Screen feature on SenoAdvantage 2.2, the Scale factor annotation printed on hard copy images is inaccurate. Secondary Capture (SCPT) images displayed on the screen and Secondary Capture hard copy images may a | Software design | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated February 8, 2010 to customers. The letter described the Safety Issue, Affected product details, Safety Instructions to prevent further inaccurate scale factor annotation, Product Correction and Contact Information. GE Healthcare will correct all affected workstations by providing a software upgrade and an updated Operator Manual. The GE service representative will contact customer to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800-437-1171, Japan: -0120-055-919. For other countries, please contact your local GE Healthcare Service Representative. |  | 73 | 73 devices | Worldwide distribution: USA including states of CA, FL, GA, MD, NJ, NM, NY, NC, OK, TX and countries of THAILAND, TAIWAN, SWEDEN SPAIN, SAUDI ARABIA, RUSSIAN FEDERATION, PORTUGAL, PAKISTAN, MEXICO, JAPAN, ITALY,GREECE, FRANCE, COLOMBIA, CHILE, CANADA, BELGIUM, and BAHRAIN. | Terminated | December 18, 2011 | 677 | 73 | Software | Display/Image | Software update | Software Update
Z-1746-2010 | 55613 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Trestle Anterior Cervical Plate 1-level assembly, Part Number: 61001-026. Alphatec Spine, Carlsbad, CA 92008. |  |  | August 25, 2009 | 2009 | Alphatec Spine, Inc. | The recall was initiated after the firm discovered that the affected lot was labeled incorrectly in that the packaginglabel indicates that the part number of the lot is 61002-026, which corresponds to a two-level Trestle Anterior Cervical Plate implant. | Labeling mix-ups | Alphatec Spine initiated a recall communication via telephone beginning August 25, 2009. Customers were given an explanation of the problem and instructed to check their inventory for the affected product. The firm indicated that they would arrange to have Federal Express contact customers for pick-up of the affected product(s) and replacement product would be sent out immediately.For further information, contact Alphatec Spine at 1-800-922-1356. |  | 13 | 13 units | Nationwide Distribution -- AL, CA, FL, TN, TX and UT. | Terminated | June 07, 2010 | 286 | 13 | Not_Computer
Z-1748-2010 | 53856 | No | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | PLAC Test Reagent Kit, immuno-turbidimetric method, Part Number 90107, 100 test kit, manufactured by DiaDexus, South San Francisco, CA. |  |  | October 29, 2009 | 2009 | DiaDexus, Inc | Inaccurate results-- some samples may return with falsely elevated results potentially as a result of an interfering substance. Some patients may be subjected to additional lifestyle changes or intensified pharmaceutical therapy; such as statins, as a result of the falsely elevated result. | Other | Notification letter issued to consignees on 10/29/2009. Customers were instructed to not report PLAC test results for samples that are greater than 360 ng/mL, and pursue confirmatory testing by contacting DiaDexus. Customers with questions please call the company technical support line toll free at 1-877-752-2837 for assistance.UPDATE: On April 30, the firm suspended marketing of the PLAC TIA method, as the limited range no longer covered the clinical data/reference range data cited in the product insert. The firm sent out a letter notifying customers on May 10, 2010. The firm set up a sample send out program, so that customers can send samples to labs running the ELISA method PLAC test. Arrangements were made so that labs could be enabled to run a CLIA complex assay, for customer credit programs for un-used products, and a fax back response form for response tracking. |  |  | all lots | Product was distributed to 123 consignees throughout the US. | Terminated | January 26, 2011 | 454 |  | Not_Computer
Z-1754-2010 | 55527 | No | CAH | Breathing circuit bacterial filter. | Anesthesiology | General Hospital | 510(k) | Xenon Diagnostic Circuit Kits labeled as "Xenon Diagnostic Circuit", "Xenon Circuit", and Needleless Xenon Diagnostic Circuit" Manufactured for RAF MEDICAL, INC., (Model number's RAF2306, RAF2304, RAF2380, RAF2300, RAF2370, RAF2372, RAF2350, RAF2341, RAF2302, RAF2340 |  |  | April 16, 2010 | 2010 | Continental Medical Labs,Inc | CML, Inc. was notified by A-M Systems in a letter dated March 31, 2010 that they were recalling a component of their Xenon Kit (ViroMax viral and bacterial filter) because the filters may have a leak at the seam where the two halves of the filter are bonded together. | Release of Material/Component prior to receiving test results | The firm, A-M Systems, sent two "URGENT: Notification of Product Recall/Correction" letters one dated March 31, 2010 and the other dated April 7, 2010 to customers. The letters described the product, problem and action to be taken by customers. Also enclosed was a list of the affected products that was sold to the consignee. The customer was instructed to cease further distribution or use of any remaining products from the listed lots, immediately notify any customers who received products from the lots, provide A-M Systems with the number of items distributed and still remaining in their inventory, once the number is provided the customer should destroy or discard any remaining product in their possession from these lot numbers.Please feel free to contact Robert Thompson by phone 360-683-8300 x302 or email: rthompson@a-msystems.com with any questions, comments or concerns. |  | 5520 | 5520 | Nationwide distribution: WI | Terminated | April 05, 2012 | 720 | 5520 | Not_Computer
Z-1755-2010 | 55693 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | The Baylis LumbarCool Pain Management Kit consisting of: 1 LumbarCool Pain Management Probe, 3 LumbarCool Pain Management introducers, and 1 Pain Management Tube Kit. Model/Catalogue Number: LUK-17-150-4 |  |  | May 12, 2010 | 2010 | Baylis Medical Corp * | Name of the device reflected on the product packing sleeve is incorrect. | Error in labeling | The firm, Baylis Medical, notified US Distributor by email and sent a "URGENT Field Corrective Action Notice" dated May 12, 2010, to customer. The notifications described the product, problem and action to be taken by customer. The customers were given the follow options: Option 1: In order to prevent confusion in identification of the product before use, please immediately remove and dispose the package sleeve that covers the tray packaging of the LumbarCool Pain Management Kit.Option 2: Alternatively, you may exchange your kit from Lot LKFA160310 with a correctly labeled replacement LumbarCool Pain Management Kit by contacting Kimberly-Clark Health Care sales representative. The customers were ask to fill out the Field Corrective Action Acknowledgment and return it.Any question please call Baylis Sales Representative at 905-602-4875 and/or contact your local sales representative for any further information you may require. |  | 58 | 58 | Nationwide distribution: TX | Terminated | October 04, 2010 | 145 | 58 | Not_Computer
Z-1757-2010 | 51812 | No | IOL | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Treadmill models TM55, TM65, and CR60 (brand names include Quinton, Burdick, and Cardiac Science). |  |  | March 30, 2009 | 2009 | Cardiac Science Corporation | 28 stress and rehab treadmills manufactured between September 2006 and January 2009 may not have received adequate electrical safety testing prior to being shipped from the factory. If these components break down in the field, a patient or operator could be exposed to excess touch currents on the surface of the product. Because safety testing may not have allowed for complete evaluation of the el | Other | On April 3, 2009 the recalling firms notified consignees by certified letter to domestic customers and via FedEx to international customers with delivery confirmation signature required. Customers were advised to remove treadmill from service until it is tested by a Field Service Engineer.For customers with questions or those that have not already scheduled a service call with Cardiac Science Field Service Engineers, please contact the firm at 1.800.426.0337, option 1. |  | 28 | 28 treadmills | Distributed domestically (including Puerto Rico) to 20 end-user accounts and 2 distributor accounts;Distributed internationally to: Canada, Mexico, Portugal, and Saudi Arabia. | Terminated | January 10, 2012 | 1016 | 28 | Hardware | Treatment/Delivery/Therapy | Remove | Remove or Replace
Z-1758-2010 | 55166 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COULTER PrepPlus, Part Number: 286600 |  |  | January 21, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports of potential dilution of reagent vials loaded onto the Preppies or PrepPlus 2 systems. Affected systems have a syringe pump with part number A46748 (see illustrations below) which cause the probe to prematurely dispense Isoflow prior to arriving at the wash station.This condition only occurs with worklist panels which include both | Other | A Product Corrective Action (PCA) letter with attached PCA Response Form was sent on the week of January 25, 2010 to all customers who purchased the pipetting and diluting system: COULTER PrepPlus, or COULTER PrepPlus 2. The letter provides the customers with an explanation of the problem identified and an instruction on proper placing antibody reagent when using a worklist panel which involves both blood specimens and antibody reagents. Customers were also instructed to complete and return the PCA Response Form in addition to sharing the information contained in the PCA with their laboratory staff. If consignees had any questions they were told to contact their local Beckman Coulter Representative. Questions can be directed to the company's Customer Service office at (800) 526-7694. |  | 193 | 193 units in total | Domestically and Canada. | Terminated | April 04, 2012 | 804 | 193 | Not_Computer
Z-1766-2010 | 55701 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DURAMAX (TM) Hemodialysis Catheters, Stacked Tip 40 cm Straight Basic Kit, Catalog Number/REF 10302807, 510 (k) 080400, Sterile --- Manufactured in USA by AngioDynamics, Inc. 603 Queensbury Avenue, Queensbury, NY 12804 |  |  | January 08, 2010 | 2010 | Angiodynamics, Inc. | The product has the potential for a cross-lumen leak. One of the two leak testers that are used to perform the leak test on the DuraMax Catheters was found to be faulty. The tester was not able to detect cross-lumen leaks. Both testers were used during the manufacturing 100% leak test. | Other | Customers were sent URGENT MEDICAL DEVICE RECALL LETTERS (dated 1/07/10) via Certified Mail, Return Receipt Requested. Customers were instructed to return all recalled product. A return authorization number was assigned to each customer in the recall notification letter and a UPS account was provided for return freight. A reply form was provided in the recall notification letter. Questions were directed to the company sales representatives or to Customer Service at 800-772-6446. |  | 208 | 208 | Nationwide and St. Thomas, VI | Terminated | July 25, 2012 | 929 | 2950 | Not_Computer
Z-1756-2010 | 55807 | No | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | Fancy i Color cosmetic lens, Contents: A pair of sterile soft color contact lens immersed in buffered isotonic saline solution, Size: B.C. 8.6mm DIA. 14.5mm. Dreamcon Co. LTD. Korea. |  |  | May 17, 2010 | 2010 | Kim's Trading Inc | Recall is being conducted because the recalled color cosmetic contact lenses are unapproved medical devices. | PMA | Kim's Trading Inc. verbally contacted the distributor to provide notification of the recall. All affected product was to be properly disposed.For further information, contact Kim's Trading Inc. at 1-484-895-3773. |  | 300 | 300 pairs of contact lens | All product was distributed to one distributor in the Bethlehem PA area. | Terminated | August 11, 2010 | 86 | 300 | Not_Computer
Z-1772-2010 | 55740 | No | FCL | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Boston Scientific Radial Jaw 4 Large Capacity w/Needle Biopsy Forceps |  |  | May 06, 2010 | 2010 | Boston Scientific Corporation | Clevis may detach from the coil preventing the device to open and/or close. Delays in the procedure may result. | Process control | Boston Scientific issued Recall notification letters via Federal Express Priority mail on May 6, 2010. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. In the letter, customers with questions are provided with a phone and email at the firm to contact. |  | 320 | 320 units | Nationwide; Panama. | Terminated | September 20, 2011 | 502 | 320 | Not_Computer
Z-1760-2010 | 55620 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Back-Up Meier Steerable Guidewires, Sterilized using ethylene oxide. |  |  | April 19, 2010 | 2010 | Boston Scientific Corporation | Boston Scientific Corporation is conducting a Medical Device Recall of certain lots/batches of its Back-up Meier Steerable Guidewires. Through their internal inspection process, they identified that the polytetrafluoroethylene (PTFE) coating on the gold plated distal coil of the Back-up Meier Steerable Guidewires of the identified lots/batches have the potential for PTFE delamination. Inconsist | Nonconforming Material/Component | Consignees were sent a Boston Scientific "Urgent Medical Device Recall" letter dated April 19, 2010. The letter was addressed to Risk Manger/Field Action Contact. The letter described the problem and product affected and recommended that to immediately discontinue use of and segregate recalled product. They also requested the consignee to complete and return the Account Reply Verification Tracking Form.Customers with questions are directed to the firm's Field Action Coordinator at (763-) 494-1417. |  | 885 | US 885 single units (177 boxes), OUS 630 single units (126 boxes) | AL, AZ, AR, CA, CT, FL, GA, IL, IN, IA, KS, KY, ME, MD, MA, MS, MO, NJ, NM, NY, NC, OH, OR, PA, SC, SD, TN TX, VT, VA, WA.SWITZERLAND, GERMANY, FINLAND, FRANCE, ITALY, JORDAN, POLAND, SWEDEN, CANADA.Complete consignee list not available at this time. | Terminated | December 10, 2011 | 600 | 885 | Not_Computer
Z-1775-2010 | 54643 | No | IMK | Stair-climbing wheelchair. | Physical Medicine | Physical Medicine | 510(k) | iBOT 4000 Mobility System Battery pack. |  |  | February 22, 2010 | 2010 | Independence Technology LLC | A thermal breaker in the battery pack has the potential to open due to overheating and disable one third of the cells in the affected battery pack. | Device Design | The company sent Medical Device Field Correction letters to all customers on February 22, 2010 by certified mail. |  | 192 | 192 devices | Nationwide and to the United Kingdom. | Terminated | October 01, 2010 | 221 | 192 | Not_Computer
Z-1780-2010 | 53611 | No | OCL | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | St. Jude Medical Epicor UltraWand LP Handheld Ablation Device, model UW-LP-2, Sterilized with Ethylene Oxide. Single-use. The Ultra Wand LP device is designed for use with the Epicor Cardiac Ablation System. St. Jude Medical 240 Santa Ana Court, Sunnyvale, CA 94085 USA |  |  | June 15, 2009 | 2009 | St. Jude Medical Atrial Fibrillation Division Inc | The firm states that overheating of cardiac tissue is caused by inadequate cooling. The potential exists when the UltraWand is run using off-label cooling methods (gravity saline flow instead of pump) that cooling flow to one or both cells can be decreased based on pressure applied by the physician to the device. To reflect this information to the user, the Atrial Fibrillation Division (AFD) has | Other | St. Jude Medical issued a Market Bulletin, dated 06/15/2009, to their Field Representatives who were instructed to contact customers to make them aware that changes were made to the Instructions for Use (IFU) by SJM and that the customer may notice an extra IFU included in the shipping package. All such worldwide regulatory approvals will be needed before a revised IFU can eventually be included directly within the device package (as opposed to the shipping box). Existing instructional materials including the Epicor LP System Set-up ACS hang sheet and the EpicorLP In-service slides are being revised to reflect the UltraWand LP IFU changes The forthcoming Epicor LP In-service video is also being updated to reflect these changes. An Important Medical Device Safety Information letter, dated 11/06/2009 was sent via FedEx beginning 11/10/2009 to all US customers and will also shipped to all OUS customers beginning 11/13/2009. A copy of the new IFU will be sent with the letter which is being sent to all customers that have purchased or consigned the SJM UltraWand LP product since its commercial release in July of 2008. The customer notification includes the SJM Marketing managers and director for questions on the action. |  | 1813 | 1813 (1155 US, 658 OUS) | AL, AZ, CA, CO, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI. OUS to include Australia, Canada, Japan, Korea, Singapore | Terminated | March 02, 2012 | 991 | 1813 | Not_Computer
Z-1777-2010 | 55638 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MoblVac Portable Wound Care Pump; an AC/DC powered portable aspirator that provides negative pressure for the application of wound drainage; Ohio Medical Corporation, Gurnee, IL 60031; Models 769600 (with Truseal), 759600 (refurbished), 769600 Rental, 769601 (rental), 769602 (evaluation pump with Truseal) and 769610 (sales unit with case). |  |  | May 10, 2010 | 2010 | Ohio Medical Corporation | Ohio Medical has received a report of an explosive event with a MoblVac Portable Woundcare Pump housing where parts were separated while in an enclosed nylon transport bag during use. | Other | Ohio Medical sent Urgent Medical Device Recall letters dated 5/7/10 to the direct accounts on 5/10/10, with follow-up telephone calls on 5/11/10. The accounts were informed of the one report of an explosion of the pump, and were requested to take the pumps out of service and return the pumps to Ohio Medical. The accounts were requested to complete and fax back to Ohio Medical the enclosed acknowledgement form indicating the number of pumps being returned. Any questions were directed to Anita Christ at 847-855-6270. Follow-up letters dated 5/13/10 were sent to the accounts on 5/14/10, providing instructions for the return of the pumps and bags. Shipping boxes will be sent to the accounts based on the number of pumps reported on the acknowledgment form. |  | 375 | 375 units | Nationwide and internationally to Qatar and Oman | Terminated | August 19, 2011 | 466 | 593 | Not_Computer
Z-1776-2010 | 55505 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | The Alternans Sensor Kit Assembly is a sealed kit that contains the components necessary for conducting a Microvolt Twave Alternans test. |  |  | April 29, 2010 | 2010 | Cambridge Heart, Inc. | Some kits contained a component with expiration date prior to the date of expiry on the outside of the kit. The user receives feedback of potential unreliable results at several stages during the test process, so any risk of unreliable results is low. | Process control | The Alternate Sensor Kit Assembly P/N 203270209 may contain a pouch of standard ECG electrodes with an expiry date which is earlier than the expiration date on the outside label of the kit. All customers were sent a letter dated April 23, 2010 with information that included instructions for returning the kits. Questions are directed to the company's Quality Manager and Recall Coordinator at phone number 978-654-7625. |  | 2740 | 2740 | Nationwide | Terminated | September 20, 2011 | 509 | 2740 | Not_Computer
Z-1779-2010 | 55637 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | RAYHACK DRILL GUIDE ANGLED, REF 4010-0206, 1 EACH, Rx ONLY, NON-STERILE, Wright Medical EMEA Krijgsman 11, 1186 Amstelveen, The Netherlands, Wright Medical Technology, 4677 Airline Road, Arlington, TN 38002 |  |  | April 08, 2010 | 2010 | Wright Medical Technology Inc | Drill guides from 2 lots may not seat properly on the compression block. | Other | The firm, Wright Medical Technology, Inc., sent an "URGENT: Medical Device Voluntary Recall Letter" on April 8, 2010 via FedEx to customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to please examine their inventory of the products and check it against the information provided, and to immediately return any affected products to Wright as soon as possible. The customers were also instructed to contact Customer service at 800-238-7117 for return instructions and replacement inventory. In addition, the customers were to immediately complete and return by fax at 901-867-7401 the attached response confirming receipt of the notice.If you have any questions regarding this matter, please contact Debby Daurer at 800-874-5630. |  | 13 | 13 units | Worldwide: USA and The Netherlands | Terminated | May 17, 2011 | 404 | 13 | Not_Computer
Z-1647-2010 | 54884 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Bard Mesh Monofilament Knitted Polypropylene; Contents: size 2" x 4", 3 units per box, Davol Inc., A Subsidiary of C.R. Bard, Inc. 100 Sockanossett Crossroad Cranston, RI 02920. |  |  | March 05, 2010 | 2010 | Ram Medical Inc | Counterfeit product mixed with authentic product. The problem was recognized by the mismatching of expiration dates and the subtle differences in packaging. | Nonconforming Material/Component | Ram Medical Inc., contacted Customers by phone beginning March 4, 2010 and issued "Urgent Device Recall" notifications beginning March 8, 2010. Further notifications were issued to expand the scope of the recall by the firm. Each notification described the affected product and asked the User to complete a Recall Acknowledgement Form and return it to the firm. All affected product should be returned to the firm.For further information, contact Ram Medical, Inc. at 1-973-633-0400 Monday through Friday, 9:00 AM to 5:00 PM EST. |  | 306 | 102 boxes of 3 units each | Nationwide Distribution -- CO, IL , NY, MO and FL. | Terminated | June 06, 2016 | 2285 | 3720 | Not_Computer
Z-1784-2010 | 55256 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Modular Foot System 3.5mm x 22mm Locking Screw M/N: MFT-021-35-22. OrthoHelix Surgical Designs, Inc. Medina, OH 44256. |  |  | March 23, 2010 | 2010 | Orthohelix Surgical Designs Inc | On 3/9/10 it was discovered that the recalling firm packaged MFT-021-35-24 Screws from lot 0964092 as MFT-021-35-22 lot 0964081. | Employee error | OrthoHelix Surgical Designs, Inc. issued an "Important Field Correction Notice" dated March 23, 2010 via FedEx overnight. They instructed their customers to inventory their stock for affected product. All affected product will be replaced by the firm.For further information, contact OrthoHelix Surgical Designs, Inc. at 1-330-247-1444. |  | 62 | 62 pieces | Nationwide Distribution -- AZ, CA, CO, FL, GA, ID, IL, MD, NC, NJ, NM, NY, OH, OK, PA, SC, TX, UT, VA and WA. | Terminated | September 10, 2012 | 902 | 62 | Not_Computer
Z-1783-2010 | 54630 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova StatStrip Glucose Test Strips, Catalog Number: 42214. Nova Biomedical Corporation. Waltham, MA 02454. |  |  | February 12, 2010 | 2010 | Nova Biomedical Corporation | Glucose strips report low glucose results | Nonconforming Material/Component | Nova Biomedical notified it's domestic Consignees by a Customer Advisory Notice by phone and follow up by email and/or fax dated February 12, 2010. The firm will replace all affected inventory at customer sites via overnight shipment of replacement Test Strips to customer site. For further information, contact your local Nova Support Department or call 1-800-545-6682. |  | 392 | 392 Cases (1800 strips each) | Worldwide Distribution -- United States (CO, MA, MN and RI), United Kingdom, Germany, France and Italy. | Terminated | July 22, 2011 | 525 | 392 | Not_Computer
Z-1782-2010 | 55497 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ReviveR AED semiautomatic external defibrillator used with the DBP-2800 Battery Pack. Manufactured by Defibtech, LLC. Seymour, CT 06483. |  |  | May 21, 2010 | 2010 | Defibtech LLC | The AED when used with an affected battery pack, may falsely detect an error condition, cancel charge and not provide therapy | Device Design | Defibtech, LLC issued an "Urgent Medical Device Safety Information and Correction" notification dated May 20, 2010 by email and hard copy via US Mail. The notification instructed Consignees on proper identification of affected product and steps to take for product replacement.For further information, contact your distributor or Defibtech Technical Support at techsupport@defibtech.com or call 1-877-453-4507 or 1-203-453-4507, 8:30 AM to 5:00 PM Monday - Friday, EST. |  |  |  | Worldwide Distribution -- United States, Canada, Australia, Jordan, Spain, Israel, UK, Poland, Italy, Netherlands, Malaysia and Ireland. | Open, Classified |  |  |  | Not_Computer
Z-1785-2010 | 55572 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Passport V Monitor: Mindray, North America. Mindray DS USA, Inc. Mahwah, NJ 07430. |  |  | April 28, 2010 | 2010 | Mindray DS USA, Inc., dba Datascope Patient Monitoring | An issue with the Passport V Monitor has been identified where the Arrhythmia Analysis, ST Analysis and Invasive Blood Pressure functions are not operational. | Software design | Mindray issued Corrective action letters directly to all customers beginning April 28, 2010 by certified mail, return receipt requested. consignees were informed of the affected product.For further information, contact Mindray DS USA, Inc. at 1-201-995-8391. |  | 41 | 24 US, 17 International | Worldwide Distribution -- United States (OH, TX, AZ, NY, MO, IL, FL and TN), Spain, Indonesia, Brazil, Italy, Columbia and British West Indies. | Terminated | June 22, 2011 | 420 | 41 | Software
Z-1788-2010 | 54276 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Triage Total Controls 5, Level 1: Five 0.25 mL tubes. Storage Temp &lt;-20C. |  |  | February 06, 2009 | 2009 | Biosite Inc | Incidence of low recovery results has exceeded post market surveillance thresholds of customer complaints in the field. | Other | The firm, Biosite Incorporated, sent an "URGENT MEDICAL DEVICE RECALL" dated February 6, 2009, to all customers. The letter describes the product, problem and action to be taken by the customers. The customers were instructed to immediately discontinue all use and/or sale of the listed products, discard all affected product in accordance with their local regulations, complete and fax the enclosed Verification Form within 10 days to confirm their receipt of the notice, to indicate the number of kits discarded from their inventory and the product replacement will be sent to them. Contact Inverness Medical Technical Service at 1-877-441-7440 and select option 2 from the phone prompt and a representative will confirm the number of kits discarded from your inventory and arrange for replacement product to be sent to you.Should you have any questions about the information contained in this notification, please contact (877) 441-7440 or email: Recalls@biosite.com |  |  |  | Nationwide distribution | Terminated | September 01, 2011 | 937 |  | Other |  |  | Other
Z-1801-2010 | 55352 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE" Monitor B850, Software version 1.0.1.13 and 1.0.2.1 |  |  | May 06, 2010 | 2010 | Ge Healthcare, Llc | GE Healthcare has become aware of a safety issue associated with the CARESCAPE" MonitorB850 when monitoring ECG. The safety issue includes the potential for a delay in treatment. | Software design | The customers were sent on May 6, 2010, a GE Healthcare "Urgent Medical Device Correction" letter dated April 22, 2010. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, Director of Biomedical Engineering. The letter listed the Safety Issue, Affected Product Details, Product Correction, Safety Instructions and Contact Information. The customers were instructed to immediately discontinue use of the "Update Lead Set" function until the device software has been upgraded.If you have any questions regarding this medical device correction or identification of affected items, please contact Technical Support at 1-800-558-7044, or your local service representative. |  | 798 | 798 units. (229 USA, 569 OUS) | Worldwide distribution: USA including CO, MA, NC, WI, and countries of UNITED KINGDOM, SWITZERLAND, SWEDEN, SPAIN, SINGAPORE, PORTUGAL, NORWAY NETHERLANDS, ITALY, INDIA, GERMANY, FRANCE, FINLAND, BELGIUM, and AUSTRALIA. | Terminated | June 05, 2012 | 761 | 798 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-1798-2010 | 54652 | No | JPI | Hematocrit measuring device. | Hematology | Hematology | 510(K) Exempt | i-STAT E3+ cartridges |  |  | February 05, 2010 | 2010 | Abbott Point of Care Inc. | Abbott has determined that there is a dependence of the i-STAT Hct assay on the vertical pitch angle of the i-STAT Analyzer during the blood analysis cycle. | Other | The firm, Abbott Point of Care, sent an "Urgent Field Safety Notice Product Recall" letter dated January 2010 to the customers. Safety Alert /Product Recall letters were also sent by Federal Express commencing on February 5, 2010. The letter describes the product, problem and the action to be taken by the customers. The customers were instructed to replace the i-STAT Hematocrit/HCT and Calculated Hemoglobin/HB /Cartridge and Test Information (CT) Sheet found in their i-STAT System manual with those included in the mailing , and to provide a copy of the letter another facility, if they have forwarded any cartridges. If you should have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-366-8020 or your Abbott Point of Care representative. |  | 40446 | 40,446 total cartridges, all varieties | Worldwide distribution: USA and Argentina; Austria; Bahamas; Bolivia; Brazil; Chile; Hong Kong; China; China; Columbia; Costa Rica; Denmark; Dominican Republic; Ecuador; El Salvador; Finland; Georgia; Greece; Honduras; India; India; Ireland; Israel; Italy; Japan; Kuwait; Mauritius; Mexico; Norway; Oman; Panama; Paraguay; Peru; Poland; Portugal; Qatar; Russian Federation; Saudi Arabia; Singapore; South Africa; South Korea; Spain; Switzerland; Taiwan; Thailand; Trinidad and Tobago; Turkey; United Arab Emirates; Uruguay; Yemen; Netherlands Antilles; Venezuela; Australia; Belgium; France; Germany; Italy; New Zealand; Spain; Sweden; and The Netherlands. | Terminated | July 25, 2012 | 901 | 40446 | Not_Computer
Z-1786-2010 | 55404 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | directCHECK Whole Blood Control; |  |  | April 15, 2010 | 2010 | International Technidyne Corporation | Indication not cleared or approved. The acceptable performance range published in the directCHECK Microcoagulation Quality Control package insert is incorrect. | Error in labeling | Notification letters were sent vis Federal Express on April 15, 2010 to all distributors and end users to return unused product or to replace the package insert. Please direct questions ITC Technical Support at phone 800-631-5945 Ext. 4707. |  | 94 | 94 boxes | Distributors and medical facilities in AZ, CA, IL, KY, MA, MD, ME, MO, PA, WA. Foreign distribution to Australia, Canada, Switzerland, Germany, Spain, Quatar, and Saudi Arabia. | Terminated | September 23, 2010 | 161 | 94 | Not_Computer
Z-1800-2010 | 55674 | No | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | InviroSnapÂż Retractable Safety Syringes, Rx Only, Single use, Sterile, Inviro Medical Devices, Inc., Duluth, GA 30096 USA. |  |  | May 10, 2010 | 2010 | Inviro Medical Devices, Inc | Syringe Defects. Group 1 syringes: The plunger pull out force was below specification. Group 2 syringes: The needle length is out of specification. | Other | Consignees were notified by letter on or about 05/10/2010. Consignees were instructed to review their inventory of InviroSnap Safety Syringes for the lot numbers and remove them from inventory and to return the affected product to Inviro Medical. They were instructed to pass on the information to anyone who uses or orders Inviro Snap Safety Syringes at their facility and to ensure a copy of the letter is provided to any other organizations to which affected devices have been transferred. A Response Card was included to be completed and faxed back to Inviro Medical Devices even if there is no product in inventory to verify receipt of notification. Questions regarding the recall were directed to the company's Customer Service number at (866) 914-2259. All product from Group 1 will be scrapped. Product from Group 2 will be retained in Quarantine until such time as the product can be exported the permission of the importing countries Ministry of Health. |  | 623000 | 623,000 units | CA, FL, IL, MI, MO, MN, NC, OH, TX; and Canada | Terminated | December 08, 2011 | 577 | 623000 | Not_Computer
Z-1787-2010 | 55697 | No | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | Difco (TM) Decarboxylase Base Moeller, packaged in 500 gm bottles, labeled in part ***Becton Dickinson Microbiology Systems, Becton Dickinson, Sparks, MD 21152*** |  |  | December 09, 2009 | 2009 | Becton Dickinson &amp; Co. | Diagnostic test reagents may not meet quality control specifications. | Other | The recalling firm notified distributors by letter flagged as "Urgent Product Recall" on 12/09/09. The letter instructed distributors to discontinue distribution and discard product. Distributors were asked to provide customer lists and the recaller subsequently notified the end users to discontinue and discard identified lots for replacement. Customers were asked to return a response form.If further assistance is needed regarding replacements, please contact BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663. |  | 209 | 209 bottles | The product was sold to distributors in the US and to affiliates internationally. A complete list is not currently available. | Terminated | June 14, 2010 | 187 | 209 | Not_Computer
Z-1799-2010 | 55658 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DT2 Mammo, Size 8 x 10 inch, Ref No. E2RN6. Agfa Corporation. Greenville, SC 29601. |  |  | April 22, 2010 | 2010 | AGFA Corp. | The RF tags on the white protection sheets of the media stacks were incorrect. | Nonconforming Material/Component | AGFA Healthcare issued an "Urgent Safety Notice" dated April 22, 2010 requesting a check of warehouse inventory and a request for a sub-recall, if necessary. All affected product should be returned to the firm for replacement.For further information, contact AGFA Healthcare at 1-877-777-2432. |  | 180 | 180 boxes | Worldwide Distribution -- United States (AL, CT, FL, KS, MA, MI, NJ, NM, NY, PA and WV), British Columbia and Guatemala. | Terminated | February 01, 2011 | 285 | 180 | Not_Computer
Z-1802-2010 | 53126 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medtronic Racer Biliary Stent System, Model number XD518YF, UPC code 00721902758164 in a box containing sterilized pouch with product, manufactured by Medtronic Vascular, Sunnyvale, CA. |  |  | July 30, 2009 | 2009 | Medtronic CardioVascular | Mislabeling-- The incorrect Instructions for use may have been included with this product. | Mixed-up of materials/components | The firm, Medtronic, Inc., sent an "URGENT- VOLUNTARY MEDICAL DEVICE RECALL" dated July 30, 2009, to custumers (Sales Representatives, to health care provider, patient, and risk manager, hospital administrator). The letter describes the product, problem and the action to be taken by the customers. The customers were instructed to immediately segregate any Racer Biliary units from model number XD518YF (L#0000705315) in their possession to prevent inadvertent use. Medtronic, Inc., representative will contact the customer to review the labeling and retrieve the device should it be in their inventory. The customers are being ask to cooperate by providing the representative with access to the product.If you have any questions, please contact 707-566-1548. |  | 13 | 13 | Worldwide distribution: USA and Greater China. | Terminated | January 31, 2011 | 550 | 13 | Not_Computer
Z-1803-2010 | 55624 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | The SwitchPoint Infinity 2 (SPI2) and SwitchPoint Element (SPE), an integrated voice, video, data router and teleconferencing interface for the operating room. |  |  | April 30, 2010 | 2010 | Stryker Communications Corp | Users may experience issues with the functionality of the SwitchPoint Infinity 2 and SwitchPoint Element that may be exhibited as a "Page Not Found" error message, missing drop down options, or a frozen touch panel. | Component design/selection | The firm, Stryker, sent an "URGENT: Medical Device Correction" letter dated May 4, 2010, to all customers. The letter describes the product, problem and action to be taken by the customers. The firm recommended that the customer continue to use their SwitchPoint Infinity 2 or SwitchPoint Element. The customer is instructed to refer to their user manual, in the event they experience any failures; drive on module failure and/or product failure during a procedure. A Stryker representative will contact the customers facility to coordinate a service event to replace the existing component with a new component at no cost. Should the customer experience any of the aforementioned failures and require urgent technical support, please call Stryker Technical Support at 1-866-841-5663.If you have any questions regarding this letter or the implementation of this correction, please contact Shanta McKinzie at (972) 410-7319 or via email at StrykerCommunicationsRA@stryker.com. |  | 1118 | 1,118 devices | Worldwide distribution: USA and countries of Australia, EMEA, Canada, China, S. Africa, India, Japan, Latin America, Korea, Mexico, New Zealand, Taiwan, and Signapore. | Terminated | April 25, 2011 | 360 | 1118 | Software | Display/Image | Replace | Remove or Replace
Z-1805-2010 | 55698 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | Calcium Sulfate Hemihydrate 0.5 gram kits. |  |  | May 14, 2010 | 2010 | Ace Surgical Supply Co., Inc. | Fast set 4% potassium sulfate solution may not be sterile. For customers who have used the Calcium Sulfate Hemihydrate kit without the Fast Set solution, then these potential risks identified do not apply. Risks of adverse health consequences with the Fast Set 4% potassium sulfate solution include tissue inflammation and infection. | Process control | Ace Surgical has become aware through internal review of possibility sterility issues with one of the components included in the Calcium Hemihydrate 0.5 and 1.0 gram kits.Customers with questions should contact ACE Surgical using the recall hot line at (800) 441-3100 Ext. 204. |  | 4133 | 4133 | Class 2 recall - Worldwide distribution . Item Number 505-2001: South Africa, Australia, Vietnam, Ontario, Canada, Italy, Lebanon, Petaling Jaya, Malaysia, Portugal, and West Indies.Item Number 505-2002: Australia, Bermuda, Canada, Lebanon, Petaling Jaya, Malaysia, and Portugal. | Terminated | February 02, 2012 | 629 | 12063 | Not_Computer
Z-1804-2010 | 55438 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Medullary Tubes |  |  | April 01, 2010 | 2010 | Synthes USA (HQ), Inc. | After repeated sterilizations product may become discolored and break including while in the medullary canal. Synthes has determined that the Medullary Tube will be a single use item that is to be discarded after one sterilization cycle. | Other | The recalling firm issued an Urgent medical Device Recall letter dated 4/9/10 to all trauma hospital accounts informing them of the problem and the need to return the products. All sales consultants were notified via email and letter.Questions are directed to the company at 800-620-7025 x 5452 or 610-719-5452, or contact your Synthes Trauma Sales Consultant. |  | 18505 | 18,505 tubes | The product was shipped to medical facilities nationwide. | Terminated | March 25, 2015 | 1819 | 18505 | Not_Computer
Z-1808-2010 | 55724 | No | MXG | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Harris Medical UV Lamp; the device does not have a label identifying the name and place of business or manufacturer, packer or distributor. It is packaged with a "Harris UV Lamp Training Manual" |  |  | May 13, 2010 | 2010 | Harris Medical Resources | The Harris UV Lamp was promoted for sub-lingual treatment of HIV/AIDS; Hepatitis B; Hepatitis C; Malaria, Typhoid, Amoeba Infections and other blood diseases. The product was marketed without a 510(k), or an approved application for premarket approval (PMA) or an approved application for an investigational device exemption (IDE). Questions are directed to the firm at 773-583-7500. | PMA | Urgent Device Recall letters were sent registered, certified mail, return receipt requested, to the direct accounts on 5/13/10 - 5/18/10. The accounts were informed that Harris could make no claims as to the safety or efficacy of the Harris Medical UV Lamp in their possession and could recommend no medical indications for use because the necessary regulatory filings have not been completed and reviewed by the FDA, and that Harris Medical has no approved investigation clinical trials underway for the Harris Medical UV Lamp. The accounts were requested to check their inventory and quarantine any Harris Medical UV Lamps they may have in their possession and return them to Harris Medical using the enclosed pre-paid UPS shipping label. If the accounts further distributed the device, they were instructed to sub-recall them to the user level. The accounts were also requested to complete and return via mail or fax the enclosed return response form acknowledging the receipt and understanding of the recall letter and indicating the disposition of the devices. Any questions were directed to Thomas Perez at the company, at 773-583-7500. |  | 20 | 20 units | Florida, Hawaii, Ohio, Texas and internationally to Canada, Dominican Republic, Israel, Kenya and Peru | Terminated | December 10, 2010 | 211 | 20 | Not_Computer
Z-1807-2010 | 55410 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ulnar Shorting Saw Blade is labeled in part: "Congruent Bone Plates, SIZE Ulnar Shortening Saw Blade - Hub Style S, REF 80-0499-S, Material: Stainless Steel, STERILE, Rx Only, ACUMED, Hillsboro, OR, USA". |  |  | April 14, 2010 | 2010 | Acumed LLC | The saw blades may not cut aggressively enough for their intended application therefore creating a risk that the saw could generate excessive heat, which may contribute to bone necrosis and/or non-union. | Other | On April 14, 2010, Acumed began contacting consignees by telephone. On April 15, 2010, Acumed sent out the URGENT NOTICE: DEVICE RECALL Initial Notification letter and Recalled Product Status Form to their consignees. The letter describes the recall product as Ulnar Shorting Saw Blade-Hub Style S, Part no. 80-0499-S and Lot Codes. The Initial Notification letter instructs consignees to return the product back to Acumed immediately and notify their sub-accounts (customers) if products have been further distributed. The consignees also are asked to fill out and send back the Recalled Product Status Form .Customers should call the Customer Service at 1-888-627-9957 or e-mail to RecallNotification@acumed.net for any questions regarding the recall. |  | 718 | 718 units | Worldwide distribution: USA and to the following countries: Spain, Great Britain, South Africa, Puerto Rico, Canada, Turkey, Ireland, Finland, Belgium, Australia, Germany, and Sweden. | Terminated | June 28, 2010 | 75 | 718 | Not_Computer
Z-1811-2010 | 55649 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | CS100 Intra-Aortic Balloon Pump. |  |  | May 10, 2010 | 2010 | Datascope Corporation | Certain MAQUET/Datascope Intra-Aortic Balloon Pumps (IABP) have the potential to experience display related issues. Although the IABP continues to deliver therapy to the patient, display related issues may cause the user to be unable to view the IABP on the information screen. | Other | Medical Device Recall Letters were sent to Hospital Administrators beginning May 10, 2010 by overnight mail.Customers with questions are directed to contact your local Sales/Service Representative or the company representative at 973-244-6314. |  | 161 | 161 total devices | Class 2 - Worldwide distribution -- US and Distribution to Australia, Belarus, Belgium, Canada, Egypt, England, Germany, Hong Kong, India, Iran, Israel, Japan, Kuwait, Latvia, Lebanon, Serbia, Singapore, Spain, Switzerland, Turkey, Venezuela. | Terminated | October 27, 2010 | 170 | 161 | Not_Computer
Z-1810-2010 | 55603 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Cardiac Science Powerheart 9390A automated external defibrillator. This fully automatic model does not require the user to press a shock button in order for the device to deliver therapy (when appropriate). |  |  | December 18, 2009 | 2009 | Cardiac Science Corporation | The K302 relay switch of the G3 series AED product line may experience an early failure. If the K302 contacts are open at the time of rescue, there is a possible delay in therapy or failure to deliver therapy. | Other | Customer notification letters were mailed to US end users and distributors (final consignees) between December 18 and 23, 2009 by USPS certified mail, return receipt requested.Notification letters were sent by FedEx to Canadian distributors (final consignees) on December 23, 2009. Secondary consignee (distributor) letters were mailed on December 23, 2009.Letters advise user to check the status indicator on front of AED to determine if it is rescue ready. Users are advised to keep the AED in service and to perform the Daily and Monthly Scheduled Maintenance as outlined in the Operator and Service Manual.Questions about this communication should be addressed to 1.877.901.1788 toll free from inside the US, or to email |  | 332 | 332 units | Domestic: 323 units to distributors and end users.Canada: 9 units to 2 distributors. | Terminated | December 08, 2011 | 720 | 332 | Other
Z-1816-2010 | 55535 | No | GHC | Erythrocyte sedimentation rate test. | Hematology | Hematology | 510(K) Exempt | Excyte Slider Barcode Tube Adapter used with Excyte Vacuum or Non-Vacuum ESR Tubes. Single Use. Product Number: EX-I0552. Manufactured by Vital Diagnostics S.R.L., Italy. |  |  | April 12, 2010 | 2010 | Vital Diagnostics, Inc. | The tube may fall out or separate from the Slider adapter | Component change control | Vital Diagnostics issued recall letter dated April 14, 2010 via Certified Mail in an envelop marked "Urgent Recall Information Inside", requesting accounts to discontinue use and destroy recalled product. A Certificate of Destruction and Refund Form was included to direct customers. A second Distributor Notice was sent by email to the distributors with Important Recall Information in the title. This was accompanied with the same Technical Bulletin, Certificate of Destruction and Refund Form.For further information, contact Vital Diagnostics at 1-800-345-2822. |  | 4723 | 4,723 units | Distributed Nationwide and Canada. | Terminated | September 13, 2010 | 154 | 4723 | Not_Computer
Z-1815-2010 | 55051 | No | LRJ | General purpose disinfectants. | General Hospital | General Hospital | 510(K) Exempt | CaviBleach Wipes; a disinfectant medical device, active ingredient Sodium hypochlorite 0.525%; a single 9" x 9" towelette, pre-moistened with 0.525% - 0.656% sodium hypochlorite, packaged in a flexible poly foil pouch, 50 wipes per dispenser box, 8 boxes per case; Manufactured for Metrex Research Corporation, 28210 Wick Road, Romulus, MI 48174. |  |  | December 02, 2009 | 2009 | Medtrol, Inc. | The disinfectant wipes were found out of specifications for the disinfectant activity prior to the expiration date. | Process control | Medtrol telephoned Metrex on December 2, 2009, informing them that Medtrol was requesting the recall of the affected product.Metrex sent Urgent Medical Device Recall letters dated December 3, 2009 to their accounts, requesting them to examine their inventory for the affected product, and return any found for credit. They also requested their accounts to sub-recall the product to the retail level. Metrex also requested their accounts to complete and return the enclosed reply sheet, acknowledging receipt of the recall letter and indicating the amount of product being returned. For further information, contact Metrex Customer Care Center at 1-800-841-1428. |  | 150 | 150 cases | Nationwide Distribution -- Including state of Michigan. | Terminated | June 23, 2010 | 203 | 150 | Not_Computer
Z-1814-2010 | 55405 | No | DPZ | Ear oximeter. | Cardiovascular | Anesthesiology | 510(k) | DATEX-OHMEDA TRUSAT PULSE OXIMETER AND ACCESSORIES, MODEL 3500 |  |  | April 29, 2010 | 2010 | Ge Healthcare, Llc | GE Healthcare has recently become aware that theTruSat Pulse Oximeter with external power supply doesnot comply with current safety standards. It is possiblethat the external power supply could leak electricalcurrent. This current could cause third degree burns oran abnormal heart rhythm that, if untreated, could leadto death. | Component design/selection | Consignees were sent on 4/29/10 a GE Healthcare "Urgent Medical Device Correction" letter dated 4/22/10. The letter was addressed to Healthcare Administrator/Risk Manager, Chief of Nursing, Director of Biomedical Engineering, and Home Healthcare Provider. The letter included the Safety Issue, Affected Product Details, Product Correction, Safety Instructions and Contact Information. |  | 28442 | 28,442 units | Worldwide distribution: USA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BANGLADESH,BELGIUM, BOLIVIA, BRAZIL, CANADA, INDIAN OCEAN TERRITORY, CHILE COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON LIBYAN ARAB JAMAHIRYA, LITHUANIA, MACEDONIA, MALAYSIA, MALTA, MEXICO, MOZAMBIQUE, NEPAL, NETHERLANDS, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAK IA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIETNAM, ZAMBIA. | Terminated | December 03, 2013 | 1314 | 28442 | Battery
Z-1890-2010 | 55676 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | TAUT ADAPt Balloon Open Access Port And Syringe 10/12mm x 100mm, REF 41244, Rx Only, Sterile, Teleflex Medical, Research Triangle Park, NC 27709. |  |  | May 12, 2010 | 2010 | Teleflex Medical | Balloon does not hold air or water once inflated. | Other | The firm, Teleflex Medical, sent an "Urgent Medical Device Recall" letter dated May 10, 2010 to customers The letter described the product, problem and action to be taken by the customers. They were instructed to return all of the affected product to Teleflex. Each distributor was instructed to forward the letter to their customers to retrieve relevant product from those locations. A second mailing to non-responding customers will be made approximately four to six weeks after the first mailing. A third mailing to non-responding customers will be made approximately four to six weeks after the second mailing or as an alternative, they may be contacted by phone, email, or fax. A record of these contacts will be maintained. Affected products received by Teleflex Medical will be destroyed upon receipt. Records will be maintained accordingly. Teleflex Medical will communicate to FDA on the status of recall activities.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 366 | 366 eaches | Worldwide distribution: USA including states of AR, CO, FL, GA, IL, IN, KY, MI, MN, MO, NC, NY, OR, PA, WI, WV and country of France. | Terminated | October 27, 2010 | 168 | 366 | Not_Computer
Z-1891-2010 | 52206 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE 2000/2500 Chemiluminescent Substrate Module |  |  | May 15, 2009 | 2009 | Siemens Healthcare Diagnostics Inc. | These substrate lots have shown a tendency to introduce a shift in both control and patient results with multiple methods. | Device Design | An Urgent Field Safety Notice dated May 2009 was sent to consignees of record for the previously mentioned lots. The Notice advised the following:These substrate lots have shown a tendency to introduce a shift in both control and patient results with multiple methods. ActionsIf you are using one of these lots with your instrument, please take the following actions:1. Upon receipt of a non-affected substrate lot, discontinue using IMMULITE 2000/2500 Chemiluminescent Substrate. Prior to receiving a non-affected substrate lot, you can continue to use the affected substrate if you follow the Interim Procedure.2. Discard all bottles of the affected Substrate lots remaining in your inventory once non-affected Substrate lots are available.3. The first time ONLY that you switch from one of the affected substrate lots in Table 1 to a non-affected lot, you should follow the procedure and steps below during instrument Startup:a. Empty the substrate reservoir completely before adding the new lot.b. Prime the instrument to cleanse the lines of the previous lot of substrate.c. Add the new lot of substrate to the reservoir and prime the instrument again.*d. Adjust all methods run on that instrument.*Note: Prime the substrate line as instructed on Page 83, Priming the Substrate Probe , of the IMMULITE 2000/2500 Operators Manual Rev. A.4. Discuss the contents of this letter with your laboratory director.5. Complete the Completion Notification Form that accompanies this notice and return it as directed on the form. Submitting this form indicates that you have received the information in the Urgent Field Safety Notice and have either discarded all lots of the affected IMMULITE 2000/2500 Substrate, or are following the Interim Procedure while awaiting the receipt of a non-affected Substrate lot and will discard the affected lots as soon as the non-affectedlots are received in your laboratory.Interim ProcedureSiemens recognizes that you |  | 41573 | 41,573 | Nationwide Distribution | Terminated | January 18, 2011 | 613 | 41573 | Not_Computer
Z-1899-2010 | 55597 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Horizon Cardiology Hemo Monitoring System |  |  | January 27, 2010 | 2010 | McKesson Provider Technologies - Medical Imaging Group | McKesson has identified a potential hardware configuration problem in some Horizon Cardiology Hemo systems which may cause delay and/or loss of the patient's physiological parameters on the Horizon Cardiology Hemo screen and on the patient record. | Software design | The firm, McKesson Israel Ltd., sent notification letter dated January 27, 2010 to customers. The letter describes the product, problem and action to be taken by customers and the firm. McKesson's will be contacting the customer to coordinate replacement of the sound component of the Horizon Cardiology Hemo system and will provide the additional sound adapter for all affected systems at your facility at not cost. The customers were instructed to pleasel immediately conduct the following instructions to disable sound operations-1) Open McKesson Management Console and locate the Windsurfer&gt;Monitor Settings; 2) Uncheck "Enable QRS beep" in the sounds section; 3) Locate Windsufer&gt;Hardware Settings in McKesson Management Console, and 4) Uncheck "Enable QRS beep" in the Analog Output, QRS Beep Section.If you have any questions, please do not hesitate to contact Ronen Gans directly at (973) 355-9900. |  | 22 | 22 devices | Nationwide distribution: CA, CT, FL, LA, NC, TN, WA, and WV. | Terminated | June 25, 2010 | 149 | 22 | Hardware | Display/Image | Replace component | Remove or Replace
Z-1896-2010 | 55939 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Trilogy 100 Ventilator, Model/Catalog Number 1054260, U1054260 (Domestic). Philips Respironics, Monroeville, PA 15146. |  |  | June 10, 2010 | 2010 | Respironics, Inc. | It is possible that the battery could electrically short, resulting in a potential for fire or a thermal event. | Nonconforming Material/Component | Philips Respironics has contacted all of its consignees by phone and instructed them to remove the potentially affected external battery packs from their inventory and any that may be on their patients/end users. These consignees have been instructed to quarantine batteries they have removed away from flammable materials. Product return/replacement instructions were provided.For further information, contact Philips Respironics at 1-877-387-3311. |  | 23 | 23 units | Worldwide Distribution -- United States (MT, IL, IA, NY, NC, TX, CA and FL), France, South Korea and Turkey. | Terminated | November 18, 2010 | 161 | 23 | Battery | Physical Safety Hazards | Return/Replace | Remove or Replace
Z-1900-2010 | 55673 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | GENESIS II TIBIAL DRILL GUIDE, REF 71440238, SIZE 5, QTY: (1), NON-STERILE, Smith &amp; Nephew, Inc., Memphis, TN 38116, USA |  |  | April 20, 2010 | 2010 | Smith &amp; Nephew Inc | Markings for the drill guides are on the incorrect side: left is printed on right side and right is printed on the left side | Process control | All affected Smith &amp; Nephew consignees (International distributors) were notified of problem and the recall via email on 04/20/2010 and an Urgent: Medical Device Recall letter, dated 4/16/2010. The letter stated the reason for the recall and distributors were requested to immediately cease distribution and use of the affected products. They were also asked to examine their stock and stock of their accounts and quarantine any affected product. The Verification Section of the letter should be completed and returned even if there is no affected stock. Anyone the product was distributed to should also be notified and also return the product. Questions should be directed to Kelly Knight at 901-399-5441. |  | 12 | 12 units | International Distribution Only -- Belgium, Spain, Germany, Australia, and Brazil. | Terminated | December 07, 2011 | 596 | 12 | Not_Computer
Z-1897-2010 | 55857 | No | HLF | Ophthalmic contact lens radius measuring device. | Ophthalmic | Ophthalmic | 510(K) Exempt | Lenstar LS 900, Biomicroscope, Slit-Lamp, AC-powered. HAAG-STREIT AG, Switzerland. |  |  | May 19, 2010 | 2010 | Haag-Streit USA Inc | During the creation of the Lenstar LS 900 software, the ability to perform a manual adjustment of the retinal thickness measurement was not disabled for products targeted for US distribution. The current software installed allows the user to manually alter the retinal thickness measurement, which could modify the automatic axial eye length measurement resulting in the selection of an incorrect in | PMA | HAAG - Streit USA, Inc. issued notification letters to their customers informing them of the problem. They instructed their customer to refrain from adjusting the gates for the measurement of retinal thickness until after the system has been updated.For further information, contact HAAG - Streit USA, Inc. at 1-513-336-7255. |  | 45 | 45 units | Nationwide Distribution -- Including AL, AZ, CA, FL, ID, KY, LA, MA, MI, MN, MO, NC, NJ, NY, OH, OR, PA, SC, VA, VT, WA and WV. | Terminated | February 24, 2012 | 646 | 45 | Software | Output/Calculation | Software update | Software Update
Z-1892-2010 | 55651 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Imaging Version Lines VB20, VB30, VB35 in combination with a RIS. |  |  | April 06, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a potential malfunction and hazard to patients when using Syngo Imaging version lines VB20, VB30, VB35A in combination with a RIS that is violating DICOM Standard by creating non-unique study instance UIDs. | Device Design | The recalling firm sent a Customer Safety Advisory Notice, dated 04-06-2010, to customers explaining when the malfunction occurs and the potential risks, steps to avoid potential risk of the issue by ensuring that the RIS works according to DICOM Standard and does not create double study instance UIDs. Siemens Service Reps will contact the consignees in a timely manner to check if there are symptoms for the reported issues. The issue will be resolved with a Software-Update for each affected version, except for version VB20. All customers on VB20 will be upgraded. |  | 3 | 3 units | Nationwide Distribution -- Including CT, NM, &amp; WI. | Terminated | July 05, 2012 | 821 | 3 | Software | Output/Calculation | Software update | Software Update
Z-1817-2010 | 55784 | No | JCT | Tracheal prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ultraflex Tracheobronchial Stent System. Uncovered - 8/20 Catalog Number: M00568920. Single Use, Rx only. Boston Scientific, Natick, MA 01760. |  |  | May 17, 2010 | 2010 | Boston Scientific Corporation | Suture related device failures during deployment and repositioning of the stent | Device Design | Boston Scientific issued an "Urgent Medical Device Recall - Immediate Action Required" notification dated May 17, 2010 to U.S. Distributors stating that use and distribution of products should cease immediately and return all affected product to the firm for replacement. Distributors were asked to complete and return reply a verification tracking form.For further information, contact Boston Scientific at 1-866-868-4004. |  |  |  | Worldwide Distribution - Only US distributed product is affected by recall | Terminated | October 20, 2011 | 521 |  | Not_Computer
Z-1901-2010 | 51498 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | VARIANT II Hemoglobin Testing System with CDM 3.5 Software Version 3.5; Model Number: 270-2000; Product is manufactured and distributed by Bio-Rad Laboratories, Diagnostics Group, Hercules, CA. |  |  | January 22, 2009 | 2009 | Bio-Rad Laboratories Inc | Sample test results for the first run appears to be correct but all subsequent results during that run were assigned incorrect barcode and tube position. | Software design | Bio-Rad Laboratories, Inc. issued and "Urgent: Medical Device Correction" dated January 22, 2009 to all affected consignees by FED EX, followed up by telephone calls for those not responding. Consignees were instructed on proper steps to take in dealing with the affected product. Further communication will follow to provide a solution.For further information, contact Bio-Rad Laboratories at 1-510-724-7000. |  |  |  | Worldwide Distribution -- United States, Argentina, Australia, Brazil, Canada, Chile, China, Columbia, Costa Rica, Czech Republic, France, Germany, Hong Kong, Hungary, India, Italy, Mexico, New Zealand, Poland, Singapore, South Korea, Sweden, Taiwan, Thailand and the United Kingdom. | Terminated | March 22, 2011 | 789 |  | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-1913-2010 | 55765 | No | OCK |  |  | Gastroenterology/Urology | PMA | FemSoft Insert, Female Urethral Insert, Sterile (R), Rx Only. Assorted Sizes Multi-Pack, REF 70200. |  |  | April 27, 2010 | 2010 | Rochester Medical Corp | Rochester Medical is voluntarily recalling specific lots of FemSoft urethral inserts manufactured and distributed in 2009 and 2010. Testing has revealed a small number of packages which contain a small void in the package seal, and it has been determined that product sterility cannot be guaranteed. If product sterility has been compromised, there is a potential for increased risk of urinary trac | Packaging | Consignees were contacted by telephone on 5/3/10 and if not reached, were sent a Rochester Medical "Urgent Medical Device Recall" letter on 5/6/10 dated April 26, 2010. The letter was addressed to Rochester Medical Corporation Customers. The letter described the problem and product involved; recommended consignees to record the quantity and lot numbers (s) on the enclosed postcard and return it to them; and to immediately destroy the product with the lots affected. In the case of distributors, the telephone call was followed by a letter sent/dated April 27, 2010 to each distributor. It advised distributors to isolate any remaining inventory with the specified lot numbers and to complete the disposition report. Letters included a customer letter for dissemination to their retail level customers. Retail customers were asked to segregate and destroy the affected products and record the action. The company will replace the destroyed product. The distributors started their contact to consignees on 4/28/10. Please contact Rochester Medical at 1-800-243-3315, or the firm's Customer Service Department at 1-507-533-9309 for additional information or for product return &amp; replacement instructions. |  | 15400 | 15,400 devices | AL, AK, AZ, CA, CO, DC, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, WY, WI,.AUSTRIA, BANGLADESH, CANADA, CYPRUS, ETHIOPIA, FRANCE, GERMANY, GREECE, INDIA, JAPAN, KOREA, NETHERLANDS, P.R. CHINA, SAUDI ARABIA, SCOTLAND, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TURKEY, UAE, UNITED KINGDOM. | Terminated | December 28, 2011 | 610 | 105980 | Not_Computer
Z-1905-2010 | 55495 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Unicel DxH 800 Coulter Cellular Analysis System, |  |  | July 28, 2009 | 2009 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed an issue with the UniCel DxH 800 during in-house testing. It wasdetermined that there is a potential for a false increase in platelet results, due to micro-bubbles. This situation could present a risk of falsely increased platelet results in patient samples as diluent containers are depleted. | Other | Beginning on July 29, 2009 Customers were contacted by Customer Technical Support (CTS) personnel and provided with the Product Corrective Action (PCA) letter. The PCA letter informed customers of the issue, the impact of the issue and actions/resolutions required by Beckman Coulter Service representative ad the affected consignee.The consignees were informed that their Beckman Coulter Service representative has made temporary changes to the Supplies setup. In addition to the Supplies setup change, consignees were instructed that one of the following workflow options must be implemented to ensure no results are affected. Please continue to follow one of the options listed below until a permanent fix is available. These options were given as: Option A: Monitor Diluent cycles or Option B: Unable to monitor Diluent cycles. The consignees were instructed that a long term solution to eliminate this risk is being developed and will be provided as soon as available. The consignees were instructed to please share this information with your laboratory staff. If they have any questions regarding this Product Corrective Action, please call 800-526-7694 in the United States or contact your local Beckman Coulter Representative. |  | 7 | 7 | Nationwide to: AZ, LA, IL, NY, PA &amp; TN | Terminated | April 04, 2012 | 981 | 7 | Not_Computer
Z-1904-2010 | 55953 | No | IZO | Diagnostic x-ray high voltage generator. | Radiology | Radiology | 510(K) Exempt | Hydradjust IV DR Urological Table Sedecal Generator (p/n 750732 - 64W generator &amp; p/n 750733 80W generator). Mallinckrodt, Inc. Cincinnati, OH 45237. |  |  | June 14, 2010 | 2010 | Mallinckrodt Inc | Covidien received complaints of the Sedecal Generator failing and producing smoke at customer sites. The failure modes are located in different parts of the generator and will lead to the generator becoming inoperative. | Process control | Covidien Issued an "Urgent Device Correction" notification dated June 14, 2010 informing their customers of the potential problem. Consignees were instructed that Field Service Engineers will be in contact arrange a service visit to perform corrections on the affected units.For further information, contact Covidien Product Monitoring at 1-800-778-7898. |  | 144 | 144 units | Worldwide Distribution -- United States, Canada and China. | Terminated | February 25, 2012 | 621 | 144 | Not_Computer
Z-1914-2010 | 53698 | No | CIB | Angiotensin I and renin test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GammaCoat Plasma Renin Activity 125 I RIA Kit, REF: CA-1533 (100 test kit ), CA-1553 ( 500 test kit), DiaSorin, Stillwater, Minnesota 55082-0285. |  |  | October 12, 2009 | 2009 | Diasorin Inc. | The kit control provided with the GammaCoat Plasma Renin Activity kit was labeled with the incorrect range. | Nonconforming Material/Component | Consignees were Emailed or faxed a "DiaSorin Customer Notification" letter dated October 21, 2009. The letter was addressed to "Dear Valued Customer". The letter described the Product name and the problem and provided the correct range for the Renin Activity Control provided with the kit. The letter also stated that replacement product for any previous failed runs. Customers are to complete the bottom portion of the letter and return it to the firm. Questions or concerns should be directed to DiaSorin Inc., Technical Services at 1-800-328-1482 or 651-439-9710. |  | 289 | 289 kits | Worldwide Distribution -- USA, including states of CA, CT, IN, LA, MI, NC, OH, TN, UT, and WA and countries of CANADA, BELGIUM, BRAZIL, INDIA, NORWAY, PORTUGAL, SPAIN, and SWEDEN. | Terminated | April 17, 2012 | 918 | 289 | Not_Computer
Z-1916-2010 | 55579 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Automated External Defibrillator(AED) of brand name "Philips HeartStart HS1 OnSite" (Model number M5066A). |  |  | April 30, 2010 | 2010 | Philips Medical Systems | Seventeen AEDs failed production line testing (Final Acceptance Test). | Nonconforming Material/Component | The firm notified U.S. distributors on 4/30/2010 via phone call and e-mail to obtain consignee location information. A website providing information was activated on 4/30/2010. On 5/3/2010 Urgent Medical Device Recall letters, dated April 29, 2010 were sent to direct customers in the US. The letters explained the reason for the recall and Customers were asked to verify that they have the product subject to recall. Replacement AEDs will be provided at no charge to the customer. Field Safety Notices were attached to the letters and provided a list of actions for customers to take. Customers should check the serial number on their devices against the list provided, contact Philips Healthcare to arrange for replacement, and ensure that necessary staff are aware of the recall. For further information and support, Philips Healthcare should be contacted at 1-800-263-3342. |  | 205 | 205 | Worldwide Distribution -- USA, Canada, Japan, United Arab Emirates, Austria, Switzerland, Germany, Spain, France, Italy, and Norway. | Terminated | December 10, 2012 | 955 | 471 | Not_Computer
Z-1922-2010 | 55552 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 3100/ 3100IQ system with InnovaIQ Table Option |  |  | June 03, 2010 | 2010 | Ge Healthcare, Llc | GE Healthcare has become aware of a potential loss of gantry and table motions affecting Innova 2100IQ, 3100, 3100IQ, 4100 and 4100IQ systems with InnovaIQ Table that may impact patient safety. | Device Design | Consignees were sent on 6/3/10 a GE Healthcare "Urgent Medical device Correction" letter dated May 12, 2010. The letter was addressed to Hospital Administrators/ Risk Mangers, Managers of Radiology/Cardiology and Radiologists/ Cardiologists. The letter included the Safety Issue, Affected Product details, Safety Instructions, Product Correction and Contact Information. Customers are to contact a service representative immediately if they notice any cable damage. |  | 13 | 13 | Worldwide Distribution -- USA, including states of AL, AZ, CA, CO, FL, GA, IL, IN, KY, MA, MI, MS, MO, MT, NE, NJ, NY, NC, OH, OR, PA, TN, TX, VT, WA, WV, and WI and countries of UNITED KINGDOM, SWEDEN, REPUBLIC OF KOREA, NORWAY, NEW ZEALAND, ITALY, ISRAEL, IRELAND, INDIA, HONG KONG, GERMANY, FRANCE, FINLAND, CZECH REPUBLIC, and AUSTRALIA. | Terminated | September 28, 2015 | 1943 | 92 | Not_Computer
Z-1919-2010 | 55818 | No | HPS | Cryophthalmic unit. | Ophthalmic | Ophthalmic | 510(k) | MIRA CR4055 25mm Finger Probe |  |  | April 08, 2010 | 2010 | Mira, Inc. | Devices distributed without an approved 510(k) | PMA | Mira contacted accounts by email between 4/08/10-05/21/10 requesting return of inventory. |  | 5 | 5 units | Worldwide Distribution -- USA, including the state of NY and the countries of Israel and Italy. | Open, Classified |  |  | 5 | Not_Computer
Z-1941-2010 | 55850 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | STERIS 4085, 5085 &amp; 5085SRT Surgical Tables |  |  | May 14, 2010 | 2010 | Steris Corporation | STERIS has learned that a limited quantity of manifold assemblies were manufactured by our supplier with incorrect washers and springs in the valves of the manifold. This supplier error has the potential to cause: a) an inability to articulate/move a table section when commanded by the operator; b) two different articulations occurring when a single articulation is commanded, or c) that the table | Nonconforming Material/Component | On May 14, 2010 the firm sent FCL letters to their customers notifying them of the problem. They informed their customers that a service representative will contact them to arrange correction of the unit. |  | 79 | 79 units | Worldwide distribution: USA, Australia, India, Korea &amp; New Zealand | Terminated | February 24, 2012 | 651 | 79 | Not_Computer
Z-1940-2010 | 55654 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Biosciences, BD FACS Sample Prep Assistant II (SPA II), blood sample preparer for flow cytometry analysis |  |  | February 19, 2010 | 2010 | BD Biosciences, Systems &amp; Reagents | Samples prepared on the BD SPA II, SN T0076 produced inaccurate sample results due to inaccurate sample dispensed as a result of the cap piercing probe | Nonconforming Material/Component | The firm filed an MDR and implemented a field action consisting of a February 2010 notification letter and product replacement of affected probes. A CAPA for the probe design has been implemented. Effectivity is being tracked via mail and follow up phone call. |  | 349 | 349 units | Worldwide distribution: USA, Australia, Belgium, China (Hong Kong), Colombia, Jamaica, Korea, Mexico, Singapore and Uruguay. | Terminated | May 18, 2011 | 453 | 349 | Not_Computer
Z-1943-2010 | 54366 | No | JEY | Bone plate. | Dental | Dental | 510(k) | BIOMET MICROFIXATION REF: 91-6103 "1.5MM" SYSTEM HIGH TORQUE (HT), SD, X-DR, SCREW 1.5MM X 3.5MM (WXL) TITANIUM LOT 829760, ***NON STERILE PRODUCT***. |  |  | October 05, 2009 | 2009 | Biomet Microfixation, Inc. | Biomet Microfixation is recalling 91-6103 1.5x3.5mm HT SD X-drive screw Lot Number 829760. This package may contain 91-6704 1.5x4mm ThinFlap SD X-drive screw. | Process design | Biomet Microfixation notified each consignee by email and fax on 10/05/2009. Each consignee was asked to use the included Response Form to provide product disposition. |  | 91 | 91 units | US distribution: CA, CO, MI, and MO. | Terminated | December 15, 2010 | 436 | 91 | Not_Computer
Z-1942-2010 | 55801 | No | DAK | Complement components immunological test system. | Immunology | Immunology | 510(k) | Inova Diagnostics BINDAZYME Human C1q binding Circulating Immune Complex EIA Kit, Part#: MK021 |  |  | October 28, 2009 | 2009 | Inova Diagnostics Incorporated | The recall was initiated after the binding site notified Inova Diagnostics, inc. that the affected Bindazyme Human C1q binding Circulating Immune Complex EIA Kits may contain an ELISA plate labeled as "tissue Transglutaminase Coated Wells" instead of "C1q Coated Wells". | Process control | Notification letters from INOVA Diagnostics Incorporated were mailed to each of the consignees via Certified mail on October 30, 2009.The notification letters inform the consignees of the reason for recall and provided instructions for final disposition of the recalled kits. In addition, a FAX back Attachment was attached to the recall notice for customers to respond to the recall. |  | 700 | 700 kits | Worldwide distribution: USA, Germany, Italy, Belgium, Romania, Portugal, Argentina, Spain, Sweden, and India. | Terminated | June 30, 2010 | 245 | 700 | Not_Computer
Z-1937-2010 | 55885 | No | FKG | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | C.A.P.D. Tubing Exchange Tray, HT80, Sterile, Micropore Surgical Tape, Single Use Rolls, Tri-State Hospital Supply Company Manufacturers of Centurion Healthcare Products Howell, MI. |  |  | March 16, 2010 | 2010 | Centurion Medical Products | The 3M micropore tape is under recall due to the tape not performing and needles becoming loose and dislodged. | Other | The firm sent Urgent Recall Product Notices to its two customers dated 3/16/2010. The letter explained that the tape may not adhere properly. The consignees were instructed to remove any kits containing the recalled tape from their inventory, if the tape was used to secure dialysis needles or pressure dressings. If the tape was used for other purposes, the kits could still be used. The recall letter included a reply form for customers to complete and email or fax back. Upon receipt of the form, a Centurion Representative will contact them to retrieve affected product. Customers should contact Matthew Price at 517-546-5400 ext. 1135, if further info is needed. |  | 24 | 24 kits | Nationwide Distribution -- AR and OH. | Terminated | February 09, 2011 | 330 | 3312 | Not_Computer
Z-1935-2010 | 55796 | No | IKN | Diagnostic electromyograph. | Physical Medicine | Neurology | 510(k) | Lumbar Matrix Scan, Diagnostic Electromyography System consisting of the CERSR Electromyography System and Electrode Array; the system consists of two devices. The first is a system of capital equipment referred to as CERSR Electromyography System, which includes a cart. The second device is disposable electrodes arrays, which are single use devices, used with the capital equipment that is adhered to the patient for monitoring the bioelectric signals produced by the muscles. The disposable is referred to as the Electrode Array. CERSR Electromyography System is packaged in a reusable crate maintained by SMI. The Electrode Array is packaged in a sealed foil pouch. Both devices are non-sterile and the packaging is for protection from damage during shipping. |  |  | January 29, 2010 | 2010 | SpineMatrix Inc | A full design review of the CERSR system was conducted in 7/2009. The audit revealed numerous design verification tests that either did not pass or did not meet the design input requirement. One requirement found that it was possible for the physician to think they were reviewing one patients scan, but were actually reviewing a scan from a different patient. This event did not happen; however, | Device Design | On 1/29/2010 all customers were provided a Voluntary Product Recall Notification letter that was sent via U.S. Postal Service Certified Mail and they were contacted by phone to arrange for pick up of the devices. The letter identified the affected product and described the problem. Customers are to check their inventory for the affected product and contact SpineMatrix Customer Service to receive instruction as to how to return the affected product, if the firm did not already contact the customer. Customers must complete the verification form and return it. Questions should be directed to Customer Service. |  | 1166 | CERSR - 14 units, Electrode Arrays - 1,152 units | Nationwide Distribution -- CA, OH, OR &amp; TX. | Terminated | September 19, 2012 | 964 | 1166 | Not_Computer
Z-1939-2010 | 55749 | Yes | DWA | Cardiopulmonary bypass pump speed control. | Cardiovascular | Cardiovascular | 510(k) | Hospital Grade Power Cord, UL, Part # 66319, 055476, manufactured by Electri-Cord Manufacturing Co. beginning in February 2003. The cords are a detachable unit, supplied with the instruments listed below: |  |  | April 21, 2010 | 2010 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | The potential risks during the use of this power cord include electrical shock, delay in setup and therapy, interruption of therapy, device failure, and fires. Although Medtronic has received no reports of adverse events related to the affected cord, a failure may potentially lead to serious adverse health consequences, depending on the device, and therapy being interrupted. This Field Notificatio | Nonconforming Material/Component | Consignees were sent on 4/21/10 a Medtronic "Urgent Medical Device Recall Notice" letter dated April 19, 2010. The letter was addressed to their customers and described the problem and the product involved. They recommended to discontinue use of the AC power cords and to destroy them. Requested consignees to fill out the attached Medtronic Recall Certificate (FCA 1005), and return it to Medtronics. |  | 699 | 619 USA, 80 OUS | Worldwide distribution: USA, Latin America, Middle East, Japan, and Canada. | Terminated | December 24, 2011 | 612 | 699 | Battery | Treatment/Delivery/Therapy | Remove | Remove or Replace
Z-1903-2010 | 55853 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK 20 defibrillator/monitor. |  |  | May 27, 2010 | 2010 | Physio Control, Inc. | A failure on the power supply assembly can result in either "No DC power" or "No DC or AC power." A failure of DC (battery) power can result in a delay of defibrillation therapy if no AC (line) power is available and the device will not operate. | Material/Component Contamination | Firm began mailing Urgent Medical Device Correction notifications to affected customers on May 26, 2010 regarding the power supply to the device. All affected power supplies will be updated; this will be accomplished in two phases. Devices at a higher risk will be updated first in phase 1. The remaining lower risk devices will be updated afterward in Phase 2. Customers are advised to keep the defibrillators in service, and to follow recommended daily Operator Checklist steps, while service updates are scheduled. Customers with questions should contact the firm at 800-442-1142 between 6 am and 4 pm Pacific Time, Monday - Friday |  | 43001 | 43,001 units | Worldwide distribution. In US, distributed domestically including to gov't accounts; and internationally to the following countries: ANGUILLA, ARGENTINA, AUSTRALIA, BAHAMAS, BAHRAIN, BARBADOS, BOLIVIA, BRAZIL, CAYMAN ISLANDS, CHILE, COLOMBIA, COSTA RICA, CYPRUS, DOMINICAN REPUBLIC, EGYPT, EL SALVADOR, GERMANY, GUAM, GUATEMALA, HONG KONG, INDIA, IRAN, ISRAEL, JAMAICA, JORDAN, KUWAIT, LEBANON, MEXICO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PUERTO RICO, SAUDI ARABIA, SOUTH AFRICA, SYRIA, TRINIDAD, TRINIDAD AND TOBAGO, TURKEY, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIRGIN ISLANDS, and YEMEN. | Terminated | October 21, 2014 | 1608 | 43001 | Battery | Device Operation | Upgrade Battery | Repair
Z-1887-2010 | 55786 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Navilyst Medical, Inc., Vaxcel(TM) Port Plastic Standard Port with 9F Polyurethane Catheter (1.8 mm/2.8 mm), REF Catalog No. 45-360, UPN Product No. M001453600, STERILE, Rx ONLY, For single use only --- Implantable Vascular Access System --- (kit contains Huber needle) -- previously distributed by Boston Scientific Corporation. |  |  | May 13, 2010 | 2010 | Navilyst Medical, Inc | FDA testing determined that non-coring needles included in the Vaxcel Port and Vaxcel Port with PASV Valve may result in coring and resulting septum damage which could lead to patient injury. | Nonconforming Material/Component | Navilyst Medical, Inc. is recalling Vaxcel Ports and Vaxcel Ports with PASV Valve with valid expiration dates worldwide. Navilyst sent "Urgent Medical Device Recall - Immediate Action Required" letter/recall notification packages on May 13, 2010, via Federal Express to 471 domestic consignees; on May 14, 2010, via e-mail to 14 international consignees and to Boston Scientific Corporation (prior owner of the affected products) via facsimile and Federal Express. On June 3, 2010, Navilyst Medical sent Recall Notification Packages on behalf of Boston Scientific Corp. via Federal Express to 182 domestic consignees. The letter instructs the customer to immediately discontinue use of and segregate recalled product, complete and return the reply verification tracking form to fax# 518-742-4450 and package and return the recalled product to the address provided.If you have any question, please call Navilyst Medical, Inc., Customer Service at 1-800-833-9973 or Michael Duerr at 1-518-742-4571. |  |  |  | Worldwide distribution:USA and countries of Belgium, Brazil, Canada, Colombia, Croatia, Germany, Greece, Hong Kong, Italy, Saudi Arabia, South Korea, Spain, Tunisia, and United Kingdom. | Terminated | August 03, 2012 | 813 |  | Not_Computer
Z-1898-2010 | 55670 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Cepheid Xpert MRSA/SA Blood Culture Test for use with the GeneXpert Dx System, manufactured by Cepheid, Sunnyvale, CA. |  |  | April 05, 2010 | 2010 | Cepheid | The firm is receiving reports of inaccurate results, which may result in delay of care for seriously ill patients. | Nonconforming Material/Component | The firm, Cepheid, issued a technical bulletin to customers on April 5, 2010 and an updated to the April 5 bulletin dated April 9, 2010 instructing them to perform a reflexive repeat test on any blood culture results initially determined as having SA.If further information on this please contact your Cepheid Sales Representative 1-888-336-2743 or www.CEPHEID.COM. |  | 5510 | 5,510 units | Worldwide distribution: USA and Europe, Taiwan, and Australia. | Terminated | May 25, 2011 | 415 | 5510 | Not_Computer
Z-1947-2010 | 55186 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Tabs Professional Monitor with Wire Bracket and Push Button Reset, Model 25223, Stanley Senior Technologies, Lincoln, NE. |  |  | March 26, 2010 | 2010 | Stanley Security Solutions, Inc. | Product might not indicate low battery alarms or exit alarms when powered by a DC power supply. | Component change control | The firm made phone calls beginning 3/26/10 to customers who purchased the 12-volt power supply informing them of the recall and instructing them to determine the number of units that will need to be returned for replacement. Important Safety Notice letters dated 3/25/10 were issued via regular mail on 3/26/10 and also providing the reason for the recall, instructions, and customer service contact information to arrange for return of the units. |  | 45 | 45 units | Nationwide Distribution | Terminated | September 16, 2011 | 539 | 374 | Not_Computer
Z-1952-2010 | 55776 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Artificial ligament fixation device # 7PE ZIPLOOP Extended Toggle, Sterile, REF 904755, Biomet Sports Medicine Warsaw, IN. |  |  | March 31, 2010 | 2010 | Biomet, Inc. | The firm has received complaints that the saddle is too long. The saddle is supposed to be between 7/8 and 1 1/8 inches long. The saddle in the complaints was up to 1.9 inches long. | Process control | Biomet Sports Medicine issued an "Urgent Medical Device Recall Notice" dated March 31, 2010 to user facilities and distributors. A second notification containing updated information was issued June 2, 2010. Consignees were instructed to discontinue use of the product and return it to the firm along with a completed Response Form. For further information, contact Biomet at 1-574-371-3755 or 1-574-371-3756, Monday through Friday, 8AM to 5PM. |  | 2466 | 2466 both products | Worldwide Distribution -- United States, Canada, Australia, The Netherlands, Belgium, Italy, France, New Zealand, Austria, United Kingdom, Chile, Greece, Poland, Denmark, Turkey, Portugal, South Africa and Spain. | Terminated | August 06, 2012 | 859 | 2466 | Not_Computer
Z-1951-2010 | 55771 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | OptifluxÂż F180NR Dialyzers, Catalog Number: 0500318E. Manufactured by Fresenius Medical Care North America, Ogden, UT 84404. |  |  | May 14, 2010 | 2010 | Fresenius Medical Care North America | Hemodialyzer may leak at the header/end cap | Process control | Fresenius Medical notified customers by telephone script beginning May 14, 2010 and follow up letter advising users to examine inventory for recalled lot, discontinue use and contact Fresenius Customer Care 1-800-325-5188 , ext. 6610 for instructions on how to return the recalled product. |  | 34152 | 34,152 units | Nationwide Distribution | Terminated | September 13, 2011 | 487 | 34152 | Not_Computer
Z-1954-2010 | 55894 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Stellant CT Dual Syringe Kits, Catalog Number SDS-CTP-QFT and SDS-CTP-SPK |  |  | May 24, 2010 | 2010 | Medrad Inc | Kits contain a T connector that may be susceptible to cracking or breaking or leaking. This could cause product leakage or, less likely, patient injury. | Other | The recalling firm issued an Urgent Medical Device Field Correction on 5/24/10 requesting customers to check their inventory for the affected lots and complete the fax back form with the amount of product in stock. The manufacturer plans to make arrangements to provide individually packaged low-pressure connector tubing to use with the product in the event that cracking or breaking and leaking of the T connector is observed. This requires vigilance by the customers over the life of all of the devices.For questions concerning this recall, please contact the company 1-800-633-7231, select option 2, and indicate that you are calling about the Stellant CT Dual Syringe Kit field correction. |  | 1382889 | 1,382,889 kits | nationwide and Worldwide distribution includes AU, BE, BG, BM, BR, CA, CH, CI, CL, CN, CO, CS, DE, DK, EC, EG, FI, FR, GB, GT, HN, IN, IT, JM, JP, KW, KY, MU, MX, MY, NL, OM, PA, PH, PK, PR, SE, SG, TW, YE, ZA. | Terminated | October 13, 2010 | 142 | 1382889 | Not_Computer
Z-1889-2010 | 55824 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | 22ga x 1" Right Angle Huber Needle |  |  | May 26, 2010 | 2010 | Multi-Med, Inc. | FDA sample determined the needles to be coring. | Pending | The firm, Multi Med, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated May 26, 2010 to their sole consignee Navilyst Medical, Inc., via email and overnight courier with signature notification. The letter describes the product, problem and action to be taken by customer. The customers was instructed to immediately examine their inventory and quarantine product subject to recall, if the product was distributed to others customers, identify and notify them at once and include recall notification letter,. Recall has been initiated due to FDAs sample analysis that these needles may core implanted port septums.If you have any questions, call me, Sue Starkey, at 603-357-8733 or via email at astarkey@multimedinc.com. |  | 5000 | 5,000 | Worldwide distribution | Terminated | May 23, 2017 | 2554 | 10169 | Not_Computer
Z-1949-2010 | 55825 | No | NBE |  |  | General &amp; Plastic Surgery | PMA | Coseal Surgical Sealant 4mL, Product Codes: 934071 &amp; 934074. |  |  | January 14, 2010 | 2010 | Baxter Bioscience | The recall is being conducted as a precautionary measure due to out of specification results at the 18 month and 21 month time points (out of the 24 month licensed shelf life to expiration) for the percent thiol substitution during a Coseal Stability study. The parameter that is out of specification is an indicator of possible failure for the product to gel appropriately. | Process control | The recall notification was initiated on January 14, 2010 with the firm sending a Product Recall letter and attached Customer Reply Form to the affected consignees. The recall letter informed the consignees of the reason for recall and requested them to immediately stop the usage/distribution of the lots noted above and to quarantine any units you may still have in your inventory. The letter provided instruction for product return.The letter also requested the consignees to please forward the recall information as appropriate to other services/facilities/departments that they may have provided this product to. Consignees were instructed to complete theCustomer Reply Form, and fax it as soon as possible to 888-345-4261. Consignees with questions were instructed to contact Stericycle at 877-870-4486. |  | 15262 | 15,262 | Nationwide Distribution | Terminated | December 22, 2010 | 342 | 26440 | Not_Computer
Z-1895-2010 | 55854 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira brand Symbiq Two-Channel Infuser; List Number: 16027; Distributed by: Hospira, Inc. Morgan Hill, CA and Manufactured by: Hospira Inc, La Aurora, Heredia, Costa Rica. |  |  | April 09, 2010 | 2010 | Hospira Inc | Hospira has received numerous customer complaints of the Symbiq infusion pump's failure to detect air-in-line conditions during operation. | Device Design | Hospira issued a "Urgent: Medical Device Recall Notification" dated June 11, 2010 to update a previously issued Clinical Bulletin. Customers were informed of the issue and instructed on proper mitigation steps to further avoid the issue associated with the affected product. Consignees were asked to complete and return the Reply Form.For further information, contact Hospira Technical Support Operations at 1-800-332-4002, 6AM to 4PM PST, Monday through Friday. |  | 28485 | 28,485 total | Worldwide Distribution -- United States (AL, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NE, NH, NJ, NY, OH, OK, PA, SC, SD, TN, TX, VA and WA), and Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Puerto Rico, Singapore and United Arab Emirates. | Terminated | December 03, 2013 | 1334 | 28485 | Not_Computer
Z-1956-2010 | 54371 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | BIOMET 3i GS300 GOLD SLOTTED SCREW-3mm, STERILE, Rx Only. Made in USA. Lots 868517, 865348, and 868515. |  |  | October 06, 2009 | 2009 | Biomet 3i, LLC | Biomet 3i, Palm Beach Gardens, FL is recalling their Gold Slotted Retaining Screw 3MM, Model Number GS300. Product is Out Of Specification - product was manufactured with a material not specified for use in this product. | Other | Each US customer was contacted by telephone using a prepared phone script. Following telephone contact, Biomet 3i performed a follow-up of each call with a faxed letter. International customers were first notified by e-mail then by fax letter. Each customer was instructed to check their respective inventory for the affected product and corresponding lot number. They were then requested to fax back a response form to Biomet 3i Regulatory Services to provide feedback concerning their inventory. They were also instructed to return the affected product to Biomet 3i.Questions or concerns are directed to Kelly Taylor at the firm at 561-776-6700. |  | 495 | 495 screws | Product was distributed to 3 distributors, 14 direct consumer/users. Product was distributed in Japan, Spain, Switzerland, Uruguay, Paraguay, Mexico, France, Netherlands, Australia, and in the US in KY, DC, IL, RI, MA, NJ and TX. | Terminated | August 30, 2010 | 328 | 495 | Not_Computer
Z-1955-2010 | 55774 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | SSH-880CV Aplio Artida System |  |  | September 16, 2009 | 2009 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems (TAMS) is initiating a field correction on the SSH-880CV Aplio Artida System due to a problem related to a wrong measurement value display in using with DICOM viewer.When an image is transferred from this system in DICOM transfer and a measurement is performed on a Doppler image in the DICOM viewer, measured values such as VTI (Velocity Time Integral) and PI/RI | Software design | The software will be replaced. To accomplish this, Toshiba America Medical Systems (TAMS) will notify consignees via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS for retention.For questions or concerns, please contact your Toshiba representative or the firm at (714) 730-5000. |  | 22 | 22 | Nationwide ((CA, GA, IA, LA, MA, MD, MN, MT, PA, PR, OH, OR, TN). | Terminated | March 21, 2012 | 917 | 22 | Software | Display/Image | Software update | Software Update
Z-1971-2010 | 55316 | No | LGC | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | Sure-Vue HSV-2 Rapid Test, Catalogue No: 23-046395. |  |  | April 14, 2010 | 2010 | Biokit U.S.A. Inc. | Negative serum was giving a positive result Herpes Simplex Type 2 (HSV-2). It is possible that misdiagnosis (i.e., positive) could be made. | Nonconforming Material/Component | Biokit USA notified US Customers by fax on 2/18/10 of the reason for the kit recall and requested to discard any remaining units. Accounts were advised to retest patients or review test results. Questions are directed to the firm at 800-926-3353. |  | 8 | 8 kits | US (CA, FL, GA, IN, KY, MA, NY, VA) and UK | Terminated | January 26, 2012 | 652 | 200 | Not_Computer
Z-1960-2010 | 55710 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker 10 MM AVS TL Trial Spacer; Catalog number: 48389410. |  |  | May 13, 2010 | 2010 | Stryker Spine | Stryker Spine has received reports of breakage of the distal end of the AVS TL Spacer Trial during the trialling step of the AVS TL procedure. | Process control | Urgent Product Recall letters, dated May 13, 2010 were sent via Federal Express to all direct accounts (branches and hospitals) with a separate letter to surgeons. The letter identified the affected product and also described the issue, potential hazards, and risk mitigations. The letter also asked customers to quarantine any affected product and return to Stryker Spine. A Customer Response Form was provided and customers should contact their Stryker Spine Sales Rep for additional information on replacement product. Questions should be directed to Tiffani Rogers on 201-760-8206. |  | 193 | 193 units | Nationwide Distribution. | Terminated | September 06, 2012 | 847 | 2110 | Not_Computer
Z-1969-2010 | 55721 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Starion Instruments TLS3 35C Thermal Ligating Shears, Electrosurgical cutting and coagulation device and accessories, manufactured by Starion Instruments Corp, Sunnyvale, CA. |  |  | February 06, 2008 | 2008 | Starion Instruments | Some jaws exhibited lack of ductility and could break if overloaded or used improperly by placing too much tissue in the bottom of the "v" of the jaw. | Nonconforming Material/Component | The firm called domestic consignees but no letter was sent to domestic cosignees and a notification letter was sent to international consignees.The firm called domestic consignees and sent a notification letter to international consignees. No records were kept of phone calls and a copy of the letter sent to international consignees was not obtained |  |  |  | Worldwide distribution | Terminated | November 27, 2010 | 1025 |  | Not_Computer
Z-1972-2010 | 55750 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Starion Instruments Real Hand ThermaSeal Thermal Ligating Shears, electrosurgical cutting and coagulation device and accessories, manufactured by Starion Instruments, Sunnyvale, CA |  |  | June 26, 2008 | 2008 | Starion Instruments | Due to an improper fixture revision, device may be missing a screw at the jaw assembly, resulting in jaw components falling off during use. | Process control | The firm submitted an MDR but determined that there was no risk to health. E-mails were sent on June 26, 2008, to consignees, describing the product, the problem, and action to be taken by consignees. The consignees were ask to provide a list of all unused ThermaSeal devices and their locations, to return the products to Starion from the lots noted and to provide a copy of the attached letter to their sales force and customers if they choose.If you have any questions, please call (408) 522-5200 ext. 326 or (408) 350-9934. |  | 217 | 217 units | Worldwide distribution. | Terminated | October 27, 2010 | 853 | 217 | Not_Computer
Z-1973-2010 | 55419 | No | HXS | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | TORQ Sternal Closure Device, SINGLE-USE |  |  | March 31, 2010 | 2010 | KARDIUM INC. | A component of the device is installed incorrectly which can result in the component falling out of the device and into the patient during use. | Other | The firm, Kardium, informed it's US distributor on March 30, 2010 by e-mail. The distributor was instructed to return all unused product to Kardium. The units remaining at hospitals will be returned to Kardium by the distributor sale reps. If you have any questions, please contact (604) 248-8891. |  | 420 | 420 units | Worldwide distribution: USA and Canada. | Terminated | October 07, 2010 | 190 | 420 | Not_Computer
Z-1974-2010 | 55598 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Horizon Cardiology Hemo Component Schiller Argus Pro PB-1000 Monitor unit. |  |  | April 15, 2010 | 2010 | McKesson Provider Technologies - Medical Imaging Group | McKesson has recently discovered that certain Horizon Cardiology Hemo systems were running specific Schiller Argus-Pro PB 1000 firmware versions that might cause a delay in the display of physiological parameters on the Horizon Cariology Hemo Real Time Monitor Screen. | Software change control | The firm, McKesson, sent an "Advisory Notice" letter dated April 15, 2010, to all customers. The letter describes the problem, products affected, and actions to be taken by customers. The customers were instructed not to use the Emergency Replacement Hemo kit until they receive a replacement Schiller Argus Pro PB-1000 and they have contacted McKesson Support department at 1-866-777-0202.If you have any questions regarding this notice, please call our Customer Support department at 1-0866-777-0202 and ask to speak to your support Manager. |  | 494 | 26 units; 19 in operational use, 8 as emergency replacement units. | Nationwide distribution: CA, CO, MD, NH, NT, NY, TN, and TX . | Terminated | July 21, 2010 | 97 | 494 | Software | Display/Image | Replacement | Remove or Replace
Z-1995-2010 | 55206 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems enhanced AccuTnl Reagent Kit; Part Number: A78803 |  |  | March 11, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because as a follow up to PCA # 12605 dated December 16, 2009, Beckman Coulter has confirmed additional reports that each of the following assays may produce lower than expected results when run in conjunction with Access AccuTnl (REF A78803):(1) Access Testosterone (REF 33560);(2) Access Thyroid Uptake (REF 33810);(3) Access Total T4 (REF 33800).This issue can occ | Other | The firm, Beckman Coulter, issued a "URGENT: PRODUCT CORRECTIVE ACTION" (PCA) letter with attached Customer Response form dated March 24, 2010 to all customers. The letter describes the issue, impact, action and resolution. In the PCA beckman Coulter recommends the following mitigation to reduce the risk of generating erroneous results: (1) Verify the APF software on the immunoassay system. If the updated APF software listed in the table above was already loaded , skip the remainder of this section and proceed to the Resolution section;(2) If the updated APF software has not yet been loaded, but is available, load the software appropriate for the system, skip the remainder of this section and proceed to the Resolution section;(3) If the updated APF software is not immediately available, please contact Technical Support at 1-800-854-3633 in the United States and Canada. Consignees outside the United States and Canada, contact their local Beckman Coulter Representative and follow the instructions below until the software has been loaded on their system. (a) Access and Access 2 users should batch process Access Testosterone, Thyroid Uptake, or Total T4 samples; (b) For UniCel DxI users, assign Access AccuTnI (REF A78803) to a different reagent pipettor(s) than Access Testosterone, Thyroid Uptake, or Total T4;(c) The Access Testosterone, Thyroid Uptake, or Total T4 results should be interpreted in light of the total clinical presentation of the patient, including: symptoms, clinical history, data from additional tests, and other appropriate information. Abnormally low Access Testosterone, Thyroid Uptake, or Total T4 results from samples tested in conjunction with Access AccuTnI (REF A78803) should be reviewed in the context of other diagnostic tests.Customers were instructed that after loading the updated APF software, the Access AccuTnI assay (REF A78803, test code TropI, test ID 275) must be recalibrated. Recalibrate of the Access Testosterone, T |  | 32407 | 32,407 units | Worldwide distribution: USA and countries of Algeria, Australia, Bangladesh, Belgium, Bulgaria, Canada, Chile, Czech Republic, France, Germany, Greece, Hong Kong, Ireland, Israel, Italy, Lebanon, Libyan Arab Jamahiriya, Monaco, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Poland, Puerto Rico, Romania, Russian Federation, South Africa, Spain, Swaziland, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | May 04, 2012 | 785 | 32407 | Other | Output/Calculation | Instructions/Software update | Software Update
Z-2020-2010 | 54688 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | Angiotech Galactography Kit. |  |  | September 01, 2009 | 2009 | Medical Device Technologies, Inc. | Medical Device Technologies doing business as Angiotech has conducted a recall on Angiotech Breast Localization Needles, Soft Tissue Biopsy Needles, Vascular Access Needles, Bone Biopsy Needles, Access Needles, Galactography Kit, Stabilization Needles and Needle Guides for packaging integrity. | Other | Angiotech sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated September 1, 2009, to all customers. on September 2, 2009 by UPS 2nd day delivery with an assigned tracking number. The letter described the product, problem and action to be taken by customers. The customers were instructed to please review your current inventory and segregate the affected lot numbers, complete the attached Customer Acknowledgement Form and return via fax to our Quality Assurance department at 1-352-338-0662 or 1-800-333-0440. If you are a distributor of this product, you are responsible for notifying any customer you may have sold the product to.If you have any questions concerning this notification, please contact Shannon Brooks at (352) 338-0440 ext.355 or Katrenia Williams at ext. 353. |  | 2630 | 2,630 | Worldwide distribution: USA and countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BERMUDA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, ESTONIA, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, GREECE, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PERU, PHILIPPINES, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, and UNITED ARAB EMIRATES. | Terminated | February 25, 2011 | 542 | 305698 | Not_Computer
Z-1999-2010 | 55790 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Microstream O2/CO2 NASAL FILTERLINE ADULT 25UN |  |  | May 26, 2010 | 2010 | Oridion Medical 1987 Ltd. | The Microstream Filterline may not be recognized by the monitor when connected. No CO2 readings will be displayed and no alarms may sound for unmonitored patients. Replacing the filterline may delay treatment. | Process design | Oridion notified customers by letter on 5/26/10 and requested that they remove the FilterLines from the related production lots from the field and an immediate replacement with FilterLines from subsequent production lots. Questions are directed to the responsible person at the supplier named in the letter, or to Oridion Capnography at 888-674-3466. |  | 2 | 2 units | AZ, CA, FL, HI, LA, , MA, MO, NJ, OH, PA, WA, WI, | Terminated | June 21, 2012 | 757 | 21212327 | Not_Computer
Z-1934-2010 | 55648 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | Constellation Vision System, Model: Constellation T, Catalogue Number: 8065751558 |  |  | April 30, 2010 | 2010 | Alcon Research LTD dba Alcon Laboratories, Inc. | Alcon identified system performance and machine settings that may impact the infusion performance of the CONSTELLATION Vision System, causing unexpected shut down, unresponsive touch screens, and fluidics issues. | Software design | The firm, Alcon, sent a "URGENT: MEDICAL DEVICE RECALL" letter dated July 2, 2010, along with training material to all customers. The letter informed the customers of the problem and indicated that Alcon Surgical Sales representative will be contacting them in the near future to follow-up on this training material and to answer any questions you may have. Alcon informed the customers that this action will not require removal of the CONSTELLATION Vision System from their facility. The letter also informed the customers that their Alcon Technical Customer Service Representative will be contacting them when we are authorized to provide you with the software and hardware updates for your system.For assistance and questions regarding the CONSTELLATION Vision System, customers were told they may contact their Alcon Surgical Sales Representative or Alcon Technical Consumer Affairs at 1-800-561-6466, Option 8. |  | 824 | 824 units total | Worldwide distribution. | Terminated | September 19, 2011 | 507 | 824 | Computer
Z-1998-2010 | 55675 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | O-Arm Imaging System |  |  | May 10, 2010 | 2010 | Medtronic Navigation, Inc. | The possibility that the real time image display can freeze up during use of a 2D continuous fluoroscopy. | Device Design | The firm, Medtronic, Inc., sent an "Urgent Field Safety Notice" letter dated May 10, 2010, to all custmers. The letter describes the product, problem and actions to be taken by customer. The firm recommended that the users of the O-Arm Imaging System reveiw the information provided to ensure patient safety remains uncompromised during use of the product. The customers were instructed to discontinue use and contact your local Medtronic Navigation representative or call our Technical Services Group at 800-595-9709 or (702) 890-3200, if the image is not updating as expected. A Medtronic Navigation representative will be contacting you within the coming months to schedule a time to evaluate and service the potentially affected components within your system. A service visit will also be scheduled starting November 2010 to update your system with a software based solution that will prevent this issue from occuring.If you have any questions about this notification, please contact your local Medtronic Navigation representative or call our Technical Services Group at 800-595-9709 or (702) 890-3200. |  | 221 | 221 | Worldwide distribution: USA and countries of Germany, Singapore, South Africa, Poland, Libya, Belgium, Switzerland, Turkey, Italy, Denmark, France, Spain, Finland, Czech Republic, Ireland, Slovakia, the NetherLands, Austria, UAE, Kuwait, Israel, Australia, Sweden, Canada, South Korea, Japan and India. | Terminated | October 07, 2011 | 515 | 221 | Software | Display/Image | Software update | Software Update
Z-1997-2010 | 55544 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | PVT-681MV Endocavitary Transducer used with the Aplio XG and Xario XG Ultrasound Systems. |  |  | January 05, 2010 | 2010 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems is initiating a field correction on the PVT-681MV Endocavitary Transducer used with the Aplio XG and Xario XG Ultrasound Systems due to an incorrect sweep direction setting. Images acquired using the endo-cavitary transducer PVT-681MV and cross-sectional images created from these volume images are reversedhorizontally on the Aplio XG (SSA-790A) and the Xario XG ( | Other | Toshiba America Medical Systems (TAMS) will issue notification to consignees via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS as proof of notification. matter. Questions regarding this letter are directed to the firm at (800) 421-1968 or to your local Toshiba representative. |  | 27 | 27 | Nationwide | Terminated | February 29, 2012 | 785 | 27 | Not_Computer
Z-1996-2010 | 55895 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | i-STAT PT/INR Cartridge |  |  | May 05, 2010 | 2010 | Abbott Point of Care Inc. | A high mean bias (12-16%) was documented in 8 lots of i-STAT PT/INR cartridges. This may lead a clinician to believe that a patient is adequately anticoagulated when they are not. Inadequate anticoagluation therapy may be given and issues with clotting and/or thrombosis could occur. No reports of injury have been received. | Other | Urgent recall notifications were sent on May 13 2010. Customers were instructed to discontinue use and to return the identified lots. Additional questions are directed to Abbott Point of Care Technical Support at 800-366Âż8020, Option 1 or your Abbott Point of Care representative. |  | 164072448 | 44976 in US, 3648 in rest of world. | Worldwide distribution: Argentina, Hong Kong, China, Finland, Greece, India, Portugal, Russian Federation, Saudi Arabia, United Arab Emirates, Italy, Spain, Australia, Germany, UK, Sweden, The Netherlands, France, Canada. | Terminated | September 22, 2011 | 505 | 164072448 | Not_Computer
Z-2025-2010 | 55867 | No | GNC | Salmonella spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Difco Salmonella 0 Group A Antigen, catalog #240731, packaged in 5 mL glass vials, labeled in part ***Becton, Dickinson and Company, Sparks, MD 21152*** |  |  | May 19, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic reagent to aid in the diagnosis of Salmonellosis may exhibit decreased or no reactivity. If proper controls are run with the lot in question, it would become apparent that the antigen is non reactive and patient specimen results should not be reported. If control testing is not performed a salmonella-infected patient could go undiagnosed, with possible adverse effects. | Nonconforming Material/Component | The recalling firm notified customers by fax and e-mail on 05/19/10. The notification advised that the referenced lot number of Salmonella 0 Group A Antigen may exhibit decreased or no reactivity when testing the patient's serum directly for homologous antibodies by either a slide or tube agglutination test to aid in the diagnosis of Salmonellosis. The recalling firm also advised that they would issue replacements for the referenced lot number under Catalog number 240731, which would not be available until approximately July 15, 2010. Consignees were requested to return enclosed response form.Customers with questions about replacements should contact BD customer service at 1-800-675-0908. For all other inquiries, please contact BD Technical Services Department at 1-800-638-8663. |  | 93 | 93 units | Product was sold to distributors nationwide and to affiliates in Belgium and Korea. | Terminated | July 21, 2010 | 63 | 93 | Not_Computer
Z-2027-2010 | 54565 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | OneTouch SureStep and SureStep Pro Test strips, for use with the OneTouch SureStep Blood Glucose Meters, in 100 count and 50 count packaging distributed by LifeScan Inc., Milpitas, CA, Manufactured by LifeScan Puerto Rico |  |  | February 24, 2010 | 2010 | Lifescan Inc | Products may provide inaccurate results at readings over 400 mg/dL | Other | The firm, Lifescan, Inc., issued two "Urgent Medical Device Recall" letters dated February 24, 2010 to all customers for the OneTouch SureStep Test Strips and the SureStep Pro Test Strips. The letters describes the product, problem and action to be taken by the customers. The customers were instructed to look at your test strip vials for the recalled lots and check the strips or if you have test strips from the lot, call LifeScan Customer Service at 800-574-6139 between 5 a.m. to 7 p.m. Pacific Time, seven days a week or visit www.SureStep.com. Replacement product will be shipped to you immediately at no charge. The customers were instructed to continue to test your blood glucose, while wafting for replacement. Also, the customers are asked to continue use of the meter with precautions if results over 400 units are observed. For more information, please visit www.SureStep.com or call our Healthcare Professional Line at 888-455-5627. |  | 1014350 | 1,014,350 | Worldwide distribution: USA and countries of Europe, Latin America and Asia. | Terminated | December 08, 2010 | 287 | 9640700 | Not_Computer
Z-2026-2010 | 54521 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ADVANTAGE Turbo 2-Button Shaver. Rx Only, Made In USA, NON STERILE. |  |  | October 29, 2009 | 2009 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec is recalling Advantage, 2-Button Turbo Handpiece Catalog Number D9924. These products may self-activate due to moisture intrusion into the housing body. This issue has been isolated to products manufactured between June 7, 2004 and June 1, 2006. | Other | ConMed Linvatech notified affected customers by letter sent by FedEx next day delivery. Customers were asked to complete a response form and indicated if they have the product or have transferred the product. A Service Request number will be assigned for the returned product.Additional questions are directed to the firm's Customer Service department at 800-535-8536. |  | 2373 | 2,373 | Worldwide distribution. Products were distributed domestically throughout the U.S. Also United Arab Emirates, Argentina, Australia, Belgium, Switzerland, Chile, China, Germany, Denmark, Spain, France, United Kingdom, Hong Kong, India, Italy, Japan, Lithuania, Morocco, Netherlands, New Zealand, Singapore, Turkey, South Africa. | Terminated | August 19, 2010 | 294 | 2373 | Not_Computer
Z-2031-2010 | 55575 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | DePuy ASR 100 Acetabular Implant, 44 mm, 999800744, sterile, DePuy International, Ltd, Leeds, United Kingdom. |  |  | March 05, 2010 | 2010 | Depuy Orthopaedics, Inc. | Notification to clinicians of new revision rate data/information regarding the use of the device. As part of the post-market surveillance, DePuy is continually evaluating data from a variety of sources including national joint replacement registries, published literature, company sponsored clinical trials, internal complaints data and unpublished clinical research reports. As part of the post-m | Other | DePuy Orthopaedics issued a Field Safety Notice to share the new data with surgeons and to reiterate the importance of correct implant positioning and patient selection in achieving optimal implant performance and survivorshop. For additional information regarding this matter, please contact the DePuy Scientific Information Office at 888 554 2482. |  | 388 | 388 | Nationwide | Terminated | May 10, 2012 | 797 | 37834 | Not_Computer
Z-2030-2010 | 55900 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic U-CLIP Double Arm Approximation Device, model V100D U-Clip, Catalog Number V100DRFRN2. Sterilized using Irradiation. The U-Clip consists of specially designed vascular clips optionally connected to needles via flexible members. These clips may or may not include coils surrounding a Nitinol wire core. This design allows precise placement of clips prior to closure and facilitation from standard medical and implantable grade materials. Accessories include Clip delivery and removal devices as well as vessel cutting and retention tools. |  |  | April 02, 2010 | 2010 | Medtronic Cardiovascular Revascularization &amp; Surgical Therap | Medtronic received reports involving the V100D U-Clip where during handling, the release mechanism has become separated from the needle assembly. These instances of separation have occurred when pulling the clip through prosthetic material, or when grasping and releasing the clip using robotic instrumentation. While this issue does not impact performance of the clip once implanted, separation of | Other | The firm, Medtronic CardioVascular, sent an "Urgent Medical Device Recall Notice", dated April 2, 2010, to Risk Managers of each affected account. The letter described the issue, identified affected product, and the action to be taken by customers. The customers were ask to immediately discontinue use of affected devices and to quarantine all unused V100D U-Clip inventory. The letter stated that a Medtronic representative will contact the account to arrange return of the V100D U-Clips devices and that credit will be issued upon return of unused devices. Medtronic has stopped manufacture of the product therefore replacements will not be available. Devices can be returned to Medtronic CardioVascular, 7611 Northland Drive, Brooklyn Park, MN 55428. The customers were ask to please fill out the Medtronic Recall Certificate and fax it to 651-367-2620.If you need additional information, please contact CardioVascular LifeLine Technical Services at 877-526-7890, or you local Medtronic sales representative. |  | 22946 | 22,946 ( 20,714 US, 2,232 OUS) | Worldwide distribution: USA including states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MT, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI and countries of Australia, Belgium, Canada, China, France, Germany, Greece, Hong Kong, Hungary, Japan, Korea, Netherlands, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, United Kingdom, and Venezuela. | Terminated | December 17, 2011 | 624 | 22946 | Not_Computer
Z-2029-2010 | 56013 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic StealthStation S7 System |  |  | June 14, 2010 | 2010 | Medtronic Navigation, Inc | Excessive wear at the power cord retractor or incorrect assembly of the power plug may cause sparking, intermittent loss of power, or electrical shock. | Device Design | Customers were notified by letter on Jun 14, 2010, advising them of the problem. Instructions were provided to continue using existing retractable power cords until they could be replaced with non-retractable cords by Medtronic, beginning on Aug 17. Further information is available from the company by calling 800-595-9709. |  | 370 | 370 units | Nationwide, including 3 VA facilities. No military distribution. Foreign accounts in Australia, Austria, Belgium, Canada, Chile, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Ireland, Italy, Latvia, Libya, Lithuania, Luxembourg, Malaysia, Namibia, Netherlands, New Zealand, Poland, Puerto Rico, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, United Arab Emirates. | Terminated | February 18, 2011 | 249 | 370 | Battery | Physical Safety Hazards | Replace | Remove or Replace
Z-2062-2010 | 56004 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Hemashield Double Velour Cardiovascular Fabrics. |  |  | May 24, 2010 | 2010 | Maquet Cardiovascular, LLC | Various Hemashield products may develop glycerol droplets on the inner blister of the package if not stored under proper conditions, which could lead to the prolongation of surgery and potentially compromise patient safety. | Package design/selection | Letters were sent by Federal Express commencing July 12, 2010.Distribution has to be obtained from Boston Scientific.Questions are directed to the local Sales/Service Representative, or to the firm's Customer Service at 1-888-880-2874. |  |  | xx | Nationwide Distribution. | Terminated | July 14, 2011 | 416 |  | Not_Computer
Z-2075-2010 | 56139 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Electrical Venous Occluder, Catalog 12-80-00, for use with Sorin HLM C5 and HLM S5. |  |  | June 18, 2010 | 2010 | Sorin Group USA, Inc. | Although accessories were designed for use with two models, the HLM C5 and the HLM S5, the submission that included references to the Mast Roller Pump, B-Care5, and the Electrical Venous Line Occluder accessories covered the new HLM S5 system. The three accessories were not cleared for use with the HLM C5. | PMA | The two customers who had received the suspect device accessories were identified and were contacted by letter on 6/18/2010. They were told to immediately discontinue use of the accessories and arrange for return of the products. They were also told that they could continue to use the CD5 Machine without the recalled accessories. A Product Recall Response Form was to completed by each consignee. Questions are directed to the firm at 800-221-7943. |  | 1 | 1 unit | CA, PA, TN. | Terminated | July 21, 2010 | 33 | 3 | Not_Computer
Z-2072-2010 | 56056 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art SystemÂż, ver. 4.0.0, 4.0.1, 4.0.2, ThomoTherapy Incorporated 1240 Deming Way, Madison, WI 53717. The TomoTherapy HI-ART SystemÂż is intended to be used as an integrated system for the planning and delivery of intensity modulated radiation therapy (IMRT). The HI-ART System provides precise delivery of radiation to tumors or other targeted tissues while minimizing the delivery of radiation to vital healthy tissue. |  |  | June 23, 2010 | 2010 | TomoTherapy Incorporated | An issue has been identified, as a result of input received from one TomoTherapy Hi-Art Treatment System customer.With a particular workflow, the Planning Station's Draft and/or final Plan Report can provide incorrect dose statistics. | Software design | Consignees were sent on 6/23/10 a TomoTherapy "Urgent Field Safety Notice Medical device Correction" letter dated June 22, 1010. The letter addressed the Issue, Product Affected, Recommended Action and Resolution.Prior to generating a draft or final Plan Report for a given plan, the plan must be closed and reopened at least once. This issue will be remedied by Hi-Art software version 4.0.3, scheduled for release in late 2010.Customers with questions were provided with a contact in the letter at the TomoTherapy Customer Interaction Center. The firm can be contacted at 608-824-2800. |  | 29 | 29 | CA, CO, IL, LA, MI, MO, NY, ND, OH, OK, PA,VA, WI, BELGIUM CANADA, KOREA, MALAYSIA, SWITZERLAND, ITALY, GREAT BRITAIN, FRANCE. | Terminated | December 17, 2011 | 542 | 29 | Software | Output/Calculation | Software update | Software Update
Z-2071-2010 | 56059 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Fathom"-14 Steerable Guidewire, Peripheral Vascular Use, Pre-Shaped 300 cm x 10 cm, REF 50-815, UPN M001508150 Sterilized with ethylene oxide gas, Rx Only, Made in 2405 Orion Circle, West Valley City, UT 84119 USA. |  |  | June 03, 2010 | 2010 | Boston Scientific Corporation | Boston Scientific is initiating a Medical Device Recall regarding specific lots of the FathomTM-14 Steerable Guidewire. For lubricity, the distal portion of these guidewires is coated with hydrophilic polymer and the proximal portion is coated with polytetrafluoroethylene (PTFE). Through internal inspection process, the firm determined that Fathom guidewires from affected lots may exhibit a P | Process design | Consignees were sent on 6/3/10 an "Urgent Medical Device Recall" letter dated 3 June 2010. The letter was addressed to Risk Manger/Interventional Radiology Manger/Material Manger. The letter described the problem and the product involved. The letter included the "Customer Instructions" which stated to Immediately discontinue use of an segregate recalled product and, requested customers to return the Reply Verification Tracking Form. Questions are directed to the firm at 510-624-2563. |  | 78 | 78 units | AL, AR, CA, CO, CT, FL, GA, ID, IL, KS, KY, LA, MD, MA, MS, NV, NY, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT,VA, WA, WT. | Terminated | December 10, 2011 | 555 | 745 | Not_Computer
Z-2080-2010 | 56084 | Yes | JOH | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Cook Medical Ciaglia Blue Rhino G2 Advanced Percutaneous Tracheostomy Introducer Set, C-PTISY-100-HC-PERC6, G13166. |  |  | April 21, 2010 | 2010 | Cook, Inc. | These set and/or tray products are being recalled because they include a Covidien Shiley 6PERC or Covidien Shiley 8PERC tracheostomy tube component that has been associated with cuff leakage as a result of the pilot balloon inflation assembly. | Nonconforming Material/Component | Cook Medical sent an Urgent Product Recall letter dated April 21, 2010, to hospitals and other user facilities receiving product direct from Cook as well as distribution centers receiving direct shipments from Cook. Consignees were directed to cease distribution of product and notify their customers immediately. |  | 28819 | 28819 total, all products | Worldwide distribution: USA, Canada, Mexico, Australia, Chile, Costa Rica, Denmark, Dominican Republic, Guatemala, India, Malaysia, Panama, Peru, Puerto Rico, Republic of Korea, Trinidad and Tobago, and Uruguay. | Terminated | January 23, 2012 | 642 | 28819 | Not_Computer
Z-2100-2010 | 56158 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | 19" Barco MFCD 1219 (Touch Screen) + Low Profile Stand and Rack with LCD Arm |  |  | June 30, 2010 | 2010 | AGFA Corp. | Monitor fell from the mounting bracket used to support the monitor. | Other | The firm, AGFA Healthcare sent an 'URGENT SAFETY NOTICE" dated June 30, 2010, via FED-EX to customers. The letter describes product, problem and actions to be taken by the firm and customers. AGFA will provide a safety check to the customers facility, have a service representative visit , in the near future, their site to inspect the monitor/s and attach a warning label on each monitor checked for preventative action. The customers are ask to distribute the information within their facility to all individuals who need to be aware, to complete the attached URGENT SAFETY NOTICE FEEDBACK FORM as soon as possible and return via fax at 864-421-1664. If you have any questions about this matter, please feel free to contact me via email at Debbie.norris@agfa.com or call me at 864-421-1754. |  | 805 | 805 units | Worldwide distribution: USA and countries of Puerto Rico and Canada. | Terminated | April 28, 2011 | 302 | 805 | Not_Computer
Z-2109-2010 | 56130 | No | FYD | Air-handling apparatus for a surgical operating room. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MICROTEK MEDICAL, INC, STACKHOUSE LASER RESISTANT SUCTION PROBE, Size: 1/4" x 9" (0.64cm x 22.9cm), REF 24014, QTY 10 EA, STERILE EO, 2015-05, Microtek Medical, Inc. , Columbus, MS 39702, www.microtekmed.com, CE0044, EC REP Microtek Medical B.V. , Zutphen, The Netherlands, Manufactured in the U.S.A. Suction probe for a smoke evacuation device |  |  | May 25, 2010 | 2010 | Microtek Medical Inc | Product labeled as sterile was distributed without sterilization. | Other | The firm notified its consignees of the problem and requested return of the product by telephone on 05/25/2010. They followed with a letter dated on 06/03/2010 re: Product Removal stating the reason for the recall and confirming the receipt of the recalled product. For further information please contact (662) 244-3220 |  | 20 | 2 cases of 10 units | Nationwide distribution including the states of Montana and Georgia | Terminated | August 12, 2010 | 79 | 20 | Not_Computer
Z-2111-2010 | 56085 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Temno Coaxial Chiba Fine Needle Aspiration; a coaxial introducer needle (20G x 15 cm) with a Chiba biopsy needle (22G x 20 cm); sterile, for single use only sets, 5 sets per case; CareFusion, McGaw Park, IL 60085 USA, Made in Dominican Republic; REF CCH2220 |  |  | June 25, 2010 | 2010 | Carefusion 2200 Inc | The introducer needle is only 10 cm in length instead of the required 15 cm. | Mixed-up of materials/components | Cardinal Health, on behalf of CareFusion, telephoned the customers on June 15, 2010 and sent a follow-up "URGENT: PRODUCT RECALL" letter on the same date via UPS next day air. The letter describes product, problem, and action taken by customers. The customers were instructed to immediately stop using affected product, return the enclosed acknowledgment form via fax to 847-686-9101 (note: Cardinal Health is required to confirm receipt of this notification from the customers), notify any customers to whom they may have distributed product to, contact CareFusion Customer Service at 800-653-6827 to arrange for credit and return of the affected product, and send the affected product to CareFusion, 1240 Waukegan Rd, Waukegan, IL 60085, Attn: Scott Fitzer, after obtaining an RMA/RGA number from CareFusion.If you have any questions, please contact 800-292-9332. |  | 65 | 65 units | Nationwide distribution: California, Florida, Georgia, Illinois, Kentucky, Massachusetts, New Jersey, New York and Pennsylvania. | Terminated | January 24, 2011 | 213 | 65 | Not_Computer
Z-2107-2010 | 56169 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | GEM Check Coag CNTRL Normal APTT Whole Blood Controls Package Insert: |  |  | July 08, 2010 | 2010 | International Technidyne Corporation | Certain lots of GEM Check Coag Whole Blood Controls contain an error in the package insert where the incorrect Acceptable Performance Ranges were published. | Error in labeling | Urgent: Medical Device Correction Notices, dated July 1, 2010, were sent by Federal Express on July 8, 2010 to all first consignee end users. The letters identified the affected product and asked customers to review and follow all instructions as outlined in the ITC in the communication. The info should be shared with laboratory staff and the notification should be retained as part of the laboratory Quality System documentation. The enclosed response form should be completed and returned to ITC within 5 days. |  | 3040 | 95 kits C9GC5001; 32 kits G9GC5003; 46 kits J9GC5004; 5 kits C0GC5001; Each kit contains 15 vials. | Worldwide Distribution -- USA, including states of NY, TN, CA, AR, IL and TX and countries of Australia, Italy, Spain, Germany, Chile, Spain, Malta, Belgium, Austria, and Switzerland. | Terminated | September 23, 2010 | 77 | 3141 | Not_Computer
Z-2110-2010 | 56114 | No | KFF | Heparin assay. | Hematology | Hematology | 510(k) | RQCHRT HRT Control Level 1 &amp; 2; |  |  | June 24, 2010 | 2010 | International Technidyne Corp. | An ingredient in the RQCHRT Whole Blood Controls for use with Hemochron Kaolin and Celite HRT Test tubes may become less potent over time and result in excessive High LQC Failures. | Other | ITC sent "URGENT MEDICAL DEVICE RECALL" letters dated June 24, 2010, via Federal Express to Laboratory Director, Clinical Laboratory Manager, or Point of Care Coordinator consignees/customers. The letter described the product, problem and action to be taken by customers. The customers were instructed to check their inventory for the product, if they have boxes or loose vials to please stop using the product, remove it from inventory and contact ITC Technical Support to arrange return. They were also instructed to complete the attached Customer Account Tracking Form, whether or not they have remaining inventory, and return it by fax, email or mail.If you have any questions, please call ITC Technical Support at 800-631-5945 (US) or 732-548-5700 (international), Extension 4707 or email techsupport@itcmed.com. |  | 114 | 38 in US, 3 Out of US | Worldwide distribution: USA including states of CA, NC, TX, FL, MD, OK, NY, and NJ and country of Italy. | Terminated | September 23, 2010 | 91 | 114 | Not_Computer
Z-2112-2010 | 56090 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Results Manager/Results Manager Administration Tool |  |  | June 11, 2010 | 2010 | AGFA Corp. | Agfa Service technician had incorrectly modified a report template at one site resulting in erroneous anatomic segment locators. | Other | A letter was sent to the consignee on 6/11/2010 indicating what steps would be taken by Agfa concerning the problem. The letter also included a feedback form which was to be returned to Agfa. The recall was expanded and a "URGENT SAFETY NOTICE" was sent to each additional consignee on January 25, 2012. The letter described the potential issue and mitigation. Acknowledgment, via FAX-Back or email, that the information was received and understood has been requested from the additional consignees. |  | 358 | 358 | AK, AZ, CA, CO, DC, FL, GA, HI, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, MS, MT, NH, NJ, NY, NC, OH, OK, OR, PA, RI, SC, SD, TN, TX, VT, VA, WA, WV, WI and Canada | Terminated | July 02, 2013 | 1117 | 358 | Not_Computer
Z-2113-2010 | 56057 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | BleaseSirius Anaesthesia System, gas machine. Blease Beech House, Chiltern Court Asheridge Road, Chesham, Bucks HP5 2PX, England |  |  | June 24, 2010 | 2010 | Spacelabs Healthcare, Llc | Potential for the touch screen on the ventilator portion of the device to go blank when touched. Possibility that this failure may also cause mechanical ventilation to stop. | Other | On 06/24/2010, Spacelabs began mailing the URGENT-MEDICAL DEVICE CORRECTION letter to their US consignees and e-mailed the same letter to their foreign consignees.The letter describes the touch screen of the BleaseSirius Anesthesia machine goes blank when touched and the possibility that this failure may cause the mechanical ventilation to stop.The consignees are asked to weigh the benefits versus the risk when deciding to use the device or not until the firm upgrades the device. The letter also instructs the consignees to notify their staff about the possibility of the display failure to prevent any adverse event that may occur. The letter states that their field service engineer will contact consignees to schedule a time to visit and upgrade all affected devices at no cost.The US consignees are advised to call 1-800-522-7025, Select 2, and the foreign consignees are advised to call 1-425-657-7200 x5089 for technical support or questions regarding the recall. |  | 80 | 80 units total | Worldwide Distribution -- USA, including states of IA, GA, LA, NC, TN, and TX and countries of AUSTRALIA, CHILE, CHINA, INDIA, ITALY, PHILIPPINES, MEXICO, POLAND, SINGAPORE, THAILAND, TURKEY, UNITED KINGDOM, and VENEZUELA. | Terminated | December 03, 2010 | 162 | 80 | Software | Display/Image | Software update | Software Update
Z-2115-2010 | 56052 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Hi-Art SystemÂż, Version 4.0.0 &amp; 4.0.1. |  |  | February 22, 2010 | 2010 | TomoTherapy Incorporated | In some cases, the patient's diagnostic CT image is narrower than the Hi-Art radiotherapy couch image. | Software design | Consignees were sent on 2/22/10 a TomoTherapy Incorporated "Urgent Medical Device Correction Field Safety Notice" letter dated Monday, February 22, 2010. The letter was addressed to TomoTherapy Customer and identified the the affected product, and included a description of the Issue, Product Affected, Recommendation Actions and Resolution. The issue will be remedied by software versions 4.0.2 and later. Questions should be directed towards TomoTherapy Customer Interaction Center by e-mail or telephone. |  | 7 | 7 | Worldwide Distribution -- MI, MO, VA, and WI and country of CANADA. | Terminated | February 10, 2011 | 353 | 7 | Software | Display/Image | Software update | Software Update
Z-2060-2010 | 55224 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Terumo HX2 Temperature Management System 120V, 50/60 Hz catalog number 809810, Terumo Cardiovascular Systems Corp., Ann Arbor MI. |  |  | June 11, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The HX2 Temperature Management System mixing valve may fail resulting in an unusable water channel and then the machine would go into standby mode. | Device Design | The firm, TERUMO Cardiovascular Systems, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 11, 2010, to the consignees. The letter describes the product, reason for correction, what to do in the event of failure, potential hazard, correction, affected population and action to be taken by the customers. The customers are instructed to review the correction notice, assure that all users are aware of this notice, and complete, sign and fax the attached Customer Response Form to Terumo CVS Customer Service at 800-292-6551. The Terumo CVS Customer Service Dept. will contact you upon receipt of the Customer Response Form to schedule the replacement. The firm plans to inspect and correct all units in the field within 5 months.We encourage you to contact Terumo CVS Customer Service at 1-800-521-2818, Monday-Friday, 8 AM-6 PM EST, with any questions or concerns. |  | 100 | 100 | Nationwide Distribution: AR, CA, CO, CT, FL, HI, IA, IL, KS, LA, MD, MI, MO, MS, NY, OH, PA, TX, UT, WA and WI. | Terminated | April 13, 2011 | 306 | 100 | Not_Computer
Z-2067-2010 | 55868 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. |  |  | May 26, 2010 | 2010 | Ge Healthcare It | There is a potential safety issue associated with the use of GE Centricity Laboratory software where results from interfaced external reference laboratories are not displayed or stored correctly. This happens when the test is set up in Centricity Laboratory as a decimal (numeric) type result and the test result sent from the reference lab includes a comma in the result field. | Software design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated May 26, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the patient safety issue involving results from interfaced external reference laboratories that are not displayed or stored correctly if an unexpected alpha character is encountered in a result message that is expected to have a numeric value. This happens when the test is set up in Centricity Laboratory as a decimal (numeric) type result and the result sent from the reference lab includes a comma in the result field. The letters provided the users with specific safety instructions to follow until the software is updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to the their GE Service Representative or the GE Helpdesk at 888-778-3375. |  | 37 | 37 units | Worldwide distribution: California, Georgia, Michigan, Minnesota, New York, Ohio, Oregon, Tennessee and Texas, and internationally to Australia, Canada, China, England, India, Malaysia, New Zealand, Qatar, Scotland and Singapore. | Terminated | July 15, 2011 | 415 | 37 | Computer
Z-2093-2010 | 55578 | No | GXY | Cutaneous electrode. | Neurology | Neurology | 510(k) | The "RS-FBG Full Back Conductive Garment" with electrode attachments for use with interferential and muscle stimulation with a Transcutaneous Electrical Nerve Stimulation device (TENS) and/or muscle stimulator. |  |  | May 17, 2010 | 2010 | International Rehabilitative Sciences, Inc. | Device distributed with labeling and IFU outside of their 510(k) clearance. | Other | The firm, RS Medical, mailed "URGENT NOTICE OF VOLUNTARY CORRECTIVE ACTION" letters dated May 14, 2010, to customers on May 17, 2010 and letters to physicians/consignees on May 19, 2010. The letters included an updated RS-FBG Operation Manual and stated what the update covered; i.e. that the device is indicated for use on the "middle and upper back". The customers were ask to please read your new manual and replace your original RS-FBG Full Back Conductive Garment Operation Manual with this updated version and discard your old RS-FBG Garment Operation Manual. RS Medical stated that your RS-FBG Garment is not affected by this Corrective Action, only the Operation Manual. RS Medical informed the customers that you may continue to use the RS-FBG Full Back Garment with your RS Medical Stimulator as prescribed by your healthcare provider.If you have any questions or concerns, please call our Toll Free Customer Service number at 1-866-417-1293, email us at RSFBG@rsmedical.com or visit our special web page at www.rsmedical.com/rsfbginfo. |  | 190073 | 108134 units, 81,939 subject to recall | Nationwide distribution. | Terminated | August 15, 2012 | 821 | 190073 | Not_Computer
Z-2094-2010 | 56033 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Rolko Trapeze handles, part 148671, for use with Care Assist bed P1170 and Hill Rom 1000 bed P1160, Hill Rom Batesville, IN. |  |  | June 11, 2010 | 2010 | Hill-Rom, Inc. | The firm received a complaint regarding the plastic housing used to support the trapeze handle. The firm investigated and found that the manufacturer of the trapeze handle didn't correctly process the plastic component. Therefore the plastic had the potential to break prematurely. | Process control | The firm, Hill-Rom, sent an "Urgent Field Safety Notice" letter dated June 11, 2010, to its customers. The letter describes the product, problem and action to be taken by the customers. Hill-Rom provided the quantity of handles needed to replace handles customers received during the affected period. The customers were instructed to immediately remove and replace all Trapeze Handles that were delivered during the affected period; to replace the Trapeze Handle Assemblies (with a manufacturing date from January 2007 to January 2009) used on their Patient Helper with the enclosed unit(s); complete and return the Customer Response Form as soon as possible, or by July 26, 2010; advise the firm immediately if they have relocated the beds and/or the accessories to another location, and forward the letter to any other facility personnel they deem appropriate. If you have any questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720. |  | 319 | 319 | Worldwide distribution: USA and Canada. | Terminated | February 18, 2011 | 252 | 319 | Not_Computer
Z-2101-2010 | 56161 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Merit Laureate Hydrophylic Guide Wire, Angled, Made In Ireland, CAT No: LWSTDA35180, 180cm, Sterile EO, Merit Medical Systems, Inc., South Jordan, UT 84095. |  |  | June 28, 2010 | 2010 | Merit Medical Systems, Inc. | Guidewires have decreased lubricity at the proximal end, inhibiting proper operation. | Process control | The firm, Merit Medical System, Inc., notified the Sales representatives by e-mail on June 28, 2010, and instructed them to contact all of their accounts and to advise them to discontinue use and return any unused products to Merit. A Product Recall Notice was to be completed by each facility under the direction of the rep and all activities were to be completed by July 9, 2010. Questions should be directed to Penny Goldman, Senior Product Manager at 801-208-4198 or Jackie Preece, Manager, Customer Service at 801-208-4365. |  | 437 | 437 units | Nationwide distribution: CA, CO, FL, GA, IL, IN, KS, MA, MD, MI, MN, NC, NH, NJ, NY, PA, and SD. | Terminated | October 19, 2010 | 113 | 964 | Not_Computer
Z-2095-2010 | 55465 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Evolis-Femoral Sizer Posterior Reference |  |  | September 11, 2009 | 2009 | Medacta Usa Inc | Medacta USA Inc is initiating a recall on the Femoral Sizer, an orthopedic manual surgical instrument, because the "L" and "R" markings are mislabeled which would cause incorrect usage. | Process design | The firm, Medacta USA Inc, sent an "Urgent Safety Notification" letter dated August 5, 2009, to the customers. The letter describes the product, problem and the action to be taken by the customers. The customers were instructed to check their inventory for devices listed, if found, to immediately sequester them and DO NOT use them in surgery. The customers were ask to please pack the affected femoral sizer(s) posterior reference for shipment to Medacta USA. Medacta informed the customers that Medacta USA personnel will contact them by telephone to provide them with an RMA number and the reworked femoral sizer(s) posterior reference will be returned as soon as they are available.If you have any questions, please feel free to contact the Medacta USA office at 805-437-7085 or contact Ms. Heather Neely directly at 805-886-1069. |  | 5 | 5 | Worldwide distribution: USA including states of Texas and Idaho and country of France. | Terminated | February 29, 2012 | 901 | 5 | Not_Computer
Z-2118-2010 | 55330 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Howmedica Osteonics Corp, |  |  | April 05, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | There is the potential for the posts on the Distal Resection Guide Stand to disassociate from the main body part. | Device Design | Stryker Urgent Product Recall letters dated April 8, 2010, were sent by Federal Express to Stryker Branches/agencies, Hospital Risk Management, Hospital Chief of Orthopaedics and Surgeons. All hospitals and surgeons that used the product in the last year were notified. |  | 1981 | 676 in US, 1305 internationally. | Worldwide distribution. | Terminated | August 28, 2012 | 876 | 1981 | Not_Computer
Z-2117-2010 | 55226 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics, Dimension Vista LOCI Reaction Vessels, catalog number KS855. For in vitro diagnostic use. |  |  | March 15, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Vessel pick up errors and vessel jams when the vessel is being loaded. The error halts the instrument and disrupts processing. Vessel flange is thin. | Process design | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice dated March 2010 informing users of the potential problem, to discard any remaining product, and if needed, to request a replacement lot. |  | 5544 | 5544 bags | Worldwide distribution: USA, Canada, Europe, Japan, Korea, and New Zealand. | Terminated | August 03, 2010 | 141 | 5544 | Not_Computer
Z-2119-2010 | 55357 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | X-Celerate Instrumentation Punch Thru Tibial Base Plate, Catalog numbers: 3750-0003. 3750-0004, 3750-0005, 3740-0006, 3750-0007, 3750-0009, 3750-0011, 3750-0013; |  |  | April 05, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The locating pins on the Punch Thru Tibial Baseplates could break or disassociate from the main body of the instrument. | Device Design | Stryker Urgent Device Correction letters dated April 5, 2010, were sent to all Stryker branches/agencies, hospital Risk Management and surgeons who have used the product by Federal Express. Questions on the bulletin should be directed to Scott Solis, Product Manager ,Knee Marketing at 201-831-5665. |  | 8276 | 8276 | Worldwide distribution: USA, China, Taiwan, India, Australia, Korea, Singapore, Germany, France, Spain, Italy, Switzerland, Poland, Romania, South Africa, United Kingdom, Netherlands, Sweden, Turkey, Japan, Chile, Brazil, Colombia, Mexico, Panama, Argentina, and Uruguay. | Terminated | August 28, 2012 | 876 | 8276 | Not_Computer
Z-2164-2010 | 55605 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | MIKA Speedblock, Size 8-12. |  |  | April 26, 2010 | 2010 | Encore Medical, Lp | A crack or complete fracture at the anterior chamfer cut slot may occur on the size 8, 10, and 12 MIKA Speedblocks. | Component design/selection | An "Urgent Field Safety Notice" Letter dated May 20, 2010, was sent to the customers. The letter described the product, problem and action to be taken by customer. The customers were instructed to locate, inspect and verify the quantity still in their possession, complete and return the field safety notice response page to Beth Onderlinde at beth.onderlinde@djosurgical.com or fax to 512-834-6313 by June 7, 2010, (completion of form is mandatory), and identify and inform customers that they may have distributed/sold product to, of the notice. The customer was also instructed to please inspect the listed part numbers prior to use and again after use, if a crack is found in the speedblock, contact your Customer Service Representative and be sure to utilize the femoral/tibial impactor (801-01-043) for impaction of the speedblock and reiterate to your physicians that the MIKA mini slaphammer (800-02-290) should be used for speedblock removal. The product will be replaced as it becomes available.If you have any questions, call Director of Commercial Logistics at (512) 834-6330 or International Customer Service at (512) 834-6275. You may also contact our EU Authorized Representative, MDSS GmbH, via email at info@mdss.com or via phone at +49 (0) 511 6262 8630. |  | 796 | 729 and 67 specials | Worldwide distribution. | Terminated | December 15, 2011 | 598 | 796 | Not_Computer
Z-2140-2010 | 55931 | No | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Exacta-Mix EVA Container, Sterile R, 2000 mL, Order No. REF: 140, Baxa Corporation. |  |  | May 26, 2010 | 2010 | Baxa Corporation | Fill port cap on TPN bags may become detached prior to use, thus compromising the sterility of the fluid pathway. | Process design | Consignees were notified by e-mail and Baxa Bag Product Notice letter on 05/26/2010 with instructions to not use any bags if the fill port cap is detached. Photos of a detached cap and properly attached cap were included. Defective bags could be returned to Baxa for credit. Information available at 800-678-2292. |  | 3306700 | 3,306,700 bags | Worldwide distribution, including USA, Argentina, Australia, Bermuda, Brazil, Canada, China, Colombia, Hong Kong, Malaysia, Mexico, New Zealand, Pakistan, Panama, South Korea, Taiwan, United Kingdom, and Uruguay. | Terminated | August 04, 2011 | 435 | 12612066 | Not_Computer
Z-2133-2010 | 55621 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | BARD DORADO PTA BALLOON DILATATION CATHETERS, PRODUCT CODE # DR135712 |  |  | May 04, 2010 | 2010 | Bard Peripheral Vascular Inc | The affected DORADO PTA Balloon Dilatation catheters contain catheter materials that, under certain conditions, are subject to premature breakage allowing the balloon to partially or completely separate from the catheter shaft. | Nonconforming Material/Component | The Bard Peripheral Vascular "URGENT MEDICAL DEVICE RECALL NOTIFICATION" with attached "RECALL AND EFFECTIVENESS CHECK FORM" was sent to the affected consignees beginning May 7, 2010 via FedEx. The recall notice informed the consignees of the reason for recall, the products affected and instructed them not to use or further distribute any affected product.The consignees were also instructed to call the firm's Recall Coordinator Silvia Carrillo at 1-800-321-4254 Option #2 Ext 2727 (M-F 7am to 4pm MST) or email at silvia.carrillo@crbard.com to facilitate the expedient return of the affected products. Additionally they were asked to fax the Recall and Effectiveness Check Form to BPV at 1-800-994-6772. |  | 72 | 72 units | Nationwide | Terminated | February 01, 2012 | 638 | 3279 | Not_Computer
Z-2120-2010 | 55367 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Howmedica Osteonics Corp., |  |  | April 05, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | One of the two posts on the Anterior Resection Guide could disassociate from the main body of the part which has been determined to be caused by a fracture weld. | Device Design | Stryker Urgent Product Recall letters dated April 5, 2010, were sent out via Federal Express to Stryker branches/agencies, Hospital Risk Management, hospital Chief of Orthopedics and Surgeons. All hospitals and surgeons that used the product in the last year were notified. Questions and/or concerns should be directed to Rich Wolyn at 201-972-2100. |  | 1974 | 682 units in US, 1292 units internationally. | Worldwide distribution: USA, Ireland, Germany, Sweden, The Netherlands, France, Italy, Switzerland, South Africa, United Kingdom, Japan, Argentina, Mexico, Miramar (city), China, Taiwan, India, Australia, Korea, New Zealand, Malaysia, Singapore, and Canada. | Terminated | August 28, 2012 | 876 | 1974 | Not_Computer
Z-2116-2010 | 55874 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | PowerPark is an accessory to the V or H Universal Stand and comprised of two parts: the stand mounted assembly (P12822) and a docking assembly (P12834). The stand mounted assembly is mounted to the V or H Universal stand and the docking assembly is plugged into a wall circuit. It provides a docking and charging station for the V and H Universal Stands. |  |  | May 27, 2010 | 2010 | Sonosite, Inc. | PowerPark system has a potential for loss of power at the corresponding wall outlet and/or circuit due to poor alignment. | Nonconforming Material/Component | On 05/27/10, SonoSite began contacting their consignees by telephone. On 06/01/10, SonoSite sent out the URGENT-MEDICAL DEVICE RECALL letter to their consignees. The letter describes the recall product as SonoSite PowerPark. Customers should call the Technical Support at 1-877-657-8118 or e-mail to service@sonosite.com for any questions regarding the recall.The letter instructs consignees to stop using the PowerPark immediately and return the product back to SonoSite in exchange for a full refund. The consignees will be contacted by the firm's Sales Representative and/or Clinical Applications Specialist to schedule an appointment to retrieve the device and deliver a refund. |  | 26 | 26 | Worldwide Distribution -- USA, including states of GA, NM, PA, TX, VA, and WA and country of Australia. | Terminated | April 18, 2011 | 326 | 26 | Battery
Z-2151-2010 | 56037 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Procedure Kit Custom |  |  | June 08, 2010 | 2010 | Terumo Cardiovascular Systems Corp | Vent port may be occluded and prevent delivery of fluid | Equipment maintenance | Terumo, Ashland, MA notified consignees on 6/8/10, byphone, fax, email and mail to advise distributors and end-users of the required recall actions. Follow-up communication sent by priority overnightmail on 6/9/10, with proof of delivery, specifying the awareness to the issue and necessary correctionactivities should a unit be discovered. Phase 2 initiated on 7/1/10 to sale rep to remove recalled product and perform a product reconciliation to ensure that all affected product is identified, returned to Terumo CVS, and that replacement product is sent immediately. |  | 48 | 48 packs | Distribution USA and Canada | Terminated | May 07, 2012 | 699 | 589 | Not_Computer
Z-2121-2010 | 56063 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MEDFUSIONÂż SYRINGE INFUSION PUMP, MODELS 3010 AND 3010a WITH SOFTWARE VERSIONS 2.0.2, 2.0.3, AND 2.0.4 MedfusionÂż Syringe Infusion Pumps are designed for use in critical care areas for the administration of fluids requiring precisely controlled infusion rates. |  |  | June 22, 2010 | 2010 | Smiths Medical ASD, Inc. | Smiths Medical is conducting a correction of MedfusionÂż Syringe Infusion Pumps, Models 3010 and 3010a, with obsolete software versions 2.0.2, 2.0.3, and 2.0.4. Smiths Medical has found that pumps with obsolete software may continue to run past the set volume limit. This can result in an over-delivery  if the syringe is over-filled. Smiths Medical has received no reports of serious patient in | Software design | Consignees were sent on 6/22/10 a Smith Medical "Urgent Medical Device Correction Notice', dated June 11, 2010. The letter was addressed to Dear Risk/ Safety Managers, Distributors, Clinical/ Biomedical Engineering, Clinicians, Critical Care Physicians, Pediatricians, Pediatric Intensivists, Neonatologists; and Nursing, Pharmacy,and Anesthesia Professionals. The letter described the problem and product involved. It provided Advise on Action to be Taken by the User which included to 1) Determine the software version of their Medfusion pump 2. Complete and return the Confirmation Form 3. Return the completed Software Update Test Form 4. If unable or unwilling to perform the software upgrade they provided information for returning the pumps to Smiths Medical for the upgrade. For further information please contact Smiths Medical at 651-628-7280. |  | 4816 | 4816 | Worldwide Distribution: AL, AK, AZ, AR, CA,CO, CT, DC, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, ME, MD, MI, MN, MS, MO, MT, NE, NV,NH, NJ, NM, NY, NC, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI. and countries of CANADA, MEXICO, NEW ZEALAND, BRASIL, HONG KONG, and GREAT BRITAIN. | Terminated | April 17, 2012 | 665 | 4816 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2172-2010 | 55871 | No | NRP | Sirolimus test system. | Toxicology | Toxicology | 510(k) | Architect Sirolimus Reagent Pack; 100 test pack containing 1 bottle (8.0 mL) Anti-Sirolimus Coated Microparticles, 1 bottle (8.0 mL) Sirolimus Acridinium-labeled Conjugate in Citrate Buffer, and 1 bottle (10 mL) Assay Diluent containing saline; Abbott Laboratories, Abbott Diagnostics Division, Abbott Park, IL 60064 USA; list 01L76-25. |  |  | May 27, 2010 | 2010 | Abbott Laboratories | The Architect Sirolimus assay lot 80162M100 is exhibiting higher than normal reports of barcode read errors. | Equipment maintenance | Abbott sent a "Product Recall" letter dated May 27, 2010, to all customers receiving the affected lot of reagent. The customers were informed of the increased frequency of barcode read errors due to the print quality of the barcode label. The following error codes have been observed for some kits within the lot: 4000 - Unable to read reagent bar code in position (x) on (y) carousel, 0208 - Reagent carousel scan error, and 2011 - Extra reagent bottle detected in position (x) on (y) carousel. The customers were requested to determine if they are currently using or have inventory of lot 80162M100, discontinue use of and destroy any remaining inventory of the lot, and retain a copy of the letter for their laboratory records. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. Customers outside the U.S., please contact your local Customer Service representative. |  | 875 | 875 kits | Worldwide distribution: USA including Puerto Rico, and countries of Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Croatia, Czech Republic, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Israel, Italy, Japan, Latvia, Lithuania, Malaysia, Mexico, New Zealand, Pakistan, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Taiwan, Thailand, Turkey, United Kingdom, Venezuela and Vietnam. | Terminated | July 15, 2011 | 414 | 875 | Not_Computer
Z-2166-2010 | 55507 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | The Well@Home Monitor by Patient Care technologies |  |  | February 02, 2009 | 2009 | Zoe Medical Incorporated | Touch panel may become unresponsive. | Component change control | Zoe Medical notified the 2 distributors by e-mail on January 26, 2009 and Feb 6, 2009 of the recall, requesting return of the units for correction. |  | 62 | 62 | Nationwide Distribution -- MA and MD. | Terminated | August 03, 2010 | 547 | 62 | Not_Computer
Z-2171-2010 | 55557 | No | NRP | Sirolimus test system. | Toxicology | Toxicology | 510(k) | IMx Sirolimus Reagent Pack; 100 test pack containing 1 bottle (9.7 mL) Anti-Sirolimus Antibody Coated Microparticles, 1 bottle (9.7 mL) Sirolimus Alkaline Phosphatase Conjugate, and 1 bottle (10 mL) 4-Methylumbelliferyl Phosphate, 1.2 mM; Produced by Axis-Shield Diagnostics, Ltd., Dundee, UK for Abbott Diagnostics Division, Distributed by Abbott Laboratories, Abbott Park, IL 60064 USA; list 5C91-21 |  |  | April 30, 2010 | 2010 | Abbott Laboratories | IMx Sirolimus reagent lot 802873106 has exhibited an increased frequency of calibration errors. | Other | Abbott sent a "Product Recall" letter dated April 30, 2010, to all customers receiving the affected lot of reagent. The accounts were informed of the increased frequency of calibration errors, including: 157 Check 1 Out of Range, 161 Check 5 Out of Range, and 173 Mode 1 deviation Factor High. The accounts were requested to determine if they are currently using or have inventory of lot 802873106, discontinue use of and destroy any remaining inventory of the lot, and retain a copy of the letter for their laboratory records. Customers requiring additional information were directed to contact Abbott Customer Service at 1-877-422-2688. Customers outside of the U.S., please contact your local Customer Service Representative. |  | 466 | 466 kits | Worldwide distribution: USA and countries including Puerto Rico, Argentina, Austria, Bahrain, Brazil, Canada, Chile, China, Croatia, Cyprus, Egypt, France, Germany, Greece, Hungary, India, Israel, Italy, Japan, Jordan, Lebanon, Lithuania, Netherlands, Norway, Philippines, Poland, Russia, Panama, Saudi Arabia, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, UAE and Venezuela. | Terminated | May 13, 2011 | 378 | 466 | Not_Computer
Z-2170-2010 | 56018 | bed, manual | FNJ | Manual adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom Basic Care Bed, P1441A (Non-electric), Hill Rom Batesville, IN. |  |  | September 16, 2009 | 2009 | Hill-Rom, Inc. | The flexing of the junction of the sleep deck's head and mid sections can create excess pressure on the pivot pin E-ring, which could cause the E-ring to come off. | Device Design | The firm sent an Urgent Field Safety Notice to its customers dated 9/16/2009. The letter identified the affected product, explained the reason for the recall, actions to be taken, and a contact reference. Customers are to locate the affected product, examine the pivot points and ensure the E-ring is present. If the E-ring is missing, then the affected bed is to be taken out of service. Customers' distributors or agents will contact them with a permanent fix. Questions should be directed towards Hill-Rom Technical Support. |  | 962 | 962 beds | Worldwide Distribution -- Argentina, Bahrain, Bhutan, Brazil, Bulgaria, Chili, China, Columbia, Costa Rica, Ecuador, Egypt, Eritrea, France, Hong Kong, India, Indonesia, Iran, Iraq, Jordan, Kenya, Kuwait, Lebanon, Libya, Malaysia, Mexico, Morocco, Oman, Pakistan, Palestine, Panama, Papua New Guinea, Paraguay, Peru, Philippines, Poland, Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Syria, Taiwan, Thailand, Trinidad &amp; Tobago, Turkey, United Arab Emirates, United States, Uruguay, Venezuela, and Vietnam. | Terminated | March 16, 2012 | 912 | 11056 | Not_Computer
Z-2165-2010 | 56055 | No | GWJ | Evoked response auditory stimulator. | Neurology | Neurology | 510(k) | SonaMed Clarity System II; evoked response auditory stimulator; SonaMed Corp., Waltham, MA 02154 |  |  | June 04, 2010 | 2010 | Natus Medical Inc | There is insufficient evidence to support any claims for safety and efficacy on the SonaMed Clarity devices. | PMA | Natus sent Urgent Medical Device Recall letters dated 6/4/10 to all known SonaMed Clarity device accounts, to the attention of the Director of Maternal Child Health and the Risk Manager. The letters referenced the letters previously sent to them by Natus in October 2008 and April 2010 informing them of Natus' concerns about whether the Clarity device could be lawfully used to comply with FDA-related requirements and other pertinent professional guidelines/standards for newborn infant hearing screening, and that they concluded that the Clarity devices should no longer be used. Natus recognized the need for the accounts to continue screening without interruption, and are providing on loan at no-charge handheld screener(s) to facilities that do not have alternate screening devices. The accounts were requested to contact Natus Technical Service group at 800-272-8075 to discuss this loaner device option. Natus requested that the accounts stop using the SonaMed Clarity devices; disable the screening function of their Clarity system(s) by returning to Natus, in the enclosed pre-addressed return envelope, either of the two probes (TPI-830- 8-pin ear probe, or the TPI-815- 7-pin ear probe) they may possess and have used in screening, making sure to return all 7- or 8-pin probes for all of your Clarity systems; and complete the attached "Device Recall Reply Form" and return it in the enclosed self-addressed stamped envelope. If you have any questions concerning any aspect of this recall notice, please call Natus Technical Support at 1-800-272-8075. |  | 284 | 284 units | Nationwide Distribution | Terminated | June 22, 2011 | 383 | 284 | Not_Computer
Z-2173-2010 | 55775 | No | KWH | Toe joint polymer constrained prosthesis. | Orthopedic | Orthopedic | 510(k) | Sgarlato Labs GAIT (Great Toe Implant Technique) Implant-- Large, Product Number 30-100-03 Large, Manufactured by Sgarlato Med LLC, San Jose, CA |  |  | March 01, 2009 | 2009 | Sgarlato Med, LLC | Products with an expiration date of August, 2009 were labeled with an expiration date of May 2013. Additionally, the outer packages were labeled with an incorrect lot number. The outer and inner packages did not match. | Labeling mix-ups | Recall initiated in March 2009 via telephone calls. |  | 27 | 27 units | Nationwide Distribution | Terminated | August 19, 2011 | 901 | 27 | Not_Computer
Z-2174-2010 | 55813 | No | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD GeneOhm Cdiff Assay, Catalog 441401, 200 tests/kit, labeled, in part ***BD Diagnostics, 2555 Boul. Du Parc-Technologique, Quebec, Canada G1P4S5*** |  |  | March 05, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic test kits do not perform to specifications. | Nonconforming Material/Component | The recalling firm notified accounts by an Urgent Product Recall letter on 03/05/2020. The letter identified the affected product and explained the reason for recall. It also advised customers to discontinue use of the affected product and discard any remaining inventory and fax back an acknowledgement form. The letter stated that BD will provide replacements for any discarded product. Further assistance can be obtained by contacting BD Technical Service at 1-888-436-3646. |  | 31 | 31 kits | Nationwide Distribution -- IL, IN, PA, WA, CA, VA, NJ, and KS. | Terminated | October 15, 2010 | 224 | 31 | Not_Computer
Z-2175-2010 | 55908 | No | GKH | Automated blood cell diluting apparatus. | Hematology | Hematology | 510(K) Exempt | CellTracks Auto Prep System, Circulating Tumor Cell Kit (epithelial, IVD |  |  | May 19, 2010 | 2010 | Veridex, LLC | Complaints of patient sample carryover. | Process control | On 19 May 2010, all consignees were notified, US and foreign, via Urgent Product Correction Notification letters. The letter identified the affected product and explained the reason for recall. Customers were asked to follow the required actions provided in the letter, including completing and returning the attached Confirmation of Receipt form. Questions should be directed to Customer Technical Services at 1-877-837-4339. |  | 188 | 101 Domestically, 87 Internationally | Worldwide distribution -- USA, China, Japan, England, France, Germany, Italy, Spain and Canada. | Terminated | May 08, 2014 | 1450 | 188 | Not_Computer
Z-2177-2010 | 56049 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Straumann Impression set for WN solid abutment H 4.00mm POM/ AI. Ref : 040.380 Kit Y4021. |  |  | June 15, 2010 | 2010 | Straumann Manufacturing Inc. | The label on the tubular packaging for the Straumann Impression set f. WN solid abutment has the incorrect description of RN solid abutment instead of a WN. | Labeling mix-ups | Field Safety Notice, dated June 15, 2010, issued to consignees states that the 040.380 Impression set f. WN solid abutment Lot Y4021 which you purchased is incorrectly labeled Impression set f. RN solid abutment. Customers were asked to return the affected products and completed Customer Confirmation Form for replacements. Questions should be directed towards Bernard McDonald at 978-747-2514. |  | 163 | 163 | Nationwide Distribution | Terminated | September 22, 2011 | 464 | 163 | Not_Computer
Z-2191-2010 | 49374 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | SonoSite TEE, 8-3 MHz Transducer labeled: "Tee Transducer... SonoSite, Inc. 21919 30th Drive SE Bothell, WA 98021". |  |  | August 08, 2008 | 2008 | Sonosite, Inc. | The TEE transesophageal transducer, when connected to a SonoSite M-Turbo Ultrasound System, may exhibit erratic operation of the Scanplane Orientation Indicator on the system display. The indicator may move erratically and not accurately represent the orientation of the TEE transducer array. | Other | On August 21, 2008, SonoSite began calling 15 consignees who have purchased the TEE/8-3 MHz Transducer and an M-Turbo 1.1 system. The firm provided instruction to discontinue use the system to 11 consignees and left voice messages to the remaining four consignees.On August 21, 2008, SonoSite sent out the URGENT-MEDICAL DEVICE CORRECTION letter and the Urgent Field Safety Notice to all affected consignees in the US and in foreign countries.The URGENT-MEDICAL DEVICE CORRECTION letter describes the recall product as TEE/8-5 MHz Transducer, SonoSite Part Numbers P05183 and P06841. The letter informed the consignees that the TEE transesophageal transducers may exhibit erratic operation of the Scanplane Orientation Indicator on the system display when it connected to the SonoSite M-Turbo Ultrasound System. The letter instructs consignees to discontinue use of the TEE/8-3MHz Transducer with the M-Turbo (or Micromaxx system) immediately, while the firm is working to create a software solution to resolve the issue and offers customers a software upgrade within 6-8 weeks.The consignees are also advised to forward the recall notice to others who need to be aware of the issue or where the potentially affected devices have been distributed.Customers should call the Technical Support at 1-1-425-951-1330 or 1-877-657-8118 from Monday to Friday, 6 am-5 pm PST or e-mail to service@sonosite.com for any questions regarding the recall. |  | 357 | 357 units | Worldwide distribution: USA and countries including Argentina, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Finland, France, Germany, India, Iran, Ireland, Israel, Italy, Korea, Netherlands, New Zealand, Poland, Slovenia, South Africa, Spain, Switzerland, and Thailand. | Terminated | August 24, 2011 | 1111 | 357 | Software | Display/Image | Software update | Software Update
Z-2183-2010 | 55139 | No | INK | Powered wheeled stretcher. | Physical Medicine | Physical Medicine | 510(k) | Stryker Atlas Transport Stretcher with Zoom, Model 660Z, Stryker Medical, Portage, MI |  |  | March 19, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The retractable cord stretchers were fitted with power cords, manufactured by Electri-Cord, whose prongs have the potential to fracture inside the molded section of the plug, causing melting and charring. | Component design/selection | The firm has sent Urgent - Medical Device Recall letters, dated March 19, 2010, to all direct accounts notifying them of the recall. The letter identified the affected product and described the reason for recall. Customers were asked to locate the stretchers identified in the letter and visually inspect each cord. If the plug shows signs of excessive wear, etc. the stretcher is to be taken out of service immediately and the plug removed. Customers are to call the number provided in order to be placed on the priority replacement list. If the cord does not show signs of damage then the stretcher can remain in service until a replacement cord is available. The enclosed post card should be completed and returned to the firm. A copy of the notice should be forwarded to any users that stretchers have been sold or loaned to. The firm should be provided with the serial numbers of any stretchers that have been disposed of. Questions should be directed towards Stryker at 800-STRYKER. |  | 61 | 61 stretchers | Worldwide Distribution -- USA, Canada, Brazil, Japan, Singapore, Spain, and the Far East. | Terminated | May 02, 2011 | 409 | 4915 | Not_Computer
Z-2188-2010 | 56091 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | CLAWÂż HEX SCREW CHARLOTTE " F&A; SYSTEM CONTENTS: 1 EACH Foot and Ankle Screw ... IMPLANT MATERIAL: Stainless Steel Ref: 4013-3520 ... Size 3.5MM, Length 20 MM, Use with ClavaÂż 3.5MM - Non-Sterile. Implantation: used in Food and Ankle surgeries. |  |  | June 07, 2010 | 2010 | Wright Medical Technology Inc | One lot of CLAW Hex Foot and Ankle Screws was incorrectly packaged containing Multi-use Compression Screws. | Packaging process control | The firm initiated their recall to their sales representatives (distributors) via telephone on 06/07/2010 and followed with a letter dated June 15, 2010 sent via Fed Ex on 06/28/2010. The hospital involved was also notified by letter on 06/29/2010. The letter described the product, the problem,and the action to be taken by the customers. The customers were instructed to immediately return any affected product to Wright as soon as possible, and to complete an attached response form confirming receipt of the notice and return a copy by fax to 901-867-7401. Customers may contact customer service for return instructions and replacement inventory at 800-238-7117.If you have any questions regarding this matter, contact Debby Daurer at 800-874-5630. |  | 8 | 8 units | Nationwide Distribution in States of FL, PA, KS, IN, MO, and CA | Terminated | November 23, 2011 | 534 | 8 | Not_Computer
Z-2190-2010 | 56115 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | Spectranetics Corporation, QuickCat Extraction Catheter, REF 60090-01, Sterile R, Working Length 145cm, for embolus and thrombus removal from arterial vessels. |  |  | July 01, 2010 | 2010 | Spectranetics Corporation | Guidewire lumen may be occluded, preventing guidewire insertion into the catheter. | Process control | Consignees were notified by letter on Jul 1, 2010, and told to not use affected lots and to contact Spectranetics to arrange a product exchange. More information is available at 800-231-0978. |  | 5080 | 5080 units (1778 US, 3302 Non-US) | Worldwide distribution: USA, Argentina, Austria, Belgium, Brazil, Chile, Czech Republic, Finland, France, French Polynesia, Germany, Hungary, Israel, Italy, Netherlands, Poland, Spain, Sweden, Switzerland, and Venezuela. | Terminated | August 20, 2010 | 50 | 5080 | Not_Computer
Z-2192-2010 | 56163 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | REF 87100 87k Arthroscopy Tubing Set. For use with the 87k Arthroscopy Pump (87000). STERILE EO. Rx ONLY. Warning: Do not use this tubing set with any Small Joint Cannula System. |  |  | July 06, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec Largo, FL is recalling 87K Arthroscopy Tubing Set, Product Number 87100, Lots 0904151 to and including 1005201. Product may have a breach in the seal that could potentially compromise the sterility of the contents. | Process control | ConMed Linvatec Corporation sent an "URGENT-MEDICAL DEVICE RECALL NOTIFICATION" letter dated July 6, 2010, via FedEx next day delivery to all customers. The letter described the product, problem and action to be taken by customers within 30 days of receipt of the letter. The customers were asked to PLEASE DISCONTINUE USE OF THE IDENTIFIED PRODUCT WITH LOT NUMBERS BEGINNING WITH 0904151 THROUGH AND INCLUDING 1005201; to immediately check their inventory for the affected lots number and Do Not Use if they have any of the products; to complete and return a Reply Form and to notify their customers if the product was further distributed and return all unused affected product with lot codes starting 0904151 to and including 1005201, to ConMed Corporation, 525 French Road, Utica, NY 13502. On receipt, ConMed Linvatec will issue credit to customers.Should you have any additional questions, please contact our Customer Service department at 800-535-8536 or fax at 727-319-5701 or email: Custerserv1@linvatec.com |  | 97830 | 97,830 | Worldwide distribution: USA and Canada. | Terminated | February 24, 2011 | 233 | 97830 | Not_Computer
Z-2193-2010 | 56321 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Crossfire Console, model number 475-000-000, Manufactured by Stryker Endoscopy San Jose |  |  | June 28, 2010 | 2010 | Stryker Endoscopy | RF energy or powered shaver handpieces may continue to operate when not intended when used with the iSwitch Wireless Foot switch. | Software design | Stryker Endoscopy sent an "Urgent: Device Correction" letter dated June 28, 2010, via FedEX Mailer to US customers. A letter dated July 9, 2010, was sent to International sites.The letter describes the product, problem, and action to be taken by customers. The Customers were instructed to follow the included instructions and use the USB dongle to install the correction and to complete and sign the enclosed "acknowledgement of receipt" form,upon completion of the software correction, and fax to (480) 754-8378 or scan in and email to crossfireconsole@stryker.com. The firm noted that-It is important to send the self addressed confirmation forms back to Stryker RAQA. If you have any questions regarding this letter, please contact us by phone at 1-800-624-4422 or via email at crossfireconsole@stryker.com. |  | 989 | 989 units | Worldwide distribution: USA and countries including Australia, France, Germany, Greece, South Africa, Iberia, Italy, New Zealand, Poland, UK, Switzerland, Canada, China, Japan and the EMEA. | Terminated | March 21, 2011 | 266 | 989 | Software | Physical Safety Hazards | Software update | Software Update
Z-2194-2010 | 48998 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | The Philips EasyDiagnost Eleva with DRF. Philips Medical Systems multifunctional x-ray system. |  |  | July 09, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Potential for bright artifacts to appear on the patient image from a previous exposure. | Other | On July 9, 2008, Philip Medical System began sending out the URGENT-Device Correction letter to their consignees. This letter describes the recalled product as EasyDiagnost Eleva with DRF system with SW release 1.0.The firm informs the consignees that under special circumstances there is a small possibility that bright artifacts appear on a patient image from a previous exposure.The consignees are advised to take care about the X-ray field limitation. If the detector is not affected by direct radiation, no problem will occur. If direct radiation reaches the detector, the collimation should not be too tight, such that the area of direct radiation is larger than 1.2 inch x 1.2 inch (or 30 mm x 30 mm). The minimum area is needed to detect and prevent memory artifacts.The firm will upgrade any affected products with the Field Change Order (FCO)70600031 and it will be implemented before the end of December 2008. The upgrade software is 1.0.3. The correction would be free of charge and the consignees should contact Philips Call Center at 800-772-9377, #5, #2. |  | 7 | 7 | Product was distributed in AZ, CA, FL and WA. | Terminated | October 27, 2010 | 840 | 7 | Software | Display/Image | Software update | Software Update
Z-2195-2010 | 55787 | No | JQX | Nephelometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMAGE Immunochemistry Systems Buffer 1 (BUF1), Part Number: 447650 |  |  | May 21, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall of IMMAGE Immunochemistry Systems Buffer 1 (BUF1) due to received reports of increased recovery of control andpatient samples using the lots of IMMAGE BUF1. Some lots of BUF1 have been reported to cause high shifts in recovery of control or patient samples. Quality Control failures may result after replacing BUF1 with a different lot. | Nonconforming Material/Component | Beckman Coulter will send a Product Corrective Action (PCA) letter on the week of April 19, 2010 to the affected customers. The letter provided the customers with an explanation of the problem and a work around. |  | 1309 | 1309 | Worldwide: U.S., Canada, Australia, Austria, Belgium, China, Czech Republic, Denmark, France,Germany,Greece,Hong Kong, Iceland,Ireland,Italy,Japan,Liechtenstein,Macao,Malaysia,Netherlands,New Zealand,Norway,Oman,Poland,Qatar,Russian Federation,Singapore,Slovakia,South Africa,Spain,Sweden,Switzerland,Taiwan,Tunisia,Turkey andUnited Kingdom, | Terminated | May 22, 2012 | 732 | 1309 | Not_Computer
Z-2202-2010 | 56092 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Intellivue Patient Monitor Model MP70 (M8007A) |  |  | June 28, 2010 | 2010 | Philips Healthcare Inc. | Additional Instruction for Use for Intellivue MP60/MP70 to prevent monitors from becoming loose and detaching from the mounting arm. | Labeling design | Philips issued on 6/28/10 an Urgent Medical Device Correction notification/Field Safety Notice. Included with the Urgent Medical Device Correction notification/Field Safety Notice is an updated instructions for use and service instructions for these monitors that remind users to follow themounting arm manufacturer's instructions for use when repositioning the monitor and instructing them toperform a periodic check that the monitor is not loose on the arm. |  | 82019 | 82,019 units | Worldwide distribution: USA, Canada, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, Bosnia, BRAZIL, BULGARIA, CHILE, CHINA, COLOMBIA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRI, LITHUANIA, LUXEMBOURG, MACAU, MALAYSIA, MAURITIUS, MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA &amp; MONTE NEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UKRAINE, United Arab Emirates, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, VIRGIN ISLANDS (BR), and YEMEN. | Completed |  |  | 93548 | Not_Computer
Z-2204-2010 | 55909 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Vanguard CR femoral 57.5 mm left interlok, Sterile, REF 183022, Knee prosthesis component. |  |  | April 20, 2010 | 2010 | Biomet, Inc. | A knee implant labeled as left was opened and the package actually contained a right knee implant. | Process control | On 4/13, 14 and 15 Biomet made phone calls to the consignees to inform them of the recall. The firm sent Urgent Medical Device Recall Notices dated 4/20/2010. The customers were instructed to immediately locate the recalled products and return them to the manufacturer. |  | 5 | 5 | 1 unit in US. 3 units in Biomet distributor inventory, 1 unit in Japan | Terminated | January 19, 2011 | 274 | 5 | Not_Computer
Z-2200-2010 | 56117 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | International Technidyne Corporation (ITC) directCheck Whole Blood Quality Controls Plasma, Coagulation Control. Catalog number: DCJLR-N; in vitro diagnostic |  |  | June 25, 2010 | 2010 | International Technidyne Corp. | The package insert for the directCheck Microcoagulation Quality Control contains an incorrect acceptable performance range regarding ACT-LR cuvettes | Nonconforming Material/Component | ITC sent Urgent Medical Device Correction letters on June 25, 2010 by Federal Express to hospital Risk Managers and Lab Directors, Clinical Laboratory Manager or Point-of-Care Coordinators. |  | 122 | 122 boxes in US, 1 out of US. | Nationwide and one customer in Belgium. | Terminated | September 23, 2010 | 90 | 122 | Not_Computer
Z-2196-2010 | 56157 | No | NMF | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | ASCENT HEALTHCARE SOLUTIONS. Reprocessed Femoral Compression Device, Model 11165. |  |  | June 23, 2010 | 2010 | Ascent Healthcare Solutions, Inc. | The FemoStop Devices may fail to inflate or hold pressure. A separation between the dome and arch base prevents the device from maintaining pressure when inflated. | Reprocessing Controls | Ascent sent an urgent medical device voluntary recall notification Letter and Response Form to each consignee on June 25, 2010. Consignees were requested to immediately examine their inventory of Ascent Reprocessed FemoStop devices with lot numbers located on an attached list; remove and quarantine all affected devices; complete the attached Recall Effectiveness Check Form; ship all affected product back to Ascent using the address indicated in the letter, and fax a completed Recall Effectiveness Check Form to 480-763-5350. |  | 175 | 175 devices | Devices were distributed to 8 consignees (end users/hospitals) located in VA, CA, FL, MI, WA. | Terminated | August 11, 2010 | 49 | 175 | Not_Computer
Z-2197-2010 | 56187 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | CERCLAGE WIRE, 0.6 MM DIA. X 175 MM LENGTH, SST, REF 71163006, QTY: (6), STERILE R, CE 008, 2020/07, Smith &amp; Nephew, Inc., Memphis, TN 38116, EC REP: Smith &amp; Nephew Orthopaedics GmbH, Tuttlingen, Germany |  |  | June 28, 2010 | 2010 | Smith &amp; Nephew Inc | Plastic trays which encase these products have burrs on the hinged packaging which can puncture the single barrier sterile pouch, compromising sterility. | Process control | All affected Smith &amp; Nephew Sales Representatives were notified of problem and the recall was sent via overnight mail on 01/23/2009. All affected Smith &amp; Nephew International Distributors were notified via email and telephone on 01/23/2009. They were instructed to immediately quarantine the product for return. |  | 1 | 1 unit | KY, TX, DC, NC, CA, AZ, GA and OH; 3 International: Australia, Malaysia, and Switzerland | Terminated | January 23, 2012 | 574 | 43 | Not_Computer
Z-2205-2010 | 56201 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Pulmonetic Systems P/N 19189-001 Patient Circuit SPU W/O PEEP 15 FT 5/Pkg. |  |  | July 02, 2010 | 2010 | CareFusion 203, Inc. | CareFusion is recalling the P/N 19189-001 Patient Circuit (SPU W/O PEEP 15ft) because some of the circuit sense lines may have been reversed during the manufacturing process. If the sense lines have been reversed, immediately upon connection of the patient circuit to the ventilator and turning the ventilator on, the ventilator will begin to autocycle (with increased pressure delivered) with both | Process control | Customers were sent a CareFusion "Urgent Medical Device Recall Notification" dated July 2, 1010. The letter was addressed to Respiratory Care Manger. The letter described the product , problem and advised consignees to remove the affected product from inventory and promptly return it to CareFusion for replacement.Please contact CareFusion Technical Support at 800-754-1914, option 2, upon receipt of this notification to arrange for prompt return and replacement of any the patient circuits in your inventory. |  | 200 | 200 boxes (5 each) | Nationwide distribution: AL, CA, FL, GA, HI, IL, IN, KS, MA,MI, MN, MT, NJ, NY, NC, OR, PA, RI, TN, TX, UT, VA, WA, and WI. | Terminated | December 17, 2011 | 533 | 200 | Hardware | Treatment/Delivery/Therapy | Return | Remove or Replace
Z-2210-2010 | 48992 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Motorized Height Ceiling Suspensions (MCS) and motorized height movement (actuator) for X-ray systems. |  |  | June 26, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Monitor ceiling carriage may separate from the monitor ceiling suspension. | Other | On 07/02/08, the firm sent a "Urgent - Device Correction" letter to all their consignees. The letter describes a product as X-Ray system with Monitor Ceiling Suspension and motorized height movement. The consignees are informed that the fixation bolts supporting the monitor ceiling carriage may come loose creating a situation which may allow the monitor ceiling carriage to separate from motorized height Monitor Ceiling Suspension and that would cause serious injury to patient and user.The firm has issued a mandatory Field Correction Action to correct the problem and the firm's technicians will inspect all ceiling suspension actuators for loose screws. If loose screws are found the actuator will be replaced.Consignees should call Philips Call Center at 800-722-9377 #5, #4, #1, and Reference FCO72200126. |  | 1111 | 1111 units | United States and Canada. | Terminated | December 02, 2010 | 889 | 1111 | Not_Computer
Z-2206-2010 | 54370 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | BIOMET MICROFIXATION "1.5MM" SYSTEM HIGH TORQUE (HT) CROSS-DRIVE SCREW, QTY 1, 1.5 X 4MM TITANIUM. REF 91-1504 1 COUNT and REF 95-1504 - 5 PACK. |  |  | March 17, 2010 | 2010 | Biomet Microfixation, Inc. | Biomet Microfixation has received a complaint from a customer who reports finding a 3.5mm length screw in one 5 pack of 4mm length screws. | Other | Biomet Microfixation is recalling their 1.5x4mm HT X-drive Screws Lot Number 148530 after receiving a customer reported finding one 3.5MM screw in a package labeled as containing 4.0MM screws. Lot 148530 was packaged into one count packages 91-1504 and 5 count packages 95-1504. The 1 and 5 count packages were distributed between 12/10/2009 and 02/01/2010. Beginning on 03/17/2010 customers identified as receiving Lot 148530 HT X-drive Screws were notified of the recall by a combination of an Urgent Medical Device Recall Notice letter, e-mail, fax and telephone calls. Customers were instructed to inspect their inventory and to return any affected products. They were instructed to complete the attached Inventory Reconciliation sheet and to fax it to Biomet Microfixation. Customers were also instructed to contact any customers to whom they may have further distributed the product. Questions should be directed to Christy Cain at 1-800-874-7711 or 904-741-4400, ext. 468. |  | 1247 | 1247 screws | Worldwide Distribution -- USA, including states of FL, GA, IN, LA, MD, MI, MO, NC, NY, OK, TX, and VA and countries of Philippines, Japan, Brazil, and Canada. | Terminated | August 11, 2010 | 147 | 1247 | Not_Computer
Z-2208-2010 | 55678 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems AFP QC, Part Number: 33219 |  |  | April 21, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed an issue with the Access AFP QC kits (REF 33219) which contain three levels of control material identified as QC1, QC2, and QC3, Beckman Coulter has confirmed that vials of QC1 in the kit lots identified above contain microbial contamination. Customers may see that the results from a contaminated vial of QC1 are within the established refere | Other | A Product Corrective Action (PCA) letter was senton the week of April 19, 2010 to all customers whopurchased the Access Immunoassay Systems AFPQC material, lot numbers referenced above. Theletter provided the customers with an explanation ofthe problem identified and instructed the customersto (1) Please discard all remaining inventory of the three Access AFP QC kit lots identified above,including all vials of QC1, QC2, and QC3.(2) Please review your ac and patient results obtained while using the affected control lots. Interpret Access AFP test results in light of the clinical presentation of the patient, including: symptoms,clinical history, data from other tests and other appropriate information.(3) Please share this information with your laboratory staff and retain this notification as part of yourlaboratory Quality System documentation. (4) Complete and return the enclosed response formwithin 10 days so we can be assured that you have received this important notification.For customers that need assistance or have any questions regarding this notification, they were instructed to contact Technical Support at 1-800-854-3633 in the United States and Canada. |  | 166 | 166 | Nationwide and Canada | Terminated | May 07, 2012 | 747 | 166 | Not_Computer
Z-2207-2010 | 56031 | No | NQK |  |  | General &amp; Plastic Surgery | 510(k) | Monarch Labs Medical Maggots with the accessory LeFlap dressing (previously known as Creature Comforts II), 250-500 per vial, Batch #: MM-090406 with CCII-0901. |  |  | May 03, 2009 | 2009 | Monarch Labs, LLC | Monarch Labs initiated the field notification regarding their Medical Maggots from Batch MM090406 with the accessory LeFlap dressing due a customer complaint from a physician of a patient developing cellulitis after using maggots from the affected batch. | Other | The customer recall communication was initiated on May 3, 2009 with Monarch Labs notifying all consignees with a via a faxed Recall Letter with fax-back attachment of the initial compliant they had received and the possible infection that may occur if instructions for use are not followed. All therapists, who used maggots from that batch were instructed to contact the firm immediately if they have experienced any adverse events, such as infection, cellulitis, or any other problem. If not, they were asked to simply return the fax-back attachment , noting "no problem" . |  | 14 | 14 vials | Nationwide Distribution -- CA, IA, KY, IL, TX, SC &amp; UT. | Terminated | December 13, 2012 | 1320 | 14 | Not_Computer
Z-2209-2010 | 56118 | Yes | NVZ |  |  | Cardiovascular | PMA | Medtronic Carelink Monitor Model 2490C. The Medtronic CareLink Monitor Model 2490C is designed to automatically gather information from the implanted heart device. This automatic wireless communication between the heart device and the monitor takes place at times scheduled by the doctor or clinic. The gathered heart device information is sent over a telephone line to the clinic using a toll-free number. The monitor can also send wireless heart device alerts to the clinic automatically. |  |  | June 11, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic has identified that a subset of 2490C CareLink Monitors recently received an incorrect software update. Patients reported that their monitor was making a "ticking sound" and they were unable to transmit device data via CareLink. These monitors are now non-functional. No patient injuries have been reported as a result of this issue. | Software change control | Starting on 6/11/2010, Medtronic began sending Australian and Canadian consignees "Medical Device Recall" letters dated 6/10/2010. The letters describe the problem and product involved. The letters instructs consignees to order replacement monitors, notify affected patients, and instruct each patient to perform a manual transmission upon receipt of the new monitor. Medtronic will provide affected patients with a prepaid return kit to faciliate the return of the monitor to Medtronic.Australian consignees should contact Scott Dunlop at 61-7 3025-3547. Canadian consignees should contact Medtronic Patient Services at 1-800-268-5346. |  | 37 | 37 Foreign, 0 Domestic | International Distribution Only: Australia and Canada. | Terminated | December 24, 2011 | 561 | 37 | Software | Device Operation | Replace monitor | Remove or Replace
Z-2212-2010 | 56131 | No | JDL | Hip joint metal/metal semi-constrained, with a cemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | DYNASTYÂż A-CLASSÂż POLY LINER, REF DLXP-LD36, 1 EACH, Rx ONLY, STERILE EO, GROUP D, I.D. 36 mm, LIP 15o, LINER STD, Wright Medical Technology, 4677 Airline Road, Arlington, TN 38002. Product is used in total hip arthroplasty. |  |  | June 07, 2010 | 2010 | Wright Medical Technology Inc | The inner and outer product packaging was missing the following information: shelf life, translations, manufacturing date, CE marking and also listed the incorrect sterilization method. | Software design (manufacturing process) | On 06/28/2010, Wright Medical Technology sent a letter to consignees. Consignees were instructed to examine their inventory for recalled product and return it to the firm. Consignees should contact customer service at 800-238-7117 for return instructions and replacement inventory. Questions concerning the recall should be directed to Debby Daurer at 800-874-5630. |  | 6 | 6 units | USA distirbution only, in states of WI, WV, IN, IA, and GA | Terminated | May 31, 2011 | 358 | 6 | Not_Computer
Z-2216-2010 | 55364 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trabecular Metal acetabular revision system column buttress left posterior/right anterior, 00-4894-300-00, sterile, Zimmer Inc., Parsippany, NJ. |  |  | April 06, 2010 | 2010 | Zimmer Inc. | The firm has determined that some of the packages could develop a breach of the seal in the sterile barrier pouches. | Packaging | The firm sent an "Urgent Device Removal" letter dated 4/6/2010. The consignees were instructed to inventory the products on hand and return the recalled product to the manufacturer. More information is available at 1-800-613-6131. |  | 1972 | 1,972 all products | US, Germany, Brazil, Singapore, Canada, Australia, England, India, and Switzerland. | Terminated | October 21, 2011 | 563 | 1972 | Not_Computer
Z-2211-2010 | 55502 | Yes | MSX | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | DATEX-OHMEDA NETWORK AND ICENTRAL '05, SALES REVISION 4.3 ( Datek-Ohmeda Network and iCentral (including iCentral Client)) The Datex-Ohmeda Network and iCentral (including iCentral Client) transfers information between networked Datex-Ohmeda devices in the Monitor Network as well as between Monitor Network and networked devices in GE Unity NetworkTM. It also allows information transfer between several iCentrals. Within one Monitor Network it allows a networked device to display, store, print and otherwise process information received from other networked devices. The iCentral maintains the network connections between the Datex-Ohmeda bedside monitors and other networked devices in Monitor Network. Network connections consist of hardwired network cables and/or Wireless LAN (WLAN) connections. Furthermore, it coordinates the transfer of information between devices in the Datex-Ohmeda Network, between the Datex-Ohmeda Network and devices in GE Unity Network as well as between the Datex-Ohmeda Network and Hospital Information Systems (HIS). The iCentral can be used for remote monitor management, storing, printing, viewing, reviewing or otherwise processing of information from several bedside monitors or other networked devices. The iCentral Client can be used for remote monitor management, printing, viewing, reviewing or otherwise processing of information from several bedside monitors or other networked devices. The Datex-Ohmeda Network will be used for patients in the hospital and it is meant for continuous use. The device is for use by qualified personnel only. |  |  | June 25, 2010 | 2010 | Ge Healthcare, Llc | The iCentral or iCentral client v5.0.2 with HP xw 4400 and 4600 may spontaneously shutdown, causing monitoring loss at the central monitoring station. Monitoring loss may result in prolonged delay of treatment of potentially lethal arrhythmia. | Software design | GE Healthcare sent a "Urgent Medical Device Correction letter" dated 6/18/10 on 6/25/10 to consignees. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. The letter instructs consignees to follow specific detailed instructions for safe use of the device pending provision of a software upgrade by the firm. Consignees will be contacted by a GE Healthcare representative to schedule the software upgrade. |  | 2962 | 2962 | AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, IA, KS, KY, LA, ME, MI, MN, MS, MO, NE, NH, NH, NY, NC, ND, OH, OK PA, RI, SC, TN, TX VT, VA, WA, WI.ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BARBADOS, BELGIUM, BERMUDA, BOLIVIA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, GUYANA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA , KUWAIT, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, NAMIBIA, NETHERLANDS, NEW ZEALAND,NICARAGUA, NIGERIA, NORWAY, PHILIPPINES, POLAND, PORTUGAL QATAR, RUSSIAN FEDERATION SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN THAILAND TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA, YEMEN. | Terminated | August 14, 2012 | 781 | 2962 | Software | Device Operation | Software update | Software Update
Z-2213-2010 | 56189 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | DSD Application Guide (MG04-0001 revision M) for the Pentax Universal Video Scope Hookup (DSD-110-HU0109) only when used with the DSD automated endoscope reprocessor (AER) and the Pentax EG-3630U endoscope. |  |  | June 15, 2010 | 2010 | Minntech Corp | The DSD hook up Application Guide incorrectly recommends use of the Medivators DSD-110-HU0109 endoscope hookup to connect the Pentax EG-3630U Ultrasound Endoscope to a Medivators DSD endoscope reprocessor or Medivators Scope Buddy Endoscope Flushing Aid. This may result in inadequate disinfection during reprocessing, and subsequent infection risk to the patient. The recall was due to the misprinte | Labeling False and Misleading | Consignees were sent a Medivators Reprocessing Systems "Urgent Medical Device Recall" letter date June 17, 2010. The letter was addressed to "Dear Medivators Customers". The letter described the problem and product involved. Recommended consignees to immediately examine the Pentax EG-3630U Ultrasound Endoscope hookups and inform all staff of the notice to ensure that the DSD-11--HU01109 is not used with the Pentax endoscope.Also requested consignees to complete and return the enclosed response form. |  | 185 | 185 | AZ, AR, CA, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, MD, MA, MI, MN, MS, MO, NH, NJ, NY, NC, ND, OH, PA, SC, TN, TX, UT, VA, WA, WV, WI, WY. | Terminated | December 17, 2011 | 550 | 185 | Not_Computer
Z-2219-2010 | 56166 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Hi-Art System, H-0000-0003 |  |  | December 01, 2009 | 2009 | TomoTherapy Incorporated | An issue was identified with the TomoTherapy HI-Art System. In the event a patient or DQA plan has a moved image, roll adjustments applied during registration will be incorrect. The Planning Station Plan Settings and DQA Setup tabs allow for images to be moved during planning. During registration when roll is applied on moved images, the Operator Station incorrectly rolls the image about | Software design | Consignee was sent a TomoTherapy "Urgent Medical device Correction Field Safety Notice' dated December 01, 2009. The letter described the Issue, Product Affected, Recommended Action and Resolution. |  | 1 | 1 | VA | Terminated | December 17, 2011 | 746 | 1 | Software
Z-2180-2010 | 56222 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical, Engage TR Introducer, 6 F - ACT (2.25 mm), 25 cm length, .035" Max Guidewire O.D., Rx, Sterile EO, REF C408513, 100017549. |  |  | June 28, 2010 | 2010 | St. Jude Medical Cardiovascular Division | Potential for a partial or complete separation of the Shaft (sheath) from the hub or a material break in the hub assembly just distal to the hemostasis seal. If either of these were to occur during use, it is likely that fluids would leak around the Introducer hub and strain relief. | Process design | Consignees were sent "Urgent Medical Device Recall Notice" dated June 24, 2010 from St. Jude Medical . The letter was addressed to Cath Lab Manger/ Risk Manager. The letter described the problem and the product involved. Advised consignees to discontinue use of the product and remove from inventory. Sales Representatives were also sent a St. Jude Medical Memorandum "Field Action Notice" dated June 24, 2010 and a reconciliation form. The memorandum described the problem and product. Advised Sales Rep to visit the affected accounts and to performed the assigned tasks.A St. Jude Medical sales representative will be contacting for removal and replacement of the affected product. For additional information, please contact St. Jude Medical at 651-491-4470. |  | 15 | 5 US, 10 OUS | Worldwide Distribution including USA, Italy, Denmark, France, Switzerland, England, Finland, Austria, Spain and Estonia. | Terminated | December 27, 2011 | 547 | 5120 | Not_Computer
Z-2220-2010 | 56246 | No | KYG | Ocular surgery irrigation device. | Ophthalmic | Ophthalmic | 510(K) Exempt | CORNEA COAT Hydroxypropylmethylcellulose 1.5ml Syringe. STERILE, Rx ONLY. |  |  | May 12, 2010 | 2010 | Insight Instruments, Inc. | Insight Instruments, Inc., Stuart, FL is recalling CORNEA COAT 2.5% Hydroxypropylmethylcellulose solution in 1.5 ml Syringe due to a potential weakness of the peel pouch seals. Products affected are Item Numbers 600100.000 sold individually and 40201.000 thru 10206.000 sold as components of BIOM Pack. | Packaging | All consignees were notified of the recall by letter sent via first class mail. Letter contained a Return Response Form. |  | 16128 | 16,128 | Worldwide Distribution -- USA, including Puerto Rico, and Frankfurt, Germany. | Terminated | February 16, 2012 | 645 | 66853 | Not_Computer
Z-1634-2010 | 55077 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCIÂż AQUA+Âż Flex, Hygroscopic Condenser Humidifier, HCH Tidal Volume Range 300-1500mL, ICU/Anesthesia Adult/Child, Catalog number 1570, Distributed by Teleflex Medical. The Aqua+ Flex is a passive humidifier indicated for use on patients under anesthesia and intensive care to effectively warm and humidify the inspired air. |  |  | March 19, 2010 | 2010 | Teleflex Medical | The patient end of the 22 cm connector on the flex tube may not fit securely within the endotracheal tube (ET) connector and therefore disconnect from the endotracheal tube. A disconnect in ventilator dependent patients without prompt response to the alarm could lead to hypoxia, organ failure, or cardio-respiratory arrest. | Device Design | Teleflex Medical sent an Urgent Medical Device Recall letter on March 22, 2010 to consignees. They were instructed to: 1) Immediately discontinue use and quarantine any products labeled as AQUA+ FLEX, Catalog number 1570 with the lot numbers provided. 2) Complete and fax the Recall Acknowledgment Form to 866-804-9881. 3) After contact by a Customer Service Representative, who will provide them with a Return Goods Authorization Number, return the product to Teleflex Medical for credit. Letters addressed to distributors instructed them to communicate the recall to any of their customers who have received the product using the customer letter template and acknowledgement form provided by Teleflex Medical. Their customers were to return the product and acknowledgment form to the distributor who would consolidate and return the product to Teleflex Medical. Affected products received by Teleflex Medial will be destroyed upon receipt. For additional instructions or information, contact the firm at 1-866-246-6990. |  | 3841 | 3841 units | Worldwide distribution in the states of CA, FL, GA, KY, KS, IL, MA, MD, MI, MN, NC, NJ, NM, NY, OH, OK, OR, PA, RI, TN, TX, VA, and WI and to the country of Japan. | Terminated | November 28, 2011 | 619 | 3841 | Not_Computer
Z-2222-2010 | 56256 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | SIEMENS AXIOM Luminos TF and Sireskop SD Solid State x-ray imager (flat panel/digital imager), Fluoroscopic x-ray diagnostic system. |  |  | July 13, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Fluid may enter the multi-functional handle (Opti Grip) when using the Axiom Luminos TF or Sireskop SD system, causing unintended reaction or movement, increased radiation or useless images. | Device Design | Siemens Healthcare Sector sent an "Urgent Field Safety Notice" to consignees describing the issue and informing them to take specific precautions when performing interventional procedures to avoid fluids penetrating the system. The multi-functional handle (Opti Grip) will be replaced through implementation of Update Instructions AX030/10/S. The modified version of the handle resists fluids from entering the handle and prevents malfunction. |  | 1296 | 1296 | Nationwide distribution: all 50 US states and PR. | Terminated | September 29, 2012 | 809 | 1296 | Not_Computer
Z-2223-2010 | 55804 | No | DSA | Patient transducer and electrode cable (including connector). | Cardiovascular | Cardiovascular | 510(k) | Disposable Surgical Cable, Product Code: FL-601-97 (100), Single use, Rx Only, Remington Medical, Inc., 6830 Meadowridge Court, Alpharetta, GA 30005 |  |  | May 19, 2010 | 2010 | Remington Medical Inc. | The Safe Connect component may separate from the cable, causing the cable to become unplugged from the EPG. | Process design | Consignees were notified on/about 05/19/2010. They were instructed to examine their inventory and quarantine product subject to the recall. If they distributed the product, they were advised to identify their customers and notify them of the recall (via the recall letter). A Return Authorization was provided. More information can be obtained by calling 770.888.8520. |  | 1488 | 1488 units | AZ, CA, CO, CT, DE, FL, GA, IL, IN, LA, MD, MI, MO, MS, NC, NM, NY, OH, OK, OR, PA, SC, TX, TN, TX, VA, WA, WI and Canada. | Terminated | November 29, 2010 | 194 | 1488 | Not_Computer
Z-2176-2010 | 56290 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | Ikaria INOMAX DS Drug Delivery System, model 10003. Ikaria 6 Route 173, Clinton, NJ 08809 USA. |  |  | July 21, 2010 | 2010 | Ikaria Holdings | The pressure switch in the INOMAX DS drug-delivery system may fail, which may interrupt or delay the administration of INOMAX (nitric oxide) for inhalation to patients. Delay in administration of INOMAX therapy may cause: -worsening of systemic oxygenation indices (i.e., hypoxemia, especially manifested as decreased arterial oxygenation saturation). -hypotension, and/or -increase in pulm | Component design/selection | Ikaria sent an Urgent Medical Device Recall letter to customers dated July 21, 2010. This letter is also posted on the Inomax webpage. The letter described the issue, how to recognize it, what to do in the event of a suspected leak, and recommended maintaining back-up cylinders and depressurization of the regulator when not in use. The replacement plan for all INOMAX DS drug-delivery systems was described. Customers can contact Ikaria Customer Care at 1-877-566-9466. |  | 2507 | 2507 | Worldwide Distribution to include: US Nationwide and Countries of Canada and Malaysia | Terminated | December 27, 2011 | 524 | 2507 | I/O | Treatment/Delivery/Therapy | Replace device | Remove or Replace
Z-2227-2010 | 54546 | Yes | FKP | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Main Power Supply Assembly (50 and 60 Hz) and AC Power cord assembly for hemodialysis machines and Granuflo 1 mixers, Distributed by Fresenius USA Inc., Walnut Creek, CA |  |  | November 20, 2009 | 2009 | Fresenius Medical Care North America | Power cords may fail if the prongs crack or fail at the plug, and may become a fire/burn/shock hazard or delay/interrupt therapy. | Component design/selection | An urgent device recall notification was sent to all consignees. Customers are asked to inspect cords for defects, and are refered to preventive maintenance requirements in the manual. Customers who identify the defective cords will be instructed to discontinue use as soon as possible without jeopardizing patient care. Customers are instructed to monitor cords regularly, and pay attention to machines that are frequently plugged or unplugged. Personnel are instructed to not use the cord to unplug the device from the wall (yanking the plug from the wall) and to refrain from rolling over the cord or plug. |  | 10267 | 10267 | Product distributed to 5404 customers throughout the US and the world. | Terminated | February 07, 2011 | 444 | 72142 | Not_Computer
Z-2229-2010 | 56193 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SpeedPass Disposable Suture Lariat with Nitinol Wire 25 Degree Left, QTY. 1, Biomet Sports Medicine, Warsaw, IN 46581. Sterile. Part number 904010. |  |  | June 18, 2010 | 2010 | Biomet, Inc. | The nitinol wire may be fractured or may have the potential to fracture making the suture lariat unusable and possibly delaying surgery. | Component design/selection | The firm notified all 29 consignees by An Urgent Medical Device Recall Notice letter on 6/18/2010. The letter identified the affected product and explained the reason for recall. Customers are asked to locate and remove the affected products, follow the instructions on the enclosed "FAX Back Response Form," and fax a copy of the form back prior to return of the affected product. Additionally, if product has been further distributed, then hospital personnel must be notified and given the "Dear Biomet Customer" letter. Questions should be directed to Mary Johnson at 800-348-9500 or 574-372-3983. |  | 301 | 301 units total | Worldwide Distribution -- USA, including the states of FL, IL, NY, CO, MD, CA, GA, AL, PA, OH, SD, KY, TX, ID, WA, NC, AK, &amp; MN, and Columbia and The Netherlands. | Terminated | August 06, 2012 | 780 | 301 | Not_Computer
Z-2233-2010 | 56262 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis MP/Artis dMP and Artis zee Multipurpose. Angiographic x-ray system. |  |  | July 13, 2010 | 2010 | Siemens Medical Solutions USA, Inc | The Axiom Artis MP, Artis dMP or Artis zee Multipurpose C-arm gearbox can become damaged. If the gearbox is damaged, the C-arm may tilt away when being repositioned and collide with nearby objects in its range of motion. | Other | Siemens has issued an "Urgent Field Safety Notice: Customer Safety Information to consignees with Update Instructions AX033/10/S. This letter informed customers of the issue, and provides an Addendum to the User Manual. The Update Instructions describe the process for installation of an additional clamping mechanism provided by Siemens into the existing system, which will be activated if the gearbox is damaged. If activated, the clamping machanism will prevent the C-arm from tilting away, stop regular operation, and alert the operator to check the system and repair the gearbox damage. |  | 131 | 131 units | Nationwide distribution: States of AL, AZ, CA, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, MT, NC, NE, NJ, NY, OH, OK, PA, PR, RI, SC, TN, TX, UT, VA, WA, WI, and WY. | Terminated | July 05, 2012 | 723 | 131 | Not_Computer
Z-2235-2010 | 56225 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | medCOMP(R) MRCTI66841 6.6F Plastic Dignity Low Profile CT Port with Pre-Attached Chronoflex(R) Polyurethane Catheter - 1 Unit. Catalog #MRCTI66841. |  |  | July 02, 2010 | 2010 | Medical Components, Inc dba MedComp | Recalled product was packaged with the incorrect port and does not match the label. The kit is labeled as a 6.6F "Dignity" Low Profile CT Port. The kit contains a 6.6F "Pro-Fuse" Low Profile CT Port. | Employee error | Medcomp sales reps and foreign distributor were notified via email on 7/1/10. Product is to be returned to Medcomp. |  | 47 | 47 units | Worldwide Distribution -- USA, including states of NC and TX and country of Israel. | Terminated | October 13, 2010 | 103 | 47 | Not_Computer
Z-2234-2010 | 56138 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Accessory Power Pack; a 4 outlet power strip with a custom pole clamp bracket to attach to the IV pole with the infusion pump; Distributed by: Hospira, Inc., 275 N. Field Drive, Lake Forest, IL 60045; part number 39003-04-01, Modo #HSP1700, Model Tripp-Lite PS-415-HG |  |  | July 02, 2010 | 2010 | Hospira Inc. | The Power Pack Accessory is not certified to UL 60601 standard or Subject 1363A for relocatable power taps intended for use with medical equipment. Potential problems could include electrical shock, electrical short or electrical fire. | Device Design | Hospira sent Urgent Device Recall letters dated 7/1/10 to their direct accounts on 7/2/10 via Federal Express 2 day delivery, informing them of the potential for electrical shock, electrical short, or electrical fire, since the power packs were not certified to UL 60601 standard or Subject 1363A of UL standard 1363 for relocatable power taps intended for use with medical equipment. The accounts were instructed to examine their inventory for the affected product and quarantine it immediately, and Stericycle at 1-800-321-7388 to obtain a return kit for returning the affected product. If they distributed the affected product further, they were instructed to sub-recall the product from their sub-accounts. They were also requested to complete the enclosed reply form and return it to Hospira by fax at 1-262-577-6921, indicating the amount of product placed on hold for return, if they further distributed the product and if they notified their retail customers. For medical inquiries, the accounts were directed to contact Hospira Medical Communications at 1-800-615-0187. |  | 1766 | 1,766 units | Worldwide Distribution -- US and Canada. | Terminated | July 20, 2011 | 383 | 1766 | Battery | Physical Safety Hazards | Return | Remove or Replace
Z-2236-2010 | 55663 | No | FDI | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus SD Snares, Reusable Electrosurgical Snares. Models: SD-9L/U-1, SD-11L/U-1, SD-12L/U-1, SD-13L/U-1, SD-13L/U-1, &amp; SD-16L/U-1. Snare tube model numbers: MAJ-13 and MAJ-15. |  |  | May 03, 2010 | 2010 | Olympus America Inc. | An investigation of consumer complaints revealed the inability of the snare wire to be retracted into the snare tube after deployment of the snare. | Other | Firm initiated this recall following investigation of customer complaints regarding the inability of the snare wire to be retracted into the snare tube. Customers were notified via certified mail, an Urgent: Medical: Device Recall letter issued on May 4, 2010. The letter identified the affected product and the reason for the recall. Customers were instructed to inspect their inventory and identify the affected product. Customers are to remove from use any identified inventory, quarantine, and discard them. Customers should contact the firm to obtain replacement products. Questions should be directed towards Laura Storms-Tyler at 484-896-5688 or laura.storms-tyler@olympus.com. |  | 108 | 108 units | Nationwide Distribution -- AR, CA, CO, CT, DC, FL, GA, IN, KS, LA, MD, MO, NC, NY, OH, PA, SC, TN, and TX. | Terminated | May 20, 2013 | 1113 | 108 | Not_Computer
Z-2218-2010 | 54687 | No | NWX |  |  | Cardiovascular | PMA | AngioSculpt Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter, cutting/scoring, Easy Exchange (EX) delivery system, Manufactured by AngioScore, Fremont, CA. |  |  | December 04, 2009 | 2009 | AngioScore Inc. | Separations of the PTCA catheter shaft at the guide wire exit port have occurred during treatment of coronary artery stenosis. Fragments of the catheter may become lodged in coronary arteries, which resulting in serious patient outcome and require emergency coronary bypass surgery and medical intervention to remove the retained device fragments. | Process design | AngioScore sent an "Urgent: Medical Device Voluntary Recall" dated December 4, 2009 addressed to "Catheterization Lab Manager" at consignees. The Recall notification letter instructed consignees to inspect their inventory to detect product from the involved lot, and segregate it from their inventory. Consignees were asked to record all affected product and fax return forms to the firm. AngioScore Customer Service will contact consignees and provide instructions for return and replacement of affected product.. Consignees can contact AngioScore at 877-264-4692. |  | 3870 | 3870 units | Nationwide distribution. | Terminated | June 21, 2011 | 564 | 3870 | Not_Computer
Z-2237-2010 | 49012 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz Dual series Computed Tomography (CT) Scanner System, X-Ray System. Part Number: 989605651321. |  |  | July 21, 2008 | 2008 | Neuisys, LLC | The reference lines for image generated for a Surview scan may appear in the incorrect position in the "FILM" display mode. | Software design | On 07/21/2008, NEUISYS Imaging Systems Solutions distributor in Greensboro, NC began sending the URGENT DEVICE CORRECTION letter on behalf of Philips and NeuSoft Medical Systems (PNMS), Co., Ltd in China.The letter informs the consignees that there is potential safety problem in the NeuViz Dual series CT scanners due to the malfunction of "Incorrect Reference Line in the Film". When a "Head first, Couch out" or a "Feet first, Couch in" Surview scan is started then paused midstream, the image generated is what the doctor expects, but the reference lines appear in the incorrect position in "FILM" display mode. The acquired images accurately represent the exact positions planned on the original view. The incorrect/unmatched reference line in film may result in incorrect diagnosis.The consignees are advised not to stop or pause during Surview acquisition while PNMS in China will provide the updated software patch to all affected systems free of charge.The Field Change Order #FCOP-08-CT-NMS-008 SP4 will be implemented by NEUISYS Imaging Systems Solutions in Greensboro, NC., who is a distributor for PNMS. Consignees are advised to e-mail Service Support Department of PNMS at helpdesk@pnms.neusoft.com regarding the recall. |  | 19 | 19 units in the US | Nationwide Distribution | Terminated | August 23, 2010 | 763 | 19 | Software | Display/Image | Software update | Software Update
Z-2239-2010 | 55043 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The Starion Instruments Thermal Ligating Shears 35C Model 102-138D and 23C Model numbers 102-137D allow the surgeon to position the distal jaws of the instrument around the region of tissue to be cut/cauterized. The surgeon squeezes the handle and depresses the switch, heating elements activates in the jaws. The heat is conducted to the tissue between the jaws to provide cutting/cauterization. For the simultaneous cutting and cauterization of soft tissue during surgery. Cutting of natural or synthetic, non-metallic, sutures during surgery. |  |  | March 09, 2010 | 2010 | Starion Instruments | The Starion Instruments Thermal Ligating Shears (STLS) 35C and the Starion Thermal Ligating Shears (STLS) 23C may turn on and remain on without the finger button being depressed. The devices may also work intermittently, incorrectly indicating the device is 'on". | Device Design | Starion distributed a "STLS Device Recall - Important Medical Device Information" letter dated March 9, 2010 to consignees. The letter identified the affected model and lot numbers. Consignees were asked to complete a inventory status form and fax to Starion. the letter included a shipping labe with a pre-assigned return authorization number. Starion will ship replacement after receipt of the inventory status form and returned product.. |  | 228 | 228 units | Worldwide Distribution: USA states of CA, OH, MI, KS and MO and countries of Italy, China, Kuwait, and Germany. | Terminated | November 27, 2010 | 263 | 228 | Not_Computer
Z-2238-2010 | 55921 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 2000, manufactured by GE Medical Systems, SCS |  |  | June 10, 2010 | 2010 | GE Healthcare, LLC | A hardware component (potentiometer and encoder)mismatch in Innova 2000 digital fluoroscopic imaging systems may cause a Source to Image Distance (SID) failure, which will cause the imager to fail to detect the distance between the imager and the patient. This will result in loss of the x-ray image if the recovery process is not performed. The error message to use the Emergency Back-Out Button for | Other | GE Healthcare sent a GE "Urgent Medical Device Correction" dated June 9, 2010letter to consignees. The letter was addressed to Hospital Administrators / Risk Mangers; Mangers of Radiology/Cardiology and Radiologists/Cardiologists. The letter described the Safety Issue, Affected Product Details; Safety Instruction; Product Correction and Contact Information. The letter detailed the method for identifying SID failure and the use of the emergency back-out buttoms to recover the X-ray image. A GE service representative will affix a label to the device on-site to remind operators how to recover from loss of X-ray from SID failure. Consignees may contact GE at 800 437-1171, option 4. |  | 885 | 885 | Worldwide distribution: States of AL, AZ, AR, CA, CO, CT,DE, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NV,NH, NJ, NM, NH,NC, ND, OH, OK OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, and WI, and countries of VENEZUELA, UNITED KINGDOM, TURKEY, TUNISIA, THAILAND, TAIWAN, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN SPAIN, SOUTH AFRICA, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF KOREA, POLAND, PAKISTAN, NORWAY, NEW ZEALAND, NETHERLANDS, MOROCCO, MEXICO, MARTINIQUE, LITHUANIA, LEBANON, JORDAN, JAPAN, ITALY, ISRAEL, ISLAMIC REPUBLIC OF IRAN, IRELAND, INDONESIA, INDIA, HUNGARY, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, EGYPT , ECUADOR, COLOMBIA, CHINA, CANADA, BULGARIA, BRAZIL, BELGIUM, BELARUS, AUSTRALIA, and ALGERIA. | Terminated | April 01, 2014 | 1391 | 885 | Hardware | Device Operation |  | Other
Z-2240-2010 | 56258 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | PoleStar FlexiCoil band D/E, revisions A and B, Catalog # 9733063, |  |  | July 01, 2010 | 2010 | Odin Medical Technologies Ltd. | Electronic box on headstrap may overheat, causing burn to patient. | Device Design | Customers were notified by letter on 07/07/2010, advising them of the overheating problem and alerting them that a service rep will be visiting to inspect and upgrade any affected FlexiCoils. Questions were referred to 800-595-9709. |  | 560 | 560 units | CA, CO, DC, IL, IN, MD, MI, MN, MO, NE, NJ, NV, NY, OH, OR, WA, and WI. Foreign distribution to Austria, Belgium, Canada, China, France, Germany, Israel, Italy, Korea, Netherlands, Poland, Spain, Russia, Switzerland and Turkey. | Terminated | November 09, 2010 | 131 | 560 | Not_Computer
Z-2241-2010 | 55153 | No | NRU | Dental chair and accessories. | Dental | Dental | 510(K) Exempt | The magnetic headrest is designed to fit the following dental chairs (THE HEADREST IS THE RECALLED ITEM, NOT THE CHAIRS): |  |  | March 23, 2010 | 2010 | Dental Equipment LLC | The magnetic headrest on certain dental chairs may affect the function/programming of some implantable pacemakers or defibrillators if the implanted device is programmed to respond to a magnet. | Other | DCI Equipment sent an "Urgent Field Safety Alert" letter dated March 15, 2010 to consignees advising them of the potential problem. The letter instructed that patients with implantable pacemakers or defibrillators should avoid dental chairs with the magnetic headrest. A warning label was sent with the letter for insertion into the product Use and Care manual for operator warning. Consignees were asked to complete a Field Safety Alert Acknowledgement/Return Form and fax to the firm. Consignees were requested to notify customers.Consignees may contact the firm at 503 537-3602. |  | 108 | 108 | Worldwide distribution in the USA and Canada. | Terminated | January 04, 2012 | 652 | 108 | Other | Device Operation | Instructions to avoid | Safety Notice/Insructions
Z-2243-2010 | 56291 | No | NUK |  |  | Clinical Chemistry | Contact ODE | Human Cardiac Specific Fatty Acid-Binding Protein (H-FABP) ELISA TEST KIT Product is an In Vitro Device (IVD) categorized as a Cardiac Marker, manufactured and distributed by BioCheck, Inc., Foster City, CA. Enzyme Immunoassay for the quantitative determination of Human Cardiac Fatty Acid-Binding Protein (H-FABP) concentration in serum, citrate-plasma, EDTA-plasma, and heparin-plasma. |  |  | July 15, 2010 | 2010 | BioCheck Inc | Products were marketed without 510(k) for diagnostic use. Unapproved device. | PMA | BioCheck sent an "Urgent Device Recall" letter dated July 15, 2010 to consignees identifying the products as unapproved and requesting prompt return to the firm. Consignees are asked to sub-recall. The letter requested a reponse form be completed and returned to the firm.Diagnostic claims have been removed from the website. Returned product will be quarantined by the firm pending determination of marketing status and labeling revision. |  | 187 | 187 kits | Worldwide Distribution - USA, including the states of CA, MI, NJ, and PA and countries of Canada, India, S. Korea, and UAE. | Terminated | December 13, 2010 | 151 | 187 | Not_Computer
Z-2242-2010 | 56285 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Surgi-Tip Transducer Cover, Ref. 610-834, Sterile (20.3 x 244cm) polyethylene telescopically-folded cover with attached (6cm) NeoGuard tip. The latex-free surgi tipped covers are utilized to cover ultrasound transducers and prevent cross-contamination from occurring with equipment and patients. |  |  | July 20, 2010 | 2010 | Civco Medical Instruments Inc | Polyethylene material did not have proper treatment applied during manufacturing. Untreated polyethylene material does not maintain a secure seal with the adhesive tape used to secure the tip to the cover. | Nonconforming Material/Component | CIVCO sent a "Urgent - Medical Device Correctdion" letter date July 19, 2010 to consignees describing the reason for recall and requesting they dispose of the affect lot number of product. A response form was enclosed for completion and return to the firm, indicating the amount of product in inventory and requesting a Return Authorization Number. The form was to be returned via FAX or e-mail. Customers may contact CIVO at 800-445-6741. |  | 22 | 22 boxes | Nationwide distribution: the states of OH, MD, NE, VA, IA, MI, WA, and WI. | Terminated | April 13, 2012 | 633 | 22 | Not_Computer
Z-2244-2010 | 56200 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKnee Baseplates, Size 8, Left, Part # 333-01-108, Lot# A1000003. |  |  | June 28, 2010 | 2010 | Encore Medical, Lp | One lot of product has an incorrect color label. The color label is a secondary visual indicator of sizing. | Labeling mix-ups | Firm notified their consignees by letter on June 28, 2010. |  | 18 | 18 | Florida, Missouri, and Italy | Terminated | October 04, 2010 | 98 | 18 | Not_Computer
Z-2245-2010 | 54701 | No | HWQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF 8570 MENISCAL REPAIR NEEDLE, STERILE, Rx ONLY. |  |  | January 28, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | This recall was due to the possibility that the products may have a breach in the poly-Tyvek pouch that could potentially compromise the sterility of the contents. | Material/Component Contamination | Each consignee will be notified of the recall by an Urgent Medical Device Recall Notification letter sent by FedEx or an equivalent method. The letter identified the affected product, explained the reason for recall, and health risk. Consignees were instructed to examine their inventory and segregate the affected product for return. Customers are to complete and return the Reply Form and to notify their customers if the product was further distributed. Consignees are instructed to return any remaining product to ConMed Linvatec. Questions should be directed to the Customer Service department at 800-535-8536. |  | 35451 | 35,451 total | Worldwide Distribution -- United Arab Emirates, Austria, Australia, Belgium, Canada, Switzerland, Germany, Denmark, Spain, Finland, France, United Kingdom, Hong Kong, Ireland, Israel, India, Italy, Republic of Korea, Netherlands, Portugal, Serbia, Singapore, Slovenia, El Salvador, Thailand, Taiwan, South Africa, and throughout the U.S. | Terminated | December 15, 2010 | 321 | 35451 | Not_Computer
Z-2267-2010 | 55156 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom Care Assist Bed, Model P1170. |  |  | April 13, 2009 | 2009 | Hill-Rom, Inc. | Pinch hazard. Potential for shearing action between the side rail arm and the bracket used to secure the side rail to the bed. | Device Design | Hill-Rom sent a customer notification letter dated 4/9/2009. A second notice was sent to all customers on 7/9/2009. The letter included instructions and a Customer Response Form, along with the required parts. Customers were instructed to do the correction on their affected beds and return the completed customer response form within 45 days. |  | 39971 | 39,971 both products | Worldwide distribution: USA, Afghanistan, Algeria, American Samoa, Argentina, Australia, Bermuda, Brazil, Canada, Chili, China, Colombia, Costa Rica, Cuba, Ecuador, Egypt, Hong Kong, India, Indonesia, Israel, Italy, Japan, Jordan, Kuwait, Kyrgyzstan, Lebanon, Libya, Malaysia, Mexico, Oman, Panama, Paraguay, Philippines, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Korea, Syria, Taiwan, Thailand, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Yemen. | Terminated | June 06, 2011 | 784 | 39971 | Not_Computer
Z-2266-2010 | 56273 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Philips M1026BT Gas Modules with manifolds |  |  | July 21, 2010 | 2010 | Philips Healthcare Inc. | The manifold seal in the Philips M1657B Water Trap used on the Philips M1013A IntelliVue G1, M1019A, IntelliVue G5, and Anesthesia Gas Modules M1026B/BT may be unintentionally removed when changing the traps. A missing seal may result in leakage and may cause incorrect monitor readings of inspiratory or expiratory gas levels. | Device Design | Philips Healthcare issued an "Urgent Device Correction" letter dated July 14, 2010 with an addendum to the Instructions for Use including additional instructions for the replacement of the water trap for the Philips M1013A IntelliVue G1, M1019A IntelliVue G5, M1026B/BT Anesthesia Gas Modules. This addendum is being sent along with the Urgent Medical Device Correction notice to all affected customers.Customers will be requested to incorporate the information contained in the Urgent Medical Device Correction notice and the Addendum to the Instructions for Use with the existing labeling (Instructions for Use) provided with the product. Customers can contact Philips at 1-800-722-9377 for further information.. |  | 367 | 367 units | Worldwide Distribution: United States, and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Bulgaria, Chile ,China ,Colombia, Czech Republic, Estonia, Finland, France, Germany, Hong Kong, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Republic of Korea, Latvia, Libyan Arab Jamahiriya, Luxembourg, Malaysia, Mexico, Mozambique, Netherlands, New Zealand, Norway, Oman, Pakistan, occupied Palestinian territory, Peru, Philippines, Poland, Portugal, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Viet Nam. | Terminated | March 09, 2012 | 597 | 4675 | Not_Computer
Z-2278-2010 | 52030 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ABC Bend-A-Beam Handpiece REF 134009 Handcontrol Malleable Handpiece, 9" (22.8 cm) QTY: 1 STERILE CONMED CORPORATION 525 FRENCH RD UTICA, NY 13205 ASSEMBLED IN MEXICO (Pouch label) |  |  | December 18, 2008 | 2008 | ConMed Electrosurgery | The internal electrode can protrude in front of ceramic insulator at the tip of the ABC Handpiece used with Conmed Electrosurgical units. | Process design | An Urgent Device Recall notification was issued 12/18/2008, which explained the problem/risk associated with the device, the discontinuance of use and the return of the device. If you have any questions, please contact 800-552-0138, ext. 5225. |  | 233 | Dom: 233 cs (2330 units); Int'l: 275 cs (2750 units) | Nationwide and International | Terminated | October 22, 2010 | 673 | 5527 | Not_Computer
Z-2270-2010 | 50226 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Intesys Clinical Suite (ICS) G2 92810 Smart Disclosure. The Smart Disclosure product collects patient waveforms, alarms, vital signs and 12-lead reports. A 24-hour waveform acquisition is standard, with an upgrade option to 72 hours. |  |  | October 17, 2008 | 2008 | Spacelabs Healthcare, Incorporated | Potential for device to print out the wrong waveform data for the time indicated in the printout or display. | Device Design | On 07/17/2008, Spacelabs Healthcare began sending the MEDICAL DEVICE CORRECTION letter to their consignees. The letter informs the consignees that there is an issue involving ICS G2 Smart Disclosure (Model 92810) which causes the display or printing of an incorrect waveform for the time indicated in the printout or display.The consignees are advised whether to discontinue to use the device until the firm upgrades the system. If the consignees decide to continue to use the device, then they should discontinue any use of the "Refresh" button. When they print or view waveform reports,t hey should utilize the Saved Event view functionality of the product rather than Arrhythmia, Alarm or Waveform Views. If the printing from the Arrhythmia, Alarm, or Waveform views, the consignees must visually inspect the print out and verify that the waveform contains what you expected to see printed.The consignees could contact Spacelabs Healthcare at 1-800-522-7052 or 1-425-657-7200, x5089 for any questions about the corrective actions of the affected device. The firm's field service engineer will make a schedule a mutually agreeable time with the consignees to update the ICS G2 system at free of charge. |  | 135 | 135 units | Worldwide Distribution -- US, Turkey, Chile, Brazil, Netherlands, Puerto Rico, Switzerland, Australia, Czech Republic, Canada, Italy, Singapore, France, Germany, and Israel. | Terminated | September 29, 2010 | 712 | 135 | Software | Display/Image | Upgrade | Repair
Z-2269-2010 | 54810 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | ESSENTIAL MEDICAL SUPPLY, INC. ENDURANCE HD, BUILT TO LAST, HEAVY DUTY FOUR WHEEL WALKER for users up to 500 pounds. |  |  | December 21, 2009 | 2009 | Essential Medical Supply, Inc. | Essential Medical Supply, Inc. is recalling W1801B and W1801R Endurance HD Walkers due to recent reports and tests that show the front caster may detach from the frame. The potential failures include a collapse of the plastic fork or shearing of the stud attaching the caster to the frame. | Other | All Consignees were mailed an Urgent: Product Recall W1801B and W1801R Endurance HD Walkers notification on December 30, 2009. All notifications were sent first class mail in a 9" X 6" envelope. Response cards were included. The letter identified the affected product and the reason for recall. Customers were asked to examine their inventory for affected product and to discontinue further use or distribution of the product. If the product was further distributed, customers should contact their customers, advise them of the recall and have them return the units to them. Customers should follow the instructions provided for return and replacement of units. |  | 2954 | 2,954 units | Nationwide Distribution | Terminated | February 21, 2012 | 792 | 2954 | Not_Computer
Z-2274-2010 | 55933 | No | KGZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Arrow Percutaneous Sheath Introducer Kit with Integral Hemostasis Valve/Side Port for use with 7.5 Fr. Catheters; Catalog number AJ-09801 |  |  | June 07, 2010 | 2010 | Arrow International Inc | Sterility not assured. The stopcock handle position was inadvertently changed from an open to a closed position, which prevents sterilization of the closed portion of the device. Sterility of the affected product cannot be guaranteed. | Process change control | June 7, 2010 Notification letter to customer instructed them to return the product to Arrow. Arrow will destroy the returned product. |  | 3000 | 3000 kits | All affected product was distributed to one foreign consignee: Arrow Japan Ltd. | Terminated | September 29, 2012 | 845 | 3000 | Not_Computer
Z-1977-2010 | 55855 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | FLEXIMA Biliary Stent System, 7Fr-10cm, M00539220, Sterile EO, Boston Scientific, Spencer, IN. The Flexima Biliary Stent System is intended for delivery of the stent to the biliary tract for drainage of the bile ducts, for splinting of a bile duct during healing, or for providing bile duct patency in a stricture or past a stone. |  |  | April 27, 2010 | 2010 | Boston Scientific Corporation | Flexima Biliary Stent System Devices were packaged with Tyvek tray lids that had incorrect adhesive, resulting in sterile barrier seal defects. The breach of the sterility barrier could lead to the device being non-sterile. | Packaging | Boston Scientific sent an "Urgent Medical Device Recall - Immediate Action Required" letter date April 27, 2010 to all consignees. The letter included instructions to immediately discontinue use and segregate the affected product for return, and complete and return by fax a Reply Verification Tracking Form. Customers with affected product were instructed to contact customer service at 866-868-4004 to obtain a Return Authorization (RGA) number, and package and ship the product to Boston Scientific. |  | 1406 | 1406 | Worldwide Distribution: United States and Puerto Rico, and countries of Arab Emirates, Argentina, Austria, Belgium, Canada, China, Columbia, Croatia, El Salvador, Estonia, Finland, France, Great Britain, Germany, Greece, Hungary, Iran, Israel, Italy, Japan, Korea, Latvia, Lebanon, Malaysia, Mexico, the Netherlands, Norway, Poland, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, Ukraine, and Uruguay. | Terminated | March 23, 2011 | 330 | 13227 | Not_Computer
Z-2275-2010 | 56141 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE MONITOR B850 |  |  | July 16, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware of four potential safety issues associated with the CARESCAPE" Monitor B850:1. Loss of user input when using certain displays or the USB remote control.2. Potential for a delay in treatment or missed alarm when the alarm light is used as a primary alarm source and audible alarms are turned down or off.3. The CARESCAPE Monitor B850 will not use user entered hemo | Software design | Consignees were sent on 7/16/10, a GE "Urgent Medical Device Correction" letter dated July 12, 2010. The letter was addressed to Healthcare Administrator/Risk Manger, Chief of Nursing, Director of Biomedical Engineering. They stated the Safety Issues, Affected Product Details, Product Correction, Safety Instructions and Contact Information. If you have any questions, please contact Technical Support at 1-800-558-7044. |  | 875 | 875 | CA, CO, FL, GA, KY, LA, MA, MI, NE, NJ, NY, NC, OH, OR, PA, TX VA, and WI, Foreign: UNITED KINGDOM, TURKEY, SWITZERLAND, SPAIN , SLOVENIA, SINGAPORE, SAUDI ARABIA, PORTUGAL NETHERLAND, ITALY, GERMANY, FRANCE,FINLAND, DENMARK, CZECH REPUBLIC, BELGIUM, AUSTRIA, and AUSTRALIA. | Terminated | September 18, 2013 | 1160 | 875 | Other | Alarm/Message | Instructions | Safety Notice/Insructions
Z-2271-2010 | 56276 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MedtronicÂż EverestÂż 20 Disposable Inflation Device, AC2200. |  |  | July 16, 2010 | 2010 | Medtronic, Inc. | Product sterility may be compromised due to breach of package seal integrity. | Packaging process control | Medtronic issued a Urgent Medical Device Recall Notice to US customer on 7/15/2010 and to foreign accounts advising of the issue and requesting that they segregate recalled product. A Medtronic Sales Rep will visit their facility to account for all affected product and initiate a credit. Foreign accounts will return product to their respective distribution centers. |  | 2875 | 2,875 units | Worldwide Distribution -- USA, including the state of MI and the countries of Austria, Belgium, Czech Republic, Finland, France, Germany, Greece, Ireland, Israel, Italy, Lithuania, Pakistan, Poland, Portugal, Slovakia, Spain, Sweden, Switzerland, Ukraine, United Arab Emirates, and United Kingdom. | Terminated | June 07, 2011 | 326 | 4897 | Not_Computer
Z-2287-2010 | 54474 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Deknatel BondekÂż Plus Polyglycolic Absorbable Surgical Sutures, Bon+VI R 0 HRX27 1N (70 cm), Teleflex Medical. |  |  | February 01, 2010 | 2010 | Telefelx Medical | Labeling error; The Bondek Suture product inside the box are of a different catalog number than labeled on the outer box. | Labeling mix-ups | Teleflex Medical issued consignees a "Urgent Medical Device Recall" letter dated February 1, 2010. Distributors were instructed to forward the letter to their consignees to retrieve relevant product from those locations. The letter described the packaging issue and potential impact. Customers were instructed to immediately discontinue use and quarantine any identified products, complete and fax the Recall Acknowledgement Form to 866-804-9881, in order to receive a Return Goods Authorization (RGA) Number, and then return the product to Teleflex Medical for replacement. Affected products received by Teleflex Medical will be destroyed upon receipt. |  | 6936 | 17 sales units (24 packs per sales unit/408 units) | Worldwide Distribution: USA and Germany | Terminated | April 19, 2011 | 442 | 6936 | Not_Computer
Z-2288-2010 | 55772 | board, bed | FPS | Bed board. | General Hospital | General Hospital | 510(K) Exempt | ReliaMed Deluxe Portable Adjusting Table. Product is a medical table used on patients for massage and chiropractic treatment. |  |  | April 23, 2010 | 2010 | MDC Acquisition Co. | Two complaints were received of the vinyl-covered, padded, plywood table surface breaking near the foot end. | Component design/selection | "Urgent: Medical Device Recall" letters were sent to consignees from Milliken Medical and Meyer Distributing Company in April 2010 describing the issue. They instructed their customers to immediately stop use of the table. customers were instructed to return the tables to the firm with an enclosed Recall Form for future replacement, or for replacement with a comparable table from another menufacturer at no cost, or for refund of the original cost of the device. Consignees may contact Meyer Distributing Company at 800-472-4221 and Milliken Medical at 800-532-1356. |  | 265 | 265 Tables | Worldwide distribution: USA, including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA and WI, and the country of Canada. | Terminated | February 16, 2012 | 664 | 265 | Not_Computer
Z-2284-2010 | 55870 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Infinity Delta/Gamma X XL Patient Monitor. Portable monitor used to monitor vital signs of neonatal, pediatric, and adult patients for bedside and transport applications. |  |  | June 04, 2010 | 2010 | Draeger Medical Systems, Inc. | Monitor keys may become inoperative or activate spontaneously, causing a membrane switch panel malfunction. This may discharge a patient automatically. | Device Design | Draeger Medical issued "Urgent-Medical Device Recall Notification" letters to affected consignees. The Delta/Gamma X XL monitor keyboards will be updated with new material at each affected facility. If the keys on affected monitors become inoperative or activate without user interaction, the customer is instructed to remove the monitor from service and contact Draeger at 978-379-8573. A Technical Service Bulletin (TSB) will be distributed globally, instructing subsidiaries to inform international customers. |  | 800 | 800 units | Worldwide Distribution - United States, including the states of AL, AR, CA, CO, DC, FL, GA, IA, IL, IN, KS, LA, MA, ME, MN, MO, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SD, TN, TX, UT, VA, WA, WI, WO, and W, and the countries of United Arab Emirates, Angola, Argentina, Austria, Australia, Bosnia, Belgium, Bulgaria, Benin, Brazil, Canada, Switzerland, Chile, China, Columbia, Costa Rica, Serbia, Czech Republic, Germany, Egypt, Spain, Finland, France, Great Britain, Greece, Hong Kong, Coratia, Hungary, Ireland, Israel, India, Iraq, Italy, Japan, Jordan, Kenya, South Korea, Kuwait, Kazakhstan, Lebanon,Lithuania, Latvia, Libya, Morocco, Macedonia, Martinique, Mexico, Netherlands, Norway, Nepal, Peru, Poland, Puerto Rico, Portugal, Qatar, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovenia, Slovak Republic, Turkmenistan, Turkey, Ukraine, Uruguay, Venezuela, Mayotte and South Africa. The government accounts are Bay Pines VAMC, Bay Pines, FL; Naval Medical Center, Portsmouth, VA; and Walter Reed Army Medical Center, Washington, DC. | Terminated | July 30, 2012 | 787 | 9475 | Not_Computer
Z-2291-2010 | 56032 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VERSYSÂż HIP SYSTEM |  |  | July 05, 2010 | 2010 | Zimmer Inc. | Zimmer, Inc. has determined that there is a potential for a package to be labeled as a 28 mm femoralhead but contain a 32 mm femoral head. In addition, there is a potential for the patient record label tobe incorrect. | Other | The firm, Zimmer, Inc., sent an "URGENT: DEVICE REMOVAL" letter dated July 9, 2010, to all consignees. The letter described the issue, clinical implications, affected product, and required actions. The consignees were instructed to carefully review the letter, ensure users of the VerSys Hip System Femoral Head device at their faculty have been informed of this removal and perform a physical count of all affected product at their facility, record the data on the Inventory Return Certification Form and/or the User Facility/HCP Form and return the form via fax at (574) 372-4265 or email: aimee.wood@zimmer.com. Zimmer Inc. will conduct a removal of all remaining (unused) units of the subject lots. For shipping assistance, questions or assistance in notifying your accounts about the device removal, please contact Zimmer, Inc. at 1-800-613-6131 or 574-372-4463. |  | 103 | 103 | Worldwide distribution: USA and countries including Bolivia, Canada, New Zealand, Taiwan, Thailand, Singapore, and India. | Terminated | January 12, 2011 | 191 | 103 | Not_Computer
Z-2292-2010 | 56329 | No | HPZ | Eye sphere implant. | Ophthalmic | Ophthalmic | 510(k) | Sterile Bio-Eye Hydroxyapatite Orbital Implant and Conformer |  |  | June 07, 2010 | 2010 | Integrated Orbital Implants Inc | Integrated Orbital Implants, Inc. has initiated a recall of it's Sterile Bio-Eye Hydroxyapatite Orbital Implant and Conformer due to a labeling issue in which the expiration date of the product was omitted. | Error in labeling | Customers were contacted by telephone and email |  | 40 | 40 units | USA (MT, AL, IN, IA, NY, KY, IL, MI, and FL), Canada, Germany, and Turkey | Terminated | August 31, 2010 | 85 | 40 | Not_Computer
Z-2285-2010 | 56542 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Suros ATEC 0912-20 Breast Biopsy and Excision System Handpiece, REF ATEC0912-12, sterile, Hologic, Inc. 6100 Technology Center Drive Indianapolis, IN. |  |  | August 11, 2010 | 2010 | Hologic, Inc. | The firm received a customer complaint that the a device labeled as ATEC 0912-12 was actually ATEC 0912-20. | Packaging | The firm, Hologic, Inc., sent out "Urgent Medical Device Recall" letters by FedEx on August 11, 2010, to all customers. The letter describes the product, problem and action to be taken by the customers. The customers were instructed to check their inventories and to quarantine any of the recalled product; if the product was distributed further, identify the users and notify them at once of this product recall; segregate all products from this lot and quarantine them in preparation for return to Hologic, and complete and return the enclosed RECALL RESPONSE FORM by fax at 877-574-3255 or mail. Note: Replacement product cannot be sent until Hologic receives your response form.If you have any questions, please contact Hologic directly at 317-344-7630. |  | 990 | 990 | Worldwide distribution: USA and countries including Canada, Belgium, Algeria, Germany, London, Lebanon and Italy. | Terminated | April 13, 2011 | 245 | 990 | Not_Computer
Z-2289-2010 | 56328 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer ITST" INTERTROCHANTERIC/SUBTROCHANTERIC FIXATION LOCKING BOLT EXTRACTOR - NONSTERILE |  |  | July 05, 2010 | 2010 | Zimmer Inc. | Instruments were manufactured with the incorrect grade of steel, increasing the potential for fracture during use. | Nonconforming Material/Component | Zimmer sent "Urgent: Device Removal" letters dated July 8, 2010 to consignees identifying the potential for fracture of the device. Consignees were instructed to inventory affected product and record on the Inventory Return Certification Form, to document users of the affected device on the User Facility/HCP Form, and return both forms by fax or email. Affected product is to be returned to Zimmer with a copy of the Inventory Return Certification Form. Returned product will be destroyed. Each of the affected consignees will be provided with a new instrument, manufactured with the correct 17-4 stainless steel.Consignees can contact Zimmer at 800-613-6131 or 574-372-4463 |  | 20 | 20 individual units | Worldwide Distribution in the countries of US, Canada, Spain, Germany, and Great Britain. | Terminated | February 25, 2011 | 235 | 20 | Not_Computer
Z-2309-2010 | 55877 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Mitek VAPR Temperature Control Electrode (integrated |  |  | June 07, 2010 | 2010 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Product does not meet the packaging specifications for pouch seal integrity which could lead to a breach in sterility | Packaging process control | The firm, Depuy Mitek, sent an "Urgent Voluntary Product Recall" letter dated June 7, 2010, to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to not use any of the identified recalled products; to isolate all inventory of the affected product; complete the Business Reply Form, and return the affected products along with the form to Depuy Mitek; ATTN: Recall Coordinator, 50 Scotland Boulevard, Bridgewater, MA 02324 or fax completed form to 1-508-828-3750. Note the IMPORTANT PRODUCT INFORMATION: Any product that contains a round green sticker has been 100% inspected, meets specifications and is acceptable for use. Please Do Not Return any product that is labeled with the green dots because this product is not affected by the recall.If you have any questions or concerns with regard to this program, please contact your local Depuy Mitek Representative or DePuy Mitek Customer Service directly at 1-800-382-4682. |  | 1414 | 1414 units | Worldwide distribution: USA and countries including Canada, Argentina, Australia, Bahrain, Belgium, Brazil, Chile, China, Columbia, Costa Rica, Czech Republic, Ecuador, Egypt, France, Germany, Greece, Guatemala, Hong Kong, India, Indonesia, Iran, Iraq, Israel, Italy, Japan, Jordan, Kazakhastan, Kuwait, Lebanon, Maylasia, Mexico, Netherlands, New Zealand, Oman, Pakistan, Palestine, Panama, Peru, Poland, Qatar, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Portugal, Sweden, Switzerland, Syria, Thailand ,Turkey, UK, Ukraine, UAE, Venezuela, Vietnam, and Yemen. | Terminated | July 12, 2012 | 766 | 813657 | Not_Computer
Z-2294-2010 | 56281 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Greatbatch OptiSeal Valved PTFE Peelable Introducer, Model 1000093-001, 1000093-002, 1000093-003, 1000093-004, 1000093-005, 1000093-006, 1000093-007. Sterilized using Ethylene Oxide. Greatbatch Medical, 2300 Berkshire Lane North, Minneapolis, MN 55441 USA. The OptiSeal Valved PTFE Peelable Introducer is designed to reduce blood loss and air intake by providing hemostatic sealing at venous pressures. The device is offered in two configurations, one without a sideport and one with a sideport attached to a segment of extension tubing terminating in a 3-way stopcock. |  |  | July 02, 2010 | 2010 | Greatbatch Medical | The outer sterile barrier (pouch) packaging for the 12 Fr device is not intact. No breaches in the inner sterile barrier (tray/tyvek lid) have been identified. The OptiSeal introducer and accessories sealed in the tray are sterile. Infection may result if a sterile field is contaminated by the outer surface of the inner tray. | Packaging | Greatbatch Medical issued an "Urgent Medical Device Recall letter dated July 12, 2010 via email to consignees. The letter described the problem and affected device. Consignees were instructed to immediately discontinue distribution of the product and return the unused units to Greatbatch Medical and to notify their customers to immediately return unused units. For shipping assistance, questions or assistance in notifying accounts, contact Shannon Springer at 763 951-8244 or Kimberly Briggs, 716 759-5702. |  | 216 | 211 5-pack boxes | Worldwide Distribution: United States, including the states of MN and CA. | Terminated | December 17, 2011 | 533 | 216 | Not_Computer
Z-2293-2010 | 55612 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator/Monitor with Q-CPR |  |  | May 06, 2010 | 2010 | Philips Healthcare Inc. | Reduced sensitivity in the force measurement sensor can result in inaccurate feedback to the user on how well he or she is releasing pressure after each chest compression. This can result in incomplete release of chest compression, which will reduce CPR effectiveness. | Device Design | Philips issued an "Urgent Medical Device Recall notification/Field Safety Notice" dated April 2010 to all consignees. The letter described the issue, how to identify affected devices, and required actions. The consignees were instructed to remove affected devices from service. Philips will contact consignees to arrange on-site replacement of affected meters with meters containing revised labeling and calibration software. The new meters will indicate that meter recalibration is required after approximately 50,000 compression cycles. Consignees returning devices to Philips for routine recalibration will receive free replacement. All CPR meters will be replaced with re-designed meters when they become available.For questions or assistance in notifying accounts about this recall, contact Philips at 1-800-722-9377. |  | 72 | 72 units | Worldwide distribution in the countries of USA, Canada, Australia, China, Germany, Iceland, India, Libyan Arab Jamahiri, Netherlands, Saudi Arabia, Spain, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | February 19, 2013 | 1020 | 72 | I/O | Output/Calculation | Replace | Remove or Replace
Z-2296-2010 | 56420 | No | GDI | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | V. MuellerÂż Peanut Sponges, Catalog number 23275-470, Sterile, For Single Use Only, Size: 3/8, X-Ray Detectable In Holders, Distributed by Cardinal Health, McGaw Park, IL 60085 USA. There are various uses including as gauze/sponge internal, x-ray detectable and as dissectors. |  |  | March 01, 2010 | 2010 | He, Inc | Sterile product packaging contains weak seals, which may result in compromised product sterility. | Packaging | Harod Enterprises, Inc. (HE) issued a "Notice of Recall" letter to Cardinal Health/V. Mueller dated March 1, 2010 identifying the issue and affected product. Cardinal Health/V. Mueller will issue a subrecall letter to all their customers (end users) instructing them to discontinue use of the affected lots and return for replacement. and will track all responses and removal of the product from their private label customers. HE will repackage and/or replace all returned product.Consignees can contact Harod Enterprises, Inc. at 1-706-228-5165. |  | 55100 | 551 cases of 100 packages per case | Nationwide: including state of IL. | Terminated | November 02, 2011 | 611 | 55100 | Not_Computer
Z-2295-2010 | 56365 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK TWO AST-N085 REF 22 225 For In Vitro Diagnostic Use - each test card is packaged into a foil pouch with 20 pouches per box. The firm name on the label is bioMerieux, Inc., Durham, NC and bioMerieux, sa, France.This is an Export Only test kit that is not marketed in the U.S. |  |  | July 29, 2010 | 2010 | Biomerieux Inc | Foil pouch was punctured by drug susceptibility cards during packaging, allowing moisture into pouch, resulting in inaccurate and/or incorrect drug susceptibility results. | Packaging | BioMerieux issued an Urgent Product Removal Notice, dated July 29, 2010 to consignees, identifying the reason for the recall and the affected lots. Customers were instructed to discontinue use of and destroy any remaining inventory of the affected lot, and to notify their customers who received the product. All consignees were instructed to complete and return the Customer Acknowledgement and Product Replacement Form. These cards will be used to distribute replacement cards. Consignees can contact BioMerieux at 1-314- 731-8667. |  | 25940 | 1297/ 20 card boxes | Worldwide Distribution: Country of Canada only. No product was distributed in the US. | Terminated | November 02, 2010 | 96 | 25940 | Not_Computer
Z-2310-2010 | 56218 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Genzyme Biosurgery Essentials Kit, Item Number 82004 |  |  | July 13, 2010 | 2010 | Genzyme Corporation | Revised Labeling of Carticel Essentials Kit clarifiesthe non-sterile packaging of the outer clear plastic tray which should not be opened in the sterile field. | Labeling design | Genzyme notiifed physicians with the Dear Doctor letter on July 13, 2010 via UPS Ground delivery. The letter informed healthcare workers of the recent incidents and advise them to heed the Essentials Kit label and exercise proper handling of the Kit.. A fax return form will be provided for the physician to sign and return to Genzyme to acknowledge receipt of the notification. ContactGenzyme Medical Information at 800-453-6948, option #3 for any additional information or contact your Genzyme sales representative with any questions. |  | 3132 | 3132 kits | Nationwide | Terminated | January 30, 2012 | 566 | 3132 | Not_Computer
Z-2312-2010 | 50214 | Yes | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | Pediatric Paddles attachments used with LIFEPAK 9, LIFEPAK 9P, LIFEPAK 12 and LIFEPAK 20/20e Defibrillator/Monitors. |  |  | November 17, 2008 | 2008 | Physio Control, Inc. | Sterilization efficacy could not be confirmed for sterilization methods on External Sterilizable Paddles and Pediatric Paddle attachments. | Labeling False and Misleading | On 11/14/08, Physio Control began sending the Urgent User Information letter via certified mail to their customers. The letter notified the customers about the modification to the cleaning and sterilization instruction of the External Sterilization Defibrillation Paddles and Pediatric Paddle attachments used with LIFEPAK 9, LIFEPAK 9P, LIFEPAK 12, and LIFEPAK 20/20e Defibrillator/Monitors The customers are informed to discontinue use of the gas sterilization methods listed as 12/88 Ethylene Oxide (EO) or 100% Ethylene Oxide (OE). They are notified that they can continue to use the sterilization method such as STERRAD per existing Operating Instructions, but limit sterilization cycles from 100 to 50 cycles.The customers are advised that prior to and after each use, they should examine cables and connector for damage or signs of wear such as loose connections, damaged pins, exposed wires and cable connector corrosion and examine paddles for scratched or pitted surfaces and stop using product if damage has occurred. Within 45 days, the firm will update the Operating Instruction that will include revised Sterilization Guidelines on a website www.physio-control.com/products/product-support. Customers can call Technical Support at 1-800-442-1142, #5 or they can visit the firm's website at www.physio-control-notices.com/extpaddles regarding the recall. |  | 1689 | 1689 units | Devices were distributed to medical facilities through out the US and WORLD WIDE to Canada, Germany, India, Netherlands, Australia, Hong Kong, Japan, Brazil, Panama, Russia, Korea, Malaysia, Venezuela, and United Arab Emirates(UAE) . | Terminated | September 28, 2010 | 680 | 39874 | Not_Computer
Z-2334-2010 | 53427 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Total Knee System Primary Tibial Baseplate; |  |  | September 01, 2009 | 2009 | Stryker Howmedica Osteonics Corp. | A lot for lot mix up occurred between the Triathlon Total Knee System Primary Tibial Baseplate # 4 Cemented, Catalog # 5520-B-400, Lot SH8NT and the #2 Cemented, Catalog # 5520-B-200, Lot SH7RP. | Mixed-up of materials/components | The firm, Stryker Orthopaedics, sent an "URGENT PRODUCT RECALL" letter dated September 1, 2009, via FedEx to all customers. The letter described the product, problem and action to be taken by customers. The customers were instructed to examine their inventory and hospital locations to identify the affected product; retrieve and return all affected products via Inter-Org Transfer and ship to Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, 07430; Attn: Regulatory Reporting, and attach the fluorescent orange PRODUCT REMEDIATION sticker and mark the outer box with the words "PRODUCT RECALL; complete and fax the attached Product Accountability Form and spreadsheet to (201) 831-6069 within 5 days of receipt of the notice. If you have any questions, please contact the Manager, Divisional Regulatory Reporting at (201) 831-5970. |  | 6 | Lot Code: SH7RP - 6 in the US. Lot Code SH8NT was not received in the US. | Nationwide distribution: AL, AZ, CA, MT, NJ and TX. | Terminated | August 27, 2012 | 1091 | 6 | Not_Computer
Z-2313-2010 | 56284 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | VP 2000 Processor Heated Reagent Basins; automated slide stainer; Abbott Molecular Inc., Des Plaines, IL 60018; The basins are sold under the the following part numbers: |  |  | July 20, 2010 | 2010 | Abbott Molecular | The VP2000 Processor Heated Reagent Basins crack along the bottom seam, have non-uniform coating of interior corners, may be warped, and some may have incorrect dimensions. | Reprocessing Controls | Abbott sent Urgent Product Recall letters dated 7/20/10 to the VP 2000 Processor customers on same date via FedEx, informing them that the heated reagent basins could crack along the bottom seam, have non-uniform coating of interior corners, warping of the aluminum housing or incorrect dimensions. These deficiencies could result in leakage or the inability to place the basins into the VP 2000 Processor, resulting in possible exposure to chemical reagents, a slippery floor hazard, and/or delay or failure to start a protocol run. An Abbott Molecular customer service representative will call them to coordinate replacement of all of their heated reagent basins before 10/31/10. Customers were to complete and return the Customer Reply form provided. Any questions were directed to their local Abbott Molecular Representative.*** A second recall letter was sent to all customers on 12/21/10 explaining that the VP 2000 heated reagent basins previously inspected and determined to be acceptable could develop a leak, and that they would be contacted by an Abbott Molecular Customer Service Representative or Field Service representative in early 2011 to coordinate replacement of all of their VP 2000 heated reagent basins. *** |  | 219 | 219 basins | Worldwide Distribution -- Australia, Belgium, France, Germany, India, Italy, Japan, Pakistan, Poland, Russia, Singapore, South Korea, Switzerland, Taiwan, Thailand, United Kingdom, Uruguay, and the United States. | Terminated | April 19, 2012 | 639 | 219 | Not_Computer
Z-2315-2010 | 56064 | No | OAH | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | E.faecalis/OE PNA FISH |  |  | June 08, 2010 | 2010 | AdvanDx, Inc. | Some strains of streptococus anginosus cross-reach with E. faecalis/OE PNA FISH to produce a false positive green signal. | Device Design | The firm, AdvanDx, sent a letter dated June 8, 2010, and email notification on June 11, 2010, to all customers. The letter described the product, problem and action to be taken by customers. AdvanDx notified the customers that there was a change to the package insert under "limitations" to include the statement "some strains of Streptococcus anginosus produce a false green positive fluorescence due to sequence similarities and enclosed a package insert for all E. facealis/OE PNA FISH (KT003) kits. The customers were instructed to replace the Package Insert (revision B) included with any kits in their inventory with the enclosed Package Insert (revision C). Note: All shipments of E. faecalis/OE PNA FISH, as of June 14, 2010, will contain this new package insert.If you have any questions or requests for further information, please contact AdvanDx Technical Support at US (Toll Free): 1-800-376-0009 or email techsupport@advandx.com. |  | 541 | 541 units | Worldwide distribution: USA and countries including Danmark, Finland and Sverige (European). | Terminated | June 06, 2011 | 363 | 541 | Not_Computer
Z-2335-2010 | 53811 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Abbott Spine PathFinder, Cannulated Poly Screwdriver II - MOD, non-sterile, Manufactured by Abbott Spine, Inc., Austin, TX; Ref. 1160-2. |  |  | January 12, 2009 | 2009 | Zimmer Inc. | The screwdriver may fracture, resulting in surgical delays and/or fragments going into the surgical site. | Other | The sales force was provided with new instructions for use on 1/12/09 and again in 05/09. On April 6, 2010 the firm sent out an Urgent Medical Device Correction to it's field representatives for distribution to it's customers. The field representatives were to distribute the forms and provide a list of the customers that recieved the recall notice.For questions or assistance, please contact Zimmer Spine at 800-326-0635. |  | 691 | 691 | Worldwide. | Terminated | December 22, 2010 | 709 | 691 | Not_Computer
Z-2321-2010 | 56331 | No | ESW | Esophageal prosthesis. | General &amp; Plastic Surgery | Gastroenterology/Urology | 510(k) | Boston Scientific Ultraflex Covered Esophageal Stent System - Proximal Release (18/23/15/12) |  |  | July 15, 2010 | 2010 | Boston Scientific Corporation | Units have excessive fraying of the deployment suture thread, as a result, these devices may experience a suture break during deployment. | Process control | Boston Scientific issued "Urgent Medical Device Recall - Immediate Action Required" letters dated July 15,2010 to consignees. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. Foreign accounts received notification translated into the local language and with changes that are applicable to local regulations.Boston Scientific can be contacted at 1-508-683-4427. |  | 10 | 10 units | Worldwide Distribution: USA, and the countries of: Austria, Germany, Finland, France, UK, Ireland, Iran, Italy, the Netherlands, Portugual, Spain, and Sweden | Terminated | October 14, 2011 | 456 | 118 | Not_Computer
Z-2314-2010 | 56402 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | PERI-LOC (R), 3.5 MM ANTEROLATERAL DISTAL TIBIA PLATE, 10 HOLE, RIGHT - 148 MM, REF 71800710, QTY: (1), STERILE R, CE 0086, Smith &amp; Nephew, Inc., Memphis, TN 38116, USA, EC REP: Smith &amp; Nephew Orthopaedics Gmbh, Tuttlingen, Germany. Adult and pediatric patients, as indicated for pelvic, small, and long bone fracture fixation. |  |  | July 20, 2010 | 2010 | Smith &amp; Nephew Inc | Inner packaging of sterile product was not sealed. Lack of sterility assurance. | Packaging | Smith &amp; Nephew issued an "Urgent:: Medical Device Recall" letter dated July 20, 2010 to consignees (International distributors), identifying the issue and affected product. Consignees were requested to identify and quarantine affected stock, and cease distribution. Consignees were instructed to contact Smith and Nephew for a Return Authorization number. The Return Authorization number and an inventory of affected devices are to be recorded on the Verification Section of the notification letter and faxed to the firm. Affected stock is to be sent to Smith &amp; Nephew. Consignees were requested to notify anyone to whom the product was distributed to also contact the firm for return of the product. Smith &amp; Nephew can be contacted at 1-901-399-5441. |  | 10 | 10 units | Worldwide Distribution: INTERNATIONAL ONLY: Countries of China, UK, Italy, Australia, and South Africa | Terminated | November 17, 2011 | 485 | 10 | Not_Computer
Z-2339-2010 | 52830 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Toshiba Aquilion ONE CT SystemTSX-301 A/2D |  |  | February 03, 2009 | 2009 | Toshiba American Medical Systems Inc | ECG Gated Dynamic Volume Prospective scan may stop if preset eXam Plan is confirmed and CTA On is changed to CTA OFF, due to exposure control PWB. | Software design | Toshiba America Medical System Inc (TAMS) sent an Urgent: Medical Device Correction dated January 14, 2009 to their consignees. The letter described the issue and provided instructions to test the device to determine if it is an affected device. If the device is affected, contact the Toshiba service representative at 800 521-1968. Consignees were requested to complete and fax a response form to Toshiba. |  | 16 | 16 | Nationwide Distribution: including the states of AL, CA, FL, GA, KY, MA, MD, MT, NV, NY, OH, PA, and Puerto Rico. | Terminated | September 02, 2011 | 941 | 16 | Software
Z-2342-2010 | 54265 | bed, manual | FNJ | Manual adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Eye surgery stretcher with crank fowler, model 1079 Stryker Medical 3800 E. Centre Ave Portage MI 49002. |  |  | February 26, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The fowler of the bed is raised by a crank. The crank consists of a steel screw that moves through a brass nut. Over time the screw can wear the threads on the brass nut. The screw and the nut can become misaligned, causing a loud squealing noise when the screw comes in contact with the drive tube. In addition the fowler can become inoperable over time with excessive wear of the brass nut. T | Device Design | The firm, Stryker Medical, sent an "Urgent Medical Device Correction" letter dated February 26, 2010, to all customers. The letter describes the product, problem and action to be taken by customers. The letter states that Stryker field service tech will contact the firm and perform device corrections within four months. The customers were instructed to locate the recalled stretchers and verify that there is no grinding noise when activating the fowler. If the fowler exhibited a grinding noise, the customers were instructed to remove the stretcher from service and contact the recalling firm at 800-STRYKER, option 3. Note: Do not put the stretcher back into service until it is repaired by our field service representative. The customers were also instructed to return the enclosed post card to confirm receipt of this notice; inform any new users and advise the firm of new locations if any of the listed stretchers were on loan or sold, and advise the firm of the obsolescence by providing the serial numbers if any of the listed stretchers were disposed of. If you have any urgent questions or concerns, contact (269) 324-6884. Normal hours are Monday-Friday 8am-5pm(EST). |  | 729 | 729 | Worldwide distribution: USA and countries including Australia, Canada, Chile, China, France, Germany, Greece, Spain, Italy, Japan, South Africa, Taiwan, Korea, Singapore, and UK. | Terminated | April 21, 2011 | 419 | 3732 | Not_Computer
Z-2337-2010 | 55699 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER LH 500 Series System |  |  | February 12, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter identified four (4) issues with the COULTER LH 500 Series System and the Unicel DxH 800 Coulter Cellular Analysis System:Issue 1:There is a potential for misidentification to occur when the system is configured in languages other than English or Chinese. The characters #,@,[,\,],`,{,¦,}, or ~ are substituted or omitted in other languages. The issu | Software design | The customer recall notification was initiated when Product Corrrective Actions letters with attached Customer Response form were sent on the week of February 15, 2010 to all customers who purchased the COULTER LH 500 Series System and the Unicel DxH 800 Coulter Cellular Analysis System. The letters will provided the customers with an explanation of the problem identified and a work around.Customers with questions regarding the Product Corrective Action were instructed to call 800-526-7694 in the UnitedStates or Canada, or contact your local Beckman Coulter Representative. |  | 1520 | 1373 units in US and 147 units in Canada | Throughout the United States and Canada | Terminated | September 10, 2010 | 210 | 3152 | Software
Z-2348-2010 | 56323 | No | NQP | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | NFlex Stabilization System 150mm |  |  | February 09, 2009 | 2009 | Synthes USA (HQ), Inc. | package integrity compromised | Equipment maintenance | The firm, SYNTHES, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated February 9, 2009 and an updated letter dated February 12, 2009, with updated expiration information to customers. The letters described the product, problem and action to be taken by customer. The customers were instructed to call Synthes at 1-800-479-6329 if they have any of the identified devices to obtain a Return Authorization Number; complete the enclosed Verification Section attachment and return this letter with the recalled product to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If they do not have the identified product ,complete the Verification Section attachment and return the document to Synthes by Fax 610-719-5120 or scan/email: lewis.lynne@synthes.com.If you have any questions, please call 800-620-7025 x 5454 or 610-719-5454 or contact your Synthes Spine Sales Consultant. |  |  |  | Worldwide distribution: USA including states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NJ, NM, NV, NY, OH, OR, PA, SC, TN , TX, UT, VA, WA, WI, and WV and countries including Denmark, Germany, Korea and Switzerland. | Terminated | September 15, 2010 | 583 |  | Not_Computer
Z-2355-2010 | 56214 | No | JNL | Thyroid autoantibody immunological test system. | Immunology | Immunology | 510(k) | Beckman Coulter Access Thyroglobulin Antibody II Reagent, Part Number: A32898. The Access Thyroglobulin Antibody II (TgAb) assay intended use is for the quantitative determination of thyroglobulin antibody levels in human serum and plasma using the Access Immunoassay Systems. |  |  | June 29, 2010 | 2010 | Beckman Coulter Inc. | The affected Access Thyroglobulin Antibody II reagent may produce inaccurate results due to a change in the thyroglobulin that was used in the kits. | Other | Beckman Coulter sent an "Urgent Product Corrective Action" letter dated June 29, 2010, with attached Customer Response form to the affected customers. The letter identified the affected product and issue.Customers were instructed to:(1) Discontinue use of Access Thyroglobulin Antibody II reagent lot numbers 916271, 917073, and 918433.(2) Review results in the context of the total clinical presentation of the patient, and rerun questionable results using another reagent.(3) Complete and return a response form within 10 days.(4) For replacement contact Beckman Coulter at (US) 1-800-526-3821, (Canada) 1-800-463-7828. Customer Technical Support can be contacted at 1-800-854- 3633. |  | 1071 | 1071 units | Worldwide distribution : United States, including states of CA, CO, FL, GA. IA, IL, KY, MA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI, and country of Canada. | Terminated | May 04, 2012 | 675 | 1071 | Not_Computer
Z-2340-2010 | 56364 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Long Nail Kit R2.0, Ti RIGHT 011x360 mm x 125 degrees; |  |  | July 23, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The nail has no thread for the set screw hindering insertion. | Process control | Urgent Product Recall letters were sent on July 23, 2010 by Federal Express to Stryker branches. A corrected letter was sent on July 26, 2010 to correct an error in the size of the screw. Additionally, letters were sent on July 26, 2010 to Hospital Risk Managers, Chief of Orthopaedics and Surgeons. The letter identified the affected product, explained the issue, discussed the potential hazard, and the risk mitigation. Customers are to examine their inventory, retrieve all affected product and return it; then reconcile all products on the Product Recall Acknowledgement Form and fax it back. Questions should be directed towards Richard Wolyn at 201-972-2100. |  | 12 | 12 units | Nationwide Distribution -- NJ, NY, DE, IN, WI, IL, PA and NM. | Terminated | August 06, 2012 | 745 | 12 | Not_Computer
Z-2350-2010 | 56309 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN 3200 CS and SL, Hematology Analyzer manufactured by Abbott Laboratories, Santa Clara, CA. Waste Reservoir Assembly Part number 8921174802. |  |  | April 16, 2010 | 2010 | Abbott Laboratories | Epoxy holding the two parts of the waste reservoir together fails, potentially resulting in biohazardous fluid leaks. | Nonconforming Material/Component | Product Recall letters were sent on 04/16/2010. The letters identified the affected product, explained the reason for recall, the patient impact/user safety, stated the necessary actions, and contact information. Users are to discontinue use until the Waste Chamber(s) have been replaced and follow the instructions in the Hazards section. In addition, appropriate protective equipment needs to be worn along with following their recommended laboratory procedures dealing with spills. Users are to contact their local Abbott Customer Support rep. US customers should be direct their questions to Customer Support. |  | 69 | 69 units total. | Worldwide Distribution -- CA, CT, ID, IN, and MA; and, countries of Australia, Brazil, Canada, China, Colombia, Mexico, Puerto Rico, Singapore, and Uruguay. | Terminated | December 06, 2010 | 234 | 69 | Not_Computer
Z-2354-2010 | 56489 | No | EFA | Dental handpiece and accessories. | Dental | Dental | 510(k) | Titan 3 5K Low-Speed Motor - Star Titan 3, 5,000 rpm- Lube Free |  |  | August 11, 2010 | 2010 | Dental EZ Stardental Division | The housing of 5K motor is labeled as a 20K motor. | Labeling mix-ups | DentalEZ group/Star Dental Division sent a "Medical Device Recall" letter to consignees on 8/11/10, informing them that the device's 5K motor was mislabeled as a 20K motor. Customers are to stop using of the device, and return to the manufacturer or to the distributor using supplied shipping labels. Replacement devices will be sent to customers at no charge. Consignees are to complete and return a form indicating whether they have any affected devices, and whether customers have been contacted about this recall. Customers can contact DentalEZ at 866-383-4636 . |  | 41 | 41 units | Nationwide Distribution: States of IN and TX | Terminated | December 13, 2010 | 124 | 41 | Not_Computer
Z-2376-2010 | 54625 | No | GXZ | Needle electrode. | Neurology | Neurology | 510(k) | 21G x 150mm Stimulation Catheter Continuous Nerve Block Tray, REF 554-021-150, Rx, Sterile, Distributed by: Stryker Instruments, Kalamazoo, MI. |  |  | February 01, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | Defects include, but not limited to: Not sterile, not properly labeled, drugs and swabs may be ineffective, and catheter tips may break in patients. | Process control | Stryker Instruments issued an "Urgent Medical Device Recall Notification" letter to consignee risk managers, directors and material managers on February 2, 2010, identifying the affected devices and product issues. The consignees were instructed to:1) Forward the recall notice and reply form to all affected locations. 2) Review all inventory and quarantine affected product.3) Destroy all affected product using attached instructions.4) Complete and return the Business Reply Form to the firm by fax.Stryker Instruments can be contacted at 1-9800-800-4236 ext 4354 |  | 18826 | 18,826 all products | Nationwide Distribution: United States, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, and UT. | Terminated | March 28, 2011 | 420 | 18826 | Not_Computer
Z-2380-2010 | 56561 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Vital Signs Inc. enFlow Disposable IV Fluid/Blood Warmer Cartridge. Medical point-of-care device designed to warm blood and IV fluids. |  |  | July 15, 2010 | 2010 | Vital Signs Inc | The enFlow Disposable IV Fluid/Blood Warmer Cartridge's male luer-lock fitting may be oversized and may not fit properly with standard female luer fittings. | Equipment maintenance | GE Healthcare issued an "Urgent Medical Device Correction" letter, dated July 15, 2010, to consignees, identifying the reason for the recall and the affected lots. Customers were instructed to discontinue use of and quarantine the affected devices, and contact Vital Signs Inc. Customer Service to arrange for return and replacement of the affected devices. Consignees were instructed to complete and return the Confirmation Form. Consignees can contact Vital Signs Inc. Customer Service at 1-973-956-5300 . |  | 406596 | 385,820 domestic - 20,776 international | Worldwide Distribution; USA (all states), and to the countries of Australia, Ecuador, Finland, India, Israel, Latvia, Lebanon, Malaysia, Mexico, Morocco, New Zealand, Panama, Philippines, Qatar, Saudi Arabia, Singapore, Switzerland, the Netherlands, and Vietnam. | Terminated | April 27, 2015 | 1747 | 406596 | Not_Computer
Z-2384-2010 | 53819 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Whatman Body Fluid Collection Paper- 903 California Newborn Screening Card Ref 110534 |  |  | October 29, 2009 | 2009 | Whatman, Inc. | Neonate newborn screening card demographic and filter portion may be detached | Process control | The firm, Whatman, issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated October 29, 2009, to the customer. The letter described the product, problem and action to be taken by the customer. The letter instructed the customer to: inspect the card and ensure the 903 Paper Portion is firmly and completely fixed to the Demographic portio, prior to recording patient demographic details on the card and obtaining a patient sample; discard the card, obtain a new card and inspect as above, if the 903 Paper Portion of the card partially or completely detaches from the Demographic portion of the card; and to please staple the card stub at two points as specified in the illustration, to prevent any opportunity for 903 paper detachment from the demographic portions.If Whatman customers/users have any questions in relation to this customer notification, please contact, either Product Manager at +1 732 885 6570,email: lili.lee@ge.com or Quality Assurance at +1 207 459 7557, email: heidi.kuell@ge.com. |  | 488000 | 488,000 cards | Nationwide distribution: CA | Open, Classified |  |  | 488000 | Not_Computer
Z-2382-2010 | 56315 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Impella 5.0 Circulatory Support System |  |  | July 05, 2010 | 2010 | Abiomed, Inc. | Pump may separate at the motor housing while device is explanted | Process control | ABIOMED issued the recall notification letter on July 5 , 2010 to the one affected customer (France) and requested that they segregate and return any unused pumps. The pumps in Canada were already used.If you have any questions, contact +49 241 8860-0 or e-mail europe@abiomed.com. |  | 10 | 10 units | International only: Canada and France | Terminated | September 10, 2010 | 67 | 10 | Not_Computer
Z-2383-2010 | 56488 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | 3M Red Dot Monitoring Electrode, Catalog #2560 &amp; Novaplus Universal Monitoring Electrode #V2560, used for ECG monitoring. |  |  | July 27, 2010 | 2010 | 3M Company / Medical Division | The sensing element of the 3M Red Dot Monitoring Electrode, Catalog # 2560 and Novaplus Universal Monitoring Electrode #V2560 may have corrosion. In an emergency situation where defibrillation is necessary and electrodes are already in use, the electrodes may malfunction and provide a delayed or no trace response. | Process control | 3M sent an "Urgent Medical Device Recall" letters and a "Distributor Product Recall Form" dated July 29, 2010 to customers and distributors of the 3M and Novaplus electrodes. The letters described the 3M Red Dot and Novaplus Universal products, problems and actions to be taken by customers. Customers were instructed to examine their inventory and quarantine the affected lots pending return to the firm. After completing and returning the enclosed Customer Product Recall Form to 3M, customers will be issued a Return Authorization Number and provided with instructions for return of the affected devices. Distributors were instructed to notify customers to the end user level and have them return affected lots to the distributor for forwarding to 3M, or forward the customer letter to end-users to instruct them on how to contact 3M for recovery of the device. The 3M Customer Helpline can be contacted at: 1-800-228-3957. |  | 553150 | 11,063 cases (20 pouches/case, 50 electrodes/pouch) | Nationwide Distribution: States of CA, FL, IL, KS, KY, MI, MO, NE, NJ, NY, OH, PA, RI, SD, TN, VA, and WI. | Terminated | April 19, 2012 | 632 | 553150 | I/O
Z-2385-2010 | 54460 | No | JMO | Iron-binding capacity test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension IBCT Flex Reagent Cartridge. The total (IBCT) method used on the Dimension system is an in vitro diagnostic test intended to quantitatively measure Total Iron Binding capacity in human serum and plasma. |  |  | December 16, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Test produces falsely elevated IBCT results and Abnormal Reaction test report messages on heparinized plasma samples. | Nonconforming Material/Component | Siemens issued an "Urgent Field Safety Notice" letter dated December 2009 to users identifying the issue and affected product. Customers were instructed not to use heparinized plasma samples for IBCT testing. IBCT Instructions for Use (IFU) has been updated to reflect that heparinized plasma is no longer an acceptable sample type. Beginning with lot ED0341, the new IFU will be packaged in the carton with the IBCT FlexÂż reagent cartridges. Consignees were requested to forward the notification letter to anyone to whom the product was distributed. The Siemens Technical Solutions Center can be contacted at 800-441-9250. |  |  | unknown - all lots involved | Worldwide Distribution: USA (all states) and the countries of Canada, Switzerland, New Zealand, Australia, Japan, and South Korea. | Terminated | October 25, 2010 | 313 |  | Not_Computer
Z-2386-2010 | 54624 | No | LAN | Digoxin test system. | Toxicology | Toxicology | 510(k) | ONLINE TDM N-acetylprocainamide, Part number 04642520190 N- acetylprocainamide reagent Roche Diagnostics Corporation Indianapolis, IN. |  |  | February 05, 2010 | 2010 | Roche Diagnostics Operations, Inc. | The N-acetylprocainamide reagent lot 14456200 shows an incorrect expiration date of 7/31/2011. The correct expiration date for the lot is 1/31/2011. | Error in labeling | Roche issued an "Urgent Medical Device Removal" letter dated February 2010 to consignees identifying the issue and affected product. Customers were requested to complete the attached form and return in to the firm to obtain replacement product. After receipt of replacement reagent, the customer should discard the affected product product. If unaffected lots are not available, customers may use the affected lot until replacement product is received. Use of the affected product before the corrected expiration date of January 31, 2011 should not produce erroneous test results. Roche Diagnostics Technical Support can be contacted at 1-800-428-2336. |  | 40 | 40 kits | Nationwide Distribution: USA states of AL, CT, NY, PA, and TX. | Terminated | November 01, 2010 | 269 | 40 | Not_Computer
Z-2332-2010 | 56162 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | BagEasy manual resuscitation devices are used for patients requiring complete or intermittent ventilatory support. It can provide positive pressure ventilation or oxygen for the spontaneously breathing patient with a face mask or through an artificial airway. |  |  | June 22, 2010 | 2010 | Westmed Inc | This recall was initiated after Westmed, Inc. became aware of a potential for disconnection at the patient port retention ring assembly of the Bag Easy manual resuscitation device. | Nonconforming Material/Component | The firm, Westmed, sent a letter dated June 21, 2010, to all customers. The letter describes the product, problem and action to be taken by customer. The customers were instructed to cease distribution immediately, notify any facilities that they may have distributed the product identified of the recall, call Customer Service Manager at 1-800-975-7987 x 1223 and/or email kpeterson@westmedinc.com to receive Return Authorization to return identified product, and complete and return the enclosed Bag Easy Recall Reconciliation Checklist.If you have any question, contact Customer Service at 1-800-975-7987 or 1-520-294-7987. |  | 22145 | 22145 | Worldwide distribution. | Terminated | March 22, 2011 | 273 | 24879 | Not_Computer
Z-2381-2010 | 56523 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira Inc. Symbiq One- and Two- Channel Infuser, List Numbers 16026 and 16027 respectively. |  |  | March 26, 2010 | 2010 | Hospira Inc | If the administrator set is removed prior to the cassette carriage fully opening without closing the slide/roller clamp, unrestricted flow and delivery may result. | Device Design | Hospitra Urgent Device Correction Notification mailed on 03/26/2010 via federal express. Tracking done by delivery tracking. Distribution of product continues and has not stopped. Additional notification was issued in 09/2010. A software correction is planned. |  | 28485 | 28485 units | Worldwide distribution: USA, Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Puerto Rico, Singapore and the UAE. | Terminated | March 30, 2012 | 735 | 28485 | Other | Treatment/Delivery/Therapy | Software update | Software Update
Z-2387-2010 | 54743 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista Chemistry 2 Calibrator, Catalog number KC120. |  |  | February 26, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | There is incorrect barcode information, which can potentially cause the Alert Message "Calibrator Insert Missing" to be displayed. | Employee error | The recalling firm issued an Urgent field Safety Notice, dated February 2010, to all affected customers informing them of the problem. Customers have been told to run calibrators from cups when calibrating or request a replacement lot. The notification should be forwarded to that this product was further distributed to. Siemens Technical Solutions Center should be contacted at 800-441-9250, if they are any technical questions. |  | 1310 | 1310 cartons | Worldwide Distribution -- USA, Austria, Germany, Netherlands, Italy, France, Denmark, Portugal, Canada, South Korea, Malaysia, Australia, and New Zealand. | Terminated | October 13, 2010 | 229 | 1310 | Not_Computer
Z-2336-2010 | 55852 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Mayfield Infinity XR2 Base Unit, model A-2079, manufactured by Integra LifeSciences Corporation. The MAYFIELD Infinity XR2 Base Unit (A-2079) is an operating room table accessory intended to be used to support a patient during diagnostic examination and/or surgical procedures where a rigid support between surgical table and headrest, or skull clamp is necessary, positional freedom is required and where x-ray imaging modalities will be used. |  |  | May 19, 2010 | 2010 | Integra LifeSciences Corporation | Storage case incorrectly labeled. The labels were for the Infinity XR2 Skull Clamp storage containers but were inadvertently placed on the Infinity XR2 Base Unit storage cases. | Labeling mix-ups | The firm, Integra Life Sciences Corporation,contacted their customers via telephone and Field Correction Letter on May 19, 2010 to notify their customers of the problem. The firm indicated that their sales representatives will visit the consignees affected, correct the case labeles according to the specific instructions indicated, revise their account tracebility records to reflect the labeling change, then, sign and return the Acknowledgement of Reciept and Return Form for confirmation. For questions regarding the correction, please contact Integra Life Sciences Corporation at (877) 44-1114 or (513) 533-7923. |  | 13 | 13 Units | Product was shipped to the following states: IN, MO, NJ, TX and VA. Product was also sold in Belgium, Canada and New Zealand. | Terminated | October 27, 2011 | 526 | 13 | Not_Computer
Z-2393-2010 | 56333 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Invacare AC-Powered Adjustable Hospital Bed. |  |  | October 16, 2006 | 2006 | Invacare Corporation | Invacare Corporation, Florida issued a device correction in October 2006 for the IVC Bed Foot Actuator used on the IVC Bed Series, Models 5890 IVC and 5490 IVC. The actuators were not in the fully extended position and could not be mounted onto the beds. All affected consignees were notified of this correction in October 2006. | Nonconforming Material/Component | Invacare Corporation issued a notice of an Urgent Field Correction to Consignees on 10/18/2006. Consignees were instructed to check the serial number of their units to see if they were potentially affected. If so, the firm provided inspection and adjustment instructions of the foot actuator to prevent a likelihood of injury. After adjustment, the consignees were requested to complete and return the enclosed tracking sheet acknowledging receipt of the recall notice and correction of the bed(s). For questions or assistance regarding this correction, please contact the Customer Relationship Management Group at (800) 333-6900, or via e-mail at CRMGroup@Invacare.com. |  | 731 | 731 | Product was distributed throughout the U.S. | Terminated | September 13, 2010 | 1428 | 731 | Not_Computer
Z-2389-2010 | 56424 | No | CGX | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SYNCHRON Systems Modular Creatinine (CREm) Reagent Kit, Part Number: 472525 |  |  | August 02, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed reports that black / brown particulates were observed in some bottles of the Creatinine Alkaline Buffer (Buffer lot M911427 and M911518) in modular Creatinine (CREm) reagent kit lots M911428 and M911519. Beckman Coulter have had reports of erroneous results when using the lots of Creatinine Alkaline Buffer listed above. Singlecreatinine d | Pending | The firm, Beckman Coulter, sent a "URGENT: PRODUCT CORRECTIVE ACTION" letter dated August 3, 2010, to all customers. The letter describes the product, problem, and action to be taken by customer. The customers were instructed to: (1) Please discontinue use of Creatinine (CREm) reagent kit lots M911428 and M911519 and discard appropriately.(2) Please contact Beckman Coulter at (800) 526-3821 in the United States or (800) 463-7828 in Canada to discussCREm kit M911428 and M911519 replacement options. Outside the United States and Canada, please contactyour local Beckman Coulter Representative.(3) Please complete and return the enclosed response form within 10 days so that we may maintain our records.(4) Please share this information with their laboratory staff and retain this notification as part of their Quality Systemdocumentation. If you have forwarded the affected products listed above to another laboratory, please provide a copy of this letter to them.If customers have any questions regarding this Product Corrective Action letter, they were instructed to please contact Beckman Coulter's Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/Customer Support/SupportRequests, or call (800) 854-3633 in the United States and Canada. |  | 3106 | 3106 units | Worldwide Distribution - USA including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, KS, LA, MA, ME, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV, and WY and country of Canada | Terminated | July 02, 2012 | 700 | 3106 | Not_Computer
Z-2390-2010 | 56464 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira brand Symbiq One-Channel Infuser |  |  | February 22, 2010 | 2010 | Hospira Inc | The clamp may not secure to the IV pole and the pump may slip or fall off. The pump may fall on a person or pull the tubing out of the IV container. | Device Design | Hospira issued an "Urgent: Device Field Correction" Letter dated February 22, 2010 to all consignees, identifying the reason for the recall and affected product. Customers were informed that they will be contacted by a Hospira representative to replace the affected pole clamps. Consignees can contact Hospira Technical Support Operations at 1-800-241-4002. |  |  |  | Worldwide Distribution: USA (all states and Puerto Rico) and the countries of Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Singapore. and the UAE. | Terminated | April 01, 2011 | 403 |  | Not_Computer
Z-2392-2010 | 56133 | No | GEN | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BIOMET MICROFIXATION REF: 02-0461 REES INSULATED FORCEPS, 21CM, (8.25") (L), STAINLESS STEEL, NYLON, ***NON-STERILE PRODUCT***. The Rees Insulated Forceps are used during surgery to grasp, dissect and coagulate selected tissue. |  |  | June 25, 2010 | 2010 | Biomet Microfixation, Inc. | The forceps insulation may have cracks or breaks in the insulation that may cause burns to patients when used. | Manufacturing material removal | Biomet Microfixation issued a letter dated July 2010 to their consignees. Each consignee was ask to discontinue use of the affected product and return the product back to the company for credit. Biomet asked their customers to complete a "Inventory Reconciliation" form and fax the completed form to Biomet Microfixation. If the consignee has further distributed the product, the consignee is asked to contact their customers. Biomet Microfixation can be contacted at 904 741-4400, ext 468. |  | 130 | 130 units | Worldwide Distribution: USA including states of CA, FL, GA, ID, IL, IN, MA, MD, MS, NC, NH, NJ, NV, NY, OH, PA, TX, VA, WA, and WI, and countries of Norway, Canada, and Mexico. | Terminated | June 21, 2011 | 361 | 130 | Not_Computer
Z-2388-2010 | 55359 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GLUCm used on Synchron LX and UniCel DxC Synchron Clinical Systems, Part Number: 472500 ( ynchron LX20, LX20 PRO, LXI 725, UniCel DxC 600, UniCel DxC 600 PRO, UniCel DxC 600i, UniCel DxC 800, UniCel DxC 800 PRO, UniCel DxC 880i, UniCel DxC 680i, UniCel DxC 660i, and UniCel DxC860i ), |  |  | April 15, 2010 | 2010 | Beckman Coulter Inc. | UniCel DxC and Synchron LX instruments produced inaccurate (low) GLUCm results. False low results could affect or delay diagnosis or treatment. | Other | Beckman Coulter issued an "Urgent: Product Corrective Action" (PCA) letter dated April 2010 to all consignees, identifying the reason for the recall, the affected lots, and the potential for inaccurate GLUCm results due to bubble retention and short samples. Customers were instructed to:1) establish specific protocols for samples from defined patient populations or sample types, which could include running specimens using GLUCm in duplicate or using cartridge glucose.2) To reduce the possibility of inaccurate results, ensure that all elements are properly installed and secured, the sample probe is adequately tightened, and proper sample volume and integrity are maintained. A copy of the PCA letter will also be included in all new instrument ship kits.Consignees were instructed to complete and return an enclosed response form. Consignees can contact Beckman Coulter at 1-973-956-5300 . |  | 84186 | 84,186 units (79,071 in US &amp; 5,115 in Canada) | Worldwide Distribution: USA and Canada | Terminated | December 31, 2013 | 1356 | 84186 | Not_Computer
Z-2398-2010 | 55700 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Clinical Chemistry Phosphorus assay |  |  | May 04, 2010 | 2010 | Abbott Laboratories Inc. | Product may give inaccurate results and aspiration errors with reagent configuration, including results outside of linear range &amp; calibration failure. | Other | Abbott Laboratories issued a "Product Recall: Immediate Action Required" letter dated May 4, 2010 to consignees. The letter described the product, problem and action to be taken by the customer. The letter instructed the customer to:1) Identify affected product in inventory2) Discontinue use and destroy any remaining inventory of the affected lot: 7D71-213) Complete and return an enclosed Customer Reply Form4) Order alternative kit sizes for Clinical Chemistry Phosphorus: LN 7D71-31 or 7D71-225) For AEROSET Systems, manually update parameters as described in the letter.6) For ARCHITECT cSystems, implement alternative kit sizes without system changes.7) Provide a copy of the letter to other laboratories that have been distributed the affected product.US customers can contact Abbott Customer Service at 1-877-4ABBOTT. Customers outside the U&gt;S&gt; are instructed to contact their local area Customer Service. |  | 3136 | 3136 | Worldwide Distribution: US, including Puerto Rico, and countries of Germany, New Zealand, Trinidad/Tobago, Cayman Island, Hong Kong, Chile, Curacao, Jamaica, Australia, and Bahamas. | Terminated | April 05, 2012 | 702 | 3136 | Not_Computer
Z-2395-2010 | 55478 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas 2020 Insulin Infusion Pumps |  |  | March 18, 2010 | 2010 | Animas Corporation | sensor could be damaged | Equipment maintenance | An Email was sent by the firm, Animas, to US patients and health care providers where email addresses were on file. Animas sent two "URGENT: MEDICAL DEVICE CORRECTION " letters one dated March 18, 2010, to all US customers, HCPs and distributors of One Touch Ping Glucose Management System and the Animas 2020 Insulin Infusion Pump and the other letter to all US patients, HCPs and distributors of Animas 2020 Insulin Infusion Pump, Animas IR 1200 and Animas IR 1250 (1200 plus). Animas also notified all global distributors and affiliates and sent the letter to customers and health care professional in those regions. The letters describe the product, problem and action to be taken by the customers. The customers were instructed to follow their User Guide instructions; Pay attention to all pump warnings; NEVER rewind the pump motor, load the insulin cartridge, or prime their infusion set while connected, and fill out the enclosed reply card and send it back to Animas as soon as possible.Should you have any questions or concerns, please call the Animas Customer Support Department at 1-866-949-1525 or call toll free number on back of pump at 1-877-937-7867. If you are traveling outside of the U.S., please call 610-644-8990 and press option 1. |  |  |  | Worldwide distribution: USA including states of CA,FL,LA, MA, MI, NY, OH, PA,TX, UT, VA and WA, and countries including Canada, United Kingdom, Finland, Germany, France, Hungary, Puerto Rico, Australia, New Zealand, Israel, Sweden, Czech Republic, Spain, Mexico and Norway. | Terminated | October 27, 2010 | 223 |  | I/O
Z-2349-2010 | 56330 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Picasso/Picasso Lite lasers |  |  | April 15, 2010 | 2010 | AMD Lasers, LLC | Product is noncompliant with the Federal laser product performance standard in certain requirements. The ON/OFF key on some of the Picasso/Picasso Lite lasers could be removed with force while the key was in the ON position. | Radiation Control for Health and Safety Act | Letters, dated April 30, 2010, addressed to Dental Professionals were sent identifying the affected product and reason for the recall. The letter asked customers to check their units and, if applicable, AMD Lasers would replace the key assembly on their lasers. The service would be provided free of charge including round trip shipping charges to the corporate office. Customers are to call to schedule the service and obtain a RMA number and shipping instructions. Questions should be directed to 1-866-999-2635 or support@amdlasers.com. |  | 3000 | 3,000 units | Worldwide Distribution | Terminated | April 21, 2011 | 371 | 3000 | Not_Computer
Z-2400-2010 | 55610 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus CS Acute Care Troponin I TestPak ( Troponin I assay) |  |  | April 13, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | False elevated Stratus CS CTNI results | Process control | Recalling firm issued " Urgent Field Safety Notice" dated April 2010, informing users of the problem and to discuss with laboratory directors regarding the need to review previous test results, conduct patient follow up, and/or repeat testing for tests conducted on or after Nov 1, 2009. |  | 7500 | 7500 cartons | Nationwide Distribution | Terminated | October 25, 2010 | 195 | 7500 | Not_Computer
Z-2399-2010 | 56271 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Peregrine 20 GA Endo Ocular Adjustable Laser Probe |  |  | July 19, 2010 | 2010 | Peregrine Surgical Ltd | Product did not have proper 510(k) clearance | Other | A recall letter is being sent to all affected accounts with instructions to return the recalled product to the recalling firm. |  | 132 | 132 boxes | Product was distributed to hospitals and eye surgery centers in Florida, Maryland, Missouri, North Carolina, Ohio, Pennsylvania, Texas. | Terminated | September 30, 2010 | 73 | 132 | Not_Computer
Z-2401-2010 | 56268 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Maquet HLM Tubing Pack; Bioline Coating |  |  | July 13, 2010 | 2010 | Maquet Inc. | The color of the tape on the arterial and the venous line were placed incorrectly, causing the connection of the arterial and the venous lines to be interchanged. | Process control | Recall notification letters were sent by Federal Express on July 13, 2010 to affected account and also an e-mail to the Maquet field representative who services the account. |  | 5 | 5 units | Distribution: Pennsylvania | Terminated | August 10, 2011 | 393 | 5 | Not_Computer
Z-2114-2010 | 56370 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Microwave Ovens, Model Number DG-2 |  |  | May 12, 2010 | 2010 | Duke Manufacturing Co. | The oven did not meet Federal Performance Standard for certain requirements. | Radiation Control for Health and Safety Act | Letters were sent to Duke Manufacturing Speed Oven Owners. The letters identified the affected product and stated the reason for recall. The letter also states that the firm instituted a repair program on all DG-2 Speed Ovens. The firm will repair of replace all affected products at no charge to the customer. Customer Service personnel will contact customers directly and coordinate the schedule of the required repairs. |  | 62 | 62 units | Nationwide Distribution | Terminated | April 17, 2012 | 706 | 62 | Not_Computer
Z-2402-2010 | 56336 | cover, mattress (medical purposes) | FMW | Mattress cover for medical purposes. | General Hospital | General Hospital | 510(K) Exempt | CareGuard Mattress Models CG10180 and CG10180A used with IVC and ValueCare Series of manual, semi-electric and electric beds. |  |  | October 11, 2007 | 2007 | Invacare Corporation | In October 2007 Invacare Corporation, Florida recalled CareGuard Mattresses, Models CG10180 and CG10180A, associated with IVC Bed Series Full Electric Beds, Semi-Electric Beds and Manual Beds and the ValueCare Bed Series Semi Electric Beds, when used with Invacare bed rails. Affected bed packages are BED24IVC, BED25IVC, BED24CA, and BED25CA. The mattresses when used in combination with the IVC a | Device Design | Consignees were notified of the recall by letter dated October 15, 2007 and entitled "CareGuard Mattresses Provided Since September 2006 In combination with IVC Bed Series Full Electric Beds, Semi-Electric Beds and Manual Beds, and the Value Care Bed Series Semi Electric Bed, and Full and Half-Length Rails". Depending on the mattress usage the consignees were instructed to either: A: 1) Move the bed rail cross bar from the 5th and 6th spring location on the head section of the bed to the 6th and 7th spring location on the head section of the bed in accordance with the enclosed instructions and document your actions on the enclosed tracking sheet. 2) Return of the Product Tracking Report indicating disposition of the recalled product and acknowledgement of receipt of the recall package. 3) The tracking sheet must be sent to Invacare by January 15, 2008 in order for you to receive the $10 credit. or B) 1) contact Customer Service and request part No. 6022, which is a kit that contains a 2" convoluted foam pad, warning label and installation instructions. Testing has shown that the addition of the foam pad on a new mattress with Invacare half rails / bed combination meets the Guidance requirements." 2) When you have completed placing the label and foam topper on all of your affected mattresses, return the tracking sheet to us and we will credit your account $10 for each mattress you have listed. the tracking sheet must be sent to Invacare by January 15, 2008 in order for you to receive the $10 credi |  | 19143 | 19143 | Nationwide Distribution | Terminated | June 08, 2015 | 2797 | 19143 | Not_Computer
Z-2403-2010 | 56335 | cover, mattress (medical purposes) | FMW | Mattress cover for medical purposes. | General Hospital | General Hospital | 510(K) Exempt | Mattress Model 5185 and 5185XL. |  |  | August 29, 2007 | 2007 | Invacare Corporation | In September 2007 Invacare Corporation recalled the innerspring mattresses associated with Invacare's IVC Bed Series Full Electric Beds, Semi-Electric Beds and Manual Beds and the ValueCare Bed Series Semi Electric Bed. These mattresses were manufactured between June 28, 2007 and July 12, 2007. These mattresses when used in combination with the IVC and ValueCare Series of manual, semi-electric a | Device Design | Consignees of the product were notified by letter dated 09/05/2007. Consignees were instructed to destroy the affected mattresses and to return the completed confirmation of the destruction form to Invacare. |  | 4896 | 4896 | Nationwide Distribution | Terminated | September 13, 2010 | 1111 | 4896 | Not_Computer
Z-2405-2010 | 46217 | No | IXR | Radiologic table. | Radiology | Radiology | 510(K) Exempt | EasyDiagnost Classic has a radiographic tilting table. This system has a cassette film option and it does not have a Digital Spot Imaging (DSI) system. |  |  | January 04, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Potential for the tower (holding the serial changer and/or image intensifier) to slide unexpectedly and potentially injure a patient or user in its path. | Other | On 1/4/08, the firm issued the URGENT PRODUCT CORRECTION letter dated 1/3/08 to their consignees via certified mail, return receipt requested. The letter informs consignees the issue of the safety switch and counter weight steel band on the Diagnost 55/66/76Plus, Easy Diagnost, and Easy Diagnost Eleva X-ray systems. However, the Easy Diagnost Eleva X-ray system was distributed and applied only to foreign consignees outside US.If the consignees hear the primary steel band clicking/rattling noise while tilting the table to 90 degrees table or positioning the serial changer, the consignees are instructed to tilt the table base towards the 0 degree tilting position and do not use the tilting drive movement until the mandatory action by a Philips service representative has been performed. The Field Change Order (FCO) would be implemented on 01/23/08 and expected to be completed by July 2008. The firm's representatives will visit sites and will inspect all systems where the switch has been replaced or adjusted for proper operation of the safety sensor and switch. Two steel bands will be inspected and replaced if any abnormity is noted.In the mean time, consignees should contact Philips Call Center at 800-722-9377, #5, #2 and reference the FCO 70600029 with any questions. |  | 1 | 1 unit | Nationwide Distribution | Terminated | September 13, 2010 | 983 | 38 | Not_Computer
Z-2407-2010 | 56145 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension Tacrolimus Flex Reagent Cartridge, manufactured by Siemens Healthcare Diagnostics Inc., Newark, Delaware. |  |  | June 15, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Low patient sample recovery results in all affected lots, with possibility of inappropriate increase in tacrolimus dosage based on false low result. | Other | Siemens issued an "Urgent Field Safety Notice" dated June, 2010 to all affected customers. Customers were requested to immediately discontinue use and discard any remaining inventory of the TACR Flex reagent lots DC0306, CA0320, DA0334 and GA0348. The letter contains the following instructions: 1) Immediately discontinue use of the affected lots; 2) Discard all remaining inventory of the affected lots; 3) Contact Siemens Healthcare Diagnostics Technical Solutions Center for replacement product; 4) Complete and return the Field Correction Effectiveness Check by fax to 302 631-8467;4) Siemens is recommending that the content of the letter should be discussed with the laboratory director. If unexpectedly low results are obtained on a patient sample, caution should be exercised before decisions regarding therapeutic dosage adjustments are made. As stated in the package insert, "Confirmation of unexpected or atypical results by an alternative methodology is recommended prior to any adjustments in tacrolimus dosage. Also individual tacrolimus values cannot be used as the sole indicator for making changes in the treatment regimen. Each patient should be thoroughly evaluated clinically before treatment adjustments are made."Siemens Technical Solutions Center may be contacted at 800-441-9250 |  | 4580 | 4580 Cartons | Worldwide Distribution to USA states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, KS, LA, MI, MO, MS, NC, NE, NH, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA and WI, and countries of Austria, Belgium, Canary Islands, Egypt, France, Germany, Great Britain, Hungary, Italy, The Netherlands, Poland, Portugal, Romania, Saudi Arabia, Slovakia, Spain, Sweden, United Arab Emirates, Argentina, Brazil, Canada, Columbia, El Salvador, Venezuela, Australia, Hong Kong, India, Japan, Malaysia, New Zealand, South Korea, Singapore, Taiwan and China. | Terminated | December 16, 2010 | 184 | 4580 | Not_Computer
Z-2420-2010 | 55923 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Stryker Bur Guard REF 2296-301, Stryker Instruments, Kalamazoo, MI (USA) |  |  | June 02, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | The reason for the field action is that the inner diameter of the Bur Guards may be out of specification. If there is excess material on the inside of the Bur Guard at the location where the guard tapers, it may come into contact with the Nose Cone of the drill. This could result in an injury or burn to the patient and/or user. | Device Design | Stryker Instruments sent an Urgent Medical Device Recall Notification dated June 10, 2010, to their customers. The letter described the product the problem and the action to be taken by the customers. The customers were instructed to:1) immediately review this recall notification.2) forward this letter and the attached Business Reply Form (BRF) to all affected locations. 3) Immediately check and quarantine any affected product found.4) Destroy all affected product by breaking off at least 3 of the tabs on the bur guard and then dispose of the affected product in normal trash.5) Complete and sign the BRF and fax to Stryker Instruments Regulatory Department, 866-521-2762.NOTE: Upon receipt of the BRF, a replacement bur guard will be sent to the customer's account for each affected bur guard destroyed.For questions regarding this recall please contact Stryker Instruments:1-800-800-4236, ext. 3808jennifer.mars@stryker.com |  | 10982 | 10982 | Worldwide Distribution: US AND COUNTRIES OF SCANDINAVIA, AUSTRALIA, CANADA, CHILE, CHINA, BRAZIL, FRANCE, GERMANY, GREECE, SOUTH AFRICA, ITALY, JAPAN, MEXICO, SWITZERLAND, PORTUGAL, SINGAPORE and UK. | Terminated | April 21, 2011 | 323 | 10982 | Not_Computer
Z-2414-2010 | 55876 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen complete knee solution MIS total knee procedure stemmed tibial component fixed bearing precoat size 7 sterile, REF 00-5950-057-01, Zimmer Inc. Warsaw, IN. |  |  | April 26, 2010 | 2010 | Zimmer Inc. | The firm has received complaints of loosening of the implanted device requiring revision surgery. There have been 114 MDRs filed all reported that the device loosened and the patient required additional surgery to replace the device. | Other | The firm, Zimmer, Inc., sent an "URGENT DEVICE CORRECTION" letter dated April 26, 2010, to all customers. The letter described the product, problem and action to be taken by customers. The customers were instructed to destroy or disregard all previous versions of the surgical technique, review the attached revised surgical technique to familiarize yourself with the modifications and warnings, and complete and return the Health Care Professional Dcoumentation Form via fax to Zimmer, Inc. at (574) 372-4265. (Note: Zimmer made a modification to the surgical technique and instructions for use. They added a warning to fully cement and pressurize the anterior and posterior surfaces of the tibial component, and to strongly recommend the use of a drop down stem extension. This change was completed and approved on April 23, 2010.) For questions please contact Zimmer at 1-877-946-2761. |  | 68384 | 68,384 all products | Worldwide distribution: USA including states of AZ, IL, IN, KY, MO, NE, NM, NY, OH, TX, VA, WA, and WI; and countries of Austria, Belgium, France, Germany, Italy, Netherlands, Spain, Sweden, Switzerland, UK, Ausralia, China, Hong Kong, India, Japan, Korea, Malaysia, Nicaragua, Singapore, Tawain and Thailand. | Terminated | March 28, 2011 | 336 | 68384 | Not_Computer
Z-2422-2010 | 47477 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | DigitalDiagnost X-ray System is a stationary Diagnostic X-ray using Digital Solid State Imaging Detector. It is identified with Part Number 4512 201 04752 Software Version 1.5.1 and Part Number 4512 201 04753 Software Version 1.5.2. |  |  | December 13, 2007 | 2007 | Philips Medical Systems North America Co. Phillips | When using DICOM print functionality, the printed images may contain data for the wrong patient. | Software design | Philips Healthcare issued an "URGENT PRODUCT CORRECTION " letter dated March 10, 2008 to their consignees. The letter described the product, problem and action to be taken by the customer. The letter instructed the customer to: 1) Validate that image and patient information are correct on hard copy prints.2) Reprint images that have mixed patient information, and re-check that image and patient information are correct.3) Do not use printed images if the information is incorrect.The firm will issue the Field Change Order (FCO) 71200025 and the field service engineers will visit each consignee and conduct a software upgrade to solves the printing problem. The consignees may contact the Philips Call Center at 1-800-772-9377, #5 #2 and reference "FCO 71200025". |  | 119 | 119 units | Nationwide Distribution: including the states of AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MN, MO, MS, NC, ND, NE, NJ, NM, NY, OH, OK, PA, SC, TN, TX, UT, VT, WA, and WI. | Terminated | September 14, 2010 | 1006 | 119 | Software | Output/Calculation | Software update | Software Update
Z-2421-2010 | 56422 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Sapphire Probe Detector PCB Assembly (also called the Aspiration Bottom Sensor) used on the Cell-Dyn Sapphire Hematology Analyzer, Manufactured by Abbott Laboratories, Santa Clara, CA |  |  | April 05, 2010 | 2010 | Abbott Laboratories | Under certain conditions in closed mode of operations, the Aspiration Probe Detector PCB assembly may fail to operate properly and provide inaccurate results. | Other | Abbott Diagnostics Division of Santa Clara, CA initiated a field correction of the CELL-DYN Sapphire Probe Detector Assembly used on the CELL-DYN Sapphire Hemotology Analyzer, on April 5, 2010. The correction is being addressed through field service representatives and a mandatory Technical Service Bulletin (TSB). The TSB requires inspection, testing, and potential replacement of the affected assembly. The firm will not be issuing a seperate notification. For questions regarding this correction please contact (408) 982-4800, or fax (408) 982-4863. |  | 926 | 926 units | Product was distributed to 926 retailers throughout the US and to argentina, Australia, Brazil, Canada, Chile, China, Columbia, Germany, Hong King, Japan, Mexico, New Zealand, Panama, Paraguay, Peru, Singapore, South Korea, Taiwan and Venezuela. | Terminated | May 17, 2011 | 407 | 926 | Hardware | Output/Calculation | Replace | Remove or Replace
Z-2424-2010 | 56320 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | NeoMarkers Rabbit Monoclonal anti-Human Cyclin D1 Antibody (Clone SP4), immunochemistry Reagent and Kit, For In Vitro Diagnostic use, manufactured by Lab Vision Products, Fremont, CA, a subsidiary of ThermoFisher Scientific |  |  | March 05, 2010 | 2010 | Lab Vision Corporation | Two lots may be contaminated with Ki-67 antibody, which could cause abnormal staining patterns, and produce false positive results. | Material/Component Contamination | Thermo Fisher Scientific issued an "Advisory Notice" letter dated March 5, 2010 to all domestic customers and international distributors. The letter described the product, problem and action to be taken. Customers were instructed to:1) discard the affected lots, and2) retest specimens using a replacement lot supplied by the firm, and 3) complete and return the Confirmation of Receipt Form via fax to 1-510-979-5489International distributors were instructed to:1) place the affected lots and replace with a replacement lot supplied by the firm, and 2) provide customers with copies of the customer letter, and3) complete and return the Confirmation of Receipt Form via fax to 1-510-979-5489For questions, contact Thermo Fisher Scientific at 1-800-828-1628 or 1-510-771-1595. |  | 103 | 103 units | Worldwide Distribution: USA, including the states of CA, SC, IL NY, WI, PA, MN, AL, WA, MA, and KS, and the coutnries of Austria, Australia, Denmark, Finland, France, Germany, Hong Kong, Italy, Japan, Malaysia, the Netherlands, Norway, Pakistan, Poland, Portugal, Russia, the UK, and Turkey. | Terminated | March 21, 2011 | 381 | 103 | Not_Computer
Z-2423-2010 | 55763 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics, Inc., Dimension Vista Intelligent Laboratory System with Alpha Pinnacle 10kVA UPS External Battery Pack. Part numbers: 070200001, 070200011, 070200021; Dimension Vista 1500 and 3000T instruments. |  |  | May 05, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Defective batteries used in the Alpha Pinnacle P10 UPS External Battery Pack may overheat and cause a strong sulfuric odor to be released and a corrosive sulfuric acid solution leak to occur. | Nonconforming Material/Component | Siemens issued an Urgent Field Safety Notice in May 2010 to customers with information on what to do if the battery leaks, overheats, or vents an odor. Firm also issued a Support Bulletin to field service personnel to inspect and replace affected units. |  | 228 | 228 units | Worldwide distribution: USA, Australia, Austria, Belgium, Canada, France, Germany, Italy, Japan, Malaysia, Netherlands, New Zealand, Portugal, Slovenia, Spain, and Switzerland. | Terminated | October 25, 2010 | 173 | 228 | Battery
Z-2419-2010 | 54682 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista ECREA Flex Reagent Cartridge |  |  | February 12, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | The ECREA test results may be falsely elevated or depressed. | Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notice dated February 2010 to all affected customers. The letter described the product, the problem, and the action to be taken by the customers. The customers were instructed to immediately discontinue the use of and discard the ECREA Flex reagent cartridge lot 09224AA. Customers are to contact their Siemens Customer Service Center at 800-241-0420 to report their additional no-charge replacement needs, and complete and return the enclosed Field Correction Effectiveness Check Form via Fax at 302-631-8467.For any technical questions regarding this Recall contact Siemens Technical Solutions Center at 800-441-9250. |  | 324 | 324 | Worldwide Distribution - USA, including the states of AZ, CA, FL, IA, IL, MT, NC, NY, OR, PA, TX, and WI and countries of Belgium, Italy, Canada, Germany and the Netherlands. | Terminated | October 13, 2010 | 243 | 324 | Not_Computer
Z-2430-2010 | 49489 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Vista Nuclear Magnetic Resonance Imaging System, Picker International, Inc. World Headquarters 595 Miner Road Cleveland, Ohio 44143, U.S.A.. |  |  | September 19, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Water may collect and freeze in the vent pipe. This can rupture the magnet vessel during a cryogen transfer, and release helium into the MRI suite. | Other | Philips issued an Urgent Correction letter dated 09/19/2008 to their consignees, identifying the reason for the recall and the affected products. The firm issued the Field Change Order (FCO) 78100232. Field service engineers will visit each consignee and conduct an inspection of the magnets venting system of Edge, Vista, Eclipse, Polaris, Infinion, and Panorama 0.6T MRI systems. The consignees should contact the Philips Call Center at 1-800-772-9377, #5 #2 and reference "FCO 78100232". |  | 16 | 16 units | Nationwide Distribution: USA., including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WY. There is no distribution to foreign consignees. | Terminated | September 17, 2010 | 728 | 374 | Not_Computer
Z-2428-2010 | 56564 | No | GDS | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Safen' simple Ostomy Appliance Deodorant (8 oz , 1 oz package sizes) |  |  | August 16, 2010 | 2010 | BioMed Laboratories | The product may contain the mold Aspergillus ssp. | Material/Component Contamination | Biomed Laboratories, Inc. of Garland, TX is conducting a recall of the Reliamed Odor Eliminator Drops, Securi-T Ostomy Deodorant, and the Safe n' Simple (Sn'S) Ostomy Appliance Deodorant. Lots being recalled were distributed between June 9, 2009 though June 28, 2010. Biomed will work with their 3 consignees to inform sub-distributors and down to any firm that keeps the product in stock to have the product removed from further distribution and returned. If there are any questions or concerns, please contact Biomed Laboratories, Inc. at (972) 282-8008 or email them at biomedlabs.net. |  | 22253 | 22,253 (18,060 8oz bottles. 4,193 1oz bottles) | OH, FL, MI | Terminated | January 24, 2011 | 161 | 72259 | Not_Computer
Z-2435-2010 | 54668 | No | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | The Provent Nasal Cannula is a nasal appliance device placed in both nostrils intended to be used for transmission of respiratory airflow signals between the Provent Professional Sleep Apnea Therapy device and physiologic recorders. |  |  | December 03, 2009 | 2009 | Ventus Medical, Inc. | The catalog number printed on two Nasal Cannula Diagnostic Kits (MRS0231) was different that the catalog number printed on the bulk pack shipper box and the other Nasal Cannula Kits in the same shipper (MRS0208). No devices were used on any patient . | Labeling mix-ups | Ventus Medical, Inc. of Belmont, CA has notified all customers who received distribution from the affected manufacturing lot (1880783) to immediately return the product in question for replacement. If there are any questions regarding this correction, please contact the firm at (650) 632-4165. |  | 60 | 60 | Product was distributed in the US through a single distribution supplier. | Terminated | March 11, 2011 | 463 | 60 | Not_Computer
Z-2436-2010 | 56626 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | The Gyrus ACMI Pleatman Sac Tissue Removal System facilitates removal of specimens while minimizing contamination of the abdominal cavity or wound tract. The system contains one (1) Sac and one (1) obturator in a sterile, ready -to -use, single- use procedure tray. |  |  | August 03, 2010 | 2010 | Gyrus Acmi, Incorporated | Gyrus Medical, Inc. has issued a Urgent Medical Device Recall for the Gyrus ACMI product 004942-903 - Pleatman SacÂż Tissue Removal System, sold as lot 186540FC with an expiration date of 06/2015. As a result of a manufacturing error in the production of the trays used in this lot, it is possible that a crack could form that would compromise the package integrity, and possibly the sterility of th | Packaging process control | Gyrus Medical, Inc. of Maple Grove, MN issued an Urgent Medical Device Recall drawing attention to a possible package integrity compromise that could affect the product sterility. USA consignees were contacted via telephone on 8/3/10. The firm directed consignees to inspect their stock to confirm if any of the affected lot of product remains in their possession. Use of remaining affected product should cease immediately. Affected products should be return to Gyrus ACMI. For questions or concerns, consignees should contact the firm at (888) 524-7266. |  | 2 | 2 cases (20 devices) | UT, PA, Canada | Terminated | August 23, 2012 | 751 | 2 | Not_Computer
Z-2437-2010 | 53044 | No | LCD | Gentamicin test system. | Toxicology | Toxicology | 510(k) | SYNCHRON Gentamicin (GEN) Reagent |  |  | February 19, 2009 | 2009 | Beckman Coulter Inc. | The reagent produces false low test results. Beckman Coulter has confirmed low recovery of values at the low end of the therapeutic range (&lt;1.5 ug/mL, &lt;3.13 umol/L) with SYNCHRON Systems Gentamicin Reagent (GEN) lot M804552. A false low result can cause a patient to receive more medication than necessary. | Other | Beckman Coulter sent an Urgent Product Corrective Action (PCA) letter dated February 27, 2009 to customers identifying the affected product and issue prompting the recall. The letter asked them to discontinue use and discard the reagent lot, and contact their Customer Service Representative at (800) 526-3821 in the United States or (800) 463-7828 in Canada or contact their local Beckman Coulter Representative to discuss reagent replacement options. They were also asked to complete and return the enclosed response form within 10 days.Customers were directed to call the Customer Support Center (Hotline) at (800) 854-3633 in the US and Canada if they had questions in regards to the PCA letter. Outside the US and Canada they were told to contact their local Beckman Coulter Representative. |  | 1547 | 1,547 Kits | Worldwide Distribution - USA, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, and WV, and countries of Canada, Belgium, Denmark, France, Germany, Ireland, Lithuania, Netherlands, Spain, Switzerland, United Kingdom, New Zealand, Singapore, Egypt, South Africa, Taiwan, and United Arab Imarets. | Terminated | July 16, 2012 | 1243 | 1547 | Not_Computer
Z-2438-2010 | 55427 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Integrated Data Management System - Version 7.1, 7.2 and 7.3; |  |  | April 09, 2010 | 2010 | International Technidyne Corporation | A tabulation error in the Integrated Data Management Systems (IDMS) data management software that results in incorrect Electronic Quality Control (EQC) Summary Reports: The Operator and the Device Performance Summary Reports. The IDMS software is not used for the clinical management of individual patients. | Software design | ITC sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 10, 2010 to Distributors and First Consignee End Users identifying the affected product and affected summary reports, the reason for recall, and the actions to be taken by the distributors and customers. New Errata Instructions were enclosed with each letter and are to be put in the idms user guide for reference. The new instructions are to be followed when running reports for the ESV or TVT tests run by the Hemochron Response. The enclosed Customer Account Tracking Form is to be completed and returned to ITC indicating that the letter has been received and appropriate questions answered as to the use of the software system. Distributors are asked to check inventory for recalled product. If affected product is found, shipment should be stopped and placed on hold. Distributors should notify customers of the correction and provide them with a copy of the correction letter and updated instructions. This should be done by traceable mail. Distributors should return the account tracking form and return affected software disks for replacement.Questions should be directed to ITC Technical Support at 800-631-5945. |  | 5544 | 84 total idms software systems; 66 in US, 18 outside of US. | Worldwide distribution: USA including the states of AL, AZ, CA, FL, GA, IL, KS, LA, MD, MI, NC, ND, NJ, NM, NY, OH, PA, TN , TX, UT, VA, WA, and WI, and the countries of Italy, Turkey, Canada, France, Belgium, Germany, UAE, and New Zealand. | Terminated | May 25, 2011 | 411 | 5544 | Software | Output/Calculation | Replace software disks | Software Update
Z-2441-2010 | 55163 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Unicel Synchron Clinical Systems |  |  | January 18, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated due to " A loose Aliquot Probe tube fitting in the Closed Tube Aliquotter (CTA), may result in erroneous Access immunoassay results or Synchron chemistry results".This may result in dilution of the sample with Wash Buffer; patient results can potentially be affected for Access immunoassays or Synchron chemistries. Access Quality Control may also be affected if run throug | Other | Beckman Coulter sent an "URGENT: PRODUCT CORRECTIVE ACTION" letter dated January 18, 2010, to all affected customers. The letter provided the customer with an explanation of the problem identified, the affected product, and instructions for the customer. The letter requested customers complete the following actions for :(1) CTA Aliquot Probe Inspection and Maintenance:a. Check daily for signs of leaks or damage to the tube fitting by looking and feeling for liquid at the fitting, the end of the probe and the surrounding areas;b. The CTA Aliquot Probe tube fitting should be checked to ensure that it is connected tightly to the Aliquot Probe. Without loosening the tube fitting, ensure it is securely tightened;c. When replacing the CTA Aliquot Probe, ensure that the Aliquot Probe tube fitting is securely tightened, prime the CTA, inspect for leaks, and run Access Quality Control through the CTA;d. If a leaking Aliquot Probe is suspected, review all sample results for both Access and Synchron in light of the total clinical presentation of the patient including: symptoms, clinical history, data from additional tests, and other appropriate information;e. As this issue may result in dilution of the sample with Wash Buffer; it may be appropriate to discard and redraw any affected sample tubes run on the CTA. The letter also informed customers that Beckman Coulter is currently evaluating potential solutions for this issue and requested they share this information with laboratory staff and retain this notification as part of their laboratory Quality System documentation. If they forwarded the affected products to another laboratory, they were asked to provide a copy of this letter to them.They were also instructed to complete and return the enclosed response form within 10 days. Customers were told to call Customer Technical Support Center (Hotline) at (800) 854-3633 in the United States and Canada with any questions. |  | 672 | 672 units for both systems | Worldwide Distribution - USA: including the states of AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, AND WV and the country of Canada. | Terminated | August 13, 2012 | 938 | 672 | Not_Computer
Z-2440-2010 | 56390 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TRABECULAR METAL Âż FEMORAL CONE AUGMENT - RIGHT, Part Number: 00-5450-012-32. |  |  | July 05, 2010 | 2010 | Zimmer Inc. | Zimmer, Inc. has determined that there is a potential for the packages to be mislabeled. Specifically, the left femoral cone augment component may be labeled as a right and the right femoral cone augment component may be labeled as a left. | Labeling mix-ups | Zimmer, Inc. sent Urgent: Device Removal letters, dated July 9, 2010, identifying the affected product and stating the issue. Specifically, the left femoral cone augment component may be labeled as a right and the right femoral cone augment component may be labeled as a left. The letter also discussed the clinical implications along with required actions. Customers are to perform a physical count of all affected product and record the data on the Inventory Return Certification Form included with the letter. In addition, customers should complete and return the Inventory Certification and User Facility/HCP Documentation Forms. Affected product along with the Inventory Return Certification Form should be returned. Zimmer, Inc. should be contacted at 1-800-613-6131 if there are any questions or need for additional information. |  | 50 | 50 total | Worldwide Distribution -- USA, Canada, Singapore, Australia, and Switzerland, Spain, UK, Sweden, and Israel. | Terminated | December 22, 2010 | 170 | 50 | Not_Computer
Z-2446-2010 | 56183 | No | IKT | Diagnostic electromyograph needle electrode. | Physical Medicine | Neurology | 510(k) | Disposable Detachable Monopolar Needle Electrode for electromyograph model number 110-725, for Physician use, manufactured by Chalgren Enterprises Inc., Gilroy, CA. |  |  | April 12, 2007 | 2007 | Chalgren Enterprises Inc | Sterility Compromised-- Investigation found that individual packaging seals were separating, which may compromise sterility. | Packaging | Chalgren Enterprises, Inc. removed all of the affected inventory from inventory and had all affected inventory retrieved from distribution. |  |  |  | Nationwide Distribution: USA states of CA, IL, NJ, OH and TX. | Terminated | March 24, 2011 | 1442 |  | Not_Computer
Z-2445-2010 | 56460 | No | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | Primary (inner) Package Label: ACUVUE OASYS BRAND CONTACT LENSES for ASTIGMATISM with HYDRACLEAR Plus (senofilcon A) Johnson &amp; Johnson Vision Care, Inc. Rx Only, STERILE. Lot B006WBS1 -9.00D -1.25 CYL 80 , 2013/05 AXIS. (Correct labeling) |  |  | July 16, 2010 | 2010 | Vistakon | ACUVUE OASYS Brand Contact Lenses for ASTIGMATISM may be mislabeled on the outer (secondary) packaging carton. The outer (secondary) cartons labeled as Lot B006X641 -8.50D -1.25cyl. 100 Axis may actually contain primary packages from Lot B006WBS1 -9.00D -1.25cyl. 80 Axis. The inner (primary) package labeling is correct. Product was only distributed to Eye Care Professionals for use in proper f | Error in labeling | Vistakon issued an Important Product Notification dated July 28, 2010 to customers identifying the affected product, the labeling issue, and actions to be taken by the customer. Vistakon issued a separate Important Product Notification dated July 28, 2010 to distribuitors identifying the affected product, the labeling issue, and actions to be taken by the distributor. Customers are instructed to check their inventory for affected products and return a business reply card. If replacement -8.50 lenses are needed, the customer can contact Customer Service at 1-800-874-5178.Distributors are instructed to check their inventory for affected products and return them using an enclosed return label. Replacement product will be sent by the firm. Customer Service can be contacted at 1-800-874-5178.Accounts will be instructed to go to their inventory and look for the B006X641, -8.50 lot and open the secondary carton and verify the product inside. If it is the -9.00 product they can keep the product and re-slot it in the proper place within the kit (product is ok, nothing wrong with primary package or lens inside) and order a replacement -8.50 if necessary. Business reply cards will be supplied for accounts to return to third party indicating if they had any mix/mislabeled product. |  | 2123 | 2,123 lenses | Worldwide Distribution: USA including all states, District of Columbia and Puerto Rico, and the countries of Canada, Bahamas, Honduras, and United Arab Emirates. | Terminated | March 28, 2011 | 255 | 2123 | Not_Computer
Z-2447-2010 | 56327 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | PSD Gel that's provided with Peri-Strips Dry Staple Line Reinforcement with Apex Processing (PSD) and Peri-Strips Dry with Veritas Collagen Matrix Staple Line Reinforcement (PSD-V) products. Sterile, Rx only, SynovisÂż Surgical Innovations. |  |  | June 30, 2010 | 2010 | Synovis Surgical Inovation Div. of | SynovisÂż Surgical Innovations is voluntarily recalling PSD Gel that's provided with Peri-Strips Dry Staple Line Reinforcement with Apex Processing (PSD) and Peri-Strips Dry with Veritas Collagen Matrix Staple line Reinforcement (PSD-V) products. Internal investigation have shown there is a potential for small defects in the PSD gel tube pouch which may compromise the sterility of the outside of t | Packaging | Consignees were sent on 6/30/10 a Recall Notification letter dated June 30, 2010. The letter was addressed to "Risk Manager". The letter described the problem and product involved. They listed both the device and gel lot numbers since it is possible that some hospitals have separated the pouched gel tubes from the PSD/PSDV products. They advised consignees to quarantine the affected lots and to complete and fax the attached "Product Return/Replacement" form to Synovis Surgical Innovations. |  | 5384 | 3,989 gel tubes and 1,395 PSD &amp; PSDV devices. | Worldwide Distribution -- USA, including states of AL, AZ, CA, CO, CT, DE, FL, GA, ID, IL,IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, and WI and countries of AUSTRALIA, GERMANY, GREECE, HONG KONG, NETHERLANDS, PUERTO RICO, ROMANIA, SAUDI ARABIA, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM. | Terminated | December 17, 2011 | 535 | 5384 | Not_Computer
Z-2443-2010 | 56121 | No | DTE | External pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | Oscor PACE 203H, Dual-Chamber External Pulse Generator |  |  | June 30, 2010 | 2010 | Osypka Medical Inc | The recall was initiated after Osypka Medical received two (2) complaints from their Japanese distributorregarding an incompatibility of the PACE 203H with some new types of 9 V batteries.The investigation revealed that PACE 203H of a certain manufacturing period malfunction(i.e., not stimulate) when a type of 9 V batteries with higher than usual open circuit voltage (OCV) is used. Failure t | Device Design | OSYPKA Medical sent "Urgent Medical Device Recall" letters dated June 30, 2010 and July 7, 2010 to the direct customers. The recall notice informed the customers of the Reason for Recall, the Potential Hazard, the Temporary Fix of the Problem, the Permanent Fix of the Problem and Recommended Actions to be taken by the customer. The customers were instructed to contact Osypka Medical via telephone (858) 454-0021 or e-mail to m.osypka@osypkamed.com to verify that the Model 3085 and/or PACE 203H is subject to this product removal, obtain a return-to-manufacturer-authorization (RMA) number ,and then return the affected devices to Osypka Medical.If you have any questions call (858) 454-0021 or e-mail to mail@osypkamed.com |  | 205 | 205 units (US Only) | Nationwide Distribution: USA including states of CA &amp; FL | Terminated | August 10, 2012 | 772 | 217 | Hardware | Device Operation | Return | Remove or Replace
Z-2461-2010 | 49594 | No | MOS | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Nuclear Magnetic Resonance Imaging System (MRI). Trace item 9896-030-14351 -1.5T Sense XL Torso coil 12nc on coil: 4535-671-4188x. |  |  | September 30, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | The cables of the coil may become too hot and cause burns when placed too close to the patient. While using the 1.5T Sense XL Torso coil, unwanted radiofrequencies (RF) interaction with the patient can occur. Especially in combination with high SAR levels and when the RF cables of the coil are too close to the patient, this RF interaction has resulted in RF burns to patients. | Other | Philips Healthcare issued URGENT-Device Correction Notices dated September 30, 2008 to their consignees. The notice identified the recalled device as 1.5 Intera and 1.5T Achieva MRI systems, described the risk to patients, and the actions to be taken by customers..The firm has issued Field Change Order (FCO) 78100242. Field service representatives will visit each consignee and conduct a hardware upgrade, applying a sleeve over the cable that will ensure proper spacing between the cable and patient, reducing the chance of RF interaction and burns.The consignees should contact the Philips Call Center at 1-800-772-9377, #5 #3, #1 and reference "FCO 78100242 - PATIENT CABLE SLEEVE FOR 1.5T SENSE XL TORSO COIL". |  | 47 | 47 units total (Achieva 1.5T and Intera 1.5T) | Nationwide Distribution: Throughout USA. No distribution to foreign consignees. | Terminated | September 30, 2010 | 730 | 47 | Not_Computer
Z-2464-2010 | 50066 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips BV Endura Mobile X-Ray system with the extended rotation option. |  |  | September 10, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | The C-arm rotation brake does not always function properly, and may not hold the C-arm position if the C-arm is not positioned in its maximum rotation and if it is accidentally bumped by medical staff or patients. Unwanted C-arm movement may cause the C-arm to collide with a patient. The brakes function properly if the C-arm is rotated to its maximum positions. | Other | Philips Medical Systems issued an Urgent Product Correction letter dated September 30, 2008 to consignees. The letter identified the affected device, what the problem is and under what circumstances it can occur, and the actions that should be taken by the customer or user in order to prevent risks for patients or users. Philips will contacdt customers to implement Field Corrective Action (FCO) #71800026), which describes a mechanical brake modification which will be installed by Philips. Until the Field Corrective Action has been completed, users should ensure that the C-arm is placed in the maximum extended position, or ensure that the C-arm is not touched during critical procedures if the C-arm is not in a maximum position.The Philips Call Center can be contacted at 800-722-9377 |  | 71 | 71 units | Nationwide Distribution: Throughout the US. | Terminated | January 07, 2011 | 849 | 1420 | Not_Computer
Z-2476-2010 | 50114 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Allura XPER FD20 Angiographic X-ray system. This system with Velara Generator can be identified as type number 9890 000 PRODUCTS 6501 * (number located on the generator cabinet). |  |  | October 23, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | A high voltage reliability issue resulting in the loss of the x-ray generation and this can cause an intermittent stop of fluoro acquisition and exposure during an examination without warning and hence no image on the display. | Other | The firm, Philips, sent "URGENT-Device Correction" letters dated November 10, 2008, to their customers and informed them about the reliability issue in the Allura XPER and Integris X-ray systems with the High Voltage Velera (HV) Generator. This reliability issue can lead to a system error causing intermittent loss of fluro and/or exposure acquisition, potentially resulting in the loss of the X-ray image on the display.The customers are advised if they experience High Voltage Generator reliability issue and the HV Generator stops working generating an O3EA error, a cold restart is required and reboot the system when the X-Ray image disappears.The firm will implement multiple mandatory Field Change Orders (FCOs) to solve the Generator reliability by updating the software for each X-ray system.The customers could contact Philips Care Center at 800-722-9377, #5, #4, #1 for any questions regarding this recall or 425-487-7469. |  | 838 | 838 units in US | Worldwide distribution. | Terminated | August 10, 2011 | 1021 | 1681 | Other | Display/Image | Software update | Software Update
Z-2448-2010 | 56093 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | MONOPLANE X-ray Imaging Systems include the Allura XPER FD10 F/C, Allura XPER FD20, Integris H5000/Allura 9 F/C, Integris CV, and Allura 12-15 Monoplane. |  |  | February 04, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | A lock-up might occur inside the Velara Generator due to intermittent shorts inside the system. The problem could only resolved by a warm or cold system restart. In exceptional cases the power converters of the KVmA unit in the generator are damaged and the system can not be restarted. | Software design | On 02/04/2010, Philips Healthcare began sending out the URGENT-Device Correction Notices to their consignees. The firm issued two notifications, one for the Allura Integris and Xper FD10C, FD10F, FD10F/C and FD20 Monoplane X-Ray imaging systems with the Velara Generator. The second notification is for the Allura Integris Xper FD10 and FD20 Biplane X-ray imaging systems with the Velara Generator.This notice informs the consignees that a system locks up during a critical interventional case and it might impose a serious injury to patients. This lock up issue could only be resolved by a warm or cold system restart. A warm restart would take about 90 seconds or a cold restart would take about four minutes, and the device could not be used during this time. If the system does not work after the restart, the consignees should contact the Philips Healthcare Customer Service Representative at 1-800-772-9377 #5 (then follow prompts). A service representative will visit each consignee and conduct a software and hardware upgrade. |  | 1470 | 1470 units | Nationwide Distribution | Terminated | September 23, 2011 | 596 | 1737 | Computer
Z-2425-2010 | 56190 | No | HAO | Nonpowered neurosurgical instrument. | Neurology | Neurology | 510(K) Exempt | Integra NeuroSciences Implants, NeuroBalloon Catheter, |  |  | July 02, 2010 | 2010 | Integra LifeSciences Corp | Certain NeuroBalloon Catheters have the potential to improperly inflate or deflate under certain conditions. | Process design | All customers and Integra Sales Specialists were notified of the recall on July 2, 2010 by fax, e-mail and letter send by Fed Ex overnight delivery. |  | 338 | 338 units | Nationwide, Canada, Taiwan and to Integra, Australia. | Terminated | May 29, 2012 | 697 | 338 | Not_Computer
Z-2465-2010 | 56179 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Spider Flex Monopolar Hook (Spider Surgical Instruments), TransEnterix, Inc., Durham, NC |  |  | June 22, 2010 | 2010 | Transenterix, INC | The hook end effecter from a monopolar device may become dislodged and fall into the patient when in use during laproscopic surgery. Retrieval from the patient may involve extended anesthesia time. | Other | TransEnterix issued a "Field Action" letter dated June 22, 2010 to consignees, identifying the affected product and actions to be taken by the customer. Customers were instructed to return the product to TransEnterix, Inc. using the address label and return information provided. The firm will replace the product at no charge. TransEnterix Customer Care can be contacted at 1-888-879-4111. |  | 46 | 46 boxes (6 units per box) | Nationwide Distribution in the states of AZ, FL, GA, NC, NY, OH, OR, SC and TX. | Terminated | January 31, 2011 | 223 | 46 | Not_Computer
Z-2466-2010 | 56303 | No | HEB | Unscented menstrual tampon. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Life Brand Regular Absorbency 40 Count Unscented Plastic Applicator Tampons |  |  | March 25, 2010 | 2010 | First Quality Hygienic Inc | Life Brand Regular Absorbency 40 Count Unscented Plastic Applicator Tampons package contained Super Plus Unscented Plastic Applicator Tampons. | Labeling mix-ups | Consignees were notified via telephone and letter of the need to withdraw product from the market and requested to return all product to manufacturer. |  | 456300 | 195 Cases/2340 Cartons | Canada | Terminated | October 06, 2010 | 195 | 456300 | Not_Computer
Z-2460-2010 | 56506 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Heavyweight Gantry Scanner, M/N 728306, S/N: 100016, 100018 -100025 |  |  | April 14, 2010 | 2010 | Philips Medical Systems (Cleveland) Inc | The gantry covers of the iCT have a latch at the base of the system which holds the cover in the closed position. During the course of maintenance the CT Field Service Engineer (FSE) will activate this latching mechanism which in turn has a potential to penetrate the insulation on cabling which carries 277 VAC. If the latch does come into contact with the cabling there is a potential that the FS | Other | Philips Healthcare sent an URGENT - Field Safety Notice letter, dated April 14, 2010, to their customers notifying them of the product, the problem, and the action to be taken. The letter asked them to identify the affected product by checking for the serial number tag (s/n 10016  10025), ensure that the electrical power was shut-down before opening or closing the gantry cover to prevent potential of an electrical shock, and that a Field Service Engineer would contact the facility to make arrangements for the implementation of a Field Change Order (FCO) before December to correct the problem. For further information or support concerning this issue, the customer was instructed to contact their Philips representative: Customer Care Center (1-800-722-9377, option 5 or their local Philips Healthcare Office). |  | 9 | 9 units | Worldwide Distribution - USA, including the states: DC, GA, IL, IN, NY and OH; and the countries of Germany and the Netherlands. | Terminated | September 24, 2012 | 894 | 9 | Not_Computer
Z-2450-2010 | 56563 | No | GAZ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Medi-Vac Non-Conductive Suction Tube with Maxi-Grip Connectors, Inner Diameter: 5 mm (3/16"), Length: 1.8 m (6 ft.); Sterile, for single use only, 50 tubes per case; Cardinal Health, McGaw Park, IL 60085 USA, Made in Mexico; catalog N56A and NEX56A (European label) |  |  | August 16, 2010 | 2010 | Cardinal Health | The packaging seals may not be intact, resulting in compromised sterility for the listed Medi-Vac Non-Conductive Tubing. | Other | Cardinal Health issued an URGENT: PRODUCT RECALL letter dated August 16, 2010 to consignees, identifying the reason for the recall, the affected lots, and actions to be taken. Cardinal has inspected all packages in their warehouses and removed the affected products. All packages that have been inspected have a 3/4" green dot on the front of each unit and on the case carton label." Cardinal Health end users were instructed to segregate and destroy their affected product and request credit or replacement" Distributor customers were instructed to destroy their product and request credit or replacement" Distributors were instructed to conduct sub-recalls and return their inventory to Cardinal Health" To arrange for credit or replacement of the product destroyed, the accounts were instructed to call customer service at 800-964-5227 or 888-444-5440." All customers were requested to complete the enclosed Recall Acknowledgment Form acknowledging receipt of the letter and the number of tubing units destroyed, and fax it to 847-689-9101.Cardinal Health can be contacted at 800-292-9332.Cardinal sent recall letters dated 8/26/10 to their PreSource kit/module customers via UPS next day air, informing them that the Medi-Vac Non-conductive Suction Tubing listed may have packaging seals that are compromised. The accounts were informed that Cardinal has inspected tubing packages in their warehouses and removed the affected products. All packages that have been inspected will have a 3/4" green dot on the front of each package. Enclosed with the recall letter were recall labels to affix to each affected Presource pack/module in the account's inventory. These labels instruct the user NOT to use the affected tubing component and to remove and discard the component upon opening the kit. The accounts were also requested to complete and fax back to Cardinal the enclosed reply sheet acknowledging receipt of the letter and the number of packs/modules on hand. Distributo |  |  |  | Worldwide Distribution, including USA. | Terminated | November 28, 2011 | 469 |  | Not_Computer
Z-2470-2010 | 56667 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | DRIVER STAR #8 CANNULATED DART-FIRE (TM), COMPRESSION SCREW, SIZE: #8, REF: DSDS-0008, 1 EACH, Rx ONLY, NON-STERILE, EC REP Wright Medical EMEA, Krijgsman 11, 1186 DM Amstelveen, The Netherlands, Wright Medical Technology, Inc., 5677 Airline Road, Arlington, TN 38002 |  |  | August 26, 2010 | 2010 | Wright Medical Technology Inc | After autoclaving, the epoxy identifier (colored) band bled and there is concern that it may flake. | Process design | The firm initiated their recall the product by letter on 08/26/2010. The letters were addressed to distributors and hospital administrators with a fax back response form. Please contact Customer Service at 800-238-7117, if you have any questions. |  | 13 | 13 units | TX, PA, TN, VA, GA, and Spain | Terminated | August 02, 2012 | 707 | 13 | Not_Computer
Z-2467-2010 | 56686 | No | ING | Wheelchair elevator. | Physical Medicine | Physical Medicine | 510(k) | Braun wheelchair lift, model NL, Braun Corporation, Winamac, IN |  |  | August 15, 2007 | 2007 | The Braun Corporation | The wheelchair lifts threshold warning system may not operate and the wheelchair user can fall from the vehicle and be seriously injured. | Device Design | The Braun Corporation issued a Service Bulletin 34273 entitled "Lift Threshold Sensitivity Bulletin 34273 on October 2007. Braun Corporation also sent a "Safety Recall" letter to the Braun lift owner and letters to all customers on or about August 15, 2007. The letters described the product, problem and the action to be taken by the customer. The customers were instructed to notify the National Highway Traffic Safety Administration (NHTSA) and conduct a safety recall if they used recalled lifts as original equipment in a system or vehicle that they built; ensure that any associated customers are notified to discontinue the use of their vehicles until they schedule a repair, and verify the threshold warning system is within the NHTSA specification. The customers were informed to contact the NHTSA with questions by calling the Vehicle Safety Hotline at 1-888-327-4236 or go to http//www.safercar.gov. If they have any problem obtaining the needed repair or wish to schedule service with an approved Braun technician, contact the Braun Corporation Product Support at 1-800-946-7513. |  | 8356 | 8,356 | Worldwide distribution | Terminated | May 12, 2011 | 1366 | 12940 | Not_Computer
Z-2482-2010 | 55360 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECTÂż i1000SR" Processing Module System For In Vitro Diagnostic Use, Abbott Laboratories Abbott Park, IL |  |  | October 21, 2008 | 2008 | Abbott Laboratories, Inc | During manufacturing, two components of the ARCHITECT i1000SR were reversed which could potentially cause falsely elevated patient results due to a sample-reagent and reagent-sample carryover. | Mixed-up of materials/components | ABBOTT issued Technical Service Bulletins to Field Service Representatives who then inspected every unit . |  | 404 | 404 units distributed worldwide (including US) | Worldwide distribution:USA including states of: CO, DC, FL, GA, IL, LA, MA, MN, NY, OH, OK, TX, UT, and WV and countries including: Australia, Brazil, Bogota, Canada, China WFOTC, Columbia, Curacao (NETH ANT), Germany, Hong Kong, Japan, Mexico, Singapore, and Taiwan. | Terminated | April 13, 2012 | 1270 | 404 | Not_Computer
Z-2480-2010 | 55911 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Terumo Advanced Perfusion System 1, Catalog number 801763, 100/120V System 1 base, Terumo Cardiovascular System, Ann Arbor MI. |  |  | September 10, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The design document states that the oxygen sensor would operate for 300,000 hours but the supplier of the oxygen sensor indicated that it has a 100,000 hour warranty. The oxygen sensor is a consumable part that slowly depletes over time. At the end of its service life, a depleted oxygen sensor results in a gas system calibration failure or a message to the user prompting service for the gas syst | Nonconforming Material/Component | Terumo issued a URGENT MEDICAL DEVICE CORRECTION letter dated September 10, 2010, identifying the affected product, the issue prompting the correction, and what actions are to be taken by the customer. The firm included the addendum to the operators manual with the recall letter. The addendum recommended replacement of the oxygen sensor every six months instead of every 12 months. The addendum provided additional user instructions in the event of a calibration failure of the sensor. Customers are to complete and return the enclosed Customer Response Form to the firm. Terumo CVS Customer Service can be contacted by 1-800-521-2818 Monday through Friday, 8 AM  6 PM EST. |  | 1340 | 1340 | Worldwide Distribution: USA, including the states of AL,AZ, CA, CO, CT, DC, FL, GA, HI, IA, IN, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WI, and WV, and the countries of Canada, Indonesia, China, Chile, UAE, Japan, Singapore, UK, Malaysia, Taiwan, Belgium, Korea, Phiippines and Thailand. | Terminated | December 20, 2011 | 466 | 2655 | Not_Computer
Z-2462-2010 | 56560 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | GCX Mountable Downloader-Recharger; |  |  | August 20, 2010 | 2010 | Abbott Point of Care Inc. | Reliability issues associated with broken connection pins within the downloader that mate with the pins from the analyzer. | Device Design | Abbott Point of Care, Inc. sent Urgent Recall Notices, dated August 20, 2010, by Federal Express to two Clarion Health locations.The letter identified the product, the issue, and the actions to be taken by the firm and the customers. 1) An Abbott Point of Care representative will contact the customers to plan the replacement of affected product at their facility.2) Replacement will occur in batches of approximately 50 units at a time.3) Abbott will provide personnel to perform the replacements to help expedite the replacement.4) The customer is required to acknowledge that all units have been replaced by signing the form provided with the Urgent Recall Notice.. If you have any questions regarding this information please contact Abbott Point of Care Technical Support at 800-366-8020, Option 1 or your Abbott Point of Care representative. |  | 277 | 277 | Nationwide Distribution - USA, including the state of Indiana | Terminated | July 25, 2012 | 705 | 277 | Not_Computer
Z-2478-2010 | 56676 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Fabius GS Anesthesia Machine, OR/Anesthesia, Draeger Medical, Inc., 3135 Quarry Road, Telford, PA 18969 |  |  | August 26, 2010 | 2010 | Draeger Medical Systems, Inc. | potential failures to Electri-Cord Manufacturing AC power cord | Other | Draeger Medical, Inc., sent an Urgent - Medical Device Recall letter dated August 2010 to all customers. The letter described the product, the problem, and actions to be taken by the customer. Customers were instructed to:1) Immediately take the cord out of service if the plug has bent or cracked prongs, burned plastic, or shows signs of excessive wear.2) Monitor the affected cords regularly and be mindful of excessive wear and tear, misuse, or abuse until the affected cord is replaced. 3) Customers were also asked to complete and return the enclosed Customer Power Cord Form via mail: Draeger Medical, Inc., 3135 Quarry Road, Telford, PA 18969; fax: 1-215-721-5808; or e-mail: kathy.benedetto@draeger.com.4) Disable the cords by cutting off the plug and dispose of the cord in accordance with their facility's disposal procedure when affected power cords are removed.NOTE: The firm will send replacement power cords free of charge. Draeger will only replace power cords originally provided by Draeger (identified with the white flag/part number 4117266.If you have any questions regarding this letter please call 1-800-543-5047 (press 1 at the first prompt, 2 at the second prompt and then 32349#). |  | 1973 | up to 1973 products (GS &amp; Tiro) | Worldwide Distribution: USA and countries of Brazil, Canada, Columbia, Costa Rica, Cuba, Ecuador, Japan, Lebanon, Nicaragua, Mexico, Panama, Philippines, Saudi Arabia, Taiwan, and Venezuela | Terminated | August 23, 2011 | 362 | 1973 | Battery |  | Replace | Remove or Replace
Z-2483-2010 | 56498 | No | HGX | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Sterile double breast personal accessory convenience kit for the Symphony Breast pump; the kit connects to a hospital grade breast pump for the removal, collection, and storage of mother's breast milk; Manufactured exclusively by and for Medela, Inc., P.O. Box 660, 1101 Corporate Drive, McHenry, IL 60051-0660; article # 67099S |  |  | August 06, 2010 | 2010 | Medela Inc | The product sterility may be compromised due to possible pinholes in the rigid tray of the package. | Packaging | Medela sent Important Device Recall letters dated August 6, 2010 to all direct accounts, identifying the affected product, the labeling issue, and actions to be taken by the customer. The accounts were informed that the sterility of the affected products may be compromised due to holes in the packaging. They were instructed to stop use, sale and distribution of the kits and return their stocks to Stericycle using the enclosed packing slip and prepaid UPS Return Service shipping label. The accounts were requested to complete and return the enclosed postage paid reply card, indicating the amount of product being returned. Distributors were requested to notify their customers. Questions about the recall process may be directed to Stericycle at 1-877-247-9970. Information on the recall is posted on Medela's website www.medelabreastfeedingus.com |  | 581296 | 581,296 kits | Worldwide Distribution: Nationwide throughout USA, including Guam, and to the countries of Barbados, Bermuda, British West Indies, Canada, Panama, Switzerland, Trinidad and Tobago, and Uganda. | Terminated | September 07, 2011 | 397 | 6924825 | Not_Computer
Z-2487-2010 | 56423 | [6 empty] | Ear Candles made from either Paraffin or Beeswax, in plain, herbal scented, or lavender scented varieties, manufactured by Wally's Natural Products, Auburn, CA. |  |  | February 17, 2010 | 2010 | Wally's Natural Products | Product is an unapproved medical device. Use as directed may result in accidental damage to the eardrum, leading to injuries up to loss of hearing. | Other | The firm conducted communications via phone, mail and website, stating that they have stopped manufacturing and distributing ear candles, and asks consumers to return for a refund. |  |  | Not Specified. All | Nationwide Distribution | Terminated | March 28, 2011 | 404 |  | Not_Computer
Z-2501-2010 | 54745 | No | LCX | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | UCG Beta Slide Monoclonal II is a latex agglutination inhibition slide test for the detection of hCG in urine. This test is used for the diagnosis of pregnancy. |  |  | February 12, 2010 | 2010 | Inverness Medical Professional Diagnostics | Specific lots no longer meet their stated limit of detection claims of 2000 mIU/mL for the UCG-Slide test and 500 mIU/mL for the UCG-Beta Slide Monoclonal II test. | Other | The recall was initiated on February 12, 2010. Letters were sent by Fed Ex on 2/12/2010. |  | 585 | 585 kits | Distributed Nationwide and to Australia | Terminated | May 14, 2012 | 822 | 1637 | Not_Computer
Z-2503-2010 | 56441 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | GATOR. REF 9263A, 4.2mm STERLING, STERILE Rx Only. |  |  | August 05, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Largo, FL is recalling ConMed Linvatec 9263A,4.2mm Sterling Blade, Gator Meniscus Cutter, Lot Number 178026 for breech in the sterile barrier. | Packaging process control | Recall Notifications will be sent to all direct consignees by FedEx, email, Fax, etc. If the account no longer has the product in their possession, they will be asked to respond as such. A Return Material Authorizations Number will be assigned for the returned product as required. Product returned will be placed into quarantine and will be dispositioned in accordance with ConMed Linvatec procedures. 100% of the consignees will be contacted. The reply form is the mechanism to verify notification effectiveness to all consignees. If a reply is not received within approximately 30 business days of the first notification ConMed Linvatec will make two additional contacts by phone and / or email, as necessary, to insure all consignees have received notification and have taken appropriate action. If a reply is not received following the third contact by email or phone, no further contacts will be made to the consignee. |  | 174 | 174 | Products were distributed to ConMed Linvatec consignees, sales representatives and affiliates within the US and Australia. US distribution was to WA, AZ, PA, KS, NM, UT, ID, IL. Products were distributed to ConMed Linvatec Australia. | Terminated | December 14, 2010 | 131 | 174 | Not_Computer
Z-2499-2010 | 56406 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Liquid Mouse Monoclonal Antibody Thyroid Transcription Factor-1; 1 mL liquid tissue culture supernatant in a 1 mL vial; an in vitro diagnostic reagent for the qualitative identification by light microscopy of thyroid transcription factor-1 molecules in paraffin-embedded tissue by immunohistochemical staining; Leica Biosystems Newcastle Ltd., Balliol Business Park West, Benton Lane, Newcastle Upon Tyne NE12 8EW, United Kingdom; Catalog No.: NCL-L-TTF-1 |  |  | August 02, 2010 | 2010 | Leica Microsystems, Inc. | In combination with a sensitive polymer detection system, the TTF-1 antibodies may stain additional tumor types other than is currently indicated on the instructions for use. | Device Design | Leica Microsystems sent Medical Device Field Correction letters dated 7/28/10 to the affected customers via first class mail on the same date, informing them that in combination with a sensitive polymer detection system, the TTF-1 antibodies may stain additional tumor types other than is currently indicated on the Instructions for Use. Currently they state that lung and thyroid tumors may stain, but Leicas investigation further characterized the expression of the TTF-1 antigen and it may also be present in other tumors, e.g. thymomas. Customers who have received the affected lot numbers of the three products were requested to refer to the new Instructions for Use to review the amendments made to the Results Expected via the following links: http://bsd.leica-microsystems.com/pdfs/products/ttf-1-ce.pdf; http://bsd.leica-microsystems.com/pdfs/products/ttf-1-l-ce.pdf; and http://bsd.leica-microsystems.com/pdfs/products/pa0364.pdf; The accounts were asked to complete the attached Acknowledgement Form, acknowledging the receipt and understanding of the letter, listing the lot numbers of the units at their facility and fax it back to Leica at 1-847-607-3147. Additional questions should be addressed to the firm's office of Regulatory Affairs and Quality Assurance. |  | 860 | 860 units Total for all three products | Nationwide and Canada | Terminated | January 29, 2012 | 545 | 860 | Not_Computer
Z-2502-2010 | 56418 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GreenLight HPS fiber optics, Model Number 10-2090, manufactured by AMS (American Medical Systems), San Jose, CA. Potentially Mislabeled as GreenLight PV 10-2079 |  |  | January 22, 2010 | 2010 | AMS Innovative Center - San Jose | Product may be packed with incorrect external or internal labels and/or missing product insert. | Labeling mix-ups | Notification issued via letter with instructions and response sheet through certified mail on 01/12/2010. |  | 11668 | 11668 units | Product was distributed to 148 medical facilities throughout the US. and 6 consignees in Great Britain, Canada and Australia. | Terminated | January 25, 2011 | 368 | 11668 | Not_Computer
Z-2504-2010 | 56591 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Modular Microplasty Cup inserter 3/8" thread , REF 31-400600 |  |  | July 28, 2010 | 2010 | Biomet, Inc. | There was a clerical error in order entry which resulted in the recalling firm recieving spring components from a supplier made of "music wire" The recalling frim conducted tests to determine the effects of theis change in materials. The testing showed that the wire would oxidize and there would be early failure of the device. | Employee error | The firm send URGENT MEDICAL DEVICE RECALL NOTICES to its distributors on 7/28/2010. The notices instructed the distributors to discontinue use of the product and return it to Biomet. |  | 331 | 331 total for both products | US, Australia, Canada, Mexico, Chile, Japan. | Terminated | October 01, 2012 | 796 | 331 | Not_Computer
Z-2496-2010 | 56363 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Two-Lumen Central Venous Catheterization Kit with Blue FlexTip ARROWg+ard Blue Catheter |  |  | July 29, 2010 | 2010 | Arrow International Inc | Incorrectly assembled 18 Ga x 2 1/2 Introducer Needle | Mixed-up of materials/components | The recalling firm issued an Urgent Medical Device Recall letter dated 7/23/2010 along with a a recall acknowledgement and stock status form. |  | 5 | 5 | The product was shipped to medical centers and hospitals nationwide and to Australia, New Zealand, and Europe. | Terminated | July 05, 2012 | 707 | 113 | Not_Computer
Z-2582-2010 | 49599 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Product is labeled in part: "CP-SM397A-02...Mono-Dox Violet Monofilament Polydioxanone Synthetic Absorbable Sterile Suture. Do Not Resterilize.... 4/0 (1.5 metric) 30" (75cm) Rx Only...LOT L0535440 EXP 2009-11...4/0 M397A...Mono-Dox CUTTING NFS-2 NEEDLE...Manufactured by: CP Medical, Portland, OR USA". |  |  | September 17, 2008 | 2008 | C P Medical | Seal integrity of inner foil pouch of the Mono-Dox Synthetic Absorbable Suture may be compromised. | Vendor change control | On 09/17/2008, the firm's Sale Representative began using the Telephone Script and calling all their customers.On 09/19/2008, the firm began sending the Recall Notice, dated 9/17/2008, letter to their customers. The letter states that the recalled product is Polydioxanone Suture, Code M397, Lot Number L0535440, Expiration Date: November 2009.The customers are advised to discontinue selling this product and return any existing stock to CP Medical. The firm also requests their consignees to notify their sub-accounts about this recall notice if the product is further distributed.The customers are instructed to provide return all accounted quantity and provide quantify information and location of any product that they are unable to recover for accountability purposes. Customers could contact CP Medical Customer Service Representative at 1-800-950-2763 for return goods authorization number and shipping account number. |  | 20208 | 1684/12 pcs. Boxes (1628/12 pcs. Boxes to veterinary market, 56/12 pcs. Boxes to human suture market) | Worldwide Distribution -- US, Australia, and Taiwan. | Terminated | September 28, 2010 | 741 | 20208 | Not_Computer
Z-2580-2010 | 56046 | No | OCL | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | Cardiovascular | 510(k) | Cardima Surgical Ablation Probe (SAP) Electrosurgical cutting and coagulation device, Model Number 01-242208, manufactured by Cardima Inc, Fremont, CA |  |  | May 20, 2010 | 2010 | Cardima Inc | Sterility compromised-- The sterile pouch in which the product is packaged may be compromised, which could result in patient infection. | Device Design | Cardima issued a Product Recall Notification dated May 18, 2010 to customers, identifying the affected product, the sterility issue, and the actions to be taken by the customer. Custormers are to quarantine and return the affected product to the firm. The firm will replace returned product as soon as it becomes available.Cardima Customer Service can be contacted at 800 354-0102. |  | 28 | 28 units | Worldwide Distribution to Distributors in USA (Pittsburgh, PA) and the countries of Ireland and Belgium, and customers in the states of CA, OH, MI, MO, AND WA, and the countries of Ireland and the Netherlands. | Terminated | March 23, 2011 | 307 | 28 | Not_Computer
Z-2584-2010 | 56379 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | WECK IMA/ENT Blade Electrode Âż (1.9 cm), REF 809335, Rx Only Sterile, Teleflex Medical, RTP, NC |  |  | July 30, 2010 | 2010 | Telefelx Medical | Complaints were received regarding peeling and melting of the insulation of the cautery tip of the IMA/ENT blade electrodes. | Other | Teleflex Medical sent an "Urgent Medical Device Recall" letter dated July 30, 2010, to all customers. The letter described the product, problem and actions to be taken by customers. The customers were instructed to immediately discontinue use, quarantine any affected products and return all of the affected product to Teleflex Medical; forward the letter to their customers to retrieve product from those locations and complete and return the Recall Acknowledgement Form via fax at 1-866-804-9881. A second mailing to non-responding customers on the consignee list will be mailed approximately four to six seeks after the first mailing. A third mailing to non-responding customers on the consignee list will be made approximately four to six weeks after the second mailing or as an alternative, they may be contacted by phone, e-mail, or fax. Note: A record of these contacts will be maintained. Affected products received by Teleflex Medical will be destroyed upon receipt. Records will be maintained accordingly.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 190344 | 190,344 eaches | Worldwide distribution: USA and countries of Canada, Australia and Singapore. | Terminated | April 21, 2011 | 265 | 190344 | Not_Computer
Z-2579-2010 | 56199 | Yes | NIK |  |  | Cardiovascular | PMA | Sorin Paradym CRT-D Model 8750 and Sorin Group Paradym sonR CRT-D Model 8770. |  |  | June 07, 2010 | 2010 | ela Medical Llc | The Paradym CRT 8750 and CRT sonR 8770 device models have a software anomaly, which will cause the device to lose the ability to sense/pace and to deliver further therapy. The software anomaly could only occurs under a specific sequence of events. | Software design | Sorin Group issued a "Medical Device Correction" letter dated June 7, 2010 to consignees. The letter described the product, the problem, and actions to be taken by customers. A Sorin Group representative will contact customers to update the software. Prior to receiving the software update, customers should program the exercise AV delay to 40 ms, which will trigger a reset in the event of an anomaly. This will restore normal operation after shock delivery. If the anolmaly occurs on an implanted unit, a RV coil or SVC coil continuity test should be started, which will ensure that the device recovers proper operation.Corin Group can be contacted at 1-615-788-1807. |  | 215 | 215 | Nationwide Distribution in the states of AZ, CA, FL, IA, LA, MA, MS,MO, NJ, NY, NC, OH, OK, PA, SC, TN, TX, and VA. | Terminated | March 29, 2012 | 661 | 215 | Software | Treatment/Delivery/Therapy | Software update | Software Update
Z-2585-2010 | 54561 | No | FZW | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CONCEPT GRAFIX TENDON STRIPPER. REF TS8850. 7.0MM, QTY 1. Rx Only, Non-Sterile, Made In USA. |  |  | January 29, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | There is a possibility that the device may exhibit Hydrogen Embrittlement in the material and are subject to potential breakage. | Process control | Recall Notification Letters were sent to customers on January 29, 2010. Customers were requested to complete and return a Customer Response Form indicating the number of products available for return and number distributed. Each distributor was asked to notifiy their customers if the product was further distributed. |  | 149 | 149 | Distributed throughout U.S. and Internationally to Arab Emirates, Belgium, Canada, China, Columbia, Cyprus, Germany, Denmark, Egypt, Spain, France, United Kingdom, Greece, Hong Kong, India, Italy, Japan, Republic of Korea, Lithuania, Morocco, Malaysia, Poland, Portugal, Saudi Arabia, El Salvador, Thailand, Taiwan and Viet Nam. Domestic Distribution to CA, FL, GA, IL, IN, KS, MI, MT, NY, OH, PA, TX, and WA. | Terminated | January 12, 2011 | 348 | 149 | Not_Computer
Z-2588-2010 | 55708 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Integra Suction Reservoir Kit for use with Integra NeuroSciences Subdural Drainage Catheter; |  |  | May 07, 2010 | 2010 | Integra LifeSciences Corp. | The sterility of the product may be compromised. Integra LifeSciences Corp. received a recall notification from Saint-Gobain Performance Plastics, France, contract manufacturer of the Integra Suction Reservoir with Anti-Reflux Valve, that the product's sterility may be affected by a sterilization issue discovered by the supplier. | Other | Integra sent Urgent Product Recall letters dated May 7, 2010 to customers identifying the affected product, the sterility issue, and actions to be taken by the customer. Customers are instructed to quarantine and return all affected inventory to Integra's distribution center in Sparks, NV, using a Returned Material Authorization Number received from Integra LIfeSciences Customer Service at 800 654-2873. All Customers are requested to return the attached "Recall Acknowledgement. and Return Form" Integra LIfeSciences can be contacted at 609-936-2495. |  | 313 | 313 units | Worldwide Distribution: USA, including the states of AK, AL, AR, AZ, CA, CO, FL, GA, IL, MD, MI, MO, NC, NH, NJ, NY, OH, PA, RI, SC, TX, UT, VA, WI, and the countries of Australia, Chile, Costa Rica, and Sri Lanka. | Terminated | August 27, 2012 | 843 | 313 | Not_Computer
Z-2609-2010 | 56160 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Engenex Non Acute Care Suction Pump for use with Engenex Negative Pressure Wound Therapy |  |  | July 09, 2010 | 2010 | Convatec Inc. | Reduced vacuum capacity has been reported with a small number of units caused by electrostatic discharge (ESD) or static electricity. | Device Design | Recall notification letters were sent to 46 customers, acute care facilities and durable medical device distributors on July 9, 2010, by Federal Express. |  | 1190 | 1190 units | Nationwide distribution. | Terminated | June 07, 2011 | 333 | 1190 | Not_Computer
Z-2586-2010 | 56507 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Stryker Epic II Critical Care Bed, Model 2030/2031, Stryker Medical, Kalamazoo, MI. |  |  | May 14, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The Fowler mechanism may malfunction, preventing it from being lowered into a flat position. | Component change control | Stryker Medical sent an Urgent Medical Device Alert letter dated May 14, 2010 to their customers. The letter identified the product, the problem, and the actions customers should take. The letter instructed users to do the following:" Return the postcard included with the letter to confirm receipt of the notification." Locate the recalled beds, verify the Fowler mechanism is functioning properly and does not display uneven motion when raising and lowering, and is not unusually noisy. " If the Fowler mechanism is not operating correctly, they were to call Stryker at 1-800-STRYKER, option 1, Reference RA-2010-014 to obtain service for the bed. " If any beds have been loaned or sold, the customer is to forward a copy of the notice to the new users and advise the firm of the new location." Provide the firm with the serial numbers for any beds that were disposed of." Retain the revised maintenance manual included with the notification.If the customer had any urgent questions or concerns, they were instructed to call the firm at 800-STRYKER, option 1, reference RA-2010-014, Monday-Friday 8am to 5pm (EST). |  | 9134 | 9,134 | Worldwide Distribution - USA (all states), and including the countries of ARGENTINA, BRAZIL, CHILE, CHINA, CANADA, MEXICO, GREECE, JAPAN, KOREA, PUERTO RICO, SPAIN, POLAND, TAIWAN, SINGAPORE, INDIA, SWITZERLAND, UK and MALAYSIA. | Terminated | June 13, 2011 | 395 | 14060 | Not_Computer
Z-2608-2010 | 56709 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Langston Dual Lumen Catheters, 6F, REF 5540, Sterile EO, RX only, Vascular Solutions, Inc, 6464 Sycamore Court, Minneapolis, MN 55369, 110 cm length, |  |  | September 02, 2010 | 2010 | Vascular Solutions, Inc. | Investigation of a recent Device Experience Report has made us aware of a potential problem with our 6F Langston dual lumen pigtail catheter (Model 5540) with the following lot numbers: 548110, 548227, 548364, 548430, 548554 &amp; 548723.Recently, it was reported that pouches containing Langston Pigtail catheter (Model 5540) were missing a seal, posing a risk of product contamination. It is possib | Process control | Consignees were sent on 9/2/10, a Vascular Solutions "Urgent Medical Device Field Action" letter dated September 1, 2010. The letter was addressed to Cath Lab Manager. The letter descried the product and the problem. Provided instructions to examine pouches of opened boxes to see if they have not been sealed. If unsealed, they recommended to remove pouches from the current inventory. For unopened boxes they recommended to remove the boxes from the inventory. Requested consignees to complete the Field Action Customer Inventory Form included with the letter. For any questions, please contact the firm at 763-656-6032. |  | 1128 | 1,128 boxes ( 5pk) | AL, AK, AZ, AR, CA, CO, CT, DE, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS,MO, NE, NY,NH, NJ, NM, NY,NC, OH, OK, OR, A, PR, SC, TN, TX, UT, VA, WA, WV, WI, and WY.SINGAPORE, MEXICO, CANADA, SPAIN, AUSTRIA, GERMANY, ITALY, ENGLAND, and SWITZERLAND. | Terminated | April 24, 2012 | 600 | 1128 | Not_Computer
Z-2589-2010 | 56543 | No | MGR |  |  | General &amp; Plastic Surgery | PMA | Apligraf is supplied as a living, bi-layered skin substitute. |  |  | August 02, 2010 | 2010 | Organogenesis, Inc. | Product is contaminated with a yeast identified as Pseudozyma antarctica. | Process control | Organogenesis notified accounts by telephone, e-mail or fax on 8/2/010 to remove from use and return recalled units. If unit was applied to patient then monitor for adverse events. The Notification Letter dated August 6, 2010, described the issue and provided instructions for customers. Customers are to place the affected unit on hold for its return to the firm. The firm will contact customers to make arrangements for its return. Replacements have already been sent. If the affected unit has already been applied, the patients should be monitored closely for any potential adverse events. An Organogenesis Medical Affairs rep will be in contact regarding patient follow up. Customers are to complete the attached form and fax it back to the number provided. Questions should be directed to the Medical Information Center at 1-888-432-5232. |  | 32 | 32 units | Nationwide Distribution -- AL, AZ, CA, CO, IL NC, PA, TN, and VA. | Terminated | March 10, 2011 | 220 | 32 | Not_Computer
Z-2598-2010 | 56522 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 8 MM DISPOSABLE CORING TREPHINE WITH PROTECTION SLEEVE, REF 900728, Sterile, BIOMET Sports Medicine, Warsaw, IN 46581. |  |  | July 01, 2010 | 2010 | Biomet, Inc. | The affected instruments are discolored and/or have a sticky residue on the end. | Packaging change control | The firm sent out "Urgent Medical Device Recall" notices dated 7/1/2010. Letters were sent to distributors, customers and OR managers. The letters stated that the devices should be located, quarantined and shipped back to the recalling firm. On 7/15/2010, the firm added one lot of one additional product to the recall (product 900788 lot 780470). The firm sent letters to physicians that had used that lot of product.Questions relating to this recall should be directed to 574-371-3755. |  | 100 | 100 | US, Canada, Australia, Chile, Argentina, Netherlands, Costa Rica and Japan. | Terminated | August 20, 2012 | 781 | 601 | Not_Computer
Z-2591-2010 | 56695 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump2 250 mL PainPump with Luer Lock Tubing Set, sterile, REF 525-350, Stryker Instruments, Kalamazoo, MI. |  |  | September 02, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | All of the failures lead to the patient receiving less medication than intended. Certain lots of the pain pump have a high failure rate associated with a cracked port housing which leads to internal or external leaks. Some lots have a high failure rate associated with E2 errors. | Nonconforming Material/Component | The firm sent our Urgent Medical Device Recall Notices dated 7/12/2010 to Risk Managers and Materials Managers. The recall notice listed specific lots of PainPump2 and PainPump2 BlockAid that were under recall. Consignees were told to check stock for the recalled product, quarantine the product, and contact Stryker to arrange return shipment of the recalled product.On 9/2/2010 the firm sent out a second Urgent Medical Device Recall Notice. This notice expanded the scope of the recall to include all products distributed within the past 12 months. The shelf life of the product is 1 year -so the recall was for all product currently in the supply chain. The letter again directed consignees to locate the recalled products and contact Stryker to arrange return shipment of the recalled product. The firm sent an additional letter dated 9/1/2010 to its customers. This letter stated that Stryker will no longer be offering any of their PainPump2 products. |  | 54568 | 54,568 all products | Worldwide Distribution -- US, Canada, and Australia. | Terminated | April 21, 2011 | 231 | 54568 | Not_Computer
Z-2607-2010 | 56485 | No | CFL | Human growth hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IGFBP-3 Immunoradiometric Assay (IRMA) Kit, Part |  |  | August 09, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed customer reports that IGFBP-3 measurementsusing IGFBP-3 IRMA kit (REF DSL-6600) lots identified above may result in IGFBP-3 levels that are not in agreement with simultaneously measured levels ofIGF-1. The investigation confirmed that above mentioned lots of IGFBP-3 kit provide results elevated by 84% on average. Erroneously elevated IGF | Pending | The recall communication was initiated with a Product Corrective Action (PCA) letter that was sent on August 09, 2010 to allcustomers who received the IGFBP-3 IRMA kits, lot numbers referenced. The letter will provide the customers with an explanation of the problem identified and an action to be taken.Customers were instructed to:(1) Discontinue using all lots identified above and discard in accordance withlocal regulations.(2) To continue testing for IGFBP-3 levels, find an alternate method.(3) Immediately notify all parties of interest (e.g., laboratory staff, clinicians) ofthe possibility of erroneously elevated IGFBP-3 results for patient samples tested with these lots, and advise all parties that clinical decision making and patient safety may be impacted.(4) Review patient IGFBP-3 results for the lots listed above, and interpret them in the light of additional tests (e.g., IGF-1) and the patient's clinical status and history.Customers were also asked to share the recall information with their laboratory staff and retain this notification as part of their laboratory Quality System documentation. Complete and return the enclosed response form within 10 days. For any questions regarding this notification, customers were to contact the DSLTechnical Support Center at 1-800-231-7970 option 4 in the United States, or contact your local Beckman Coulter Representative. |  | 63 | 63 | Nationwide in US and Canada | Terminated | August 13, 2012 | 735 | 63 | Not_Computer
Z-2596-2010 | 56366 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MODEL P8400 HANDLE WITH INTEGRATED CABLE, FOR USE WITH MODEL P8400 CONTROLLER, REF 7013-8050, QTY (1), STERILE EO, HEMOSTATIX MEDICAL TECHNOLOGIES, LLC, BARTLETT, TN 38133 USA, CE 0473, EU REP : QFI LTD, STRATFORD, LONDON E15 2ST UK |  |  | July 27, 2010 | 2010 | Hemostatix Medical Technologies LLC | Lack of sterility assurance | Other | On July 27, 2010, the firm, Hemostatix Medical Technologies, LLC, notified their international distributor in Japan of the recall by email. Hemostatix sent an "URGENT: MEDICAL DEVICE RECALL" letter dated July 29, 2010, to U.S. customers delivered via FedEx Priority Overnight. All non-responding customers were to be notified by telephone after 2 weeks. The other international customer, Italy, were to be notified by email on July 30, 2010. The letter describes the product, problem and action to be taken by customers. The customers were instructed to immediately check their inventory and remove any affected product found; ensure that satellite facilities/clinics receive the recall notice; complete and return the Recall Response Form via fax to Hemostatix's international fax number at 901-261-0017, and contact the Customer Care Department's recall hotline at 901-261-0012 for further instructions, if they do have affected stock.If you have any questions or concerns, please feel free to contact our Customer Care Department at: 1-800-722-5735 in the USA or 1-901-261-0012 outside the USA. |  | 40 | 40 units | Worldwide distribution: USA including CA, GA, LA, MA, MD, MO, NC, OH, OR, PA, SC, TX, VA, and WI; and countries of Japan and Italy. | Terminated | November 29, 2010 | 125 | 259 | Not_Computer
Z-2606-2010 | 56322 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Centrifugal Pump with X-Coating, packaged within custom ordered Cardiovascular Procedure Kits, P/N 3ZZ164275X, 3ZZ164275, CEOT028, CEOT032; and as single sterile units with product codes CV-164275, 164275 and 164275X. |  |  | March 15, 2010 | 2010 | Terumo Cardiovascular Corporation | Medical device component may malfunction and require user intervention during surgical bypass procedures. | Other | The recalling firm notified end users, perfusion groups, and distributors by phone, fax, email and mail on 03/15/10 to advise of the required recall actions with instructions for labeling and returning the involved product from the customer level. The firm expanded their recall to include additional lots on 05/17/10. If you have any questions, please contact Customer Service at 800-521-2818. |  | 28257 | 28257 pumps | The product was sold to medical device distributors and to direct accounts nationwide and to a subsidiary in Japan. | Terminated | November 04, 2011 | 599 | 28257 | Not_Computer
Z-2624-2010 | 50095 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GMRS Extension Piece, 60 mm; Sterile |  |  | May 05, 2005 | 2005 | Stryker Howmedica Osteonics Corp. | Mislabeled package: Package labeled as GMRS Extension Piece 60 mm, lot ZYCOC1 actually contains a GMRS Extension Piece 70 mm extension piece. | Labeling mix-ups | Stryker issued Important Market WIthdrawal letters dated May 5, 2005 to customers and distributors. The letters identified the affected product, the labeling issue, and actions to be taken by the customer or distributor. Customers were instructed to identify and quarantine affected product, which will be reconciled and returned by the Stryker sales representative. Distributors were instructed to identify affected product, reconcile all product on an attached Product Accountability Form to be faxed to Stryker at 201 831-4825 within two days of receipt of the letter. The affected product was to be returned using the enclosed pre-paid mailing label. Contact the Stryker sales representative to re-order product that is being returned.Stryker can be contacted at 201 831-5825. |  | 5 | 5 units | Nationwide Distribution: USA, including the states of AZ, FL, MD, and PA.G | Terminated | October 01, 2010 | 1975 | 5 | Not_Computer
Z-2621-2010 | 55188 | No | CGR | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems CORTISOL, Part Number: 33600 |  |  | February 17, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated by Beckman Coulter after the firm confirmed customer reports that the affected lot of Access Cortisol reagent (REF 33600) identified above may produce suppressed patient results. Beckman Coulter also confirmed customer reports that the two lots of Access Progesterone reagent (REF 33550) identified above may produce suppressed patient results and an increased incidence of I | Other | Two (2) Product Corrective Action (PCA) letters with attached Customer Response form were sent on the week of February 15, 2010 to the affected customers. The letters provide the customers with an explanation of the problem identified and instruct the customers to discontinue use of Access Progesterone and Access Cortisol Reagents, lot numbers referenced. For product replacements consignees were instructed :(1) In the United States, please contact Customer Service at 1-800-526-3821; (2) In Canada, please contact Customer Service at 1-800-463-7828;(3) Outside of the United States and Canada, contact your local Beckman Coulter Representative.If consignees need assistance or have any questions regarding this notification, they were instructed to contact TechnicalSupport at 1-800-854-3633 in the United States and Canada. Outside the United States and Canada, contact your local Beckman Coulter Representative. |  | 7282 | 7282 units | United States &amp; Canada | Terminated | April 19, 2012 | 792 | 13694 | Not_Computer
Z-2618-2010 | 54617 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Jazzy Select with XLR Port Harness |  |  | February 01, 2010 | 2010 | Pride Mobility Products Corp | Improper set-up of the wire harness can allow wires to be pulled from their intended position and become caught under the motor mounting bracket located under the powered wheelchair. This can cause the wire connector and/or the wire itself to be pulled or pinched which can lead to a short or open circuit. | Other | The recalling firm issued recall notifications with response cards to their dealers. Included with the notification was fasteners and an example letter for customer notification of the problem. Dealers were instructed to contact the end users to make arrangements with the customer/end user to verify that the devices were properly set up. If the dealers' inspection revealed that additional fasteners were required they dealer would install them. The recalling firm also issued a recall letter to consumers. The consumer recall letter included instructions that if the correction was completed by anyone other than the wheelchair provider, then the consumer should schedule an appointment with the wheelchair provider to ensure the correction was completed properly. |  |  |  | The product was shipped to dealers nationwide who further shipped the product to end users. The product was also shipped to distributors in Canada, Australia and United Kingdom | Terminated | November 20, 2014 | 1753 |  | Not_Computer
Z-2622-2010 | 56110 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Stryker Scope Warmer, disposable. Endoscopic accessory composed of polyvinyl Chloride (PVC) and Sodium Acetate, Model Number 502-360-000, manufactured by Stryker Endoscopy, San Jose, CA. |  |  | June 21, 2010 | 2010 | Stryker Endoscopy | Outer packaging seals have inconsistencies that have resulted in compromised seals and may result in compromised sterility. | Process control | Stryker issued an Urgent Device Removal letter, dated June 21, 2010, to customers which identified the affected product, the sterility issue, and the actions to be taken by the customer. Customers were instructed to identify and quarantine any affected product in their inventory, and contact Stryker Customer Service at 800 624-4422 and request an RMA#. The RMA# is to be written on the outside of the package used to return products, using an enclosed pre-paid shipping label. All cusotmers are to complete and sign an enclosed "acknowledgement of receipt" and fax to (408) 754-8278.Stryker can be contacted at 800-624-4422. |  | 191855 | 191855 units | Worldwide Distribution: USA, Including the states of AK, AL, AR, CA, CO,CT, DE, DC, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MS, MO, MT, NE, NJ, NM, NH, NV, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY, and the countries of Australia, Brazil, Canada, Chile, China,France, Germany, Iberia, India, Italy, Netherlands, New Zealand, Japan, Korea, Latin America, Polska, Singapore, South Africa, Switzerland, Thailand, and United Kingdom. | Terminated | December 16, 2010 | 178 | 191855 | Not_Computer
Z-2533-2010 | 56548 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | All Med Medical 22ga x 1.5" Standard Bore Winged Administration Set |  |  | August 30, 2010 | 2010 | Multi-Med, Inc. | Huber needles may core and result in port leakage or emboli being flushed our of the port. | Device Design | Multi-Med issued an Urgent: Medical Device Recall notification on 8/30/10 via fax and followed by hard copy letter. Users were advised of the problem and asked to immediately examine their inventory and quarantine the affected product. They were to also notify their customers of the recall. The firm stated in the letter that they would contact customers to determine the number of product that they had and the amount they expected from the field. The firm will coordinate a plan for the return of the affected product. Questions should be directed to Sue Starkey at 603-357-8733. |  | 672 | 672 units | Worldwide Distribution -- USA and Mexico. | Terminated | January 31, 2013 | 885 | 546873 | Not_Computer
Z-2623-2010 | 56261 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Strattice Reconstructive Tissue Matrix for Stoma Reinforcement. |  |  | July 12, 2010 | 2010 | LifeCell Corporation | Mislabeled -One lot of product labeled as 8 cm x 8 cm were actually cut using a 6 cm x 6 cm template. | Employee error | LifeCell issued Recall Notice letters dated July 12, 2010 to customers, identifying the affected product, the labeling issue, and actions to be taken by the customer. Two versions of the recall notification were sent, one for implanted units and one for those not yet implanted. Both letters stated that the mislabeling has no impact of the safety or effectiveness of the medical device. Replacement units were offered for all affected units. Those with unused units were to contact LifeCell Customer Solutions for return and replacement of the unit at 1-866-423-2433.Those with implanted units may contact LifeCell Customer Solutions at 1-866-423-2433 with any additional questions or concerns. |  | 32 | 32 units | Nationwide Distribution: USA, including the states of AZ, DC, FL, GA, ID, IL, KY, LA, MO, NC, OH, OK, OR, TX, WA, and WI. | Terminated | October 01, 2010 | 81 | 32 | Not_Computer
Z-2626-2010 | 56646 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips M8102A IntelliVue MP2 |  |  | August 31, 2010 | 2010 | Philips Healthcare Inc. | Speakers on the Philips IntelliVue X2 and MP2 patient monitors may fail, causing absence of an audible alarm and delaying patient treatment. | Process design | Philips issued a URGENT - MEDICAL DEVICE letter dated August 26, 2010, to each affected customer. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue and gives instructions on actions to take while awaiting their replacement speaker assemblies. The speaker replacement assembly is expected to be available in 8 weeks.Affected products may continue to e used while awaiting replacement of the speaker, provided that the user routinely checks for the display of the "Speaker Malfunct." INOP at power-up.If you experience this INOP or there is no sound from your X2/MP2 speaker while it is not connected to another audible alarm annunciation system, remove the device from use and contact your local Philips service representative at 1-800-722-9377. |  | 2832 | 2,832 UNITS | Worldwide Distribution - USA and the countries of : ARGENTINA AUSTRALIA AUSTRIA BAHRAIN BANGLADESH BELGIUM BRAZIL CHINA CZECH REPUBLIC DENMARK EGYPT ESTONIA FINLAND FRANCE GERMANY GREECE HONG KONG HUNGARY ICELAND INDIA INDONESIA IRAQ IRELAND ISRAEL ITALY JAPAN JORDAN KENYA KUWAIT LATVIA LIBYAN ARAB JAMAHIRI LITHUANIA MALAYSIA MAURITIUS MEXICO MOROCCO NAMIBIA NETHERLANDS NEW ZEALAND NORWAY OMAN PAKISTAN PALESTINE PARAGUAY POLAND PORTUGAL QATAR ROMANIA RUSSIA SAUDI ARABIA SINGAPORE SLOVAKIA SLOVENIA SOUTH AFRICA SOUTH KOREA SPAIN SRI LANKA SWEDEN SWITZERLAND SYRIAN ARAB REPUBLIC TAIWAN THAILAND TUNISIA TURKEY TURKMENISTAN UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN and VIET NAM | Completed |  |  | 22232 | Not_Computer
Z-2627-2010 | 56505 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Gemstar Pump Set, Latex-Free Extension Set with Integral Pressure-Activated Anti-Siphon Valve and Y-Extension with Backcheck Valve-SL, Nonvented, 110 inch; Rx, sterile, 24 individually packaged sets per case; Hospira, Inc., Lake Forest, IL 60045 USA, Made in Costa Rica; list number 12640-01 |  |  | August 12, 2010 | 2010 | Hospira Inc. | Underdelivery of infusion during clinical use at low rate settings (less than 10 mL/hour). Patient may not receive the intended amount of medication. | Nonconforming Material/Component | Hospira issued Urgent Device Recall letters dated August 12, 2010 to Hospira's direct accounts, informing them of the potential for under delivery under low rate settings using the affected lots of pump sets. The accounts were instructed to examine their inventory for the affected product, quarantine it immediately, and complete the attached reply form indicating the amount of product being returned, faxing it to Stericycle at 1-877-523-9110. If they distributed the affected product further, they were instructed to sub-recall the product from their sub-accounts. Accounts may contact Hospira Customer Care at 1-877-946-7747 or their Hospira Account Manager to obtain replacement product. For medical inquiries, the accounts were directed to contact Hospira Medical Communications at 1-800-615-0187. Customers may contact Stericycle, which is managing this recall, at 877-274-7163 for further information concerning this recall. |  | 22680 | 22,680 sets | Worldwide Distribution: USA, (all states), and the countries of Canada, Costa Rica and Taiwan. | Terminated | June 21, 2012 | 679 | 340728 | Not_Computer
Z-2617-2010 | 56413 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Cardiovascular Procedure Kit |  |  | June 08, 2010 | 2010 | Terumo Cardiovascular Corporation | Medical device vent port may be occluded and prevent delivery of priming fluid. | Equipment maintenance | Consignees were contacted on June 8, 2010, by phone, fax, email and priority mail to advise distributors and end-users of the potential occluded vent port defect and provide instructions for the replacement product recall. Consignees were mailed follow-up communications by priority overnight mail on June 9, 2010. |  | 60 | 60 kits | The products were distributed nationwide. | Terminated | February 29, 2012 | 631 | 3979 | Not_Computer
Z-2590-2010 | 55173 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel DxC 880i, 860i and 600i Systems; Software Part Number: A84500, A86646; Software |  |  | February 25, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall on their UniCel DxC Clinical Systems (880i, 860i,680i and 660i Systems with serial number &lt;2760 that have software version 4.9.01 installed and have the VME Motorola ICS board) because they may exhibit issues with stat sample rack loading, causing a possible delay in results. | Other | Beckman Coulter sent an URGENT: FIELD SAFETY NOTICE letter, dated March 25, 2010, to all customers. The letter identified the affected product, the problem, and the actions customers take.Customers were instructed to:if the UniCel DxC system serial number was less than 2760 the customer should contact their local representative to arrange for an upgrade to software version 4.9 build 02.Complete and return the enclosed response form within 10 days of the notice.Share the information with their laboratory staff and retain the notification as part of the Quality System documentation. If affected products were forwarded to another laboratory, provide a copy of the notification letter to them.Customers may contact the Customer Support Center (Hotline) at 1-800-854-3633. For questions or comments please contact the firm at (714) 961-3634 |  | 22 | 22 | Class II Recall - USA, including the states of CA, IN, GA, NC, NY, OH, SC, and WI | Terminated | July 02, 2012 | 858 | 22 | Software | Output/Calculation | Software update | Software Update
Z-2653-2010 | 56247 | No | DEW | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | SYNCHRON System(s) Immunoglobulin A (Ig-A) Reagent, Part Number: 467920 |  |  | June 28, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter Inc. confirmed reports of high recovery of immunoglobulin A (IgA) in proficiency surveys or patient samples using Synchron Ig-A reagent lots M812164 or M902345. Using these reagent lots, some samples deficient in IgA have yielded results in the normal range.Other samples with normal levels of IgA have yielded falsely high results. | Other | A Product Corrective Action (PCA) letter was sent on the week of June 28, 2010 to the affected customers. The letter provided the customers with an explanation of the problem identified and advised customers to discontinue use of the affected lots of the SYNCHRON System(s) Immunoglobulin A (Ig-A) Reagent.Customers were instructed to:(1) Discontinue use of Ig-A reagent lots M812164 and M902345. Contact your Customer Service Representative at (800) 526-3821 in the United States or (800) 463-7828 in Canada or contact your local Beckman Coulter Representative to discuss reagent replacement options.(2) You may need to re-evaluate samples for IgA that were tested with either of the listed lots of Ig-A reagent.(3) Please complete and return the enclosed response form within 10 days so that we may maintain our records.Customers were also instructed to share this information with your laboratory staff and retain this notification as part of yourQuality System documentation. If you have forwarded the affected products listed above to another laboratory, please provide a copy of this letter to them. If they have any questions regarding this Product Corrective Action letter, customers were told to contact our Customer Support Center (Call Center) on the Beckman Coulter website at www.BeckmanCoulter.com/Customer Support/Support Requests, or call (800) 854-3633 in theUnited States and Canada. |  | 852 | 852 units | US Nationwide &amp; Canada | Terminated | July 31, 2012 | 764 | 852 | Not_Computer
Z-2634-2010 | 55680 | No | LJX | Leukocyte peroxidase test. | Hematology | Hematology | 510(K) Exempt | Hypoguard, DiaScreen 4pH, Reagent Strips for Urinalysis, 100 strips, Cat. No. D11400. |  |  | March 24, 2010 | 2010 | ARKRAY USA INC. | This recall has been initiated due to a labeling issue with the Hypoguard DiaScreen 4-way pH 100 strips. The front and back panel of the bottle carton lists five reagents when it should only list four reagents. The Ketone reagent should not be included on the label. The test strips contained in the bottle and all other labeling are correct and properly align with 4pH part number D11400 listed on | Labeling False and Misleading | Distributors were sent a Arkray "URGENT: Voluntary Medical Device RECALL " letter dated March 24, 2010. The letter described the problem and the product being recalled. The letter requested the distributors to immediately examine their inventory , quarantine and return the product subject to the recall. If the distributors had further distributed the recalled product, then they were requested to notify the customers of the recall. Arkray enclosed a customer recall letter that the distributors were to use for notifying their customers.Distributors were also contacted by phone on March 24, 2010, and were informed of the recall issue, the product involved, instructed to quarantine the product and to notify their customers. |  | 5504 | 5504 | FL, NJ, MI, NV, FL, OH, GA, WI, NY, KY, LA, CA, CT, TN, ILISRAEL | Terminated | March 08, 2012 | 715 | 5504 | Not_Computer
Z-2636-2010 | 56220 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Glenoid Head Inserter (RSP Inserter/Impactor), Device Part #804-03-041, All Lots. |  |  | July 12, 2010 | 2010 | Encore Medical, Lp | Complaints indicate product's distal threaded tip may fracture under certain circumstances. | Device Design | DJO Surgical sent an "Urgent Field Safety Notice" Letter dated July 12, 2010, to all of their consignees. The letter described the product, the problem, and the action consignees should take. Consignees were instructed to: Contact their Customer Service Representative for an R-RPR (RMS) number and return all RSP Glenoid Head Inserters, 804-03-041, under the assigned number to receive credit. Complete and return the Field Safety Notice Response to the firm. Identify and inform any customers who were sold or received the product of the Field Safety Notice Any questions consignees should call the firm at (512) 834-6330. |  | 104 | 104 | Worldwide Distribution - USA, including the states of AR, AZ, CA, FL, GA, ID, IL, IN, LA, MA, ME, MI, MO, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, WA, and WI; and the countries of Puerto Rico, Germany, and Saudi Arabia. | Terminated | June 30, 2011 | 353 | 104 | Not_Computer
Z-2651-2010 | 56304 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GYRUS ACMI, dissector PlasmaKnife" , REF 7035-3005, Rx only, STERILE R, CE 0344, 161416-CA, MANUFACTURED FOR; GYRUS ACMI, INC., 136 TURNPIKE ROAD, SOUTHBOROUGH, MA 01722-2104, USA, EC REP, GYRUS MEDICAL LIMITED, FORTRAN ROAD, ST. MELLONS, CARDIFF, CF3 0LT, UK |  |  | August 06, 2010 | 2010 | Gyrus Acmi, Incorporated | Lack of sterility assurance due to compromised package seals. | Packaging | The firm initiated their recall to their direct accounts by Priority Mail with delivery confirmation on 08/06/2010. The letter requested return of the product and was extended to the retail level. |  | 4244 | 4244 units | WorldwideIn the United States: AL, AK, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, ME, MI, MN, MO, MS, NC, NE, NJ, NV, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV and the District of Columbia. Foreign: UK, TAIWAN, SOUTH KOREA, GERMANY, CANADA, AUSTRALIA, SPAIN, ITALY, ISRAEL, NORWAY, LITHUANIA, SWEDEN, SWITZERLAND, SAUDI ARABIA, TURKEY, and GREAT BRITAIN | Terminated | January 26, 2012 | 538 | 4244 | Not_Computer
Z-2638-2010 | 56463 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra CUSA Ultrasonics NXT Ultra Surgical Aspirator System |  |  | July 29, 2010 | 2010 | Integra LifeSciences Corp. | A ground fault malfunction at the rectifier retaining screw may cause the aspiration pump in the CUSA NXT Service module to run at maximum speed, and continue to run until the entire system is turned off. This may result in excess aspiration that requires medical intervention. | Device Design | Integra issued Urgent: Product Recall Notification letters dated July 29, 2010 to all customers on July 29, 2010 identifying the affected products, issue prompting the recall, and actions to be taken by the customer. Letters were also forwarded to all Integra Account Reps.Customers were advised to discontinue use of and quarantine the affected service modules. The Surgical Aspirator may be operated without the Service Module, utilizing the hospitals wall aspiration. Company representatives will contact customers to arrange product return and repair of the affected service modules. Customers are to complete and return an enclosed Recall Acknowledgement and Return Form . The firm can be contacted at 609 936-2485. |  | 208 | 26 units in US, 8 in Europe, 1 in Asia (both types of units) | Worldwide Distribution: USA , including the states of AZ, CA, CO, FL, GA, IA, KY, MN, MO, NV, NY, OH, OR, RI, TX, VA, WI, and the countries of Germany, Hungary, Indonesia, Italy, Netherlands, and Switzerland. | Terminated | May 15, 2012 | 656 | 234 | Not_Computer
Z-2641-2010 | 56188 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS Panels Test Strips, CHOL +HDL+GLU Panel for CardioChek PA brand analyzers, 15 tests 1 MEMo Chip, manufacturer Polymer Technology Systems, Inc., 7736 Zionsville Rd., Indianapolis, IN 46268. |  |  | July 06, 2010 | 2010 | Polymer Technology Systems, Inc. | The test strips appear to under-recover Cholesterol, HDL and glucose. | Other | Distributors and direct customers were notified by letter on July 2, 2010. They were told to discontinue use and dispose of the affected lot of test strips. The firm also asked customers to complete an acknowledgement and confirmation document . |  | 5805 | 5805 | US accounts in: CA, MA, FL, NY, KY, VA, UT, MN, OH, and MD. Direct foreign accounts in: Belgium, England, Sweden, and Italy. | Terminated | May 11, 2011 | 309 | 5805 | Not_Computer
Z-2635-2010 | 56292 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EnsealÂż TRIO Product Code ETRIO335H, Lot Code: G4TC14, Exp. Date 2012/05 |  |  | July 09, 2010 | 2010 | Ethicon Endo-Surgery Inc | The weld that connects the electrode to the active rod was separated. This connection is critical to the proper function of these devices. If the weld is separated, the product will not deliver energy to the jaws, resulting in the potential for a cut/transection without a seal. Ethicon Endo-Surgery identified this issue during a routine finished goods quality inspection. | Other | Ethicon Endo-Surgery issued an Urgent: Device Recall, Event 2193 letter dated July 19, 2010 identifying the affected device, the issue prompting the recall, and actions to be taken by consignees.Consignees were instructed:Do not use the affected product.Identify and quarantine affected product from inventory.Complete the Business Reply form and fax to Stericycle at 877-597-9580 within three daysTo return product, place it in a box with a copy of the Business Reply form and mail to Stericycle using the pre-printed UPS shipping label. Replacement product should be received by the customer within 60 days of return of affected product. Customers can contact Stericycle about this recall at 877-243-7314 and referencing Event 2193. |  | 396 | 396 instruments | Nationwide Distribution to the following states: AZ, AL, AZ, CA, CO, CT, FL, GA, IN, KY, MA, MD, MI, MO, MS, NC, NV, NY, OH, PA, TN, TX, WA, and WV. | Terminated | February 14, 2011 | 220 | 396 | Not_Computer
Z-2652-2010 | 56481 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Endo Ocular Directional Laser Probe, Lumenis Inc. Salt Lake City, UT STERILE EO |  |  | July 27, 2010 | 2010 | Lumenis, Inc. | Firm did not have proper 510(K) clearance. | Component change control | Lumenis sent letters to all customers on 7/27/10 instructing them to review their inventory, isolate affected product, complete a recall verification form and return all affected product to them. |  | 2500 | 250 probes (25 boxes, each containing 10 probes) | FL, IL, MD, CO and TX | Terminated | October 05, 2010 | 70 | 2500 | Not_Computer
Z-2654-2010 | 49387 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Laparoscopic Ultrasound Transducer (LAP Transducer) is used on the SonoSite MicroMaxx ultrasound system. The LAP Transducer label is labeled in part: "MicroMaxx LAP...Manufactured for SonoSite, Inc., Bothell, WA 98021". |  |  | November 02, 2007 | 2007 | Sonosite, Inc. | Sterilization of the SonoSite LAP surgical Transducer with the Sterrad NX System may result in damage to the transducer. | Process control | On November 2, 2007, SonoSite sent out the Important Information Regarding Sterrad NX Incompatibility letter and the Technical Bulletin 005/07 Rev A to their consignees. The letter advises the consignees that SonoSite has recently found that sterilization of the SonoSite SLT and LAP surgical ultrasound transducers with the Sterrad NX system may result in damage to the transducer.The Technical Bulletin 005/07 Rev B dated November 2, 2007 informs the consignees that the damage has been observed primarily on the circuit board in the transducer connector, and cable and laparoscope damage may also be occurring. All damages have occurred after processing using the Sterrad NX systems in its "Advanced" cycle and the potential damage from the use of "normal" cycle is being evaluated by the firm.Consignees should call the Technical Support at 1-425-951-1330 or 1-877-657-8118 or e-mail to service@sonosite.com for any questions regarding the recall. |  | 16 | 16 units | Devices were distributed to medical centers and hospitals in AR, CA, IN, KS, MA, MD, MI, NC, NE, NJ, NV, OH, SC, TN, TX, WI, Washington DC and to FOREIGN COUNTRIES Australia, China, France, Germany, Greece, Hong Kong, India, Iraq, Italy, Korea, Republic of, Libyan Arab Jamahiriya, Mexico, Spain, United Arab Emirates, United Kingdom, and United States. | Terminated | October 05, 2010 | 1068 | 108 | Not_Computer
Z-2639-2010 | 54382 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | One Touch Ultra Test Strips, Part Number 020-245-32, lot number 2964512, Manufactured by LifeScan Inc., Milpitas, CA. |  |  | January 07, 2010 | 2010 | Lifescan Inc | One Touch Ultra Test Strip products may actually contain a different product, One Touch Select strips. | Mixed-up of materials/components | Notification was issued by the firm during the first week of February. This letter asked consignees to share the information with patients, and are asking users to contact the firm and to examine their stock for the recalled product. |  | 38524 | 9631 cartons, 4 vials per carton | Product was distributed to one wholesale dealer and one distributor and 2 medical facilities in CA, OH and PA, as well as the UK. | Terminated | March 14, 2011 | 431 | 38524 | Not_Computer
Z-2658-2010 | 55196 | No | CEC | Free thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Access Immunoassay Systems Free T4, Part Number: 33880 |  |  | March 24, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter determined that kit boxes for Access Free T4 reagent (REF 33880), lot number 916293 may contain reagent packs labeled as Access Vitamin B12, (REF 33000), lot 916277. Using an Access Vitamin B12 reagent pack, lot 916277, filled with Access Free T4 reagents may generate incorrect patient results. | Other | A Product Corrective Action (PCA) letter was sent on March 24, 2010 to the affected customers. The letter provides the customers with an explanation of the problem identified and instructs them to discard the FT4 kit, Lot 916293 in their stock if the FT4 packs within the kit packaging are labeled asVitamin B12 Lot 916277.The PCA requested customers:(1) Inspect all Access Vitamin B12 reagent packs that have been removed from their kit boxes, including reagent packs currently loaded on your system and discard any reagent packslabeled as Access Vitamin B12, lot 916277;(2) Inspect your Access Free T4 reagent kit inventory for any reagent packs labeled as Access Vitamin B12, lot 916277 and discard any reagent packs labeled as Access Vitamin B12, lot916277; (3) Review your historical quality control to ensure that assay performance for Access Vitamin B12 lot 916277 has been within your laboratory specifications.In addition customers were instructed to share this information with their laboratory staff and retain this notification as part of your laboratory Quality System documentation. Complete and return the enclosed response form within 10 days so Beckman Coulter can be assured that you have received this important notification.If customers need assistance or have any questions regarding the notification, they were instructed to contact Customer Technical Support Center at 1-800-854-3633 in the United States and Canada. |  | 10826 | 10826 | Nationwide In the US &amp; Canada | Terminated | May 07, 2012 | 775 | 10826 | Not_Computer
Z-2656-2010 | 56565 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | AVI Sterile Impad Rigid Sole Foot Cover, Right Foot |  |  | August 11, 2010 | 2010 | Covidien LP | Covidien issued an Urgent Product Recall of the Sterile ImPad Rigid Sole Foot Cover. The product is indicated for prevention of blood pooling in the limb. 120 cases of each Lot code 6066 (right foot) and 6067 (left foot) were shipped to the distributor before sterilization. The customer is instructed to quarantine the product and return it to the manufacturer for replacement. If there are any | Release of Material/Component prior to receiving test results | Covidien issued an Urgent Product Recall of the Sterile ImPad Rigid Sole Foot Cover. The product is indicated for prevention of blood pooling in the limb. 120 cases of each Lot code 6066 (right foot) and 6067 (left foot) were shipped to the distributor before sterilization. The customer is instructed to quarantine the product and return it to the manufacturer for replacement. If there are any questions please contact, Covidien Customer Service at (800) 962-9888, Option 1. |  | 1984 | 1984 units (496 cases) | Nationwide | Terminated | June 12, 2012 | 671 | 6840 | Not_Computer
Z-2642-2010 | 56593 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Camera Head OTV-S7H-1D-L08E and Camera Head OTV-S7H-1D-F08E (designed to be used with Olympus endoscopic equipment) |  |  | August 20, 2010 | 2010 | Olympus America Inc. | Mislabeled. Endoscopic body-type camera heads labeled as cardiac-type. | Labeling mix-ups | Asked consignees to return camera heads to correct the labeling. |  | 217 | 217 camera heads | nationwide; camera heads were serviced and returned to consignees in March 2003 - September 2007 timeframe | Terminated | September 27, 2012 | 769 | 217 | Not_Computer
z-2640-2010 | 56652 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Sonopet Ultrasonic Surgical System, Model UST-2001, Manufacturer Miwatec Co. LTD, Tokyo, Japan. |  |  | July 15, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments acquired the Sonopet Ultrasonic Surgical System from Mutoh co. Ltd on 3/29/2010. Stryker determined that the Sonopet consode did not meet the IEC-60601-1 standard. Stryker decided to recall the Sonopet consode. | Device Design | The firm to send URGENT MEDICAL DEVICE RECALL NOTIFICATIONS to customers dated 7/15/2010. The customers were instructed to continue to use the device. Stryker technicians were dispatched to replace a board with a new board. |  | 1161 | 1161 | Wordwide | Terminated | October 19, 2011 | 461 | 1161 | Not_Computer
Z-2583-2010 | 56343 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | SynFix-LR Spinal Implant System (Implant Holder for SynFix-LR and surgical technique brochure) |  |  | July 23, 2010 | 2010 | Synthes USA (HQ), Inc. | A potential for the implant holder tip to break off and lodge in the SynFix-LR implant, which pose a potential risk to the patient. | Error in labeling | Synthes Spine Hospital sent an ' URGENT: MEDICAL DEVICE LABELING CORRECTION" letter dated August 20, 2010 to all customers. The letter described the labeling issue and actions to be taken by the customers. The customers were instructed to review the supplementary labeling provided; contact their Synthes Spine Sales Consultant to arrange for provision of the Holding Forceps (p/n 388.401) prior to each SynFix-LR procedure; and complete and return the enclosed Verification Section form to Synthes via fax at 610-251-9005 or email: FieldAction@Synthes.com.If you have any questions, please call 1-800-620-7025 x 3500 or contact your Synthes Spine Sales Consultant. |  |  |  | Worldwide Distribution: USA and the countries of Canada, Japan, Switzerland, Australia, Mexico, Malaysia, New Zealand, and Singapore. | Terminated | December 18, 2014 | 1609 |  | Not_Computer
Z-0003-2011 | 56411 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | BVI 9600 Bladder Volume Instrument, Aorta Scan Mode |  |  | August 23, 2010 | 2010 | Verathon, Inc. | The BladderScan BVI 9600 with AortaScan devices may experience the loss of a factory installed software configuration file which results in a loss of the aorta measurement function. | Software design | Verathon sent an URGENT Medical Device Recall Notification letter dated August 12, 2010 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Verathon provided an addendum to the User Manual with instructions on how to reset the instrument. Customers were advised to:"Add the enclosed addendum to their current Users Manual."Update the application software of the device to v2.7.0 or higher. If the customers had questions or needed support for the software update process, they were to contact Verathon Medical Customer Care at (425) 867-1348 or (800) 331-2313 orcservice@verathon.com. |  | 351 | 351 units | Worldwide Distribution - USA and the countries of Australia, Canada, China, Colombia, European Union, United Kingdom, Saudi Arabia, and Singapore. | Terminated | September 08, 2011 | 381 | 684 | Software | Device Operation | Software update | Software Update
Z-0014-2011 | 50100 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Duracon Total Knee Distal Femoral Locking Screw for 5 mm Spacer; Product No. 6632-5-600. |  |  | October 05, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics became aware that a box labeled as Duracon Total Knee Distal Femoral Locking Screws for 5 mm spacer, may actually contain a 10 mm screw. | Labeling mix-ups | Important Market Withdrawal Letters were sent via Federal Express October 5, 2006. The letters identified the affected product and the reason for recall. Customers were asked to examine their inventory to identify the product; and retrieve all affected product lots and return it their branch or agency warehouse for reconciliation. All product should be reconciled on the Product Accountability Form and the form should be faxed back. Customers are to contact a Customer Service Rep to re-order the product that is being returned. Questions can be directed to Rita Intorrella on 201-831-5825. |  | 7 | 7 units | Nationwide Distribution -- Including Ohio. | Terminated | October 07, 2010 | 1463 | 7 | Not_Computer
Z-0012-2011 | 49138 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Essenta DR is a digital multifunctional X-ray systems, suitable for all routine radiographic exams, including specialist areas like trauma or pediatric work, excluding mammography. |  |  | August 15, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | When a patient is selected in the patient list and examination tab pressed with background query running at the same time, the wrong patient may be selected. | Other | Philip Medical System sent an "URGENT-Device Correction" letter dated August 13, 2008 to their customers. The letter describes the recalled product as Essenta DR Patient and Image Data Software (SW) release 1.0. The firm informs the customers that when patient is selected in patient list and examination tab pressed with background query running at the same time, there is a risk that the wrong patient is selected on the Essenta DR Computed Radiography System.Customers are advised to verify the selected patient's information before releasing an exposure while the firm works on upgrading the Software. All affected units will receive a software upgrade to correct the issue.Customers should contact Philips Customer Care Service at 800-772-9377, #5, #4, #1 and reference FCO 712000031 for any questions or support concerning the recall. |  | 23 | 23 units | Nationwide distribution: CA, GA, FL, KS, MN, NY, OH, OR, SC, TN, TX, UT, and WA. | Terminated | October 07, 2010 | 783 | 23 | Software | Output/Calculation | Software update | Software Update
Z-0013-2011 | 56026 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Summit Broach Handle, NON-STERILE, MFG: DEPUY ORTHOPAEDICS, INC., PO BOX 988 WARSAW, IN 46581, 1 800 366 8143 |  |  | June 15, 2010 | 2010 | Depuy Orthopaedics, Inc. | Firm discovered that some of their straight broach handles are experiencing early failure due to a weakened handle locking mechanism (Leaf Spring). | Process control | On June 3, 2010, an "Alert: Femoral Broach Handle Repair" went out to the entire sales force/customers via "Sales Mail" and posted to the DePuy Internal website. The Alert stated that "We have discovered that some of our straight broach handles are experiencing early failure due to a weakened handle locking mechanism (Leaf Spring).As you perform routine maintenance on your sets, if you have any of these broaches stamped with the lot numbers listed, it is recommended you have them sent in to be reworked with a new spring device."Send the broach handles to: DePuy OrthopaedicsAttn: Repairs700 Orthopaedic DrWarsaw, IN 46581Please include return shipping informationIf you have any questions regarding this information or need additional information, call 574-372-7333. |  | 8358 | 8358 | Nationwide distribution: AR, AZ, CA, FL, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, NC, NV, NY, OH, PA, TN, VA, and WI. | Terminated | November 30, 2011 | 533 | 8358 | Not_Computer
Z-0011-2011 | 56728 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | MAQUET HLM Tubing set with Bioline Coating (BEQ-TOP 21400 with 1/2 x 1/2 HG 0286 Connector; Device Part Number BEQ-TOP 21400 |  |  | September 07, 2010 | 2010 | Maquet Inc. | Tubing packs may contain defective 1/2 inch x 1/2 inch tubing connectors. Leaks at the shank of the connector may result in air aspiration in the venous line. | Device Design | Maquet Getigne Group issued Product Recall/Urgent Medical Device Corrective Action letters dated September 7, 2010 to customers, identifying the affected product, the issue prompting the recall, and actions to be taken by customers.Customers are to check their inventory for affected product. Customers with affected product are to contact Laura Olilver at 201-995-8849 for an RMA number and shipment instructions. All firms are to coomplete and return the Corrective Action Response Form.Maquet Customer Service can bre contacted at 800-777-4222. |  | 22 | 22 items | Worldwide Distribution: USA State of IL and country of the Netherlands. | Terminated | September 25, 2012 | 749 | 22 | Not_Computer
Z-0010-2011 | 56599 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Datex-Ohmeda Amingo Anesthesia System, manufactured by GE Healthcare, P.O. Box 900, FIN-00031 GE, Finland |  |  | August 30, 2010 | 2010 | GE Healthcare, LLC | Users may inadvertently leave a test plug in the inspiratory flow sensor after completing Low P Leak Check on the SCGO (Switched Common Gas Outlet) configuration of Avance, Amingo, and Aisys anesthesia machines that may impact patient safety.If the test plug remains in place and therapy is initiated, the machine will not provide gas to the patient. This would inhibit mechanical or manual ventila | Employee error | GE Healthcare issued an " Urgent Medical Device Correction" letter dated August 30, 2010 to consignees. The letter was addressed to Chief of Anesthesia, Health Care Administrator/Risk Manager and Director of Biomedical/ Clinical Engineering. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Instructions.Customers were instructed to:1. Ensure the plug is removed at the completion of the Low P Leak check.2. Review the Preoperative Checkout procedure listed in section 4 of the User Reference Manual with appropriate personnel. Ensure that the Preoperative Checkout procedure is completed in its entirety prior to each case, or as required.All facilities will be contacted by their local GE GE Field Technical Service Team to schedule installation of a software update and a revised test plug for your machines configured with SCGO. Customers are to contact their local GE Healthcare Customer or Technical Support Representative if they need additional information. |  | 131 | 131 | Worldwide Distribution: USA: All States, including DC, except Delaware, and the countries of YEMEN, VIETNAM, VENEZUELA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, SYRIA, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI, ARABIA, RUSSIA, PORTUGAL POLAND, PERU, PANAMA, NORWAY, NETHERLAND, MACEDONIA, LITHUANIA, LEBANON, LATVIA, KUWAIT, KOREA, KENYA, JORDAN, JAPAN, ITALY, IRELAND, INDIA, HUNGARIA, HONG KONG, HONDURAS, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CZECH REPUBLIC, CROATIA, COLOMBIA, CHINA, CHILE, CANADA, BULGARIA, BRAZIL, BERMUDA, BELGIUM, AUSTRIA, AUSTRALIA, and ARGENTINA. | Terminated | September 13, 2012 | 745 | 7711 | Computer
Z-0021-2011 | 56707 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | The BABIX holder is an accessory used with the BuckyDiagnost X-ray system. It is used to hold the baby either in a secured or immobilized standing position in front of the X-ray detector during the examination. |  |  | September 09, 2010 | 2010 | Philips Medical Systems North America Co. Phillips | Potential of broken strands in the main steel cable supporting the accessory holder used with BuckyDiagnost and DigitalDiagnost X-ray systems. | Equipment maintenance | Philips issued URGENT-Device Correction Notice letters dated September 9, 2010 to their consignees identifying the affected products and action to be taken by customers. The letter states that Babix holder used with Philips BuckyDiagnost and DigitalDiagnost X-Ray System - VS Wall Stand Risk of steel cable fracture in the support holder. The consignees are informed that there is a potential of broken strands at the main steel cable within the holder which cannot be recognized by the user. The firm will issue the Field Change Order (FCO) and the Service Engineer will replace the Babix holder.Consignees can call Philips Call Center at 800-722-9377, #5 (then follow prompts) and reference to FCO 70400038 for the BuckyDiagnost X-ray system and FCO 71200046 for the DigitalDiagnost X-ray system. |  | 4 | 4 units | Nationwide Distribution in the states of CA, GA, and TN. | Terminated | July 22, 2011 | 316 | 5 | Not_Computer
Z-0023-2011 | 56663 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BD 30G1/2 PrecisionGlideÂż Needle;Catalog/Lot Number: 305106. |  |  | August 24, 2010 | 2010 | Becton Dickinson &amp; Company | There have been reports of blocked/clogged needles for the 30 gauge PrecisionGlide Needles | Pending | BD sent out Product Correction letters to all end users of the product on 8/24/10. Letters were mailed by BD via UPS second day air. The letter stated the issue and said that customers should assess the patency prior to injection. The letter also stated that a BD 30 gauge PrecisionGlide Needle from another lot can be used in place of a clogged needle. If customers require replacement needles or further assistance, they can contact BD at 1-888-237-2762. Finally, the letter states that BD has implemented corrective actions to reduce the potential for further recurrence of the blockage issue. |  | 2599000 | 2,599,000 units | Worldwide Distribution -- USA and Canada. | Terminated | August 25, 2011 | 366 | 2599000 | Not_Computer
Z-0020-2011 | 56536 | No | GBS | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | HydraGlide Silicone Chest Tube Catheter, 32 Fr Straight |  |  | August 10, 2010 | 2010 | Atrium Medical Corporation | Mislabeled: One lot of 32Fr Straight Catheters packaged and labeled as 24Fr Straight Catheters. | Packaging process control | Atrium Medical isued Product Recall Notice letters dated August 10, 2010 to customers, identifying the affected product and actions to be taken by customers. Customers were instructed to remove the affected products from inventory and return them to Atrium, and complete and fax a recall notice acknowledgement to Atriuim at 603-821-1420. Customers can contact Atrium at 800 528-7486, x5350 (Joe DePaolo) or x5351 (Karen Hall) for further information. |  | 10 | 1 case, 10/case | Nationwide Distribution: USA, including the states of AK, CT, OH, and UT. | Terminated | April 22, 2016 | 2082 | 110 | Not_Computer
Z-0017-2011 | 56630 | No | KRG | Pacemaker programmers. | Cardiovascular | Cardiovascular | PMA | Guidant, Vitality HE ICD, Model: T180DDR, Sterile Lot, Guidant Corporation, Cardiac Rhythm Management, 4100 Hamline Avenue North, St. Paul, MN 55112-5798. |  |  | July 22, 2010 | 2010 | Boston Scientific CRM Corp | Exposure to a magnet (typically in a clinic/hospital environment) may cause a magnetic reed switch to become permanently stuck in a closed position in certain Boston Scientific CRT-Ds and ICDs (CONTAK RENEWALÂż 3 / 3 RF / 4 / 4 RF and VITALITYÂż HE implantable defibrillators). This prevents delivery of programmed tachy therapy, and use of the Patient Triggered Monitor feature. This issue may require | Nonconforming Material/Component | Boston Scientific Issued Urgent Medical Device Correction and Physician Device Advisory Notice letters dated July, 2010 to customers identifying the affected devices, the issue prompting the recall, clinical implications, and actions to be taken by customers.Customers were instructed:1) Physicians should continue routine follow-up sessions, and patients should be reminded to contact their clinic or go to the hospital emergency room immediately if they hear tones/beeps from their device. 2) In a hospital/clinic/surgery setting, if tones are heard upon magnet application but do not cease upon magnet removal, the device should be interrogated with a programmer and checked per normal standard of care. 3) In the United States, use of the LATITUDE remote patient monitoring system may help identify loss of daily measurements and thereby facilitate timely detection of a stuck reed switch. [NOTE 1: A pop-up message and/or LATITUDE alert do not appear for missing Daily Measurements. NOTE 2: Daily Measurements may be unavailable for reasons other than a stuck switch.] 4) Prophylactic explant is NOT recommended.5) If a stuck magnetic switch is confirmed, program the Enable Magnet Use feature to Off, which ensures that programmed therapy will be provided to treat tachyarrhythmias. However, if Enable Magnet Use is programmed Off, a magnet will no longer inhibit tachy therapy, and the Patient Triggered Monitor feature will no longer be available. Do not routinely program Enable Magnet Use to Off in the absence of a confirmed stuck magnetic reed switch because the benefits of magnet use to disable tachy therapy in emergent situations outweigh the probability of patient harm associated with a stuck reed switch.In the US, Boston Scientific Cardiac Rhythm Management can be contacted at 800-227-3422, and in Europe, Boston Scientific Europe S.A. can be contacted at +32 2 416 7222. |  | 5821 | 5821 | Worldwide Distribution: All States in the continental USA including PR and DC , but excluding VT, the countries of Canada, Mexico, and in Europe.ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BERMUDA, BRITISH VIRGIN ISLANDS, CANADA, CAYMAN ISLANDS, CHILE, COSTA RICA, CYPRUS, CZECH REPUBLIC, CZECHOSLOVAKIA, DENMARK, EGYPT, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KOREA, KUWAIT, LEBANON, LUXEMBOURG, MALAYSIA, MEXICO, NETHERLANDS,NEW ZEALAND, NORWAY, POLAND, PORTUGAL , QATAR, SAINT KITTS &amp; NEOIS, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM,VENEZUELA, URUGUAY, ZIMBABWE, IRAN. | Terminated | April 03, 2012 | 621 | 38377 | Not_Computer
Z-0018-2011 | 56600 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Oxylog 3000, Emergency and Transport Ventilator; Catalog #2M86300; Drager Medical Inc., 3135 Quarry Road, Telford, PA 18969. |  |  | August 19, 2010 | 2010 | Draeger Medical Systems, Inc. | Patients received insufficient ventilation. The Oxylog 3000 Instructions are inadequate relating to dead space volume. | Error in labeling | The firm sent Important Safety Notice Letter, dated August 2010, and Instructions for Use supplement to consignees. The letter explained the issue with the affected product. The letter asks customers to attach the Instructions for Use amendment sheet supplied with the letter to their corresponding Instructions for Use. The firm also states that a they will provide a dedicated pediatric breathing circuit for the Oxylog 3000. Questions regarding the letter should be directed to Michael Kelhart at 1-800-543-5047. Questions regarding the use of the device should be directed to DragerService Technical Support at 1-800-543-5047 or your local Drager Sales Representative. |  | 78 | 78 potentially afftected units | Nationwide Distribution | Terminated | November 18, 2010 | 91 | 78 | Not_Computer
Z-0015-2011 | 56219 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Exodus(TM) 8F (2.7 mm)/25 cm Standard Loop Multipurpose Drainage Catheter with Hydrophilic Coating, REF Catalog No. 10002, UPN Product No. H965100201, Rx ONLY, STERILE/EO --- Made in USA: 2301 Centennial Blvd., Jeffersonville, IN 47130 --- Manufactured for: Navilyst Medical, Inc. 26 Forest Street, Marlborough, MA 01752 USA --- Each box unit contains 5 catheters. --- The Exodus Standard Loop Multipurpose Drainage Catheter consists of a radiopaque polyurethane catheter with a locking pigtail distal end and is coated with a hydrophilic coating. Catheter markings are provided in centimeters on the shaft to indicate the distance form the distal tip. The catheter is provided with a trocar, a metal stiffening cannula and a plastic cannula. |  |  | May 24, 2010 | 2010 | Navilyst Medical, Inc | One box containing five catheters did not complete the sterilization process prior to distribution. | Employee error | On/about 5/24/10, NMI telephoned the consignee. The recall notification package was sent by NMI to the sole consignee on 6/11/10, followed by a telephone call by NMI Quality staff. On 6/15/10, the recall notification package was sent via FedEx to the sole consignee. Subsequently the consignee was provided a return goods authorization (RGA) number for return to NMI. The letter described the product, problem and action to be taken by the consignee/customer. The consignee/customer was instructed to immediately discontinue use of and segregate recalled product; remove the affected lot from their inventory; complete and return the Reply Verification Tracking Form via fax to the QA Manager; Navilyst Medical, Inc., at 1-518-742-4414; to please call the Navilyst Medical Customer Service at 1-800-833-9973 between 8:30 a.m. and 7:00 p.m. (Eastern Standard Time) to obtain a Return Authorization (RGA) Number, and package/ship the recalled product to Navilyst Medical, Inc, Distribution Center, Attn: QA Returns, 10 Glens Falls Technical Park, Glens Falls, NY 12801. If you have any questions, call Navilyst Medical Customer Service at 1-800-833-9973. Subsequent to the initial notification letter, NMI identified that a typo was made to the lot number on the letter. NMI telephoned the consignee and sent a revised letter via fax on 6/25/10 indicating the correct lot number. |  | 1 | 1 box unit (contains 5 catheters) | Nationwide distribution: Ohio only | Terminated | October 07, 2010 | 136 | 1 | Not_Computer
Z-0034-2011 | 50215 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | The device's marketing brochure states: "IntelliVue M3150 Information Center...Philips Medical Systems". |  |  | November 21, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Potential for Internal Sound Amplifier to malfunction which could delay treatment. | Component change control | On 11/20/08, Philips began sending the URGENT-Medical Device Correction letter to their consignees via e-mails. On 11/21/08, the firm sent out the recall letter via United State Postal Service.The letter states the affected device as Philips IntelliVue Information Centers, running on the HP rp5700 PC platform equipped with an active speaker solutions using a DB-15 to 3.5 mm cable set.The consignees are advised to ensure that the audible alarm system is clear, free of noise and distortion and adjustable to different sound levels which is explained in p. 5-47 "Adjusting Alarm Tone Volume" of IntelliVue of the Instruction for Use, Release K Manual, supplied with the hardware, Philips Part Number 453564062301.Consignees can call Philips Healthcare Customer Care Service Center at 1-800-722-9377, #3, #1 and reference FCO 99862012. |  | 424 | 424 units in US. | Nationwide Distribution | Terminated | October 12, 2010 | 690 | 424 | Hardware | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-0027-2011 | 54527 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | SpeedLink 5.5 Transverse Connector Medium 43-51mm, non-sterile, 1302-55M |  |  | January 01, 2009 | 2009 | Zimmer Inc. | The firm received complaints that the device was damaged or fractured while the surgeon was bending it prior to implantation. There were also complaints that the cam may dislodge during tightening. The firm added additional instructions for use to reduce the probability of these two types of events. | Other | The firm provided a revised surgical technique document to the Zimmer spine sales force in January, 2009. The firm will provide an Urgent Medical Device Correction notification to speedlink users through the sales force. |  | 23205 | 23,205 Units for ALL PRODUCTS IN THIS RECALL | The devices are contained in implant cases stocked by Zimmer spine and made available to the Zimmer sales force for surgeries. They are not in general distribution. | Terminated | January 05, 2011 | 734 | 23205 | Not_Computer
Z-0025-2011 | 56767 | No | MJG | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Tru-Core II Disposable Automatic Biopsy Instrument and Tru-Core II URO Disposable Automatic Biopsy Instrument. The Tru-Core Disposable Automatic Biopsy Instrument is used to obtain multiple core samples from soft tissue such as liver, kidney, prostate, breast, ect. The recalling lot's for the product are as follows: |  |  | July 02, 2010 | 2010 | Medical Device Technologies, Inc. | Medical Device Technologies, Inc., dba Angiotech, Gaineville, Florida is recalling Tru-Core II Disposable Biopsy Devices due to compromised packaging integrity. The Tru-Core Disposable Automatic Biopsy Instrument is used to obtain multiple core samples from soft tissue such as liver, kidney, prostate, breast, ect. The recall is for the following Model Numbers: 763114100X, 763114160X, 763114 | Other | Angiotech sent an "URGENT PRODUCT RECALL NOTIFICATION" letter dated July 2, 2010 and response card to all customers including distributors. The letter described the product, problem and action to be taken by the customers. The customers were instructed to review their current inventory and segregate the affected lot numbers; arrange for return and replacement of the affected product, and complete and return the Customer Acknowledgment Form via fax to the Quality Assurance department at 1-352-338-0662 or 1-800-333-0440.If you have any questions concerning this notification, contact Shannon Brooks at (352) 338-0440 ext 355 or Katrenia Williams at ext 353. |  | 11123 | 11,123 boxes (111,230 units) | Worldwide distribution: USA and countries including: Argentina, Australia, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, Israel, Japan, New Zealand, Peru, South Africa, , Uruguay, and Venezuela. | Terminated | June 28, 2011 | 361 | 11123 | Not_Computer
Z-0024-2011 | 56645 | No | KRL | Cardiopulmonary bypass bubble detector. | Cardiovascular | Cardiovascular | 510(k) | Air bubble detection cable assembly, catalog numbers 78-8067-7375-6 and 149892, Terumo Cardiovascular System, Ann Arbor, MI |  |  | October 15, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | Malfunction of the air bubble detection cable assembly causes a false air alarm or a check sensor message during system set-up. In some instances, users were not able to reset the alarm to allow for further air bubble detection. . | Other | The firm sent URGENT MEDICAL DEVICE CORRECTION notices dated September 17, 2010. The notices told consignees that a Terumo representative would visit, inspect the cables and replace the recalled cables. if the customer has a service contract, the visits would occur at the next scheduled preventative maintenance visit. Customers without a service contract will be contacted by Terumo representative to schedule an appointment.Customers can contact Terumo Customer Service at 800-521-2818. |  | 5000 | 5,000 estimated | Worldwide Distribution: USA states and territories of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VA, WA, WI, WV, and WY, and the countries of Australia, Belgium, Canada, Columbia, Chile, Costa Rica, Germany, Dominican Republic, Ecuador, Hong Kong, Indonesia, Japan, Korea, Mexico, Malaysia, Panama, Philippines, Saudi Arabia, Singapore, South Africa, Thailand, Taiwan, UAE and Vietnam. | Terminated | June 29, 2012 | 623 | 5000 | I/O | Alarm/Message | Replace cables | Remove or Replace
Z-0040-2011 | 55602 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | Smith &amp; Nephew Aimer Endofemoral 3mm Offset (Blue) |  |  | April 29, 2010 | 2010 | Smith &amp; Nephew, Inc. Endoscopy Division | Endofemoral Aimer tip may separate due to failure of the soldered joint | Component design/selection | Smith &amp; Nephew sent a "Medical Device Recall Information" letter dated April 29, 2010, with an attached "Urgent-Product Recall 1st Notification - Urgent" response form via Federal Express on April 29, 2010, to all customers. An e-mail was sent to the contacts in the identified OUS countries alerting them of the recall. ( Note: On June 24, 2010, the firm expanded the recall to include 4 additional catalog numbers 7220l713F, 72201714F, 72201715F, 722017l6F that are refurbished and one additional lot number to catalog Number: 72201714.) The letter described the product, problem and action to be taken by customers. The customers were instructed to immediately check stock for affected product; immediately remove the items from stock and set aside for quarantine; complete and return the attached Urgent-Product Recall 1st Notification - Urgent" response form, even if no product is being returned, via fax at +1-508-261-3636 attention Cindy Burns or email a PDF to Cindy.Burns@smith-nephew.com; call for a return Authorization number at 1-800-343-5717 Option #3, and return product to Smith &amp; Nephew, Inc., Endoscopy Division as soon as possible. If you have any questions or require additional information, please contact Cindy Burns at 1-508-261-3655. |  | 577 | 577 units | Worldwide distribution: USA and countries including: Australia, Austria, Belgium, Brazil, Canada, Costa Rica, Denmark. Dubal, Finland, France. Germany, Greece, Hong Kong, India, Italy, Japan, Malaysia, Netherlands. Norway, Portugal, Puer To Rico, Singapore. South Africa, Spain. Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Venezuela. | Terminated | May 09, 2012 | 741 | 3607 | Not_Computer
Z-0036-2011 | 55802 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | HLD SYSTEM 520 is a two-tank Medical Device Cleaning and High Level Disinfection Systems (Washer/Pasteurizers). Each unit has a metal serial tag that is riveted to the enclosure of the unit and it is labeled in part: "HLD SYSTEMS 520***FOR SALES.SERVICE.SUPPLIES CALL***Cenorin*** Kent, WA 98032***P:800.426.1042***F: 253.395.2650***www.cenorin.com". |  |  | June 07, 2010 | 2010 | Cenorin | The water level sensor in the Medical Device Cleaning and High Level Disinfection (Washer/Pasteurizer) is sensitive to externally induced vibration noise, and it produces a false signal that the water level in the tank is full while it is only 75% full. It could present a health risk of biological hazard because the medical devices are not being adequately cleaned or pasteurized with enough wat | Component design/selection | On June 7, 2010, firm, CENORIN, sent "URGENT: MEDICAL DEVICE RECALL" Initial Notification letter with Recall Response Form, Series 500 Water Level Check Instructions, and the Water Level Switch Upgrade Kit Instruction (service kit P/N # 082395-00) to their consignees. The firm also sent these documents in Spanish language to their foreign consignees. The letter describes the recall product as HLD Systems Medical Device Cleaning and High Level Disinfection Series 500 Washer/Pasteurizer-Model 540, Model 540HT and Model 520. Customers are advised to call the Customer Service at 1-800-426-1402 or e-mail the firm.The Initial Notification letter instructs consignees to conduct and quarantine the recall products and notify their sub-accounts (customers) if products have been further distributed. While waiting for a replacement kit, the consignees are asked to follow the Series 500 Water Level Check Instructions to perform a test to ensure the water tank is completely filled up the expected water level. The firm also advised the consignees to add a cautionary label on the product to avoid any sharp or bump impact that may create any externally induced vibration noise to the device.The firm is developing a replacement kit that should be ready to send out to their consignees by June 30, 2010. A response card would be also sent out with a new kit, and it is used to verify that the kit has been successfully implemented. If you have any questions, please email jbennett@cenorin.com or call 253-395-2400 or toll free: 1-800-426-1042. |  | 9 | 9 units | Worldwide distribution:USA and countries including: Costa Rica, Ecuador, Japan, Mexico, and Taiwan. | Terminated | May 07, 2012 | 700 | 93 | Not_Computer
Z-0001-2011 | 56425 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter Colleague Single Channel Volumetric Infusion Pumps. Baxter Healthcare Corporation, Medication Delivery Division. Product Codes: 2M8151, 2M8161, and 2M9161. |  |  | August 04, 2010 | 2010 | Baxter Healthcare Corp. | The FDA sent a letter to Baxter on April 30, 2010, ordering the company to recall and destroy all models of its Colleague Volumetric Infusion Pumps currently in use in the United States. FDA determined that this action is necessary, as Baxter has failed to adequately correct, within a reasonable timeframe, the deficiencies in the Colleague pumps still in use. This recall is a consolidation of al | Device Design | Baxter Healthcare Corporation sent "Urgent Product Recall" letters dated August 4, 2010 to all Colleague Infusion Pump customers in the U.S., to the attention of the Hospital Administrator, Chief Executive Officer, Chief Operating Officer, Chief Medical Officer and Chief Nursing Officer, referencing Baxter's May 3, 2010 announcement to recall all Colleague pumps from the U.S. market pursuant to an order under its existing consent decree with the FDA. The letters provided the customers with important information regarding the actions they must take prior to November 14, 2010 in order to continue using their Colleague pumps during the transition period. Customers were instructed to complete the attached Certificate of Medical Necessity Form and return it to Baxter as soon as possible, but no later than November 14, 2010, by scanning and e-mailing to Baxter at FCA@Baxter.com. Any questions were directed to their local Baxter Sales Representative or by calling the Center for One Baxter 1-800-422-9837, between 8AM and 5 PM Central Standard Time (CST).*** On July 20, 2011 Baxter sent recall letters to those Colleague customers who had not completed a decision form regarding the status and replacement of their Colleague Pumps. The accounts were asked to inform their Baxter sales representative as soon as possible of their pump decision, and to complete the attached customer reply form and return it to Baxter via fax or scanned e-mail. *** |  | 221097 | 221,097 pumps | Nationwide, including Puerto Rico | Terminated | November 26, 2012 | 845 | 302872 | Not_Computer
Z-0006-2011 | 56257 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CareFusion Alaris PC Unit Model 8015 (formerly Medley |  |  | July 13, 2010 | 2010 | CareFusion Corporation | Under certain wireless network conditions a communication error can occur, which freezes the PC Unit screen, which may result in a delay of therapy. A delay of therapy may result in serious injury and/or death. CareFusion identified this potential risk associated with the Alaris PC unit model 8015 only (PC Unit) manufactured between December-2008 and September-2009. | Component design/selection | CareFusion sent the customer notification letter to all customers using the affected Alaris PC Unit beginning August 24, 2010. The customer notification letter is addressed to the Directors of Nursing, Risk Management, and Biomedical Engineering. Customers will be required to confirm receipt of the notification by returning the Recall Response Card to CareFusion by postage-paid, self-addressed mail, fax or email. The recall notice informed the customers of the reason for the notice, the units affected, the potential risk, a step-by-step description of the issue and the required actions for users.The customers were also told that CareFusion will contact their facility through phone or in person within 60 days to initiate implementation of the corrective action. The corrective action will require a hardware update to all affected units. CareFusion does not require that you return your devices. In the interim, if you observe this issue, remove the device from service and contact CareFusion Recall Center immediately.Customers were also told that any adverse reactions experienced with the use of this product, and/or quality problems should also be reported to the FDA's MedWatch Program. The recall notice provided contact information for CareFusion's Recall Center, Customer Advocacy and Technical Support units. |  | 17081 | 17,081 units (17,050 in US) | Nationwide and worldwide to: Canada, Netherlands, New Zealand and Singapore | Terminated | December 12, 2011 | 517 | 17081 | I/O | Display/Image |  | Other
Z-0050-2011 | 56680 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Scorpio/ NRG 4in1 Cutting Block |  |  | August 23, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Stryker s has received reports in which the Cutting Blocks have fractured on impaction during surgery. | Device Design | Stryker Orthopaedics sent an "URGENT PRODUCT RECALL" letter dated August 30, 2010, to all consignees/customers (branches/agencies, hospital risk management, hospital Chief of Orthopaedics, surgery) including foreign countries. The letter described the product, problem, and action to be taken by customers. The customers were instructed to complete and return the attached PRODUCT RECALL ACKNOWLEDGMENT FORM via fax to 201-831-6069.If you have any questions, feel free to contact the Manager, Divisional Regulatory Compliance (201) 972-2100. |  | 839 | 839 (total) | Worldwide distribution: USA and countries including: Asia, Canada, Europe, Middle East and Africa | Terminated | August 30, 2012 | 738 | 839 | Not_Computer
Z-0058-2011 | 56694 | No | DSQ |  |  | Cardiovascular | PMA | HeartMate II System Controller, a component of the HeartMate II Left Ventricular Device System (LVAS), distributed both as stand alone devices and as components of the LVAS Implant kit. Manufactured/Distributed by Thoratec Corporation, Pleasanton, CA |  |  | September 01, 2010 | 2010 | Thoratec Corp | Connection failure-- product may malfunction because of bent pins or broken wires in power cord connection | Device Design | Thoratec Corporation sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated August 30, 2010, to all customers. The letter described the affected product, description and symptoms of problem, and immediate action to be taken by the customers. The customers were instructed to:1) contact all ongoing patients using the affected Systems Controllers (serial numbers starting with "EPC") to determine if they are experiencing any unexplained intermittent alarms or have any bent pins in the connectors between the System Controller and the power sources, if so, have them return to the hospital so that their power leads can be examined. If further assistance is needed with the evaluation, please contact Thoratec Technical Service (800-456-1477).2) review the attached addendum to the HeartMate II labeling with all VAD clinicians and ongoing HeartMate II LVAS patients, and3) complete and sign the attached Acknowledgment Form via fax to Thoratec Regulatory Affairs at (952) 847-8571.If you have any questions, please contact 925-847-8600. |  | 8965 | 8965 units | Worldwide distribution: USA including states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX , UT, VA, WA, and WI; and countries including: Austria, Belgium, Canada, Czech Republic, Denmark, France, Germany, Hong Kong, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Lithuania, Netherlands, Singapore, Slovenia, Sweden, Switzerland, and the UK. | Terminated | April 07, 2011 | 218 | 8965 | Other
Z-0060-2011 | 56651 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Trabecular metal Modular Acetabular system 56 mm Porous shell with cluster holes, Catalog 00-6202-056-22, Zimmer Inc., Warsaw, IN. |  |  | August 18, 2010 | 2010 | Zimmer Inc. | The threads on the device were not completely machined and therefore would not mate with athe corresponding cup positioner/shell inserter. | Other | The firm, Zimmer, made telephone calls using a phone script document entitled "Urgent: Medical Device Recall" Phone Script Call Communication, to the customers that received this product starting on 8/18/2010. The telephone calls were to determine the status of the recalled product. A Zimmer Field Action Strategy notification dated August 9, 2010 was also provided. The notifications/phone script described the product, problem and actions to be taken by the customers. The customers were instructed to remove the 00-6202-056-22, lot 61091653 product from their inventory and return to Zimmer Product Service Department, Attn: Product Service, 1777 West Center Street, Warsaw, IN 45680. Note: If the product had been used, no action was necessary. The returned product will be evaluated to determine how many of the devices had the defect.If you have any questions, please call (574) 371-8852. |  | 24 | 24 | Worldwide distribution:USA including states of: MI, IN, WI, MO, CA, VA, MD, TX, NC, PA, OR, WA, FL, GA, NY, ME, AZ, and UT and country including Canada. | Terminated | February 07, 2011 | 173 | 24 | Not_Computer
Z-0059-2011 | 56589 | No | HXO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Closed knot pusher REF 902813, Biomet Sports Medicine, Warsaw, IN. |  |  | August 02, 2010 | 2010 | Biomet, Inc. | The firm has received four complaints that the closed knot pusher is missing a radius which contributes to sutures being cut during surgery. | Device Design | Biomet Sports Medicine sent URGENT MEDICAL DEVICE RECALL NOTICES dated 8/2/2010 to its customers and distributors. The notices stated that the use of the device should be discontinued and the products returned to Biomet.Healthcare professionals and consumers were to report serious adverse events (side effects) or product quality problems with the use of this product to the FDA's MedWatch Adverse Event Reporting program either online at www.fda.gov/MedWatch/report.htm, by regular mail to MedWatch, 5600 Fishers Lane, Rockville, MD 20852-9787 or by phone at (800)FDA-1088 or Fax (800)FDA-0178.Questions related to this notice should be directed to 574-371-3755 or 574-371-3756, Monday through Friday, 8AM to 5PM. |  | 101 | 101 | Worldwide Distribution - USA, including FL, NY, VA, CA, GA, AL, PA, MI, LA, KS, SD, KY, TX, AR, and WV, and the countries of Belgium, Italy, Argentina, Poland, Netherlands, Turkey and Portugal. | Terminated | August 20, 2012 | 749 | 101 | Not_Computer
Z-0071-2011 | 55157 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Continuum Acetabular Systems Shell Inserter Adapter with Rotational Control, 00-8790-003-00, Zimmer Inc., Warsaw, IN. |  |  | March 19, 2010 | 2010 | Zimmer Inc. | The firm has received reports of users encountering difficulties in usage leading to surgical delays or minor damage to a threaded surface of the shell implant. | Device Design | The firm sent Urgent: Device Removal letters dated 3/15/2010. The letter identified the affected product, described the issue, discussed the clinical implications, and provided a list of required actions. The consignees were instructed to perform an inventory of the recalled product and report the amount on hand to Zimmer by completing the Inventory Return Certification Form. After Zimmer received the report, they would send out new product to the consignee. After new product was received the consignees were instructed to return the recalled product to Zimmer. The firm expects to have all product replaced by 4/30/2010. If there are any questions or additional info needed, consignees should contact Zimmer, Inc. at 1-800-613-6131 or 574-372-4463. |  | 1125 | 1125 all products | Worldwide Distribution -- US, Singapore, Australia, Portugal, Germany, Switzerland, Spain, Finland, United Kingdom, Israel, Italy and Sweden. | Terminated | January 12, 2011 | 299 | 1125 | Not_Computer
Z-0068-2011 | 56719 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Architect Tacrolimus Reagent Kit; an in-vitro diagnostic chemiluminescent microparticle immunoassay (CMIA) for the quantitative determination of tacrolimus in human whole blood; Abbott Laboratories, Abbott Diagnostics Division, Abbott Park, IL 60064 USA; list 1L77-25 |  |  | September 07, 2010 | 2010 | Abbott Laboratories | There are reagent barcode read errors when using reagent lot 86599M500 due to the print quality of the barcode. | Other | Abbott Laboratories sent a "Product Recall Immediate Action Required" letter dated September 7, 2010, to all customers. The letter described product, problem and action to be taken by customers. The customers were instructed to determine if they are currently using and/or have inventory of lot 86599M500; discontinue use of and destroy any remaining inventory; inform and provide copies of letter to other laboratories if appropriate; retain the letter for laboratory records and complete and return the Customer Reply Immediate Action Required form via fax at 1-800-777-0051 or email QAGCO@abbott.com. IMPORTANT NOTE: If you do not currently have replacement material than you may continue to use this reagent lot until your laboratory receives replacement material. Upon receipt of the replacement material, discontinue use of and discard lot 86599M500 and switch to the replacement material. Contact Abbott Customer service for replacement material (1-877-4ABBOTT, Option 1).If you or any of the health care providers you serve have any questions regarding this information, U.S. customers should call Customers Service at 1-877-4ABBOTT. Customers outside of the U.S., please contact your local Customer Service representative. |  | 1995 | 1,995 kits | Worldwide distribution: USA and countries including Puerto Rico, Argentina, Australia, Austria, Belarus, Belgium, Bosnia, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Czech Republic, Egypt, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Lithuania, Montenegro, Netherlands, New Zealand, Norway, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom and Vietnam. | Terminated | May 13, 2011 | 248 | 1995 | Not_Computer
Z-0070-2011 | 56699 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System Buretrol Solution Set, 83" (2.1 m), 150 mL Burette Drip Chamber Filter Valve, 2 Injection Sites, Male Luer Lock Adapter, Large Bore 4-Way Stopcock with Rotating Male Luer Lock Extension Set, 38" (97 cm), Vol. 5.8 mL, Injection Site, Male Luer Lock Adapter, 60 DPM; a sterile Rx IV fluid pathway device; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product code 3C0166 |  |  | September 03, 2010 | 2010 | Baxter Healthcare Corp. | There are leaks in the Drip Chamber Check Valve Subassembly of the Solution Set. A leak in a sterile pathway could allow microbial entry, which could lead to a blood stream infection. | Employee error | Baxter sent an "Urgent Product Recall" letter dated September 2, 2010, via first class mail to the direct account, to the attention of the Director of Materials Management, with a copy sent to the Director of Nursing, on the same date, informing them that leaks were detected in the Drip Chamber Check Valve Subassembly used in the affected lot of Interlink System Buretrol Solution Sets. A leak in a sterile pathway could allow microbial entry, which could lead to a blood stream infection. The account was requested to examine their inventory in all locations to determine if they have any of the affected lot and remove the affected product from inventory; call Baxter Healthcare Center for Service at 1-888-229-0001 to return the affected product, request a credit, or request replacement product; complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 948 | 948 units | Nationwide distribution: Wisconsin only. | Terminated | February 25, 2011 | 175 | 948 | Not_Computer
Z-0069-2011 | 56278 | No | LNX |  |  | General Hospital | 510(k) | Hospital Information System Results Reporting Interface Versions 4.0.2.38.1, 4.0.2.39.2, 4.02.40.1, and 4.0.2.41.1. |  |  | October 05, 2007 | 2007 | SCC Soft Computer | SCC Soft Computer, Clearwater, FL issued a Correction for Hospital Information System Results Reporting Interface Versions 4.0.2.38.1, 4.0.2.39.2, 4.0.2.40.1, and 4.0.2.41.1 on 10/05/2007. A software anomaly in the result reporting interface caused result comments entered after an empty line | Software design | SCC Soft Computer sent a RISK-TO-HEALTH NOTIFICATION TASK VERBIAGE dated October 5, 2007, through their Task Management System to all affected clients on October 17, 2007. The notice identified the product, the problem and the action clients should take to correct the problem. Clients were to acknowledge receipt of the task notice and to identify if the Workflow or functionality is not used and if they Agree to use alternative working solutions. Each client was asked to grant permission to load the required software to correct the issue.For any questions regarding this recall call (727) 789-0100. |  | 35 | 35 | Worldwide Distribution - USA, including AL, CO, FL, GA, IN, KS, KY, MA, MO, NC, NH, NJ, NY, OH, PA, TN, and WA and the country of Canada | Terminated | October 19, 2010 | 1110 | 35 | Software | Output/Calculation | Software update | Software Update
Z-0066-2011 | 56756 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Control Syringe, 12ml, CAT No: CCS201E Version B, Sterile R, Merit Medical Systems, Inc. |  |  | September 09, 2010 | 2010 | Merit Medical Systems, Inc. | Small holes in the packaging may render the product non-sterile. | Employee error | Merit Medical sent an Urgent Product Recall Notice dated September 9, 2010, to all sales representatives. The notice identified the product, the problem, and the action to be taken by the sales representatives and their customers. Sales representatives were instructed to contact all customers via telephone and instruct them to immediately discontinue use and quarantine any unused product. Recall Notification Forms were to be completed and signed by the site representative and the customer and faxed to Merit customer service at 801-208-3378. If customers were returning affected products they were to place the completed original Recall Notification Form in the box with the products to be returned, write the RGA# on the shipping box and return by Federal Express 2nd Day air to: RGA DepartmentMerit Medical Systems, Inc.1600 West Merit ParkwaySouth Jordan, UT 84095-2416For any questions regarding this recall call 801-316-4822 or 801-208-4365. |  | 199 | 199 units | Worldwide Distribution - USA, including IL, PA, WI and WV and the countries China, Greece, Hong Kong, Japan, New Zealand and Spain. | Terminated | March 21, 2011 | 193 | 16574 | Not_Computer
Z-0067-2011 | 56275 | No | LNX |  |  | General Hospital | 510(k) | SoftPath GUI Releases 3.17 through 4.1.0. |  |  | March 28, 2007 | 2007 | SCC Soft Computer | SCC Soft Computer, Clearwater, FL issued a Correction for SoftPath GUI Releases 3.17 and 4.1 in March 2007 for clients with high volume output of final reports who use the Batch Print for Final Reports in the Scheduler, there is a potential for some reports not printing. This issue was corrected for the one affected user in October 2003. | Software design | SCC Soft Computer sent a SAFETY NOTIFICATION TASK VERBIAGE dated January 5, 2004, to one client. The Notice identified the product, the problem, and the action to be taken by the client. SCC recommended the client set the Search Date Range to a small enough range to assure that the number of cases scanned during the job would be less than 10,000.The client was asked to acknowledge receipt of the task and to indicate if the Workflow or functionality was used/not used and Agree to use the alternative working solution. For any questions regarding this recall call (727) 780-0100, ext. 4235. |  | 84 | 84 | Nationwide - USA, including the states of CA, CO, CT, FL, GA, IL, IN, LA, MD, MI, MO, MS, NC, NH, NJ, NV, NY, OH, OK, PA, SD, TX, WI, and WV | Terminated | October 19, 2010 | 1301 | 84 | Software | Output/Calculation | Software update | Software Update
Z-0075-2011 | 54550 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Pathfinder End Screw Extender Sleeve, REF 1165-10, Rx, non-sterile, sold under label Abbott Spine, Inc. Austin TX (prior to 2008) and Zimmer Spine Austin TX (7/2009 to present) |  |  | April 13, 2010 | 2010 | Zimmer Inc. | The tip of the PathFinder End Screw Extender Sleeve has fractured during use. In some cases, this instrument fracture led to surgical delay or additional surgical steps necessary to remove the fragments of the broken tip. | Device Design | Zimmer Spine issued an "Urgent Medical DeviceCorrection" letter through the Zimmer Spine sales force to be provided to PathFinder system users in February 2010. the letter identified the affected product, safety issue, and actions to be taken by customers. Customers will receive replacement devices with a design change for the recalled devices, as they are available from the manufacturer. |  | 1471 | 1471 (888 subject to removal) | Worldwide Distribution: USA including states of AZ, CA, CO, FL, GA, LA, MI, MN, MO, NC, NH, NY, OH, OK, PA, TN, TX, WA, and WI, and the countries of Australia, Belgium, Brazil, Canada, China, Colombia, Cyprus, France, Germany, Greece, Italy, Japan, Mexico, Thailand, Turkey, UK and Switzerland. | Terminated | February 10, 2012 | 668 | 1471 | Not_Computer
Z-0078-2011 | 55150 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyRA Cuvette Segment 1, P/N 007000-001 |  |  | March 29, 2010 | 2010 | Medica Corporation | Chemistry analyzers may report Calcium results below the sample's actual concentration. Reported results may be ten to fifteen percent lower than expected sample results.The defect is a shortened path length at specific cuvette positions of those segments manufactured using molding cavity 1. The shortened path length was determined by measuring the cuvette dimensions with a Co-ordinate Measuri | Other | Medica issued an Urgent Product Recall letter dated March 29, 2010 to customers. Users were requested to cease using and discard the affected product, and complete and return a Response Form in order to obtain replacement product.Customers can contact Medica Technical Support at 877-777-5895. |  | 14926 | 14,926 | Worldwide Distribution: USA, and the countries of Bulgaria, Canada, India, Mexico, Panama, Phillipines, Russia, Turkey, New Zealand, South Africa, United Arab Emirates, and Venezuela. | Terminated | June 07, 2011 | 435 | 14926 | Not_Computer
Z-0079-2011 | 56805 | No | KWZ | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | R3 (TM) 22 MM I.D., 52 MM O.D., CONSTRAINED LINER, UHMWPE, CO-CR, TI-6AL-4V, REF 71339152, STERILE EO, QTY 1, Smith &amp; Nephew, Inc., Hip prosthesis acetabular component. |  |  | August 13, 2010 | 2010 | Smith &amp; Nephew Inc | The firm received reports of a femoral head popping out of the liner intraoperatively. | Other | Smith &amp; Nephew notified all distributors via e-mail and sent a letter dated August 13, 2010 to all hospitals notifying them of the product, the problem, and the action to be taken. 1) Distributors were to immediately cease distribution and use of the product, examine their stock and the stock of the accounts they supplied if applicable and remove any affected devices from all inventories and place them into quarantine to prevent accidental distribution. 2) Identify customers that received these products and immediately notify them of the recall.3) Carry out a physical count and record the data on the Verification Section included in the letter.4) Contact their inventory specialist for an RA# and record this on the Verification Section.5) Return the Verfication Section even if they do not have the recalled product.6) Return the recalled product to: Smith &amp; Nephew, Inc. Attn: Return Goods 3303 Holmes Road Memphis, TN 38118Smith &amp; Nephew, Inc. sent a second notice via e-mail on 09/03/2010 to non-responding distributors and by Fed Ex to the hospitals 09/21/2010.For any questions regarding this recall customers were to call (901) 399-5441. |  | 118 | 118 units | Worlwide Distribution, USA including AL, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, MA, MI, MN, MO, NC, NJ, NY, OH, OK, PA, TN, TX, WA, and WI and the countries of Australia and Portugal | Terminated | September 21, 2011 | 404 | 716 | Not_Computer
Z-0074-2011 | 56274 | No | LNX |  |  | General Hospital | 510(k) | Soft Query Results Reporting Versions 2.2.0.6.3 and 2.2.0.7.1. |  |  | November 30, 2004 | 2004 | SCC Soft Computer | Soft Query Results Reporting Versions 2.2.0.6.3 and 2.2.0.7.1 HIS may display incorrect collection data for a procedure or test result. Clients using HBO Star RR interface to query for results on a canceled test may see the previous occurrence of the test with the current date and time displayed as the collection date and time instead of the correct collection date and time sent by the SCC interf | Software design | Client Notification Task Verbiage was sent through the Task Management System on December 3, 2003, identifying the software issue, the clinical impact, and corrective action. Customers were informed that the HIS may display incorrect collection data. Customers received the SQRR Service Pack with revised software, which will will not produce inaccurate information if there are no verified results. Customers were to acknowledge receipt of the task notification and indicate whether they would like the corrected software.SCC Soft Computer proprietary Task Management System is used to manage client communication. Clients use this system to request services, receive and provide software updates for any issue.Customers can contact SCC Soft Computer at 727-789-0100. |  | 5 | 5 | Nationwide Distribution: including states of GA, LA, and NH. | Terminated | October 19, 2010 | 2149 | 5 | Software | Display/Image | Software update | Software Update
Z-0073-2011 | 56789 | No | EZQ | Ostomy pouch and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | ActiveLife One-Piece Drainable Pouch with Stomahesive Skin Barrier; |  |  | September 17, 2010 | 2010 | Convatec Inc. | Mislabeled: the market units were incorrectly labeled as # 125339 ActiveLife Drainable Pouch 25 mm; the product inside was actually ActiveLife Pouch Small Tan 32 mm, ICC # 125333. | Labeling mix-ups | ConvaTec, Inc. sent a Returned Goods Authorization form by e-mail dated September 3, 2010, to one customer. The form identified the product, the problem, and the action the customer should take. The customer was instructed to detach the label and affix it to the outside of the box, copy the RGA # and retain a copy for their records, enclose the completed form and return with the product, For questions call (908) 904-2675. |  | 1 | 1 case (20 market units) | USA, including IL | Terminated | October 19, 2010 | 32 | 1 | Not_Computer
Z-0076-2011 | 56736 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Sequencer Verify &amp; Record System, for use with radiotherapy treatment machines, Class IIb, manufactured by Impac Medical Systems, Sunnyvale, CA. |  |  | September 02, 2010 | 2010 | Impac Medical Systems Inc | Software issue-- matching criteria can be configured that permits an invalid patient match, resulting in overwriting of patient records with information from a different patient. | Software design | Elekta Impac Software issued a Important Safety Notice letter dated September 2010 to all customers identifying the software issue and clinical impact. Customer Support representatives will contact each site and modify site configuration or upgrade existing software to the version containing the corrective fix. Customers can contact Impac Software at 800 488-4672 |  | 797 | 797 units | Worldwide Distribution: throughout the US and the countries of Australia, Austria, Belgium, Canada, China, Hong Kong, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Ireland, Israel, Italy, Japan, Jordan, Morocco, Netherlands, New Zealand, Poland, Portugal, Singapore, Spain, Sweden, Switzerland, and UK. | Terminated | June 24, 2011 | 295 | 797 | Software | Output/Calculation | Software update | Software Update
Z-0077-2011 | 56729 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | DDC-6-AA Data Card, an accessory to the Sentry DDU-1 00 semiautomatic external defibrillator (AED). |  |  | August 27, 2010 | 2010 | Defibtech, LLC | Data cards may include an incorrect format which may cause the AED to not function correctly. | Mixed-up of materials/components | Defibtech Initially notified customers by telephone on 8/27/10. Customers were notified of the product, the problem and the action to be taken. Customers were instructed to locate affected data cards and send them back to Defibtech for evaluation. If the data card was inserted into the AED, then the customer was instructed to check the AED software version. If the software version was at 2.005, the customer was instructed to keep the AED out of service until Defibtech could evaluate the returned data card. (Any other software version indicated that the device was not affected by a faulty data card). If the software version was any other version other than 2.005, then the customer was instructed to keep the AED in service unless the original service inquiry required it to remain out of service.If the data card was not inserted into the AED, then the customer was instructed to return the data card for evaluation.For any questions customers were to contact Customer Service at (866) 333-4241 from 8:30-5pm EST, M-F. |  | 38 | 38 | USA - Nationwide Distribution and Puerto Rico | Terminated | October 05, 2011 | 404 | 38 | Other |  |  | Other
Z-0088-2011 | 56851 | No | LNX |  |  | General Hospital | 510(k) | SoftMic Versions 4.0.4.7, 4.5.1.3. |  |  | July 19, 2010 | 2010 | SCC Soft Computer | Softmic system may delay or omit reporting of clinically significant results including organisms and drug sensitivities. Incomplete culture results may appear complete. SCC Soft Computer is issuing a Correction for SoftMic Versions 4.0.4.7-4.5.1.3 software due to this potential hazard. | Software design | SCC Soft Computer issued Risk-to-Health notifications to clients on July 19, 2010 via proprietary communication software. Corrective actions will be via programming correction via hot fix (software modification) or such so that Global variable is not used. Clients are asked to acknowledge receipt of this task and update the task with permission to load the required correction. |  | 35 | 35 | Worldwide Distribution: USA, including the states of CA, CT, DE, FL, IA, KY, MA, MA, MI, MN, NJ, NY, OH, PA, TX, VA, and WI, and the country of Canada. | Terminated | July 16, 2012 | 728 | 35 | Software | Output/Calculation | Software update | Software Update
Z-0090-2011 | 56280 | No | LNX |  |  | General Hospital | 510(k) | SoftPath GUI version 3.17.6.3.1. |  |  | March 11, 2004 | 2004 | SCC Soft Computer | For clients using the SoftPath GUI version 3.17.6.3.1 Results Reporting interface with parsed text and having the auto correct and auto formatting functions active in MS Word, there is a possibility for some ANSI characters/symbols to be translated incorrectly to ASCII. SCC Soft Computer initiated a Correction in March 2004 for all affected clients. | Software design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System on March 11, 2007. SCC recommended turning off the autocorrect and autoformatting functions in MS Word if custom configuration parses text to an ASCII file (HIS) for the Result Reporting interface. If this function was turned on at the time of creating the canned message in the setup file, edit and re-save all canned messages in SoftPath. Client was asked to acknowledge receipt of the task and grant permission to load the required software correction. |  | 1 | 1 | Nationwide distribution: One hospital in NY. | Terminated | November 03, 2010 | 2428 | 1 | Software | Output/Calculation | Software update | Software Update
Z-0087-2011 | 56502 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Manta Ray: Anterior Cervical Plate (ACP) System, 510 (k) #K080690. The labeling for the plate component reads in part: 'REF: 22-'**** MAT'L: Titanium per ASTM F-136, Lot: W****, Description: Plate - 40 mm, Qty: 1.' The labeling for the screw component reads in part: 'REF: 22-**** MAT'L: Titanium per ASTM F-136 LOT W**** Description: Screw-Fixed Tapping 4x10mm QTY: 1. |  |  | July 29, 2010 | 2010 | Theken Spine Llc | The Manta Ray Cervical System was subject to corrective action by the firm after two (2) customer complaints were received alleging that screws were 'backing out' of the device during surgical procedures. | Other | Integra sent a Technical Bulletin dated July 29, 2010 via Fedex to their customers. The bulletin identified the product, the problem, and action to be taken by the user/customers (surgeons) of the two (2) reported instances of screws backing out of the plates. The Bulletin provided instructions outlining the proper surgical techniques for implantation of screws using the Manta Ray ACP System. The bulletin also included an attached acknowledgement form which the user/customers were asked to complete and FAX to (609) 275-5363 or E-mail back to the E-mail address provided. The form verified that each customer: 1) received the technical bulletin; and 2) has read and understood the proper technique for surgical placement of the screws as outlined in the bulletin.For technical support customers were to contact the Customer Service Department at (866) 942-8698. |  | 1081 | 1081- includes totals distributed for both plates &amp; screws | Nationwide Distribution - USA including AL, AZ, CA, CT, FL, GA, HI, IL, ID, IN, KS, LA, MI, MD, MN, MA, MD, MS, NE, NY, NC, NJ, NV, NM, OH, OR, OK, PA, TN, TX, UT, WI, and WA.WA. | Terminated | September 21, 2012 | 785 | 1081 | Not_Computer
Z-0091-2011 | 56808 | No | FHN | Hemorrhoidal ligator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Ten Shooter Saeed Multi-Band Ligator, Rx Only, Wilson Cook Medical 4900 Bethania Station Road, Winston-Salem, NC 27105 |  |  | September 03, 2010 | 2010 | Cook Endoscopy | A section of the ligator barrel may separate from the barrels friction fit adapter that attaches to the patent end of the endoscope. This could result in a section of the ligator barrel detaching inside the patients gastrointestinal tract. | Other | COOK Endoscopy sent an URGENT: PRODUCT RECALL dated September 3, 2010, via USPS Certified Mail and by email to all consignees. The letter identified the products, the problem, and the action to be taken by the consignees. All consignees were instructed to quarantine any affected product and return them to Cook Medical. Distributors were instructed to contact their customers and coordinate the return of the affected product to Cook Medical. Within 2 weeks of the initial recall notification, Level B effectiveness checks (40% of the total number of consignees) will be conducted. These customers will receive a recall effectiveness survey. All non-responding consignees will receive a total of 3 recall notifications (2 follow-up requests for a response in addition to the initial recall notification. Monthly status reports regarding this recall will be provided to FDA. All devices returned as part of this recall will be destroyed by the recalling firm upon receipt of authorization to do so by FDA. Forty percent (40%) of total customers affected received a second notice on 9/17/2010 as part of the effectiveness checks. As of 9/21/2010, 38% of all affected customers have responded.For questions contact the Customer Relations Department at CustomerRelationsNA@cookmedical.com or call 1-800-457-4500, press 4 then enter extension 2146. |  | 2011 | 2011 units (total including both 6 &amp; 10 Shooters) | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, WA, WI, WV, and WY; and the countries of Algeria, Australia, Austria, Belgium, Canada, Chile, Cyprus, Denmark, France, Georgia, Germany, Greece, Hong Kong, Ireland, Israel, Italy, Jordan, Korea, Mexico, Malaysia, Morocco, Netherlands, New Zealand, Peru, Poland, Portugal, South Africa, Slovakia, Spain, Switzerland, United Arab Emirates, UK and Venezuela. | Terminated | March 31, 2011 | 209 | 2012 | Not_Computer
Z-0089-2011 | 56277 | No | LNX |  |  | General Hospital | 510(k) | SoftPath GUI Versions 3.17 and 4.1. |  |  | December 30, 2003 | 2003 | SCC Soft Computer | SoftPath Cases with Word Perfect documents that were not signed out in ASCII are printing in GUI as a converted document from Word Perfect to PDF document instead of the current MS Word Document. This affects clients upgrading from ASCII to SoftPath GUI 3.17.7.6 Version. All affected clients have been notified and corrections have been made. | Software design | SCC Soft Computer issued Risk-to-Health notifications with Client Notification Verbiage to affected clients using SCC Soft Computer's proprietary Task Management System. beginning on December 30, 2003. Clients were instructed to sign out all Word Perfect documents before upgrading to GUI. The correction for this anomaly is to upgrade to the latest release of version 3.17.7 or 3.17.8. Clients were asked to acknowledge receipt of this Task and indicate whether they would like to take the correction at this time. When the client decides to accept the correction a new task will be created to initiate the patch using the Task Management System. |  | 5 | 5 | Nationwide Distribution to the states of IN, NC, NH, and NJ. | Terminated | November 03, 2010 | 2500 | 5 | Software |  | Software update | Software Update
Z-0107-2011 | 55195 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active I-PTH Enzyme-Linked Immunosorbent Assay(ELISA), Part Number: DSL-10-8000 |  |  | March 04, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed that the stated stability of the DSL-10-8000 I-PTH ELISA assay maynot be achieved for lot numbers: 891198, 990980 and 991448.The kit lot numbers listed above passed the stability testing performed at 6 months, but failed the testing performed at 12 months. Therefore the affected I-PTH ELISA kit lots may produce falsely elevated patient r | Other | Beckman Coulter sent an URGENT: PRODUCT CORRECTIVE ACTION letter dated March 17, 2010 with attached Customer Response form to all customers who received Active I-PTH ELISA kits with the affected lot numbers. The letter identified the product, the problem, and the actions to be taken.Customers were instructed to:(1) Discontinue use of all materials provided with I-PTH ELISA kit lot numbers 891198, 990980 and 991448;(2) Review historical external quality control results to ensure the assay performance of these lots was within laboratory specifications. Retesting may be necessary if your laboratory uses only the control materials provided with the I-PTH ELISA kit;(3) Abnormally elevated patient results that are obtained with these lot numbers and that are inconsistent with the patient's clinical status should be reviewed in the context of other clinicaldiagnostic tests; (4) I-PTH ELISA kit lot numbers 891198, 990980 and 991448 will no longer be shipped.Alternate lots will be provided for current and future replenishment orders or requests.in addition customers were instructed to share this information with your laboratory staff and retain this notification as part of yourlaboratory Quality System documentation. Complete and return the enclosed response form within 10 days.For questions regarding this recall contact (800) 231-7970. |  | 70 | 70 kits | Nationwide Distribution including the states of CA, CT, IA, IL, MI, NY, SC, and TX | Terminated | April 27, 2012 | 785 | 70 | Not_Computer
Z-0099-2011 | 56282 | No | LNX |  |  | General Hospital | 510(k) | SoftWeb version 4.0.4.10.1. |  |  | June 08, 2006 | 2006 | SCC Soft Computer | If a case number is re-used in SoftWeb with SoftPath, the SoftPath case is linked to the original MRN in SoftWeb. This causes SoftPath cases being linked to the wrong patient in SoftWeb. Incorrect results on the patient with original MRN and no results on the patient with the second MRN (and the reused case number) are displayed in SoftWeb Results Query. SCC Soft Computer issued a Correction for | Software design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System on June 8, 2006. SCC recommended SoftPath/SoftWeb users NOT re-use SoftPath case numbers when cases are cancelled. The generic option setting "Specimen Registration - Standard Settings - Minimum Time to Reactivate Case Number" should be set to "0" so the user is not prompted to reuse the same case number. Client was asked to acknowledge receipt of the task and grant permission to load the required software correction. |  | 1 | 1 | Nationwide Distribution: 1 hospital in Alaska. | Terminated | November 03, 2010 | 1609 | 1 | Software | Output/Calculation | Software update | Software Update
Z-0098-2011 | 56687 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Stryker PainPump2 blockAid 400 mL Painpump for continuous Nerve Block, REF 575-100, Sterile, Rx only, Stryker Instruments, Kalamazoo, MI. |  |  | August 16, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | The instructions for use (IFU) has been revised for health care professionals. The indication for use has been changed to remove site specific pain management. The firm has revised the warning to not use the pump around joint spaces. A warning has been added regarding the use of the pump with anticoagulants in epidural applications. A special reminder to users to always evaluate the potential | Other | Stryker Instruments sent Urgent Medical Device Recall Notification letters dated 8/16/2010 to healthcare professionals. The recall letters outlined the changes for the instructions for the use of these products. The letters informed the healthcare professional to review the new IFU, to distribute the new IFU to all healthcare professionals that use the device, and to complete and return the business reply form. The letters included a complete IFU that was to be used with any remaining stock of devices. Customers should dispose of the old version of the IFU. There is no product being returned for this recall. |  | 11976 | 11,976 all products | US, Canada and Australia. | Terminated | April 21, 2011 | 248 | 11976 | Not_Computer
Z-0103-2011 | 56675 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | LIBERTY Automated Peritoneal Dialysis System Line Cycler Set with stay safe connectors (Dual Patient Connector) PIN: 050-87212, manufactured for Fresenius Medical Care, Waltham MA by Erika de Reynoza, Reynosa Tampa, Mexico. |  |  | September 03, 2010 | 2010 | Fresenius Medical Care Holdings, Inc. | Cycle set cassettes may leak due to holes/scratches in the film, contaminating dialysate and potentially resulting in peritonitis. | Nonconforming Material/Component | Fresenius Medical Care issued an Urgent FMCNA LIBERTY Cycler Set Recall letter on September 3, 2010 identifying the specific affected part numbers and product lots and requesting that customers identify and discontinue the use of the affected lots immediately. Customers were requested to return a receipt of notification. Customers may arrange return unused product to FMCNA by contacting customer service at 1-800-323-5188. |  | 363265 | 363,265 units | Nationwide Distribution: Throughout USA. | Completed |  |  | 370588 | Not_Computer
Z-0108-2011 | 56802 | No | JEY | Bone plate. | Dental | Dental | 510(k) | 2.0/1.0 4 hole long plate (titanium plate ).Lot 120150. |  |  | August 31, 2010 | 2010 | Biomet Microfixation, Inc. | Mislabeled size: The 2.0/1.0 4 hole long plate, Lot 120150 is labeled as 01-9205 (2.0/1.0 4 Hole Long Plate), however, the package contains a 01-9204 (2.0/1.0 4 Hole Regular Plate). | Other | Biomet Microfixation sent URGENT MEDICAL DEVICE RECALL NOTICES dated August 31, 2010 to its customers and distributors, identifying the affected product and actions to be taken by customers. Customers were requested to check their inventory for product and return any affected product to Biomet Microfixation. Customer questions or concerns should be directed to Laura Sabo, VP Global QA/RA at 904-741-4500. |  | 56 | 56 | Worldwide Distribution: USA including states of Hawaii and Kansas, and the countries of Australia, Denmark, Italy, Jordan, and South Africa. | Terminated | November 19, 2010 | 80 | 56 | Not_Computer
Z-0093-2011 | 56279 | No | HBA | Neurosurgical paddie. | Neurology | Neurology | 510(k) | Codman 5/16" Round Surgical Patties |  |  | July 16, 2010 | 2010 | Codman &amp; Shurtleff, Inc. | Surgical Pattie x-ray stripe (radio-opaque marker) became detached | Process control | Codman notified customers by letter dated 7/16/10 issued July 23 through July 26, 2010 via Fed'x overnight advising users that the x-ray marker on Codman Surgical Patties may become detached. Customers are instructed to return product from affected lots. An error in the original letter was discovered and a corrected letter dated July 28, 2010 issued. If there are any questions, please contact Codman at (800) 225-0460. |  | 6730 | 673 boxes (20 pouches of 10 patties per box) | NationwideForeign: Canada, Singapore, Belgium, Mexico, Madrid, Dubai, Tokyo, Brazil, New Zealand, Egypt, South Africa, Singapore, | Terminated | July 09, 2012 | 724 | 53090 | Not_Computer
Z-0102-2011 | 56272 | No | LNX |  |  | General Hospital | 510(k) | SoftPath GUI version 4.2.3.4 2008-06-30 (Manuf Date). |  |  | January 26, 2009 | 2009 | SCC Soft Computer | For SoftPath GUI versions 4.2.3.4, 4.2.4.1, 403.0.3, and 4.3.0.4., the Previous Patient History and some QA Management reports may have missing or incorrect diagnosis text. However, the patient final report, HIS resulting and Query display are all correct. All affected clients have been provided corrected software.. | Software design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to affected clients using SCC Soft Computer's proprietary Task Management System on January 26, 2009. SCC issued a software upgrade with corrective patch electronically to all clients. Client was asked to acknowledge receipt of the task and grant permission to load the required software.SCC will contact each client to coordinate the running of the utility and the loading and validation of the correct patch. |  | 52 | 52 | Worldwide Distribution: Throughout USA and Canada. | Terminated | November 03, 2010 | 646 | 52 | Software | Output/Calculation | Software update | Software Update
Z-0106-2011 | 56346 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand TrueBeam and TrueBeam STx, |  |  | May 24, 2010 | 2010 | Varian Medical Systems Oncology Systems | Due to a software anomaly, guidance-based couch shift values may not be applied as expected when the operator selects "Apply Shift" and presses Motion Enable buttons. This can result in the actual patient position differing from that position indicated by the imaging application affecting poistion accuracy and precision. | Software design | Varian Medical Systems issued an Urgent Medical Device Correction letter date dated May 24, 2010 to users identifying the affected device, the issue prompting the correction, and the actions to be taken by the user.Customers were asked to set couch tolerance values to small numbers, and for radiosurgical treatment tolerance values to be set to zero. A Varian representative will contact customers to schedule installation of a software version correcting the issue. US Customers can contact Varian at 1-888-827-4265,. European customers can contact Varian at +41 41 749 8844. |  |  |  | Worldwide Distribution: Two consignees, one in the US and one in Switzerland. | Terminated | July 20, 2011 | 422 |  | Software | Output/Calculation | Software update | Software Update
Z-0101-2011 | 56682 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | NOVAPLUS (TM), SUCTION CANISTER, 1200cc, RIGID, WITH ELBOW, Cat. No. V71-3001, Rx only, NON-STERILE, Made exclusively for Novation by DeRoyal, 200 DeBusk Lane, Powell, TN 37849 |  |  | July 01, 2010 | 2010 | DeRoyal Industries Inc | Suction canisters may crack during use | Other | DeRoyal sent a Recall letter dated July 1, 2010 to all their consignees via UPS Air. The letter identified the product, the problem, and action to be taken by the consignees. Consignees were to complete the Recall Response Form by checking their inventory for DeRoyal Suction Canisters V71-3001, Lot #22622679.Indicate the number of replacement products needed on the Recall Response Form, including the address, reference PO#, etc. to be used for the replacement products and fax the form to DeRoyal at 865-362-3755 no later than July 7, 2010.Recalled products were to be returned to DeRoyal via UPS Ground Attn: Quality Control 300 DeBusk Lane Powell, TN 37849Once the recalled product was received by DeRoyal, the replacement product would be shipped per the instructions on the Recall Response Form.Distributors were to forward the recall notice to their end-users, or provide customer listing so that DeRoyal may contact them.For any questions or assistance regarding this recall call 865-362-1037. |  |  |  | Nationwide | Terminated | April 10, 2012 | 649 |  | Not_Computer
Z-0035-2011 | 56685 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Excelsior Disposable Syringe W/Normal Saline (0.9% sodium Chloride) |  |  | September 02, 2010 | 2010 | Excelsior Medical Corp | There is the potential for a dimensional issue in 6 ml syringe products, codes E0100-50, 10056-1000, 10056-240, 14056-240, 910056-1000, S5, to cause leakage and.or loss of sterility | Device Design | Excelsior Medical sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 2, 2010, to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to discontinue using syringes from the listed lots and return the attached Customer Reply Form via fax (866) 688-3185 as confirmation that notification had been received and that proper arrangements can be made to have the product returned to the firm. The customers were also instructed to inform their customers, if they have resold products, of the recall and instruct them to return the form, and verify that they have returned all unused syringes from any of the listed lots . NOTE: The customers were informed that they may alternatively have the products destroyed and certified to Excelsior that they have done so. They must include a copy of the certificate of destruction with the attached Customer Reply Form via fax. - Excelsior issued a press release on October 14, 2010. The press release advised consumers of the recall and requested that use of the product be immediately discontinued and returned to the point of purchase.Should you have any questions regarding the recalled product, please contact Excelsior Medical Customer Service at (800) 487-4276 or if you prefer by fax (866) 688-3185. You may also contact us at recall10@excelsiormedical.com. |  | 13734640 | 13,734,640 | Worldwide distribution: USA and countries including Canada and Brazil. | Terminated | August 23, 2012 | 721 | 13734640 | Not_Computer
Z-0100-2011 | 56283 | No | LNX |  |  | General Hospital | 510(k) | SoftLab version 4.01 ASCII, ASCII Base 3.1.3.0-3.1.3.32, 3.1.4-3.1.4.27, 3.1.6-3.1.6.4, 4.0.0.0-4.0.0.9; LAB/MIC Add-on 4.0.1, 4.0.1.0-4.0.1.13, 4.0.2.0 in 2006. |  |  | August 04, 2006 | 2006 | SCC Soft Computer | For SoftLab version 4.01 ASCII, the qualifying test result does not reach the Call List because the order is locked by either another application, or two Hosparams (HISRR_RBS_sync and Auto_sync_RBS). This creates a lock on the RBS server. The problem was identified in the RBS "C" rule which sends qualifying results to the Call List for SoftLab. All affected clients received the upgrade for Soft | Software design | SCC Soft Computer issued a Risk-to-Health notification with Client Notification Verbiage to the affected client using SCC Soft Computer's proprietary Task Management System in June 2006. SCC recommended SoftLab version 4.01 ASCII users change program settings to prevent locking of the RBS server. A software upgrade with corrective patch will be delivered electronically to all clients. Client was asked to acknowledge receipt of the task and grant permission to load the required software. |  | 131 | 131 | Nationwide Distribution: USA in the states of AL, AR, CA, CT, FL, GA, IA, IL, IN, KS, KY, LA, MD, ME, MI, MO, MS, NC, ND, NH, NJ, NY, OH, OK, PA, SD, TN, TX, VA, WA, WI, and WV. | Terminated | July 16, 2012 | 2173 | 131 | Software | Device Operation | Software update | Software Update
Z-0110-2011 | 55922 | Yes | NIK |  |  | Cardiovascular | PMA | Medtronic Concerto II CRT-D, model D274TRK and Model D294TRK not available in the US. Digital implantable cardioverter defibrillator with cardiac resynchronization therapy (DDE-DDR). Sterilized using ethylene oxide. Medtronic, Inc., Minneapolis, MN 55432 USA. Manufactured in: Juncos, Puerto Rico, USA. |  |  | May 05, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Medtronic is informing you of a rare device software issue in certain ICD and CRT-D devices. A software update will be available to correct his issue. There have been 5 reports of this issue occurring with no reported patient injuries or deaths. Medtronic has identified the root cause to be a rare and specific sequence of events that must occur within a few milliseconds of each other. In the u | Software design | Medtronic, Inc. sent an IMPORTANT MEDICAL DEVICE CORRECTION LETTER, dated May 2010 to physicians beginning May 05, 2010. The letter informed doctors of the issue and informed them that a software update would be available to correct this issue. The software was to be installed on the programmer, with a Medtronic representatives assistance, so that on a patient's next in-clinic interrogation the device would receive the upgrade automatically and their device would be corrected.For any questions regarding this recall physicians were to contact their local Medtronic Representative or Medtronic Technical Services at 800-723-4636. |  | 18778 | 18,778 | Worldwide Distribution - USA, Puerto Rico, Virgin Islands, and the countries of Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Brunei Darussalam, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Saint Lucia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Vatican City, Venezuela, and Viet Nam. | Terminated | March 15, 2012 | 680 | 210237 | Computer
Z-0121-2011 | 56484 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | EasyLink Informatics Systems QCFIRST Custom Rule |  |  | July 23, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | the EasyLink QCFIRST custom rule does not operate as intended. | Other | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice dated July 2010 to all customers that have the EasyLink QCFIRST custom rule installed. This Urgent Field Safety Notice included information on the problem and confirmed that a Siemens representative would contact them to facilitate the removal of the rule from the EasyLink System.For questions contact the Siemens Healthcare Diagnostics Technical Solution Center at 1-800-441-9250. |  | 82 | 82 | Worldwide Distribution - USA including AZ, CA, FL, GA, IA, IL, IN, KS, KY, LA, MO, MS, NC, NH, NJ, NM, NY, OH, OR, PA, TX, VA, WA, WI, and WY and the country of Canada. | Terminated | March 19, 2013 | 970 | 82 | Not_Computer
Z-0120-2011 | 56701 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂż Complete Knee Solution Stemmed Tibial |  |  | January 26, 2010 | 2010 | Zimmer Inc. | Zimmer had determined that two lots of these tibial implants have incorrectly positioned or missing flange plugs that were not seated in the device upon receipt to the customer. A missing flange plug could lead to a delay of surgery or possibility of wear leading to more surgery. | Employee error | The firm sent an URGENT MARKET WITHDRAWAL to its sales force and distributors on 1/26/2010. The letter identifies the affected product and explains the reason for the Market Withdrawal. The letter also explains the clinical implications and instructs consignees to locate any unused devices and quarantine them immediately. A physical count should be carried out of all affected devices and recorded on the Inventory Return Certification Form that's included with the letter. A copy of the letter should be faxed to Zimmer, Inc. and the affected product should be returned along with the form to the firm. |  | 61 | 61 units | Worldwide Distribution -- US, Canada, Singapore Australia, and India. | Terminated | October 23, 2010 | 270 | 61 | Not_Computer
Z-0109-2011 | 56762 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | in2it (I) A1c Test Cartridges (Catalog 281-0001EX) and in2it (II) A1c Test Cartridges (Catalog 281-0001), manufactured by Bio-Rad Laboratories Deeside Ltd., Flintshire, UK, distributed by Bio-Rad Laboratories, Hercules, CA. |  |  | September 07, 2010 | 2010 | Bio-Rad Laboratories Inc | Test Cartridges may leak and have the potential to generate error messages and biased inaccurate results. | Component design/selection | BIO-RAD Laboratories Deeside sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated September 7, 2010, via registered mail to all subsidiaries/customers. The letter described the product, problem and actions to be taken by the customer. The customers were instructed to stop using and discard the test cartridges from the listed lots; complete and return the attached CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM via fax +44 1244 833401; record the number of affected kits of product, confirm that they have discarded the cartridges and record the method of disposal.If you have any questions, contact +44 1244 288888. |  | 4056 | 4056 | Worldwide distribution: USA and countries including: Australia, New Zealand, Taiwan, France, the UK, Singapore, Germany, Italy, Thailand, China, Korea, Sweden, Sri Lanka, Maldives, Canada and the Philippines. | Terminated | May 17, 2011 | 252 | 4056 | Not_Computer
Z-0123-2011 | 56627 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | GE Centricity Laboratory Instrument Interface software; GE Healthcare Integrated IT Solutions, Barrington, IL 60010. The intended use: the Centricity Laboratory is used to facilitate the general clinical, anatomic pathology and cytology laboratory workflow such as order entry, results entry, instrument interfacing, results reporting and patient record retrieval. |  |  | August 27, 2010 | 2010 | Ge Healthcare It | Instrument may incorrectly report the result of an antimicrobial sensitivity test. | Software design | The firm, GE Healthcare, sent "Urgent Medical Device Correction" letters dated August 27, 2010 to GE Centricity Laboratory software customers on the same date. The letters advised the users of the patient safety issue associated with the use of GE Centricity Laboratory instrument interface to Biomerieux's Vitek analyzer where results may be incorrectly reported. The letters provided the users with specific safety instructions to follow until the software is updated. The customers were requested to complete the enclosed customer reply form and fax the completed form to 847-620-2780. Any questions should be directed to their GE Service Representative or the GE Help desk at 888-778-3375. |  | 22 | 22 units | Worldwide Distribution -- USA, including states of Georgia, Michigan, Minnesota, New York, Ohio, Oregon, Tennessee, and Texas and countries of Australia, Canada, England, India, New Zealand, Qatar, and Scotland. | Terminated | July 15, 2011 | 322 | 22 | Software | Output/Calculation | Software update | Software Update
Z-0133-2011 | 56508 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System MX 8000 IDT 16; Model # 728120; Serial #'s: 2513, 2518, 3147, 3217, 3014, 3050, 3068, 3188, 3091, 3051, 3099, 3126, 3183, 3303, 2539 &amp; 3056. |  |  | July 27, 2010 | 2010 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare has decided to recall the Computed Tomography X-Ray System MX 8000 IDT due to the fact that the software upgrade for the MX 8000 IDT 16 v4.2.1 may not provide the optional collimation of 4x4.5 (18mm). If the customer does not verify the increment and thickness coincide with the newly selected collimation suggested by the conversion report, there is potential for a 6mm gap to ap | Software change control | Philips Medical Systems (Cleveland), Inc. sent an URGENT - Field Safety Notice dated July 29, 2010 and August 23, 2010 to all customers. The notices identified the product, the problem, and the steps to be taken.Customers were instructed to locate and correct the error. Field Service Engineers were to visit each site to update the softwareFor further information customers were to contact their local Philips Customer Care Center at 1-800-722-9377 |  | 16 | 16 units | Product was sold in the following states: CA, DE, IN,OR, MA, NV, TN. Product was also sold in Canada and Australia | Terminated | September 21, 2012 | 787 | 16 | Software | Output/Calculation | Software update | Software Update
Z-0131-2011 | 56791 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | ARRAY and POLARIS Spinal System Medium Cross Connector |  |  | September 16, 2010 | 2010 | Ebi, Llc | One of the two connector ends of the medium cross connector may not be fully secured to the rod after tightening. This may not be readily apparent to the operating surgeon, resulting in inadequate torsional stability. | Other | Biomet Spine issued an Urgent Medical Device Recall letter dated September 16, 2010 to distributors and physicians, identifying the affected device and actions to be taken.Customers and distributors were requested to identify, remove, and immediately return the affected product to Biomet, following enclosed Product Return Instructions, and complete and return the Removal Tracking and Verification Form.Customers and distributors can contact Biomet at 1-800-526-2579. |  | 82 | 82 units | Worldwide Distribution: USA and the countries of Australia, Columbia, Great Britain, The Netherlands, New Zealand, and Switzerland. | Terminated | June 01, 2012 | 624 | 82 | Not_Computer
Z-0122-2011 | 56503 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | in2it (I) and (II) System A1c Test cartridges and in2it Self Test System A1c Test Cartridges, for use with the in2it (I) and (II) self test system analyzer, manufactured by Bio-Rad, Flintshire, UK |  |  | July 26, 2010 | 2010 | Bio-Rad Laboratories Inc | Inaccurate results with high bias-- Inaccurate results may be read by the machine, creating the risk for untreated hypoglycemia at elevated altitudes. | Device Design | BIO-RAD Laboratories Deeside Ltd., sent an URGENT: MEDICAL DEVICE CORRECTION letter, dated July 26, 2010, to their subsidiaries and to all customers identifying the product, the problem, and the action to be taken. Subsidiaries were to ensure that all customers received the letter with the Customer Medical Device Correction Response Form and collect the completed form from every customer. Once forms were received from all customers they were to be faxed to BIO-RAD Deeside Regulatory Affairs at +44 (0) 1244833401. Customers were instructed to run a test and visually inspect the test cartridge to detect any residual clear liquid. If the cartridge did not run properly the test was to be rejected.Customers were to run Quality Control test with each new shipment of test cartridges. Patient's results were not to be reported if the QC failed.Customers were to complete the CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM and fax it to the number provided on the form. For questions customers were to contact their regional BIO-RAD office. |  | 90703 | 90703 | Product was distributed to Bio-Rad's US subsidiary, along with 18 other distribution subsidiaries outside the US. These distributed further to other consignees. POroduct was distributed to Mexico, Korea, Singapore, Italy, South Africa, India, Australia, Spain, Greece, France, Thailand, Germany, New Zealand, China, Hong Kong, Portugal, Czech Republic, United Kingdom, Canada, Russia, Netherlands, Taiwan, Austria, Denmark, Israel, Finland and Belgium as well as to other African and Latin American countries. | Terminated | October 25, 2011 | 456 | 90703 | Not_Computer
Z-0124-2011 | 49488 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura Xper FD10 Monoplane X-ray system. |  |  | August 21, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Due to a potential geometry software lock up during a Digital Rotational Angiography (DRA) run, the stand and table movement may no longer be possible. | Software design | Philips Medical Systems North America Co. sent an URGENT - Device Correction, dated September 15, 2010, to their consignees. The letter identified the product, the problem, and the action to be taken. The consignees were informed that there was a potential geometry software lock up during a Digital Rotational Angiography (DRA) run, the stand and table movement may no longer be possible.The consignees were advised that as geometry software errors happen intermittently, no actions can be performed by the consignees to minimize the risk except for not performing DRA procedures until the firm upgrades the software.For question consignees were to contact the Philips Call Center at 800-722-9377 #5, #4, #1 and reference FCO 72200111 or FCO 72200125. |  | 47 | 47 units | Devices were distributed to hospitals and medical facilities through out the U.S. | Terminated | November 19, 2010 | 820 | 73 | Software | Device Operation | Software update | Software Update
Z-0132-2011 | 56610 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | PD Access Vascular Access Device. Escalon Vascular Access, Inc., 2440 South 179th Street, New Berlin, WI 53146: |  |  | May 05, 2010 | 2010 | Escalon Medical Corp | The packaging of the vascular access device may potentially become damaged, compromising the sterility of the product. Use of a non-sterile product increases risk of patient infection. This recall involes: Model 78050 18 Gauge PD Access Vascular Access Device, Model 78060 Extended 18 Gauge PD Access Vascular Access Device, Model 75010 18 Gauge SmartNeedle Vascular Access Devices with Expiration Da | Package design/selection | Escalon Vascular Access issued an "Urgent: Recall" letter dated May 5, 2010 to customers. The letter identified the affected product being recalled, described the problem, and the actions to be taken.. Consignees were recommended to immediately examine their inventory, quarantine the product, notify their customers, and return product affected. They also advised them that if they have further distributed the product, to please identify the customers and notify them at once of the product recall. Replacement product can be ordered from Vascular Solutions, Inc at 888-240-6001. Customer service can be contacted at 800-676-0043. |  | 188375 | 188,375 | Worldwide Distribution: USA, including the states of AL, AR, AK, AZ, CA, CO, CT, DE, DC, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, ME, MD.MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, MN, NY, NC, ND, OH, OK OR, PA, RI, SC, SD,TN, TX UT, VA, WA, WV, and WI, and the countries of: Australia, Austria, Bahamas, Canada, Colombia, Denmark, England, Germany, Greece, Hong Kong, Italy, the Netherlands, New Zealand, Norway, Portugal, Spain, Sweden, and Switzerland. | Terminated | March 28, 2013 | 1058 | 188375 | Not_Computer
Z-0127-2011 | 56462 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Hospira brand Symbiq One-Channel Infuser |  |  | February 22, 2010 | 2010 | Hospira Inc | Hospira has identified motor encoder failures in the Symbiq pumping mechanism that causes the infuser to cease operation during infusion therapy, resulting in delay or interruption of infusion therapy. | Device Design | Hospira issued a Recall Notification letter dated October 4, 2010, updating the Urgent Device Field Correction letter dated February 22, 2010, ideentifying the affected products, the issue prompting the recall, and actions to be taken by customers.Customers are to examine their inventory to determine if they have any of the affected products, and recommended to remove them from service in critical care areas, and contact Hospira to schedule removal and replacement of the motor. Replace affected devices with a loaner or corrected back-up pump if available.Custoemrs can contacvtd Hospira Global Product Safety and Complaints at 1-800-441-4100, and Technical Support Operations at 1-800-241-4002. |  |  |  | Worldwide Distribution: USA and the coutnries of Australia, Canada, Colombia, Costa Rica, Malaysia, Mexico, Puerto Rico, Singapore and the UAE. | Terminated | December 03, 2013 | 1380 |  | Not_Computer
Z-0126-2011 | 56730 | No | DNN | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | EndoTool Drug Dose Calculator-Software Version 7.2.1800.3 and 7.2.1825.5 |  |  | July 26, 2010 | 2010 | Hospira, Inc. | The EndoTool software did not always provide an audible alert when a blood glucose measurement was due. This could result in a patient not receiving an appropriate dose of insulin. | Other | Hospira issued an Urgent Device Field Correction letter dated July 23, 2010 to consignees, providing details of the potential issue and the associated patient risk. Hospira is finalizing software enhancements which will resolve this situation within the next 90 days ,and will contact each consignee when it is available to be downloaded from the Hospira website. The notification letter also included a clinical bulletin, outlining mitigation instructions for clinicians to use until the software enhancement is available. Consignees were requested to complete and return an attached Reply Form by fax.Customers can contact the Hospira EndoTool support line at 1-877-349-4582 for clinical inquiries about this recall. |  | 97 | 97 units total (96 version 7.2.1800.3 and 1 version 7.2.1825.5) | Nationwide Distribution in states of CA, FL, IL, IN, KY, KS, MI, MO, MT, NC, NJ, NM, NV, OH, SC, SD, TN, TX, VA, WA, and WY. | Terminated | February 25, 2014 | 1310 | 97 | Software | Alarm/Message | Software update | Software Update
Z-0140-2011 | 49605 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips BV Pulsera R2.2 Mobile X-ray system. The Type Product Code 718094 can be found under the X-ray trolley cover on the front lower right side of the module or under the stand cover. The device's brochure is labeled in part: "Philips...Mobile C-Arm systems BV Pulsera". |  |  | September 05, 2008 | 2008 | Philips Medical Systems North America Co. Phillips | Reliability of the Stand-Trolley cable may result in loss of images on the monitor during an examination. This intermittent problem can lead to System lock-up, where no radiation is possible and images are not visible on the monitor | Component design/selection | On 10/04/2008, Philips began sending the URGENT-Device Correction BV Pulsera R2.2 and Endura R2.2 letters to their customers. On 09/03/2008, the firm issued the Field Change Order (FCO) 71800024 and the firm's representative will replace the stand-trolley cable free of charge.Customers should contact Philips Call Center at 1-800-722-9377, #5, #2, and reference "FCO 71800024" with any questions. |  | 1331 | 1331 units | Nationwide distribution. | Terminated | December 07, 2010 | 823 | 1458 | Not_Computer
Z-0135-2011 | 56657 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | ACL TOP (Base) |  |  | August 13, 2010 | 2010 | Instrumentation Laboratory Co. | Software: New data check, "Multiple Threshold Check", was added to identify and fail abnormal (bimodal) clot curves | Software design | Instrumentaion Laboratory notified the US Distributor/Canada via BCI Service Modification Notice (#10793) on August 13, 2010. The Mandatory software upgrade to Verison V4.3.0 in the United States and Canada will be conducted by BCI service representatives and tracked through the distributor, Beckman Coulter (BCI). International Affiliates/Dealers:Worldwide Regulatory Notices were e-mailed to IL international affiliates and dealers, notifying them of the mandatory software upgrade, providing issue description and a risk summary and the support. The upgrades were to be handled through the local representatives. |  | 177 | 177 units | Worldwide Distribution, USA and Canada | Terminated | May 23, 2012 | 649 | 1076 | Software
Z-0139-2011 | 56703 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista V-LYTE Standard A, Catalog No K820 |  |  | August 19, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | High potassium content in the affected lots results in the potential for low potassium recovery for Dilution Check, QC and patient samples when run on the Dimension Vista System. | Other | Siemens sent an Urgent Field Safety Notice (dated "August 2010") to customers identifying the product, the problem and the action to be taken. Customers were instructed on how to continue testing until new lots of the recalled product were available and asked to complete and fax a Field Correction Effectiveness Check to 301-631-8467 to confirm receipt of the Urgent Field Safety Notice. For technical questions or concerns customers were instructed to contact the Siemens Technical Solutions Center at 800-441-9250. |  | 3706 | 3,706 cartons | USA, Nationwide Distribution | Terminated | October 26, 2010 | 68 | 3706 | Not_Computer
Z-0143-2011 | 56790 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | The label (which is affixed to ScanPoint Communications Cradle) states: "...VERATHON ScanPoint Docking Station...Part No. 0570-0168...Bothell, WA USA..." |  |  | September 20, 2010 | 2010 | Verathon, Inc. | Incorrect programming causes QuickPrint to fail to recognize the communication Cradles. As a result, ScanPoint Communications Cradles could not communicate with ScanPoint Online Tool while using the QuickPrint interface. This will affect the transfer of exams and device calibration through QuickPrint interface. | Software design (manufacturing process) | On 09/20/10, Verathon issued an URGENT: Medical Device Accessory Recall Notification letter dated September 20, 2010 to consignees, identifying the affected product, the issue prompting the recall, and actions to be taken by consignees.The consignees are instructed to return any cradles within the serial number range between CC012000-CC012084, and the Part Number 0570-0168.The firm will provide a replacement communication cradle at no cost. Consignees can contact Verathon Customer Care at 425-867-1348 or 800-2331-2313 or send e-mail to cservice@verathon.com. |  | 70 | 70 units | Worldwide Distribution: U.S. nationwide including states AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, KY, LA, MA, MD, MI, MN, MO, ND, NY, OH, PA, SC, TN, TX, UT, WI, and Washington DC, and the countries of Australia, Canada, France, South Korea, Taiwan, The Netherlands, and United Kingdom. | Terminated | April 01, 2011 | 193 | 70 | Other | Device Operation | Replace | Remove or Replace
Z-0413-2011 | 56580 | No | KOX | Surgical skin degreaser or adhesive tape solvent. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Detachol adhesive remover, 4 fl oz (118 mL), Marketed by Eloquest Healthcare, Inc., a Ferndale Pharma Group company, Manufactured by Ferndale Laboratories, Inc., Ferndale, MI 48220 U.S.A. |  |  | July 07, 2010 | 2010 | Ferndale Laboratories, Inc. | There is a defect on the neck surface of the bottles, causing bottles to leak. Bottle defect may potentially cause contamination of solvent stored in bottles. | Other | Ferndale Pharma Group issued an Urgent Medical Device Recall letter dated July 7, 2010 to customers identifying the affected product. Customers were instructed to identify and quarantine all affected product and return it to Ferndale Laboratories for a credit. Customers are requested to complete and return an enclosed response form.Customers can contact Ferndale at 800-621-6003. |  | 55000 | 55,000 bottles | Nationwide Distribution: Throughout the US. | Terminated | March 02, 2011 | 238 | 55000 | Not_Computer
Z-0145-2011 | 56893 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Vertex Max Drill Guide, Catalog Number 9732355, to be used with a Medtronic Computer-Assisted Surgery System to help assist with screw placement during spine procedures. |  |  | September 24, 2010 | 2010 | Medtronic Navigation, Inc | Drill guide handle may detach unintentionally during use. | Other | Medtronic Navigation sent a "Product Correction Notification" letter dated September 22, 2010 to all customers. The letter described the product, problem and action to be taken by the customers. The customers were instructed to set aside and do not use the Vertex Max Drill Guide 9732355, Lot Number 100526 and complete and return the Verification of Field Action Completion form via email: steve.murray@medtronic.com or fax: 720-890-3517. A Medtronic representative will contact the customer, visit their facility, remove the defective product, and provide a replacement Vertex Max Drill Guide that is correctly manufactured.For further information or if you have questions regarding use of the Vertex Max Drill Guide, please contact Medtronic Navigation Technical Services at 1-800-595-9709. |  | 8 | 8 units | Worldwide distribution: USA including states of DC, IN, PA, and VA and country of Japan. | Terminated | April 13, 2011 | 201 | 8 | Not_Computer
Z-0144-2011 | 56754 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis Systems VB23A-VB23H with Perivision functionality, Model Numbers: 5904433,5904649, 5904466, 5904441, 5904656, 7728350, and 7007755 |  |  | September 09, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Sporadic system crash while Perivision is performed. | Software design | Siemens issued an Urgent Field Safety Notice to all customers. The letter informed customers of the potential issues and provided additional information on the implemented resolution. To resolve the problem, a system update was to be installed on each affected system under Update Instructions AX054/10/S. For questions phone (610) 448-3237. |  | 286 | 286 units | USA - Nationwide Distribution | Terminated | September 29, 2012 | 751 | 286 | Not_Computer
Z-0147-2011 | 56844 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Recalled chairs include the following: |  |  | September 07, 2010 | 2010 | Winco, Inc. | Winco Mfg., LLC, Ocala, Florida is recalling their Care Cliner. Due to malfunctioning axles which become lose from the caster assembly. This chair is a multi position recliner with a steel frame used primarily in clinics, hospitals and treatment centers for recovery, dialysis, oncology, examination, infusion. Recalled chairs include the following: Care Cliner, XL Care Cliner, Drop Arm Care Cliner | Device Design | Winco sent an Urgent Device Recall letter dated October 1, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Replacement parts were sent out to the consignees and customers were asked to dispose of the defective casters. Any consignee not reponding by November 15, 2010, would receive an additional letter. The firm placed a follow up phone call to the non responding consignees after two weeks.Customers were advised to call 800-237-3377 or email customer service at customerservice@wincomfg.com with questions or concerns.For questions or concerns regarding this recall call 352-854-2929 x 110. |  | 18522 | 18,522 casters | Worldwide Distribution - USA (all states except Vermont and Nebraska) and the countries of Belgium, Canada, Chile, Ireland, Kuwait, Netherlands, Saud Arabia, Singapore and United Arab Emirates. | Terminated | May 20, 2011 | 255 | 18522 | Not_Computer
Z-0158-2011 | 56718 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Catheter Extension Set, Luer Activated Valve for I.V. Access, Male Luer Lock Adapter, Volume 1.0 mL, Length 6" (15 cm); an Rx sterile device; Baxter Healthcare Corporation, Deerfield, IL 60015 USA; product 2N9191 |  |  | September 09, 2010 | 2010 | Baxter Healthcare Corp. | During packaging testing, Baxter observed tears and pinholes in Chevron-style pouches for certain infusion disposable products. The pinholes and tears could compromise the sterile barrier properties of the pouch. | Package design/selection | Baxter Healthcare Corporation sent an Urgent Product Recall letter dated September 9, 2010, to all customers. The letter identified the product, the problem and the action to be taken by the customer. The customers were requested to examine their inventory in all locations to determine if they had any of the affected lots of the affected product codes and to remove the affected product from their inventory; they were to call the Baxter Healthcare Center for Service at 1-888-229-0001 to return the affected product, request a credit, or request replacement product; complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they could comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  |  | data not available | Worldwide Distribution - U.S., Canada, Japan, Australia, New Zealand, and Chile. | Terminated | September 07, 2011 | 363 |  | Not_Computer
Z-0161-2011 | 56820 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Licox Brain PMO Probe Kit; |  |  | September 15, 2010 | 2010 | Integra LifeSciences Corp. | Incorrect expiration date - The outer boxes of the IP2P Licox Complete Kit for Cerebral Pressure, Temperature and ICP Monitoring Probes, Lot 250210A are labeled with an incorrect shelf life date (10 months too long). | Employee error | Integra issued Urgent Product Recall Notification letters dated September 16, 2010 to US customers and foreign consignees, identifying the affected lots and actions to be taken by customers.Affected product is not to be further distributed and should be quarantined pending return to the firm. Customers and Distribors should contact Integra Customer Service for a Returned Material Authorization Number, and then return the product to the Distribution Center. In addition, a Recall Acknowledgement and Return Form is to be completed and returned.Customers can contact Integra at 609 936-2495. |  | 23 | 23kits | Worldwide Distribution: USA, including the states of FL, ID, IL, and NJ, and the countries of Argentina and Costa Rica. | Terminated | April 18, 2012 | 581 | 23 | Not_Computer
Z-0160-2011 | 56148 | No | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Edwards' MHANM8 Pressure Slave Cable, Part Number:693657001 |  |  | June 25, 2010 | 2010 | Edwards Lifesciences, Llc | The recall was initiated because the Edwards' MHANM8 Pressure Slave Cable was wired incorrectly, thus will not work as stated. | Process design | Edwards Lifesciences contacted customers via telephone on June 22, 2010 informing them of the product, the problem and the action they should take. Customers were asked to locate their cables and send them back overnight using the provided Edwards FedEx account number.Edwards Lifesciences indicated their personnel would follow-up with each customer until all units were returned.For any questions regarding this recall call (949) 250-2007. |  | 20 | 20 units | Worldwide Distribution - USA including CA, PA, and MI and the countries of Switzerland and Thailand | Terminated | May 24, 2011 | 333 | 20 | Not_Computer
Z-0159-2011 | 55712 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Impac Software Sequencer Verify &amp; Record System, for MOSAIQ, manufactured by Elekta Impac Software, Sunnyvale, CA. |  |  | February 19, 2010 | 2010 | Impac Medical Systems Inc | Mistreatment-- The database conversion utility used during the upgrade errantly changes the start gantry angle to zero. Any affected treatment fields, if not already treated, would start delivering from zero instead of the original prescribed value. | Process change control | ELEKTA IMPAC SOFTWARE sent a USER NOTICE dated February 19, 2010, to all customers. The notice identified the product, the problem, and the action to be taken by the customers.Customers were urged to verify the start angle of the gantry against the treatment field definition prior to delivery.Review all treatment plans and charts with conformal arcs. If a treatment plan with an errant gantry start angle was identified, customers were to cease using that plan immediately. A workaround was to be re-imported into their current version of MOSAIQ.IMPAC Support was to follow-up with each site and execute a script to detect and correct all affected treatment plans.For questions, concerns, and requests for upgrade, customers were to email support@impac.com or call their support center. |  | 19 | 19 units. | Worldwide Distribution -- USA, Australia, Belgium, France, Germany, Hong Kong, Ireland, Israel, and Italy | Terminated | March 28, 2011 | 402 | 19 | Software | Treatment/Delivery/Therapy | Upgrade | Repair
Z-0167-2011 | 56975 | No | KPT | Radionuclide dose calibrator. | Radiology | Radiology | 510(k) | Captus 3000 Thyroid Uptake System with Portable Stand option. Capintec item number 5430-2130, serial numbers: 91001, 91003 through 91010. Manufactured by Capintec, Inc, Pittsburgh, PA. |  |  | September 20, 2010 | 2010 | Capintec Inc | The spring arm on portable stand may have been over tightened during assembly, which may result in failure. If the spring arm fails, it will not balance the collimator weight, and the arm will drop to the lowest position. | Employee error | Capintec issued an Alert letter to all customers identifying the affected device, informing them of the problem and what actions are to be taken by customers. Capintec will send a replacement arm for the system, which is to be installed by the customer using included instructions. The firm included a correction acknowledgement form to be completed and returned by the customer.Customers can contact Capintec concerning this recall at 201 825-9500. |  | 9 | 9 units | Worldwide Distribution: USA in the states of NC and TN, and in the countries of Pakistan, Paraguay, Portugal, Saudi Arabia, and Spain. | Terminated | July 05, 2012 | 654 | 9 | Not_Computer
Z-0162-2011 | 56826 | No | KGZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | IV Catheter - Optiva IV Catheter. This is found in box labeled 'Graft Jacket Xpress Flowable Soft Tissue Scaffold'; labeled by Wright. |  |  | September 13, 2010 | 2010 | LifeCell Corporation | Graft Jacket Xpress kits contain catheters that will expire or have already expired prior to the labeled expiration of the human tissue component of the product on the outer box. | Labeling mix-ups | Lifecell Corporation, Branchburg, NJ initiated a recall to their distributor Wright Medical Technology, Arlington, TN on 9/13/2010. |  | 1507 | 1507 kits | Nationwide Distribution: USA, to distributor in TN | Terminated | August 17, 2012 | 704 | 1507 | Not_Computer
Z-0163-2011 | 56852 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Leksell Gamma Knife Perfexion with EXTEND Frame System, manufactured by Electa Instrument AB, Stockholm, Sweden. |  |  | September 17, 2010 | 2010 | Elekta, Inc. | The anterior fixation bar of the EXTEND mouthpiece frontpiece may separate from the dental impression tray. This may result in reduced treatment accuracy. | Other | Elekta issued an Important Notice: Urgent Field Safety Notice dated September 17, 2010 concerning Field Change Order 10000202003-A342 "Important Notice regarding recall of mouthpieces for EXTEND system, to all users of the EXTEND system for LGK Perfexion. the letter identified the issue and associated potential hazard, and requested the identification and return of all affected mouthpieces, and confirmation of receipt of the letter. Elekta will arrange return and replacement through local service representatives.Customers can contact Elekta concerning this recall at 770 300 9725. |  | 3 | 3 units | Nationwide Distribution: to the states of MO, VA, and WA. | Terminated | November 10, 2010 | 54 | 3 | Not_Computer
Z-0164-2011 | 56908 | No | HBT | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LeMaitre Vascular AnastoClip GC Vessel Closure System-Medium Size |  |  | September 22, 2010 | 2010 | Lemaitre Vascular, Inc. | Mislabeled size on external packaging: The contents section listed on the Tyvek lid labels of the Medium and Large sizes incorrectly identified the size of the clip as XL, while the device inside was either a Medium or Large size. | Packaging change control | LeMaitre Vascular issued an Urgent AnastoClip GC Device Field Safety Notice letter dated Sepember 28, 2010 identifying the affected devices and actions to be taken by customers. The letter contains a form that is requested to be returned to LeMaitre Vascular as record of notification and reconciliation. The letter provides an option for the notified hospital to return the affected product for replacement if unused or if they simply wish to have it corrected. Customers can contact LeMaitre Customer Service for replacement at 800 628-9470. |  |  |  | Nationwide Distribution: to the states of CA, FL, GA, IL, IN, LA, MD, MS, NJ, OK, OR, and TX. | Terminated | September 15, 2011 | 358 |  | Not_Computer
Z-0166-2011 | 56352 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Siemens brand 550 TxT Patient Table, Part No. 7346534, Distributed by and/or Manufactured by Siemens AG, Medical Solutions, Germany |  |  | March 31, 2010 | 2010 | Siemens Medical Solutions USA, Inc | An increase in voltage may cause a short circuit of a power transistor that may cause a smoldering and smoking effect of the electrical wiring and also an unexpected lowering of the Table. Both issues present the potential for injury to a patient or the operator. | Other | Siemens sent a Customer Information Letter on March 31, 2010, to all consignees. The letter identified the product, the problem, and the action that should be taken. Siemens implemented minor modifications to address the issues. No further action was required for the continual operation of the 550TxT Table. For questions regarding this recall call 925-602-8083. |  | 326 | 326 units | Worldwide Distribution including USA | Terminated | March 01, 2012 | 701 | 326 | Hardware | Physical Safety Hazards | Correction | Repair
Z-0170-2011 | 56909 | No | LPB |  |  | Cardiovascular | PMA | Boston Scientific Peristaltic Tubing Kit, Model 2104, UPN M0043521040. For use with Boston Scientific CHILLI Cooled Ablation Catheters and Model 8005 Continuous Operation Pump. Sterilized using Ethylene Oxide. Contents: peristaltic pump fluid delivery extension tube set. Manufactured for: Boston Scientific Corporation, One Boston Scientific Place, Natick, MA 01760-1537, USA. |  |  | September 27, 2010 | 2010 | Boston Scientific Corporation | Boston Scientific is conducting a Recall on CircuCoolÂż Tubing Sets used with the Chilli IIÂż Cooled Ablation System because there may be a potential for a sterile barrier breach which could compromise device sterility and potentially the sterile field. Even though there have been no reported complaints. | Packaging | Boston Scientific Corporation sent an Urgent Medical Device Recall letter, dated September 27, 2010, to all customers via overnight Federal Express. The letter described the product, the problem, and the action the customer should take. The letter instructed customers to cease distribution of affected devices and remove them from their inventory, segregate in a secure location for return to Boston Scientific. A Reply Verification Tracking Form was to be completed and returned via email or faxed to (510) 624-2475 even if the customer did not have any of the affected product. If affected units were located the customer was to call (510) 624-2563 to obtain a Returned Good Authorization (RGA) Number and return to:Boston Scientific CorporationUS Distribution CenterBoston Scientific Marina BayCustomer Fulfillment Center500 Commander Shea Blvd.Quincy, Massachusetts 02171RGA: _________________For additional questions call (510) 624-2563. |  | 4694 | 4031 US , 663 OUS | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NY, OH, PA, RI, SC, TN, TX, VA, WA, and WV and the countries of France, Germany, Great Britain, Hungary, Israel, Italy, Netherlands, Saudi Arabia, Spain, and Taiwan | Terminated | April 17, 2012 | 568 | 4694 | Not_Computer
Z-0171-2011 | 56921 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobilett XP and Mobilett XP Hybrid, Mobile radiology device |  |  | October 04, 2010 | 2010 | Siemens Medical Solutions USA, Inc | Units not equipped with protective plastic cover over electronics. | Device Design | Siemens Medical Solutions USA, Inc issued a mandatory field update. The letter identified the product, the problem and the action to be taken by the customer. All systems that were not yet equipped with protective cover over the D972 board would be upgraded.For questions regarding this recall call (610) 448-3237. |  | 177 | 177 units | Nationwide Distribution including AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, LA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, PA, TX, VA, WA, WI, WY AND PUERTO RICO. | Terminated | September 29, 2012 | 726 | 177 | Not_Computer
Z-0134-2011 | 56902 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Medtronic Octopus Nuvo Tissue Stabilizer, model TSMICS1. The Octopus Nuvo Tissue Stabilizer is a disposable tissue stabilization system consisting of a rigid shaft, detachable tissue stabilizer headlink and a flexible articulating arm. The device is assembled within the thoracic cavity. Sterilized using ethylene oxide, non pyrogenic, disposable, single use only. |  |  | September 13, 2010 | 2010 | Medtronic Perfusion Systems | Medtronic has received reports regarding the Octopus Nuvo Tissue Stabilizer stating that during use, the collet (a component of the device that stabilizes the head link) fractured, thereby causing immediate separation of the head link from the shaft of the device. The resulting potential hazards are that fragments of the collet could fall into the patient's chest cavity and/or damage the heart ti | Component design/selection | Medtronic sent an "Urgent Medical Device Recall Notice" dated September 14, 2010, to Risk Managers at each affected account. The letter described the issue, identified affected lot numbers, asked for immediate discontinued use of the device, quarantine all unused devices and to return unused devices to Medtronic CardioVascular, 7611 Northland Drive, Brooklyn Park, MN 55428 and requested a recall certificate be faxed back to Medtronic indicated acknowledgement of the Recall Notice and quantity to be returned.If you need additional information, please contact Medtronic Cardiovascular Lifeline Technical Services at 877-526-7890, or your local Medtronic sales representative. |  | 571 | 571 units | within the US to: AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, MA, MI, MN, MS, NJ, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI and Washington D.C. OUS to include: Belgium, Canada, Germany, Italy, Poland, Switzerland, United Kingdom. | Terminated | April 16, 2012 | 581 | 571 | Not_Computer
Z-0168-2011 | 56794 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Volcano s5/s5i imaging system Family Software version 3.2, Part Numbers 435-0602.01 to 05, including ChromaFlo and Virtual Histology, manufactured by Volcano Corporation of Rancho Cordova, CA. |  |  | June 29, 2010 | 2010 | Volcano Corporation | Software defect results in improper image brightness, which may result in introduction of artifacts, which may cause misclassification of results and affect classification of tissue types. This may result in misdiagnosis and inappropriate treatment. | Software design | Volcano issued a Voluntary Recall of Volcano s5 Family Software version 3.2 letter dated June 30, 2010, identifying the software defect and actions to be taken by Volcano. Field Service Engineers will visit each site and reinstall the previous software version, which will resolve the issue.Customers can contact Volcano at 916 281-2790. |  | 79 | 79 software kits | Worldwide Distribution: USA, and the countries of France, Germany, Japan, and the UK. | Terminated | December 22, 2010 | 176 | 79 | Software | Display/Image | Reinstall previous version | Software Update
Z-0169-2011 | 56858 | Yes | NOX |  |  | Orthopedic | Contact ODE | Bone Void Filler, sterile orthopedic product, 510 (k) #K053228. The product is delivered in a syringe which is packaged in an aluminum-backed polyethylene pouch and is distributed one (1) pouch per box. The product affected by this recall was distributed by the recalling firm under six (6) different names. They are: Form-Putty; SBvFShape; Valeo BP; X-Putty; TheriGraft Putty; and SynFlow. |  |  | August 26, 2010 | 2010 | Theken Spine Llc | The firm's quality control testing determined that a separation was occurring between the flowable carrier and the B-TriCalcium Phosphate within some of the product syringes. | Other | Theken Spine Llc sent Recall Notification Letters dated August 26, 2010, to all customers. One (1) document, titled an 'Urgent Product Recall Notification' letter was addressed: 'ATTENTION; PRODUCT RECALL COORDINATOR/RISK MANAGER.' The other letter was titled: 'RECALL NOTIFICATION.' The recall letters were accompanied by an attached 'Recall Acknowledgement and Return Form' and were sent out via Federal Express. The letters identified the product, the problem, and the action to be taken by the customer. The letters requested that the customers review their product in inventory and identify any product remaining from the affected lot numbers; part numbers; and/or item names for immediate return of these items to the recalling firm. The letters also asked that the customer complete and return the attached Recall Acknowledgement and Return form according to the instructions provided on the form. In addition, the 'RECALL NOTIFICATION' letter instructed the distributor/customers provide information concerning their sub-account customers to Theken Spine so they could issue a Recall Notification letter to these customers as well. To obtain instructions from the recalling firm concerning product return and issuance of credit, customers were instructed to call (866) 942-8698. For questions regarding this recall call (330) 825-8258. |  | 2904 | 2904 units (syringes) | Nationwide Distribution - including UT, SD, TX, AL, PA, CA, MI, OH, NJ, NH, MD, ID, FL, LA, NM, TN, IA, and OR. | Terminated | October 27, 2011 | 427 | 2904 | Not_Computer
Z-0288-2011 | 53863 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | The product is a 19" Spacelabs Medical Flat Panel Display, Model 94260-19, manufactured by ELo Touchsystems |  |  | October 23, 2009 | 2009 | Spacelabs Healthcare, Incorporated | The Spacelabs Medical Flat Panel Display, Model 94260-19 may fail prematurely and causing the screen to go blank. This inability to view data may result in delayed patient treatment. | Other | On 10/23/09, Spacelabs issued URGENT: MEDICAL DEVICE RECALL notification letters dated October 23, 2009 to their consignees.The notification states the recalled product as Spacelabs Medical 19" Flat Panel Display, Model 94260-19. The consignees are informed that there is apossibility that the internal power distribution circuit board will fail prematurely, causing the display to go blank. Audio alarms will continue to function normally.The firm's Field Service Engineer will be contacting the consignees to set up a convenient time to schedule the upgrades of the affected displays. Displays with part number ET1920L-7SKR-1-35S-SPC are not part of this recall and should not be sent in for upgrading.The firm recommends that their consignees brief their staff on the possibility of the display going blank to help prevent an adverse event from happening. The firm will be providing loaners while the affected displays are upgraded at free of charge.If the consignees have any questions about the recall, they could contact Spacelabs Medical at 1-800-522-7025 and select 2 for Technical Support . Consignees outside the United States may call 1-425-657-7200 x 5089. |  | 2560 | 2560 units (2020 domestic, 540 international) | Worldwide Distribution: USA in the states of AK, AL, AR, CA, CO, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WY, and the following countries: Australia, Austria, Canada, Czech Republic, Egypt, France, Germany, Grand Cayman, Greece, India, Israel, Italy, Kuwait, Lebanon, Mexico, Netherlands, Pakistan, Panama, Poland, Puerto Rico, Saudi Arabia, Singapore, Slovakia , South Africa, Switzerland, Trinidad and Tobago, Turkey, United Kingdom, and Venezuela. | Terminated | November 03, 2010 | 376 | 2560 | Hardware | Display/Image | Replace | Remove or Replace
Z-0308-2011 | 55242 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Panta Support Device PantaÂż Arthrodesis Nail Instrumentation Tray One unit per package. |  |  | March 01, 2010 | 2010 | Integra LifeSciences Corporation | Inability to properly align the calcaneal screw placement on a consistent basis, due to a manufacturing defect of of the targeting frame (Panta Support Device). | Nonconforming Material/Component | OIntegra issued Urgent Product Recall Notification letters dated 3/1/2010 to customers and distributors to notify them of the problem. Customers were requested to check their inventory to determine if they have an of the affected devices. If so, custo0mers should stop using the product. Custoemrs with the affected product should contact their local Sales Representative for assistance, or contact Integra at 800-654-2873 - 3 to obtain a returned material authorization number, and return the product to their local representative.All customers were requested to complete and return a Recall Acknowledgement and Return Form.. |  | 108 | 108 Units | Worldwide distribution: USA including states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL, LA, MA, MD, MI, MO, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SD, TX, UT, VA, WA and WI; and countries of: Australia, Austria, Belgium, Czech Republica, Denmark, Finland, France, Germany, Ireland, Italy, Portugal, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | February 16, 2012 | 717 | 108 | Not_Computer
Z-0302-2011 | 56938 | No | LTI |  |  | Gastroenterology/Urology | PMA | SAGB QuickClose Swedish Adjustable Gastric Band including Velocity" Injection Port and Applier. Product Code BD2XV, Lot #: 1040602, , 1040603, , 10270503, , 11020504, , 11030502, , 11090510, , 11090511, 1090512, 12030502, 12070502, 12070503, 12070504, 12200501, 12200502, 12200503, 12290505, 12290506, 01120603, 01200602, 01260601, 01260602, 01260603, 01260604, 02210605, 02210606, 02210607, 02280622, 02280623, 03080601, 03080602, 03080605, 03220611, 03220612, 05150601, 05150602, 05150603, 05150604, 05300613, 05300614, 06190606, 06190607, 06190608, 06190614, 06190615, 06190616, 06190617, 06190618, 06260601, 07060601, 07100605, 08140607, 08140608, 09050601, 09050602, 09050603, 09050604, 09050605, 09130601, 09130602, 09130603, 09130604, 09130605, 09180602, 09180603, 09180604, 09250602, 10110601, 10110602, 10110603, 20002024, 20002082, 20002185, 20002186, 20002193, 20002438, 20002439, 20002440, 20002441, 20002484, 20002485, 20003152, 20003153, 20003154, 20003155, 20003196, 20003362, 20003363, 20004437, 20004438, 20005244, 20005659, 20006819, 20006921, ZGNB43, ZGNB44, ZGNB45, ZGPBCC, ZGPBCD, ZGPBCJ, ZGPBCK, ZHBBBN, ZHBBBP, ZHBBBT, ZHCBB0, ZHCBB1, ZHCBB2, ZHCBB3, ZHCBB5, ZHCBB6, ZHCBB7, ZHCBB8, ZHCBB9, ZHDBF7, ZHFB76, ZHFB77, ZHFB79, ZHFBBW, ZHFBBY, ZHGBBJ, ZHGBDJ, ZHGBDK, ZHGBDL, ZHHBD6, ZHHBD7, ZHHBD8, ZHHBD9, ZHHBFB, ZHHBFC, ZHJBBJ, ZHJBBK, ZHJBBL, ZHJBBM, ZHJBBN, ZHJBBP, ZHMBDP, ZHNBC1, ZHNBC2, ZHPBBT, ZJBBBF, ZJBBDP, ZJCBB6, ZJCBDP, ZJCBGL, ZJDBBD, ZJDBBF, ZJDBBG, ZJDBBH, ZJDBBJ, ZJDBBK, ZJFBL6, ZJFBML, ZJHBCH, ZJHBK7, ZJJBC3, ZJKBC1, ZJKBHC, ZJKBK2, ZJMBB4, ZJMBN9, ZJNBBV, ZKDBM6, ZKDBM7, ZKFBDC, ZKFBDH, ZKGBDY, ZKHBBJ, ZKHBCR, ZKHBFG, ZKNBCK, ZKNBHC, ZLBBFK, ZLDBB8, ZLFBBC, ZLFBGG, ZLGBFV &amp; ZLGBG5. |  |  | September 27, 2010 | 2010 | Ethicon Endo-Surgery Inc | The Tubing Strain Relief component of the Injection Port may detach from the Locking Connector, and migrate from its original position on the Locking Connector. This may result in: Tubing Strain Relief migrating on the Tubing; Kinking of the Tubing resulting in difficulty/inability to adjust the fluid volume within the Band; or a free foreign body in the abdominal wall or abdominal (peritoneal) ca | Other | Ethicon Endo-Surgery issued an Urgent : Devoced Recall Event 2240 letter dated October 8, 2010 to all consignees identifying the affected devices and actions to be taken.Consignees were instructed:DO NOTt use product from the affected lots, but implanted devices do not need to be removed as part of this recall.Examine inventory and remove the affected product.If they DO have the affected product, complete the Business Reply Form, keeping a copy for records, and place the original form in the box with the product to be returned, affix the pre-printed label, and return to Stericycle.If they DO NOT have the affected product, complete the Business Reply Form, and fax it to Stericycle.A redesigned replacement product will be sent out to consignees within 30-60 days of product return.Ethicon Endo-Surgery can be contacted at 513 337-8208 concerning this recall. |  | 26966 | 26,966 pieces | Worldwide Distribution: USA including the following states: AK, AL, AR, AZ, CA, CO, CT, Washington DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, Puerto Rico, RI, SC, TN, TX, UT, VA, VT, WA, WI and WV, and the following countries: ARAB EMIRATES, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, EGYPT, FRANCE, GERMANY, GREAT BRITAIN, GREECE, GUATEMALA, HONG KONG, HUNGARY, INDIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LIBYA, MEXICO, NETHERLANDS, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, RUSSIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY and VENEZUELA. | Terminated | November 18, 2011 | 417 | 127130 | Not_Computer
Z-0130-2011 | 56735 | No | LPN | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | AMO Total Care (1) Contact Lens Solution, Formula: 8772x; Product Code: 90275 (240 ml). |  |  | July 23, 2010 | 2010 | Abbott Medical Optics Inc (AMO) | The recall was initiated after AMO determined, through internal testing, that a limited number of the flip top caps used during production of these solutions may leak and, although unlikely, the sterility of the product may be compromised. Products that are non-sterile have the potential to cause eye infections, which may be sight threatening. | Pending | AMO prepared and issued on July 28, 2010 both a customer and distributor notification letter as part of their customer communication process. The letter informed the customers of the reason for recall, the products involved, the hazard involved and actions to be taken regarding the affected products.Customer with questions related to the COMPLETE Multi-Purpose Solution and/or Total Care (1) Contact Lens Solution return process, were instructed to contact an AMO Customer Service Representative at 1800 931 663. |  | 144 | 144 units | International Distribution Only -- United Kingdom, Lebanon, Dubai, Bahrain, Russia, Kazakhstan, Ireland, and France. | Terminated | October 13, 2011 | 447 | 34224 | Not_Computer
Z-0309-2011 | 56837 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 3.5mm LCP Periarticular Proximal Humerus Plates |  |  | September 16, 2010 | 2010 | Synthes USA (HQ), Inc. | screw can pass through plate hole | Nonconforming Material/Component | Synthes (USA) notified all Synthes Trauma Hospital accounts of record of this recall via an Urgent Medical Device Recall letter dated September 27, 2010. The notification included instructions for hospital accounts to immediately cease use of the product, examine inventory, remove the 3.5 mm LCP Periarticular Proximal Humerus Plates from their shelves and complete and return the attached Verification Section form via mail: Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132; fax: 610-215-9005 or Scan/email: FieldAction@synthes.com. If they have any of the identified devices, they were to call Synthes at 1-800-479-6329 to obtain Authorization number prior to completing Verification form.If you have any questions, please call 1-800-620-7025 x5403 or contact your Synthes Trauma Sales Consultant. |  | 2849 | 2849 units | Worldwide distribution: USA and countries including: Bermuda, Korea, Italy, Japan, Georgia, Qatar and Turkey | Terminated | April 16, 2015 | 1673 | 2849 | Not_Computer
Z-0287-2011 | 56324 | No | DTB | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | PMA | enpath medical Myopore Sutureless Myocardial Pacing Lead, Bipolar (labeled as Unipolar) Connector, , REF 511212. 54 cm. Distributed by Guidant/CPI. Contents in unopened, undamaged package have been sterilized by ethylene oxide gas. Manufactured by enpath medical, Inc., 2300 Berkshire Lane North, Minneapolis, MN 55441 USA. |  |  | May 13, 2010 | 2010 | Greatbatch Medical | The Myopore Bipolar Epicardial Leads packaging is labeled as unipolar leads. The lead body itself is correctly labeled as bipolar. This could result in a bipolar lead being implanted with a pulse generator operating in unipolar mode. | Employee error | Greatbatch Medical issued an Urgent Medical Device Recall letter dated May 13, 2010 to their distributor. The letter described the problem and product involved, and gave the action required which included:1) immediately discontinue distribution of the units and return any unused units to Greatbatch Medical.2) Reconcile medical recordsbetween BSC and Greatbatch Medical to confirm implant data3) Confirm pulse generator type (bipoloar, unipolar, or unknown) used with each serial number listedGreatbatch will work with customers to contact physicians as necessary regarding patient monitoring. Greatbatch Medical can be contacted at 763-951-8312. |  | 34 | 34 | Worldwide Distribution: to USA and Europe, through distributor in MN. | Terminated | May 10, 2012 | 728 | 34 | Not_Computer
Z-0312-2011 | 55287 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | STar Drive System, |  |  | March 30, 2010 | 2010 | FHC, Inc. | Fixation thumbscrew on the DBS depth stop adaptermay be overtightened and damage the implantable lead. | Device Design | FHC issued an "URGENT: MEDICAL DEVICE CORRECTION" letter dated March 29, 2010, to all customers. The letter described the product, problem and actions to be taken by the customers. FHC informed the customers that that they had returned to the original design of the fixation screw in the Depth Stop Adapter. FHC sent replacement Depth Stop Adapter with an original fixation screw. The customers were instructed to discard the recalled Depth Stop Adapter (s) and replace with the enclosed corrected version(s) and sign and fax the receipt notification to 207-666-8309 as proof of their receipt of notification. For assistance with the product replacement, contact the FHC's technical support line at 800-326-2905. This line is staffed 24 hours a day.Should you have any further questions regarding this matter, please contact Dorothee Sasse, product manager, or Keri Seitz, recall coordinator at 207-666-8190. |  | 7 | 7 | Worldwide distribution: USA in the states of AZ, CA, FL, ID, NC, MA, MI, MN, NV, RI, UT, VA, and WA, and countries of Columbia and South Africa. | Terminated | February 25, 2015 | 1793 | 63 | Not_Computer
Z-0316-2011 | 55748 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACS Sample Prep Assistant III; (BD SPA III), Manufactured by BD Biosciences, San Jose, CA. |  |  | February 19, 2010 | 2010 | BD Biosciences, Systems &amp; Reagents | Due to inaccurate sample dispense, use of product may result in getting incorrect results on samples prepared using the device. Having a piece of the lid partially or fully clogging the Probe during piercing a hole in the lid of tube may cause inaccurate sample dispense, resulting in incorrect test results.The piece of the lid clogging the Probe can be in the processed sample tube. It has the | Device Design | BD Biosciences sent an "Important Product Information" letter dated February 2010 and an "URGENT: Follow-Up to Letter Dated February 2010" letter dated July 2010 to customers identifying the affected device and actions to be taken: Customers were instructed to stop use of the SPA probes and to contact BD for replacement of the affected probes and assistance in the event of probe clogging. BD has arranged replacement of affected probes. BD's followup communication detailed the potenial hazards and reiterated remedial actions in the event of error messages from FACS flow cytometers.Customers with questions about the performance of theirr SPA instrument or the content of the letters may contact BD Customer Support at 1.877.232.8995, prompt 2. |  | 349 | 349 units | Worldwide distribution: USA including OH, NC, AL,TX, FL, CA, NY, and MN and the country of Belgium. | Terminated | May 18, 2011 | 453 | 349 | Not_Computer
Z-0315-2011 | 56499 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Distal Centralizer 10 mm |  |  | July 28, 2010 | 2010 | Ortho Development Corporation | The hip stem did not fit the distal centralizer because the hole was not drilled with the appropriate taper. | Nonconforming Material/Component | Ortho Development sent a Field Notification Memo dated July 28, 2010 to Japan Medical Dynamic Marketing instructing them to reconcile parts that may have been used and return unused parts to Ortho. Customers can contact Ortho Development at 1-801-619-3419. |  | 39 | 39 units | Worldwide Distribution: Japan | Terminated | November 04, 2010 | 99 | 39 | Not_Computer
Z-0319-2011 | 56668 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT BNP cartridge |  |  | January 01, 2008 | 2008 | Abbott Point Of Care Inc. | i-Stat cartridges may not fill, cartridge labels or channels may appear discolored, cartridges may have fluid visible in the cartridge channel, and cartridge packaging may be distended. | Other | Abbott Point Of Care Inc.sent notification letters with a customer reply form to all affected customer in January, 2008. All distributors outside the US who were sent shipments from the affected lots were notified of the lot numbers that were affected. Affected product is to be quarantined and returned for replacement. Customers with questions about this recall can contact Abbott Point Of Care at 609 454-9000. |  | 482500 | 482,500 (total) | Worldwide Distribution: USA and Canada. | Terminated | September 14, 2011 | 1352 | 482500 | Not_Computer
Z-0310-2011 | 56825 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | smith&nephew; EZCARE Negative Pressure Rx only. |  |  | January 29, 2010 | 2010 | Smith &amp; Nephew Inc. | Smith&Nephew; Inc., Advance Wound Management, St Petersburg, FL is recalling their Ezcare Negative Pressure Wound Therapy pumps manufactured by BlueSky Medical Group Inc., between July 17, 2007 and December 13, 2007 and distributed by Smith&Nephew; between July 18, 2007 and January 14, 2010. Affected products codes are 66800187, 66800322, 66800323, 66800324, and 66800325 with Serial Numbers startin | Other | Smith &amp; Nephew Inc. FEDEXed overnight an URGENT DEVICE RECALL letter dated January 27, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were to check their inventory and identify any affected product, fill out and return the response form by both mail and fax, and return the affected product with the prepaid Federal Express (FedEx) Ground label provided by Smith &amp; Nephew. For any questions regarding this recall contact the Customer Care Center at (800) 876-1261. |  | 2717 | 2717 units | Worldwide Distribution - US including AL, AR, AZ, CA, GA, IL, IN, MA, MD, MD, MO, NC, NJ, NY, OH, OR, PA, TN, TX, and UT and the countries of Australia, UK, South Africa, Canada, Puerto Rico, and New Zealand. | Terminated | November 10, 2010 | 285 | 2717 | Not_Computer
Z-0320-2011 | 52231 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Straumann T-Sleeve 2.8mm, height 6mm, guided |  |  | June 04, 2009 | 2009 | Straumann Manufacturing Inc. | Incorrect Inner diameter of the T-sleeve | Process control | Strauman USA LLC notified the consignee by telephone on June 4, 2009, and sent an Urgent Device Recall letter dated, June 8, 2009 as a follow-up. The letter described the product, the problem, and the action to be taken by the customer. The customer was asked to return any affected product, in the prepaid preprinted mailer pouch as soon as possible to:T-sleeve Recall, Straumann USA, LLC60 Minuteman RoadAndover, MA, 01810Customers were asked to sign and indicate on the Recall Acknowledgement Form if they would like a replacement product or a credit to their account for the product returned and fax the form to (978) 747-0023.For any questions regarding this recall call (978) 747-2575. |  | 3 | 3 units | Worldwide Distribution - Canada | Terminated | February 20, 2012 | 991 | 3 | Not_Computer
Z-0321-2011 | 56785 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Single-stage 3.5mm Healing Abutment, 4 mm, REF SYHA40, Rx Only, STERILE R, gamma irradiated, expires 2014-44, CE 0473, BIOHORIZONS, BIRMINGHAM, AL 35244. |  |  | September 07, 2010 | 2010 | BioHorizons Implant Systems Inc | Healing abutment may not fully seat on the implant. | Other | Consignees were notified by an Urgent: Medical Device Recall letter on 09/07/2010. The letter identified the affected product and gave some background regarding the issue. It also stated the possible impact and the cases that were not affected. Consignees were instructed to contact BioHorizons Returns &amp; Complaints Manager Rick Self at 205-986-1250, if they have unused product from the affected lot, for a Return Authorization Number and to receive replacements. The attached Notification Certification form is to be completed and returned via fax. |  | 396 | 396 | Worldwide Distribution -- USA, including PA, MI, NY, AL, WA, and WV, and countries of Canada and Chile. | Terminated | December 13, 2010 | 97 | 396 | Not_Computer
Z-0305-2011 | 56164 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | B. Braun addEASE Binary Connector, Catalogue #N7990, 200 units (50 units per bag with 4 bags per shipping box) |  |  | June 24, 2010 | 2010 | B. Braun Medical Inc | During this insertion, fragmentation of the PAB container stopper may occur resulting in a small amount of visible particulates in the solution. | Device Design | On July 8, 2010 B Braun sent an Urgent - Medical Device Recall letter, dated June 28, 2010, to consignees informing them of the recall and that further distribution or use of the affected product should cease immediately. The attached "Product Removal Acknowledgement" form should be used to record the number of units and full-unopened cases that customers have in their inventory. The completed form should be returned via mail or fax within 2 weeks of receipt. If consignees have any full or partial cases of these affected products, they are to please call the Customer Support Department at (800) 227-2862 to arrange for return. Canadian customers may call (800) 624-2920. International customers may call (610) 691-5400. A Customer Support Representative will provide consignees with instructions for handling the affected product. The firm will arrange for unopened cases to be returned to B. Braun Medical, Inc. for proper disposal. Individual units (out of shipping cases) may be destroyed at the consignees facility.Consignees were informed that B. Braun is currently working to resolve this issue and provide an acceptable alternative to the market. They expect this product to be available in August 2010. In the interim, B. Braun PAB containers can continue to be used safely with a standard syringe and needle in accordance with the Direction for Use. If customers have additional questions regarding with a matter, they are to call B. Braun Customer Support Department at the number listed above. Should consignees identify product where this condition has occurred, consignees were asked to report it promptly to B. Braun Medical Inc. by contacting the B. Braun Clinical and Technical support at 1-800-854-6851. |  | 33905600 | 33,905,600 units | Worldwide Distribution -- USA and Canada. | Terminated | October 19, 2011 | 482 | 36405600 | Not_Computer
Z-0307-2011 | 56750 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | WalkMed Triton Pole Mount Infusion Pump, Catalog 300000. |  |  | June 17, 2010 | 2010 | Walkmed Infusion LLC | Pump door may be in a near shut position, but unlatched, and the "Door Open" alarm may not sound. | Device Design | All customers were notified by letter on 06/17/2010, describing the affected product and advising them of the possible malfunctioning alarm in the case a user does not follow the instructions as specified in the Triton Operation Manual. The letter also reminded all users to follow the procedures in the IFUs. The firm stated that it was best if customers re-educated all nurses on the proper door closure procedure to help prevent the potential bypass issue from occurring. Attached to the letter were two pages explaining the proper process for closing the pump door. Follow-up phone calls were made to coordinate an upgrade to all affected pumps. Questions or concerns should be directed to Ross Kurz at 720-531-4933. |  | 1565 | 1565 pumps | Nationwide Distribution -- AL, CA, FL, and PA. | Terminated | March 14, 2012 | 636 | 1565 | Software | Alarm/Message | Software update | Software Update
Z-0146-2011 | 56833 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Sigma Spectrum VOLUMETRIC INFUSION PUMP, with or without Master Drug Library, Part Number 35700 -- |  |  | September 15, 2010 | 2010 | Sigma International General Medical Apparatus, Llc | The infusion pumps have the potential to fail causing inaccurate flow conditions during use. These conditions range from back flow to free flow, which could result in over-infusion. | Device Design | Urgent: Medical Device Recall Notification letters and forms (dated 9/14/10) were sent to the consignees via Certified Mail on/about 9/15/10. The letters inform the consignees of the recall of the pumps since these units may fail suddenly causing inaccurate flow conditions during use, ranging from back flow to over-infusion, including free flow. Consignees are instructed to check their inventory to verify whether their serial numbers fall within the range of pumps being recalled. Due to possible bearing failure, SIGMA is taking the precautionary measure of servicing or replacing all affected Spectrum Infusion Pumps and is requiring the return of the recalled devices to its facility. SIGMA has provided the following instructions for users: (1) Do not use the affected Spectrum Infusion Pumps on patient populations where inaccurate flow, ranging from back flow to over-infusion, including free flow, could result in serious adverse health consequences or death. (2) Do not use on neonatal patients. -- If the user facility has no acceptable alternative methods for delivering therapy, the user should weigh the risks and benefits for the continued use of the affected pumps. If you must use these pumps, SIGMA recommends the following actions which may mitigate the risk. (1) for pediatric and adult patients, an add-on Buretrol may limit the amount of an over-infusion if the failure occurs. Always follow the Buretrol manufacturer's instructions. (2) consider using air-eliminating filters. Under free flow conditions pumps cannot stop fluid flow. If air is in the line with fluid, the air will also infuse. (3) Frequently observe that the drip rate (mL/hr) compares to the programmed pump flow rate. Also, make an observation of the container volume to make sure it is correct. If a discrepancy is observed, immediately remove the pump from service and treat the patient as appropriate. There are additional instructions to Asset Management Customers: SIGMA requests that each A |  | 17557 | 17,557 Infusion Pumps | Nationwide Distribution | Terminated | June 28, 2017 | 2478 | 17557 | Not_Computer
Z-0323-2011 | 55655 | No | IPF | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | The device is labeled in part: " RS Medical RS-4i Stimulator". The device is sold in a kit that includes a carrying case, connector cables, pads (electrodes), battery charger, patient document, instructional DVD, return mail bag, and the operation manual. |  |  | June 11, 2008 | 2008 | International Rehabilitative Sciences, Inc. | The charge current is leaking through two Diodes, which are supposed to block the reverse current, when the device is warm during the main battery charge operation, it then causes the lithium battery to swell around two years of use and it causes the device to be inoperable. | Other | Since June 11, 2008, RS Medical has taken the following actions when the device was sent back through a consumer complaint, or when the device was returned for routine maintenance, or through the rental service. The firm has upgraded the blocking diodes with higher temperature rated diodes, replaced the Lithium coin cell battery, and the software register bit has been changed to denote for subsequent automated testing that the diode upgrade/Lithium coin cell battery replacement has occurred. The firm reported that the device must be returned for the rechargeable NiMH battery replacement around two years of use, and the battery should not be replaced by patients.On 10/29/10, the firm began sending out URGENT NOTICE of VOLUNTARY RECALL letters to their customers. The firm will continue to send out the letters through 2011. The letter states the recall product as RS-4i Sequential Stimulator, Affected Serial Numbers 4i300000 to 4i467119. The customers are informed that "RS-4i Sequential Stimulators, only those within the affected serial number range indicated above, are susceptible to early expiration of the lithium coin cell battery that provides back-up power to the device memory when the stimulator is powered OFF. Customer experience indicates that the 3V lithium coin cell battery can be depleted as early as one third of its rated life. The RS-4i Stimulator will perform as intended until the lithium coin cell battery is fully depleted. When this happens, the Stimulator will turn OFF and fail to turn ON, even if it has been charged by battery charger connection to a standard electrical wall outlet".Customers are advised to make a schedule for a free Return and Repair of a RS-4i Sequential Stimulator and should call the Toll Free Service Line at 877-363-5753 (English) or 877-363-5754 (Spanish) from 6AM to 5PM (Pacific Time) Monday-Friday.If the customers have any other questions or to review additional information before scheduling a device |  | 166131 | 166, 131 units | Nationwide Distribution | Terminated | September 10, 2012 | 1552 | 166131 | Battery | Device Operation | Replace | Remove or Replace
Z-0322-2011 | 56255 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator/Monitor |  |  | July 16, 2010 | 2010 | Philips Healthcare Inc. | Updated Instructions for Use- HeartStart MRx monitor/defibrillator. The update provides more detailed information on:" Default ECG Lead Behavior" Synchronized Cardioversion using External Paddles | Labeling design | Philips issued an Urgent Voluntary Medical Device Correction notification/Field Safety Notice and an Instructions for Use Addendum on July 20, 2010, to affected customers. The Notice identified the product, the problem, and the action customers should take. In the USA, the notification sent to all affected customers via UPS and confirmation of delivery would be tracked. Outside the USA, Philips representatives in the affected geographies would manage their mailings locally. Customers were instructed to review the HeartStart MRx Instructions for Use addendum and keep a copy with their equipment - Instructions for Use. Additionally, they were instructed to insure that all users are fully trained on the MRx and, in particular, ensure that they review and understand have reviewed and understand the additional detail provided in the notification.For questions regarding this recall call (800) 722-9377. |  | 41119 | 41,119 units | Worldwide Distribution - US Nationwide and the countries of: CANADAAFGHANISTANALBANIAARGENTINAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHRAINBELGIUMBOSNIA&HERZEGOVINA;BRAZILBRUNEI DARUSSALAMBULGARIACHILECHINACOLOMBIACYPRUSCZECH REPUBLICDENMARKECUADOREGYPTESTONIAFINLANDFRANCEGERMANYGREECEGUADELOUPEHONG KONGHUNGARYICELANDINDIAINDONESIAIRELANDISRAELITALYJAPANJORDANKENYAKUWAITLATVIALEBANONLIBYAN ARAB JAMAHIRILITHUANIALUXEMBOURGMACEDONIAMALAYSIAMEXICOMOROCCONAMIBIANETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANORWAYOMANPAKISTANPALESTINEPANAMAPAPUA NEW GUINEAPERUPHILIPPINESPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIASAUDI ARABIASERBIA&MONTE; NEGROSINGAPORESLOVENIASOUTH AFRICASOUTH KOREASPAINSRI LANKASWEDENSWITZERLANDTAIWANTHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYUGANDAUKRAINEUNITED ARAB EMIRATESUNITED KINGDOMVENEZUELAVIET NAMYEMEN | Terminated | September 05, 2012 | 782 | 41119 | Not Computer
Z-0328-2011 | 56628 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Stryker M1 Ambulance cot, Model 6100, Stryker Medical, Kalamazoo, MI. |  |  | May 15, 2010 | 2010 | Stryker Medical Div. of Stryker Corporation | The firm has received 42 complaints regarding the head end lock rod, head end slide tube, brushing, and upper cross tube failures. | Device Design | Stryker Medical sent an Urgent Medical Device Alert letter dated May 15, 2010 to all customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter stated that there is a possibility for fatigue related failures of certain components of the cot and that a Stryker representative will visit the firm to correct the device. A representative from Stryker field service would contact the customer to perform the device corrections, free of charge.Immediate Action required by the customer:Locate the ambulance cots listed in this notice.Return the enclosed post card to confirm receipt of this notification.If any of the cots listed in the letter were loaned or sold the customer was to forward a copy of the notice to the new users and advise Stryker of their new location.If the customer disposed of any of the affected stretchers and they were no longer in use, customers were to advise Stryker of their obsolescence by providing the serial numbersFor any questions or comments regarding this issue, please call (800) 869-0770, ext. 6902. |  | 7314 | 7,314 | Worldlwide Distribution - US including AZ, CA, CO, FL, LA, MS, NC, NJ, OH, PA, and TN and the countries of Australia, China, France, Germany, Hong Kong, Italy, Japan, New South Wales, Poland, Portugal, Spain, Sweden, Switzerland, and the United Kingdom. | Terminated | February 07, 2012 | 633 | 7314 | Not_Computer
Z-0333-2011 | 55301 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Excel Care ES Bariatric Bed, P 610B, Hill-Rom, Batesville, IN. |  |  | March 05, 2010 | 2010 | Hill-Rom, Inc. | The firm received complaints regarding the siderails not operating correctly or becoming inoperable. The siderails could experience false latching, unintentional lowering or become inoperable. There have been reports of 5 patient falls and 2 injuries to caregivers. | Device Design | Hill-Rom sent out "Urgent Medical Device Correction" letters dated March 5, 2010. The letters described the product, problem and actions to be taken by the customers. The letters instructed the customers to locate the beds and examine the siderails for any visible damage (cracks, bending, loose parts) and conduct a physical assessment of all four siderails on each bed to determine if the sideral could be placed in all positions. If the siderails were found to be inoperable, the bed was to be taken out of service until a qualified Hill-Rom technician can make any necessary correction. The customers were instructed to contact Hill-Rom to make arrangements and inform them of all beds that they have loaned, sold, no longer use or disposed of at 800-445-3720, option 2. The letter states the firm is evaluating a design change to the siderail.If you have any questions concerning this request or the procedure as outlined, please contact Hill-Rom Technical Support at 800-445-3720, option 2. |  | 1406 | 1,406 beds both models. | Worldwide distribution: USA states including: AK, FL, HI, IN, KS, LA, MS, OK, PA, TN, VT, and WA and countries including: Canada, Mexico, Angola, Australia, Saudi Arabia, Qatar, and United Emirates. | Terminated | October 06, 2014 | 1676 | 1406 | Not_Computer
Z-0329-2011 | 55840 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | Titan XL Handpiece Infrared Therapeutic Heating Lamp, Manufactured by Cutera Inc, Brisbane, CA |  |  | May 08, 2010 | 2010 | Cutera Inc | Risk of burns-- under certain conditions, devices containing a gold plated end block could produce energy above what is stated on the device setting. | Device Design | Cutera sent Certified letters sent to all consignees on May 8, 2010. The letter identified the product, the problem, and the action the consignees should take. Consigness were instructed to immediately stop using the handpiece and complete the attached Recall Tracker Form. A second letter was sent to customers who received a replacement handpiecce but did not return the affected devices. All affected products were to be exchanged for new or reworked devices.For questions contact customer service at (888) 428-8372, ext 5561. |  | 1344 | 1,344 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OK, OR, PA, OH, SC, TN, TX, UT, VA, VT, WA, WI, WV, AND WY and the countries of Argentina, Australia, Austria, Barbados, Belgium, Bermuda, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Cypress, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Greece, Guadeloupe, Hong Kong, India, Indonesia, Ireland, Israel, Italy, Japan, Lle de la Reunion, Mexico, New Zealand, Phillipines, Poland, Puerto Rico, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Thailand, Turkey, UAE, and Ukraine. | Terminated | December 28, 2010 | 234 | 1344 | Not_Computer
Z-0330-2011 | 56800 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | External Compression Device; 1 unit per package |  |  | August 30, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | Variations in the outer diameter of the shaft of the External Compression Device and the inner diameter of the Nail Holding Screw may make assembly of these 2 instruments not possible. | Other | Stryker sent Recall notification letters by Federal Express on September 9, 2010 to all Stryker branches/agencies and Risk Managers, Chief of Orthopaedics of affected hospitals and Surgeons on September 10, 2010, that used the product. The letter identified the product, the problem, and the action that customers should take. Customers were instructed to:Examine their inventory and hospital locations to identify the affected product and return it to their branch or agency warehouse for reconciliation.Reconcile all product on the attached Product Recall Acknowledgment Form and fax a copy to (201) 831-6069 within 5 days of receipt of notice.A Stryker representative would contact the customer with details for return of the product.For any questions regarding this recall call (201) 972-2100. |  | 11 | 11 | Nationwide Distribution including AZ, CA, CO, FL, GA, IL, KY, MD, OH, OK, and TX | Terminated | August 21, 2012 | 722 | 33 | Not_Computer
Z-0335-2011 | 57024 | No | LKK |  |  | General Hospital | PMA | Medtronic Intrathecal Catheter Spinal Segment Revision Kit, 8598A. Method of Sterilization: Ethylene Oxide. Manufacturer Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604 USA. The Medtronic Model 8598A Spinal Segment Revision Kit is used when a revision to the spinal segment of a Model 8731 or Model 8731SC catheter is required. The catheter is part of an infusion system that stores and delivers parenteral drugs to the intrathecal space. The implanted infusion system components consist of a Medtronic pump and a Model 8731 or Model 8731SC catheter. The catheter connects to the pump at the catheter port. |  |  | September 29, 2010 | 2010 | Medtronic, Inc. - Neuromodulation | Medtronic is recalling specific lots of Model 8731SC intrathecal catheter kits and Model 8598A intrathecal catheter spinal revision kits because there is the potential that the 11.4 cm introducer needles within these kits may exceed USP requirements for bacterial endotoxin. | Material/Component Contamination | Medtronica sent an "Urgent: Medical Device Recall" letter, dated September 2010, was sent to medical facilities/customers beginning September 29, 2010. The letter described the issue, identified affected product, requested that unused inventory be segregated and returned to Medtronic Neuromodulation per the enclosed Product Return Instructions form. A Recall Reply Card was also asked to be returned. Additionally an "Urgent: Medical Device Recall" letter was addressed to Healthcare Providers/customers to inform them of the recall of the Catheter and Revision Kits, identified patients by means of registration information, Informed them of the issue with the needle inside the kits, and requested they stop using the specific lots and to monitor any patients for symptoms. If you have any questions or comments, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 420 | 420 | AL, CA, CO, CT, FL, GA, HI, ID. IL, IN, KS, KY, MA, MD, MI, MN, MO, MT, NC, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, VA, WA, WI, WV, WY. OUS to include: Argentina, Austria, Belgium, Canada, Chile, Czech Republic, France, Germany, Greece, Hungary, Ireland, Israel, Italy, San Marino, Jordan, Lithuania, Netherlands, Norway, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, United Arab Emirates, United Kingdom. | Terminated | December 10, 2011 | 437 | 1743 | Not_Computer
Z-0337-2011 | 56973 | Yes | NIM |  |  | Cardiovascular | PMA | RX ACCULINK Carotid Stent System 9 X 20mm, Part Number: 1011341-20, Lot Number: 0071361 |  |  | September 10, 2010 | 2010 | Abbott Vascular | The recall was initiated because Abbott vascular has discovered through internal testing that the affected lots may not meet the firm's quality specifications for catheter shaft tensile strength. If the affected device is used it may result in acute stent deployment difficulties and subsequent intervention. | Pending | Abbott Vascular sent an URGENT DEVICE RECALL letter dated September 10, 2010, to all consignees. The letter identified the product, the problem, and the action to be taken. The Sales Representatives were to personally visit each account to deliver the recall letter and assist in identifying and returning anyunused devices to Abbott Vascular. Also, Abbott Vascular Regulatory Compliance initiated contact with the Risk Management department at each account by phone to verify their mailing address and provide a courtesy copy of the recall letter.The consignees were instructed to work with their local account representative to review their inventory, complete the Recall Effectiveness Check and return identified products to Abbott Vascular.For questions regarding this recall call (951) 914-3324. |  | 42 | 42 units total for all products | Nationwide Distribution including AL, CA, CT, DC, FL, GA, IA, IL, IN, MA, ME, MI, MN, MO, OH, OK PA, TX, VA, and WI | Terminated | November 16, 2010 | 67 | 42 | Not_Computer
Z-0343-2011 | 56976 | No | JHX | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Stratus(R) CS Rotors, catalog number CROTOR, lot number NH42-251-09 |  |  | December 07, 2009 | 2009 | Siemens Healthcare Diagnostics, Inc. | Low frequency defect in the gel filling of this lot that may result in inadequate separation of cells in centrifugation. | Process control | The recalling firm, SIEMENS, issued an "Urgent Field Safety Notice" dated December 2009 to all affected customers. The notice described the product, problem and the action to be taken by the customers. The customers were instructed to review the situation with their laboratory director to determine the need to evaluate previously reported test results; immediately examine their CROTOR inventory and discard all bags that are from lot NH42-251-09 including rotors that were not stored in their original packaging (bag); contact Siemens Customer Service Center at 800-241-0420 to request additional no-charge replacements (Note: A "no-charge" order has already been placed so that one bag of replacement rotors would be recieved shortly); forward this notice to anyone they may have distributed this product, and complete and return the FIELD CORRECTION EFFECTIVENESS CHECK via fax at 302-631-8467.If you have any questions regarding this information, please contact the Siemens Technical Solution Center at 800-405-6473. |  | 330 | 330 bags | Worldwide distribution: USA including states of AK, AZ, AR, CA, FL, GA, HI, IA, IL, IN, KY, LA, MI, MN, MO, MS, MT, NC, ND, NH, NM, NY, OH, TX, UT, WA, WI, and WV and country of Canada | Terminated | November 10, 2010 | 338 | 330 | Not_Computer
Z-0344-2011 | 56863 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab/Mic ASCII Add-on Versions 4.0.1.14, 4.0.1.15, and 4.0.1.16. |  |  | February 16, 2010 | 2010 | SCC Soft Computer | When a test is being verified, the wrong (expired) ranges may be associated to that test result, causing an incorrect flag to be posted for that test result. | Software design | Clients were notified via the proprietary communication software on 02/16/2010. The Notification Verbiage identified the affected product and informed clients of the potential hazard, root cause, corrective action, and requested action. Clients were instructed to acknowledge receipt of the notification. If the client wishes to receive the correction, they should document a request to load the software. |  | 31 | 31 | Nationwide Distribution - US and Puerto Rico. | Terminated | July 17, 2012 | 882 | 31 | Software
Z-0341-2011 | 56751 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | The box label states: "ATF 40 Fast Start Kit 9108491...STERILE...FRESENIUS KABI...Fresenius Kabi AG 61345 Bad Hamburg Germany...Federal (USA) Law restricts this device to sale by or on order of a physician". |  |  | March 29, 2010 | 2010 | Fresenius Kabi, LLC | Using AT1 Autotransfusion Set from specific batches in combination with the auto transfusion device CATS Continuous Autotransfusion has led to insufficient suction of shed blood from the collection reservoir. The results in the failure message "alarm blood flow" and the CATS device stops in a secure mode. | Other | Fresenius Kabi sent out a Technical Notification letter dated March 29, 2010, to one distributor located in Ann Arbor, Michigan. The distributor sent the technical notification with a cover letter to each of their customer's that had ordered AT1's, ATF120's and ATF40's over the last year. The letter identified the product, the problem, and the action to be taken. Customers were notified that the autotransfusion prime and / or washing procedure can be continued after alarm blood flow by following Alternative procedures. The Technical Notification informs the distributor about the problem of the AT1 (item #9005101), but also mentions the other two products ATF120 (item #9108501) and ATF40 (item #9108491) which have no known problem.The distributor is advised to contact Fresenius Kabi, LLC for any question regarding the Technical Notification.For questions regarding this recall call (425) 242-2074. |  | 461 | 17, 444 units distributed in US | Worldwide Distribution - USA and the countries of Australia,Austria, Brazil, Canada, China, Denmark, France, Germany, Hong Kong, India, Italy, Japan, Malaysia, Netherlands, Phillipines, Portugal, Switzerland, Thailand, UK, and Vietnam. | Terminated | June 15, 2011 | 443 | 1444 | Not_Computer
Z-0285-2011 | 56760 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | VITEK 2 Gram Negative Susceptibility Card, AST-N107, REF 22 265, For In Vitro Diagnostic Use, 20 cards/box. The firm names on the label are bioMerieux, Inc., Durham, NC, and bioMerieux sa, 69280 Marcy-l'Etoile France. |  |  | September 16, 2010 | 2010 | Biomerieux Inc | AMENDED REASON FOR RECALL: The piperacillin/tazobactam drug on the antimicrobial susceptibility cards is providing false susceptible and false resistant results for E. coli, as well as an additional 7 organisms; and the same drug is providing false resistant results for K. pneumoniae. | Nonconforming Material/Component | bioMerieux issued a "Product Correction Notice" letter dated September 13, 2010, to their entire customer base (regardless of whether they ever ordered the recalled cards). The letters explained the reason for the recall and instructed the account to confirm all results for E. coli using a second method, and for K. pneumoniae, to confirm with a second method only for resistant results. These alternate test confirmations are to continue until further notice. The customers were to complete and return the enclosed Acknowledgement Form via fax at (314) 731-8689 or (919) 620-6350 which indicate that the notice has been read, the instructions provided in the "Actions Required" section were followed, and notes if they have received any reports of illness or injury related to this issue. All future shipments will include an orange multilanguage insert to go into the boxes reminding customers they must perform an alternate test for TZP/E. coli combinations and that customers must also perform an alternate method prior to reporting TZP results when a resistant result is obtained for the TZP/K. pneumoniae. Additionally, this recall letter will accompany shipments of instrument software kits for the new customers.The recalling firm issued a press release on 10/27/10 via the Associated Press notifying the public of the recall.The recalling firm expanded their recall to add 7 more organisms to the list of organism testing that were to be suppressed. As a result the recalling firm issued a letter dated 3/30/11 via Fed Ex on 4/4/11 flagged "URGENT PRODUCT CORRECTION NOTICE (EXPANDED RECALL), VITEK 2 Piperacillin/Tazobactam Test, Document #9300786." The letter describes the issues, how the issues impact the test results, describes the product changes that are going to take place, and provides instructions on required actions to be taken.The product changes will involve the eventual changeover to new cards that will be created which do not contain Piperacillin/Taz |  | 90 | Product not distributed as of 9/10/10 | Worldwide Distribution: USA (all states), and the countries of:SUBSIDIARIES - Argentina, Canada, Switzerland, Chile, Colombia, Germany, Spain, Great Britain, Hong Kong, Japan, Korea, Mexico, Poland, Brazil, Sweden, Singapore, and Thailand. DISTRIBUTORS - Dutch Antilles, Australia, Bermuda, Uruguay, Costa Rica, Dominican Republic, Ecuador, France, Guatemala, Honduras, Indonesia, Israel, Italy, Singapore, Malaysia, Nicaragua, The Netherlands, Panama, Peru, Philipines, Paraguay, El Salvador, Venice, and Viet Nam. | Terminated | May 08, 2012 | 600 | 5709147 | Computer
Z-0349-2011 | 51883 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | ViewSonic 19" flat panel LCD display Model VG930M which is used as part of the Welch Allyn Acuity Central Monitoring System. The affected serial numbers are those less than QC2082880001. |  |  | May 01, 2009 | 2009 | Welch Allyn Protocol, Inc | The 19" Model VG930M ViewSonic flat panel LCD display may have faulty capacitors at certain locations in the power supply section. The faulty capacitors may cause a loss of power to the display resulting in the display screen going dark and loss of both visual and secondary patient alarms. | Other | On May 1, 2009, Welch Allyn Protocol, Inc. sent out recall notification to their consignees in via UPS and US Postal Certified Mail.The letter describes the recalled product as 19" model VGM930M ViewSonic flat panel LCD display with serial numbers less than QC2082880001. For the Welch Allyn service contract consignees/customers, the firm's Field Service Engineer will replace the LCD monitor for the consignees. The consignees could contact Technical Support at 1-800-289-2501.For the Warranty/non-Warranty contract consignees/customers-For the Welch Allyn service contract consignees, Stericycle Return Management Services (RMS) will manage all recall and recovery initiatives. Consignees/customers can call Welch Allyn Customer Care at 1-800-535-6663 for any questions or concern about this recall. |  | 637 | 637 units | Worldwide distribution: USA including states of: AK, AZ, CA, CO, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MD, MI, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SD, TX, UT, VT, WI, and WV; and countries including: AUSTRALIA, CANADA, ESTONIA, GERMANY, ICELAND, IRELAND, ITALY, JORDAN, MALAYSIA, POLAND, and RUSSIA. | Terminated | July 25, 2011 | 815 | 637 | Hardware | Display/Image | Replace LCD | Remove or Replace
Z-0367-2011 | 55839 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRAD Sterile Disposable Syringe Kit Recall, Catalog Number 150-FT-Q; Kit includes angiographic injector and syringe |  |  | May 17, 2010 | 2010 | Medrad Inc | Some kits have a hole in the packaging which could compromise the sterility of the components. | Packaging | MEDRAD issued an URGENT MEDICAL DEVICE RECALL letter dated May 17, 2010, to all affected customers. The letter identified the product, the problem, and the action the customer should take.Customers were instructed to identify the product with the affected lot numbers.Contact MEDRAD Customer Support at 1-800-633-7231 to make arrangements for the return of the product.Complete and sign the form attached to the letter and fax it back to MEDRAD at 412-406-0942.For any questions regarding this recall call (412) 767-2400, ext 6851. |  | 39448 | 39,448 syringe kits | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO MS, NC, ND, NE, NJ, NM, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, and WV and the countries of Belgium, China, Columbia, Germany, Denmark, France, Great Britain, Hungary, Ireland, India, Israel, Italy, Mexico, Netherlands, Sweden, and Singapore. | Terminated | November 17, 2010 | 184 | 39448 | Not_Computer
Z-0357-2011 | 53048 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific brand Neuro Renegade Hi-Flo Microcatheter, 150cm/20cm; Rx Only; |  |  | August 24, 2009 | 2009 | Boston Scientific Corporation | Sterility Loss-- The sterile barrier pouch seal (top seal) may be breached by the product, resulting in loss of sterility which may result in transmission of disease or infection. | Packaging process control | Boston Scientific issued an Urgent Medical Device Recall letter dated August 26, 2009 identifying the affected devices, the issue prompting the recall, and actions to be taken by customers.Customers were instructed to determine if they have affected product in their inventory, and if so, immediately quarantine and return to Boston Scientific.Customers can contact Boston Scientific at 510 624-1734. |  |  |  | Distribution Worldwide, including throughout the US. | Terminated | February 03, 2011 | 528 |  | Not_Computer
Z-0355-2011 | 56792 | device, vein stabilization | LBJ | Vein stabilizer. | General Hospital | General Hospital | 510(K) Exempt | Baxter INTERLINK System, Non-DEHP Catheter Extension Set Kit; 7.1" (18.0 cm), Vol. 0.85 mL, 1 Catheter Extension Set with injection site and male luer lock adapter with retractable collar, 1 VITAL-HOLD Medium Catheter Stabilization Device, 2 foam tape strips; an Rx sterile convenience kit for IV infusions; 50 units per case; Manufactured by an affiliate of Baxter Healthcare Corp., Deerfield, IL 60015 U.S.A., Made in Costa Rica; catalog 2N8221 |  |  | September 17, 2010 | 2010 | Baxter Healthcare Corp. | The removable paper liner covering the adhesive strip that is used to secure the catheter stabilization device to the patient is pyrogenic. The device is labeled as nonpyrogenic. | Component design/selection | Baxter issued Urgent Product Recall letters dated September 17, 2010 to the affected direct accounts. The accounts were informed that the paper lining covering the adhesive strip that is used to secure the catheter stabilization device to the patient was found to be pyrogenic. The accounts were requested to discontinue use and segregate the affected lots of products from inventory and call Baxter Healthcare Center for Service at 1-888-229-0001, Monday through Friday, 7:00 AM to 6:00 PM CT, to arrange for the return of the affected product for credit. The accounts were also requested to complete and fax back to Baxter the enclosed reply sheet confirming receipt of the letter and the amount of each product being returned. If the accounts further distributed the products, they were requested to forward the communication to those accounts. Clinical questions regarding the communication are directed to the Medical Information Services at Baxter at 1-800-933-0303. Any other questions may be directed to the Center for One Baxter at 1-800-422-9837. |  | 12500 | 12,500 units | Nationwide Distribution: Throughout the US. | Terminated | October 04, 2011 | 382 | 172300 | Not_Computer
Z-0368-2011 | 56814 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE LightSpeed VCT Scanner System, models 5212920-300, 5212920-310, 5311595-10. Composed of a gantry, patient table, operator console, computer, and PDU and includes image acquisition hardware, image acquisition and reconstruction software, associated accessories and connections/interfaces to accessories. GE Healthcare, 3000 North Grandview, Waukesha, WI 53188 USA. |  |  | September 03, 2010 | 2010 | GE Healthcare, LLC | GE has become aware of a potential set of circumstances that could cause x-ray continuation during an unexpected table stop on certain LightSpeed VCT and BrightSpeed systems. | Software design | An Urgent Medical Device Correction letter dated September 2, 2010 was sent to consignees beginning September 3, 2010. This letter addressed the LightSpeed VCT systems. Another Urgent Medical Device Correction letter, dated September 24, 2010 was sent September 24, 2010 for the BrightSpeed Systems. The letters described the issue and provided safety instructions for the health care provider to be aware if the table should stop unexpectedly and the x-ray remains on to manually stop the scan until the devices software is upgraded.If you have any questions, please contact the US (800) 437-1171, Japan 0120-055-919, Korea 1544-6119, Australia/New Zealand 800 659 465, China 800-810-8188; for other countries, please contact your local GE Healthcare Services Representative. |  | 68 | 68 | Worldwide distribution: USA including states of: MI, NM, OH, TX, and Puerto Rico and countries including: Algeria, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Germany, Greece, Guatemala, Honduras, Hungary, India, Indonesia, Ireland, Islamic Republic of Iran, Israel, Italy, Jamaica, Japan, Kenya, Lebanon, Malaysia, Mauritania, Mexico, Morocco, Netherlands, New Caledonia, New Zealand, Nigeria, Norway, Pakistan, Paraguay, Philippines, Poland, Portugal, Republic of Korea, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Togo, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Viet Nam. | Terminated | May 15, 2012 | 620 | 1014 | Computer
Z-0364-2011 | 56865 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SA BASE with SoftLab Version 3.1.4.x. |  |  | June 01, 2010 | 2010 | SCC Soft Computer | Functionality issue: For clients using the instrument to generate results on the interface menu, two different result records could be displayed under the same order and sequence numbers. This could cause incorrect results to be populated to a patient record. | Software design | Clients were notified via the proprietary communication software on 06/01/2010. The communication discussed the risk-to-health notification task, potential hazard, root cause, corrective action, and requested action. Clients were instructed to acknowledge receipt of the notification. If the client wishes to receive the correction, they should document a request to load the software. |  | 40 | 40 | Worldwide Distribution -- US, including states of CT, FL, IL, KY, GA, LA, IN, MS, NJ, NY, PA, OH, and TN and country of Canada. | Terminated | July 16, 2012 | 776 | 40 | Software | Display/Image | Software update | Software Update
Z-0363-2011 | 56461 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Transport ProÂż Monitor with the CARESCAPE" Patient Data Module. |  |  | October 05, 2010 | 2010 | GE Healthcare, LLC | Transport Pro Monitor stops communication with the CARESCAPE Patient Data Module (PDM) after 414 days of continuous run time. This time will equate to different amounts of real time depending on how much the units is actually in service per day. Transport Pro contains an internal timer that is used to control the software and remind users to perform preventive maintenance. When this internal timer | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated September 17, 2010 to consignees. The letter was addressed to Hospital Administrator, Head of Biomedical Engineering and Nursing Manager. The letter described the product, Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Service representatives will updates all of the affected Transport Pros with PDM that were distributed.Customers may contact GE at 262-548-2731 about this correction. |  | 3256 | 3256 | Worldwide Distribution: Domestic in the states of AL, AZ, AR, CA,CT, CO, DC, FL, GA, ID, IL, IN, IA, KY, LA, MD, MA, ME, MI, MN, MS, MO, NE, NV, NJ, NM, NY, NC, OH, OK, PA, SC, TN, TX, VA, VT, WA, WI, and WY, and Internation in the countries of BRAZIL, BERMUDA, BELGIUM, AUSTRIA, AUSTRALIA, CANADA, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY GREECE, INDONESIA, IRELAND, ITALY, JAPAN, JORDAN, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MAURITIUS, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, REPUBLIC OF KOREA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD 7 TOBAGO, and the UNITED KINGDOM. | Terminated | May 17, 2012 | 590 | 3256 | Software | Display/Image | Software update | Software Update
Z-0360-2011 | 56873 | No | LNX |  |  | General Hospital | 510(k) | SoftID.NET versions 2.0.0.6 - 2.0.0.19. |  |  | August 28, 2009 | 2009 | SCC Soft Computer | SCC Soft Computer issued a Correction for SoftID.NET versions 2.0.0.6 - 2.0.0.19. When using the Cancel functionality, if the collection list refreshes during the transaction and changes the position of the patient on the list, the wrong patient order can be cancelled. All affected clients have been notified. | Software design | SCC Soft Computer notified Customers on August 28, 2009, via the proprietary communication software entitled "Risk-to-Health Notification Task Verbiage" dated August 17, 2009. The notification described the product, problem and action to be taken by customer. The customers were instructed to acknowledge receipt of the notification; document a request to load the software, if they wish to receive the correction; document that the Alternative Working Solution will be used by the customers facility until the correction has been taken live.If you have any questions, contact (727) 789-0100. |  | 41 | 41 | Nationwide distribution: AR, CA, FL, IL, IN, KS,LA, MA, MI, MN, MO, NC, ND,NH, NJ, NY, OH, PA, RI, TN, TX, VA and WI. | Terminated | November 16, 2011 | 810 | 41 | Software | Output/Calculation | Upgrade | Repair
Z-0359-2011 | 56664 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Power supply boards on Dash 3000, 4000 and 5000 Monitors. |  |  | October 11, 2010 | 2010 | GE Healthcare, LLC | The power supply board in GE Healthcare Dash 3000, 4000 and 5000 patient monitors may fail due to insufficient solder of a component. Insufficient solder on the power supply board could result in excess EMC emissions, smoking of components and eventual failure of the power supply, tripping of the fuse and immediate shutdown resulting in a loss of monitoring that could delay recognition of a clinic | Nonconforming Material/Component | GE Healthcare issued an "Urgent Medical Device Correction" letter dated October 8, 2010, which described the loss of monitoring issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. Customers were instructed to remove the units from use. Local service representatives will contact firms to schedule installation of replacement power supply boards.GE Healthcare can be contacted at 1-800-558-7044, Option 2, Technical Support, about this correction. |  | 137 | 137 | Worldwide Distribution: USA in the states of IL, NV, TX, and VA, and the countries of Australia, Brazil, China, Japan, South Korea, and Taiwan. | Terminated | May 29, 2012 | 596 | 137 | Battery | Device Operation | Replace boards | Remove or Replace
Z-0361-2011 | 56940 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ENDOPATHÂż ETS-FLEX45 Articulating Endoscopic Linear Cutters, Model # ATS45. Manfuactured by Ethicon Endo-Survery, LLC; 475 Calle C; Guaynabo PR 00969. |  |  | January 13, 2009 | 2009 | Ethicon Endo-Surgery Inc | The mechanism that connects the articulation joint to the device shaft may cause the jaws of the cutters to remain closed and clamped down on tissue after the device is fired. If the jaws of the device remain clamped down on the tissue, there is a risk the surgery will need to be modified to remove the device. This may lead to an increase in procedure time, tissue manipulation, and a change in t | Other | Ethicon Endo-Surgery issued "Urgent Device Recall Event 1952" notification letters dated January 21, 2009 to their customers. They instructed their custormers to immediately discontinue use of the product. They are to complete and fax back the Business Reply Form. Affected product is to be returned with a copy of the completed Business Reply Form to Stericycle. Replacement product should be received within 30-60 days after return.Customers can contact Stericycle concerning this recall at 1-866-861-6921, reference Event 1952. |  | 4113 | 4,113 UNITS | Worldwide Distribution: USA only, in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI. IA. ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | May 10, 2012 | 1213 | 4113 | Not_Computer
Z-0362-2011 | 56872 | No | MMH | Automated blood grouping and antibody test system. | Hematology | Hematology | 510(k) | SoftBank Versions 23.1.2.x, 23.2.0.x, and 25.0.0.x with SoftReports. |  |  | June 22, 2010 | 2010 | SCC Soft Computer | SCC Soft Computer is issuing a Correction for SoftBank II Versions: 23.1.2.x, 23.2.0.x, and 25.0.0.x software applications. Functionality issue - For clients generating patient history reports using Patient&gt;Reports&gt;Batch Reports&gt;History Report option in SoftBank with SoftReports activated for patient reports, when the report is generated for more than one patient ALL of the results for the qualif | Software design | SCC Soft Computer notified Customers on June 22, 2010, via the proprietary communication software entitled "Risk-to-Health Notification Task Verbiage" dated June 11, 2010. The notification described the product, problem and action to be taken by customer. The customers were instructed to acknowledge receipt of the notification; to grant permission to load the required hot fix, and to make the manual correction. For the manual correction to the datapath for patient report layouts, your SoftBank Support Specialist will Coordinate those changes and your validation of affected environments.If you have any questions, contact (727) 789-0100. |  | 30 | 30 | Worldwide distribution: U.S. and country of Canada. | Terminated | September 05, 2012 | 806 | 30 | Software | Output/Calculation | Software update | Software Update
Z-0381-2011 | 57052 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sorin OptiPack, Perfusion Tubing Set, Sterile EO, REF 084500900, Manufactured by Sorin Group USA, Inc. 14401 W. 65th Way Arvada, CO 80004 |  |  | October 07, 2010 | 2010 | Sorin Group USA, Inc. | Defective product packaging may compromise product sterility. | Packaging process control | Sorin Group sent an "URGENT: FIELD SAFETY NOTICE" letter dated October 7, 2010, to all customers. The letter described the product, problem and action to be taken by customer. The customers were instructed to check all affected inventory to exam for holes. Any damaged products were to be returned to Sorin. If no damage was found, product could be used. The customers were also instructed to contact Sorin Cardiopulmonary Customer Service at 1-800-650-2623 to return product and receive replacement; they were to share this information with all personnel who may use this product; and to complete and return the attached Customer Response Form via the self-addressed prepaid envelope provided or fax to (303) 467-6502 as soon as possible.For further information, contact 1-800-650-2623. |  | 86 | 86 | Nationwide distribution: CA, FL, IL, IN, KS, MD, MO, NE, NJ, NM, TN, TX, and VA. | Terminated | December 17, 2010 | 71 | 3021 | Not_Computer
Z-0366-2011 | 57001 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | BD Phoenix" ID/AST Panels, Item # PMIC/ID-107, Catalog number 448607, and labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, MD 21152 USA (800) 638-8663*** |  |  | May 14, 2010 | 2010 | Becton Dickinson &amp; Co. | The diagnostic reagent packaging may not be completely sealed, which may cause substrate degradation, and produce incorrect test results, including false positive and false negative results. | Packaging process control | BD Diagnostic Systems issued an "Urgent Product Recall" letter dated May 2010 to direct accounts, identifying the affected product, potential impact, and actions to be taken by the customer. Accounts were requested to respond by returning the enclosed reply form indicating the number of replacements needed for the affected product to be discarded.For assistance iwth replacements, contact BD Customer Service at 800 675-0908. For other inquiries, contact BD Technical Services at 1 800-638-8663. |  | 1675 | 1675 units | Nationwide Distribution: To direct accounts in CA, MD, NJ, PA. and WI. | Terminated | November 29, 2010 | 199 | 1675 | Not_Computer
Z-0395-2011 | 55999 | No | IXQ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | CIVCO Type-S Extension, REF MTIL6605, Reusable non-sterile extension for use with Universal Couchtop. The firm name on the label is CIVCO, Orange City, IA. |  |  | June 21, 2010 | 2010 | Med Tec Inc | Under certain conditions, the extensions can become detached from the main body of the couch. | Device Design | The recalling firm, CIVCO, issued an 'Urgent Field Safety Notice" dated June 15, 2010, via overnight mail on June 21, 2010, regarding the first issue for the three extension models. The letter described the product, problem, and actions to be taken by the customers. The customers were instructed to load the patients on the couch from the bottom of the couch and not from the superior end with the extension and do not let patients crawl to the top of the couch with their full upper body weight at the end of the couch; notify others within their organization that may have the extension, and complete and return the enclosed response form via fax: Attention: CIVCO MTL6600 Rectangular Extension at 877-271-2310. The letter informs the customer that they will be receiving a kit to replace the locking tab so that the extension will be less likely to come loose. Once the kit is received, the customers are ask to return the replaced part and sign a confirmation form for their records.CIVCO sent a second "Urgent Field Safety Notice" dated July 20, 2010 via overnight mail on July 21, 2010, and by phone on July 22, 2010, regarding the second issue which covered only MTIL6600 regarding the weight being dropped onto the corner of the rectangular extension. The letter described the product, problem and action to be taken by the customers. The customers were instructed to accept a one-piece Universal Couch Top replacement; sign the attached form stating the they will stop using the rectangular extension and return it to CIVCO or sign the form stating they have been advised of the recall and they will exercise the necessary precautions to prevent the event from occurring by always using the thumb screw lock and never allowing the patient to crawl on the couch top or extension, and to notify others within their organization that have the extension. The customer was informed that their current system would be replaced and sent back to the recalling firm. Please contac |  | 3 | 3 extensions | Worldwide distribution: USA including states of: CA, FL, IA OH, and WI; and countires of: United Kingdom, Australia, and Switzerland. | Terminated | April 17, 2012 | 666 | 30 | Not_Computer
Z-0365-2011 | 56867 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab GUI Version 4.0.1.0 - 4.0.1.16, 4.0.2.0 - 4.0.2.10, 4.0.3.0 - 4.0.3.13, and 4.0.4.0 - 4.0.4.5. |  |  | April 01, 2009 | 2009 | SCC Soft Computer | Clients using Patient Maintenance to perform Moves of Stays may send the wrong billing number during an ADT update to SoftWeb, SoftPath or SoftMedia. The wrong billing number can come from a different patient in Order Entry. Only Moves of Stays are affected. | Software design | Clients were notified via the proprietary communication software on 06/01/2009. This is a MANDATORY CORRECTION. The Risk-To-Health Notification Task gives the background on the issue, the potential hazard, root cause, corrective action, alternative working solution, and requested action. Clients were instructed to acknowledge receipt of the notification. If the client wishes to receive the correction, they should document a request to load the software. |  | 49 | 49 | Nationwide Distribution | Terminated | August 07, 2013 | 1589 | 49 | Other | Output/Calculation | Software update | Software Update
Z-0392-2011 | 56988 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 4100/4100IQ GE Innova 4100 / 4100 IQ. , GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | August 06, 2009 | 2009 | GE Healthcare, LLC | GE Healthcare has become aware of a potential video splitter failure with certain Innova 2000 and 4100 X-ray systems that may impact patient safety.A potential video splitter failure inside the GE equipment cabinet could occur with no advance warning and thereby prevent x-ray image display. Should this occur, the systemcannot be recovered without service intervention.GE Healthcare has obse | Nonconforming Material/Component | GE Healthcare sent an " Urgent Medical Device Correction" letter dated July 9, 2009, to Hospital Administrators / Risk mangers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter identified the product, described the Safety Issue, Affected Product Details, Safety Instructions Product correction and Contact Information.Hospitals were instructed to establish emergency workflow procedures should a component stop operating during a patient examination. A GE Healthcare service representative would contact each customer and correct all affected systems.For questions regarding this recall call (262) 548-2731. |  | 43 | 43 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MO, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX UT, VA, WA, WV, and WI, and the countries of UNITED KINGDOM, TUNISIA, TAIWAN, PROVINCE OF CHINA, SPAIN, SOUTH AFRICA, SAUDI ARABIA, RUSSIAN D=FEDERATION, POLAND, NORWAY, NEW ZEALAND, NETHERLANDS, MEXICO, MARTINIQUE, LITHUANIA, LEBANON, REPUBLIC OF KOREA, JORDAN, JAPAN, ITALY, IRELAND, ISLAMIC IRAN, INDONESIA, INDIA, GREECE, GERMANY, FRANCE, FINLAND , EGYPT, CHILE, CANADA, BULGARIA, BELGIUM BELARUS, AUSTRALIA. | Terminated | May 17, 2012 | 1015 | 408 | Not_Computer
Z-0403-2011 | 56876 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Vortex TR, AngioDynamics, Inc., One Horizon Way, Manchester, GA 31816 |  |  | September 27, 2010 | 2010 | Angiodynamics, INC | Labeling Error; The labeling for the specific products incorrectly identified that the packaged device kit contains a Lifeguard Safety Infusion Set. | Other | AngioDynamics issued a Notification of Medical Device Correction letter dated September 27, 2010 to customers. The notification served to make the customer aware of the mislabeling. Customers were instructed not to return the affected products.Customers can contact AngioDynamics at 1 800 772-6446 about this correction. |  |  | see line one. | Nationwide Distribution : Throughout the US. | Terminated | June 21, 2011 | 267 |  | Not_Computer
Z-0410-2011 | 56967 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Innova 4100/4100IQ GE Innova 4100 / 4100 IQ. , GE Healthcare, Waukesha, WI 53188 ( The system consists of an a monoplane positioner, a vascular or cardiac table, an X-ray system and a digital detector) |  |  | August 05, 2009 | 2009 | GE Healthcare, LLC | GE Healthcare has become aware of a potential issue regarding Innova systems with GE Fluoro UPS 20KVA that may impact patient safety.GE Fluoro UPS 20KVA (Uninterrupted Power Supply) is used to maintain fluoroscopy availability for approximately 5 minutes should there be a main power failure. Once power is restored, the equipment power will be returned to the main hospital power.UPS fluoro ba | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter on August 5, 2009 to all Hospital Administrators / Risk Mangers, Mangers of Radiology / Cardiology and Radiologists / Cardiologists. The letter was dated August 4, 2009, and described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information.For questions regarding this recall call (262) 548-2731. |  | 117 | 117 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, DC, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK ,OR, PA, SC, TN, TX, VT, VA, WA, WV, and WI and the countries of ANGOLA, BELARUS, BULGARIA, CANADA, EGYPT, FRANCE, GERMANY, HUNGARY, IRELAND, MEXICO, ITALY, NIGERIA, NORWAY, POLAND, PORTUGAL, RUSSIAN FEDERATION, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, TUNISIA, TURKEY, and the UNITED KINGDOM | Terminated | January 16, 2014 | 1625 | 578 | Not_Computer
Z-0401-2011 | 57015 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Millar Disposable Angiographic Catheter (Left Heart), manufactured for Millar Instruments, Inc 6001-A Gulf Freeway, Houston, Texas 77023 |  |  | October 14, 2010 | 2010 | Millar Instruments, Inc | The sterile packaging was found to be compromised when tested after shipping conditions. This may result in infection or transmission of disease. A test for transportation of sterile package was conducted and resulted in several samples failing. | Packaging | Millar issued a Recall Report letter on October 14, 2010 to consignees identifying the affected product, sterility issue, and actions to be taken. Customers are to return all unused catheters to Millar for replacement. Customers can contact Millar about this recall at 832-667-7131. |  | 12 | 12 Units | Worldwide Distribution: USA, including the states of MA, OH, PA, and D.C., and the countries of Canada and Germany. | Terminated | February 15, 2011 | 124 | 12 | Not_Computer
Z-0400-2011 | 57005 | No | JBP | Activated whole blood clotting time tests. | Hematology | Hematology | 510(k) | i-STAT ACT Kaolin Cartridges |  |  | October 06, 2010 | 2010 | Abbott Point Of Care Inc. | Certain lots of i-STAT Celite ACT and Kaolin ACT cartridges may be difficult to fill or will not fill when attempting to transfer blood to the sample entry port of the cartridge. | Other | Abbott Point of Care, Inc. sent an URGENT RECALL NOTICE dated September 2010 on October 6, 2010 by Federal Express to all US customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of cartridges from the identified lots. If customers did not currently have replacement cartridges then they should continue to use the cartridge lot until their facility receives replacement material.Upon receipt of replacement material, customers were to discontinue use of and return the affected cartridges and switch to the replacement material. All unused cartridges will be credited and will be returned.If you have any questions regarding this information, please contact Abbott Point of Care Technical Support at 800-366-8020, Option 1 or your Abbott Point of Care representative. |  | 95325 | 95,325 cartridges | Worldwide Distribution | Terminated | August 22, 2012 | 686 | 144350 | Not_Computer
Z-0398-2011 | 57029 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | SKYLight Gamma Camera Gearbox brake, a component of the SKYLight Gamma Cameras 3/8" SPECT, Model number 4535 600 66661; and 5/8" SPECT, Model number 4535 602 20381, manufactured by Philips Medical Systems (Cleveland), Cleveland OH. |  |  | October 15, 2010 | 2010 | Philips Medical Systems | Detector arm assembly may slide to hardware limit, potentially causing impact. | Component change control | Firm will notify consignees via Urgent - Medical Device Correction, dated Nov 3, 2010, via FedEx. The letter identified the affected product, described the problem, the hazard involved, how to identify affected products, actions to be taken, and actions planned by Philips. Customers are to contact their Field Safety Engineer immediately to schedule an inspection of their system. If customers suspect that the identified problem has occurred with their device, they should immediately stop using the system until Philips has confirmed it as safe to use. If customers need further information concerning this issue, they are to contact their local Philips representative. |  | 24 | 24 units | Worldwide Distribution -- USA and the Netherlands. | Terminated | March 21, 2012 | 523 | 24 | Not_Computer
Z-0412-2011 | 52108 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips EasyVision MM workstation with software R10.2 and R11.1 |  |  | January 30, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | Review by firm of Philips EasyVision MM workstations found that to assure compliance with 21 CFR 900.12(c)(5) corrected software was needed so that all sites are in compliance. | Other | Philips Medical Systems North America Co. sent an URGENT-Field Safety Notice dated January 30, 2009, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were advised to ascertain the quality of a hard or soft copy explicity before further distribution to make sure that no part of the tissue is covered by image identification information. Philips will issue the Field Change Order (FCO) 83000129 and release a level on EasyVision MM R10.2 L6 and EasyVision MM R11.1 L3 to resolve the issue.Customers could contact Philips Healthcare Call Center at 1-800-722-9377, #2, #3 and Reference FCO 83000129 for any questions. |  | 11 | 11 units in US | Nationwide Distribution including DC, FL, MI, MO, NM, NY, OH, AND TX | Terminated | November 19, 2010 | 658 | 11 | Software |  | Software update | Software Update
Z-0415-2011 | 56978 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Rascal LiteWay 214 manufactured by Giant Manufacturing, Taiwan. |  |  | September 27, 2010 | 2010 | Electric Mobility Corp | Scooter seat bases were found to have insufficient welding which may lead to the seat base becoming loose and causing the seat to wobble. This may result in complete failure of the seat base, causing the user to fall off the vehicle. | Device Design | The Rascal Company (Electric Mobility Corporation) sent Urgent Medical Device Recall Notice Scooter Seat Base letters dated October 7, 2010 and November 11, 2010 letter to their affected retail customers and affected dealers. The letter instructs the reader that the seat base may become loose, causing the seat to wobble and if not replaced the wobble could lead to a complete failure of the seat base and cause an individual to fall off their vehicle and become injured. The letters also instruct the reader to replace seat bases immediately. Respone forms are attached. Seats with new seat bases will be shipped with return shipping labels to return the original base to the factory. Service representatives will replace the seat base for customers.Customers can contact the Rascal Company about this recall at 1-800-662-4548. |  | 172 | 172 units | Worldwide Distribution: Throughout the US and Canada | Terminated | April 03, 2014 | 1284 | 653 | Not_Computer
Z-0346-2011 | 56874 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | ATM100, LB-AP-M0001 |  |  | March 15, 2010 | 2010 | Micromedics, Inc. | Micromedics, Inc is conducting a product recall on a number of medical device products because of weak seals of the sterile pouches potentially resulting in a non-sterile product which may cause transmission of disease or infection. | Packaging process control | Micromedics issued a Product Recall letter dated March 15, 2010 to customers, describing the product being recalled, the product sterility issue, and actions to be taken. The letter advised customers to contact Customer service to return any affected product to Micromedics for replacement.Customers can contact Micromedics at 1-800-624-5662 or 651-452-1977 ext. 226 concerning this recall. |  | 620 | 620 (800 pouches) | Nationwide Distribution in the states of CA, CT, DE, FL, GA, ID, IL, IN, MD, MA, MI, MN, NJ, NY, NC, PA, TX, UT, and WI.FRANCE, NETHERLANDS, BELGIUM, UNITED KINGDOM, CANADA, ISRAEL, PERU, SPAIN, IRELAND. | Terminated | June 21, 2011 | 463 | 4880 | Not_Computer
Z-0419-2011 | 57041 | No | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew FAST FIX 360, STRAIGHT NEEDLE |  |  | October 14, 2010 | 2010 | Smith &amp; Nephew, Inc. Endoscopy Division | FasT-Fix 360 Needle Delivery System may not allow the T-2 (anchor) to advance for implantation. | Device Design | Smith &amp; Nephew notified customers by an Urgent - Product Recall 1st Notification - Urgent letter sent via Federal Express October 14, 2010. The letter identified the affected product and explained the problem. The letter asked customers to call for a Return Authorization (RA) Number. Customers are to complete the form and return of the affected product along with a copy of the letter to the address provided being sure to reference the RA number. Questions were to be directed to Cindy Burns at 508-261-3655. Consignees out side the US were notified by e-mail. |  | 53 | 53 units | Worldwide Distribution - USA, Dubai, and Great Britain. | Terminated | May 09, 2012 | 573 | 804 | Not_Computer
Z-0418-2011 | 57094 | No | LTI |  |  | Gastroenterology/Urology | PMA | LAP-BAND Adjustable Gastric Band System; Product codes: |  |  | September 16, 2010 | 2010 | Allergan | In a LAP-BAND System adjustment, failure to inject the needle perpendicular to the access port septum, as described in Directions for Use, may lead to port damage, resulting in subsequent port leakage. Leakage results in deflation of the LAP-BAND to its widest, open position requiring a procedure to replace the port. The patient may have symptoms of reduced satiety and increased appetite. Stalling | Pending | Allergan sent an URGENT FIELD CORRECTIVE ACTION letter dated September 15, 2010, to all affected customers. The recall communication was initiated on 09/16/2010 when the firm begin forwarding Customer letters with attached current Directions for Use distributed via Federal Express, with tracking verification provided by Federal Express, to the direct consignees.The Customer letters informed consignees of the products affected, description of the Issue, clinical implications and firm recommendations.Customers were instructed to follow the Directions for Use for inserting the needle. Customers were instructed to report any malfunction or adverse event related to the device to Allergan Product Support at 1-800-624-4261 ext. 5972 or 1-805-961-5972 and FDA's Medwatch Program (1-800-FDA-1088 or www.fda.gov/medwatch). For questions, consignees were instructed to contact their Allergan field representative or contact Allergan Product Support at 1-800-624-4261. |  | 152000 | Approx. 152,000 | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DC, FL, GA, ID, IL, IN, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, AND WY | Terminated | December 22, 2010 | 97 | 152000 | Not_Computer
Z-0414-2011 | 57039 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliVue Clinical Information Portfolio (ICIP), Clinical Management System |  |  | October 19, 2010 | 2010 | Philips Healthcare Inc. | Software: Scheduling function of the Philips Intellivue Clinical Information Portfolio Releases D.00 through D.03. Edits to a scheduled even frequency order may result in pending interventions not being generated after the first 24 hours. | Software design | Philips Healthcare issued an Urgent Medical Device Correction letter dated October 15, 2010 to customers. Customers are asked to follow the guidelines in the "Action to be taken by Customer/User" section of the Urgent Medical Device Correction Notice until they receive their software correction. The correction will be in the form of a software upgrade and will be performed by Philips service personnel. Contact Philips Customer Care Center with questions or concerns about this correction at 1-800-722-9377 |  | 200 | 200 units | Worldwide Distribution: USA, including states of CA, DC, FL, IL, IN, KY, MA, MD, ME, NJ, NY, OK, RI, TX, VA, VT, and WV; and the countries of: AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHINA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, HONG KONG, IRELAND, ITALY, KENYA, LATVIA, MALAYSIA, NETHERLANDS, NORWAY, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, UNITED ARAB EMIRATES, and the UNITED KINGDOM. | Terminated | August 01, 2017 | 2478 | 200 | Software | Device Operation | Software update | Software Update
Z-0411-2011 | 57047 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | smith &amp; nephew TRIGEN SURESHOT Distal Targeting System, Trauma Interface, V2.0.2, Trauma Launcher V1.01.1 |  |  | October 15, 2010 | 2010 | Smith &amp; Nephew Inc | Targeting software has been found to provide erroneous targeting when used with the 13 mm TAN/FAN Left nails. | Other | Smith &amp; Nephew Inc. sent an URGENT MEDICAL DEVICE CORRECTION letter by FedEx on October 15, 2010, to all customers. The letter was dated October 4, 2010, and identified the product, the problem, and the action to be taken by the customers. Customers were instructed to perform a software upgrade and complete the form provided and return by fax to (901) 566-7835. For questions regarding this recall call (901) 399-5441. |  | 445 | 445 units | Worldwide Distribution - USA including AL, AR, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, MS, NC, NJ, NM, OH, PA, TN, TX, AND WA and the countries of Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Italy, Japan, Netherlands, Norway, South Africa, Spain, Sweden, Switzerland, Turkey, and the UK. | Terminated | February 17, 2011 | 125 | 445 | Software | Device Operation | Software update | Software Update
Z-0423-2011 | 57010 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BARDÂż VentrioÂż Large Oval Hernia Patch 5.4" x 7.0" |  |  | October 15, 2010 | 2010 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: Hernia Patch may be a different size than on the product label | Packaging process control | Davol notified accounts by an Urgent: Medical Device Recall letter on 10/15/10 via overnight delivery. The letter stated that the product maybe a different size than indicated on the product label. It also identified the affected product. Accounts are requested to examine their inventory, remove, and return recalled product. Customers are to notify any of their customers that they may have further distributed the affected product to by including a copy of the recall notice. Customers should contact Davol Customer Service 1-800-556-6275 for instructions on how to return the affected product. Customers are also requested to complete and fax back the attached fax sheet to 1-401-825-8753. If there are any additional questions, customers should contact Davol Customer Service Department at 1-800-556-6275 or C.R. Bard's Medical Service Support Department at 1-800-562-0027. |  | 184 | 184 units | Nationwide Distribution | Terminated | October 27, 2011 | 377 | 406 | Not_Computer
Z-0426-2011 | 56380 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Signa Ovation, 0.35T MR Systems (MFO 1-4). |  |  | August 23, 2010 | 2010 | GE Healthcare, LLC | If the lateral table adjustment crank handle breaks off, it has sufficient attractive force to the magnet and may result in patient injury. | Nonconforming Material/Component | Consignees were sent on 8/23/10 a GE "Urgent Medical Device Correction" letter dated August 20, 2010. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Managers and Radiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction, and Contact Information. Customers were asked to discontinue use of their MR system and to contact their local service representative, if the lateral table adjustment crank handle on the table becomes loose. |  | 202 | 202 | Worldwide Distribution -- USA, including states of AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, LA, KS, KY, LA, MD, MA, MI, MN, MS,MO,MT, NV, NH, NJ, NY, NC, OH, OK, OR, PA, PR, SC, SD, TN, TX, VA, WA, WV, and WI and countries of VENEZUELA, ECUADOR, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, SPAIN, RUSSIAN FEDERATION, REPUBLIC OF KOREA, POLAND, PHILIPPINE,S OMAN, MEXICO, KUWAIT, JAPAN, ITALY, HUNGARY, GERMANY, EGYPT, DENMARK, CHILE, CHINA, and BRASIL. | Terminated | July 13, 2012 | 690 | 378 | Not_Computer
Z-0432-2011 | 49384 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart FRx Defibrillator, Automated External Defibrillator (AED), catalog number 861304. |  |  | August 07, 2008 | 2008 | Philips Medical Systems | The affected device may not deliver appropriate defibrillation therapy. | Employee error | Philips notified the customer with the affected device by phone on August 8, 2008 . A replacement device was shipped to the customer on August 11, 2008. |  | 1 | 1 unit | Nationwide Distribution: to one account in Arkansas. | Terminated | December 01, 2010 | 846 | 1 | Not_Computer
Z-0429-2011 | 56017 | No | NOE | Low-density lipoprotein immunological test system. | Immunology | Immunology | 510(k) | diaDexus PLAC Test ELISA Kit; Manufactured and distributed by diaDexus , Inc., South San Francisco, CA; Model Part Number: 90123 |  |  | March 15, 2010 | 2010 | DiaDexus, Inc | The reason for the recall is that Real Time Stability Test program showed the kit lot failed stability prior to expiration, affecting samples with results&gt; 500 ng/ml. Results up to 500 ng/ml exhibit passing stability results and are not impacted by this issue. | Other | DiaDexus notified all customers by phone, including distributor's customers, beginning on March 15, 2010. Customers were told to re-test samples that were tested from 2/1/2010 through 2/28/2010 which have values above 500 ng/ml with a new lot of product.For questions regarding this recall call (650) 246-6477. |  | 239 | 239 kits | Nationwide Distribution - US including NJ, OH, KS, IN, FL, CA, UT, OK, PA, WA and MI. | Terminated | January 11, 2011 | 302 | 239 | Not_Computer
Z-0437-2011 | 56215 | No | KYI | Wrist joint carpal trapezium polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | The Ascension PyroHemiSphere (aka NuGrip), Model #30-L NUG-443-30L, Catalog #NUG-443-30L, Lot #10-0731, is a single-use uncemented, one-component prosthesis for the basal thumb joint. |  |  | July 12, 2010 | 2010 | Ascension Orthopedics, Inc | Two different product lots were found to be incorrectly labeled and placed into the incorrect box. | Mixed-up of materials/components | Ascension Orthopedics notified by phone on June 25, 2010 and sent a letter dated July 22, 2010, to all customers. The customers were instructed to return the devices to the firm.Should you have any questions or concerns please feel free to contact (512) 836-5001 ext 1585. |  | 9 | 9 | Worldwide distribution: USA including states of: IL, IN, MD, OH, and TN; and countries of : Puerto Rico and Netherlands. | Terminated | June 30, 2011 | 353 | 19 | Not_Computer
Z-0431-2011 | 56416 | No | CDZ | Testosterone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Sex Hormone Binding Globulin (SHBG) Reagent for use on Access Immunoassay Systems, Part number(s): A48617 |  |  | July 26, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because the Instruction for Use (IFU) for Access SHBG does not specify the units of measure to be usedwhen calculating the Free Androgen Index (FAI %). Access SHBG (nmol/L) values and Access Testosterone (ng/mL) values are reported in different default units. The testosterone value (when reported in ng/mL) requires conversion to nmol/L before the FAI % can be calculated. | Pending | Beckman Coulter sent an Urgent Product Corrective Action (PCA) letter on the week of July 26, 2010 to the affected customers. The letter was dated July 30, 2010, and identified the product, the problem and provided the appropriate equation to recalculate the FAI% values. The customers were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Customers were also asked to complete and return the enclosed response form within 10 days to confirm that they had received the important notification.The customers were instructed to contact Beckman Coulter Technical Support at 1-800-854-3633 in the United States, or contact their local Beckman Coulter Representative if they had any questions regarding the notification. |  | 1113 | 1,113 Units ( 894 in US &amp; 219 in Canada) | Worldwide Distribution - USA including AL, AR, CA, CO, FL, GA, IL, IN, LA, MA, MI, MN, MO, MS, NC, NE, NV, NY, OH, OK, OR, TN, TX, UT, VA, WA, AND WI and the country of Canada | Terminated | May 04, 2012 | 648 | 1113 | Computer
Z-0430-2011 | 57098 | Yes | LWP |  |  | Cardiovascular | PMA | Medtronic Dual Chamber Temporary Pacemaker, model 5388. Manufacturer: Medtronic, Inc., 710 Medtronic Parkway, Minneapolis, MN 55432-5604, USA. |  |  | October 11, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | A subset of Model 5388 Dual-Chamber External Temporary Pulse Generators worldwide may be unable to power up or may power down unexpectedly. The issue presents itself in one of two ways. First, during startup, and prior to initiating patient therapy, the instrument may power down in 1 - 2 seconds. Second, the instrument may power up correctly, but power down at a later time while in use. | Nonconforming Material/Component | Medtronic issued an Important Medical Device Correction letters, dated October 2010 and November 2010 to hospital's biomedical engineering department. The letter described the issue, how it presents, and stated that Medtronic will contact the hospitals to return affected product for servicing. The second letter reiterated that Medtronic would notify customers when FDA approves their design upgrade for returning devices for service. In the interim, hospitals should continue to use normally functioning devices. Cusotmers may contact Medtronic Instruments Technical Service at 800-638-1991 concerning this action. |  | 5595 | 5,595 ( 2,696 US, 2,899 OUS) EXPANDED: 944 (852 US, 92 OUS) | Worldwide Distribution: throughout USA and territories of Guam and Puerto Rico and to the countries of: Algeria, Argentina, Australia, Austria, Barbados, Belarus, Belgium, Bolivia, Bosnia And Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, French Guiana, Georgia, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Kuwait, Latvia, Lebanon, Luxembourg, Macedonia, Malaysia, Martinique, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Oman, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, and Yemen NO ADDITIONAL COUNTRIES ARE AFFECTED BY THE EXPANSION. | Terminated | May 29, 2012 | 596 | 5595 | Hardware | Device Operation | Return for servicing | Remove or Replace
Z-0427-2011 | 56913 | No | KDN | Isolated kidney perfusion and transport system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Belzer UW Cold Storage Solution, 1000 mL. Sterile and non-pyrogenic. Formulated for flushing and cold storage of kidney, liver and pancreas organs. Manufactured by: Preservation Solutions, Inc., Elkhorn, WI 53121. Manufactured for: BTL Solutions LLC, a subsidiary of Bridge to Life, Ltd., 1225 Laurel Street, Columbia, SC 29201. |  |  | September 24, 2010 | 2010 | Preservation Solutions Inc | Mislabeled volume - 2 liter bags of Belzer UW Cold Storage Solution were labeled as 1 liter. | Labeling mix-ups | Bridge to Life, Ltd, issued a Notification of Product Recall letter dated October 18, 2010 to all direct accounts identifying the affected product and labeling issue. Customers were informed that all product from the affected lot has been replaced by the firm. Customers with questions about this recall can contact Bridge to Life at 803 673-4548.. |  | 55 | 55 units | Nationwide Distribution - in US to the states of CA, MO, and PA. | Terminated | April 02, 2012 | 556 | 55 | Not_Computer
Z-0439-2011 | 56985 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cholestech LDX Multi-Analyte Controls, Level 1 &amp; 2, 1 vial set, Catalog Number 12-712. |  |  | October 06, 2010 | 2010 | Biosite Inc Dba Innovacon Inc. | The recall was initiated because results were being reported as being out of range. Firm is reassigning ranges for HDL cholesterol in product. | Pending | All customers were notified and given instructions, requirements, and expectations for their role in the Correction activities using the Notice of Correction with attached Returned Verification Forms which were supplied via fax, email or direct mail beginning on October 6, 2010.Customers are to share the information provided in the Notice with their laboratory staff and retain the letter as part of their laboratory Quality System documentation. If any affected product had been further distributed to another laboratory, then a copy of the letter should be provided to them. Customers are to complete and fax the enclosed Verification Form within 10 days to confirm their receipt of the notice.Consignees with questions about the information contained in the notification were instructed to contact:Alere San Diego, Inc.9975 Summers Ridge RoadSan Diego, CA 92121U.S.A.Phone: 877 441 7440FAX: 858 805 8457E-mail: Response.ts@alere.com |  | 3382 | 3,382 kits | Worldwide Distribution -- USA, Australia, Canada, Great Britain, India, Italy, Mexico, Norway, Puerto Rico, South Korea, Spain, Switzerland, and Trinidad &amp; Tobago. | Terminated | August 01, 2012 | 665 | 6529 | Not_Computer
Z-0433-2011 | 57028 | No | FTY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LeMaitre VascuTape Radiopaque Tape, LeMaitre Stent Guide, 20 strips |  |  | October 22, 2010 | 2010 | Lemaitre Vascular, Inc. | Box label (Top Part). Lot and Expiration Date information were switched. | Labeling Change Control | Lemaitre Vascular notified customers by an Urgent: VascuTape Radiopaque Tape Device Field Safety Notice letter on 10/22/10 via USPS with delivery confirmation. The Notice advised users of the incorrect lot and lot number on the outer box and also provided an image of the box and label for ease in identifying the product in their inventory. The letter states that the use and functionality of the device are not affected by the problem; however, if customers feel uncomfortable continuing to use the affected device, they can contact Customer Service for replacement. The firm recommended discarding the outer box of the affected product. The letter contains a form that is requested to be returned to LeMaitre Vascular as record of notification and reconciliation. Accounts can contact customer service at 800-628-9470. |  | 160 | 160 units | Nationwide Distribution -- AL, CA, CT, IL, ME, MI, MS, NC, NJ, NY, PA, PR, TX, and WA. | Terminated | September 15, 2011 | 328 | 1940 | Not_Computer
Z-0444-2011 | 55889 | No | OFY | Endosseous dental implant accessories. | Dental | Dental | Enforcement Discretion | Zimmer Dental, Drill stop kit, REF DSKIT, for use with dental implants. |  |  | May 18, 2010 | 2010 | Zimmer Inc. | The instructions for use contain pre-vacuum sterilization instructions that may lead to compromised sterility at the lid-kit interface location. | Labeling design | Zimmer sent out "Urgent Device Correction" letters dated 5/18/2010. The consignees were provided with new instructions for use. The letter informed users that the current pre-vacuum cycle of 3 to 18 minutes is not enough time to assure sterility between the lid-kit interface and could lead to cross contamination of multiple patients. The users were informed that a pre-vacuum time of 30 minutes was required to assure sterility. |  | 158 | 158 | Distribution Nationwide, to France, and Spain. | Terminated | May 18, 2011 | 365 | 158 | Not_Computer
Z-0448-2011 | 55200 | No | CHP | Estradiol test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Estradiol, Part Number: 33540 |  |  | March 04, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated after Beckman Coulter confirmed that three lots of Access Estradiol reagent (REF 33540) contains an incorrect raw material which may produce incorrect patient results. | Other | A Product Corrective Action (PCA) letter dated March 17, 2010, with attached PCA Customer Response Form was sent to the affected customers. The letter describes the product, problem and actions to be taken by the customers. The customers were instructed to discontinue use of the three Access Estradiol Reagent lots; discard all remaining inventory of affected product; share this information with laboratory staff and retain the notification as part of their laboratory Quality System documentation, and complete and return the enclosed response form within 10 days via fax to 786-639-4000 or mail: Beckman Coulter, Inc., 1000 Lake Hazeltine Drive, Chaska, MN 55318, Attn: Regulatory Affairs.The following product replacement contacts were provided:In the United States, please contact Client Services at 1-800-526-3821.In Canada, please contact Customer Service at 1-800-463-7828.Outside the United States and Canada, contact your local Beckman Coulter Representative.Questions and assistance could be obtained through Technical Support at 1-800-854-3633 in the United States and Canada. Outside the United States and Canada, contact consignees were told to contact their local Beckman Coulter Representative. |  | 20209 | 20,209 units (5475 in US) | Worldwide distribution: USA and Internationally to: Algeria, Argentina, Armenia, Australia, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, China, Czech Republic, El Salvador, France, Georgia, Germany, Ghana, Greece, India, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahiriya, Malaysia, Mexico, Morocco, Namibia, Netherlands, New Zealand, Nigeria, Oman, Panama, Philippines, Poland, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Serbia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab, United Kingdom, Venezuela, Vietnam, and Zambia | Terminated | May 07, 2012 | 795 | 20209 | Not_Computer
Z-0446-2011 | 57008 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BARDÂż Mesh Pre-shaped (1.8" x 4.0"/4.5cm x 10cm) |  |  | October 15, 2010 | 2010 | Davol, Inc., Sub. C. R. Bard, Inc. | Mislabeled: The labeled BARDÂż Mesh Pre-shaped product code 0112700 (1.8" x 4.0"/4.5cm x 10cm), but the contents are the BARDÂż Mesh Large Pre-shaped product code 0113700 (2.4" x 5.4"/6.0cm x 13.7cm). | Process control | Davol issued an "URGENT MEDICAL DEVICE RECALL" letter on October 15, 2010, via Overnight Delivery to all customers. The letter describes the product, the problem and the action to be taken by the customers. The customers were instructed to immediately examine their inventory and quaratine the product; identify and notify customers they may have further distributed the product to of this recall; contact Davol Customer Service at 1-800-556-6275 for instructions on how to return product, and complete and return the attached EFFECTIVENESS CHECK form via fax to 1-401-825-8753. If you have an questions, please contact Davol Customer Service Department at 1-800-556-6275 or C. R. Bard's Medical Service Support Department at 1-800-562-0027. |  | 144 | 144 units | Nationwide distribution | Terminated | October 27, 2011 | 377 | 144 | Not_Computer
Z-0445-2011 | 57282 | [6 empty] | Lumen Biomedical Xtract Aspiration Catheter system, intravascular catheter, consisting of One Catheter, Two 30 mL Syringes, One Extension Tube with Stopcock, and One 40 um Cell Strainer. Single-user, 0.014" guidewire compatible intravascular extraction and aspiration catheter, Part number 1450-0001 size 4.2F, and part number 1451-003 size 5.8F, distributed by Volcano Corporation. |  |  | November 03, 2010 | 2010 | Volcano Corporation | Potential for product failure. Packaging/Storage issue results in compromised integrity and potential product failure during use, leading to potential patient injury. | Device Design | Firm initiated recall on 11/03/2010, and sent notifications to its consignees via letter. The firm has requested that consignees provide the notification to all personnel within the organization and letters will be sent directly to known customers on 11/04/2010 by Volcano. Customers have been instructed to cease use and contact customer service to arrange returns, and credit will be issued. |  | 4724 | 4724 units | Worldwide distribution: USA, Greece, Egypt, Kuwait, Netherlands, Saudi Arabia, Italy, France, Sweden, Portugal, Great Britain, Austria, Switzerland, Germany, Romania, Poland, and Bulgaria | Terminated | December 07, 2011 | 399 | 4724 | Not_Computer
Z-0447-2011 | 57048 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Medtronic Navigation O-Arm Imaging System, Door Assembly |  |  | October 15, 2010 | 2010 | Medtronic Navigation, Inc. | Safety Notice: Door Cable Periodic Maintenance Schedule and Cable Wear for the O-ARMÂżÂż Imaging System Breakage of the door cable during opening and closing may occur and cause physical harm or delay of therapy | Equipment maintenance | Medtronic Navigation sent all consignees/customers an "Urgent Field Safety Notice" letter dated October 15, 2010. The letter described the product, problem and action to be taken by the customers. The customers were instructed to review and follow the attached information that describes the process to properly detect and proactively manage cable wear; to ensure that the door is completely open when positioning the patient into or removing them from the gantry and when placing the O-ARM Imaging System Tube Drape (BI-900-00003) onto the unit, and report any visual or audible (cracking noise) wear of the cable to Medtronic Representative ASAP. Note: A Medtronic Representative will be contacting the customers within the coming weeks to schedule a time to inspect the current condition of the cable, explain and train the customers on the recommended workflow and inspection procedure and place a permanent label on their O-ARM Imaging System to reinforce these actions. If you have any questions about this notification, please contact your local Medtronic Navigation representative or call our Technical Services Group at 800-595-9709 or 720-890-3200. |  | 244 | 244 units | Worldwide distribution: USA including states of: AL, AZ, CA, CO, CT, DC, FL, GA, IA, IL IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, and WI; and countries including: Canada, Austria, Australia, Belgium, Brazil, Libya, South Africa, India, Japan, Singapore South Korea, Haiti, Kuwait, Saudi Arabia, Turkey, UAE, Germany, Netherlands, Great Britain, Italy, Poland, Netherlands, Spain,Sweden, and Switzerland. | Terminated | June 04, 2012 | 598 | 244 | Not_Computer
Z-0461-2011 | 55604 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | The Dimension NTP and LNTP Flex reagent cartridge (N- Terminal Pro- Brain Natriuretic Peptide and N-Terminal Pro brain Natriuretic Peptide (low volume) |  |  | April 12, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Some LOCI cartridges have sub-optimal seal of the lidstock that can cause rupture of the package at high elevations, causing the cartridge to leak. Use of leaking cartridges may result in inaccurate test results. | Other | Siemens issued an "Urgent Field Safety Notice" dated April 2010, informing users of the problem and instructing them to inspect cartridges for leaking fluid. Cartridges with leaking fluid should be reported to the Siemens Technical Solutions Center. The Technical Solutions Center can be contacted at 800 441 9250. |  | 3364 | 3364 | Worldwide Distribution: USA, Europe, Australia, China, Japan, Korea, and Saudi Arabia. | Terminated | March 19, 2013 | 1072 | 14037 | Not_Computer
Z-0456-2011 | 55178 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT CREA Cartridge, Abbott List Number: 06F10-01, 02; IVD. Manufactured by Abbott Point of Care, Inc., Princeton, NJ. |  |  | March 17, 2010 | 2010 | Abbott Point of Care Inc. | The i-STAT CREA Cartridges may produce inaccurate (depressed) results at or near the end of the labeled expiry date when stored above 25 degrees for 2 weeks. | Other | Abbott Point of Care issued Urgent Product Recall Notification letters dated March 2010 to all direct customers identifying the affected product and the actions to be taken. Customers were instructed to determine if they had cartirdges that had been stored above 25 degrees Celsius, and if so, contact Abbott for replacement or credit. Abbott Technical Support can be contacted at 800-366-8020 option 1 about this recall. |  | 34525 | 34525 cartridges | Worldwide Distribution: Throughout USA and Canada. | Terminated | September 22, 2011 | 554 | 34525 | Not_Computer
Z-0449-2011 | 56345 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens brand ARTISTE MV, Artiste Digital Linear Accelerators in combination with the Flat Panel, Part No. 8139789, |  |  | January 22, 2010 | 2010 | Siemens Medical Solutions USA, Inc | The robotic lift armature shaft may break and allow the Flat Panel Positioner with the digital flat panel imager to fall into the Flat Panel cradle. If the positioner is over a patient, there is a potential for injury. | Other | Siemens sent an Urgent Medical Device Correction letter to all on January 27, 2010, via phone, Fax, certified mail and/or delivery via Siemens Customer Service Group to all its direct consignees, informing them of the affected products, the problem, and provided instructions on the recall.Consigneees were to be contacted by an ARTISTE Service Representative for an appointment for corrective action. Consignees were instructed to include Customer Safety Advisory Notice in their ARTISTE System Owner Manual.For questions regarding this recall call (925) 246-8407. |  | 106 | 106 | Worldwide Distribution - USA including CA, DE, FL, LA, MA, MI, MO, NC, NE, NJ, OH, PA, TN, UT, WI, AND WV and Egypt, Germany, Japan, Lebanon, Malaysia, New Zealand, Poland, Russia, Spain, Sweden, and Turkey | Terminated | April 12, 2011 | 445 | 106 | Not_Computer
Z-0457-2011 | 56147 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics V60 Ventilator, manufactured by Respironics California Inc. in Carlsbad, CA. |  |  | July 02, 2010 | 2010 | Respironics California Inc | Respironics California Inc has received customer complaints of an odor or burning smell coming from the V60 Ventilator, and identified a component on the Power Management PCB that exhibited thermal damage. | Device Design | Phililps Healthcare issued an Urgent Medical Device Field Correction letter dated July 2, 2010 concerning the RespironicsV60 Ventilator issue. Customers will be contacted by US Field Service Technicians and international distributors to schedule and perform inspections and replacement of the affected component. The unit may continue to be used pending completion of the field inspection.Customers may contact Respironics US Customer Service about this action at 1 800 345 6443, option #5, then option #1. |  | 504 | 504 | Worldwide Distribution: Throughout the US and Puerto Rico,, and to the countries of: Australia, Austria, Belgium, Canada, Denmark, France, Germany, Hong Kong, India, Italy, Kenya, Kuwait, Libya, Malaysia, Oman, Netherlands, Norway, the United Kingdom, Saudi Arabia, and Spain. | Terminated | December 01, 2010 | 152 | 504 | Hardware | Physical Safety Hazards | Replace component | Remove or Replace
Z-0450-2011 | 56912 | No | GXQ | Dura substitute. | Neurology | Neurology | 510(k) | DuraGen XS Dural Regeneration Matrix: 2 inches x 2 inches. |  |  | September 29, 2010 | 2010 | Integra LifeSciences Corp. | Test results no longer support a 3 year shelf life. | Other | Integra LifeSciences Corporation sent an URGENT Product Recall Notification letter dated September 29, 2010, to all affected customers. Integra NeruoSpecialists were informed of the recall via telephone conference call and recall notification letters sent by e-mail. Letters were also sent to customers by Federal Express with copies of e-mail provided to Interga Neuro Specialists on October 1, 2010.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately quarantine any affected product and contact Integra Customer Service at 1-800-654-2873, select option 5, followed by option 3.Customers were instructed to fill out and return the enclosed "Recall Acknowledgement and Return Form" indicating whether or not they possessed any of the affected product.For any questions related to this recall call 1-800-654-2873 and selection option 1. |  | 118 | 118 | Nationwide Distribution including AK, AL, AZ, CA, CO, DC, DE, FL, GA, HI, IL, IN, LA, MA, MD, ME, MN, MO, MT, NC, ND, NH, NJ, NM, NY, OK, PA, TN, TX, UT, VA, VT, WA, WI, AND WV. | Terminated | April 23, 2012 | 572 | 2766 | Not_Computer
Z-0428-2011 | 57166 | No | NRY | Percutaneous catheter. | Cardiovascular | Neurology | 510(k) | Penumbra System Reperfusion Catheter 032, catalog number PSC032, lot number lot F15020, manufactured by Penumbra, Alameda, CA. |  |  | October 25, 2010 | 2010 | Penumbra Inc. | Reperfusion catheter may fail at the mid-shaft joint and may break off. If the segment cannot be removed, it may result in blocked blood flow, requiring emergency medical treatment. | Process control | Penumbra issued an Urgent Voluntary Field Removal letter dated October 28, 2010 to consignees, identifying the affected device and actions to be taken by consignees. Customers were instructed to inspect their inventory for the affected product and quarantine it, pending return to Penumbra. Penumbra employees will contact customers to arrange return and replacement of affected units. Customers were also requested to complete and return a product identification form and return form to Penumbra within 5 business days.Customers can contact Penumbra at 1 510 748-3223 concerning this action. |  | 61 | 61 | Worldwide Distribution: 17 consignees in the US, including in the states of CA, FL, HI, ID, IL, KY, MA, MS, NJ, NY, PA, and TX, 1 in Canada and 21 in the EU countries of France, Germany, Hungary, Italy, Norway, Slovenia, Spain, and the United Kingdom. | Terminated | June 01, 2011 | 219 | 61 | Not_Computer
Z-0455-2011 | 56012 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Inc. Eclipse/VARiS Vision, an Eclipse radiotherapy Treatment Planning System (Eclipse TPS); Eclipse Version 7.3.10 (multiple workstations) VARiS Vision System 7. x; |  |  | May 13, 2010 | 2010 | Varian Medical Systems Oncology Systems | Incorrect treatment plan or record: A software design deficiency has the potential for corrupted treatment plan records or treatment records in cases where multiple users editing a record simultaneously can result in a corrupted (but accessible) treatment plan record. | Software design | On May 13, 2010, the firm issued letters to all its direct consignees, informing them of the affected products and providing instructions on the recall.***12-20-12, AMENDED INFO ***: An additional letter was sent to consignees on 11/5/12 that Varian identified as also being affected by the recall.*** |  | 2466 | 2417 units***12-20-12 AMENDED TO: 49 units*** | Nationwide and Worldwide | Terminated | February 24, 2012 | 652 | 2466 | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0462-2011 | 57142 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SignaÂż Excite" 1.5 T Magnetic Resonance Systems, manufactured by GE Healthcare, P.O. Box 44 Milwaukee, WI 53201. |  |  | May 11, 2009 | 2009 | GE Healthcare, LLC | SignaÂż MR Excite" 1.5T system containing 11.1_M4_0818.a software, associated with 2D Fast Spin Echo based pulse sequences (2D FSE-XL, 2D FRFSE-XL, 2D FSE-IR, 2D T2FLAIR and 2D T1FLAIR) may experience an image flip along the phase encode direction and incorrect annotation, which may result in incorrect localization of anatomy in oblique axial images with the pulse sequences. | Software design | GE Healthcare issued an "Urgent Medical Device Correction" letter dated April 21, 2009. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information. GE Field Engineers will contact customers to schedule a field modification of the affected products to address this issue.Customers can contact GE Healthcare at 262 521-6100 about this issue. |  | 1464 | 1464 | Worldwide Distribution: Throughout USA except states of AK, VT and WY, and the countries of: Argentina, Australia, Bulgaria, Brazil, China, Canada, Switzerland, Chile, Colombia, Czech Republic, Ecuador, Egypt, Spain, Finland, France, Great Britain, Guatemala, Honduras, Croatia, Hungary, Indonesia, Ireland, India, Iran, Iceland, Italy, Japan, Kenya, South Korea, Kuwait, Lebanon, Luxembourg, Libya, Morocco, Malta, Mexico, Malaysia, Netherlands, Norway, New Zealand, Panama, Pakistan, Poland, Puerto Rico, Portugal, Reunion, Romania, Russian Federation, Saudi Arabia, Sweden, Singapore, Syria, Thailand, Tunisia, Turkey, Taiwan, Venezuela, and South Africa. | Terminated | November 24, 2015 | 2388 | 1464 | Software | Display/Image | Software update | Software Update
Z-0481-2011 | 50117 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | T2 Ankle Arthrodesis System Nail and Nut Adapter; |  |  | August 02, 2006 | 2006 | Stryker Howmedica Osteonics Corp. | The thread of the nut which should fix the nail adaptor to the targeting arm may seize on the threat of the nail adaptor, interrupting surgery. | Other | Stryker Orthopedics issued an Important Stock Recovery letter dated August 2, 2006 The customer was instructed to inspect inventory for the affected product and return it to the firm. The customer was asked to complete and return a Product Accountability Form and contact their Stryker representative for replacement product.Stryker can be contacted about this issue at 201 831-5118. |  | 8 | 8 sets | Nationwide Distribution: To one location in New Jersey. | Terminated | November 22, 2013 | 2669 | 8 | Not_Computer
Z-0482-2011 | 55760 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Neuroform 3 Microdelivery Stent System; Distributed and manufactured by: Boston Scientific Corporation, Fremont, CA |  |  | March 01, 2010 | 2010 | Boston Scientific Corporation | The product may contain manufacturing aids (mandrels) which should not be present in the final product. If the mandrel remains in place, the catheter cannot be loaded over the required guidewire, so cannot be used. | Process control | Boston Scientific issued an Urgent Customer Notice dated March 1, 2010 to all US customers, with personal delivery by sales representative where applicable outside the US, identifying the affected device. Customers were instructed to return the device for replacement, and to return an acknowledgement form.Customers can contact Deborah Springer at Boston Scientific at 800-845-4598 about this recall. |  | 638 | 638 units | Worldwide Distribution: Throughout the US and to the countries of: Austria, Australia, Belgium, Brazil, Canada, China, Columbia, Czech Republic, Germany, Denmark, Ecuador, Spain, France, Great Britain, Greece, Hungary, India, Ireland, Italy, South Korea, Luxemburg, Mexico, the Netherlands, New Zealand, Peru, Russian Federation, Sweden, Slovenia, Taiwan and Ukraine. | Terminated | November 26, 2010 | 270 | 638 | Not_Computer
Z-0479-2011 | 52100 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Product's marketing brochure is labeled in part: "Philips Easy Diagnost Eleva Conventional R/F system...PHILIPS". manufactured by Philips Medical Systems Gmbh, Hamburg, Germany. |  |  | January 21, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | Patient and image data can be exchanged when a patient is selected in the patient list and the examination tab is pressed with background query running at the same time. There is a risk that the data may then be linked to the incorrect patient, which may adversely affect diagnosis and treatment. | Software design | Philips Medical Systems issued URGENT-Device Correction letters dated January 27, 2009 to consignees describing the issue with the Easy Diagnost Eleva GXR RF systems. The consignees were informed of the potential for incorrect diagnosis or treatment if an incorrect image is used for diagnosis. Consignees should contact the Philips Healthcare Call Center at 800-722-9377, #5, #2 and reference "FC0 70600030" with any communication. |  | 314 | 314 units | Nationside Distribution: USA, including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NE, NH, NJ, NV, NY, OH, OK, OR, PA, PA, RI, SC, SD, TN, TX, TX, UT, VA, VT, WA, WI, and WV. Devices are not distributed to any foreign consignees. | Terminated | December 30, 2010 | 708 | 314 | Software | Output/Calculation | Notification | Safety Notice/Insructions
Z-0469-2011 | 52040 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | The product's marketing brochure states: "1.5T coils and accessories...Philips Achieva and Intera 1.5T MR systems". |  |  | May 01, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | Combined use of the Synergy Flex-M / Shoulder Coil 1.5T increases the chance of RF interaction and heating up of the coil. This may result in possible burns of the patient. | Other | Philips issued URGENT-Field Safety Notice letters dated May 1, 2009 and October 14, 2009 to their consignees. The letter informed customers that the firm will replace the "old" Flex M/Shoulder Coils with a new Flex-M coil.Consignees could contact Philips Healthcare Call Center at 800-722-9377, #5, #3, #1 and reference FCO 78100290 for any questions. |  | 389 | 389 units | Nationwide Distribution: in the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VA, WA, WI, WV, WY, and Washington DC. There were no foreign consignees. | Terminated | January 06, 2011 | 615 | 389 | Not_Computer
Z-0473-2011 | 55682 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart FR2+ automated external defibrillator (AED) model: M3841: LAERDAL brand; text prompts only on screen (NO ECG); no configurable manual charge. |  |  | May 28, 2010 | 2010 | Philips Medical Systems | Failure of the voltage detector, U8 may cause the AED battery to drain more rapidly than normal or make the AED unusable. | Nonconforming Material/Component | Philips Medical Systems issued an Urgent Medical Device Recall letter dated May 28, 2010 identifying the affected device and actions to be taken. Customers were instructed to review their inventory for affected devices. If affected devices are found, the customer should contact Philips to arrange for a replacement device. The customer should also run a battery insertion test. If the device passes the test, the AED may remain in service until a replacement is received. For further information, contact Philips at 1 800 263-3342, option 6. |  | 1478 | 1,478 units | Worldwide Distribution: Throughout the USA, and to the following countries: ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BERMUDA, BOSNIA AND HERZEGOVINA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, FINLAND, FRANCE, GERMANY, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, REPUBLIC OF KOREA, KUWAIT, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PERU, PORTUGAL, PUERTO RICO, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, and VIETNAM. | Terminated | September 13, 2013 | 1204 | 16573 | Other
Z-0478-2011 | 56428 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | The software is ARM version 60.80.101.025 (referred to as Nano 1.0.5). |  |  | August 13, 2010 | 2010 | Sonosite, Inc. | Software ARM version 60.80.101.025 (referred to as Nano1.0.5) on the NanoMaxx ultrasound system results in errors when the "print" command is selected. As a result, the displayed MI or TI values may be incorrect. | Software design | Sonosite issued Urgent Medical Device Correction letters dated August 9, 2010 to consignees. Consignees were advised of the software problem and given directions to verify the software version they are using. They are advised to contact SonoSite at 877-657-8118 to obtain a free software upgrade. |  | 69 | 69 | Worldwide Distribution: USA and to the countries of Australia, Great Britain, India, Israel, the Netherlands, New Zealand, Norway, and Spain. | Terminated | July 21, 2011 | 342 | 69 | Software | Output/Calculation | Software update | Software Update
Z-0480-2011 | 57023 | No | KZJ | Gas-generating device. | Microbiology | Microbiology | 510(K) Exempt | GasPak EZ Large Incubation Container, a component of the GasPak EZ Gas Generating Container System, catalog # 260672, labeled in part ***Becton Dickinson and Company Loveton Circle Sparks, MD 21152 USA*** |  |  | August 13, 2010 | 2010 | Becton Dickinson &amp; Co. | In vitro diagnostic reagent containers may be defective and could cause incorrect test results in patient samples. | Nonconforming Material/Component | BD Diagnostic Systems issued an Urgent Product Recall letter dated August 2010 to consignees. Consignees were advised of the defect and were requested to discard product for replacement. Distributors were additionally requested to provide distribution lists of end users for direct BD notification. Accounts were requested to return a response form.BD Customer Service can be contacted at 800 675-0908, and BD Technical Services can be contacted at 800-638-8663. |  | 496 | 496 units | Worldwide Distribution: USA and the countries of Australila, Belgium, Canada, India, Mexico, New Zealand. Singapore, and South Korea, | Terminated | January 05, 2011 | 145 | 496 | Not_Computer
Z-0465-2011 | 57117 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Rectal Tubes, 22 French, 20 Long (51 cm), X-Ray Opaque Rubber with Funnel end, Open Tip, One Eye , Non sterile, 12 Each, C.R. Bard, Inc., Covington, GA 30014 |  |  | October 28, 2010 | 2010 | C.R. Bard, Inc., Urological Division | The product is misbranded. The product is labeled as 22 French, however the product is actually 20 French. | Employee error | Bard sent an Urgent Product Recall Notification letter dated October 28, 2010 to the consignee. They were instructed to return any outstanding inventory of the affected product. They were advised to notify their consignees of the recall and to retrieve relevant product from those locations. Affected products received by Bard Medical will be destroyed.BMD can be contacted at 1 770 784-6969 for questions or assistance. |  | 108 | 108 eaches | Nationwide Distribution: In the state of Ohio. | Terminated | March 31, 2011 | 154 | 108 | Not_Computer
Z-0463-2011 | 57115 | No | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE Stainless Steel Primary Molar Crowns, Made in USA by 3M ESPE Dental Products, St Paul, MN. |  |  | October 06, 2010 | 2010 | 3M Company / Medical Division | Unitek Stainless Steel Permanent Crowns with product Codes 902150, 900213, and 900223, is mislabeled.The crown contained within the package is incorrectly labeled; series UL3 crowns are labeled series UR3 and vice versa. The lid of the package correctly identified the product number and lot. This mislabeling could result in placement on the right side of a crown intended for use on the left, and | Packaging process control | 3M ESPE issued a Device Recall letter dated October 6, 2010 to distributors and dentists. The letter listed the product being recalled including code and lot codes, background information including a picture of the product problem, and the recall procedure. The letter included a template letter to be used by distributors to inform their customers who have received the product. Mislabeled product is to be quarantined and returned to 3M ESPE.3M ESPE can be contacted at 651 733-1110. |  | 9819 | 9819 | Nationwide Distribution: Throughout USA, including the states of AZ, CA, GA, IL, IN, IA, LA, MA, MN, MO, NV,NJ, NY, NC, PA, TN, TX, WA, and WI.Canada and Dominican Republic | Terminated | February 04, 2012 | 486 | 9819 | Not_Computer
Z-0464-2011 | 57109 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | DX-G Digitizer, manufactured by Agfa-Gevart NV, Mortsel, Belgium. |  |  | October 08, 2010 | 2010 | AGFA Corp. | Packaging was switched between a DX-G Digitizer and a DX-M Digitizer, resulting in the wrong product, which is not sold in the US, being shipped to the consignee. | Other | Agfa halted Installation of the incorrect unit, and the unit was sent to Agfa warehouse in Carlstadt, NJ, where it is being held in quarantine for export. |  | 1 | One unit | Nationside Distribution: to NJ. | Terminated | December 30, 2010 | 83 | 1 | Not_Computer
Z-0468-2011 | 57095 | No | HWX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CSS-059-40-4 Modified 4.0mm Modified Depth Gauge, Lot 0985011 |  |  | October 15, 2010 | 2010 | Orthohelix Surgical Designs Inc | The print for modified CSS depth gauges has inaccurate definition of the etched screw length. The error results in a length reading that is longer (5 mm in the case of the 7.0 and 4mm in the cases of the 5.5 and 4.0) than the actual length of the buried guide wire segment and even longer (10 mm in the case of the 7.0 and 8mm in the cases of the 5.5 and 4.0) than the design intention readings. This | Device Design | OrthoHelix issued an Important Recall Notice letter dated October 15, 2010 to their customers, identifying the affected products. Customers were requested to immediately return the afected devices. OrthoHelix can be contacted at 330 247-1444 concerning this recall. |  |  |  | Nationwide distribution: to the states of GA and MO. | Terminated | September 06, 2012 | 692 |  | Not_Computer
Z-0475-2011 | 57090 | Yes | LWS |  |  | Cardiovascular | PMA | Medtronic Sprint Quattro Secure, model 6947. Steroid eluting, quadripolar, screw-in, ventricular lead with Right Ventricular (RV) and superior vena cava (SVC) defibrillation coil electrodes. Contents Sterile. Manufactured in: Villalba, Puerto Rico, USA. Medtronic, Inc., Minneapolis, MN 55432 U.S.A. |  |  | October 25, 2010 | 2010 | Medtronic Inc. Cardiac Rhythm Disease Managment | Over-retraction of the helix during initial implant or subsequent repositioning may result in the inability to extend the helix. This does not impact acute or chronic performance of successfully implanted leads. | Device Design | Medtronic issued an Important Medical Device Correction letter, dated October 2010, to doctors. The letter contains a Performance Note. The letter describes the issue and the performance note provides information on how to mitigate the issue through replacement with a new lead.Customers can contact Medtronic Technical Services at 800 723-4636 |  | 154000 | 154,000 total | Worldwide Distribution: Throughout the US, including Guam, Puerto Rico, and US Virgin Islands, and to the countries of: Algeria, Argentina, Armenia, Australia, Austria, Bahamas, Bangladesh, Barbados, Belgium, Bolivia, Bosnia And Herzegovina, Brazil, Brunei Darussalam, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, El Salvador, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Malta, Martinique, Mexico, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Saint Lucia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Viet Nam. | Terminated | February 04, 2012 | 467 | 154000 | Not_Computer
Z-0477-2011 | 57077 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare IDI Mammo Workstation |  |  | August 28, 2009 | 2009 | GE Healthcare, LLC | The measurement values provided by the GE Workstation may be incorrect when applied to magnification images acquired on non-GE digital mammography systems.If not noticed by the caregiver, this could lead to an overestimate of the size of the breast lesion that may impact patient treatment and safety. The measurement values for GE images are not affected. | Software design | GE Healthcare issued "Urgent Medical Device Correction" letters dated August 28, 2009 or August 31, 2009 to either Seno Advantage Customers (Hospital Administrators/Risk Mangers, Radiology Department Managers), and Image Diagnostic Customers (Hospital Administrators/Risk Managers, Radiology Department Managers). The letters described the Safety Issue, Affected Product Details, safety Instructions, Product correction and Contact Information. GE will linstall new software with the release of Field Modification Instruction 12129.Customers can contact GE Healthcare about this issue at 800-437-1171. |  | 158 | 158 | Worldwide Distribution: USA in the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI,MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WV, WI, and WY, and to the countries of UNITED KINGDOM, TURKEY, TAIWAN, SWITZERLAND, SPAIN, SOUTH, KOREA, SAUDI ARABIA , PORTUGAL, NEW ZEALAND, NETHERLANDS, KUWAIT, JAPAN, ITALY, GERMANY, FRANCE , CHINA, CANADA, BRAZIL, BELGIUM, AUSTRIA, AUSTRALIA, VENEZUELA, UNITED ARAB EMIRATES, THAILAND, SWEEDEN, SOUTH AFRICA, SLOVENIA, SINGAPORE, RUSSIA, PHILLIPNES, NORWAY, POLAND, MOROCCO, MEXICO, MARTINIQUE, MALAYSIA, LUXEMBOURG, LEBANON, JORDAN, JAMAICA, ISRAEL, IRELAND, INDIA, ICELAND, HUNGARY, HONG KONG, HONDURAS, GUADOPUO, GREECE, GEORGIA, FINLAND, EGYPT, DOMINICAN REPUBLIC, DENMARK, CZECH REPUBLIC, CROATIA, COLOMBIA, BULGARIA, BOTSWANA, and ALGERIA. | Terminated | June 05, 2012 | 1012 | 2672 | Computer
Z-0485-2011 | 56816 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart FR2+ AED with ECG Display, Model 3840, Laerdal brand, configurable manual charge in advance mode. |  |  | September 22, 2010 | 2010 | Philips Medical Systems | A flash memory component in the HeartStart FR2+ and Automated External Defibrillator was improperly relabeled by a third party. Therefore the component could not be identified by its source, quality, and date of manufacture. | Nonconforming Material/Component | On 10/5/10, Philips began sending out the Voluntary Medical Device Removal notification letter to their consignees.This letter informs the consignees that Philips Healthcare is conducting a voluntary product removal of a limited number of HeartStartFR2+ automated external defibrillators (AEDs), model numbers M3860A, M3861A, M3849A, and M3841A, manufactured by Philips and shipped between March 2007 and August 2010. The owners of affected AEDs are advised to run a battery insertion test. If it passes, the AED should remain in service until its replacement is received. Philips will be replacing affected units.The consignees can visit the firm's website www.philips.com/FR2AEDAction or contact Philips at 1-800-263-3342 for any information or support concerning the issue. |  | 3244 | 3244 units | Devices were distributed to hospitals, medical facilities, and distributors through out the US and to foreign consignees. Foreign consignees including countries:Algeria, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Ecuador, El Salvador, Finland, France, French Guiana, Georgia, Germany, Ghana, Gibraltar, Guatemala, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Japan, Jordan, Kampuchea, Kenya, Korea, Latvia, Lebanon, Libya, Luxembourg, Malaysia, Malta &amp; Gozo, Mexico, Netherland, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestinian Territory, Panama, Peru, Philippines, Portugal, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, United Arab Emirates, United Kingdom, Vietnam, and Zimbabwe. | Terminated | September 05, 2013 | 1079 | 27270 | Not_Computer
Z-0489-2011 | 56775 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | CELL-DYN Vent Needle used on the CELL-DYN Sapphire and CELL-DYN 4000 Hematology Analyzers; |  |  | July 23, 2010 | 2010 | Abbott Diagnostics Division, 5440 Patrick Henry Drive, Santa Clara, CA | The Vent Needle, which is built into the Vent Head Assembly on the CELL-DYN Sapphire Instrument and which is used directly on the CELL-DYN 4000 Instrument, may be defective, which may result in an undetected short sample, generating incorrect low results, or which may bend, preventing aspiration of the sample. | Process control | Abbott Laboratories issued a Product Recall letter dated July 23, 2010 to all its direct consignees, informing them of the affected products and providing instructions on the recall. Customers were instructed to discard all affected vent needles and replace them with the needle included with the recall letter. Customers may contact customer support at 1-877-422-2688 concerning this recall. |  |  |  | Worldwide Distribution: Thoroughout the US, and in the countries of Argentina, Canada, Germany, Hong Kong, Japan, and Singapore. | Terminated | September 01, 2011 | 405 |  | Not_Computer
Z-0490-2011 | 57002 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx E. coli/P. aeruginosa PNA FISH Culture Identification Kit for the diagnosis of E. coli and/or P. aeruginosa bacteremia |  |  | October 06, 2010 | 2010 | AdvanDx, Inc. | Strains of Acinetobacter radioresistens cross-react with both E. coli/P. aeruginosa PNA FISH and EK/P. aeruginosa PNA FISH to produce a false positive red signal. | Device Design | AdvanDx issued a customer letter dated October 6, 2010 with a PDF file of a new package insert and an acknowledgement form to fill out and return to AdvanDx. Customers were instructed to replace the package insert in any existing inventory with the enclosed package insert.Customers can contact AdvanDx technical support about this action at 1-866-376-0009. |  | 260 | 260 units | Worldwide Distribuition: Throughtout the US, and to the countries of Denmark, Italy, Sweden, and UK. | Terminated | June 06, 2011 | 243 | 260 | Not_Computer
Z-0493-2011 | 55119 | No | HWN | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer Reconstruction System Tenaculum Left Tine Straight 9 inch, Item # 00-1179-021-00, Zimmer Inc., Warsaw, IN. |  |  | February 16, 2010 | 2010 | Zimmer Inc. | The firm determined that the recalled instruments were manufactured using the wrong grade of material. As manufactured, the instruments are more brittle, increasing the potential for fracture. | Nonconforming Material/Component | Zimmer sent an Urgent Device Recall letter dated February 16, 2010, to all affected customers. The letter identified the product, the problem, and the action the customers should take. Customers were instructed to monitor all potentially affected patients through regular office consultations. Ensure Zimmer Reconstruction System Tenaculum System Instrument Users at their facility were informed of the urgent device recall.Locate all instruments and quarantine them immediately.Count all affected product in their territory and record the data on theInventory Return Certification Form included with the letter.Fax a copy of the completed Inventory Return Form to:Zimmer, Inc. at (574) 372-4265.Return the recalled product along with the original completed Inventory Return Certification Form to:Zimmer Distribution CenterAttn: Product Service1777 West Center StreetWarsaw, IN 46580For questions customers were instructed to contact their Zimmer Sales Representative.For questions regarding this recall call 1 - (800) 613-6131. |  | 217 | 217 all products | Worldwide Distribution - USA including AL, FL, HI, IA, IL, KS KY, LA, MI, MO, MS, NC, ND, NJ, OH, OK, OR, PA, TN, TX, UT, VA, WA, AND WI and the countries of Australia, Canada, Central or South America, China, Germany, India, Italy, Saudi Arabia, and UK, | Terminated | October 05, 2011 | 596 | 217 | Not_Computer
Z-0502-2011 | 56892 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Interlink System, Micro-Infusion Manifold; and Interlink Manifold. An Rx sterile device with 3 injection sites, volume per injection site 0.1 mL, priming volume 0.8 mL, length 8.7 cm. Product code 2N3410 and A2N3410 - Japan |  |  | October 21, 2010 | 2010 | Baxter Healthcare Corp. | If the manifolds are connected in series, there may be a risk of the manifolds loosening or disconnecting at the luer connection between them. | Other | Baxter sent an "Urgent Device Correction" letter dated October 20, 2010, via first class mail to the direct accounts, to the attention of the Director of Nursing and the Director of Materials Management, informing them that there may be a risk of the manifolds loosening or disconnecting at the luer connection between them. The loosening or disconnection can result in an interruption of therapy, potentially resulting in patient injury and death. The accounts were requested to place the notification where the inventory of affected product is dispensed, so that users are made aware of the procedures to follow: "If these devices are connected in series, and the luer connector is tightened, do not twist the manifolds. This may lead to loosening or disconnection, which may cause a leak. Additionally, Baxter recommends checking all connections before starting and during infusion therapy." The accounts were requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at the number provided on the form. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  | 752965 | 752,965 units in US and all lots in Japan | Worldwide Distribution -- United States and Japan. | Terminated | October 16, 2011 | 360 | 752965 | Not_Computer
Z-0491-2011 | 56997 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx EK/P. aeruginosa PNA FISH- Culture Identification Kit for identification of E. coli and/or K. pneumoniae and P. aeruginosa bacteremia (In-Vitro Diagnostic) |  |  | October 08, 2010 | 2010 | AdvanDx, Inc. | Strains of Acinetobacter radioresistens cross-react with both E. coli/P. aeruginosa PNA FISH and EK /P. aeruginosa PNA FISH to produce a false positive red signal | Device Design | AdvanDX sent an e-mail notification along with a PDF file to all customers on October 8, 2010, followed by written notification dated October 6, 2010. The letter identified the product, the problem, and the action the customers should take. Customers were instructed to replace the package insert(s) included with any kits in their inventory with the enclosed package insert(s). Customers were also instructed to fill out the enclosed Acknowledgement of Receipt Form and return it to AdvanDx to confirm that they received the new package insert. For questions regarding this recall contact AdvanDx Technical Support at 1-888-376-0009. |  | 27 | 27 kits | Worldwide Distribution - USA including IL, MI, NY, and PA and the country of Denmark | Terminated | June 06, 2011 | 241 | 27 | Not_Computer
Z-0492-2011 | 56405 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit Model 8000 only when used with Alaris Pump module or Alaris Syringe module |  |  | August 02, 2010 | 2010 | CareFusion Corporation | Than recall was initiated because CareFusion has identified identified a potential risk associated with the Alaris PC unit formerly known as Medley PC unit (PC Unit) model 8000 only when used with Alaris Pump module (Pump module) or Alaris Syringe module' (Syringe module), This recallnotification details the potential risk and recommended steps for users to take if they encounter a specific PC | Device Design | CareFusion will send by registered return receipt mail the customer notification letter to all customers usingthe affected Alaris PC Unit beginning August 24, 2010. The customer notification letter will be addressed tothe Directors of Nursing, Risk Management, and Biomedical Engineering. Customers will be required toconfirm receipt of the notification by returning the Recall Response Card to CareFusion by postage-paid, selfaddressedmail, fax or email. The recall notice informed the customers of the reason for the notice, the unitsaffected, the potential risk, a step-by-step description of the issue and the required actions for users.The customers were also told that CareFusion is working to release a software update to address this display error and will contact your facility as soon as the software update is available. CareFusion does not require that you return your devices. Clinicians should weigh the risk/benefits to the patients beforecontinuing the use of the device.Customers were also told that any adverse reactions experienced with the use of this product, and/or quality problems shouldalso be reported to the FDA's MedWatch Program. The recall notice provided contact information for CareFusion's Recall Center, Customer Advocacy and Technical Support units. |  | 94146 | 94,146 units | Nationwide in the United States, Canada, Mexico, and the European Union | Terminated | July 25, 2013 | 1088 | 94146 | Computer
Z-0350-2011 | 56753 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta Synergy XVI |  |  | September 01, 2010 | 2010 | Elekta, Inc. | It is possible to do couch corrections twice if both Couch Move Assistant in MOSAIQ and Remote Automatic Table in XVI are configured when in SYNERGISTIQ mode. | Other | Elekta sent an "Important Notice A340" dated September 1, 2010 to all customers. The notice described the product, problem and actions to be taken by the customer. The customers were instructed to make a decision if XVI or MOSAIQ does the corrective table movement; configure and ensure that the system only uses RATM or CMA for the corrective table movement, when in SYNERGISTIQ mode, and complete and return the IMPORTANT NOTICE Confirmation of Receipt form via fax at +44(0)1293 654401. If you have any questions contact your local service representative at +44(0)1293 654200. |  | 23 | 23 units | Nationwide distribution: AZ, CA, DC, IL, KY, MA, MI, MS, NV, NY, OH, OR, PA, VT, and WV. | Terminated | December 14, 2010 | 104 | 23 | Not_Computer
Z-0336-2011 | 56939 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | ENDOPATHÂż XCEL" with OPTIVIEW" Technology trocar |  |  | September 27, 2010 | 2010 | Ethicon Endo-Surgery Inc | This recall has been initiated because Ethicon Endo-Surgery is aware of an increase in reports of leaking in the ENDOPATHÂż XCEL" with OPTIVIEW" Technology trocar, which potentially can lead to loss of pneumoperitoneum during a surgical procedure. Pneumoperitoneum leaks may lead to conversion from laparoscopic to open surgery. | Device Design | Ethicon Endo-Surgery sent an URGENT: DEVICE RECALL EVENT 2244 letter, dated October 8, 2010. The letter identified the product, the problem, and the action to be taken by the customer. The customers were instructed to stop using the product, examine all inventory to determine if they had the affected product on hand, and remove the affected products from inventory. Ethicon Endo-Surgery requested that each customer complete the Business Reply Form and fax it back to Stericycle at (877) 272-6141 with in 3 business days and also return any affected product to Stericycle.For question regarding this recall call (877) 598-5706 and reference 2244. |  | 671 | 671 lots totalling 2,479,134 instruments | Worldwide distribution: USA, ARGENTINA, AUSTRALIA,BELGIUM,BRAZIL,CANADA,CHILE,CHINA,COLOMBIA,COSTA RICA, DOMINICAN REPUBLIC,ECUADOR, EL SALVADOR, GREECE, GUATEMALA, HONG KONG, ISRAEL, ITALY, JAPAN,JORDAN, KOREA, REPUBLIC OF, LEBANON, NETHERLANDS ANTILLES, NICARAGUA, PANAMASAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICASPAIN, TRINIDAD AND TOBAGO, UNITED ARAB EMIRATESURUGUAY, and VENEZUELA. | Terminated | January 25, 2012 | 485 | 671 | Computer
Z-0506-2011 | 57102 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Merit MEDALLION Syringe, 10ml, REF: MSSW11-Y, Sterile EO, Single Use, Merit Medical Systems, Inc., South Jordan, Utah 84095. |  |  | October 22, 2010 | 2010 | Merit Medical Systems, Inc. | Packaging for syringes may have holes which would compromise the sterility of the package contents. | Packaging process control | Merit Medical notified product managers of the product recall on October 22, 2010 following a voice mail script. Frequently Asked Questions, Sales Rep Letter, and Recall Notification Forms were then sent.. Sales reps were instructed to contact all of their customers and have them identify and quarantine any unused inventory and either return it to Merit or discard. A response form was to completed by the sales rep and the firm's representative, indicating amount of product quarantined, returned , or destroyed.Also, on 10/22/2010, authorized distributors were contacted by phone, mail, and / or fax. They were also provided a Distributor Letter, FAQs, and Recall Notification Forms instructing them to identify and quarantine any unused inventory and return it to Merit or discard. They were instructed to continue the recall to customers to whom they had shipped the affected Merit product.Additional information is available from Merit at 801-316-4822. |  | 25 | 25 units | Worldwide Distribution: USA including the states of AL, CA, GA, OH, PA, TX, UT, and WA, and to the countires of Japan and UK. | Terminated | January 18, 2011 | 88 | 960 | Not_Computer
Z-0397-2011 | 56048 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | BD Bioscience CD45 APC-H7 Analyte Specific Reagent, manufactured by BD Biosciences, San Jose, CA. |  |  | November 23, 2009 | 2009 | BD Biosciences, Systems &amp; Reagents | Incorrect expiration date: product with an expiration date of 03/31/2010 was marked with an incorrect expiration date of 06/30/2010. | Labeling mix-ups | A Product Field Action Notification letter, dated November 2009, was sent to customers who purchased the affected device. The letter identified the affected device and explained the reason for recall. Customers were asked to not use the current CD45 APC H-7 lot number 58633 past 2010-03-31 until they receive their new lot of CD45 APC H-7 at no cost to them. Customers were also asked to complete and return the attached fax form. Questions should be directed to BD Customer Support Center at 1-877-232-8995. |  | 87 | 87 units | Nationwide Distribution -- CA, WA, NJ, UT, TX and OH. | Terminated | December 03, 2010 | 375 | 87 | Not_Computer
Z-0528-2011 | 57178 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer, NexGen Complete Knee Solution LPS Femoral Component, Size G Left, sterile, REF 00-5996-017-01. |  |  | September 15, 2010 | 2010 | Zimmer Inc. | These lots of Zimmer NexGen Complete Knee Solution LPS Femoral Component and Zimmer NexGen Complete Knee Solution LPS Flex Gender Femoral Component exhibit a nonconforming internal CAM radius. The condition is intermittent and not all devices have nonconforming geometry. | Process design | The firm sent URGENT DEVICE CORRECTION AND REMOVAL notices to distributors and sales staff, Risk Managers, and implanting surgeons dated 9/15/2010. Sales staff and risk managers were instructed to return any product on hand. Surgeons were informed of the issue but there was no action to be taken by surgeons. |  | 1151232 | 192 all sizes of 5996 products | US, Germany, Spain, France, UK, Italy, Sweden, South Africa, Czech Republic, Bulgaria, Poland, Romania, and Austria. | Terminated | July 27, 2011 | 315 | 1151232 | Not_Computer
Z-0537-2011 | 56818 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | XELERIS (1.0, 1.1, 2.0 &amp; 2.1) PROCESSING AND REVIEW WORKSTATION |  |  | October 27, 2010 | 2010 | GE Healthcare, LLC | GE Healthcare has become aware of inaccurate user protocol setup associated with the Thyroid Uptake Index application on certain Xeleris workstations(1.0, 1.1, 2.0, 2.1) that may impact patient safety.Inaccurate protocol setup by the user while using Thyroid Uptake Index application has been identifiedon certain Xeleris workstations. Incorrect protocol setup during clinical data acquisition a | Other | GE Healthcare sent an "Urgent Medical Device Correction" letter dated October 27, 2010, to all consignees. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology and Radiologists/Cardiologists and Identifed the product, the problem and the action to be taken by the consignees. Consignees were instructed to replace the original protocol instructions provided with their system with the updated protocol instructions provided with the letter. For any questions consignees were to contact their local GE Healthcare Service Representative or contact the GE Healthcare Call Center:United States or Canada (800) 437-1171Other countries: (262) 896-2890 |  | 6020 | 6020 | All States in continental USA including DC and PR. Viet Nam, Venezuela, Uruguay, United Kingdom, Ukraine, Turkey, Thailand, Taiwan, Switzerland, Sweden, Spain, South Africa, Slovakia, Singapore, Saudi Arabia, Russian Federation, Republic of Korea, Portugal, Poland, Philippines, Panama, Pakistan, Norway, New Zealand, Netherlands, Morocco, Mexico, Martinique, Malaysia, Luxembourg, Kuwait, Kenya, Japan, Jamaica, Italy, Israel, Ireland, Indonesia, India, Hungary, Hong Kong, Guadeloupe, Greece, Germany, France, Finland, Ecuador, Dominican Republic, Cyprus, Colombia, China, Canada, Bulgaria, Brazil, Bosnia &amp; Herzegovina, Belgium, Azerbaijan, Austria, Australia, Argentina, Algeria. | Terminated | August 28, 2012 | 671 | 6020 | Not_Computer
Z-0538-2011 | 57111 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Elekta VBH Head FIX |  |  | October 14, 2010 | 2010 | Elekta, Inc. | The HeadFIX Baseplate is not screwed down to the adapter to allow to compensate for roll and pitch with the HeadFIX leveling screws of the HeadFIX Baseplate. | Other | Elekta, Inc. sent an Important Notice A343 dated October 14, 2010 to all affected consignees. The Notice identified the product, the problem, and the action to be taken by the consignees. Consignees were instructed constantly monitor the patient before and during the treatment. Check the patient's position after every patient movement (e.g. coughing). Advise all appropriate personnel of the content of the notice place a copy of the Notice In the Important Notice Section of the HeadFix User Manual P10103-108.For questions regarding this recallservice.schwabmuenchen@elekta.com |  | 34 | 34 units | Worldwide Distribution - USA including AZ, CA, CT, FL, GA, IL, KY, MI, MS, NY, NC, OK, OR, RI, TN, TX, VA, WA and the country of Canada | Terminated | December 28, 2010 | 75 | 34 | Not_Computer
Z-0540-2011 | 56047 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Autoimmune EIA ANA Screening Test Kit, 576 Tests, |  |  | March 01, 2010 | 2010 | Bio-Rad Laboratories Inc | The ANA Screening kit generates a high percentage of false positive results, associated with a low check control reading. | Other | Bio-Rad issued an Urgent Medical Device Recall Notification letter dated March 1, 2010 to customers describing the issue and actions to be taken. Responses will be tracked via response forms. The letter instructed customers that affected products should be identified, quarantined, and discarded, or reworked with shorter expiration dates until newer lots are available.Bio-Rad can be contacted at 510-724-7000 concerning this recall. |  |  |  | Worldwide Distribution: Throughout the US, and to the countries of: Australia, Canada, France, Hong Kong, India, Italy, Japan, New Zealand, Panama, Singapore, South Korea and the UK. | Terminated | January 05, 2011 | 310 |  | Not_Computer
Z-0545-2011 | 57069 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Allura XPER X-ray system. |  |  | May 26, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | The Printed Circuit Board (PCB) in the Main-cabinet. Further investigation revealed that a manufacturing error in the PCB is responsible for the reliability issue. As a result, the complete system may shut down and not re-start. | Other | Philips began mailing out an URGENT-Device Correction letter on May 26, 2009. Philips sent out a second URGENT-Device Correction letter dated June 4, 2009, to all consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were advised to contact their local Philips representative immediately if their system completely shut down and would not restart. A mandatory Field Change Order 72200144 was issued for replacement of the Printed Circuit Board (PCB) in the Allura XPER FD system. A mandatory Field Change Order FCO 70800091 was issued for replacement of the Printed Circuit Board (PCB) in the MultiDiagnost Eleva with FD system. The consignees were to be contacted by their Philips Representative for implementation of the corrective action. The firm's Field Engineer would replace the power tray of the systems.If the consignees needed any further information or support concerning this issue, they were to contact the call Center at 800-722-9377, # 5 and reference FCO 72200144 or FCO 70800091 with any questions. |  | 16 | 16 units | Nationwide Distribution including the states of AL, AR, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KY, LA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, OK, OR, PA, RI, TN, TX, VA, and WI. | Terminated | December 09, 2010 | 562 | 100 | Not_Computer
Z-0543-2011 | 57143 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Edwards Protection Cannulae - Monitoring and Infusion Set, Model MIS004, Sterile EO, Edwards Lifesciences LLC, One Edwards Way, Irvine, CA 92614. |  |  | October 19, 2010 | 2010 | Edwards Lifesciences Llc | Monitoring and Infusion Set connectors may break during manipulation. | Employee error | Edwards Lifesciences issued a Urgent Product Recall letter dated October 13, 2010 to customers identifying the affected devices, requesting the return of unused devices and completion and return of a response form.Further information is available at 800-424-3278. |  | 570 | 570 sets | Nationwide Distribution: in the states of FL, KS, MO, TX, VA, and PR. | Terminated | May 18, 2011 | 211 | 892 | Not_Computer
Z-0563-2011 | 57241 | No | JOR | Cardiopulmonary bypass pulsatile flow generator. | Cardiovascular | Cardiovascular | PMA | Sarns 8000 Roller pump for Sarns System 8000 Heart Lung Machine console, 220V, Catalog No 16407, Terumo Cardiovascular System Corp Ann Arbor, MI. (not distributed in the US) |  |  | November 02, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The system 8000 roller pumps may fail resulting in a persistent pump stop. | Other | The firm sent "URGENT MEDICAL DEVICE CORRECTION" dated 11/02/2010 to customers. The letter describes the product, problem and corrective actions taken and/or needed to be taken by the customers. The letter states the system responds to a failed inductor by displaying an overspeed message and stopping the roller pump without an audible alert. The pump cannot be restarted. The customers/users are instructed to stop the pump, replace the pump with a backup or if necessary use the handcrank. The letter states that Terumo will replace the roller pump at the next preventative maintenance visit - or sooner- if the customer has a service contract. If the customer does not have a service contract, Terumo will contact the customer within 2 weeks to set up an appointment for the service. Terumo CVS expects all affected inductors to be replaced to be replaced within three months. Call Terumo CVS Customer Service upon completion of the case at 800-521-2818. The customer was also instructed to complete and return the enclosed CUSTOMER RESPONSE FORMPlease note that we are not recommending that you stop using the roller pumps for your Sarns system 8000.For any questions, call toll free 800-262-3304 or phone 734-663-4145. |  | 95 | 95 devices both products | Worldwide distribution: USA including states of CA, FL, NC and NY and countries of United Arab Emirates, China, Ecuador, and Panama. | Terminated | October 19, 2011 | 351 | 95 | Not_Computer
Z-0561-2011 | 57064 | No | CBL | Carbon dioxide absorbent. | Anesthesiology | Anesthesiology | 510(K) Exempt | Grace SODASORB LF Carbon Dioxide Absorbent, Medical Grade (16%) Soda Lime USP-NF with an ethyl violet indicator which turns from white to violet as absorbent is depleted or upon moisture loss; Manufactured by: W.R. Grace &amp; Co., 6050 West 51st St., Chicago, IL 60638; the product is packaged in pre-paks pre-filled 2.5 lb. canisters, 12 per carton |  |  | October 20, 2010 | 2010 | Darex Container Products Div of W.R. Grace &amp; Co. | Pre-Pak units of SODASORB CO2 absorbent manufactured prior to August 2010 may have missing or illegible expiration dates and/or batch numbers. | Packaging process control | Darex/W.R. Grace sent a Customer Notification to their sole direct account of the recall and the account provided Darex/W.R. Grace with lists of their customers who received the recalled products. Darex then sent undated recall letters to those accounts on 10/20-22/10. The accounts were informed of the product, the problem, and the action that should be taken. The accounts were warned not to use Pre-Pak units with missing or illegible expiration dates and/or batch numbers and were requested to examine their inventories of Pre-Pak units, placing all units that had missing or illegible expiration dates and/or batch numbers on hold for destruction and replacement. The accounts were instructed to call W.R. Grace at 1-800-492-7632 to arrange for replacement. The accounts were also requested to complete and return the reply postcard, indicating how much product was destroyed. |  | 8096 | 8,096 cartons | Nationwide Distribution | Terminated | September 21, 2011 | 336 | 103638 | Not_Computer
Z-0557-2011 | 57176 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Custom Kit, Sterile EO, Rx Only, Catalog No. K09-09815, Manufacturer: Merit Medical Systems, Inc. South Jordan, Utah 84095, USA, |  |  | October 25, 2010 | 2010 | Merit Medical Systems, Inc. | Defective custom kits may draw in air during use which may compromise patient safety and affect treatment. The valve's fixed luer may be over- or under-tightened, causing leaking at the bond joint. | Process control | Merit issued an Urgent Product Recall Notice letter dated October 25, 2010 to consignees. Copnsignees were instructed to quarantine the product immediately for return and to notify customers of recall activity and instruct them to return product.. All distributors and customers are to complete and return a Recall Notification Form.For further information about this recall, contact Merit at 800 356-3748. |  | 1033 | 1033 units | Nationwide Distribution in the states of CA, IL, IN, and UT. | Terminated | March 23, 2012 | 515 | 6943 | Not_Computer
Z-0584-2011 | 56250 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Active I-PTH ELISA, Part Number: DSL-10-8000 |  |  | April 22, 2010 | 2010 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that the stated stability of the DSL-10-8000 I-PTH ELISA and DSL-8000 I-PTH IRMA assays may not be achieved for the lot numbers affected by the recall. The affected I-PTH ELISA and the I-PTH IRMA kit lots may produce falsely elevated patient results. Internal stability testing showed an average 16% elevation for I-PTH ELISA lot 99 | Pending | Beckman Coulter sent an Urgent Product Corrective Action Letter to their consignees dated June 18, 2010. The letter identified the product, the problem, and the action to be taken. Consignees were instructed to:(1) Discontinue use of all materials provided with I-PTH ELISA kit lot number 992751 and I-PTH IRMA lot numbers 090013, 090121, and 090372.(2) Review historical external quality control results to ensure the assay performance of these lots was within laboratory specifications. Retesting may be necessary if their laboratory uses only the Control materials provided with the I-PTH ELISA and IRMA kits.(3) Abnormal elevated patient results that are obtained with these lot numbers and that are inconsistent with the patient's clinical status should be reviewed in the context of other clinical diagnostic tests.(4) I-PTH ELISA kit lot number 992751 and I-PTH IRMA lot numbers 090013, 090121, and 090372 will no longer be shipped. Due to the calibrator instability recently encountered with DSL-10-8000 I-PTH ELISA and DSL-8000 I-PTH IRMA assays, Beckman Coulter has decided to immediately discontinue these two part numbers.Consignees were asked to share this information with their laboratory staff and retain the notification as part of their laboratory Quality System documentation. Complete and return the enclosed response form within 10 days so Beckman Coulter can be assured that they received this important notification.If they have any unused product the were asked to contact the numbers listed below:(1) In the United States, contact Customer Service at 1-800-526-382.1(2) Outside of the United States, consignees were asked to contact their local Beckman Coulter RepresentativeConsignees were told to contact DSL Technical Support Center at 1-800-231-7970 in the United States, or their local Beckman Coulter Representative with any questions.For questions regarding this recall call (714) 961-3634. |  | 3 | 3 kits | Worldwide Distribution - USA including MD, NY, GA, NJ, CA, MI, TX, PA, IA, and IN. Internationally to Austria, Brazil, China, Colombia, Czech Republic, France, Germany, Hong Kong, India, Japan, Mexico, and the Russian Federation | Terminated | May 07, 2012 | 746 | 199 | Not_Computer
Z-0574-2011 | 56731 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ZimmerÂż Trabecular Metal Humeral Stem Surgical Technique. Catalog Number: 97-4309-002-00, Rev. 3 |  |  | September 10, 2010 | 2010 | Zimmer Inc. | The results of an investigation into complaints revealed that Trabecular Metal Humeral and/or ZimmerÂż Trabecular Metal Reverse Humeral implants engaging in the canal earlier than expected could lead to incomplete seating of the implant, delay in surgery, humeral fracture, or the need for cement in a press-fit application. | Device Design | Zimmer, Inc. sent an URGENT DEVICE CORRECTION letter dated September 10, 2010 to surgeons, risk managers, and distributors. The letter identified the product, the problem, and the action to be taken. Sales Representatives were instructed to:a. provide the Dear Surgeon (Urgent Device Correction) letter and revisedsurgical technique directly to surgeons and risk managers at facilities usingthe device orb. provide a list of facilities and surgeons using the TM Humeral and TMReverse shoulder systems to the Zimmer Recall Coordinator, andc. destroy previous versions of the surgical technique, andd. inform users of when the new surgical technique, instrument case, distalreamers, and distal pilots will be shipped to them based on communicationfrom Zimmer, Inc.Handling of returned product:HCPs were notified to return all of the distal pilots in the scope back toZimmer which were to be be destroyed once the recall has been completed. They willalso be instructed to either return the reamers back to Zimmer or to remove these reamers from the TM Reverse shoulder systems sets and place into theBigliani/Flatow Complete Shoulder Solution set. New surgical techniques will also be introduced and the user will be asked to remove previous versions of the surgical technique from all points of use. This removal will be verified byZimmer Sales Representatives.For questions regarding this recall call (574) 371-8852. |  |  |  | Worldwide Distribution - USA, Nationwide, Switzerland and Eschbach, Germany, Switzerland, Australia, China, India, Japan, Korea, Malaysia, Singapore, Taiwan, Brazil, Venezuela, El Salvador, Colombia, Canada, and Thailand. | Terminated | September 14, 2012 | 735 |  | Not_Computer
Z-0582-2011 | 57157 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | McKesson Medical Surgical 107-101 Urine Analyzer, Catalog Number: 107-101 |  |  | October 01, 2010 | 2010 | Teco Diagnostics | This recall has been initiated due to the CLIA-101 Urine Analyzer not having 510(k) clearance by the FDA for use in measuring urinary glucose or occult blood with the CLIA waived, Inc. CLIA-10 test strips. Teco Diagnostics has been notified bythe FDA that the CLIA-101 requires 510(k) clearance and is not CLIA Waived for use with the CLIA-10 test strips. Without FDA approval or clearance, Teco Di | Other | Teco Diagnostics initiated an "Urgent: Medical Device Recall" letter dated October 14, 2010, with an attached "Recall Return Response Form" to the affected consignees via US Postal on October 15, 2010. The recall notice informed the consignees of the reason for recall, the products affected and the actions required.Consignees were instructed to complete and return the enclosed Recall Return Response Form as soon as possible via mail, fax, or email. Upon receipt of the Recall Return ResponseForm, Teco Diagnostics would work with the consignees/distributors to process any refunds or reimbursements for the returned product and return shipping. If consignees/distributor had questions, they were to call Teco's Technical Support Department at (800) 222-9880 or (714) 463-1111. |  | 2140 | 2,140 units total for all products | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NM, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI, AND WV and the country of Canada. | Terminated | August 03, 2012 | 672 | 2140 | Not_Computer
Z-0587-2011 | 57089 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter Three-Way Large Bore Stopcock with Male Luer Slip; a sterile, nonpyrogenic fluid pathway; Baxter Healthcare Corporation, Deerfield, IL 60015; product code 2C6202. |  |  | October 28, 2010 | 2010 | Baxter Healthcare Corp. | Rotating the handle on the 3-Way Large Bore Stopcock past the stop can cause the stopcock to leak. A leak can lead to under-infusion or interruption of critical life-sustaining therapy. | Device Design | Safety Alert letters dated October 28, 2010 were mailed via first class mail to the direct accounts, to the attention of the Director of Nursing and the Director of Materials Management, on the same date, informing them that rotating the handle on the 3-Way Large Bore Stopcock past the stop can cause the stopcock to leak. A leak can lead to under-infusion or interruption of critical life-sustaining therapy. The accounts were requested to post the enclosed attachment, which illustrates the correct direction to turn the handle of the stopcock, where the stopcocks/sets are dispensed. The accounts were also requested to complete the attached customer reply form confirming their receipt of the letter and fax it to Baxter at (847) 270-5457. Dealers, wholesalers and distributors were requested to notify their customers of this action so they can comply with the procedures listed above. The Baxter Field Corrective Action Gatekeepers were also notified of this action via e-mail on 10/27/10, and instructed to notify their customers and ministries of health of the recall.Any questions regarding the communication were directed to The Center for One Baxter at 1-800-422-9837 during the hours of 8:00 am to 5:00 pm CST. |  |  |  | Worldwide distribution: USA and countries including Puerto Rico, Guam and American Samoa, and internationally to Canada, Bermuda, Jamaica, Guatemala, Panama, Bahrain, United Kingdom, Hungary, Lebanon, Pakistan, Saudi Arabia and New Zealand. | Terminated | November 01, 2012 | 735 |  | Not_Computer
Z-0569-2011 | 57315 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | MONOJECT Prefill Heparin Lock Flush Syringe |  |  | November 15, 2010 | 2010 | Covidien LP | The heparin sodium USP used to manufacture was produced with crude heparin containing Over-Sulfated Chondroitin Sulfate. | Material/Component Contamination | COVIDEN sent an URGENT PRODUCT RECALL letter dated November 12, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and in-use stock to determine if they had any units of the affected product and quarantine the product immediately. Distributors were instructed to notify their customers immediately of the URGENT PRODUCT RECALL letter. Customers were to fill out the enclosed response form and return it by fax to (203) 822-6009 or email to adfeedback@coviden.com following the instructions on the form. For questions customers were to contact Customer Service at 1-800-962-9888 |  | 1539540 | 1,539,540 | Nationwide Distribution including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, LA, KS, KY, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and Bermuda | Terminated | December 04, 2015 | 1845 | 3510660 | Not_Computer
Z-0551-2011 | 57045 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Dixtal Handheld Pulse Oximetry Model 512 |  |  | September 27, 2010 | 2010 | Dixtal Medical, Inc. | Unit may overestimate Sp02 levels in a high humidity environment. | Equipment maintenance | Dixtal Medical notified accounts by email on 9/27/10 and followed by hard copy Urgent Product Recall letter, dated September 27, 2010. The letter identified the affected product and stated the reason for the recall. The letter also requested that customers to cease distribution and/or use of the suspect product and to return product for replacement. The attached Confirmation Notification should be completed and returned. Customers are to contact Customer Service: sales@dixtal.com, Phone 1-203-269-1112, Fax 1-203-269-1760, with questions or to arrange to have a replacement product provided. |  | 4 | 4 units | Worldwide Distribution -- USA, including states of AK, KY, NJ, and NY and countries of Australia, Austria, Bolivia, Canada, Finland, Indonesia, Italy, Poland, Sweden, and Switzerland. | Open, Classified |  |  | 39 | Not_Computer
Z-0589-2011 | 55977 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | GE Responder 2000 Defibrillator/Monitor, models 2025653, 2026109, and 2026114 that were manufactured between July 25, 2008 and October 2, 2009 by GE Healthcare. |  |  | April 28, 2010 | 2010 | Cardiac Science Corporation | Some units may be non-operational during a rescue, due to a defective PCB interconnect on traces connected to the reed relay. This will result in an inability to analyze the patient's heart rhythm or deliver therapy during a rescue. | Nonconforming Material/Component | Cardiac Science issued a Urgent Medical Device Corrective Action letter dated April 28, 2010 requesting GE Healthcare cease distribution of affected devices and to instruct end users of distributed devices to schedule a service call to have the device corrected. Cardiac will send rework kits no later than June 30, 2010. CSC can be contacted at 44 161 926 0011 to arrange return material authorization and replacement of affected devices. |  | 130 | 130 units | Worldwide Distribution: Throughout the US, and in the countries of Germany, Korea, Peru, Singapore, and Spain. | Terminated | December 08, 2011 | 589 | 130 | Hardware | Treatment/Delivery/Therapy | Replace | Remove or Replace
Z-0602-2011 | 54143 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Spinal Concepts polyaxial screw driver, AO, Manufactured by Spinal Concepts, Austin, TX, REF 775-01. |  |  | December 01, 2005 | 2005 | Zimmer Inc. | The devices may fracture, resulting in surgical delay or the need for intervention to remove fragments from the surgical site. | Device Design | The affected instruments were removed and replaced in the U.S. Abbott Spine and Spinal Concepts instrument kits between December 2005 and October 2007. User accounts were not advised of the removal Zimmer has confirmed that none of these instruments are currently available for use in the U.S.For further information about this action, contact Dale Miller of Zimmer at 574 372-4962. |  | 220 | 220 | Worldwide Distribution, including US. | Terminated | June 01, 2011 | 2008 | 273 | Not_Computer
Z-0588-2011 | 55571 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension LIP Flex Reagent Cartridge, Catalog# DF55A |  |  | March 10, 2010 | 2010 | Siemens Healthcare Diagnostics, Inc. | Open Well instability: After calibration, QC shifts high (out of range) after the Flex Well is hydrated. Shifts of up to 20% have been observed in both QC and patient samplesas soon as two hours post hydration. | Other | Recalling firm issued "Urgent Field Safety Notice", dated March 2010, informing users of the problem and to transition to the revised Lipasae method. |  | 21171 | 21171 | Nationwide Distribution throughout the US. | Terminated | August 04, 2011 | 512 | 21171 | Not_Computer
Z-0594-2011 | 56832 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FUSO Analyzer 300F |  |  | September 20, 2010 | 2010 | Abbott Point Of Care Inc. | i-STAT Analyzers batteries may overheat due to a circuitry failure, resulting in the analyzer becoming uncomfortably hot in the area of the battery component. | Other | Abbott Point of Care issued Urgent Recall Notice letters dated September 20, 2010 to customers. The letter informs users to of the possibility of the unit becoming hot to touch, to remove the unit if this is the case, not to remove batteries unless instructed to do by APOC and to forward notification of the recall to customers if it has been further distributed.Abbott Point of Care can be contacted at 800 366-8020, option 1, about this action.Recall extended in May 2011 to included distribution dates of 9/20/2010-6/6/2011. |  | 2202 | 2202 | Worldwide Distrbution: Throughout USA and the countries of Argentina, Australia, Austria, Belgium.Bolivia, Bosnia-Herz, Canada. Croatia, Chile, Colombia, Costa Rica, Denmark, Ecuador, El Salvador, Finland, Georgian Republic, Greece, Honduras, France, Germany, Italy, Netherlands, New Zealand, Hong Kong, India, Ireland, Israel, Japan, Kingdom of Saudi Arabia, Korea, Kuwait, Mauritius, Mexico, Nigeria, Norway, Panama, Paraguay, Peoples Republic of China, Peru, Poland, Portugal, Qatar, Russia, Singapore, South Africa, Spain, Sweden, UK, Sultanate of Oman, Switzerland, Taiwan, R.O.C., Thailand, Turkey, UAE, Uruguay, and Yemen. | Terminated | January 28, 2013 | 861 | 1295561 | Battery
Z-0592-2011 | 56864 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microtek Medical, Inc., Microscope Drape for Zeiss Standard |  |  | August 23, 2010 | 2010 | Microtek Medical, Inc. | This product is being recalled due the product is miss-labeled as being sterilized by EO processing; however, the product was processed using Gamma sterilization. | Error in labeling | On 08/23/10 Microtek Medical, Inc., Jacksonville, Florida notified all of their customers via a telephone call and requested immediate return of the product . The firm followed up with each of the customers by sending out an Urgent: Medical Device Recall letter (written notice), dated September 02,2010, to each on 9/8/10. The letter identified the affected product, stated the reason for the recall, and asked that the product be returned immediately. Customers are to complete and return the enclosed response form. Anyone needing further assistance is asked to contact Elizabeth Jolly, V.P. Legal at 1-888-844-0988. |  | 60 | 5 Cases, 12 drapes per case, 60 Drapes total distrubuted | Nationwide Distribution | Terminated | December 14, 2010 | 113 | 60 | Not_Computer
Z-0600-2011 | 57096 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Medefil Heparin I.V. Flush Syringe 10 units/mL; a polypropylene luer locking syringe containing Heparin Lock Flush Solution, USP, a sterile non-pyrogenic, isotonic injection; each mL in a Heparin I.V. Flush Syringe contains 10 units of heparin sodium derived from porcine intestinal mucosa and 9 mg of sodium chloride in water for injection; Medefil, Inc. Glendale Heights, IL 60139. The pre-filled syringes were packed as 60 syringes per dispenser box, 16 boxes per shipping carton in the following syringe configurations: a) NDC# 64253-222-21, catalog # MIH-2221: 1 mL fill in 6 mL syringe; b) NDC# 64253-222-23, catalog # MIH-2223: 3 mL fill in 6 mL syringe; c) NDC# 64253-222-35, catalog # MIH-2235: 5 mL fill in 12 mL syringe |  |  | November 01, 2010 | 2010 | Medefil Incorporated | The heparin lock flush solution was manufactured with contaminated Heparin Sodium USP Active Pharmaceutical Ingredient (API). | Material/Component Contamination | Medefil faxed/emailed the direct accounts on 11/1/10, and sent the accounts a hard copy of the Urgent Product Recall letter, dated 11/1/2010, on 11/2/10. The accounts were informed of Medefil's recall of 18 lots of Heparin I.V. Flush Syringes produced from the active pharmaceutical ingredient (API) Heparin Sodium, USP, lot 1035-0780, that has been recalled by the manufacturer, Scientific Protein Laboratories, due to the presence of a contaminant, Oversulfated Chondroitin Sulfate. The accounts were instructed to stop use of the 18 lots of Heparin I.V. Flush Syringes immediately, place them under quarantine, and contact Medefil to arrange for the product to be returned. The accounts were requested to complete and return to Medefil the enclosed response form, indicating the quantity on hand of each recalled lot. Any questions were directed to Pradeep Aggarwal, President at 630-682-4600. |  | 583500 | 583,500 syringes | Nationwide Distribution --- including states of Arizona, California, Florida, Illinois, Ohio, Louisiana, Massachusetts, Michigan, Minnesota, New York, Texas, and Virginia. | Terminated | February 13, 2012 | 469 | 2919180 | Not_Computer
Z-0590-2011 | 57011 | No | LWE |  |  | General Hospital | Enforcement Discretion | Orthofix Bone Marrow Aspiration Needle, Catalog number 21-5000, 21-5011, Distributed By: Orthofix Inc, 1720 Bray Central Drive, McKinney, TX 75068-8207. Product includes handle, stylet, cannula, probe, probe guide, and universal cap. Cannula is available as an 8 gauge needle or 11 gauge needle. |  |  | October 18, 2010 | 2010 | Orthofix Inc | Outer cartons for the product may be labeled with incorrect needle gauge. | Other | Orthofix issued an Urgent Product Recall letter dated October 14, 2010 to consignees. Customers were instructed to inspect their inventory, and remove and return affected product to the firm. Replacement needles will be provided. Orthofix account representatives can be contacted at 1 888 298 5700, and Orthofix Regulatory Affairs can be contacted at 1 214 937 2061. |  | 274 | 274 | Worldwide Distribution - USA, including the states of AL, AR, CA, CT, FL, LA, MI, MO, NC, NJ, NY, PA, SC, and TX, and the country of Spain. | Terminated | February 28, 2011 | 133 | 274 | Not_Computer
Z-0591-2011 | 57216 | No | FTD | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CHROMOPHARE Surgical Lighting System &amp; CHROMOVIEW Flat Panel Monitor, Mechanical Spring Arms. |  |  | November 16, 2009 | 2009 | Berchtold Corp. | A mechanical failure of a welded joint within the spring Arm may result in the rapid drop of a monitor to the mechanical stop. | Other | Berchtold issued a Product Safety Notification letter on November 16, 2009 identifying the affected device and issue prompting the action. Consignees were instructed to:Exercise caution when adjusting the position of the surgical lights or flat panel monitors. While an event is highly unlikely, the likelihood does increase when the spring arm is being repositioned. Identify a specific contact person for their facility to facilitate the communication and field action process with BERCHTOLD. The following action is being planned: To further identify if their units are part of the affected population, determine the supplier serial number identification and specific location for each spring arm in their facility. Report the facility information, supplier spring arm serial numbers and specific location or spring arms to a BERCHTOLD. Detail instructions on serial number identification and the method of reporting to BERCHTOLD will also be provided in a subsequent communication. Once BERCHTOLD has received the serial numbers of their spring arm assembly, customers will be notified whether their particular units are affected and the further required action.Berchtold Corporation can be contacted at 1 800 968 6097 about this action. |  | 7031 | 7031 arms | Worldwide Distribution: Throught the U.S,.and the countries of Canada and Germany | Terminated | July 06, 2012 | 963 | 7031 | Not_Computer
Z-0614-2011 | 51242 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Heartstring II Proximal Seal / VASOVIEW 6 Endoscopic Vessel Harvesting System, |  |  | February 05, 2009 | 2009 | CTS, Inc. dba Guidant Cardiac Surgery | Sterility may be compromised by a loss of package integrity. | Packaging | Maquet issued an Urgent Device Removal letter dated February 5, 2009 to customers, explaining the reason for the recall and requesting the product be returned for replacement. Customers were instructed to examine their inventory and discontinue distribution and use of the affected devices. Customers also were requested to complete and return a Field Action Response Form.Maquet can be contacted at 888-880-2874 concerning this recall. |  |  |  | Distribution: Nationwide and Internationally. | Terminated | January 19, 2011 | 713 |  | Not_Computer
Z-0608-2011 | 55815 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Quest EIA ANA Screen Bulk Kit, model number 96AN-BPU-QUEST. Manufactured by Bio-Rad Laboratories, Inc., Hercules, CA. For labeling and distribution by Quest International Inc. |  |  | February 26, 2010 | 2010 | Bio-Rad Laboratories Inc | The value for the ANA Cutoff and ANA Positive Controls will decrease overtime resulting in an unusually elevated positivity rate for patient samples tested. | Other | Bio-Rad Laboratories, Inc. sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letter dated February 24, 2010, to their customer. The letter identified the product, the problem, and the action to be taken. The customer was instructed to:Examine inventory and quarantine affected product. Identify all customers that have received the affected product and immediately notify them. Instruct them to discontinue use of the affected product.Complete and return the enclosed MEDICAL DEVICE RECALL RESPONSE FORM.For questions regarding this recall call (510) 724-7000. |  | 453 | 453 kits | Nationwide Distribution including FL | Terminated | January 05, 2011 | 313 | 453 | Not_Computer
Z-0609-2011 | 55363 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | NexGen Complete Knee Solution Stemmed tibial broach impactor, Catalog # 00-5977-011-00, Zimmer Inc., Warsaw, IN. |  |  | April 16, 2010 | 2010 | Zimmer Inc. | The impactor may disassemble or detach during use. The detached piece may not be observed intraoperatively and may be left in the surgical site. Complications may include pain, infection, and prolongation in surgery. These complications may also require revision surgery. | Device Design | Zimmer issued "Urgent Medical Device Safety Alert" letters dated April 16, 2010 to Risk Managers, user facilities, Health Care Professionals and Zimmer Sales force. The users were informed to inspect the device for loosening. The alert also informed users to inspect the device after each use to verify that the device was intact. Customers were instructed to complete and return a User Facility Documentation Form. Zimmer will repair and replace all affected devices.Customers can contact ZImmer, Inc. at 800 613-6131 concerning this recall. |  | 10666 | 10,666 all products. | Worldwide Distribution: Throughout US including Puerto Rico, and the countries of Argentina, Australia, Belgium, Bolivia, Brazil, Canada, Chile, China, Columbia, Denmark, Ecuador, El Salvador, France, Germany, Guatemala, Hong Kong, India, Italy, Japan, Korea, Mexico, Netherlands, Nicaragua, Panama, Peru, Portugal, Singapore, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine, and United Kingdom. | Terminated | September 12, 2012 | 880 | 10666 | Not_Computer
Z-0613-2011 | 56386 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio-Rad brand Meridian ANA Screen ELISA Test Kit, 96 Test, |  |  | February 24, 2010 | 2010 | Bio-Rad Laboratories Inc | False positive test results caused by decrease in the OD values of ANA Cutoff Controls and ANA Positive Controls. The increased positivity rates were associated with the Meridian ANA Screening kit. False positive ANA values could result in misdiagnosis or mistreatment . | Other | Bio-Rad issued an Urgent Medical Device Recall NOtification letter dated February 24, 2010 to their distributor, identifying the affected product and actions to be taken. The distributor was instructed to examine inventory and quarantive affected kit lots, and to identify and notify customers with affected lots to discontinue use and discard the product. Replacement kits will be distributed. Customers were asked to complete and return a recall response form.Bio-Rad can be contacted at 510 724-7000 concerning this recall. |  |  |  | Nationwide Distribution: to Meridian Diagnostics in Ohio. | Terminated | January 10, 2011 | 320 |  | Not_Computer
Z-0607-2011 | 56752 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems, Part Number: 81600 ; Access 2 Immunoassay Systems, Part Number: 81600N; |  |  | August 13, 2010 | 2010 | Beckman Coulter Inc. | This recall was initiated because Beckman Coulter has received reports of erroneous test results being reportedas a result of assay reagent packs being transferred between systems. Erroneous results will be generated as a result of missing or incorrect assay reagents. In these casesboth false positive and false negative results are possible, depending upon the assay configuration. | Pending | Beckman Coulter, Inc. sent two (2) Product Corrective Action (PCA) letters (Access letter and DxI letter) with attached Customer Response form the week of August 09, 2010 to the affected customers. The letters were dated August 13, 2010, and identified the product, the problem, and the action to be taken by the customers. The letter reminded customers that it is critical that they load reagent packs according to the procedure provided in their Access or Access 2Operators Guide, Instructions for Use, or Help system. A copy of the procedure was included with the letter. Customers were reminded that all operators should be properly trained on these important procedures andunderstand the impact on reported results if these procedures are not followed.Customers were instructed to share this information with their laboratory staff, and retain this notification as part of their laboratory Quality System documentation. Customers who needed assistance or had any questions regarding this notification, were instructed to contact Technical Supportat 1-800-854-3633, option 1 in the United States or Canada. |  | 9571 | 9571 units (instruments) | Nationwide Distribution &amp; Canada | Terminated | July 11, 2012 | 698 | 9571 | Not_Computer
Z-0611-2011 | 56849 | No | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | RotoProne Therapy System, Model 209500, Device List Number E232632 |  |  | September 09, 2010 | 2010 | KCI USA, Inc. | Product complaints indicating that while attempting to resolve an alarm or error condition, operator was unable to return product to the prone position for patient. | Other | KCI USA, Inc. sent an URGENT - Medical Device Safety Alert letter, dated September 8, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the cusotmer. Customers were instructed to have a practiced back-up plan, such as a plan for the manual proning of patients in the event that clinical circumstances dictate the need for more immediate prone therapy. For questions, customers in the US should call KCI's Technical Support Center at 1-800-275-4524, option 5. For customers outside the US they should contact their local sales representative. Customers should contact their accounts with affected product and advise them of the safety notice. |  | 240 | 240 units | Worldwide Distribution - USA Nationwide, United Kingdom, Germany, and Saudi Arabia | Terminated | March 21, 2011 | 193 | 240 | Not_Computer
Z-0615-2011 | 57214 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Incorrect Label: Endocare 2.4mm VARIABLE ATTACHED CRYOPROBE, REF: CVA 2400, LOT 2751. Correct Label: Endocare 2.4mm CRYOPROBE, REF: CRYO-44F (Sharp Tip). |  |  | November 04, 2010 | 2010 | Endocare, Division Of Healthtronics | Mislabeled : The label on the outer box and inner pouch of the EndoCare 2.4mm CryoProbe (CRYO-44F) inaccurately reference the CVA2400 (2.4mm Variable Probe aka: V-Probe). | Other | Healthronics issued an Urgent Device Recall letter dated November 2, 2010 to consignees. The letter described the product mislabeling, and requested customers to examine their inventory, quarantine all affected product, and complete and return the verfication section of the letter to the firm. A return authorization number and prepaid return label were provide to facilitate return and replacement.Healthronics can be contacted at 877 367-4873 concerning this recall. |  | 33 | 33 Units | Nationwide Distribution: to the states of FL, MI, PA, TN, and WA. | Terminated | April 14, 2011 | 161 | 33 | Not_Computer
Z-0612-2011 | 56126 | No | LKJ | Antinuclear antibody immunological test system. | Immunology | Immunology | 510(k) | Bio-Rad Laboratories Autoimmune EIA ANA Screening kit, 96 test, model 96AN, Distributed and Manufactured by Bio-Rad Laboratories, Inc. |  |  | June 24, 2010 | 2010 | Bio-Rad Laboratories Inc | Product mislabeled-- the ANA cutoff Control and Prostiva controls were mislabeled. Use of the incorrect controls will cause run failure - patient results cannot be reported because controls will not pass.. | Labeling mix-ups | Bio-Rad issued Medical Device Recall Notification letters dated June 24, 2010 to consignees, identifying the affected products and requested that all affected product be disposed of in accordance with local procedures. Response forms are to be completed and returned to the firm. Bio-Rad can be contacted for replacement product and about this recall at 510 741-6114. |  | 37 | 37 units | Nationwide Distribution: including the states of California, Ohio, Massachusetts, and Missouri.Product delivered to 11 consignees only in the US. | Terminated | August 08, 2011 | 410 | 37 | Not_Computer
Z-0618-2011 | 53182 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Titanium Intraline Anchor, Model Numbers 3910400050, 3910400055, 3910400060, 3910400065, Soft Eyelet Anchor, manufactured by Stryker Endoscopy, Puerto Rico |  |  | August 20, 2009 | 2009 | Stryker Endoscopy | Potential breakage -- product may break while inserting it into the bone during surgery. This may prolong surgery, or, if undetected during surgery, result in retained fragments that may cause metal sensitization or immune response, or necessitate further intervention to retrieve the titanium fragments. | Material/Component Contamination | Stryker issued an Urgent Device Removal notification letter dated August 21, 2009 to consignees. Consignees were asked to acknowledge receipt and distribute the letter, track reponses, and destroy any returned material and document through weekly summary and records.Stryker can be contacted at 408 754-2000 about this recall. |  | 447 | 447 units | Nationwide Distribution: to one distributor in NJ. | Terminated | December 14, 2010 | 481 | 447 | Not_Computer
Z-0619-2011 | 57304 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Inter-Op Acetabular System Hooded Trial Insert size 32mm/57mm, non-sterile, REF 9363-32-057, Zimmer, Warsaw, IN. |  |  | November 15, 2010 | 2010 | Zimmer Inc. | Mislabeled - Device packages labeled as hooded trial inserts contain standard trial inserts. | Employee error | Zimmer contacted the two affected accounts directly by phone on November 15, 2010 to inform them of the recall. Zimmer will monitor the return of the affected instruments and account for the outstanding instruments until the removal action is complete. Each of the affected accounts will be provided with a new trial insert, manufactured and packaged with the correct component. The accounts will be asked to return the affected units back to Zimmer where the returned product will be destroyed. Customers can contact Zimmer at 574 371-8852 concerning this recall. |  | 3 | 3 | Worldwide Distribution: US in the states of KS and UT, and the country of Canada. | Terminated | May 23, 2011 | 189 | 3 | Not_Computer
Z-0620-2011 | 57370 | No | DXJ | Medical cathode-ray tube display. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Flat Panel Display is a touchscreen display, 19-inch thin film transistor liquid crystal display (TFT-LCD). The PNs are 010-1619-00 and 010-1655-00, Models 94260-19 and 94256. |  |  | March 20, 2009 | 2009 | Spacelabs Healthcare, Llc | The power supply of the Spacelabs Medical Flat Panel Display may fail, causing the display to go blank. | Other | Spacelabs Healthcare sent an URGENT: MEDICAL DEVICE RECALL Spacelabs Medical Flat Panel Power Supply PNs: 010-1619-00 &amp; 010-1655-00 letter dated March 20, 2009 to US consignees. On March 27, 2009, Spacelabs sent the same recall letter with translation option to their international consignees.The letter stated the firm would be sending a replacement power supply. While waiting for the replacement of the power supply, Spacelabs Medical recommended their consignees brief the staff on the possibility of the display going blank to help prevent any adverse event from happening. The consignees were advised to dispose the old power supply per their premise protocol or local regulations. The consignees from US can contact Spacelabs Medical at 1-800-522-7025, #2 for Technical Support , and consignees outside the United States can call 1-425-657-7200 x5089 for any questions about the recall. |  | 3608 | 2877 units in US and 731 units outside US | Worldwide Distribution: USA in the following states: AK, AL, AR, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, PA, RI, SC, SD, TN, TX, VA, WA, WI, and WY, and including the following countries: AUSTRALIA, AUSTRIA, CANADA, Caribbean, CZECH REPUBLIC, EGYPT, ESTONIA, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, ISRAEL, ITALY, JORDAN, KUWAIT, LEBANON, MEXICO, NETHERLANDS, PAKISTAN, PANAMA, POLAND, PUERTO RICO, SAUDI ARABIA, SINGAPORE, SWITZERLAND, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED KINGDOM, and VENEZUELA. | Terminated | June 13, 2011 | 815 | 3608 | Battery | Display/Image | Replace power supply | Remove or Replace
Z-0617-2011 | 57223 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The PercuNav MP008R-B Button Probe is designed to be used with Philips (formerly Traxtal) PercuNav Image Guided Intervention System. Some of the clinical applications of the Button Probe require sterility and others may not. However, it is possible that patient contamination could occur if a non-sterile Button Probe is used in clinical applications where aseptic technique is required. |  |  | November 10, 2010 | 2010 | Philips Healthcare Inc. | The protective cover over the tip of the MP008R-B Button Probe may fall off, causing the probe tip to damage the packaging and compromise the sterility of the probe. | Other | On 11/10/10, Philips issued URGENT - Medical Device Recall notification letters dated November 10, 2010 to their consignees.Consignees are advised to examine their inventory of MP008R-B Button Probes to confirm the probe tip cover is properly placed over the tip of the probe. They should discard any Button Probes that have been confirmed to have the protective probe tip cover removed from the Button Probe tip. Consignees are also reminded to follow the MP008R-B Button Probe and PercuNav Instructions for Use prior to using the device. Do not use the probe if the unit package is damaged, broken or soiled.The firm's representative will contact the consignees to confirm quantities and Lot Codes of MP008R-B Button Probes that the firm may have examined or discarded. The firm estimates replacement probes will be available in the first quarter of 2011.Consignees can contact Philips Customer Service at 800 722-9377 and reference the FCO 79500189. |  | 25 | 25 units | Nationwide Distribution: in the states of FL, IL, MA, NE, NY, SC, and VA. There are no foreign consignees and US government consignees. | Terminated | August 15, 2011 | 278 | 25 | Not_Computer
Z-0621-2011 | 57269 | No | KGO | Non-powdered surgeon's glove. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ENCORE (R) Microptic (R) Powder-Free Latex Surgical Gloves, 1 Pair, Size 5Âż, Sterile R, CE 0086, EC REP Ansell (UK) Ltd., Tamworth, Staffordshire, B79 7UL UK, Ansell Lanka (pvt) Ltd., Biyagama Export Processing Zone, Biyagama, Sri Lanka |  |  | September 03, 2010 | 2010 | Ansell Healthcare Products LLC | open package seams on sterile gloves | Other | The firm, Ansell, sent a letter dated September 3, 2010 to all customers. The letter described the product, problem and actions taken. An Ansell representatives will contact the customer to remove the gloves bearing the noted lot number from their facility; Ansell will ensure that the gloves removed will be replaced shortly with products from a different manufacturing batch and the recall was extended to the retail (hospital and clinic) level requesting that the product be returned. If you have any questions, contact US (334) 615-2563, international telephone 01.34.24.52.52 and email: infofrance@eu.ansell.com. |  | 4000 | 4000 units | International distribution only: France and Germany | Terminated | November 28, 2011 | 451 | 4000 | Not_Computer
Z-0678-2011 | 55861 | No | JEY | Bone plate. | Dental | Dental | 510(k) | 2.8 mm Inion CPS, 2.8 x 18 mm screw, REF SCR-1216, sterile, Manufacturer, INION Oy, Finland. |  |  | April 14, 2010 | 2010 | Stryker Craniomaxillofacial Division | Inadequate Instructions for Use. The stated strength retention is ambiguous and the maximum immersion time of the plates/meshed is not defined. | Other | Stryker issued an "URGENT MEDICAL DEVICE CORRECTION" letter dated April 14, 2010. The notice informed users of the change in the instructions for use (IFU), and included a copy of the revised IFU. The consignees were informed that the affected product does not need to be returned.Stryker can be contacted about this action at 877-534-2464 (option "Regulatory Affairs", reference Product Field Action #2009-445) |  | 8631 | 8631 all products | Nationwide Distribution: Throughout the US | Terminated | June 06, 2012 | 784 | 8631 | Not_Computer
Z-0624-2011 | 56659 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Cmax Surgical Tables, Model #'s: Cmax110, Cmax 220 manufactured by Steris Corporation in Montgomery, AL |  |  | August 12, 2010 | 2010 | Steris Corporation | Cmax Surgical Tables may lose calibration if the table is operated only on battery power and if the table is in a low battery charge condition. This could result in the table moving into an unexpected tilt position when the Level button is depressed on the hand control keypad. | Other | Steris issued Urgent Field Correction Notice letters dated August 26, 2010 to each affected consignee notifying them of the voluntary field correction. A Steris Representative will contact customers to arrange for STERIS to upgrade the software in the table.For further information about this corrective action, contact Steris at 1 800 333-8828. |  | 1635 | 1,635 tables | Worldwide Distribution: USA in the following states: AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and WV, and in the countries of Australia, Canada, China, Columbia, Dominican Republic, India, Israel, Mexico, New Zealand, Saudi Arabia, Singapore, South Korea and Taiwan. | Terminated | September 27, 2012 | 777 | 1635 | Software | Physical Safety Hazards | Software update | Software Update
Z-0627-2011 | 56716 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Wingspan Stent System, Model Number M003WE0300150, Manufactured by Boston Scientific, Fremont, CA |  |  | August 25, 2010 | 2010 | Boston Scientific Corporation | Products may not meet radial force specifications. | Process control | Recall initiated August 25, 2010. "Urgent Medical Device Recall - Immediate Action Required" notifications, dated 25 August 2010, were delivered by FedEx where possible, and by personal delivery assigned to sales representatives. The letter identified the affected product and the reason for recall. The letter stated that further distribution or use of any remaining product should cease immediately. Customers were asked to read the enclosed instructions and told that their local Sales Rep could answer any questions that they may have regarding the recall. |  | 10 | 10 units | International Distribution Only -- China, Germany, Italy, the Philippines, Portugal and Turkey. | Terminated | December 17, 2010 | 114 | 10 | Not_Computer
Z-0629-2011 | 56757 | No | DLJ | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | ARCHITECT iVancomycin Reagent; an in vitro chemiluminescent microparticle immunoassay (CMIA) for the quantitative measurement of vancomycin in human serum or plasma on the ARCHITECT i System with STAT protocol capability; Abbott Laboratories, Diagnostics Division, Abbott Park, IL 60064 USA; list 1P30-25 |  |  | September 13, 2010 | 2010 | Abbott Laboratories | The assay may produce falsely elevated vancomycin results. Inaccurate test results may result in inappropriate patient treatment. | Other | Abbott Laboratories issued a Product Correction letters dated September 13 and 15, 2010 to all customers that received ARCHITECT iVancomycin reagents in the last 12 months. In the September 13, 2010 letter, the accounts were informed of the falsely elevated results and were instructed to continue to follow all sample handling instruction in the package insert and to take the following additional actions to mitigate the risk of obtaining an incorrect result:1. After following the primary tube manufacturer's recommended centrifugation instructions transfer serum or plasma for all specimens to a microcentrifuge tube for centrifugation at 150,000-300,000 g-minutes, following centrifugation guidelines in the chart in the letter.2. Prior to testing, transfer the serum or plasma to a sample cup or secondary tube for testing. Care must be taken to transfer only the clarified specimen and not the lipemic material or pellet.The accounts were also requested to complete and fax back to Abbott the enclosed customer reply sheet, acknowledging the receipt and understanding of the recall. On September 15, 2010, Abbott discovered that the calculation for the determination of RCF (xg) given in the September 13, 2010 letter was incorrect. Incorrect Calculation previously provided: Relative Centrifugal Force (RCF) = 0.0000118 x rotational radius in centimeters x RPM2 Correct Calculation within revised letter: Relative Centrifugal Force (RCF) = 0.00001118 x rotational radius in centimeters x RPM2 Abbott stopped distribution of the previous letter and initiated redistribution of the revised letter to all impacted customers, globally, on September 15, 2010. Additionally, all customers which had already received the initial letter were sent a cover letter identifying the error.Customers can contact Abbott at 1-877-422688 about this action. |  | 2411 | 2,411 kits | Worldwide Distribution: Throughout USA, including Puerto Rico, and the countries of: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Finland, France, Germany, Greece, Hungary, Ireland, Israel, Italy, Japan, Latvia, Malaysia, Mexico, Netherlands, Oman, Pakistan, Poland, Portugal, Saudi Arabia, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates, United Kingdom and Vietnam. | Terminated | June 22, 2011 | 282 | 2411 | Not_Computer
Z-0625-2011 | 56957 | Yes | MVK |  |  | Cardiovascular | PMA | Model 4000 Monitor component of the LifeVest Wearable Defibrillator |  |  | September 24, 2010 | 2010 | Zoll Lifecor Corporation | The response button function may become intermittent. This can result in the device not progressing beyond the response button test during startup, which may prevent a conscious patient from delaying a treatment shock. | Other | Zoll Life Vest contacted patients that are expected to be in the device more than 60 days by letter and phone to arrange replacement monitors. Instructions for the return of the recalled monitor accompany the replacement monitor The firm will also contact the prescribing physician via telephone for long-term use patients to inform the physician of this voluntary recall and to assist in contacting any patients the firm was unable to reach by telephone.Zoll Life Vest can be contacted at 1-800-543-3267 about this recall. |  | 494 | 494 units | Nationwide Distribution: Throughout the US. including the states of AL, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, and WV. | Terminated | July 05, 2012 | 650 | 494 | Hardware | Treatment/Delivery/Therapy | Replace monitor | Remove or Replace
Z-0623-2011 | 56679 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Vacuette 2 ml Lithium Heparin, Collection Tube, Manufactured by: Greiner-Bio One North America, Inc., 4238 Capital Drive, Monroe, NC. |  |  | August 25, 2010 | 2010 | Greiner Bio-One North America, Inc. | Lithium heparin solution was contaminated by EDTA K (potassium) solution due to an operator error and a misinterpretation of a test result. Contaminated tubes will produce inaccurate test results. | Other | Greiner Bio-One instructed customers to stop distribution and to complete an attached Product Disposition form and return it via fax to Greiner Bio-One. They were further instructed to destroy any product they may have. Distributors were instructed provide a list of their customers that have received the affected item via Excel file so that they can be notified of the recall. |  | 18 | 18 cases (total 21,600 pieces) | Nationwide Distribution: USA, including the states of CA, FL, IL, IN, KS, KY, ND, NE, NH, OH, SC, SD and WI. | Terminated | May 15, 2012 | 629 | 18 | Not_Computer
Z-0622-2011 | 56855 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | ThermaCare Menstrual 8 Hour Heat Wraps, 3 count, NDC 0573-3020-02, SKU: 302002, UPC: 3-0573-3020-02-9, Wyeth Consumer Healthcare, P.O. Box 26609, Richmond, VA 23261-6609 |  |  | September 24, 2010 | 2010 | Pfizer Global | The chemical components contained in the heat wrap may leak directly onto the skin, which could cause skin injury, including burns. | Packaging process control | Wyeth Consumer Healthcare issued an Urgent Voluntary Medical Device Recall letter dated September 24, 2010 to direct accounts, and issued a press release. The firm is removing the product in question from store shelves and asking consumers who have purchased the product to record the lot number, throw the product away in its entirety without opening the foil pouch and call 1 800 232-3383 for replacement. Direct accounts were instructed to remove the affected product from sale and return it immediately. If they further distributed the product they were instructed to notify their customers of the recall and request return of the products to them credit. Instructions were provided in the event consumers inquired about the recall. A prepaid post card was enclosed to record the quantity on hand to be returned. |  | 192000 | 192000 cartons (3 wraps per carton) | Nationwide Distribution: Throughout the U.S. and Puerto Rico. | Terminated | September 16, 2011 | 357 | 192000 | Not_Computer
Z-0616-2011 | 57290 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | Fresenius Red Blood Cell Set, Catalog Number 9007601. |  |  | October 18, 2010 | 2010 | Fresenius Kabi, LLC | Potential for patient hematocrit dropped to levels lower than expected post-treatment. | Device Design | On October 21, 2010, Fresenius Kabi, LLC in Redmond, WA sent a "Field Safety Corrective Action" Letter dated October 18, 2010 to their US Consignees/Customers. On October 22, 2010, the firm followed up with their customers by telephone call.The letter informed the consignees/customers that the Red Blood Cell (RBC) exchange sets for use on the AS104 Blood Cell Separation Device are being recalled.The consignees/customers were advised to examine their stock and determine if they have any affected products on hand. They were instructed to discontinue distributing, using, dispensing the affected products, and return the product to Fresenius Kabi, LLC, 8635 154th Avenue, NE, Redmond, WA 98052.The consignees/customers were also instructed to notify their sub-account customers if the products were further distributed and to complete the Product Recall Response Form and fax it back to 425-242-2101. TIf you have any further questions concerning the recall, please contact the Quality Assurance Manager at 800-909-3872. |  | 255 | 255 units. | Nationwide distribution: including states of: AL, AR, AZ, CA, CT, GA, KY, MN, NC, NJ, OH, TN, and TX. | Terminated | October 05, 2011 | 352 | 255 | Not_Computer
Z-0606-2011 | 57104 | No | LZP | Intraocular fluid. | Ophthalmic | Ophthalmic | PMA | Beaver-Visitec International (Beaver-Visitec) BD 1% OVD Sodium Hyaluronate, 10 mg/ml 1.0 mL. Sterile. Rx |  |  | October 04, 2010 | 2010 | Becton Dickinson and Company | Finger flange on the BD 1% OVD syringe can disengage during cataract surgery and and may lead to injury. | Component design/selection | Beaver Visitec issued an Urgent: Field Safety Notice letter, dated 10/4/10 to direct accounts, identifying the affected product and informing users of the potential for the finger grip to disengage during use. The firm offers solutions as follows:Allow the finger grip to remain attached to the syringe and continue to use the product. Position the open side of the finger grip towards the palm to avoid detachment of the grip from the syringe. Remove the finger grip and continue to use the product. Return the unused product to Beaver-Visitec for reimbursement.Users were requested to complete the Customer Response Form. Questions can be directed to: Inside Sales at 781-906-7952. |  | 72278 | 72278 units | Nationwide Distribution | Terminated | February 16, 2012 | 500 | 72278 | Not_Computer
Z-0626-2011 | 57112 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ev3 Inc. 4600 Nathan Lane, Plymouth, MN ProtÂżgÂż " GPS" Self-Expanding Biliary Stent System, Sterile EO, 14 mm 60mm 80 cm, REF SERB65-14-60-80, Rx only. |  |  | October 30, 2010 | 2010 | Ev3, Inc. | Mislabeled size: Specific ProtÂżgÂż GPS lots have product labeling that does not match the actual stent sizes contained in the packages. The affected product outer packaging is labeled as a 12 mm x 60 mm ProtÂżgÂż GPS stent but contains a 14 mm x 60 mm stent. This situation can potentially result in vessel damage. Conversely, a lot was labeled as a 14 mm x 60 mm ProtÂżgÂż GPS stent but actually contai | Packaging process control | ev3 issued a "Medical Device Recall" letter dated September 29, 2010 to consignees. The letter described the product, Issue Summary, &amp; Required Action. Consignees are to review inventory and remove all affected devices. ev3 representatives will contact customers to arrange for the return and credit of unused product.ev3 can be reached at 800 716-6700 concerning this recall. |  | 30 | 30 | Nationwide Distribution: including the states of AZ, DE, C, FL, GA, ID, IL, IN,LA, MN, MO, NJ, NY, NC, OH, VA, and WA. | Terminated | March 29, 2012 | 516 | 30 | Not_Computer
Z-0711-2011 | 57230 | No | CHH | Cholesterol (total) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Genzyme Diagnostics Total Serum Cholesterol Kit, IVD |  |  | November 10, 2010 | 2010 | Diagnostic Chemicals, Ltd. | Cholesterol Reagent bottles labels reversed for R1 and R2 | Labeling mix-ups | Genzyme Diagnostics P.E.I. Inc. sent an "IMPORTANT PRODUCT FIELD CORRECTION" letter dated November 10, 2010, via Federal Express to customers. The letter described the label error and requested that the reagents not be interchanged with another reagent lot. The firm ask that the customers complete and return the attached IMPORTANT PRODUCT FIELD CORRECTION confirmation of notification via fax at 902-628-6504. Contact a Genzyme Diagnostics Technical Services Representative at 1-800-565-0265, if you require further information, or if you have any questions. |  | 30 | 30 kits | Worldwide distribution: USA including states of: KY, PA, NV, and WA and country of Canada. | Terminated | September 19, 2011 | 313 | 30 | Not_Computer
Z-0712-2011 | 57220 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Comprehensive Reverse Shoulder Humeral Tray with Locking Ring, 49 mm standard, sterile, REF 115350, Biomet Orthopedics, Warsaw, IN |  |  | September 13, 2010 | 2010 | Biomet, Inc. | The firm received complaints regarding fracturing of the device. | Pending | The firm sent "URGENT MEDICAL DEVICE RECALL NOTICE" dated 9/13/2010 to their distributors and customers. The letters stated the reason for the recall and what may happen as a result of the fracture of the device. The notices stated that the recalled devices should be immediately located, discontinued from use, and shipped back to Biomet. Customers should follow the instructions on the "FAX Back Response Form." If the affected product was further distributed, hospital personnel must be notified of the recall via the enclosed "Dear Biomet Customer" notice. The firm sent "URGENT MEDICAL DEVICE RECALL NOTICES' to implanting surgeons. These letters informed the surgeons of the recall and instructed them to continue to monitor the patient for fracture of the device.Customers are to confirm receipt of the notice by calling 800-348-9500 ext 3755 or 3983. Questions should be directed to 574-371-3755 or 574-372-3983. |  | 45 | 45 both products | Nationwide Distribution -- NY, VA, MD, AL, CO, SD, TX, and OK. | Terminated | July 27, 2011 | 317 | 45 | Not_Computer
Z-0628-2011 | 57281 | [6 empty] | ASP Automatic Endoscope Reprocessor (AER), Product Code: 20300 &amp; 20301 |  |  | February 19, 2010 | 2010 | Advanced Sterilization Products | The recall was initiated because Advanced Sterilization Products (ASP) received three customerreports of elevated temperature of the high-level disinfectant solution used in the ASP AER as a result of apossible temperature monitoring system malfunction. This recall communication advised consignees that if high-leveldisinfectants are exposed to elevated temperatures, fumes may occur and resul | Pending | Advanced Sterilization Products sent an URGENT: PRODUCT CORRECTION letter dated February 19, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. The recall communication was initiated with ASP forwarding a "Urgent: Product Correction" letter with attached customer fax-back response form to the affected customers. Customers were also instructed to complete and sign the enclosed customer fax-back response form indicating that they have disconnected the heater on the ASP AER according to the instructions provided and fax the form to 1-888-345-1435.If customers have any questions or need assistance, they were instructed to call ASP Professional Services at 888-783-7723, option 3 then option 1. |  | 1881 | 1,881 units (1514 in US) | Worldwide Distribution -- United States, Argentina Brazil, Canada, Chile, Colombia, Costa Rica, Ecuador,Hungary, Mexico, Panama, Guatemala, Honduras, Poland, Puerto Rico, Virgin Islands, Singapore, Taiwan,Venezuela, Croatia, and Bulgaria. | Terminated | July 30, 2013 | 1257 | 1881 | Not_Computer
Z-0714-2011 | 57236 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System, medical charged-particle radiation therapy system, manufactured by Accuray Inc., Sunnyvale, CA. |  |  | October 28, 2010 | 2010 | Accuray Inc | If users modify the dose calculation box to cover just the contoured anatomy and not the entire CT volume, dosage may not be properly calculated. | Other | Accuray Inc., sent an "URGENT DEVICE CORRECTION" letter dated October 21, 2010, to all customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to increase the size of the dose box to include CT volume, not just the contoured anatomy; recalculate dose at high resolution, and examine the treatment plan to ensure that applicable organs at risk are accounted for and all portions of the patient receive the desired level of doses. If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at +1-877-668-8667 (USA), +1-408-716-4700 (non-USA) or email: customersupport@accuray.com. |  | 219 | 219 units | USA and countries of: Athens, Belgium, Canada, China, Czech Republic, France, Germany, India, Italy, Japan, Korea, Malaysia, Russia, Saudi Arabia, Spain, Switzerland, Taiwan, Thailand, The Netherland, Turkey, Ukraine, United Kingdom, Vietnam, Venezuela. | Terminated | April 01, 2011 | 155 | 219 | Other |  |  | Other
Z-0007-2011 | 54154 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | HiSpeed X/i Computed Tomography operator consoles with the Model Number 2153675 and 2247802. |  |  | May 01, 2009 | 2009 | GE Medical Systems, LLC | It was discovered that the operator console did not have a label affixed in accordance with 21CFR 1010.2 indicating certification to applicable performance standards. | Radiation Control for Health and Safety Act | GE Healthcare will bring customer systems into compliance with 21 CFR by having a GE Healthcare field service representative, who will perform a field corrective action, visit each customer site. This will be carried out via a Field Modification Instruction planned to be released in October 2009 and planned to be completed by February 2010. Additionally, all forward production operator consoles are manufactured so as to comply with 21 CFR and to include a manufacturing check for appropriate labeling. |  | 45 | 45 units (13 Nationwide) | Worldwide Distribution -- USA, including states of ME, FL, IA, AL, CT, CA, TX, KS, KY, OH, and WV and countries of Japan, Russian Federation, India, Argentina, Brazil, Italy, Poland, Puerto Rico, Romania, and Saudi Arabia. | Terminated | May 16, 2012 | 1111 | 45 | Not_Computer
Z-0717-2011 | 57311 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Luminos dRF |  |  | October 25, 2010 | 2010 | Siemens Medical Solutions USA, Inc | patient leg can be pinched | Other | SIEMENS sent a Customer Safety Advisory Notice dated October 25, 2010, to affected customers via Update Instructions AX072/10/S. The notice described the issue and the actions to be taken by customers. The customers were instructed to retain the annex together with their operating instructions and bring the hazard to the attention of their employees; inform all users and other persons of this notice; if they sold the device, forward this notice and attachment to the new owners of the device and provide SIEMENS with details of the new owners.If you have any questions, please contact +1-888-826-9702. |  | 24 | 24 units | Nationwide distribution: CA, CO, FL, GA, IL, IN, KY, MI, MN, MO, NC, NJ, and NY. | Terminated | September 29, 2012 | 705 | 24 | Not_Computer
Z-0716-2011 | 57242 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | SDS Kerr Revolution Formula 2 Flowable Light Cure Composite. Net Wt. 4 g, Contents 4 Syringes (1g) each. This Package Contains 4 - 12 g Syringes A1, 20 - Dispensing tips, 29493. |  |  | October 15, 2010 | 2010 | Patterson Dental Supply, Inc. | Patterson Dental is conducting a recall because they have been advised by the manufacturer of Revolution Formula 2 (Kerr Corporation) that some syringes and tips marked with the brand name are not manufactured and distributed by Kerr and, thus, not authentic product. | Other | Consignees were sent a Patterson Dental "Urgent Device Recall" letter dated October 15, 2010. The letter was addressed to "Valued Customer". The letter described the product and the problem. It also advised the consignees to quarantine and return the product with the specific distribution dates. Consignees were to complete and return the enclosed Recall Return Form along with the affected product. Questions should be directed to Nancy Strand at 651-686-1966 or nancy.strand@pattersondental.com. |  | 3204 | 3,204 packages | Nationwide Distribution -- Including DC except VT. | Terminated | September 11, 2012 | 697 | 3204 | Not_Computer
Z-0715-2011 | 57277 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Polaris Spectra System Control Unit, a component of the StealthStation S7 System. Medtronic Navigation, Louisville, CO 80027. |  |  | November 05, 2010 | 2010 | Medtronic Navigation, Inc | Electronic surgical navigation aid may lose tracking ability or stop working. | Component design/selection | Customers were notified by a Product Correction Notification letter on 11/05/2010. The letter advised customers of the affected product, the potential problem, and the description of the correction. Customers were told that a Medtronic rep will visit them to install a replacement control unit. Further information is available at 800-595-9709 or 720-890-3200. |  | 48 | 48 devices | Worldwide Distribution -- US, including states of AL, AZ, CO, FL, GA, KY, MD, MN, NE, PA, TX, WI, and MD and countries of Australia, Belgium, Canada, Czech Republic, Italy, Netherlands, Poland, Portugal, Saudi Arabia, Singapore, South Korea, Spain, Turkey, and UAE. | Terminated | May 03, 2011 | 179 | 48 | Other
Z-0718-2011 | 57312 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobilett XP Hybrid/Digital/CR |  |  | October 04, 2010 | 2010 | Siemens Medical Solutions USA, Inc | battery fuse | Other | SIEMENS released a field update dated October 25, 2010, via SP011/10/S to the representatives. The update described issue and corrective action plan. The representatives are to visit each account to upgrade the systems with a new battery fuse and a new holder.If you have any questions regarding this report, please contact Quality Engineer by phone 610-448-3237 or email: meredith.adams@siemens.com and Director, Product complaints &amp; Field Support by phone 610-448-4634 or email: josephine.mcbride@siemens.com. |  | 170 | 170 units | Nationwide distribution: AZ, CA, DC, MD, PA, SC and VA. | Terminated | September 29, 2012 | 726 | 170 | Battery |  | Repair | Repair
Z-0794-2011 | 57351 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Accessory Cart for The ACHIEVA 3.0T and Panorama 1T Magnetic Resonance Imaging unit. |  |  | May 12, 2009 | 2009 | Philips Medical Systems North America Co. Phillips | The Accessory Cart contained magnetic material and is attracted to the magnet when bringing it close to the system. | Nonconforming Material/Component | On 05/12/2009, Philips began sending out the URGENT-Field Safety Notice MR Accessory Cart to their consignees. The letter identified the affected product and stated that the drawer rails were made from magnetic material and were not according specification. The hazards involved were also provided along with how to identify the affected product. The firm issued the FCO 78100296 and the firm's Field Service Engineer replaced the magnetic rails with the non-magnetic rails at the consignee's site.The consignees can contact contact the Philips Healthcare Call Center at 800-722-9377 #5 (then follow prompts) and reference FCO 78100296 with any communication. |  | 3 | 3 units | Nationwide Distribution -- IL, MI, and MD. | Terminated | December 28, 2010 | 595 | 3 | Not_Computer
Z-0797-2011 | 57219 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vital, Vitrea fX and Vitrea Enterprise Suite fX |  |  | October 14, 2010 | 2010 | Vital Images, Inc. | Vital Images Inc., reported that a software defect exists with Vitrea fX 2.1, 3.0 and 3.1 CT Brain Analysis whereby the annotated numeric ROI values on a batched perfusion map are incorrect when created using the auto batching function (from the right click menu) with the auto ROI templates enabled. | Software design | Vital Images, Inc sent a Vital "Urgent Software Update Alert" letter dated 14 October 2010 to all Consignees/Customers. The letter described the product and problem. They advised, until new software is installed, that the customers can avoid obtaining incorrect values by tuning off ROI templates prior to selecting the batch option from the right click menu in the CT Brain Analysis application; and recommended that the customers re-examine any ROI perfusion values that were previously generated using automatic collage batching with ROI templates.Please direct any questions or comments to Customer Support at support@vitalimages.com or at 1-800-208-3005. Any help needed to install this software update can also be directed to Customer Support. |  | 182 | 182 | Worldwide distribution: USA including states of: AL, AZ, AR, CA, CO, FL, GA, IA, KY, LA, MD, MA, MI, MN, MO, MT, NY, NJ, ND, OH, PA, PR, TN, TX, WA, WI, and WY, and countries of: CANADA, NETHERLANDS, TAIWAN, CHINA, SINGAPORE, UNITED KINGDOM, BRAZIL, GERMANY, RUSSIA, THAILAND, FRANCE, REPUBLIC OF KOREA, SPAIN, PAKISTAN,IRELAND, JAPAN, AUSTRALIA, INDIA, CZECH REPUBLIC, HONG KONG, EGYPT, ITALY, AUSTRIA, SLOVAK IA, BULGARIA, LEBANON, ALGERA, TAIWAN, SOUTH ARICA, BELGIUM, MALAYSIA, DENMARK, and JORDAN.CANADA, NETHERLANDS, TAIWAN, CHINA, SINGAPORE, UNITED KINGDOM, BRAZIL, GERMANY, RUSSIA, THAILAND, FRANCE, REPUBLIC OF KOREA, SPAIN, PAKISTAN,IRELAND, JAPAN, AUSTRALIA, INDIA, CZECH REPUBLIC, HONG KONG, EGYPT, ITALY, AUSTRIA, SLOVAK IA, BULGARIA, LEBANON, ALGERA, TAIWAN, SOUTH ARICA, BELGIUM, MALAYSIA, DENMARK, JORDAN. | Terminated | May 15, 2012 | 579 | 182 | Software | Output/Calculation | Software update | Software Update
Z-0795-2011 | 57349 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Advanced Bionics HiRes 90K Cochlear Implant device (unimplanted units), Model Numbers: CI-1400-01 |  |  | November 23, 2010 | 2010 | Advanced Bionics Corporation | The recall action is being taken in response to two confirmed instances where the product experienced a malfunction requiring explantation. These recipients experienced severe pain, overly loud sounds and/or shocking sensations at 8-10 days after initial activation of their device. | Pending | Advanced Bionics (AB) Personal contacted all centers with known surgeries about the recall by telephone on November 22, 2010, and sent Blastfax to affected surgeons and clinicians.AB forwarding a Press Release on November 23, 2010 to the Business Wire News Trak. The firm also faxed Health Care Provider letters to the affected Health Care Providers and placed a Patient notification letter on the AB website on November 23, 2010. Health Care Provider letters were FedEx'ed November 24, 2010 to the affected Health Care Providers. Patient Letters were sent via FedEx'ed beginning Monday November 29th and 30th to approx. 10,500 patients. A Recall Acknowledgement Form was included for each recipient to sign as proof that the notification has been read.The press release &amp; recall letter informed the patients/consignees of the reason for recall, the product affected and the potential hazard involved. The Health Care Provider letter also requested providers with recipient under their care who reports pain, overly loud sounds, or sudden shock sensations while the implant is powered to follow these steps:1. Advise them to remove the external equipment immediately,2. Schedule an appointment to complete standard troubleshooting steps listed in the AB Sound Processor User Guide,3. If standard troubleshooting does not resolve the issue, please contact Advanced Bionics for additional troubleshooting steps that may assist in determining the presence of recall problem, 4. If your evaluation of the patient's condition suggests that the implant exhibits this problem, device explantation and replacement is advised. In addition, the Health Care Provider letter informed providers that if AB records show that you have any unregistered implants in their possession, a list of serial numbers has been provided for devices that must be returned to Advanced Bionics. If no unregistered implants have been identified, no list will be enclosed. The Patient letter also |  | 332 | 332 units | Nationwide distribution: AL, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT and VA. | Terminated | March 23, 2012 | 486 | 414 | Not_Computer
Z-0719-2011 | 57291 | No | EXS | Urine flow or volume measuring system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | BardÂż CritiCoreÂż Monitor, Software Version 2.1, Catalog number 000002N, C.R. Bard, Inc., Covington, GA 30014. |  |  | November 17, 2010 | 2010 | C.R. Bard, Inc., Urological Division | Error can occur in the reading of the Present Interval Output (PIO) reading when the device is subjected to bumping or forceful replacement of the fluid container. | Other | Consignees were notified by an Urgent: Product Advisory Letter and Medica Device Recall letter on/about November 18, 2010. The letter identified the affected product, stated the reason for recall, the interim precautions, and actions to be taken. The letter requested acknowledgement of the recall and to return the entire affected product. Each consignee that is a distributor was instructed to forward the letter to their consignees to retrieve relevant product from those locations. Affected product received by Bard Medical will be reworked with the software upgrade and place back into distribution. Consignees will either receive a software upgrade on site by a third party or replacement product with the software upgrade. For international distribution, the hospital will be informed of this issue through Bard affiliate and Notified Body Bard Europe and managed according to the requirement. If there are questions or assistance is needed, customers can contact BD at 1-770-784-6969. |  | 234 | 234 units | Worldwide Distribution -- U.S., Netherlands, and Japan. | Terminated | February 02, 2012 | 442 | 234 | Software | Output/Calculation | Software update | Software Update
Z-0507-2011 | 57376 | No | MCM |  |  | Ear Nose &amp; Throat | PMA | Aimshot HeatSeeker Model HS3500, contains Class IIIb infared laser illuminator. |  |  | August 17, 2010 | 2010 | Amstech Incorporated | Product contains infra red laser, which can pose a potential hazard to users who unknowingly shine the beam into people's eyes. | Radiation Control for Health and Safety Act | Firm sent a letter dated August 5, 2010 to customers stating that a recall was being conducted on the AimSHOT HeatSeeker and specified model. The letter stated the reason for the recall and that the firm has redesigned the HeatSeeker and removed the infra red laser feature. The letter also asks customers to check their inventory and return the affected products to the firm. The firm states that they will send customers their new model and pay for the shipping. Questions should be directed towards 770-840-8039. |  |  | All units | Nationwide Distribution | Terminated | November 09, 2011 | 449 |  | Not_Computer
Z-0802-2011 | 57209 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | klinoPORT, AC powered Operating Table |  |  | October 26, 2010 | 2010 | Kreuzer Gmbh &amp; Co. KG | The motorized adjustable pendant arm may suddenly descend to its lowest position. | Other | Consignees were notified by a Warning Notice on/about 10/26/2010. Medical facilities were instructed to adjust the pendant arm to the lowest height position and disable the motor and to place the accompanying warning label on the pendant. They will, therefore, not be able to adjust the height of the pendant until corrective action is developed and completed. TRUMPF Medical Systems, Inc. will receive corrective action repair kits by 11/12/2010. |  | 15 | 15 units | Nationwide Distribution -- CT, MA, and MO. | Terminated | March 20, 2012 | 511 | 15 | Not_Computer
Z-0800-2011 | 57210 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Turon Shoulder System Instrumentation, Turon Trial Pegged Glenoid, DJO Surgical |  |  | November 05, 2010 | 2010 | Encore Medical, Lp | Complaints of fracturing were reported for affected trial products. | Device Design | Encore Medical notified consignees by letter on 11/5/2010. |  | 85 | 85 units | TX, FL, SC, GA | Terminated | April 05, 2011 | 151 | 578 | Not_Computer
Z-0790-2011 | 56965 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Arrow Ultra 8 IAB (Intra-Aortic Balloon) Catheter, 8 Fr, 30cc. |  |  | October 11, 2010 | 2010 | Arrow International, Inc., Division of Teleflex Medical Inc. | SuperArrow Flex IAB catheter becomes stuck in the sheath, unable to move the IAB catheter forward or backward, causing a delay in therapy, bleeding or arterial injury. | Device Design | Arrow international initiated a voluntary recall on 12/21/10 by letter requesting users to discontinue use and return all unused 5800 Series IAB with SuperFlex Introducers to Arrow. This field action supercedes the Safety Alert IAB 'Stuck in Sheath' letter issued on 10/8/2010. Questions should be directed towards their local sales representative or the IABP support Line at 1-866-396-2111 or 1-617-389-6400. |  | 5289 | 5289 units | Worldwide Distribution -- USA, Canada, Australia, Belgium, Czechoslovakia, Germany, Russia, Spain, Finland, France, Great Britain, Hungary, Italy, Netherlands, Poland, Saudi Arabia, Switzerland, Brazil, Australia, Chile, Colombia, Ecuador, India, Indonesia, Japan, Korea, Malaysia, Mexico, New Zealand, Panama, Peru, Philippines, Taiwan, Thailand, and Venezuela. | Terminated | February 05, 2015 | 1578 | 69523 | Not_Computer
Z-0789-2011 | 57335 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ev3 NanoCross 0.14 OTW PTA Dilatation Catheter, sterilized using ethylene oxide gas. REF #'s: AB14W040040150, AB14W030040150, AB14W020040150, AB14W025080150, AB14W020080150, AB14W020120150, AB14W030080150, AB14W035120150, |  |  | November 12, 2010 | 2010 | Ev3, Inc. | ev3 has determined that specific lots of the NanoCross .014 OTW PTA Dilatation catheters may crack or break during use. Cracking or breaking of the catheter shaft can result in the inability to inflate or deflate the balloon, and can result in material separation and potential embolization. | Process control | A Medical Device Recall letter addressed to Risk Managers, dated 11/10/2010, was sent to hospitals beginning 11/12/2010. The letter identified the affected devices, described the issue and requested that affected devices be located and removed from use. ev3 sales representatives will contact customers to arrange return of all unused devices. A credit or replacement will be provided for any unused devices. Questions regarding this issue may be directed towards Customer Service at 800-716-6700. |  | 4962 | 4,962 (3,858 US; 1,104 OUS) | Worldwide Distribution -- USA, AUSTRALIA, BELGIUM, BRAZIL, BULGARIA, CANADA, CHILE, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, KUWAIT, LEBANON, MALTA, NETHERLANDS, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SYRIA, TAIWAN, UNITED ARAB EMIRATES, and UNITED KINGDOM, | Terminated | May 17, 2011 | 186 | 4962 | Not_Computer
Z-0818-2011 | 56894 | No | LRP | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Depuy Spine Mountaineer OCT Spinal System Screw Caddy Set, , Non Sterile |  |  | October 02, 2010 | 2010 | DePuy Spine, Inc. | The scale which serves as a secondary means of identifying screw length is incorrectly etched on the screw caddy | Labeling design | Depuy Spine issued an "URGENT-VOLUNTARY PRODUCT RECALL NOTIFICATION" via e-mail on October 2, 2010 to customers. The notification describes the product, problem and action to be taken by customers. The customers were instructed to please distribute this information to appropriate personnel at their facility and return any caddies that they may have to their local distributor for DePuy Spine products. Depuy will return devices to customer promptly (within two business days), after reworking the devices by removing the scale altogether. Once the product is fully corrected, Depuy will provide the customers with a new and updated unit with the corrected scale markings. If you should have any questions or concerns regarding this action, please contact DePuy Spine Customers Service at 800-227-6633. |  | 121 | 121 units | Nationwide distribution: USA including states of: AR, AZ, CA, FL, HI, IA, IL, KY, LA, MA, MI, MN, MO, NC, NY OH, PA, TN, TX, VA, and WA. | Terminated | April 15, 2011 | 195 | 441 | Not_Computer
Z-0820-2011 | 55976 | No | FAF | Testicular prosthesis. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | Coloplast, Saline Filled Testicular Prosthesis, Sterile R, Manufacturer Coloplast A/S 3050 Denmark. |  |  | April 30, 2010 | 2010 | Coloplast Manufacturing US, LLC | On April 22, 2010, a decision was made by Coloplast to conduct a field correction based on a report that the Saline Filled Testicular Prosthesis is considered mislabeled. This is solely due to a difference in expiration dates between the Terumo Butterfly Needle included in the kit and the testicular implant. The needle has a 3yr expiration date &amp; the Saline Testicularhas a 5yr expiration date. | Labeling False and Misleading | The recall will extend to field personnel who have control of all SFTP Systems. The product is ordered by the customer (physician and/or hospital) directly from a field sales representative who ships the device to them or drops it off in person.Coloplast sent all Field Personnel an email on April 30, 2010, with a copy of a letter dated April 30, 2010 and a follow-up letter dated May 1, 2010. The letters described the problem and the product involved in the recall. The letters included a list of Corrections Required. This included the field personnel receiving a stock of sterile Butterfly Needles based on their home consignment inventory to be provided for use with all Saline Testicular Implants; Copies of customer letters to be provided with each replacement needle; Changes in all Saline Testicular orders, and Field personnel to administer the details of the correction and make the direct customer contacts. For questions, contact the Customer Service/Field Inventory Representative at 612-337-7859 or email: usbam@coloplast.com |  | 1018 | 1018 | Worldwide distribution: USA including states of AL, AK, AZ, AR, CO, CT, CA, FL, GA, HI, IL, IN, IA, LA, MD, MA, MI,MN, MS, MO, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT and WA; and countries of Australia, Canada, Hong Kong and Latin America (Chile, Columbia, Ecuador and Venezuela). | Terminated | December 17, 2011 | 596 | 1018 | Not_Computer
Z-0816-2011 | 56972 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | BD Vacutainer Push Button collection Set; |  |  | October 05, 2010 | 2010 | Becton Dickinson &amp; Company | There is the possibility for certain lots of the BD Vacutainer Push Button Blood Collection sets to have separation of the cannula from the hub. | Process control | BD Diagnostics sent "RECALL NOTIFICATION" separate letters October 1 and 5, 2010 and follow up letter dated October 8, 2010, as well as a letter dated December 1, 2010 to all distributors and customers commencing October 5, 2010 by federal express. The letters describe the product, the problem and the action to be taken by the customers. The customers were instructed to immediately examine their inventory for specific catalog and lot numbers, and return product from the lots following the instructions on the enclosed packing slip (BD will send replacement upon receipt of returned product). If you have any questions on returning and replacing product, please contact BD logistics at 1-866-380-0046. If you have any further detail questions regarding this recall, please contact your BD Sales Consultant or BD Technical Services via email at www.BD.com/Vacutainer/Contact or via telephone at 1-800-631-0174. |  | 4223200 | 4,223,200 | Nationwide Distribution | Terminated | June 06, 2012 | 610 | 4223200 | Not_Computer
Z-0808-2011 | 56846 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Model # 720088-01: AMS AdVance Male Sling System, US; Model # 720088-02: AMS AdVance Male Sling System, INTL; Model # 72404224: Male TO Sling System, INTL. |  |  | September 07, 2010 | 2010 | American Medical Systems, Inc. | American Medical Systems (AMS) is recalling all lots of the AdVance male Sling System with an expiration date between July 2012 and August 2013. Internal testing has indicated the potential for degradation of the bioabsorbable tensioning sutures when the product is stored for longer than 15 months after the date of manufacture. | Package design/selection | Consignees were sent an AMS "Urgent Product Recall" letter dated September 7, 2010. The letter described the problem and the product involved. They requested consignees to return any existing stock product and to return the Product Recall Acknowledgement form enclosed. Customers may request a credit or a full refund or also, arrange for replenishment of AdVance product. Questions regarding the recall should be directed towards AMS Customer Service at 1-800-328-3881. |  | 7069 | 7,069 | Worldwide Distribution -- All states in the Continental USA including DC and Puerto Rico and countries of ARGENTINA, AUSTRIA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, SWITZERLAND, CHILE, CHINA, COLOMBIA, COSTA RICA, GERMANY, ESTONIA, SPAIN, FINLAND, FRANCE, UNITED KINGDOM, GREECE, HONG KONG, CROATIA, IRELAND, ISRAEL, INDIA, ITALY, REPUBLIC OF KOREA, LEBANON, MOROCCO, MEXICO, MALAYSIA, NETHERLANDS, NORWAY, NEW ZEALAND, PANAMA, POLAND, PORTUGAL, RUSSIAN FEDERATION SAUDI ARABIA, SWEDEN SINGAPORE, SLOVENIA, THAILAND, TURKEY, TAIWAN, VENEZUELA, and SOUTH AFRICA. | Terminated | June 28, 2012 | 660 | 7069 | Not_Computer
Z-0813-2011 | 57348 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Blease Sirius Anaesthesia System is a anaesthesia Gas Machine. The system has the capability to mount two anaesthetic vapourisers on the back bar via a demountable system. The delivery of anaesthetic agent is achieved by the rotation of a control knob which actuates a feature designed "shot bold" mechanism. The purpose of which is to prevent two vapourisers being turned on simultaneously.. |  |  | June 25, 2009 | 2009 | Spacelabs Healthcare, Llc | Users may experience difficulty turning the control knob to an active vapour delivery setting, or may not be able to turn on the Left Hand mounted vapouriser on the Blease Sirius Anaesthesia Systems. | Other | Spacelabs Healthcare issued an URGENT MEDICAL DEVICE CORRECTION letter dated June 25, 2009 to the two consignees in the US. Spacelabs Healthcare UK notified consignees outside the US.The consignees were advised to brief their staff about the issue and test the Anesthesia Systems in advance of a procedure. If there is an issue with turning on a vaporizer in the left hand position, simply move the vaporizer to the right hand position.The firm contacted their consignees for a schedule to replace the flowmeter mouldings. The US consignees can contact Spacelabs Healthcare Representatives at 425-657-7200 ext. 5089 about this recall. Foreign consignees can contact Spoacelabs at 44 1494 784422. |  | 41 | 5 units to US consignees and 36 units Internationally. | Worldwide Distribution: US in the states of LA and SD, and internationally by Spacelabs Healthcare, Ltd in the countries of China, Czech Republic, Hong Kong, India, Poland, Saudi Arabia, Singapore, and the United Kingdom. | Terminated | December 28, 2010 | 551 | 41 | Not_Computer
Z-0804-2011 | 57274 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | ACL 9000 Coagulation Analyzer, PN 00000990011 |  |  | November 11, 2010 | 2010 | Instrumentation Laboratory Co. | Sporadically lowered APTT results were observed with the PT/APTT Test Groups. | Device Design | On 11/11/10, An Urgent Product Notification informing customers to discontinue use of all IL locked PT/APTT-based Test Groups was mailed to U.S. and Canadian customers, instructing them to run in Single Test or Profile Mode until further notice. The letter identified the affected product, the issue, risk and stated a list of mandatory actions. Customers are to immediately acknowledge the receipt of the notification by signing and returning it via email or fax. A Worldwide Regulatory Notice was sent on November 8, 2010 to international affiliates (non-U.S. or Canada), notifying them of the issue, providing a risk summary and the support documentation for the notification. Secondary international dealers were notified Nov 8 and a fax response of notification of the dealers will be tracked by International Laboratories. The customer notification of the secondary dealers will be handled per the local regulatory requirements of the countries where they are located. |  | 87 | 87 distributed in USA &amp; Canada | Worldwide Distribution -- USA, Canada, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Sri Lanka, Sweden, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, Turkmenistan, United Arabian Emirates, Uruguay, Venezuela, Vietnam, Austria, Belgium, China, France, Germany, Holland, Italy, Japan, Spain (and Portugal), and UK. | Terminated | August 23, 2013 | 1016 | 87 | Not_Computer
Z-0815-2011 | 57292 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CuttingEdge Advantage Rasp Handle, Straight Rasp Handle, &amp; Specialty Head Center Accolade Broach Handle; Stryker Howmedica Osteonics, Non-sterile Instruments; Howmedica Osteonics Corp., 325 Corporate Drive, Mahwah, NJ 07430; |  |  | November 01, 2010 | 2010 | Stryker Howmedica Osteonics Corp. | The CuttingEdge Advantage Rasp Handle, Straight Rasp Handle, and Specialty Head Center Accolade Broach Handle may fracture in two pieces along the width of the body under the striking plate. This may extend surgical time under anesthesia, cause trauma to the surgeon or patient, or contaminate the surgical site. | Device Design | Stryker issued a Product Correction Bulletin dated November 2010 to their consignees, alerting surgeons of the need to inspect the instruments before and after each surgical use, and, if a defect is detected, to contact Stryker for return of the product. Stryker can be contacted at 201 831-5288 concerning this issue.In response to their revised strategy, Stryker issued notification letters and product recall acknowledgement forms to their branches and agencies on 2/17/2011 for 2/18/2011 delivery via Fed Ex and to their hospital risk managers on 2/18/2011 for 2/21/2011 delivery via Fed ex. |  | 459 | 459 units | Worldwide Distribution: Throughout the US, the countries of Canada and Japan, and the regions of Asia Pacific, Europe, Middle East, Africa, and Latin America. | Terminated | March 26, 2014 | 1241 | 459 | Not_Computer
Z-0814-2011 | 57337 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ThomoTherapy Hi-Art System |  |  | November 08, 2010 | 2010 | TomoTherapy Incorporated | TomoTherapy Inc. is sending this Field Safety Notice to make you aware of an anomaly which may affect the performance of the Hi-Art System and make it so the user cannot generate a Completion Procedure. | Software design | Consignees were sent a TomoTherapy "Urgent Field Safety Notice" dated November 8, 2010. The letter addressed to TomoTherapy Customers described the Issue, Product Affected, Recommended Action and Resolution. Customers are to contact the TomoTherapy Customer Interaction Center for further instruction and assistance in generating a completion procedure, if they experience the issue. Questions should be directed towards the TomoTherapy Customer Interaction Center via email or telephone. |  | 40 | 40 (21US; 19 OUS) | Worldwide Distribution -- USA, including the states of AZ, CO, IL, KS, MI, MO, NV, NY, ND, OH, OK PA, UT, and VA and countries of MALAYSIA, SWITZERLAND, KOREA, ITALY, SPAIN, CANDA, TAIWAN, JAPAN, BELGIUM | Terminated | February 04, 2012 | 453 | 40 | Software
Z-0811-2011 | 57317 | No | FCG | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific Expect Endoscopic Ultrasound Aspiration Needle, 22GA (0.72 mm) sterile, UPN M00550011, 5 pack, 428 Technology Drive East, Menomonie, WI. |  |  | October 20, 2010 | 2010 | Boston Scientific Corporation | The firm has received a high number of complaints related to difficulty advancing and/or retracting the needle. | Other | "Urgent Medical Device Recall - Immediate Action Required" notification letters were sent via Federal Express Priority mail on October 20, 2010. The letters identified the affected product and stated the reason for the recall. The notification instructs the account to immediately discontinue use of, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. |  | 567 | 567 all products. | Worldwide Distribution -- US and Canada. | Terminated | August 03, 2011 | 287 | 567 | Not_Computer
Z-0831-2011 | 51260 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Manual Wheelchairs with Transit System Option, |  |  | February 02, 2009 | 2009 | Sunrise Medical Inc | The stabilizer bar was inadvertently omitted during the manufacturing process. | Process change control | An Urgent: Medical Device Correction letter, dated February 2009, with certified return receipt was sent to customers in the United States; Canadian consignees were notified via e-mail. The letter identified the affected product and the reason for the recall. The letters asked customers to contact their Quickie dealer to make arrangements to have their device wheelchair corrected at no charge and to also verify their wheelchair serial number on the form. Questions should be directed towards their dealer or Sunrise Medical at 800-456-8168. |  | 23 | 23 | Worldwide Distribution -- United States and Canada. | Terminated | January 03, 2011 | 700 | 23 | Not_Computer
Z-0830-2011 | 53183 | No | FAT | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Stryker Infravision Esophageal Kit, Model Numbers 220180548, 220180540, 220180560, manufactured by Stryker Endoscopy San Jose. |  |  | August 20, 2009 | 2009 | Stryker Endoscopy | The silicone sheath component of the product detaches from the catheter during removal of the e-kit from the body. | Device Design | Stryker issued Urgent Product Recall letters dated August 20, 2009 to all its direct consignees, informing them of the affected products and providing instructions on the recall. Customers were requested to return the affected product to Styrker. |  | 2985 | 2985 units | Worldwide Distribution: USA in the states of: Alabama, Arkansas, California, Colorado, Connecticut, Delaware, Florida, Georgia, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Michigan, Minnesota, Missouri, North Carolina, New Hampshire, Nebraska, New Jersey, New Mexico, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Virginia, Vermont, Washington, Wisconsin, West Virginia, and Wyoming, and the countries/region of Australia, Canada, Scandinavia, Italy, and Switzerland. | Terminated | January 03, 2011 | 501 | 2985 | Not_Computer
Z-0825-2011 | 54147 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Stryker brand End Cutter Shaver Blade F-Series, 4.0 mm; Model Number -375-747-000, manufactured by Stryker Puerto Rico |  |  | November 12, 2009 | 2009 | Stryker Endoscopy | Potential breakage of the internal shaft of the End Cutter Shaver Blade; F-Series, 4.0mm. if breakage occurs, the device becomes non-functional and the defect could be observed through the camera/console system. | Device Design | Stryker issued Urgent Product Recall letters dated November 12, 2009 to all its direct consignees, informing them of the affected products and providing instructions on the recall. Customers were requested to return the product to Stryker.Stryker can be contacted concerning this recall at 408 754-2124. |  | 295 | 295 units | Nationwide Distribution includes: Arizona, California, Colorado, Denver, Florida, Hawaii, Illinois, Kansas, Kentucky, Massachusetts, Michigan, New York, Ohio, and South Carolina. | Terminated | January 03, 2011 | 417 | 295 | Not_Computer
Z-0823-2011 | 56245 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | Steris Reliance EPS Endoscope Processing System , manufactured by Corporation STERIS Canada, Quebec, Canada. |  |  | July 02, 2010 | 2010 | Steris Corporation | Reliance EPS had internal water leaks and interrupted cycles when air leaks triggered false Fault 38 alarm that DG/HLD cup has not been replaced. | Other | Steris issued an "Urgent Field Correction Notice" dated July 2, 2010 to each affected US consignee, notifying them of the issue and affected device. Steris Field Service Representatives and Dealers will contact each affected customer to schedule a service visit to install system upgrades. In the event of an interrupted cycle prior to the upgrade, consignees were instructed that the load should be handled in accordance with the Reliance EPS Operators Manual. Steris can be contacted at 440-392-7043. |  | 936 | 936 | Worldwide distribution: USA, including states of AK, AL, AZ, CA, CO, DE, FL, ID, IL, KS, KY, MD, MI, MN, MO, NC, NV, NY, OH, PA, SC, VA, WA, WI, and WY, and countries of Algeria, Australia, Belarus, Canada, France, Germany, Greece, India, Italy, Japan, Kuwait, Lebanon, New Zealand, Russian Federation, Saudia Arabia, Spain, Turkey, United Arab Emirates and United Kingdom. | Terminated | September 27, 2012 | 818 | 936 | Not_Computer
Z-0821-2011 | 56774 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand Trilogy System with RapidArc, Model Numbers: |  |  | July 20, 2010 | 2010 | Varian Medical Systems Inc. | The system may deliver an inaccurate dose that does not match the treatment plan. As a result of a system configuration problem, the VMAT (Volumetric Modulated Arc Therapy) plans may be delivered as Dynamic Arc therapy plans. If this occurs, the VMAT plan will not be recognized. This issue affects Trilogy machines with the Rapid Arc Treatment delivery option. | Software design | Varian issued Urgent Medical Device Correction letters dated July 20, 2010 to its direct consignees, informing them of the affected product and providing instructions on the actions to be taken. These instruction include: if the unit does not display "RapidArc" or "VMAT" under Treatment Type during dose delivery, discontinue use of the device and call Varian service.Varian Service Help Desk can be contacted at 1-888-827-4265. |  |  |  | Worldwide Distributiuon: Throughout the US, and the countries of Australia, Canada, China, Japan, and New Zealand, and in the regions of North and Latin America, Europe, and South East Asia. | Terminated | July 18, 2011 | 363 |  | Software
Z-0822-2011 | 56170 | [6 empty] | BD Beaver Slit Blade 3.0mm Straight |  |  | June 24, 2010 | 2010 | Bd Medical Systems | Blades may contain defects at the blade tip or cutting edge. | Process control | BD Medical notified customers via letter dated 6/24/10 to discontinue use of product and remove from inventory. Customers were instructed to complete the customer response form and fax to BD Medical. BD Medical was to contact customers upon receipt of the customer response form to arrange for replacement.Unused product was to be returned to:BD MedicalOphthalmic SystemsAttn: WT-12582411 Waverly Oaks RoadBuilding 2, Suite 229Watham, Massachusetts 02452-8402For questions regarding this recall call 1-781-906-7952. |  | 524 | 524 pieces | Nationwide Distribution | Terminated | June 29, 2012 | 736 | 524 | Not_Computer
Z-0829-2011 | 57273 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Vanguard Knee Instruments offset reamer bushing - angled, REF 32-488519, Biomet Bridgend, UK. |  |  | September 10, 2010 | 2010 | Biomet, Inc. | The firm has initiated the action following a report that during surgery the Offset Reamer Brushign would not mate with the prepartion mask. After a 5 minute delay in surgery the procedure was completed using another attachment that was on hand. There was no injury to the patient. | Other | The firm, BIOMET, sent a "FIELD SAFETY NOTICES" dated September 10, 2010, addressed to the head of orthopaedic departments, trauma departments, operating departments, sterile service departments, procurement, supplies and risk management. The notice described the product, problem and action to be taken by the customers. The letters instructed the customers to locate and immediately quarantine the recalled device; to complete and return the attached "FAX BACK RESPONSE FORM" via fax to +44 (0) 1656 645454, and the recalled device for credit via mail to Biomet UK Ltd, Waterton Industrial Estate, Bridgend, CF31 3XA. The customers were also instructed to inform all in their organization who are using or ordering the products and send a copy to the organization to which potentially affected products may have been transferred.If you have any questions regarding this communication, please contact the Quality Manager at +44 (0) 1656 655221 or email: david.broderick@biomet.com. |  | 16 | 16 units | International distribution only: Netherlands, Germany, France, Spain, Italy and Denmark | Terminated | January 19, 2011 | 131 | 16 | Not_Computer
Z-0828-2011 | 57286 | No | HWB | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | CLP Instrumentation Removal Hook, distributed by DJO Surgical, manufactured by Encore Medical, Austin, TX. |  |  | November 10, 2010 | 2010 | Encore Medical, Lp | Revision A and B CLP Extraction Hooks may fracture at the distal end, resulting in surgical delay. | Other | DJO issued an Urgent Recall Notice letter dated November 10, 2010 to consignees identifying the affected devices and requesting return of all affected devices for replacement. Customers are also requested to complete and return an enclosed Recall Response Form.Customers can contact David Harris of DJO at 512 834-6330 with questions about this recall. |  | 27 | 27 Units | Nationwide Distribution to the states of AZ, CA, FL, GA, ID, KS, MO, NY, TX, and WI. | Terminated | April 25, 2011 | 166 | 27 | Not_Computer
Z-0827-2011 | 57289 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC 2000 ultrasound system. |  |  | September 29, 2010 | 2010 | Siemens Medical Solutions USA, Inc. | There is the possibility of a short occurring in the transducer which could cause focal heating on the probe face. | Other | SIEMENS issued a filed correction in the form of a "Customer Safety Advisory Notification" initiated on September 29, 2010 and mailed to customers by Federal Express on the same day. The notification described the product, problem and actions to be taken by the customers. The customers were instructed to examine the 4Z1c transducer prior to use (Note: if transducer does not show a blister on its front face or does not feel unusually when it is safe for use); to immediately discontinue use and contact their local Siemens representative to get issue resolved, if the transducers developed a hot spot that exhibits a visible blister on the patient contact surface and feels hot to the touch during operation, and to pass this notice on to all those within their organization who need to be aware.If you have any questions, please contact (650) 969-9112 or (800) 422-8766. |  | 196 | 196 total units, transducers and SC2000 | Worldwide Distribution: USA and countries including: Australia, Belgium, Brazil, China, Czech Republic, Finland, France, Germany, India, Ireland, Italy, Japan, Kenya, Malaysia, Netherlands, Norway, Portugal, South Korea, Spain, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom. | Terminated | October 04, 2011 | 370 | 2392 | Not_Computer
Z-0824-2011 | 57341 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Zimmer Natural Nail System Cephalomedullary Lag Screw Inserter Long, Catalog 00-2490-003-50, Zimmer, Warsaw, IN. |  |  | October 29, 2010 | 2010 | Zimmer Inc. | The firm has determined that two lots of this instrument were inappropriately manufactured by the supplier. The manufacturing process created a step in the lumen which impedes the Lag Screw Retaining Shaft from passing through the lumen without the use of excessive force. | Nonconforming Material/Component | The firm sent Correction and Removal Notification Immediate Response Required letters via email to consignees. The letter identified the affected product and the reason for the recall. The letters also discussed the clinical implications and required actions. Customers are to complete the inventory certification form and user facility documentation form. Each of the affected accounts will be provided with a new instrument, manufactured with the correct lumen diameter and will secure and return the affected instruments per the attached notification, along with the completed inventory certification form. If any further information is needed, customers are to contact Aimee Wood at 574-372-4463. |  | 34 | 34 | Worldwide Distribution -- USA, including states of MI, NJ, MS, OR, KS, TN, A, UT, and CA and countries of Belgium and Germany, Spain, Italy, Greece, France, UK, and Switzerland. | Terminated | September 16, 2011 | 322 | 34 | Not_Computer
Z-0836-2011 | 56024 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems FastPlan software, Versions 5.5 and 5.5.1, Model Number HZ3, manufactured by Varian Medical Systems, Palo Alto, CA |  |  | March 23, 2010 | 2010 | Varian Medical Systems Oncology Systems | Due to software Issue, an error occurs when performing cone planning using MRI, which may result in the patient receiving a higher radiation dose than expected, potentially to healthy tissue. | Device Design | Varian sent Urgent Medical Device Correction letters on March 23, 2010, by Federal Express to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to run test cases or otherwise determine whether the MR images that they receive have the 'Series Date' parameter populated.Customers should contact their MR system manufacturer to get a confirmation that the 'Series Date' field is populated.Customers were instructed to confirm by visual inspection that the fusion results are loaded appropriately in Cone Planning before starting to create or edit a plan.For questions regarding this recall contact your local Varian Customer Support Distirct or Regional Manager.Phone: USA and Canada - 1-888-VARIANS (888-827-4265) Europe -+41 41 749 8844E-mail: North America: support-america@varian.com Australia/New Zealand: support-anz@varian.com Europe: support-emea@varian.com South East Asia: seasia.apps.helpdesk@varian.com China/Asia: china.apps.helpdesk@varian.com Japan: Japan.Apps.Helpdesk@varian.com Latin America: soporte.al@varian.comInternet: Oncology Systems customer site - www.myvarian.com Varian Medical System public site - www.varian.com |  | 218 | 218 units | Worldwide distribution - USA nationwide and the countries of Australia, Brazil, Canada, Chile, Colombia, Georgia, Germany, Hong Kong, India, Ireland, Japan, Macau, Malaysia, Russia, Singapore, Spain, Sri Lanka, and Taiwan | Terminated | February 24, 2012 | 703 | 218 | Software | Treatment/Delivery/Therapy | Instructions | Safety Notice/Insructions
Z-0834-2011 | 56763 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GreenLight MoXy Fiber Optic, part number 0010-2400, for use with the GreenLight XPS Laser System, manufactured by American Medical Systems Inc., San Jose, CA |  |  | July 14, 2010 | 2010 | AMS Innovative Center - San Jose | Product not intended for release was distributed, which may result in old components, not built to specifications, or undergo final testing | Labeling mix-ups | American Medical Systems sent oral notifications that the product used for patient case evaluations was from the 2 year shelf life study. Products have been returned, as the product was with AMS field personnel.For questions regarding this recall call 847-855-6270. |  | 77 | 77 units | Worldwide Distribution, USA including TX, OH, MO, FL, MA, CA, IL, NY, VA, and the country of Australia. | Terminated | January 03, 2011 | 173 | 77 | Not_Computer
Z-0833-2011 | 57196 | No | BTA | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Care E Vac 3 Portable Suction; Battery/AC Portable Aspirator; Ohio Medical Corporation, 1111 Lakeside Drive, Gurnee, IL 60031; Model 758000 |  |  | November 08, 2010 | 2010 | Ohio Medical Corporation | The D-14 diode on the circuit board may fail, resulting in a damaged, overheated circuit board. Also a screw used to attach one of the feet of the unit may damage the battery inside the device if the screw is over-tightened. | Device Design | Ohio Medical sent Urgent Medical Device Correction Notification letters dated November 4, 2010, to all affected customers on November 8, 2010. The letters identified the product, the problems, and the action to be taken by the customer.Customers were instructed to take the affected product out of service, remove diode 14, and replace the foot screw contained therein. Customers were instructed to visit the micro-sites established for these repair procedures for step-by-step instructions on performing the repairs. Accounts with questions after viewing the video demonstration were directed to the Ohio Medical technical assistance line at 224-430-1522. Customers were asked to complete the attached Customer Acknowledgment Form and return the form to the Quality Assurance Department at (fax) 1.847.855.6304, or e-mail to:Complaints@OhioMedical.com, or mail to:Ohio Medical CorproationQuality Assurance1111 Lakeside DriveGurnee, Illinois 60031.For questions regarding this recall call 847-855-6270. |  | 4183 | 4,183 units | Worldwide Distribution - USA nationwide, Puerto Rico, and the coutries of Canada, Israel, Jordan, Kenya, Kuwait, Saudi Arabia, Taiwan, Thailand and Vietnam. | Terminated | March 16, 2012 | 494 | 4183 | Hardware | Physical Safety Hazards | Repair | Repair
Z-0835-2011 | 57110 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Mammography Keypad |  |  | October 22, 2010 | 2010 | AGFA Corp. | Product did not receive UL Certification | Other | AGFA sent an Urgent Safety Notice letter dated October 22, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to unplug existing keypad(s)  including USB cable. Plug in loaner Keypad(s). Print out the Fed Ex shipping label that they received via the e-mail notification. Use that label to ship their existing keypad(s) and their corresponding USB cables to the manufacturer. The manufacturer would perform the upgrades and return the keypad(s) along with the USB cables.Upon receipt of the original keypad(s) customers would also receive another e-mail notification with a shipping label so that they may return the loaner keypad(s) to AGFA HealthCare. Customers were to distribute this information to all those who need to be aware of it, and complete the feedback form as soon as possible and return it to AGFA HealthCare.For questions regarding this recall call 838-342-3989.. |  | 26 | 26 units | Worldlwide Distribution - USA including FL, GA, KY, ME, MD, MS, OH, UT and the country of Canada | Terminated | February 01, 2011 | 102 | 26 | Not_Computer
Z-0832-2011 | 57254 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Urethral Warmer System Model number FPRCH3001 |  |  | November 11, 2010 | 2010 | Galil Medical | This voluntary recall has been initiated because our supplier (Gaymar Industries, Inc.) notified us that these products contain power cords that may crack or fail at or inside the plug due to a prong design problem. Affected plugs can be identified by a round ground pin and black plastic bridges. Cords with plugs that are not affected have U-shaped ground pins and no plastic bridges. Although Gali | Other | Gaymar sent an URGENT: MEDICAL DEVICE RECALL letter dated September 30, 2010, to all consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to: Customers were instructed to:" Inspect units in their possession to determine if they contain affected power cord plugs or if there were any signs of excessive wear or damage. " Immediately discontinue use and quarantine the affected product. " Contact Customer Service Department at 800-828-7341 to arrange for replacement. " Complete the attached form for each unit and follow the instructions given to return the form by fax or e-mail. " If customers further distributed the product to another person or entity, they were to identify their customers and notify them at once of this product recall by including a copy of the recall notification letter and the attached form. For questions regarding this recall call 800-828-7341 or 716-662-2551. |  | 4 | 4 | Worldwide Distribution - USA, including CA, GA, TX and the country of Canada. | Terminated | December 10, 2011 | 394 | 4 | Not_Computer
Z-0541-2011 | 52039 | No | KLC | Dental chair and accessories. | Dental | Dental | 510(k) | A-dec Decade Plus 1221 Dental Chair |  |  | April 01, 2009 | 2009 | A-Dec Inc | The design of the Decade Plus 1221 chair backrest creates a possible pinch-point between the chair backrest and the underside of the armrests. | Device Design | A-dec sent a letter dated March 23, 2009 to consignees/customers (doctors and network dealers). The consignees were informed that the design of the Decade Plus 1221 chair backrest creates a possible pinch-point between the chair backrest and the underside of the armrests. The potential for this occurs when returning the chair to the upright position. The consignees were advised to reprogram the Position 0/Entry-Exit preset to stop the back before it reaches a full upright position. This increases the gap between the armrest and the back slightly. To reprogram a chair, the firm included excerpts from a chair Instructions for Use (IfU) guide that detail the steps to reprogram a chair. Consignees can contact the firm's Customer Service at 503-538-9471, 1-800-547-1883, or via fax at 503-538-0276.The firm will replace the chair back assembly for all affected dental chairs. |  | 2723 | 2723 units | Worldwide distribution: US and Canada. | Terminated | December 16, 2010 | 624 | 2723 | Not_Computer
Z-0790-2011 | 56965 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Arrow Ultra 8 IAB (Intra-Aortic Balloon) Catheter, 8 Fr, 30cc. |  |  | October 11, 2010 | 2010 | Arrow International, Inc., Division of Teleflex Medical Inc. | SuperArrow Flex IAB catheter becomes stuck in the sheath, unable to move the IAB catheter forward or backward, causing a delay in therapy, bleeding or arterial injury. | Device Design | Arrow international initiated a voluntary recall on 12/21/10 by letter requesting users to discontinue use and return all unused 5800 Series IAB with SuperFlex Introducers to Arrow. This field action supercedes the Safety Alert IAB 'Stuck in Sheath' letter issued on 10/8/2010. Questions should be directed towards their local sales representative or the IABP support Line at 1-866-396-2111 or 1-617-389-6400. |  | 5289 | 5289 units | Worldwide Distribution -- USA, Canada, Australia, Belgium, Czechoslovakia, Germany, Russia, Spain, Finland, France, Great Britain, Hungary, Italy, Netherlands, Poland, Saudi Arabia, Switzerland, Brazil, Australia, Chile, Colombia, Ecuador, India, Indonesia, Japan, Korea, Malaysia, Mexico, New Zealand, Panama, Peru, Philippines, Taiwan, Thailand, and Venezuela. | Terminated | February 05, 2015 | 1578 | 69523 | Not_Computer
Z-0845-2011 | 57192 | No | HCG | Neurovascular embolization device. | Neurology | Neurology | 510(k) | Micrus Endovascular Microcoil Delivery System; |  |  | October 21, 2010 | 2010 | Micrus Endovascular Corp., 821 Fox Lane, San Jose, CA 95131 | The firm has received reports of difficulty connecting to the detachment control cable. Loss of the electrical connection will result in failure to detach the coil for deployment inside an aneurysm. | Other | Micrus Endovascular Corp sent Urgent: Medical Device Recall Letters dated October 21, 2010, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine inventory (including consignment) and quarantine the product to prevent it from being used or shipped. If the product was further distributed, customers were to notify their customers by sending a copy of the recall letter and were instructed to work with all customers to ensure that all product was returned and that pre-use protocol was followed.The response form was to be completed and returned indicating the number of devices that had been used (or shipped if distributor) and the number that would be returned. Customers were to call the Micrus Customer Service Department at 1-888-550-4129 from 8:00 am to 5:00 Pacific time for arrange for a Return Authorization number and replacement. Recall letters also stressed the importance of always following the pre-use protocol and inspecting the device immediately prior to use. Additionally, connector functionality should be verified and instructions were provided.Questions or concerns should be directed to Micrus Customer Service at 1-888-4120.Consignees who did not respond to the initial recall notification after 15 days were notified again by letter. A third letter was sent after another 15 days if there was still no response. |  | 6755 | 6755 | Worldwide Distribution - USA including AK, AL, AZ, CA, CO, DC, DE, FL, GA, ID, IL, IN, KY, LA, MA, MD, MI, MN, NC, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV and the countries of Australia, Austria, Belgium, Canada, Chile, Czech Republic, Denmark, France, Germany, Hungary, Ireland, Italy, Japan, Netherlands, Poland, Republic of Korea, Russia, Singapore, Slovakia, Spain, Sweden, Switzerland, Turkey, and the United Kingdom | Terminated | January 11, 2011 | 82 | 6755 | Not_Computer
Z-0846-2011 | 54638 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The device is the NeuViz Dual series Computed Tomography (CT) Scanner System. The device is identified as NeuViz Dual, PN: 989605651321. |  |  | February 10, 2010 | 2010 | Philips And Neusoft Medical Systems Co., Ltd. | Potential for the CT patient table top to become detached from the carrier due to missing washers. | Other | Neusoft Medical Systems Co., China sent an URGENT DEVICE CORRECTION letter dated Feb. 10, 2010 to their distributor in NC. On 02/12/10, The distributor sent the same letter to their consignees.The firm informed this distributor that there was report of the four screws that secure the patient table top to the carrier that did not have the proper washers in place which caused the table top to detach. If the issue would reoccur, there is a potential for the table top to fall and cause patient injury.Customers were instructed to check the patient table top fixing. If it is found that the the assembly is improper customers were instructed to top using the product and contact local Service personnel for repair. Any customer can contact Service Support Department, Neusoft Medical Systems Co, LTD at nms-service@neusoft.com or Service Support Department at helpdesk@neusoft.com for questions about this recall. |  | 26 | 26 devices | Nationwide Distribution including NC | Terminated | April 15, 2011 | 429 | 26 | Not_Computer
Z-0849-2011 | 56829 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Video Cart Multi-Specialty product labeling is as follows: |  |  | February 23, 2010 | 2010 | Linvatec Corp. dba ConMed Linvatec | ConMed Linvatec, Endoscopy Division, dba. ConMed Linvatec, Largo, Florida is recalling five products: (1.) Video Cart Premium 120 VAC (VP6501), (2.) Video Cart Multi-Specialty 120 VAC (VP6502), (3.) Video Cart Articulating Arm 120V (VP6504), (4.) Video Cart (VP8500), and (5.) Power Cord Assembly Self Retracting (VP6547). These products are being recalled due to there is a possibility that the Vid | Component design/selection | ConMed Linvatec sent URGENT MEDICAL DEVICE CORRECTION NOTIFICATION letters to their customers by FedEx. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to review their plugs for damage and determine if they were affected. Distributors were asked to contact affected customers and schedule delivery of their replacement plugs within 30 days. If the account no longer has the product in their possession, or if they have transferred the product, they were asked to respond as such via a Reply form.Customers were asked to document delivery of replacement plugs on the Sales Rep Reply Form and mail back or return by fax to 1-888-370-3062.For questions regarding this recall call the Customer Service Department at 1-888-292-0100 or e-mail Custserv1@Invatec.com. |  | 173 | 173 carts | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, AND WY and the countries of Canada, Chile, Colombia, Denmark, Korea, Mexico, Singapore, South Africa, United Arab Emirates, and Venezuela. | Terminated | March 22, 2012 | 758 | 1761 | Not_Computer
Z-0775-2011 | 56400 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | HiSpeed FX/I (2200997 with Gantry mo. #2249696) |  |  | February 01, 2010 | 2010 | GE Healthcare, LLC | A defect with the hardware and software displaying the incorrect dose data. | Radiation Control for Health and Safety Act | Important Electronic Product Radiation Warning letter was sent to customers identifying the affected product, explaining the defect and related hazards, giving instructions, and the product correction. Customers are to decrease the CTDI value by a factor of 30% when using 5, 7, 10 mm slice thickness, and by a factor of 20% when using 3 mm slice thickness for more accurate estimate of the CTDI. GE Healthcare will remedy the defect free of charge and provide an updated Operator Manual and configuration update to correct the CTDI display to all affected systems. Questions should be directed to the contact numbers provided in the letter. |  | 102 | 102 total | Nationwide Distribution | Terminated | May 21, 2014 | 1570 | 102 | Computer
Z-0838-2011 | 57271 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | AngioScore, Inc. AngioSculpt PTA Scoring Balloon Catheter; |  |  | November 16, 2010 | 2010 | AngioScore Inc. | Reports of distal bond failure which has resulted in peeling of the bond and/or detachment of the distal end of the scoring element. May result in vessel injury or embolism. | Device Design | Medical Device Recall notifications were sent to all domestic consignees on Nov 16, 2010 by overnight mail. Foreign distributors, other than Japan, were sent recall letters by e-mail or international overnight delivery on November 16, 2010. The Japanese distributor received notification and notified their customers on November 10, 2010. |  | 17682 | 17682 total units, all sizes and lots | Worldwide Distribution: USA, China, Finland, Italy, Japan, Norway, Saudi Arabia, Spain, Sweden, Switzerland, UAE, United Kingdom | Terminated | October 04, 2011 | 322 | 17682 | Not_Computer
Z-0859-2011 | 57179 | No | HGC | Obstetric-gynecologic general manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Boston Scientific brand Matrix2 360 Soft SR Coils. |  |  | October 19, 2010 | 2010 | Boston Scientific Corporation, Neurovascular, 47900 Bayside Parkway, Fremont, CA 94538 | The product appears to have premature degradation of the PGLA coating on the Matrix Coil. | Process control | US consignees were notified by letter sent by Federal Express on October 19, 2010. Customers out of the United States were notified by Federal Express if possible and by personal delivery by locally assigned sales representatives as needed. |  | 77364 | 77,364 total units | Worldwide distribution. USA, Argentina, Austria, Australia, Belgium, Bahrain, Brazil, Canada, Switzerland, Chile, China, Colombia, Serbia, Czech Republic, Germany, Denmark, Spain, Finland, France, Great Britain, Greece, Hong Kong, Hungary, Ireland, India, Israel, Italy, Jordan, Japan, South Korea, Lithuania, Latvia, Macau, Mexico, Malaysia, Netherlands, Norway, New Zealand, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovakia, Thailand, Taiwan, Turkey, Ukraine, Vietnam, and Yemen. | Terminated | January 27, 2012 | 465 | 77364 | Not_Computer
Z-0852-2011 | 57352 | No | FAD | Ureteral stent. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Cook Medical, Universa Soft Ureteral Stent and positioner, 6 Fr x 24 cm, REF USH-624-R, sterile, Cook Urological, Inc., Spencer IN. |  |  | October 18, 2010 | 2010 | Vance Products Incorporated | The firm received a customer complaint that a box labeled as Universa Firm Stents but the box actually contained Universa Soft Stents. The inner product label is correct but the outer box is wrong. | Packaging | The firm sent URGENT: RECALL NOTICES dated 10/18/2010 to its direct accounts. The notice identified the affected product and instructed firms to pull the recalled product from stock and isolate it. Customers are to complete the form and e-mail or fax the form back to Cook. Product should be returned to the recalling firm using the Return Goods Authorization number. Cook Urological will credit customer accounts upon receipt of the returned product. |  | 25 | 25 units | Worldwide Distribution -- USA, including states of WI, TX, NY, SC, AZ, TN, FL, AR, CA, AL, and WA and country of Canada. | Terminated | July 27, 2011 | 282 | 25 | Not_Computer
Z-0864-2011 | 55541 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Pioneer Plus Catheter PPlus 120 dual lumen device, inserted through a 6F introducer sheath and placed percutaneously into a peripheral vessel. The device tracks to its intended site in the vasculature over a standard length, commercially available 0.014" (0.36 mm) Rapid Exchange (RX) tracking guide wire. It utilizes an extendable, hollow Nitinol guide tip (needle) to facilitate the redirection and placement of a 0.014 (0.36 mm) Over the Wire (OTW) guide wire into peripheral vessels. Specification Developer: Medtronic, Minneapolis, MN. Manufactured by: Accellent, Laconia NH |  |  | April 30, 2010 | 2010 | Medtronic CardioVascular | Needles may fail to retract and thus cause injury. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Safety Information letter to all affected customers. The letter identified the product, the problem,and the action to be taken by the customer. Sales Representatives were to contact each account and complete an acknowledgement form indicating the Physicians had been notified. For questions regarding this recall customers were to contact their local representative. |  | 1080 | 1,080 units | Worldwide Distribution - USA, Germany, Netherlands, Slovakia, and Spain | Terminated | January 06, 2011 | 251 | 1080 | Not_Computer
Z-0865-2011 | 57457 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Groshong 3 Fr. Single-Lumen PICC Intermediate Tray with Safety Introducer, Reorder Number 7717300, Bard Access Systems, Inc., Salt Lake City, UT 84116. |  |  | December 10, 2010 | 2010 | Bard Access Systems | Prepackaged sterile latex gloves beyond their expiration date were packaged in surgical kits. | Employee error | BARD Access Systems sent an URGENT: GROSHON* 3 FR SINGLE-LUMEN PICC INTERMEDIATE TRAY WITH SAFETY INTRODUCER RECALL letter dated December 10, 2010. The letter identified the product, the problem,and the action to be taken by the customer. Customers were instructed to examine their inventory and quarantine the product subject to recall, discontinue use, and cease further distribution of any the affected product. If the product was further distributed, those customers were to be notified immediately with a copy of the recall notification letter. Any remaining product was to be returned to Bard. For information regarding this recall call 800-290-1689. |  | 195 | 195 units | Nationwide Distribution including AL, AZ, CA, GA, IA, MN, MO, NE, NJ, NY, OH, OR, SD, and WI. | Terminated | May 04, 2011 | 145 | 195 | Not_Computer
Z-0882-2011 | 57416 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Air bubble detect module, catalog 802110, Terumo Cardiovascular System Corp, Ann Arbor, MI. |  |  | December 02, 2010 | 2010 | Terumo Cardiovascular Systems Corporation | The firm identified that a component within the System 1 ABD and Occluder Modules that may malfunction. The Air Bubble Detector Module (ABD) is used with an air sensor to detect gross air bubbles in the perfusion circuit. The Occluder module is used with the Occluder head to provide a computer controlled tube clamping mechanism to regulate the flow in the venous line.The firm identified a fai | Component design/selection | Terumo Cardiovascular Systems Corporation issued an URGENT MEDICAL DEVICE CORRECTION letter dated 12/2/2010 to the Director of Cardiac Operating Room Services or Operating Room Nurse Manager. The letter identified the product, the problem, and the action to be taken. The letter stated the firm will replace all modules in the affected population, and that a representative would contact the consignee regarding the replacement schedule.Customers were to review the Medical Device Correction.Assure that all user were aware of the notice.Confirm receipt of the communication by faxing the attached Customer Response Form to the fax number indicated on the form. For questions regarding this recall call 1-800-521-2818. |  | 508 | 508 devices | Worldwide Distribution - US, Austrailia, Belgium, Canada, Chile, Columbia, Costa Rica, Dominican Republic, Guatamala, Hong Kong, Indonesia, Japan, Korea, Malaysia, Mexico, Phillipines, Saudi Arabia, Singapore, Taiwan, Thailand, United Arab Emirates, and Vietnam | Terminated | April 17, 2012 | 502 | 729 | Hardware | Device Operation | Replace component | Remove or Replace
Z-0885-2011 | 54685 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific Guider Softip TM XF Guide Catheter |  |  | December 18, 2009 | 2009 | Boston Scientific Corporation | Complaint about product quality: degradation of polymer portion of the distal section of the catheter. | Other | Boston Scientific sent an Urgent Medical Device Recall and Correction letter dated December 18, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product and fill out the Customer Instructions and REPLY VERIFICATION TRACKING FORM indicating the status of the affected product and send a copy of the form by fax to 510-624-1600 or e-mail to Joe.Nesci@beci.com. |  | 94296 | 94,296 units | Worldwide. US Domestic including US Federal Government, and foreign countries. | Terminated | December 19, 2011 | 731 | 94296 | Not_Computer
Z-0867-2011 | 54651 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Stryker Digital Capture Pro 2 device (240-050-800, CD version) and (240-050-810, DVD version) designed for digital documentation of medical procedures. Manufactured by Stryker Endoscopy, San Jose, CA. |  |  | December 11, 2009 | 2009 | Stryker Endoscopy | Device will not able to assign the correct time/date stamp for saved filed after December 31, 2009. On Jan 1, 2010, the device system date for all SDC Pro 2 units will re-start from 1 Oct 2001 and begin incrementing from that point onwards. The system date will be incorrect and users will not be able to correct this issue. Problem will not affect how pictures or videos are captured using SDC Pro | Software design | Stryker Endocoscopy sent an URGENT: Device Correction letter dated December 11, 2009, to the OR Supervisor/Risk Manager and all Stryker Sales Representatives. The letter identified the product, the problem, and the action to be taken by the Sales Representatives. Sales Representatives were to identify affected accounts. Go to customer accounts and install the upgrade for each affected unit. Upon completion of the upgrade, the sales rep was to sign the acknowledgment form indicating that all of their affected accounts received the appropriate software fix and mail to Stryker Endoscopy or fax both sides of the postcard to 408-754-8378 or scan in both sides of the postcard and email to sdcpro2recall@stryker.com.For questions regarding this recall call 408-754-2738 or 408-754-2000. |  | 3070 | 3070 units | Worldwide Distribution - USA including AL, AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, ID, IL IN, IA, KS, KY, LA, MA, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NV, NY, NC, ND, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WI, WV, and WY and the countries of Africa, Australia, Brazil, Cnanda, China, France, Germany, Hong Kong, Iberia, Italy, Japan, Korea, Latin America, Mexico, Netherlands, Poland, Puerto Rico, Scandinavia, Singapore, Switzerland, Thailand, and the UK | Terminated | January 08, 2011 | 393 | 3070 | Software | Output/Calculation | Software update | Software Update
Z-0874-2011 | 56388 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific brand Synchro2 0.14" Guide Wire, 300 cm, Standard; |  |  | June 07, 2010 | 2010 | Boston Scientific Corporation | The Guidewire has the potential to flake off polytetrafluroethylene (PTFE) particulate, which has the potential to embolize and result in infarction when introduced into a patient's circulatory system. | Other | Boston Scientific Corporation sent an Urgent Medical Device Recall letter dated June 7, 2010, by Federal Express to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use of and segregate the affected product. Complete and return the Reply Verification Tracking Form even if they did not have any affected product. Contact the Customer Service Call Center to obtain a Returned Goods Authorization (RGA) Number at (800) 811-3211, Option #1. E-mail or Fax the completed and signed Reply Verification Tracking Form to email Bob.Ware@bsci.com or the Field Action Fax (510) 624-2475.PackageShip and return the affected product along with the original Reply Verification Tracking Form by Federal Express to:Boston Scientific CorporationUS Distribution CenterBoston Scientific Marina BayCustomer Fulfillment Center500 Commander Shea Blvd.Quincy, Massachusetts 02171RGA: ____________For questions regarding this recall call (510) 624-2563. |  |  |  | Worldwide Distribution - USA including CA, CO, IN, MD, MI, OR, TX, and VA and the countries of Australia, Austria, Belgium, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Georgia, Germany, Great Britian, Greece, Honduras, Hong Kong, India, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Lithuania, Malaysia, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Puerto Rico, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Korea, Sweden, Switzerland, Thailand, Tunisia, Turkey, and Ukraine. | Terminated | January 10, 2011 | 217 |  | Not_Computer
Z-0884-2011 | 56986 | No | DSA | Patient transducer and electrode cable (including connector). | Cardiovascular | Cardiovascular | 510(k) | CARTO 3 System Interface Cable, 34-pin to 10-pin, Red, 10' - Part Number: D-1286-18; Catalog Number CR3410CT |  |  | October 05, 2010 | 2010 | Biosense Webster, Inc. | When connecting this cable to the mapping system, loss of Body Surface ECG signals may occur. This may result in inaccurate locations appearing for catheters displayed 3D on the CARTO monitor. | Device Design | Biosense Webster issued an Urgent Field Safety Notice letter dated October 5, 2010. The letter identified the product, the problem,and the action to be taken by the customer. Customers were instructed to return the affected cable, accompanied by the attached Return Authorization form.For questions regarding this recall call Biosense Webster at 909 839-8874. |  | 1692 | 1692 | Worldwide Distribution: USA, including the states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, KS, KY, MA, MD, MI, MN, NC, NE, NJ, NV, NY, OH, OR, PA, SC, TN, TX, UT, WA, WI, and WV, and the countries of Belgium, China, Israel, Japan, Singapore, and Thailand. | Terminated | May 07, 2012 | 580 | 1692 | Not_Computer
Z-0866-2011 | 57418 | No | GBX | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Merit One Step Centesis Drainage Catheter |  |  | November 17, 2010 | 2010 | Merit Medical Systems, Inc. | The catheter's valved hub may fail resulting in an inability to drain fluid. | Device Design | Merit Medical Systems, Inc. sent a Product Safety Notice letter dated November 18, 2010, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to provide a copy of the Safety Notice form to their applicable customers. Advise them to immediately stop using the affected product and to discard it.Customers were asked to complete and sign the Safety Notice forms and fax them to 1-801-208-3378 and send the original Safety Notice Form with both signatures via Federal Express 2nd day air to:Customer ServiceMerit Medical Systems, Inc.1600 West Merit ParkwaySouth Jordan, UT 84095-2416entirety and faxed them For questions regarding this recall call 801-316-4929 or Customer Service at 801-208-4365. |  | 20863 | 20,863 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, AND WV and the countries of Canada, Finland, Germany, Netherlands, Singapore, Sweden, Switzerland, and the UK | Terminated | March 23, 2012 | 492 | 20863 | Not_Computer
Z-0886-2011 | 56347 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Varian brand C-series Clinac, Trilogy and Novalis Tx, Software Versions 6.X and 7.X, Model Numbers: H14, H27, H29, HCX, Product is manufactured and distributed by Varian Medical Systems Inc., Palo Alto, CA |  |  | June 10, 2010 | 2010 | Varian Medical Systems Oncology Systems | The Auto goto or Auto setup functions in C-Series version 7 software ignore the couch angle and exceed motion zone (unless a tolerance is defined by the user); it has the potential for collision with the patient on the couch. | Other | Varian sent a Customer Technical Bulletin letter to all affected users on May 27, 2010, and was also posted to Varian's customer support site: http://www.MyVarian.com. The letter identified the product, the problem, and provided instructions on the recall. Customers were instructed to return the proof of notification (via completed "Proof of Notification Form or appropriate Postal system certified receipt of delivery to Varian. For questions regarding this recall call 650-483-3153. |  |  |  | Worldwide Distribution - USA including CA, TX, IL, GA, NJ, and NV and the countries of Australia, Belgium, Brazil, Canada, China, Denmark, France, Germany, Hong Kong, India, Italy, Japan, Middle East, Netherlands, Russia, Spain, Switzerland, and the UK | Terminated | December 23, 2011 | 561 |  | Software | Output/Calculation | Instructions | Safety Notice/Insructions
Z-0888-2011 | 51217 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems Floorstand. An accessory of the Optical Guidance Platform, Model number HZ5 for use with a charged particle accelerator (such as the Clinac 600C). |  |  | January 26, 2009 | 2009 | Varian Medical Systems Oncology Systems | Potential patient impact -- limit switches for Floorstand may be inadequate for some models and allow for potential impact. | Device Design | Varian sent an urgent device correction notification letter 01/26/2009 to all affected users, with a description of the problem and user corrective action steps via certified mail. The letter will also be distributed to Varian Sales, Marketing and Service organizations. |  | 3 | 3 units | Product was distributed to 3 consignees in Florida and Wisconsin. | Terminated | February 24, 2012 | 1124 | 3 | Not_Computer
Z-0891-2011 | 54146 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console, versions: 8.0.12, 8.0.23 SP2, 8.1 SP2, 8.1.1.3, 8.1.1.4, 8.1.2, 8.1.3, 8.1.4, 8.3, 8.6, 8.6.1, 8.6.2, 8.8.0; Model Number: H51; Distributed and manufactured by: Varian Medical Systems Inc., Palo Alto, CA |  |  | October 16, 2009 | 2009 | Varian Medical Systems Oncology Systems | Software anomaly may result in a design deficiency causing Plan Edit Settings to malfunction under specific circumstances, where the user selects to use a Cone Beam CT set up field. | Device Design | The firm, Varian Medical Systems, sent a Product Notification Letter/Field Safety Letter entitled "URGENT MEDICAL DEVICE CORRECTION" dated October 16, 2009, to all affected customers. The letter described the product, problem, actions to be taken by the customers and Varian. The customers were instructed to use caution when creating CBCT imaging field and limit use of the 'select all' button when acquiring treatment parameters; review departmental policies and procedures regarding the acquisition of any gantry angle at the 4D ITC, and advise the appropriate personnel of this letter. The firm, Varian, recommended that only vertical, longitudinal and lateral couch parameters are acquired when creating a CBCT field at the 4D ITC; notified all affected customers of this document, and will correct this issue in a future release of the 4D ITC software. If owners of $D ITC, version 8.0.12 and higher, have any questions regarding this product notification letter, please contact Varian Medical Systems Help Desk at 1-888-VARIANS (888-827-4265) for USA and Canada, +41 41 749 8844 or support-emea@varian.com for Europe;Emails for North American: support-americas@varian.com; Australia/New Zealand: support-anz@varian.com; South East Asia: seasia.apps.helpdesk@varian.com; China/Asia: china.apps.helpdesk@varian.com; Japan: Japan.Apps.Helpdesk@varian.com and Latin American: soporte.al@varian.com, and Internets for Oncology Systems- www.myvarian.com and Varian Medical Systems- www.varian.com. |  | 3629 | 3629 | Worldwide distribution: USA and countries including: Australia, Asia, China, Japan, Latin American, New Zealand , North America, South East Asia. | Terminated | July 20, 2011 | 642 | 3629 | Software | Device Operation | Software update | Software Update
Z-0890-2011 | 56326 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Treatment Planning System, a subsystem of the CyberKnife Robotic Radiosurgery System radiation therapy device, manufactured by Accuray Inc., Sunnyvale, CA. |  |  | July 06, 2010 | 2010 | Accuray Inc | Latches responsible for securing the cover to the device may come lose if not properly secured and the cover may unexpectedly drop off. A design change is planned. | Device Design | Accuray sent an "URGENT DEVICE CORRECTION" letter dated July 6, 2010, to all customers. The letter described the product, problem and actions to be taken by customers. The customers were instructed to assure the LINAC cover is secure following reattachment and adhere to and follow the instructions included in the letter. Note: A new design is being developed, and will be distributed to user base once available.If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at +1-877-668-8667 (USA) or +1-408-716-4700 (non-USA) or customersupport@accuray.com. |  | 207 | 207 units | Worldwide distribution: US including: AK, AL, AR, AZ, CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, NC, ND, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, WI and WV, and countries including: Japan, China, Vietnam, Taiwan/China, South Korea, Hong Kong/China, Thailand and Malaysia. | Terminated | October 04, 2011 | 455 | 207 | Not_Computer
Z-0887-2011 | 57338 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | TiTLE 2 Polyaxial Spinal System, Counter Torque Assembly Tool, Part ET1132-01, Zimmer, Warsaw, IN. |  |  | October 21, 2010 | 2010 | Zimmer Inc. | There is a chance that the ball plunger could fall out during surgery and fall into the surgical site. There could be a delay in surgery to remove the ball plunger. If the problem was undetected, the ball plunger could remain in the patient causing pain and potentially necessitating subsequent surgery for removal. | Other | Zimmer sent an URGENT MEDICAL DEVICE RECALL LETTER dated October 21, 2010, to all affected customers. The letter identified the product, the problem, the action to be taken by the customer.Customers were instructed to:Inspect their inventory, stop using the device, and place it in quarantine. Contact Zimmer Customer Service at 1-800-774-6368 to obtain a return authorization number. Complete the Inventory Return Certification form and fax a copy to (952) 837-6958, or scan and email to qualitycompliance@zimmer.com.Return the recalled product along with the original completed Return Certification Form to:Zimmer Spine, Inc.Attn: RA Receiving5301 Riata Park Ct.Austin, Texas 78727For questions regarding this recall call Zimmer Spine Customer Service at 1-866-774-6368. . |  | 68 | 68 | Nationwide Distribution including AL, AZ, CA, CO, FL, MI, MN, NC, PA, OH, OK, TN, TX, WA, and WI. | Terminated | October 12, 2011 | 356 | 68 | Not_Computer
Z-0843-2011 | 57293 | Yes | FJK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Medical CombiSet BVM Hemodialysis Blood Tubing Set |  |  | November 19, 2010 | 2010 | Fresenius Medical Care Holdings, Inc. | Hemodialysis Blood tubing Sets may develop kinking of the arterial line. | Device Design | Fresenius Medical issued recall notification dated 11/19/10, "URGENT RECALL FMCNA CombisetÂż BVM Hemodialysis Blood Tubing Sets (Part Numbers 03-2695-9 and 03-27957)", by certified mail with return receipt. Customers have been instructed to check their stock immediately to determine if they have any affected lots on hand. If customers have the affected lots, they have been instructed to discontinue use immediately and place the product in a secure segregated area for return to Fresenius Medical Care North America. Customers have been instructed to contact Fresenius Medical Care Customer Service Team Representative for instructions on how to return the recalled product. |  | 190080 | 190,080 units (7920 cases x 24 units) | Distribution Nationwide and Canada | Terminated | November 03, 2012 | 715 | 190090 | Not_Computer
Z-0837-2011 | 57099 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook 400ES Safety Infusion System, Model Number 621-400ES. For use with B. Braun Medical Inc. Horizon Pump IV Sets |  |  | August 24, 2010 | 2010 | B Braun Medical, Inc | Possible pump stop with audible alarm, but no visual error indication that the pump is not delivering the medication. Under certain conditions in a network environment, Management Processor may become non-responsive, ceasing normal operation. | Component design/selection | B. Braun Medical Inc. notified consignee by letter on August 24, 2010. |  | 2107 | 2,107 Pumps (582 Pumps to Winchester) | Virginia and New York | Terminated | May 29, 2014 | 1374 | 2107 | Software | Alarm/Message | Notification | Safety Notice/Insructions
Z-0889-2011 | 57414 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | TandemHeart Transseptal Cannulas-EF set (THTC-EF) - Dilator assembly |  |  | December 02, 2010 | 2010 | Cardiac Assist, Inc | Component (Dilator of TandemHeart Transseptal Cannula Set-EF) does not meet specifications | Other | CardiacAssist, Inc. sent a Voluntary Recall letter dated November 12, 2010, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignees. Consignees were instructed to examine inventory for the affected product and remove from inventory and identify it as a recalled product by applying the enclosed "CAUTION: DO NOT USE" sticker(s) and quarantine to prevent usage.Consignees were to complete the form included with the letter and fax a copy of the form to (412) 963-9739 to coordinate the return and replacement of any affected product. For questions regarding this recall call (412) 963-7770, ext. 271. |  | 164 | 164 | Nationwide distribution including AZ, CA, CO, CT, DC, FL, GA, IL, LA, MA, MD, MI, MN, NC, NH, NJ, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, and WI. | Terminated | December 01, 2011 | 364 | 164 | Not_Computer
Z-0853-2011 | 53047 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | RoboCouch Patient Support System, model number 025007, a subsystem of the Cyberknife Robotic Radiosurgery System, manufactured by Accuray Incorporated, Sunnyvale, CA. |  |  | July 21, 2009 | 2009 | Accuray Inc | Mechanical Failure-- Patient Support System's table top may unexpectedly descend several inches. | Device Design | Consignees were notified verbally about the issue when an inspection was scheduled. |  | 26 | 26 | Worldwide Distribution -- USA, including states of CA, FL, VA, AL, CT, MA, NC, TN, PA, RI, and NE and countries of Turkey, France, UK, Saudi Arabia, and Switzerland. | Terminated | January 12, 2011 | 540 | 26 | Not_Computer
Z-0892-2011 | 56350 | [6 empty] | Varian brand C-Series Ximatron Radiotherapy Treatment Stimulator, |  |  | January 19, 2010 | 2010 | Varian Medical Systems Oncology Systems | There is a potential for failure of a lead screw drive system or image intensifier lead screw drive assembly, if normal maintenance inspections are not followed correctly. | Other | Varian sent an URGENT MEDICAL DEVICE CORRECTION letter via certified mail receipt on January 19, 2010, to all direct consignees. The letter identified the product, the problem, and the action the consignee should take.Consignees were instructed to have maintenance personnel perform the recommended measurements as described in the Varian Customer Technical Bulletin.Ensure that all measurements made as part of the procedure are recorded in the 'wear table' that is part of the procedure.Previous wear measurements should be compared to new measurements and any large increase in wear should be investigated immediately.Pay special attention to the "Recommendations based on wear measurements" table.Ensure that new wear measurements comply with the operating parameters described in the table. For any questions consignees are to contact Varian Medical Systems Customer Support Services at one of the locations described in the CTB.For questions regarding this recall call (650) 424-5731. |  |  |  | Worldwide Distribution including USA, Australia, New Zealand, Europe, Southeast Asia, China, Japan, Latin America, and Canada. | Terminated | February 24, 2012 | 766 |  | Not_Computer
Z-0894-2011 | 57463 | No | OFM | Impedance phlebograph. | Cardiovascular | Cardiovascular | Enforcement Discretion | Blood Center Kit w/Chloraprep One-Step, Sterile/Disposable/Single Patient Use, Reorder No. Q921CL, Manufactured for : Blood Center of Wisconsin, Milwaukee, WI. |  |  | December 06, 2010 | 2010 | Continental Medical Labs,Inc | Continental Medical labs, Inc, is recalling various lots of Blood Center Kit and Port VAD Kits because Covidien reported that testing could not pick up an Over-Sulfated Chondroitin Sulfate (OSCS) contaminant in the purified heparin sodium USP. | Material/Component Contamination | Covidien sent an URGENT PRODUCT RECALL letter dated November 12, 2010, to their distributors. Continental Medical Labs,Inc telephoned on 12/06/2010 and later sent a letter dated December 7, 2010, with an enclosed copy of the recall notice. The letters were addressed to McCormick Medical Distribution, Inc and to Medico Mart, Inc. The letter described the problem and product affected. Advised consignees to inform all their customers to return the product as well as the distribution information as soon as possible so CML could notify Covidien for replacements.For questions regarding this recall call (262) 723-6715, ext. 313. |  | 900 | 30 cases, 30 kits per case | Nationwide Distribution including WA and WI | Terminated | July 05, 2012 | 577 | 900 | Not_Computer
Z-0899-2011 | 51401 | No | GBS | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Concentric 7F Balloon Guide Catheter (percutaneous catheter per 21 CFR 870.1250), REF 90072, manufactured by Concentric Medical, Inc., Mountain View, CA |  |  | February 27, 2009 | 2009 | Concentric Medical Inc | Product malfunction-- hub may leak at connection to the shaft due to insufficient adhesive curing. | Nonconforming Material/Component | Concentric Medical products sent an URGENT VOLUNTARY MEDICAL DEVICE REMOVAL letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer.Customers were instructed to examine their inventory and identify any affected product.Quarantine the product and call Concentric Medical Customer Service at 877-471-0076 (USA toll-free) or 650-938-2100 to receive instructions for returning the devices.For questions regarding this recall call 877-471-0076. |  | 36 | 36 units affected. | Worldwide Distribution - USA including AL, FL, GA, IL, MN, NY, OH, OR, PA, TX, VA, WA, and WV and one consignee in Turkey and one in Switzerland | Terminated | January 13, 2011 | 685 | 36 | Not_Computer
Z-0897-2011 | 51274 | No | INQ | Powered table. | Physical Medicine | Physical Medicine | 510(K) Exempt | Precedence SPECT/CT System, Precedence 16 3/8, Upper Patient Pallet, A single photon nuclear medicine and image from an X-ray computed tomography system, |  |  | January 26, 2009 | 2009 | Phillips Nuclear Medicine | A problem has been identified with some upper patient pallets on Precedence SPECT/CT System, which may reduce its load bearing capacity and may result in its cracking. If this were to happens, could pose a risk to patient. | Process change control | The firm, Philips Healthcare, sent an "Urgent - Field Safety Notice" dated January 26, 2009 to all consignees/customers. The notice described the product, problem and action to be taken by customers and by Philips. The customers were instructed to reference their system's System Operation and System information Manuals regarding loading and unloading of patient; monitoring the patient during a study ,and careful monitoring during unsupported (cantilevered) table translation. Philips has ordered replacement parts and they will have a Philip Healthcare Field Service Engineer contact and visit the customers site within two months to replace the affected upper patient pallet.If you need any further information or support concerning this issue, please contact your local Philips representative at: Customer Care Center (1-800-722-9377, option 5: Diagnostic Imaging, option 5: Radiation Therapy, option 3: and Oncology or your local Philips Healthcare office. |  | 2 | 2 units | International distribution only: Mexico and Saudi Arabia. | Terminated | January 13, 2011 | 717 | 2 | Not_Computer
Z-0896-2011 | 54718 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Acuson Antares Ultrasound System |  |  | September 30, 2009 | 2009 | Siemens Medical Solutions USA, Inc. | Software safety guard measures on the device malfunctioned to prevent the temperature from going over its limits when the device is used. | Software design | Field notification in the form of a Customer Safety Advisory Notice was initiated on 09/30/2009 by the firm, and the notice was mailed to all affected customersSiemens sent a CUSTOMER SAFETY ADVISORY NOTIFICATION letter on September 30, 2009. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disconnect and reconnect the V5Ms transducer between uses until their local Customer Service Engineer installs the software update to their system. For any questions customers were to contact their local service support person.For questions regarding this recall call 650-694-5398. |  | 452 | 452 | Worldwide Distribution - United States (nationwide) and the countries of Australia, Austria, Azerbaijan, Brazil, Belarus, Canada, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Guatemala, Croatia, India, Indonesia, Italy, Jordan, Japan, Latvia, Malaysia, Mexico, Moldova, Netherlands, New Zealand, Norway, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Taiwan, Turkey, Ukraine, United Kingdom, Uruguay, Uzbekistan, Venezuela, and Vietnam. | Terminated | June 18, 2012 | 992 | 452 | Software | Physical Safety Hazards | Instructions on reconnecting the transducer | Safety Notice/Insructions
Z-0898-2011 | 57218 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow-Howes (TM) Multi Lumen Central Venous Catheterization Set with Blue FlexTip(R) Catheter. Catalog #: CV-15703. |  |  | November 09, 2010 | 2010 | Arrow International Inc | One lot of product was inadvertently mislabeled with a product lidstock identifying the product as a triple-lumen catheter with product number CV-15703, but a double-lumen catheter with part number CV-16702 was actually packaged in the kit. | Process control | The recalling firm issued Urgent Medical Recall letters dated 11/9/10 to inform their customers of the problem and the need to return any unused product. Customers are to check their stock for the affected product, contact Arrow Customer Service for a Return Authorization Number, and complete the enclosed Recall Acknowledgement &amp; Stock Status Form. For customers with no affected product in stock, the form should still be completed and also faxed to Arrow Customer Service. Affected product should be returned along with the original completed form to Arrow International, Inc. Arrow will replace the affected product once it has been returned. Additional information or clarification can be obtained from Arrow Customer Service at 1-919-361-3964. |  | 1150 | 1150 units | International Distribution Only -- Argentina and Columbia. | Terminated | September 29, 2012 | 690 | 1150 | Not_Computer
Z-0895-2011 | 57462 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Disc Space Distractor |  |  | December 07, 2010 | 2010 | Stryker Spine | Stryker Spine has received reports of the Disc Space Distractors failing during surgical procedures. | Other | Stryker Spine sent an URGENT PRODUCT RECALL letter dated December 7, 2010, via Fedex to all affected customers.. The letter identified the product, the problem, and the action they should take. Customers were instructed to examine their inventory and hospital locations to identify the product.Reconcile IBD Disc Space Distractors with the catalog and lot numbers referenced. Complete the Customer Response Form and distribution sheet and fax a copy of the Customer Response Form to Regulatory Compliance at 201-760-8370.Return the affected product using the enclosed pre-paid mailing label to Regulatory Compliance,Stryker Spine 6 Pearl CourtAllendale, New Jersey 07401For questions regarding this recall call 201-760-8206. |  | 21 | 21 units | Nationwide Distribution including AZ, FL, MA, NY, PA, and TX | Terminated | November 14, 2012 | 708 | 21 | Not_Computer
Z-0900-2011 | 57436 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Horizon Cardiology ECG Management |  |  | June 16, 2010 | 2010 | McKesson Provider Technologies - Medical Imaging Group | When importing ECG files from Mortara ECG carts using DICOM interface, the Horizon Cardiology ECG Management System does not display the amplitudes of the waveform properly. | Device Design | McKesson Medical sent an advisory notice letter dated June 16, 2010, to all affected customers. Advisory notice ADV01646 was distributed to 13 affected clinical sites from August 30, 2010 to September 28, 2010. The letters identified the product that is affected, the circumstances under which patient safety issue may occur, the risk, and the immediate actions that need to be taken along with a temporary fix. Customers were instructed to use the ECG cart printout to verify results of procedures acquired by Mortara ECG carts. Users should be aware of the problem when referrring to past ECG procedures and when performing a visual comparison with ECG procedures acquired by Mortara ECG carts.Customers should disable the Glasgow Serial Comparison for ECG procedures.McKesson Support would contact the customer to coordinate the deployment of the software update on all the affected applications and Horizon Cardiology client applications. For any questions regarding this recall, customers should call the Customer Support department at 1-866-777-0202 and speak to their Support Manager. |  | 16 | 16 units | Nationwide Distribution including AZ, CA, FL, GA, MA, PA, UT, and WI. | Terminated | August 17, 2012 | 793 | 16 | Software | Display/Image | Software update | Software Update
Z-0904-2011 | 52716 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | VISU 200 on S8 Floor Stand, Manufactured by Carl Zeiss Surgical GmbH, Germany |  |  | May 18, 2009 | 2009 | Carl Zeiss Meditec Inc | Potential Retina Damage-- An error in the user manual may result in misapplied filter and injury during treatment. | Labeling mix-ups | Carl Zeiss Meditec Inc sent a SAFETY ALERT letter dated May 18, 2009, via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were reminded to use the correct retinal protection filter (GG475) when using the OPMI VISU 200 system during ophthalmic surgery.Verify the position of the retinal protection filter and inform their staff of the customizable feature. Verify that the correct filter was in place. Customers were asked to complete the acknowledgement form indicating receipt of the alert. Sign and return the form to:Carl Zeiss Meditec in Dublin, California via the enclosed self-addressed, pre-paid Federal Express envelope.For any questions customers should call (925) 557-4422. |  | 2540 | 2540 units | Worldwide Distribution including USA (nationwide) and the UK | Terminated | February 09, 2011 | 632 | 2540 | Not_Computer
Z-0901-2011 | 57377 | No | LSZ |  |  | Anesthesiology | PMA | One Cat. Life Pulse High Frequency Ventilator, Catalog Number: 203A, Manufactured by: Bunnell Incorporated Salt Lake City, Utah, Rx Only. |  |  | October 25, 2010 | 2010 | Bunnell, Inc. | Product software coding error may compromise patient safety. | Software design | Customers were notified by an "Urgent: Medical Device Recall or Correction" letter starting 10/22/10. The letter identified the product and described the issues. Customers were given the option of either returning their ventilator(s) for rework or to receive installation/re-test instructions with an upgrade kit to rework themselves. The letter states that all Life Pulse ventilators in clinical use must be upgraded and there is no charge for the software upgrade. |  | 660 | 660 total | Worldwide Distribution -- USA, Australia, Canada, Czech Republic, Korea, Malaysia, Trinidad &amp; Tobago, and Uruguay. | Terminated | April 17, 2012 | 540 | 660 | Software | Physical Safety Hazards | Return/Update | Remove or Replace
Z-0903-2011 | 57021 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PROCEED Surgical Mesh; Product Code PCDG1; :Lot BMG618; Ethicon, Inc. a Johnson &amp; Johnson Company. |  |  | October 18, 2010 | 2010 | Ethicon, Inc. | There is a potential for delamination in one lot of Proceed surgical mesh. | Nonconforming Material/Component | Ethicon sent Urgent Voluntary Product Recall letters via Fed Ex on October 18, 2010. The letters described the issue with the affected product and instructed customers not to use it. They are to immediately remove the recalled product from their inventory and return it according to the instructions provided. A reply card is attached and must be filled out and returned to Stericycle whether or not the customer has product in their possession. Separate correspondence has been sent to end customers stating that if they have purchased the recalled product through a distributor, to return the product directly to Stericycle. Replacement product will be provided at no additional charge for all recalled product. |  | 87 | 87 products | Worldwide Distribution | Open, Classified |  |  | 87 | Not_Computer
Z-0909-2011 | 51550 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife System Robotic Radiosurgery System, (Robocouch Patient positioning System) Synchrony Respiratory Tracking System, a Radiation therapy device, Model Numbers: 026430, 028186, Product is manufactured and distributed by Accuray, Inc., Sunnyvale, CA |  |  | March 16, 2009 | 2009 | Accuray Inc | The product has the potential for the extension coupling of the Synchrony extension mounting hardware to become loose over time. The assembly weighs approximately 17 pounds; an unexpected contact with users or patients may result in serious injury if the boom descends unexpectedly. | Device Design | Accuracy Inc., sent an URGENT ADVISORY NOTIFICATION letter, dated March 16, 2009, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their system to see if it was affected. Customers were instructed to contact Accuracy Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700 or e-mail customersupport@accuracy.com to schedule service. |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Turkey, Russia, France, Italy, the Netherlands, the UK, Spain, Greece, Switzerland, Germany, India, Taiwan, China, Hong Kong, Korea, Thailand, Malaysia and Japan | Terminated | January 19, 2011 | 674 |  | Not_Computer
Z-0910-2011 | 53778 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | C-Series Clinac, Trilogy and Novalis Tx used for SRS treatments, Model Numbers H14, H18, H27, H29, HCX, manufactured by Varian Medical Systems, Palo Alto, CA. |  |  | October 08, 2009 | 2009 | Varian Medical Systems Oncology Systems | Product may deliver radiation treatment to areas larger than intended,to healthy tissue. | Nonconforming Material/Component | Varian sent an URGENT MEDICAL DEVICE CORRECTION letter dated October 8, 2009, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Varian recommended to consignees that test runs be performed prior to actual patient treatment to ensure that the entire system is behaving as intended and that appropriate quality assurance procedures should be followed prior tothe delivery of any treatment plan involving radiation therapy. Operators should monitor the patient and the treatment console continuously during treatment. For questions call USA and Canada - 1-888-827-4265, Europe +41 41 749 8844. |  | 5208 | 5208 units | Worldwide Distribution | Terminated | February 24, 2012 | 869 | 5208 | Not_Computer
Z-0914-2011 | 56624 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Instruments, SmartPump Tourniquet System, Report Printer Kit, REF 5920-012-000, Rx only, Stryker Instruments, Kalamazoo, MI (USA), 49001, (269) 323-7700, (800) 253-3210. |  |  | April 13, 2010 | 2010 | Stryker Instruments Div. of Stryker Corporation | This recall is related to the Electri-cord recall. There is a risk when excessive force is applied to the plugs, a fracture can develop on the prongs inside the molded section of the plug. Over time the fracture may lead to arcing inside of the molded section of the plug and a build up of heat which can cause charring, electric shock or fire. To this date there have been zero complaint for this | Component design/selection | Stryker Instruments sent an Urgent - Medical Device Recall Notification letter dated April 13, 2010, to all affected customers. The letter identified the product, the problem,and the action to be taken by the customer. Stryker notified customers again on May 24, 2010, and again on September 14, 2010, via email, fax, or telephone.Customers were instructed to locate the units and visually inspect each affected cord. Remove the affected cord and dispose of it using their healthcare facilities disposal procedure. If the cord did not show signs of damage, the unit could remain in service until a replacement cord was available. Customers were to return the enclosed business reply form to confirm receipt of the notification letter and identify how many affected units were currently in their inventory. Stryker instruments would send out replacement cords.If the customer loaned or sold any of the units listed in the notification, the customer was to forward a copy of the notice to the new users and advise Stryker of the new user's name and their new location. If products were disposed of and were no longer in use, the customer was to advise Stryker of their obsolescence by providing them with their serial numbers. For questions regarding this recall call 800-800-4236, ext 3808 or ext 3584 Monday -Friday 8am-5pm (EST). |  |  |  | Worldwide Distribution - USA (nationwide) and the countries of Portugal, Canada, China, Taiwan, France, Germany, Greece, Spain, India, Italy, Korea, Switzerland, Singapore, and England. | Terminated | May 16, 2011 | 398 |  | Not_Computer
Z-0906-2011 | 56871 | No | FGS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | InfraVision IR Illuminator Console, model number 220-180-521, manufactured by Stryker Endoscopy San Jose, for use with the Stryker InfraVision system. |  |  | September 29, 2010 | 2010 | Stryker Endoscopy | Product may not adequately visualize IR light emitted from U- and E- kits due to a manufacturing error. | Nonconforming Material/Component | The firm, Stryker Endoscopy, sent an "Urgent: Device Correction" letter dated September 28, 2010, t o all consignees/customers via a Fed Ex Mailer that included notification letter, acknowledgment receipts and updated instructions for use. International Stryker sites received product advisory notification. Stryker reps at those sites are responsible for notifying customers. The letter described the product, problem and actions to be taken by the customers. The customers were instructed to discard any old IFUs from the InfraVision Console and ensure that they only use InfraVision compatible scopes; read the newly updated IFU provided with the letter; and complete and sign the enclosed "acknowledgement of receipt" form and fax to (408) 754-8378 or scan in or email to infravision@stryker.com. Note: Not returning the "acknowledgement of receipt" card will result in additional notices.Please contact your local Stryker sales representative to obtain information on scopes that are specific to the InfraVision system.If you have any questions, please contact us at 1-800-624-4422 or email at infravision@stryker.com. |  | 861 | 861 units. | Worldwide distribution: USA and countries including: Australia, Africa, Brazil, Canada, China, France, Greece, India, Italy, Korea, Latin America Mexico, New Zealand, Portugal, Romania, Singapore, Sweden, Switzerland, Taiwan, Thailand, and United Kingdom. | Terminated | February 09, 2011 | 133 | 861 | Not_Computer
Z-0907-2011 | 56773 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE and ONCOR Digital Linear Accelerators configured with the 160 leaf Multi-Leaf Collimator (MLC) and MEVATRON and PRIMUS Digital Linear Accelerators configured with the High Performance Defining Collimator (HPD); Product is manufactured by Kemnath, Germany-Siemens AG OCS and Siemens AG, Erlangen, Germany-Siemens AG OCS and distributed by Siemens Medical Solutions, Oncology Care Systems, Concord, CA; |  |  | August 13, 2010 | 2010 | Siemens Medical Solutions, Oncology Care Systems, 4040 Nelson Ave., Concord, CA 94510 | The potential safety issue with the 160MLC or the HPD collimator of mechanical noises and/or sticking of the collimator when rotating the collimator. The potential for mistreatment to a patient should the collimator not be aligned correctly if the bearings are damaged. | Pending | Urgent: Medical Device Correction Safety Advisory Notice letters were sent by certified mail commencing August 3, 2010. The letter described the issue and when it occurs, the measures users could take, and how Siemens would address the issue. It is recommended that customers include a check of the field alignment at 90 and 270 degrees gantry angle in the daily check prior to start of patient treatments. A copy of the Safety Advisory Notice should be included in the Owner's Manual and questions should be directed towards the local Siemens Service Engineer or Application Specialist. |  | 1167 | 1167 | Worldwide Distribution. | Terminated | January 31, 2012 | 536 | 1167 | Not_Computer
Z-0911-2011 | 56772 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian Medical Systems, Eclipse Client software versions with build numbers from 7.3.10 through 8.6.15 (i.e. excluding 8.6.17); Model Number: H48; Radiation Treatment Planning System |  |  | August 17, 0010 | 10 | Varian Medical Systems, Palo Alto, CA | Possible incorrect dose calculation: An anomaly in certain Eclipse versions where, under certain specific sequences of user actions, the dose distribution may fail to be re-calculated even though changes have been made which affect dose. | Software design | Varian Medical Systems Urgent Medical Device Correction letters were sent to all affected users on August 6, 2010. Letters were also sent to Varian Sales, Marketing and Service organizations to inform them of the recall. |  | 7019 | 7019 | Worldwide Distribution. | Terminated | July 21, 2011 | 730823 | 7019 | Software | Device Operation | Software update | Software Update
Z-0919-2011 | 54229 | No | JFC | Intermittent pressure measurement system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Product labeling of the X3 Display reads in part: |  |  | July 20, 2009 | 2009 | Xsensor Technology Corporation | The plastic enclosures for the X3 Display and X3 PRO Pressure Mapping System has a UL94 HB flammability rating while the medical equipment specification UL 60601-1 requires V-2 rated plastic for electronic enclosures. | Component change control | On July 6, 2009, Xsensor sent an "URGENT: MEDICAL DEVICE RECALL" notifications to their consignees/customers.The consignees/customers were informed that the incorrect material of the plastic enclosures was used. The incorrect plastic had a UL94 HB flammability rating while the medical equipment specification UL 60601-1 required V-2 rated plastic for electronic enclosures. The firm provided prepaid shipping box to consignees/customers with instructions for returning the components. The components were refurbished and returned to consignees for free of charge.For further information please contact XSENSORs Customer Support at 1-866-927-5222. |  | 66 | 33 units of X3 Display and 2 units of X3 Pro | Devices were distributed to one distributor, medical facilities, and rehab centers through out the US, including following states:AL, AR, AZ, CA, CO, FL, GA, IA, IL, KS, ME, MI, MN, MO, NC, NY, OH, OK, OR, PA, SC, TN, TX, VA, VT, and WI. | Terminated | January 31, 2011 | 560 | 66 | Not_Computer
Z-0920-2011 | 57535 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ultra Surgical Gown, Large with Polypropylene Fabric and Towel, Reference number 95111-35, Single use only, Sterile, Made in Honduras. |  |  | December 23, 2010 | 2010 | Kimberly-Clark Corporation | Product was not properly sterilized. | Pending | Kimberly-Clark contacted all ten (10) Distributors with a Distributor Recall Letter (sent FedEx Priority overnight) on 12/23/2010, as well as via telephone. This letter instructs the Distributor to extend the Recall to end-user level by notifying each of their customers of the Recall via the provided Customer Recall Letter. Both the Distributor Letter and Customer Letter include Response Sheets that allows for both Distributor and Customer (end-user) to indicate the quantity of affected product they have in their inventory or indicate whether they do not have any affected product. The Distributor is instructed to immediately cease further shipment of the product and to physically quarantine the product, and fax the completed Response Sheet to our firm'sRecall Coordinator within 5 business days of receipt of the letter. The Customer is instructed to immediately cease further use of the gowns and physically quarantine the product, and fax the completed Customer Product Response Sheet to Kimberly Clark's Recall Coordinator within 5 business days of receipt of the letter. Upon receipt of the Response Sheets, the firm will provide instructions for the return of the impacted gowns for delivery to a quarantine location at our warehouse. |  | 200 | 200 cases (32 units per case) | Nationwide Distribution -- CA, CT, LA, NJ, NM, RI, CO and TX. | Terminated | July 27, 2011 | 216 | 240 | Not_Computer