Z-1096-2013 | 64725 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare LOGIQ S8 (with software revisions R1.1.1 and R1.5.1) and LOGIQ S7 (with software revisions R1.0.1, R1.0.2 and R1.0.3) diagnostic ultrasound system with the S4-10-D Ultrasound Probe. The device is intended for use in ultrasound evaluation. |  |  | February 08, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue due to the S4-10-D Ultrasound Probe used in conjunction with the LOGIQ S8 and LOGIQ S7 Ultrasound systems. It was also discovered that this issue also leads to a specification issue of the acoustic output of this probe. There is a potential for a skin burn when the flow model in the cardiac application is activated with the S4-10-D Ultraso | Software design | Consignees were sent on 2/8/2013 a GE Healthcare " Urgent Medical Device Correction" letter dated February 8, 2013. The letter was addressed to Hospital Administrators / Risk Manager, Biomedical Engineering and Managers of General/Diagnostic Imaging. The letter described the Safety Issue, Safety Instructions, Affected Product details, Product correction and Contact Information. US customer contact number is (1)800-437-1171. |  | 2061 | 2061 | Worldwide Distribution, USA, including the states of AL, AZ, AR, CA, CO,K CT, DC, FL, GA, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT NE, NV N NJ, NM, NY, NC, ND, OH, OK, PA, RI ,SC, SD, TN, TX, VA, WA, WI and the countries of AUSTRALIA, AUSTRIA, BRAZIL, CANDA, CHILE, CHINA, COSTA ARICA, CROATIA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY INDIA, INDONESIA ITALY JAPAN, KOREA, LATVIA, LEBANON, MALAYSIA, MEXICO, MOLDOVA, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PHILIPPINES, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA , SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VIETNAM. | Terminated | November 21, 2013 | 286 | 2061 | Not_Computer
Z-1101-2013 | 64670 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MRT-1504/A4 - Titan X MRI System |  |  | October 15, 2012 | 2012 | Toshiba American Medical Systems Inc | It was found that there is a problem with the LCD monitor used in the console. The LCD monitor maybecome blank during use. Specifically, the backlight of the LCD monitor may turn off while the systemis in use, causing the LCD monitor to become blank. However, the other parts of the system continue tooperate normally, and the risk of a serious health hazard is therefore very low. The LCD Moni | Nonconforming Material/Component | Toshiba sent an Urgent Medical Device Correction letter dated October 1, 2012 via USPS return receipt mail to all affected customers . The letter identified the affected product, problem and actions to be taken. The letter contains a return reply form that is to be faxed to TAMS for retention. For questions contact your local Toshiba representative at 800-521-1968. |  | 27 | 27 units | USA Nationwide Distribution including the states of : IL, NH, MA, MI, MO, PR, NV, NM, MN, ID, NY, OH, NC, AZ, MS, LA, OR, PA, CA, and CT. | Terminated | November 06, 2013 | 387 | 27 | Not_Computer
Z-1052-2013 | 64308 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ShapeMatch Cutting Guides |  |  | November 08, 2012 | 2012 | Stryker Howmedica Osteonics Corp. | With regards to ShapeMatch Cutting Guides using MRI Imaging, Stryker Orthopaedicshas become aware that there are potential issues associated with internal processesfor planning cases and producing ShapeMatch Cutting Guides. | Device Design | Stryker sent Product Correction Bulletins on 11/8/2012 dated 11/9/2012 and 11/19/2012 to all affected customers. Stryker sent an updated Urgent Medical Device Product Notification letters and Acknowledgement letters dated 1/24/2013 through 1/28/2013 to all affected customers via Fed Ex with return receipt: The letter identified the affect product, problem, potential hazards, risk mitigating factors and actions to be taken. For questions call 201-831-6878.Stryker issued press release on 4/10/2013. |  | 7868 | 7868 units | USA Nationwide Distribution | Open, Classified |  |  | 7868 | Not_Computer
Z-1074-2013 | 64624 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | All OneTouch Verio IQ Blood Glucose Meters sold as: |  |  | March 11, 2013 | 2013 | Lifescan Inc | The Verio IQ meter will shut off and revert to set up mode at glucose values above 1023 instead of displaying EXTREME HIGH GLUCOSE. | Software design | Lifescan sent an Urgent Medical Device Voluntary Recall letter dated March 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to identify and hold all affected product in inventory and discontinue distributing. The letter states that a Lifescan sales representative will contact customers to arrange for pick up and replacement. |  | 710413 | US 710,413 | USA Nationwide Distribution | Terminated | August 03, 2015 | 875 | 710413 | Not_Computer
Z-1104-2013 | 64693 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Aplio 500/400/300 Diagnostic Ultrasound System (TUS-500/400/300). |  |  | December 11, 2012 | 2012 | Toshiba American Medical Systems Inc | Toshiba Medical Systems has received reports of startup issues with the Toshiba TUS-500 andTUS-300 ultrasound systems. The startup issue is predominant when attempting to bring thesystem up from standby mode when the system is used remotely. It has been determined that thismay potentially cause damage to the system. | Device Design | Toshiba sent an "URGENT: MEDICAL DEVICE CORRECTION" letter dated December 11, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at (800) 421-1968 for questions regarding this letter. |  | 233 | 233 units | Nationwide Distribution. | Terminated | April 10, 2014 | 485 | 233 | Not_Computer
Z-1103-2013 | 63856 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integra XKnifeÂż Stereotactic Radiosurgery and Radiotherapy Treatment Planning, |  |  | December 10, 2012 | 2012 | Integra LifeSciences Corp. | Depending on the system configuration, a software error message in versions 5.0.1 and 5.0.2. occurs if a beam plan is transmitted from XKnife using DICOM-RT. | Software design | Integra initiated a voluntary recall on December 10, 2012 by providing a written Urgent: Medical Device Recall Notification to US consignees who have experienced the software issue due to the configuration of their system. The letter identifies the affected product, as well as, advises customers of the nature of the issue and how it affects treatment records.- Advises customers that they will be contacted when V 5.0.3 software, where the software error has been addressed, is available to schedule an upgrade.- Requests that customer consignees complete and return the Product Recall Acknowledgment Form.If customer have additional questions, they should contact the firm's service hotline at 1-888-772-7378. |  | 12 | 11 XKnife V5.0.1 and 1 XKnife V5.0.2 | Distributed in the states of IL, ID, LA, MA, NJ, NY, OH, PA, and SD. | Terminated | August 01, 2013 | 234 | 12 | Computer
Z-1102-2013 | 64547 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Intellispace Portal (ISP) Philips Medical Systems (Cleveland), Inc., |  |  | February 12, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips was notified of a problem using the IntelliSpace Portal with software version 4.0.2 with the URL PACS integration. Philips was informed that when loading a study to an analysis application while using URL PACS integration configuration, after another study was opened by a review application, the Analysis monitor is updated with the study of the current patient, but the Review monitor conti | Software design | Philips sent a Urgent Medical Device Correction letter dated February 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They advised their customers that when reviewing two disparate image datasets, not to only be focused on the actual image content but to also be alerted to the patient's study name and/or ID prior to interpretation.They recommend to manually exit the Review and Analysis applications before loading a new study to Intellispace Portal v4.0.2 from the PACS.For North America and Canada contact the Customer Care Solutions Center at (1-800) 722-9377. |  | 313 | 313 units | Worldwide Distribution - (USA) Nationwide Distribution including the states of AL, AR, AZ, CA, FL, GA, IA, IN, LA, MI, MN, MO, NC, NJ, NY, OH, OK, OR, PA, TN, TX, WA, WI &amp; WV. and the countries of Australia, Austria, Belgium, Canada, China, Czech Republic, Denmark, Egypt, Finland, France, France, Germany, Greece, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lebanon, Monaco, Netherlands, Norway, Poland, Portugal, Russia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Thailand &amp; United Kingdom | Terminated | November 04, 2013 | 265 | 313 | Computer
Z-1105-2013 | 64737 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo WorkFlow SLR, Radiological Information System, Version VA31A SP1/SP2/SP3 |  |  | October 05, 2012 | 2012 | Siemens Medical Solutions USA, Inc | Under certain conditions, when using the Search functionality, the potential exists for the wrong patient's Interactive Documents to display when viewing a signed report in the single patient view. This can occur when the Interactive Document is viewed or auto displayed, dismissed by the user, and then a different patient and signed exam is selected in single patient view. The Notes Viewer will | Software design | Siemens Medical sent a Customer Safety Advisory Notice letter dated October 5, 2012 , to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.There will be a Hot Fix available.If you have any further questions please contact Siemens Customer Product Support at (1-888) 767-8326. |  | 37 | 37 | Worldwide Distribution - USA including the states of AL, CA, CO, DC, DE, FL, GA, HI, IL, IN, KS, KY, LA, MA, MD, MO, NC, NE, NH, NJ, NV, NY, OH, OR, PA, TN, TX, VA, WI, WV, and WY. and the countries of Canada and Australia. | Terminated | July 16, 2013 | 284 | 37 | Not_Computer
Z-1106-2013 | 64577 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Product Code 6801375 and VITROS 5,1 FS Chemistry System Refurbished, Product Code 6801890. |  |  | February 07, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho-Clinical Diagnostics, Inc. (OCD) determined that extension springs on some VITROSÂż Systems may beout of its manufacturing specifications. Extension springs are part of the slide alignment guide assembly. During an internal evaluation in its manufacturing facility, it was discovered that an extension spring had insufficient force. If the extension spring does not provide sufficient force, MicroSlides may not be properly positioned within theMicroSlide incubator which has the potential to cause imprecise results. Extension springs are just onecomponent in the slide alignment assembly. OCD's evaluation determined that the probability of an out ofspecification extension spring alone adversely affecting results is unlikely. | Nonconforming Material/Component | URGENT MEDICAL DEVICE CORRECTION letters were sent to the consignees on 2/07/13. On/about 2/07/13, OCD trained service personnel started to perform an assessment of all the VITROSÂż System(s) in End-User facilities worldwide to determine if the extension springs are out of manufacturing specifications. In addition, the slide alignment guides will also be inspected. If the extension springs are out of specification they will be replaced. Consignees will also be provided with a communication that explains the issue. |  | 2075 | VITROS 5,1 FS: USA: 670 units, Foreign: 1249 units; VITROS 5,1 FS Refurbished: USA: 52 units, Foreign: 2 units | Worldwide Distribution- USA (nationwide) including Puerto Rico and countries of: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Republic of Panama, Singapore, Spain, and Venezuela. | Terminated | April 20, 2017 | 1533 | 3811 | Not_Computer
Z-1109-2013 | 64755 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | IS2000 USER MANUAL for Intuitive Surgical da Vinci S Surgical System, Model IS2000. |  |  | March 15, 2013 | 2013 | Intuitive Surgical, Inc. | Revised instructions for users of the da Vinci S, Si and Si-e systems with additional information regarding sterile adapter engagement and instrument installation. | Nonconforming Material/Component | Intuitive Surgical sent an "Urgent Medical Device Correction" letter dated April 1, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 800-876-1310 for questions regarding this letter. |  | 1642 | 1642 customers | Worldwide Distribution.-USA (nationwide) and countries of: Austria, Belgium, Denmark, Finland, France, Germany, Ireland, Netherlands, Norway, South Korea, Sweden, Switzerland, and United Kingdom. | Terminated | September 04, 2014 | 538 | 1642 | Not_Computer
Z-1111-2013 | 64782 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Duraflo coated Femoral Cannula. |  |  | March 14, 2013 | 2013 | Edwards Lifesciences, LLC | Two unsealed pouches were found at an Edwards' distribution location in Japan. Incorrectly sealed pouches may result in a sterility barrier breach. | Packaging process control | Edwards Lifesciences sent an Urgent Field Safety Notice letter dated March 28, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use of any affected product remaining in inventory and to return any unused devices to Edwards. A completed Acknowledgement Form should be faxed to Edwards at 800-422-9329 within three days of receiving the notice. Questions regarding the recall should be directed to Edwards Customer Services at 800-424-3278, 6:00AM - 4:30PM PST. |  | 379 | 379 units | USA Nationwide Distribution including the states of: FL, MA, MI, NY, PA, TX. | Terminated | March 08, 2016 | 1090 | 379 | Not_Computer
Z-1115-2013 | 64810 | No | GFA | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Product is distributed by Zimmer (Warsaw, Indiana)/Synvasive Technology, Inc. (El Dorado Hills, CA) &amp; manufactured under Synvasive Technology, Inc. El Dorado Hills, CA |  |  | April 02, 2013 | 2013 | Synvasive Technology Inc | Various reciprocating and oscillating saw blades were incorrectly labeled. The labeling errors include: (1) The pouch says oscillating saw blade when it actually contains a reciprocating saw blade; (2) The box bar code date does not match the human readable expiration date; and (3) The label incorrectly states use with Zimmer/Synthes handset and not the Zimmer Universal Power System. | Packaging process control | Zimmer sent a Urgent Device Removal letter dated April 2, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers that received affected product from both Synvasive and from Zimmer Surgical will receive two customer letters, as each letter sets out specific item and lot numbers dependent upon distribution entity. Zimmer Distributors will also receive an e-mail notification. All notifications include a response form for both Synvasive and Zimmer distributors as well as customer consignees to account for product disposition (disposal in the field per facilities' sharps protocol).(1) Notifications will be mailed via Fed Ex to Risk Managers and Zimmer Distributors with inventory.(2) Notifications will be emailed to international geographies.(3) Notifications will be emailed all Zimmer Distributors.(4) A second Fed Ex mailing will be sent to domestic non-responding consignees.(5) Effectiveness phone calls will be made to any remaining domestic non-responding consignees. (6) If a customer reports they have further distributed the product, the secondary consignee identified will receive a notification from Zimmer.The scope of the recall is limited to certain confirmed lots which were shipped to the affected consignees..E-mails, certified mail receipts, and telephone records will be the firm's course of action in monitoring effectiveness checks. All returns will be stored in a separate area away from regular inventory. Final disposition of returned product is yet to be determined.The firm's recall strategy appears to be adequate.For further questions please call (916) 939-3913. |  | 3442 | 3442 units per the part numbers listed | Worldwide Distribution - USA (nationwide) and countries of: Austria, Australia, Belgium, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Netherlands, Norway, Panama, Spain, Sweden, Switzerland, Tunisia, United Kingdom and Yemen. | Terminated | July 15, 2014 | 469 | 3442 | Not_Computer
Z-1124-2013 | 64679 | No | NRA |  |  | Orthopedic | PMA | REF 32-422760 |  |  | March 20, 2013 | 2013 | Biomet U.K., Ltd. | Investigation of complaints found excessive use of the instrument could result in cracking on the weld between the head and the handle, which could lead to part of the head becoming detached from the handle. | Device Design | BIOMET sent an Urgent Medical Device Recall Notice on March 20, 2013, to all affected customers with directions to forward the notice if product was further distributed to Hospital Operating Room managers. The notice identified the recalled device, reason for the recall, and possible adverse events directing consignees to locate and discontinue use of the device. A FAX back Response form was to be returned to 574-372-1683 within 3 days of receipt prior to return of the product to Biomet, USA. Questionsregarding this recall should be addressed to 574 372-1570 M-F, 8a.m. -5 p.m. |  | 430 | 430 | Worldwide Distribution - USA (nationwide), Europe, Japan, Australia, Canada, Singapore, South Africa, and Costa Rica. | Terminated | May 04, 2015 | 775 | 430 | Not_Computer
Z-1116-2013 | 64841 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients. |  |  | March 21, 2013 | 2013 | Hospira Inc. | During infusions of 2.0 mL/hr the motor assembly may rotate backwards capturing additional medication resulting in over-infusion. | Nonconforming Material/Component | Hospira sent an Urgent Device Field Correction letter dated March 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to refrain from using the affected produc and contact Hospira at 1-800-441-4100 to arrange for return of their device for repair. Customers were directed to notify their customers if they further distributed the product and confirm they have done so by return of the reply form to Stericycle.For questions regarding this recall call 224-212-4892. |  | 38801 | 1) List Number 13000: 25,119 pumps; 2) List Number 13100: 185 pumps; 3) List Number 13150: 13,497 pumps | Worldwide Distribution - USA (nationwide) and Internationally to Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Open, Classified |  |  | 107955 | Not_Computer
Z-1112-2013 | 64653 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015 with software version 9.12 |  |  | March 06, 2013 | 2013 | Carefusion Corporation | The recall was initiated because Carefusion has identified potential risk associated with bolus programming with the Alaris PC unit (model 8015) software version 9.12. Automating infusion pump parameter input from either the electronic medical record (EMR) or the Alaris Auto-ID module for a continuous infusion with a bolus dose option may result in a loss of information from the Guardrails Data s | Software design | Carefusion sent an Urgent Medical Device Recall Notification letter dated March 6, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are required to confirm receipt of the notification by returning the Recall response card to Carefusion by postage-paid, self-addressed mail, fax, or email. The letter stated that a Carefusion representatives will contact customers by telephone to schedule field remediation. For questions call Carefusion Support Center at 888-562-6018, Customer Advocacy at 888-812-3266, and/or Technical Support at 888-812-3229. |  | 4090 | 4,090 total units | Worldwide Distribution USA Nationwide and the countries of Australia and Canada. | Terminated | May 08, 2014 | 428 | 4090 | Computer
Z-1118-2013 | 63945 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Clinical Chemistry Phenobarbital, list number 1E08 |  |  | November 16, 2012 | 2012 | Abbott Laboratories, Inc | Abbott Diagnostics is recalling Clinical Chemistry Phenobarbital reagent lots 62299UN12 and 85773UN12 due to imprecision issues which increase with the age of the reagent lot. The imprecision of the Phenobarbital assay is d 7% Total CV%; however, internal testing confirmed CVs up to approximately 12% for in-date reagent lots over five months old. Although lot 52803UN12 is now expired, it exhibit | Nonconforming Material/Component | Abbott Laboratories sent an Immediate Action Required Product Recall letter dated November 16, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use and destroy any remaining inventory of the affected product in accordance with their institution's policies and procedures, order replacement lot 27426UN12 which will expire 18 DEC 2012, and is the only lot currently available (Additional lots will be available in December and January), and complete and return the enclosed Customer Reply Form. If lots were forwarded to another laboratory, a copy of the letter should be forwarded to them. For questions regarding this recall call 1-877-4ABBOTT. For customers outside of the US they should contact their local area Customers Service. |  | 1289 | 1289 | Worldwide Distribution - USA (nationwide) and Internationally to Internationally distributed to Argentina, Australia, Austria, Bahamas, Brunei, Canada, Chile, Curacao, Denmark, France, Germany, Greece, Ireland, Italy, Lebanon, Mexico, Netherlands, Norway, Oman, Paraguay, Portugal, Puerto Rico, Russia, Saudi Arabia, Serbia, South Korea, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | June 18, 2013 | 214 | 1289 | Not_Computer
Z-1113-2013 | 64460 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Inform II Blood Glucose Monitoring System bar code scanner. |  |  | March 06, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche investigation into complaints found Accu-Chek Inform II and CoaguChek XS Pro System barcode reading meters had the potential of intermittent erroneous decoding of patient identification if barcodes being read fail to meet size and/or quality standards. | Device Design | Consignees of both devices were notified by Urgent Medical Device Correction letters sent via UPS Ground (receipt required) beginning 3/06/2013. The notice contained recommendation not to use EAN-13 barcode system and ensure barcodes used meet quality standards; use operator and patient lists (downloaded from a data manager) in order to check any decoded identification against the respective list. The notice directed forwarding of the notice if the product was further distributed. Customers were asked to complete the attached fax form. Contact the Customer Care Service Center, 24 hours a day, seven days a week at 1-800-440-3638 for Accu check questions and 1-800-428-4674 for CoaguChek questions about the information contained in this UMDC. |  | 4467 | Both devices 4467 | Nationwide Distribution. | Terminated |  |  | 4467 | Not_Computer
Z-1123-2013 | 64700 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVOTECH Endoscope Cleaner &amp; Reprocessor System, P/N 50004 |  |  | March 22, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) has determined that a limited number of EVOTECH ECRs manufactured between 2011 and 2012 did not receive sufficient High Potential (HIPOT) DC voltage testing. | Nonconforming Material/Component | An Urgent Medical Device Field Correction letter, dated 3/22/13, was sent to customers who purchased the EVOTECH Endoscope Cleaner and Reprocessor (ECR) HIPOT Testing. The letter informs the customers of the problem identified and the actions to be taken. Customers are instructed on what to do if electrical shock occurs. Customers are instructed that they will be contacted in the next 2 weeks on behalf of ASP to schedule HIPOT testing. Customers with questions are instructed to contact Stericycle at (888) 965-5816. Customers are instructed to report complaints or suspected problems with their EVOTECH ECR system by contacting ASP Professional Services directly at (888) 783-7723 and select option #2. |  | 273 | 273 units | Nationwide Distribution-USA including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | November 18, 2013 | 241 | 273 | Not_Computer
Z-1119-2013 | 64610 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Vitamin D Assay (100 tests); Reference Number: 10491994. |  |  | March 13, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | Negative bias in patient samples for Vitamin D | No Marketing Application | Siemens Healthcare issued Urgent Medical Device Recall was sent to all affected Siemens Healthcare Diagnostics customers in the United States on March 13th, 2013. This notice informs the customer of the recall and the actions that need to be taken. The Urgent Medical Device Recall instructs customers to discontinue the use of ADVIA Centaur VitD Assay Kit Lots (REF 10491994 and 10631021) and ADVIA Centaur Systems VitD Cal (REF 10493589 and 10630911) released prior to the recalibration of the assay and request replacement lots. This notice also directs customers to review the contents of the recall notice with their laboratory or medical director. A confirmation fax-back form is included to ensure that all customers in the United States have been notified and have complied with the mandatory action. |  | 3871 | 3871 | Nationwide Distribution. | Terminated | March 05, 2015 | 722 | 8489 | Not_Computer
Z-1098-2013 | 64728 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015 with software version 9.12 |  |  | March 06, 2013 | 2013 | Carefusion Corporation | The recall was initiated because Carefusion has received reports of a communication error on the Alaris PC unit model 8015 with software version 9.12 when the Alaris EtCO2 module or Alaris SpO2 module is attached. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated March 6, 2013 to all affected customers. The letter identified the affected product, problem and action to be taken. Customers are required to confirm receipt of the notification by returning the Recall Response card to Carefusion by postage-paid, self-addressed mail, fax, or email. Carefusion representatives will contact customers by telephone to schedule field remediation. For questions contact CareFusion Support Center at 888-562-6018, Customer Advocacy at 888-812-3266 and/or Technical Support at 888-812-3229. |  | 9241 | 9,241 units | Worldwide Distribution - USA (nationwide) and the countries of Australia and China. | Terminated | September 02, 2014 | 545 | 9241 | Computer
Z-1125-2013 | 64568 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Instrumentation Industries BE183-SU and BE 183 SU-12 Manual Jet Ventilators; BE 183-SUR and BE 183 SUR-12 Manual Jet Ventilators with Regulator and Gauge; BE 183-2 4 ft. Small Bore Tubing Assembly w/Inline Filter; and BE 183-R Regulator |  |  | September 07, 2012 | 2012 | Instrumentation Industries Inc | Firm initiated an update to Directions for Use necessitating replacement of previously distributed Directions for Use. | Error in labeling | Instrumentation Industries, Inc. initiated a recall of this product on September 7, 2012 by sending a letter via priority mail to all consignees. The letter identified the product, the problem, and the action to be taken by the customer. Consignees were instructed to replace the package insert of any affected product with the enclosed corrected package insert (now at Rev. H). The Directions for Use can alsbe downloaded from their website if they should need additional copies. If the product was further distributed consignees were asked to notify their consignees about this Field Correction.For questions regarding this recall call 412-854-1133. |  | 2998 | 2998 | Worldwide Distribution - USA including AK, AL, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV. Shipped internationally to Canada, Ireland, Hong Kong, the Netherlands, New Zealand, and Thailand. | Terminated | July 05, 2013 | 301 | 2998 | Not_Computer
Z-1126-2013 | 64712 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Accuray CyberKnife System Medical Physicist, Robotic Radiosurgery System, Medical charged-particle radiation therapy system, Model: Iris Variable Aperture Collimator P/Ns 028986 and 031777. The CyberKnife System is indicated for radiation treatment. |  |  | March 19, 2013 | 2013 | Accuray Incorporated | A defect on the master-side connector block of the pneumatic tool changer could introduce a tilt in the collimator housings, resulting in a potential shift to the central axis of the radiation beam. The issue could pose significant clinical impact to plans created using the smallest collimators. | Component design/selection | On 3/19/13, an "Urgent Device Correction Letter" was distributed to all affected users, with a description of the problem and user corrective action steps. US customers can contact 1-877- 668-8667 if questions. |  | 176 | 176 units | Worldwide distribution, including Nationwide in the US and the countries of Korea, Myanmar, Canada, Belgium, England, Czech Republic, Finland, France, Germany, Greece, India, Italy, Poland, Russia, Spain, Switzerland, Turkey, Columbia, Mexico, Japan and Saudi Arabia. | Terminated | October 23, 2013 | 218 | 176 | Not_Computer
Z-1138-2013 | 64486 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | OsteoSelect DBM Putty |  |  | February 21, 2013 | 2013 | Bacterin International, Inc. | The firm's retesting procedure was not validated. | Process control | Bacterin sent a Voluntary Product Recall Notification letter via e-mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are being asked to return any remaining product by obtaining a Return Authorization number from them. Bacterin is also asking customers to return the verification form mailed with the recall letter regardless of whether or not they had any product on hand. For questions contact Bacterin sales representative or Bacterin Customer Service. |  | 165 | 165 | USA Nationwide Distribution | Terminated | July 01, 2013 | 130 | 165 | Not_Computer
Z-1020-2013 | 64734 | No | EAP | Intraoral source x-ray system. | Dental | Radiology | 510(k) | DeGotzen XRay unit XGenus |  |  | September 26, 2012 | 2012 | DeGotzen | It was discovered at the WEAC Labs method sample did not meet requirements for labeling and certification | Radiation Control for Health and Safety Act | DeGotzen will add new labeling or change existing labeling to correct for the 5 mentioned deficiencies. This correction of the noncompliant products must be made at no cost to the customers.For further questions please call 800-289-6367 ext. 390. |  | 159 | 159 units | US (nationwide) | Open, Classified |  |  | 159 | Not_Computer
Z-1132-2013 | 64669 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | TRIGEN (TM) L-P SCREW, 4.5 MM X 37.5 MM, TI-6AL-4V, QTY: (1), STERILE R, REF 71645437, Smith &amp; Nephew, Inc. Memphis, TN 38116 USA |  |  | March 06, 2013 | 2013 | Smith &amp; Nephew Inc | One batch of 4.5 mm diameter Trigen Low-Profile Bone Screws were incorrectly anodized with a gold coating (indicating a 5 mm diameter) instead of a gray coating (indicating a 4.5 mm diameter). | Process control | Smith &amp; Nephew sent an Urgent Product Recall Notification letter dated March 6, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease use and distribution of the affected product and quarantine for return. Non-responding consignees were notified again on March 15, 2013, by email. Surgeons were to be notified by letter beginning on March 20, 2013.For questions regarding this recall call 978-749-1330. |  | 23 | 23 units | Nationwide Distribution including CA, FL, GA, MA, MI, NC, SC, OH, and MO. | Open, Classified |  |  | 23 | Not_Computer
Z-1136-2013 | 64395 | No | ODE | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EndoGastric Solutions EsophyX2 Plus Device with Serosa Fastener and Accessories Part number C02033-01. Label reads in part "EsophyX2 2-LINK Device Qty: (1) Fastener Delivery System ***" |  |  | February 04, 2013 | 2013 | Endogastric Solutions Inc | EsophyX2Plus device was not validated under actual or simulated use conditions prior to distributing the product to physicians. | Release of Material/Component prior to receiving test results | EndoGastric Solutions, Inc. sent an Urgent Medical Device Recall letter dated January 31, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to segregate and record lot numbers for any affected product. If no product remain in inventory customers were instructed to report such to customer service. Upon receipt of the product a credit will be issued for each device returned. For any questions customers were instructed to call 425-307-9226.For questions regarding this recall call 425-307-9200. |  | 190 | 190 | Nationwide Distribution including NC, VA, WI, TN, KY, IN, FL, MI, IA, IL, TX, LA, AR, HI, CA, NV, OR, and WA, | Terminated | July 02, 2013 | 148 | 190 | Not_Computer
Z-1128-2013 | 64698 | No | FRW | Patient scale. | General Hospital | General Hospital | 510(K) Exempt | SR Instruments FG3105 Patient Scales Used on various Joerns Healthcare Patient Lifters. Model# FG3105 Patient Scales. |  |  | December 13, 2012 | 2012 | Sr Instruments Inc | SR Instruments Inc. is implementing a field correction for the Joerns Healthcare Patient Lifts after they received two incident reports regarding two field failures of a tensile type load cell involving the FG3105 Patient Scale. The failures resulted in patient falls. | Use error | SR Instruments, Inc. sent a Notification letter dated December 13, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions call 716-693-5977. |  | 806 | 806 patient scales | USA Nationwide Distribution in the state of Texas. | Terminated | March 12, 2015 | 819 | 806 | Not_Computer
Z-1135-2013 | 64747 | Yes | NPT |  |  | Cardiovascular | PMA | Edwards Lifesciences Crimpers, Model 9100CR26. |  |  | March 26, 2013 | 2013 | Edwards Lifesciences, LLC | Holes in the sterile barrier pouch may be caused by handling of the device during manufacturing or once pouches have been removed from the unit box and stored on the shelf by customers. | Package design/selection | Edwards Lifesciences notified customers via letter and instructed them to maintain the device in its protective unit carton until ready for use and to inspect the pouch for holes or tears prior to use as instructed in the product's labeling. If any holes or tears are found, customers are instructed to contact Edwards Customer Service to arrange for return and replacement of damaged product. |  | 6117 | 6,117 units | Nationwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | September 16, 2013 | 174 | 11589 | Not_Computer
Z-1130-2013 | 64478 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | LifeScan brand OneTouchÂż Select Control Solution, Part Number/Model #: 02168902. |  |  | February 25, 2013 | 2013 | Lifescan Inc | When OneTouchÂż Ultra and OneTouchÂż Select test strips are tested with Ultra and Select control solution the meter may display a result outside the control solution test range printed on the strip vial label even though the system is working properly. | Nonconforming Material/Component | LifeScan sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 25, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers may contact OneTouch Healthcare Professional Line at 866-461-0089 for questions about this notice. |  | 132301 | 132,301 units total | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, 10, IL, IN, KS, KY, LA, MA, MD, ME, Ml, MN, MO, MS, MT, NC, NO, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SO, TN, TX, UT, VA, VI, VT, WA, WI,WV, and WY, and the countries of Russia and Ukraine. | Terminated | October 23, 2013 | 240 | 132301 | Not_Computer
Z-1133-2013 | 64748 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Muse v7 cardiology information system. Model number 2026443-001: MUSE 7.0 app cd, Model Number 2026443-015: MUSE 7.0.2 app cd, Model number 2026443-017: MUSE 7.1 app cd, Model number 2026443-029:MUSE 7.1.1 A app cd, and Model Number 2026443-031: MUSE 7.1.1 app cd. |  |  | January 04, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare conducted a recall on the Muse v7 because incorrect medications were displayed on the MUSE, which could result in incorrect treatment. | Software design | Consignee letters were not sent for this recall. GE Healthcare's recall strategy was to update all affected systems that were distributed. Updating of the devices was carried out from February 2012 to August 2012 by GE Field personnel. |  | 963 | 963 | Worldwide Distribution -- USA, including all states except ME, and the U.S. Virgin Islands; and, the countries of AUSTRALIA, AUSTRIA, BELGIUM, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, INDIA, IRELAND, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, MALAYSIA, NETHERLAND, NEW ZEALAND, NORWAY, PHILIPPINES,S QATAR, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | July 19, 2013 | 562 | 963 | Not_Computer
Z-1003-2013 | 64708 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Computed Tomography X-Ray System. |  |  | April 05, 2010 | 2010 | Philips Healthcare Inc. | Philips Healthcare discovered the customer was performing CTDI measurements for facility accreditation and received results that were 15 to 25% higher than Philips protocol parameters stated. | Radiation Control for Health and Safety Act | Philips Healthcare proposed corrective action plan (CAP) is comprised of the following elements:* Updating procedures in manufacturing to verify, either through software calibration or through gage-pin measurement, that A-plane collimation is appropriate to the design specification. It is our understanding that these procedures have been implemented.* For refurbished PET/CT and SPECT/CT systems distributed by Philips, themanufacturing process now requires that CTDI testing be conducted in addition to air testing and that each system meet updated CTDI tolerance limits to be published in the product's instructions for use.* Reassessing and correcting the correlation between air dose and CTDI specifications for all CT systems or sub-systems, resulting in updates of the CTDI specifications (tolerances) to be published in the instructions for use. It is our understanding that this reassessment and correction have been implemented, that the updated CTDI specifications will be made available as new CTDI tolerance tables with the customer information letter to be distributed immediately, and that these updated CTDI specifications will also be distributed as an addendum to the instructions for use manualsto the installed base during the next 12 months.* For all Brilliance CT, Gemini PET/CT, and Precedence SPECT/CT systems, in the installed base as well as in forward production, instructions for field service will be updated to provide clear directions on how to verify that the A-plane collimator has been appropriately adjusted to ensure expected product performance after its replacement and/or the replacement of an x-ray tube.This plan appears to adequately address the problem, and it is hereby approved. For further questions please call (978) 687-1501. |  | 6125 | 2089 US 4036 ROW | Worldwide Distribution including USA (nationwide) | Open, Classified |  |  | 6125 | Computer
Z-1139-2013 | 64671 | No | CKD | Alanine amino transferase (ALT/SGPT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alere Cholestech LDX ALT AST Test Cassette, Model #12-788. |  |  | March 06, 2013 | 2013 | Alere San Diego, Inc. | The recall was initiated because Alere San Diego has an update for the Alere Cholestech LDX ALT AST cassette, and it has the potential for humidity changes to impact results for certain analytes on the Alere Cholestech LDX System. | Other | Alere sent an "URGENT MEDICAL DEVICE RECALL UPDATE" letter dated March 6, 2013 to all affected customers. The letter informed the customers of the problem identified and the actions to be taken. Customers with any questions about the recall notification, are instructed to contact Alere San Diego 9975 Summers Ridge Rd, San Diego, CA 92121; (877) 308-8289; Fax: (858) 805-8457. Customers are instructed to complete the customer response form even if they do not have any involved product. |  | 38571 | 38,571 kits | Worldwide Distribution-USA (nationwide) and the countries of Canada, South Africa, Netherlands, Italy, Brazil, Singapore, Russia, Hong Kong, Argentina, South Korea, Philippines, El Salvador, Indonesia, and Trinidad. | Terminated | August 13, 2013 | 160 | 38571 | Not_Computer
Z-1137-2013 | 64900 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, AW Server. |  |  | June 28, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to the Results Viewer in the AW Server.After starting the Results Viewer with the selected series of a patient, the pixel and certain annotation data from a different patient is displayed in one of the viewports. This occurs on an occasional basis, when multiple users are simultaneously accessing the Results Viewe | Software design | GE Healthcare sent an "Important Product Information" letter dated June 27, 2012. The letter was addressed to Hospital Administrators , Radiology Managers and Risk Managers. The letter described the Subject, Instructions and Affected Product Details. GE Healthcare will correct all affected product by providing a software upgrade at no cost. For questions call: US 800-437-1171, Japan 0120-055-919, and/or local GE Healthcare Service Representative. |  | 429 | 429 | Worldwide Distribution - USA Nationwide including the states of: AK, AZ, CA, CO, CT, DE, DC, FL, GA, HI, IL, IN, IA, LA, MD, MI, ,N, ,S, MO, MT, NJ, NM, NY, NC, OH, OK, PA, SC, SD, TN,TX, UT, VA, WA, WI. and the countries of : AUSTRALIA, AUSTRIA, BELGIUM, BOSNIA, BRAZIL, CANADA, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, GREECE, GUADELOUPE, HONG KONG, INDIA, ISRAEL, ITALY, JAPAN, KOREA, LIBYA, MALAYSIA, NETHERLANDS, NEW ZEALAND, NORWAY, PANAMA, POLAND, QATAR, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TURKEY and UNITED ARAB EMIRATES. | Terminated | February 04, 2014 | 586 | 429 | Computer
Z-1127-2013 | 64646 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Oxford Partial Knee; Oxford Knee Resection Procedure 3 Pack CEMENTED |  |  | March 14, 2013 | 2013 | Synvasive Technology Inc | Biomet part # 506076, lot 928182 was received from Synvasive Technology containing the incorrect blade. | Packaging change control | An Urgent Medical Device Recall Notice letter dated March 21, 2013 was sent to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use; remove the products from inventory and complete and return the attached Response Form. For questions call 574-372-1570. |  | 100 | 100 kits | USA Nationwide Distributor including the state of IN | Terminated | January 13, 2014 | 305 | 100 | Not_Computer
Z-1149-2013 | 64641 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | MAQUET CARDIOHELP Support System: CARDIOHELP-I Software (up to version 3.3.0.0) |  |  | March 07, 2013 | 2013 | Maquet Cardiovascular, LLC | It has come to the attention of MAQUET that there have been episodes of brief unexpected shutdown, followed by an automatic device restart of the human machine interface (touch screen) of the CARDIOHELP device. The human machine interface (HMI) is the central display of the CARDIOHELP device, where measured values, settings and alarms are displayed and adjusted/confirmed, with the exception of bl | Software change control | MAQUET sent an Urgent Medical Device Field Correction letters dated March 7, 2013 via Federal Express to all affected customers. The letter identifies the affected product, problem and actions to be taken. Customers are asked to complete and submit the included Customer Fax-back Form. Maquet Service will exchange all affected products with the updated version. For questions MAQUET Customer Service at 888-627-8383 (press option 2, followed by option 2). |  | 154 | 154 units affected | Worldwide Distribution - USA Nationwide and the following countries: Australia, Austria, Belarus, Belgium, Canada, Columbia, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Iceland, Iran, Ireland, Italy, Jordan, Kazakhstan, Korea, Kuwait, Luxembourg, The Netherlands, Norway, Oman, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates and the United Kingdom. | Terminated | August 22, 2013 | 168 | 154 | Not_Computer
Z-1143-2013 | 64538 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | Elana Catheter Sizer 2.0*** |  |  | February 14, 2013 | 2013 | Elana, Inc. | Distribution of an unapproved device into interstate commerce. | No Marketing Application | Elana Inc. issued an Urgent: Medical Device Recall letter on February 14, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the Elana Catheter Sizer 2.0 as soon as possible. The Sizer 2.0 was intended as an optional accessory to the Elana Surgical Kit. The Elana Surgical Kit may be used properly without the Elana Catheter Sizer 2.0. Please acknowledge receipt of these instructions immediately by Email: Claartje@elana.com or Fax: 301-215-7038. The return shipping address for the Sizer 2.0 is Elana, Inc., 10480 Little Patuxent Parkway, Suite 400, Columbia, MD 21044 USA. Contact information for questions: Phone: 301-215-7010 or 443-538-4923 or +1 31 30 2537276 www.elana.com. |  | 3 | 3 devices | Distributed in the states of Pennsylvania, Texas, and New York. | Terminated | May 09, 2013 | 84 | 3 | Not_Computer
Z-1145-2013 | 64840 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | ACIST Medical Systems Inc. |  |  | March 15, 2013 | 2013 | Acist Medical Systems | ACIST Medical Systems is issuing a voluntary recall notice for three specific lots of BT2000 Automated Manifold Kits. For these specific lots, ACIST has received field reports related to ingress of air into the manifold body. This has been detected by the Air Column Detect Sensor halting the injection, and normally replacement of the BT2000 Automated Manifold Kit is necessary in order to proceed w | Nonconforming Material/Component | ACIST Medical Systems sent an Urgent Medical Device Recall letter dated March 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter advised customers that if they had inventory of any of the lots under the scope of the recall, to discontinue use or further distribution of the product. If you possess inventory containing any of the lot numbers listed above, please discontinue use or further distribution of these products. Please return this inventory to ACIST. Upon receipt of the returned BT2000 Automated Manifold Kits, ACIST will arrange for a shipment of replacement kits at no charge. To order additional or immediate replacement product, please contact BRACCO customer service at 1-877-BRACCO9 (1-877-272-2269), option 1.Please also complete and return the enclosed Medical Device Recall Form (see enclosure). Even if you do not possess any of the affected product in your inventory, the form needs to be completed and returned, indicating that you have no remaining units containing these lot numbers in your inventory. |  | 6640 | 6640 units | Nationwide Distribution including the states of AL, AK, AZ, AR, CA, CO, CT, DC, GA, HI, ID, IL, IN, IA, KS, KY, MD,MI, MN, MS, MO, NV, NH, NJ, NM, NY,NC, OR, OK, OH, PA, TN, TX, VA and WA. | Terminated | August 07, 2013 | 145 | 6640 | Not_Computer
Z-1151-2013 | 64596 | No | KXT | Digoxin test system. | Toxicology | Toxicology | 510(k) | Siemens Dimension Vista(R) System Vista Drug 4 Calibrator. The device is used as an in vitro diagnostic product. |  |  | February 21, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed complaints of low QC and patient recovery with the Dimension Vista(R) LOCI(R) Digoxin(DIGXN) Flex(R) reagent cartridge when calibrating DIGXN with either Drug 4 Cal (KC460) lot 2KD052 or Drug 4 Cal (KC460A) lot 2KD053. Internal testing has shown an average 15% low bias in the therapeutic range of 0.9 to 2.0 ng/mL. | Process control | The firm initiated their recall of this product by sending an Urgent Medical Device Recall letter dated February 2013 to all Dimension Vista Drug 4 Calibrator customers who received the affected lots. The letter notified them of the issue, the potential risk to health, and actions to be taken by customers. |  | 1638 | 1,638 | Distributed within the US to AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | January 07, 2014 | 320 | 1638 | Not_Computer
Z-1144-2013 | 64683 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Vitrea CT Multi-Chamber Cardiac Functional Analysis application on Vitrea version 5.2 (Vitrea Enterprise Suite 1.3 and Vitrea fX 3.1) and subsequent versions of VitreaWorkstation, Vitrea fX, VitreaAdvanced and Vitrea Enterprise Suite, through version 6.4. |  |  | March 15, 2013 | 2013 | Vital Images, Inc. | Vital Images has made a decision to perform a field correction to the Vitrea CT Multi-Chamber Cardiac Functional Analysis application. Vital Images has identified a defect in which numerical values may be incorrect in the Results Table under the Cardiac Analysis portion of the user interface. | Software design | Vital Images sent an Urgent Field Safety Notice dated March 15, 2013, to all affected consignees. The letter described the problem and the product affected by the recall. Letter provided a series of steps to follow which included to cease use of the device and to send the signed enclosed Effectiveness Check form to fieldnotices@vitalimages.com or fax to 1-952-540-3717. For questions customers were instructed to call Customer Support at 1-800-208-3005.For questions regarding this recall call 952-487-9500. |  | 102 | 102 | Worldwide Distribution - USA including AL, AR, CA, CO, IA, MI,MN, MO, NY, NC, OH, OR, PA, TX, and UT.Internationally to AUSTRIA, BELGIUM, CANADA, CHINA, COLOMBIA, GEORGIA, JAPAN, KOREA, MACEDONIA, NETHERLANDS, POLAND, PORTUGAL, RUSSIAN FEDERATION, SINGAPORE, SWEDEN, SWITZERLAND, TURKEY, UKRAINE, UNITED KINGDOM. | Terminated | February 27, 2015 | 714 | 102 | Not_Computer
Z-1131-2013 | 64848 | No | KST | Vacuum-assisted blood collection system. | Hematology | Hematology | 510(K) Exempt | Heritage Labs Path Study Extra Supplies Kit, Product # A 2376, kit contains 3 zippered plastic bags which hold support supplies for health care providers when treating patients. Bag #3 contains 5 different blood collection tubes (5 each) and 2 different gauge needles--all for drawing blood. One of the tubes is the recalled BD blue top citrate tube (5 tubes). These tubes were manufactured by BD, Franklin Lakes, NJ. |  |  | February 06, 2013 | 2013 | Heritage Labs International LLC | The product has a Lithium Heparin contaminant which could adversely affect coagulation results potentially causing erroneous lab results. | Material/Component Contamination | The recalling firm contacted 25 of their 26 customers by phone on February 5, 2013, to notify them of the recall and sent a follow up memo on February 6, 2013, During the call, the firm reminded their customers how many recalled kits they recieved and instructed them to go through the kits and pull the recalled lots. The recalled kits are to be returned to the recalling firm and the firm sent a follow-up memo on February 6, 2013, with the shipping label their customers to place on the returning kit shipping box. This memo was sent to the customers by overnight Fed-ex mail and the ship label for the returned product was an overnight UPS label. On February 7, 2013, the final customer was notified by phone and instructed to return the recalled kits.For further questions please call (913) 764-1045. |  | 25 | 25 kits | Nationwide Distribution including the states of AL, CA, GA, IL, IN, KY, MA, MD, NC, NY, OK and VA. | Terminated | November 25, 2013 | 292 | 25 | Not_Computer
Z-1147-2013 | 64770 | No | DXG | Single-function, preprogrammed diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | HCS3011 AV Loop Kit For use with Transonic CO status System Only, -HCS3021 AV Loop Kit For use with Transonic CO status System Only. |  |  | March 28, 2013 | 2013 | Transonic Systems Inc | Transonic Systems Inc. have received complaints of blood leakage involving the AV Loop Kit. | Component design/selection | Transonic Systems Inc. sent an Urgent: Medical Device Recall dated March 27, 2013 to all affected customers. The letter identified the recalled product, problem and actions to be taken. Customers were instructed to examine inventory, quarantine, discontinue use and return any recalled product via FedEx. Transonic Systems Inc. has offered to reimburse the firm for the return of these recalled products by check or credit memo for the returned goods and postage. Transonic Systems Inc. has requested that if a firm has further distributed any of the above lots, that they immediately contact their accounts, inform them of the recall, and have them return their outstanding recalled stocks to you. Transonic Systems Inc. have included copies of the product labeling and a Recall Return Response Form in their packet. Completed Recall Return Response Forms should be faxed to 607-257-5310. |  | 195 | 195 units total (140 units Domestic, 55 units Foreign) | Worldwide Distribution - USA Nationwide including the states of AZ, CA, FL, GA, NM, NY, TX, VA and WA and countries of Netherlands, Canada and Russia. | Terminated | March 05, 2014 | 342 | 195 | Not_Computer
Z-1141-2013 | 64723 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | HORIBA ABX PENTRA 400 I.S.E. Module, Model Numbers: P400ISE110EN02 and P400ISE110US02, All Serial Numbers, Software Versions 5.0.7 and Below. |  |  | March 20, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is informing all ABX PENTRA 400 Analyzer with I.S.E. Module customers that there are noLinearity Limit Flags for the I.S.E. assays currently programmed into the ABX PENTRA 400 Analyzer software. Allother assays run on this analyzer do have Linearity Limit Flags to alert the user to values that fall below or above thevalidated linearity range for each assay. | Software design | A recall letter, Field Safety Notice &amp; Acknowledgement mailed to the customers March 20, 2013. The letter included the product information, reason for recall, instructions on what to do with the recalled product, and contact information:"If you have any questions or concerns, please contact your local HORIBA Medicalrepresentative. We sincerely apologize for any inconvenience this may cause your laboratory. Please contact your Customer Service Representative at 1-888-903-5001 (option 3) for any additional questions related to this customerletter." |  | 905 | 905 units | Nationwide Distribution. | Terminated | September 26, 2014 | 555 | 905 | Computer
Z-1142-2013 | 64884 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2100IQ, Innova 3100IQ, Innova 4100IQ, |  |  | October 21, 2011 | 2011 | GE Healthcare, LLC | As part of GE Innova IQ table introduction, a label was designed to be affixed on the front Cover of theGantry Pivot section. This safety label is missing on systems manufactured between March 2010 and July 2011. In the situation a staff member (doctor, nurse, operator, anesthetist, etc&amp;) stands in the region between the C-arm gantry pivot and the head section of the table, there is a risk of co | Labeling design | A field action was initiated on October 21, 2011, and completed March 9, 2012, . The issue was corrected to all affected systems by applying the missing label onto the gantry. No recall letter was issued.For further questions please call (262) 513-4122. |  | 48 | 48 | Worldwide Distribution - Nationwide Distribution including the states of FL, IL, IN, LA, MI, MN, NJ, NM, NY, OH, PR, SC, TX and VA and the countries of Vietnam, Latvia, Italy, Israel, Indonesia, Hong Kong, United Kingdom, Finland, Spain, Czech Republic, Canada, Bahrain and Australia. | Terminated | February 04, 2014 | 837 | 48 | Not_Computer
Z-1140-2013 | 64499 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Duane Reade TRUEtrack monitor kit, UPC# 6 39194 01679 7 Part # A4006-81 |  |  | February 25, 2013 | 2013 | Nipro Diagnostics, Inc. | Meter kit outer boxes containing 10 ct. test strips state incorrect storage temperature range. | Environmental control | NIPRO Diagnostics sent a Voluntary Product Field Action Notice dated August 12, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check with all Duane Reade retail locations and Duane Reade Disgtribution Center to determine if they have any of the affected product ath their location. If product is located customers were instructed to call 1-800-803-6049 for a Return Authorization Number and instructions for returning the affected product.For questions regarding this recall call 954-677-9201. |  | 1627 | 1,627 Kits | Nationwide Distribution only in NY. | Terminated | July 08, 2013 | 133 | 1627 | Not_Computer
Z-1148-2013 | 64790 | No | KTN | Infectious mononucleosis immunological test system. | Immunology | Immunology | 510(k) | IMO-402 Mononucleosis Rapid Test Device (Whole Blood/Serum/Plasma). Catalog #IMO-402/94201 |  |  | February 26, 2013 | 2013 | Alere San Diego, Inc. | Alere San Diego is recalling the Clearview Mononucleosis Cassette because the positive control included in the rapid test may produce false negative results. | Material/Component Contamination | Alere sent an Urgent Medical Device Recall letter dated February 26, 2013, to all affected customers. The letter informs the customers of the problems identified and the actions to be taken. Customers with unused product remaining are instructed to immediately discard it. Investigations of customer complaints indicate that the positive control included with this lot may produce false negative results. Customers are instructed to return the attached Verification form. Customers that have consumed all product are instructed to indicate zero under "Quantity Discarded" and complete and fax back the form. Replacement product will be provided for any unused portion of the affected lot based on completion and return of the attached Verification Form. Customers with any questions are instructed to contact Alere Technical Services at (866) 216-0094 or by email at later.flow.support@alere.com.For questions regarding this recall call 858-805-2000, ext 3015. |  | 78 | 78 kits | Nationwide Distribution including TN, WA, UT, KS, KY, MN, GA, MA, PA, NY, VT, and the Dist. of Columbia | Terminated | August 29, 2013 | 184 | 78 | Not_Computer
Z-1150-2013 | 64781 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVOTECH Endoscope Cleaner and Reprocessor System, P/N 50004. |  |  | March 28, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the EVOTECH Endoscope Cleaner and Reprocessor (ECR) because they have received reports of smoke emanating from the system due to a circuit board malfunction that may occur in the Channel Control Manifold sub-assembly. | Under Investigation by firm | An Urgent Medical Device Field Correction letter dated 3/28/13 was sent to all customers who purchased the EVOTECH Endoscope Cleaner and Reprocessor (ECR). The letter informed the customers of the problems identified and the actions to be taken. Customers are informed that the failure can potentially cause the following outcomes for the cycles listed below:1. In the Full Reprocessing Cycle (Cleaning and HLD): there is a potential for residual rinse water and/or alcohol to remain in the endoscope channels after the cycle completes. 2. In the HLD-Only Cycle: there is potential for an adequate amount of high-level disinfectant solution to be sent through some endoscope channels.If customers observe smoke, then they are instructed to:a. Power OFF the device.b. Report your incident to ASP Professional Services at (888) 783-7723 and select option #2.c. Do not use the device until it has been serviced by ASP.Customers are instructed to return attached buyer return card acknowledging their receipt of this communication. Customers are instructed to contact Stericycle at (866) 761-9566 for any questions related to the field correction. |  | 761 | 761 units | Nationwide Distribution | Terminated | August 22, 2013 | 147 | 761 | Not_Computer
Z-1146-2013 | 64229 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Pressure Monitoring sets with VAMP Plus closed blood sampling system model # T001691A, PXVP0566, PXVP2260, PXVP0826, PXVP0827, PXVP2284AT3, PXVP23X3AT3, T398807D, T100605B, T001762A, T391T00A, VP2. |  |  | January 30, 2013 | 2013 | Edwards Lifesciences, LLC | The device is being recalled because the firm has received multiple complaints related to incorrect assembly that may result in dilution of blood samples. | Process change control | The firm initiated a recall in Europe via customer letters starting January 30, 2013 to notify the recall of model number T001691A lot number 59326333. The recall was expanded and a worldwide recall via customer letters was initiated on February 11, 2013. The letter identified the product, problem, and actions to be taken by the customers. |  | 30467 | 30,467 | Worldwide Distribution-USA (nationwide) including the states of AR, CA, FL, HI, ID, IN, KS, KY, LA, MA, MI, MN, MO, NC, ND, NM, NV, NY, OH, OK, PA, SD, TN, TX, VA, WA, and WI) and Europe. | Terminated | March 10, 2016 | 1135 | 30467 | Not_Computer
Z-1153-2013 | 64638 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova StatStrip Xpress Glu-Test Strips (Glucose test strips) |  |  | February 25, 2013 | 2013 | Nova Biomedical Corporation | Glucose test strips reports no result message when tested | Process design | The firm, Nova Biomedical, sent an "Urgent Field Safety Notice 01-13SS" dated March 18, 2013 and US customers notification by phone contact and to the International customers including a Faxback Form to be communicated to and sent to Nova Biomedical Subsidiary by two methods of delivery (Email and Phone Call). The customers were instructed to INVALIDATE the Nova StatStrip Glucose Test Strip lots indicated in the list on all StatStrip meters within your institute; REPLACE all Nova StatStrip Glucose Test Strips from the indicated lots with replacement test strips provided by Nova; REMOVE all unused boxes of the Nova StatStrip Glucose Test Strip lots indicated from all locations within your institute; RETURN all unused boxes of the Nova StatStrip Glucose Test Strips lots indicated to Nova Biomedical U.K. or your local Nova distributor following the country specific instructions provided by your authorized Nova representative, and complete and return the Fax-Back Form via email or FAX to your Regional Service Manager at 781-8914806.On April 08, 2013, the firm expanded the recall notification to include one additional test strip lot #0313021309 which was shipped under two catalog numbers(Catalog # 42214 and Catalog# 51493). US Notification by phone contact with key customer contact facilitated by Senior Manger of Technical Support Services. International - Field Safety Notice and Faxback Form is being communicated to and sent to Nova Biomedical Subsidiary by two methods of delivery (Email and Phone Call).Questions regarding this notice should be directed to the following Nova Biomedical EU Authorized Representative at: Nova Biomedical Innovation House, Aston Lane South Runcorn,WA7 3FY Telephone: + 44 1244 287087Email: Europe@novabiomedical.co.uk. or US at (781) 894-0800. |  | 7433 | 7433 cases | Worldwide distribution: USA (nationwide) including states of: CA, IL, KY,NY, WV and countries of: United Kingdom and Italy. | Terminated | September 26, 2014 | 578 | 9633 | Not_Computer
Z-1152-2013 | 64707 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products PHBR (Phenobarbital) Slides, REF 822 1384. |  |  | March 14, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) received customer complaints of intermittent imprecision, outliers, or shifts in accuracy and control fluid values outside of expected intervals when using VITROS Chemistry Products PHBR Slides. OCD's investigation confirmed that positively biased values for patient samples could be obtained that may not be detected by quality control fluids. | Unknown/Undetermined by firm | On 3/14/13, URGENT PRODUCT CORRECTION NOTIFICATION letters (dated 3/14/13), Questions &amp; Answers, and Response Forms were sent to consignees via FedEx courier or US Postal Service Priority Mail. Consignees were informed of the issue and advised to immediately discontinue using VITROS PHBR Slides. On 3/14/13, distributor letters were sent via US Postal Service Priority Mail to the distributors. They were advised of the issue and instructed to discontinue distributing and discard their remaining inventory of VITROS PHBR Slides. On 3/14/13, foreign affiliates were informed by e-mail of the issue and instructed to notify their consignees of the issue and actions. |  | 11758 | Domestic: 11,164 units, Foreign: 594 units | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Terminated | June 23, 2015 | 831 | 11758 | Not_Computer
Z-1155-2013 | 64809 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | RoboCouch Patient Positioning System |  |  | March 27, 2013 | 2013 | Accuray Incorporated | A5 gearboxes responsible for pitch (head up and head down) and roll (patient left and right) motions of the RoboCouch manipulator may be defective. | Nonconforming Material/Component | Accuray issued an Urgent Advisory Notification letter dated March 27, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that Field Service Engineers will be replacing the affected product. Questions and concerns are to be directed to Accuray Customer Support at 1-877-668-8667 |  | 16 | 16 devices are subject to correction. | Worldwide Distribution - USA Nationwide including the states of: OH, MO. OK, CA and VA and the countries of: Finland, France, Luxembourg, Poland, Russia, Turkey and Japan. | Terminated | August 20, 2013 | 146 | 16 | Not_Computer
Z-1156-2013 | 64584 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Orthopedic Alliance Spine System (alias "Blue &amp; Gold") |  |  | March 07, 2013 | 2013 | Orthopedic Alliance LLC | Orthopedic Alliance is recalling Blue &amp; Gold Implants and Instruments because the Blue &amp; Gold system is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product peformance failures that could cause patient harm due to implant breakage, loosening, or inadequate sterilization. | Device Design | A recall letter dated March 7, 2013 was sent to one customer who purchased the Blue &amp; Gold Implants and instruments. The letter informed the customer that the Blue &amp; Gold system is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product performance failures that could cause patient harm due to implant breakage, loosening, or inadequate sterilization. Customer is informed of the actions to be taken. Customer is instructed to complete the customer response form and return it by FAX to (858) 764-9739. |  | 3300 | 3,300 units | Distributed in Texas. | Terminated | March 26, 2014 | 384 | 3300 | Not_Computer
Z-1157-2013 | 64578 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | APLIF Implants and Instruments, |  |  | March 08, 2013 | 2013 | Spinal Solutions, LLC | Spinal Solutions is recalling the APLIF system because it is not supported by adequate testing and documentation to demonstrate that it meets performance or safety standards. These inadequacies might result in product performance failures that could cause patient harm due to implant breakage, movement, or inadequate sterilization. | No Marketing Application | Spinal Solutions LLC sent an Urgent Medical Device Recall letter dated April 12, 2013, to all affected customers. The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to discontinue use of the affected product and return to the firm. Customers were also instructed to complete the bottom of the form and to return it by fax to (858) 764-9739.For questions regarding this recall call 951-304-9001. |  | 220 | 220 units | Nationwide Distribution including NV, WI, MD,and CA. | Terminated | March 26, 2014 | 383 | 220 | Not_Computer
Z-1080-2013 | 64594 | No | GWQ | Electroencephalograph. | Neurology | Neurology | 510(k) | Harmonie software versions 5.1 uo to 6.2e : |  |  | April 06, 2009 | 2009 | Stellate Systems | The firm is initiating this recall because the implied or stated indications in their labeling associated with Harmonie software have not yet been cleared by the FDA for sale in the USA. | No Marketing Application | Stellate sent an "DEVICE CORRECTION" letter dated April 6, 2009, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Customers were instructed to check the product code referenced in the service manual to verify if your system is affected and check in "Help" &amp; "About Harmonie" in the menu to determine the software version on these systems. Stellate will shortly release updates to disable the noted features that were present in your particular version of the Harmonie software and revise labeling and Instructions for Use. All customers in possession of an affected system will be contacted by our Customer Support team to set up an appointment for installation of the updated version.If you have any questions or concerns regarding this notification, please do not hesitate to call our customer support line at 1-888-742-1306. |  | 2660 | 2660 Software licenses | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CT, DE, FL, HI, IA, IN, IL, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, ND, NJ, NY, OH, OR, PA, RI, SC, TX, VA, VT, WI, WV, and WY; and the countries of: Argentina, Australia, Bahrain, Brasil, Canada, China, Denmark, Finland, France, Germany, Greece, Haiti, Hong Kong, Hungary, Indonesia, Iran, Israel, Italy, Japan, Korea, KSA, Mexico, Netherlands, Norway, Pakistan, Portugal, Rep. DE Panama, Saudi Arabia, Spain, Switzerland, Taiwan ROC, Thailand, Turkey, and United Kingdom. | Terminated | April 22, 2013 | 1477 | 2660 | Computer
Z-1158-2013 | 64715 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Belly Bag Urine Collection Bag with Hip Belt. |  |  | March 28, 2013 | 2013 | Teleflex Medical | Sterile packaging may be compromised. | Nonconforming Material/Component | Teleflex Medical sent an Urgent Medical Device Recall Notification dated March 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed if you have any of the affected stock, immediately discontinue use and quarantine any products with the catalog numbers listed.If you and your customers have no affected stock please complete the enclosed Recall Acknowledgement Form and fax it to 1 (866) 804-9881 Attn: Customer Service. This will allow us to document your receipt of this letter.To return product complete the enclosed Recall Acknowledgement Form and fax it to 1 (866) 804-9881 Attn: Customer Service. This will allow us to document the amount of product you have on hand for return.If you have any other questions feel free to contact your local sales representative or Customer Service at 1 (866) 246-6990. |  | 190103 | 190,103 | Worldwide Distribution - USA and the countries of Singapore, Canada, New Zealand, Germany, Ireland, Japan and Australia. | Terminated | May 01, 2014 | 399 | 190103 | Not_Computer
Z-1164-2013 | 64804 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Xeleris 2 Processing &amp; Review Workstation. |  |  | October 06, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has become aware of a software issue in your Xeleris 2.1 workstation (that is used for processing Discovery 530/570 data), that may impact patient safety.On imported CT scans, Myocardial Perfusion Imaging (MPI) studies may create artifacts in attenuation corrected images. This may lead to a potential for misdiagnosis. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated October 5, 2011, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Consignees were instructed to contact their local GE Healthcare Service Representative if they have questions or concerns. They may also contact GE Healthcare Call Center , United States or Canada at 800-437-1171.Other countries 262-896-2890.For questions regarding this recall call 262-513-4122. |  | 11 | 11 | Worldwide Distribution - USA including GA, and NY; and countries of: Canada, Israel, Italy and Switzerland. | Terminated | February 04, 2014 | 852 | 11 | Computer
Z-1162-2013 | 64508 | No | NAW |  |  | Radiology | PMA | Sirtex SIR-yttrium 90 microspheres (Radionuclide); SIR-Y001. |  |  | February 27, 2013 | 2013 | Sirtex Medical Limited | Incorrect expiration date | Labeling False and Misleading | Sirtex notified medical end users on 2/27/13 via email correspondence with a copy of the recall letter attached. This was done to ensure advice was received at the same time that the goods were arriving. Hard copies of the recall letter, which were stamped "copy", were also sent to the individual facilities via UPS, processed for delivery Thursday, 28 February, 2013. Medical end users were advised in the recall letter to contact the US Customer Service staff at (888) 474-7839 x 702 to arrange for the pick-up if they decide to return the unused product. |  | 21 | 21 doses | Nationwide Distribution-USA (nationwide) including DC and the states of AZ, CA, CT, GA, ID, IL,KY, LA, MD, ME, MO, MT, NJ, NV, WA, and WV. | Terminated | March 30, 2015 | 761 | 21 | Not_Computer
Z-1163-2013 | 64828 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian brand RPM Respiratory Gating System, RPM Gating v1.7.5 with 3D Option, Model Number: HS2 |  |  | March 26, 2013 | 2013 | Varian Medical Systems, Inc. Oncology Systems | Varian has identified an anomaly with the RPM Respiratory Gating System version I.7.5 with 3D Option. While operating with the 6-dot marker block, this version of RPM Gating may produce an inverted respiratory waveform in amplitude and phase-gating sessions. | Device Design | Varian sent an Urgent Medical Device Correction - Urgent Field Safety Notice Letter dated March 26, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to retain a copy of this document along with their most current product labeling, advise the appropriate personnel of the content of this letter, and for future reference this document will be posted to the Varian Medical Systems customer support website: http://www.MyVarian.com. The letter stated that a Varian Medical Systems has developed a technical correction and a Customer Service representative will contact customers to schedule installation. For question contact your local Varian medical Systems Customer Support District or Regional Manager. |  | 2508 | 2508 units | Worldwide Distribution - USA (Nationwide) and countries of: Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Estonia, Finland, France, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macau, Malaysia, Mexico Morocco, Netherlands, New Zealand, Norway, Peru, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab E, and United Kingdom. | Terminated | July 09, 2014 | 470 | 2508 | Computer
Z-1166-2013 | 64783 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | NeuViz Dual series CT Scanner System, Part number (PN): 989605651321. This is a computed X-ray, Tomography system. |  |  | March 04, 2013 | 2013 | Philips And Neusoft Medical Systems Co., Ltd. | User error of the NeuViz 16 or NeuViz Dual may cause personnel injuries. There was a report where a finger was pinched when the table was moved backward while the tabletop was controlled by another person from the gantry left control panel. | Use error | The firm, Neusoft Medical Systems Co., Ltd., sent an "URGENT DEVICE CORRECTION" NeuViz Dual and NeuViz 16 letter, dated Mar. 4, 2013, to their US consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to review this information with all members of their staff and retain a copy of the letter with the equipment Instruction For Use.Customers with questions should contact the Service Support Department by email to nms-service@neusoft.com |  | 25 | 25 units | Nationwide distribution: USA including states of: California, Connecticut, North Carolina, Nebraska, Ohio, South Carolina, Texas, and Puerto Rico. | Terminated | November 22, 2013 | 263 | 48 | Not_Computer
Z-1161-2013 | 64739 | No | LOM |  |  | Microbiology | PMA | VITROSÂż Immunodiagnostic Products Anti-HBs (total antibody to hepatitis B surface antigen (anti-HBs)) Reagent Packs |  |  | November 01, 2012 | 2012 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling certain lots of the Vitros Anti-HBs Reagent packs an inability to obtain a valid calibration. | Under Investigation by firm | Ortho Clinical Diagnostics (OCD) sent an Urgent Product Correction Notification letter dated November 1, 2012, to all affected customers via USPS. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to discontinue and discard any remaining recalled lots of VITROS Anti-HBS Reagent Packs. Foreign affiliates were informed by email on November 1, 2012, of the issue and instructed to notify their consignees of the issue and actions . On March 22, 2013, OCD sent customer letters via FedEX overnight and/or via USPS to US consignees for two additional lots of product being recalled. The letter provided similar recall instructions as the initial notification. Ortho Clinical Diagnostics has requested that their clients discontinue and discard any remaining inventory of VITROS Anti-HBS Reagent Packs subject to this recall. OCD has also requested that recall product recipients also complete a Confirmation of Receipt and return that document to OCD, fax to: 1-888-557-3759 or 585-453-4110.We apologize for any inconvenience this may cause your laboratory. If you have technical questions please contact Customer Technical Services at 1-800-421-3311. If you have questions about your replacement order, please call our Customer Service Center at 1-800-828-6316. |  | 1464 | 1464 units | Worldwide Distribution- USA (nationwided) and countries of: Australia, Brazil, Canada, Chile, Chao Yang District, Beijing, Columbia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain and Venezuela. | Terminated | March 30, 2017 | 1610 | 1464 | Not_Computer
Z-1168-2013 | 64694 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ventri Discovery NM 530c, Ventri with Premium Table Discovery NM 530c, models H3000ZW &amp; H3000SA. The system is used to perform nuclear imaging procedures. |  |  | January 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of an incident which caused a patient injury during an unload following an exam on a Ventri scanner. During a patient unload following a scan, the patients fingers may be pinched between the moving cradle edge and the stationary base of the table while the table is moving outward automatically. | Device Design | Consignees were sent on 1/9/2013 a GE Healthcare "Urgent Medical Device Correction" letter dated January 9, 2013. The letter was addressed to Managers of Nuclear Medicine, Hospital Administrators / Risk Managers andManagers of Radiology / Head of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 556 | 556 | Worldwide Distribution - US, including the states of AL, AZ, AR, CA, CO, CT, DE, DC, FL, IL, IN, IA, KY, LA, ME, MD, MA, MI, ,N,MS, MO, NE, NV, NH, NJ, NM, NY,NC, OH, OK,OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, WI and the countries ofTAIWAN, TURKEY, SINGAPORE, SWD EN, PORTUGAL, POLAND, NEW ZEALAND, NORWAY, NETHERLANDS, MALAYSIA, MEXICO, SOUTH KOREA, JAPAN, ITALY, INDIA, ISRAEL GREAT BRITAIN, FRANCE, DENMARK, GERMANY, SWITZERLAND, CANADA, BRAZIL BELGIUM, AUSTRALIA, and ARGENTINA. | Terminated | August 06, 2013 | 209 | 556 | Not_Computer
Z-1167-2013 | 64731 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU5800 Clinical Chemistry Analyzer |  |  | March 20, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter, Inc. initiated this recall because there is a software error that has the potential to cause the reporting of erroneous patient results because of misidentification by the AU5800 system. | Software design | Beckman Coulter sent an Urgent Product Correction letter during the week of March 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were informed of the software limitation and provides instructions to be taken (temporary workaround) until the software is upgraded and released. Customer contact information provided: "Call Center Hotline at 800 854-3633. New customers will receive a copy of the customer notification at the time of installation by the Field Service Personnel until revised software that corrects this issue is released. |  | 490 | 490 | Worldwide Distribution - USA, including Canada and Puerto Rico and the states of NJ, TX, CA, UT, KS, MA, NY, MI, GA, VA, IN, WI, IL, OH, NV, NC and PA. and the countries of Australia, Belgium, Czech Republic, Japan, Germany, Brazil, China, Italy, Ireland, Israel, India, Hong Kong, Malaysia, Netherlands, Spain, Switzerland, France, Taiwan, Turkey, United Kingdom and Vietnam. | Terminated | September 16, 2014 | 545 | 490 | Computer
Z-1178-2013 | 64701 | No | DCK | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | The Alere Cholestech LDXÂż high sensitivity C-Reactive Protein (hsCRP) Test Cassette, Model Number 12-807. |  |  | March 06, 2013 | 2013 | Alere San Diego, Inc. | The Alere Cholestech LDXÂż hsCRPcassette (PN 12-807) may have increased imprecision relative to the performance data in the package insert,which may cause differences in quantitative results for hsCRP compared to a reference method orrepeat testing on the same patient sample. | Nonconforming Material/Component | Alere sent an Urgent Medical Device Recall letter dated March 6, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to block all stock of the affected lot in their warehouses; discontinue use and discard all stock on hand, complete the provided Customer Verification form to document disposal, contact their primary consignees and advise them of the required discontinuation and disposal of the affected lot. For question call 877-308-8289. |  | 2333 | 2333 kits | USA Nationwide Distribution | Terminated | March 12, 2014 | 371 | 2333 | Not_Computer
Z-1176-2013 | 64706 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ORTHOCORD One Violet and One Blue Braided Composite Suture with OS-6 Reverse Cutting Needles, Product Code: 223116. |  |  | March 22, 2013 | 2013 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | The IFU provided with the free strand ORTHOCORD sutures currently indicates that approximately 100% suture strength remains at six weeks after implantation. While this statement is true for Violet ORTHOCORD sutures, it is an inaccurate statement for Blue ORTHOCORD free strand sutures. Based on in-vivo strength retention study, a Blue ORTHOCORD sutures maintain approximately 80% of strength at six | Labeling Change Control | On March 22, 2013, firm notified consignees of recall via letter. Consignees were informed of the following: The IFU provided with the free strand ORTHOCORD sutures currently indicates that approximately 100% suture strength remains at six weeks after implantation. While this statement is true for Violet ORTHOCORD sutures, it is an inaccurate statement for Blue ORTHOCORD free strand sutures. Based on in-vivo strength retention study, a Blue ORTHOCORD sutures maintain approximately 80% of strength at six weeks after implantation. As Blue and Violet ORTHOCORD free strand sutures share the same IFU, IFU-108147 will be revised. Consignees were requested to confirm the receipt of the letter by March 29, 2013. |  | 46501 | 46501 | Nationwide Distribution. | Terminated | February 24, 2014 | 339 | 449009 | Not_Computer
Z-1180-2013 | 64972 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD CYCLESURE 24 Biological Indicator (BI) |  |  | April 22, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products is recalling certain lots of STERRAD CYCLESURE 24 Biological Indicator product because it may not have adequate data to support the entire duration of the labeled shelf-life. | Under Investigation by firm | Advanced Sterilization Products sent an Urgent Medical Device Recall letter dated April 22 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers that have installed and validated STERRAD system between Feb 2012 and March 2013, are instructed to contact Stericycle at (888) 943-4901 to determine if their STERRAD System is eligible for an optional revalidation. If customers would like to have their system validation repeated, ASP will revalidate their system without charge. Customers are instructed that the revalidation can occur as part of a routine Planned Maintenance (PM) within 12 months. Customers are also notified that the optional procedure will require that their Sterrad System will be unavailable for use for approximately 31 hours (depending on their system) while revalidation is performed. Customers are instructed to contact Stericycle at the number above with their STERRAD System serial number by May 31 2013. Customers are instructed to contact Stericycle at (888) 943-4898 or contact their ASP representative. Stericycle's hours of operations are Mon-Fri , 8am-5pm, ET. |  | 69994 | 69,994 cases | Worldwide Distribution - Nationwide Distribution including the states of AR, AL, AK, AZ, CA, CT, CO, DE, FL, GA, HI, IA, IL, IN, ID, KY, LA, MA, ME, MD, MI, MS, MN, MO, MT, NC, ND, NE, NJ, NY, NV, NH, NM, OH, OR, OK, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV., and the countries of AR, AU, BE, BR, CA, CL, CN, CO, DO, EC, GT, HK, ID, IN, JP, KR, MV, MX, PA, PE, PH, PR, RU, SG, TH, TW, UY, VE and VN. | Terminated | August 01, 2013 | 101 | 69994 | Not_Computer
Z-1159-2013 | 64839 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Pump - Battery |  |  | March 18, 2013 | 2013 | Hospira Inc. | Battery level lower then 2.4 volts results in corrupt history log and loss of data for preview therapy's parameters and causes an 11/004 error resulting in the pump not being able to be used. | Device Design | Hospira sent an Urgent Device Field Correction letter dated March 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify their consignees if they further distributed the product and confirm they have done so by return of the reply form to Stericycle. Lithium batteries that are older than three (3) years should be replaced. Customers should contact their Hospira Advanced Knowledge Center at 1-800-241-4002, option 4, to determine if their battery need to be replaced and if necessary, arrange for the return of their device to perform battery replacement. Customers with questions should call 1-877-907-7516. For questions regarding this recall call 224-212-4892. |  | 38801 | List Number 13000: 25,119 pumps; List Number 13100: 185 pumps; List Number 13150: 13,497 pumps | Worldwide Distribution - USA (nationwide) Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Terminated | March 07, 2017 | 1450 | 107955 | Not_Computer
Z-1177-2013 | 64645 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Maquet Getinge Group battery modules for use with Maquet SERVO-i ventilator systems (Part #6487180). |  |  | March 07, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | When using the SERVO-i on battery power, in a very small number of cases, some battery modules distributed after January 31, 2010 have a shorter battery run time than expected. | Device Design | The firm, Maquet Getinge Group" sent "Field Safety Notices"/"Field Safety Notice Confirmation Forms" dated March 7 &amp; 11 2013 via Federal Express and an updated Field Safety Notices dated April 15, 2013 to its customers. The Field Safety Notices identifies the problem, product, risk factors and actions to be taken by the customer. The customers were instructed to reconnect the ventilator to mains power immediately or add additional battery modules in case of 'low battery voltage' alarm; ensure that at least two fully charged batteries are installed at all times, and that sufficient resuscitation equipment is accessible to the ventilator; identify and immediately replace battery modules with date codes 1005-1243, that have generated either: a "low battery voltage" alarm or a "replace battery" message, prior to the expiry date; and complete and return the enclosed Field Safety Notice Confirmation Form by e-mailing a scanned copy to FieldActions@maquet.com or via fax at (973) 807-1832.If you have any questions or require replacement batteries, please contact your local MAQUET representative, or MAQUET Customer Service at 888-627-8383 (press option 2, followed by option 3) Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST. |  | 90000 | 90,000 units (batteries) | Worldwide distribution: USA (nationwide) and countries of: ALGERIA, ANDORRA, ANTILLES (Netherlands), ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBADJAN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOTSWANA, BRAZIL, BRUNEI, BULGARIA, CANADA, CHILE, CHINA,COLOMBIA, CONGO, COSTA RICA, CROATIA, CUBA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, YUGOSLAVIA (Former), FIJI, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUATEMALA, GUYANA, HONDURAS, HONG KONG,HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, KENYA, KUWAIT, KYRGYZSTAN, LATVIA, LIBYA, LITHUANIA, LUXEMBURG, MACAU, MACEDONIA, MALAYSIA, MALDIVES, MARTINIQUE (French), MEXICO, MOLDAVIA, MOROCCO, NEPAL, NETHERLANDS, NEW CALEDONIA (French), NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PAPUA NEW GUINEA, PARAGUAY, PERU, PHILIPPINES, POLAND, POLYNESIA (French), PORTUGAL, QATAR, REUNION (French), ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SEYCHELLES, SINGAPORE, SLOVAK REPUBLIC, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SUDAN, SURINAM, SWEDEN, SWITZERLAND, SYRIA, TADJIKISTAN, TAIWAN, TANZANIA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VATICAN city state, VENEZUELA, VIETNAM and YEMEN. | Terminated | August 11, 2016 | 1253 | 90000 | Not_Computer
Z-1182-2013 | 64944 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Software correction applies to customers who use the RT Therapist RTT4.2.108 and RT Therapist Connect RTT 4.2.108 |  |  | April 04, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Software update to fix multiple safety related issues. | Software design | Siemens sent updated instructions of TH004/13/S on April 4, 2013. The SW Update will be delivered by the Siemens Customer Service Group.For questions regarding this recall call 925-293-5442. |  | 245 | 245 active devices | Worldwide Distribution - USA (nationwide) and internationally to Poland, India, Canada, Germany, United Kingdom, New Zealand, Belgium, Malaysia, Republic Korea, Ireland, Australia, Norway, Japan, Turkey, Croatia, Saudi Arabia, Colombia, P.R. China, Brazil, Thailand, Romania, Lebanon, Mexico, Italy, Sweden, Czech Republic, Russian Federation, and South Africa. | Terminated | August 12, 2013 | 130 | 245 | Computer
Z-1181-2013 | 64833 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | VitreaÂż CT Brain Perfusion made available in various configurations under the names Vitrea, Vitrea fX, VitreaAdvanced and Vitrea Enterprise Suite. |  |  | March 15, 2013 | 2013 | Vital Images, Inc. | Vital Images, Inc. (Vital Images) received one report about potential user confusion when viewing a brain perfusion scan on Vitrea CT Brain Perfusion 2D when the scan was performed using a wide coverage shuttle acquisition technique. This may result in patient treatment delays, misdiagnosis or mistreatment. | Labeling design | Consignees were sent on 3/15/2013 a VITAL Images, Inc., "Urgent Field Safety Notice" dated March 15, 2013. The letter described the problem and the product involved in the recall. The letter listed the limitations of the product and the steps to take. Consignees were requested to return the Product Recall Effectiveness Check form. Questions should be directed to Vital Images Customer Support at 1-800-208-3005 or via e-mail, support@vitalimages.com. |  | 1978 | 1,978 | Worldwide Distribution -- USA, excluding RI and SD and including DC and PR, and the countries of ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BELARUS, BOLIVIA, BRAZIL, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, CYPRUS, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, CHANA, GREECE, GUAM, HONDURAS, HONG KONG, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOSCOW, MYANMAR, NAMIBIA, NETHERLANDS, NEW ZEALAND, NIGERIA, NORTHERN IRELAND, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SERBIA AND MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, VIETNAM, and YEMEN. | Terminated | April 08, 2015 | 754 | 1978 | Not_Computer
Z-1183-2013 | 64504 | Yes | FKQ | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Diacap(R) Ultra Dialysis Fluid Filter, Catalog Number 7107367. |  |  | March 01, 2013 | 2013 | B. Braun Medical, Inc. | There is the potential of the residual moisture of the Diacap Ultra membrane being reduced over time, which results in the reduction of the membrane permeability (ultrafiltration coefficient). | Process design | B. Braun Medical Inc. sent and "URGENT RECALL NOTICE" dated February 28, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Customer Support Department at (800) 848-2066 for questions regarding this issue.This recall was expanded on/about June 18, 2013 to include additional lot #s 305010113 and 305000113. Customers were notified via certified mail on/about June 14, 2013. |  | 1580 | 1580 | Distribution and manufacturing dates for the expansion lots was reported as 5/23-6/14/2013 and 1/22-24/2013 respectively. Nationwide Distribution including the states of AZ, CA, DE, FL, GA, HI, IA, ID, IL, KS, KY, LA, MN, MO, NJ, NM, NY, OH, OR, PA, SC, SD, TN, TX, UT, WA and WV. | Terminated | September 03, 2014 | 551 | 1580 | Not_Computer
Z-1171-2013 | 64851 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport Medical Instruments HT70 |  |  | April 05, 2013 | 2013 | Newport Medical Instruments Inc | Newport Medical Instruments is conducting a voluntary recall on certain Newport Medical Instruments HT70and HT70 Plus ventilator Power Pac batteries due to customer reports of Newport HT70 and HT70 Plus ventilators alarming and going to internal backup battery sooner than expected while the ventilator is being operated on Power Pac battery. | Component design/selection | Newport Medical Instruments sent an Urgent Medical Device Recall letter dated April 4, 2013, via FedEx to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers who sold the affected product to another party, or placed the affected product with a patient or healthcare provider were instructed to immediately inform the other party, patient, or healthcare provider of this action and provide them with a copy of the letter. If customers were using the affected product, they should ensure that an alternate source of power is available nearby. Customers were instructed to report any issues with Newport" HT70 and HT70 Plus ventilators or Power Pac batteries to the Technical Support group to ensure proper device reporting procedures are followed. Please call 800-451-3111, extension 500, for U.S. customers; and 714-427-5811, extension 500 for international customers and you will be given further instructions).For questions regarding this recall call 714-427-5811, ext 344.A press release was issued by the firm on 04/18/13. |  | 4392 | 2528 ventilators and 1864 power pac batteries | Worldwide Distribution - USA including US (AL, AR, AZ, CA, CO, CT, DE, FL, GA, lA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, NH, NJ, NV, NY, OH, OK, OR, PA, Puerto Rico, RI, SC, TN, TX, UT, VA, WA, and WI. Internationally to Argentina, Australia, Bangladesh, Belgium, BoliVia, Brazil,Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Ecuador, Egypt, France, Greece, Hong Kong, Hungary, India, Indonesia, Israel, Ireland, Italy, Japan, Kazakhstan, Kosova, Lebanon, Lithuania, Malaysia, Mexico, Nambia, Nepal, Norway, Oman, Paraguay, Peru, Philippines, Poland, Qatar, Saudi Arabia, Singapore, South Africa, South Korea, Sri Lanka, Switzerland,Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Viet Nam. | Terminated | March 10, 2015 | 704 | 4392 | Not_Computer
Z-1187-2013 | 64812 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Rad-8 Pulse Oximeters |  |  | March 13, 2013 | 2013 | Masimo Corporation | Masimo has identified a very small number (0.02%) of Rad-8 devices with 20-pin patient cable connector that, when moved, can power off without the operator pressing the power button. | Device Design | Masimo Corporation sent an Urgent Recall Notification letter dated March 8, 2013, to all affected customers The letter identified the product the problem and the action needed to be taken by the customer.1.Customers were instructed to please promptly remove all Rad-8 device(s) from use and your inventory and segregate these devices.2. Please call 1-800-326-4890 and select option 2 for Technical Services. Determine with Technical Services whether you will return the affected device(s) to Masimo for repair or if eligible, receive a repair kit. If you will return the device(s) for repair, obtain a Return Material Authorization Number (RMA number).3. Complete the Tracking/Verification Form in Attachment 2 and fax it to Masimo at 1-949-297-7700, even if you do not have or intend to use the affected device(s).4. If you are eligible and intend to repair units onsite fax the completed original form in Attachment 2 to Masimo at 1-949-297-7700.5. If you have elected to return the affected device(s), and have an RMA number include the completed original Form in Attachment 2 with the returned device(s) and indicate the RMA number on the outside of the return-shipping box.We apologize for the inconvenience. Please be assured that Masimo is committed to consistently providing high quality products and services to you, our customers. We thank you for your patience and cooperation while we actively work to resolve this issue... |  | 42908 | 42,908 units | Worldwide Distribution - USA (nationwide) including the countries of Africa, Asia, Australia/New Zealand, Europe, Canada, Latin America, and Middle East / SW Asia. | Terminated | June 10, 2015 | 819 | 42908 | Not_Computer
Z-1189-2013 | 64860 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defribillator; Product Codes: M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6. |  |  | April 08, 2013 | 2013 | Philips Healthcare Inc. | Device may fail to deliver defibrillation therapy in either Manual Defib or AED mode. | Software design | The firm, Philips Healthcare, sent an "Field Safety Notice-URGENT - MEDICAL DEVICE RECALL" letter dated April 2013 to its customers. The letter describes the product, problem and actions to be taken. In addition, letter provides instructions for using the MRx prior to receiving the software upgrade and informs the customer that Philips is initiating a correction consisting of a software upgrade for all units affected. The customers were instructed to follow the "ACTION TO BE TAKEN BY CUSTOMER/USER" section of the Field Safety Notice; send a copy of the attached package including customer letter and Field Safety Notice to any customer to whom you have distributed one of the affected device; and provide your local Philips organization with the names and addresses of customers to whom you have sold affected devices, so that arrangements can be made to perform the software updates.Should you have any questions or concerns about this device correction, please contact your local Philips representative at 1-800-722-9377.Philips updated the Field Safety Notice dated May 2013 to users who have not had the software upgrade performed. It provided additional information in the "Problem Description" and "Action to Be Taken By Customer/User". |  | 64972 | 64972 | Worldwide distribution: USA (nationwide) and countries of: Afghanistan, Albania, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Bolivia, Bosnia and Herzegowina, Brazil, Brunei Darussalam, Bulgaria, Canada, Cambodia, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Finland, France, Germany, Greece, Guadeloupe, Hong Kong, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Korea, Republic Of, Kuwait, Laos, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Macau, Macedonia, Malaysia, Malta, Mexico, Mongolia, Morocco, Namibia, Nepal, Netherlands (Antilles), New Caledonia, New Zealand, Nicaragua, Nigera, Norway, Oman, Pakistan, Palestine, Panama, Papua New Guinea, Paraguay, Peru, Philippines, Poland, Portugal, Puerto, Rico, Qatar, Reunion, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, SRI Lanka, Sweden, Switzerland, Taiwan (Province of China), Tanzania, Thailand, Trinidad and Tobago, Tunisia, Turkey, Turkmenistan, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Veit Nam, and Yemen. | Terminated | January 10, 2017 | 1373 | 64972 | Computer
Z-1184-2013 | 64868 | Yes | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | Medi-Trace Cadence Adult Radiotransparent Defibrillation Electrode |  |  | April 08, 2013 | 2013 | Covidien LLC | Defibrillation Electrodes may arc/spark result in thermal damage to the leadwire, which could render the device incapable of delivering the appropriate energy or shock to the patient | Process control | Covidien sent an "URGENT MEDICAL DEVICE RECALL" letter dated April 8, 2013 to all affected customers. The letter identifies the product, problem and requests that customers immediately stop using and segregate any affected product from inventory. Questions:Call 1-800-962-9888 option 8, then extension 2500. |  | 133470 | 133470 (sets of 2) | Worldwide Distribution-USA (nationwide) and the countries of Korea, Japan, Canada, Australia, Brazil, Belgium, Singapore, and South Africa. | Terminated | November 19, 2015 | 955 | 140955 | Not_Computer
Z-1186-2013 | 64621 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SL3 SOFT-TISSUE Laser (Model LR2002); Manuf: Discus Dental, LLC, 8550 Higuera Street, Culver City CA 90232. |  |  | July 21, 2011 | 2011 | Discus Dental LLC | The firm is initiating this recall of the SL3 Soft Tissue Laser equipped with a fiber optic Unifiber system which contains a fragile glass filament because the fiber broke near the hand piece. | Device Design | DISCUS Dental, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated July 21, 2011. The letter described the product, problem and action to be taken. Customers were instructed tocontact Discus Dental at 1-888-427-9279 for a no-cost replacement.For questions regarding this recall call 310-845-8200. |  | 150 | 150 units | Nationwide Distribution including AL, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, MA, MD, MI, MO, MT NH, NJ, NM, NV, NY, OH, PA, SD, TX, UT, VA, and WA. | Terminated | April 29, 2013 | 648 | 150 | Not_Computer
Z-1188-2013 | 64889 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE, VCT for Discovery VCT, Discovery RX VCT, and Discovery PET/CT 690, Lightspeed VCT, Discovery CT750HD. |  |  | July 25, 2011 | 2011 | GE Healthcare, LLC | Fastener bolts on the heat exchanger were not properly torqued and part of the heat exchanger can detach from the gantry and hit the inside of the covers and other components of the gantry possibly resulting in the oil conduit breaking, spilling oil within the gantry and outside the system. This issue affects GE Healthcare VCT for Discovery VCT, Discovery RX VCT, and Discovery PET/CT 690, Lightspe | Process control | No consignee letter was sent out. At customer site Field Engineer replaced heat exchanger, the tube and the pump. Field Engineer also verified torque on all bolts. Field action was initiated on 7/25/2011 and has been completed as of 3/29/2013. |  | 166 | 166 | Worldwide Distribution-USA (nationwide) including DC and the states of AZ, CA, CO, CT, FL,GA, IL, IN, KY, MA, MI, MS,MO, NH, NJ, NM, NY, NC, OK, OR, PA, SC, SD, TX, UT, VA, WA, and WI and the countries of ALBANIA ARGENTINA AUSTRALIA BELGIUM BRAZIL CANADA CHINA COLOMBIA COSTA RICA CZECH REPUBLIC DENMARK ECUADOR EGYPT FINLAND FRANCE GREECE HONG KONG HUNGARY INDIA IRELAND ITALY JAPAN KOREA, REPUBLIC OF MEXICO MOROCCO NIGERIA NORWAY POLAND RUSSIAN FEDERATION SINGAPORE SPAIN SWEDEN SWITZERLAND TURKEY UKRAINE UNITED KINGDOM. | Terminated | February 04, 2014 | 925 | 166 | Not_Computer
Z-1170-2013 | 64837 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients. |  |  | March 18, 2013 | 2013 | Hospira Inc. | Pump shutting off during use without warning. | Device Design | Hospira sent an Important Device Information letter dated March 18, 2013, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect the affected product for signs of leakage, corrosion or other damage prior to each use. In addition, each time the batteries are replaced the battery compartment should be inspected for damage. If a device exhibits damage caused by leakage, it should immediately be removed from clinical service and Hospira should be contacted at 1-800-441-4100 to arrange for return of the device for repair. If the product has been further distributed they should notify their consignees and confirm they have done so by returning the reply form to Stericycle.For questions regarding this recall call 224-212-4892. |  | 69604 | List Number 13086: 286 pumps; List Number 13087: 45,376 pumps; List Number 13088: 23,942 pumps | Worldwide Distribution - USA (nationwide) and internationally to Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Terminated | March 10, 2017 | 1453 | 108405 | Not_Computer
Z-1208-2013 | 64918 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, GE 1.5T SIGNA HDe MR SYSTEM. |  |  | June 08, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a Reformat Image Flip Issue associated with GE MR Products.The axial and coronal reformatted images from a sagittal 3D series (such as with applications of FSPGR-3D T1, BRAVO T1, or CUBE T1) of over 126 slices may display a left-to-right image flip .In addition, the acquired sagittal images may be annotated with left-right location annotation revers | Software design | GE Healthcare sent an "Important Product Information" letter dated June 8, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers should contact their local service representative for questions regarding this notice. |  | 822 | 822 (154 US; 668 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, excluding the state of RI and the countries of Yugoslavia YEMEN Viet Nam Venezuela Uzbekistan Uruguay United Arab Emirates Tunisia Thailand Tanzania Taiwan Turkey Syria Switzerland Sweden Spain South Korea South Africa Slovenia Slovakia Singapore Saudi Arabia Russia Romania Reunion Qatar Portugal Poland Philippines Peru Pakistan Paraguay Philippines Oman Norway Nigeria New Zealand Netherlands Morocco Mexico Martinique Malaysia Macedonia Mauritania Luxembourg Lithuania Libya Lebanon Latvia Kuwait Kenya Korea Korea (Republic Of [South] Korea) KUWAIT Kazakhstan Jordan Japan Jamaica Italy Israel Ireland Iran Iraq Indonesia India Hungary Hong Kong Honduras Guadeloupe Greece Germany Georgia, Republic of French Polynesia France Finland Estonia Ecuador Egypt Denmark Dominican Republic Czech Republic Croatia Colombia China Chile Canada Bulgaria Brazil Bosnia And Herzegovina Belgium Belarus Barbados Bahrain Austria Australia Argentina Arab Emirates Algeria and United Kingdom. | Terminated | November 03, 2014 | 878 | 12048 | Not_Computer
Z-1192-2013 | 64919 | applicator, absorbent tipped, non-sterile | KXF | Absorbent tipped applicator. | General Hospital | General Hospital | 510(K) Exempt | MediChoiceÂż Rayon Tipped OB/GYN Applicator, 8 inch, Plastic Shaft, Non-Sterile, Item Numbers WOD5001 and WOD5002. |  |  | March 26, 2013 | 2013 | Owens &amp; Minor Distribution, Inc. | Owens &amp; Minor Distribution, Inc. is conducting a market withdrawal of MediChoice Applicator OB-GYN 8 inch product that does not meet our quality standards because the rayon may loosen or unravel from the applicator during use. | Process change control | Owens&Minor; sent an Urgent: Medical Device Recall letter to the first set of customers on March 26, 2013 and to their second set of customers on April 9, 2013. The letter identified the affect product, problem and actions to be taken. Customers were instructed to examine their inventory and quarantine all affected products, contact service representative to discuss whether affected products should be returned or destroyed, and complete and return the Recall Reply Form. For questions call 804-723-7417. |  | 4298.5 | 4298.5 cases | US Nationwide Distribution. | Terminated | October 31, 2014 | 584 | 4298.5 | Not_Computer
Z-1191-2013 | 64439 | No | MRU |  |  | Neurology | HDE - Humanitarian Device Exemption | Medtronic DBS Therapy for Dystonia Kit, models 3317, 3319, 3337 and 3339. Sterile and Non-Pyrogenic. |  |  | February 08, 2013 | 2013 | Medtronic Neuromodulation | There is a potential for lead damage due to the use of the lead cap provided in DBS and Dystonia kits. Medtronic has received reports of DBS leads being damaged at the connector end of the lead when the lead cap is used. The connector end of the lead is the end of the lead connected to the lead extension. Tightening or loosening of the setscrew may twist the setscrew connector block and may damag | Process control | Medtronic sent an Urgent Medical Device Correction letter dated February 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to follow the modified instructions if the lead cap is used during the implant procedure. For questions call Medtronic Neuromodulation Technical Services 1-800-707-0933. UPDATE: Medtronic mailed an Urgent: Medical Device Removal notification to customers on June 07, 2013, who may have unused DBS lead kits that were manufactured prior to the process change to correct the issue. On June 10, 2013 Medtronic representatives began visiting the identified accounts to facilitate removal of any unused product. The Medtronic Rep documented the hospital visits on an Account Specific Customer Confirmation Form and left a copy with the Hospital. |  | 30000 | 30,000 | Worldwide Distribution - USA Nationwide and the countries of: Algeria, Australia, Austria, Belarus, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Latvia, Lebanon, Lithuania, Luxembourg, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Norway, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam. | Terminated | May 22, 2014 | 468 | 30000 | Not_Computer
Z-1210-2013 | 64942 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Howmedica Osteonics Corp. AxSOS System Instrument. |  |  | March 28, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics has received a report indicating that the scale of the AxSOS Calibrated Drill Bit 4.3,,X263mm, flat , manufactured by Stryker Osteosynthesis, is incorrect. The incorrect scale could result in the surgeon selecting a screw that is 10mm longer than intended. | Device Design | Stryker Orthopaedics sent a Urgent Medical Device Removal letter dated March 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Please inform users of this Medical Device Removal and pass thisnotice to all those individuals who need to be aware within yourorganization.2. Return all affected products available at your location toStryker Osteosynthesis or Contact Stryker Customer Service.3. Complete and sign the enclosed Business Reply Form and fax a copyto: 1-865-252-3635 .4. Keep a copy of the completed and executed Business Reply Form foryour records.Report any adverse events or product quality problems to StrykerOrthopaedics: 1-866-0R-ASSIST. (1-866-672-7747).As we strive for products that meet your expectations for quality and reliability, please do not hesitate to contact us, in case you have any further questions at (201) 831-5970.. |  | 107 | 107 units | Nationwide Distribution including the states of AL, AZ, IL, KY, LA, NJ, NC, FL, CA, NM, MO, TN, VA, CO, MA, TX, IN and MT. | Terminated | November 17, 2015 | 964 | 107 | Not_Computer
Z-1212-2013 | 64885 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Optima MR430s MRI Scanner; 5000-0002 |  |  | March 28, 2013 | 2013 | GE Healthcare, LLC | Coolant supply lines may leak liquid onto the floor resulting in a slippery surface and fall hazard. | Process control | GE Healthcare sent an Urgent Medical Device Correction letter dated April 23, 2013, to all affected customers. The letter informed customers of a potential safety issue due to leakage from gradient coolant lines. Consignees were instructed to monitor the floor for leakage and continue to use if leakage is not observed. If leakage is observed consignees are instructed to stop using the system and notify their local service representative. GE Healthcare will correct all affected scanners by providing an update at no cost.Customers with questions were instructed to contact their local service representataive.For questions regarding this recall call 262-513-4122. |  | 81 | 81 | Worldwide Distribution - USA including AL, CA, CO, FL, ID, IL, MA, NH, NJ, NY, OH, PA, SC, SD, TN, TX, UT, VA, WA, WI; Worldwide: Argentina, Australia, Austria, Brazil, Canada, Chile, Denmark, Egypt, Finland, Germany, India, Italy, Japan, Korea, Malaysia, Netherlands, Norway, Peru, Poland, Russia, Singapore, South Africa, Spain, Sweeden, Switzerland, UK, Ukraine, and UAE. | Terminated | October 25, 2013 | 211 | 302 | Not_Computer
Z-1214-2013 | 64820 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage PureFlow-B Solution, Model/Catalog Number: RFP-401, packaged in a case |  |  | March 18, 2013 | 2013 | NxStage Medical, Inc. | Product may be mislabeled. | Labeling mix-ups | NxStage sent a letter dated March 18, 2013, via email to all consignees instructing them to check their entire inventory and to segregate and quarantine lot F023444 pending further instruction from NxStage.On 4/2/13 NxStage sent a letter via email to all consignees requesting a return of all RFP-400 and RFP-401 labeled with lot F023444.. Consignees were instructed to complete the recall reply form and provide a list of customers (name address, phone number) and fax them to NXStage Medical, Inc.Customers with questions were instructed to call 1-866-697-8243.For questions regarding this recall call 1-866-697-8243. |  | 1480 | 1480 cases (2 bags per case) | Nationwide Distribution including IL, SC, TX, VA, FL, AR, and NY. | Terminated | September 17, 2014 | 548 | 1480 | Not_Computer
Z-1213-2013 | 64859 | No | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 3DKnee e+ Tibial Insert. The insert is used along with the other components that comprise the 3DKnee System. |  |  | April 01, 2013 | 2013 | Encore Medical, Lp | U.S. Agent contacted Director Commercial Logistics, notifying him of incorrect color coded labels on two packages of e+tibial inserts. There were 8 e+ tibial inserts of the same lot number in Finished Goods as the two from the agency. These devices were inspected and all were incorrectly labeled. | Labeling mix-ups | The firm sent their US customer a "Urgent Field Safety Notice" letter on April 1, 2013. The letter explained the product problem and provided recommended actions as well as customer contact numbers. |  | 5 | 5 | US Distribution to the state of California and International Distribution to the country of Venezuela. | Terminated | July 01, 2013 | 91 | 5 | Not_Computer
Z-1000-2013 | 64589 | No | EAP | Intraoral source x-ray system. | Dental | Radiology | 510(k) | CareStream Dental CS 2200 X-ray system |  |  | September 18, 2012 | 2012 | Trophy Sas | There have been approximately 2% of the x-ray generators manufactured between February 2012 and November 2012 returned due to errors. These generators showed the type of defect that indicates poor contact of HT transformer power pin. The defect is a bent connector associated with the generator. | Radiation Control for Health and Safety Act | Trophy, A Carestream Dental Company, - A Subsidiary of Carestream Health Inc. The Center for Devices (CDRH) understands that the corrective action plan (CAP) involving 167 units will consist of implementing rework procedure TL008, which describes the procedure for a bending the metallic strip the female connector to insure strong connection with the male counterpart.The CAP appears to adequately address the problem and is hereby approvedFor further questions please call (585) 781-1997. |  | 167 | 167 installed in the US | USA Nationwide Distribution | Open, Classified |  |  | 167 | Not_Computer
Z-1215-2013 | 64444 | No | JTY | Culture medium for pathogenic Neisseria spp. | Microbiology | Microbiology | 510(k) | BD ProbeTec Neisseria gonorrhoeae (GC) Qx Amplified DNA Assay Reagent Pack, REF 441124, contains 12 each Microwells labeled in part ***BD ProbeTec Neisseria gonorrhoeae (GC) Qx Amplified DNA Assay Reagent Pack REF 441124 ***For use with the BD Viper System *** Becton, Dickinson and Company 7 Loveton Circle, Sparks, MD 21152 USA www.bd.com/ds EC REP Benex Limited, Pottery Road, Dun Laoghaire, Co. Dublin, Ireland. |  |  | February 08, 2013 | 2013 | Becton Dickinson &amp; Co. | Increased levels of false positives in BD Probe Tech Neisseria gonorrhoeae Q Amplified DNA assay kits. | Environmental control | BD Diagnostic Systems sent an "URGENT PRODUCT RECALL" letter via UPS overnight delivery on February 8, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to verify that these kits have been consumed, returned or destroyed by completing the enclosed form. The firm will issue a no charge replacement for customers' product. Any product remaining in customers' inventory should be discarded or destroyed per their facility procedures. Complete the attached form regardless if any inventory remains, so that the firm may acknowledge customers' receipt of this notification. It is not necessary to contact BD by telephone to receive replacements. Complete and fax the form to: Attention: Regulatory Compliance, Fax: 410-316-4258 OR E-mail to RegulatoryComplianceFax@BD.com. If further assistance is needed regarding replacements, please contact BD Customer Service Department at 1-800-675-0908. |  | 77 | 77 kits | Distributed in the states of CA, FL, GA, NJ, and TX. | Terminated | May 13, 2013 | 94 | 77 | Not_Computer
Z-1218-2013 | 64682 | No | GCX | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Stryker Irrigator Reusable Tips: |  |  | March 27, 2013 | 2013 | Stryker Endoscopy | The pre-vacuum steam sterilization (wrapped method) temperature in the Instructions for Use incorrectly listed 121-123ÂżC , instead of the correct 132-133ÂżC. | Error in labeling | Stryker sent "Urgent Device Correction" letters dated March 27, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 147921 | 147,921 - all sizes | Worldwide Distribution: USA (nationwide) and countries of: Mexico, Portugal, Latin America, India, Greece, Taiwan, Singapore, New Zealand, Phillipines, Netherlands, France, Australia, Canada, Japan, Scandinavia, Iberia, Brasil, Romania, South Africa, Chile, Korea, China. | Terminated | September 24, 2013 | 181 | 147921 | Not_Computer
Z-1238-2013 | 64771 | No | KWI | Elbow joint radial (hemi-elbow) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Synthes Radial Head Prosthesis: Trial Radial Head |  |  | February 22, 2013 | 2013 | Synthes USA HQ, Inc. | The Trial Head may come loose from the implant stem during manipulation of the arm during surgery. To date, no adverse events have been reported related to this issue and this recall is not being initiated as a result of adverse events. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated February 22, 2013, to all affected consignees. The letter requested consignees examine their inventory, remove them from use and return them to Synthes. Customers were asked to complete the attached Verification Section at the end of the letter indicating whether or not they have any of the affected product. For questions customers were instructed to call 610-719-5450 or email FieldAction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 417 | 417 | Nationwide Distribution: USA including states of: IL, IN, MA, MI, MO, NJ, NY, OH, OR, PA, SC, TX, VA, WA and WI. | Terminated | August 31, 2015 | 920 | 417 | Not_Computer
Z-1206-2013 | 64736 | No | OFP | Anesthesia breathing circuit. | Anesthesiology | General Hospital | Enforcement Discretion | Presource PBDS, Knee Arthroscopy, Kit, Circulator |  |  | March 26, 2013 | 2013 | Cardinal Health, Medical Products &amp; Services | Various Presource Kits contain a pre-assembled Filter and Anesthesia Circuit that may contain outer plastic packaging on one or more components. If the packaging is removed without disassembling the components, remnants of the plastic from the packaging material may become lodged in the filter potentially causing an obstruction of airflow. | Packaging process control | Affected consignees were notified with a letter, dated March 26, 2013, sent via FedEx Priority, of this action. Cardinal Sales Representatives will also be advised and will verbally communicate the action to affected customers. The letter identified the affected product and the problem. Customers are to locate any affected product and quarantine them. In addition, clinicians should be informed of the issue; and, warning labels should be affixed to the front of each kit instructing users to remove and discard the Anesthesia Circuit and Filter Assembly. In order to confirm receipt of this action and completion of labeling the product, the enclosed Acknowledgement Form should be completed and returned to the number provided. If customers do now wish to accept overlabeled product or to overlabel the product in their possession, they are to contact Presource Sales Operations at 800-766-0706, or their sales representative for further instruction. Customers are to contact Cardinal Health Quality Systems at 800-292-9332, if they have any questions or need special assistance. |  | 4571 | 4,571 Kits total | Distributed in the states of AR, MS, SC, and TN. | Terminated | April 12, 2017 | 1478 | 4571 | Not_Computer
Z-1244-2013 | 64705 | No | MQF | Assisted reproduction catheters. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Wallace Sure-ProÂż Single Stage Embryo Replacement Catheter |  |  | March 28, 2013 | 2013 | Smiths Medical ASD, Inc. | Unable to pass catheter through outer sheath during embryo transfer | Process control | The firm, Smiths Medical, sent an "URGENT FIELD SAFETY NOTICE" dated March 28, 2013 to its customers. The notification describes the product, problem and actions to be taken. The customers were provided with two options for proceeding: 1) continue to use product if the product passes a pre-test; or 2) return affected product to Smiths Medical for replacement or credit. In addition, the customers were instructed to inspect their inventory and quarantine all affected products; perform a count of affected products currently in inventory; complete and return the attached Confirmation Form via fax to +1-866-258-6068 or email to pamela.harrell@smiths-medical.com; send a copy of this notice to your customers identified having received affected product; and provide email confirmation of your customers notifications to Smiths medical at pamela.harrell@smiths-medical.com. Customers should report any issues with these products to Smiths Medicals Global Complaint Department at +1-866-216-8806 or globalcomplaints@smiths-medical.com. If you should have any questions regarding this information, please contact Smiths Medical's Customer Service Department at +1-800-258-5361. |  | 800 | 800 units | Worldwide distribution: USA (Nationwide) and countries of: Arab Emirates (AE) Ireland (IE), Albania, Israel (IL), Argentina (AR),Italy (IT) Armenia (AM) , Japan (JP), Austria (AT) Jordan (JO), Bahrain (BH) Kuwait (KW) Belgium (BE), Lebanon (LB), Belarus , Netherlands (NL), Brazil (BR),Nigeria (NG), Bulgaria (BG) , Pakistan (PK), Canada (CA) Poland (PL), China (CN) Romania (RO), Columbia (CO), Russia (RU), Cyprus(CY) , Saudi Arabia (SA) Czech (CZ) , Singapore (SG), Denmark (OK), South Korea (KR), Egypt (EG), Spain (ES), Finland (FI), Sri Lanka (LK),Germany (DE) Turkey (TR) Ghana (GH) , Uganda (UG), Great Britain (GB) Ukraine (UA), Greece (GR) Uruguay (UY), and Hong Kong (HK). | Completed |  |  | 136349 | Not_Computer
Z-1267-2013 | 64962 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Codman(R) Collins Radioparent Sternal Blade Nylon,4 3/4" (121 mm) long, 3" (76 mm) wide, 1 1/2" (38 mm) deep, REF 50-8081 |  |  | April 10, 2013 | 2013 | Symmetry Medical/SSI | There is a possibility that the Collins Radioparent Sternal Blades Nylon may break during use. | Under Investigation by firm | Symmetry Surgical sent an Urgent Medical Device Recall letter dated April 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory immediately to determine if they have any of the affected product. If so, customers were instructed to discontinue use of the product immediately and promplty call 800-251-3000 for a return material authorization (RMA) number and to identify if they would like credit or an alternative product. Customers were asked to fill out the attached Recall Response Form and fax the completed form to 615-964-5567. Customers with questions were instructed to call 615-964-5515.For questions regarding this recall call 615-964-5515. |  | 5240 | 524 boxes of 10 | Worldwide Distribution - USA including CA, MS, FL, PA, TN, CO, IL, TX, OH, MI, MA, IL and internationally toBelgium, Canada, Germany, UK | Terminated | May 12, 2014 | 397 | 5240 | Not_Computer
Z-1248-2013 | 53487 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Mercury Spinal System DIA 5.5 x 50 MM Lordosed Rod, |  |  | April 17, 2009 | 2009 | Spinal Elements, Inc | Spinal Elements has initiated a recall for the Mercury Spinal System due to a raw material issue. | Nonconforming Material/Component | Spinal Elements sent Urgent Medical Device Recall notification letters to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Chief Technology Officer at Spinal Elements for questions about the recall letter. |  | 68 | 68 units total | Distributed in OH and PA. | Terminated | May 07, 2013 | 1481 | 68 | Not_Computer
Z-1269-2013 | 64696 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | A) Product Code 2N3371: Non-DEHP Y-Type Catheter Extension Set, 5.7", Vol. 0.8 mL, 2 Injection Sites, Male Luer Lock Adapter, Single use only |  |  | March 28, 2013 | 2013 | Baxter Healthcare Corp. | Customers have reported that when separating an individual package from its attached grouping, the adjacent package has opened, compromising its sterile barrier properties. | Packaging process control | Baxter sent an "Urgent Product Notification" dated March 28, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taked. Customers are advised to locate and remove all of the affected products from their inventory. If the products were further distributed, consigneees are requested to notify their customers of the recall. Products can be returned for credit by contacting Baxter. An acknowledgement form is required to be compelted by customers. |  | 228200 | A) Product Code 2N3371: 228,200 units | Worldwide Distribution - USA Nationwide and the countries of Puerto Rico, Brazil, Canada and Columbia. | Terminated | July 31, 2017 | 1586 | 844300 | Not_Computer
Z-1276-2013 | 64988 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ExoShape Soft Tissue Fastener 12 mm x 30 mm REF 1101-00-1230. |  |  | April 04, 2013 | 2013 | Medshape Solutions, INC. | Expiration date on the patient label and date on the product carton were not the same. | Error in labeling | Medshape sent an "URGENT: MEDICAL DEVICE RECALL" letter to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 26 | 26 | Nationwide Distribution including the states of CO, FL, GA, IA, MD, MO, OR, TX, VT, VA, and WI. | Terminated | August 07, 2013 | 125 | 26 | Not_Computer
Z-1216-2013 | 60555 | No | OAY |  |  | Gastroenterology/Urology | PMA | KARL STORZ Photodynamic Diagnosis (PDD) D-Light C System, |  |  | August 16, 2012 | 2012 | Karl Storz Endoscopy America Inc | The recall was initiated because Karl Storz has confirmed that the Photodynamic Diagnosis (PDD) D-Light C System instruction manuals have important safety-related changes. | Process control | A Safety Alert letter, dated 8/16/12, was sent to all their customers to inform them that there is a labeling error in which important safety-related changes were made to the KARL STORZ Photodynamic Diagnosis (PDD) D-Light C System instruction manuals. The correct Ethylene Oxide (EtO) Sterilization set point for Relative Humidity should read, =70% RH. The correction applies to the PDD Camera Head instruction manual, PDD Rigid Telescopes instruction manual, and Tricam SL II instruction manual, the three component manuals indicated for use within the KARL STORZ Photodynamic Diagnosis (PDD) D-Light C System. The letter instructed the customers to place each of the three enclosed stickers, in the appropriate instruction manual. Customers are instructed to cover the erroneous information with the provided stickers as indicated. Customers with any questions regarding the proper use of the device, are instructed to contact Karl Storz's Technical Support Department at 800-421-0837, Ext. 5350, during office hours; or Ext. 7222, 24 hours/7 days a week for emergencies. |  | 15 | 15 units | Nationwide distribution: USA including states of: IL, CA, WV, TX, SC, WI, NY, PA, NE, NJ, MD, OH, IA, and NC. | Terminated | May 08, 2013 | 265 | 15 | Not_Computer
Z-1272-2013 | 64699 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Siemens Healthcare Diagnostics Inc. MIcroScan MICroSTREP plus Type 1 B1027-201 panels. |  |  | March 06, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Customer complaint investigation confirmed the issue of false skipped wells and false susceptible misreads with A. agalactiae affecting multiple antimicrobial agents on MicroScan Microbiology Systems. | Under Investigation by firm | Siemens Healthcare sent a Urgent Field Safety Notification letter dated March 6, 2013, to all affected customers of MicroScan MICroSTREP plus type 1, who process panels in a Walk Away System. The letter identified the product the problem and the action needed to be taken by the customer.As stated in our product Instructions for Use, test results should be interpreted in conjunction with the patients medical history, clinical presentation and other findings. We recommend discussing the content of this letter with your laboratory director regarding the need to review previous test results, conduct patient follow-up, and/or repeat testing for testing conducted in the last twelve months.Confirmation of receipt of this letter is required. Attached you will find a form indicating you have received and understood the information. We would greatly appreciate your assistance in notifying us that you have received the information by taking a moment to complete and return the form below.If you require further information or assistance, please contact your local Siemens Healthcare Diagnostics Representative. Please forward this notification to anyone to whom you may have distributed this product.We apologize for the inconvenience that this situation has caused. For further questions please call 1-800-677-7226 Option 1 (USA/Canada)Thank you for your patience and continued support. |  | 432160 | 21,608 boxes = 432,160 panels | US Nationwide Distribution including Puerto Rico. | Terminated | October 28, 2013 | 236 | 432160 | Not_Computer
Z-1273-2013 | 64898 | No | HFW | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Cypress Medical Products LLC, MediChoice Umbilical cord clamp, sterile; single use; Distributed by Owens &amp; Minor, Mechanicsville, VA. The product is used to clamp the umbilical cord after childbirth. |  |  | April 15, 2013 | 2013 | Cypress Medical Products LLC | Cypress Medical Products LLC is voluntarily recalling one lot of the Owens and Minor (O&M;) Umbilical Cord Clamps because recent complaints indicate the clamps are breaking upon use. The product may not perform as intended. | Unknown/Undetermined by firm | Cypress Medical Products sent an "Urgent Voluntary Medical Device Recall" notification and an associated response form dated April 15, 2013 to their only customer for this device - Owens &amp; Minor. The product problem was identified and recommended actions provided. |  | 150 | 150 Cases (75,000 Units) | US Distribution to the state of VA only. | Terminated | August 10, 2015 | 847 | 150 | Not_Computer
Z-1277-2013 | 64935 | No | KNS | Endoscopic electrosurgical unit and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus SurgMaster UES-40 electrosurgical unit ("UES-40") |  |  | November 10, 2009 | 2009 | Olympus America Inc. | Olympus initiated a software upgrade and enhancement in 2009 and completed the action in 2010. The upgrade and enhancement were initiated due to complaints of overheating, smoke vapors and failure to operate associated with the Olympus SurgMaster electrosurgical unit, model UES-40 ("UES-40") when used in conjunction with the Olympus WA22557C electrode. The complaint investigations revealed that | Use error | Olympus America Inc sent a Field Corrective Action letter dated November 10, 2009 to all affected consignees. The letter identified the affected product, problem and actions to be taked. Consignees were directed to contact Gyrus at 888-524-7266 to schedule return of the UES-40 generator for the software upgrade. The letter also directed consignees to contact OAI Customer Service at 800-848-9024 if they needed a new copy of the UES-40 Instruction Manual. For questions call 484-896-5688. |  | 608 | 608 | Nationwide Distribution - USA Nationwide including the states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NV, NY, OH, OR, PA, RI, SC, TN, TX, UT, VA, WA, WV, and WY. | Terminated | July 05, 2013 | 1333 | 608 | Computer
Z-1278-2013 | 64913 | No | NBH | Arthroscope. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Mitek Tissue Liberator -Blade Up |  |  | April 15, 2013 | 2013 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Labeled incorrectly as a Blade Up configuration instead of a Blade Down configuration. | Process control | DePuy Synthes sent an Urgent Voluntary Product Recall letter dated April 15, 2013, to all affected customers. The letter iidentified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check all inventories to locate and return the affected product. For returns customers were instructed call 1-877-379-4871 to obtain a goods authorization (RGA) number. Affected product should be returned to:Mitek Sports MedicineATTN: Mitek Complaints325 Paramouont DriveRayham, MA 02067For questions customers were instructed to call 1-877-379-4871. |  | 27 | 27 units | Worldwide Distribution - USA including KY, NJ and Internationally to Argnetina, Australia, Belgium,Columbia, India, France, Japan, Jordan, Latvija, Israel, Korea, and Great Britain.For questions regarding this recall call 508-977-3813. | Terminated | March 11, 2016 | 1061 | 27 | Not_Computer
Z-1245-2013 | 64876 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle InsuLinx Blood Glucose Monitoring System: |  |  | April 15, 2013 | 2013 | Abbott Diabetes Care, Inc. | AT rare, extremely high glucose levels (1024 mg/dL and above), the FreeStyle InsuLinx Blood Glucose Meter may provide an inaccurate reading. | Software design | Abbott Diabetes Care sent an Urgent Product Recall letter dated April 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take one of the following options immediately to address the issue. 1) Customers can access a software update to resolve the issue at www.freestyleinsulinx.com/swupdate.2) Customers can contact Abbott Diabetes Care Customers Service at 1-866-723-2697 to expedite return and replacement of the affected product at no charge. For questions customers were instructed to call Abbott Diabetes Care Customers Service at 1-866-723-2697.For questions regarding this recall call 510-239-2775. |  | 544609 | 335,273 meters; 209,336 remain at firm. | Nationwide Distribution | Terminated | August 03, 2015 | 840 | 544609 | Computer
Z-1282-2013 | 64778 | No | GXQ | Dura substitute. | Neurology | Neurology | 510(k) | Dural Graft Matrix |  |  | March 28, 2013 | 2013 | Integra LifeSciences Corporation | Due to the process deviation, product lots in question may have been released with higher levels of pyrogens than permitted by the specifications for the products. | Process control | The firm, Integra, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated April 9, 2013, to all US, Canada, and non-EU Integra consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to verify if they have any lots of the listed product, if so, STOP using that product; remove the product from service; and promptly complete and return the PRODUCT RECALL RETURN ACKNOWLEDGMENT FORM via Email: FCA@integralife.com or FAX to : 1-609-275-9445. Customer Service department will contact you to return and replace the product."However, if you have already implanted or used collagen sponge products affected by this recall, we recommend you monitor the patient for a fever in the immediate postoperative period according to the standard hospital or clinician protocol."For assistance or any other questions that you may have, please contact Customer Service at 1-855-532-1723. |  | 15607 | 15,607 boxes | Worldwide distribution: USA (nationwide) including Puerto Rico and countries of: Arab Emirates, Argentina, Belgium, Brasil, Canada, Chile, China, Costa Rica, Denmark, Egypt, Spain, France, Great Britain, Italy, Japan, Peru, Singapore, Thailand, and South Africa. | Terminated | June 03, 2015 | 797 | 587490 | Not_Computer
Z-1281-2013 | 64887 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Insight Instruments MIS ( Multi-Port Illumination System) |  |  | March 15, 2013 | 2013 | Insight Instruments, Inc. | Insight Instruments, Inc., Stuart, FL is recalling M.I.S. (Multi-Port Illumination System), Ophthalmic Cannula due to a misclassification of the medical device. | Use error | Insight Instruments, Inc. sent a recall letter dated March 15, 2013, to all affected customers The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to do the following: 1)Perform a physical count of their inventory immediately for any of the recalled lots and complete and return the Return Response Form. 2) If they have any products listed, contact the firm at (800) 255-8354, to arrange for return and prompt replacement. 3) Discontinue using and dispensing all M.I.S. devices of the recalled lots. |  | 776 | 776 | US Distribution IN and IL | Terminated | August 21, 2014 | 524 | 776 | Not_Computer
Z-1280-2013 | 64875 | No | EZY | Implanted mechanical/hydraulic urinary continence device. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | AMS 800Âż Urinary Control System. |  |  | March 01, 2013 | 2013 | American Medical Systems, Inc. | American Medical Systems, Inc. is initiating a recall on two components of the AMS 800Âż Urinary Control System, Control Pump &amp; Pressure-Regulating Balloon. These products were released despite not meeting all final product release criteria. | Release of Material/Component prior to receiving test results | American Medical Systems sent two written communications distributed within the United States. The first is a letter that is being delivered to customers who purchased the affected lots using their ship to address from the original sales order. The second is a letter being delivered to implanting physician customers who are known by AMS to be regular users of the devices. Similar letters will be sent to international customers, including both end users and distributors (where applicable). American Medical Systems sent a "Urgent Medical Device Recall" letter dated February 28, 2013, to all affected customers. The letters identified the product, the problem, and the action needed to be taken by the customer. Customers were instructed to return the recalled product and to complete the Product Recall Acknowledgement Form. The letters addressed to "Physicians" advised consignees against prophylactic removal of implanted control pumps and pressure-regulating balloons contained within the scope of the product recall.For further questions please call AMS Customer Service at 1 (800) 328-3881 |  | 77 | 77 | Worldwide Distribution -- Nationwide Distribution including the states of AL, AZ, AR, CA, CO, CT, DE, FL, GA, IL, IN, IA, KS, KY LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NH, NJ, NM, NY,NC, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VT, VA, WA and WI., and the countries of AUSTRIA, AUSTRALIA, CANADA, BELGIUM, SWITZERLAND, CZECH REPUBLIC, GERMANY, DENMARK, ESTONIA, SPAIN, FINLAND, FRANCE, GREAT BRITAIN, IRELAND, ITALY, LEBANON, NETHERLANDS, NORWAY, SWEDEN, TURKEY, ARGENTINA, BRAZIL, IVORY COST, COLOMBIA, ECUADOR, INDIA, JAPAN, MEXICO, NEW ZEALAND, PANAMA and VENEZUELA. | Terminated | April 24, 2014 | 419 | 682 | Not_Computer
Z-1305-2013 | 55358 | No | FJL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Electrode, cutting, 24FR (27050G/6). |  |  | June 01, 2009 | 2009 | Karl Storz Endoscopy America Inc | Possible damage to the sterile packaging (blister tray) of some Karl Storz sterile products, which may result in airborne contamination. | Packaging | A Safety Alert dated May 29, 2009 was sent to affected customers. The letter informed them of possible damage to the sterile packaging (blister tray) of some Karl Storz sterile products which may result in airborne contamination. Karl Storz requested consignees to inspect their inventory of the subject products thoroughly and look for small holes in the blister trays. Any identified units should be returned to Karl Storz, and a replacement package (6 units per package) will be provided to at no charge. The Karl Storz Customer Support number 1-800-421-0837 ext. 7577 was provided to obtain a Return Authorization Number (RMA) and arrange replacement. Questions and comments were forwarded to Technical Support at 1-800-421-0837 ext. 5350. |  | 9432 | 9432 units | Distributed Nationwide and in Canada. | Terminated | May 13, 2013 | 1442 | 9432 | Not_Computer
Z-1304-2013 | 64908 | No | KDD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Curity Staple Remover Kit; Product Code: 66701. |  |  | April 12, 2013 | 2013 | Covidien LLC | On March 26, 2013 Covidien received a recall notice from one of their suppliers CareFusion. CareFusion is recalling a variety of OTC drug products including the 10% Povidone Iodine SEPP Applicator which is utilized in specific convenience kits. | Error in labeling | US Customers were notified of the recall by letter sent by Federal Express on April 12, 2013. The letter requests that they immediately stop using any affected product, quarantine it, and return it to Covidien. Consignees were asked to return a verification form and return it to Covidien via fax. |  | 52363 | 52363 distributed | Worldwide distribution: USA (nationwide) and in the country of: Panama. | Terminated | June 02, 2015 | 781 | 70103 | Not_Computer
Z-1317-2013 | 65112 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | PadPro(R) Adult Radiotransparent Multifunction Electrodes, REF/Catalog No. 2001M-C. |  |  | May 03, 2013 | 2013 | ConMed Corporation | ConMed was notified by the contract manufacturer that their wire/connector supplier reported damage to the connector tooling which could pinch or compress the lead wire filaments and potentially render the device incapable of delivering the appropriate energy or shock to the patient and/or could cause potential injury to user or patient. | Component design/selection | ConMed sent a URGENT: Field Safety Notice letter dated May 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to please review your inventory for any of the devices listed on Attachment I.Business Reply Forms were sent to the domestic and foreign consignees on May 3, 2013. The domestic notices were mailed by USP Priority Mail pouches. The recall is being conducted to the end user level.We ask that you contact all of those organizations within your facility and any other facilities that you may have supplied or given these affected products to. It is imperative that all end users of these devices receive this notice and respond immediately. If you have questions, please call 1 315-624-3237 orfax to +1 315-624-3225. |  | 600 | Domestic: 49,010; Foreign: 600 | Worldwide distribution: USA (nationwide) and the country of: Saudi Arabia. | Terminated | August 25, 2015 | 844 | 600 | Not_Computer
Z-1319-2013 | 65034 | No | JYO | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Summit Medical, BL-2006, Myringotomy Lance (Juvenile) Blade, Sterile P, Summit Medical, Inc., 815 Northwest Parkway, Suite 100, St. Paul, MN 55121 |  |  | March 27, 2013 | 2013 | Summit Medical, Inc. | Summit Medical has initiated a recall of Myringotomy Lance Blade (Juvenile) because the box contains the wrong blade. The box contains Myringotomy Spear Blade (Juvenile). | Process control | Summit Medical sent a Medical Device Recall letter dated March 27, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to immediately examine your inventory, and quarantine the product subject to this recall. Please complete and return the enclosed "Recall Response Form" as soon as possible. The fax number is (651) 789-3961.For further questions please call (651) 789-3921. |  | 9 | 9 boxes | Distributed in MS, NY, PA and TX. | Terminated | August 06, 2013 | 132 | 9 | Not_Computer
Z-1320-2013 | 64650 | No | LKQ | Cytomegalovirus serological reagents. | Microbiology | Microbiology | 510(k) | QIAGEN artus CMV RG PCR ASR (96) (Catalog number 4503225) |  |  | January 30, 2013 | 2013 | QIAGEN Gaithersburg, Inc. | CMV RG PCR kits may contain a mixture of incorrect vials among the correct vials. | Mixed-up of materials/components | Qiagen Gaithersburg, Inc. issued a PRODUCT FIELD ACTION letter dated February 4, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect the vials prior to use to ensure that the vials are labeled "CMV RG Master" Vials labeled "EBV TM Master" should NOT be used and count the vials and discard per internal laboratory procedures. Qiagen will replace free of charge or credit any kits discovered with incorrect master mix vials. Please complete page 2 of this notice as soon as possible and email or fax it back to us at: (661) 702-3854, techservice-na@qiagen.com. If you have additional questions or concerns polease do not hestiate to call QIAGEN Technical Services at 1-800-344-3631 option 2 or by e-mail techservice-na@qiagen.com. |  | 90 | 90 kits in total | USA Nationwide Distribution including the states of: CA, DC, FL, GA, MA, MD, MI, and IL. | Terminated | July 01, 2013 | 152 | 90 | Not_Computer
Z-1318-2013 | 65029 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova 2100 IQ, Innova 3100 IQ, Innova 4100 IQ, Innova 2121 IQ, Innova 3131 IQ , Innova 3100, Innova 4100. |  |  | December 20, 2011 | 2011 | GE Healthcare, LLC | While performing a fluoroscopic examination on the GE Innova systems, there is a potential of nonrecoverableloss of imaging modes with no advance warning, due to a relay failure in the power distribution box (PDB). Although no injuries have been reported due to this issue to date, there is riskof complications during the procedure as a result of imaging functionality loss. | Nonconforming Material/Component | GE Healthcare did not send out customer letters. GE Healthcare completed the correction by replacing the faulty relay in all affected systems.As an effectiveness check, they have tracked all service requests for the field action to completion. These service requests recorded the repair of the affected units. Exceptions to completion of the service request were recorded as well. Records of field action completion will be retained per their processes and procedures.Should you have any questions or require additional information, please call (262) 548-2757. |  | 713 | 713 | Worldwide Distribution -- Nationwide including the states of AL, AZ, AR, CA, CO, CT, DC, FL, GA, GUAM, HI, ID, IL , IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NH, NM, NY, NC, ND, OH, OK, OR, PA, PR, SC, TN, TX, UT, VT, VA, WA, WY and WI., and the countries of Philippines, Peru, Paraguay, Panama, Mexico, Malaysia, Korea, Japan, Guatemala, Ecuador, Dominican Republic , Columbia, China, Chile, Canada, Brazil.,Argentina and Venezuela | Terminated | March 13, 2014 | 814 | 713 | Not_Computer
Z-1322-2013 | 64221 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Proplege Coronary Sinus Catheter model PR9, 11 French Introducer, kitted with PR9. |  |  | January 25, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain Proplege Coronary Sinus Catheter, model PR9 due to confirmed reports of blood leakage through the hemostasis valve in the introducer that is kitted with the catheter. | Device Design | Edwards Lifesciences sent an Urgent Field Safety Notice letter dated February 1, 2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review entire inventory for the affected product, quarantine and return to Edwards. Contact Customer Services at 800-424-3278 for any question, to obtain a Returned Goods Authorization number and replacement product. |  | 807 | 807 | Worldwide Distribution - USA Nationwide including the states of: MD, NY, PA, CA, FL, AZ, AL, TX, VA, OH, MO, NC, GA, TN, CO, OR, SC, ME, IL, HI, NJ, WA, WI, CT, KY, LA and Canada. | Open, Classified |  |  | 807 | Not_Computer
Z-1323-2013 | 64390 | No | DSQ |  |  | Cardiovascular | PMA | Heartware HVAD Pump Implant Kit (Heartware Ventricular Assist Device) |  |  | December 27, 2012 | 2012 | Heartware Inc | HeartWare, Inc. of Miami Lakes, FL is recalling their Ventricular Assist System due to the loosening of the HVAD Pump driveline connector. | Device Design | HeartWare, Inc. sent an Urgent Medical Device Correction letter to all their customers on February 7, 2013 instructing that at each routine clinic visit, they should inspect the patient's driveline connector as described in the attached Technical Bulletin. They were asked to sign and return the acknowledgement form to HeartWare. For questions contact your HeartWare representative or Customer Service at 1-877-367-4823. |  | 3544 | 3,544 | Worldwide Distribution-USA (nationwide) including the states of MN, IL, MI, FL, WI, TX, MO, CA, MA, TN, OH, NC, GA, VA, UT, KY, MD, IN, PA, NY, NJ, LA, OR, WA, AL and CO, and the countries of Poland, Australia, Austria, Italy, Germany, New Zealand, South Africa, Czech Republic, France, Brazil, Switzerland, Singapore, United Kingdom, Saudi Arabia, Netherlands, China, Chile, Belgium, Canada, Norway, Lebanon, Israel, Malaysia, Turkey, India, Argentina, Greece, Japan and Lithuania. | Terminated | July 09, 2013 | 194 | 3544 | Not_Computer
Z-1324-2013 | 64986 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Segmented Cylinder Applicator Set, Model # GM11004150, when used for Pulsed Dose Rate (PDR) |  |  | April 11, 0013 | 13 | Varian Medical Systems, Inc. | Segmented cylinder applicator set may slip during treatment, causing the delivery of radiation to areas outside the target. | Device Design | Varian Medical Systems, Inc. notified customers of the recall with 'Urgent Medical Device Correction Urgent Field Safety Notice," dated 4/11/13, and delivered by mail to customers. The notice identified the product, problem, and instructions for the customers. Varian Medical Systems, Inc. informed customers of the possible malfunction of the GM11004150 Segmented Cylinder Applicator Set for pulsed dose rate brachytherapy treatment, with codes H640383, H640277, H640384, H640788, H640D126, H640385, H640372, H640303, H640387. The notice instructed customers to cease use of the Segmented Cylinder Applicator Set and return the devices along with a completed proof of notification form (included with the recall notice) to the recalling firm. |  | 21 | 21 devices | Worldwide distribution: USA (nationwide) and in the countries of: Belgium, France, Martinique, Slovenia, Sweden, and the United Kingdom. | Terminated | January 13, 2014 | 730762 | 21 | Not_Computer
Z-1274-2013 | 64975 | Yes | NIU |  |  | Cardiovascular | PMA | ZilverÂż PTXÂż Drug-Eluting Peripheral Stent - 6 mm X 80mm; 125 cm, Rx, Sterile |  |  | April 18, 2013 | 2013 | Cook, Inc. | Cook Medical has received a small number of complaints related to the delivery system for the Zilver PTX Drug-ElutingPeripheral Stent involving fractures of the delivery system inner catheter after stent deployment, and separation of theinner catheter tip section (13 complaints with an occurrence rate of 0.043%) Two adverse events, including one death,occurred in cases where a tip separation | Process control | Cook Medical sent an Urgent Medical Device Recall letter dated April 24, 2013 to all affect customers. The letter identified the affected products, description of the problem, and actions to be taken. The letter instructed customers to stop use, quarantine, and return all affected products. Customers were asked to complete product reply form and return affected products for credit to Cook Medical, ATTN: Retrun Goods/RGA # 2013C0003. For questions contact Cook Medical Customer Relations at 1-800-457-4500 or 1-812-339-2235. |  | 1057 | 1,057 units in USA | Worldwide Distribution - USA Nationwide; in the countries of Japan, Europe, Brazil, and Australia | Terminated | December 05, 2013 | 231 | 1592 | Not_Computer
Z-1328-2013 | 64910 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Iris Diagnostics Division iQ Control/Focus Set, Part Number 800-3104, Lot Number 153-12. |  |  | September 18, 2012 | 2012 | Iris Diagnostics | The firm initiated this recall because complaints from customers reported that the recovery compared to assigned assay values is significantly below specification. This results in the instrument not passing controls and the customer not being able to run patient samples. | Nonconforming Material/Component | IRIS sent a Medical Device Recall letter dated September 18, 2012 to all affected customers. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to dispose of all unopened or partially used inventory, complete the attached worksheet certifying the disposal of the affected product and email or fax to Iris Diagnostics. |  | 942 | 942 boxes | Worldwide Distribution - USA Nationwide and the countries of: Canada, Malaysia, Spain, Switzerland, Taiwan, Peru, France, Russia, Japan, United Kingdom, Germany, and Paraguay. | Terminated | January 27, 2014 | 496 | 942 | Not_Computer
Z-1325-2013 | 64997 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Varian High Energy Clinacs, High Energy Accelerator, Radiation Treatment System, Model Numbers: H14, H27, H29, HCX. |  |  | May 31, 2011 | 2011 | Varian Medical Systems, Inc. Oncology Systems | Under certain conditions, photon beams in High Energy Clinacs may experience a gradual change in beam symmetry, potentially reaching an asymmetry of up to about 7% before interlock occurs. At maximum asymmetry, this may result in no more than about 3.5% dose deviation from expected at any point in the beam, and may result in minor injury to the patient or slightly decreased local disease control. | Process design | Varian sent a Customer Technical Bulletin to all affected customers. The bulletin identified the affected product, problem and actions to be taken. For questions, issues, or concerns contact Varian at 1-888-827-4265 or go to http://my.varian.com |  | 3943 | 3943 units | Worldwide distribution - USA (nationwide) including Puerto Rico, and countries of: Algeria, Angola, Australia, Austria, Bahrain, Bangladesh, Belarus, Belguim, Brazil, Bosnia and Hezregovina, Bulgaria, Canada, Chile, China, Columbia, Cyprus, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, Frances, French Polynesia, Georgia, Germany, Greece, Gaudaloupe, Gaum, Hong Kong, India, Indonesia, Ireland, Italy, Israel, Latvia, Lithuania, Luxemburg, Japan, Kazarhstan, Kuwait, Lebanon, Macau, Macedonia, Malaysia, Martinique, Mauritania, Mexico, Moldova, Monaco, Morroco, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Peru, Phillipines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudia Arabia, Serbia, Singapore, Slovakia, Spain, South Africa, South Korea, Sri Lanka, Syria, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezula, Vietnam, and Zimbabwe. | Terminated | May 15, 2013 | 715 | 3943 | Not_Computer
Z-1330-2013 | 64959 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquichek" ToRCH Plus IgM Control - Positive |  |  | April 15, 2013 | 2013 | Bio-rad Laboratories, Inc. | The Liquichek" ToRCH Plus IgM Control - Positive, contains detectable levels of HCV RNA which originatedfrom the non-specific human IgM component used in the mousemonoclonal antibody conjugation process for this product. | Material/Component Contamination | Bio-rad Laboratories initiated this recall on April 18, 2013, via telephone call and certified mailing to customers. The notification letter titled "Important Product Information Notice of Correction" was accompanied by a "Field Correction Notification Acknowledgement" form that notified customers of the product for recall (description and codes), issue (This product contains detectable levels of HCV RNA which originated from the non-specific human IgM component used in the mouse monoclonal antibody conjugation process for this product), Customers were instructed to immediately discontinue use this lot of Liquichek ToRCH Plus IgM Control. Discard all remaining inventory of this lot of Liquichek ToRCH Plus IgMControl that you have at your location in accordance with therequirements of your local waste management authorities. Handle this product with the same precautions used with patientspecimens. In accordance with good laboratory practice, all humansource material should be considered potentially infectious. Have the appropriate laboratory personnel complete theacknowledgment form and email or fax it to us. It is very important to usto know that you have received the information provided in this Notice.Contact information (1-800-2-BIORAD) |  | 216 | 211 boxes and 5 MiniPaks | Worldwide Distribution - Distributed in San Juan Puerto Rico and the states of NM, NY, MN, OR and WI., and the countries of Australia, Brazil, France, Italy, Korea, and New Zealand. | Terminated | August 01, 2013 | 108 | 216 | Not_Computer
Z-1329-2013 | 64800 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | PROFILE-V MEDTOXScan Drugs of Abuse Test System. |  |  | April 05, 2013 | 2013 | Medtox Diagnostics Inc | Review of validation information and customer complaints which indicated preliminary screen positive results for cocaine that were not positive on confirmatory testing. | Nonconforming Material/Component | MEDTOX issued "URGENT MEDICAL DEVICE RECALL" letters dated April 3, 2013 to consignees informing them of the recall. The letter requested that they examine their inventory and immediately discontinue all use and distribution of the affected lots. The letter also requested that they immediately inform clinicians of the recall at the testing sites that have utilized the recall product. The letter also included a Customer Verification Form which was to be returned. |  | 221625 | 221,625 | Worldwide distribution: USA (nationwide) and in the country of: Canada. | Terminated | September 12, 2013 | 160 | 221625 | Not_Computer
Z-1331-2013 | 65078 | No | FYE | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DePuy LifeLiner Stick &amp; Cut Resistant Gloves Rt, Lg |  |  | May 08, 2013 | 2013 | DePuy Orthopaedics, Inc. | LifeLiner Stick and Cut Resistant Gloves are being recalled because cases and individual glove packages were mispackaged, with right handed, large sized glove packaging containing left handed, large sized gloves. | Labeling mix-ups | DePuy Synthes Joint Reconstruction, a division of DePuy Orthopaedics, Inc.,sent a " Urgent Information Medical Device Recall Notice" dated May 8, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please cease further distribution or use of the affected gloves immediately. Our records indicate your hospital purchased the affected lot. The purpose of this communication is to inform you of this recall and requestacknowledgement of receipt of this letter by signing and returning the Hospital/User Facility Reconciliation Form within 10 business days of receiving this notice.Response cards should be faxed to (574) 371-4964. Please contact your DePuy Synthes Joint Reconstruction Sales Representative for assistance returning any unused affected gloves.For product related questions, please contact your local DePuy Synthes Joint Reconstruction Sales Consultant. For clinical questions from surgeons, please contact DePuy Synthes Joint Reconstructions Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m.- 5 p.m. EDT). For questions about device recall information provided, please call (574)371-4917 (M-F; 8 a.m. - 5 p.m. EDT). |  | 180 | 180 gloves (120 gloves distributed in the US) | Worldwide Distribution - US (nationwide) including the states of IL, OH, NC, SC, MN, NV, TN and WA., and the countries of UK, Netherlands and Canada. | Terminated | January 14, 2014 | 251 | 180 | Not_Computer
Z-1334-2013 | 64883 | No | JSG | Multipurpose culture medium. | Microbiology | Microbiology | 510(K) Exempt | BBL(tm) CDC Anaerobe 5% Sheep Blood Agar. |  |  | March 11, 2013 | 2013 | Becton Dickinson &amp; Co. | Microbiological media may be contaminated with bacteria. | Material/Component Contamination | BD Diagnostics notified consignees with a "Field Correction Notification" sent on 3/11/13 to distributors, direct accounts, and distributors' customers as available, via fax and email, and followed by a letter sent via UPS. The notice identified the product, the problem, and the action to be taken by customers. BD Diagnostic Systems informed customers of the bacterial contamination of some microbiological media, including BBL Trypticase Soy Agar with 5% Sheep's blood, part number 221239, lot 2348199; part number 221261, lots 2339019, 2348108, 2348208, 2348074, 2348157, 2355145, 2348202, and 2361011; BBL Trypticase Soy Agar with 5% Sheep's blood plus MacConkey II-I plate, part number 221291, lots 2355138 and 2355164; and BBL CDC Anaerobe 5% Sheep blood agar, lot 2361003. The notice instructed customers to follow product insert directions for inspecting media before use to identify and eliminate contaminated media, and to verify Listeria findings with clinicians. The notice included a response form to indicate receipt of the notice. BD Diagnostic Systems contacted distributors and end-users via fax, email, and letter on 4/25/13 to notify them of a recall expansion to include lot 3038371 from part/catalog number 221261. The recall notice included an acknowledgement form to document receipt. The notice instructed customers to follow product insert directions for inspecting media before use to identify and eliminate contaminated media, and to verify Listeria findings with clinicians. |  | 521160 | 521,160 total pieces (all three recalled products) | Worldwide Distribution-USA (nationwide) including Puerto Rico, Virgin Islands, and Guam and the countries of Columbia, Japan, Singapore, Canada, Turks, Calicos Islands. | Terminated | June 24, 2013 | 105 | 521160 | Not_Computer
Z-1326-2013 | 65028 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Fluoroscan Mini C-arm InSight, software versions 1.x, 2.x, 3.x |  |  | April 25, 2013 | 2013 | Hologic, Inc. | The audible alarm to alert the operator when cumulative irradiation time reaches five minutes is not accurate and the total cumulative irradiation time is not accurate. | Software design | The firm, Hologic, Inc. sent an "Urgent: Medical Device Recall" letter dated April 25, 2013, via FedEx to consignees/customers notifying them of the software defect and providing them with corrected user installable software version. The customers were instructed to complete and return the Enclosure to Customer Notification, Confirmation of Notification form via fax to Hologic at 866-652-8674 or mail the document to Hologic using the enclosed pre-addressed, pre-paid return envelope.Should you have any questions regarding this recall, require further information or assistance with the software installation call Hologic's Help Desk at 1-800-321-4659. |  | 287 | 287 | Worldwide distribution: USA (nationwide) and countries of: DE, RU, AU, IT, GB, ES, FR, KR, ZA, CA, AR, LT, TW, CN, BE, AT, JP, QA, SA, SG, HK, SI, NL, and SE. | Terminated | August 21, 2014 | 483 | 3067 | Computer
Z-1336-2013 | 64468 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ACUVUE ADVANCE Brand Contact Lenses with HYDRACLEAR, Soft, disposable contact lenses. |  |  | December 21, 2012 | 2012 | Johnson &amp; Johnson Vision Care, Inc. | Limited number of individual contact lens packages may not have been completely sealed. | Packaging | The firm, Johnson &amp; Johnson, Vision Care, Inc. sent a letter dated January 9, 2013 to all their consignees stating the following actions:" Please contact your patients that may have received any of the affected product and instruct them to discontinue use immediately and return product to you for replacement " Please return the enclosed Business Reply Card (BRCl confirming for our records that you have received this notice, and indicating whether or not you have in stock, or have distributed to patients,any product from the lots noted above. " Please return any remaining affected product in your possession using the enclosed pre-paid mailer. " Pass this notice on to anyone in your practice that needs to be informed " Customer service will arrange for the return and replacement of any affected product. If you have any questions and/or if replacement lenses are needed, please contact customer Service (1-800-87 4-5278). |  | 3546 | 3546 lenses | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, KY, LA, MA, MD, MI, MO, NC, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI and Puerto Rico; and countries of: Austria, Belgium, Canada, Costa Rica, Denmark, France, Germany, Hungary, India, Ireland, Israel, Italy, Korea, Japan, Luxembourg, Netherlands, Portugal, Norway, South Africa, Slovenia, Spain, Sweden, and United Kingdom. | Terminated | July 05, 2013 | 196 | 39876 | Not_Computer
Z-1337-2013 | 64901 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products CK-MB Calibrators, REF 128 6293, IVD, Ortho-Clinical Diagnostics High Wycombe HP12 4DP, UK |  |  | April 03, 2013 | 2013 | Ortho-Clinical Diagnostics | Internal testing demonstrated that opened-reconstituted VITROS CK-MB Calibrators stored at -20 degrees Celsius (-4 degrees Farenheit) generated invalid calibrations that were detected by positively biased quality control results. | Process control | Ortho Clinical Diagnostic sent a " Important Product Correction Notification" letter dated April 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers were informed of the issue and advised they must discontinue storage of opened-reconstituted vials of VITROS CK-MB Calibrators frozen at -20 degrees Celsius (-4 degrees Farenheit). The same day, the Distributor letters were sent via US Postal Service Priority Mail or FedEx overnight courier to US Distributor consignees who were informed of the issue and instructed to notify their customers of the issue and actions. Also on April 3, 2013, the foreign affiliates were informed of the issue by e-mail and instructed to notify their consignees of the issue and actions. Confirmation of Receipt - Important Response Required so that we can complete our records, please return this form to us no later than April 19, 2013.FAX: 1-888-557-3759 or 1-585-453-4110We apologize for any inconvenience this may have caused your laboratory. If you have any additional questions, please call Customer Technical Services at 1-800-421-3311. |  | 3924 | Domestic: 3019 units, Foreign: 905 units | Worldwide Distribution - Nationwide including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 3924 | Not_Computer
Z-1340-2013 | 64647 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Ref: VP1600 HD Still Capture System ConMed Linvatec . |  |  | March 18, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | The VP1600 HD Still Capture Image System is recalled due to an anomaly between the software code in the VP1600 console and the communications interface for various keyboard key strokes (Caps lock and Num Lock) which may lock up and requires the user to reboot to continue. | Software design (manufacturing process) | The firm, ConMed, sent an "URGENT: MEDICAL DEVICE RECALL NOTIFICATION" letter dated 3/8/2013 to its customers after initial phone notifications. The letter identified the product, problem, and actions to be taken by the customers. |  | 267 | 267 units | Worldwide Distribution-USA (nationwide) including the Virgin Islands and the states of CA, CT, FL, GA, IN, KY, KS, LA, MD, MI, MN, MO, MS, MT, NE, NC, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, and WI, and the countries of Austria, Belgium, Canada, Denmark, Hong Kong and Zambia. | Terminated | November 06, 2013 | 233 | 267 | Computer
Z-1338-2013 | 64713 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 16-Slice Water, System Code: 728245; Brilliance CT 10-Slice Water, System Code: 728250; Brilliance CT 6-Slice Water, System Code: 728255; MX8000 IDT 16, System Code: 728120 &amp; MX8000 IDT 10, System Code: 728125 Computed Tomography X-ray systems. Distributed by Philips Medical Systems (Cleveland), Inc., 595 Miner Road, Cleveland, OH 44143 |  |  | March 12, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received a report from the field stating when they viewed the CTDIVol information from the PACS system, the calculations were too high. Instead of storing the CTDIVol value for each slice in the DICOM tag, the system took the sum of all CTDIVol values from all the images in the dataset and stored this value in the DICOM tag, resulting in incorrect information. | Use error | The firm, Philips Healthcare, sent an "Urgent - Medical Device Correction" letter dated March 15, 2013, to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to disregard unusually higher CTDIvol value for Type-3 DICOM tag 0018:9345 in the PACS system, if seen. The firm is implementing software update versions 4.2.2 or 2.2.2 to correct the issue. A Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customers Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 515 | 515 units | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, VA, VT, WA, WI, WV and WY; countries of: Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Denmark, France, Germany, Gibraltar, India, Iraq, Ireland, Israel, Italy, Japan, Mexico, Netherlands, New Zealand, Nicaragua, Norway, Philippines, Poland, Portugal, Singapore, South Korea, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom and Venezuela. | Terminated | June 30, 2015 | 840 | 515 | Computer
Z-1339-2013 | 65096 | Yes | LWS |  |  | Cardiovascular | PMA | LATITUDEÂż Patient Management System, Model 6488, Version 7.3. |  |  | April 10, 2013 | 2013 | Boston Scientific CRM Corp | Boston Scientific CRM is conducting a recall on the LATITUDE Patient Management System Model 6488 Version 7.3 because between March 4 and 27, 2013, a Siebel release caused ~5000 distributed PMR clinic faxes to display with illegible characters and not the required, Patient non-compliance information that was intended. | Software change control | Boston Scientific CRM Corp began re-faxing the 4,912 previously illegible PMR notifications on April 10, 2013 and completed on April 15, 2013, successfully fulfilling labeling expectations that these notifications be communicated to the clinic.For questions regarding this recall call 651-582-4000. |  | 1 | 1 | Nationwide Distribution including DC and PR | Terminated | November 26, 2013 | 230 | 1 | Computer
Z-1341-2013 | 64943 | No | OTO | Surgical mesh. | General &amp; Plastic Surgery | Obstetrics/Gynecology | 510(k) | Rostorelle L, Polypropylene Mesh, Catalog number 501440, Manufacturer Coloplast A/S 3050 Denmark. |  |  | April 12, 2013 | 2013 | Coloplast Manufacturing US, LLC | Coloplast is recalling one lot of Restorelle L prosthesis because it is mislabeled. The incorrect Instructions For Use (IFU) pamphlet is included in the packaging. | Labeling mix-ups | Coloplast sent a letter to dated April 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included an attachment with the correct IFU to replace the incorrect IFU. Customers with questions were instructed to call Customer Service at 800-258-3476.For questions regarding this recall call 612-302-4992. |  | 111 | 111 | Nationwide Distribution including AL, CA, CO, DC, IL, IN, KY, MD, MA, MO, NH, OH, RI, and VA. | Terminated | January 21, 2014 | 284 | 111 | Not_Computer
Z-1344-2013 | 49026 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cardinal Health hCG Cassette Rapid Test, Part Number: B1077-22 (Mfg PN: 4582105026) |  |  | January 21, 2013 | 2013 | Alere San Diego, Inc. | Alere San Diego Inc. is recalling the Cardinal Health hCG Cassette Rapid Test because some of them were incorrectly packaged into Cardinal Health hCG Combo Rapid test kit boxes. The kit box incorrectly indicates the product may be used for both urine and serum sample types. | Labeling mix-ups | The firm, Alere, sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 21, 2013 to all their customers. The letter describes the product, problems identified and the actions to be taken. Customers are instructed to discontinue use and/or sale of the product; immediately discard all units of the the product in accordance with your local regulations; share this information with your laboratory staff and retain this letter for your records; if you have forwarded the product to another laboratory, provide a copy of this letter to them and include the name, address and a contact for the laboratory on the enclosed Customer Verification Form; complete and return the Customer Verification Form within 10 days via fax to: Technical Service at 858-805-8457 or e-mail to responses.ts@alere.com. Customers with any questions are instructed to contact Alere Technical Services at (866) 216-0094 or by email at lateral.flow.support@alere.com. |  | 2167 | 2,167 kits | Nationwide distribution: USA including states of: AL, AK, AR, AZ, CA, CO, CT, DE, HI, IA, ID, IL, IN, FL, GA, KY, KS, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, UT, TN, TX, VA, WA, WI, WV AND WY. | Terminated | August 28, 2013 | 219 | 2167 | Not_Computer
Z-1345-2013 | 56447 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TD-Synergy. |  |  | July 29, 2010 | 2010 | Technidata S.A. | Technidata has discovered that using the F8 key in the ERM session to add a combined test can result in the modification of the value of a result that has already been entered for a test of the request. | Software design | The firm, Technidata, sent an "P1 DEFECT INFORMATION" notice dated July 2, 2010 to their consignee/customers. The notice describes the product, problem and actions to be taken. The notice directs them to find the correction to the software issue on the Technidata website where they are able to install the correction kit depending on the version and operating system used. Note "You must back up the database and the installation directory before installing this correction".Should you need any further information, do not hesitate to contact the Support team at +33 (0) 4 76 04 13 00. |  | 10 | 10 | Worldwide Distribution-USA (nationwide) including Washington DC and the states of California, Nevada, Kentucky, Alabama, Rhode Island, Pennsylvania, and Louisiana and the country of Italy. | Terminated | December 05, 2013 | 1225 | 10 | Computer
Z-1342-2013 | 64512 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Molift Smart 150 |  |  | February 15, 2013 | 2013 | Moller Vital | The leg spreading mechanism will fail to hoist and collapse if assemble incorrectly by the user. | Employee error | Molift Group AS issued a Field Safety Notice letter to all affected customers. The letter identified the affected products, problem and actions to be taken to the Smart hoists that are already in use. Customers were instructed to make sure that all affected customers receive this information and encorage to increase attention until the affected product has been upgraded. For questions contact your local Molift representative or Customer Service at: groupsales@molift.com. |  | 15110 | 7,555 units subject to review; US: 1,280 units - Foreign: 6,275 units | US Nationwide Distribution: CA, FL, GA, ID, IL, KY, MA, ME, MI, MN, NC, NH, NJ, NM, NV, NY, OH, OK, PA, TN, TX, VT, WA and Hawaii.Foreign Distribution:: Australia, Austria; Belgium, Denmark, Finland, France, Germany, Iceland, Israel, Italy, Japan, Netherlands, New Zealand, Norway, Spain, Sverige, Switzerland, United Kingdom. | Terminated | July 11, 2014 | 511 | 15110 | Not_Computer
Z-1346-2013 | 64930 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobasÂż 4800 system z480 instrument. |  |  | April 01, 2013 | 2013 | Roche Molecular Systems, Inc. | In rare instances, channel shifted results have been generated with cobasÂż 4800 assays run on v1.1.1 of the cobasÂż 4800 system. Raw data review indicates that, in these cases, all signals are shifted by one channel:channel 1 results report as channel 2,channel 2 results report as channel 3,channel 3 results report as channel 4, andchannel 4 results report as channel 1.In the reported | Software design | Roche sent an "Urgent Medical Device Correction" (UMDC) notification with response form dated April 1, 2013, to all affected customers via UPS Ground on April 1, 2013. The letter identified the product, the problem, and the action to be taken by the customer. A contact number (Roche Diagnostic Technical Support) is provided for technical support 24 hours a day, seven days a week at phone number: 1-888-912-7090. |  | 425 | 425 instruments | Nationwide Distribution. | Terminated | August 07, 2014 | 493 | 425 | Not_Computer
Z-1179-2013 | 65018 | No | EAP | Intraoral source x-ray system. | Dental | Radiology | 510(k) | Fona Srl IntraOs 70 intra-oral x-ray system |  |  | July 03, 2012 | 2012 | Chicago X-Ray Systems, Inc. | The IntraOs 70 system was found to have missing or incomplete system labels as required by 21 CFR 1010.2 &amp; 1010.3. Also, the user's manual was found to have missing or improperly labeled information as required by 21 CFR 1020.30(h). | Radiation Control for Health and Safety Act | FONA will have upgraded copies of the manuals that will be made available on the firms website for free download and may also be mailed to each customer or end user should they request it. They will also provide users and distributers with corrected system labels to be placed on the x-ray control, beam limiting device, and tube housing assembly along with instructions for the proper application of these labels. They will provide users and distributers with an amended "Technical Data" page for insertion int the user's manual.Further questions please call (847) 459-3889. |  | 1960 | 1960 units | Worldwide Distribution- USA (nationwide) and Latin America. | Open, Classified |  |  | 1960 | Not_Computer
Z-1347-2013 | 64618 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PROGRESSIVE MEDICAL Laparoscopic Electrode PTFE Coated L Hook 33cm. |  |  | February 15, 2013 | 2013 | Unimed Surgical Products, Inc. | Firm is recalling due to concerns regarding changes that were made to the PTFE coating materials. | No Marketing Application | Unimed Surgical Products, Inc. issued a Medical Device Correction and Removal letter dated February 15, 2013 to its customers. Customers were instructed to quarantine any undistributed product, and notify its customers of the recall to retrieve any unused product. For further information, contact Unimed Surgical Products Inc. at 1-800-531-2627. |  | 120416 | 120,416 units | Distributed in the states of MO, TN, and PA. | Terminated | August 07, 2013 | 173 | 120416 | Not_Computer
Z-1355-2013 | 64931 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Progressive Medical PTFE Coated Disposable Electrosurgical Electrodes -Laparoscopic Electrode PTFE Coated Curved Blade 33cm Length, Product number P0019 Rx Only. Sterilized by Radiation. Manufactured for Progressive Medical, Inc., St. Louis, MO 63128. |  |  | April 12, 2013 | 2013 | Progressive Medical Inc | Product coatings require 510K approval. | Manufacturing material removal | Progressive Medical Inc. sent an Urgent Medical Device Recall letter dated April 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and segregate the affected units prior to return to Progressive Medical, Inc. If product was further distributed customers should notify their accounts of ther ecall and return the products. Customers were asked to complete and return the attached form by fax to 314-961-4535, even if they had no inventory remaining. Customers with questions were instructed to call 1-800-969-6331.For questions regarding this recall call 314-961-5786. |  | 228 | 38 boxes with 6 units each. | Nationwide Distribution including AR, AZ, CA, CO, FL, ID, IL, IN, KS, KY, LA, MA, MI, MN, MO, NE, OH, OK, OR, TN, TX, WA, and WI. | Terminated | May 07, 2014 | 390 | 6317 | Not_Computer
Z-1365-2013 | 65158 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | MEDITECH Bi-Directional interactive Communication Analyzer Interface; |  |  | May 07, 2013 | 2013 | Medical Information Technology, Inc. | Incorrect transmission of laboratory results. | Device Design | Beginning on May 7, 2013, MEDITECH distributed a notification to MAGIC and Client Server customers who have the MEDITECH Bi-Directional Interactive Communication Analyzer Interface in use with Vitros 3600, 4600, or 5600 Analyzer in TEST and/or LIVE. This notification was made via remailed task updates that can be printed by the customer. Task updates sent via e-mail are immediately transmitted to the customer.This notification informs customers of the potential patient safety issue that could result in reporting inaccurate qualitative results. Customers will receive a corrective change, in which the previous test's results will no longer be associated with the test not performed. Until customers receive the corrective change, they are instructed to monitor all No Result values from their analyzer and confirm that the result in MEDITECH matches. For further questions please call 781-821-3000, ext. 5742. |  | 56 | 56 devices | Worldwide Distribution - USA, including Puerto Rico and the country of Canada. | Terminated | October 01, 2014 | 512 | 56 | Not_Computer
Z-1366-2013 | 61315 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Pinnacle Hip Solutions ALTRX Polyethylene Acetabular Liner, Ref 1221-36-452 |  |  | March 07, 2012 | 2012 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. (DePuy) is issuing a voluntary recall of one lot of the AltrX" liner because the product was incorrectly labeled on the box. The box was labeled for a 36IDX52OD +4 Neutral Liner, but the product enclosed was a 36IDX56OD Neutral Liner. Three ( 3 ) complaints were received in January 2012. | Packaging process control | DePuy Orthopaedics sent an URGENT INFORMATION- MEDICAL DEVICE RECALL NOTICE dated March 7, 2012 via email to all affected distributors and customers. The letter identified the affected product, problem and actions to be taken. Distributors and customers are instructed to cease further distribution or use or of recalled product and to return them to DePuy. The sales representatives are expected to aid customer in returning of recalled product. Distributors and sales representatives are to return the recalled product through the normal DePuy Returns process. Reconciliation form should be completed and returned to your DePuy Sales Representative or faxed to 574-372-7567. For questions about recall information contact the Manager of Customer Quality 574-372-7333. |  | 20 | 20 liners | Worldwide Distribution - USA Nationwide including the state of VA and the country of Ireland. | Terminated | May 24, 2013 | 443 | 20 | Not_Computer
Z-1217-2013 | 65036 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | MyRay RXDC dental unit |  |  | August 10, 2012 | 2012 | CEFLA DENTAL GROUP | The RX DC dental x-ray unit was found to be noncompliant with 21CFR 1010,2, 1020.30(e) &amp; 1020.30(h). Specifically, 1. The tube housing did not contain a Certification label.2. The collimator did not contain an identification label.3. The user's manual did not contain leakage technique factors, tube housing cooling curves and rating charts, and technique factors which lead to operation at ma | Radiation Control for Health and Safety Act | The firm, CEFLA Dental Group will update the user's manual to include the information required by 21 CFR 1020.30(h). Future production will ensure that system labels required by 21 CFR 1010.2, and 1020.30(e) are properly applied including the 53 units manufactured awaiting shipment. The updated user's manual will be made available online for free download. Alternatively, system purchasers, subsequent transferees, dealers and distributors may elect to receive a hard copy of the updated user's manual at no charge. All corrections will be made at no cost to the purchaser, subsequent transferee, dealer, or distributor.The above actions will be completed within 4 weeks. For further questions please call 1-800 323-2690. |  | 173 | 173 shipped to US | FOREIGN: CANADA; DOMESTIC: GA, TX, IL, CA, FL, NY, MA, AL, NJ, SC, WA, LA, MD, WY, CT, NM, WI, CO, NC, VA, PA and MI. | Terminated | March 04, 2015 | 936 | 173 | Not_Computer
Z-1368-2013 | 64850 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CobasÂż 8000 modular analyzer series software version 03-01, 02, 03 and 04, impacting the cobas e 602 module only. |  |  | April 04, 2013 | 2013 | Roche Diagnostics Operations, Inc. | When using the cobas e 602 module, Diluent Multiassay may be used for autodilution of certain assays instead of Diluent Universal. This may cause incorrect results under specific circumstances if all of the following conditions are met simultaneously: More than a total of 2000 reagent kits or diluents have been registered on the cobas e 602 modules within one system core, and two different diluent | Software design | Roche sent an Urgent Medical Device Correction (UMDC) letter dated April 4, 2013 via UPS to all affected customers. The letter idenified the affected product, problem and instruction to be followed. Customers were instructed to consult with the physician or pathologist at their facility to determine specific clinical implications for patients. A future software version will correct this issue. Customers were asked to complete and fax back the UMDC attached fax form to 1-88-628-0730. Questions should be addresssed to Roche Diagnostics Technical Support 24 hours a day at 1-800-428-2336. |  | 72 | 72 units | USA Nationwide Distribution | Terminated | February 24, 2015 | 691 | 72 | Computer
Z-1367-2013 | 64718 | No | IFB | Cytocentrifuge. | Pathology | Pathology | 510(K) Exempt | SLIDERPREP PLUS SLIDE PREPARATION CENTRIFUGE. |  |  | March 13, 2013 | 2013 | Separation Technology, Inc. | Separation Technology, Inc. recalled SlidePrep Plus Cytology Centrifuges manufactured prior to January 1, 2013 for having the potential to start up unexpectedly when encountering an electrostatic discharge. | Device Design | On March 13, 2013, Separation Technology, Inc. sent their customers letters instructing them to complete and return the attached Acknowledgement &amp; Receipt form via fax to (407) 788-3677 and requesting the return of the SlidePrep Plus Cytology Centrifuges to STI for correction. For any questions regarding this recall, call STI Customer Service at (800) 777-6668. |  | 102 | 102 | Worldwide Distribution-USA (nationwide) including the states of PA, NC, IL, KS, and FL, and the countries Canada and Turkey. | Terminated | December 24, 2013 | 286 | 102 | Not_Computer
Z-1370-2013 | 64941 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Triathlon Femoral Distal Augment- Left; 1 single unit per package |  |  | March 25, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics received a report that a label on the external packaging (box) of a Triathalon Femoral Distal Augment 10mm- Size 5 Left depicted a glyph of an acetabular shell rather than a femoral distal augment. In addition the label was missing a reference to the "10mm" augment thickness which per the label print is required to appear below the size reference. | Software design | The firm, Stryker Orthopaedics, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated March 25, 2013 and Product Accountability Forms via Fed Ex to its Consignees/ customers. The letter described the product, problem and actions to be taken. The customers were instructed to complete and return the attached Notification Acknowledgment Form within 5 days via fax to 855-251-3635; contact the hospitals in your territory that have the affected product to arrange return of product; and return affected product to the attention of Regulatory Compliance Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430, attach flourescent orange PRODUCT REMEDIATION sticker to your return indicating the PR #RA2013-011 and mark outer box with the words "Product Recall".If you have any questions, feel free to contact Manager, Divisional Regulatory Compliance at (201) 972-2100. |  | 11 | 11 units | Nationwide distribution: USA including states of: HI, IN, MS, NC, NJ, OK, TX, and WI. | Terminated | May 06, 2015 | 772 | 11 | Not_Computer
Z-1378-2013 | 64588 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Femoral component, CR, cemented, #1, left2103-1310 |  |  | February 20, 2013 | 2013 | Orthopedic Alliance LLC | The recall of the Loaner Bin and Sterilization Trays for implants and instruments was initiated by Orthopedic Alliance because the directions for use and surgical techniques were not being provided to every physicianand for every surgery. | Labeling design | A recall letter dated 3/8/13 was sent to all customers who purchased loaner bins and sterilization trays for implants and instruments from Orthopedic Alliance. The letter informed the customers of the problem identifiedand the actions to be taken. Customers are instructed to call the office at the number (951) 304-9001 if they are requesting paper copies of any of the literature. |  | 136 | 136 units | Nationwide Distribution including the states of CA, MD, TX, NV, and WI. | Terminated | March 26, 2014 | 399 | 821 | Not_Computer
Z-1377-2013 | 64886 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | ***REF S-3D*** 3mm Diamond Ball. |  |  | March 04, 2011 | 2011 | The Anspach Effort, Inc. | Due to three batches of product may have diamond grit size larger than intended applied to the head of the cutting burr. | Labeling mix-ups | Anspach sent an URGENT: Medical Device Correction letter dated March 4, 2011 to its customers identifying the affected product and actions to be taken by the customers. Customers questions or concerns should be directed to Anspach Customer Support at (800) 327-6887. |  | 1827 | 1,827 burrs | Worldwide Distribution-USA (nationwide) including DC and the states of OH, TX, IL, CA, GA, KY, FL, WI, NC, VA, WI, MN, MD, KS, OK, PA, OR ,HI, AZ, WA, TN, AR, IA, MI, CO, LA, NJ, IN, WV, MO, AL, NY, VT, and SC and the countries of France, Canada, Great Britain, Mexico, Australia, Sweden and Netherlands. | Terminated | December 17, 2013 | 1019 | 1827 | Not_Computer
Z-1376-2013 | 65042 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | NEXGEN Complete Knee Solution Monoblock Tibial Drill with stop 10.7 mm. |  |  | March 27, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the Trabecular Metal" Tibia Stop Drill due to the potential of the outside diameter being oversized. As a result, there is a potential for the drill to not pass through the applicable drill guide. | Process control | Zimmer sent an "URGENT MEDICAL DEVICE RECALL LETTER" dated April 19, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 69 | 69 units | Nationwide Distribution. | Open, Classified |  |  | 69 | Not_Computer
Z-1375-2013 | 65031 | No | IQA | Powered environmental control system. | Physical Medicine | Physical Medicine | 510(K) Exempt | QuietCare. |  |  | April 04, 2013 | 2013 | Intel-GE Care Innovations LLC | Customers at 4 facilities reported receiving duplicate alerts using the QuietCare-Networked product. | Software design | Care Innovations contacted their customers by phone and e-mail regarding the potential problems and fixes on April 4, 17, and 23, 2013. The information provided stated the product, problem, and actions to be taken by the customers. Contact Technical Support at 855-885-2273 for questions regarding this issue. |  | 99 | 1722S; 99 Out of US | US and Great Britain | Terminated | June 04, 2013 | 61 | 145 | Not_Computer
Z-1400-2013 | 63367 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Synchron LX System(s) and UniCel DxC 600/800 System(s) Lactate Reagent, Part Number: A95550 |  |  | April 17, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter initiated the recall for the Lactate Assay (LACT) Lot M206209 as they received a customer complaint of error message (RXN RATE HI), i.e. no numerical values reported for samples containing Lactate concentration greater than 6.5 mmol/L. | Under Investigation by firm | Beckman Coulter sent an Urgent Product Correction letter dated April 17, 2013 to all affected customers. The letter identified the affected product, problem actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Beckman Coulter Customer Support Center (Call Center) on the Beckman Website at http://www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and Canada. |  | 8774 | 8,774 units | Worldwide Distribution - USA Nationwide and the countries of Canada, Australia, Belgium, Czech Republic, France, French Polynesia, Germany, Greece, Hong Kong, Ireland, Italy, Japan, Jordon, Kuwait, Lebanon, Malawi, Morocco, Namibia, Netherlands, oman, Panama, Paraguay, Qatar, Reunion, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom | Terminated | October 22, 2014 | 553 | 8774 | Not_Computer
Z-1395-2013 | 63204 | No | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthobiologics Aliquot Delivery System Manual Surgical Instrument - Aliquot Reamer, 8", Part Number 2110-0525 |  |  | November 10, 2012 | 2012 | Orthovita, Inc., dBA Stryker Orthobiologics. | There have been a few reported cases of the catheter breaking when removing the Aliquot catheter from the needle during vertebroplasty procedures. This has been attributed to bending of the catheter during the retraction from the needle. It was determined to be use errors. The correction is limited to a revision of the Directions for Use. | Labeling design | Stryker sent an Urgent Medical Device Correction letter dated November 10, 2012 to all affected customers. The letter identified the affected product, problem issue, potential hazards and actions to be taken. Customers were instructed to immediately review this notification and the attached IFU. Customers were asked to complete and sign the enclosed Business Reply Form and fax a copy to 1-484-323-8803. For questions contact Stryker Orthobiologics. |  | 5 | 5 | USA Nationwide Distribution in the states of AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KS, KY, LA, MD, ME, MI, MO, NC, ND, NH, NJ, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA, and WI. | Terminated | October 29, 2014 | 718 | 1413 | Not_Computer
Z-1408-2013 | 65150 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Merit Medical PerfOrma Anglographic Catheter 5F Multipack |  |  | May 02, 2013 | 2013 | Merit Medical Systems, Inc. | Merit Medical Systems is recalling various lots of Performa Cardiac Multipacks containing a JR4.0 catheter due to a slight variation in tip shape. | Process control | Merit sent an Urgent Product Recall Notice letter dated May 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were asked to identify, quarantine, discontinue using products, and contact their sales representative to arrange product return and replacement or credit. For questions call 1-801-208-4468. |  | 1465 | 1,465 | USA Nationwide Distribution. | Terminated | July 22, 2013 | 81 | 1465 | Not_Computer
Z-1403-2013 | 65067 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | PTS PANELS, Glucose Test Strips for use with CardioChek Brand Analyzers |  |  | April 04, 2013 | 2013 | Polymer Technology Systems, Inc. | Internal investigation demonstrated under-recovery for glucose when tested against a reference method. This could result in a low to moderate health risk. | Under Investigation by firm | Polymer Technology Systems sent a Field Correction Action letters dated April 4, 2013 to distributors and customers. The letter identified the affected product, problem and actions to be taken. The letter directed customers to discontinue use and dispose of remaining inventory of recalled product. Customers were asked to acknowledge and confirm recall directions were follwed by fax to 1-317-870-5608. |  | 7048 | 7048 total, 3499 in USA | Worldwide Distribution - USA Nationwide including Puerto Rico and countries of: UK, CHL, ENG, BEL, ITA, TKY, BEL, LIT, LTV, EGY, SPG, HOK, SWE, MEX, UAE, COR, ISR, MAL, AUS, POR, POL, SLN and STQ. | Terminated | April 16, 2014 | 377 | 7048 | Not_Computer
Z-1405-2013 | 65119 | No | DQR | Catheter cannula. | Cardiovascular | Cardiovascular | 510(k) | AngioDynamics Micro Introducer Kits under the following labeling: |  |  | April 30, 2013 | 2013 | Angiodynamics, Inc. | AngioDynamics Inc., is recalling Micro Introducer Kits and VenaCure EVLT NeverTouch Procedure Kits, because the potential exists that the micro Introducer (sheath/dilator) packaged within the kit could be the incorrect French size. | Labeling mix-ups | AngioDynamics sent an Urgent Medical Device Recall letter dated April 29, 2013, to all customers by Fed Ex. The letter identified the affected product, problem and actions to be taken. Customers were instructed to lmmediately remove the affected product from inventory and call AngioDynamics, Inc. Customer Service at 1-800-772-6446 to obtain a replacement or credit for your returned product and to obtain UPS Account Number to return the package via second day delivery. Customers are instructed to segregate and return all affected devices to: AngioDynamics 603 Queensbury Avenue Queensbury, NY 12804 Attn: Micro Introducer and VenaCure EVLT NeverTouch Procedure Kits Recall Coordinator. Write the RMA number on the shipping box (Provied on the Recall Verification Tracking Form). Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return); following the directions on the notification form and the Reply Verification Form. Send the Reply Verification form to AngioDynamics either via Fax or Emai. Fax Reply Verification Tracking Form to: Attn: Micro Introducer and VenaCure EVLT NeverTouch Procedure Kits Recall Coordinatory Fax number 1-518-798-1360 or Email Reply Verification Tracking Form to: rdenino@angiodynamics.com. |  | 806 | Domestic(1,760 units): 1)260 units/26 boxes, 2)310 units/31 boxes,3)260 units/26 boxes,4)880 units/88 boxes,5)60 units/6 boxes,Foreign:10 units/1 box | USA Nationwide Distribution including the states of: IL, GA, CA, NY OH, AZ, TX, NE, FL, PA, NC, WI, AL, IN, CO, MO, MT, NE, MA, VA, AK, IN, MI, NJ, MD, TN, CT, WA, MN | Completed |  |  | 1557 | Not_Computer
Z-1402-2013 | 65060 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integrity 3.0 |  |  | April 30, 2013 | 2013 | Elekta, Inc. | Unexpected movement during beam delivery may not be detected if there are incorrect static tolerances for Linac Geometric Axes loaded in Integrity R3.0. | Nonconforming Material/Component | Elekta sent an Important Field Safety Notice (IFSN) letter dated April 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are instructed to complete and return the FCO Action Notification Report to local Elekta Office or Representative. |  | 31 | 31 | USA Nationwide including the states of CA, FL, IL, MD, MA, NJ, NY, NC, ND, OH, OR, PA, TN, TX, VA, WA, DC and Puerto Rico | Terminated | July 26, 2017 | 1548 | 31 | Not_Computer
Z-1401-2013 | 65048 | No | JHI | Human chorionic gonadotropin (HCG) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere" hCG Combo Cassette (20/10 mIU/mL) Kit. |  |  | April 16, 2013 | 2013 | Alere San Diego, Inc. | The firm initiated recall of Alere" hCG Combo Cassette (20/10 mIU/mL) kit PN 92215, lots hCG1110133 and hCG1110135, because an investigation has indicated a portion of these Alere" hCG Combo Cassette (20/10 mIU/mL) PN 92215 kit lots were incorrectly packaged with the Alere" hCG Cassette (20 mIU/mL) PN 92210 Product Insert. The Alere" hCG Cassette (20 mIU/mL) Product Insert contains information for | Packaging | Alere initiated this recall by sending recall notification letters, titled "URGENT MEDICAL DEVICE RECALL", dated April 16, 2013 via via fax, email or direct mail. The letter informed customers of the product recalled, reason for recall, customer required action, and contact information, "Alere Technical Services at 866-216-0094 or by e-mail at lateral.flow.support@alere.com.". A verification form was attached to the notification letter. |  | 3975 | 3975 kits | United States Nationwide Distribution | Terminated | October 18, 2013 | 185 | 3975 | Not_Computer
Z-1404-2013 | 65082 | No | HKO | Ophthalmic refractometer. | Ophthalmic | Ophthalmic | 510(K) Exempt | Canon Full Auto Ref-Keratometer RK-F2, Catalog Number 6937B002AA. |  |  | December 05, 2012 | 2012 | Canon Inc. | Canon Inc. had received 17 claims from our foreignsales companies between October to November2012 that measurement value of RK-F2s refractivepower (SPH value) sometimes indicated abnormalvalue which occurs approximately 5% in probability. | Software design | Canon initiated recall by sending letters on November 22, 2012 to all affected customers. The letter identified the affected product, problem, actions to be taken and contact information. |  | 384 | 384 units | Worldwide Distribution - USA Nationwide and countries of: Australia, Japan, Korea, Singapore and United Kingdom. | Terminated | September 20, 2013 | 289 | 384 | Not_Computer
Z-1407-2013 | 65171 | No | GWF | Evoked response electrical stimulator. | Neurology | Neurology | 510(k) | Natus neurology Dantec KEYPOINT.NET 2.20. |  |  | April 09, 2013 | 2013 | Alpine Biomed Aps | Alpine Biomed ApS is recalling Natus neurology Dantec KEYPOINT.NET ver. 2.20 Software due to the presence of a software bug. | Software design | Alpine Biomed ApS sent a Recall Notification/ Field Safety Notice dated April 8, 2013 to all affected customers via e-mail and courier. Alpine Biomed ApS advised their customers to upgrade all installations of Keypoint.NET v2.20 to Keypoint.NET v2.21. Alpine BioMed ApS has recommended that their customers complete and return the verification form to QA/RA Engineer Michael Jacobsen at e-mail michael.jakobsen@natus.com. These actions and the return of the verification form to Alpine Biomed ApS should be completed within 3 months of receipt of the recall notifications. If any assistance is required to complete these actions, please contact Technical Support Engineer Michael Jensen via e-mail michael.jensen@natus.com, alternatively QA/RA Engineer Michael Jacobsen by e-mail michael.jakobsen@natus.com. |  | 871 | 871 licenses installed | Worldwide Distribution - US Nationwide and the countries of: Germany, France, Bulgaria, India, Colombia, Tunisia, Saudi Arabia, Finland, Denmark, Brasil, Italy, Estonia, South Korea, Russia, Greece, Thailand, Israel, Poland, Turkey, Slovakia, Portugal, Switzerland, Hungary, Great Britain, Netherlands, China, Indonesia, Australia, Sweden, Spain, Iran, Singapore, Norway, Austria and Czech Republic. | Terminated | January 06, 2014 | 272 | 871 | Computer
Z-1409-2013 | 64586 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | 45 mm Screw, Item #BG7045 |  |  | February 20, 2013 | 2013 | Spinal Solutions, LLC | The recall of the Loaner Bin and Sterilization Trays for implants and instruments was initiated by Spinal Solutions because the directions for use and surgical techniques were not being provided to every physician and for every surgery. | Labeling design | Spinal Solutions sent a recall letter dated March 8, 2013 to all customers who purchased loaner bins and sterilization trays for implants and instruments from Spinal Solutions. The letter identified the affected product, problem identified and actions to be taken. Customers are instructed to call the office at the number (951) 304-9001 if they are requesting paper copies of any of the literature. |  | 24 | 24 units | USA Nationwide Distribution in the states of MS: TX, NV, CA, MD, and WI. | Terminated | March 26, 2014 | 399 | 503 | Not_Computer
Z-1419-2013 | 65136 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Therapy Treatment Planning System, Model Numbers 2.0, 2.5, and 3.0; CLASSIFICATION NAME: System, Planning, Radiation Therapy Treatment. |  |  | February 15, 2013 | 2013 | RAYSEARCH LABORATORIES AB | Under some circumstances, imported CT, MR and PET images are offset from their true positions by one pixel in one or two directions. Users must be aware of this information to avoid incorrect dose calculations during treatment planning. | Software change control | RaySearch Laboratories sent an Urgent Field Safety Notice, Medical Device Correction letter dated January 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to educate their planning staff and all users about the recommended workarounds contained in the letter. Customers were also asked to inspect their product and identify all installed units that were infected and then complete the attached Reply Form and e-mail or scan a copy of the completed form to the firm. Customers were instructed to pass along this information to all those who need to be aware of this Medical Device Correction Customers with questions were instructed to contact the RaySearch support email at support@raysearchlabs.For questions regarding this recall call +46 8 545 06130. |  | 97 | Domestic: 70 units; Foreign: 27 units | Worldwide Distribution - USA including Connecticut, Florida, Massachusetts, Missouri, New York, Ohio, and foreign distribution to(Canada, Italy, Australia, Germany, Netherlands, and Switzerland. | Open, Classified |  |  | 97 | Not_Computer
Z-1421-2013 | 65165 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Leucadia Autolok" Polyaxial Pedicle Screw System. |  |  | May 06, 2013 | 2013 | Alphatec Spine, Inc. | Alphatec Spine initiated this medical device recall for the following reason: due to tolerancing, the flange of the Phygen Leucadia" Autolok set screw may interfere with the top of the screw head prior to fully locking the rod. If this condition occurs, the polyaxialscrew head and rod may fail to provide a rigid construct under loading. To date there have beenno reports of patient injury, adve | Device Design | Alphatec Spine initiated recall by contacting customers via telephone and/or email on 05/06/2013. An Urgent Medical Device Recall Notification letter along with a response form were sent via FedEx to all affected accounts. The letter identified the product, problem, and actions to be taken by the customers. |  | 1814 | 1,814 | Nationwide Distribution. | Terminated | September 25, 2014 | 507 | 1814 | Not_Computer
Z-1372-2013 | 65070 | No | CCQ | Medicinal nonventilatory nebulizer (atomizer). | Anesthesiology | Anesthesiology | 510(K) Exempt | AsthmaNefrin Starter Kit contains ten (10) AsthmaNefrin vials and one (1) EZ Breathe Atomizer. |  |  | April 30, 2013 | 2013 | Nephron Pharmaceuticals Corp. | Possible dislodgement of the "Plate A", if this instance occurs, it presents a potential choking hazard. | Nonconforming Material/Component | The firm, Nephron Pharmaceuticals Corporation, notified its consignees on April 30, 2013, and May 3, 2013, of the problem via a recall notice and a Press Release, respectively. The firm directed consignees to discontinue use of any EZ Breathe atomizer units contained in the Asthmanefrin Starter Kits with the affected lot numbers and also those that were sold individually. Customers were directed to return all affected product. |  | 241538 | 241,538 kits | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AL, AZ, CA, CO, CT, FL, IL, IN, IO, GA, KS, KY, LA, MD, MN, MS, MO, NK, NV, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN , TX, VT, VA, WA, and WV, and the countries of Guatemala, Poland, Panama, and United Kingdom. | Terminated | June 28, 2016 | 1155 | 242892 | Not_Computer
Z-1423-2013 | 65089 | No | JTO | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BBL DrySlide Oxidase Kit, catalog number 231746 |  |  | April 12, 0013 | 13 | Becton Dickinson &amp; Co. | In-vitro diagnostic agents were placed into incorrect packaging for distribution to microbiological laboratories. | Mixed-up of materials/components | BD Diagnostic Systems sent an Urgent Product Recall" letter dated April 2013 to all affected consignees and end users. The letter identified the affected product, problem and actions to be taken. The notice instructed consignees and end users to discontinue use of the products, discard remaining products, and return a completed response form. Consignees were requested to provide contact information for end users who received recalled product in distribution, to allow for recall notice. The letter included instructions to report injuries and/or illnesses to Medwatch, and address recall inquiries to BD Technical Services at 1-800-638-8663. |  | 610 | 610 total test kits (PYR and Oxidase) | Worldwide Distribution - USA Nationwide and tthe countries of Australia, Canada, Europe, Hong Kong, Mexico, New Zealand, Singapore, and Taiwan. | Terminated | September 28, 2013 | 730654 | 610 | Not_Computer
Z-1425-2013 | 65073 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Tip Cover Accessory, 8 MM, MCS, box of 10 |  |  | September 23, 2011 | 2011 | Intuitive Surgical, Inc. | Clarification of existing labeling and related cautions for the tip cover accessory of the da Vinci Surgical System. | Device Design | Intuitive Surgical sent an Important Product Notification letters dated October 10, 2011 to all affected customers. A second letter was sent to those that had not responded on December 4, 2011. The letters identified the affected product, problem, and actions to be taken, The letter included recommendations and suggestions for proper use and prevention of damage to the Tip cover during surgery. Also attachments include recommendations and further information for the following areas: avoid instrument collisions; carefully install tip cover; straighten wrist prior to removal; inspect cannula prior to use and do not exceed maximum monopolar cautery settings.Attached is a list of validated ESU settings and suggestions and instruction on being aware of critical anatomy in contact with the instrument and to change the tip cover in cases with extended cautery use.Questions may be referred to Customer service at 1-800-876-1310 in the US and at + 41 21 821 2020 or +800-0821-2010. |  | 500000 | 50,000 boxes of 10 tip covers | Worldwide distribution - USA Nationwide and the countries of Australia, Austria, Belgium, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Malaysia, Mexico, Netherlands, New Zealand, Norway, Pakistan, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Slovenia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom. | Terminated | May 30, 2013 | 615 | 500000 | Not_Computer
Z-1426-2013 | 65072 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Thyroidectomy Indication for the da Vinci Surgical Systems; |  |  | October 13, 2011 | 2011 | Intuitive Surgical, Inc. | Promotional literature for use of the da Vinci system for thyroidectomy, which had not been approved with a change to the 510(k). | No Marketing Application | Intuitive sent a notification letter on October 17, 2011 by email to all consigness whose email address were on file. Also a copy of the letter was sent by registered mail on February 8, 2012 to all surgeons whose addresses were on file. The letters informed consignees that Intuitive Surgical were in the process of reviewing labeling for the da Vinci System in regard to use for thyroidectomy procedures. During this period, Intuitive will not provide or facilitate any promotional, training, case observations, proctoring or in case procedural support for use of the da Vinci System in conjunction with thyroidectomy procedures. The letters recommend consulting the institution policy regarding use of approved or cleared medical devices for procedures that are not specified in the labeled indication for use. Questions may be directed to customer service at 1-877-408-3872. |  | 2423 | 2,423 consignee | USA Nationwide distribution | Terminated | May 30, 2013 | 595 | 2423 | Not_Computer
Z-1424-2013 | 65094 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | User Manual addendum, TORS for da Vinci Surgical System IS 1200, IS 2000 and IS 3000. |  |  | January 24, 2013 | 2013 | Intuitive Surgical, Inc. | Change to User Manual to remove one indication for use and amend warnings to emphasize restrictions for use. | Labeling False and Misleading | Intuitive sent an Important Medical Device Correction letters dated January 24, 2013 via Federal Express to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate and discard the existing TORS addendum, place the new TORS addendum included with this letter with your System User Manual, and complete and return the attached Acknowledgement Form to Intuitive Surgical. For questions call Intuitive Surgical Customer Service at 1-800-876-1310. |  | 1374 | 1,374 | US Distribution Only | Terminated | May 30, 2013 | 126 | 1374 | Not_Computer
Z-1428-2013 | 65083 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Flush Port (used in almost all da Vinci instruments) |  |  | September 19, 2011 | 2011 | Intuitive Surgical, Inc. | Clarification of existing labeling information and operating procedures. | Labeling False and Misleading | Intuitive Surgical sent an "IMPORTANT PRODUCT NOTIFICATION" letter dated October 17, 2011 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Customer Service at 1-800-876-1310 for questions regarding this notice. |  |  |  | Worldwide Distribution- USA Nationwide and countries of: Argentina, Australia,Canada, Chile, China, Czech Republic, Egypt, India, Israel, Italy,Korea-Republic of, Poland, Qatar, Romania, Saudi Arabia, Singapore, Spain, Thailand, Turkey, and Venezuela. | Terminated | May 30, 2013 | 619 |  | Not_Computer
Z-1430-2013 | 65118 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU2700 Clinical Chemistry Analyzer |  |  | April 26, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling AU2700/AU2700 Plus &amp; AU5400 Clinical Chemistry Analyzers becuase it was determined that the tubing/joint connections associated with these MODs may leak concentrated detergent. Beckman Coulter has received complaints that the joint connecting the tubing appears to disintegrate over time, causing a leak. | Component design/selection | Beckman Coulter sent an Urgent Product Correction letter on April 26, 2013 along with a response form. The letter notified customers of the product, problem, and actions to be taken by the customers. Contact the Call Center at 800-854-3633 for questions regarding this notice. |  | 108 | 108 units | Worldwide Distribution-USA Nationwide and countries of: AU, CA, DE, ES, GB, IT, PT, and JP. | Terminated | August 04, 2014 | 465 | 108 | Not_Computer
Z-1431-2013 | 65194 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 8mm Precise Bipolar Forceps used in conjunction with the da Vinci Surgical System, Model IS1200. |  |  | May 08, 2013 | 2013 | Intuitive Surgical, Inc. | Reports of dislodging of the Cautery Plug Insert from the Cautery Plug Riser, preventing a bipolar cautery cable from being properly connected which results in loss of cautery function of Bipolar Instruments. | Device Design | Intuitive Surgical sent an "Urgent Medical Device Correction" letter dated May 8, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Customer Service at 800-876-1310 for questions regarding this notice. |  | 697 | 697 units - all products | Worldwide Distribution-USA (nationwide) including the states of AL, CA, CO, FL, IL, KY, MI, MS, NJ, NY, PA, TX, UT, VA, WA, and WI, and the countries of Austria, Australia, Belgium, Canada, Czech Republic, Denmark, France, Germany, Italy, Netherlands, Norway, Qatar, Romania, Spain, Sweden, Switzerland, Taiwan, United Kingdom, Saudi Arabia, Singapore, and Venezuela. | Terminated | February 18, 2014 | 286 | 697 | Not_Computer
Z-1437-2013 | 49033 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Fisher &amp; Paykel Healthcare ICON CPAP (continuous positive airway pressure). |  |  | April 22, 2013 | 2013 | Fisher &amp; Paykel Healthcare, Ltd. | Fisher and Paykel Healthcare is recalling the ICON CPAP units because there is a potential power issue on certain ICON CPAP. These ICON CPAP units could potentially become inoperable, which would result in a lack of treatment to the patient's obstructive sleep apnea symptoms until a replacement unit is obtained. | Nonconforming Material/Component | The firm, Fisher &amp; Paykel Healthcare Limited, sent a "Medical Device Recall" letter dated April 22, 2013 to all their customers who purchased the Fisher &amp; Paykel Healthcare ICON CPAP Replacement. The recall letter described the product, notified the customers of the problem identified and the actions to be taken. Customers are instructed to locate and verify if any of the affected products are in stock; check the part and lot numbers of your device; complete and return the enclosed Product Recall Response form, even if you have not received or distributed any affected products via Email: FA-2013-001@fphcare.com; Fax: (949) 453-4001; or Post: Fisher &amp; Paykel Healthcare,15365 Barranca Parkway, Irvine, CA 92618; transfer this notice immediately to all those who need to be aware and if affected product has been distributed to any other organizations or distributors, notify them regarding the recall within 10 business days upon receipt of this letter; prepare any products to be returned to Fisher &amp; Paykel Healthcare and contact the field action coordinator via email at FA-2013-001@fphcare.com or directly at 1-800-792-3912 ext. 1426 to arrange for a return shipping label to be sent to you. If you have any questions related to this recall, contact the field action coordinator via email at FA-2013-001@fphcare.com or directly at (800) 792-3912 ext. 1426. |  | 1407 | 1,407 units | Worldwide distribution: USA (nationwide) and countries of: Canada, Japan, and Taiwan. | Terminated | October 18, 2013 | 179 | 1407 | Not_Computer
Z-1438-2013 | 65120 | Yes | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | adult defibrillation, cardioversion, pacing, monitoring Disposable Electrodes, Model P-211-M1 and P-214-M1. |  |  | April 30, 2013 | 2013 | TZ Medical Inc. | TZ Medical Inc is recalling its P-211-M1 and P-214-M1 Medtronic / Physio Control Compatible transparent defibrillation electrodes because it may contain a defective white wire on the connector assembly.If a failure were to occur, the end user would be able to identify the failure by inspecting the white wire 1/4 inch from the Gray Medtronic / Physio adapter. The consignees will notice discolor | Process control | TZ Medical Inc. started notifying their customers by telephone on April 30, 2013. The firm followed up with an Urgent: Medical Device Recall notification letter dated May 14, 2013, to their consignees. TZ Medical Inc., Representatives have will contact all affected customers and arrange for the replacement of the effected product in stock with inspected product. For any questions about this recall, please fax to 503-639-0239. |  | 10510 | US: 10,410 pieces; Outside US: 100 pieces | Worldwide distribution - USA (nationwide) and Chile. | Terminated | September 05, 2013 | 128 | 10510 | Not_Computer
Z-1435-2013 | 65085 | Yes | LDF | Cardiovascular permanent or temporary pacemaker electrode. | Cardiovascular | Cardiovascular | 510(k) | Vermed, Inc Radiotransparent Defibrillation Electrodes |  |  | April 26, 2013 | 2013 | Heart Sync, Inc | On 4/18/13 HEARTSYNC became aware of a recall from their wire/connector sub-assembly supplier which is used in the C100-Physio Electrodes. There were reports of equipment damage due to misalignment from tooling installed in August 2012. The result of this damage could render the device incapable of delivering the appropriate energy or shock to the patient. | Process design | HeartSync sent an URGENT MEDICAL DEVICE RECALL notification letter dated April 26, 2013, to all consignees of Heart Sync for C100-Physio Radiotransparent Defrillation Electrodes. On April 29, 2013, an URGENT MEDICAL DEVICE RECALL notification was sent to Vermed Inc for the D304 Defrillation Electrodes . The letters identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return the affected product to the firm for full credit. Customers were asked to complete the form and send back via fax to 734-213-5640 or call 734-213-5530.For questions regarding this recall call 734-213-5530.. |  | 17180 | 17180 units for all part numbers | Worldwide Distribution - USA including AL, GA, MI, MN, NC, ND, OH, PA, SD, VA, VT, OH, TX, MA, and Internationally to Korea. | Terminated | April 16, 2014 | 355 | 17180 | Not_Computer
Z-1440-2013 | 65109 | No | EPT | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zeiss brand OPMI Pentero C, Surgical microscope and accessories, Model: OPMI Pentero C, Product Numbers: 302583-9001-000 (PENTERO C), and 305961-0000-00 (Ceiling Mount for PENTERO C), |  |  | April 18, 2013 | 2013 | Carl Zeiss Meditec, Inc. | Carl Zeiss Meditec is conducting a field corrective action for its OPMI Pentero C surgical microscope system. During a routine inspection, one microscope installation was found to have 3 (three) broken screws out of 6 (six) in the lower section of the ceiling mount. | Under Investigation by firm | Carl Zeiss Meditec, Inc sent a Customer Notification letter to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that firm's sales representative will be scheduling an on-site visit, informing them of the affected product and providing instructions on the recall. For question call 925-580-5371. |  | 6 | 6 units | US Nationwide Distribution. | Terminated | June 26, 2013 | 69 | 6 | Not_Computer
Z-1441-2013 | 65191 | No | HLI | Ophthalmoscope. | Ophthalmic | Ophthalmic | 510(K) Exempt | Optovue iVue with Normative Database (NDB) with Software Version 3.0 and 3.1. |  |  | May 14, 2013 | 2013 | Optovue, Inc. | Colors from the NDB comparison for Ganglion Cell Complex thickness in the singular OU report in iVue version 3.0 (US) and 3.1(international) are different from those in the individual iWellness ID and OS reports. | Software design | Optovue sent a Field Correction Notice letter dated May 16, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 329 | ~329 users | Distributed Nationwide and in Canada. | Terminated | October 24, 2013 | 163 | 329 | Computer
Z-1442-2013 | 65216 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 8 mm Monopolar Curved Scissors (a.k.a. Hot Shears) |  |  | May 08, 2013 | 2013 | Intuitive Surgical, Inc. | Intuitive Surgical has identified a potential for some units of the Monopolar curved Scissors version 09 and 10 instruments to develop very small cracks near the distal end. This may create a pathway for electrosurgical energy to leak to tissue and potentially cause thermal injury. | Nonconforming Material/Component | Intuitive Surgical sent an Urgent Medical Device Recall letter dated May16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that they did not need to quarantine or return the affected device at that time. a Intuitive Surgical representative would contact them to arrange for replacement of the affected product once they became available. In the interim customers were instructed to follow the precautions and warnings outlined in Attachment A in the letter as well as the Instructions for Use. Customers were also asked to ensure all affected personnel were fully informed of the contents of the Recall Notification. The letter should be forwarded to other surgeons and/or managers within their facility. Customers were also asked to complete and return the attached Acknowledgement Form. For further questions customers were instructed to contact Customer Service:North and South America: 800-876-1310 Option 3Japan: 0120-56-5635 or 003-5575-1362Korea: 02-3271-3200Europe, Middle East, Asia and Africa +800 0821 2020 or +41 21 821 2020For questions regarding this recall call 408-523-2244. |  | 132441 | ~132,441 instruments | Worldwide distribution - USA (nationwide) Foreign Distribution to Argentina, Austria, Australia, Belgium, Brazil, Bulgaria, Canada, China, Chile, Czech Republic, Denmark, Finland, France, Germany, Cyprus- Greece, Egypt, Guadeloupe, Israel, Ireland, India, Indonesia, Italy, Japan, Kuwait, Lebanon, Malaysia, Mexico, Monaco, Netherlands, New Zealand, Norway, Pakistan, Philippines, Portugal, Poland, Qatar, Russia, Romania, Slovenia, South Korea, Spain, Sweden, Switzerland, Thailand, Taiwan, Turkey, United Kingdom, Saudi Arabia, Singapore, South Korea, and Venezuela. | Terminated | May 01, 2014 | 358 | 132441 | Computer
Z-1439-2013 | 65255 | No | LRC | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | GYRUS ACMI, MYRINGOTOMY KNIFE DISPOSABLE, SICKLE BLADE FLAT HANDLE, REF 130715, Rx only, STERILE EO, GYRUS ACMI, INC., 136 TURNPIKE ROAD, SOUTHBOROUGH, MA 01722-2104 |  |  | May 24, 2013 | 2013 | Gyrus Acmi, Incorporated | lack of sterility assurance | Employee error | Cyrus ACMI notified their direct accounts of the recall by Priority Mail with delivery confirmation on May 24, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and immediately cease any further use of any affected product they may have, remove it from inventory, and quarantine it until they ship it back to the firm. For questions customers were instructed to contact their Olympus customer service representative at 1-800-848-9024 to obtain a Returned Goods Authorization so that they may return the product with no charge to them. Customers were asked to note on the enclosed questionnaire that they have received this information and fax it to 484-896-7128.For questions regarding this recall call 508-804-2739. |  | 600 | 100 boxes of 6 units | Worldwide Distribution - USA (nationwide) and countries of: Canada and United Kingdom. | Terminated | May 30, 2014 | 371 | 600 | Not_Computer
Z-1369-2013 | 64803 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Shimadzu Corporation - Mobile DaRt Evolution |  |  | March 18, 2013 | 2013 | Shimadzu Medical Systems | The recall was initiated because Shimadzu Corporation has identified a potential issue in certain manufacturing lots of the Shimadzu MobileDaRt Evolution in combination with CXDI-70C, CXDI-55C, and CXDI-55G. As a result of Shimadzu's investigation, there is a potential risk for screws becoming loose inside the Shimadzu Mobile X-ray systems. | Radiation Control for Health and Safety Act | Shimadzu sent a recall letter dated March 18, 2013, to all customers who purchased the Shimadzu MobileDart Evolution in combination with CXDI-70C, CXDI-55C, or CXDI-55G. The letter informs the customers of the problem identified and the actions to be taken. Customers are instructed to follow the instructions of the recall notice step by step. Customers with questions are instructed to contact their local Shimadzu representatives or contact Shimadzu National Technical Support at (800) 228-1429.A Field Safety Notice dated March 18, 2013, was sent to Service Managers in regards to the Shimadzu Mobile DaRt Evolution in combination with CXDI-70C, CXDI-55C, or CXDI-55G. The Safety Notice informs the Service managers of the problems identified and the actions to be taken. |  | 200 | 200 units | US Nationwide Distribution | Terminated | February 13, 2014 | 332 | 200 | Not_Computer
Z-1443-2013 | 65164 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Canister: CO2 canister with GE part numbers 1407-3200-000, 1407-7004-000, or M1084850. |  |  | May 14, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the reusable CO2 absorbent canister accessory used with Aespire, Aespire View, Avance, Avance CS2, Aisys, ADU, 9100, M900, M904, 9300 systems. Reusable CO2 absorbent canisters accessories may not seal properly. | Component change control | GE Healthcare sent an "URGENT MEDICAL DEVICE CORRECTION" letter April 30, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. A Fax back form was included with the letter requesting the consignees to complete and return. |  | 800 | 800 | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, CT, ID, MK, MI, NJ, NY, OH, TN, TX, and VA, and the countries of INDIA, AUSTRALIA, BAHRAIN, BANGLADESH, BOLIVIA, BOSNIA &amp; HERZEGOWINA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, HONG KONG, HUNGARY, INDONESIA, IRAQ,ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA,LEBANON, MALAYSIA, MEXICO, NETHERLAND,NEW ZEALAND, OMAN, PARAGUAY, PERU., PHILIPPINES, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, THAILAND, TURKEY, UNITED ARAB EMIRATES , UNITED KINGDOM, and VIET NAM. | Terminated | March 12, 2014 | 302 | 800 | Not_Computer
Z-1446-2013 | 64968 | cassettes, tissue | IDZ | Tissue processing equipment. | Pathology | Pathology | 510(K) Exempt | LPC Fine Mesh Cassette for the Leica LPC Printer. Part No: 14060546848; 14060546846; 14060546845; 14060546933; 14060546932; 14060546852; 14060546850; 14060546847. The intended use of this device is for storage of specimens/tissues during processing. |  |  | April 04, 2013 | 2013 | Leica Microsystems, Inc. | An LPC Fine Mesh Cassette failed to stay closed during processing operations. | Device Design | A Medical Device Correction letter dated April 17, 2013 was sent to customers via USPS 1st class mail. The product issue was described and recommended actions were provided. Customers were requested to acknowledge receipt of the letter by fax-back. Customer questions were directed to (847) 405-5413. |  | 20160 | 20,160 units | Worldwide distribution: US (nationwide) to states of AZ and KY; and the country of New Zealand. | Terminated | July 21, 2017 | 1569 | 20160 | Not_Computer
Z-1444-2013 | 64826 | No | BSO | Anesthesia conduction catheter. | Anesthesiology | Anesthesiology | 510(k) | I-Flow ON-Q Pain Relief System, T-Block Needle/Catheter Tray, Model #TBTB1089T |  |  | June 01, 2012 | 2012 | I-Flow LLC | I-Flow initiated two voluntary product removals of the ON-Q Pain Relief System: T-Block Needle/Catheter Tray and T-Block Needle/Catheter Kit because they may have potentially malfunctioned. | Use error | IFlow sent the customer notification letter on May 29, 2013, to 28 customers via FedEx along with a copy of the Service Guide to inform customers of the issue that can cause the Sting Ray to disconnect from the Catheter and how to prevent this problem from occuring. Customers are informed that in certain circumstances, ultrasound gel on the catheters significantly reduces the amount of force required to pull the catheter out of the Stingray connector. The In Service Guide clarifies that the method of catheter securement. It also instructs that wiping off any excess ultrasound gel from the catheter prior to inserting it into the Stingray connector will help with connecting the catheter to the Stingray connector. Customers are instructed that there is no action required for the I Flow ON Q Pain Relief System and ON Q T Bloc. The recall letter is only to provide further instructions on successful set up of the SPIROL catheter with the Stingray Connector included in the ON-Q T-bloc trays. Customers with any questions are instructed to call IFlow Customer service at (800) 448-3569 or IFlow 24 hr Tech support at (800) 444-2728.On June 1, 2012, an I-Flow sales representative removed the catheters and connectors from the one consignee/customers.For questions regarding this recall call 949-206-2700. |  | 450 | 450 units | Nationwide distribution: USA including states of: CA, CT, FL, IL, IN, KY, NY, OH, OR, TN, TX, VA and WI. | Terminated | June 25, 2013 | 389 | 525 | Not_Computer
Z-1420-2013 | 64631 | No | ETN | Surgical nerve stimulator/locator. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | REF 8229705 x1 NIM TriVantage EMG Endotracheal Tube 5.0 mm I.D. x 6.5 mm O.D. Medtronic Xomed, Inc., 6743 Southpoint Drive North, Jacksonville, FL 32216-0980 USA STERILE Rx Only |  |  | March 04, 2013 | 2013 | Medtronic Xomed, Inc. | In March 2013, Medtronic issued a recall of the NIM TriVantage EMG Endotracheal Tube due to cuff leakage in the intraoperative cuff deflation which then requires re-inflation or replacement of the deflated tube in order to ensure the continued ventilation of the patient. | Use error | The firm, Medtronic Xomed, Inc. sent an "URGENT Product Recall Notification" letter dated March 14, 2013, to its U.S. and International customers. The letter described the product, problem and actions to be taken. The customers were instructed to return any affected devices still in their possession per the following procedure: 1. Check inventories for the affected products listed on the enclosed checklist. 2. Fill-in the "quantity on-hand" column on the checklist. 3. Fax checklist to Medtronic ENT at 1-904-296-2386. 4. If you have any of the subject product, contact Medtronic ENT Customer Service at 1-800-874-5797 to arrange for returns and replacement or credit. 5.When returning products, clearly mark the outside of the container Returned Goods Authorization (RGA) number. If you have any questions regarding the subject action or the content of the letter, contact Gabriela Anchondo at 904-279-7550. |  | 2591 | 2,591 units | Worldwide distribution: AL, AR, AZ, CA, CO, FL, GA, HI, IL, IN, KS, KY, MA, MD, MI, MN, MT, NC, NH, NJ, NY, OH, PA, RI, SC, TX, UT, VA, and WA; and countries of : Australia, Canada, Germany, Italy, Norway, Poland and United Kingdom. | Terminated | January 30, 2014 | 332 | 2591 | Not_Computer
Z-1451-2013 | 65160 | No | FXX | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Comfort brand Ear-Loop Face Mask, Item #: 3100, Model Numbers: # 3100B, # 3100P, # 3100G, # 3100Y, # 3100W; Product is manufactured and distributed by Maytex Corporation, CA |  |  | May 15, 2013 | 2013 | Maytex Corp | Recent Lab testing conducted in March 2013 indicates that the product does not meet the standard of "Differential Pressure of only 1.6 Delta P (mm H2O/cm2)" as indicated on the box/package of the product; and also product manufactured between 2008 - 2012 cannot be guaranteed as Fluid Resistant. | Labeling False and Misleading | The firm, Maytex Corporation, sent a "MEDICAL DEVICE SAFETY NOTICE" dated May 15, 2013, to all its customers. The notice described the product, problem and actions to be taken. The customers were instructed to immediately contact your medical professional for treatment and contact Maytex Corporation to file a complaint should any illness or injury occur; discontinue use of the product if you detect any potential health hazards; if this product has been further sold or distributed, notify your customers to the retail level and send copies of safety notice to customer/users; and complete and promptly return the enclosed Medical Device Safety Notice Return Response Acknowledgement and Receipt form via mail to: Maytex Corporation 23521 Foley Street, Hayward CA 94545; fax at; 510-786-0209; email: www.maytex18@aol.com or by delivery to Maytex Corporation.Maytex Corporation has temporarily taken down their Website at www.maytexcorp.com to correct the wrong product specifications as described in the notice.If you have any questions, contact Maytex Corporation at 1-800-462-9839 or email: maytex18@aol.com Monday through Friday, 8:30 AM to 5:00 PM Pacific Time. |  | 43915 | 43,915 Cases | Worldwide distribution: USA (nationwide) including states of: AL, AZ, CA, CO, FL, GA, HI, ID, IL, IN, MA, MD, ME, MI, MN, MO, MT, NE, NH, NJ, NM, NY, OK, OR, PA, PR, TN, TX, UT, WA, and WI; and country of: Mexico. | Terminated | October 29, 2013 | 167 | 43915 | Not_Computer
Z-1477-2013 | 64164 | No | GQO | Herpes simplex virus serological assays. | Microbiology | Microbiology | 510(k) | Plastic bottle containing 10.5 mL of IgG Conjugate |  |  | January 23, 2013 | 2013 | Bion Enterprises Ltd | Affected products have a labeled expiration date of one year too long. | Error in labeling | MBL Bion sent an Urgent Product Recall letter dated January 28, 2013 to all affected customers via email. The letter identified the affected product, problem and actions to be taken. Customers were requested to destroy the affected product and to notify Bion of the number of units consumed and destroyed. For questions call 847-544-5044. |  | 207 | 207 units | Worldwide Distribution - United States (nationwide) in the states of IN, SC and the countries of Canada (Toronto); South Africa (Randburg) and Brazil (Rio de Janeiro) | Terminated | May 05, 2015 | 832 | 912 | Not_Computer
Z-1485-2013 | 65258 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer Air Dermatome II, REF 00-8851-001-01 |  |  | May 07, 2013 | 2013 | Zimmer Surgical Inc | This removal is in response to complaints alleging that the Zimmer Air Dermatomes and Zimmer Air Dermatome II products would either not operate oroperate intermittently. Zimmer's investigation determined that the planetary gear teeth were broken. This may cause a delay or postponement of the surgery or result in a suboptimal graft retrieval with potential of requiring an additional donor site. | Process control | The firm, Zimmer, sent an "URGENT: DEVICE REMOVAL" notice dated May 7, 2013, to all consignees/customers via FED Ex and Distributors via E-mail. Consignees who had devices repaired received telephone calls as well as the mailed Recall Notice. The Notice identified the devices being recalled; possible adverse events and directions to immediately stop use of and remove affected product, and obtain an RGA (Return Goods Authorization) number by calling (800)348-2759 or email information to dover.repairs@zimmer.com . to return the devices for repair to:Zimmer Surgical, Attn: QA/RA Dept.- Recall, 200 West Ohio Avenue Dover, Ohio 44622 USA; and to complete and return the Response Form-Immediate Response Required via email to CorporateQualitv.PostMarket@zimmer.com. If you have further distributed affected product(s), please provide the customer's information on the Business Response Form to Zimmer.If you have any questions regarding this action, please email us at CorporateQuality. PostMarket@zimmer.com or call 1-330-364-0989. |  | 13 | 13 units | Worldwide distribution: USA Nationwide including Puerto Rico; UAE, TAIWAN, SWEDEN, THAILAND, KOREA, SINGAPORE, HONG KONG CHINA, CHILE, CANADA, AUSTRALIA, JAPAN, ENGLAND, GERMANY, and EGYPT. | Terminated | December 15, 2014 | 587 | 223 | Not_Computer
Z-1466-2013 | 65257 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Hill-Rom 1000 Bed, Product No. P1160. The Hill-Rom 1000 Bed is intended for low to moderate acuity patient in the medical surgical area of the hospital. |  |  | April 15, 2013 | 2013 | Hill-Rom, Inc. | Complaints have been reported for brakes disengaging when the bed is subjected to a hard jolt or force to the frame. | Process control | On 4/16/2013 an URGENT MEDICAL DEVICE CORRECTION notification letter was sent to all consignees. The firm plans a visit to all facilities to correct each bed. Call Hill-Rom Technical Support at 800-445-3720 if you have any questions. |  | 8737 | 8,737 devices | Worldwide distribution: US (Nationwide) and countries of: Argentina , Bahamas, Bahrain, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Ecuador, Egypt, Ethiopia, France, Ghana, Guam, Hong Kong, India, Indonesia, Japan, Jordan, Kuwait, Lebanon, Libya, Malaysia, Mexico, N. Mariana Islands, Pakistan, panama, Paraguay, Puerto Rico, Qatar, Saudi Arabia, Singapore, South Africa, South Korea, Taiwan, Thailand, United Arab Emirates, United States, US Virgin Islands, Venezuela and Vietnam. | Open, Classified |  |  | 8737 | Not_Computer
Z-1476-2013 | 64550 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | SlingBar Standard. |  |  | April 30, 2013 | 2013 | Hill-Rom, Inc. | Retrospective review found 4 sling bars that may present a potential for impalement if a patient, caregiver or bystander were to fall onto the hooks based on a simulated test method. In effort to eliminate any risk of injury, Hill-Romis requesting products be removed from field use andreplaced with an updated design. | Device Design | On 4/30/13 an URGENT MEDICAL DEVICE RECALL CORRECTION was issued to all consignees detailing the hazards and actions that are needed. Included with this notification is a Response Form/Receipt and a Replacement Guide, complete and return the form to Hill-Rom as soon as possible. Hill-Rom will send you the applicable number of sling bars to you at no cost. After you receive the new sling bars, we request that you discard the old ones. |  | 726 | 726 | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, FL, IL, IN, KS, LA, MA, MI, MN, MO, NC, NH, NJ, NV, OH, OK, PA, SC, TN, UT, WA, and WI, and the countries of Canada, Australia, Austria, Belgium, Switzerland, Germany Denmark, Spain, Finland, France, United Kingdom, Greece, Hong Kong, Ireland, Israel, Italy, Japan, Netherlands, Norway, New Zealand, Portugal, Sweden, and Turkey. | Terminated | November 28, 2014 | 577 | 4710 | Not_Computer
Z-1472-2013 | 65093 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Versaport Bladeless Optical Fixation Cannula - 5mm |  |  | May 06, 2013 | 2013 | Covidien LP | Seals may disengage from the cannula which may result in a component inadvertently disengaging into the patients cavity | Device Design | Covidien sent an Urgent Medical Device Recall letter dated May 6, 2013,via federal express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the product and return all inventory to Covidien Surgical Solutions, 195 McDermott Road, North Haven, CT 06473. Customers were requested to return completed Response Form.Customers with questions were instructed to contact their Covidien representative or Covidien Customer Service, Monday through Friday, Sam - 6:30pm ET, at (800) 962-9888, option 1, and then option 2.For questions regarding this recall call 203-492-8165. |  | 57461 | 57461 untis | Worldwide Distribution - USA (nationwide) and internationally to the following countries:CanadaAustriaBahrainBenoniCenturionChatsworthCzech RepublicDenmarkEgyptFinlandFranceGermanyIrelandIsraelItalyJapanKuwaitNetherlandsNorwayPortugalQATARSaudi ArabiaSerbiaSingaporeSouth AfricaSpainOMANSwedenSwitzerlandUAEUnited KingdomVenezuela | Terminated | August 12, 2014 | 463 | 218221 | Not_Computer
Z-1481-2013 | 65140 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView SPECT gamma camera for Emission Computed Tomography, Philips Medical Systems, Cleveland, OH 44143 |  |  | April 25, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips has received one report from the field that there was an unexpected motion of the detector due to a failure of the mechanical assembly. | Other | Philips Healthcare sent an Urgent Medical Device Correction letter dated April 26, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customer were not required to take any specific action. Customers who need further information or support concerning the affected product, they should contact their local lPhilips representative or local Philips Healthcare office. For North Americ and Canada they should contact the Customer Care Solutions Center (1-800-722-9377, option 5).For questions regarding this recall call 440-483-7600. |  | 505 | 505 units | Worldwide Distribution - USA including AK, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, VA, WA, WI, WV &amp; WY.Internationally to Australia, Austria, Bahrain, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Dominican Republic, Egypt, France, Germany, India, Indonesia, Iran, Italy, Japan, Japan, Korea, Lebanon, Lithuania, Malaysia, Mexico, Portugal, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Syria, Thailand, Turkey, Taiwan, Ukraine, United Arab Emirates, United Kingdom, Venezuela &amp; Viet Nam. | Terminated | February 12, 2014 | 293 | 505 | Not_Computer
Z-1465-2013 | 65209 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Active Articulation E1. |  |  | April 26, 2013 | 2013 | Biomet, Inc. | Biomet is recalling Active Articulation E1 Hip Bearing implants. Biomet has initiated this action following an investigation which identified that outer label indicates two conflicting O.D. sizes while the inner labels indicate an incorrect O.D. size.The implant itself is etched with the correct O.D. size. | Labeling mix-ups | Biomet sent an Urgent Medical Device Recall letter dated April 26, 2013 to consignees detailing that the outer label indicates two conflicting O.D. sizes while the inner labels indicate an incorrect O.D. size. This action requires the immediate location and discontinued use of the product and its return to Biomet. Ciustomers were advised to:Immediately locate and remove the identified device(s) listed from circulation .-Carefully follow the instructions on the enclosed "FAX Back Response Form" .- Fax a copy of the Response Form to 574-372-1683 prior to return of product.-Use priority carrier for your shipment.- If you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. However, you are charged with the location and return of these products. |  | 8 | 8 units | Distributed in New York. | Terminated | August 29, 2013 | 125 | 8 | Not_Computer
Z-1463-2013 | 65168 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci 5 mm Instrument Cannulae for use with the da Vinci S and Si Surgical Systems. |  |  | May 06, 2013 | 2013 | Intuitive Surgical, Inc. | Intuitive Surgical is implementing a voluntary Field Action to implement a modified inspection process for the 5 mm Cannula (PN 420011) and 5 mm Flared Cannula (PN 420262). | Use error | Intuitive Surgical sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 14, 2013 to all affected customers. The letter identified the product,problem, and actions to be taken by the customers. Contact Customer Service at 800-876-1310 for questions regarding this notice. |  | 2101 | 2101 consignees | Worldwide distribution: US (nationwide) and countries of: Austria, Belgium, Denmark, Finland, France, Germany, Ireland, Netherlands, Norway, South Korea, Sweden, Switzerland, and United Kingdom. | Terminated | April 15, 2014 | 344 | 2101 | Not_Computer
Z-1486-2013 | 65314 | No | GAM | Absorbable poly(glycolide/l-lactide) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Digitex Absorbable Polydiaxanone Suture, Size 2-0. |  |  | May 13, 2013 | 2013 | Coloplast Manufacturing US, LLC | Coloplast is conducting a recall on three lots of Digitex Absorbable Polydiaxanone Suture, Size 2-0 because it is mislabeled with a three year expiration date. | Incorrect or no expiration date | The firm, Coloplast, sent a letter dated May 8, 2013 to its customers. The letter described the problem, the product and actions to be taken. The customers will be contacted by Coloplast Customer Service with instructions for the handling and return of any unused product.If you have any additional questions, the dedicated Customer Service number regarding this effort is 800-258-3476. |  | 772 | 772 units | US Distribution to: KY, MI, OH, PR, TX, and VA. | Terminated | April 10, 2014 | 332 | 772 | Not_Computer
Z-1482-2013 | 64991 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Straight Impactor - Dorado IBC, Part Number: SI02140. The device is used to position implants after they have been inserted into the disc space during implantation of intervertebral body fusion devices. |  |  | May 10, 2008 | 2008 | SpineFrontier, Inc. | During two separate occasions an impactor plate on the handle broke off while being struck by a mallet. | Process design | Firm states they notified consignees of recall between 5/10/08 and 5/22/2008. Zero product is currently in distribution. |  | 10 | 10 | Worldwide distribution: US (nationwide) to states of: CO, MA, and TX; and country of: Jamaica. | Terminated | October 08, 2014 | 2342 | 10 | Not_Computer
Z-1480-2013 | 64927 | No | OSI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | E-Z WIDER BARIATRIC CHAIR BED; Single Motion Expandable Width Technology |  |  | August 02, 2012 | 2012 | Kreg Medical Inc. | An EZ Wider bed (capacity of 1000 lb. safe working load collapsed during patient use, leaving the patient in a forward tilt position. | Device Design | Kreg Medical sent an Urgent Voluntary Medical Device Recall letter dated May 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Kreg Medical started replacing potentially defective beds with refurbished beds in August 2012. All of the potentiall defective beds were refurbished by November 20, 2012. Customers were requested to send back the Medical Device Recall Return Response to confirm receipt of the notification letter.For questions regarding this recall call 312-275-7070. |  | 59 | 59 beds | Nationwide Distribution including IL and IN. | Completed |  |  | 59 | Not_Computer
Z-1516-2013 | 65113 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | American Optisurgical's TX1 Tissue Removal System Console. The console is a software unit that is indicated for use in surgical procedures. |  |  | March 30, 2013 | 2013 | American Optisurgical Inc | The recall has been initiated due to an overheating protection voltage suppressor on the redundant safety protection circuit board. Secondary to the recall, American Optisurgical, Inc. will be making a correction to add spacers between the valve and cassette to correct a small irrigation leak that can occur after the priming cycle. Use of this product may cause the device to lose power and stop wo | Device Design | "Urgent Medical Device Recall" notifications letters were sent to customers on May 20, 2013 via USPS registered mail. Included in the Recall Notification Letter were the recommended actions and instructions for customers. Questions were referred to (949) 580-1266. |  | 194 | 194 units | Worldwide distribution, US, including the states of United States: Alabama, Arkansas, Arizona, California, Colorado, Florida, Georgia, Iowa, Illinois, Indiana, Kansas, Kentucky, Massachusetts, Maryland, Michigan, Minnesota, Missouri, North Carolina, Nebraska, New Jersey, New York, Ohio, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, Washington Wisconsin, West Virginia and the countries of Singapore and the Philippines. | Terminated | March 24, 2014 | 359 | 194 | Not_Computer
Z-1505-2013 | 65098 | No | OJU | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | FAST Procedure Kit. |  |  | April 30, 2013 | 2013 | American Optisurgical Inc | American Optisurgical, Inc initiated a recall of ChloraPrep Hi-Lite Orange 3mL Applicator, in response to a sub-recall letter from Medline Industries, Inc due to lack of sterility assurance caused by an incomplete seal on the individualunit packaging. | Nonconforming Material/Component | American Optisurgical Inc. sent an "URGENT: VOLUNTARY MEDICAL DEVICE RECALL AND PRODUCT EXCHANGE" letter dated May 17, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 949-580-1266 for questions regarding this notice. |  | 63 | 63 individual applicators | Nationwide Distribution-USA (nationwide) including the states of AZ, CA, FL, GA, IL, KS, MA, MI, NC, NJ, OH, PA, TN, TX, VA, and WI. | Terminated | March 24, 2014 | 328 | 63 | Not_Computer
Z-1503-2013 | 65321 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker Instruments 0400-850-000S2 T4 Toga, XL, w/Inside Tie) Sterile, EO; single use; |  |  | May 23, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | Based on an accelerated aging test for this product packaging, there is an indicator of potential for a breach in the sterile barrier in the header bag and potentially the blue sterilization wrap for togas within the second and third year of shelf life. The test method used involved high heat to accelerate the aging of the packaging. Stryker is conducting additional testing and assessing this pr | Nonconforming Material/Component | Stryker sent an Urgent Urgent Recall noitifications letter dated May 7, 2013 via certified US mail and E-mail l to all affected customers, distributors and direct accounts. The letter identified the affected products, problem, risk to health, and actions to be taken. Customers were instructed to check inventory, immediately quarantine any affected product, and open each package before discarding to prevent use. Customers were directed to fill out and return a Business Reply Form to Stericycle, Phone: 888-943-2396 Fax: 888-943-4192, Email: stryker5893@stericycle.com verifying directions and confirming recalled product was destroyed. For questions call 269-389-4354 or 269-389-4518. |  | 2000 | 2,000 units ( 200 boxes) | USA Nationwide Distribution | Terminated | April 14, 2014 | 326 | 599490 | Not_Computer
Z-1506-2013 | 65006 | No | HQE | Vitreous aspiration and cutting instrument. | Ophthalmic | Ophthalmic | 510(k) | The Constellation Vision System Pneumatic Handpiece. |  |  | April 17, 2013 | 2013 | Alcon Grieshaber AG | Alcon is conducting a Medical Device Removal for the Constellation Pneumatic Handpiece, DSP, product number 725.01. This action is being taken because the forceps or scissors tips (sold separately) may dislodge from the hand piece, if not attached by the user according to the Directions for Use. | Use error | Alcon sent an "URGENT: MEDICAL DEVICE REMOVAL" letter on April 17, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your local Alcon Sales Representative for questions regarding this notice. |  | 8455 | 8,455 units | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Belgium, Canada, Chile, China, Egypt, France, Germany, Great Britain, India, Ireland, Israel, Italy, Japan, Kazakhstan, Kuwait, Kyrgyzstan, Lebanon, Korea, Malaysia, Mexico, Pakistan, Peru, Russia, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, Turkistan, Venezuela, and Veit Nam. | Terminated | January 07, 2015 | 630 | 8455 | Not_Computer
Z-1488-2013 | 65062 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Hand Controls for Operon D750, D752, D760, D820 and D850 |  |  | March 12, 2013 | 2013 | Berchtold Corp. | The firm received complaints for unintended movements. | Device Design | Berchtold sent an Urgent Medical Device Recall letter in March 2013, to all customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were recommended to clean the affected product and use the instructions to reduce the possibility of unintended movements until the firm could replace the affected product. Customers with questions were advised to contact the Berchtold Operon Pendanct Recall Coordinator at 800-243-5135, or 843-569-6100, ext 306.For questions regarding this recall call 843-569-6133. |  | 3763 | 3763 | Nationwide Distribution | Terminated | May 08, 2017 | 1518 | 3763 | Not_Computer
Z-1487-2013 | 65223 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | AGFA Digital Radiography X-Ray System DX-D 100 |  |  | May 15, 2013 | 2013 | AGFA Corp. | Customers could potentially experience intermittent, unintended and illogical movement when using the product. | Under Investigation by firm | Agfa sent an "URGENT SAFETY NOTICE" letter dated May 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Customers were instructed if you experience unintended movement, the unit has to be taken out of operation. Please notify you Agfa service contact at once.Please distribute this information within your facility to all those who need to be aware of this. Please complete the feedback form as soon as possible and return to us.For further questions please call 1-877-777-2432. |  | 104 | 104 | Worldwide Distribution - US Nationwide including the states of AL, FL, IL, IN, IA, KY, MA, MI, MN, MS, MO, NE, NJ, NY, NC, OH, OR, SC, TN, VA, WV, WY and the country of Canada | Terminated | December 31, 2014 | 595 | 104 | Not_Computer
Z-1458-2013 | 65192 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Part 510450 Disposable Infusion Pump, ...Elastomeric Pump 600 mL( 10mL/hr ),... found in the following kits |  |  | May 10, 2013 | 2013 | Symbios Medical Products, LLC | Symbios Medical Products, Indianapolis, Indiana, is performing a voluntary recall of all GoPumps found in the GoPump Rapid Recovery System kits and all GoBlock kits manufactured with flow control components assembled prior to July 2012. The recall is being expanded after receiving additional reports relating to excessively high flow rates on the 150 mL reservoir, 2 mL/hr GOPump (PN510031), and the | Under Investigation by firm | Symbios sent UrgentMedical Device Recall notices to all consignees beginning May 10, 2013. A second corrected/amended notice was sent on May 14, 2013 to consignees and a National Press release was issued. May 30,2013, a third amended Urgent Recall Notice was sent to all consignees correcting the pump description and lot number range table on page 5 of the notice. The notices and press release provided identification of the recalled devices including devices in the previous February 2013 recall, described possible adverse events, and included directions to to quarantine product in a secure location, in preparation for return. Consignees were asked to call 317 225-4447 ext. 25, 8:00 a.m to 4:00 p.m. EDT to receive a Returned Goods Authorization Number prior to returning product. Consignees were asked to complete and return the Attached Verification Form. Question and Concerns should be addressed to 317-225-4447 ext 25. |  | 33543 | 33, 543 for all kits | US Nationwide distribution. | Terminated | March 07, 2014 | 301 | 33543 | Not_Computer
Z-1507-2013 | 65182 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | FLEXSTENT Biliary Self-Expanding Stent System. |  |  | April 23, 2013 | 2013 | Cordis Corporation | Holes and/or tears can be introduced in the FLEXSTENTÂż inner pouch under some conditions, and to a lesser degree, in the outer pouch. | Package design/selection | Cordis sent an Urgent Medical Device Correction and Potential Removal letter dated April 24, 2013 e-mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately identify and set aside all affected products to prevent further use of the affected product, until inspected. Âż Use only product identified with a blue dot next to the lot number on the carton face label, carton end label, and outer pouch label, or product with upgraded sterile barrier system (not yet available, as of the date of this letter), review, complete, sign and return the enclosed Acknowledgement Form and maintain a copy of this notice with the affected product. For questions related to the issue, please contact our FSS/Cordisrepresentative, or please contact Cordis Direct at (800) 781-0282, Mondaythrough Friday from 9:00 AM to 5:00 PM EDT. |  | 23 | 23 units | Worldwide Distribution - USA (nationwide) and the states of: AZ, CA, CT, FL, GA, IA, IL, LA, MD, MI, MS, NC, NJ, NY, OH, PA, RI, SC, TX, WA, WI, WV, and the countries of: Czech Republic, Italy, Netherlands, and New Zealand | Terminated | June 27, 2014 | 430 | 2724 | Not_Computer
Z-1515-2013 | 65224 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA CentaurÂż iPTH Calibrator 6 pk; REF: 10492387, 10630628. |  |  | May 14, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | iPTH Assay- reagent and calibrator lots may have demonstrated higher than anticipated variability | No Marketing Application | Siemens sent an Urgent Medical Device Correction letter dated May 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For questions, contact your Siemens Technical Solutions Center or your local Siemens technical support. |  | 198 | 198 (6 pk) | Nationwide Distribution. | Terminated | March 05, 2015 | 660 | 3908 | Not_Computer
Z-1517-2013 | 65289 | No | DXE | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | lnvatec SpA brand Diver C.E., Model/Catalog Numbers: DVR010014145, DVR011014145, |  |  | March 04, 2009 | 2009 | Invatec Llc | Devices labeled for an intended use was not included in the existing 510(k) or PMA, in that, the Indications for use was modified. | Labeling False and Misleading | Invatec sent an Urgent - Device Field Action letter dated March 4, 2009, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were required to return a confirmation form, acknowledging the action taken, and identifying affected units in their inventory for return of the product for replacement. The firm indicated that this recall was fully executed in 2009, and this 806 captures the events that took place at that time.For questions regarding this recall call 707-566-1143. |  | 177 | 177 units | Nationwide Distribution | Terminated | June 10, 2013 | 1559 | 177 | Not_Computer
Z-1436-2013 | 65248 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Brand Name: AFX" Introducer System, Model Number S17-45. |  |  | May 13, 2013 | 2013 | Endologix Inc | Endologix, Inc. initiated this voluntary recall of certain lots of Endologix, Inc. due to reports of dilator breakage during clinical procedures. | Nonconforming Material/Component | Endologix Inc sent an "URGENT-MEDICAL DEVICE RECALL NOTICE" letters dated May 13, 2013 to all affected customers. The letter identified the affected product, reason for recall, and actions to be taken. The letter instructed customers to do not use or further distribute any affected product. The letter instructed customers to share this information with physicians that perform these procedures at the facility. The firm will call to confirm that this device is located at the facility. After confirmation the firm will retrieve the device. For questions call (800) 983-2284. |  | 53 | 53 | Worldwide Distribution - USA Nationwide in the states of: NY, NH, IN, MI, FL and NJ and in the countries of: Europe and Latin America. | Terminated | August 19, 2013 | 98 | 53 | Not_Computer
Z-1522-2013 | 65283 | No | NME | Dental bur. | Dental | Dental | 510(k) | StarDental Classique(TM) Diamond Dental Bur. |  |  | May 28, 2013 | 2013 | Dental EZ Stardental Division | Recall was initiated because a manufacturing issue was found that prevented the diamond coating from adhering properly to the shank of the diamond bur for these lots. | Equipment maintenance | Dental EZ Inc. sent an "URGENT: MEDICAL DEVICE RECALL" letter on May 28, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the firm at 866-383-4636 ext.4350 for questions regarding this recall. |  | 3560 | 3,560 | Worldwide Distribution-USA (nationwide) including the states of CA, FL, IA, IN, NJ, NV, NY, PA, TN, TX, and WI, and the country of Japan. | Terminated | March 18, 2014 | 294 | 3560 | Not_Computer
Z-1520-2013 | 65081 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | DePuy Glenosphere orientation guide |  |  | May 03, 2013 | 2013 | DePuy Orthopaedics, Inc. | The Glenosphere Orientation Guide Instrument Used with the Delta XTEND Reverse Shoulder is being recalled because an arrow that is etched on the instrument was incorrectly placed on the opposite side of the instrument. Using the improperly etched arrow as a guide could potentially lead to incorrect orientation of the glenosphere implant. | Process control | DePuy Synthes sent and Urgent Information - Medical Device Instrument Recall Notice letter dated May 3, 2013 via email to affected customers. The letter identified the affected product, problem, clinical implications, and potential adverse events. Arrangements will be made for all affected Glenosphere Orientation Guides (part #230795000/lot #5120443) to be returned. Customers were asked to return the response form via fax (574-371-4964). Product related questions should be directed to DePuy Synthes Joint Reconstruction (1-888-554-2482). Recall inquiries should be directed to Kim Earle, Recall Coordinator, 574-371-4917 (M-F; 8a.m.-5 p.m. EDT). |  | 20 | 20 units | Worldwide Distribution - USA Nationwide and countries of: Australia, Canada, Czech Republic, England, Germany, Switzerland and Turkey. | Terminated | August 28, 2013 | 117 | 20 | Not_Computer
Z-1518-2013 | 65108 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | TX1 Tissue Removal System disposable handpieces. |  |  | April 30, 2013 | 2013 | American Optisurgical Inc | American Optisurgical, Inc. is recalling certain lots of TX1 Tissue Removal System disposable handpiece due to inventory being released into distribution without a cleared FDA 510(k) to cover recent product changes. Use of this product may lead to an underperformance of the device. | No Marketing Application | American Optisurgical Inc. (AOI) sent an Urgent Medical Device Recall letter along with a response sheet via USPS registered mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine your inventory and quarantine product subject to recall. Once the product has been identified, return all products by FedEx to AOI using the AOI FedEx account 30655659-2, and referencing the Returned Goods Authorization (RGA) number:RGA# (to be assigned). For questions contact American Optisurgical Inc. Customer Service 949-580-1266. |  | 1436 | 1,436 units | Worldwide Disitribution - USA Nationwide in the states of Alabama, Arkansas, Arizona, California, Colorado, Delaware, Florida, Georgia, Iowa, Illinois, Indiana, Kansas, Kentucky, Massachusetts, Maryland, Michigan, Minnesota, Missouri, North Carolina, Nebraska, New Hampshire, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, West Virginia, Wyoming and the countries of Quezon City, Philippines, Singapore, and Panama City, Panama. | Terminated | March 24, 2014 | 328 | 1436 | Not_Computer
Z-1521-2013 | 65217 | No | NWV | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Applied Medicals EpixÂż and Direct Drive Laparoscopic Graspers, Model Numbers C4130 and C4140. |  |  | May 01, 2013 | 2013 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall on specific lot numbers of its EpixÂż and Direct Drive LaparoscopicGraspers. When the ratchet trigger is activated, if a great enough force is applied, there is a potential for the triggerto fracture and the jaws to remain in the closed position. The likelihood of this situation to occur and result inpermanent patient injury is highly unlikely; | Nonconforming Material/Component | Applied Medical notified customers of the recall beginning on May 13, 2013. The "URGENT: MEDICAL DEVICE RECALL" letter identified the product, problem, and actions to be taken by the customers. For questions regarding this notice, contact the firm at (949) 713-8041. |  | 127305 | 127,305 units | Worldwide Distribution: USA Nationwide and countries of: Argentina, Austria, Australia, Bahrain, Belgium, Canada, Chile, Columbia, Denmark, Ecuador, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Jordan, Kuwait, Malaysia, Lebanon, Lybia, Mexico, Netherlands, New Zealand, Pakistan, Peru, Poland, Portugal, Qatar, Saudi Arabia, Serbia, Singapore, Spain, South Africa, Switzerland, Taiwan-Province of China,Turkey, United Arabs Emirates United Kingdom, Uruguay, and Venezula-Bolivarian Republic. | Terminated | February 06, 2014 | 281 | 127305 | Not_Computer
Z-1525-2013 | 65188 | No | LTT | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Siemens brand MicroScan Dried Negative Breakpoint Combo 47 panels (B1017-417), SMN #: 10483099), |  |  | May 02, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | The product has the potential for false negative or delayed reaction for Arginine (ARG) results. In addition, 42 hour identification panel holds may occur due to a weak ARG reaction. | Nonconforming Material/Component | Siemens sent an Urgent Field Safety Notice dated May 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact your local Siemens Healthcare Diagnostics Representative for information concerning this notice. |  | 683 | 683 units | Nationwide Distribution. | Terminated | September 17, 2013 | 138 | 683 | Not_Computer
Z-1523-2013 | 65296 | No | LKN |  |  | Gastroenterology/Urology | 510(k) | AMICUS Exchange Kit; |  |  | May 24, 2013 | 2013 | Fenwal Inc | Fenwal has initiated a voluntary Urgent Product Recall for lot FA12C07070 of product code R4R2339, AmicusExchange Kits. Fenwal identified a labeling issue with this batch of Product Code R4R2339 (which was cleared for European use only) in which this batch was distributed within the United States market. The problem is identified as labeling that did not include the written description for the s | Process change control | Fenwal sent an Urgent Product Recall letter dated May 24, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately check inventory, discontinue the use of this prod.uct, return the affected product to Fenwal, and complete and return the attached reply form. For questions call 1-800-333-6925. |  | 19 | 19 Cases (114 Units) | USA Nationwide Distribution | Open, Classified |  |  | 19 | Not_Computer
Z-1524-2013 | 65046 | No | FZP | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BARD Biopsy Systems UltraClip Dual Trigger breast tissue marker. |  |  | April 16, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) has confirmed that some devices with the recalled product code / lot numbercombination may have the metal tissue marker separated from the needle prior to use. | Nonconforming Material/Component | Bard Peripheral Vascular, Inc. initiated a recall of certain product code and lot number combinations of Bard UltraClip Dual Trigger Tissue Marker by contacting customers on 04/16/2013. On 04/18/2013, BPV sent a Customer Recall notification Letter via FedEx with proof of delivery notice. The letter included reason for recall, a table of affected product code and lot numbers, instructions (Do not use or further distribute any affected product), a response form (titled Recall and Effectiveness Check Form) enclosed, and contact information. |  | 3085 | 3,085 | Worldwide Distribution: USA Nationwide and Internationally to Europe. | Terminated | January 30, 2014 | 289 | 3085 | Not_Computer
Z-1526-2013 | 64276 | No | NEU | Implantable clip. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BARD Biopsy Systems (A Business Unit of Bard Peripheral Vascular, Inc.), UltraClipÂż Dual Trigger Breast Tissue Marker 17g X 10cm needle, Ultrasound Enhanced Coil Marker 5 Pack; Sterilize, Non-pyrogenic, Rx only, Single Use; Bard Peripheral Vascular, Inc., 1625 West 3rd Street, Tempe, AZ 85281 USA. |  |  | January 23, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) is initiating this recall because a combination of BardÂż UltraClipÂż Dual Trigger Tissue Marker's and support catheters (Microsheath and Usher) were inadvertently distributed to customers without completing the sterilization process (non sterile). | Process control | The firm, BARD Peripheral Vascular, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter on January 23, 2013, to its customers. Bard EXPANDED its recall to include (MicroSheath and Usher support Catheters) and sent another "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated March 5, 2013 describing the additional products, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed as follows: Do not use or further distribute any affected product; check all inventory locations within your institution; remove any identified product from your shelves and return to Bard Peripheral Vascular, Inc. 1415 W. 3rd Street, Tempe, AZ 85281; if you have further distributed any of the product, immediately contact that location and advise them of the recall and have them return to BPV; If you have used the affected product, consider notifying, educating and monitoring those affected; and complete and return the Recall and Effectiveness Check form via fax to: ATTN; Recall Coordinator in Customer Service at 1-800-994-6772, even if you no longer have possession of the product; if unable to FAX call BPV at 1-800-321-4254 Option #2 Ex 2727 and report verbally.Please call our Recall Coordinator at 1-800-321-4254 Option #2 Ext 2727 (M-F 7am to 4pm MST) or email at HUsilvia.carrillo@crbard.com with any questions. |  | 815 | 815 | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KY, KS, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, OK, PA, TX, VA, WA, WV and WY. | Terminated | June 10, 2013 | 138 | 1370 | Not_Computer
Z-1531-2013 | 65341 | No | LZF | Infusion pump. | General Hospital | General Hospital | 510(k) | VIA Medical Collection Bag; The collection bag is an accessory waste bag used in conjunction with the sterile LVM blood gas sensor with the LVM Monitor. The collection bag is packaged in a tyvek pouch which is included in either replacement kit (777-3103) or LVM sensor kit (777-3102). |  |  | May 24, 2013 | 2013 | International Biomedical, Ltd. | International Biomedical has received reports of interference between the collection bag and the extension set connectors. Evaluation of the affected components indicates that the collection bag connector can become occluded due to an over-tightening of the luer lock by the end user. To alleviate the occurrence of this issue in the fixture, the connector on the collection bag has been improved to | Device Design | International Biomedical decided to recall and sent "Product Bulletin" notification letters to their consignees on 05/24/2013 or 05/30/2013. The letter described the product issue and provided customers with recommended actions. Questions were directed to (512) 873-0033. |  | 7000 | 7,000 | Nationwide Distribution. | Terminated | July 29, 2013 | 66 | 7000 | Not_Computer
Z-1530-2013 | 65330 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ECHELON 60mm Endoscopic Linear Cutter Reloads Black in Flexible Blister pack with Tyvek lid. |  |  | May 13, 2013 | 2013 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery is initiating a voluntary worldwide recall for ECHELON 60mm Black Reload (ECR60T) due to the potential for incomplete staple line formation from reload damage during the firing sequence. | Device Design | Ethicon sent a Urgent: Medical Device Recall letter dated May 29, 2013, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.Customers were instructed1.To examine your inventory immediately to determine if you have affected product manufactured by Ethicon and remove the affected product(s).See Attachment A for assistance in identifying affected product.2. Complete the Business Reply Form (BRF) Attachment C and fax or email it to Stericycle at 1-888-943-4190. Please reference Event 4706. If you have product to be returned, keep a copy of this form for your records.Your JJHCS number has been pre-populated on the BRF. Replacement product will be issued for all returned recalled product.3. All affected product must be returned by August 30, 2013, to receive replacement product. To return affected product, photocopy your completed Business Reply Form, placeit in the box with your product, and affix the pre-printed UPS shipping label to return to Stericycle. Returned ECHELON 60mm Black Reload (ECR60T) will be replaced with corrected ECHELON 60mm Black Reload (ECR60T) as soon as supply becomes available.4. If you do not have affected product in your inventory, you should still complete and return the Business Reply Form, indicating you have no affected product.5. Please share this information with all of the appropriate staff at your facility.6. If you need assistance identifying an alternative product code, please speak directly with your Ethicon Sales Representative. If you need clinical or product support, please contact your local sales representative or call ourCustomer Support Center, 1-877-ETHICON (877-384-4266 OPTION 6 to speak with a registered nurse). If you need additional shipping labels or a communications package, contact the Stericycle support center at 1-888-943-2397 and reference Event 4706. |  | 57540 | 57,540 Units | Worldwide Distribution - US Nationwide all fifty states and including the countries of Canada, Belgium, India, Malaysia, Saudi Arabia, Brazil, Israel, Mexico, Singapore, Jordan, New Zealand, Taiwan, Chile, Korea, Puerto Rico, Turkey, Columbia, Kuwait, Qatar, United Arab Emirates, Egypt, Lebanon and Russia. | Terminated | February 11, 2014 | 274 | 57540 | Not_Computer
Z-1532-2013 | 65175 | No | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | embla Embletta Gold (It functions as a Polysomnography Recorder and Ventilator Effort Recorder). |  |  | April 24, 2013 | 2013 | Embla System Llc | Calibration error for Embla Embletta Gold units calibrated through service center between June 15, 2012 to February 27, 2013. This calibration error could cause a clinician to over titrate the patient and prescribe a therapy pressure higher than is necessary. | Nonconforming Material/Component | The firm, Embla Systems, LLC, sent an "URGENT: MEDICAL DEVICE RECALL" letter on April 24, 2013, to their customers via courier. The letter described the product, problem and actions to be taken. The customers were instructed to return all recalled units for recalibration and to contact Technical Support at 1-888-662-7632, option #1, at your earliest convenience; Reference the EBG Field Corrective Action in order to obtain your return number. Embla will provide overnight shipping both ways and a three business day turn around for affected units. If you have transferred any of these products to another location, please forward a copy of this urgent notification to that location and notify Manager, Quality and Regulatory Affairs, at 613-254-8877 ext 1105 or via email at EBGFieldCorrectionProgram@natus.com. |  | 519 | 519 units were calibrated at the service center | Worldwide distribution: US (nationwide) including states of: CO, AK, WA, CA, WI, MI, MA OH, DE, FL CA, CT, TX, IL, GA, PA, AZ, DC, VA, MD, UT, IN, MO, NY, NC, OR, NV, NJ, NH, DE, and country of: Canada. | Open, Classified |  |  | 519 | Not_Computer
Z-1533-2013 | 65395 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | BARD PERIPHERAL VASCULAR VACCESSÂż PTA Balloon Dilatation Catheters, 8mm x 4cm x 80cm. |  |  | May 28, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular (BPV) has confirmed that some product code/lot number combination of devices may have the incorrect balloon diameter labeled on the inflation hub. Specifically, the balloon size printed on the hub may read 7mm x 4cm, when the actual balloon size is 8mm x 4cm. All other product labeling is consistent with the actual 8mm x 4cm balloon size. | Mixed-up of materials/components | BARD sent an URGENT MEDICAL DEVICE RECALL NOTIFICATION letters dated May 30, 2013 via FedEx overnight with proof of delivery notice to all affected customers. The letter identified the affected product, reason for recall, actions to be taken, and contact information. Customers were instructed to complete the enclosed Recall and Effectivenss Check Form and fax to the attention of Recall Coordinator in Customer Service at 1-800-994-6772. |  | 415 | 415 | US Nationwide Distribution | Terminated | October 23, 2013 | 148 | 415 | Not_Computer
Z-1536-2013 | 65091 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 90 Degree Cannulated Infant Blade Plate 35mm x 5mm x 3 Hole |  |  | April 19, 2013 | 2013 | OrthoPediatrics Corp | Post Market Surveillance of the 90 degree Cannulated Infant Blade Plate indicated that there have been five reports concerning bending of the plate intra-operatively. | Device Design | OrthoPediatrics sent an Urgent Voluntary Recall notifiation via telephone and email to all affected consignees. The letter instructed consignees to take possession of any affected devices and quarantine immediately. A Mandatory Reply Form is included with the notification letter to be completed and returned. Please contact Greg Teghtmeyer if you have any questions (574-268-6379). |  | 130 | 130 devices were shipped | USA Nationwide Distribution | Terminated | September 19, 2013 | 153 | 130 | Not_Computer
Z-1543-2013 | 65374 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Deltran Disposable Pressure Transducer Kit (DPT Kit) |  |  | June 04, 2013 | 2013 | Utah Medical Products, Inc | Utah Medical Products, Inc. is recalling certain Deltran Disposable Pressure Transducer Kits due to complaints related to packaging integrity of the sterile packaged devices. | Packaging | Utah Medical Products (UTMD) sent a Recall Notice letter dated June 3, 2013 via fax and email to all affected customers and will be completing follow up telephone calls by June 7, 2013. The letter identified the affected product, problem, health risk, and actions to be taken. UTMD instructed customers to locate and return affected products for exchange. For questions call Customer Service at 800-533-4984. |  | 1800 | 1,800 kits | Worldwide Distribution - United States Nationwide in the states of (MO, NE, KY, FL, OH, NY), and countries of Columbia, Brazil, Venezuela. West Indies, Latin America, Ecuador are distributed through a US direct consignee exporter. | Terminated | September 19, 2013 | 107 | 1800 | Not_Computer
Z-1542-2013 | 65173 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | QuickDraw Venous Cannula with Insertion Kit. ThruPort Systmes. Models QD22 (22 Fr) and QD25 (25 Fr). The product is packaged sterile and non-pyrogenic in a sealed, peel-type pouch. Two pouches (cannula and accessories) go into one shelf carton. One shelf carton goes in one shipper carton. |  |  | May 14, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards is recalling the QuickDraw Venous Cannula due to complaints of cannula separation during withdrawal of the device when the Percutaneous insertion method is used. | Process design | The firm, Edwards Lifesciences, sent an "URGENT FIELD SAFETY NOTICE PRODUCT RECALL - ACTION REQUIRED" notice dated May 14, 2013 to its customers. The letter described the product, problem, and actions to be taken. The customers were instructed to review our entire inventory for any QuickDraw Venus Cannula that has not expired; quarantine affected product from your inventory; discontinue use and return unused devices; complete and return the acknowledgment form via fax to Edwards Customer Service on 800-422-9329 within three days of receipt of this Field Safey Notice; contact Customer Service at 800-424-3278 to obtain a Returned Goods Authorization number and replacement product; return product to Edwards Lifesciences, Attn: Santosh Bhagat, 12050 Lone Peak Drive, Draper, UT 84020 Attention: RECALL, RGA#xxxxxx.If you have any questions that have not been answered by this notice, please call Edwards Customer Service at 800-424-3278 from the hours of 6:00AM - 4:30PM PST. |  | 22616 | 22,616 | Worldwide distribution: US Nationwide and country of Canada. | Terminated | October 14, 2014 | 518 | 22616 | Not_Computer
Z-1541-2013 | 65285 | No | LZS |  |  | Ophthalmic | PMA | STAR S4 IR Excimer System (STAR S4 IR) |  |  | May 28, 2013 | 2013 | Abbott Medical Optics, Inc. | Under specific conditions, there is an unlikely potential to result in inadvertent laser firing of the laser. | Other | AMO Manufacturing sent an Advisory Notice dated May 28, 2013, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to distribute the attached advisory notice to all affected customers in their assigned region. Customers were instructed to notify each surgeion at their site and complete and return the attached form to AMO via fax as an an acknowledgement of receipt of this advisory notice and that they understand the information contained within the letter. Customers with questions were instructed to contact their regional contact.For questions regarding this recall call 408-273-4016. |  | 1882 | 1,882, total all systems | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Canada, China, Denmark, Finland, France, Germany, Great Britian, Greece, India, Ireland, Italy, Netherlands, New Zealand, Norway, Poland, Puerto Rico, Singapore, Sweden, Turkey, Japan, Spain, Chile, Argentina, Colombia, Brazil, Taiwan, Mexico, South Korea, Venezuela, Paraquay, South Korea, Russian Federation, Costa Rico, Portugal, Uruquay, Martinique, Israel, Bolivia, Ecuador, Lebanon, Thailand, Saudi Arabia, Egypt, Dominican Republic, United Arab Emirates, Czech Republic, Indonesia, Kuwait, Ukraine, Guatemala, Phillipines, Trinidad, Tobago, Mongolia, Jordan, Malaysia, Hong Kong, Bulgaria, Oman, Kazakhstan, Bulgaria, Peru, Bolivia, Croatia, Romania, Yemen, Vietnam, Uzbekistan, Libyia, Slovakia, Algeria, El Salvador, and Cyprus.. | Terminated | November 12, 2014 | 533 | 1882 | Not_Computer
Z-1544-2013 | 64216 | No | EKR | Dental hand instrument. | Dental | Dental | 510(K) Exempt | SybronEndo Elements Obturation Unit; REF 973-0322; SN 91-1234; 2005-10; SybronEndo, 1332 S. Lone Hill Ave. Glendora, CA 91740; EU Rep: Kerr Italia, S.r.l, Via Passanti, 332, 1-84018 Scafati (SA), Italy |  |  | January 17, 2013 | 2013 | Ormco/Sybronendo | A recall has been issued by SybronEndo because a Canadian Standards Association (CSA)-directed cautionary statement label was inadvertently omitted from the Elements Obturation Unit and its Replacement A/C Power Cord. | Labeling mix-ups | The firm SybronEndo sent a "FIELD CORRECTION ACTION" letter dated January 30, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were requested to review and determine if they have any of the affected equipment in their inventory; Note: the equipment itself has not been modified in any way, and therefore there is no change in the performance; customers should continue to use the equipment in the same manner they have in the past:; Note: SynbronEndo has enclosed the cautionary label along with instructions for proper placement of the affected equipment; the customers should apply the enclosed cautionary statement label to Elements Obturation Unit and/or Replacement AC Power Cord as per instructions and complete and return the FIELD CORRECTION ACKNOWLEDGEMENT FORM ,whether or not they have any affected equipment in their inventory, via Fax to: 909-962-5605.For questions regarding this recall call 714-516-7602.Note: If you are an authorized SybronEndo international distributor, we request that you identify those customers that may have been shipped the affected lots of equipment and contact these customers to inform them of the issue within forty-eight (48) hours of receipt of this notification.If you have any questions, contact SybronEndo Customer Care at 1-800-346-3636. |  | 3615 | 3615 units total | Worldwide Distribution - USA (nationwide) and the countries of Australia, Azerbaijan, Belgium, Canada, China, Cyprus, Czech Republic, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Iceland, India, Italy, Japan, Kuwait, Latvia, Lebanon, Malaysia, Netherlands, New Zealand, Pakistan, Poland, Romania, Russia, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Thailand, Turkey, Ukraine, United Arab Emirates, and United Kingdom. | Terminated | November 07, 2013 | 294 | 3615 | Not_Computer
Z-1546-2013 | 64424 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Zimmer Patient Helper Bed Frame Adapter Assembly |  |  | February 14, 2013 | 2013 | Zimmer, Inc. | The bed frame adapter assemblies may have a weak weld where the patient helper adapter mounting bars attach to the tube. | Device Design | Zimmer sent an Urgent Device Removal letter dated February 14, 2013, to all affected consignees. The letter identified the product, the problem,and the action to be taken by the consignee. Consignees were advised to review the notification letter, follow the instructions contained in the letter, remove from use and quarantine the affected product. Consignees were also instructed to complete and return the fax back from to 877-787-0375 or zimmer3632@stericycle.com. Consignees were instructed to return the affected product back using the shipping label provided. If affected product was further distribution, consignees were asked to provide the customer's information to the firm on the back of the form. For questions regarding this recall call 866-201-9081. |  | 344 | 344 Units | Worldwide distribution: USA (nationwide) including states of: AL, AR, AZ, CA, CT, FL, HI, IL, KY, MA, MD, MN, MO, NE, NJ, NY, OH, OK, PA, TN, TX, WA, and WI; and country of Canada. | Terminated | May 01, 2014 | 441 | 344 | Not_Computer
Z-1547-2013 | 65210 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Trauma Phoenix Retrograde Femoral Connecting Bolt Trauma Fixation Systems. |  |  | April 29, 2013 | 2013 | Biomet, Inc. | Biomet Trauma ("Biomet") has initiated a recall of Retrograde Femoral Connecting Bolt, which involves Part Number 14-442021. The connecting bolt has an undersized diameter specification that may cause an interference fit with the 4mm hex driver (Part Number: 41024) near the edge of the tolerance. If the 4mm hex driver becomes stuck in the connecting bolt it cannot be taken apart and a delay in su | Device Design | Biomet sent an URGENT MEDICAL DEVICE RECALL notification letter dated April 29, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. The letter instructed customers to: immediately locate and remove the identified affected device(s), carefully follow the instructions on the enclosed "FAX Back Response Form", fax a copy of the Response Form to 574-372-1683 prior to return of product, use priority carrier for your shipment, if you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. For questions call 574-372-1570. |  | 75 | 75 | Worldwide Distribution - USA Nationwide: CA, TX, NY, PA, NJ, FL, OH, MI, CO, SD, IN, MT, GA, and WI. and the countries of Netherlands, Costa Rica, Japan, and San Juan. | Terminated | May 06, 2015 | 737 | 75 | Not_Computer
Z-1549-2013 | 33728 | microscope, phase contrast | IBM | Microscopes and accessories. | Pathology | Pathology | 510(K) Exempt | Aperio ScanScope System. Aperio ScanScope Models, XT, AT, AT Turbo, CS and CSO. |  |  | February 22, 2013 | 2013 | Aperio Technologies Inc | Aperio initiated this recall because the User Guides incorrectly state the devices had been cleared for specific applications. The User Guides should state the devices are for Research Use Only and are not for use in diagnostic procedures. | No Marketing Application | Aperio sent an Urgent Medical Device Recall letter dated April 22, 2013 to all affected consignees via certified USPS mail. The letter included the reason for the recall, recall product information, instructions, acknowledgement form, contact information ( Aperio Technical Support at 1-866-478-3999), and product labeling. The consignees were requested to acknowledge receipt of the Medical Device Correction letter. In addition, an Acknowledgement Form is required to be signed and e-mailed back to the Manufacturer asevidence of letter receipt.For questions regarding this recall call 760-304-6211, ext 6211. |  | 364 | 364 | Nationwide Distribution - including all states, none internationally. | Terminated | September 25, 2013 | 215 | 364 | Not_Computer
Z-1449-2013 | 64128 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips BuckyDiagnost Radiographic, a Diagnostic X-ray system. |  |  | January 16, 2013 | 2013 | Philips Healthcare Inc. | An inspection shows several IATDs (Installation Acceptance Test could not be located or different templates have been used | Release of Material/Component prior to receiving test results | Philips Healthcare issued an Urgent Field Safety Notification on January 11, 2013, to advised users of issued and perform corrective action to the units. Field Service Engineer (FSE) will execute corrections and attach record to IATD file.Your proposed corrective action plan (CAP) for the Bucky Diagnost system involves two mandatory field actions including new testing for sites with missing documentation and correcting the table height measurements for systems where the incorrect measurement procedure was used. You have also revised your IATD document to improve instructions to the field engineer for table height measurements. In your March 28, 2013, letter, you proposed a 6 month timeframe for this corrective action, and in your May 9, 2013 email, you communicated that the field corrections were ongoing and would be completed by July 16, 2013. In your May 9, 2013 email, you also clarified that the recalls are to be completed free of charge to any affected customer, whether or not they have a service agreement with Philips Healthcare.Further questions please call (978) 687-1501. |  | 66 | 66 units | Worldwide Distribution - US Nationwide and the country of Canada. | Terminated | October 15, 2013 | 272 | 182 | Not_Computer
Z-1537-2013 | 65376 | Yes | MNT | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Respironics V60 Ventilator, Model #V60 |  |  | June 04, 2013 | 2013 | Respironics California Inc | Respironics California, Inc. has initiated a recall on the V60 ventilator because of an issue with the software on the V60 Power Management Board Assembly. If a component fails on the Power Management Board Assembly, it may cause ventilator support to be lost with potentially no audible alarm from the ventilator. | Software design | The firm, Philips Healthcare, sent a "MEDICAL DEVICE CORRECTION" letter dated June 3, 2013 to all customers who received the V60 Ventilators. The letter describes the product, problem identified and the actions to be taken. The customers were informed that the V60 ventilator may continue to be used in accordance with its directions for use, pending the completion of the update. In addition, the customers were instructed to refer to their manual for additional information on warnings. A Philips Field Service Engineer, Approved Sevice Provider or Distributor will be contacting the customers to schedule a no-cost update and replacement of the Power Management Board Assembly Software on all V60 ventilators shipped from the manufacturer prior to April 1, 2013.If you need any further information or support concerning this issue, contact your local Philips Respironics representative at (800) 722-9377. |  | 19200 | 19,200 units total (12,060 units in US) | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belgium, Bolivia, Brazil, Canada, Chile, China, Croatia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt , Finland, France, Germany, Hong Kong, Iceland, India, Indonesia, Iran, Ireland, Israel, Italy, Kenya, Kuwait, Lebanon, Libya, Malaysia, Malta, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Oman, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Russian Fed., Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Utd.Arab.Emir. and Vietnam. | Terminated | April 24, 2014 | 324 | 19200 | Computer
Z-1548-2013 | 65212 | No | PAO | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Superior Approach Resection Guide Assembly. |  |  | May 15, 2013 | 2013 | Biomet, Inc. | Biomet has initiated this action following an investigation which identified that the slot on the resection guide is offset in the wrong direction.When the slot on the resection guide is offset in the wrong direction, the Steinman pin that is intended to hold the guide into place may not come into contact with the bone. | Device Design | On 5/15/2013 an URGENT MEDICAL DEVICE RECALL notification was sent to all consignees detailing to immediately locate and discontinue use of the product and return to Biomet.The following actions are REQUIRED:-Immediately locate and remove the identified device(s) listed below from circulation .-Carefully follow the instructions on the enclosed "FAX Back Response Form" .-Fax a copy of the Response Form to 574-372-1683 prior to return of product.-Use priority carrier for your shipment.- If you have further distributed this product. you MUST notify hospital personnel of this action via the enclosed "Dear O.R. Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. |  | 17 | 17 | Worldwide Distribution-USA (nationwide) including the states of UT, TX, AL, KY, and WI, and the countries of Australia, The Netherlands, Korea, and Canada. | Terminated | June 09, 2014 | 390 | 17 | Not_Computer
Z-1447-2013 | 64130 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Diagnost Stationary Radiographic System |  |  | January 16, 2013 | 2013 | Philips Healthcare Inc. | Installation Acceptance Test Document (IATD) non compliant with regulations IATDs (Installation Acceptance Testcould not be located or different templates have been used | Radiation Control for Health and Safety Act | Philips Heathcare issued an Urgent - Field Safety Notice letter dated January 11, 2013 to all affected customers to inform them of the systems afffected by this Field Action. The letter identified what the issue is, under what circumstances it may occur, the actions to take to avoid or minimize the occurrence of the issue and the actions planned by Philips to correct the issue. Customers were inform that if they have multiple BuckyDiagnost systems, they will be contacted by a Field Service Engineer to clarify which systems are affected. The Field Service Engineer (FSE) performs a full IATD (Installation PHILIPS Acceptance Test Document) and file the IATD record. Your proposed corrective action plan (CAP) for the Bucky Diagnost system involves two mandatory field actions including new testing for sites with missing documentation and correcting the table height measurements for systems where the incorrect measurement procedure was used. You have also revised your IATD document to improve instructions to the field engineer for table height measurements. In your March 28, 2013, letter, you proposed a 6 month timeframe for this corrective action, and in your May 9, 2013 email, you communicated that the field corrections were ongoing and would be completed by July 16, 2013. In your May 9, 2013 email, you also clarified that the recalls are to be completed free of charge to any affected customer, whether or not they have a service agreement with Philips Healthcare. For questions contact your local Philips representative. 1-800-722-9377. |  | 177 | 177 units | Worldwide Distribution - USA Nationwide and the country ofCanada | Open, Classified |  |  | 278 | Not_Computer
Z-1551-2013 | 65229 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, PET VCAR. |  |  | May 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to the automated PERCIST response criteria calculation applied in the PET VCAR application on Advantage Workstation or AW Server. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 14, 2013, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers with questions were instructed to call one of the following numbers:United Sates 800-437-1171Japan 0120-055-919For other countries they were instructed to contact their local GE Healthcare Service Representative.For questions regarding this recall call 407-475-4717. |  | 231 | 231 (38 US; 193 OUS) | Worldwide Distribution - USA including CA, CO, CT, FL, GA, IL, NY, NY NC OH, TX, VA, WV, WI.Internationally to: AUSTRALIA, AUSTRIA, BAHARIN, CANADA, CHILE, CHINE, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GEORGIA, GERMANY, HONG KONG, INDIA, INDONESIA, ITALY, JAPAN, KOREA, MEXICO, MOROCCO, NIGERIA, POLAND, ROMANIA, RUSSIA,SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM. | Terminated | January 21, 2014 | 257 | 231 | Not_Computer
Z-1550-2013 | 65245 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | IDEV Technologies SUPERA VERITAS Interwoven Self-Expanding Nitinol Stent Transhepatic Biliary System. Indicated for palliative treatment of biliary strictures produced by malignant neoplasms |  |  | May 29, 2012 | 2012 | IDEV Technologies, Inc. | There was an error on the outer labeling of this device. The inner label was correct. | Labeling mix-ups | The firm decided to rework the products that were in their control at the time and the one that was not in their control was corrected by an IDEV employee. That hospital was called and notified of the situation in May 2012. |  | 3 | 3 | Nationwide Distribution to the states of FL, NY, and PA. | Terminated | July 29, 2013 | 426 | 3 | Not_Computer
Z-1556-2013 | 64874 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products-HCY Reagent packaged in a kit carton. Catalog# 6802002. |  |  | April 01, 2013 | 2013 | Ortho-Clinical Diagnostics | The firm is recalling the specific lot due to lower than expected results. | Process control | Ortho Clinical Diagnostics sent an "Important Product Correction Notification" letter dated April 1, 2013 to all affected customers. Ortho Clinical Diagnostics recommends that their customers immediately discontinue using and discard all VITROS HCY Reagent remaining in your inventory for the affected lot. They have requested that their customers complete and return the attached Confirmation of Receipt form no later than April 12, 2013. Ortho Clinical Diagnostics requested that these notifications be forwarded to all laboratories and posted beside each VITROS System that utilizes VITROS HCY Reagent within their customers facilities. To place a product order for an alternate lot, please contact Ortho Clinical Diagnostics customer service representatives at 1-800-828-6316. The VITROS Chemistry Products HCY Reagent is used on the VITROS 5,1 FS Chemistry System, the VITROS 4600 Chemistry System and the VITROS 5600 Integrated System to quantitatively measure total homocysteine concentration in human serum and plasma. |  | 478 | 478 total (339- Domestic) (139- International) | Worldwide Distribution-USA (nationwide) including the states of AR, AZ, CA, CO, CT, FL, HI, ID, IL, KY, LA, ME, NJ, OK, TN, TX, WA, and WV, and the countries of Colombia, India, Japan, Mexico, Puerto Rico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, and Spain. | Terminated | March 22, 2017 | 1451 | 478 | Not_Computer
Z-1553-2013 | 65263 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | BC Hammock Mesh Sling (Large), product number, 537030. |  |  | May 23, 2013 | 2013 | Prism Medical Services USA | The product does not meet the firm's size standards, the product was produced in the incorrect size. | Employee error | Prism Medical contacted their US consignees on May 28, 2013 with an e-mail notification of the recall. The notification explained the Field Safety Notice Letter dated May 22, 2013. The notice identified the affected product, problem and actions to be taken. Customers were instructed to contact Customer Service at Prism Medical to receive a return goods authorization to replace the affected product at no charge. For questions call 905-850-9003 Ext. 249. |  | 5 | 5 slings | Worldwide distribution: USA (nationwide) including states of: CA, IL, OH, and TX. and country of: Canada. | Terminated | November 25, 2013 | 186 | 21 | Not_Computer
Z-1559-2013 | 65438 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Echelon, Echelon Oval and Oasis MR imaging device systems; one unit per box. |  |  | January 31, 2013 | 2013 | Hitachi Medical Systems America Inc | Hitachi discovered a software error that can occur when simultaneously scanning a patient while processing images previously acquired from a different patient. The software error can occur randomly. The error causes an image in the processed dataset to have the same unique identifier (UID) as in the scanning dataset. The processed image is stored in the place of the original image, which is lost. | Software design | Hitachi sent a Device Correction Notice beginning on January 31,2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Software correction is targeted to start by March 1, 2013. All sites will be visited by Hitachi Service to perform corrections. All letters are recorded by tracking number and date received by the customers.For further questions please call 800-800-3106 x 3720. |  | 45 | 45 system units | Worldwide Distribution - US Nationwide including the states of AZ, CA, CO, DE, FL, IL, IN, KS, KY, MD, NC, ND, NY, OH, OR, PA, TX, VA, WA, and WI. One (1) system unit was also distributed to Mexico. | Terminated | September 20, 2013 | 232 | 45 | Computer
Z-1557-2013 | 64830 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 11-Hole, 1/3 Tubular Plate DCL Plate with Collar, Part Number 241.000.011 |  |  | February 22, 2013 | 2013 | Synthes USA HQ, Inc. | The Synthes One-Third Tubular DCL Plate (11-Hole, 1/3 Tubular Plate w/Collar), was released to the United States market in error. This action is not being initiated as a result of adverse events. | Under Investigation by firm | The firm initiated their recall of the product on February 22, 2013 by sending an "Urgent Notice: Medical Device Recall" letter to the firm's sales representatives. The letter requested they examine their inventory, and/or customers inventories for the Part Number and if found, return it to Synthes for appropriate disposition. |  | 10 | 10 | Nationwide distribution. | Terminated | May 21, 2015 | 818 | 10 | Not_Computer
Z-1560-2013 | 65331 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Tecan Freedom EVO with Touch Tools Suite Version 3.0 |  |  | February 19, 2013 | 2013 | Tecan US, Inc. | WIth TouchTools 3.0 after the execution of a Rich User Prompt in Display and Wait mode, door lock errors are ignored. If an operator presses "Continue" following a Rich User Prompt and the door has been left open, the script will progress (continue without warning). This could lead to the instrument moving unexpectedly and potentially cause an injury to the operator if they are physically inside t | Software change control | The firm, Tecan, sent an "Urgent Field Corrective Action" letter dated February 15, 2013 to its consignees/customers and also contact them by phone February 19, 2013. The letter describes the product, problem and actions to be taken. The customers were instructed to not use Rich User Prompts in Display and Wait mode until TouchTools 3.0 Patch 1 or a higher version has been installed. Tecan is currently testing and releasing a patch that will correct this issue. The patch will be available by the end of February. Tecan will contact you upon its release as installation of the patch is required. If you have any further questions, please contact your local Tecan Helpdesk for assistance or call 919-361-5200 ext 19519. |  | 12 | 12 units | Worldwide distribution: US (nationwide) including states of: CA, KY, NC, OH, TX, and WI; and countries of: France, Germany and Italy. | Terminated | June 20, 2013 | 121 | 12 | Not_Computer
Z-1561-2013 | 64856 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | MEDICYCL-E- Lite Portable Oxygen System |  |  | March 08, 2013 | 2013 | Linde Gas North America Llc | A number of MEDICYL-e Portable Oxygen Systems may have misapplied valve flow knobs. | Nonconforming Material/Component | The firm, The Linde Group, sent a "Voluntary Medical Device Recall" letter dated June 21, 2013 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to view the enlarged photo on the reverse side of the notice for guidance. If you or your customers have a Unit with lot number 25/07, then please return it with the Product to your local LifeGas fill center as soon as possible. No other lot numbers are subject to this removal action.Customers with questions should contact Linde's Customer Care Center at 1-866-LIFEGAS (543-3427). |  | 52264 | 52,264 units (1106 units) | Nationwide distribution. | Terminated | February 03, 2015 | 697 | 52264 | Not_Computer
Z-1558-2013 | 65129 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | GlideScope Video Laryngoscope (GVL) 3, 4, 5 and GlideScope (AVL) 2, 3, 4 and 5 reusable blades |  |  | May 08, 2013 | 2013 | Verathon, Inc. | Verathon GlideScope Reusable GVL and AVL Blades are being recalled due to the potential risk of premature failure/breakage of the blade tip that may not be readily visible during routine inspection before or after intubation. | Device Design | Verathon issued an Urgent - Safety Alert and Medical Device Recall letter, dated May 10, 2013, to all affected customers The letter identified the affected product, problem and actions to be taken. The letter advises customers to discontinue the use of the specific serial numbers listed in the letter and asked that the affected product be returned to Verathon for replacement. The customers are also directed to fill out a return response form (included in the letter) regarding this recall.Consignees with questions are directed to contact their Verathon representative or Verathon Customer Care at 800-331-2313 (US and Canada) and 1-425-867-1348 for those consignees outside of US and Canada. the support team will be available Monday through Friday from 6 AM - 5 PM PST for the US and from 6 AM to 4:30 pm PST outside the US. Alternately they can e-mail Verathon at customerservice@verathon.com |  | 4784 | 4784 units worldwide (3,381 in the US and 1403 international) | Worldwide distribution: USA (nationwide) including states of : AK, AL, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, Washington DC, Guam, and countries of : Argentina, Australia, Brazil, Canada, China, Ecuador, France, Germany, Hong Kong, India, Indonesia, Italy, Japan, Japan, Malaysia, Panama, Peru, Philippines, Singapore, South Korea, Taiwan, Thailand, The Netherlands, United Kingdom, and Venezuela. | Terminated | December 17, 2013 | 223 | 4784 | Not_Computer
Z-1563-2013 | 65187 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Thermo Scientific MAS Bilirubin for Siemens and Dimension Vista Systems, , Product is manufactured and distributed by Microgenics Corporation, Fremont, CA |  |  | May 13, 2013 | 2013 | Microgenics Corporation | It has been confirmed from in-house testing that Direct Bilirubin (DBIL), and Total Bilirubin (TBIL) claimed in Thermo Scientific MAS Bilirubin Systems are displaying reduced stability. | Under Investigation by firm | ThermoFisher Scientific sent a Product Recall Notification letter dated May 13, 2013, by Certified US mail to all affected customer. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed1. Determine if you are using or have inventory of the recalled product.2. Discontinue use of the product immediately.3. Evaluate quality controls. They may vary among laboratories and should be evaluated according to laboratory procedures.4. Retain a copy of the letter for laboratory records.5. If any of the affected lots of MAS Bilirubin had been forwarded to another laboratory, please provide a copy of the recall letter to them.6. Complete and return the response form within 5 days.Letters also provide an alternative product - MAS Bilirubin part number BC103.Questions or concerns may be directed to Technical Services at 1-800-232-3342 or 510-979-5417. |  | 188 | 188 | Nationwide Distribution: US including the states of WI, UT, FL NY, PA, IL, CO, AZ and NJ. | Terminated | August 27, 2013 | 106 | 188 | Not_Computer
Z-1567-2013 | 65379 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE Monitor B850 |  |  | July 20, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare become aware of 4 potential issues associated with the CARESCAPE Monitor B850. These issues include CARESCAPE Monitor B850 uDOM failure, CARESCAPE Monitor B850 unstable USB communication, CARESCAPE Monitor B850 remote printing issue with currently printing PRN 50 recorder and CARESCAPE Monitor B850 reboots due to an X server freeze. | Software design | GE Healthcare sent "Important Product Information" letter dated 12 July 2012 to all affected customers. The letter described the Subject (issues) , Instructions, Affected Product Details, Product Correction and actions to be taken. For questions contact Technical Support 1-800-558-7044 or your local Service Representative. |  | 12779 | 12,779 | Worldwide Distribution - USA Nationwide including the states of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS, KY, IA, ME, MD,MA, MI, MN, MS, MO, MT, NE,NV,NH, NJ, NM, NY, NC, OR, OK OH, PA, PR, SC, TN, TX, UT, VA, WA, WY, WI, WY, and the countries of: ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK,, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, HONG KONG, INDIA, IRAQ, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MALTA, MEXICO, NETHERLAND, NEW ZEALAND, NORWAY, PANAMA, PERU, POLAND, PORTUGAL, QATAR, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY. | Terminated | June 30, 2014 | 710 | 12779 | Not_Computer
Z-1562-2013 | 65435 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE with syngo RT Therapist version 4.3 |  |  | May 29, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The potential for a safety issue, for the image quality degradation, that is addressed by the Customer Safety Advisory Notice when verifying the patient position based on 2D images that were acquired as single-exposure MV or In-Line kView images using the syngo RT Therapist 4.3. | Device Design | Siemens sent an Urgent: Medical Device Correction Customer Safety Advisory Notice letter to all affected customers. The letter identified the affected product, problem, what preventive measures can the user take and actions to be taken. Customers were advised to perform the preventive measures provided and inform all affected personnel immediately. The letter states that Siemens is currently working on a new version of the software. Customers are asked to include the Customer Safety Advisory Notice in the System Owner Manual chapter "Safety Advisory Letters". For question call 925-602-8157. |  | 8 | 8 units | Worldwide distribution | Terminated | November 26, 2013 | 181 | 8 | Computer
Z-1569-2013 | 64530 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Sarns " TCM II Temperature Control Module |  |  | May 30, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Internal testing found certain Sarns" TCM II systems were released with incorrectly assembled heatsink assemblies which could cause the units to fail. | Employee error | Terumo sent beginning May 30, 2013, an URGENT MEDICAL DEVICE RECALL NOTIFICATION were sent to all customers with instructions to notify all customers of the recall requesting all affected units to be quarantined until the heatsink units are replaced Confirm receipt by filling out the response form containing communication by faxing or emailing the attached Affiliate Response Form to the fax number or email address indicated on the form.Questions should be directed to Terumo CVS customer service in the U.S.:Telephone: +1-734-663-4145; Fax: +1-734-741-6149 |  | 7 | 7 units | Worldwide Distribution including the countries of UAE (United Arab Emirates) HONG KONG and THAILAND. Not distributed in USA | Terminated | September 17, 2013 | 110 | 7 | Not_Computer
Z-1572-2013 | 65417 | No | FGY | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | MAQUET GETINGE GROUP Arterial Cannula OD |  |  | May 14, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | The specified part numbers and lot numbers of Maquet Pediatric Arterial Cannula Packs may contain an incorrectly sized connector. | Packaging process control | Maquet sent an "Urgent Product Recall"/"Recall - Device Corrective Action Response" forms dated May 14, 2013, via Federal Express to all affected customers. The notification identifies affected product; informs customers of the nature of the issue; advises of potential adverse patient consequences that may occur; and advises customers to follow instructions on the "Recall-Device Corrective Action Response" forms included as part of the notification.If you have any further questions please contact your local MAQUET representative or call Customer Service at (888) 627-8383. |  | 100 | 100 units (60 Domestic; 40 International) | Worldwide Distribution - US Distribution including the states of Delaware, Illinois, Minnesota , New York., Ohio and Pennsylvania and the countries of Germany, Turkey and France. | Terminated | September 18, 2014 | 492 | 100 | Not_Computer
Z-1577-2013 | 65243 | No | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | BD SafetyGlide Allergy. REF 305950 Sterile Allergy Treatment Syringe Tray. The device is used for aspiration and injection of medications. |  |  | May 01, 2013 | 2013 | Becton Dickinson &amp; Company | There has been an increase in complaints regarding the needle/safety mechanism disengaging from the syringe. | Process control | Direct consignees were contacted by phone and email on 05/01/2013 and instructed to return the affected product. End users were mailed "Urgent: Product Recall" letters/Response Forms dated 05/01/2013. The notice advises customers of the product issue, identifies affected product and provides instructions on how to handle affected product. |  | 1326000 | 1,326,000 units | Worldwide Distribution including US Nationwide and the countries of Canada and Suriname. | Terminated | September 25, 2013 | 147 | 1326000 | Not_Computer
Z-1568-2013 | 65362 | No | KYI | Wrist joint carpal trapezium polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | CMC Saddle |  |  | June 03, 2013 | 2013 | Integra LifeSciences Corp. | As a result of a quality review, labelling content discrepancies were identified between the Instruction for Use (IFU) distributed in the US with the Ascension CMC and the content that is cleared by FDA. | Labeling Change Control | Integra sent an Urgent: Medical Device Recall letter to all affected product. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory, discontinue use, and complete and return the attached form. When the form is received, Customer Service Returns and Repairs will send an RMA number to return the affected inventory. For quwstions contact your regional managers or call 512-520-7697. |  | 467 | 467 devices | USA Nationwide Distribution | Terminated | October 08, 2014 | 492 | 467 | Not_Computer
Z-1564-2013 | 65458 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, Launcher Guiding Catheter, Model # LA8EBU45SH, Lot # 0006722866, 8F and Model # LA6IMAD, Lot # 0006724224, 6F |  |  | June 12, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic has identified a labeling issue for 24 units of the Launcher coronary Guide Catheter. The two affected catheter lots are mislabeled and do not accurately reflect the catheters contained within each package. | Labeling mix-ups | Medtronic "Urgent Field Safety Notice" letter dated June 2013 to all affected customers. The letter described the affected product, problem and actions to be taken. The letter instructed customers to remove and quarantine all affected product within their inventory and return the affected product to Medtronic. For questions contact your local Medtronic representative. |  | 15 | 15 | Internationally distributed: to countries of Germany and Poland. | Terminated | November 19, 2013 | 160 | 15 | Not_Computer
Z-1565-2013 | 65254 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter, MiniCap with Povidone-Iodine solution, 5C4466P, Baxter Healthcare Corporation. For dialysis. |  |  | May 21, 2013 | 2013 | Baxter Healthcare Corp. | leaking pouches | Packaging change control | The firm, Baxter Healthcare Corporation sent an "URGENT PRODUCT RECALL" letter dated May 21, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For general questions regarding this communication, contact The Center for One Baxter at 1-800-422-9837. |  | 629280 | 629,280 units | Nationwide Distribution. | Completed |  |  | 629280 | Not_Computer
Z-1566-2013 | 65343 | No | OZT |  | Cardiovascular | Cardiovascular | 510(k) | TRUE Dilatation Balloon Valvuloplasty Catheter, |  |  | March 18, 2013 | 2013 | Loma Vista Medical | Slow deflation time: Slower than normal balloon deflation time of greater than 10 seconds. The balloon normally deflates in less than 5 seconds. The device performance specification for deflation time is 10 seconds maximum. | Process design | The firm sent an "URGENT VOLUNTARY DEVICE REMOVAL" letter dated March 18, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 2379 | 2379 units | Worldwide Distribution-US Distribution in New Jersey and the countries of Germany, Switzerland, Italy, Norway, The Netherlands, Spain, Portugal, Ireland, and Austria. | Terminated | July 17, 2013 | 121 | 2379 | Not_Computer
Z-1578-2013 | 65193 | No | FNG | AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Prism Medical C-Series Carry Bar, Class I, model number 360755. |  |  | January 27, 2013 | 2013 | Prism Medical Services USA | The firm replaced the recalled carry bar with a plastic insert with a carry bar with a metal insert due to improper installation of the recalled carry bars. | Use error | The firm, Prism Medical, sent a letter dated January 27, 2013 to is customers.The letter described the product, problem and actions to be taken. The customers were instructed to perform an immediate safety check checking that both the pin and the plug have been inserted into the carry bar as outlined in the attached User Instruction guide. Prism offered to replace the recalled product with a different product with greater functionality.If you have any questions, contact VP-Global Product Development at 416-726-3107. |  | 1131 | 1131 carry bars/New England | Worldwide distribution: US (nationwide) and country of: Canada. | Open, Classified |  |  | 1131 | Not_Computer
Z-1574-2013 | 65342 | No | LKK |  |  | General Hospital | PMA | Medtronic Intrathecal Catheter Pump Segment Revision Kit, model 8596SC. Contents: 60-cm pump segment with attached sutureless pump connector, Spinal Segment Strain-relief sleeves, Pump segment strain-relief sleeves, Connector pin. Contents of inner package are STERILE. |  |  | June 03, 2013 | 2013 | Medtronic Neuromodulation | The Sutureless Connector (SC) Intrathecal Catheter connector has been redesigned to reduce the potential for occlusion at the catheter to pump interface. Medtronic is removing the unused products from the market that were manufactured with the previous design, and recommend the previous design no longer be used due to greater potential for misalignment and subsequent occlusion. | Device Design | The firm, Medtronic, sent an "Urgent: Medical Device Removal" letter dated May 2013 to its customers. The letter described the product, problem and actions to be taken. Representatives (Rep) are visiting all locations, beginning June 3, 2013, to retrieve devices with a Use By date of 2014 08 14 (August 14, 2014) or sooner. The Rep is leaving a letter with the hospitals to tell them that the Sutureless Connector Intrathecal Catheter connector has been redesigned and that they are removing unused devices. They also do not recommend using any of the old design. An Account Specific Customer Confirmation Form will be completed by the Rep. A copy of the completed form will be left with the Hospital along with the Urgent Medical Device Removal letter.If you have questions, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7am - 6pm CST. |  | 115722 | 115,722 total | Worldwide distribution: US (nationwide) and countries of: Aruba, Australia, Austria, Belgium, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Georgia, Germany, Greece, Guadeloupe, Hong Kong, Hungary, Iceland, India, Ireland, Israel, Italy, Jordan, Korea, Kuwait, Lebanon, Lithuania, Luxembourg, Malta, Martinique, Morocco, Netherlands, Netherlands Antilles, New Zealand, Norway, Panama, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, San Marino, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Terminated | April 07, 2014 | 308 | 115722 | Not_Computer
Z-1587-2013 | 65371 | No | PAA |  |  | Radiology | PMA | Soma v Automated Breast Ultrasound System (ABUS), Automated Breast Ultrasound. |  |  | May 16, 2013 | 2013 | U-systems Inc | Devices labeled for an intended use not included in the existing 510(k) or PMA, in that, the user manual contains the Indication for Use of the soma v Platinum ABUS [a 510(k) product]. | Software design | On 6/5/13, an "URGENT MEDICAL DEVICE CORRECTION" letter was distributed to all affected users, with a description of the problem and user corrective action steps. One consignee received an Urgent Notice to Customer - Medical Device Correction letter on 5/16/13. Contact the firm at 408-328-7925 for questions regarding this notice. |  | 3 | 3 units | Distributed in California and Illinois. | Terminated | November 04, 2013 | 172 | 3 | Not_Computer
Z-1584-2013 | 65246 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | CoaguChek XS System (HCP) |  |  | May 16, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche has confirmed the potential for an undetected elevated INR test result with the CoaguChek XS, CoaguChek XS Plus, and CoaguChek XS Pro meters. In rare cases, instead of a value, an ERROR 6 message is displayed. | Software design | Rochet sent an Urgent Medical Device Correction letter dated May 16, 2013, via UPS Ground to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were issued the following instructions: If you are using Warfarin therapy in combination with antibiotics and/or chemotherapy drugs, which could potentially lead to an extremely high test result (&gt; 10 INR), you may in rare cases receive repeated ERROR 6 messages instead of a test result. If the ERROR 6 message is displayed repeatedly on your CoaguChek XS meter, immediately contact your treating physician.The CoaguChek XS PT Test Strip package insert; the online version of the CoaguChek XS PST system users manual; and the CoaguChek XS systems are being updated to provide additional ERROR 6 message information and actions. Updated versions will be available on the www.poc.roche.com website by the end of June 2013.Health Care Professionals are instructed to use an alternative method to determine INR results if an ERROR 6 message is displayed repeatedly on the CoaguChek XS, CoaguChek XS Plus, or CoaguChek XS Pro meters for the same patient. Self testing patients are instructed to contact their Health Care provider immendiately if the ERROR 6 message is displayed repeatedly on the CoaguChek XS.Please contact your Roche Account Manager or Roche Diagnostics Point of Care Technical Service, 24 hours a day, seven days a week at 1-800-428-4674 if you have questions about the information contained in this UMDC.For questions regarding this recall call 317-521-3711. |  | 86553 | 86553 meters | Nationwide Distribution | Terminated | April 16, 2014 | 335 | 89896 | Not_Computer
Z-1585-2013 | 64789 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | November 20, 2012 | 2012 | GE Healthcare It | If environment variables MICRO_CORRECT_MSG_TYPE and CORRECT_MSG_TYPE are set to different values, in Microbiology result entry when released results (Final Report released) are changed and the Corrected Result messages are edited on the free-text form, some of the originally entered panel free-text is lost. For this issue to occur the original result must be amended and the environmental variabl | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated November 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Software correction is available. Please contact your GE service representative to request or install this update. Please contact your GE Service Representative with any question regarding this issue. For other questions please contact GE Helpdesk.Help Desk: North America: 1-888-778-3375 Help Desk: Europe: + 44 1603 877595 Help Desk: Asia Pacific + 61 3 5327 4000 |  | 9 | 9 Units | Worldwide Distribution - US Distribution including the states of OH, NY, GA, TX, TN and MN and the countries of Australia, Qatar and England | Open, Classified |  |  | 9 | Computer
Z-1586-2013 | 65411 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | DC power cable (p/n 3010556) supplied with the MEDRAD(R) Veris MR Vital Signs Monitor: |  |  | April 15, 2013 | 2013 | Medrad Inc | The firm is recalling these power cables due to a latent design reliability issue and the potential for shorting which can result in heating/melting of the cable jacket. | Device Design | Medrad Inc sent an Urgent Medical Device Component Recall letter dated April 15, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that replacement products will be available in June and shipped in the coming months along with additional instructions for returning the affected product. Customers were instructed to complete the enclosed recall response form and fax it back to Bayer at 412-406-0942. For questions contact customer support team at 1-877-229-3767 and select option 2. |  | 1066 | 1,066 | Worldwide Distribution - US Nationwide including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY and the countries of: AE, AU, AZ, BN, BR, CA, CH, CL, CO, DE, DK, DO, DZ, EG, ES, FR, GB, HK, HN, HR, HU, IN, IT, JO, JP, KR, LB, MU, MX, MY, NL, PH, PS, SA, SG, TH, TR, TW, UY, VE, and ZA. | Terminated | April 03, 2014 | 353 | 1066 | Not_Computer
Z-1570-2013 | 65135 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II, Model 8637, (The SynchroMed II Pump is supplied in 20 ml or 40 ml reservoir size.) Sterilized using ethylene oxide. |  |  | June 03, 2013 | 2013 | Medtronic Neuromodulation | Medtronic is providing Healthcare Professionals with important safety information and patient management recommendations regarding the unintended delivery of drug during the priming bolus function for the SynchroMed implantable infusion pump. The unintended delivery of drug can contribute to patient overdose or underdose symptoms which may be clinically relevant. | Device Design | Medtronic sent an Urgent Medical Device Correction letter dated May 2013, to all affected customers. The letter described the background and nature of the issue and provided recommendations for patient management and monitoring after initiation of Intrathecal Therapy. The letter directed customers to continue to use the priming bolus precedure; to monitor all patients following start or restart of intrathecal therapy; consider priming the pump prior to implant; and educate caregivers and family members to recognize the signs and sypmtoms associated with intrathecal drug therapy complications. For questions contact your local Medtronic field representative or Medtronic Neuromodulation Technical Services at 1-800-707-0933. |  | 261109 | 261,109 total devices estimated implanted Worldwide | Worldwide Distribution - USA Nationwide | Terminated | May 14, 2015 | 710 | 261109 | Not_Computer
Z-1588-2013 | 65137 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 64, Ingenuity CT, Ingenuity Core, and Ingenuity Core128 scanners are whole body Computed Tomography X-ray systems. |  |  | March 21, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | If the operator selects a patient name from the worklist using the mouse and then selects a different patient name using the keyboard (arrow keys then press enter) the patient name that populates the exam information page is the patient name selected with the mouse. Also, it is possible to begin a normal clinical scan while the anti-virus scan is running. However, if the user is running a bol | Software design | Philips Medical Systems sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated Mar 21, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 50 | 50 units | Nationwide Distribution-including the states of AR, AZ, DE, IN, KS, LA, MA, MN, MS, MT, NJ, NY, NY, OH, OR, PA, TN, TX, VT, WV, and WY. | Terminated | November 27, 2013 | 251 | 50 | Not_Computer
Z-1590-2013 | 65297 | No | NEK |  |  | Orthopedic | PMA | INFUSE(R) Bone Graft, MEDIUM KIT, REF 7510400, Medtronic Sofamore Danek USA, Inc., 1800 Pyramid Place, Memphis, Tennessee 38132 |  |  | May 28, 2013 | 2013 | Medtronic Sofamor Danek USA Inc | Kits contain a collagen sponge recalled by Integra due to possible elevated endotoxin levels resulting from a process deviation. | Nonconforming Material/Component | Medtronic on behalf of Integra LifeSciences Corporation sent an Urgent Voluntary Removal (Recall) letter on May 24, 2013, directly to all Risk Managers and Sales Representatives, with Recall Questionnaires and Confirmation forms included. Foreign consignees were notified by email on May 28, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the product and discontinue use. The letter should be forwarded to surgeons or attending physicians who have used the affected products. Customers were instructed to use the attached questionnaire to account for product in their possession, document which physicians they notified and report any patient injuries that may have occurred related to the use of the affected product by calling the phone number in the questionnaire. Customers were asked to return the questionnaire using the enclosed self-addressed stamped envelope, via their Medtronic Sales Representative or via fax to 763-367-8134For questions regarding this recall 901-399-3182.Customers were instructed to contact their Medtronic Sales Representative to coordinate the return and replenishment of affected product . Customers with questions should call 800-876-3133, ext 3197. |  | 8884 | 8884 units | Worldwide distribution: USA (nationwide including Puerto Rico) and Internationally to Australia, Brazil, Canada, Chile, Hong Kong, India, Mexico, Netherlands, , New South Wales, Panama, and Paraguay. | Terminated | May 13, 2014 | 350 | 57508 | Not_Computer
Z-1580-2013 | 65443 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed EL Programmable pumps, Models 8626, 8626L, 8627, 8627L, (The SynchroMed EL Pumps are supplied in 10 ml or 18 ml reservoir size.) Sterilized using ethylene oxide. The pump is part of an infusion system that stores and delivers a prescribed drug to a specific site. The drug is stored in the pump reservoir. Per a programmed prescription, the drug moves from the pump reservoir, through the pump tubing, catheter port, and catheter, to the infusion site. |  |  | June 03, 2013 | 2013 | Medtronic Neuromodulation | Medtronic Neuromodulation is providing Healthcare Providers with safety information and patient management recommendations related to the potential for electrical shorting internal to they SynchroMed infusion pump. An electrical short circuit in a feedthrough may present as a motor stall or low battery reset/alarm and lead to a loss of or reduction in therapy which may result in the return of und | Device Design | Medtronic sent an Urgent Medical Device Correction letter dated May 2013, to all affected customers. The letter described the issue of the pump's potentially shorting, how to identify if a pump is affected, and provided ongoing patient management recommendations. The Medtronic Rep would complete a Physician Notification Confirmation Form. A copy of the recall notification letter and the completed Confirmation Form is to be left with the physician. Customers with questions were instructed to call 1-800-707-0933.For questions regarding this recall call 763-526-9738. |  | 261109 | 261,109 total devices estimated implanted Worldwide | worldwide | Terminated | November 16, 2015 | 896 | 261109 | Not_Computer
Z-1519-2013 | 65372 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | GE Healthcare Senographe Essential System. |  |  | January 14, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential issue due to inaccessible Operator Instructions associated with the defective Operator Manual CD of the Senographe Essential. | Radiation Control for Health and Safety Act | GE Healthcare will notify customers with affected systems via letter and enclose a new CD- Rom, along with notification that a GE Representative will schedule an appointment to install the User's Manual along with a GE Service Engineer will go on site to install it on the control station of the Senograph Essential system. 16- H LA 6.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter.GE Healthcare will secure and summarize documentation that the manual has been installed and is available on each affected system and provide summarize of the documentation to the district recall coordinator.For further questions please call ( 262 ) 513-4122. |  | 49 | 49 installed in the US. | Nationwide Distribution including the states of NJ, CO, LA, AL, OH, MA, NC, ME, WI, GA, PA, TX, IN, NY, MO, KY, KS, CA, MN, FL and IL. | Terminated | November 26, 2013 | 316 | 49 | Not_Computer
Z-1599-2013 | 65407 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Multilite Dye Handpieces, an optional accessory of the RevLite SI Laser. |  |  | May 09, 2013 | 2013 | ConBio, a CynoSure Company | An inconsistency was identified related to electronic display identification of the optional Multilite Dye Handpieces of the RevLite SI Laser system that may confuse user. | Nonconforming Material/Component | The firm, Cynosure Inc., sent a "MEDICAL DEVICE FIELD ACTIONS NOTIFICATION" letter dated May 9, 2013 to its customers via UPS next day. The letter described the product, problem and actions to be taken. The customers were instructed to immediately stop using the Revlite SI Multilite Dye Handpieces upon receipt of this letter and complete and return the enclosed Field Corrective Action Acknowledgement Form via e-mail to: mflynn@cynosure.com or FAX to: 978-256-6556. A Cynosure representative will be contacting customers to schedule and arrange an exchange of the affected handpieces.If additional information is required please contact VP of Regulatory Affairs and Quality Systems at 978-367-2350 or via email at ikulinets@cynosure.com. |  | 72 | 12 (6 of model 659-1050; 6 of model 659-1100) | Worldwide distribution: USA (nationwide) including states of: IL, ME, NM and VA; and countries of: Australia and Chile. | Terminated | June 26, 2013 | 48 | 72 | Not_Computer
Z-1596-2013 | 65138 | No | MED | Liquid chemical sterilants/high level disinfectants. | General Hospital | General Hospital | 510(k) | Quick Connect QLC1677E, QRC1699E, QPC1721E and QFC1683E. STERIS Corporation, Mentor, OH 44060 |  |  | March 04, 2013 | 2013 | Steris Corporation | During a routine label review, the firm identified four SYSTEM 1E Quick Connects that will be updated to reflect the removal of obsolete device models. This labeling review process also highlighted the opportunity to combine two existing quick connects into a single product while eliminating the obsolete models. | Labeling Change Control | STERIS sent a Customer Notification letter dated February 19, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to replace their existing labels with labels they provided in their customer notification package.We apologize for any inconvenience this may have caused you and as always STERIS Corporation is dedicated to support our products and valued Customers. If you have any questions regarding this matter, please contact your STERIS Capital Account Manager. For further questions please call (440)392-7348 |  |  | QLC1677E: 38; QRC1699E: 128 &amp; QPC1721E/QFC1683E: 0 | Nationwide Distribution including the states of AK, AZ, CA, CO, FL, IL, IN, KY, LA, MA, MD, ME, MI, MO, MS, NC, NE, NH, NY, OH, OK, OR, PA, SD, TN, TX, VA, WA, WI and WY. | Terminated | July 02, 2013 | 120 | 0 | Not_Computer
Z-1597-2013 | 65361 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | Code Alert Advanced 4-Way Care Solution, CA520 System, |  |  | May 17, 2013 | 2013 | RF Technologies, Inc. | RF Technologies, Inc. has initiated a recall involving CA520 Code Alert Integrated Care Management Advanced 4-Way Care Solution . Due to a firmware error, the alarm may not be transmitted to the system server immediately. | Nonconforming Material/Component | Consignees were sent on 5/17/2013 a "Medical Device Recall" letter dated 3 May 2013. The letter described the problem and the product involved in the recall. The letter advised consignees to discontinue use of the device and return to RF Technologies. |  | 167 | 167 | Nationwide Distribution in US including the states of AK, AZ, CA, GA, IA, KY, MT, NE, NJ, ND, and WI. | Terminated | December 03, 2013 | 200 | 167 | Not_Computer
Z-1615-2013 | 65247 | No | DCK | C-reactive protein immunological test system. | Immunology | Immunology | 510(k) | VITROS Chemistry Products CRP Slides (Reagent), REF/Catalog Number 192 6740 and 809 7990, |  |  | May 09, 2013 | 2013 | Ortho-Clinical Diagnostics | While performing interval testing Ortho Clinical Diagnostics (OCD) determined that an excess amount of surfactant was unintentionally added to VITROS Chemistry Products CRP Slides during the manufacturing process. | Employee error | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated May 9, 2013 via FedEx or US Priority Mail to all US customers. Foreign Affiliates were notified by e-mail on May 9, 2013 informing them of the issue. Distributors were instructed to notify their customers of the recall. Customers were instructed to discontinue use, discard the affected product, comp[lete and return the attached Confirmation of Recipt Form, and post this notification by each VITROS System. For questions call Customer Service Center at 1-800-828-6316. |  | 923 | Domestic: 4,257; Foreign: 923 | Worldwide Disribution - USA Nationwide including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain and Venezuela. | Terminated | June 12, 2015 | 764 | 9262 | Not_Computer
Z-1608-2013 | 65464 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Cordis a Johnson &amp; Johnson Company, RadialSource Spring Wire Transradial Access Kit. Manufacturer Greatbatch Medical 2300 Berkshire Lane North, Minneapolis, MN 55441, Distributor Cordis Corporation 14201 North West 60th Avenue, Miami Lakes, Florida 33014. |  |  | May 31, 2013 | 2013 | Greatbatch Medical | Cordis Corporation, Inc and Greatbatch Medical are recalling one lot of the cordis RADIALSOURCE Spring Wire Transradial Access Kit product. Some units were manufactured with an undersized vessel dilator tip, too small to slide over the guidewire. | Process control | Greatbatch Medical and Cordis sent an "Urgent Field Safety Notice" dated May 31, 2012, to all affected customers. The letter identified the product the problem and the action to be taken by the customer. Customers were instructed:1.Read the Description of the problem section.2. Immediately identify and set aside all product listed above in a manner thatensures the affected product will not be used.3. Review, complete, sign and return the enclosed Acknowledgement Form inaccordance with the directions on the form.4. Either return any affected product per the attached instruction, or contact yourlocal sales representative to facilitate return of the affected product. Credit willbe provided.5. Pass on this notice to anyone in your facility that needs to be informed.6. If any product listed above has been forwarded to another facility, contact thatfacility to arrange return.7. Maintain awareness of this notice until all product listed above has beenremoved from the facility and returned to Cordis.8. Maintain a copy of this notice with affected product until product is returned.We apologize for any inconvenience this communication may cause. We knowthat you place high value in our products and we appreciate your cooperation inthis matter. Cordis and Greatbatch are committed to maintaining your confidence in the safety and quality of the products that Greatbatch manufactures and Cordis supplies. Further questions please call (763) 951-8181. |  | 550 | 550 kits (OUS) | Worldwide Distribution - AUSTRIA, BELGIUM, SWITZERLAND, GERMANY, SPAIN, GREECE, HUNGARY, IRELAND, ISRAEL and ITALY. | Terminated | August 04, 2014 | 430 | 550 | Not_Computer
Z-1611-2013 | 65397 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | PinnacleÂż DestinationÂż Peripheral Guiding Sheath, 8 French, 90cm, Straight Tip, Cross Cut Valve. |  |  | May 30, 2013 | 2013 | Terumo Medical Corporation | Medical devices were incorrectly labeled with extended expiration dates. | Error in labeling | The recalling firm notified consignees by letter on 5/30/2013. The letter was flagged as an urgent product bulletin addressed to risk managers. The notification advised customers to immediately discontinue use and quarantine any available stock for subsequent instructions to be provided for pick-up and replacement. |  | 20 | 20 units | Nationwide Distribution. | Terminated | October 30, 2013 | 153 | 570 | Not_Computer
Z-1600-2013 | 65365 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Brand Name: Infusor SV and LV Elastomeric Infusion Devices. |  |  | June 07, 2013 | 2013 | Baxter Healthcare Corp. | Baxter Healthcare Corporation (Baxter) is issuing an Urgent Product Recall for the products listed above due to an increase in rupture complaints. Baxter has implemented product and process improvements which have substantially improved product quality. Complaint rates for ruptures have since decreased. Baxter wants to ensure that no affected product remains in the field. | Nonconforming Material/Component | An Urgent Product Recall notification was sent to all affected consignees on 06/13/2013. The letter states the affected products, problem description, hazard involved and actions to be taken by customer/user, with contact information. The letter was accompanied by a Customer Reply Form and a table of Affected Product List. The recall notification instructs customers to segregate affected product and contact Baxter for instructions for return of product. |  | 3255096 | 3,255,096 total | Distributed Nationwide and in Puerto Rico. | Terminated | June 04, 2014 | 362 | 3255096 | Not_Computer
Z-1609-2013 | 65364 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Brand Name: Coiled Tube Infusors (Portable Elastomeric Infusion Systems). Baxter's Infusor SV Elastomeric Infusion Device |  |  | June 07, 2013 | 2013 | Baxter Healthcare Corp. | Due to an increase in complaints for leaks at the distal male Luer and Luer cap. Baxter implemented enhancements to the blue winged cap to improve the seal between the cap and the Luer. | Device Design | Baxter sent an Urgent Produd Recall notification letter dated June 13, 2013, to all affected consignees via USPS.The letter was accompanied by a Customer Reply Form and Affected Product List (table). The letter stated affected products, problem description, risk to health (or health risks), actions to be taken by customers/users and contact information.The recall notification instructs customers to segregate affected product from current inventory and contact Baxter for instructions to return affected product. Customers with questions were instructed to call Baxter at 1-800-437-5176,For questions regarding this recall call 847-270-5556. |  | 601230 | 601,230 total | Worldwide Distribution - USA (nationwide), Puerto Rico, and Internationally to Brazil, Colombia, Panama, Mexico, Singapore Japan, Australia, New Zealand, Hong Kong, India, Thailand, Taiwan, and Korea. | Terminated | June 04, 2014 | 362 | 601230 | Not_Computer
Z-1613-2013 | 65377 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SYNCHRON System(s) Lactate (LACT) Reagent. |  |  | May 28, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter has confirmed customer reports of LACT Lot M210077 failing calibration witherror condition OCR LOW. The calibration failures appear to be related to improper shipping orstorage conditions (frozen reagent). Failed calibration of LACT may cause a delay in reporting results. | Under Investigation by firm | Beckman Coulter Inc. sent an Urgent Product Correction letter dated May 28, 2013 to affected customers. The letter identified the affected product, the issue, impact, actions to be taken, and resolution. The letter instructed customers to discontinue use of the affected product and request a replacement product. For questions contact Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support/ or call 800-854-3633 in the United States and Canada. Outside the United States or Canada, contact your local Beckman Coulter Representative. |  | 3353 | 3,353 units | Worldwide Distribution - USA Nationwide and country of Canada. | Terminated | June 03, 2014 | 371 | 3353 | Not_Computer
Z-1607-2013 | 65433 | Yes | NKE |  |  | Cardiovascular | PMA | Medtronic, ConsultaÂż CRT-P, Model C4TR01, C3TR01 (not approved for US), and SyncraÂż CRT-P, Model C2TR01, Sterile EO. |  |  | June 07, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic recently identified an issue with a subset of ConsultaÂż CRT-P and SyncraÂż CRT-P devices during production. As of May 30, 2013, there have been no reported or confirmed device failures. However, because of the potential for malfunction, Medtronic is requiring the return of non-implanted devices manufactured between April 1 and May 13, 2013 for re-inspection. The issue is unique to specifi | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter dated June 7, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to immediately return non-implanted devices to Medtronic and to register the implanted ones immediately by calling the Medtronic Patient Registration Department at 1-800-848-9300. Customers with questions were instructed to contact their local Medtronic representative or Medtronic Customer Service at 1-800-848-9300.For questions regarding this recall call 763-526-6000. UPDATE: An Urgent Medical Device Correction letter, dated June 2013, was sent to affected implanting and following physicians beginning June 28, 2013. In the US these letters were hand delivered by a Medtronic representative. Outside the US, the notice will be mailed or hand delivered as determined by the geographic center. The letter informed physicians that they have implanted or are following a patient implanted with a device potentially affected by the connector bracket weld issue. The letter described how a loss of hermeticity may affect the device functionality and what symptoms a patient may experience. The letter also offers patient management recommendations. |  | 1912 | 1912 (159 USA &amp; 1137 OUS) UPDATED: 1128 affected | Worldwide Distribution - USA including CA, FL, GA, KY, LA, LMN, MO, NE, NY, OK, PA, TN, TX, WI, and PR and Internationally to Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Denmark, Ecuador, Egypt, Finland, France, Germany, Hungary, Ireland, Italy, Japan, Latvia, Lebanon, Netherlands, Norway, Poland, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Tunisia ,United Arab Emirates, United Kingdom, Uruguay , and Yemen. UPDATED Physicians location in USA: AL, AR, FL, GA, IL, IN, KS. LA, ME, MI, MN, MO, NY, OH, PA, SC, TN, TX and VA. OUS: Luxembourg | Terminated | March 07, 2014 | 273 | 1912 | Not_Computer
Z-1617-2013 | 65506 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | 20-Pole LASSO Nav Catheter Eco, Catalog # D-1349-02S. |  |  | June 19, 2013 | 2013 | Biosense Webster, Inc. | Biosense Webster is recalling the 20-Pole LASSO NAV Catheter Eco because it has been incorrectly calibrated causing them to be recognized and displayed as a 10-Pole LASSO NAV Catheter Eco by the CARTO 3 System. | Use error | An Urgent Field Notice letter was sent to all their customers who purchased the 20-Pole LASSO NAV Catheter Eco. The letter informed the customers of the problem identified and the actions to be taken. Customers are instructed to complete and return the attached Voluntary Field Removal Certification Form in accordance with the instructions listed on the form. Customers are instructed to contact Biosense Webster sales representatives or call (866) 473-7823, Monday-Friday from 7:00AM to 8:00 PM EST if they have any questions.For questions related to the Voluntary Field Removal Certification Form and its return, customers are instructed to contact the Field Action Coordinator, at (909) 839-8468. |  | 10 | 10 units | Distributed in the states of MN, FL, and UT. | Terminated | February 07, 2014 | 233 | 10 | Not_Computer
Z-1616-2013 | 65022 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 3100 and Innova 4100 systems. |  |  | May 16, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare reported a potential safety issuedue to temporary loss of X-Ray Imaging Modes during Collimator Auto-Calibration which affects Innova 3100 and Innova 4100 Systems. | Software design | A Customer Letter was not originally sent. GE Healthcare has retrospectively re-assessed the reportability of the field action and has determined that the action should be reported to the FDA.GE Healthcare has already updated all of the affected Innova 3100 and Innova 4100 systems that were distributed.As an effectiveness check, they have tracked all service requests for the field action to completion. These service requests have recorded the repair of the affected units. Exceptions to completion of the service request have been recorded as well. Records of field action completion have been retained per their processes and procedures.If you have any questions, please call 262-548-2757.For questions regarding this recall call 262-513-4122. |  | 48 | 48 | Worldwide distribution - USA including AZ, CA, CO, FL, ID, IL, ME, NV,NY, NC, OK, TN, UT, and WI; and Internationally to Canada, Chile, France, Germany, Italy, Japan, Korea, Norway and Spain. | Terminated | March 13, 2014 | 1032 | 48 | Not_Computer
Z-1618-2013 | 65202 | No | KMK | Intravascular catheter securement device. | General Hospital | General Hospital | 510(K) Exempt | StatLock CV Plus w/Pigtail, Catheter Stabilization Device Sterile EO Single Use Only, C.R. Bard, Inc. The StatLockÂż device provides stabilization for compatible medical tubes and catheters. |  |  | June 10, 2013 | 2013 | Bard Access Systems | BARD Access Systems is recalling lot JUWJF283 of StatLock CV Plus Stabilization Butterfly Devices due to incorrect pad assembly used during manufacture. The absence of the hydrocolloid strip on the bottom may pose an incremental risk including localized skin irritation. Missing directional arrow may lead to mis-orientation of the StatLock stabilization which may impact the securement of the cathet | Packaging | On 6/18/13, an "Urgent Kit Recall" notification letter was mailed to all customers informing them of the recall and instructing the consignee to remove and return any affected product lot numbers to Bard Access Systems through return goods authorization. Customers may contact 800-290-1689 if questions. |  | 3875 | 3,875 units. | National and Global distribution (Europe, European Economic Area, Argentina, China, and Brazil). | Terminated | May 21, 2014 | 345 | 3875 | Not_Computer
Z-1622-2013 | 65159 | No | LGX | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General Hospital | 510(K) Exempt | MAQUET Getinge Group Mobile Operating Table YUNO OTN GA 1433.01 EN 05. |  |  | April 26, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | MAQUET has identified a potential issue with the extension device handle (part number 1433.62A1) and will make enhancements to the YUNO table (part number 1433.01F0).Identification of the Extension device issue: By fully actuating the lever of the extension device handle, the operator's skin can enter a gap between the handle lever and a recess in the handle bar. When the lever is released, | Device Design | MAQUET Getinge Group issued an Urgent Medical Device Recall letter/"Field Action Response Form dated April 26, 2013, via Federal Express to all affected customers. The recall notification identifies the product, issue with the product, and measures to be taken by the user. Customers were instructed that the firm would contact them to arrange for a loaner YUNO table and extension device while the table and extension device are upgraded. Upon completion of the upgrades, MAQUET would reinstall their upgraded YUNO table and extension device and remove the loaner device. Customers were instructed to complete and return the enclosed Field Action Response Form, acknowledging their receipt and understanding of the communication by e-mailing a scanned copy to FieldAction@maquet.com or via fax at 973-807-1658. Customers with questions were instructed to contact their MAQUET representative or MAQUET Technical Support at 888-627-8383, option 3, followed by option 1.For questions regarding this recall call 973-709-7691. |  | 37 | 13 Tables; 24 Extension devices | Worldwide Distribution - USA including Arizona, Georgia, North Dakota, South Carolina, Texas and Tennessee, Internationally to United Arab Emirates, Austria, Brazil, Canada, Switzerland, China, Germany, Spain, France, United Kingdom, Italy, Japan, Singapore, Thailand, Turkey, Serbia, and South Africa. | Terminated | April 29, 2014 | 368 | 37 | Not_Computer
Z-1623-2013 | 65205 | No | EMF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | -Instratek Carpal Tunnel Release Blades, (1) Triangle Blade, (1) Hook Blade. REF R2056 - LOT 122279 |  |  | May 03, 2013 | 2013 | Surgical Instrument Service And Savings, Inc. | The seal which maintains a sterile barrier for reprocessed medical devices became compromised to the point where product may fall out of the pouch. | Packaging process control | MEDISISS sent a Correction Immediate Action Request letter dated May 2, 2013, to all affected customes. The letter identified the product, the problem, and the action to be taken by the customer. The letter directed the customer to inspect the products prior to use for seal integrity as specified on the product labeling: "Sterile unless package has been damaged or opened. Do not use if package is damaged."This will be followed up by a visit from the customer's sales representatives who are tasked with inspecting all the lots that may be remaining in the customer's inventory. The sales representative is directed to complete a "Verification of Seal Inspection" form and return it to MEDISISS Customer Relations. Devices with unacceptable seals are to be returned to MEDISISS. The sales representative is responsible for obtaining a shipping label from customer service to enable the devices to be returned to the recalling firm via Fed Ex.Customers with questions were instructed to call 866-866-7477. |  | 56 | 56 units | Nationwide Distribution. | Terminated | July 11, 2013 | 69 | 4540 | Not_Computer
Z-1619-2013 | 65166 | No | MAI | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | GENESYS Cross FT Suture Anchor with Two #2 (5 metric) Hi-FI Registered Trade Mark Sutures Rx Only STERILE EO CONMED LINVATEC REF CFBC-4502. |  |  | March 06, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | Incorrect blue/white suture was used. | Mixed-up of materials/components | ConMed Linvatec sent an "URGENT: Medical Device Recall Notification" letter by FedEx next day delivery on March 6, 2013, to notify all direct consignees of the recall. If the account no longer had the product in possess, they were asked to respond via a reply form or to contact ConMed Linvatec Customer Service to obtain a Service Request number for return and credit of unused unopened product. All returned product for this action would be identified, segregated and quarantined per the firm's procedures. Any returned product not part of the recall will be dispositioned in accordance with the firm's procedures.Please accept our sincere apology for any inconvenience this action may have caused . Should you have any questions please contact our Customer Service department at (800) 237-0169. |  | 60 | 60 each of both units | US Distribution including the states of FL, KY, LA, MI, NC, and NY. | Terminated | July 22, 2013 | 138 | 60 | Not_Computer
Z-1620-2013 | 65301 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Bigliani/FlatowÂż The Complete Solution Fukuda Retractors Large Nonsterile 4305-19-40. |  |  | May 23, 2013 | 2013 | Zimmer, Inc. | This is a recall expansion from the 2010 recall which now includes seven additional lots that were not previously included. Zimmer, Inc. is initiating a recall of small and large Fukuda-type retractors due to fractures occurring at the grooves around the perimeter of the blades while being used during surgery. | Device Design | In 2010, Zimmer initiated a recall of the Bigliani/FlatowÂż Fukuda Retractors due to fractures occurring at the grooves around the perimeter of the blades. Zimmer is expanding this recall to include seven lots that were not previously included. On 5/23/13, an Urgent Medical Device Recall notification letter went to all consignees explaining the recall expansion and risk associated with the recall. All consignees are to locate and quarantine the device immediately. Please return the product along with the completed inventory certification form to Zimmer. If you have any questions regarding this recall, call Zimmer at 877-946-2761. |  | 722 | 722 units (includes 2010 market withdrawal) | Worldwide Distribution-USA (nationwide) including the states of NY, MI, MN, OH, IN, IL, VA, MD, PA, WA, KY, CT, RI, NH, TX, MT, and CA and the countries of Chile, China, South Korea, and Australia. | Terminated | December 18, 2014 | 574 | 722 | Not_Computer
Z-1645-2013 | 65423 | No | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Lithium Thionyl non-rechargeable disposable battery pack P/N 350-0294-00 used with ER920 Wireless Series Arrhythmia Event Recorder. |  |  | May 02, 2011 | 2011 | BRAEMAR, INC. | Braemar Inc., became aware of a battery related incident that occurred with an Able battery pack being used in a Braemar ER920W wireless event monitor. The event did not result in any patient related injury or present any clinical impact. | Process control | Braemar began notifying customers by telephone on 5/2/2011. Braemar issued Advisory Notice letter P/N 600-0662-00 RevA dated May 16, 2011. The letter was sent by UPS to all consignees.The letter described the problem and the product involved in the recall. Consignees were advised not to use the battery pack in the Braemar ER920W or Braemar Fusion device and requested consignees to complete and return the Verification Form.For questions regarding this recall call 651-286-8620. |  | 10785 | 10,785 packs | Nationwide Distribution including FL, IL, MD, MN, NY, OH, PA, SC, TN, and TX. | Terminated | July 15, 2013 | 805 | 10785 | Not_Computer
Z-1646-2013 | 65139 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-ray systems |  |  | April 15, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received reports from the field that certain Brilliance CT 64, Ingenuity Core, and Ingenuity Core128 systems running software versions 2.6.1 or 3.5.4 exhibited intermittent swirl-like ring artifacts that may appear on reconstructed images, which may impair the diagnostic quality of affected images. | Software design | Philips Healthcare sent an Urgent Medical device Correction letter dated April 15, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter states that Philips Healthcare is implementing software update versions 2.6.2 or 3.5.5 to resolve the issue and a Philips Field Service Engineer will contact each site to schedule the software installation. For questions contact your local Philips representative or local Philips Healthcare office. |  | 2348 | 2348 total units: 2182 Brilliance CT 64 units; 95 Ingenuity Core units; and 71 Ingenuity Core128 units. | USA Nationwide Distribution - including the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. | Terminated | January 12, 2015 | 637 | 2348 | Computer
Z-1647-2013 | 65207 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | 1) VITROS 5,1 FS Chemistry System (product code 6801375) and 2) VITROS 5,1 FS Chemistry System Refurbished (product code 6801890) |  |  | May 06, 2013 | 2013 | Ortho-Clinical Diagnostics | There is evidence that an electrical short within the systems connector has caused an accumulation of heat which has radiated back through the wiring harness causing the wires and connectors to melt. | Device Design | Ortho Clinical Diagnostics (OCD) sent an Urgent Product Correction Notification letters dated May 6, 2013 and response forms to all affected customers. Ortho Clinical Diagnostics informed their customers of the issue and advised if they detect the smell of melting plastic or observe smoke originating from your VITROSÂż System, to immediately power down the system and contact OCD for assistance. They were also advised to consider performing the backup procedure more frequently to help prevent the loss of data on their Systems. In the near future, an OCD Field Engineer will contact customers who are affected and perform a Modification (MOD) procedure that will replace their potentially affected Wire Connector(s). Foreign affiliates were informed by email on May 6, 2013 of the issue and instructed to notify their customers of the issue and actions. OCD also found that this wire connector is commonly used in many types of computers (i.e., medical devices, personal computers, etc.). On 29 May 2013, OCD expanded their recall to incllude ORTHO Assay Software (OAS) Servers. As a precaution, OCD sent notifications to encompass all ORTHOÂż Assay Software (OAS) Servers that potentially contain this type of wire connector. The letters (Ref. CL13-154) were sent via US Priority Mail to all customers who were shipped an affected OAS Server. Ortho Clinical Diagnostics has requested that customers with any questions or in need of additional information, please call Customer Technical Services at 1-800-421-3311. |  | 101 | Total 101 units (84 units for # 6801375- 26 units domestic &amp; 58 units foreign) and (17 units for # 6801890-15 units domestic, 2 units foreign) | Worldwide Distribution - USA Nationwide including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI &amp; WV, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, PR, Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain | Open, Classified |  |  | 264 | Computer
Z-1654-2013 | 64888 | No | LRY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | HealthLink***BIOPSY PUNCHES*** |  |  | March 18, 2013 | 2013 | Clorox Healthcare Holdings, LLC (dba HealthLink) | Due to lack of sterilization. | Packaging change control | On March 19, 2013, HealthLink sent their customers letters instructing them to do the following: 1) Screen their inventory and quarantine the product. 2) Notify all customers that received the affected products and ask them to return the lots to their firm. 3) Complete the attached form and return the lots to HealthLink Distribution center. For further assistance, customers were directed to call customer service at 1-800-638-2625. |  | 1044600 | 1,044,600 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AZ, CA, CO, FL, GA, IA, IL, IN, MI, MO, MS, MT, NC, NJ, NY, OH, SC,TX ,VA ,WA, WI, and HI, and the countries of Canada, Mexico, and New Zealand. | Terminated | February 25, 2014 | 344 | 2849900 | Not_Computer
Z-1652-2013 | 65410 | No | DRG | Radiofrequency physiological signal transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Lithium Thionyl non-rechargeable disposable battery pack P/N 350-0294-01 used with ER920 Wireless Series Arrhythmia Event Recorder. |  |  | February 24, 2012 | 2012 | BRAEMAR, INC. | Braemar has recently became aware of a battery related incident that occurred with EWT battery packs that are used in a Braemer ER920W wireless event monitor and Fusion device. The event did not result in any patient related injury or present any clinical impact. | Material/Component Contamination | Braemar began notifying customers by telephone on 2/24/2012. Braemar issued Advisory Notice letter P/N 600-066201 RevA on 3/02/2012. The letter described the product, the problem, and the action to be taken by the customer. Customers were requested to only use TESTED batteries that they have receivewd fromBraemar starting on February 23, 2012. Complete and return the attached Verification Form. If customers have any untested product in their inventory, they were instructed to segregate those battery packs and await further instruction from Braemar to coordinate the return of those untested units. Upon receipt of the attached verification form, Braemar will continue to send replacement battery packs that have passed the additional testing protocol. Customers were instructed to contact their Braemar sales representative if they have any questions regarding the notice.For questions regarding this recall call 651-286-8620. |  | 3286 | 3286 packs | Nationwide Distribution including FL, IL, LMD, MA, MN, NY, OH, PA, TN, and TX. | Terminated | July 15, 2013 | 507 | 3286 | Not_Computer
Z-1653-2013 | 65286 | No | CAD | Anesthetic vaporizer. | Anesthesiology | Anesthesiology | 510(k) | Delta Vaporizer |  |  | May 14, 2013 | 2013 | Penlon, Ltd. | It has come to Penlon Ltd's attention of an instance where the installation and pre-use checks listed in the User Manual may not have been followed prior to the clinical use of the vaporizer. Penlon Ltd has updated the Installation and daily Pre-Use Check sections of the User Manual to ensure that the instructions are explicit and clear. | Nonconforming Material/Component | Penlon sent an Urgent Field Safety Notice dated May 14, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were reminded and requested to follow the instructions for daily pre-use checks, initial checks, and servicing. Customers were instructed to distribute the notice to all personnel who need to be aware within their organization and to any organization where the devices are used.For questions regarding this recall call 01235 547000. |  | 5303808 | 9208 (of which 576 were within the US) | Worldwide Distribution - USA including FL, KY, PA, ND, MN, IL, NJ, TX, TN, and VA. Internationally to Belgium, China, Czech Republic, France, Germany, Switzerland, Australia, Finland, Holland, Lithunia, Malta, Poland, and Slovakia. | Terminated | May 12, 2014 | 363 | 5303808 | Not_Computer
Z-1658-2013 | 64973 | Yes | NAV |  |  | Cardiovascular | HDE - Humanitarian Device Exemption | ***REF SC2101***eMax 2 Plus Console, Single Port, with Irrigation***SN G17307901802*** |  |  | June 24, 2011 | 2011 | The Anspach Effort, Inc. | The RPM of the Motor Handpiece may be less than 80,000 RPM. | Process control | Anspach Effort sent an "URGENT: Medical Device Correction" letter dated June 24, 2011 to their customers via mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete a reply form promptly confirming receipt of their letter and return it by fax to the number provided on the form. If any of the affected products were distributed to other services or facilities, please forward this information as appropriate. For questions contact Anspach Customer Support at 800-327-6887 or email customer.support@synthes.com. |  | 61 | 61 units | Worldwide Distribution - USA Nationwide including the states of: AZ, CA, FL, KY, ME, MI, MO, OH, PA, TN, TX , VA, and the countries of: Australia, Egypt, Italy, Japan, Switzerland, Taiwan, and Turkey; | Terminated | December 24, 2013 | 914 | 61 | Not_Computer
Z-1660-2013 | 64078 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Custom Procedural Trays: Merit's Custom Angio Kits |  |  | November 21, 2012 | 2012 | Merit Medical Systems, Inc. | Custom Procedural trays contain Medrad Hand Controller Sheaths which were recalled because the seal between the white polyethylene and the clear plastic may be missing on the package resulting in a potential breach of the sterility of this accessory. | Packaging | Merit Medical sent an "URGENT PRODUCT RECALL NOTICE" letter dated November 21, 2012, to affected customer. The letter identified the product, the problem, and the action to be taken by the customer: Ensure all personnel within their institution to which the affected Merit trays were distributed are notified, quarantine all affected trays and return the affected product to Merit - Maastricht. Complete and sign the Customer Response Form attached to this letter and perform the following steps:(1) Immediately fax a copy of the completed Customer Response Form to: 1-804-416-1031. (2) Mail the completed original Customer Response Form from (1) above using the provided postage-paid, self-addressed envelope to Merit. Product Return Instructions: Please return the affected products by shipping them back to Merit via Federal Express 2nd Day Air (Acct #1129-5275-6) to: RGA Department, RGA #NC-54948 Merit Medical Maastricht, Amerikalaan 42, 6199 AE Maastricht - Airport, The Netherlands. NOTE: A copy of the Customer Response Form must accompany the shipment. If you have any questions, please call 1-804-416-1048. |  | 134 | 134 trays | U.S. Military Hospital on base in Landstuhl, Germany | Terminated | July 12, 2013 | 233 | 134 | Not_Computer
Z-1661-2013 | 65493 | No | JOW | Compressible limb sleeve. | Cardiovascular | Cardiovascular | 510(k) | VenaFlow Elite 30B-B and 30BI-B |  |  | May 28, 2013 | 2013 | DJO, LLC | DJO, LLC has recently identified a product safety issue with the battery assembly for the VenaFlow 30B-B and 30 BI-B products. There is a potential for battery leakage which could result in thermal meltdown (fire). | Device Design | DJO sent a Customer Notification letter dated May 28, 2013 to Domestic customers via USPS certified return receipt and to International customers via UPS on June 14, 2013. The Customer Notification informed customers of the product safety issue, product description, and instructions to follow. DJO plans to install replacement batteries at no charge. For questions contact your local DJO Customer Service Representative. |  | 2491 | 2491 | Worldwide Distribution - USA Nationwide and countries of: Australia, Belgium, China, Guatemala CentroAmeria, France, Hong Kong, India,Italy, Korea, Lebanon, Malaysia, Morocco, Pakistan, Paraguay, The Netherlands, Romania, Singapore, Spain, Sweden, Thailand, Turkey, United Arab Emirates, United Kingdom and Venezula. | Terminated | February 13, 2014 | 261 | 2491 | Not_Computer
Z-1663-2013 | 65418 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens MEVATRON series, PRIMART (limited release)PRIMUS and/or ONCOR or ARTIST Linac Systems with BEAMVIEW capability. |  |  | May 29, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The information regarding the FILM Mode calibration may have been considered difficult to interpret when translated. SIEMENS HEALTHCARE CR / RADIATION ONCOLOGY had become aware that this issue may have resulted in mistreatment of patients. | Process change control | Siemens sent an "URGENT: MEDICAL DEVICE CORRECTION" letter and Update instruction TH022/13/S beginning on May 29, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.July 19, 2013 - TH023/13/S was distributed is now in progress. This is the FILM Mode Physics Primer Addendum. |  | 1956 | 1956 customers | Worldwide Distribution: USA Nationwide and countries of: Algeria, Angola, Argentina, Austria, Australia, Bahamas, Belgium, Bosnia Herzeg, Bulgaria, Brazil, Canada, Chile, Colombia, Croatia, Ecuador, Egypt, El Salvador, France, Germany, Greece, Hungary, India, Indonesia, Italy, Iran, Iraq, Ireland, Japan, Jordan, Kenya, DPR of Korea, Kosovo, Kuwait, Lebanon, Lithuania, Malaysia, Mexico, Nepal, Netherlands, New Zealand, Norway, Pakistan, Palestinian Ter, Peru, Philippines, P.R. China, Poland, Portugal, Romania, Saudi Arabia, Slovakia, Spain, Sri Lanka, South Africa, Sweden, Switzerlan, Taiwan, Thailand, Trinidad Tobago,Turkey, United Kingdom, Venezula, Vietnam, Rep of Yemen, and Zambia. | Terminated | January 14, 2014 | 230 | 1956 | Not_Computer
Z-1662-2013 | 65431 | No | NJR | Streptococcus spp. serological reagents. | Microbiology | Microbiology | 510(K) Exempt | BD MAX Instrument, catalog #441916 |  |  | May 31, 2013 | 2013 | Becton Dickinson &amp; Co. | Electrical power module installed on some in vitro diagnostic instruments does not meet industry standards as labeled. | Component change control | Becton Dickinson (BD) notified their customers in the US by phone on May 31, 2013 and by an Urgent Field Correction Action letter dated May 2013. Customers were advised of the affected product, problem and actions to be taken. The letter informed customers that an onsite visit by a BD Service representative is planned to perform a system check so to determine if the instrument is affected. BD's field service team plans to immediately replace the module on site if the affected part is found. For questions contact BD Technical Services Department at 1-800-638-8663. |  | 71 | 71 instruments | Worldwide Distribution - USA Nationwide and in the countries of: Japan, Singapore, Hong Kong, China, Australia, Europe, and Chile. | Terminated | August 19, 2013 | 80 | 71 | Not_Computer
Z-1665-2013 | 65429 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-LIFT Inserter Outer Sleeve |  |  | May 24, 2013 | 2013 | SpineFrontier, Inc. | The outer sleeve would not fit over the S-LIFT Inserter assembly because the hole at the top of the sleeve handle was too small. | Device Design | Spine Frontier sent an Urgent Advisory Notice letter dated On June 7, 2013 to all consignees via email. The letter identified the affected product, problem and actions to be taken. The letter informed consignees that a modification has been implemented to the Outer Sleeve component of the S-LIFT Inserter and that SpineFrontier has initiated removal of the affected product from the field.. An acknowledgement form is to be signed and returned from each recipient of the Advisory Notice. Upon receipt of the acknowledgement form and RMA request for part return, replacement parts will be shipped. For questions call 978-232-3990. |  | 12 | 12 devices | USA Nationwide Distribution including the states of: FL, TX, OK, MO, NJ, MA | Terminated | October 08, 2014 | 502 | 12 | Not_Computer
Z-1664-2013 | 65432 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BASE, S-LIFT Instrument Case |  |  | May 17, 2013 | 2013 | SpineFrontier, Inc. | The S-LIFT Instrument Case contained a bracket with peeling nylon coating. | Component design/selection | Spine Frontier sent an Urgent Advisory Notice letter dated On June 7, 2013 to all consignees via email. The letter identified the affected product, problem and actions to be taken. The letter informed customers that a modification has been implemented for the improved durability of S-LIFT Case Brackets and that SpineFrontier has initiated removal of these cases from the field. An acknowledgement form is to be signed and returned from each recipient of the Advisory Notice. Upon receipt of the acknowledgement form and RMA request for part return, replacement parts will be shipped. For questions call 978-232-3990. |  | 12 | 12 | USA Nationwide including the states of: FL, CO, Wash DC, VA, KS, TX, OH, TN, OK, CA, IL, MO, and MN. | Open, Classified |  |  | 12 | Not_Computer
Z-1666-2013 | 65442 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | InSpan Inserter, Right Assembly; InSpan Inserter, Left Assembly |  |  | May 07, 2013 | 2013 | SpineFrontier, Inc. | Potential for set screw to be come deformed. | Device Design | Spine Frontier sent an Urgent Advisory Notice letter dated May 30, 2013 to all affected consignees. The letter identified the affected product. problem and actions to be taken. The letter informed consignees that a correction was implemented for InSpan Inserters. Consignees were instructed to remove parts from the field immediately for modification and that replacement Inserters will be shipped upon receipt of completed acknowledgement form and RMA request for part return. For questions call 978-232-3990. |  | 61 | 61 individual inserter assemblies | Worldwide Distribution - USA Nationwide: in the states of: FL, CO, VA, KS, TX, OH, TN, OK, CA, IL, MO, MN, Wash DC, including PR and the countries of Dominican Republic, Medico, Panama | Terminated | January 11, 2016 | 979 | 61 | Not_Computer
Z-1669-2013 | 65462 | Yes | LDD | DC-defibrillator (including paddles). | Cardiovascular | Cardiovascular | 510(k) | Cardio Medical Products, Inc. Defibrillation/Cardioversion/Monophasic or Biphasic Pacing/ECG Electrode Rx Only Single Patient Use Only |  |  | April 26, 2013 | 2013 | Cardio Medical Products | Covidien received customer reports of Arcing/Sparking on the defibrillation electrode leadwire which may affect Cardio Medical Products Inc. defibrillator electrodes Model Number O101. | Process control | Cardio Medical Products, Inc. issued an Urgent Medical Device Recall letter dated April 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and segregate any product with the affected lot numbers and return affected product according to the instructions for product credit by contacting the Customer Service group at 1-800-962-9888, option 1, then option 2, and reference the recall letter to request a Return Goods Authorizing (RGA). Customers were instructed to complete the enclosed Verification Form and return it to Covidien by fax to 800-895-6140, even if they did not have any of the affected lots. Customers were also instructed to report any issues to the Quality Assurance Group to ensure proper device reporting procedures are followed. Customers with questions were instructed to call 1-800-962-9888, option 8, then ext 2500.For questions regarding this recall call 973-586-9500. |  | 657 | 657 units | Nationwide Distribution | Open, Classified |  |  | 657 | Not_Computer
Z-1667-2013 | 65449 | Yes | BTO | Tracheostomy tube and tube cuff. | Anesthesiology | Anesthesiology | 510(k) | Tracoe Mini 3.0mm Tube, for pediatric use. |  |  | April 17, 2013 | 2013 | Bryan Medical Inc | The pediatric tracheostomy tubes were mislabeled as the smaller infant tracheostomy tubes. | Labeling Change Control | The consignees were contacted by telephone and asked to return the recalled product. |  | 5 | 5 units | Distributed in KY and AL. | Terminated | October 29, 2013 | 195 | 5 | Not_Computer
Z-1670-2013 | 65292 | No | HQC | Phacofragmentation system. | Ophthalmic | Ophthalmic | 510(k) | SovereignÂż Compact Phacoemulsification System, v 5.3, Part Number: CMP680300. The system is used for cataract surgery. |  |  | May 23, 2013 | 2013 | Abbott Medical Optics, Inc. | Abbott Medical Optics Inc. (AMO) initiated this Urgent Notice of Medical Device Correction and Advisory Notice to all customers who use the AMO SovereignÂż Compact Phacoemulsification System, Version 5.3, because it has the potential to lose the vacuum tone mode during surgery. | Software design | AMO issued an "Urgent Notice of Medical Device Correction and Advisory Notice" on May 23, 2013 via email, to all customers with a SovereignÂż Compact Phacoemulsification System v 5.3 to inform them of the issue and provide them with actions to take in the event of sound loss during use of the SovereignÂż Compact Phacoemulsification System, v 5.3. |  | 534 | 534 systems | Worldwide Distribution, including Nationwide (US) and countries in Latin America, EMEA, and APAC. | Terminated | October 06, 2014 | 501 | 534 | Computer
Z-1668-2013 | 65430 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Insulin Infusion Sets |  |  | June 07, 2013 | 2013 | Medtronic MiniMed | Medtronic is recalling the Medtronic MiniMed Paradigm Medtronic is recalling the Medtronic MiniMed Paradigm Insulin Infusion sets because of a potential safety issue that can occur if insulin or other fluids come in contact with the inside of Medtronic Paradigm infusion sets. Under certain conditions, the infusion set may malfunction and deliver an incorrect volume of insulin. | Device Design | Medtronic, Inc. sent an Urgent Medical Device Safety Notification letter dated June 10, 2013, to all affected consignees. The letter identified the product and notified healthcare professionals of the potential for over or under delivery of insulin if insulin or other fluids contact the inside of Medtronic Paradigm Tubing Connectors. The letter informed the consignees that it applies to all Medtronic infusion sets designed for use with Medtronic Paradigm family infusion pumps. Paradigm Insulin Pump users, distributors, and healthcare professionals were all instructed to contact Medtronic's 24 hour helpline at (888) 204-7616.For questions regarding this recall call 818-567-4700. |  | 11000000 | 11,000,000 units | Worldwide distribution: US (nationwide) distribution: Algeria, Argentina, Australia, Bahamas, Bahrain, Bermuda, Boliva, Bosnia and Herzegovina, Brazil, Canada, Chile, China, Columbia, Costa Rica, Croatia, Cyprus, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, French Southern Territories, Germany, Great Britain, Greece, Guatemala, Hong Kong, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Macedonia, Malaysia, Mexico, Morocco, Netherlands, Norway, New Zealand, Oman, Panama, Paraguay, Peru, Poland, Province of China, Taiwan, Qatar, Republic of Korea, Romania, Russia Federation, Saint Lucia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Thailand, Trinidad and Tobago, Turkey, Ukraine, United Arab Emirates, Uruguay, andÂżVenezuela. | Terminated | June 10, 2015 | 733 | 11000000 | Not_Computer
Z-1672-2013 | 65317 | No | LJE |  |  | Gastroenterology/Urology | 510(k) | NephroMax high Pressure Nephrostomy Balloon Catheter. Material Number/Catalog Number: M0062101400/210-140, M062101440/210144 |  |  | May 28, 2013 | 2013 | Boston Scientific Corporation | Incorrect sized renal sheath was packaged with the NephroMax high pressure nephrostomy Balloon Catheter. | Process control | The firm, Boston Scientific, sent an "Urgent Medical Device Recall Removal - Immediate Action Required" letter dated May 28, 2013, consignees via Federal Express Priority Mail to its consignees/customers.OUS - Customer notification letters has also been sent to consignees/customers in the EU. The letter instructs the customers to Immediately discontinue use of and segregate recalled product; check their inventory; immediately remove all affected product from their inventory, and return it to Boston Scientific at Boston Scientific Corporation, US Distribution Center, Customer Fulfillment Center, 500 Commander Shea Blvd., Quincy, Massachusetts 02171 include original Verification Tracking Form and RGA #. The letter also requests the completion and return of the Account Reply Verification Tracking Form via email to: MapleGroveFieldActionCenter@bsci.com or Fax to: Field Action Center 1-866-213-1806. Follow up communications will be made to the customers that do not respond to this inquiry. Distributors are instructed to forward the recall notification to their customers.If you have any questions call 763-494-1133 or email to MapleGroveFieldActionCenter@bsci.com. |  | 27 | 27 | Worldwide distribution: USA (nationwide) including states of: NJ, NY, OH, IL, MI, MD, MA and countries of: Denmark, Great Britain, and Spain. | Terminated | January 23, 2014 | 240 | 42 | Not_Computer
Z-1674-2013 | 65332 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Liberty Cycler Set Single Conn./Ext DL |  |  | June 03, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Cassette may leak during set up or during peritoneal dialysis treatment | Unknown/Undetermined by firm | Fresenius Medical Care sent a Urgent Recall letter dated June 11, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to check their stock immediately and discontinue use and return product. Please contact your Fresenius Medical Care Customer Service Team at 1-800-323-5188 for instructions to return the product. |  | 5480 | 5480 units | Nationwide Distribution | Terminated | June 18, 2014 | 380 | 5480 | Not_Computer
Z-1681-2013 | 65320 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | ***REF C7122***QTY 1***APEX Arthroscopy Tubing Set. |  |  | May 30, 2013 | 2013 | Linvatec Corp. dba ConMed Linvatec | Conmed Linvatec received complaints that their products C7120 APEX Arthroscopy Tubing Set, One Connection and C7122 APEX Arthroscopy Tubing Set, Two Connection, contained a scratch/grooved area on the inside diameter of the pressure fitting that may cause the APEX Irrigation (pump) Console to sense a pressure drop and increase fluid flow to the surgical site. | Equipment maintenance | ConMed Linvatec sent an "URGENT: Medical Device Recall Notification" letter by FedEx next day delivery or equivalent on 5/30/2013 to all domestic direct consignees. ConMed Linvatec is translating customer notifications letters into the necessary languages and will send International notification letters no later than 6/4/2013. |  | 4691 | 4,691 units, total of each | Worldwide Distribution-USA (nationwide) including the states of AR, CA, FL, GA, IL, IN, KS, MO, NC, NJ, NM, NY, OH, OK, PA, TN, TX, UT, VA, and WI, and the countries of Austria; Belgium, Canada, China, Czech Republic, Finland, France, Great Britain, Italy, and Japan. | Terminated | September 04, 2013 | 97 | 4691 | Not_Computer
Z-1671-2013 | 65347 | No | KNQ | Esophageal dilator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | RigiflexTM 11 Single-Use Achalasia Balloon Dilator, Material Number (UPN) M00554510 |  |  | May 28, 2013 | 2013 | Boston Scientific Corporation | Product labeled with incorrect expiration date. | Incorrect or no expiration date | Boston Scientific sent an Urgent: Medical Device Recall letter to all affected customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to immediately discontinue use, check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The letter also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory. Follow-up communications will be made to the accounts that do not respond. For questions contact your local Sales Representative. |  | 91 | 91 | Worldwide Distribution - USA Nationwide including the state of NY, NE, MD, OH, TX, PA, VA, PR, CA, NJ, NY, CA and the countries of: AT, FR, DE, GB, IL, IT, LB, NL, ES, SE, TN, Argentina, and Japan. | Terminated | February 10, 2014 | 258 | 91 | Not_Computer
Z-1678-2013 | 65577 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | Baxter's elastomeric infusion pumps are non-electronic, ambulatory, disposable devices that deliver continuous medication for periods of 30 minutes up to 7 days. They are designed for medication therapies requiring continuous infusion. |  |  | July 31, 2009 | 2009 | Baxter Healthcare Corp. | There are reported incidents of leakage on the INFUSOR and FOLFUSOR pumps at the connection of the blue winged cap and the distal male Luer. Investigation of these reports indicates the incidence of leakage has only occurred after pumps are filled, primed, re-capped (using the blue winged Luer cap) and stored for a period of time, typically overnight. In addition, the reported instances of leakage | Equipment maintenance | Baxter sent an "Important Product Information" notification dated July 31, 2009, with a follow-up "Important Product Information" notification mailed on March 5, 2010. The notifications informed customers of the potential issue, and included instructions to ensure that the winged Luer cap is tightly secured after filling and priming. Customers were also instructed to follow normal clinical practices to reduce contact with any leaking medication, and to store pumps with pre-dispensed medication in a sealed pouch. If customers noticed a leak they were instructed to place the pump in a plastic bag, quarantine the device and call Baxter's Corporate Product Surveillance at 1-800-437-5176 for further instructions. Customers were asked to return the attached reply form via FAX to 1-847-270-5457. If customers further distributed the affected product they were instructed to forward the notification to other facilities. Customers with questions were instructed to call 1-800-933-0303.For questions regarding this recall call 224-270-4667. |  | 112860 | Code 2C1700KP: 7,128 units; Code 2C1702KP: 105,732 units | Worldwide Distribution - USA (nationwide), Puerto Rico and Internationally to Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, Columbia, Czech Republic, Denmark (including Faroe Islands), Ecuador, Finland, France (including Martinique and Reunion Islands), Germany, Greece, Iceland, India, Ireland, Isreal, Italy, Luxemburg, Japan, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Korea, Spain, Sweded, Switzerland, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Terminated | March 08, 2016 | 2412 | 559758 | Not_Computer
Z-1683-2013 | 65183 | No | LJH | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Two Lead Arthroscopic Irrigation Set, for use with ARTHROMATIC Plastic Containers, Product Code 2C4030 - Approximate Length 98" (2.5. m) |  |  | May 15, 2013 | 2013 | Baxter Healthcare Corp. | A firm's internal investigation led to the discovery of pinholes in the packaging of the products. The pinholes are not likely to be detected by the end user prior to use and could compromise the sterile barrier properties of the packaging. | Packaging | Baxter sent an Urgent Product Recall letter dated May 15, 2013 to all affected custorners. The letter identified the affected product, problem and actions to be taken. Customers were requested to remove all affected products and return for credit by contacting Baxter Healthcare Center for Service at 888-229-0001. Customers were instructed to complete the attached Customer Reply Form and return it to Baxter by either faxing to 224-270-5457 or scanning and e-mailing the reply form to fca@baxter.com |  | 530488 | 530,488 units | US Nationwide Distribution | Terminated | June 16, 2015 | 762 | 1136792 | Not_Computer
Z-1676-2013 | 65290 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Bigliani/FlatowÂż The Complete Shoulder Solution, Humeral Stem Provisional 12mm x 170mm |  |  | June 11, 2013 | 2013 | Zimmer, Inc. | The firm is initiating a removal of one lot of the Bigliani/Flatow Humeral Provisional Stem (00-4301-012-17, lot 62283991) as the stems manufactured under the lot are 14mm x 170mm devices incorrectly etched and packaged as 12mm x 170mm devices. | Packaging process control | Zimmer sent an URGENT MEDICAL DEVICE RECALL notification dated June 11, 2013, to all consignees via email and courier describing the affected device and problem and the risks associated with using the affected device. The notifications provided instructions for returning the affected device to Zimmer Product Service Department, 1777 West Center Street, Warsaw, IN 46580. Any questions or concerns concerning the recall, please call the customer call center at 1-877-946-2761. Customers may report adverse events or quality problems experienced with the use of these products to Zimmer by email at zimmer.per@zimmer.com.For questions regarding this recall call 800-613-6131. |  | 1 | 1 | Nationwide Distribution including IL and MO. | Terminated | January 30, 2014 | 233 | 1 | Not_Computer
Z-1675-2013 | 65439 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | AngioDynamics RITA Main Cable, 9 Pin to 14 Pin, REF 700-101892. RITA Main Cables are indicated for use with the RITA RF Generators and RITA electrosurgical devices. |  |  | June 11, 2013 | 2013 | Angiodynamics, INC | The RITA Main Cable product contains the incorrect Instructions for Use. | Packaging process control | Consignees were notified via certified mail on 06/11/2013 with an "Urgent Medical Device Field Safety Correction" notification. This included the correct Instructions for Use and the recommended actions. |  | 13 | 13 units | Worldwide Distribution, including the US, Austria, China, Italy, Russia and the UK. | Terminated | June 12, 2015 | 731 | 13 | Not_Computer
Z-1598-2013 | 65049 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | TOSHIBA Kalare Diagnostic X-Ray System |  |  | April 03, 2013 | 2013 | Toshiba American Medical Systems Inc | Firm initiated recall because when radiographic conditions with the small focus are set and the radiography ready switch is held down, the large focus is also heated after a few seconds or more. This occurs when the Automatic Exposure Function (AEC) function is used. As a result, a radiographic tube current exceeding the set value flows and a brighter image than expected is acquired. | Component design/selection | Firm initiated recall by sending recall notification letters along with a customer reply form on April 3, 2013 via USPS return mail receipt. A follow-up letter (with customer reply form) was sent on April 11, 2013 to inform customers that they should continue to use the AEC function as usual. The updated letter along with the customer reply form dated April 11, 2013 was also sent via USPS return mail receipt. The letter titled, "URGENT: MEDICAL DEVICE CORRECTION", notified customers of the reason for recall, product description under recall, corrective action, instructions to customer under 1. A new circuit board will be installed in each involved unit by a Toshiba service representative, without any charge to the user.2. In the meantime while waiting for service, the user is to continue usingthe device with the AEC function. If the AEC function is not available, theuser should push the radiography READY and EXPOSURE switches atthe same time."Note to Customer", and contact information "(800) 421-1968 or your local Toshiba Representative at (800) 521-1968". |  | 153 | 153 units | US Nationwide Distribution | Terminated | February 05, 2016 | 1038 | 153 | Not_Computer
Z-1692-2013 | 65554 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | GEO-MED ORTHO TOTAL JOINT TRACECART(R), REF 53-1831.07, Rx Only, NON-STERILE |  |  | May 30, 2013 | 2013 | DeRoyal Industries Inc | DeRoyal recalled kits that contained Stryker togas which were recalled for a potential problem with sterilization. | Nonconforming Material/Component | DeRoyal sent an Urgent: Voluntary Product Recall letter dated May 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory, remove any affected product from their inventory, open its package to prevent use, and dispose in a standard trash receptacle. Customers were asked to complete the enclosed notice of destruction form and return it to DeRoyal by faxing it to 865-362-3716 or emailing it to recalls@deroyal.com. For question call 865-362-1034 or contact your DeRoyal representative. |  | 87 | 87 units | USA Nationwide Distribution in the state of PA | Terminated | July 12, 2013 | 43 | 87 | Not_Computer
Z-1695-2013 | 65459 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 154200 Oxford Fixed Bearing Partial Knee |  |  | June 07, 2013 | 2013 | Biomet, Inc. | Biomet investigation found the design of the tibial bearing creates the potential for anterior impingement between the bone and the poly bearing component in terminal extension. | Device Design | The firm, Biomet, sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letter dated June 7, 2013 to its consignees/customers. The letter identified the device being recalled, the reason for the recall, potential adverse events associated with the recall and actions to be taken. The consignees/customers were instructed immediately locate and remove the identified device(s) listed in the letter; carefully follow the instructions on the enclosed "FAX Back Response Form"; complete and return the FAX Back Response Form to 574-372-1683 prior to return of product within three (3) business days; return Goods to Biomet, Inc 56 East Bell Drive Warsaw, IN 46580. To confirm receipt of this notice by call 574-372- 1570. Obtain an RGA # by phone (574)372-6677 email rqarequest@biomet.com or through FAST; use priority carrier for your shipment; and if you have further distributed this product you MUST notify personnel of the letter. Questions should be addressed to 574-372-1570 M-F 8 a.m. to 5 p.m. |  | 26 | 26 | Worldwide distribution: USA including NC, IN, PA, and OH; and countries of: Canada and Netherlands. | Terminated | July 17, 2014 | 405 | 794 | Not_Computer
Z-1686-2013 | 65590 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Si Surgeon Side Cart, model number SS3000; |  |  | June 08, 2013 | 2013 | Intuitive Surgical, Inc. | Factory testing on da Vinci Si Surgical Systems may not be in compliance with UL Standards as one of the testing devices was found to be working incorrectly. | Software design | Intuitive Surgical sent an Urgent Device Correction notice dated July 15, 2013, to all affected customers by Federal Express.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that all affected personnel were fully informed of the notification and that the notification should be forwarded to other managers within their facility. Customers were also instructed to complete and return the attached Acknowledgment Form to acknowledge receipt of the notification. Customers with questions were instructed to call Customer Service at 800-876-1310, Option 3.For questions regarding this recall call 408-523-2244. |  | 30 | 30 total, all devices | Worldwide Distribution - USA including UT, CA, AZ, CO, FL, HI. MT, NY,TX and WA and one customer in Russia. | Terminated | January 22, 2014 | 228 | 30 | Not_Computer
Z-1693-2013 | 65090 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Service Replacement 0-Ring. |  |  | May 24, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Incorrect rubber nitrile O-rings distributed instead of the correct material EPDM. (Ethylene Propylene Dine Monomer (M-class) rubber) | Nonconforming Material/Component | Fresenius notified customers on May 24, 2013, by telephone and by certified mail with signature confirmation and fax-back form. Customers were instructed to examine their stock immediately to determine whether they have any of the affected parts on hand. If customers have the affected parts, they will be instructed to immediately discontinue use and place all 0-rings in a secure area for return to Fresenius Medical Care North America.Please contact your Fresenius Medical Care Technical Service Team at 1-800-227-2572, (choose option 4, then option 2) for instructions how to return the recalled product. |  | 725 | 725 units | Nationwide Distribution including the states of AR, CA, FL, GA, KY, LA, MA, MD, NC, NJ, NY, PA, PR, SC, TX and VA. | Completed |  |  | 725 | Not_Computer
Z-1685-2013 | 65503 | No | LRZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Integra Hallu-Fix cannulated screw driver (model number P/N 129733ND) and cannulated screwdriver tip (model number P/N 129735ND). |  |  | February 01, 2013 | 2013 | Integra LifeSciences Corp. | The firm has received numerous complaints of the cannulated screw driver tip breaking during surgery. | Nonconforming Material/Component | The firm sent recall notification letters to all consignees via traceable email or traceable fax on 02/01/13. The firm requested that the consignees fill out and return the Recall Acknowledgement and Return Form. |  | 221 | 154 screwdrivers were distributed; 67 screwdriver tips were distributed. | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NV, NY, OH, PA, RI, SC, TN, TX, UT, VA, and WA and the countries of Canada, Austria, Belgium, Switzerland, Czech Republic, Germany, Denmark, Spain, Finland, France, Ireland, Israel, Italy, Lebanon, Luxembourg, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Sweden, United Arab Republic, United Kingdom, and South Africa. | Terminated | July 25, 2013 | 174 | 221 | Not_Computer
Z-1694-2013 | 65487 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Siemens Mammomat Inspiration - full-field digital mammography |  |  | February 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Firm became aware of an unintended behavior when using the Mammomat Inspiration. The possibility exists that a malfunction may occur during Stereo or Tomo acquisitions if the tomo clutch was incorrectly adjusted. Although the possibility is rare, the system could move while a patient's breast is compressed resulting in a possible bruise. Firm released Update Instructions XP031/12/S for this iss | Device Design | Siemens issued an Update Instructions XP031/12/S by letter to the affected customers. For affected sites, a Siemens Service Engineer will inspect the unit and make any adjustments needed. Following the inspection and adjustments, a proper calibration will be completed. For questions call 610-448-3237 and 610-446-4634. |  | 3 | 3 | United States Nationwide Distribution in the states of ND, OH and SC. | Terminated | February 21, 2014 | 358 | 3 | Not_Computer
Z-1725-2013 | 65148 | No | MNR | Breathing frequency monitor. | Anesthesiology | Anesthesiology | 510(k) | Finger Sensor for Remmers Recorder (Formerly SnoreSat Sleep recorder) |  |  | May 08, 2008 | 2008 | SagaTech Electronics, Inc. | The outer jacket is receding away from the black sensor prematurely. | Process design | SagaTech Electronics Inc, sent a letter to their Customers dated May 2008. The letter identified the product the problem and the action needed to be taken by the customer.Please take particular attention when cleaning these sensors by holding the cable somewhere in the middle and clean towards each end. This should become standard practice with all sensors to prolong their lifespan.If you have further questions please call (403) 228-4214 ext. 234. |  | 27381 | 152 units 77 devices from lot 13367 and 75 devices from lot 13710 | Worldwide Distribution: US (nationwide)Âżand countries of: ÂżCanada and Mexico. | Terminated | July 22, 2013 | 1901 | 27381 | Not_Computer
Z-1726-2013 | 65670 | No | MQN | Bone plate. | Dental | Dental | 510(k) | Synthes CMF Distraction System |  |  | June 12, 2013 | 2013 | Synthes USA HQ, Inc. | The firm initiated a voluntary recall of the BC Distractor Body which is part of the Craniomaxillofacial (CMF) Distractor System, due to a mis-alignment issue. There have been instances reported within the impacted lots in which the slot in the barrel of the BC distractor was rotated such that it was not possible to attach the B-type and C-type foot plates. | Nonconforming Material/Component | SYNTHES sent an Urgent Notice: Medical Device Recall letter dated June 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to examine their inventory for affected product, remove it from use, and return it to Synthes. If they have affected product they were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Complete the Verification Section at the end of the letter and indicate the number of devices found and note the Return Authorization Number and return to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If customers did not have the affected product they were asked to complete the Verification Section and faxt to 610-251-9005. Customers with questions were instructed to call 610-719-5450 or e-mail FieldAction@synthes.com.For questions regarding this recall call 610-719-5450. |  | 10 | 10 | Nationwide Distribution including PA, CT, DE, and NY. | Terminated | September 21, 2015 | 831 | 10 | Not_Computer
Z-1727-2013 | 65163 | No | JMM | Hydroxyproline test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | B-CrossLaps/serum (B-CTx in serum) |  |  | April 25, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche Diagnostics internal investigations in R&D; have shown that the claims for biotin interference and the sample stability given in the Elecsys Âż-CrossLaps (Âż-CL) assay package insert version 5, could not be reproduced. The Âż-CrossLaps test results contribute to the monitoring of antiresorptive therapy. False low or false normal results may lead to delayed diagnosis and treatment of osteoporosi | Device Design | The firm, Roche, sent an "URGENT MEDICAL DEVICE CORRECTION" (UMDC)letter dated May 2, 2013, to its consignees/customers via UPS Ground. The letter describes the product, problem and actions to be taken. The consignees/customers were instructed to apply the correct sample storage conditions as stated in the UMDC and to inform all relevant parties about the correct biotin interference claim. The consignees/customers were instructed to keep the UMDC on file for reference until the revised package insert is available. The consignees/customers were asked to complete and return a form via fax (1-888-943-4205). Consignees/customers were instructed to call Roche Diagnostics Technical Support, 24 hours a day, seven days a week, at 1-800-428-2336 if they had any questions about the recall. Non-responding accounts are monitored on an ongoing basis. |  | 3056 | 3056 units | Nationwide distribution: USA including states of: AL, CA, GA, FL, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NY, OH, PA, TX, VA, VT, and WA. | Terminated | September 27, 2013 | 155 | 3056 | Computer
Z-1730-2013 | 60480 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System 500 with Arterial and Venous blood parameter modules and Hematocrit/Saturation probe. |  |  | February 27, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo discovered through its own testing that CDI 500 Blood Parameter Monitoring System software contains inaccuracies associated with a software algorithm, and a software update has been developed to address this issue and a premarket notification (K130425) has been submitted to the agency. Once clearance is received for this premarket notification from the US FDA, Terumo CVS will implement the | Software design | On 2/27/13, all consignees will receive a recall notification letter via Federal Express. Terumo CVS will correct all affected CDl Monitors. Correction to the devices will be completed concurrently with BPM replacements related to Z-2185/2189-11 RES 57194.Terumo continues to manufacture and ship devices affected by this recall. The long term corrective actions are associated with Recall# Z-2185/2189-11 RES 57194. As committed previously, we will re-notify all customers once we are ready to implement the fix for the BPM in the field.All units produced prior to the change in manufacturing will be included in thescope of this correction. Terumo continues to manufacturer units affected by this correction.All units returned for this correction will be tagged with a service traveler to indicate reason for correction and units will be properly segregated to prevent unintended re-shipment. |  | 15888 | 3972 units since 4/26/12 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV and the countries of Jordan, CANADA, AUSTRALIA, Japan ,Hong Kong ,UNITED ARAB EMIRATES (UAE), BELGIUM, South Korea, Malaysia, Mexico, Philippines, Singapore, Taiwan, and Thailand. | Terminated | February 17, 2016 | 1085 | 15888 | Computer
Z-1728-2013 | 65505 | No | GBA | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 27023WU Balloon Catheter, 3 Fr., 2/pkg, sterile, contains latex. |  |  | May 31, 2013 | 2013 | Karl Storz Endoscopy America Inc | Karl Storz Endoscopy-America has initiated the recall of the 27023WU Balloon Catheter because the packaging may be compromised, and therefore the sterility of the product could be affected. | Packaging | The firm, Karl Storz, sent a "A SAFETY ALERT-ATTENTION NEEDED" letter dated May 31, 2013 to all customers who purchased the 27023WU Balloon Catheter. The letter described the product, problem and actions to be taken. The customers were instructed to check their inventory and to return the recalled product back to Karl Storz.. The customers were also instructed to complete and fax back the Recall Response Form attached with the safety alert letter at Fax #424-218-8559. If you have any questions, please feel free to call 424-218-8289. |  | 40 | 40 units | US distribution to: AZ, FL, IL, IN and TX. | Terminated | July 15, 2013 | 45 | 40 | Not_Computer
Z-1729-2013 | 65664 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Safety Lancets - Odyssey, Button Activated; |  |  | July 03, 2013 | 2013 | KooJoo Trading Company | Lancets were not sterilized within the dose range established by sterilization validation. | Process control | The firm, KooJoo, sent a "MEDICAL DEVICE VOLUNTARY T RECALL" letter dated July 3, 2013 via first class mail to its customer. The letter described the product, problem and actions to be taken. The customer was instructed to identify and hold all safety lancets you have in inventory that was received before June 2013; notify your retail customers of this product recall and complete and return the attached Customer Reply Form via email to heegonkim@hotmail.com and also return form along with any unused/undistributed product for replacement.If you have further questions, please email heegonkim@hotmail.com or call 209-403-9191 during 9:00am to 5:00pm M-F PST. |  | 7 | 7, 712, 000 pieces (1928 cartons) | US Distribution GA only. | Terminated | March 13, 2015 | 618 | 7 | Not_Computer
Z-1735-2013 | 65088 | No | KIS | Animal and human sera. | Hematology | Pathology | 510(K) Exempt | Various types of Fetal Bovine Serum |  |  | April 12, 2013 | 2013 | Paa Laboratories Inc | Current product labeling band data sheets state that FBS Pharma Grade and FBS/FBS Nativeare pure Fetal Bovine Serum (FBS). These products may contain added adult Bovine SerumAlbumin (BSA)of United States origin, water, and/or cell growth promoting additives. All otherFBS products, as listed below, including FBS Gold and FBS Standard, described as specialtysera', may contain adult BSA of Un | Under Investigation by firm | GE Healthcare notified customers via USPS on April 12, 2013 and April 14, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers being contacted by phone as well.Customers were asked to contact their local customer service representative if the product does not meet their needs. Customers were instructed to call (732) 457-8455 or 301-660-0256. For questions regarding this recall call 609-514-6888. |  | 280000 | approximately 280,000 liters | Worldwide Distribution - USA (nationwide), Novia Scotia, and Canada | Terminated | September 14, 2015 | 885 | 280000 | Not_Computer
Z-1739-2013 | 63109 | No | LGW |  |  | Neurology | PMA | Eon Mini - 3721 (Charger). |  |  | July 26, 2012 | 2012 | St. Jude Medical | St. Jude Medical has informed your doctor that a number of cases have been reported in which discomfort associated with heating occurred at the device site while patients were using the Charging System to charge their spinal cord stimulator. In three cases, patients received a burn to the skin (one second-degree and two first-degree burns) at the implant site. As of June 30, 2012, St. Jude Medica | Device Design | ST. Jude Medical sent an Important Medical Device Safety Information dated July 26, 2012 to affected customers. The letter identifed the affected product, problem and actions to be taken. For questions or concerns contact your ST. Jude Medical representative or St. Jude Medical Neuromodulation Division at 1-866-240-6741. |  | 70638 | 70,638 (combined total of the Eon and the Eon Mini systems) | Worldwide Distribution - USA Nationwide Including Puerto Rico and countries of: Argentina, Australia, Belgium, Brazil, Canada, China, Chile, Colombia, Ireland, Israel, India, Malaysia, Mexico, Saudi Arabia, Singapore, South Africa, South Korea, Switzerland and United Kingdom. | Terminated | November 18, 2015 | 1210 | 70638 | Not_Computer
Z-1690-2013 | 65420 | No | LKK |  |  | General Hospital | PMA | MedStream Programmable Infusion Pump- Implantable for the intrathecal delivery of Baclofen or Morphine |  |  | June 07, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Miscalibrated Fill Level Sensor may affect dosing level | Device Design | Codman issued a Medical Device Correction Notification to Medical Facilities dated June 7, 2013 via FedX to inform them of the potential problem and a Worksheet (instructions) to Identify Pumps with Miscalibrated FLS during each patient's next scheduled refill session or sooner if patient is symptomatic. If a pump is identified as having a miscalibrated FLS, notify Codman Neuro using the same worksheet. In addition a sample letter was provided to Physi cans to provide to patients on the system.Follow the recommendations for managing patients who are identified as having a pump with a miscalibrated FLS.For additional information,contact your Codman Neuro representative, or contact Codman Neuro Clinical Support at 1-800-660-2660. Please report any malfunction or adverse event related to this device to Codman Neuro at 1-866-491-0974, Option 2. |  | 737 | 2 units US and 735 OUS | Worldwide Distribution-USA (nationwide) including the states of FL, NC, NJ, and TX, and the countries of Austria, Belgium, Spain, Great Britain, Germany, Italy, Ireland, Luxemburg, Russia, Saudi Arabia, Portugal, Greece, Egypt, Switzerland, and Turkey. | Terminated | April 21, 2015 | 683 | 1102 | Not_Computer
Z-1737-2013 | 64989 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | KRD1 PedFuse Rods, Model: IM80100-13 Rev B |  |  | December 30, 2011 | 2011 | SpineFrontier, Inc. | One of the rods in each of two test constructs did not meet fatigue requirement as outlined in ASTM F1717-04 dynamic compression testing 5 million cycles. | Vendor change control | SpineFrontier issued a Recall Notification to the affected distributor via telephone. The distributor was instructed to return the affected product for a replacement product. |  | 2 | 2 rods | USA Nationwide Distribution including the state of TX | Terminated | January 08, 2014 | 740 | 2 | Not_Computer
Z-1742-2013 | 65456 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | RemB Electric Universal Driver REF 6400-099-000 For use with the CORE and TPS System sRx Only |  |  | July 01, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | The safety margin values detailed in the Engineering Design were entered incorrectly into the programming software. Handpieces that have been programmed with the incorrect parameters could result in unintended activation. There is a potential for injury to the patient or user as a result of unintended forward or reverse activation. There is also the potential that the handpiece may not reach 100 | Process change control | The firm, Stryker Instruments, sent emails beginning May 30, 2013 "URGENT MEDICAL DEVICE CORRECTION" letter dated July 1, 2013 to its customers. The letter identified the device, the reason for the recall, possible adverse events and actions to be taken. The customers were instructed to immediately review this letter, locate the units listed, complete and return the business reply form via FAX to 866-521-2762; or email Kara.spath@stryker.com; if you have further distributed this product, please forward this letter and the attached Business Reply Form to all affected locations. After receiving this form, Stryker will send you a pre-paid shipper to send your affected device(s)back to Stryker to be reprogrammed with the correct specifications. Send back all affected devices using the pre-paid shipper provided to you by Stryker. If you require a loaner, please contact the Stryker Service Department to arrange for a loaner at 888-308-1983. For questions regarding this recall please Stryker Instruments: Monday-Friday 8am-5pm (EST) 269-389-4518 / 269-389-3808 kara.spath@stryker.com / jennifer.mars@stryker.com. |  | 9 | 9 | Worldwide distribution: US (nationwide) including states of: AZ, FL, GA, NC, NY, SC, TX and VA; and countries of: India, Netherlands and Japan. | Terminated | September 16, 2013 | 77 | 20 | Computer
Z-1736-2013 | 65406 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Integra PANTA Arthrodesis Nail System. |  |  | May 09, 2013 | 2013 | Integra LifeSciences Corporation | As a result of the investigation into an adverse trend of complaints for reports of misalignment of the calcaneal or tibial screws during use of the Panta Nail Support Device, Integra has decided to issue this voluntary medical device correction. The misalignment can result in the inability to properly insert the screws into the Panta Nail and create the potential need for revision surgery. | Device Design | Integra initiated a voluntary correction on May 9, 2013 by providing a written notification and copy of the revised Surgical Technique Manual as well as the specific highlighted sections of the manual addressing calcaneal and tibial screw alignment to all consignees that Integra records indicate have been shipped a PANTA Nail. The written notifications and Surgical Technique Manual have been delivered to US Consignees by traceable courier service and by traceable email or facsimile for Canada and Non-US Consignees. The notification informs customers of the nature of the issue; advises of potential adverse patient consequences that may occur; advises customers to review highlighted revisions of the Surgical Technique Manual to minimize the potential for misalignment of the screws and requests customers to complete and return the attached Acknowledgement Return Form. |  | 692 | 692 | Worldwide Distribution-USA (nationwide) and the countries of Austria, Australia, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Ireland, Israel, Italy, Netherlands, Portugal, Russia, Saudia Arabia, South Africa, Spain and United Arab Emirates. | Terminated | November 07, 2013 | 182 | 692 | Not_Computer
Z-1748-2013 | 65303 | No | OOY |  |  | Anesthesiology | PMA | Alair Bronchial Thermoplasty Catheter, Bronchial Thermoplasty System; Material Number: M005ATS25010, Catalog Number: ATS 2-5. Product is manufactured and distributed by Boston Scientific Corporation (BSC), Sunnyvale, CA |  |  | May 28, 2013 | 2013 | Boston Scientific Corporation | Boston Scientific has discovered an inconsistent "Use By" expiration date on a single lot/batch of Alair Bronchial Thermoplasty Catheters. The expiration date indicated on the inner tray, 2012-04, is incorrect. The correct expiration date is 2014-04 as indicated on the outer box label. | Error in labeling | The firm, Boston Scientific, sent a letter dated May 28, 2013, via express mail to all its direct consignees. The letter described the product, problem and actions to be taken. The customers were informed that you can continue to use the device until the expiration date on the outer box lablel, 2014-04; No product is being recalled and the you are not required to return product to Boston Scientific. The customers were instructed to read the enclosed information carefully and post the information in a conspicuous location near the affected product and/or take such other steps, consistent with your internal policies and procedures, to ensure that the information is easily accessible to all users of this device; and to complete and return the Customer Acknowledgment Form even if you no longer have any units from the affected lot/batch via email: MapleGroveFieldAc@bsci.com or Fax to: Field Action Center - 1866-213-1806. If you have any questions, call 763-494-1133 or email: MapleGroveFieldAc@bsci.com . |  | 35 | 35 units | Worldwide distribution: US (nationwide) including states of: CA, FL, IA, ID, MI, OK, TA and WA; and countries of: Canada and Thailand. | Terminated | August 28, 2013 | 92 | 35 | Not_Computer
Z-1749-2013 | 65492 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Suction Wand, models S099 and S0998B |  |  | June 19, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards is recalling certain lots of Rigid Suction Wands because they discovered plastic particulate inside the sterile pouch and within the internal diam ether of the wand. | Component design/selection | Edwards sent an Urgent Field Safety Notice Product Recall letter, dated June 19, 2013, via FedEx to all affected customers. Two additional attempts will be made to notify customers to return products. The letter idenditifed the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory for any affected device in their inventory, quarantine all affected product from their inventory and return all devices to Edwards. Customers were instructed to acknowledge that they have reviewed the Recall Notice and confirm that they have taken appropriate action by completing, sign and dating the enclosed Recall Response Form and returning the form to Edwards by fax to Edwards Customer Service at 800-422-9329. Customers should contact Customer Service at 800-424-3278 to obtain a Returned Goods Authorization number to return their current product to and receive replacement product. Customers with questions were instructed to call Edwards Customer Service at 800-424-3278. |  | 676694 | S099: 171,587 units and S099B: 505,107 units | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | November 04, 2014 | 503 | 676694 | Not_Computer
Z-1750-2013 | 65620 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens AXIOM Artis zee/zeego systems equipped with generator A100 Plus |  |  | January 31, 2013 | 2013 | Siemens Medical Solutions USA, Inc | During the course of product monitoring, Siemens became aware of a tolerance problem in the power supply of the generator control above a specific value, thus, the generator may fail. | Device Design | Siemens sent an Urgent Customer Information letter to their consignees. The letter identified the affected product, problem , update Instructions and actions to be taken. Siemens began to equip affected systems with a modified board assembly in January, 2013. |  | 23 | 23 | United States Nationwide Distribution - including the states of: CA, CO, FL, IA, IL, IN, MA, MD, MI, NC, NY, PA, TX, VA and WA and Puerto Rico. | Terminated | November 12, 2014 | 650 | 23 | Not_Computer
Z-1745-2013 | 65561 | No | DZL | Intraosseous fixation screw or wire. | Dental | Dental | 510(k) | Multi Vector Distractor Pin Holding Clamp, limited bone stock |  |  | May 31, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes is initiating a voluntary Medical Device removal for specified part numbers and lot numbers of the Mutis Vector Distractor Pin Holding clamp, limited bone stock. The firm received complaints that the clamp and nut construct would not engage the distractor arm threads and subsequently bind. It is not likely that a serious adverse event will result due to a nonconforming clamp, however there | Device Design | Synthes sent a Urgent Notice: Medical Device Removal letter dated May 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. The firm received complaints that the clamp and nut construct would not engage the distractor arm threads and subsequently bind. It is not likely that a serious adverse event will result due to a nonconforming clamp, however, there is a remote potential that additional intervention would be required due to a non-functioning clamp.Synthes requested the product to be removed from inventory and returned to Synthes by calling 1-800-479-6329 to obtain a Return Authorization Number. Direct Accounts were asked to complete and return the Verification Section at the end of the letter indicating the quantity of affected product was found. Negative responses were also requested.If you have any questions please call (610) 719-5450. |  | 121 | 121 | US Nationwide Distribution including the states of MA, TX, MN, SC, UT, VA, and NY. | Terminated | August 31, 2015 | 822 | 121 | Not_Computer
Z-1743-2013 | 65652 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TDHistology/Cytology |  |  | June 20, 2013 | 2013 | Technidata S.A. | A recall was initiated by Technidata because use of symbols (less than or equal to or greater than or equal to or any other symbol) in Microsoft Word reports when using electronic transmissions of results by of HPRIM (including HPRIM Medecin via Kermit) or HL7 must be absolutely prohibited. Such symbols are not supported by the Microsoft Word interface used by the Histology/Cytology software and | Software design | TD Technidata Medical Software sent a Medical Device Recall letter dated June 20, 2013, to one affected customer who received the TD-Synergy Histology/Cytology module. The customer was informed of the problems identified and the actions to be taken. Customers were instructed to complete the attached Reply Form and return it to the firm by fax to 514-270.2770. The customer were also instructed to contact Technidata Canada at (514) 270-7777 or at support-canada@technidata-web.com. |  | 1 | 1 site | US Distribution CA only. | Terminated | July 29, 2013 | 39 | 1 | Computer
Z-1747-2013 | 65558 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Afinity CQ2015A Intraocular Lens |  |  | June 10, 2013 | 2013 | Staar Surgical Co. | STAAR Surgical Company is recalling a limited number of nanoFLEX CC4204A and Afinity CQ2015A intraocular lenses due to the potential presence of glass particles in the storage vial. | Material/Component Contamination | The firm, STAAR Surgical, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated June 10, 2013 to all the customers that received a limited number of nanoFLEX CC4204A and Afinity CQ2015 Intraocular lenses. The letter described the product, problem and actions to be taken. The customers were instructed to check your inventory; remove any lenses from the attached list; contact STAAR Surgical Customer Service for Returns Goods Authorization Number via phone: 1-800-352-7842; email: customerservice@staar.com; and/or fax: 1-800-952-4923; return listed lenses to STAAR Surgical; complete and return the attached response form by fax to Customer Service at (800) 952-4923, and pass this notice on to all those who need to be aware within your organization or any organization where the devices have been transferred.If you have any questions or concerns regarding the product return, contact STAAR Surgical Customer Service at (800) 352-7842. |  | 3361 | 3,361 units total | Nationwide distribution: USA including states of: AL,AZ, CA, CO, CT, FL, GA, HI, KY, IA, IL, IN, LA, MA, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, PA, SC, TN, TX, UT, VA, WI, and WV. | Terminated | October 17, 2013 | 129 | 3361 | Not_Computer
Z-1744-2013 | 65355 | light, microscope | IBJ | Microscopes and accessories. | Pathology | Pathology | 510(K) Exempt | Polarizer D Fixed w/Rotating Red 1 lambda plate. |  |  | May 28, 2013 | 2013 | Zeiss, Carl Inc | Carl Zeiss Microscopy, LLC is recalling the Polarizer D component used in All upright ZEISS microscopes (Axio Imager, Axio Scope.A1, Axio Lab.A1) due to Misaligned Full-Wave (Lambda) Plate. | Nonconforming Material/Component | On May 28, 2013, Carl Zeiss Microscopy, LLC recalled a Polarizer Component due to Misaligned Full-Wave (Lambda) Plate. Carl Ziess Microscopy distributed notices to their customers via e-mail. Carl Zeiss Microscopy, LLC requests that all customers immediately discontinue the use of the component for any applications which rely on accurate color-change sequence results. Your service provider will contact you to schedule a visit in which a factory trained service technician will confirm the alignment of your Polarizer D component;s lambda plate. If it is defective, a new Polarizer D component will be installed. There will be no charge for this action. Should you have any questions related to this issue, please do not hesitate to the Carl Zeiss facility in New York by telephone 914-681-7815 or via email at ed.mancini@zeiss.com |  | 10 | 10 | Distributed in the states of MI, TX, SD, FL, GA, NY, and NY. | Terminated | October 29, 2014 | 519 | 10 | Not_Computer
Z-1753-2013 | 65176 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Hydra Vision DR 60/80 System, Models 700539 and 700540, Liebel-Flarshiem Company. The System is used in radiologic imaging procedures. |  |  | April 25, 2013 | 2013 | Leibel-Flarsheim Company LLC | A complaint was received in which a hospital physicist indicated the dosage was too high on the HUT system. During the service call the systems Maximum Fluoro Dose was found to be 30 R/min which is three times the Maximum Fluoro Dose limit of 10 R/min. | Employee error | The customers received an "Urgent Device Correction" notification letter on May 12, 2013 explaining the product issue and providing recommended actions. Customers were contacted to set up a service call to test and recalibrate equipment as necessary. Questions were directed to 800-778-7898. |  | 12 | 12 units | Nationwide Distribution (US), including the states of TN, KY, CO, KS, CT, HI, and AZ. | Terminated | July 26, 2013 | 92 | 12 | Not_Computer
Z-1752-2013 | 65215 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ExactaMix 1200 and 2400 Compounders. |  |  | June 04, 2013 | 2013 | Baxa Corporation | Baxter Healthcare is recalling certain ExactaMix automated compounding devices due to the complementary metal oxide semi-conductor battery failure resulting in resetting the internal clock of the device to midnight December 31, 2010. | Nonconforming Material/Component | On 6/04/13, customers were notified by 1st class mailing with mitigating instructions including reminding customers that the battery failure results in a failure of the ExactaMix system to alert the user to replace the disposable set after 24 hours of use. Users should not rely on the alert to trigger the daily replacement of disposable sets. Users should also confirm the correct time and date at the setup of each use. If the compounder displays an incorrect time and date, customers are to contact Baxter Technical Support. |  | 753 | 753 units | Worldwide Distribution-USA (nationwide) including Puerto Rico, and the countries of Canada, Latin America, Europe/Middle East/Africa, and Taiwan. | Terminated | October 16, 2015 | 864 | 753 | Not_Computer
Z-1751-2013 | 65647 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | LifeShield; Sterile; Rx only; List No. 19666-28; SYMBIQ PUMP SET, CONVERTIBLE PIN, PIGGYBACK, 2 CLAVES PORTS AND OPTION LOK, NON-DEHP; NOMINAL LENGTH: 97 IN (249 CM); PRIMING VOLUME: 14 ML; For the administration of I.V. fluids; Made in Costa Rica; Hospira, Inc., Lake Forest, IL 60045 USA |  |  | June 18, 2013 | 2013 | Hospira Inc. | Hospira has received reports of customers experiencing air-in-line alarms related to administration sets that have a drip chamber with a ball valve. Preliminary investigation findings found that there is a potential for the ball valve to not seal the drip chamber as intended, leading to air in the line of the administration set. | Under Investigation by firm | Hospira sent a Urgent Device Recall letter dated June 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customer were instructed to check their inventories, quarantine any affected products, complete an attached reply form, and return the product to Stericycle for credit using a label included in the mailing. Customers who further distributed the products are requested to forward the recall notification information to their customers.A 100% effectiveness check will be performed for direct Hospira customers who do not return the reply forms and/or the product. Any direct or indirect accounts that return a reply form indicating they have product but do not return it will also be contacted as part of the effectiveness check. In addition, customers who further distributed the product will be requested to provide confirmation that they have notified their customers who purchased the product.Please contact Hospira Customer Care at 1-877-946-7747 ( M-F 7am - 6pm CT) or your Hospira representative regarding replacement product availability and for questions regarding this recall. |  | 14400 | 14,400 units | Worldwide Distribution - US (nationwide) including the states of IA, IL, MN, SD and WA., and Costa Rica. | Open, Classified |  |  | 14400 | Not_Computer
Z-1771-2013 | 65630 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Spacelabs ARKON Anesthesia Workstation. |  |  | June 07, 2013 | 2013 | Del Mar Reynolds Medical, Ltd. | Spacelabs ARKON Anesthesia System is recalled due to software defect. The System Status Computer may incorrectly determine that a communication error has occurred. If this situation occurs, a buzzer will activate and a yellow triangle will be displayed on the System Status Computer screen that alerts the user of this error. | Software design | Spacelabs Healthcare sent an Urgent Medical Device Correction" letter, dated June 20, 2013, to all affected consignees. The letter identified the product, th problem, and the action to be taken by the consignee. Consignees were instructed to advise their staff of the situation. Consignees were also advised that Spacelabs Field Service personnel would be contacting them at their facility to schedule installation of a software update to resolve the issue at not cost to the consignee. Consignees with questions were instructed to call Technical Support at 1-800-522-7025 and select option 2.For questions regarding this recall call 800-522-7025. |  | 9 | 9 units (4 in the US and 5 outside US) | Worldwide Distribution - USA including South Carolina and internationally to Canada and Great Britian. | Terminated | August 29, 2013 | 83 | 9 | Computer
Z-1778-2013 | 65488 | No | FTC | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dermalight 80 Model Number UVB-080, one unit per box. |  |  | April 09, 2013 | 2013 | National Biological Corp | The incorrect manual may have been packaged with this product. UVB-080 devices may have included the manual for a UVA-083 device. | Labeling mix-ups | The firm, National Biological, sent an "VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated May 10, 2013 to its consignees/customers by UPS or other delivery service with a signature required. The letter described the product, problem and actions to be taken. Consignees/customers who do not respond to the letter will be contacted via telephone. The customers were instructed to check within your household to determine if you have any of the listed product; immediately destroy the incorrect User Manual, it is not necessary to return the product; record the Customer Name and or the Dermalight-80 serial number and complete and return the enclosed Verification Form, even if you do not have any User Manuals to destroy.Please contact your National Biological Corporation sales representative if you have any questions regarding this recall, any of our products, or would like assistance with the recall call 216-831-0600 or 800-338-5045 email: infor@PhototherapyExperts.com. |  | 124 | 124 units | Worldwide distribution: US (nationwide) including states of: AL, CA, CO, CT, FL, GA, IA, ID, IL, LA, MA, MN, MO, MS, NE, NJ, NY, OH, PA, SC, TN, TX, VA, WA, and WI; and countries of: Canada, Brazil, Malaysia, and Mexico. | Terminated | February 24, 2014 | 321 | 124 | Not_Computer
Z-1774-2013 | 65688 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | PASCAL and PASCAL Streamline (532 nm.577 nm); |  |  | March 27, 2012 | 2012 | Topcon Medical Laser Systems, Inc | When using the "Favorites" function for saved settings, there is potential for a mismatch between displayed power and the actual power output. This may result in over or under treatment. | Software design | Initial notification was sent on March 27, 2012 to all TMLS Distribution Partners and Sales Representative. This was followed by a follow up notification on June 4, 2012. A final notice was sent on February 11, 2013 to all those who had not sent documentation that the field action had been implemented. |  | 40 | 7 US, 33 foreign | Worldwide Distribution-USA and the countries of China, Hong Kong, Brazil, France, Japan, Netherlands, Pakistan, Malaysia, United Kingdom, South Korea, Thailand, Dubai, and South Africa. | Terminated | July 19, 2013 | 479 | 40 | Not_Computer
Z-1776-2013 | 65141 | No | GDO | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LIGACLIPÂż Endoscopic Rotating Multiple Clip Applier Convenience Kit (Flex Trays), 10mm M/L |  |  | April 12, 2013 | 2013 | Ethicon Endo-Surgery Inc | Ethicon Endo-Surgery is initiating a voluntary recall for LIGACLIPÂż 10mm M/L Endoscopic Rotating Multiple Clip Applier (ER320) due to potential clip formation and feeding issues which may result in improper clip formation, vessel damage and/or insufficient occlusion of the vessel or other structure | Under Investigation by firm | Ethicon sent an Urgent : Medical Device Recall letter dated April 30, 2013 to all primary and secondary consignees via overnight UPS to allow tracking of the receipt. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine inventory immediately and return affected products for credit. This credit is intended to facilitate replacement purchase. For questions contact your local representative or call Ethicon Customer Support Center at 1-877-384-4266 option 6. |  | 49357 | 49,357 | Worldwide Distribution - USA Nationwide to all fifty (50) states, and internationally including the following countries: ARGENTINA, ARGENTINA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CYPRUS, ECUADOR, EGYPT, EL SALVADOR, GREECE, HONDURAS, INDIA, INDONESIA, ISRAEL, JAMAICA, JAPAN, JORDAN, KOREA, REPUBLIC OF, KUWAIT, LEBANON, MEXICO, NETHERLANDS ANTILLES, NICARAGUA, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, PUERTO RICO, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA and YEMEN. | Terminated | May 01, 2014 | 384 | 798201 | Not_Computer
Z-1779-2013 | 65494 | No | FTC | Ultraviolet lamp for dermatologic disorders. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dermalight 80 Model Number UVB-083, one unit per box. |  |  | April 08, 2013 | 2013 | National Biological Corp | The incorrect manual may have been packaged with this product. UVB-083 devices may have included themanual for a UVA-080 device. | Labeling mix-ups | The firm, National Biological, sent an "VOLUNTARY PRODUCT RECALL NOTIFICATION" letter dated May 10, 2013 to its consignees/customers by UPS or other delivery service with a signature required. The letter described the product, problem and actions to be taken. Consignees/customers who do not respond to the letter will be contacted via telephone. The customers were instructed to check within your household to determine if you have any of the listed product; immediately destroy the incorrect User Manual, it is not necessary to return the product; record the Customer Name and or the Dermalight-80 serial number and complete and return the enclosed Verification Form, even if you do not have any User Manuals to destroy.Please contact your National Biological Corporation sales representative if you have any questions regarding this recall, any of our products, or would like assistance with the recall call 216-831-0600 or 800-338-5045 email: infor@PhototherapyExperts.com. |  | 6 | 6 units | US distribution including the states of: AL, CA, MN, and NC. | Terminated | February 24, 2014 | 322 | 6 | Not_Computer
Z-1772-2013 | 65504 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Endopath Electrosurgery Probe Plus II, twelve (12) different devices, including hooks, spatulas, right angles, curved dissectors, suction/irrigation devices, accessory ports, pool/sump devices and stone retrieval cannula, packaged and sterilized individually in a transparent flexible blister sealed to a Tyvek lid. 6 flexible packages are placed into a sales unit carton, with 4 cartons per corrugated shipper. |  |  | April 08, 2013 | 2013 | Ethicon Endo-Surgery Inc | The firm is recalling the ENDOPATH Probe Plus II Shafts because the possibility that a tear in the Tyvek packaging may affect the sterility of the device. | Packaging | The firm, Ethicon Endo-Surgery, Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter dated April 22, 2013 to all consignees/customers via overnight UPS to allow tracking of the receipt. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to do the following: EFFECTIVE IMMEDIATELY-DO NOT USE ANY OF THE FOLLOWING TWELVE (12) PRODUCT CODES within the noted Expiration Dates provide in the letter; examine your inventory immediately to determine if you have affected product on hand; remove the affected product; complete and return the Business Reply Form (BRF) and fax to: 1-888-912-7089 or email to: ees3118@stericycle.com within 3 business days, even if you do not have affected product; return affected product with a photocopy of the completed BRF; and share this information with all of the appropriate staff at your facility. If you need additional shipping labels or a communications package, contact the Stericycle's support center at 1-888-943-4896 and reference Event 3118."Returned ENDOPATH Probe Plus II Shafts will be replaced with corrected ENDOPATH Probe Plus II Shafts once supply becomes available. If replacement product is not available and/or has been discontinued, a credit will be issue."If you need clinical or product support, please contact your local representative or call our Customer Support Center, 1-877-ETHICON (877-384-4266). |  | 327648 | 327,648 instruments | Worldwide distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, lA, ID, IL, IN, KS, KY, LA, MA, MD, ME, Ml, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY and countries of: Canada, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Canada, Chile, China, Columbia, Czech Republic, Dominican Republic, Ecuador, Egypt, India, Indonesia, Israel, Italy, Jamaica, Japan, Korea, Kuwait, Libya, Malaysia, Mexico, Pakistan, Panama, Philippines, Poland, Puerto Rico, Russia, Singapore, Slovenia, South Africa, Spain, Saudi Arabia, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Uruguay, &amp; Venezuela. | Terminated | December 24, 2013 | 260 | 327648 | Not_Computer
Z-1773-2013 | 65565 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | The LIFEPAK 500 Automated External Defibrillator is a semi-automatic defibrillator that uses a patented Shock Advisory System. This software algorithm analyzes the patient's electrocardiographic (ECG) rhythm |  |  | June 12, 2013 | 2013 | Physio Control, Inc. | The LIFEPAK 500 Automated External Defibrillator (AED) could be susceptible to malfunction due to an assembly error. If affected, the malfunction would be the result of internal shorting on a printed circuit board assembly and could render the device inoperable. | Employee error | The firm, Physio-Control Inc., contacted their consignees by telephone or email starting on June 13, 2013 and sent and/or delivered an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 2013 to its customers. The letter described the product, problem and actions to be taken. A Physio-Control representative will be contacting the customers by direct visit, telephone, or email you to coordinate the replacement of their device(s).Customers with questions regarding this notification, please call Physio's Technical Support at 1-800-442-1142, 6:00 a.m. to 4:00. pm. (Pacific), Monday-Friday. |  | 139 | 139 total (125 units in the US and 14 units outside US) | Worldwide distribution: US (nationwide) including states of: IL, MA, NC, ND, NY, PA, RI, TX, WA, and WI; and country of: Canada. | Terminated | September 19, 2013 | 99 | 139 | Not_Computer
Z-1781-2013 | 65542 | No | GXH | Cryogenic surgical device. | Neurology | Neurology | 510(k) | Myoscience brand iovera Smart Tip, Myoscience Cryo-Touch IV, a Cryogenic Surgical device, REF Catalog Number: STA0311-10. |  |  | June 21, 2013 | 2013 | Myoscience Inc | The outer box label of the product was labeled with a down revision label which did not include the US labeling requirements, specifically the symbol descriptors and also the "Caution: Federal Law (USA) restricts this device to sale by or on the order of a physician." This product was previously labeled for EU distribution only. | Labeling mix-ups | Myoscience representatives visited both affected sites on June 14, 2013 and relabeled the product in person. Neither site had used the product and was unaffected by the relabeling. No formal communication was issued. |  | 3 | 3 units | US distribution including CA and ID. | Terminated | July 15, 2014 | 389 | 3 | Not_Computer
Z-1780-2013 | 65622 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Radiohead Access Point 1.4 GHz Radiohead Access Point |  |  | July 08, 2013 | 2013 | Philips Healthcare Inc. | IntelliVue Smart-hopping 1.4 GHz Access Point may experience unexpected signal loss of severalminutes impacting wireless connectivity | Software Design Change | Philips sent an Urgent Medical Device Correction letter dated July 2013 to affected customers. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer. Philips will perform a firmware upgrade (version C.00..07 for all affected customers at no cost. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice.If you need any further information or support concerning this issue, please contact Philips Customer Support at 1-800-722-9377 |  | 20105 | 20105 units | USA Nationwide Distribution | Terminated | May 17, 2017 | 1409 | 20105 | Not_Computer
Z-1784-2013 | 65399 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | HemosIL RecombiPlasTin 2G, Part Number 0020002950 (8 ml size). The product is used for the evaluation of the extrinsic coagulation pathway and the monitoring of Oral Anticoagulant Therapy. |  |  | April 09, 2013 | 2013 | Instrumentation Laboratory Co. | Instrumentation Laboratory is recalling HemosIL RecombiPlasTin 2G (RPT2G), Part No. 0020009580 (8 mL Size) due to being Out of Specification. | Nonconforming Material/Component | Between April 9, 2013 and April 11, 2013, Instrumentation Laboratory distributed "Important Product Correction" notification letters to their U.S. and foreign customers informing them that the on-board instrument stability claim has been reduced to 3 days for all current and future product lots of HemosIL RecombiPlasTin 2G. A second notification to non-respondents was initiated and completed on May 7, 2013. Technical questions were referred to 1-800-678-0710, Option # 2. |  | 5553475 | Kits per product:1)11080, 2)11603, 3)11487, 4)11483, 5)11290, 6)11562, 7)11535, 8)11577, 9)11505, 10)11708,11)11579,12)11580, 13&amp;14) not distributed | Worldwide Distribution- US (nationwide) including states of: MA, ME, CT, NJ, NY, PA, MD, VA, NC, SC, GA, FL, AL, TN, MS, KY, OH, IN, MI, IA, WI, MN, SD, MT, IL, MO, KS, NE, LA, AR, OK, TX, CO, ID, AZ, CA, OR, WA, AK, RI, VT and the countries of: Angola, Argentina, Australia, Austria, Belgium, Bermuda, Brazil, Bulgaria, Canada, Caribbean, Chile, Colombia, Costa Rica, Czech, Denmark, Dubai, Ecuador, Egypt, France, Germany, Gibraltar, Great Britain, Greece, Guatemala, Hungary, India, Indonesia, Iran, Israel, Italy, Jamaica, Jordan, Korea, Kosovo, Kuwait, Lebanon, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad, Turkey, Ukraine, Uruguay, Vietnam, and Japan. | Terminated | May 06, 2015 | 757 | 5553475 | Not_Computer
Z-1786-2013 | 65573 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | STERIS Fine Traction Device, one unit per box |  |  | May 20, 2013 | 2013 | Steris Corporation | A bearing used in the assembly of the Fine Traction Device did not perform successfully in pre-release reliability testing; a replacement bushing passed reliability testing and was substituted per a change order. The change order authorized use of the substitute bushing in assembly of the Fine Traction Accessory; however, inadvertently the new part was not implemented. During complaint evaluati | Component change control | STERIS Account Managers notified affected customers via onsite visits. All consignees were notified as of 5/29/2013. Customers were instructed to return affected product for a replacement unit. |  | 16 | 16 units | USA Nationwide Distribution in the states of CO, KY, LA, VA, and WI. The affected product was also distributed to Canada. | Terminated | August 28, 2013 | 100 | 16 | Not_Computer
Z-1785-2013 | 63947 | No | LGW |  |  | Neurology | PMA | Eon and Eon Mini charging system, model numbers Eon -3701 (Wall Charger), 3711 (Portable Charger) and Eon Mini -3721 (Charger); This is a continuation of the firm's July 2012 recall. |  |  | November 30, 2012 | 2012 | St. Jude Medical | There is issues of warmth or heating at the implant site during charging for the Eon Mini implantable pulse generators (IPGs). In the July 2012 letter St. Jude Medical informed customers that they would be implementing design improvements to the charger to address possible increased energy dissipation when the charger is misaligned with the IPG or the IPG is implanted too near the surface of the | Device Design | St Jude Medical sent an Important Medical Device Safety Information letter dated November 30, 2012 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that a replacement charger for the Eon Mini IPGs is now available and will be provided at no charge. For questions or concerns call St. Jude Medical at 877-650-0367. |  | 75663 | 75,663 | Worldwide Distribution - USA Nationwide including Puerto Rico and countries of: Austria, Australia, Belgium, Brazil, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Ireland, Israel, Italy, Japan, Luxembourg, Martinique, Mexico, Netherlands, Poland, Russian Federal, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | November 04, 2015 | 1069 | 75663 | Not_Computer
Z-1783-2013 | 65679 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Rultract Retractor Ratchet Cover Plate Distributed by Pemco, Inc., Cleveland, OH 44131 |  |  | July 10, 2013 | 2013 | Pemco Inc | The firm became aware of the problem when a ratchet cover plate on a device was returned for repair showed signs of corrosion. | Nonconforming Material/Component | Rultract Skyhook sent a Recall Notification letter dated July 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify the device and return preferably just the gearbox portion of the device to:Pemco, Inc.Attn: Cover Plate Replacement5663 Brecksville RoadIndependence, OH 44131For customers who are unable to send the gearbox for recall service, Pemco would replace the Cover Plate free of charge.Customers with questions were instructed to call 216-524-2990, ext 103.For questions regarding this recall call 216-524-2990. |  | 87 | 87 units | Worldwide Distribution - USA including AL, CA, CT, FL, GA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, NC, NJ, NM, NV, OK, PA, TN, TX, UT, VA &amp; WA. Internationally to Czech Republic, Italy, Japan, Korea &amp; Singapore. | Terminated | February 11, 2014 | 216 | 87 | Not_Computer
Z-1782-2013 | 65651 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | ARTIS zee and zeego x-ray, angiographic system with Software Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed. |  |  | June 26, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a customer safety advisory notice and field safety corrective action for a potential issue on Artis systems with Artis VC20x/VC21A that have Patch 7 and a 3D and DSA license installed. The issue affects-DSA rotational acquisition. The restriction is related to any rotational 3D-acquisition program with real-time subtraction (e.g. 3D DSA). During the fill phase of the rotational aqui | Software Design Change | Siemens sent a Customer Safety Advisory Notice dated June 26, 2013, to all affected customers. The letter identified the product, the problem, and the action needed to be taken by the customer. Customers were strongly recommend to carefully evaluate one of the following options before running 3D rotational acquisition programs that include real-time subtraction until a patch is installed on their system: 1) position the monitor in a way such that contrast advancement can be monitored during the fill phase on the live monitor in the control room.2) For applications not requiring actions based on the contrast flow during the fill phase, the contrast may not need to be monitored during the fill phase. Although the monitor is not showing the expected subtraction, the acquisition is working correctly. There is no need to interrupt, but continue and finish the workflow.3) Perform 3D rotational programs that include real time subtraction from the exam room. The display on the live monitor in the exam room is not affected by this problem. To resolve the issue, the update AX028/13/P (VC20x/VC21B), will be made available around July 2013, and scheduling an appointment for implementation of the update with Siemens Customer Service organization was recommended. If this device/equipment is no longer in theirr possession, customers were asked to forward this safety notice to the new owner and inform Siemens of the new user's identity. |  | 29 | 29 | Nationwide Distribution and Puerto Rico | Terminated | February 05, 2015 | 589 | 29 | Not_Computer
Z-1759-2013 | 65553 | No | MAL | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | AlboGraft Knitted with Collagen, Bifurcate |  |  | June 17, 2013 | 2013 | LeMaitre Vascular, Inc. | Blood blushing/leaking from the surface of the graft after implantation | Process control | LeMaitre Vascular, Inc.sent an Urgent Field Safety Notice daed June 19, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify and return the affected device to MeMaitre Vascular which they will replace with a new unit. Customers with questions were instructed to call 781-221-2266, ext. 183. For questions regarding this recall call 781-425-1670, ext 108. |  | 3 | 3 units | Worldwide distribution: USA including PA and countries of: BRAZIL, BULGARIA, FRANCE, GEORGIA, GERMANY, GREECE, ITALY, LEBANON, NETHERLAND, POLAND, SPAIN, and TURKEY. | Terminated | August 15, 2016 | 1155 | 33 | Not_Computer
Z-1794-2013 | 65259 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | TotalCare Bed used with AccuMax Quantum VPC mattress |  |  | May 30, 2013 | 2013 | Hill-Rom, Inc. | When the bed is put into the chair egress position, the magnets on the mattress may not hold the bottom of the mattress flat to the bed frame, and the mattress foot section will extend away from the bed at an angle. A patient attempting to get back into the bed may miss the seat and sit on the foot section instead. The patient can slide down to the floor, or fall, with possible injury. | Device Design | The firm, Hill-Rom, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated May 29, 2013 was sent to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: Until the correction is implemented, after placing the bed into chair position, regularly check the foot section of the mattress. If necessary, push it back against the bed frame until the magnets hold it in place before allowing the patient to get back into or out of bed. In addition, forward a copy of this letter to any other facility personnel you deem appropriate.If you have any questions concerning this notification, please contact your distributor, your Hill-Rom representative, or Hill-Rom Technical Support at 800-445-3720 (Mon-Thurs, 8 am-6:30 pm; Friday, 8 am-6 pm). |  | 4121 | 4,121 Units | Worldwide distribution: USA (nationwide) and countries of: Canada, Colombia, Nigeria, Kuwait and Saudi Arabia. | Terminated | April 17, 2014 | 322 | 4121 | Not_Computer
Z-1796-2013 | 65635 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | TEG 5000 Thromboelastograph Coagulation Analyzer |  |  | June 20, 2013 | 2013 | Haemonetics Corporation | Research Use Only (RUO) parameters were included in TEG 5000 User Manual. | Labeling False and Misleading | Haemonetics sent a Medical Device Correction letter dated June 20, 2013 to all affected customers. The letter identified the affected product, problem and action to be taken. The letter instructs customers to discard the current manuals and replace with revised manuals included with the Customer Notification Letter. A Consignee Response Form has been included with the letter, which is to be signed and returned for effectiveness purposes. Follow up will be made to customers who don't return the Response Form. For questions contact TEG Technical Support team at 1-800-438-2834. |  | 1590 | 1590 devices | US Nationwide | Terminated | June 04, 2015 | 714 | 1590 | Not_Computer
Z-1799-2013 | 65614 | No | JSJ | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | BD BBL Schaedler K-V Agar with 5% Sheep Blood, carton of 20 plates, catalog # 221555, labeled in part ***Becton, Dickinson and Company 7 Loveton Circle Sparks, MD 21152 USA 800-638-8663 www.bd.com/ds*** |  |  | June 06, 2013 | 2013 | Becton Dickinson &amp; Co. | Microbiological identification media may exhibit reduced levels of Vancomycin. | Nonconforming Material/Component | The firm, BD Diagnostic Systems, sent an "URGENT PRODUCT RECALL" letter dated June 2013 to Distributors by fax, US customers via letters by UPS and OUS BD sites via email and teleconference. The letter described the product, problem and actions to be taken. All customers were instructed to discontinue use of the listed product; discard any remaining packages, and complete and return the attached form, whether or not you have any inventory remaining, via Fax to: Attention: Regulatory Compliance, 410-316-4258. BD will issue replacements for the discarded material. It is not necessary to contact BD by telephone to receive replacements.If you have any questions, call 410-316-4000. |  | 92 | 92 cartons | Worldwide distribution: US (nationwide) and countries of:Belgium, Canada, Columbia, Mexico, Singapore, and Taiwan. | Terminated | September 03, 2013 | 89 | 1744 | Not_Computer
Z-1795-2013 | 65604 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Axiom Aristos MX Solid State Flat Panel Digital Imager, Model number 5895003. The device is used for diagnostic X-ray imaging. |  |  | June 05, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Field Safety Notice about the potential hazard to patients or operators when using the function "table top float" of the AXIOM Aristos MX. When repositioning a patient lying on the table top and releasing the table top, it is possible that the fingers of the patient or operator are between the table top and the detector tray and if the table top is moved, the fingers can be squee | Device Design | On 6/5/13, consignees were notified by letter of a "Field Safety Notice" and received instructions about how to avoid the potential risk of this issue while using the device. Siemens also rolled out a modification of AXIOM Aristos MX that will avoid these potential injuries in the future. |  | 401 | 401 | Nationwide Distribution. | Terminated | February 24, 2015 | 629 | 401 | Not_Computer
Z-1797-2013 | 65468 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator, |  |  | June 19, 2013 | 2013 | Philips Healthcare Inc. | The MRx may deliver a non-synchronized cardioversion shock resulting in the delivery of incorrect therapy, which may induce ventricular fibrillation. | Software design | Philips sent an Urgent Medical Device Correction letter on June 19, 2013, to all affected customers. The letter informed customers of the issue, identified details of the units affected, gave instructions on actions to be taken by the customer and identified what action Philips planned to take to remedy the issue.Philips will perform a software upgrade and replace the label on the unit that outlines the steps for Synchronized Cardioversion with a revised label for all affected units at no charge to the customer.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice: You can continue to use your MRx prior to receiving the software upgrade and device label revision, provided that you follow the instruction provided.Customers with questions were instructed to contact their local Philips representataive or call 1-800-722-9377.For questions regarding this recall call 978-687-1501. |  | 66150 | 66,150 units | Worldwide Distribution - USA (nationwide) and countries of: AFGHANISTAN ALBANIA ARGENTINA ARUBA AUSTRALIA AUSTRIA AZERBAIJAN BAHRAIN BANGLADESH BELGIUM BOLIVIA BOSNIA AND HERZEGOWINA BRAZIL BRUNEI DARUSSALAM BULGARIA CAMBODIA CHILE CHINA COLOMBIA COSTA RICA CROATIA (local name: Hrvatska) CYPRUS CZECH REPUBLIC DENMARK ECUADOR EGYPT EL SALVADOR ESTONIA ETHIOPIA FINLAND FRANCE GERMANY GREECE GUADELOUPE HONG KONG ICELAND INDIA INDONESIA IRAQ IRELAND ISRAEL ITALY JAPAN JORDAN KAZAKHSTAN KENYA KOREA, REPUBLIC OF KUWAIT LAOS LATVIA LEBANON LIBYAN ARAB JAMAHIRIYA LITHUANIA LUXEMBOURG MACAU MACEDONIA MALAYSIA MALTA MEXICO MONGOLIA MOROCCO NAMIBIA NEPAL NETHERLANDS NETHERLANDS ANTILLES NEW CALEDONIA NEW ZEALAND NICARAGUA NIGERIA NORWAY OMAN PAKISTAN PALESTINE PANAMA PAPUA NEW GUINEA PARAGUAY PERU PHILIPPINES POLAND PORTUGAL PUERTO RICO QATAR REUNION ROMANIA RUSSIAN FEDERATION SAUDI ARABIA SERBIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SRI LANKA SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA TANZANIA, UNITED REPUBLIC OF THAILAND TRINIDAD AND TOBAGO TUNISIA TURKEY UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN VENEZUELA VIET NAM and YEMEN. | Open, Classified |  |  | 66150 | Computer
Z-1791-2013 | 65579 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Discovery NM 630, dual detector free-geometry integrated nuclear imaging system. |  |  | June 13, 2013 | 2013 | GE Healthcare, LLC | GE became aware of an incident at a VA Medical Center facility in the US in which a patient died due to injuries sustained while being scanned on an Infinia Hawkeye 4 Nuclear Medicine System. UPDATE: On July 03, 2013 GE notified hospitals that they are recalling several Nuclear Medicine Imaging Systems because serious injuries or deaths could occur due to the failure mode associated with this re | Device Design | GE sent an Urgent Medical Device Correction letter dated June 17, 2013 to all affected customers. The letter identified the affected product, recommended that qualified service personnel maintain the equipment and that Preventative Maintenance procedures are executed according to labeling. Also the Safety Chapter Sections should be re-reviewed with personnel to ensure proper operation of the equipment. GE notified customers again on July 03, 2013, via Urgent Medical Device Recall letter (including confirmation of delivery for US customers) and follow-up telephone call. Customers were instructed to cease using the affected system until GE Healthcare Field Engineer can complete an inspection of the system and perform any necessary repairs at no cost. A GE Healthcare representative will contact the hospitals to arrange for the inspection. |  | 108 | 108 | Worldwide Distribution - US Nationwide including Guam, Puerto Rico and the countries of : Afghanistan, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea (Republic of), Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Malta, Martinique, Mexico, Morocco, Namibia, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Uruguay, Venezuela, Vietnam. | Terminated | June 19, 2014 | 371 | 3454 | Not_Computer
Z-1810-2013 | 65600 | syringe, irrigating (non dental) | KYZ | Irrigating syringe. | General Hospital | General Hospital | 510(K) Exempt | Medtronic Preimplantation Test Kit, Catalog No. 21047. The kit is indicated for use with Medtronic CSF-Flow Control Valves and Delta Valves, when a simple preimplantation test is desired to verify that the valve conforms to labeled product specifications. |  |  | June 28, 2013 | 2013 | Medtronic Neurosurgery | Medtronic Neurosurgery is recalling the Medtronic Preimplantation Test Kit because the outer carton had expiration dates that were incorrect. The date on the outer carton is later than the expiration date for one or more of the individually packaged sterile components in the kit. | Incorrect or no expiration date | An "Urgent Field Safety Notice" recall letter dated 6/28/13 was sent to customers who purchased the Medtronic Preimplantation Test Kit, catalog number 21047, to inform them of the recall due to the incorrect outer carton expiration dates. The recall letter informed the customers of the problems and actions identified. Customers were instructed to contact the Quality Department at (805) 571-8725 for questions. |  | 915 | 915 units | Nationwide Distribution in the US. | Terminated | September 23, 2013 | 87 | 915 | Not_Computer
Z-1808-2013 | 65654 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Stryker brand Neuroform 3 Microdelivery Stent System, 4.5 mm X 20 mm, REF Catalog Number: E345020, Model/Product Number: M003450200, Product is manufactured by Boston Scientific Corporation, Fremont, CA and distributed by Stryker Neurovascular, Fremont, CA |  |  | June 24, 2013 | 2013 | Stryker Neurovascular | The device is labeled incorrectly, in that, the device Model Number on the outer carton and inner pouch labels did not match, although both referenced the same lot number. The Model Number on the outer box states UPN M003E345020 (4.5mm x 20 mm Neuroform 3 Stent) while the pouch label states UPN M003E3450300 (4.5mm x 30 mm Neuroform 3 Stent). | Packaging process control | Stryker Neurovascular sent an Urgent Field Safety Notice on June 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately check their internal inventory and quarantine all subject devices pending return to Stryker. Circulate the Field Safety Notice internally to all interested/affected parties. Maintain awareness of this notice internally until all requred actions have been completed within their facility. Inform Stryker if any of the subject devices were distributed to other organization. Complete the attached customer response form and return to their Stryker Representataive.For questions regarding this recall call 510-413-2500. |  | 23 | 23 units | International Distribution including Austria, China, and Germany. | Terminated | September 30, 2013 | 98 | 23 | Not_Computer
Z-1809-2013 | 65391 | No | LXM |  |  | Physical Medicine | 510(k) | Advanced Orthogonal Percussion Adjusting Instrument |  |  | March 06, 2013 | 2013 | Advanced Orthogonal Equipment, Incorporated | Unapproved medical device | No Marketing Application | The firm, Advanced Orthogonal Equipment, sent a "CORRECTION NOTICE" to its consignees/customers. The notice described the product, problem and actions to be taken. The customers were instructed that " none of these devices be used pending final clearance by FDA".Questions and concerns regarding this notice should be directed to the President of Advanced Orthogonal Equipment by phone at: 727-528-8700 or email: Atlasortho@aol.com. |  | 54 | 54 | Nationwide distribution: US including states of: AL, CA, CO, FL, GA, ID, KY, MD, NC, NJ, NY, OH, PA, TN, VA, and WA. | Open, Classified |  |  | 54 | Not_Computer
Z-1806-2013 | 65703 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | ProSpace Peek Implant 5 degree x 8.5 x 22 mm |  |  | June 18, 2013 | 2013 | Aesculap, Inc. | Aesculap Implant Systems, LLC initiated a recall of the Aesculap Prospace PEEK 5 degree Implant - 8 x 8.5 x 22 mm (SN038P) because the device is incorrectly etched with an 8 degree lordosis instead of a 5 degree lordosis. The device is labeled correctly, only the etching of the implant may be incorrect. | Employee error | Aesculap notified customers via email and phone on June 18, 2013, and were instructed to check their inventory and return affected product to Aesculap.Aesculap believes that the risk to patient is low.For further questions please call 1-800-234-9179. |  | 9 | 9 | US Distribution including the states of NV and MI | Terminated | January 13, 2014 | 209 | 9 | Not_Computer
Z-1807-2013 | 65401 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Swift Mobil shower chair, Swift Mobil 160 shower chair, and Swift Mobil Tilt shower chair. |  |  | May 23, 2013 | 2013 | Etac Supply Center Ab | Etac is voluntarily conducting a field correction of the Etac Swift Mobil shower chair. If the Swift Mobil shower chair loses a caster and thus becomes instantly unstable, there is potiential risk for harm or injury. | Component change control | Etac sent an "URGENT FIELD CORRECTION" letter dated May 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 7200 | 7200 internationally | International distribution: Austria, Australia, Belgium, Brazil, Canada, Czech Republic, Denmark, Esonia, Finland, France, Great Britain, Germany, Greece, Hong Kong, Ireland, Isreal, Iceland, Italy, Kingdom of Bahrain, Kuwait, Lithuania, Netherlands, Norge, New Zealand, Spain, Switzerland, Sweden, and United Arab Emirates. | Terminated | October 24, 2014 | 519 | 7200 | Not_Computer
Z-1811-2013 | 65686 | [6 empty] | Sodium Bicarbonate, USP |  |  | June 11, 2013 | 2013 | Church &amp; Dwight Inc | Church and Dwight Co., Inc. has initiated a recall for Sodium Bicarbonate, USP produced by the Old Fort, OH Sodium Bicarbonate Production facility due to metal and other contamination. | Material/Component Contamination | Church and Dwight Inc., issued an Urgent Recall letter dated June 11, 2013, and return response form to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to quarantine and return the affected product. For questions regarding this recall call 609-279-7748. |  | 585000 | 234 sacks at 2500 lbs. each of Dialysis Grade 1.5 ; 880 bags at 50 lbs. each of Dialysis Grade 2 | Nationwide Distribution including IL, PA, IN, OH, MI, TN, TX, MN, MO, NJ, KY,and MD. | Terminated | June 30, 2015 | 749 | 585000 | Not_Computer
Z-1812-2013 | 65619 | No | MLA | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD Affirm VPIII Microbial Identification Tests, packaged in kits, 120 test\kit, Catalog # 446257 and 24 test\kit, Catalog # 446252; kits are labeled in part ***Becton, Dickinson and Company, 7 Loveton Circle Sparks, MD 21152 USA, 800-638-8663, www.bd.com/ds*** |  |  | June 13, 2013 | 2013 | Becton Dickinson &amp; Co. | In vitro diagnostic test kit may exhibit invalid internal negative control results. | Under Investigation by firm | Becton Dickinson Diagnostic Systems notified distributors by fax on June 13, 2013, and advised consignees to discontinue use and discard affected product for credit. Distributors were requested to return their distribution list to BD for direct notification by the recaller. End users were contacted by letter on June 13, 2013 and the notification addressed to Microbiology Lab Director/Risk Managers advised that a small number of tests associated with the referenced lots of product may exhibit invalid internal negative control results. If encountered, the issue would result in an invalid test, as described in the package insert. End users were requested to discontinue use of these lot numbers, discard any remaining packages for replacement and return the enclosed response form.Customers with questions were instructed to contact BD Technical Services Department at 1-800-638-8663. For assistance regarding replacements customers were instructed to contact BD Customer Service at 1-800-675-0908.For questions regarding this recall call 410-316-4054. |  | 177 | 177 kits | Nationwide Distribution | Terminated | September 28, 2013 | 107 | 177 | Not_Computer
Z-1828-2013 | 65473 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Medline Med-Checks BI Test packs with Instant Readout Integrator, SteriTec Product Code PL 200, Distributor Product Code MDS200500IR |  |  | June 10, 2013 | 2013 | Steritec Products Mfc. Co., Inc. | Mesa Labs, based out of Bozeman, MT, informed SteriTec Products, Inc. of a product recall involving the EZTest Steam and Smart-Read Biological Indicators (BI) because the readout time is not in specification with the label claims. | Nonconforming Material/Component | The firm, SteriTec, sent all customers recall notifications via e-mail. The e-mail included a .pdf file of the recall notification letter entitled "URGENT: MEDICAL DEVICE RECALL" dated June 11, 2013. The letter described the product, problem and actions to be taken. The letter also informed customers of MesaLabs, "URGENT: MEDICAL DEVICE RECALL" letter dated June 7, 2013 to its customers. The customers were instructed to notify SteriTec products of receipt of this notification via email at: Jon@SteriTecproducts.com or fax to: 303-660-4213 Attn: Director Regulatory Affairs; examine the lot number printed on the packaging components; place affected items in hold and do not sell or distribute these lots; record inventory of all items remaining in your warehouse and all items that have been distributed; advise SteriTec of the status of all inventory you received used or unused, if you have distributed the affected product to other locations within your healthcare system, notify them immediately of this recall, ask them to provide inventory status and proceed with destruction or return via instructions provided in letter-If you have shipping questions, please contact SteriTec directly at 303-660-4201, or email us at: Info@SteriTecproducts.com, and also provide this notice to anyone in your facility that needs to be informed.Should you have any questions or concerns, please do not hesitate to contact Director of Regulatory Affairs and Quality Assurance or our offices at 303-660-4201. SteriTec's hours of operation are 7:30am Mountain Standard Time (MST) through 5:00pm MST. |  | 3535 | 3535 units | Worldwide distribution: US (Nationwide) and countries of: Argentina, Australia, Canada, British Virgin Island, Czech Republic, Italy, and Turkey. | Terminated | December 19, 2013 | 192 | 57665 | Not_Computer
Z-1813-2013 | 65328 | No | DSQ |  |  | Cardiovascular | PMA | Heartware HVAD Pump Implant Kit (Heartware Ventricular Assist Device) |  |  | May 16, 0013 | 13 | HeartWare Inc | Potential to electrostatic discharged (ESD) | Device Design | HeartWare sent an Urgent Medical Device Correction Safety Notification letter on May 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to educate their patients about the HeartWare Controller and instructions for equipment handling stated in the IFU and recognized practices to avoid electrostatic discharge. The direct accounts were also told to scan and return the signed Acknowledgment Form via fax to 305-364-2665 or e-mail to quality@heartwareinc.com.Customers with questions were instructed to contact their HeartWare Representative or HeartWare Customers Service.For questions regarding this recall call 305-364-1575. |  | 76 | 76 (US) and 82 (International) | Worldwide Distribution - USA including AL, AZ, CA, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NY, OH, OR, PA, TN, TX, UT, VA, WA, and WI. Internationally to:Australia, Austria, Belguim, Brazil, Canada, Chile, Croatia, Czech Republic, France, Germany, Greece, India, Israel, Italy, Japan, Lebanon, Lithuania, Luxemburg, Malaysia, Netherlands, New Zeland,Norway, PolandSingapore, South Africa, Switzerland, Turkey, and United Kingdom. | Open, Classified |  |  | 76 | Not_Computer
Z-1818-2013 | 65718 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | ViSi Mobile Chest Sensor (3 lead-wire ECG, AAMI), Catalog No. 93-10014 |  |  | May 23, 2013 | 2013 | Sotera Wireless, Inc. | Sotera Wireless, Inc. is recalling the ViSi Mobile Monitoring System due to the potential for the Chest Sensor to cause a blister wound. | Device Design | Sotera sent an Urgent Product Recall letter dated May 23. 2013, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to report any adverse reactions experienced with the use of this product and or quality problems to the FDA's MedWatch Program by: www.fda.gov/medwatch, (800) FDA-1008, fax (800) FDA-0178. Customers were also instructed to contact Sotera Wireless Technical Support at (866) 794-5526 for questions. |  | 75 | 75 units | US distribution including CA and UT. | Terminated | September 09, 2013 | 109 | 165 | Not_Computer
Z-1821-2013 | 65668 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Synapse SD Drill Guide, Part Number SD389.477, Lot Number 6913900 |  |  | October 12, 2012 | 2012 | Synthes USA HQ, Inc. | A complaint was received from the field, that the drill guide was etched incorrectly, 12mm etch instead of 18mm drill guide. The firm identified that this constituted a field action and retrospectively reported this action. | Labeling Change Control | The firm, Synthes USA, sent an email to its customer. The email identified the product, problem and the actions to be taken. The customer was instructed to do the following: do not use this product; do not remove any product from your account or return any product without further instructions. If you have identified that you have the product, please have the product located in an area that ensures it will not be used. Further instructions will be sent out with regards to any action that needs to be taken next.Please call Synthes if you have any questions at 610-719-5364. |  | 2 | 2 | US distribution: PA only.. | Terminated | June 11, 2015 | 972 | 2 | Not_Computer
Z-1814-2013 | 65666 | No | NSX |  |  | General Hospital | Contact ODE | Picis ED Pulsecheck - EMR Software Application - 2125, Software Versions: 5.2 and 5.3. The application stores patient information in a database, and it may analyze and/or display the data in different formats for evaluation by healthcare professionals for informational purposes. |  |  | June 21, 2013 | 2013 | Picis Inc. | Notes associated with prescription are not printed to the prescription or to the patient chart. | Software change control | Initial customer notifications were sent via email on June 21, 2013 informing consignees of the recall and providing further instruction regarding the software solution. Consignees were provided with recommended actions until they receive the necessary update. |  | 35 | 35 | Nationwide Distribution, including the states of: AK, AR, AZ, CA, CO, DC, DE, FL, GA, ID, IN, MA, MD, MO, NH, NJ, OH, OR, SC, TN, WA, and WV. | Terminated | December 19, 2014 | 546 | 35 | Computer
Z-1819-2013 | 65068 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Custom Spine ISSYS LP Inverted Screw System. 8.5MM X 45 MM ISSYS LP Screw Rx Non-Sterile. Custom Spine, Inc. 9 Campus Dr. Parsippany, NJ 07054 USA. TEL: 973-808-0019 |  |  | March 28, 2013 | 2013 | Custom Spine, Inc. | On March 28, 2013, Custom Spine, Inc. issued a minor removal of a lot of 18 screws that were found to be out of specification. The screw length was 5 mm shorter in length than the required specification for the screw size. The product is laser marked 8.5mm X 45mm, but its length measured 40mm. | Process control | Custom Spine initiated a voluntary removal On 3/28/2013, Custom Spine initiated a voluntary removal of all 10 pieces of the ISSYS LP (8.5mm X 45mm screws) Lot 363. Distributors were contacted via phone and/or email and were asked to immediately return these parts to Custom Spine. |  | 18 | 18 units | Nationwide Distribution including Nebraska and Virginia. | Terminated | October 24, 2013 | 210 | 18 | Not_Computer
Z-1815-2013 | 63935 | No | IEO | Automated tissue processor. | Pathology | Pathology | 510(K) Exempt | Leica ASP6025 Vacuum Tissue Processor |  |  | November 13, 2012 | 2012 | Leica Microsystems, Inc. | The Instruction for Use 1v9 Ref F and all previous versions specify incorrect dimensions for the Ready To Use (RTU) bottles from other suppliers than Leica. In case of a pressure failure of the device in combination with a bottle with incorrect dimensions, overflow can cause contamination of reagents. Additionally, the current instructions for use does not highlight the importance to check the filling levels of all reagent bottles (RTU and system bottles) prior to every process run.Without this important information, risk of tissue damage or loss is possible. Therefore the missing information will be distributed to all customers by a Field Service Notice. | Device Design | Leica Biosystems sent an Urgent Field Safety Notice dated December 12, 2012, to all affected customers. The notification included a description of the problem and a recommendation that the personnel who use the equipment only use Leica Ready To Use (RTU) bottles that are already supplied as a standard delivery item with the ASP6025 machine. The personnel are further reminded that reagent bottle levels need to be checked and corrected when necessary. Customers with questions or questions regarding this recall call 847-405-5413. |  | 128 | 128 machines | Worldwide Distribution - USA (nationwide) and internationally to Australia, Belgium, Brazil, Chile, France, Germany, Ireland, Italy, Japan, Malaysia, Mexico, Netherlands, Philipines, Singapore, South Africa, Spain, Switzerland, Thailand, Turkey, and United Kingdom. | Terminated | January 05, 2017 | 1514 | 128 | Computer
Z-1830-2013 | 65576 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AIA Analyzer Pipette Tips used on Tosoh AIA-2000, AIA-1800, AIA-900 and AIA-600II Analyzers, 1000 tips/bag |  |  | May 15, 2013 | 2013 | Tosoh Smd Inc | Visual inspection of the pipette tips indicated that some of the tips were not uniform in size/length which was causing error messages on the analyzers. | Nonconforming Material/Component | Tosoh Bioscience sent an Urgent Recall Notification letter dated May 15, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to respond on the amount of product that was destroyed by completing the attached form and faxing it to 614-317-1941.Customers with questions were instructed to call Tosoh Technical Support at 1-800-248-6764.For questions regarding this recall call 614-317-1909. |  | 11309 | 11,309 bags | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WY.Internationally to Canada, Japan, Mexico, Chile, and Uruguay. | Terminated | November 27, 2013 | 196 | 11309 | Not_Computer
Z-1831-2013 | 65084 | No | DCF | Albumin immunological test system. | Immunology | Immunology | 510(K) Exempt | uTASWako DCP, Lot TG340 2013.11, REF 995-60701, Wako Pure Chemical Industries, Ltd. 1-2 Doshomachi 3-chome, Chuo-Ku Osaka 540-8605 Japan Distributed by Wako Diagnostics Wako Life Sciences, Inc. 1025 Terra Bella Ave., Suto A. Mountain View, CA 94043 USA. The product is a single reagent cartridge run on an automated system and each assay performed on a |  |  | April 30, 2013 | 2013 | Wako Life Sciences, Inc. | Product fails to meet the accuracy at the low end of the measuring range. | Material/Component Contamination | Wako Life Sciences, Inc. sent an email and issued a Recall Notice letter via FED EX on April 30, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue the use of these kits, fill out the attached RMA form, and return unused kits using Wako's FedEx account (provided on RMA form). Partially used reagent cartridges should be disposed of according to your local and state requirements. Wako Life Sciences will relplace your existing stock of uTASWako DCP kits with kits from a different log number that has been tested and has shown to not be affected. If you have any questions regarding the Recall, please submit your inquiries to liver@wakousa.com or call 443-535-0702. |  | 3600 | 36 cartridges -100 tests per cartridge | Distribution to CA, GA, MN, NC, UT and the District of Columbia. | Terminated | August 19, 2013 | 111 | 3600 | Not_Computer
Z-1834-2013 | 65440 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | AU5800 Clinical Chemistry Analyzer with Automation Ready Connection (All Software Versions), Model Numbers AU5811, AU5821, AU5831 and AU5841. This system is for in vitro diagnostic use only. |  |  | June 03, 2013 | 2013 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has identified a software error on AU5800 analyzers with the Automation Ready (AR) Sampler Kit. This software error results in the failure of the concentrated wash solution pump to stop, creating an overflow or leak on each analyzer unit. | Software design | An Urgent Product Correction letter dated 5/29/13 was sent on the week of June 3, 2013 to all the customers who received the AU5800 Clinical Chemistry Analyzer with Automation Ready (AR) Connection (All Software versions). The letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to complete and return the enclosed Response Form within 10 days and to share the information with their laboratory staff. Customers with questions are instructed to contact the 24 hours Call Center (Hotline) at (800) 854-3633 in the US and Canada. |  | 76 | 76 units | Worldwide Distribution, including Nationwide (US) and the countries of China, Italy, Japan, Canada, Turkey, United Kingdom, Singapore, Israel, and the Netherlands. | Terminated | August 23, 2014 | 446 | 76 | Computer
Z-1846-2013 | 65646 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | GAMMA3 System Long Nail Kit, R2.0, Ti, Left |  |  | May 21, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker Osteosynthesis has become aware that LEFT GAMMA3 nails are marked as right nails. | Process control | Stryker sent notification letters and product accountability forms dated May 21, 2013. via Fed Ex with return receipt to all branches/agencies and on May 22, 2013, via Fed Ex with return receipt to all hospital risk management, chief of orthoapedics, and surgeons.Our records indicate that you may have received and/or used the above referenced product(s). It is Stryker'sÂżÂż responsibility as the manufacturer to ensure that customers who may have received and/or used these affected products also receive this important communication.Please contact your StrykerÂżÂż Sales Representative to arrange for return of the product if you have any of the above listed items. Please assist us in meeting our regulatory obligation by faxing back the attached Product Recall Acknowledgment Form as soon as possible to 855- 251-3635.Please note that your signature on the following form only confirms that you received this notification and does not obligate you to take any additional action beyond what is called for in this notification letter.We regret any inconvenience this action may cause you and if you have any questions, please call (201) 972-2100. |  | 5 | 5 units | US Distribution including the states of IL, KY and TN. | Terminated | October 14, 2014 | 511 | 5 | Not_Computer
Z-1836-2013 | 63108 | No | LGW |  |  | Neurology | PMA | Eon Mini - 3788 (IPG), |  |  | July 26, 2012 | 2012 | St. Jude Medical | As part of St. Jude Medical's routine tracking and product monitoring, they have received a total of 214 reports, out of 34,617 Eon Mini IPGs, that lost the ability to communicate or recharge due to an inner battery weld issue resulting in loss of pain relief and subsequent explant. Explant surgery, as with any surgery, presents a risk to patient health. Adverse events associated with an unplanne | Device Design | St. Jude Medical sent a Important Medical Device Recall Advisory dated July 26, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.To further assist in your patient care we are providing you with a list of all serial numbers we show have been distributed to you. ( See Attachment A ).Further questions please call (972) 309-2154. |  | 34617 | 34,617 | Worldwide Distribution--USA (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Belgium, Brazil, canada, Colombia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Mexico, Netherlands, Saudi, South Africa, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | August 21, 2015 | 1121 | 34617 | Not_Computer
Z-1832-2013 | 65568 | No | GFC | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Exactech Articulated Driver |  |  | May 30, 2013 | 2013 | Exactech, Inc. | When torque is applied to the Articulated Driver at the maximum angle, the cross pin may shift out of position, potentially leading to disassembly of the instrument. | Under Investigation by firm | Exactech sent a" Important Product Recall Notice" dated June 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were requested to Immediately cease distribution or use of the products. Extend this information to your accounts that may have this product in their possession. Verify if you have any of the subject ARTICULATE DRIVER (catalog number 181-00-04) in the specified lots. FAx back the attached form.Further questions please call 1-800-392-2832. |  | 57 | 57 | US Distribution including the states of AR, CA, FL, NJ, NY, OK, and TX | Terminated | September 19, 2013 | 112 | 57 | Not_Computer
Z-1833-2013 | 65755 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, VOLUSON E6, E8, E8 EXPERT, E10 DIAGNOSTIC ULTRASOUND SYSTEM. |  |  | July 13, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware that during use of the Voluson E6 or E8 systems, that an error message Out of Memory can be displayed which could cause the system to freeze. This is an intermittent behavior that also requires a certain workflow. This issue can be resolved by rebootingthe system. | Software design | No consignee letter was sent out.Starting on July 13, 2012, Field Engineers visited consignees and carried on instructions as per Field Modification Instruction (FMI 79044A). The Engineer returned the system to the customer and ensure acceptance. Before leaving the customer site have the staff operate the system. Ensure that: all settings are customized, all peripherals are functioning and all communications are working properly. We plan to update the software on all of the affected GE VOLUSON E6, E8, E8 EXPERT, E10 DIAGNOSTIC ULTRASOUND SYSTEM that were distributed.As an effectiveness check, we will track all service requests for the field action to completion. These service requests will record the repair of the affected units. Exceptions to completion of the service request will be recorded as well. Records of field action completion will be retained per our processes and procedures.Further questions please call 262-513-4122. |  | 375 | 375 | Worldwide Distribution - US Distribution including the states AK, AZ, AR, CA, CO, CT, DE, FL, GA, IL, LA, MD, MI, MN, MO, NE, NV, NM, NY, NC, OH, OR, PA, SC, SD, TN,TX, VA, WA and WI., and the countries of ALGERIA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRIL, CANADA, CHINA, COLOMBIA, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, INDIA, IRAQ, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, REPUBLIC OF MOLDOVA, NETHERLANDS, NORWAY, POLAND, ROMANIA, RUSSIAN FEDERATION, KINGDOM OF SAUDI ARABIA, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA. | Terminated | November 10, 2014 | 850 | 375 | Computer
Z-1837-2013 | 65498 | No | BXB | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Omnicycle Elite; US Importer and Distributor: Accelerated Care Plus, Reno NV. Manufactured by Medica Medizintechnik GmbH, Hochdorf, Germany. The product is indicated for powered exercise equipment intended for medical purposes. |  |  | June 12, 2013 | 2013 | Accelerated Care Plus Corporation | An anomaly in the firmware that may cause Omnicycle Elite to operate at a speed other than what was previously set when in reverse mode. | Software design | Letters were issued on June 12, 2013 that included a memory stick with self loading software so consumers can update the device. Letters were sent by trackable method. Customers with questions or concerns should call Customer Support at 800-350-1100. |  | 1001 | 1001 units | Nationwide Distribution. | Terminated | October 23, 2013 | 133 | 1001 | Computer
Z-1820-2013 | 65591 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LPS Lower Extremity Dovetail Intercalary component. |  |  | July 11, 2013 | 2013 | DePuy Orthopaedics, Inc. | The LPS Lower Extremity Dovetail Intercalary component has the potential for fracture of the female component when exposed to certain physiological loads. | Device Design | On 7/11/13, URGENT MEDICAL DEVICE RECALL NOTIFICATIONS were sent to consignees: Depuy Distributors, hospitals, and surgeons. End-user Hospital Customers will submit reconciliation forms to DePuy. Possible clinical implications of device failure are stated in these notifications, including poor mechanics and loss of function, pain, component dislocation, adverse tissue reaction, bone damage, potential need for revision surgery and the risks associated with this type of surgery. A patient letter will be included with the surgeon letter to assist in surgeon's notification and discussion with patients. The devices will be returned to DePuy Warsaw through the normal DePuy Returns process, to attention of Returns and marking H13-13 on the outside of the box.If a patient presents with a fractured LPS" Lower Extremity Dovetail Intercalary component with well-fixed proximal and distal stems and the surgeon determines LPS" Lower Extremity Dovetail Intercalary component is the best treatment option, DePuy plans to make the LPS" Lower Extremity Dovetail Intercalary component available through authorized approvals. Directions on how to purchase the component are provided in the Recall Communications.Any questions or concerns about the recall, please contact Kim Earle, Recall Coordinator, at 574-371-4917 (M-F; 8am-5pm EDT) |  | 2262 | 58 U.S. / 39 OUS | Worldwide Distribution-USA including the states of FL, GA, ID, KS, KY, MI, MO, NC, NV, NY, OH, OK, PA, TN, TX, VA, and WA and the countries of AUSTRALIA, CANADA, ENGLAND, GERMANY, INDIA, COUNTY CORK, IRELAND, and TURKEY. | Terminated | July 10, 2014 | 364 | 2262 | Not_Computer
Z-1859-2013 | 65603 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Information System (sDIS) versions 9.5 and VA10A. |  |  | June 10, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens initiate this recall due to a potential issue when using the sDIS matching rules of the syngo Dynamics Information System version 9.5 or VA10A with modality device configured to end study at association close. Images or DICOM SR objects sent from a modality device to syngo Dynamics may not be saved when the modality is configured at syngo Dynamics to end study at association close, and whe | Software design | Siemens sent a Customer Safety Advisory Notice dated June 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. 1) With sDIS implemented, use the DICOM Modality List (DMWL), without sDIS matching rules enabled OR 2) Do not implement sDIS. Update instructions were also provided. The reported issue will be resolved in a modification to syngo Dynamics version 9.5 and syngo Dynamics VA10A which will be released in the near future.We regret any inconvenience that this may cause , and we thank you in advance for your understanding.For further questions please call (610) 219-6300. |  | 9 | 9 | US Distribution only including the states of PA, OH, UT, CA, NY, MO, and TX. | Terminated | February 24, 2015 | 624 | 9 | Computer
Z-1869-2013 | 64987 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Caddy - PedFuse, Screw Caddy, 8MM and 9MM, Part Number: 11-80009 Rev. B |  |  | October 05, 2012 | 2012 | SpineFrontier, Inc. | Two sections that are designated for the 8mm and 9mm screws were switched; the holes that hold the 8mm screws are designed for 9mm screws, and the holes that hold the 9mm screws are designated for 8mm screws. | Device Design | The firm, SpineFrontier, sent an Advisory Notice entitled "URGENT May 1, 2013, PedFuse Screw Caddy Removal" dated May 1, 2013 to their consignees/customers. The notice described the product, problem and action to be taken. Consignees/customers were informed that a removal has been initiated to remove the parts from the field to ensure that all sets contain correctly marked screw caddies. Consignees/customers wre instructed to complete and return the URGENT PedFuse Screw Caddy Removal ACKNOWLEDGEMENT FORM via fax to: SpineFrontier, Inc., 978-232-3991 or email: customerservice@spinefrontier.com. Replacement caddies will be shipped upon receipt of completed acknowledgement form and RMA request for part return.Please feel free to contact SpineFrontier at 978-232-3990 with any questions. |  | 11 | 11 caddies | US distribution in states of: CO, DC, MA and TX. | Terminated | July 29, 2015 | 1027 | 11 | Not_Computer
Z-1847-2013 | 65610 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Hospira Symbiq Infusion System is intended for the delivery of fluids, solutions, drugs, agents, nutritionals, electrolytes, blood and blood products via parenteral, enteral, intravenous, intra-arterial, subcutaneous, epidural, or irrigation routes of administration. |  |  | February 04, 2013 | 2013 | Hospira Inc. | The devices secondary audio signal fails to activate. | Process control | Hospira sent a Urgent Device Field Correction letter dated February 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. In addition, letters were sent by e-mail to ECRI, NRAC, Noblis and ISMP for notification of their members and posting on their websites February 5, 2013. The letter was also posted on the Hospira website February 4, 2013.Customers were requested to return a reply form indicating they have received the letter and notified their users.For further inquiries please contact Hospira Global Complaint Management at1-800-441-4100 ( 8am - 5pm CST, M-F) to report adverse events or product complaints, or Hospira Advanced Knowledge Center at 1-800-241-4002, option 4 for additional information or technical assistance. ( 24hrs a day / 7 days a week. |  | 2791 | 2,791 units | US Distribution including the states of AL, CA, CT, DE, FL, HI, IA, IL, IN, LA, MA, MD, ME, MI, MN, NC, NJ, NY, PA, SD, TX and WA. | Open, Classified |  |  | 2791 | Not_Computer
Z-1864-2013 | 64625 | No | KTI | Bronchoscope (flexible or rigid) and accessories. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Inspira AIR Balloon Dilation system. |  |  | January 11, 2013 | 2013 | Acclarent, Inc. | Labeling correction for all sizes of the Inspira AIR Balloon Dilation System to include additional language and warnings as well as to clarify some instructional steps. | Device Design | Urgent Medical Device Labeling correction letters will be sent to all consignees through a delivery system that can track delivery. After CDRH review and corrections, Acclarent began issuing customer notification on August 7, 2013 |  | 1217 | 1217 | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, and Saudi Arabia. | Terminated | October 25, 2013 | 287 | 28660 | Not_Computer
Z-1867-2013 | 65544 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Mesa EZtest BIOLOGICAL INDICATOR MONITORING SYSTEM. |  |  | June 04, 2013 | 2013 | Mesa Laboratories, Inc., Bozeman Manufacturing Facility | Mesa EZTest , Self-contained steam biological indicator, are recalled because Mesa recently determined that the read out time is not in specification with the label claims. | Nonconforming Material/Component | Mesa Labs sent an "URGENT: MEDICAL DEVICE RECALL" letter dated June 7, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. For questions contact Mesa Labs at BI-Support@mesalabs.com. |  | 21625 | 865/25 ct boxes, 414/100 ct boxes, 307/500 ct boxes | Worldwide Distribution: US (nationwide) and countries of: Australia, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Egypt, Greece, India,Israel, Japan, Jordan, Malaysia, Mexico, Milanese, Netherlands, Peru, Portugal, Pozan, Romania, Spain, Sweden, Taiwan, Thailand, United Kingdom, and Uruguay. | Terminated | November 22, 2013 | 171 | 532650 | Not_Computer
Z-1870-2013 | 65033 | No | FZY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | S-LIFT Slap Hammer, Part Number: 11-50204, Revision A |  |  | April 25, 2013 | 2013 | SpineFrontier, Inc. | Malfunction of slap hammer. | Device Design | The firm, SpineFrontier, sent an advisory notice entitled "URGENT April 25, 2013 S-LIFT Slap Hammer Removal" dated April 25, 2013 via email to its customers. The notice described the product, problem and action to be taken. The notice requested the removal of all distributed SLIFT Slap Hammer with listed lot number. The customers were instructed to complete and return the URGENT S-LIFT Slap Hammer Removal ACKNOWLEDGEMENT FORM via fax to: SpineFrontier, Inc., 978-232-3991 or email to customerservice@spinefrontier.com.Please feel free to contact SpineFrontier at 978-232-3990 with any questions. |  | 3 | 3 | US distributions to states of: KS and TX. | Terminated | January 09, 2014 | 259 | 3 | Not_Computer
Z-1849-2013 | 65759 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, LOGIQ e Ultrasound System. |  |  | November 09, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare became aware of a potential issue associated with the LOGIQ e system. Image data from one patient has the potential to be linked to the patient ID of another patient. | Software design | Starting on November 9, 2012, GE Healthcare, will plan to update the software on all of the affected GE LOGIQ e Ultrasound systems that were distributed at no cost to the customer.As an effectiveness check, we will track all service requests for the field action to completion. These service requests will record the repair of the affected units. Exceptions to completion of the service request will be recorded as well. Records of field action completion will be retained per our processes and procedures.No Customer letter was sent.Further questions please call (262) 513-4122. |  | 588 | 588 | Worldwide Distribution - USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, DE, FL, GA, ID, IL , IA, MD, MA, MI, MS, MO, NJ, NM, NY, NC, OH, PA, RI, SC, TN,TX, UT,VT, VA and WI., and the countries of ARAB EMIRATES, AUSTRALIA, AUSTRIA, BELGIUM, DENMARK, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, HONG KONG, INDIA, ITALY, JAPAN, NETHERLANDS,NEW ZEALAND, NORWAY, POLAND, RUSSIA, SAUDI ARABIA, SINGAPORE,SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND and UNITED EMIRATES. | Terminated | September 15, 2014 | 675 | 588 | Computer
Z-1848-2013 | 65491 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Luminos dRF Fluoroscopic X-Ray system |  |  | May 07, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens became aware of an unintended behavior when using the Luminos dRF with software version VD10A. The possibility exists that a software error may occur causing the table to tilt into an upright position and may collide with the floor. Four consignees were notified of this recall by Update Instruction XP067/12/S -- all four units have been corrected. | Software design | Siemens initiated their correction of this product in May, 2013 by sending an Safety Update Instruction XP067/12S to resolve the software error to all affected customers. All units have been corrected.Further questions please call (610)-219-6300. |  | 4 | 4 | US Distribution including the states of CA, IN, NJ and OK. | Terminated | December 24, 2013 | 231 | 4 | Computer
Z-1835-2013 | 65564 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit, model 8015, with version 9.12-Keyboard Processor and logic board assembly part number TC10006584. |  |  | June 17, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the Alaris PC units model 8015 (PC unit), version 9.12, because it is operating at an incorrect voltage and it could potentially experience a loss of communication between the PC Unit main Processor and Keyboard Processor. | Process design | The firm, Carefusion, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated June 17, 2013 to all customers who purchased the Alaris PC unit (model 8015) with version 9.12-Keyboard Processor. The recall letter informs the customers of the problems identified and the actions to be taken. The customers were instructed to either discontinue use of affect PC units until remediated or continue use on patients and follow the steps provided in the letter; fully charge Alaris PC units prior to use; during operation and when not in transport, keep Alaris PC units plugged in at all times; consider having additional back up Alaris PC units available and if you experience the described error code messages on the Alaris PC unit during infusion, you may continue the infusion(s) as programmed unless a programming change is required. Obtain a back-up Alaris PC unit to replace the device. In addition, if the customer experiences a SYSTEM ERROR (111.2000 or 111.2002), or a CHANNEL DISCONNECTED error on the Alaris PC unit contact CareFusion Customer Advocacy at 888-812-3266; 24 hours a day, 7 days a week; email: customerfeedback@carefusion.com. The customers were also instructed to promptly complete and return the enclosed Customer Response Card.CareFusion will provide an update to correct the voltage on the keyboard processor for those affected Alaris PC units. Carefusion will contact all affected customers within 60 days to initiate the scheduling process for the remediation.Customers are instructed to contact CareFusion Support Center at (888) 562-6018 or at supportcenter@carefusion.com for recall related questions; hours 7:00am to 4:00pm PT, Monday-Friday. For adverse reports, customers are instructed to contact Customer Advocacy at (888) 812-3266; 24 hours a day, 7 days a week or at customerfeedback@carefusion.com. For Technical questions, customers are instructed to call (888) 812-3229; hours 6:00pm to 5:00pm PT, Monday-Friday or at dl-us-inf-techsupport@ca |  | 26564 | 26,564 total units (26,220 units for US) | Worldwide distribution: US (nationwide) and countries of: Australia, Canada and Europe. | Terminated | May 20, 2014 | 337 | 26564 | Computer
Z-1872-2013 | 65724 | No | FZZ | Skin marker. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Non-sterile Viscot Surgical Skin Markers |  |  | July 10, 2013 | 2013 | Precision Dynamics Corporation | PDC Healthcare initiated this recall of Non-Sterile Viscot Part Number 1437-100 Surgical Skin Markers, becuase they were inadvertently mislabeled as PN STER-VIS Sterile Skin Marker Pen with Labels and Flex Ruter. | Component change control | Precision Dynamics Corp. (PDC) sent a Urgent: Skin Marker Pen Recall letter dated July 10, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.All affected customers have also been contacted by phone and email as of July 12, 2013We want to take every precaution to ensure patient safety.1. Immediately examine your inventory2. Quarantine the product subject to recall3. Complete and return the enclosed response form as soon as possibleYour PDC Healthcare customer service representative will contact you directly to arrange the return of the product. In addition, if you may have distributed this product, please identify your customers and notify them at once of this product recall. Your customer notification may be enhanced by including a copy of this recall notification letter.If you have any questions please call the PDC Healthcare Returns Department at 800-435-4242 ext. 5003.Thank you for your understanding and assistance related to this matter. |  | 1200 | 1,200 pens | US Distribution including the states of SC, TX, TN, WA, OK. | Terminated | January 06, 2014 | 180 | 1200 | Not_Computer
Z-1874-2013 | 65778 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MRidium 1145 Dose Reduction System (DERS) drug library kit. (This is an option for the MRidium 3860+MR Infusion Pump System). |  |  | July 01, 2013 | 2013 | Iradimed Corporation | The Dose Error Reduction System (DERS) can indicate an incorrect recommended value for the rate via a specific key sequence during the initial infusion setup. This can result in a risk of over-infusion or under-infusion if the infusion is started with this rate value. | Software Manufacturing/Software Deployment | The firm, Iradimed Corporation, notified by phone and sent an "URGENT MEDICAL DEVICE RECALL NOTICE" via email dated July 1, 2013 all consignees/customers. The notice described the product, problem and action to be taken. The customers were instructed to identify the MRidium 3860+ MR Infusion Pumps which have the 1145 DERS Library card (AM05) installed. If these are not installed, no further action is required at this time; however, if they are installed-remove the 1145 DERS Library Card from the 3860 Pump using the instructions provided. Product will be reconditioned by installing a software upgrade that corrects this problem. The updated software (version 3.5.1) can be seen during the 3860+ Pump's power on cycle.Should your users have any questions following these instructions, please contact Iradimed Customer Service; 866-677-8022 (toll free) or 407-677-8022 ext. 100 or via e-mail - recallsupport@iradimed.com |  | 176 | 176 individuals kits | Worldwide distribution: US (nationwide) including states of: AL, CA, CO, FL, GA, IL, KY, LA, MD, MA, MI, MN, NJ, NY, NC, OH, PA, RI, TN, TX, UT, VA, and WA; and countries of: Australia, Austria, Denmark, Epsom, France, Germany, London, Netherlands, Saudia Arabia, and United Kingdom. | Terminated | July 13, 2015 | 742 | 176 | Computer
Z-1871-2013 | 64985 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Drill Guide - Chameleon FacetFuse Fixation System, Model: 11-00028 Rev. A |  |  | November 09, 2012 | 2012 | SpineFrontier, Inc. | Prior to use in surgery, the Drill Guides were presenting noticeable degrees of denting and/or bending. As a result, drills and taps could not be directed down guides during surgery due to mechanical interference. | Device Design | Spine Frontier sent an Advisory Notice dated November 9, 2012, via email to all affected customers. Customers were advised that their FacetFuse Drill Guides need to be replaced. Customers were instructed to write their number ( included in the email) on the outside of the package and provide us with the tracking information. A replacement part has been sent out to you.If you have any further questions please call (978) 232-3990. |  | 13 | 13 drill guides | US Distribution including Puerto Rico and the states of CO, KS, MA and TX. | Terminated | June 05, 2015 | 938 | 13 | Not_Computer
Z-1873-2013 | 63605 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter AU5800 Clinical Chemistry Analyzer, All Serial Numbers. |  |  | October 26, 2012 | 2012 | Beckman Coulter Inc. | Beckman Coulter is recalling the AU8500 Clinical Chemistry Analyzer because there is no labeling on the diluted detergent tank to show the proper disposition of the on/off valve. | Software design | A recall notification letter was forwarded to all their customers on 11/1/12 who purchased the AU5800 Clinical Chemistry Analyzer. The customers were informed of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Reponse form within 10 days. Customers with questions are instructed to call (800) 854-3633 or contact their local BCI Representative. |  | 338 | 338 units (11 units in the US) | Worldwide Distribution-USA (nationwide) and Puerto Rico and the countries of Germany, China, Austria, Brazil, Italy, Argentina, Korea, Vietnam, Hungary, Ireland, France, Spain, Turkey, Switzerland, Israel, Australia, India, United Kingdom, Netherlands, Malaysia, Russian Federation, Japan, Taiwan, and Canada | Terminated | October 18, 2013 | 357 | 338 | Not_Computer
Z-1875-2013 | 65605 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobiletta Mira |  |  | May 02, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Field Safety Alert about the potential risk to users when operating the Mobilett Mira. When a Mobilett Mira is regularly exposed to extensive mechanical strain (driven over high thresholds for doors or elevators), it is possible that one of its main wheel shafts could break because of material fatigue, resulting in the chassis falling on the ground on the affected wheel side. This | Device Design | Siemens sent a Field Safety Notice dated May 2. 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed to :1) Never step or sit on the unit or otherwise put substantial weight on it.2) Avoid moving the unit over high thresholds for doors or elevators where possible, or do so carefully and slowly. 3) Make sure that no person had his/her extremities under the chassis, also when the unit is not moving. Update instructions were also released to alert customers of the potential risk and to provide a hardware change which is intended to improve corrosion resistance.An acknowledgement of receipt form was also sent out for the field safety notice and requested to be returned to Siemens.For further questions please call (888) 826-9702. |  | 39 | 39 | US Distribution including the states of IA, SD, NC, WI, OH, NE, TX, KS, ID, MI, GA, AR and IL. | Terminated | February 24, 2015 | 663 | 39 | Not_Computer
Z-1851-2013 | 65689 | No | BRZ | Intravascular administration set. | General Hospital | General Hospital | 510(k) | List No. 14203-28; BLOOD SET; 200 Micron Filter, 80 Inch, Non-DEHP; For Blood Product Administration; Made in Costa Rica; Product Inquiries should be directed to Hospira., Inc., Lake Forrest, IL 60045 USA |  |  | April 01, 2013 | 2013 | Hospira Inc. | It was discovered that there was a potential for the piercing pin on certain Hospira blood sets to puncture the outer wall of blood bags during insertion of the pin into the blood bag. | Device Design | Hospira sent an Important Safety Information letter dated April 1, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to exercise EXTREME CAUTION when using Hospira blood bags and to make sure that all instructions for use are included with the blood bag and their facility's protocol for spiking blood bags are completely followed in order to minimize the possibility of puncturing the outer wall of the blood bag. Customers with issues were instructed to call Hospira at 1-800-441-4100 or email ProductComplaintsPP@hospira.com. Customers were asked to complete the attached replay formand return it to the fax number on the form. If customers distributed the product further they were instructed to notify their accounts that received the affected product and ask them to contact Stericycle at 1-866-608-3936 to receive a reply form. Customers with questions were instructed to contact Hospira at 1-800-441-4100.For questions regarding this recall call 224-212-2000. |  | 1152 | 1,152 units | Worldwide Distribution - USA (nationwide) and Puerto Rico and Internationally to Costa Rica. | Terminated | May 24, 2017 | 1514 | 701870 | Not_Computer
Z-1876-2013 | 65757 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Mullins Transseptal Adult 7F, REF 008551. |  |  | July 08, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic identified a specific labeling issue for specific lots of Mullins Transseptal Catheter Introducer Sheath, model 008551, in which packages labeled as Mullins 7F (French) may actually contain an 8F(French) dilator. Medtronic is requesting the return of potentially mis-labeled Mullins Transseptal Catheter Introducer 7 French Sheaths. Due to the differences in size, an 8F dilator will not | Packaging change control | Consignees were sent on 7/8/2013 a Medtronic, Inc " Urgent Medical Device Recall" letter dated 8 July 2013. The letter was addressed to Healthcare Professionals. The letter identified the problem and the product involved in the recall. Advised consignees to remove and quarantine the product from their inventory and return the product. Requested consignees to complete the attached Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612. If the product has been further distributed advised consignees to please notify them of the recall. |  | 269 | 269 | Worldwide Distribution-USA (nationwide) including DC and the states of AZ, AK, CA, DE, GA, IL, KY, MA, MI, MN, MS, NV, NY, NC, OH, PA, TN, and TX, and the countries of AUSTRALIA, BRAZIL, CHILE, INDIA, and SINGAPORE. | Terminated | March 06, 2014 | 241 | 269 | Not_Computer
Z-0476-2013 | 63187 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Enterprise 8000 Bed; an AC-powered adjustable hospital bed; ArjoHuntleigh AB, Eslov, Sweden; Models 8000, 8001, 8002, 8003, 8500, 8600, 8700 and 8800 |  |  | November 16, 2012 | 2012 | Arjo, Inc. dba ArjoHuntleigh | ArjoHuntleigh has received reports of unintended movement where the backrest and/or thigh sections of the bed have moved without anyone consciously pressing the control button. These incidents have occurred in High Dependency and Intensive Therapy environments where unintended movement of the patient may have serious consequences. | Under Investigation by firm | Please be aware that this is not a new recall. Posting to the Internet was delayed due to a technical error.ArjoHuntleigh sent an Urgent Field Correction Recall letters dated November 16, 2012, to all affected customers, informing the accounts of reports of unintended movement of the backrest and/or thigh sections of the bed without intentional activation of the control buttons, which could have serious consequences due to the patient's links to other life supporting devices. The accounts were requested to disseminate the information to all users of the affected beds, to attach the enclosed warning labels to the outside of both head end safety sides immediately adjacent to the controls, to include a copy of the notice in the Instruction for Use Manuals for their beds, and to complete the Customer Response Form online, using the enclosed card's instructions. Questions were directed to the ArjoHuntleigh at 1-800-323-1245, ext 57985. |  | 26135 | 94 beds in U.S., 31 beds in Canada, 26,010 beds in ROW | Worldwide Distribution-USA (nationwide, including Puerto Rico) and the countries of Algeria, Angola, Australia, Austria, Bahrain, Belgium, Bosnia and Herzegovia, Botswana, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cuba, Czech Republic, Denmark, Ecuador, Egypt, Falkland Islands, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jordan, Kenya, Kuwait, Lebanon, Libyan Arab Jamahira, Lithuania, Malaysia, Malta, Mauritius, Mexico, Morocco, the Netherlands, New Zealand, Nigeria, Nordic, Norway, Oman, Panama, Poland, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sri Lanka, Switzerland, Syrian Arab Republic, Thailand, Turkey, United Kingdom, United Arab Emirates, Uruguay, Venezuela and Yemen. | Open, Classified |  |  | 29474 | Not_Computer
Z-1881-2013 | 65162 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Systems TSH3 Ultra ReadyPack ((REF 06491072 - 100 Tests, REF 06491080 - 500 Tests, REF 04862625 - 2500T). |  |  | May 10, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | A rare variant of TSH, identified in a small cluster of patients, is not detected. | Device Design | On May 10, 2013, an Important Safety Information notice was sent to all affected Siemens Healthcare Diagnostic customers in the United States. On May 10, 2013, an Urgent Field Safety Notice was sent to all customers outside of the United States. These notices inform the customer that there is a rare variant of TSH, identified in a small cluster of patients that is not detected by some Siemens assays. This notice informs customers that they should continue to use the affected TSH assays as the observed rate of occurrence during a 30-month time period was 0.6 X 10 -7 and there are strong mitigating factors. These mitigating factors include co-interpretation of TSH results in conjunction with thyroxine and TS, patient history and clinical signs and symptoms. Although Siemens believes a look back is not required as the expected frequency of the variant is extremely rare, these notices state that this decision should be made by the medical director. |  | 172913 | 172913 | Worldwide Distribution-USA (nationwide) and the countries of Canada, AE, AR, AT, AU, BA, BD, BE, BG, BO, BR, BY, CH, CL, CM, CN, CO, CR, CY, CZ, DE, DK, DO, DZ, EC, EE, EG, ES, FI, FR, GB, GR, GT, HK, HN, HR, HU, ID, IE, IL, IN, IR, IS, IT, JO, JP, KA, KR, KW, KZ, LB, LK, LT, LV, LY, MA, MK, MT, MX, MY, NI, NL, NO, NP, OM, PA, PE, PH, PK, PL, PT, PY, QA, RO, RU, SA, SE, SG, SK, SV, SY, TH, TM, TN, TR, TW, UA, UY, UZ, VE, VN, XS, and ZA. | Open, Classified |  |  | 330020 | Not_Computer
Z-0834-2013 | 64244 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | VENTANA (a member of the Roche Group) DISCOVERY ULTRA; REF/MU 750-601; Research Use Only |  |  | September 28, 2012 | 2012 | Ventana Medical Systems Inc | Potential for a waste fluid overflow condition to accur while running decontamination cycles on the BenchMark ULTRA automated staining platform or the DISCOVERY ULTRA automated staining platform.The waste fluid drainage system may become restricted by build-up of debris in the tubing and in-line filter, thereby inhibiting optimal flow rates. This type of overflow event can lead to migration of flu | Device Design | The firm, Roche sent a "Technical Bulletin" and an letter dated October 3, 2012 to its customers. The letter and bulletin described the product, problem and actions to be taken. The customers were informed that Ventana is aggressively working on solutions to this issue. Until mitigations to the overflow condition are available for implementation; inform all personnel in your organization that work with the BenchMark ULTRA or DISCOVERY ULTRA instrument about the potential for fluid overflow to occur during decontamination cycles; follow the instructions provided in the letter to avoid the documented failure mode and follow the instructions provided in the technical bulletin.Contact your local support center with any questions or concerns or call 520-877-7191. |  | 54 | 54 instruments | Worldwide distribution: USA (nationwide) and country of: Germany. | Terminated | February 24, 2016 | 1244 | 1789 | Not_Computer
Z-1907-2013 | 65470 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | bioMerieux Vitek 2 Systems Software version 6.01 software, used with the VITEK 2 automated system. Kit, V2S 6.01 RP5800-WES7 Media. Part number: 413861. |  |  | June 14, 2013 | 2013 | Biomerieux Inc | A product issue has been identified for customers using software version 6.01 and are interfaced with a Laboratory Information System (LIS). During the update of software from version 5.xx to 6.01, the Bid-directional Computer interface (BCI) antibiotic codes are not updated properly for Cefovecin, Prulifloxacin and Ceftaroline. For the referenced antibiotics, the firm code in the reference table | Software change control | The firm, bioMerieux, sent an "Urgent Product Correction Notice" letter dated June 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure the Urgent Product Correction Notice and the Instructions for modifying BCI Translation Codes are distributed to all appropriate personnel in your organization; follow the instructions in the attached file (Instructions for modifying BCI Translation Codes) to resolve the discrepancy associated with Cefovecin, Prulifloxacin and Ceftaroline in the translation table and complete and return the Attachment A: Acknowledgement Form via Fax to: Attn: RA North American ComOps at (919)287-2542.If you require additional assistance or have any questions, please contact your local bioMerieux Clinical Customer Service organization at (800) 682-2666 or Industrial Customer Service organization at (800) 634-7656. |  | 12384 | 12,384 total kits | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Bahrain, Belarus, Benelux, Brazil, Bulgaria, Canada, Chile China, Costa Rica, Croatia, Czech Republic, Dominican Republic, Dutch Antilles, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Israel, Italy, Japan, Kazakhstan, Korea, Kuwait, Libya, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, Nicaragua, Nigeria, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Serbia, Singapore, Slovenia, Spain, Sweden, Switzerland, South Africa, Thailand, Tunisia, United Arab Emirates, United Kingdom, Uruguay, and Vietnam. | Terminated | June 01, 2015 | 717 | 12384 | Computer
Z-1879-2013 | 65486 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Sonopet Ultrasonic Aspirator Console |  |  | June 24, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | A supplied component exhibiting corrosion was built into the power supply boards of the Sonopet Consoles listed above. Theissue is limited to 2 lots of the supplied component. The presence of the corrosion on the relay of the power supply board canpotentially lead to a conduction defect which could cause the irrigation pump to stop working. The purpose of the irrigationpump is to supply sali | Component change control | Stryker sent an URGENT Medical Device Recall Notice letter dated June 24, 2013 to all affected customers. The letter identified the affected product, product issue, risk to health and actions to be taken. Customers were instructed to locate affected product and hold for upgrade, use enclosed response form to document inventory and receipt of notice, and in case of further distribution forward the recall notice. Stryker will dispatch a technician to perform the upgrade to the affected product. For questions regarding this recall please contact Stryker Instruments: Monday-Friday 8am-5pm (EST)Kara Spath, Regulatory Affairs269-389-4518kara.spath@stryker.com. |  | 215 | 215 units total | Worldwide - USA Nationwide including the states of CA, CO, FL, HI, IA, IL, KS, KY, LA, MD, MI, MN, MS, NC, NH, NJ, NV, NY, OH, PA, TX, VA, ITALY, and the countries of AUSTRALIA, CANADA, UNITED KINGDOM, NETHERLANDS, INDIA, JAPAN, NEW ZEALAND, SWITZERLAND, HONG KONG, and TAIWAN. | Terminated | January 30, 2014 | 220 | 215 | Not_Computer
Z-1888-2013 | 65819 | No | MGO | Hydrophilic wound dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CVS Pharmacy Wound Closure Adhesive Surgical Tape Strips |  |  | July 26, 2013 | 2013 | ASO, LLC | This recall has been initiated by Aspen Surgical Products, Inc., due to the inability to confirm sterility. | Package design/selection | ASO sent a letter on July 26, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine your inventory and stop using/distributing the product. Place in quarantine to be returned or destroyed. ASO customers will receive credit for returned product or a returned completed Certificate of Destruction. Please use Return Authorization Number 1675 when returning the product.Please complete and return the enclosed response form as soon as possible. Return the form to:Address:ASO LLCRegulatory Affairs300 Sarasota Center Blvd.Sarasota, FL 34240Fax:941-378-9040Attn: Regulatory AffairsFor questions regarding this recall call 888-364-7004, ext 146. |  | 2689560 | Total 2,689,560 pouches (2 cards/pouch) | Nationwide Distribution including FL, IA, IL, MI, NC, NY, OH, OR, PA, RI, TX, WA and WI. | Terminated | August 04, 2014 | 374 | 2689560 | Not_Computer
Z-1922-2013 | 65781 | No | FPD | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Nestle Nutrition, Compat, Enteral Feeding Tube, sterile, 8 French (REF 08310400), 10 French (REF 08310600), and 12 French (REF 08310500). |  |  | July 24, 2013 | 2013 | Nestle HealthCare Nutrition | Nestle HealthCare Nutrition, Inc. is initiating a recall on some of Compat NG Tubes because they have been delivered in an insufficiently sealed pouch, thus compromising the sterility claimed on the label. | Packaging process control | Nestle Health Science sent a "Medical Device Recall" letter dated July 24, 2013. The letter was addressed to Risk Manger. The letter described the product involved, problem and actions to be taken. Customers were advised that products currently in use should not be returned. Customers were requested to examine their inventory, quarantine the product, carry the recall to the user facility level and to complete and return the Acknowledgment / Response Form.For questions call Nestle HealthCare Nutrition Customer Service at 1-877-463-7853. |  | 37310 | 37,310 | US Nationwide Distribution - including the states of: AK, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NE, NJ, NM, NV, NY, OH, OK, PA, RI, SC, TN, TX, UT, VA, WA. | Terminated | June 13, 2014 | 324 | 37310 | Not_Computer
Z-1925-2013 | 65152 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension(R) Mass Creatine Kinase MB Isozyme Calibrator (RC420) |  |  | June 05, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received complaints that the lyophilized cake is not fully dissolved after the stated time in the MMB calibrator Instructions For Use (IFU). The frequency of this occurrence is low but if it occurs, gel-like clumps may be observed. | Material/Component Contamination | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated May 2013, to all consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed as follows: Until a permanent solution is identified, please follow the instructions listed in the letter in order to optimize calibration preparation; complete and return the attached Correction Effectiveness Check form via fax to the Siemens Technical Solutions Center at (302) 631-8467 and forward this notification to anyone to whom you may have distributed this product.If you have any questions, please contact the Siemens Technical Solutions Center or your local Siemens technical representative or call 800-441-9250. |  | 7470 | 7,470 | Nationwide distribution: US including states of : AK, AL, AR, AZ, CA, CO, CT, DC, GA, GU, FL, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | March 18, 2015 | 651 | 7470 | Not_Computer
Z-1924-2013 | 65706 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension Tacrolimus Flex reagent cartridge (DF107) and Dimension TACR CAL (DC107) |  |  | July 01, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens initiated a recall due to confirmed complaints of low patient sample recovery with Dimension TACR Flex Reagent cartridge (DF107) lot BB4087 and its linked calibrator (DC 107) lot 3BD029. Falsely low Tacrolimus results across the assay range may lead to alterations in dosing. | Process control | Siemens sent an Urgent Medical Device Recall letter dated July 2013 to affected customers The letter identified the affected product, problem, potential risk, and actions to be taken by customers. Customers were requested to discuss the letter with thier laboratory director to determine if additional follow up is appropriate for patient tests conducted with the affected device lots. Additionally, customers were asked to discard any remaining inventory and to complete and return the field correction effectiveness check for replacement product. They were also asked to forward the letter to anyone wha had received this affected product.For questions contact your Siemens Technical Solutions Center or your local Siemens Technical support representative. |  | 5423 | 5423 | USA Nationwide Distribution including the states of: TX, NC, GA, TN, UT, NJ, MI, VA, AR, LA, MO, FL, AZ, CA, DC, PA, NH, WI, SD, SC, NM, MD, and IN. | Terminated | August 26, 2014 | 421 | 5423 | Not_Computer
Z-1887-2013 | 65482 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Cordis OPTEASE Retrieval Inferior Vena Cava (IVC) Filter***REF Cat. No.***466F220A and 466F220B***Vena Cava Filter and Introduction Kit***(BRITE TIP Catheter Sheath Introducer and Angiographic Vessel Dilator)***Cordis***a Johnson and Johnson Company***ENDOVASCULAR*** |  |  | March 29, 2013 | 2013 | Cordis Corporation | Correction to provide clarification and additional labeling to minimize likelihood on implanting filter backwards. | Labeling design | Cordis Corporation sent an "URGENT Medical Device Correction" notification to their consignees in the US by letter dated April 3, 2013 and an e-mail was sent to the Canadian consignees. The letter identified the affected product, problem and actions to be taken. Consignees were instructed to read the description and recommendations sections carefully, then sign and return the enclosed Acknowledgement Form directly to Cordis to their fax number(908) 429-8287 or Scan and e-mail to MiamiQA@crdus.jnj.com. The letter states to pass on the notice to anyone in the consignees' facilities that need to be informed. Additionally, the notification request consignees to maintain awareness of this notification until the information has been incorporated in the Cordis OPTEASE Vena Cava Filter labeling. For questions related to the issue contact Cordis Endovascular Medical Information Line 877-338-4235. |  | 33000 | 33,000 units | Worldwide Distribution - USA Nationwide in the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, Ny, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, including Hawaii and Puerto Rico and the countries of Canada &amp; Panama | Terminated | October 27, 2014 | 577 | 33000 | Not_Computer
Z-1923-2013 | 65322 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer Multiple Sample and Direct Draw Luer Adapter |  |  | May 01, 2013 | 2013 | Becton Dickinson &amp; Company | BD Vacutainer Multiple Sample and Direct Draw Luer Adapter may have blood leakage and low draw volume. | Under Investigation by firm | BD sent a Product Advisory Notice dated May 3, 2013, via UPS to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discard filled tubes since it is possible they will contain less blood volume (low draws), replace the Luer Adaptor unit being used with a fresh unit, and continue to draw blood samples using the fresh unit. Customers with questions were instructed to contact BD Technical Service at 1-800-631-0174, Option 4, or contact them through website: www.bd.com/vacutainer/contact.For questions regarding this recall call 201-847-6800. |  | 37906400 | 37,906,400 units | Nationwide Distribution | Terminated | December 04, 2015 | 947 | 37906400 | Not_Computer
Z-1932-2013 | 64993 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | T-Handle, Strike Plate, Push Connection, Part Number: SI50023 Rev A |  |  | August 29, 2012 | 2012 | SpineFrontier, Inc. | During an operation the strike plate of a T-Handle loosened and separated from the T-Handle. | Process design | On August 29, 2011, The Supply Chain Team was notified of the recall via email. They were informed that all handles are being pulled from the field. This resulted in the return of all T-Handles to SpineFrontier.Further questions please call (978) 232-3990. |  | 11 | 11 | US Distribution including the states of MA, NV and TX. | Terminated | October 18, 2013 | 415 | 11 | Not_Computer
Z-1926-2013 | 64984 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Rev E PedFuse Return Screw Inserter, Part Number: 11-80030 |  |  | March 14, 2013 | 2013 | SpineFrontier, Inc. | Inserters have a set screw that fixes a collar to the inserter. The screw set can loosen and fall out of the inserter which allows the collar to slide off the proximal end of the inserter. When the collar is not in place the locking sleeve can also slide off the inserter. | Device Design | Spine Frontier sent an email dated December 10, 2012, to the affected customers that a Pedfuse Inserter with part number PM3979 needed to be returned.Further questions please call (978) 232-3990. |  | 8 | 8 | Distribution only in Texas. | Terminated | June 03, 2015 | 811 | 8 | Not_Computer
Z-1933-2013 | 65653 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Indus Invue Screw Caddy, Model: SI70097 |  |  | June 21, 2013 | 2013 | SpineFrontier, Inc. | Difficulty removing screws from the Invue caddies. | Device Design | SpineFrontier sent a Urgent Invue Caddy Advisory letter dated July 10, 2013, to all affected customers. Customers were advised the removal of the Invue Screw Caddy, SI70097 from the field for modification. Customers were asked to return the parts to SpineFrontier and to complete an Acknowledgement Form indicating that they have read the notice and will initiate the immediate return of the parts noted above.Please feel free to contact us at (978) 232-3990.On 8/21/2013 the recall was expanded. |  | 13 | 13 | Worldwide Distribution - USA, including the states of CT, VA, MA, FL, TX, CO and MS. and the country of Jamaica. | Terminated | January 11, 2016 | 934 | 13 | Not_Computer
Z-1930-2013 | 65843 | No | LPP | Cutaneous oxygen (PcO2) monitor. | Anesthesiology | Anesthesiology | 510(K) Exempt | TCM4 series Base Units, and spare part Power Supply, Radiometer Medical ApS, Akandevej 21, BrÂżnshÂżj, Denmark, www.radiometer.com. The devices are intended for continuous monitoring of carbon dioxide and oxygen partial pressures, oxygen saturation, and pulse rate. |  |  | November 27, 2012 | 2012 | Radiometer America Inc | RADIOMETER became aware that the TCM4 series Base Unit does not comply with IEC 60601-1 Sub Clause 15B. | Device Design | On 11/30/2012 the firm sent an "Urgent Field Safety Notice" to their customers. The notification explained the issue and provided recommended actions. Customers may contact 1-800-736-0600 option 4 with questions. |  | 5874 | 5,874 units | Worldwide distribution: US (nationwide) AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, TX, UT, VA, VT, WA, WI, WV &amp; WY; and country of: Canada. | Terminated | June 16, 2014 | 566 | 5874 | Not_Computer
Z-1928-2013 | 65370 | No | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | R3 (TM) XLPE ACETABULAR LINER UHMWPE, 20o OVERHANG, 36 MM ID, 62 MM OD, QTY: (1), STERILE H2O2, REF 71335762. |  |  | May 29, 2013 | 2013 | Smith &amp; Nephew Inc | One batch of 60 mm 20o R3 XLPE liner may be laser etched and packaged as a 62 mm 20o liner and one batch of 62 mm 20o R3 XLPE liner may be laser etched and packaged as a 60 mm 20o liner. | Packaging process control | All affected Smith &amp; Nephew Sales Representatives were notified of problem and the recall via email on 5/29. They were instructed to cease use and distribution of the product and quarantine for return. Surgeons were notified by letter beginning on 6/12. International consignees were notified on 05/29 via email. |  | 53 | 53 units | Worldwide Distribution-USA including the states of PA, WA, TX, FL, IA, MI, WI, TN, KY, IN, NY, MO and OR, and the countries of India, Canada, Chile, Venezuela, Dubai, South Africa, Australia, Poland, Romania, Belgium, Germany and United Kingdom. | Terminated | March 22, 2016 | 1028 | 53 | Not_Computer
Z-1931-2013 | 65786 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System, Software Version 3.0.0 revision 13 or earlier; Proprietary and Trade Name: MHI-TM2000/ VERO (TM); Common Name: Medical Linear Accelerator. |  |  | June 21, 2013 | 2013 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Software anomaly: Mitsubishi Heavy Industries, Ltd. became aware of a potential safety issue associated with the VERO (TM) Linear Accelerator System. The VERO (TM) may save treatment delivery record into wrong location under certain conditions. | Software Design Change | The firm, Mitsubishi Heavy Industries, Ltd. (MHI), sent an "URGENT FIELD SAFETY NOTICE" letter dated June 21, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: If any minor failure occurs, please always press [ALARM RESET] button on operator console first; and if [EXIT] button is clicked after a minor failure occurs, please click [YES] on the pop up dialog to finish treatment completely. MHI is developing a correction for this issue. Tentatively planned availability date: End of October 2013. The correction will be free of charge and a MHI Customer Service representative will contact affected customers to schedule its installation as the correction becomes available.If you need any further information or support concerning this issue, please feel free to contact our Customer Service representative: VERO GmbH, Kapellenstrasse 12, 8566 Feldkirchen Germany, at +49-89-99-1568-44 or email: franz.gum@vero-sbrt.com. |  | 14 | 14 (2 in US) | Worldwide distribution: US (nationwide) including states of TX and FL; and countries of: Belgium, Italy, Japan and Korea. | Terminated | March 23, 2017 | 1371 | 14 | Computer
Z-1927-2013 | 65444 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Invue Split Driver, Assembly with Handle, Model SI70006 Rev G. Used in conjunction with the Invue Anterior Cervical Plate System under 501(k) clearance number K121060. |  |  | May 10, 2013 | 2013 | SpineFrontier, Inc. | Device design. | Device Design | SpineFrontier sent an URGENT Advisory Notice dated June 27, 2013, to all affected consignees regarding the use of Invue Drivers. The notice advises users that the Split Tip Stylet Driver Assembly (11-70099) and Sold Driver, AO Connection (11-70070 or 11-70077) are being shipped in place of the Split Driver Assembly w/Handle (SI70006). The uses of the replacement parts are explained. Consignees are instructed to complete the Acknowledgement Form and return it to SpineFrontier.On June 28, 2013, customers were sent a second Advisory Notice requesting the removal of the Split Driver Assy w/ Handle Part Number SI70006 Lot Numbers 027104 and 027105. Consignees were instructed to complete the Acknowledgement Form and return it to SpineFrontier.Consignees with questions were instructed to call 978-232-3990. |  | 12 | 12 devices | Nationwide Distribution including Texas, Ohio, and Missouri. | Terminated | July 20, 2015 | 801 | 12 | Not_Computer
Z-1817-2012 | 61253 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | TX1 Tissue Removal System Console. Model: TX1 |  |  | February 22, 2012 | 2012 | American Optisurgical Inc | Two reasons for recall.1. Incorrect pressure relief valve installed inside console, leading to a remote chance of a risk to health. 2. Technical Specifications, under Irrigation, the Operating Pressure value is incorrect. | Nonconforming Material/Component | The firm, American Optisurgical Inc., sent an "URGENT: MEDICAL DEVICE RECALL" letter via email on February 23, 2012. An updated "URGENT: MEDICAL DEVICE RECALL" recall letter dated March 1, 2012 was sent via email on March 1, 2012 to its customer. This updated recall letter described the product, problem and actions to be taken. The letter notified the two reasons for recall: " An incorrect pressure relief valve has been installed inside the console. Should the regulator that controls pressure delivered to the inflation cuff fail, the pressure relief valve would not open until a pressure of 10psi was achieved, which may present a health risk". "The Operator's Manual included with the system, Rev. 1, has a typographical error. Specifically, in Section E. Technical Specifications, under Irrigation, the Operating Pressure value is incorrect". The updated recall included the above second reason for recall. The firm instructed the customer, for replacement of the pressure relief valve, and to return the revision 1 operator manual to the company's address using a pre-paid FedEx account. For questions call 949.580.1266 M-F, 7:00am to 5:00pm PST. |  | 5 | 5units | Nationwide distribution: MN only. | Terminated | June 15, 2012 | 114 | 5 | Not_Computer
Z-1943-2013 | 65133 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | SIGNA INFINITY MR SYSTEM WITH EXCITE TECHNOLOGY |  |  | November 21, 2011 | 2011 | GE Healthcare, LLC | Screen save images are included in wrong exam due to duplicate UID creation. This may cause a safety issuedue to a potential patient data mismatch.While the issue has only been observed at one customer site on the Signa HDxT 1.5T system, internal investigation revealed that the issue also impacts multiple GE healthcare MR Systems. | Software design | GE Healthcare communicated internally to their Service Team the directions for completing the correction on the customers' systems. GE Healthcare's recall strategy was to update all affected systems that were distributed. At this time,Week of 02/06/2012, deployment of this recall is considered complete.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2608 or email paulcorrigan@ge.com. |  | 7016 | 7016 (2919 US; 4097 OUS) | Worldwide distribution: USA (nationwide) including DC and Puerto Rico and countries: ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIDJAN, BELARUS, BELGIUM, BOLIVIA, BOSNIA-HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREAT BRITAIN, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, KAZAKHSTAN, KENYA, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIAN, LUXEMBOURG, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU. PHILIPPINES, POLAND, PORTUGAL, REPUBLIC OF MACEDONIA, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAK REPUBLIC, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TUR,EKY, UKRAINE, UNITED ARAB EMIRATES, URUGUAY, VENEZUELA, VIETNAM, YUGOSLAVIA, MALAWI, MAURITANIA, NIGERIA, UKRAINE, IVORY COST, OMAN, ZAMBIA, UZBEKISTAN, ESTONIA, IRAQ, and JAMAICA. | Terminated | October 22, 2014 | 1066 | 7016 | Not_Computer
Z-1938-2013 | 65698 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | C-QUR V-Patch Mesh (All sizes and shapes). |  |  | July 19, 2013 | 2013 | Atrium Medical Corporation | Additional Instructions for Use and Storage Conditions as Coated mesh can adhere to the inner packaging liner due to exposure to high humidity conditions | Package design/selection | Atrium Medical issued Recall Letter via UPS and e-mail on 7/19/13 to the accounts and field representatives. The notification identifies the problem, product, and risk factors. If the Product is exposed to excessive humidity for an extended period of time, then the increased humidity occurring inside the pouch can potentially cause the coating on the mesh to strongly adhere to the inner handling sleeve. A reply for is requested to be completed to acknowledge receipt to the notification.Additional language to the instructions for use (IFU) to include:Prolonged exposure to high humidity may result in increased rate of adherence of the C-QUR mesh to its handling sleeve. Store in a Controlled Room Temperature (25 C / 77 F ) or less. Brief exposure to up to 40 C (104 F ) is acceptable.Questions please contact Atrium Medical Customer Service at 1- 800-528-7486 Monday through Friday 9:00 am to 5:00 pm EDT. |  | 95286 | 95,286 units | USA (nationwide) including Puerto Rico and the countries of Australia Austria Bahrain Brazil Canada Chile Colombia Dominican Republic Ecuador El Salvador France Germany Great Britain Greece Honduras Hong Kong India Ireland Israel Italy Japan Jordan Korea Malaysia Mexico Netherlands New Zealand Nicaragua Norway Panama Peru Portugal Romania Saudi Arabia Singapore South Africa Spain Sri Lanka Switzerland Taiwan Thailand Turkey and Venezuela. | Terminated | March 18, 2016 | 973 | 145251 | Not_Computer
Z-1940-2013 | 65333 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE LightSpeed CT750 HD |  |  | April 05, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare conducted a recall involving 2 events affecting the GE LightSpeed CT750 HD. Event 1 - There was a report regarding CT number variance in time line at posterior fossa region of head perfusion. When beam hardening and helical artifacts in this region are reconstructed into the perfusion map, these artifacts will give subtle differences in the CT numbers.Event 2 - The software does no | Software design | GE Healthcare's recall strategy was to update CT750HD CT Systems with software 10MW25.6 on all systems that were distributed. Updates started on 4/5/2011. At this time, deployment of this recall is considered complete. The recall was completed 2/2/2012.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2198 or email paulcorrigan@ge.com. |  | 244 | 244 | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CT, FL, GA, HI, IL, IN, IA, KY,ME, MD,MA, MI,MN, MS, NV, NH, NJ, NY, NC, PA, TN, TX, VA, WA,WV, and WI; and countries of: AUSTRALIA, BELGIUM, BRAZIL, CANDA, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN,JORDAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, NEW ZEALAND, POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SWEDEN, SWITZERLAND, TUKEY, and UNITED KINGDOM. | Terminated | August 11, 2014 | 1224 | 244 | Computer
Z-1935-2013 | 65563 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | 5.0mm Unit Rod 270mm |  |  | June 05, 2013 | 2013 | Synthes USA HQ, Inc. | DePuy Synthes has issued a Medical Device Correction notice for the 5.0mm Stainless Steel rods. This product was produced using a finishing process not identified as part of the manufacturing specification. No adverse event reports associated with the use of this product were received, and an extensive engineering and medical safety review concluded that the product poses no incremental risk to t | Nonconforming Material/Component | On June 5, 2013, De Puy Synthes notified by letter the direct acc acccounts which were identified to have received 5.0mm stainless steel rods and requested return of the product by calling 1-800-479-6329 to obtain a return authorization number. Completion of the verification section of the letter, including negative responses was also requested. |  | 492 | 492 | Nationwide Distribution. | Terminated | September 21, 2015 | 838 | 492 | Not_Computer
Z-1944-2013 | 65583 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Healthcare, LOGIQ S8 Ultrasound System |  |  | August 13, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare became aware of a potential issue related to incorrect image measurement associated with the software of the LOGIQ S8 system. This issue was found internally. Users may notice that the system shows incorrect image measurement results on Pulsed Wave (PW) or TimeIntensity Curve (TIC) or depth measurement when using Virtual Convex mode and putting a cursor (Sample Volume, Depth measur | Software design | GE initiated a field action on 8/13/12. GE plans to update the software on all of the affected GE LOGIQ S8 Ultrasound systems that were distributed, at no cost to the customer.Should you have any questions or require additional information, please contact RA Leader, US Region by telephone at 262-548-2089 or email James.Giles@med.ge.com. |  | 755 | 755 | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CO, CT, DC, FL, GA, IL, IN, IA, LA, MA, MI, MS, MO, MT, NE, NV, NJ, NY, NC, OH, OK, PA,RI, SC, TN, TX, VA, WA, and WI; and countries of: ALGERIA, AUSTRALIA, AUSTRIA, BRAZIL, CANADA, COLOMBIA, CZECH REPUBLIC, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, IRELAND, ITALY, JAPAN, LATVIA, LEBANON, MALAYSIA, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, REPUBLIC OF MOLDOVIA, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and YUGOSLAVIA. | Terminated | September 15, 2014 | 763 | 755 | Computer
Z-1942-2013 | 64992 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Optima XR220amx, Optima XR200amx, and Brivo XR285amx. |  |  | August 07, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare conducted a recall on various mobile x-ray units (Optima XR220amx, Optima XR200amx, and Brivo XR285amx systems) because of an Erratic Drive Performance. | Software design | GE Healthcare recall strategy was to update all of the affected Optima XR220amx, Optima XR200amx, and Brivo XR285amx systems that were distributed. At this time ( 2/14/2013), deployment of this recall is considered complete.Should you have any questions or require additional information, please contact Regulatory Affairs Leader by telephone at 262-548-2198 or email paulcorrigan@ge.com |  | 458 | 458 (306 US; 152 OUS) | Worldwide distribution: USA (nationwide) including states of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, GUAM, IL, IN, IA, KY, LA, ME, MD, MA, MI, MN, MS, MO, NV, NH, NH, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, and WY; and countries of: Albania, Australia, Bahrain, Belgium, Canada, Chile, Czech Republic, Dominican Republic, France, Germany, Great Britain, Honduras, Hungary, Ireland, Israel, Italy, Japan, Korea, Kuwait, Netherlands, Nigeria, Norway, Romania, Saudi Arabia, Spain, United Arab Emirates and Venezuela. | Terminated | February 04, 2014 | 546 | 458 | Not_Computer
Z-1903-2013 | 65617 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova Max Plus Monitor System (Kit w/10 Count Vials) |  |  | July 15, 2013 | 2013 | Nova Biomedical Corporation | Reports of false abnormally high glucose readings from test strips | Material/Component Contamination | Nova Biomedical U.S.A issued an Urgent Medical Device Voluntary Recall Letter dated July 19, 2013 via registered/certified letters to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use of the affected product; contact Nova Diabetes Care Customer Service for replacement strips; do not stop testing blood glucose while waiting for replacement strips; if affected strips must be used follow precautions provided in notification letter. For questions call Nova Diabetes Care Customer Service at 1-800-681-7390.. |  | 1316420 | 1,316,420 kits | Worldwide Distribution - USA Nationwide and the countries of Chile, England , Peru, Nicaragua | Terminated | August 01, 2017 | 1478 | 58957354 | Not_Computer
Z-1946-2013 | 65337 | No | IOQ | Powered flotation therapy bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | IsoGel AIR 84, Support Surface. |  |  | July 12, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | Complaints were received from the field indicating that some Air Pumps were flashing and alarming when connected to the IsoGel AIR" surface upon turning the pump on and the setting it to the IsoGel AIR Mode. Upon further investigation, it was determined that the flashing/alarm being experienced was due to the air pressure being out of the expected range; indicating a blockage within the LAL tubi | Nonconforming Material/Component | On 7/12/13, an URGENT MEDICAL DEVICE RECALL NOTIFICATION was mailed to consignees. The notification addressed the affected product and issue, potential hazard, actions needed including faxing (269-488-8691) or emailing (ProductFieldAction@stryker.com) the attached reply form to confirm receipt of notification, and risk mitigation. Please contact the firm at 269-389-8129 if you have questions or concerns regarding this recall. Business hours are M-F 8am-5pm(ET). |  | 586 | 586 units | Nationwide Distribution-USA including the states of OH, ME, IL, CO , FL, TX, MI, GA, and VA. | Terminated | September 11, 2014 | 426 | 586 | Not_Computer
Z-1941-2013 | 65828 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.2mm Drill Bit. This is a component within the PediLoc Locking Plate System utilized to pre-drill a hole in bone prior to placement of a bone screw. |  |  | July 24, 2013 | 2013 | OrthoPediatrics Corp | OrthoPediatrics has initiated a voluntary recall of Part # 01-1050-002, 2.5mm Drill Bit, Lot #91300335. Only drill bits that are in their original packaging are being recalled. The description on the package of this drill bit has a wrong size in the description (3.2mm). The recall affects only packaged devices still in original packaging due to misbranding of the product. The drill bit itself has | Process design | On 7/24/13, OrthoPediatrics notified their consignees of this URGENT MEDICAL DEVICE RECALL via phone call. A follow-up email message was then sent with the recall notification, and then a hard copy of the notice was sent via certified letter. The notification addresses the affected product, recall issue, reply form instructions, and directs consignees to return affected product to OrthoPediatrics via a Fed-Ex return prepaid pouch. The firm directs consignees to forward the notification and reply form if the product was further distributed. |  | 68 | 68 | Nationwide distribution, including the states of MO, MI, WA, FL, NY, TN, CO, AZ, VA, TX, CA, MN, LA, NC, and UT. | Terminated | August 26, 2013 | 33 | 68 | Not_Computer
Z-1945-2013 | 65721 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens AXIOM Artis zeego |  |  | May 13, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a customer safety notice about the overlay of system information after a functional check of the brakes on the Artis Zeego system. | Software design | The firm, Siemens, sent an "Important safety notice" by letter on May 13, 2013 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to immediately notify the service department to clarify what further measures need to be taken, if after carrying out the regular brake test, the system can only be moved at a reduced speed. If this device has been sold and is therefore no longer in your possession, send this safety notice to the new owner and provide the new owner to Siemens.Update instruction AX012/13/S was issued, and a software update will be released as AX006/13/S and AX007/13/S which will modify the system software so that it will not be possible for information to be overlaid.If you have any questions call +49 (9191) 18 0. |  | 132 | 132 | Nationwide distribution: AL, AK, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KY, LA, MA, MD, MN, MO,MS, NC, NE, NH, NJ, NY, OH, PA, PR, SD, TN, TX, UT, VA, WA, and WI. | Terminated | November 06, 2014 | 542 | 132 | Computer
Z-1934-2013 | 65837 | No | MMO | Ear, nose, and throat manual surgical instrument. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(K) Exempt | Tray Surgical, ENT (various instrument trays) |  |  | July 09, 2013 | 2013 | Medtronic Xomed, Inc. | The firm will recall and replace the identified instrument trays in the field, with trays that have improved labeling, including lids that contain the statement indicating 'PREVACUUM ONLY" | Device Design | Medtronic sent a Customer Commincation Recall Letter to all affected customers with the following information: Subject: URGENT PRODUCT Recall  Immediate Attention RequiredMEDTRONIC XOMED INSTRUMENT TRAYS- model numbers provided on Attachment One- all lots distributed prior to July 2013Dear Purchasing Manager, Risk Manager and Central Supply Manager/Supervisor:This letter is to inform you that Medtronic Surgical Technologies (ENT Division) is initiating a voluntary recall to remove certain instrument trays from the field and replace them with trays that have modified labeling. Medtronic Xomed Instrument Trays are intended for the storage and transportation of reusable surgical instruments. These trays are not intended and not designed to maintain sterility. We are aware, however, that some users may also choose to use the trays for holding instruments during steam sterilization. If the user chooses to sterilize instruments within these instrument trays, they must validate the sterilization process. These users should be aware of the following issue:Issue Description:During the validation of a new instrument, testing revealed that use of a GRAVITY steam cycle with one of the listed instrument trays did not produce the Sterility Assurance Level (SAL) of 10-6 recommended by industry standard (ANSI/AAMI ST77:2006). A subsequent investigation of all Medtronic Xomed instrument trays identified additional trays as having certain locations within thetray that would not meet the recommended sterilization level of SAL10-6 during GRAVITY steam sterilization cycles. The trays identified through the testing are the subject trays involved in this field action. Instruments that are sterilized in the worst-case locations of the identified trays have a higher potential to be inadequately sterilized. An instrument that is inadequately sterilized can become the source of cross-contamination or cross-infection when used in multiple surgical procedures. At this time |  | 3562 | 3,562 trays | Worldwide Distribution - USA (nationwide) Nationwide, Hawaii, Puerto Rico and Guam.Internationally to: Canada, Austria, Belgium, Cyprus, France, Germany, Greece, India, Iran, Ireland, Israel, Italy, Kuwait, Lebanon, Luxembourg, Morocco, Netherlands, Nigeria, Portugal, Qatar, Reunion, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Turkey, United Arab Emirates, United Kingdom, Australia, New Zealand, Singapore, Thailand, Vietnam, China, Hong Kong,Taiwan, Korea, Japan, Brazil, Chile, Colombia, Costa Rica, Ecuador, Guatemala, Mexico, Nicaragua, and Venezuela. | Terminated | September 25, 2014 | 443 | 3562 | Not_Computer
Z-1949-2013 | 64304 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CELLTRACKS AUTOPREP System (Software Version 2.4.1 &amp; Below. |  |  | December 28, 2012 | 2012 | Veridex, LLC | Veridex, LLC has received a report that a patient's demographics were incorrectly associated with another patient's identification (PID) number. The purpose of this notification is to inform you of the issue and provide instructions to help prevent an occurrence on your CELLTRACKS AUTOPREP System and CELLTRACKS ANALYZER II. | Software design | Veridex sent an Important Product Correction Notification letter dated December 28, 2012, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Confirmation of Receipt form. For questions contact Customer Technical Services at 1-800-421-3311. |  | 160 | US: 82 Units; Ex-US 78 Units. | Worldwide Distribution - USA Nationwide and the countries of Australia, Canada, China, India, Japan, Singapore, England, France, Germany, Italy and Spain. | Terminated | August 11, 2015 | 956 | 306 | Not_Computer
Z-1951-2013 | 65655 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ISE Electrolyte Reference reagent (for Synchron LX Systems and UniCel DxC 600/800 Systems), Part No. A28937, A28945, 467935, 467915. |  |  | September 22, 2010 | 2010 | Beckman Coulter Inc. | Beckman Coulter is recalling the Synchron Systems LX20/UniCel DxC ISE Reference Reagent because it has a potassium (K) response near the instrument calibration flag limits and may exceed the limits on some instruments. | Process control | Beckman Coulter Inc sent an Urgent Product Corrective Action (PCA) letter dated September 22, 2010 to all affected customers. Customers were notified of the the problems identified and the actions to be completed. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact Customer Support Center (Call Center) on the Beckman Coulter website at http://www.beckmancoulter.com/customersupport/support/ or call (800) 854-3633 in the US and in Canada. |  | 718 | 718 units total (132 units in US) | Worldwide Distribution - USA Nationwide and the countries of Belgium, France, Germany, Netherlands, Poland, Spain, Sweden, Switzerland, and United Kingdom | Terminated | August 10, 2013 | 1053 | 718 | Not_Computer
Z-1954-2013 | 65502 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | REF 00-4309-071-40 Trabecular Metal Reverse Shoulder System Instrumentation, Glenosphere Helmet, 40 mm Diameter Nonsterile, Qty -1 |  |  | June 28, 2013 | 2013 | Zimmer, Inc. | There is potential for one or both tabsof theTrabecular Metal Reverse Glenosphere Helmet to fracture during use and become disassociated from the device. There have been 47 reported complaints of tab fracture. | Device Design | Zimmer Inc sent Urgent Medical Device Recall notifications on June 28, 2013, to all affected customers. The initial letter identified the issue, possible adverse consequences and informed customers that a replacement would be available in the coming months. In addition, the letter included references to both the TM Reverse Surgical Technique and to the video available on Zimmer TV and Zimmer.com. Both resources explain in detail the proper technique for utilization of the Glenosphere Helmet. Proper utilization of the Glenosphere Helmet dramatically reduces the likelihood of fracture. Distributors will receive a copy of the surgeon letter for their reference along with the list of affected surgeons. The affected surgeons will be identified using invoice data and in cases where the surgeon is unknown, the distributor will be instructed to assist with the surgeon list. In addition, information regarding the return process for the damaged devices will be included. Questions and concerns are directed to Zimmer's customer call center (1-877-946-2761).Phase 2: As replacement devices with the new design became available for distribution, a 2nd letter was sent to the distributors. Estimated replacement device availability will begin in July 2013. The distributors will be instructed to order replacement devices as a means of completing the affected instrument cases in circulation within their territories. Instructions for the return of devices were included.For questions regarding this recall call 800-613-6131. |  | 3163 | 3,163 both sizes | Worldwide Distribution - USA (nationwide) and internationally to Australia, Austria, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Israel, Italy, Korea, Lebanon, Malta, Netherlands, Norway, Portugal, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, United Kingdom, Utd.Arab.Emir. | Terminated | March 02, 2015 | 612 | 3163 | Not_Computer
Z-1877-2013 | 65369 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX Evolution System, X-Ray System --- |  |  | June 03, 2013 | 2013 | Carestream Health Inc. | Due to potential for an x-ray exposure technique change. | Radiation Control for Health and Safety Act | Carestream will make the correction which includes the installation of a SW patch to the affected customer sites, and to any other affected customer sites utilizing the DRX Evolution with V5. 5 SW outside of the US. --- Once receiving FDA approval on the corrective action plan, the correction will be implemented to all affected devices at identified customer sites: 1. A notification letter was sent to all customers identifying the issue and highlighting the approved procedures of the Corrective Action Plan. The letters were sent on 8/16/13 via Certified Mail, Return Receipt Requested. 2. Carestream Health will contact the customer to schedule a service visit to install software updates to correct this issue. The installation time will be approximately one hour.3. As stated in the customer notification letter, this software update will be provided to the customer at no expense.4. All software updates are estimated to be completed within 1-3 months.5. Carestream Health will provide monthly progress reports on the effectiveness of the plan, including the number of units corrected.Further questions please call (800) 328-2910. |  | 49 | US: 49 units; Foreign: 183 units (revised) | Worldwide Distribution - USA (nationwide) including the countries of Austria, Australia, Belgium, Canada, China, Finland, France, Germany, Italy, Korea, Kuwait, Saudi Arabia, Spain, Sweden, Switzerland, Taiwan, United Kingdom. | Terminated | January 29, 2014 | 240 | 49 | Computer
Z-1955-2013 | 65822 | applicator, absorbent tipped, sterile | KXG | Absorbent tipped applicator. | General Hospital | General Hospital | 510(K) Exempt | Fisherbrand Sterile Swab, Calcium Alginate Fiber Tipped, Wood Applicator Swab 100 pouches/boxes, 10 boxes/case. |  |  | June 05, 2013 | 2013 | Fisher Scientific Co | The firm is recalling Fisherbrand Sterile Swabs due to 4 customer complaints reporting the presence of non-pathogenic organisms in two lots. The finding was confirmed in lot 61513 and lot 8710 is also being recalled as a precautionary measure. | Material/Component Contamination | Customers were notified by letter on June 18, 2013 and requested to discontinue use and return all affected product. The firm also asked that any additional affected customer base be contacted and notified of the recall if the lots had been further distributed. If you have any questions, contact Stericycle (reference "Event 4070") at 1-877-410-5558. |  | 1230 | 1230 | Nationwide and the country of Canada. | Terminated | October 20, 2014 | 502 | 1230 | Not_Computer
Z-1952-2013 | 65560 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Flexible Medullary Reamer. |  |  | June 11, 2013 | 2013 | Synthes USA HQ, Inc. | Due to the coiled design for this product, the product is difficult to clean and the potential for corrosion on the device exists. The potential for harm exists if there is an introduction of foreign material at the operative site originating from the presence of corrosion and/or the retained undefined material that could not be adequately cleaned. | Device Design | Sythes notified direct accounts by letter on 6/11/13 and requested them to check their inventory and remove affected lots from stock. Synthes requested that affected lots be returned and to call Synthes at 1-800-479-6329 for a Return Authorization Number. Completion and return of the Verification section at the end of the letter was requested, including negative responses. |  | 13050 | 13,050 | Nationwide Distribution. | Terminated | August 31, 2015 | 811 | 13050 | Not_Computer
Z-1957-2013 | 64994 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Invue Solid Driver, Tipped, AO. Part Number: 11-70083. |  |  | May 09, 2012 | 2012 | SpineFrontier, Inc. | The nipple of the tipped driver broke off when surgeons were attempting to insert the screw. | Device Design | On May 9, 2012 an email was sent to consignees requesting that the product be returned for servicing and that a replacement part would be sent. |  | 9 | 9 | Distributed in DC and the states of MA, and OK. | Terminated | January 10, 2014 | 611 | 9 | Not_Computer
Z-1958-2013 | 64970 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Case Lid and Base (Inspan, Invue, SLIFT) Caddy Lid (Inspan, Arena-C); Inspan Lid: 12-60039A, 11-60042B, 11-60043B, 11-60045B, 11-60046C, 11-60045A, 11-60046A, 11-60053A; Invue Lid: SI70095A; S-Lift Lid: SI50070A; S-Lift Base: SI50067A; Arena-C Lid: 11-72013D, 11-72014D; Generics Lid: 12-99005A. |  |  | April 22, 2013 | 2013 | SpineFrontier, Inc. | Non-toxic inks, used on the silk screen graphics of the case bases and lids, and caddy lids, were found to be either delaminating or running off from the substrate after repeated washing and sterilization cycles. | Labeling design | On 2/15/2013, SpineFrontier emailed distributors informing them that they needed specific parts returned to work towards completing the conversion of product graphics. |  | 258 | 258 units (30 in distribution at the time of the recall) | Worldwide Distribution-USA including the states of FL, OH, MA, and TX, and the countries of Mexico, and Dominican Republic. | Terminated | November 06, 2015 | 928 | 258 | Not_Computer
Z-1959-2013 | 65857 | No | FLL | Clinical electronic thermometer. | General Hospital | General Hospital | 510(k) | Disposable Temperature Probes, (sterile), Sensor diameter: 3mm/0.12 in; 25 units per box. Vital Signs Inc. a GE Healthcare Company, 20 Campus Road, Totowa, NJ 07512 USA, www.gehealthcare.com/vitalsigns. |  |  | July 18, 2013 | 2013 | Vital Signs Devices, a GE Healthcare Company | GE Healthcare has become aware of a potential safety issue associated with the use of GE General Purpose 9 French Temperature Probe, Disposable, 400 Series [M1024229]. GEs Disposable General Purpose 9FR Temperature Probe M1024229 is intended for oro-esophageal and rectal use. As it pertains to esophageal placement, it has been brought to our attention that some users are inserting the probe | Use error | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated July 23, 2013 via Fed Ex beginning on July 29, 2013 to customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. GE will update the Instructions for Use (IFU) and pouch label with new content related to insertion locations. GE has will also place the probe (M1024229) on quality hold until the new labeling changes are implemented.If you have any questions or concerns regarding this notification, please contact GEMSIT Customer Service at 1-800-588-7044 (Domestic) or Vital Signs Customer Service at 1-800-932-0760 (international). Hours of operation: 8:00 am EST to 6:00 pm EST. |  | 12239 | 12,239 units (5,002 units US, 7,237 units OUS). | Worldwide distribution: US (nationwide) and countries of: Australia, Belgium, Canada, Czech Republic, Finland, France, Germany, Italy, Lebanon, Netherlands, Norway, Poland, Saudi Arabi, Singapore, South Africa, Spain, Sweden and United Kingdom. | Terminated | January 02, 2014 | 168 | 12239 | Not_Computer
Z-1962-2013 | 64711 | No | GNX | Influenza virus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | Remel X/pect Flu A&B; Control Swabs, For In Vitro Diagnostic Use, Ref R246003, each package contains 20 Flu A Control Swabs and 20 Flu B Control Swabs, each swab is individually packaged in a foil pouch. The firm name on the label is Remel Europe, Ltd., Dartford, Kent, United Kingdom. |  |  | February 19, 2013 | 2013 | Remel Inc | Flu A+ Control Swabs incorrectly give Flu A- results when used for quality control testing with the Xpect Flu A&B; test kit. | Employee error | The firm, ThermoFisher Scientific, sent a "MEDICAL DEVICE RECALL" letter dated February 20, 2013 to its customers via regular mail. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory for affected product lots; discard the affected product; notify any personnel who needs to be notified of the potential for a false negative indication when using the swab, and complete and return the attached Medical Device Recall Return Response Acknowledgment &amp; Receipt Form via Fax to: Attn: Technical Service &amp; Regulatory Affairs at 1-877-428-1924.If you have any questions, please contact our Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 1920 | 96 kits of 20 swab sets | Nationwide distribution: AL, AR, CA, CO, CT, FL, GA, IA, IL, KS, MA, MD, MI, MO, NM, NV, NY, OH, PA, TX, WA, and WI. | Terminated | August 14, 2013 | 176 | 1920 | Not_Computer
Z-1960-2013 | 65005 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Biomet 3i Dental Implant, Parallel Walled 6mm (D) x 10mm (L) |  |  | May 10, 2013 | 2013 | Biomet 3i, LLC | A small number of outer implant boxes may be mislabeled. The inner tray, containing the implant, is label correctly. This condition, if present and not recognized could potentially result in the dental implant not being able to be placed into the osteotomy or seated to the desired depth if placement is attempted. | Labeling mix-ups | The firm issued "Urgent Medical Device Recall" notification letters to their customers dated May 10, 2013. The notification described the issue and provided recommendations actions regarding affected product. Customers with questions may contact 1 800-342-5454. |  | 51 | 51 | Distribution to the state of Massachusetts, and the foreign countries of Australia, Belgium, Spain, Denmark, France, Luxembourg, Switzerland, Germany, Bonaire, Sweden, and Italy. | Terminated | August 13, 2014 | 460 | 84 | Not_Computer
Z-1947-2013 | 64971 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | 1) Adult Manual Resuscitator, 2K8004, Oxygen Reservoir Bag, Adult Mask; |  |  | April 22, 2013 | 2013 | Carefusion 2200 Inc | CareFusion has received four (4) customer reports of a component (oxygen connector) within the resuscitation bag being occluded which disables the flow of supplemental oxygen to the resuscitation bag and thus may pose a serious patient safety risk. | Nonconforming Material/Component | CareFusion sent an URGENT RECALL NOTICE letter dated April 22, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are requested to complete and return an attached Distributor Acknowledgement Form within fifteen (15) days of receipt indicating their acknowledgement of the URGENT RECALL NOTICE and indicate if they have any affected product on hand. They are further requested to perform a 100% physical inventory to verify if any of the affected product codes/lot numbers are in stock within their distribution network/facility. In addition, customers are requested to forward the Customer Acknowledgement Form, the Customer Notification Letter, the Certificate of Destruction Form, the Pictures of Location of Lot Numbers on Case/Product, and the Frequently Asked Questions (FAQs) to all end users which they have distributed the affected product. For questions call 847-473-7097. |  | 14112 | 14,112 units | US Nationwide in the states of AZ, CA, CO, DE, FL, GA, IL, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NM, NY, OH, PA, TX, UT, WA | Terminated | June 16, 2015 | 785 | 14112 | Not_Computer
Z-1956-2013 | 65743 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | StealthStation S7 Framelink Software version 5.4. |  |  | July 15, 2013 | 2013 | Medtronic Navigation, Inc. | Medtronic Navigation, Inc. has initiated a field correction to their Framelink software version 5.4 which can be used on Medtronic Navigation StealthStation 57 and i7 systems or a Medtronic Planning Station due to the potential of discrepancy between the exam image and the displayed measurement or overlay under certain conditions of use. | Software design | A Field Safety Notice was sent on 7/15/13 to users informing them that the Framelink software version 5.4 has been revised to correct the software anomalies described in this letter. A new version will be installed on their system(s) by a Medtronic Navigation representative. Until the software upgrade has been installed on their system, do not use high resolution examinations (greater than 256 x 256). If users are not sure of the exam resolution being used, firm recommend contacting the hospital imaging department or local Medtronic Navigation representative for assistance. Medtronic representatives will visit all identified hospitals and upgrade all systems to version 5.4.1. This will account for all consignees. |  | 572 | 572 | Worldwide Distribution. | Terminated | January 29, 2014 | 198 | 572 | Computer
Z-1973-2013 | 65381 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | TerumoÂż Advanced Perfusion System 1 (APS1, System 1) |  |  | July 26, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (TCVS) has become aware that an Advanced PerfusionSystem 1 (APS1) Centrifugal Pump 'Service Pump' message may occur due to user inducedmotor movement. In this instance, the APS1 will identify the movement (as it isdesigned) and display a Service Pump message error. The user might interpret this as apump which requires service rather than a message displayed d | Device Design | Terumo Cardiovascular Systems (TCVS) notified their consignees of the Urgent Medical Device Recall letter dated 7/26/13 and was mailed 7/29/13. The notification provided reason for the device correction, affected product, how the device will be corrected, and the potential health hazard associated with a user's inaccurate interpretation of the service pump error message. Instructions for consignee responses to the recall notice communication may be faxed to 734-741-6149 or emailed to the address on the return form. If you have any questions or concerns, call Terumo Cardiovascular Systems (TCVS) Customer Service at 1-800-521-2818 M-F, 8am-6pm ET. Terumo Cardiovascular Systems (TCVS) will provide an addendum to the APS1 Manual to their consignees. The addendum will be placed at the beginning of Chapter 6, Centrifugal Control Unit. |  | 1657 | 1657 total | Worlwide Distribution - US Nationwide: AL,AR,AZ, CA, CO, CT, DC, DE, FL,GA, HI, IA, ID, IL, IN, KY, LA, MA , MD, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and the countries of: AUSTRALIA,UNITED ARAB EMIRATES (UAE),Hong Kong, onesia,Singapore,Taiwan,Thailand, Philippines,CHILE, Malaysia,USA,South Korea,BELGIUM,Japan,Mexico,CANADA | Open, Classified |  |  | 1657 | Not_Computer
Z-1972-2013 | 65744 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards EndoReturn Arterial Cannula, model ER21B and ER23B |  |  | July 16, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain lots of theEndoReturn Arterial Cannula that were built with an incorrect component. This component can create aperformance problem with the IntraClude IntraAortic Occlusion device. | Nonconforming Material/Component | Edwards sent an Urgent Field Safety Notice letter dated July 16, 2013 via Fed-Ex. The letter identified the affected product, problem and action to be taken. The letter included a Customer Acknowledgement Form to be completed and faxed to 800-422-9329 Attn: Recall Coordinator. If the customer does not send back the acknowledgement form, two additional contacts are made to notify the customer and request return of affected product. This is tracked by Edwards customer service and reported each week. Customer were requested to assess their inventory to assure that all affected product is identified and returned. Returned products will be destroyed. For questions call Edwards Customer Service at 800-424-3278. |  | 855 | 855 | US Nationwide including Canada and VA sites. | Open, Classified |  |  | 855 | Not_Computer
Z-1966-2013 | 65732 | Yes | DTM | Cardiopulmonary bypass arterial line blood filter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Affinity Pediatric Arterial Filter, model number S4014, sterilized using ethylene oxide. |  |  | June 14, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic is recalling 148 Affinity Pediatric Arterial Filters from 2 manufacturing lots because a limited quantity of filters may have a small breach in the filter media that may allow unwanted particulate or gaseous emboli to pass through the filter and the outlet of the device, which could result in serious injury to a pediatric patient including neurological damage, or potential patient death. | Employee error | Customer communication was initiated verbally on June 14, 2013, to the 4 affected customers to quarantine any un-used units. An Urgent Medical Device Recall letter, dated June 26, 2013. was mailed to customers on June 28, 2013. The Urgent Medical Device letter referenced that the letter is in follow-up to the previous phone call. It identified the affected product, stated the issue with the 2 affected lot numbers and health risks, and asked that the affected devices be quarantined for return to Medtronic.If you have any questions please contact your Medtronic Sales representative or Lifeline Technical Services at 1-877-526-7890. 1-763-526-7890.Update: Customer communication to the newly identified customer was initiated verbally on August 20,2013, with the customer letter delivery beginning AUG 27, 2013. A Medtronic representative is available to facilitate the return of affected product. |  | 153 | 153 devices. | Worldwide Distribution - US Distribution only to MS and IL, and the countries of Argentina, South Africa and Singapore. | Terminated | March 07, 2014 | 266 | 153 | Not_Computer
Z-1970-2013 | 65626 | No | JIS | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur Calibrator A. Cat. #04800735, SMN 10285904 (6 pack); Cat. # 04800840, SMN 10285905 (6 pack - REF). |  |  | June 28, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | Calibrator A lot CA77 is under-recovering to the assigned target for the ADVIA Centaur Total T4 Assay. | Nonconforming Material/Component | An Urgent Medical Device Correction (UDMC) was sent to all affected Siemens Healthcare Diagnostic customers in the United States on June 28, 2013 via FedEx. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on June 27, 2013. FedEx Return Receipts will be used to confirm that the customers received the UMDC. If confirmation of receipt by any customer cannot be verified through a return receipt, a Siemens Technical Solutions Center Representative will contact those customers by telephone to ensure that they received the Urgent Medical Device Correction and completed the required actions. These notices inform the customer that they should discontinue the use of Calibrator A, lots ending in 77, for use with the ADVIA Centaur Systems T4 assay. Customers are informed that they may continue to use Calibrator A lot CA77 with the other assays FT3, FT4, Total T3, T-Uptake. For questions or additional information contact your local Technical Support Provider or Distributor. |  | 265 | 265 Cal A 6 packs (US and OUS) | Worldwide Distribution-Nationwide including PR and the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NJ, NV, NY, OH, PA, TN, TX, VA, VT, and WV and the countries of Slovakia, Germany, Turkey, TW, HK, CH, TH, BR, CN, SG, and ZA. | Terminated | July 25, 2014 | 392 | 265 | Not_Computer
Z-1971-2013 | 65126 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | epoc BGEM Test Card (Catalogue number CT-1004-00-00). |  |  | March 13, 2013 | 2013 | Epocal | This recall is being initiated because the affected lots of this product may yield results that are not consistent with stated precision performance criteria at concentration values at the high-end of the Glucose test sensor (&gt;400 mg/dl). Specifically, after a period of storage of three (3) to four (4) months, the Glucose sensor may report low results at the high end of the measurement range. | Nonconforming Material/Component | On March 13th 2013, Epocal placed all CT-1004-00-00 test card on shipment hold. Affected product in Epocal control has been segregated and quarantined. Beginning on March 26 2013, Epocal Inc. plans to begin distribution of the Recall Notification Letter to the 4 direct consignees (customers/distributors) via certified mail. Included with the Recall Notification Letter is a faxback verification form. Notification will include all distribution channels to the end-user level. Direct consignees are instructed to discontinue use and discard all stock on hand, complete the provided faxback verification form to document disposal and contact their consignees to advise them of the notification and disposal of affected product. |  | 828750 | 828,750 Test Card | Worldwide Distribution-USA (nationwide) including the District of Columbia, US Virgin Islands and Puerto Rico, and the states of CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MD, MI, MN, MS, NC, ND, NE, NH, NJ, NM, NY, OH, PA, RI, SC, SD, TN, TX, UT, VA, VE, WA, WI, and WV and the coutries of India, Argentina, Australia, Azerbaijan, Belarus, Canada, China, Colombia, Egypt, Germany, Guatemala, Israel, Japan, Malaysia, Mexico, Nepal, Netherlands, Peru, Philippines, Russia Fed., Singapore, South Africa, South korea, Thailand, and Vietnam. | Open, Classified |  |  | 828750 | Not_Computer
Z-1965-2013 | 65829 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow International, Inc. Two-Lumen CVC Kit with Blue Flex Tip Catheter. The multiple lumen catheter permits venous access to the central circulation of patients. |  |  | July 31, 2013 | 2013 | Arrow International Inc | Arrow is recalling the product due to a labeling inconsistency. One portion of the lidstock states the kit contains a 3-Lumen catheter while the rest of the lidstock correctly identifies the catheter as 2-Lumen. | Labeling Change Control | An "Urgent Medical Device Recall Notification" letter dated July 31, 2013 was sent to direct accounts requesting immediate discontinuation of use and quarantine of all affected product on hand. Negative responses were also requested and the immediate return of product to Arrow was authorized. Questions were directed to Customer Service at 800 343-2935. |  | 156 | 156 | AL, PA, CT, NC, VA, KS, RI, AZ, NJ, and IL. | Terminated | April 28, 2015 | 636 | 156 | Not_Computer
Z-1963-2013 | 65838 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Anastaflo Intravascular Shunt, Model code IVS12512, IVS1512, IVS17512, IVS2012, IVS2512, IVS3012, IVS1530, IVS2030 and IVS25 |  |  | July 30, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain lots of the Anastaflo intravascular Shunt due to concerns of excessive adhesive on the shunt body that may interfere with suturing of by-pass grafts. | Process control | Edwards Lifesciences sent a Urgent Field Safety Notice letter dated July 30, 2013, via Fed-Ex to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letters are tracked through customer service and as part of the customer acknowledgement letter. Customers are asked to return products.If the customer does not send back the acknowledgement letter, three additional contacts are made to notify the customer and request return of affected product. This is tracked by Edwards customer service and reported each week. We reconcile the shipments made to customers against the returned product to determine impact to distribution. Once you have verified your inventory, please complete the attached acknowledgment form and fax it back to Edwards Customer Service at 800.422.9329 within three days of receipt of this Field Safety Notice. The return of this form allows us to confirm that you have reviewed this notice and have taken appropriate action. Please contact Customer Service at 800.424.3278 to obtain a Returned Goods Authorization number and replacement product.We sincerely regret the inconvenience caused by this action and appreciate your immediate attention to this matter. The Customer Service organization can answer questions about when replacement IVS Shunts will be available. If you have questions that have not been answered by this letter, please call Edwards Customer Service at 800.424.3278 from the hours of 6:00AM - 4:30PM PST or contact your Edwards sales representative concerning the recall. |  | 29618 | 29,618 units | Nationwide and the country of Canada | Terminated | October 14, 2014 | 441 | 29618 | Not_Computer
Z-1969-2013 | 65354 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | TerumoÂż Advanced Perfusion System 1: 220/240V AC, Advanced Perfusion System 1 |  |  | July 26, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | The operators manual does not fully describe the sequence of expected events that will occur when a specific operating condition exists. Specifically, if the user is adjusting the System 1 centrifugal pump speed via the Central Control Monitor at the same time an alarm condition occurs with configuration for a pump coast response, the labeling in the operators manual does not fully describe how th | Software design | Terumo CVS, sent an URGENT MEDICAL DEVICE RECALL notification letter dated July 26, 2013, to all affected customers explaining the correction to the operators manual, potential hazards and what to do in response to the issue.If you have any questions or concern please contact Terumo CVS Customer Service 1-800-521-2818 Recall Fax 1-734-741-6149 Customer Service Hours:Monday  Friday, 8 AM - 6 PM ET |  | 1657 | 1657 | Worldwide Distribution - USA (nationwide) including the countries of Australia, Dubai, Hong Kong, Indonesia, Singapore, China, Mexico, British Columbia, Taiwan, Thailand, Chile, Malaysia, Japan, Belgium, United Arab Emirates (UAE), South Korea, Manila and Canada. | Open, Classified |  |  | 1657 | Not_Computer
Z-1975-2013 | 65719 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza SW version VA20 with Mammography license. |  |  | March 04, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens issued a Customer Safety Notice to inform customers of a potential problem when using syngo.plaza. In some instances, it can happen that the laterality information in mammography images is not displayed and these images may not be properly aligned to mammography segments. | Software design | The firm, Siemens, notified their customers on March 4, 2013 and sent a "Customer Safety Advisory Notice" letter dated April 03, 2013. The letter described the product, problem and actions to be taken. The customers were instructed to always check whether the laterality information is displayed in the image text of mammography images; if not, the DICOM header is accessible and can be checked for the tag "Laterality" (0020, 0060). If a system is not able to provide the DICOM tag "Image Laterality" (0020, 0062), configuring the image text by adding the tag was proposed. If you have sold the device/equipment and it is no longer in your possession, forward this safety notice to the new owner of the device/equipment and inform us about the new owner of the device/equipment. This issue will be resolved in the syngo.plaza software version VA20D_HF04 which is planned to be released in 02/2013. Customers were also instructed to contact the local Siemens service office to plan the update.If you have any questions, contact the SYNGO CEO or Vice President Quality at +49 (9131) 84-4874 or email to: michael.mlb.braun@siemens.com. |  | 6 | 6 | Nationwide distribution: FL, MA, KS, NY, OH, PA, TX and WI. | Terminated | March 09, 2015 | 735 | 6 | Computer
Z-1977-2013 | 65380 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | 100/120V AC, Advanced Perfusion System 1 |  |  | July 26, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | TCVS is taking action to update the Operators Manual with the specific instructions to respond to this condition and will issue a labeling addendum for current users. In the event of a false high pressure alarm condition occurs that displays a "999" pressure reading and results in stopping a pump, this pressure display can be cleared and the pump re-started by disconnecting the pressure transduce | Software design | Terumo sent an Urgent Medical Device Recall notification letter dated July 26, 2013 to all affected customers. The letter identified the affected product, reason for correction, potential hazards of the recall issue, the correction to be implemented, and customer instructions for replying to the recall notification. Terumo CVS is providing an addendum to be placed in the Operator's Manual at the beginning of Chapter 7, Modules. Questions or concerns are directed to Terumo CVS Customer Service at 1-800-521-2818, Recall Fax 1-877-275-6860. Customer Service hours are Monday to Friday 8am-6pm ET. |  | 1657 | 1657 | Worldwide Distribution - US Nationwide in the states of :AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MSMT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV and countries of AUSTRALIA, UNITED ARAB EMIRATES (UAE),Hong Kong,Indonesia, Singapore, Taiwan, Thailand,Philippines, CHILE,Malaysia,South Korea, BELGIUM,Japan,Mexico,CANADA | Open, Classified |  |  | 1657 | Not_Computer
Z-1979-2013 | 65798 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific, ENERGEN DR ICD, Model E143. The device is an Implantable Cardioverter Defibrillator. |  |  | July 15, 2013 | 2013 | Boston Scientific CRM Corp | Boston Scientific CRM manufacturing quality system recently discovered test artifacts (related to shock charge time) in a small number of implantable defibrillators (ENERGEN DR ICD model E143) which require further investigation. While these devices successfully passed all manufacturing tests and met all requirements, they showed a slightly longer charge time than other devices during a manufactu | Under Investigation by firm | Consignee was sent on 7/15/13 a Boston Scientific "Medical Device Retrieval" letter dated July 12. 2013. The letter described the product problem and informed the consignee that they will replace the product removed from their inventory. For further questions, contact Boston Scientific Technical Services at 1 800 227-3422. |  | 1 | 1 (4 devices were implanted and are not part of this action) | US distribution in the state of Tennessee. | Terminated | November 12, 2013 | 120 | 1 | Not_Computer
Z-1980-2013 | 65552 | No | INO | Electric positioning chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | CHAMPION Medical Recliner, AC Powered with heat and massage, one chair per carton, or one chair wrapped in a blanket if delivered in person. |  |  | June 18, 2013 | 2013 | Invacare Corporation | There is a possibility of elevated temperatures in the seat pad. In addition, the hospital cord set does not agree with the labeling (information in manual). | Process control | Champion Manufacturing plans to initiate this recall within approximately 10 business days after the submission of this report through a phone campaign and notification letters. |  | 285 | 285 units | Nationwide distribution: USA including states of: CA, FL, IN, MN, NJ, NY, OH, RI, SC, SD, TX, and VA. | Terminated | January 09, 2014 | 205 | 285 | Not_Computer
Z-1982-2013 | 65855 | No | ODE | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EndoGastric Solutions EsophyX2 Device with Serosa Fastener and Accessories. Models R2001 and R2002. |  |  | June 06, 2013 | 2013 | Endogastric Solutions Inc | Endogastric Solutions, Inc. has received a limited number of reports relating to the loss of tissue mold control when operating the R2001 or R2002 EsophyX2 device. In one case, surgical intervention was necessary to remove the device. | Nonconforming Material/Component | Endogastric sent the Safety Alert: EndoGastric Solutions (EGS), EsophyX2 Device letter, dated June 5, 2013, to their consignees.Endogastric sent the second letter URGENT: MEDICAL DEVICE RECALL EsophyX2, dated July 26, 2013. This letter advised customers that the firm is voluntarily recalling EsophyX2 Devices with SerosaFuse Implantable Fastener and Accessories (Models R2001 and R2002) manufactured before February 2012.Customers who have UN-EXPIRED devices in their inventory are advised to discontinue use and fill out the Medical Device Recall Return Response form and return it to Endogastric. The Customer Service will contactcustomers with instructions on how to return the product to the company.Customers who may have EXPIRED devices are advised to with their central supply departments to ensure all identified products have been removed from inventory and destroyed. They should fill out the Medical Device Recall Return Response form with the lot numbers and quantity destroyed and return thecompleted form to the firm.Customers can call the Customer Service and Support at 425-307-9269, Monday through Friday, 8:00AM to 5:00 PM, Pacific Time for questions. Adverse reactions or quality problems experienced with the use of this product may be reported to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax. |  | 5233 | 5192 units in the US and 41 units outside the US | Distributed nationwide and Italy. | Terminated | April 09, 2014 | 307 | 5233 | Not_Computer
Z-1981-2013 | 65833 | No | JDG | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Moore Hip Prosthesis Long Fenestrated Stem, several head Diameter, 203 mm Stem length. |  |  | July 16, 2013 | 2013 | Zimmer, Inc. | Zimmer recently conducted a review of historic packaging validations completed. Based on this review, it was determined that a subset of product packaged at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. This testing determined that a single packaging tray configuration, which | Packaging process control | Zimmer is removing the devices from the field including distributor and hospital inventories. A notification was issued to known users of the product on 07/16/2013. |  | 1710 | 1,710 units are in distribution | Worldwide Distribution-USA (nationwide) including the states of CT, FL, GA, IL, IN, KS, KY, MA, MI, MN, MT, NC, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VI, and WI, and the countries of Singapore, Canada, Mexico, Arab Emirates, Saudi Arabia, and Hong Kong. | Terminated | April 30, 2015 | 653 | 1710 | Not_Computer
Z-1983-2013 | 65785 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GYRUS ACMI, REF 3740, Everest BiCOAG Bipolar Dissecting Forceps 5mm/45 cm, Sterile EO. |  |  | July 24, 2013 | 2013 | Gyrus Medical, Inc | On April 29, 2013 Olympus Surgical Technologies America R&D; received results from an accelerated aging study conducted as part of a new market submission. The study found that one of 60 Dissector units that underwent 3 year accelerated aging failed the dye penetration test indicating that the seal on sterile packaging was compromised. The dissector products are labeled with a 3 year shelf life. | Packaging process control | Olympus sent a "Medical Device Recall" letter dated July 24, 2013, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer.. Customers were advised to immediately cease any further use of the product and remove it from inventory and quarantine it until is shipped back to the company. To return the product, call 1-888-524-7266 Olympus customer service representative. Customers were requested to complete and return by fax the enclosed questionnaire to 484-896-7128. If the product was further distributed then they need to identify their customers and notify them at once of the product recall. |  | 1140 | 1140 | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, CA, FL, GA, IL, IN, KS,KY, MD, MA, MI, MN, MS, MO, NY, NC, OH, PA, SD, TN, TX, VA and WI., and the countries of BRAZIL, CHINA, AUSTRALIA, CHINA, GERMANY and INDIA. | Terminated | June 20, 2014 | 331 | 1140 | Not_Computer
Z-1984-2013 | 65908 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Product Name: nanoPOINT Injector System, item number 37-0000-08. |  |  | August 05, 2013 | 2013 | Medicel Ag | The firm initiated this recall of one lot (121402) of nanoPOINT Injectors after they noticed that Tyvek lids might have been damaged and, therefore, the stability of the sterile barrier might be negatively affected. | Use error | Medicel AG sent recall notification to its one direct consignee via email on 08/05/2013. The recall notification letter was titled "URGENT-Medical Device Correction Notice for nanoPOINT Injectors LOT 121402". The recall notification letter informed the customer of the product device description and lot number recalled, reason for recall, instructions to "not use this affected lot...and return them", contact information and an area on the bottom of the letter to acknowledge receipt of the correction notice.The firm's direct consignee who further distributed this product to the end users will send out notifications. |  | 748 | 68 boxes or 680 units | Nationwide Distribution. | Terminated | March 11, 2014 | 218 | 748 | Not_Computer
Z-1991-2013 | 65879 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ConvaTec Flexi-Seal SIGNAL +, Rx Only |  |  | June 25, 2013 | 2013 | Convatec Inc. | ConvaTec notified their customers on 6/25/2013 regarding the manufacturing irregularity that occurred related to the packaging of certain lots of Flexi-Seal Signal FMS. In certain lots, the lot number convention was duplicated on the packaging. | Under Investigation by firm | ConvaTec issued a notification letters to their customers dated June 25, 2013 via UPS. The letter informed the customer of the issue with the related product and the action to be taken. If the customer has any question regarding the notification they can contact customer service at 1-800-422-8811. |  | 22896 | 22,896 units | US Nationwide Distribution | Terminated | January 08, 2014 | 197 | 22896 | Not_Computer
Z-1985-2013 | 65869 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Suction Wand Model S099B, packaged in Sorin Group Heart/Lung Perfusion Packs. |  |  | July 08, 2013 | 2013 | Sorin Group USA, Inc. | Sorin Group is recalling certain lots of Edwards Lifesciences Suction Wand Model S099B due to presence of plastic particles inside the pouch. | Material/Component Contamination | On 7/8/13, customers and distributors were sent a "Field Safety Notice Product Recall" letter and requested to return unused Edwards Suction Wands found in the Perfusion Packs to the Sorin Group. Customers are to contact Sorin Group Customer Service at 1-800-650-6502 to receive instructions for returning the affected Edwards Suction Wand to Sorin. Customers were also advised to not discard affected product. |  | 18927 | 18,927 packs | Nationwide Distribution | Terminated | January 28, 2014 | 204 | 18927 | Not_Computer
Z-1986-2013 | 65644 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Reservoirs, Model No: MMT-326A (1.8mL) and MMT-332A (3.0mL) |  |  | July 03, 2013 | 2013 | Medtronic MiniMed | Medtronic is recalling certain lots of Medtronic MiniMed Paradigm Reservoirs MMT-326A (1.8mL) and MMT-332A (3.0mL) used with Medtronic Paradigm insulin pumps because they may have increased risk for leaking. A leak in the reservoir may result in delivery of less insulin than intended. In addition, if there is a leaky reservoir and an insulin blockage occurs in the infusion set, the pump may not | Under Investigation by firm | An Urgent Medical Device Recall notice was sent to customers on July 3, 2013, who purchased the Medtronic Paradigm Reservoirs MMT-326A and MMT-332A. The recall letter informed the customers of the problems and the actions to be taken. Medtronic instructed customers that they will replace all reservoirs from the affected lots at no cost. Customers are instructed to refer to www.medtronicdiabetes.com/support/product-updates for a list of frequently asked questions. Customers are instructed to report adverse reactions or quality problems experienced with the use of Medtronic Diabetes insulin pumps, infusion sets or reservoirs to FDA's MedWatch Adverse Event Reporting Program. Customers are instructed to call for replacements at (866) 222-2584 Monday through Friday (8am-6pm) central time. Customers with an urgent issue may use the 24 hour helpline. |  | 669672 | 669,672 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Bermuda, Brazil, Canada, Chile, Columbia, Ecuador, El Salvador, Germany, Japan, Mexico, Netherland, Paraguay, Puerto Rico, Singapore, Uruguay and Venezuela. | Terminated | January 16, 2014 | 197 | 669672 | Not_Computer
Z-1987-2013 | 65678 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | PleuraGuide Disposable Chest Tube Kit; Product codes: 17100, 17150, and 17200. |  |  | July 16, 2013 | 2013 | Atrium Medical Corporation | Individual component's IFUs are not included with the kit. | Labeling Change Control | Notice of correction was sent on July 18, 2013 via UPS and e-mail to the field representatives that represent these accounts. The notification identifies the problem, product, and risk factors. It provides instructions how to handle the product that may be affected. Additional labeling will be added to the kit to include instructions for use for the additional components. Atrium Medical will confirm through tracking of the UPS letters that each consignee has received the letter of correction. This shall be 100%. Where a UPS delivery cannot be confirmed, Atrium Medical will make good faith effort to contact the consignee to ensure that the letter was received. |  | 6516 | 6,516 units | Nationwide Distribution. | Terminated | July 15, 2014 | 364 | 6516 | Not_Computer
Z-1967-2013 | 65797 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Discovery NM/CT 570c CT Standalone Scan |  |  | April 10, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of an incident where the wrong part of the patient body wasscanned during a CT standalone exam with the Discovery NM/CT 570c | Radiation Control for Health and Safety Act | GE Healthcare will remedy the issue or bring the product into compliance with each applicable Federal standard. GE Healthcare has provided field modification instructions (FMI) to update software for the Discovery NM/CT570c hybrid system. The software update will correct the incorrect positioning of the table not prescribed by the operator. The FMI includes:A software upgrade with foreign language support to be installed in the affected models listed above,Labels (FMI label and packing label),Software upgrade workflow procedure with functional verification for all modes of operation, andCertification Report, Postponement of Medical Device Correction, Rejection of Medical Device Correction, or Refusal and/or Location Unknown Documentation.For further questions please call (262) 513-4122. |  | 10 | 10 Systems | US Distribution in the state of CT. | Terminated | January 21, 2014 | 286 | 10 | Computer
Z-1989-2013 | 65862 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Precision 500D |  |  | May 02, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the overhead Video Monitor Suspension involving Precision 500D Systems. There is a potential for the locking nut to loosen out of the Video Monitor Suspension yoke which could result in impact. This will not lead to an immediate fall of the suspended monitor since it will be held in place by the remaining bushing and | Device Design | GE Healthcare "Urgent Medical Device Correction" letter dated May 1, 2013 to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The letter informed customers that GE Healthcare will correct all affected systems at no cost. A GE Healthcare service representative will contact you to arrange for this correction. For questions or concerns regarding this notification call: US 800-437-1171, Japan: 0120-055-919. |  | 445 | 445 (421 US, 24 OUS) | Worldwide Distribution - USA Nationwide in the states of AL, AK, AR, CA, CO, CT, FL, GA, HI, ID, IL IN, IA, KS, KY, LA, MD, MA, NMI, MN, MS, MO, NE, NV,NH,NJ, NM, NY, NC,ND, OH, OK OR, PA,SC, SD, TN, TX, UT, VA, WA, WV, WI, WY and the countries of CANADA, FRANCE GERMANY, GREAT BRITAIN ITALY, PANAMA, SAUDI ARABIA. | Terminated | November 26, 2013 | 208 | 445 | Not_Computer
Z-1988-2013 | 65745 | No | FPX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(k) | Distributed in USA: Product Number: 40-2894: Reinforced Wound Closure, Latex Free Sterile EO Rx, 1/4" X 4 ", 6 mm X100 mm, (2 Cards/pouch, 1,000 Pouches/box); |  |  | July 18, 2013 | 2013 | Aspen Surgical Products, Inc. | Aspen Surgical is recalling Wound Closures because firm could not confirm sterility, after FDA sampling indicated the product may not be sterile. Using the recalled product could lead to infection which may require further treatment including the use of antibiotics. There have been no reports of injuries or illnesses as of 7/18/2013. | Process design | Aspen sent an Urgent Medical Device Recall letter dated July 18, 2013 via certified mail to all customers with directions to forward the notification and provide Aspen with sub-account information. Aspen has given the customers the option to either 1.) Return the product 2.) Destroy the product with a Certificate of Destruction to be provided to Aspen. Any returned product to Aspen will be put on MRB and dispositioned by Quality. Questions or concens should be addresses to Kelli Jona Directo of QA / RA at 888-364-7004 ext 177 jonask@aspensurgical.com or Michelle Pulaski, Quality System and Compliance Manager 888-364-7004 ext 146 pulaskim@aspensurgical.com 8-5 EST M-F. |  | 3286 | 3286 boxes ( 6,353,000) | Worldwide Distribution - US Nationwide in the state of Florida and in the countries of , Netherlands, and Reditch Worcestershire England | Terminated | April 17, 2014 | 273 | 3286 | Not_Computer
Z-1964-2013 | 65839 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare MobileDiagnostwDR system. Mobile Digital radiography X-Ray System |  |  | June 11, 2013 | 2013 | Philips Healthcare Inc. | The Instructions for Use (IFU) for the MobileDiagnost wDR fails to comply with a Federal standard. The IFU does not provide tolerance levels for several generator values as required by 1020.30(h)(3). | Radiation Control for Health and Safety Act | Phillips Healthcare sent an Important Notification letter dated July 10, 2013 to all affected customer. The notification informed customers that Philips will implement an update to the Instruction for Use (IFU). Philips will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services. For questions contact your local Philips representative or local Philips Healthcare office. |  | 240 | 240 systems US -- 1 system within Puerto Rico | US Nationwide Distribution including Puerto Rico | Terminated | January 12, 2017 | 1311 | 240 | Not_Computer
Z-1994-2013 | 65866 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | Getinge Assure Accufast Biological Indicators, Catalog Number 61301606636 (50 Pack) and 61301606637 (100 Pack), STEAM --- Manufactured for: Getinge USA, Inc. 1777 E. Henrietta Rd., Rochester, NY 14623. |  |  | August 02, 2013 | 2013 | Getinge USA Inc | AccuFast Biological Indicator Lot SR-430 may have been exposed to conditions affecting its viability. This could result in a false negative (or pass) when exposed to sub-lethal sterilization parameters. | Under Investigation by firm | GETINGE sent an DEVICE CORRECTION NOTICE letter dated July 31, 2013, and response forms were sent to customers on August 2, 2013, via Certified Mail Return Receipt Requested. The letters instruct customers to immediately discontinue use, quarantine, and destroy the affected lot. In addition, they instruct customers to: Check the appropriate box and record the quantity of affected unused product on the enclosed Correction Response Form OR If if they do not have any lots of affected unused inventory check the box indicating such. Destroy the unused product in accordance with your institution's policy and instructions for use. Record the method of destruction on the Correction Response Form. Be sure to include the Facility Name, Address, Contact Name and Phone Number. Sign, date and e-mail the form back to SR430responseform@getinge.com within 10 days or sooner of receiving this notice. Credit will be issued upon receipt of the Correction Response Form. Replacement product of unaffected lots is currently available. The warehouse shipping the affected lot is no longer being used (due to business decision, not as result of this recall). The current warehouse is maintaining the product in a controlled temperature segregated room away from any potential sterilants.If you have any questions regarding this notification please call 1-800-475-1400 ext. 5026 |  | 78 | Catalog # 61301606636: 78 Packs (3900 BIs); Catalog # 61301606637: 43 Packs (4300 BIs) | Worldwide Distribution - USA and the country of Canada. | Open, Classified |  |  | 78 | Not_Computer
Z-1992-2013 | 65888 | No | KZF | AC-powered medical examination light. | General Hospital | General Hospital | 510(K) Exempt | Stryker Visum 300 Ceiling-mounted Exam Light. |  |  | July 25, 2013 | 2013 | Stryker Communications | Potential incomplete final weld on the drop tube flange. | Process control | The firm decided to recall and notified consignees with an "Urgent Medical Device Correction/Removal" notification letter dated 07/29/13. The product issue was described and recommended actions were provided. Customers may contact Stryker Technical Support at 866-841-5663. |  | 1375 | 1,375 units | Worldwide Distribution: US natiowide and to the country of Canada. | Terminated | April 03, 2015 | 617 | 1375 | Not_Computer
Z-1995-2013 | 65834 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Blue Belt Technologies, Inc. Monitor Drape REF: PFSD-V0001, Rx only, Sterile EO. |  |  | July 19, 2013 | 2013 | Blue Belt Technologies MN | Blue Belt Technologies is recalling sterile Monitor Drapes used with their Navio PFS system during orthopedic knee surgery. The drapes are used to cover the Navio PFS system computer monitor and arm assembly and provide a sterile barrier in the surgical suite. The defect is evident when the computer monitor is being draped. The front section of the drape separates from the rest of the drape at | Package design/selection | Blue Belt Technologies contacted consignees by telephone using a script on July 19, 2013. The consignees were informed of the problem and the product involved in the recall. Consignees were advised to immediate stop using the product and to locate and quarantine all unused product. Additionally, an "Urgent Medical Device Recall" letter dated July 24, 2013 was issued on July 24, 2013 and August 12, 2013. The letter includes the description of the Issue, description of safety concerns, affected product, and action to be taken by user. For questions contact Blue Belt Technologies, Lindsey Bring (+1 763.452.4922) or Robin Gonzalez (+1 763.452.4910) and get a Return Authorization number. Lindsey or Robin will arrange a pickup of the product. BBT will send replacement product at no charge. |  | 2400 | 60 cartons (20 pieces/carton) 40 in EU, 20 in US. | Worldwide Distribution and US Nationwide in the states of: CA and PA.United Kingdom. | Terminated | November 14, 2013 | 118 | 2400 | Not_Computer
Z-1976-2013 | 65777 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Stryker Oasys Midline Occiput Plate-1) Small 2) Medium 3) Large 4) Large Long |  |  | May 30, 2013 | 2013 | Stryker Spine | Stryker has received reports from customers indicating post-operativefracture of the pin that connects the tulip head to the plate body. | Other | Stryker Spine sent a 'Urgent Medical Device Removal Letter'/'Customer Response Form' to Branch Managers, Agent Principals, Quality Contacts, Hospital OR Risk Managers, OR supervisor on 5/30/2013 via Fed Ex. Inventory should be examined and returned.The attached form should be completed and returned as well.Stryker Spine sent 'Urgent Medical Device Removal Letters' to Spinal Implant Surgeons on 6/20/2013 via Fed Ex. |  | 1536 | 1536 units (US) 880 (Foreign) | Nationwide Distribution. | Terminated | December 30, 2014 | 579 | 1536 | Not_Computer
Z-1998-2013 | 65413 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00597001101 Femur CR PRECOAT FEM COMP SIZE AML, Rx, Sterile; |  |  | June 07, 2013 | 2013 | Zimmer, Inc. | Surgeons are implanting components of the NexGen Cruciate Retaining (CR) Complete Knee System in combinations that are not approved (micro with standard). Includes CR Micro Femurs, Articular Surface, Patellae, instruments and provisionals combined with standard components. A March 2012 recall reiterated the proper combination labeling, but the mis match problem continues. Zimmer is removing the d | Labeling design | Beginning June 7, 2013 all distributors will be sent Urgent Device Recall notices via electronic mail, with distributors, hospital risk managers with affected inventory also receiving the notice via courier. The Notice outlined the products recalled; the reason for the recall, possible adverse events related to the recall and directed consignees to locate, quarantine and count all affected product recording the information on the Inventory Return Form acknowledging that recall directions were followed. Further questions should be directed to 1-877-946-2761. Affected product is to be returned to Zimmer Product Service Department , 1777 West Center St. Warsaw, IN 46580 or internation return at Zimmer International Logistics GmbH, Max - Immelmann Allee 12 79427 Eschbach Germany."Communications outside of the United States will occur approximately two weeks after the United States communications. |  | 192355 | 192,355 all devices | Worldwide Distribution-USA (nationwide) and the country of Japan. | Terminated | December 15, 2014 | 556 | 192355 | Not_Computer
Z-1997-2013 | 65832 | No | KHE | Occult blood test. | Hematology | Hematology | 510(k) | Hemoccult ICT (Immunochemical Fecal Occult Blood Test) is a rapid, visually read, qualitative immunochemical chromatographic method for detection of human hemoglobin in fecal samples. |  |  | July 17, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter has received a customer complaint that suggests trending ofpositive results has increased from approximately 4% to 1 0%. As a result of ourinvestigation, Beckman Coulter has confirmed an increase in sensitivity at theborderline positive level." This testing also demonstrated that the issue does not affect results reported asnegative. | Under Investigation by firm | Beckman Coulter sent an Urgent Product Correction letter dated July 9, 2013, to all affected customers. The letter was accompanied with a cover letter and response form, dated July 9, 2013. Customers were informed of the product under recall with product description, actions to take, contact information, the issue, impact, and resolution. Beckman Coulter conducted a risk assessment for this issue. Per the risk documentation, there is only a remote probability that use of the product will cause any adverse health consequences,and the product is not likely to lead to serious adverse health consequences.If you have any questions please call (714) 961-3777. |  | 68116 | 68,116 units | Worldwide Distribution - USA and the country of Canada. | Terminated | October 24, 2014 | 464 | 68116 | Not_Computer
Z-1996-2013 | 65738 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS(R) II CRUCIATE RETAINING ARTICULAR INSERT, TRIAL, SIZE 3-4, A/P 48 MM, M/L 68 MM, QTY 1, NON-STERILE, REF 71430490, |  |  | July 08, 2013 | 2013 | Smith &amp; Nephew Inc | The products were produced using the wrong color material, resulting in the devicesbeing red (used for deep dished inserts) instead of blue (used for cruciate retaining inserts). | Labeling mix-ups | Smith &amp; Nephew sent an Urgent Recall Notification letter dated July 8, 2013 to all affected customers. Surgeons were notified by letter beginning on July 16, 2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory, locate unused devices, quarantine immediately, and return affected product to the address provided in the notification. For questions email to FieldActions@smith-nephew.com. |  | 9 | 9 units | International only: Germany, France, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 9 | Not_Computer
Z-2045-2013 | 65810 | No | OUT |  |  | Neurology | PMA | Pipeline Embolization Device (PED) |  |  | July 19, 2013 | 2013 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | Covidien Neurovascular has initiated the recall of several lots of Pipeline Embolization Device (PED) because the Instructions For Use (IFU) do not include the following contraindication: Patients in whom a pre-existing stent is in place in the parent artery at the target aneurysm location. Additionally, the designation of aspirin and clopidogrel was missing from the following contraindication " | Employee error | Covidien sent an Customer Notification letter dated July 19, 2013, to all affected customers who received affected lots of Pipeline Embolization Device (PED). The customer notification letter informed the customers of the problems identified and the actions to be taken. Customer are instructed to reply and fax response form to Covidien Customer Service at (763) 398-7001. Customers are instructed that Customer Service will issue an RGA number for the return of the affected product back to Covidien. Customers are also instructed that upon receipt of a Returned Goods Authorization (RGA) number, to write theRGA number on the FEDEX slip and send the product back to:CovidienAttention: Product Recall14005 13th AvePlymouth, MN 55441Prepaid and preaddressed FEDEX slips are enclosed for their convenience. Once Covidien receives the returned units, product replacement units will be shipped directly to customers. Local Covidien Neurovascular representative is also available to assist them in expediting the exchange of product.Customers with any questions, are instructed to call (949) 839-3700. |  | 62 | 62 units | US Distribution including the states of AZ, KY, MO, NC, GA, WI, OR, FL, NY, CA, IL, TN, NV, TX, CO, MA, IA, UT and VA. | Terminated | October 17, 2013 | 90 | 62 | Not_Computer
Z-2048-2013 | 65841 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | SARKEN STAT! Multi-Drug Test Panel |  |  | July 29, 2013 | 2013 | Sarken, Inc. | These products are being recalled because they were distributed without adequate labeling and instructions for use. Inadequate labeling could lead to misuse of the Test Kits and incorrect test results. | Other | Firm notified customers via telephone, followed up with MEDICAL DEVICE RECALL letter dated July 29, 2013 sent 08/02/13, via USPS regular mail to all affected customers. The recall notification letter is accompanied by a return response, acknowledgement and receipt form. The letter notify customers of the voluntary recall, product description (with prodcut images), reason for recall, potential impact, instructions/actions to follow, and contact information. |  |  | various quantities | Worldwide Distribution - US Nationwide in the country Virgin Islands | Terminated | July 16, 2014 | 352 |  | Not_Computer
Z-2047-2013 | 65609 | [6 empty] | PSA Assays for use in the IMMULITE/IMMULITE 1000 Systems Analyzers, Catalog Numbers LKPS1, LKPS5 and LKPTS1(D); and in the IMMULITE 2000/2000XPi Systems Analyzers, Catalog Numbers L2KPS2, L2KPS6, L2KPTS2(D) and L2KPTS6(D). IVD Immunoassay kits in the following sizes: 100, 200, 500 and 600 test kits; For in vitro diagnostic use with the IMMULITE/IMMULITE 1000 and the IMMULITE 2000/ 2000 XPi Systems Analyzers. PMA # P930027/ Device Listing # D002374. |  |  | June 26, 2013 | 2013 | Siemens Healthcare Diagnostics | Customers complained of a positive bias in patient results on the IMMULITE/ IMMULITE 1000 and the IMMULITE 2000/ 2000 XPi PSA assay when compared against the IMMULITE /IMMULITE 1000 3rd Generation PSA assay. Siemens has confirmed an overall average positive bias of approximately 20% to 23% across the assay range relative to the WHO 96/670 with the IMMMULITE Systems PSA Assay. This positive bias i | Software design | Seimens sent an Urgent Medical Device Recall Letters to US consignees via FedEx and an Urgent Field Safety Notices were sent to the foreign consignees starting on June 26, 2013. In addition, Support Bulletins were distributed to Siemens Healthcare Diagnostics service personnel describing the issue and instructing them how to deal with customer questions. Customers were informed to discontinue use, discard the kits remaining in inventory and there are no replacement PSA kits available at this time. Siemens is instructing customers to contact their local Siemens representative for assistance with determining appropriate PSA testing solutions for your laboratory. |  | 24361 | 24,361 kits | Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of: Albania, Afghanistan, Algeria, Argentina, Austria, Bangladesh, Belgium, Bolivia, Bosnia- Herzegovinia, Brazil, Bulgaria, Camaroon, Canada, Canary Islands, Chile, China, Colombia, Coast Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Germany, Great Britain, Greece, Guatemala, Honduras, Hungary, India, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Latvia, Lebanon, Lithuania, Macedonia, Malaysia, Malta, Mexico, Morocco, Nicaragua, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Slovakia, Slovenia, Spain, South Africa, Rep. of Korea (S), Sweden, Switzerland, Syria, Taiwan, Thailand, The Netherlands, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, Uruguay, White Russia, Venezuela, and Vietnam. | Open, Classified |  |  | 24361 | Not_Computer
Z-2046-2013 | 65865 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | NAVIS-EX , Image Filing Software for NIDEK AFC Fundus camera, Ver 1.1.0 to 1.3.5. NAVIS-EX is indicated for use as file management software for ophthalmic images. |  |  | August 01, 2013 | 2013 | Nidek Inc | Under certain circumstances, the Left eye image may be saved as the Right eye image. | Software design | "Medical Device Correction Notification" letters dated August 5, 2013 were sent to consignees through their US contracted distributor - Marco Opthalmic. The letter described the product issue and provided information on the correction. Customers can contact 800-223-9044 for questions. |  | 86 | 86 | Nationwide US distribution only. | Terminated | February 25, 2014 | 208 | 86 | Not_Computer
Z-2084-2013 | 48997 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Continuous Ventilator/Continuous Ventilator with Respiratory Mechanics, Model No. V1000 |  |  | October 29, 2007 | 2007 | Respironics California Inc | The power to the Esprit Display's Backlight is interrupted, causing the GUI to flicker or go dark and making the display unreadable. | Device Design | A Medical Device Field Correction letter dated 10/26/2007 was sent to all customers who received the Esprit Ventilator. The recall letter informed the customers of the problems identified and the actions to be taken. Customers with questions are instructed to contact Respironics' Customer Service Dept. at (800) 345-6443, select Option#5, then select Option #4. |  | 3255 | 3,255 units (2302 units in US) | USA Nationwide and Worldwide to Algeria, Argentina, Australia, Bangladesh, Brazil, Cambodia, Canada, Chile, Colombia, Ecuador, Egypt, France, Germany, Guatemala, Hong Kong, India, Italy, Japan, Libya, Malaysia, Mexico, Pakistan, Peru, Philippines, Puerto Rico, Russia, South Africa, South Korea, Spain, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Terminated | August 26, 2013 | 2128 | 3255 | Not_Computer
Z-2091-2013 | 65513 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Free Lock Femoral Hip Fixation System Compression Tube/Plate, different length (59, 75, 91, 107, 123 mm), holes (2, 3, 4, 5, 6) and degree angles (130, 135, 140, 145 and 150). Item codes starting with 00-1181-130, 00-1181-135, 00-1181-140, 00-1181-145, 00-1181-150). |  |  | June 19, 2013 | 2013 | Zimmer, Inc. | Single packaging configuration used for the VERSA-FX-II Femoral Fixation System failed the leak test compromising the sterility of the product. | Packaging process control | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE NOTIFICATION" letter dated July 11, 2013 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to inspect the device according to the insert instructions; not use the device if damage to the corner of the inner tray is observed, and contact the firm for return and questions. Customers should also keep this notification in mind and inform Zimmer of product with this condition using zimmer.per@zimmer.com.If after reviewing this notification you have questions or concerns, please contact Zimmer at 1-877-946-2761. |  | 85 | 85, 794 units in total | Worldwide distribution: US Nationwide (including PR and USVI), and countries of: Israel, Argentina, Brazil, France, Mexico, Saudi Arabia, Beirut, China, Korea, San Salvador, Honduras, Canada, Australia, Switzerland, Germany, Arab Emirates, UK, Belgium, Egypt, Spain, France, Italy, Iran, Lebanon, Jordan, Iraq, Canary Islands. | Terminated | June 16, 2015 | 727 | 85 | Not_Computer
Z-2085-2013 | 65770 | No | NYQ |  |  | Immunology | PMA | HSV 2 Primers, Catalog #18081020. Analyte Specific Reagent. The HSV-2 primer pair is specific for HSV-2 (Herpes Simplex Virus 2). For use in Clinical laboratories when developing their own diagnostic tests. |  |  | May 13, 2013 | 2013 | Life Technologies Corporation | Firm received complaints that HSV-2 primers failed to amplify HSV-2 DNA. | Process design | Life Technologices sent an Urgent Product Recall letter dated May 13, 2013 by certified mail to all affected customers. The letter identified the product, description of problem, and actions to be taken by the user. Customers were instructed to examine inventory immediately, discontinue use and dispose all affected units. Customers were advised to complete the Customer Response Sheet included for product/quantities consumed, and send e-mail to techsupport@lifetech.com. For questions about the details of this product issue, were instructed to contact Technical Support at 800-955-6288 option 4. |  | 15 | 15 kits | USA Nationwide Distribution in the states of DE, IA, MN, and NC. | Terminated | August 28, 2013 | 107 | 15 | Not_Computer
Z-2078-2013 | 65800 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | TRIAL STEM for MP(R) RECONSTRUCTION PROTHESIS, REF 99-0155/20, Stem size S-4, prox. Stem dia. 20 mm, Length 180 mm, WALDEMAR LINK GmbH &amp; CoKG, Barkhausenweg 10 . D . 22339 Hamburg |  |  | July 16, 2013 | 2013 | Waldemar Link GmbH &amp; Co. KG (Corp. Hq.) | The trial stems were manufactured using the wrong design. | Device Design | The firm notified their consignees of their product recall by parcel post on 07/16/2013. |  | 8 | 8 units | US Distribution to the state of Tennessee only. | Terminated | June 04, 2014 | 323 | 127 | Not_Computer
Z-2056-2013 | 65824 | No | HBG | Manual cranial drills, burrs, trephines, and their accessories | Neurology | Neurology | 510(k) | Integra Leyla Ball Joint Clamp, also as included in Leyla Table Mounting Hardware. |  |  | July 17, 2013 | 2013 | Integra Limited | As a result of complaint investigations it was identified that there is a potential that the Leyla Ball Joint Clamp (R2383) may not provide the level of stability needed to maintain retraction in all cases. | Device Design | Integra sent an Urgent Medical Device letter dated July 17, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were requested to review your inventory and determine if you have any of the affected products. If so, stop using them immediately and remove them from your service. When your form is received, Customer Service will you an RMA number and directions for returning the products.Should you have any questions regarding these instructions, please contact Integra Customer Service at 1-800-997-4868. |  | 65 | 65 | Worldwide Distribution - USA including Puerto Rico and the states of AZ, CA, FL, GA, IL, IN, KY, MA, MI, MS, NE, NM, NY, OH, PA, TN, TX, UT, VA, WA, WI and WV., and the countries of Spain, Singapore, South Africa, Canada, Libya,Pakistan and Cayman Islands. | Terminated | October 01, 2014 | 441 | 65 | Not_Computer
Z-2096-2013 | 65874 | [6 empty] | OLA Buffer |  |  | July 12, 2013 | 2013 | Celera Corporation | The package insert for FG OLA Buffer, GPR displayed the wrong concentration of ÂżM instead of the correct mM unit. | Error in labeling | Celera sent a customer notification letter dated July 12, 2013 by e-mail and mailed on July 19, 2013 via first class mail to all affected customers. The letter identified the affected product, problem and actions to be taken. The information is being provided for records and assures customer that there were not quality issues with the product and it may be used as usual in laboratory test procedures. For questions call 1-810-749-4200 or contact your local authorized Celera representative. |  | 43 | 43 units | USA Nationwide in the state of CA | Terminated | October 01, 2013 | 81 | 43 | Not_Computer
Z-2083-2013 | 65773 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | BRAINLAB; FRAMELESS SRS QA TARGET POINTER |  |  | July 16, 2013 | 2013 | Brainlab AG | The Frameless SRS QA Target Pointer - Pointer Cap with the engraved cross hairs may become loose even when only a moderate force is applied to the Pointer Cap. | Device Design | BrainLab sent an "FIELD SAFETY NOTICE / PRODUCT NOTIFICATION" dated July 9, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to discontinue use of the SRS QA Target Pointer. Consignees can instead use the Phantom Pointer feature until Brainlab can provide replacement hardware. Brainlab anticipates being able to provide all impacted facilities with replacement hardware by the end of July 2013.We sincerely apologize for any inconvenience and thank you in advance for your co-operation. If you require further clarification, please feel free to contact your local Brainlab Customer Support Representative. Customer Hotline: +49 89 99 15 68 44 or +1 800 597 5911 (for US customers) or byE-mail: support@brainlab.com. Fax Brainlab AG: + 49 89 99 15 68 33 |  | 6 | 6 units | Worldwide Distribution - US Distribution only to AR., and the countries of Australia, Germany and Japan. | Open, Classified |  |  | 6 | Not_Computer
Z-2058-2013 | 65631 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Trauma Nail System. The devices are indication for bone fixation. |  |  | November 02, 2012 | 2012 | Synthes USA HQ, Inc. | Recall was initiated due to the possibility that the outer pouch was compromised. The product is packaged in two pouches; a sterile, internal pouch within an external pouch. It is the outer, external pouch that may be compromised. | Packaging change control | The firm initiated their recall of this product on November 2, 2012 by sending an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter to all consignees. The letter explained the problem and gave instructions for consignees to examine their inventory and urged consignees to return the identified devices. The letter stated If returning the product would adversely impact their ability to provide necessary medical care to patients, they should re-sterilize the product per Pre-Vacuum Sterilization instructions contained in the Instruction for Use. |  | 1011783 | 1,011,783 | Nationwide distribution. | Open, Classified |  |  | 1011783 | Not_Computer
Z-2088-2013 | 65607 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista 500 and Dimension Vista 1500 System Aliquot Well Double Dispense with software versions 3.5.1 and 3.6. |  |  | June 19, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens is conducting a field correction for Dimension Vista 500 and Dimension Vista 1500 systems using software versions 3.5.1 and 3.6. Under specific rare conditions which require a system reset, there is a potential for a sample to be dispensed into an aliquot well that has previously been used for a QC or calibrator product from a Vista vial. If two samples are dispensed into the same aliquot | Software Design Change | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated June 2013 to its consignees/customers. The letter describes the product, problem and actions to be taken. The customers were instructed to take the following actions: 1) In the event that an instrument Reset is required for any reason, review the Current Activities Screen to determine if there are any pending QC or Calibration tests.2) If there are no pending QC of Calibration tests, Rest the instrument or troubleshoot the instrument per the Operators Guide.3) If there are pending QC or Calibration tests, restarting the software is necessary before resuming patient processing. Go to Operation -&gt; Shutdown, confirm Yes, then on the blue console menu screen select start Vista. Consignees/customers were also asked to complete the Field Effectiveness Check form attached to the letter via fax to Siemens Healthcare Diagnostics at (302) 621-8467, and retain the letter with laboratory records, and forward the letter to those who may have received the product.If you have any questions, contact your Siemens Customer Care Center at (800-441-9250)-Technical Solutions or your local Siemens Technical Support representative. |  | 1937 | 1937 | Worldwide distribution: US (nationwide) including PR and countries of: Australia, Austria, Belgium, Canada, Czech Republic, Denmark, Finland, France, Germany, Italy, Japan, Malaysia, Netherlands, New Zealand, Norway, Portugal, Republic Korea, Saudia Arabia, Slovakia, Slovenia, South Africa, Spain, Switzerland, and United Kingdom. | Terminated | August 27, 2014 | 434 | 1937 | Computer
Z-2057-2013 | 65556 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4mm VA-LCP Two Column Volar Distal Radius Plate and Torque limiting attachment, 1.2Nm |  |  | May 31, 2013 | 2013 | Synthes (USA) Products LLC | The firm is initiating a Field Safety Correction for the labeling of the 2.4mm VA-LCP Two Column Volar Distal Radius Plate and 1.2 Nm Torque Limiting Attachment. The US Technique guide lists 1.2Nm TLA as an option for use with the 2.4mm Variable Angle LCP (VA-LCP) Two-Column (Standard &amp; Narrow) Volar Distal Radius Plate. The potential exists for the VA locking screw to be driven through the 2.4mm | Device Design | Synthes sent an Urgent Notice: Medical Device Labeling Correction letter dated May 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm is initiating a Field Safety Correction for the labeling of the 2.4mm VA-LCP Two Column Volar Distal Radius Plate and 1.2 Nm Torque Limiting Attachment. The US Technique guide lists 1.2Nm TLA as an option for use with the 2.4mm Variable Angle LCP (VA-LCP) Two-Column (Standard &amp; Narrow) Volar Distal Radius Plate. The potential exists for the VA locking screw to be driven through the 2.4mm VA-LCP Two Column Volar Distal Radius Plate when used with a 1.2Nm Torque Limiting Attachment. A screw that passes through the VA-LCP plate is not likely to result in a serious adverse event although should it occur, additional medical intervention during the index procedure could be needed.If you have any questions, please call (610) 719-5450, or contact your Synthes sales Consultant. |  | 119891 | 119, 891 | Nationwide Distribution including the states of AK. AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MD, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI SC, SD, TN, TX, UT, VA, VT, WA WI and WV. | Terminated | September 21, 2015 | 843 | 119891 | Not_Computer
Z-2081-2013 | 64852 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Test Discs F300 (Nitrofurantoin), Cat. #CT0036B, packaged in cartridges of 50 discs, 5 cartridges/pack. The firm name on the label is Oxoid Ltd., Basingstoke, Hampshire, England |  |  | March 25, 2013 | 2013 | Remel Inc | Use of affected lots may give false indication of susceptibility to Nitrofurantoin. | Unknown/Undetermined by firm | ThermoFisher sent an Medical Device Recall letter dated March 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please notify any personnel who needs notified of the potential for a false indication of susceptibility to Nitrofurantoin. Please review your inventory for affected product lots and discard the affected product. Return the attached Medical Device Recall Return Response acknowledging the disposal of the affected product.If you have any questions, please contact our Technical Services Department at 800-255-6730 (US) or 913- 888-0939 (International). |  | 60 | 60 packs | US Distribution including the states of AL, AR, CA, FL, GA, IL, KS, LA, MA, MD, NY, OK, PA, RI, TX, VA, and WA. | Terminated | June 03, 2014 | 435 | 60 | Not_Computer
Z-2087-2013 | 65890 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4 VA-LCP 2 CLMN VLR DSTL Radius PL 7H HD/3H Shaft/Right |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | The possibility exists that 2.4 mm VA-LCP 2 column distal radius plates from a specific lot are etched with "R" for right plate when they should be etched with "L" for left plate. | Process control | Synthes GMBH sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory and immediately remove the affected product from stock. Customers were advised to complete the Verification Section at the end of the letter. For questions call 610-719-5450 or contact your Synthes Sales Consultant. |  | 11 | 11 | USA Nationwide Distribution in the states of FL, GA, NC, OH, KY, NJ, ID, IN, and MN. | Terminated | May 21, 2015 | 660 | 11 | Not_Computer
Z-2089-2013 | 65769 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | Mc Kesson Anesthesia-Rx Model Number 8.0.2 &amp; 8.0.1 with RapiDispense |  |  | July 22, 2013 | 2013 | MCKESSON TECHNOLOGIES INC | The firm issued a clinical alert after users reported an issue where the Anesthesia-Rx system was interrupted ANF became temporarily unavailable during use. | Software design | McKesson sent a Clinical Alert letter dated July 22, 2013 to all affected customers. The letter identified the affected product, symptoms, and actions to be taken. McKesson committed to releasing 8.0.2 Update 1.5 to affected customers which will include enhanced code to prevent the unhandled excpetion errors. Additional resolution steps were included in the letter. For questions call 1-800-700-8737. |  | 106 | 106 | USA Nationwide Distribution in the states AR, MD, NJ, NY, PA, and WV. | Terminated | June 19, 2014 | 332 | 106 | Not_Computer
Z-2082-2013 | 65080 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008 MeDS Pump, Venofer(R) Pump |  |  | April 01, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. dba Renal Solutions | Product is being recalled because the combination of the current 50 mg vial, the 2008 MeDS pump, and the 2008 MeDS pump tubing Set for CombiSet Bloodlines can lead to a situation where Venofer(R) delivery does not meet dosage accuracy requirements. This may lead to a slight over or under delivery of Venofer(R) that does not pose an acceptable risk to the patient. Please note that this recall | Device Design | Renal Solutions sent an Urgent Notification letter dated April 1, 2013 via certified mail to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to cease further use of the device. The letter informed customers that a Fresenius Medical Care Technical Service/Customer Service representative will contact customers to arrange for removal of the device. For questions contact Fresenius Medical Care Customer Service at 800-227-2572. |  | 132 | 132 | USA Nationwide Distribution in the states of CA, MA, and NY. | Terminated | March 02, 2015 | 700 | 132 | Not_Computer
Z-2054-2013 | 65787 | No | GXP | Methyl methacrylate for cranioplasty. | Neurology | Neurology | 510(k) | Stryker Part Numbers 79-43903 (3cc), 79-43905 (5cc), 79-43910 (10 cc), 79-43915 (15cc), 397003 (3cc), 397005 (5cc), 397010 (10 cc), 397015 (15 cc) HydroSet Injectable HA Bone Substitute Sterile R/EO, nonpyrogenic, Product of Ireland |  |  | July 30, 2013 | 2013 | Stryker Craniomaxillofacial Division | Stryker received reports of the desiccant bag in the kit leaking silica gel granulate. This could increase exposure to anesthesia during surgical delay up to 60 minutes or inadvertent implantation of silica gel granulate which could cause foreign body reaction including inflammation and might lead to the need for revision surgery. | Package design/selection | Stryker sent an Urgent Medical Device Recall Notices with attached Product Correction Bulletin dated, July 30, 2013, describing the product, problem and possible adverse events associated with the recall were sent to all affected customers. Customers were instructed to pass the notice and bulletin to all individuals who may be using the device, provided instructions on procedures which would help mitigate risk and return the customer response form to our Customer Service Team by email, fax or letter.Email: CMF-custserv@stryker.comFax: 1-877-648-7114Address: 750 Trade Centre Way, Suite 200Portage, MI 49002 Customers responses will be tracked and non-responders will be contacted.Further questions please call (877) 534-2464. |  | 12719 | 12,719 | Worldwide Distribution - USA (nationwide) including the countries through Stryker Distribution: Austria, Belgium, Denmark, Estonia, Finland, France, Germany, Ireland, Italy, Latvia, Luxembourg, Netherlands, Norway, Poland, Portugal, Romania, Serbia, Spain, Sweden, Switzerland, United Kingdom, Argentina, Australia, Brazil, Jordan, Kingdom of Saudi Arabia, Kuwait, Latin America (others), New Zealand, Russia, Singapore, South Africa, and United Arabian Emirates. | Terminated | April 17, 2014 | 261 | 12719 | Not_Computer
Z-2059-2013 | 65950 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Easy Upgrade DR, Code No: 712086 |  |  | August 08, 2013 | 2013 | Philips Healthcare Inc. | The mirror icon on the bottom of each image could be misinterpreted as a RIGHT lead marker for side indication if the image is mirrored within a PACS system. | Software design | Philips Healthcare sent an Urgent Field Safety Notice letter dated July 24, 2013, to all affected customer. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed : When reading an image, a physician should not use the mirror icon for image side interpretation. 2. Clinical staff should follow the Instructions for Use and use lead letters to indicate body orientation of patient or equivalent electronic markers to provide Left/Right indications, during all X-ray image acquisitions. Philips will upgrade the system, which involves software and hardware upgrades of the system. The software provides a replacement icon that reduces the risk of the potential misinterpretation of the original mirror icon. Philips will also provide updated Instructions for Use for the work spot and the system (including an addendum).If you need any further information or support concerning this, please contact your local Philips representative Technical Support Line at 1-800-722-9377.We apologize for any inconvenience this may cause and trust that this information is adequately addressing any concerns you may have. |  | 89 | 89 | Worldwide Distribution - US Nationwide including the states of AZ, CA, CO, DE, FL, GA, MO, MS, NC, NH, NY, PA and TX., and the countries of Canada, Australia, China, Czech Republic, France, Germany, Greece, Italy, Netherlands, Portugal, Russia, Saudi Arabia, Slovenia, Spain, Sweden, Switzerland, Turkey, United Arab Emirates and United Kingdom. | Completed |  |  | 89 | Computer
Z-2055-2013 | 65677 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00 4309-049-00 Trabecular Metal Reverse Shoulder Glenosphere Distractor Rx, Non sterile (sterilize before use) |  |  | August 01, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the Trabecular Metal" Reverse Glenosphere Distractor due to the potential of the device not functioning properly should it be utilized without proper lubrication. As a result, there is a potential for the device to fail to provide impact to the Glenosphere when the trigger is pulled. In addition, there is a potential for the foot of the device to fract | Device Design | Zimmer sent a Urgent Medical Device Recall Notification letter August 1, 2013, and via E-mail and courier to all affected customers. The initial recall notice identified the reason for the recall, possible adverse risks and injuries and informed customers that a replacement will be available in the coming months. In addition, this letter will remind that the instrument is to be lubricated before each use and to inspect the foot. Customers are instructed to inspect any recalled product before and after use for cracks at the foot and ensure the foot is intact. Also ensure all moving parts are lubricated prior to each use according to the Recommendations for Care, Cleaning, Maintenance, and Sterilization Manual 97-5000-170-00. If you find a device which does not function properly or has a damaged foot, provide it to your Zimmer sales representative for return to Zimmer on a Product Experience Report (PER). Questions and concerns should be addressed to the Zimmer customer call center at 1-877-946-2761 between the hours of 8 a.m. and 5 p.m. EST. Phase 2: As devices become available for distribution, a 2nd letter will be sent to the distributors and risk managers. Distributors will be instructed to order replacement devices as a means of completing the affected instrument cases in circulation within their territories. Instructions for the return of devices will be included. |  | 771 | 771 | Worldwide Distribution - USA (nationwide) including the countries of Singapore, Brazil, Taiwan, Australia, Canada, Germany(with further distribution to UAE, Austria, Belgium, Switzerland, Czech Republic, Denmark, Spain, Finland, France, Greece, Israel, Italy, Malta, Norway, Portugal, Sweden and South Africa, | Terminated | March 02, 2015 | 578 | 771 | Not_Computer
Z-2095-2013 | 65848 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | PRIMUS, ONCOR, ARTISTE Linac systems with component Tabletop, 550TxT, TT-A s/n 1001-1233 and Tabletop 550TxT, TT-S s/n 1001-1083. |  |  | July 18, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens has become aware of a customer complaint concerning the rail threaded holes of the 550TxT treatment tabletop. Due to a manufacturing error, the screws holding the side rails in position on the tabletop might break off. | Device Design | Siemens released Updated Instruction containing the Medical Device Correction Field Safety Notice of Improvement letter, which was delivered by the Siemens Customer Service Group to all affected customers. The release of this UI TH002/12/S to distribute the update instruction began on July 18, 2013. This document includes completion protocol, which will be completed at each facility and returned to Siemens. A copy of the notification should be kept in the Digital Linear Accelerator System owners manual under Safety Advisory letters.UPDATE: a secnd update intruction, TH003/12/S has been released to begin the replcement of tabletops identified through the inspection carried out through TH002/12/S. |  | 179 | 179 active devices. | Worldwide Distribution - USA Nationwide in the states of NJ, IN, PA, FL, OH, MN, MO, CA, AL, NE MI, LA. and countries of:Argentina, China, South Africa, Pakistan, Spain, Switzerland, Italy, Algeria, Croatia, Turkey, Canada, Germany, France, Poland, Japan, Russian Federation. Lebanon, Egypt, Slovakia, Latvia, Kenya, Brazil, Kuwait, Saudi Arabia, Greece, Belgium, Taiwan, Portugal. | Terminated | November 24, 2014 | 494 | 179 | Not_Computer
Z-2086-2013 | 65672 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Illico Ti Cannulated Polyaxial Screw 7.5mm X 40mm. The Illico MIS Posterior Fixation System is intended to facilitate the surgical correction of noncervical spinal deformities by providing temporary internal fixation and stabilization during bone graft healing and/or fusion mass development. |  |  | June 28, 2013 | 2013 | Alphatec Spine, Inc. | The firm initiated this recall of certain lots of Illico Ti Cannulated Polyaxial Screws 7.5mm x 40mm implants, due to an off-center cannulation hole at the proximal end of the bone screw. There is a potential risk that when the screw is mated to a polyaxial screwdriver, the guide wire may not pass through the screwdriver. | Equipment maintenance | The firm, Alphatec Spine, sent an "URGENT-MEDICAL DEVICE RECALL NOTIFICATION" with response sheet dated June 28, 2013 via a combination of e-mail, Fedex, and telephone contact to its customers. The notification described the product, problem and actions to be taken. The customers were instructed to immediately return any and all affected lots of the illico Ti Cannulated Polyaxial Screw 7.5 mm x 40 mm implants that have been sent to them; if the device has previously been implanted, there is no further or additional risk post-implantation; immediately contact Alphatec Spine Customer Service Department - customerservice@alphatecspine.com on how to replace the implants; and complete and return the last page of the notification via mail to Alphatec Spine, Inc. 5818 El Camino Real, Carlsbad, CA 92008; email: sarrhenius@alphatecspine or fax to: (760) 431-0289.If you have any questions or need any additional information, please contact Director of Compliance at 760-494-9636 or email at dmcgrath@alphatecspine.com. |  | 333 | 333 screws | Worldwide distribution: US (nationwide) including states of:US states: AZ, CA, FL, GA, IL, MD, NJ, NY, OK, PA, TN, TX,and WA ; and countries of: Spain, France, Japan.International: Spain, France, Japan. | Terminated | October 09, 2014 | 468 | 333 | Not_Computer
Z-2099-2013 | 65891 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker |  |  | July 08, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report that the length of the cannulated screw asnis III, article # 604640 Lot R11482, was incorrect. | Process control | Stryker sent an Urgent Medical Device Removal letter dated July 8, 2013 via Fed Ex to all affected customers. The letter identified the affected product, product issues, potential hazards, risk mitigation and actions to be taken. Customers were instructed to return all affected products available to Stryer Osteosynthesis or contact Stryer customer service and complete and sign the enclosed Business Reply Form and fax a copy to: 1-865-251-3635. Questions can be directed to Stryker at 201-972-2100. |  | 23 | 23 units | USA Nationwide Distribution in the states of: FL. OH, TX, &amp; VA | Terminated | March 31, 2015 | 631 | 23 | Not_Computer
Z-2107-2013 | 65914 | No | GAA | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SpineNet Bone Marrow Aspiration Needle Kit (which is the Ranfac Bone Marrow Aspiration Needle #74150-01M). |  |  | February 04, 2013 | 2013 | SpineNet | SpineNet was marketing this device to be used with SpineNet ACC device without FDA clearance or approval. | No Marketing Application | SpineNET sent a "VOLUNTARY PRODUCT RECALL" letter dated 02/04/2013 to Halifax Medical Center. The letter identified the product, problem, and actions to be taken by the customer. Contact the firm at 407-539-2483 for questions regarding this notice. |  | 100 | 100 units | Distributed to one customer in Florida. | Terminated | November 15, 2013 | 284 | 100 | Not_Computer
Z-2103-2013 | 65917 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | SpineNet Anterior Cervical Cage, Lordotic, Peek (originally submitted as Daytona Anterior Cervical Cage) |  |  | February 05, 2013 | 2013 | SpineNet | The recall decision has been made because because it has recently been determined that the instructions for use (IFU) had certain statements that should have been removed prior to initial product marketing. | Error in labeling | SpineNET sent a Customer Notification letter dated 02/05/2013 to their only customer (Halifax Medical Center). The letter identified the affected product, problem and actions to be taken. Customers were informed to remove any copies of the original IFU-ACC-01 form from their records and replace with the IFU-ACC-02 form attached to the letter. For questions contact SpineNET office directly. |  | 400 | 400 | USA Nationwide Distribution in the state of Florida | Terminated | May 02, 2014 | 451 | 400 | Not_Computer
Z-2100-2013 | 65913 | No | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE Unitek Primary Stainless Steel Crowns (First Primary UL4), Product code 902150 (Kit- 112 crowns/kit) and 900114 (Refill- 5 crowns/box). |  |  | August 05, 2013 | 2013 | 3M Company/3m Espe Dental Products | 3M ESPE is recalling Unitek Primary Stainless Steel Crowns (First Primary UL4) because these products have been manufactured with the incorrect marginal contour edge. | Process control | 3M ESPE sent a 3M "Voluntary Recall" letter dated August 1, 2013 to all affected customers. Separate letters were addressed to distributors and doctors. The letters described the product affected by the recall, problem and actions to be taken. The doctors were advised to complete the attached inventory checklist and call the 3M ESPE Dental Products Customer Care Center at 1-800-634-2249. Select option #1 to arrange for return and replacement of your product.Distributors were advised to quarantine and return the product and not to further distribute it. For product return they should contact 3M ESPE Dental Products Customer Service at 1-800-237-1650. If the product was further distributed they ask that they inform them of the recall. A letter was prepared that can be used for their communication to customers. for additional questions they should contact 651-733-7767. |  | 3409 | 3409 crowns (2949 US and 460 OUS) | Worldwide Distribution - USA Nationwide in the states of CA, FL, GA, IN, IA, MO, NV, NJ, NY, PA, SC, TN, TX, WA. and the country of CANADA. | Terminated | December 11, 2013 | 128 | 3409 | Not_Computer
Z-2097-2013 | 65995 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | E-Z CleanÂż Disposable Laparoscopic Electrode Model Number 0100L and 0100LS. Expiration dates of 05/2018 to 07/2018. Sterile. RX ONLY. This product is packaged individually in a chevron TyvekÂż - polyester peel pouch, 6 pouches per box. |  |  | August 07, 2013 | 2013 | Megadyne Medical Products, Inc. | Megadyne Medical Products is recalling certain lots of E-Z Clean, L-Wire laproscopic Electrode Standard Stem and Split Stem due to potential breach of insulation. | Process control | Megadyne sent an Urgent: Medical Device Recall Field Safety Notice dated August 7, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customer were also notified with follow-up phone calls and email beginning on August 7, 2013. Customers were instructed to examine their stock immediately to determine if you have any of the products from the affected lots on hand. If so, please remove form service/discontinue use and contact Megadyne Customer Service at 1 (801) 576-9669 or (800) 747-6110 (USA) for return authorization and replacement or credit. |  | 1816 | 1816 | Class II Recall - Worldwide Distribution - US including the states of CA, HI, IL, MN, NC, NJ, OH, PA, TX, UT and VA. | Terminated | January 28, 2014 | 174 | 1816 | Not_Computer
Z-2104-2013 | 65831 | No | DWE | Cardiopulmonary bypass pump tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Custom Perfusion System. Model Numbers: 1D80R8, 1E34R1, 1F75R1, 1G49R2, 1G75R1, 1P07R8, 1P91R4, 5B15R11, and 5P51R9. |  |  | July 17, 2013 | 2013 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic was notified that Edwards Lifesciences has initiated a product recall for the Rigid Suction Wand, model S099B. The Wands are included within certain Perfusion Tubing Packs that are manufactured and distributed by Medtronic. | Process control | Metronic sent an "Urgent Field Safety Notice" letter dated July, 2013 to all affected customers. The letter described the affected product, problem and actions to be taken by User. Customerss were advised to remove and destroy all affected products. At the same time customers were informed that replacement products can be ordered by contacting their Medtronic representative. Customers were requested to complete the Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612. For any questions they can contact Medtronic Lifeline Technical Services at 877-526-7890 or their Medtronic representative. |  | 13223 | 13223 | Worldwide Distribution - USA Nationwide CA, IL, LA, MI and the country of CANADA | Terminated | February 04, 2014 | 202 | 14198 | Not_Computer
Z-2101-2013 | 65850 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | ZIMMER PERIARTICULAR LOCKING PLATE SYSTEM |  |  | July 30, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific voluntary recall of the Distal Medial Humeral Locking Plate, Short due to the incorrect information printed on the Chinese label. The original label indicates that implants are RIGHT sided, but the Chinese label indicates that the implants are LEFT sided. | Labeling mix-ups | Zimmer sent an Urgent Medical Device Recall Notification letter dated July 30, 2013 to Distributors and Hospitals in China. The letter described the affected product, problem, labeling issue, directions on how to respond if affected product is discovered, and potential health risks associated with the labeling error. An inventory return certification form is included with the recall notification. Questions or concerns are directed to Zimmer Shanhai at 021-2211 5196, Monday to Friday, between the hours of 9 a.m. and 5 p.m. |  | 11 | 11 units | Internationally Distributed in China only. | Terminated | April 15, 2014 | 259 | 11 | Not_Computer
Z-2098-2013 | 65880 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Ultrasound ViewPoint 6. |  |  | August 07, 2013 | 2013 | GE Healthcare, LLC | GE has recently become aware of a potential safety issue due to the incorrect calculation of the pulsatility index value (PIV) for venous flow measurements associated with obstetric reporting of the GE ViewPoint 6 system. The ViewPoint calculated Doppler PIV for veins could be incorrect. An incorrect calculation of the ductus venosus PIV could lead to false high risk for trisomy 13, 18, and 21, | Software Design Change | GE Healthcare sent An Urgent Medical Device Correction letter dated August 7, 2013 to all affected customers. The letter described the affected product, problem and instructed customers to discontinue use of venous PIV for making any and all clinical decisions. The letter informed customers that a GE service engineer will install a correction at no cost. For questions contact your local GE Healthcare Representative or call 877-644-3114 (US) and 49 (0)8153-9311-11(Europe/Asia) |  | 225 | 225 | Worldwide Distribution - USA Nationwide including Puerto Rico and the countries of: ANDORRA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BRAZIL, BULGARIA, CANADA, CHILE, COLOMBIA, CZECH REPUBLIC, ESTONIA, FINLAND, GERMANY, HONG KONG, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, KUWAIT, MALAYSIA, Norway, OMAN, PALESTINIAN TERRITORY, OCCUPIED, POLAND, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TURKEY, and UNITED KINGDOM. | Terminated | April 25, 2014 | 261 | 225 | Not_Computer
Z-2109-2013 | 65886 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ACCU-CHECK FastClix Lancing Device. It is intended for use in conjunction with a blood glucose monitoring system. |  |  | August 19, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Internal Roche inspections have revealed that after the lancet is used for a finger stick, the needle may not fully retract back within the lancet drum or back into the cap. This allows the needle to protrude inside the cap or outside of the cap of the ACCU-CHEK FastClix lancing device. | Equipment maintenance | On 8/19/ 2013, Roche sent Distributor/Retailer URGENT MEDICAL DEVICE CORRECTION notifications which described the recall issue and health consequences, provided product description, and provided actions required of consignees. Roche sent a CUSTOMER BULLETIN to their direct customers. URGENT MEDICAL DEVICE CORRECTION faxback form for consignee acknowledgement of recall is to be faxed to 1-888-240-2804. Consignees may contact the ACCU-CHEK Customer Care Service Center, 24 hours a day, seven days a week at 1-800-440-3638 if they have questions about the information contained in the Distributor/Retailer Letter. |  | 36332 | 36,332 | Nationwide Distribution | Terminated | May 12, 2014 | 266 | 36332 | Not_Computer
Z-2114-2013 | 65852 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | ADVIA 120 Hematology Analyzer, Catalog Numbers SMN 10360958 and 10360959 . |  |  | June 26, 2013 | 2013 | Siemens Healthcare Diagnostics | The ADVIA 120 optics cover or hood is located on the very top of the ADVIA 120 analyzer. When the optics cover is raised, it is supported by two gas struts (SMN 10309266) attached to the cover on the left and right sides. Over time, it is possible that these struts may lose their effectiveness and fail to support the optics cover in the open position. This could cause the cover to fall during m | Equipment maintenance | Siemens sent an Urgent Device Correction Letter dated June 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed of the ADVIA 120 system displays a problem when raising and supporting the optics cover, immediately contact your service provider to arrange to have the cover struts serviced. If the optics cover will not stay in the up position, as required, do not perform maintenance procedures on the optics area until the hood struts have been replaced. Siemens field support will be proactively replacing the optics cover hood struts on a regular preventive maintenance cycle based on the age of the system.Please forward this notification to whomever you may have distributed this productWe apologize for the inconvenience that this situation has caused. Thank you for your patience and continued support.This response form is to confirm receipt of the enclosed Siemens Healthcare Diagnostics Urgent Device Correction dated June 2013 regarding Optics Cover Struts (SMN 10309266).Please read each question and indicate the appropriate answer. Fax this completed form to Siemens Healthcare diagnostics at the fax number indicated at the bottom of this page.PLEASE FAX THIS COMPLETED FORM TO THE TECHNICAL SOLUTIONS CENTER AT (302) 631-7597. |  | 4319 | Domestic: 950 units; Foreign: 3369 units | Worldwide Distribution - USA (nationwide) and the countries of Austria, Australia, Belgium, Bosnia - Herzegovina, China, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Ireland, Iraq, Israel, Italy, Japan, Kenya, Republic of Korea, Kuwait, Lebanon, Lithuania, Luxembourg, Martinique, Malaysia, Netherlands, Norway, New Zealand, Philippines, Poland, Portugal, Romania, Russian Republic, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom and Vietnam. | Open, Classified |  |  | 4319 | Not_Computer
Z-2118-2013 | 49516 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Benchmark XT/LT and Discovery XT |  |  | June 12, 2008 | 2008 | Ventana Medical Systems Inc | Ventana Medical System is initiating the recall of the BenchMark and DISCOVERY series instruments because there is a possibility that the heating element could separate from the underside of the heater cap and cause a non-staining or inconsistent staining event. Additionally, the system may not detect this anomaly and therefore it may not generate an error message for the slide position. | Device Design | Ventana sent an Urgent Medical Device Correction Notification letter dated June 16, 2008 to all affected customers. The letter identifed the affected product, problem and actions to be taken. The letter instructed customers to carefully read and perform all of the actions outlined in the notification letter concerning potential safety issue. Customers with concerns regarding the staining results on any patient specimen are instructed to follow their internal Quality Procedures regarding the review of patient reports. For question contact your local Customer Service Center. |  | 1217 | 1217 units total | Worldwide Distribution - USA Nationwide and the countries of Canada, Europe, Asia, Japan. | Terminated | August 29, 2013 | 1904 | 1217 | Not_Computer
Z-2111-2013 | 65779 | No | MKQ |  |  | Pathology | PMA | PrepStain Tecan US-I. |  |  | July 08, 2013 | 2013 | Becton Dickinson &amp; Co. | PrepStain system rack setting was incorrectly entered on one PrepStain system during a routine preventive maintenance visit by a BD Field Service Representative. | Software design | Ex-US BD sites were notified via email and teleconference, and will communicate directly with customer sites in their region as required by local regulation. All US customers were initially contacted via telephone and letters were mailed. |  | 890 | 890 instruments | Worldwide Distribution-USA (nationwide) and Puerto Rico and the country of Canada. | Terminated | May 01, 2014 | 297 | 890 | Not_Computer
Z-2117-2013 | 66016 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, TRAM MODULE, High-acuity patient monitoring. |  |  | April 03, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to intermittent noise on ECG and/or Respiration waveforms associated with movement of the ECG connector on your DASH, PDM, and/or Tram Module. ECG and Respiration signal noise caused by DASH, PDM, and/or Tram ECG cable connector movement may result in reduced ECG, Respiration and arrhythmia detection performance. | Nonconforming Material/Component | GE Healthcare "Urgent Medical Device Correction" letter dated April 3, 2013 to all affected customers. The letter was addressed to Healthcare Administrator / Risk manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Product, Safety Instructions, Affected Product Details, and Product Correction. For questions contact Technical Support at 1-800-558-7044. |  | 4807 | 4807 (2167 US, 2207 OUS, &amp; 434 unknown. | Worldwide Distribution - USA Nationwide including DC and PR except ND and the countries of: ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOSNIA-HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, CHANA, GREECE, HONDURAS, HUNGARY, ICELAND, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MEXICO, MOROCCO, NEPAL, NETHERLANDS, NEW ZEALAND, NIGERIA, NORWAY, PANAMA, PARAGUAY, PERU, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAK IA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNSIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, BOLIVIA, HONG KONG, JAMAICA, PHILIPPINES, SOUTH KOREA, TRINIDAD TOBAGO, URUGUAY, VENEZUELA. | Open, Classified |  |  | 77509 | Not_Computer
Z-2113-2013 | 65764 | No | NKC | Laparoscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | DeRoyal Insufflation Tubing sets: |  |  | July 15, 2013 | 2013 | DeRoyal Industries Inc | DeRoyal Industries is recalling their product Insufflation Tubing Sets due to materials incompatibility which could possibly result in occlusion at or near the filter housing which could prevent the flow of gas through the device. | Packaging process control | The firm notified customers via letters dated 7/15/2013 sent UPS and via DeRoyal Representatives. Customers were asked to notify DeRoyal of any remaining inventory to return for replacement product or credit. Non-responders will receive additional notifications. Effectiveness checks will be performed by contacting non-responders by phone. |  | 524486 | 506,632 sets in US; 17,854 sets International | Worldwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and the countries of Dubai, Arab Emirates, Quebec, Canada, Japan, and Ireland. | Terminated | May 20, 2015 | 674 | 524486 | Not_Computer
Z-2108-2013 | 65793 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Axiom Artis systems. C-arm angiographic x-ray system |  |  | April 26, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm became aware of a potential issue with Axiom Artis systems installed by a third party installer that may involve incorrect channel nuts being used. | Nonconforming Material/Component | Siemens sent Update Instructions dated December 2012, to all affected customers. This update was initiated on behalf of the US service organization to re-check all systems installed by MEC Group. A call center representative or designated engineer called each site on April 26, 2013, and described the work being done, amount of time needed and number of engineers that would be on site. At that time the call center or engineer scheduled the update with the customer.For further questions please call (610) 219-6300. |  | 178 | 178 | US Distribution including the states of PA, NJ, DC, VA, MD and DE. | Terminated | November 06, 2014 | 559 | 178 | Not_Computer
Z-2130-2013 | 65902 | No | GDJ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Howmedica Osteonics Triathlon tibial alignment ankle clamp extramedullary (EM) referencing instrument (Cat. No. 6541-2-609), for use in Total Knee Arthroplasty (TKA) when the surgeon elects to use Stryker's Triathlon Total Knee System in conjunction with EM referencing. Non-Sterile Instrument. |  |  | July 02, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Cracks or fracture of the Triathlon Tibial Alignment Ankle Clamp, an instrument. | Device Design | Stryker Orthopaedics sent an Urgent Medical Device Correction letter dated July 2, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by t he customers. Customers are advised of the issue, potential hazards and risk mitigation factors. Customers are asked to complete the attached form and fax to 855-251-3635. Any questions should be directed to (201) 972-2100.All letters included a Product Correction Bulletin. |  | 10537 | 10,537 units | Nationwide Distribution | Open, Classified |  |  | 10537 | Not_Computer
Z-2131-2013 | 65923 | No | JTC | Microtiter diluting and dispensing device. | Microbiology | Microbiology | 510(K) Exempt | WASP (Walk Away Specimen Process). Automatic Planting and Streaking Instrument |  |  | August 09, 2013 | 2013 | Copan Italia | The recall of the WASP (Walk Away Specimen Process) instrument was initiated because Copan Italia SpA has upgraded the software concerning the broth inoculation table feature. The upgrade has been considered necessary from the manufacturer in order to eliminate the very low risk of rotation of the broth table when the door is open. | Device Design | Copan sent an Urgent WASP recall letter dated July 31, 2013, to all affected customers that received the WASP (Walk Away Specimen Process) instrument to notify them of the software upgrade concerning the broth inoculation table feature. The upgrade has been considered necessary from the manufacturer in order to eliminate the very low risk of rotation of the broth table when the door is open. This event may happen only if damage in the door locking system happens. This could result in a minor, transient injury to the user due to the impact with the rotating table.The WASP instruments involved in this upgrade are all whose serial numbers show the configuration with the broth inoculation feature activated.The machine installed in your laboratory is one of this serial number involved in the SW upgrade. The Software upgrade will be performed on site by personnel trained by Copan and no action is required by yourself.Customers will be contacted by the WASP Technical Support in order to define a date that will fit with their planning. Customers with questions are instructed to call +39 0303666130 or write to regulatory.wasp@copanitalia.com. |  | 63 | 63 units | Worldwide Distribution - USA including Oregon and California., and the countries of France, Belgium, Italy, Denmark, Norway, Switzerland, Romania, Finland, Germany, South Africa, Canada, Ireland, UK, Sweden and Japan | Terminated | December 12, 2013 | 125 | 63 | Computer
Z-1948-2013 | 65472 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Colorimetric (CM)/ Rate (RT) incubator wear pads. Product Codes: 1) 6801375, 2) 6801890, 3) 6802445, 4) 6802413 |  |  | June 06, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling CM/RT wear pads which are a component in the CM/RT incubator due to becoming worn before their replacement. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated June 6, 2013 via FedEx to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that an OCD trained service personnel will contact them and schedule a replacement of the CM/RT wear pads. OCD requested customers to complete and return the enclosed Confirmation of Receipt form no later than June 17, 2013. Customers are asked to place a copy of the notification by each VITROS System in their facility. Foreign affiliates were notified by email on June 6, 2013 of the issue and instructed to notify their consignees of the issue and actions. If you have any questions or need additional information, please call Customer Technical Services at 1-800-421-3311. |  |  |  | Worldwide Distribution - USA Nationwide in the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, including Puerto Rico and the countries of Mexico, Singapore, Venezuela, England, Germany, Italy &amp; Spain. | Open, Classified |  |  |  | Not_Computer
Z-2136-2013 | 65823 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SC2316, Distal Femoral Augment with Screw |  |  | June 24, 2013 | 2013 | Stelkast Co | The firm became aware of an incident relating to a breach of sterility in the sterility barrier packaging of SC2316, Distal Femoral Augment with Screw. | Process control | Stelkast called and emailed all customers on June 24, 2013, to notify them of the recall. Customers were asked to recover all affected products from their inventory and return them to Stelkast. Customers were instructed to contact Stelkast Customer Service for a Return Authorization (RA) Number prior to shipment to Stelkast. Customers with questions were instructed to call 1-888-273-1583. For questions regarding this recall call 724-941-6368. |  | 11 | 11 | Nationwide Distribution including TX, VA, OK, and PA | Terminated | October 08, 2014 | 471 | 11 | Not_Computer
Z-2135-2013 | 65868 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | ADVIA Chemistry A1c_3 Calibrators, Catalog Number 10491408, IVD --- COMMON/USUAL NAME: ADVIA Chemistry A1c_3 Calibrators |  |  | July 30, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens confirmed that the ADVIAÂż Chemistry Systems HbA1c method exhibits a positive bias of up to 12% for patient and CAP survey samples due to over-recovery when using A1c_3 Calibrator lots 1MD014, 1BD063, and 2GD014. The established ranges of CAP Survey Samples GH2-04, GH2-05 and GH2-06 and IFCC sample recoveries use Âż 7% from the established mean as their acceptance criteria. This is the crite | Under Investigation by firm | Siemens sent an Urgent Medical Device Correction letter dated July 30, 2013, to all affected customers via FedEx (for US consignees). The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use and discard any of the affected reagent kit lots. Customers werealso requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within thirty (30) days.Field service personnel were sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions.Customers were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative for questions.For questions regarding this recall call 914-524-2955. |  | 2064 | US: 1546 units; Foreign: 518 units | Worldwide Distribution - USA (nationwide) Puerto Rico and internationally to Argentina, Austria, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Czech Republic, Denmark, Estonia, Finland, France, Germany, Great Britain, Greece, Guatemala, Hungary, Indonesia, Italy, Kazakhstan, Lithuania, Norway, Poland, Portugal, Romania, Russian Federation, Serbia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, South Africa, South Korea, Thailand, Turkey, Uruguay, Vatican City, and Vietnam. | Open, Classified |  |  | 2064 | Not_Computer
Z-2139-2013 | 65485 | No | OBC |  |  | Cardiovascular | PMA | LeGooÂż Endovascular Occlusion Gel-0.25 mL LeGoo with 3 Cannula Pouch |  |  | June 21, 2013 | 2013 | Genzyme Corporation, A Sanofi Company | Product contains particulates (fibers) | Material/Component Contamination | Genzyme sent a Medical Device Recall Notification letter dated June 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were avised to immediately discontinue use of all lots of LeGoo Endovascular Occlusion Gel and return all LeGoo and return product with the shipping label and packing slip enclosed. syringes.Complete the reply card via fax (414) 459-8791. Customer with questions were instructed to call GENCO at 877-674-2087.For questions regarding this recall call 508-270-2285. |  | 30 | 30 units | Worldwide Distribution - USA (nationwide) and internationally to Austria, Belgium, Canada, Denmark, Egypt. Europe, Finland, France, Germany, Hong Kong, Ireland, Italy, Netherlands, Poland, Saudi Arabia, Slovena, Spain, Sweden, Switzerland, The Netherlands, Turkey, and UK. | Terminated | August 07, 2014 | 412 | 5312 | Not_Computer
Z-2134-2013 | 64937 | No | NDN | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | CONFIDENCE SPINAL CEMENT SYSTEM KIT, Product Code: 2839-13-000 |  |  | April 17, 2013 | 2013 | DePuy Spine, Inc. | During injection of cement, the water in the hydraulic pump leaks past the piston within the pump body resulting in the loss of pressure and inability to continue to inject cement. | Process control | DePut Synthes Spine sent an Urgent Voluntary Product Recall Notification letter dated April 15, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the con signee. Consignees were instructed to review their inventory and quarantine for return any products listed in the recall notice., complete the enclosed Business Reply Form, indicating level of inventory of the affected product and return the form to Depuy via fax to 888-943-4897or email to DePuy7862@stericycle.com and to return the product with the enclosed UPS label to Stericycle.For questions regarding this recall call 508-977-6606. |  | 3896 | 3896 units (Overall quantity distributed appears greater than quantities manufactured due to the distribution pathways) | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IN, KY, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WV, WY; Worldwide: Algeria, Australia, Austria, Belgium, Brazil, Canada, Colombia, Cyprus, Czech Republic, France, Germany, Hong Kong, India, Israel, Italy, Latvia, Luxembourg, Malaysia, Mexico, Netherlands, Norway, Poland, Portugal, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UAE, and UK | Terminated | August 13, 2014 | 483 | 3896 | Not_Computer
Z-2137-2013 | 65919 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Small Peripheral Cutting BalloonTM, Sterilized using irradiation: |  |  | July 31, 2013 | 2013 | Boston Scientific Corporation | Boston Scientific is initiating a recall removal of the 4.0mm x 15mm size of Small Peripheral Cutting Balloon (Monorail and Over-the-Wire) Microsurgical Dilatation Devices and Flextome Cutting Balloon (Monorail and Over-the-Wire) Microsurgical Dilation Devices. Boston Scientific has determined that users may experience significant difficulty or inability to remove the protector cap from the devic | Process change control | Boston Scientific sent an Urgent Medical Device Recall Removal letter dated July 30, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were advised that further distribution or use of the product should cease immediately, and to segregate the product and return it to Boston Scientific. Consignees were asked to complete and return the "Recall Removal Reply Verification Tracking Form". For questions consignees were instructed to contact their local Sales Representative.For questions regarding this recall call 763-494-2042. |  | 1331 | 1,331 | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, IA, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV,NH,NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC,SC, TN, TX, UT, VA, WA, WV, and WI. Internationally to Algeria, Finland, Japan, Spain, Australia, France, Latvia, Sweden, Austria, Germany, Netherlands, Switzerland, Belgium, Great Britain, Norway, Taiwan, Brazil, Greece, Panama, Turkey, Canada, Hong Kong, Poland, Unit.Arab Emir., China, India, Russian Fed., Colombia, Ireland, Saudi Arabia, Czech Republic, Israel, South Africa, Denmark, Italy, and South Korea. | Terminated | January 27, 2015 | 545 | 1671 | Not_Computer
Z-2158-2013 | 66029 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Steinmann Pin, packed in a clear tube with end caps, tubes packed in a sterile barrier pouch, 6 pouches/sales unit. Pouches are labeled in part ***MicroAire Surgical Instruments 3590 Grand Forks Blvd Charlottesville, VA LLC 22911*** |  |  | August 12, 2013 | 2013 | MicroAire Surgical Instruments, LLC | A medical device used in surgical procedures was packaged in a pouch in which a defective seal may compromise sterility | Packaging process control | MicroAire Surgical LLC sent an Urgent Medical Device Recall letter dated August 12, 2013, to all affected customers. The letter advised consignees recover and quarantine the affected product so that it does not make it into the surgical process. Customers were instructed to either return affected product to MicroAire for inspection or they could inspect the product for the defect by looking at the seal area on the inside sterile package (short non-chevron side). The recalling firm further advised distributors to contact all of their customers to recover and quarantine the product so that it does not make it into the surgical process. Customers with questions were instructed to call 434-975-8370.For questions regarding this recall call 434-975-8370. |  | 15786 | 15786 total units | Worldwide Distribution - USA (nationwide) and internationally to Australia, Ireland, Canada, Netherlands, Sweden, Japan, and United Kingdom. | Terminated | February 12, 2014 | 184 | 15786 | Not_Computer
Z-2157-2013 | 66048 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | DMLC IV-ERGO - Beam shaping block for Radiation Therapy. |  |  | August 15, 2013 | 2013 | Elekta, Inc. | Clinical mistreatment may occur when there are errors in the clinical workflow, or settings are incorrect. | Device Design | Important Field Safety Notices 200 01 406 030 (Changes to your system to improve the safe operation of DMLC) and 200 01 602 031 (Changes to your system to improve the safe operation of 3DLINE DMLC) letters were sent to customers on 8/15/2013 advising them of the issue and recommendations to resolve either by upgrading their system or to discontinue a clinical use of the device. |  | 31 | 31 | Nationwide distribution: AL, CT, FL, GA, IA, MS, MO, NJ, NY, ND, OH, OK, PA, RI, TX, VA and Puerto Rico. | Terminated | July 28, 2015 | 712 | 31 | Not_Computer
Z-2162-2013 | 65905 | No | FQZ | Clinical color change thermometer. | General Hospital | General Hospital | 510(K) Exempt | Tempa DOT Single-Use Clinical Thermometer |  |  | July 10, 2013 | 2013 | Medical Indicators, Inc. | TempaDot products may have been mislabeled. | Labeling mix-ups | The firm, Medical Indicators Inc.sent an "URGENT" Recall letter dated July 10, 2013 and Acknowledgement Form to all distributors/customers via email and Priority Mail with delivery confirmation. The letter described the product, problem and actions to be taken. The customers were instructed to check full &amp; partial cases of the tempadot; open and inspect all 5122NS cases; pull out and quarantine any of the lots found in the 5122NS cases, pending return to medical indicators; if individual Tempadot 100-count boxes of 5122 and 5122NS are stored in different warehouse follow the instructions provided in the letter; if Tempadot 100-count boxes of 5122 were sold to a customer with the affected lots, contact the customer and have them return the product; and complete and return the Recall Return Response Acknowledgement and Receipt Form via mail to: 16 Thomas J Rhodes Industrial Drive, Hamilton NJ, 08619; fax: (609) 587-5636 or email to: www.medicalindicators.com or General Manager at: pbaker@medicalindicators.com.Contact information for any question(s) related to this recall may be addressed to General Manager of Medical Indicators, Inc. at 609-737-1600, extension 204 or email to: pbaker@medicalindicators.com, Mondays through Fridays from 8:00am to 5:00pm EST. |  | 111800 | 111,800 boxes (5,590 full boxes) | Nationwide distribution: USA including states of: CA, GA, FL, IN, NY, NV OH, PA, TN,TX, VA, and WI. | Terminated | January 16, 2014 | 190 | 111800 | Not_Computer
Z-2161-2013 | 65970 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Tec 6 continuous Flow Vaporizer and Tec 6 Plus Anesthesia Vaporizer. The Datex-Ohmeda Tec 6 Plus and Tec 6 Plus NAD Variant is an electronic vaporizer which delivers the anesthetic agent desflurane. The Tec 6 Plus attaches only to Datex-Ohmeda anesthesia machines equipped with the proprietary Selectatec manifold. The Tec 6 Plus NAD Variant attaches only to NAD anesthesia machines equipped with the NAD exclusion type manifold. |  |  | July 29, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the pre-operative check, low pressure leak test at 1% dial setting of the Tec 6 and Tec 6 Plus Vaporizers. The user manual pre-operative check, low pressure leak test may not detect the full range of leaks from seal wear degradation in the vaporizers. If the leak is not detected, it will result in fresh gas leaking to | Component change control | GE Healthcare sent an Urgent Medical Device Correction letter dated August 1, 2013, to all affected customers. The letter described the issue, asked that the Tec 6 and Tec 6 Plus Vaporizers user manual be updated with the attached addendum. Also Users are to perform the user manual recommended pre-operative check which includes the low pressure leak test on every vaporizer with the dial turned to the 12% setting instead of the specified 1% setting as per the attached addendum. If the device does not pass the new low pressure leak at 12%, it should be removed from service and consignee should ontact a GE Healthcare Representative for a repair/replacement. If the vaporizer passes the new low pressure leak test at the 12 % setting, it can continue to be used, however, as detailed in the user manual, continue to routinely perform the pre-operative check which includes the lowpressure leak test.A Medical Device Correction Confirmation form was asked to be sent back to GE.Customers with questions or concerns were instructed to call 1-855-700-4350 or send an email to vaporizer.admin@ge.com. For countries outside the US, customers were instructed to contact their local GE Healthcare Service Representative.For questions regarding this recall call 262-513-4122. |  | 31443 | 31,443 devices (23,947 US, 7496 OUS) | Worldwide Distribution - USA (nationwide) Puerto Rico and internationally to AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COSTA RICA, CROATIA local name-Hrvatska, CYPRUS, DENMARK, ECUADOR, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY, JORDAN, KOREA-REPUBLIC OF, KUWAIT, LITHUANIA, MALAYSIA, MALTA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, POLAND, PORTUGAL, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, and VENEZUELA, | Terminated | April 17, 2015 | 627 | 31443 | Not_Computer
Z-2164-2013 | 64146 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering |  |  | October 31, 2012 | 2012 | B Braun Medical, Inc. | This notice is being provided in follow up to the field correction initiated by B. Braun Medical Inc. on 26 August, 2011 due to an issue in the Outlook ES Safety Infusion System where the Care Area can inadvertently be exited thereby removing DoseGuard dose limit protection. This occurs when a sequence of HOLD - HOLD key strokes are performed on the pump key panel while the user is in the Dose Dat | Software design | The firm decided to conduct a field correction to correct the issue and carried it out through routine field services. The firm was notified by FDA that this is a reportable event, so they notified their consignees via letters dated 10/31/2012. |  | 1970 | 1,970 pumps added for this expansion | Nationwide Distribution. | Terminated | February 03, 2016 | 1190 | 1970 | Not_Computer
Z-2166-2013 | 65993 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook Safety Infusion System Burette Set. |  |  | August 05, 2013 | 2013 | B. Braun Medical, Inc. | Reports were received of an incorrectly assembled Outlook Safety Infusion System Burette Set (Product Code 375038), and the investigation determined that one of the ULTRASITE injection sites and corresponding low pressure backcheck valve was assembled upside down. | Process control | On August 4, 2013, a letter was sent to direct accounts asking them to check their inventory for the affected product. These customers were also instructed to notify their customers if the product was further distributed and to return all affected product and/or the product removal acknowledgement form. |  | 900 | 900 | Distributed in the states of FL, NY, NJ, and TX. | Terminated | September 03, 2014 | 394 | 900 | Not_Computer
Z-2165-2013 | 66082 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxa Vial Adapter, 14mm - Luer Lock, Order No. REF: 90, Quantity: 50 each, Single Use. Product Code H93890. |  |  | August 22, 2013 | 2013 | Baxa Corporation | Baxter Healthcare Corporation is voluntarily recalling one lot of Baxa Vial Adapters, 14mm - Luer Lock, due to incorrect expiry date. | Labeling Change Control | An Urgent Product Recall notification was mailed to one customer via USPS First Class mail on 8/22/13. The one customer affected by this action was instructed to locate and remove the affected product from inventory and to return to Baxter for credit and/or replacement. Any customer returns received will be discarded per the Corporate Hold. |  | 34 | 34 cases | Distributed to One Customer in NC. | Terminated |  |  | 34 | Not_Computer
Z-2163-2013 | 64147 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering |  |  | October 24, 2012 | 2012 | B Braun Medical, Inc. | B. Braun Medical Inc. has become aware of an issue with the Outlook ES Safety Infusion System. The key panel may fail to function as intended if the underlying circuitry fails due to intermittent contact. This issue applies specifically to the HOLD button. If this occurs, the Door Processor (DP) will not receive the desired input from the user. This issue can be observed in pumps with key panels | Component design/selection | B Braun sent an URGENT: MEDICAL DEVICE FIELD CORRECTION letter dated October 24, 2012 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact Customer Support at 1-800-627-7867 for questions regarding this notice. |  | 28427 | 28,427 | Nationwide Distribution. | Terminated | July 27, 2015 | 1006 | 28427 | Not_Computer
Z-2168-2013 | 65960 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | OPERON D850 Surgical Table |  |  | April 16, 2012 | 2012 | Berchtold Corp. | Table was not lowering. | Nonconforming Material/Component | Berchtold sent a Customer Notification letter dated April 16, 2012 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that BERCHTOLD will upgrade the OPERON Surgical tables with a new firmware version that will eliminate the potential for experiencing problems. For questions call 843-569-6100 ext. 322. |  | 263 | 263 | USA Nationwide Distribution in the states of AL, AZ, CA, CO, CT, FL, GA, HI.IL, IA, KY, LA, MD, MA, MI, MN, MS, NH, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WI and including Canada. | Terminated | September 10, 2013 | 512 | 263 | Not_Computer
Z-2167-2013 | 64965 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Xlumena brand NAVIX Access Device, Model/Catalogue Number: NVX-10-03, Part Number: FG-00666, Product is manufactured and distributed by Xlumena, Inc., Mountain View, CA |  |  | April 17, 2013 | 2013 | Xlumena, Inc. | A complaint investigation has found that product fractures can occur at the distal end of the catheter under load. | Process design | Xlumena sent an Urgent Medical Device Recall letter dated April 17, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and segregate the affected product. Customers with questions were instructed to contact dhasker@xlumena.com. For question regarding this recall call 650-961-9901. |  | 148 | 148 units | Worldwide Distribution - USA including AZ, CA, FL, GA, IN, NY,SC, WA, PA, IL, CO, and internationally to Netherlands, Germany, Spain, Denmark, and Switzerland. | Terminated | September 23, 2013 | 159 | 148 | Not_Computer
Z-2177-2013 | 64601 | No | FSO | Surgical microscope and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Surgical microscope |  |  | December 18, 2012 | 2012 | Leica Microsystems (Schweiz) Ag | The brake in the Z axis and/or brakes in the AB and C axes failed during installation. | Device Design | The firm, Leica Microsystems, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 19, 2012 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to DISCONTINUE USE until your system software is updated and to complete and return the attached Acknowledgement Form via email at LMGRA@leica-microsystems.com or fax to: 847-236-3747 or +41 71 726 32 49 within 10 working days after receipt of field safety notice. Your Leica Mycrosystems will contact you to make the necessary arrangements to complete an upgrade at no charge to the customers. If you have any further questions, contact Regulatory Affairs/Quality Assurance Manager at 847-405-5413 or +41 71 726 32 16. |  | 28 | 28 units | Worldwide distribution: USA (Nationwide) including states of: AK, CA, CO, NY, and OR; and countries of: Belarus, China, Hong Kong, India, Indonesia, Italy, Norway, South Africa, Spain, Sweden, Turkey, and Vietnam. | Open, Classified |  |  | 33 | Computer
Z-2178-2013 | 65978 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Monoject 0.9% Sodium Chloride Flush Syringe, 12 mL Syringe with 10 mL Fill |  |  | August 16, 2013 | 2013 | Covidien LLC | Monoject prefill flush syringes may contain non-sterile water and labeled as saline or heparin | Process control | Covidien sent an Urgent Medical Device Recall letter dated August 16, 2013 via Federal Express to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter requested that customers immediately quarantine and discontinue use of the affected products, return any affected product by following the instructions provided in the letter, and complete and return the Recalled Product Return Form to Covidien. The letter states that if the affected products have been distributed, assure that all customers are notified and affected product is removed from the distribution channel as well as from customer stock. Questions or concerns, please do not hesitate to contact your Covidien representative or Covidien Customer Service, Monday through Friday, 8am- 6:30pm ET, at (800) 962-9888, option 1, and then option 2.Firm issued Press on 8/19/13 |  | 3235860 | 3,235,860 units | Worldwide Distribution - USA Nationwide and the country ofBermuda | Terminated | October 21, 2015 | 796 | 3565980 | Not_Computer
Z-2174-2013 | 65957 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100NX Cassettes, Product Code: 10144 |  |  | July 25, 2013 | 2013 | Advanced Sterilization Products | The recall of the STERRAD 100NX Cassettes was initiated because Advanced Sterilization Products (ASP) does not have adequate data to support the effectiveness of the packaging's leak indicator for the entire duration of the STERRAD 100 NX System cassette's 15 month shelf life. | Under Investigation by firm | ASP sent an Urgent Medical Device Field Safety Notification letter dated July 25, 2013, to all affected customers. The recall letter informed the customers of the problem identified and the actions to be taken. Customers were instructed to contact Stericycle directly at (866) 562-5937 to report any complaints or suspected problems with their STERRAD 100NX System cassette. For questions regarding this recall call 949-453-6400. |  | 1923 | 1,923 units | Worldwide Distribution - USA (nationwide) and internationally to RU, VE, KW, SA, HU, GB, IT, HK, IL, MX, ZA, PH, GR, AE, IN, CO, KR, AU, IN, ES, SG, MY, TH, TR, TW, SI, BE, PL, EC, CN, JP, PA, DE, VN, CL, and Switzerland, United Kingdom. Israel. Slovenia, Qatar, Panama, Venezuela. | Terminated | November 27, 2013 | 125 | 1923 | Not_Computer
Z-2175-2013 | 64150 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-100ES, and 621-200ES. To deliver a broad range of drugs with use with the B. Braun Medical Inc. Horizon Pump IV Sets. |  |  | December 05, 2012 | 2012 | B Braun Medical, Inc. | B. Braun Medical Inc. is informing our customers of a field correction previously performed to address the loss of the drug library in the Outlook Pump Models 620-100 and 620-200. This could have occurred after periods of lost DC (battery) power or out of tolerance low power. This could have led to a delay in therapy, or to delivery of an over or under-dose of a therapy if the user was to rely on | Component design/selection | The firm, B. Braun Medical Inc., sent an "URGENT: MEDICAL DEVICE FIELD CORRECTION" letter dated December 05, 2012 to its customers. The letter describes the product, problem and actions taken. The firm finished performing "field upgrades" on March 20, 2010 for this issue.Should you require any additional information, please do not hesitate to contact B. Braun Medical Inc., Infusion Systems Customer Service at 1-800-627-7867 (627-PUMP). |  | 14161 | 14,161 | Worldwide Disribution: USA (nationwide) and country of: Canada. | Terminated | October 24, 2014 | 688 | 14161 | Not_Computer
Z-2173-2013 | 66036 | No | MAQ |  |  | Microbiology | PMA | KIT cobas 4800 HPV AMP/DET 240T / 960T US-IVD, CE-IVD |  |  | July 25, 2013 | 2013 | Roche Molecular Systems, Inc. | During an internal investigation, a formulation error that occurred was identified during the manufacture of the cobasÂż 4800 HPV Master Mix reagent that was packaged in four (4) cobasÂż 4800 AMP/DET kits, US- and CE-IVD (240 and 960 test kit configurations). Examination of the batch production record of the material revealed there was an error in the tare weight used for the carboy during the fina | Device Design | Roche Molecular Diagnostics sent an Urgent Medical Device Correction letter dated July 25, 2013, to all affected customers via UPS return receipt. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to:1. Immediately discontinue use of the product lots listed above.2. Discard affected product according to local guidelines.3. To receive replacement inventory, complete the entire attached faxback form and fax it to 1-866-918-8741.4. The faxback form must be completed and returned even if you do not have affected inventory.5. File this Urgent Medical Device Correction (UMDC) for future reference.Please contact Roche Molecular Diagnostics Technical Support, 24 hours a day, seven days a week at 1-800-526-1247, if you have questions about the information contained in this UMDC. |  | 791 | US-IVD: 350 Kits; CE-IVD: 441 KITS | Worldwide Distribution - USA (nationwide) including the countries of Austria, Australia, Belgium, Brazil, Czech Rep, Canada, Croatia, Columbia, Denmark, Dominican Republic, Germany, Ecuador, France, Japan, Greenland, Mexico, Italy, New Zealand, Netherlands, Peru, Norway, Singapore, Romania, South Africa, Spain, South Korea, Sweden, Switzerland and UK. | Terminated | August 06, 2015 | 742 | 791 | Not_Computer
Z-2172-2013 | 65806 | No | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Folate (2500 tests), REF SMN 10340209, Cat No. 09132781 |  |  | July 25, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | Calibration failures and/or significant negative shifts in quality control (QC) and patient results. | Material/Component Contamination | Siemens sent an Urgent Medical Device Correction letter on July 26, 2013, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. an Urgent Field Safety Notice was sent to all domestic customers via Federal Express. On July 25, 2013 this notice was sent to all OUS customers via hard copy, email, and/or fax as determined by each country's local regulations and procedures. These notices inform the customer that a portion of Folate ReadyPacks in kit lots ending in 218 are exhibiting calibration failures and/or significant negative shifts in quality control (QC) and patient results. If customers have an alternate lot of the ADVIA Centaur or ADVIA Centaur CP Folate assay, they are instructed that this lot should be used for future testing. However if no alternative lot is available, quality control (QC) results should be obtained for each individual ReadyPack as a QC or calibration is an indication of an affected pack. An affected pack must be replaced with a different Folate ReadyPack where a valid calibration and acceptable QC values are obtained. These notices should be reviewed by the facility's laboratory or medical director, though no look back is required.Customers were instructed that Fedex Return Receipts will be used to confirm that the customers received the UMDC. Customers with questions were instructed to contact their local Technical Support Provider or Distributor.For questions regarding this recall call 508-668-5000. |  | 44 | 44 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA,MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, and WY. Internationally to Canada and Mexico. | Terminated | June 12, 2015 | 687 | 6486 | Not_Computer
Z-2191-2013 | 66133 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Vision One Laser System Model GA-0025020, Serial No: 10050, 10051, 10053. |  |  | August 19, 2013 | 2013 | Lumenis, Inc. | Lumenis has initiated a recall on certain models of Vision One System due to a potential for unintended laser exposure to the user. | Software design | Customer's were notified via letter on 8/19/13. Service visits to consignees were scheduled with anticipated completion on 8/26/13 to replace the control board. Consignee monitoring was performed by use of return reply verification tracking cards returned by the Lumenis service engineer after completion of CPU board replacement. |  | 3 | 3 units | Distributed in the states of NJ, PA, and IL. | Terminated | December 19, 2013 | 122 | 3 | Not_Computer
Z-2187-2013 | 65968 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | StaXxÂż IB Convex Expandable Implant (25 x 12 x 9 mm), Intervertebral Fusion Device, Lumbar; Catalog Number 11 -1902. |  |  | May 14, 2013 | 2013 | Spine Wave, Inc. | A component Superior endplate manufactured from a different polymer than specified | Component change control | Spine Wave notified surgeons verbally on 5/14/13 and followed by written notification on June 6, 2013 of the incorrect polymer used for the endplate and the health risks. |  | 17 | 17 units | Distributed in the states of FL and OR. | Terminated | February 10, 2014 | 272 | 17 | Not_Computer
Z-2122-2013 | 66058 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | Microscan Synergies Plus Negative Urine Combo 2 |  |  | August 21, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Confirmed increase in false susceptible and false intermediate misreads for Imipenem and/or meropenum on Synergies plus Negative and rap ID/S plus Negative panels read on Walk Away System. | Process control | Siemens sent an Urgent Field Safety Notice letters dated August 2013 to all US customers via over night delivery, receipt requested. Letters for foreign customers are being translated and will be sent as soon as possible. The letter identified the affected product, reason for field correction, risk to health problem and actions to be taken by customer. Customers were instructed to suppress all reporting of sensitive and intermediate results for imipenem and meropenem. Siemens also recommends that the customer consider the need to review previous test results, conduct patient follow-up, and/or repeat testing. Customer were requested to complete the attached Field Correction Effectiveness Check form and fax it to 302-631-8467. For questions contact your Siemens Technical Support Provider or Distributor. |  | 27804 | 1,324 boxes, 26,480 panels | Worldwide Distribution - USA Nationwide in the states of: AK, SC, IA, MO GA, WI, TX, KS, CA, IL, CO, AZ, SD, IN, MD, FL, MN and the countries of:Japan, France, Algeria, Germany, Austria, Romania, Russian Federation, South Africa, Kazakhstan, Belgium, Great Britain, Turkey, Spain, Italy, Saudi Arabia, Qutar, Poland, Iraq, Unite Arab Emrites, Algeria, Oman, Greece, Portugal, and Ireland. | Terminated | May 15, 2014 | 267 | 200172 | Not_Computer
Z-2192-2013 | 66031 | No | FMK | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BD Microtainer Quikheel Lancet 1.0 mm incision Depth 2.5 mm Incision Length REF 368101 Rx Only Becton, Dickinson and Company, Franklin Lakes, NJ 07417 USA Made in USA. |  |  | July 10, 2013 | 2013 | Becton Dickinson &amp; Company | The Lancet will not retract after activation thus creating a potential for exposure to a used sharp. | Component change control | BD notified their customers and distributors by phone, e-mail,and/or letter/notice of return/ customer packing list on 7/10/2013. |  | 58200 | US - 39,800 units ; Canada - 18,400 units | Distributed in the states of CT, KY, IL, NY, PA, TX, and VA and the country of Canada. | Terminated | March 19, 2014 | 252 | 58200 | Not_Computer
Z-2193-2013 | 66123 | No | GCB | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | *** 1) Cat. 4341B; Qty. 10; Thoracentesis Tray; With Catheter; Sterile; CareFusion; 1500 Waukegan Road; McGaw Park, IL 60085 USA; |  |  | August 28, 2013 | 2013 | Carefusion 2200 Inc | Catheters included in Thoracentesis Tray and Safe-T Thoracentesis Tray - Drug Free may occlude, preventing drainage of excessive fluid from the pleural cavity leading to a potential delay in the medical procedure and potentially serious patient complications. | Process change control | Recall notification letters dated 8/27/13 were distributed via courier to direct accounts. Consignees were/are requested to confirm receipt of the notification, examine their inventory for the recalled products and to destroy the products. After the products are destroyed, the consignees are requested to submit a completed response form to CareFusion for replacement with functional devices. Consignees who cannot destroy the affected products are requested to contact CareFusion for additional information. |  | 197913 | *** 1) 174,580 units; *** 2) 23,330 units | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MD, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, and WV and Puerto Rico, and the country of Iceland. | Terminated | June 16, 2015 | 657 | 197913 | Not_Computer
Z-2189-2013 | 65984 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | "***Flo Trac Sensor Model MHD85***" |  |  | July 31, 2013 | 2013 | Edwards Lifesciences, LLC | The Edwards Lifesciences FloTrac Sensor is being recalled due to a manufacturing non-conformance related to incorrect packaging configuration where the IFU was placed on top of the Tyvek instead of being placed at the bottom of the box. There is no evidence available at this time that could validate the sterility of this incorrect packaging configuration. | Packaging process control | Edwards Lifesciences initiated a US recall via an Urgent Product Recall letter dated July 31, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that an Edweards Lifesciences representative would come to their facility to retrieve the affected device. A device replacement would be provided to the customer. Customers were asked to complete the enclosed Confirmation Form and give it to the Edwards Lifesciences representative that will come to their facility. For customers with questions were instruct4ed to call Edwards Customers Service at 1-800-424-3278.For questions regarding this recall call 972-410-7100. |  | 35 | 35 units | Nationwide Distribution including MO and MI. | Terminated | April 14, 2017 | 1353 | 35 | Not_Computer
Z-2190-2013 | 65854 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | HemoCueÂż Glucose 201 Microcuvettes, 25 Microcuvettes Individually Packaged. |  |  | July 15, 2013 | 2013 | HemoCue AB | During a complaint investigation HemoCue AB discovered punctured HemoCue Glucose 201 single pack pouches. | Packaging | HemoCue sent an Urgent Field Safety Notice dated July 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine any affected product and Return the Field Safety Notice Verification Form to e-mail: 2013glufsn@hemocue.com or fax 562-668-5794 Customers would be contacted to coordinate return and replacement of the affected microcuvettes. Customers with questions were instructed to call 800-426-7256.For questions regarding this recall call 440-925-3485. |  | 2221500 | 22,215 boxes of 100 cuvettes distributed in US, 34, 841 distributed outside the US | Worldwide Distribution - USA (nationwide) and internationally to AUSTRALIA, BELGIUM, CANADA, COTE D'IVOIRE, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, IRELAND, ITALY, KENYA, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, QATAR, RUSSIAN FEDERATION, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM, VIET NAM &amp; YEMEN.. | Terminated | September 02, 2014 | 414 | 2221500 | Not_Computer
Z-1993-2013 | 66046 | No | MNO |  |  | Cardiovascular | PMA | Cardiogenesis Corporation - Cardiogenesis Transmyocardial Laser. |  |  | May 01, 2011 | 2011 | Cardiogenesis Corporation | Following acquisition of Cardiogenesis, we began receiving complaints of sparks, burns, or smoke incidents related to the handpieces. | Device Design | Cardiogenesis is responsible for the correction of all noncompliant products. CDRH understands that the corrective action plan (CAP), involving 2,437 Sologrip III handpieces and 61 PEARL 5.0 handpieces, consists of the following:1. Notification to purchasers of the handpiece removal and replacement procedures,2. Replace existing hospital inventory of handpieces with a redesigned handpieceCDRH approves the CAP subject to the following conditions:1. Notification of all distributors and purchasers will be successfully implemented, and2. Corrections will be made at no cost to the purchasers by October 1, 2013.For further questions please call (949) 420-1800. |  | 2498 | 2498 | US Distribution | Terminated | September 15, 2014 | 1233 | 2498 | Not_Computer
Z-2195-2013 | 66110 | No | DYJ | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | AOA Mini RPE Screw Assembly , a component of |  |  | June 06, 2013 | 2013 | Sybron Dental Specialties | The firm initiated the recall of AOA Mini RPE Screw Assembly, because it may be missing a complete weld joining one side of the arm with the screw body. Herbst, MARA and Fixed Metal devices may be affected, because they contain an AOA Mini RPE Screw Assembly. | Device Design | AOA sent a recall notification letter titled "Urgent-Medical Device Correction", on June 6, 2013 via USPS 1st class mail. The letter identified the affected product, reason for recall, images, and actions to be taken. Customers were instructed to complete and return the attached Acknowledgement Form to AOA. For questions contact the AOA Customer Support Department at 1-800-262-2551. |  | 1091 | 1,091 | Worldwide Distribution - USA Nationwide and the countries of Australia, Canada, and EU. | Terminated | September 13, 2013 | 99 | 1091 | Not_Computer
Z-2194-2013 | 66173 | No | OAT | Endosseous dental implant. | Dental | Dental | 510(k) | Brand Name: Vector TAS 8mm Screw. |  |  | July 12, 2013 | 2013 | Ormco/Sybronendo | Ormco Corporation is voluntarily recalling one lot of 8 MM Vector TAS Screws because some screws of this lot were machined undersized at the top of the threads may potentially cause a weakening of the screw. Use of the undersized product may result in breakage of the screw during insertion or removal when torqued, possibly requiring surgical removal of any remaining pieces. | Device Design | Ormco sent an "Urgent Medical Device Recall letter dated July 12, 2013, via USPS 1st class mail, along with a return form to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were informed of the recall, with product description, product codes, ship dates, contact information and instructions to return any affected product identified in stock, Additionally, doctors are instructed to proceedwith treatment if the screw has already been placed and to take special care not to over-torque the screw upon removal. Periodic effectiveness checks will be conducted in order to follow up with non-responders.Please fax this acknowledgement form to the the following number to confirm receipt of this notification whether or not you have any affected products. (714) 288-4603 |  | 4786 | 4,786 | Worldwide Distribution - USA (nationwide) including the countries of Canada, Australia, New Zealand, EU and ROW consignees. | Terminated | September 11, 2014 | 426 | 4786 | Not_Computer
Z-2160-2013 | 66024 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Juno DRF System. Mobile X-ray system |  |  | August 02, 2013 | 2013 | Villa Radiology Systems LLC | It has been discovered that the Air Kerma rate information is inconsistent in Continuous Fluoro modes 30x30 (12"x12") and 15x15 (6"x6). This could result in a deviation of the displayed values of Air Kerma Rate and Cumulated Air Kerma from the actual values greater then +/-35% and then in a failure to comply with the requirements of 21 CFR 1020.32(k)(6 | Radiation Control for Health and Safety Act | The Center for Devices and Radiological Health (CDRH) understands that the Villa Sistemi Medicali (U.S. Agent: Villa Radiology Systems) is responsible for the correction of all 22 noncompliant units distributed within the U.S. Should the CAP prove ineffective in bringing the noncompliant unit(s) in the field into compliance with the applicable performance standard, the CDRH can require more stringent measures to obtain compliance.The system will be modified by upgrading the generator power firmware, which will bring Juno DRF x-ray system into compliance. You provided a verification and validation test report (Juno DRF System AK fix-DAM Prom V&V; Test Report) to FDA demonstrating that these proposed upgrades addressed the problem. Field service engineers will schedule an appointment with customers to install the new software.2. The customer notification letter (URGENT: Electronic Product Radiation Warning), dated August 02, 2013, includes the statement that "the manufacturer will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services, the details of which will be included in a subsequent communication to you."3. The CAP will be completed within a six month timeframe (by March 1, 2014).4. The manufacturer will notify the Secretary on monthly basis providing progress with the field correction.CDRH approves the CAP subject to the following conditions:Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22.For further questions please call (203) 262-8836. |  | 22 | 22 units within US | US Distribution including the states of AZ, MI, IN, UT, CT, AL, NY, PA, IL, MD, CA, MA, IA, OH, MS and MO. | Completed |  |  | 22 | Computer
Z-2196-2013 | 65980 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | DiaScreen Liquid Urine Controls: Positive and Negative Assayed Liquid Quality Controls for use with Dipsticks in Routine Urinalysis; Product Number D76004 (4x12 mL bottles) consists of: 2x12mL Positive controls [and] 2x12mL Negative controls; Manufactured for ARKRAY USA, Inc.; Minneapolis, MN 55439 |  |  | July 16, 2013 | 2013 | Assuramed | The products may have been stored at temperatures outside their required storage conditions. | Storage | Independence Medical (im) sent an Urgent Product Recall letter dated July 16, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to identify and immediately cease the distribution of the recalled products, quarantine the recalled products, and contact the recalling firm at 800-860-8027 for instructions on how to return the recalled products. Consignees who further distributed the recalled products were requested to notify their customers of the recall. Consignees were also instructed to complete the Product Recall Acknowledgement Form and return the document by email to recall@indemed.com indicating receipt of the notification.Customers with questions were instructed to contact the product manufacturer for ProTime Disposable Cuvettes at 800-579-2255, DiaScreen Liquid Urine Controls at 800-818-8877, and Hemoccult II SENSA elite Dispensapak at 800-854-3633.For questions regarding this recall call 800-860-8027. |  | 10 | 10 units | Nationwide Distribution including AR, AZ, CA, CO, FL, IL, KY, LA, MI, NJ, NY, OH,OK, OR, SC, TX, UT, and WA. | Open, Classified |  |  | 154 | Not_Computer
Z-2188-2013 | 64152 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Outlook ES pump is intended for use with B. Braun Medical Inc. Horizon Pump IV Sets. Used to regulate the flow of primary and secondary fluids when positive pressure is required. The infusion system is capable of delivering fluid from a negative head height (when the IV fluid container is lower than the pump), and provides clinically accepted volumetric accuracy for all standard IV fluids, including blood, lipid, and Total Parenteral Nutrition (TPN). Model number: 621-300ES and 621-400ES. To deliver a broad range of drugs with use with the B. Braun Medical Inc. Horizon Pump IV Sets. |  |  | November 14, 2012 | 2012 | B Braun Medical, Inc. | This notice is being provided in follow up to the field correction initiated by B. Braun Medical Inc. on August 26, 2011 due to an issue in which the Outlook ES Safety Infusion System may halt infusion but the "RUN" Light Emitting Diodes (LEDs) on the front display continue to advance as if the pump were infusing. The pump emits a backup alarm, but there are no visual indicators that the infusion | Software design | The firm, B Braun Medical Inc, sent an "Urgent Expanded Medical Device Field Correction" letter dated November 14, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that pumps that have been upgraded (software version 151599 or higher), no additional action is required on their part at this time. For the pumps that have not yet been upgraded, within three weeks of the letter, customers should confirm the serial numbers using the enclosed response form or contact B. Braun Medical, Inc. Infusion Systems Customers Support at 1-800-627-7867 (800-627-PUMP) to schedule their upgrade.For questions regarding this recall call 972-245-2243. |  | 6919 | 6,919 | Nationwide distribution: USA including states of: AR, CA, FL, IL, LA, NC, NJ, NY, MD, ME, MN, SC, TN, TX, VA, WI, and WV. | Terminated | March 06, 2015 | 842 | 6919 | Computer
Z-2199-2013 | 66057 | No | HTZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | AEQUALIS Reversed II or AEQUALIS Reversed Fracture Impactor. These products are contained within the Aequalis reversed II Instrument set tray no.YKAD83 and the Aequalis reversed fracture instrument set tray no.YKAD95. The items affected are MWD21, MWD023, MWD024 and MWD025. |  |  | August 13, 2013 | 2013 | Tornier, Inc | Following the identification of the disassociation of an impactor tip during surgery, Tornier is initiating a voluntary recall of all lots of Aequalis Reversed II and Reversed Fracture impactors. Upon review of the event, Tornier determined that a redesign of the instrument was warranted. These products are contained within the Aequalis Revered II instrument set tray no YKAD83 and the Aequalis R | Device Design | Tornier sent an "Urgent Medical Device Recall" letter dated August 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customer were advised to cease further use and distribution of the product and to contact all end users. Also requested the completion and return of the attached questionnaire via fax to 952-236-4007, e-mail to CustomerServices@Tonier.com or using the envelope provided.We apologize for the incovenience and thank you in advance for your cooperation in this matter. If you have any questions please contact our Customer Service Department at 1-888 494-7950. |  | 2298 | 2,298 (1045 US, 1253 OUS) | .Worldwide Distribution - USA including Puerto Rico and the states of AL, AZ, CA, CO,CT, FL, GA, IL, KY, MD, MA, MI, MN, NE, NJ, NC, OH, OR, PA and TX., and the countries of CANADA, MEXICO, HONG KONG, SOUTH AFRICA, ARGENTINA, COLOMBIA, ISRAEL, AUSTRALIA, BELGIUM, FINLAND,FRANCE, GERMANY, ITALY, NETHERLANDS, SLOVENIA, SPAIN, SWITZERLAND and UNITED KINGDOM | Terminated | March 03, 2015 | 567 | 2298 | Not_Computer
Z-2200-2013 | 66010 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | T8 REMOVAL HEAD, REF 8910641, QTY: 1 EA, |  |  | August 12, 2013 | 2013 | Medtronic Sofamor Danek USA Inc | One lot of the T8 Removal Head had incorrect hardness specification, which could result in the T8 Removal head failing during removal of the CD Horizon Spire and Spire Z Spinal System implant. There is a potential for tissue damage to occur and/or minor surgical delay to remove the remaining fragments of the tip. | Nonconforming Material/Component | Medtronic sent an Urgent Voluntary Market Removal Recall letter dated August 9, 2013 with Recall Questionnaires and Confirmation forms included to Risk Managers and Sales Representatives. The letter identified the affected product, problem and action to be taken. Customers were instructed to remove the affected product, discontinue use and contact your Medtronic Sales Representative to coordinate the return and replenishment of affected product. For questions call your Medtronic Sales Representative or Global Quality Department at 800-876-3133. |  | 9 | 9 units | USA Nationwide Distribution in the state of GA, AK, WA, and CA. | Terminated | December 27, 2013 | 137 | 9 | Not_Computer
Z-2205-2013 | 66143 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | C2244.4308 CONELOG Titanium Base CC Âż4.3GH2 0.8 inkl HS |  |  | August 28, 2013 | 2013 | Camlog Usa | Product not cleared for sale/use in the United States. | No Marketing Application | Camlog Usa notified four US consignees of the recall by telephone on August 5, 2013. The notification identified the recalled product, which did not have proper marketing clearance for distribution in the United States and asked customers to return all unused product. Cusotmers were sent an overnight UPS label via e-mail and asked customers to return the device to CAM:LOG USA Attn: Debby Callahan, 1822 Aston Avenue Carlsbad, CA 92007. |  | 6 | 6 to US customers | USA Distribution in the states of NH, MD, CO, and IL. | Terminated | January 29, 2014 | 154 | 12 | Not_Computer
Z-2209-2013 | 66155 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Blanketrol III, Model Number 233, Cincinnati Sub-Zero Products, Inc., 12011 Mosteller Road, Cincinnati, OH 45241 |  |  | September 03, 2013 | 2013 | Cincinnati Sub-Zero Products Inc | During a label review it was discovered that both the Blanketrol III Operations (57201-N) &amp; Operations &amp; Technical (56201-Q) manuals had verbiage added back into the manual during ECN M111-4260 (released 12/29/2011) which was previously removed for clarification of the intended use and subsequent 501(K) approval. | Labeling Change Control | Cincinnati Sub-Zero (CSZ) sent an Urgent Medical Device Field Action letter dated September 5, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact CSZ at 513-772-8810 or 1-800-989-7373 to request new versions of the manuals. At that time customers are instructed to provide the quantity and serial numbers of the affected devices that are at their facility. Upon receipt of the new versions of the manual, immediately collect and destroy/discard all the affectred devices Operation and Technical Manuals their facility is currentlyin possession of and replace the discarded manuals with the new versions of the manuals provided by CSZ. After the manuals have been discarded and replaced, customers were asked to complete and return the enclosed response form as soon as possible to acknowledge receipt of this notification and to inform CSZ that they have performed and completed the requested actions. Customers should return the form by fax to 513-772-9119 scan and email the information to B3manual@cszinc.com or mail to Cincinnati Sub-Zero, 12011 Mosteller Road, Cincinnati, Ohio 45241. Customers with questions were instructed to contact their Blanketrol III Model 233 field action coordinator at 1-800-989-7373 or 513-772-8810.For questions regarding this recall call 513-772-8810, ext 5295. |  | 1126 | 115V - 396 units; 230V - 385 units | Worldwide Distribution - USA including AK, AR, AZ, CA, CO, DC, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MT, , NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, TN, TX, UT, VA, WA, WI &amp; WV. Internationally to ARGENTINA, AUSTRALIA, AZERBAIJAN, BELGIUM, CANADA, CHILE, CHINA, COSTA RICA, CROATIA, GERMANY, HONDURAS, HONG KONG, INDIA, INDONESIA, ISRAEL, KOREA, KUWAIT, MEXICO, NEW ZEALAND, PERU, SAUDI ARABIA, SINGAPORE, SPAIN, TAHITI, THAILAND, THE NETHERLANDS, THE PHILIPPINES, UNITED ARAB EMERITES &amp; UNITED KINGDOM. | Terminated | August 22, 2014 | 353 | 1126 | Not_Computer
Z-2202-2013 | 66027 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Sonogage EyeScan, A-Scan Biometer, |  |  | August 09, 2013 | 2013 | Sonogage Inc | During an FDA inspection on 7/15-26/2013, it was discovered that the firm lacked registration and 510(k) clearance for the device. | No Marketing Application | Sonogage sent an Urgent: Medical Device Recall letter dated August 9, 2013 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately examine inventory and quarantine the affected product for return. Customers were asked to complete and return the enclosed response form. For question call 216-464-1119 or 800-798-1119. |  | 564 | 564 units | Worldwide Distribution - USA Nationwide in the following states: AK, AZ, CA, FL, GA, IL, KY, LA, MA, MD, MN, NE, NJ, NM, NY, OH, OR, PA, TX, VA and WA. Product was also shipped to Canada and Mexico. | Terminated | September 19, 2015 | 771 | 564 | Not_Computer
Z-2207-2013 | 66204 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Osseotite 2 Certain Prevail 5mm(D)X 4.1mm (P) X X 11.5mm (L) Dental Implant |  |  | June 10, 2013 | 2013 | Biomet 3i, LLC | One of the dental implants do not have the internal hex feature. One of the dental implants do not have the internal thread feature. | Device Design | Biomet 3i sent an Urgent: Medical Device Recall letter dated June 10, 2013 to affected customer via email, facsimile, or postal mail. The letter identified the affected product, problem, and actions to be taking. Customer were instructed to check their respective inventory for the affected product and return to Biomet 3i. Biomet 3i will issue replacement product. For questions or concerns call 1-800-342-5454. |  | 215 | 215 | Worldwide Distribution - USA Nationwide in the states of MA, CT, NJ, DE, NY, FL, OK, TX, PA, DC, IL, KS, LA, MD, CA, RI, TN, MN, and SC and the countries of Canada, Japan, Australia, Italy Germany, France, Switzerland, Apain, Brazil, Uruguay, Singapore, Dominican Republic, Taiwan, Austria, Portugal, and the United Kingdom. | Terminated | November 06, 2014 | 514 | 393 | Not_Computer
Z-2210-2013 | 65691 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 6, M/N 728256; Brilliance CT 16 M/N 728246; Brilliance CT 40, M/N 728235 &amp; Brilliance CT 64 M/N 728231, Philips Healthcare, Cleveland, OH 44143. |  |  | July 05, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The patient support may move in an unintended manner if the footswitch cover impinges on the footswitch. | Component design/selection | Philips Healthcare sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated June 4, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Contact the Customer Care Center at 1-800-722-9377, option 5 for information concerning this issue. |  | 2913024 | M/N: 728256, 4 units; M/N: 728246, 89 units; M/N: 728235, 9 units &amp; M/N: 728231, 207 units | Worldwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, DE, FL, GA, IN, KY, MA, MD, MI, MN, MS, NC, NM, NV, NY, OH, OK, PA, SC, TN, TX, VT, WI &amp; WV, and the countries of Australia, Austria, Bosnia &amp; Herzegovina, Chile, China, Croatia, Denmark, Equatorial Guinea, France, Germany, Greenland, Iceland, Indonesia, Israel, Japan, Kazakhstan, Netherlands, Nigeria, Norway, Poland, Romania, South Africa, South Korea, Sweden, Switzerland, Taiwan, United Arab Emirates, United Kingdom &amp; Vietnam. | Terminated | September 15, 2015 | 802 | 2913024 | Not_Computer
Z-2201-2013 | 66156 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B850. |  |  | July 20, 2011 | 2011 | GE Healthcare, LLC | GE Healthcare has recently become aware of an issue associated with the CARESCAPE Monitor B850 when used with the iCentral network. The iCentral must have the alarm management software license (L-NETM-00) and the alarm limit change option enabled to experience this issue. When changing the HR, SpO2 or Invasive pressure alarm limits from the iCentral over the network, the alarm priority on th | Software design | Consignees were sent on 7/20/2011 a GE Healthcare "Important Product Information" letter. The letter was addressed to Healthcare Administrator / Risk Manager, chief of Nursing and Director of Biomedical Engineering. The letter described the Issue, Affected Product Details, Instructions, Product Correction, and Contact Information (Technical Support 1-800-558-7044). |  | 636 | 636 (53 US, 583 OUS) | Worldwide Distribution-USA (nationwide) including the states of CA, CO, CT, HI, IL , IA, KY, LA, MA, MI, NE, NY,OH, OK, TX, and WI, and the countries of UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, THAILAND, SYRIAN ARAB REPUBLIC, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FEDERATION, REPUBLIC OF SERBIA, PORTUGAL, POLAND, PARAGUAY, NORWAY, NETHERLANDS, NEPAL, MEXICO, MALAYSIA, KUWAIT, REPUBLIC OF KOREA, JORDAN, JAPAN, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, HONG KONG, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, DENMARK, CZECH REPUBLIC, COSTA RICA, COLOMBIA, CHINA, CHILE, CANADA, BOSNIA AND HERZEGOVINA, BELGIUM, BAHRAIN, AUSTRIA, and AUSTRALIA. | Terminated | October 25, 2013 | 828 | 636 | Computer
Z-2215-2013 | 66037 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | BIORAPTOR 2.9 mm Suture Anchor with 2 ULTRABRAID Suture; Part Number: 72200775 |  |  | August 06, 2013 | 2013 | Smith &amp; Nephew, Inc. Endoscopy Division | Pinholes detected in a small number of primary pouches that may constitute a breach of the sterile barrier. | Packaging | Smith&Nephew; sent an Urgent Product Recall Notification letter dated August 6, 2013, via Federal Express to all affected customers. Addendum customer letters for 4 additional part numbers were sent via Federal Express on August 19, 2013, The letters identified the product, the problem, and the action to be taken by the customer. Customers were asked to inspect their inventory, locate unused devices and quarantine them. Customers were also asked to complete the last three columns in the inventory Return Certification form indicating the quantities that need to be returned, and include their contact information in the spaces provided. Customers were instructed to contact Smith &amp; Nephew's Demo Service Departmentat 800-343-5717 for a Return Authorization (RA#), and provide the information response form. Replacement product would be provided. For questions regarding this recall call 978-749-1073. |  | 43385 | 43385 US | Worldwide Distribution - USA (nationwide) Canada, Europe: Austria, Belgium, Cyprus, Czech Republic, Denmark, Germany, Greece, Finland, France, Italy, Ireland, Luxembourg, Netherlands, Norway, Poland, Portugal, Russia, Slovenia, Spain, Sweden, Switzerland, Turkey, Ukraine and United Kingdom, Asia-pacific: Australia, China, Dubai, Hong Kong, India, Japan, Malaysia, New Zealand, Philippines, Sri Lanka, Singapore, South Korea, Taiwan, Thailand, Argentina, Bermuda, Brazil, Chile, Colombia, Costa Rica, Israel, Mexico, Puerto Rico, South Africa, and Venezuela. | Terminated | April 20, 2016 | 988 | 224222 | Not_Computer
Z-2214-2013 | 66098 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Cypher MIS Screw System, Biomet Spine Parsippany, NJ 07054. The Cypher MIS Screw System is a non-cervical spinal fixation device. |  |  | August 13, 2013 | 2013 | Ebi, Llc | Several complaints were reported to exhibit screw head splay in the Cypher MIS Screw System. | Device Design | Biomet Spine &amp; Bone Healing Technologies sent "Urgent Medical Device Recall Notice" letters with Fax-Back Response Forms dated 8/13/2013 to distributors and OR managers via Fed Ex. Follow-up e-mails with a copy of the Recall Notice were sent to distributors on 8/13/2013. Customers with questions may contact 973-299-9300 x2162 Monday to Friday from 8:00 am to 5:00 pm EST. |  | 427 | 427 units | Nationwide Distribution (US), including the states of IN, MD, MS, NC, OH, PA, SC, and TX. | Terminated | March 11, 2015 | 575 | 427 | Not_Computer
Z-2211-2013 | 66208 | No | FWP | Chin prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Device Name: Terino Square Chin-Style II-Medium; Catalog Number: TSCII-M. |  |  | August 30, 2013 | 2013 | Implantech Associates Inc | Implantech initiated a voluntary recall of certain lots of Terino Square Chin-Style II-Medicum (Lot # 848488 and 850156), after determining that the devices from the lot numbers do not conform to the dimensional requirements for the TSCII-M. | Nonconforming Material/Component | Recall was initiated by sending recall notification letters to customers on 08/30/2013 via a combination of 1st class mail, email, and telephone. The recall notification letter, dated August 30, 2013, informs customers of the reason for recall, product description and codes, instructions on what to do with the recalled product and contact information. Within the recall notification letter, a section is provided for responses and signatures. |  | 8 | 8 | Worldwide distribution: US (nationwide) including states of NY, KY, and CO, and the country of Israel. | Terminated | February 06, 2014 | 160 | 8 | Not_Computer
Z-2212-2013 | 66176 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Zeppessis' reprocessed single-use manifolds (filters) for the Stryker Neptune 1 and Neptune 2 Waste Management Systems: |  |  | August 16, 2013 | 2013 | Zeppessis Reprocessing, LLC | Zeppessis' reprocessed single-use manifolds (filters) for the Stryker Neptune 1 and Neptune 2 Waste Management Systems are recalled due to lacking 510K and and without a validated cleaning protocol for reprocessing the used manifolds. | Reprocessing Controls | Zeppessis Reprocessing, LLC sent an Urgent Medical Device Recall letter dated September 4, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are advised to stop using the Zeppessis reprocessed single-use manifolds (filters) for the Stryker Neptune 1 and Neptune 2 Waste Management Systems immediately and instructed to destroy the products in a proper manner.Customers should fill out the Customer Response form and fax or email it back to Zeppessis.For any questions about this recall, please call Zeppessis Reprocessing, LLC at 208-559-7064. |  |  | undetermined the distributed amounts | Distributed in the states of California, Idaho, Oregon, and Washington. | Terminated | November 08, 2013 | 84 |  | Not_Computer
Z-2213-2013 | 65657 | No | ILQ | Powered communication system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Responder 5 is a nurse call communication system that connects patients with staff in a hospital. The nurse call system also connects team members and departments from radiology, physical therapy, transport, environmental services and more. |  |  | May 10, 2013 | 2013 | Rauland-Borg Corp | A counterfeited electronic component was identified in the products affected. The component, a data transceiver (transmitter/receiver), is associated with the intercommunication of these devices in the Responder 5 system across the L-Net data path. Failure of a data transceiver results in the loss of nurse call functionality in the associated room, and may disrupt communications on the L-Net run | Counterfeit | Tech Bulletins sent a letter dated May 10, 2013, to all affected customers notifying them of the counterfeit parts and requesting that customers uninstall and return affected products to Rauland-Borg for replacement. A subsequent recall notification letter dated August 5, 2013 was mailed to customers. The letter included instructions to both direct accounts and end users on the process of replacing affected units with units manufactured with authentic parts.For additional information please call 847-590-5920. |  | 385 | Product Code 351000: 385 units; Product Codes 352000 and 352020: 14,808 units (total) | Worldwide distribution - US: (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KS, KY, LA, MA, MI, MN, MO, NC, NE, NV, NY, OH, OK, OR, PA, SD, TN, TX, UT, VA, WA, WI and WV., and the countries of : Australia, Canada, Columbia, Lebanon, Mexico and Saudi Arabia. | Terminated | May 06, 2015 | 726 | 385 | Not_Computer
Z-2241-2013 | 66111 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Advantage Workstation, READY View Application. |  |  | March 27, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has become aware that when analyzing fMRI images using the READY View application software, incorrectfMRI activation maps may be computed and displayed. | Software design | GE provided Urgent Medical Device Correction letters to consignees in the following countries: Canada, Finland, France, Germany, Guadeloupe, Italy, Japan, Latvia, Martinique, Netherlands, Norway, Poland, Reunion, Romania, Saudi Arabia, Singapore, Spain,Sweden, Switzerland, Turkey and the United Kingdom.The letters were dated March 27, 2012 and were addressed to Hospital Administrators, Radiology Mangers and Risk Managers. The letters described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (US 800-437-1171, Japan 0120-055-919 and for other countries contact your local GE Healthcare Service Representative).*Letters to other consignees is being planned for September 2013. |  | 318 | 318 (28 USm 290 OUS) | Worldwide Distribution-USA including the states of IL, MN, NJ, NY, and VA, and the countries of ALGERIA, AUSTRALIA, BELGIUM, BRAZIL, CANADA, DENMARK, FINLAND, FRANCE, GERMANY, GUADELOUPE, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN,KOREA, LATVIA, MARTINIQUE, NETHERLANDS, NORWAY, PHILIPPINES,S POLAND, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED KINGDOM, and VIETNAM. | Terminated | November 21, 2014 | 969 | 318 | Computer
Z-2252-2013 | 66060 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Central Dressing Change Kit; AMS-8464CP- 6. Vygon Central Dressing Change Kit contains: 1 Pair Nitrile Gloves W/Wallet, Powderfree 1 Face Mask 1 Polylined Drape, 13" X 18" 1 ChloraPrep***, 3ml Applicator W/Insert 1 Pkg. Alcohol Swabsticks (3/Pkg.) 1 14" Extension Set W/Bionector*** 2 Dressing Sponges, 2" X 2" 1 Foam Strip, 8/5cm X 1cm 1 3M*** Tegaderm *** IV Dressing W/CHG, 2 1/2" X 4 1/2" *** Vygon MFG Dover, NH 03820***" |  |  | August 12, 2013 | 2013 | Churchill Medical Systems, Inc. | Microbore tubing disconnecting from male luers on the extension sets. | Under Investigation by firm | "Urgent Medical Device Recall" letters were sent to each customer via courier service on August 12, 2013. Consignees were provided with a description of the issue and instructions on how to handle product that they received and that was further distributed by them. Questions were directed to customer service at 1-800-473-5414. |  | 3040 | 3040 | Nationwide Distribution (US), including the states of NY, MN, NJ, PA, FL, MD, and VA. | Terminated | June 16, 2014 | 308 | 16140 | Not_Computer
Z-2254-2013 | 66223 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | POWERWAND Safety Introducer with Extended Dwell Catheter; |  |  | August 16, 2013 | 2013 | Access Scientific Inc | Access Scientific is recalling the POWERWAND Saftey Introducer with Extended Dwell Catheter because it is mislabeled. | Labeling Change Control | On 8/16/13, Access Scientific visited 2 consignees and initiated a field correction. This is the communication given to the two(2) accounts by our field personnel: "Access Scientific has inadvertently shipped mislabeled devices to your facility. The products affected were 5 Fr. POWERWANDs, Catalog # 92006, lot # I666473."The consignees were instructed to "Quarantine the devices until Access Scientific personnel arrived with the appropriate labeling and performed the corrective action of applying the correct labels to the devices and inserting the DFU into the boxes." |  | 160 | 160 units | Distributed in MI and TX. | Terminated | September 24, 2013 | 39 | 160 | Not_Computer
Z-2242-2013 | 66052 | bandage, cast | ITG | Prosthetic and orthotic accessory. | Physical Medicine | Physical Medicine | 510(K) Exempt | 3M Scotchcast, Wet or Dry Cast Padding, Catalog Numbers WDP2, WDP3, WDP4, and WDP6. |  |  | August 05, 2013 | 2013 | 3M Company - Health Care Business | 3M is conducting a voluntary product recall of selected lots of 3M" Scotchcast" Wet or Dry Cast Padding. Upon investigation of a health care provider feedback, and confirmed with internal testing, it was found that the cast padding material does not consistently repel water. Lack of repellency may result in prolonged cast dry times after exposure to wet conditions. There have been no repo | Nonconforming Material/Component | 3M sent a Voluntary Product Recall Notification letter dated July 31, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to examine their inventory and set aside any affected lots subject to the recall notice. Customers were asked to inform all departments of any affected lots and provide them with a copy of the Notification letter that instructs them to contact their 3M distributor for alternative options. All affected products were to be returned to the 3M distributor they were purchased from. Customers with questions were instructed to contact the 3M Customer Helpline at 1-800-228-3957. For questions regarding this recall call 651-733-1110. |  | 24140 | 1207 cases (4 puches per case, 20 rolls per pouch) | Worldwide Distribution - USA including AL, AZ, AR, CA, CO,CT, FL, GA, HI, IL, IN,KS,KY, LA, MD, MA, MI, MN, MO, NE, NV, NJ, NY, NC, OH, OR,PA, TN, TX, UT,VA, WA, and WI. Internationally to GERMANY, CANADA, BELGIUM, NETHERLANDS, HONG KONG, NEW ZEALAND, AUSTRALIA, INDIA, KOREA, SINGAPORE AND UNITED ARAB EMIRATES, JAPAN, JAMAICA.. | Terminated | June 16, 2014 | 315 | 24140 | Not_Computer
Z-2248-2013 | 65893 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | ACL TOP 700 CTS; PN 0000280050 |  |  | July 19, 2013 | 2013 | Instrumentation Laboratory Co. | Potential carryover issue that can can cause shortened APTT clotting times on the instrument. | Device Design | Instrumentation Laboratory sent an Important Product Correction letter on July 19, 2013, to all affected customers. On July 23, 2013, the same communication was circulated by the IL International Group in Milan, Italy to the affected international distributors and dealers. The interim field notification was followed by the release of corrective Software V5.1.0 and ParametersP-16.3.00 (Phase II) on September 30,2013.Consignees were advised of a potential carryover issue and a workaround was provided. Consignees were asked to fax the Mandatory Response Tracking Form to 781-861-4207 or email to ra-usa.ilww.comFor technical questions customers were instructed to contact the IL Technical Support Center at 1-800-678-0710, option 2. For product availability questions, customers were instructed to contact their IL Customer Service Representative at 1-800-955-9525, Option #2.For questions regarding this recall call 781-861-4467. |  | 185 | 185 | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV. Internationally to Canada, Argentina, Australia, Austria, Belarus, Belgium, Brazil, Brunei, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Dubai, Finland, France, Germany, Great Britian, Greece, Guatemala, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Korea, Kuwait, Luthuania, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Uruguay, Venezuela, and Viet Nam. | Terminated | October 27, 2014 | 465 | 4799 | Computer
Z-2243-2013 | 64406 | No | HTY | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Cannulated Distal Femoral Nail (Trauma Nail), Part Number 450.851S, 420MM-Sterile. Orthppedic implant device. |  |  | April 18, 2013 | 2013 | Synthes USA HQ, Inc. | Firm discovered that a nail that was within the scope of a previous Trauma Nail recall had been distributed. | Packaging process control | A letter dated April 18, 2013 was sent to one consignee who was identified as receiving the affected device. A review of inventory and return of any affected product on hand was requested. Negative responses were also requested. |  | 1 | 1 | US distribution to MA. | Terminated | March 30, 2015 | 711 | 1 | Not_Computer
Z-2253-2013 | 66168 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery IGS 730 Cardiovascular X-Ray Imaging System. |  |  | June 19, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to Discovery IGS 730 Imaging Systems.A patient data management error may prevent the Discovery IGS 730 Imaging System from booting on start-up or re-booting during procedure. The boot issue may occur when a largeamount of patient data remains in the browser and is not deleted. The inability to boot the system ma | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated June 19, 2013 to its Consignees/customers. The letter described the Safety Issue, Safely Instructions, Affected Product Detail, Product Correction and Contact Information. The consignees were instructed to please ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call : United States: 800 437 1171. For other countries, please contact your local GE Healthcare Service Representative. |  | 4 | 4 (2 US, 2 OUS) | Worldwide distribution: US (nationwide) including states of : PA and TN; countries of: FRANCE and SINGAPORE. | Terminated | October 25, 2013 | 128 | 4 | Not_Computer
Z-2258-2013 | 65999 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | VisuALIF,System consisting of a 'U' shaped PEEK block in two footprint configurations, four heights and three lordosis angles., part number: 0301-1410. |  |  | July 31, 2013 | 2013 | Spine Smith Holdings, LLC | The product has the potential to be laser marked as a Medium when it actually is a Small. | Process control | The firm decided to remove this lot and verify all of the markings. A stock sweep for mis-marked parts was performed at the time and it is not known why this lot is identified, however an additional stock sweep of all PEEK products from all product lines was performed beginning 07/31/2013 and no additional lots were identified. |  | 32 | 32 | Distributed in FL. | Terminated | March 18, 2014 | 230 | 32 | Not_Computer
Z-2256-2013 | 66148 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650. |  |  | July 12, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware of a number of potential issues associated with the CARESCAPE Monitor B650.No injuries or illnesses have been reported. | Software design | GE Healthcare sent an Important Product Information letter dated July 12, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.GE Healthcare will provide a software and hardware correction at no charge once it is available. We will contact you to arrange for this correction.If you have any questions regarding this notification, please contact Technical Support (1-800- 558-7044) or your local Service Representative |  | 9466 | 9466 (3261 US, 6205 OUS) | Worldwide Distribution - USA (nationwide) including the states of : AL, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, NE, NV,NH, NH, NM, NY, NC, OH, OK OR, PA, PR, SC, TN, TX, VA, WA, WV, WI and WY., and the countries of : ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, BULGARIA, CANADA, CHILE, COLOMBIA, COSTA RICA. CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HONG KONG, HUNGARIA, ICELAND, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MALAYSIA, MOROCCO, NETHERLAND, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIA, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM and VENEZUELA. | Terminated | April 21, 2014 | 648 | 9466 | Computer
Z-2257-2013 | 66126 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, Seno Advantage, 5183685, 5183685-2, 5183685-3, 5183685-4, 5183685-5, 5183685-6, 5183685-7, 5183685-8, 5183685-9, 5342425, 5342425-2, 5342425-3, 5342425-4, 5342425-5. |  |  | September 20, 2012 | 2012 | GE Healthcare, LLC | GE Healthcare has recently become aware that patient mismatch without a warning message could occur on Seno Advantage (SA) workstations under certain circumstances. This issue may impact patient safety. | Software design | GE Healthcare sent an Important Product Information letter dated September 19, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. GE Healthcare has recently become aware that patient mismatch without a warning message could occur on Seno Advantage (SA) workstations under certain circumstances. This issue may impact patient safety.The Contact Information (US 800-437-1171, Japan 0120-055-919, Other countries contact local GE Healthcare Service Representative). |  | 179 | 179 | Worldwide Distribution - USA (nationwide) including the states of : AL AZ, CA, CO, CT, FL, GA, IL, IN, MD,MI, MO, NJ, NY, NC, OK, PA, TX and WI.,and the countries of : AUSTRIA, BELGIUM, CANADA, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, IRAQ, JAPAN, LEBANON, NORWAY, PORTUGAL, SPAIN, SWEDEN and TAIWAN. | Completed |  |  | 179 | Not_Computer
Z-2255-2013 | 66185 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King LT-D Oropharyngeal Airways |  |  | September 03, 2013 | 2013 | King Systems Corp. | Report that size 5 KLTD airways were packaged in pouches indicating a size of 2.5 and shipped to your facility. The correct size of the KLTD airway is clearly marked on the actual airway product and the physical size difference between 2.5 and 5 is distinct and obvious. | Labeling mix-ups | King Systems notified one consignee who received the affected product verbally on August 7, 2013, via email and an Urgent Medical Device Recall - Immediate Action Required notification on September 3, 2013. The Urgent Medical Device Recall - Immediate Action Required notification described the product and issue as well as includes a response form for the consignee to complete and send back to King Systems via FAX to 317-776-5175, or email to wnet@ambu.com. Questions regarding the notice are directed to: King Systems customer service representative at 800.642.5464, Monday through Friday 8:00am- 5pm EDT.For questions regarding this recall call 317-776-6823. |  | 54 | 9 cases; 45 products | Distributed in NY. | Terminated | April 15, 2014 | 224 | 54 | Not_Computer
Z-2260-2013 | 66186 | No | HGB | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Cranial Access (CRAK) Kits, Product / Catalogue No.s: INSHITH, INSHITHND, INSHITHRZN |  |  | September 04, 2013 | 2013 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra is recalling certain lots of Cranial Access Kit due to a possibility that the seal integrity of the sterile package may become compromised during shipping. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall letter dated September 4, 2013, to all affected customers that have been shipped unexpired affected product lots of IntegraÂż Cranial Access Kits, Catalogue No.'s INSHITH, INSHITHND, INSHITHRZN.The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to review their inventory and immediately stop the use or distribution of any affected product. Customers were instructed to identify and report to Integra if they do or do not have any of the affected product lots. Customers were advised how to return any affected product lot numbers in their possession.Effectiveness Checks will be conducted to determine the level of customers who:- Integra has verified, have been notified, and / or acknowledged the recall either by returning the acknowledgement form.- Indicate they will be returning / have returned unexpired affected product lots of kits.Any affected product lot numbers that are returned from consignees will be forwarded to the Integra Salt Lake City, UT facility for final disposition per its site QMS processes.Customers with questions were advised to contact Customers Service at 1-855-532-1723.For questions regarding this recall call 801-886-9505. |  | 8648 | 8,648 | Nationwide Distribution including CA, CO, FL, IN, KS, NC, NV, NY, OK, TN, TX, VA, WA, and WI. | Terminated | January 28, 2014 | 146 | 8648 | Not_Computer
Z-1538-2013 | 62624 | No | REA |  |  |  |  | Laserworld RGB &amp; Piko Nano RGB Models |  |  | July 19, 2012 | 2012 | Laserworld Usa Inc | A potential lack of certification, and/or missing laser certification, laser safety warning and administrative labels, and potentially incomplete Performance Feature set and incomplete User Information were discovered. | Radiation Control for Health and Safety Act | CDRH approves the CAP subject to the following conditions:1. Laserworld shall expand the scope of the CAP to include all projectors distributed between 2008 and 20l2, including, but not limited to, the Proline Series AG, Revolution Series AG, Club Series AG, Club Series, Pure Light Series, and Revolution Series AG, and Laserworld will continue to evaluate whether these additional projector models have similar defects.2. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.3. Laserworld will maintain records to document the number of projectors returned under the CAP, and will document the specific items of non-compliance addressed for each product tested, repaired, relabeled and/or replaced.4. Laserworld will provide CDRH with quarterly status updates on the progress of the CAP.5. The correction of the non-compliant products must be made at no cost to thecustomers.You may proceed with implementation of the CAP (subject to the conditions |  | 1004000 | Total distribution of more than 500 units between July1, 2008, and June 30, 2012 | US Distribution including the states of CA, CO, FL, KS and PA. | Open, Classified |  |  | 1004000 | Not_Computer
Z-2259-2013 | 66276 | No | EBF | Tooth shade resin material. | Dental | Dental | 510(k) | NX3 Try-In Gel. The product is used as a tooth shade resin material. |  |  | September 13, 2013 | 2013 | Kerr/Pentron, DBA Kerr Corporation and Pentron Clinical | Kerr Corporation is voluntarily recalling one lot of NX3 Try-In Gel, because some of the NX3 Try-In Gel syringes in the affected lot contain a different product material. The material in the affected syringe does not match the shade of the cement as it is labeled. | Mixed-up of materials/components | Kerr Corporation initiated the voluntary recall by sending out notifications on 09/13/2013 via USPS 1st class mail. The recall notification letter, titled "URGENT: MEDICAL DEVICE RECALL", informed customer of the recall's product description, reason for recall, potential risk, instructions to customers and contact information. |  | 197 | 197 units | Worldwide Distribution, including Nationwide in the US and the countries of Canada, Great Britain, Australia, France, Germany, Italy, Denmark, Hong Kong, China, Brazil and Saudi Arabia. | Terminated | February 11, 2014 | 151 | 197 | Not_Computer
Z-2261-2013 | 65981 | Yes | DTM | Cardiopulmonary bypass arterial line blood filter. | Cardiovascular | Cardiovascular | 510(k) | Terumo Pall AL6X Arterial Blood Line Filter, labeled in part Terumo Cardiovascular Systems Corporation, packaged as part of a Terumo Cardiovascular Procedure Kit or as a single, sterile filter unit |  |  | July 26, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Medical device used in cardiovascular procedures is defective. | Nonconforming Material/Component | Terumo Cardiovascuolar Services sent recall letters to Chief of Perfusion, Director of Operating Room Services, Director of Biomedical Services, and Risk Management departments on July 26, 2013. The notifications advised that filters are being recalled because some of the filters may exhibit a partial or full occlusion of an outlet or purge port. Terumo CVS advised users to stop using Terumo AL6X arterial filters from the affected lot numbers for replacement. A supply of replacement filters for customer's immediate use accompanied the notification. For users ordering the AL6X filters as part of a Cardiovascular Procedure Kit, Terumo CVS is providing a preliminary supply of replacement Terumo AL6X arterial filters in which users may follow the instructions to replace the filter or users may return the Cardiovascular Procedure Kit and Terumo CVS will replace the filter for them. Customers with questions were advised to contact Terumo CVS Customer Service at 800-521-2818.For questions regarding this recall call 800-521-2818. |  | 15817 | 15,817 filters | Nationwide Distribution including WA, IN, OK, TX, MI, CA, NY, CT, AZ, FL, WV, NC, VA, PA, TN, GA, AL, MA, SC, OH, MS, MO, SD, WI, AR, AND NV. | Terminated | July 28, 2015 | 732 | 15817 | Not_Computer
Z-2262-2013 | 66049 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | QUADROX iD HMOD 2030 Diffusion Membrane Oxygenator; QUADROX-i ADULT HMO 70000 / HMO 71000; QUADROX-iD Pediatric HMOD 30000. The QUADROX oxygenators are intended for use in extracorporeal circulation during cardiopulmonary bypass in cardiac surgery. |  |  | August 15, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | The manufacturer has conducted a review of the results of product testing performed during the last twelve months and has identified specific lots of product which may not meet specifications for bond strength on the inlet and/or outlet connector. | Nonconforming Material/Component | Maquet sent a Urgent Product Recall letter/Response Form, dated 8/15/2013 via Fed Ex. On 8/26/2013 Maquet sent a letter to their customers informing them that all shipments of oxygenators and stainless steel hold clamps have been placed on a temporary hold due to an agency review in connection with FDA clearance. Customer service may be contacted at 1-888-627-8383, (option 2 followed by option 2) Monday though Friday between the hours of 8:00 am and 6:00 pm EST.Maquet issued a Certificate of Medical Neccesity (CMN) on September 17, 2013 to notify Quadrox customers that they may continue to receive and use their current inventory of Quadrox oxygenators and clamps by acknowledging the conditions provided in the CMN. |  | 1182 | 1,182 units | Worldwide Distribution, including Nationwide (US) and foreign countries. | Terminated | July 11, 2017 | 1426 | 1182 | Not_Computer
Z-2263-2013 | 66247 | No | NZA | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Reliance Endoscope Processing System (REPS), STERIS Corporation, 5960 Heisley Road, Mentor, OH 44060-1834 USA www.steris.com. |  |  | August 29, 2013 | 2013 | Steris Corporation | The air pipe that delivers unfiltered ambient air from the procedure room into the water bottle on the listed FUJINON G5 &amp; G8 type flexible GI endoscopes does not connect to the main air/water channel within the light guide (LG) connection head and, as a result, the air pipe does not receive forced flow during the Reliance EPS processing cycle. | Use error | Steris sent an "URGENT PRODUCT SAFETY NOTICE" dated September 5, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. |  | 266 | 266 units | Worldwide Distribution-USA (nationwide) including the states of AK, CO, FL, GA, HI, ID, IL, KS, LA, MD, MN, MO, MT, NE, NM, OH, PA, SC, UT, &amp; WY and the countries of Australia, Canada &amp; Spain. | Terminated | July 03, 2014 | 308 | 266 | Not_Computer
Z-2264-2013 | 66063 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | ENDOPLUS; Precision Laparoscopic Instruments; REF 346VA; Description 5mm Atraumatic Grasper S/A, Standard; Non-sterile; Manufactured by Endoplus; 750 Tower Rd., Suite A, Mundelein, IL 60060 USA. |  |  | July 15, 2013 | 2013 | Endoplus, Inc. | Multiple complaints reporting that the moveable jaw has broken during laparoscopic procedures. | Process control | Letters were sent on 7/15/2013 to all customers notifying them of the recall. The letters requested that customers return the recalled devices via private courier using an enclosed account number. Replacement devices will be made available upon customer request. |  | 81 | 81 units | Distributed in the states of FL, TX, and VA. | Terminated | February 08, 2016 | 938 | 90 | Not_Computer
Z-2266-2013 | 65992 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE 1.5T Signa HDxt, 3.0T Signa HDxt, Signa Vibrant, HD23 Software and the following model numbers: 2395001, 2395001-2, 2395001-3, 2395001-4, 5269997-2, 5269997-4, 2377062-61, 5127452-3, 5127452, 5337894-3, 5337894-4, 5911000-4, 5911000-5, 53085735308573-2. |  |  | April 09, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to non-unique series identifiers embedded in image headers of certain GE MR Products.For ClariView and Screen Save from GVTK (Gems Visualization ToolKit) Viewer, the MR Product can generate image headers with non-unique series identifiers, which when exported to some PACS systems, results in the series being associated wit | Software design | Customers were sent two GE Healthcare "Urgent Medical Device Correction" letters one dated April 9, 2013, and the other May 22, 2013, each addressing one of the 2 groups of MR Products involved in the recall. The letters were addressed to Hospital Administrators / Risk managers, Radiology Department Mangers and Radiologists. The letters described the Safety Issue and provided Safety Instructions, Affected Product Details, Product Correction and Contact Information.GE Healthcare will correct all affected systems by providing a software update at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.Contact your local service representative if you have any questions concerning this notification.For further questions please call (262) 513-4122. |  | 1246 | 1,246 (521 US and 725 OUS) | Worldwide Distribution - All states in continental USA including DC except AK, AZ, RI and WY. and the countries of : AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KOREA, KUWAIT, LEANON, MARINIQUE, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE,SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, LUXEMBOURG, PORTUGAL, VIETNAM and URUGUAY. | Terminated | May 05, 2014 | 391 | 1246 | Computer
Z-2267-2013 | 65883 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | DRI Cocaine Metabolite Assay, labeled as Multigent Cocaine Assay for Abbott Laboratories. |  |  | July 23, 2013 | 2013 | Microgenics Corp | Surveillance of reagent stability for specific lots of DRI Cocaine Metabolite Reagent show an unexpected decreasing change in the reaction rate, indicating a faster than expected reagent degradation. | Under Investigation by firm | Thermo Fisher Scientific sent an Urgent Medical Device Recall letters dated July 23, 2013, except Abbott labs who was mailed a letter on August 2, 2013. Abbott had previously been notified by telephone of the recall.The letter identified the product the problem and the action needed to be taken by the customer.Customers advised users on reagent stability issue with specific lots and the fact that it indicated a faster than expected degradation of the reagent. Microgenics asked that all use of the affected lots be stopped. Customers were informed that since each patient run is monitored with the controls at +/-25% of the cutoff, a possibility of reporting incorrect patient results is minimal to nonexistent. The use of appropriate controls addresses an risks of incorrect results.In addition to stopping use of the affected reagent, customers are asked to contact Thermo Fisher Scientific for replacement and to keep a copy of the recall letter for laboratory files. If the affected lots of reagent were forwarded to any other laboratories, they should be contacted immediately.. The enclosed acknowledgement form is to be completed and returned indicating that the letter was received, how much product, if any was on hand and destroyed.Questions and concerns may be addressed to Microgenics Technical Service at 1-800-232-3342 or 510-979-5417. |  | 1473 | 1473 units in distribution. | US Distribution including the states of DC, MA, NJ, TX. FL, CA, WA, NY, TX, CT, KY, LA, AR, TN, AL, GA, CT, VA, NC, KS, MD, OH and IL | Terminated | December 02, 2013 | 132 | 1473 | Not_Computer
Z-2270-2013 | 65348 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Invacare Power Wheelchair, one per box |  |  | June 04, 2013 | 2013 | Invacare Corporation | Unintended slowing or deceleration, and on rare occasion,unanticipated and unattended acceleration of power wheelchairs equipped with SPJ and MPJ type joysticks. | Process control | Invacare sent an Urgent Patient Safety Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact their wheelchair povider to have their joystick replaced. Customers with questions were instructed to call their provider at Invacare's recall support line at 877-413-6008. Go to www.invacare.com/joystickrecall for more information regarding this recall |  | 166230 | 166,230 | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. Internationally to Canada, Australia, Bermuda, Brazil, China, Costa Rica, Germany, France, United Kingdom, Guam, Hong Kong, Israel, India, Kuwait, New Zealand, Puerto rico, Qatar, Saudi Arabia, Singapore, Thailand, Trinidad, Tobago, and South Africa. | Terminated | January 04, 2016 | 944 | 166230 | Not_Computer
Z-2268-2013 | 66234 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Smith Medical, Deltec, REF 21-2386-24, 11 Fr Introducer Set with Peel-Away Sheath, Sterile EO, Rx Only |  |  | September 10, 2013 | 2013 | Smiths Medical ASD, Inc. | Smiths Medical is conducting a voluntary recall of PORT-A-CATHÂż and PORT-A-CATHÂż II Implantable Venous and ArterialAccess Systems and Introducer Sets.Smiths Medical has become aware of a small number of customer reports regarding the introducer sheath and the introducer opening being too small for the width of the catheter with the use of these Sets.Smiths Medical has received no reports | Mixed-up of materials/components | Smiths Medical ASD, Inc. sent an Urgent Field Safety Notice dated September 6, 2013, to all affected customers. The letter described the problem and the product involved in the recall. The letter included the "Advice on Action to be Taken by the User" which instructed consignees to inspect and remove affected devices and to complete and return the "Urgent Field Safety Notice Confirmation Form. The letter also included information on the "Transmission of this Urgent Field safety Notice" which included information where customers should report issues related to the product to Smiths Medical's Global Complaint Department at 1-866-216-8806 or globalcomplaints@smiths-medical.com. For questions on the information provided they should contact 800-258-5361 (Smiths Medical's Customer Service Department).For questions regarding this recall call 1-800-332-0178. |  | 190 | 190 | Worldwide Distribution - USA including AZ, AR, CA, GA, IL, IN, KY, MD, MA, MI, NV, NJ, NY, NC, OH, OR, PA, RI, SC, TN, TX, VA, WA, WV, WY.Internationally to DENMARK, BELGIUM CANADA. | Terminated | October 21, 2014 | 406 | 562 | Not_Computer
Z-2271-2013 | 66211 | cart, emergency, cardiopulmonary (excluding equipment) | BZN | Cardiopulmonary emergency cart. | Anesthesiology | Anesthesiology | 510(K) Exempt | Stryker Vision Mounting Arm, for use with other Stryker Endoscopy medical device products for human patients. Model number 0240-095-200, manufactured by Stryker Endoscopy San Jose |  |  | August 30, 2013 | 2013 | Stryker Endoscopy | When positioning the display located on top of the video cart, it is possible for the user to apply excessive force and push the display up and out of the mounting arm. This can result in the display falling upon the user/patient. | Device Design | Stryker Endiscopy sent an Urgent Device Correction letter dated August 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inform individuals within their organization who needed to be aware of this Medical Device Field Correction, contact Stericycle at 1-877-546-0482 or email stryker4645@stericycle.com to schedule a time for Stericycle to perform the correction, or perform the correction themselves by requesting the materials and following the correction instructions included. Customers were also instructed to report all adverse events or product quality problems to Stryker Endoscopy at 1-800-624-4422, press 3 for Customer Service. --UPDATE--On September 17, 2013, the recall was expanded to include additional customers who may have received the screw and washer included with the design change update. Notification was sent to these consignees on September 17, 2013. . |  | 14271 | 14271 units | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and internationally to Australia, Canada, Chile, China, Argentina, Guam, Taiwan, India, Japan, Korea, Latin America, Philippines, Mexico, New Zealand, Hong Kong, Singapore, and Switzerland. | Terminated | April 30, 2014 | 243 | 14271 | Not_Computer
Z-2278-2013 | 65009 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 3.5 MM LCP(R) Distal Humerus System |  |  | March 08, 2013 | 2013 | Synthes USA HQ, Inc. | Recall is being initiated due to the part being mislabeled (Part Number 241.267 was mislabeled with Part Number 241.276). | Labeling mix-ups | Synthes sent an Urgent Notice Medical Device Recall letter dated March 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to obtain a Return Authorization Number, complete the Verification Section at the end of the letter by the checking the appropriate box indicating affected product located. Return the Verification Form with the product to: Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. Customers with questions were instructed to call 610-719-5450 or email Fieldaction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 2 | 2 | Nationwide Distribution including CA,and VA.. | Terminated | May 21, 2015 | 804 | 2 | Not_Computer
Z-2275-2013 | 66051 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | TEGRIS system |  |  | August 19, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | MAQUET has identified a potential issue that might occur because of the lack of an emergency stop button on a secondary touch screen display on the TEGRIS System. | Device Design | Maquet sent an Urgent Medical Device Recall Letters/Field Action Response Forms dated August 19, 2013, via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were informed that a MAQUET authorized service technician would contact them to arrange an appointment to equip their system with a second emergency stop button that would be located in close proximity to the secondary touch screen. Customers were instructed to complete and return the enclosed Field Action Response Form acknowledging their receipt and understanding of the communiction by emailing a scanned copy to FieldActions@maquet.com or via fax at 973-807-1658. Customers with questions were asked to contact their MAQUET representative or MAQUET Technical Support at 888-627-8383, option 3.For questions regarding this recall call 973-709-7361. |  | 21 | 21 units | Worldwide Distribution - USA including NY and PA and internationally to Australia, Brazil, Canada, Germany, Hong Kong, Japan, Sweden. | Terminated | April 17, 2014 | 241 | 21 | Not_Computer
Z-2276-2013 | 65996 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System |  |  | July 02, 2013 | 2013 | Elekta, Inc. | Upgrade installation script is designed to overwrite the same pouch files in the Treatment Couch Library with the original default values. | Device Design | Monaco 3.3 upgrades have been temporarily halted and Important Safety Notice (LINMON0007) has been issued to effected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that couch files are checked for releases 3.20.00 and 3.30.00 to verify that the correct relative electron density values are being used in the dose calculation. If customers have edited the sample couch files, and previously upgraded to Monaco 3.20.00 or 3.30.00, their couch files have already been overwritten. Open the couch files, make the necessary updates to the couch parameters, and use the "Save As Treatment Couch" option to save the modified sample couch files to a new, unique name.Please note that the couch files can also be edited in Focal and then the patients opened in Monaco. Comparable Focal releases would be Focal4.64.00(Monaco 3.10.00), Focal4.70.00(Monaco3.20.00) and Focal4.80.00 (Monaco 3.30.00).For questions regarding this recall call 404-993-5886. |  | 45 | 45 | Nationwide Distribution including AK, CA, CO, CT, FL, IL, IN, MI, MN, MO, NC, NJ, NY, OH, OK, OR, PA, SC, TN, TX, WA and Puerto Rico. | Terminated | May 08, 2014 | 310 | 45 | Not_Computer
Z-2277-2013 | 65998 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco RTP System Radiation Therapy Treatment Planning System. |  |  | May 22, 2013 | 2013 | Elekta, Inc. | Patient needs to be positioned as indicated by the treatment plan. | Labeling design | Computerized Medical Systems issued an Important Safety Notice dated May 22, 2013 to all affected customers. This letter idenitfies the problems the products and the actions needed to be takem by the custom. |  | 72 | 72 | Distributed in Puerto Rico and the states of AK, AZ, CA, CO, CT, FL, IL, IN, MA, MD, MI, MN, MO, NC, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, and WA. | Terminated | December 23, 2013 | 215 | 72 | Not_Computer
Z-2283-2013 | 65922 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | The FL23SE Bed is an AC-powered adjustable hospital bed with four built-in electric DC motors and remote controls that is intended for medical purposes. It can be operated by the patient or the caregiver to adjust the height and surface contour of the bed. The bed includes movable and latchable siderails. The FL23SE Bed is intended for use with patients for procedures, therapy, recovery in healthcare environments, and to transport patients between bays and procedural rooms. |  |  | August 19, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | The nurse call cable (Part Number: QDF23-0571) on the FL23SE Bed is located right under the opening of the bed head box. If an accessory (i.e. IV pole) is inserted into this slot, the circuit directing the nurse call system can be damaged. | Device Design | Stryker sent an Urgent Medical Device Recall Notification letter dated August 19, 2013, to all affected consignees. The notification described the affected product, reason for recall, potential hazards, and recall customer actions and reply instructions. A Stryker Representative will contact the consignees to arrange a time to repair the unit(s) and eliminate the risk of the potential interference . Urgent questions or concerns are directed to Blaine Burnett (269)389-6954. Stryker Medical business hours are Monday-Friday 8 a.m-5 p.m. (ET).For questions regarding this recall call 269-324-6609. |  | 61 | 61 units | Worldwide Distribution - USA including MI, WI, VA and internationally to Canada. | Terminated | January 29, 2014 | 163 | 61 | Not_Computer
Z-2279-2013 | 66150 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set, Model#72313E |  |  | August 30, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set, Model 72313E, Lot Number13016314, because of an incorrect filter was assembled on this lot. The correct filter is a 1.2 micronfilter (blue color) but a 0.2 micron filter (clear color) may have been added instead. | Employee error | CareFusion sent an Urgent Medical Device Recall Notification dated August 30, 2013, to all affected customers. The letter informed the customers of the problem identified, issues, potential risk, and the immediate actions to be taken. Customers were instructed to immediately check their inventory for the affected product, if any affected product is encountered, customers were instructed not to use. Customers were instructed to complete the customer response form and return the form to CareFusion with the affected product for replacement.. A distributor letter was also sent to distributors on August 30, 2013, to inform them of the problems identified and the actions to be taken. Distributors were instructed to contact the Customer Support Team, return customer response form, and if they have questions then they were instructed to contact the CareFusion Support Center at 1-800-562-6018.For questions regarding this recall call 858-617-4000. |  | 600 | 600 units | CA, FL, AZ, and TN | Terminated | October 28, 2013 | 59 | 600 | Not_Computer
Z-2280-2013 | 66209 | No | CIC | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABX PENTRA Calcium CP Reagent, Model # A11A01633 |  |  | September 06, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is informing all PENTRA 200 and PENTRA 400 customers of a potential for reportingincorrect calcium results when using the ABX PENTRA CALCIUM CP reagent. | Under Investigation by firm | Horiba Medical sent an Urgent Safety Notice dated September 6, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customer notification letter dated September 6, 2013 will be sent to all PENTRA 200 and PENTRA 400 customers of a potential for reporting incorrect calcium results when using the ABX PENTRA CALCIUM CP reagent (Reference #: A11A01633) Lot#: 014664503. The recall letter informs the customers that the initial functional tests indicate that there is no interference identified with regards to the patient results. However, with time, there may be an increase in the quantity of precipitates and this could potentially interfere with the performance of this reagent. Customers are informed of the actions to be taken and to complete the customer response form within 10 days.Further questions please call (949) 453-0500 Ext. 208 |  | 2658 | 2,658 units ( 1240 US 1418 OUS ) | Worldwide Distribution - USA including the country of Canada. | Terminated | August 23, 2014 | 351 | 2658 | Not_Computer
Z-2282-2013 | 64236 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Oracle Spacer System Slap Hammer, Part Number 03.809.690, Lot Number 3723847 |  |  | January 11, 2013 | 2013 | Synthes USA HQ, Inc. | This recall is being initiated in response to complaints received where it was reported that the proximal end of the Slap Hammer broke during use. | Device Design | SYNTHES sent an Urgent Notice Medical Device Recall letter dated January 11, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. The letter requested consignees examine their inventory, remove the recalled product from use and return them to Synthes. Customers were also instructed to completed the attached Verification Section at the end of the letter indicating whether or not they had any of the affected product and the number of devices found. Customers with questions were instructed to call 610-719-5450 or email Fieldaction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 27 | 27 | Worldwide Distribution - USA including CO, PA, TN, VA, TN, WA, and internationally to Czech Republic. | Terminated | August 31, 2015 | 962 | 27 | Not_Computer
Z-2284-2013 | 66355 | No | ETA | Total ossicular replacement prosthesis. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Universal Titanium Prosthesis, Round Head, with Flex H/A titanium shoe. The ear prosthesis is implantable and indicated for hearing loss. |  |  | August 23, 2013 | 2013 | Medtronic Xomed, Inc. | One lot of the product was assembled using the Cam Head version of the same prosthesis instead of the Round head Shaft. | Process control | An "URGENT PRODUCT RECALL" letter was faxed to customers following the initial telephone notification. The letter described the product issue and provided recommended actions. Questions were directed to customer service at 800-874-5797. |  | 60 | 60 | Nationwide Distribution (US), including the states of CA, NE, GA, and IL. | Terminated | November 07, 2013 | 76 | 60 | Computer
Z-2274-2013 | 66070 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | BD Epilor Syringe |  |  | August 23, 2013 | 2013 | Becton Dickinson &amp; Company | BD has received several reports indicating that the plunger of the 7 mL BD Epilor plastic BD Luer-Lok Loss of Resistance Syringe may stall or stick within the syringe preventing the user from detecting a loss of resistance upon entering the epidural space. | Unknown/Undetermined by firm | BD sent a product advisory notice to distributors/customers dated August 23, 2013 via UPS 2nd day mail. Consignees will be contacted by phone email and letter. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to be aware of the issue with the product and to follow the standard routine procedure to ensure that there is no resistance in the syringe prior to use. Any syringe that shows sticking or stalling should be discarded. Customers who distributed this product, should identify their customers and notify them at once of the product advisory. Customers with questions were instructed to call BD Customers Service at 1-800-237-2762.For questions regarding this recall call 201-847-6800. |  | 1335885 | 1,335,885 units | Worldwide Distribution - USA including FL, VA, TX, NY, IL, AZ, MI, RI, NJ, SC, MO, TN, CA, MN, UT, GA, NE, OH, MD, PA, MT and internationally to Brazil, Australia, Canada, and Puerto Rico.. | Terminated | March 19, 2014 | 208 | 1335885 | Not_Computer
Z-2281-2013 | 66152 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CareFusion Gravity Set, Model #44000-07 |  |  | August 30, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the Gravity Set (Model 44000-07) because of an incorrect expiration date. The affected lot numbers of the Gravity Set are labeled witha 5 year expiration instead of three (3) years. | Employee error | CareFusion sent an Urgent Medical Device Recall Notification letter dated August 30, 2013, to all affected customers to inform them that CareFusion is recalling the Gravity Set, Model 44000-07, Lot Numbers 12086930, 12106215, 13016408, 13016834, 13025446 and 13025672 as a result of an incorrect expiration date. The letter informs the customers of the problem identified, issues, potential risk, and the immediate actions to be taken. Customers are instructed to complete the customer response form and return the form to CareFusion. Customers are instructed to return recalled products directly to distributors. A distributor letter will also be sent to distributors on August 30, 2013, to inform them of the problems identified and the actions to be taken. Distributors are instructed to contact the Customer Support Team, return customer response form, and if they have questions then they are instructed to contact the CareFusion Support Center at 1-800-562-6018.For questions regarding this recall call 858-617-4000. |  | 2850 | 2,850 units | IL | Terminated | December 26, 2013 | 118 | 2850 | Not_Computer
Z-2285-2013 | 66151 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set, Model# 10015048. Intravascular administration set. |  |  | August 30, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set, Model# 10015048, because there is a potential for separation at the tubing sleeve below the drip chamber. | Process design | A customer letter dated August 30, 2013 will be sent to customers to inform them that CareFusion is recalling the SmartSiteÂż Low Sorbing Infusion Set Models 10015048, Lot Numbers 130153883 and 13025342 as a result of a potential for separation at the tubing sleeve below the drip chamber.The letter informs the customers of the problem identified, issues, potential risk, and the immediate actions to be taken. Customers are instructed to complete the customer response form and return the form to CareFusion. Customers are instructed to return recalled products directly to distributors. A distributor letter will also be sent to distributors on 8/30/13 to inform them of the problems identified and the actions to be taken. Distributors are instructed to contact the Customer Support Team, return customer response form, and they are instructed to contact the CareFusion Support Center at 1-800-562-6018 with any questions. . |  | 15860 | 15860 units | Nationwide in US: PA, OK, TX, FL, NC, CA, IN, WI | Terminated | April 16, 2014 | 229 | 15860 | Not_Computer
Z-2291-2013 | 66242 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery MR750 3.0T, Optima MR450w 1.5T, and Discovery MR450 1.5T . |  |  | August 13, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to exposed yellow foam on the end bells of Discovery MR450, Discovery MR750, and Optima MR450w products.The yellow foam strip used between the end bell covers and patient bore of certain GE MR Products may extend past the surface of the covers. The foam itself is not hazardous. However, any fluid or body substance absorbed | Component design/selection | GE Healthcare sent and Urgent Medical Device Correction letter dated August 13, 2013, to all affected consignees. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Mangers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product details, Product correction and contact Information. Customers were instructed that the system may continue to be used as normal even with the yello foam protruding in the bore. Customers were told not to attempt to remove the yellow foam. Customers with questions were instructed to contact their local service representative if they have any questions concerning the notification.For questions regarding this recall call 262-513-4122. |  | 828 | 828 (443 US, 385 OUS) | Worldwide Distribution USA including DC, all states except HI, ND, RI and WY and Internationally to Australia, Bahrain, Belarus, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, India, Indonesia, Israel, Italy, Japan, Korea (Republic of), Kuwait, Lebanon, Martinique, Mexico, Netherlands, New Zealand, Norway, Oman, Poland, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom, Uruguay, .OUS: AUSTRALIA, AUSTRIA, BAHRAIN, BELARUS, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, MARTINIQUE, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, POLAND, QATAR, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY. | Terminated | April 25, 2014 | 255 | 828 | Not_Computer
Z-2272-2013 | 66184 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | Universal Flex2 Breathing Circuit |  |  | August 23, 2012 | 2012 | King Systems Corp. | King Systems received a customer complaint indicating that the customer had ordered a breathing circuit with a Latex breathing bag. The product they received contained a Latex breathing bag, but was labeled as non-latex. | Labeling mix-ups | Ambu, Inc. sent a letter dated June 24, 2013, informing the customer that they have decided to discontinue the manufacture and sales of the King Systems latex breathing bags and to conver all customers to their King Systems latex-free breathing bags. The customer was advised to adjust their inventories accordingly and transition to the latex free part number as their inventory levels reach their minimums. For questions the customer was advised to call 317-403-8677.For questions regarding this recall call 317-776-6823. |  | 41 | 1 case; 40 units | Nationwide Distribution - NE only. | Terminated | November 18, 2013 | 452 | 41 | Not_Computer
Z-2292-2013 | 65043 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Solitaire-35 Peek Solitaire Anterior Spinal System. The Solitaire-35 is an intervertebral body fusion device. |  |  | April 09, 2013 | 2013 | Ebi, Llc | Several complaints were recorded regarding one of the three screws driving through the cage of the Solitaire 35 implant. | Process control | Biomet LLC sent "Urgent Medical Device Recall" letters/fax back response forms on Tuesday April 9, 2013 via Fed Ex to all distributors. The notification identified the problem, product involved, and risk factors. In addition, the notification provides instructions to the representatives on how to return the product to EBI, LLC. Surgeons that have implanted the product were contacted on April 12, 2013. |  | 1730 | 1730 units | Nationwide distribution: US including states of: CA, GA, IN, KY, MO, NY, OR, TX, and WI. | Terminated | June 17, 2014 | 434 | 1730 | Not_Computer
Z-2273-2013 | 66219 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Product Name: AVEAÂż Ventilator. |  |  | September 05, 2013 | 2013 | Carefusion 211 Inc dba Carefusion | CareFusion has identified a potential risk associated with AVEAÂż ventilators when used at higher altitudes. A proactive complaint review identified an error in the barometric pressure compensation.CareFusion is voluntarily initiating a field correction of the affected devices to correct the error. | Component design/selection | CareFusion sent customer notification letters on September 05, 2013 via USPS mail return receipt carrier. The customer notification letter dated September 3, 2013 was titled "URGENT PRODUCT RECALL", and was addressed to Biomedical Engineering, Respiratory Care, and Risk Management. The customer notification letter informed customers of the product under recall with product name, description, and serial numbers enclosed as list unique to each customer. The customer notification letter also states the Problem and Affected Devices: Issue, Potential Risk, Actions to be Taken by CareFusion, Action to be Taken by the Customer, and Contact Information.Care Fusion has initiated a project to develop a new version of software that will resolve this issue. Once this software is available to download you will receive a notification letter.The customers were instructed to promptly return the enclosed Response Card to expedite the correction process and acknowledge receipt of this Notification and follow the listed instructions noted in the letter.For Recall Related and Support contact CareFusion Recall Support Center at 888.562.6018 or email: SupportCenter@carefusion .com Questions. For Product Technical Support Adverse Event Reporting contact Support/Customer Advocacy at 800.231.2466, 714.283.2228 or email to support.vent.us carefusion.com. from (6:30am -5:00pm POST).UPDATED:Firm issued press release of the recall on 09/30/2013. |  | 19546 | 19,546 units | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Albania, American Samoa, Argentina, Australia, Austria, Bahrain, Barbados, Bangladesh, Belgium, Bermuda, Belize, Bolivia, Bosnia and Herzegovina, Brazil, Canada, Chile,China, Colombia, Costa Rica, Croatia, Cyprus, Czech Rep, Dominican Republic, Ecuador, Egypt, Estonia, El Salvador, Finland, France, Germany, Georgia, Greece, Guatemala, Hong Kong, Honduras, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Lithuania, Lybia Jamahiriya, Malaysia, Mexico, Moldova, Morocco, Netherlands, Nepal, Nicaragua, Norway, Oman, Pakistan, Palastine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Rwanda, Saudi Arabia, Singapore, Slovenia, Slovakia, Spain, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, and United Kingdom. | Terminated | June 04, 2015 | 637 | 19546 | Computer
Z-2301-2013 | 66195 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Selenia Full Field Digital Mammography System with software versions 5.0.x and 5.1.x |  |  | September 06, 2013 | 2013 | Hologic, Inc. | Software: When the Selenia upon meeting the hard drive threshold of 4,000 images, may delete the most recently dated images | Software design | Hologic sent an Urgent Medical Device Recall letter dated September 5, 2013, by Certified Mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised how to avoid loss of images until Hologic corrects the problem. A software upgrade to customers who are connected to Hologic Connect who have already been connected remotely and customers who have already been connected via a Service Call will receive notification of the problem and the date the system was remotely corrected notified by Certified Mail.A Dealer notification issued by certified mail to inform Dealers of the need to notify their customers of the problem and to correct the identified Selenia system using Hologic Technical Bulletin TB-00072.Selenia software version 5.2 will implement a double safeguard to both guarantee a 28 day retention period, as well as ensure images are removed in a least recently used order.If there are any questions regarding this field correction, please call Hologic Customer Support at 877 371-4372. |  | 62 | 62 systems | Worldwide Distribution - USA including the states of AL, AR, CA, CO, FL, IL, KY, LA, NE, NC, NJ, NY, OH, PA, TN, TX, UT, WA and WV., and the countries of Canada, Denmark, Iran, Italy, Mexico and Taiwan. | Terminated | December 13, 2013 | 98 | 62 | Computer
Z-2302-2013 | 66334 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | NC Sprinter; Rapid Exchange Balloon Dilatation Catheter; |  |  | September 16, 2013 | 2013 | Medtronic Vascular, Inc. | Compliance chart included in lot 207002011 of the NC Sprinter RX 2.25 x 12 mm products lists the size of the device as 2.5mm instead of the correct 2.25mm size. | Mixed-up of materials/components | Medtronic sent an Urgent Medical Device Recall letter beginning September 16, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that your facility has received affected NC Sprinter RX Balloon Dilatation Catheters as noted in the attached customer notification detail report. Consequently, Medtronic is asking you to take the following actions:1. Remove and quarantine all affected catheters that remain in your inventory.2. Return the affected catheters to Medtronic. A Medtronic representative can assist in facilitating the return and replacement of this product as necessary.3. Complete the attached Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality.Medtronic has taken the necessary steps to prevent any future shipment of mislabeled product. Regulatory agencies will be notified about this recall as applicable.Please share this notification with others in your organization as appropriate. If any NC Sprinter RX Balloon Dilatation Catheters within the scope of this recall have been forwarded to another facility, please notify that facility accordingly and facilitate the retrieval of the affected product.We appreciate your cooperation and apologize for the inconvenience that this issue may cause. Should you have any questions, please contact your local Medtronic representative. |  | 122 | 122 | Worldwide Distribution - USA and the countries of Australia, Belgium, Canada, Hong Kong, Malaysia, New Zealand, Singapore, Spain, Switzerland, Taiwan and Thailand. | Terminated | February 10, 2014 | 147 | 122 | Not_Computer
Z-2303-2013 | 66108 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Lactate Membrane units for E7077 Electrode Radiometer Medical ApS. The intended use is in vitro testing of samples of whole blood. |  |  | October 08, 2012 | 2012 | Radiometer America Inc | RADIOMETER became aware that some membranes may have enzyme residue on the outer membrane. The enzyme residue may cause an initial negative bias on the reported Lactate result upon replacement of the Lactate membrane. | Process control | On 11/07/2012, the firm sent "Urgent Field Safety Notices to their customers. The product issue was described and recommended actions were provided. Customers may contact 1 (800) 736-0600, option 4 if questions. |  | 54528 | 13,632 boxes/4 membranes | Nationwide Distribution, including the states of: AK, AL, AR, AZ, BM, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and WV. | Terminated | December 24, 2013 | 442 | 54528 | Not_Computer
Z-2295-2013 | 65821 | No | LKK |  |  | General Hospital | PMA | MedStream Refill Kit, U.S: 91-4287 |  |  | August 13, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Drug flow rates exceeding programmed flow rates. | Use error | Codman &amp; Shurtleff, Inc. sent a Medical Device Field Safety Notice letter dated August 13, 2013 to US Consignees via FedEx overnight delivery and International Consignees were notified of the recall via email. The letter identifed the affected products, potential clinical impact, and updated IFU pump filling warning. The letter informed consignees that the Instructions for Use (IFU) and product training materials had been updated by adding a warning statement about potential over infusion if air gets inside the pump reservoir and additional clarifications to reinforce the proper filling technique. Consignees are asked to complete the Field Safety Notice Acknowledgement Form and return it to your local Codman Neuro Sales Representative or fax to: 508-977-6665.Consignees are asked to contact your Codman Neuro representative if they have questions, or contact Codman Neuro Clinical Support at 1-800-660-2660. Please report any malfunction or adverse event related to this device to 1-866-491-0974, Option 2. |  | 5 | 5 units | Worldwide Distribution - USA Nationwide in the states of NJ, NC, TX, FL and in the countries of Austria, Belgium, Egypt, Germany, Greece, Ireland, Israel, Italy, Luxemborg, Portugal, Russia, Saudi Arabia, Slovenia, Spain, Sweden, Switzerland, Turkey, and UK. | Terminated | September 11, 2014 | 394 | 4116 | Not_Computer
Z-2318-2013 | 66301 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Horizon Medical Imaging (HMI) Radiological Image Processing System, product versions 4.6.1 to 11.6., is a medical image and information management application intended to receive, transmit, store, retrieve, display, print and process digital medical images, digital medical video, and associated medical information from medical imaging systems. |  |  | September 04, 2013 | 2013 | Mckesson Information Solutions LLC | Software malfunction may occur resulting in a certain combination of server configurations which has the potential to cause image loss. | Software change control | Consignees were notified via Urgent Field Safety Notice dated September 4. MMIC Customer support will be contacting customers known to have a multi-server configuration and will implement the configuration change. |  | 329 | 329 units | Worldwide Distribution - United States including the territories of Guam and Puerto Rico and the countries of Canada and France. | Open, Classified |  |  | 329 | Computer
Z-2311-2013 | 65697 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Advanta SST Grafts &gt;or= 6mm |  |  | July 12, 2013 | 2013 | Atrium Medical Corporation | Inner tray label may adhere to wall of outer tray which make it difficult to remove the tray onto the sterile field | Package design/selection | Atrium Medical Corporation issued an Urgent Medical Device Safety Alert letter on July 16, 2013, via UPS to customers and e-mail to the field representatives that represent these accounts. The notification identifies the problem, product, and risk factors. In addition, it provides instructions how to handle product that may be affected.For questions customers were instructed to contact their Atrium Medical Customer Service at 603-880-1433 Prompt 1 or by e-mail at returns@atriummed.com |  |  |  | Worldwide Distribution - USA (nationwide) and internationally to the following countries:American Samoa AntillesArgentina Australia Austria Bangladesh Belgium Bolivia Brazil Canada Chile Colombia Egypt El Salvador Finland France Germany Great Britain Greece Guatemala Hong Kong India Indonesia Ireland Italy Jamaica Japan Korea Lebanon Libya Malaysia MauritiusMexico Neal New Zealand Nicaragua Norway PakistanPanama Peru Philippines Poland Portugal Puerto Rico Romania Russia Saudi Arabia Singapore Slovenia South Africa Spain Switzerland Taiwan Thailand the Netherlands Trinidad Tunisia | Terminated | July 15, 2014 | 368 |  | Not_Computer
Z-0002-2014 | 66189 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King LTS-D Oropharyngeal Airways |  |  | December 13, 0012 | 12 | King Systems Corp. | On December 13, 2012, King Systems initiated a voluntary recall of one (1) case(s) (5) individual products) from one (1) lot [(IV1V3) 62 cases, 310 products]of King LTSD Oropharyngeal Airways because the package label indicated that the airway was a size 5 when in fact the airway in the package was a size 4. When King manufacturers a lot, all of the components are the same size, part number etc. I | Error in labeling | King Systems notified the only one (1) consignee who received the affected product of the problem via a customer complaint. Five (5) of the five (5) affected products were destroyed by consignee. |  | 6 | 1 case; 5 units | USA Nationwide Distribution in the state of NY | Terminated | November 08, 2013 | 730815 | 6 | Not_Computer
Z-0004-2014 | 66357 | No | MXE |  |  | General &amp; Plastic Surgery | 510(k) | SHOPKO Waterproof Adhesive Pads Antibacterial. |  |  | August 21, 2013 | 2013 | ASO, LLC | No antibacterial bandages were packed in boxes labeled as antibacterial. | Labeling Change Control | A letter was sent via e-mail and mail to the only customer on August 23, 2013 that a recall of two lots was in progress and to examine their inventory and stop distribution of the product. Product was to be either place in quarantine or destroyed. A credit for all returned product or by returned completed certificate of destruction. |  | 37 | 37 cases (72 retail boxes in each case) | Distributed in the states of ID, WI, IN, and NE. | Terminated | August 04, 2014 | 348 | 37 | Not_Computer
Z-0006-2014 | 66128 | Yes | BSK | Inflatable tracheal tube cuff. | Anesthesiology | Anesthesiology | 510(k) | H&H; Emergency Cricothyrotomy Kit (Part No. HHECT01/ HHECTK01, NSN 6515-01-573-0692), vacuum sealed and packaged in a 5 mil poly bags, labeled in part ***H &amp; H Associates, Inc. CAGE 1NNH6 www.gohandh.com Ordinary, VA 23131 (800) 326-5708***. |  |  | August 27, 2013 | 2013 | H &amp; H Associates | Medical device tubing used for endotracheal airway application may become weakened and rendered ineffective. | Process change control | H&H; Medical Corporation notified their customers by "Urgent Medical Device Recall" letter dated August 27, 2013 and advised that the cuff balloon wall may become weakened and rendered ineffective. The firm advised their customers to discontinue use and take steps to remove affected units from further use. The manufacturer will replace the current endotracheal airway and cuff within the kits identified with a new cuffed airway in its own sterile packaging. The recaller included two response forms for customers to return: The Acknowledgement of Recall Form is provided for those who can quarantine the Cricothyrotomy Kits indicated in the recall. The form confirms notification of this recall and discontinued use of affected devices and provides details regarding the accounting and quarantine of the product lots in question for return and correction. The Product Recall Replacement Part Request Form is only required when inventory quarantine is not feasible. This form acknowledges the receipt of this notification and the undersigned accepts the responsibility for affixing the replacement airways along with the instructions that H&H; Medical Corporation will provide. The undersigned also assumes the responsibility to inform users regarding the subject of this recall, requires the use of the replacement airways over the airways provided in the affected kits and requires the disposal of the airways within the kits upon opening. For questions regarding the corrective request, please contact Ms. Michelle Morgan at 804-642-3663 or at mmorgan@gohandh.com during the hours of 8:00 a.m.-4:30 p.m. Eastern Standard Time, Monday through Friday. |  | 6619 | 6,619 kits | Nationwide distribution. | Terminated | January 13, 2015 | 504 | 6619 | Not_Computer
Z-0003-2014 | 66026 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis ADROIT 6F Guiding Catheter |  |  | August 12, 2013 | 2013 | Cordis Corporation | Due to a labeling error related to the conversion of the inner diameter of the catheter from inches to millimeters. The English dimension is correct and is the primary dimension. | Error in labeling | Cordis sent an Urgent Medical Device Recall letter dated August 12, 2013 Acknowledgement Form to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to immediately identify and set aside all affected product in a manner that ensures the product will not be used. Customers were asked to review, complete, sign and return the enclosed Acknowledgement form directly to Cordis to the fax number on the form. For questions call the Customer Service number at 800-561-7683. |  | 1238 | 1238 units | Worldwide Distribution - USA Nationwide in the states of AZ, CA, FL, ID, LA, MA, MD, MI, MN, MO, NC, NH, NY, OH, OK, PA, SC, TX, VA, WI, and WV and the countries of: Austria, Belgium, Czech Republic, France, Germany, Greece,Hungary, Israel, Italy, Latvia, Netherlands, Poland, Portugal, Russian Federation, Slovenia, Spain, Sweden, Switzerland, United Arab, and United Kingdom | Terminated | September 30, 2014 | 414 | 1238 | Computer
Z-2319-2013 | 66347 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | DeVilbiss Vacu-Aide QSU Suction Unit (Disposable Suction Container/Filter), 800 ML DISP Container, Unassembled, 800 ML DISP Container w/6' patient tubing, and 800 ML DISP Container w/6' patient tubing, assembled. |  |  | March 07, 2013 | 2013 | DeVilbiss Healthcare LLC | End users of the DeVilbiss Disposable Suction Container/Filter had exposed the bacterial filter cartridge to fluid, thereby occluding the filter and rendering it unusable. | Use error | DeVilbiss Healthcare sent End users a field corrective action letter dated March 7, 2013, to all affected customers to reinforce the proper cleaning method.This letter is to inform you that we have come across cases where end users had exposed the bacterial filter cartridge in the disposable suction container to fluid, thereby occluding the filter and rendering it unusableSpecifically what do you have to do when you receive a correction/recall notice?" The US FDA states in its correction/recall regulations that: "Consignees that receive a recall communication should immediately carry out the instructions set forth by the recalling firm and, where necessary, extend the recall to its consignees.""Enclosed you will find two Urgent Medical Device Correction letters addressed to DeVilbiss Dealer/Provider" and "DeVilbiss User" that we ask you read carefully."It is important that you follow the actions outlined in these letters to ensure we can effectively disseminate this information to end-users."To reach your direct accounts we ask that you send a copy of the "DeVilbiss User" letter to each of these end-users.We sincerely appreciate all of your efforts to ensure the safety of your customers that depend on DeVilbiss products. We thank you in advance for your assistance and apologize for any inconvenience that you may experience as a result of this situation. Should you have any questions or concerns, please contact DeVilbiss Healthcare Customer Service at 1-800-338-1988. |  | 112849 | 112,849 | Worldwide Distribution - USA and the countries of Mexico and Canada. | Terminated | November 06, 2014 | 609 | 112849 | Not_Computer
Z-0005-2014 | 64454 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The product is labeled in part "***Surgical Laser Technologies, Inc. Montgomeryville, PA***SLT Fiber Delivery System***Sterile***Non-Pyrogenic***. |  |  | January 04, 2013 | 2013 | PhotoMedex, Inc. | Firm became aware that some products within the specified lot may not have been sealed appropriately, thus compromising the sterile barrier. | Process change control | The firm initiated their recall of these products beginning on January 4, 2013 by sending notification letters to their consignees via electronic mail and/or facsimile. Customers were asked to return the affected product for a replacement. |  | 549 | 549 | Worldwide Distribution-USA including the states of AR, KS and OH, and the countries of Germany, South Africa and the United Kingdom. | Terminated | April 16, 2014 | 467 | 549 | Not_Computer
Z-0001-2014 | 66365 | No | OFA | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | Enforcement Discretion | Ultima OPCAB System, Sterile, Rx Only, |  |  | September 20, 2013 | 2013 | Maquet Cardiovascular, LLC | MAQUET has identified a potential issue with the ULTIMA OPCAB (part numbers OM-2003S &amp; OM-2001D). Labeling for the 2 blade lots had been incorrectly applied. The standard blades from Lot #25052176 were incorrectly labeled as deep blades, and the 20 deep blades from Lot #25052200 were incorrectly labeled as standard blades. | Error in labeling | MAQUET sent an "Urgent Medical Device Recall" letter and "Field Recall Response Forms" dated September 20, 2013 to customers via Federal Express. The notification identifies the affected product; informs the customers of the nature of the issue; and advises of potential adverse patient consequences that may occur. The customer is asked to complete and return the Field Recall Response Form. If the customer has questions they are directed to contact their field representative or MAQUET Technical Support (888.627.8383 between 8 am and 5 pm EST Monday through Friday). |  | 20 | 20 units (Lot 25052176); 12 units (Lot #25052200) | USA Nationwide Distribution in the states of: Florida, Maryland, New Jersey, New York and Texas. | Terminated | June 10, 2014 | 263 | 20 | Not_Computer
Z-0007-2014 | 66374 | Yes | NIP |  |  | Cardiovascular | PMA | Bard LifeStent Solo Vascular Stent, Item Numbers: EX062001CL, EX072001CL, EX062003CL, EX072003CL. |  |  | September 30, 2013 | 2013 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular is recalling specific lots of Bard Lifestent Solo Vascular Stent because they have determined a higher incidence rate of partial stent deployment associated with their use. | Device Design | The firm, Bard Peripheral Vascular, sent a "Urgent: Medical Device Recall Notification" letter dated September 30, 2013 via FedEx to all the customers who received the specific lots on recall. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to fill out the Recall and Effectiveness Check Form and fax to: BPV at 1-800-994-6772, even if you no longer have possession of the recalled product; call Bard Peripheral Vascular's Recall Coordinator at (800) 321-4254 Option #2 Ext 2727 (M-F 7am to 4pm MST) or email silvia.carrillo@crbard.com. The Recall Coordinator will issue you either a Return Authorization (RCL) Number or Consignment Recall Number (CRC) to facilitate the expedient return of the product; and to return the product to Bard Peripheral Vascular, Inc.,1415 W. 3rd Street,Tempe, AZ 85281. BPV will provide replacement product for your returned product.If you have any questions, call 1-800-321-4254 Option #2 Ext. 2727. |  | 4112 | 4112 units | US Nationwide Distribution. | Terminated | February 05, 2014 | 128 | 4112 | Not_Computer
Z-0008-2014 | 66025 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | Draeger Fabius GS Premium, Fabius OS, Fabius Tiro, Fabius Tiro D-M |  |  | August 14, 2013 | 2013 | Draeger Medical, Inc. | During final testing of the Draeger Fabius GS Premium, Fabius OS, and Fabius Tiro anesthesia machines, some units did not pass the high voltage test and the minimum clearance between an electrical component and the unit housing was not maintained for some from a specific batch. | Device Design | The firm, Drager, sent an "Urgent Medical Device Recall" letter dated August 2013 to its customers. The letter described the product, problem and actions to be taken. The letter informed the customers that in the event of a failure, manual ventilation mode should be used and that Draeger reps would be contacting customers to schedule the replacement of the power supply. The customers were instructed to notify personnel accordingly.If you have any questions, contact Director, Quality at 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your Draeger anesthesia machine please contact DragerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 1127 | 1127 units | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Azerbaijan, Bangladesh, Brazil, Bulgaria, Canada, Chile, China, Columbia, Ecuador, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Israel, Italy, Ivory Coast, Japan, Kazakhstan, Latvia, Libya, LithuaniaMalaysia, Morocco, Myanmar, Nepal, Philippines, Poland, Rumania, Russian Fed., Saudi Arabia, Senegal, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom, Uruguay, Vietnam, White RussiaTurkmenistan. | Terminated | September 23, 2014 | 405 | 1127 | Not_Computer
Z-0009-2014 | 66381 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aisys, Amingo, Avance, and Avance CS2. |  |  | September 16, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the Avance, Avance CS2, and Aisys anesthesia devices.While operating in Pressure Control Ventilation  Volume Guarantee (PCV-VG) mode, the affected products may produce an over delivery tidal volume when given a unique sequence of user inputs and a collapsed bellows (partially filled circuit).Exposure to excessiv | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated September 17, 2013. The letter was addressed to Chief of Anesthesia, Health Care Administrator / Risk Manager, and Director of Biomedical / Clinical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (US call 1 800-345-2700, for other countries contact your local GE Healthcare Service Representative). |  | 10263 | 10,263 (3,194 US, 7,069 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico except DE, ME, NM, RI, and VT, and the countries of UNITED ARAB EMIRATES, JAPAN, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACEDONIA , MALAYSIA, MEXICO, MOLDOVA, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERATION,SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TANZANIA, THAILAND, TURKEY, UNITED KINGDOM, VIETNAM, URUGUAY, and VENEZUELA. | Terminated | October 01, 2015 | 745 | 10263 | Not_Computer
Z-0010-2014 | 66250 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Avance, Avance CS2, Aisys, Aespire View, and Engstrom. |  |  | August 28, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving unresponsive buttons on the display of the EngstrÂżm ventilator and the Aespire View, Aisys, and Avance anesthesia machines.As a result of a manufacturing issue, the buttons on the left, right, and bottom keypads may not always detect button presses. This may result in the inability to access certain menu functions wh | Nonconforming Material/Component | Consignees were sent on 8/28/2013 a GE Healthcare "Urgent Medical device Correction" letter dated August 28, 2013. The letter was addressed to Chief of Anesthesia, Health Care Administrator / Risk Manager and director of Biomedical / clinical engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (1-800-345-2700 US). |  | 911 | 911 (252 US, 659 OUS) | Worldwide Distribution-USA (nationwide) including the states of US: AK, AZ, CA, CO, DE, FL, GA, IL, KY, LA, MD, MI, MN, MS, MT, NV, NY, NC, OH, OR, PA, SC, TX, WY, and WI and the countries of AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COST ARICA, CROATIA, CZECH REPUBLIC, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, KAZAHHSTAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, LITHUANIA, MAURITITUS, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NIGERIA, PAKISTAN, PERU, PHILIPPINES, POLAND, RUSSIAN FEDERATION, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VIETNAM, and VENEZUELA. | Terminated | January 30, 2015 | 520 | 911 | Not_Computer
Z-0011-2014 | 65898 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | ADVIA 2120/2120i Hematology Systems with ADVIA Autoslide connected to it, SMN 10361162 and SMN 10361798 --- CLASSIFICATION NAME: Counter, Differential Cell, Automated (Particle Counter) |  |  | August 12, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens determined that if an autosampler rack jam error occurs during operation of an ADVIA 2120/2120i connected to an ADVIA Autoslide, it is possible that the next slide processed by the Autoslide could be labeled with the wrong sample identification information. | Software design | Siemens sent an Urgent Medical Device Correction Letter dated August 8, 2013, and Effectiveness Checks Forms were sent to customers on August 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised if they do not have an ADVIA Autoslide connected to their system, no action is required. To avoid the possibility of mislabeling a slide when an ADVIA 2120/2120i rack jam error occurs: Eject the sample rack. Troubleshoot and correct the reason for the rack jam error. Rerun the sample rack. Field service personnel were sent a support bulletin describing the issue and instructing them how to deal with customer questions.If you need additional assistance, please contact Global Product SupportTarrytown, New York @ osggps.hschematology.healthcare@siemens.com |  | 311 | 311 units total (US: 51 units; Foreign: 260 units) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Belgium, Canada, Croatia, Denmark, France, Germany, Greece, Italy, Japan, Malaysia, Netherlands, Norway, Philippines, Portugal, South Korea, Singapore, Slovenia, Spain, Switzerland, Thailand, Turkey, United Kingdom and Vietnam. | Terminated | November 18, 2015 | 828 | 311 | Not_Computer
Z-0019-2014 | 66237 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-Slice CT Scanner System. |  |  | July 22, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | A pinch point at the end of the table top can cause injury. | Use error | Field Safety Notice (FSN 72800601_72800603, will be sent to all US consignees via certified mail. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 326 | 326 | Worldwide Distribution-USA including the states of AL, AR, CA, FL, GA, IL, IN, KY, MS, NC, NJ, NV, NY, OH, PA, and VA; and countries of: Albania, Algeria, Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Cambodia, China, Ghana, Greece, India, Indonesia, Ireland, Italy, Jordan, Latvia, Lebanon, Libya, Lithuania, Malawi, Mexico, Nepal, Netherlands, Oman, Palestine, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Switzerland, Thailand, Turkey, Uganda, Ukraine, Uzbekistan, Venezuela and Vietnam. | Terminated | September 25, 2014 | 430 | 326 | Not_Computer
Z-0012-2014 | 66277 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Precision 500D, Legacy, RFX, and SFX X. |  |  | August 30, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the spotfilm device power assist handle.There is a potential for the internal bolts on the power assist handle to be loose or to loosen and/or fall out. This issue, if present, can result in a loss of calibration and offset of the power assist handle and/or, under extreme condition, sticking of the handle in the ON p | Employee error | GE Healthcare sent an "Urgent Medical Device Correction" letter dated August 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were requested to inspect your power assist handle for either of the two conditions listed below. Immediately contact your local GE Healthcare Service Representative if any issue is found.1. If you observe that the power assist handle plastic grip is loose relative tothe mount, discontinue use.2. If plastic grip is secure, but you observe erratic performance, (power assistunresponsive, drifting of spotfilm device with no command, or lock releasealways ON), you may continue to use the device with care. Use all safetyprecautions relative to application of the myelographic stop and the inhibit,table horizontal stop OFF.GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following phone numbers:United States: 800 437 1171Japan: 0120 055 919For other countries, please contact your local GE Healthcare Service Representative. |  | 810 | 810 (783 US, 27 OUS) | Worldwide Distribution - USA (nationwided) including DC., and the countries of AUSTRALIA, CANADA, CHILE, GERMANY, CHANA, GUATEMALA, ICELAND, ISRAEL, ITALY, JAMAICA, JAPAN, RUSIAN FEDERATION, UNITED ARAB EMIRATES and UNITED KINGDOM. | Terminated | May 22, 2014 | 265 | 810 | Not_Computer
Z-0024-2014 | 65958 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE PET Discovery 610, PET Discovery 710. |  |  | March 26, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue related with the DMPR function and Screen Saved Images on your scanner which involves Optima CT660, Brivo CT385, Discovery PET/CT 610, Discovery PET/CT 710. The DMPR and Screen Save functionality may possibly generate the sameSeries UID. It has been seen that some PACS may combine series under a patient using only the Series UID | Software design (manufacturing process) | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated March 26, 2013. The letter was addressed to Director of Radiology, Emergency Medicine Physicians and Radiologic Technologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. |  | 28 | 28 (4 US &amp; 24OUS) | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, CO, FL, GA, IL, IA, IS, LA, MA, MO, NH, NM, NY, OH, OK PA, SC, TX, UT, VA, WA, and WI, and the countries of UNITED KINGDOM, DENMARK, CHILE, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, CANADA, CZECH REPUBLIC, ECUADOR, EGYPT, ETHIOPIA, FINLAND, FRANCE, FRANCE GUIANA, GERMANY, GUADELOUPE, GUATEMALA, GUYANA, INDIA, INDONESIA, IRAQ, ITALY, JAPAN, JORDAN, KAZAHSTAN, KOREA, KUWAIT, LATVIA, MYANMA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, POLAND, ROMANIA, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWEDEN SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, UZBEKISTAN, and VIETNAM. | Terminated | December 05, 2013 | 254 | 408 | Not_Computer
Z-0014-2014 | 65955 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes Midfoot Fusion Bolt 6.5 mm. It is indicated for fracture, fixation, osteotomies, nonunions, and fusions of large bones in the foot and ankle. |  |  | March 08, 2013 | 2013 | Synthes USA HQ, Inc. | A post market safety review following several customer complaints indicated additional clarification was necessary in the Midfoot Fusion Bolt Technique Guide to reduce potential for bolt back-out or migration by emphasizing the need for supplemental fixation. | Labeling design | An "Urgent Notice: Medical Device Recall" letter dated March 8, 2013 was sent to customers asking them to check their inventory, remove all affected product technical guides from use, and return them to Synthes along with the Verification form. Customer questions were directed to (610) 719-5450. |  | 62 | 62 | Nationwide Distribution, including the states of MI, PA, SC, IA, MA, CA, WI, IL, and KS. | Terminated | August 12, 2015 | 887 | 62 | Not_Computer
Z-0017-2014 | 64808 | No | JGQ | Total protein test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista (R) Urinary/Cerebrospinal Fluid Protein (UCFP) Flex (R) Reagent Cartridge Lot 11349BB |  |  | December 19, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed an increase on absorbance errors when calibrating Dimension Vista Urinary/Cerebrospinal Fluid Protein (U CFP) Flex Reagent Cartridge lot 11349 BB. | Process control | The firn, Siemens, sent an "Urgent Medical Device Correction" letter (dated December 2012) was provided to all Dimension Vista (R) UCFP customers who have received UCFP lot 11349BB. The letter described the product, problem and actions to be taken. The customers are instructed to discontinue use and discard any remaining inventory of Dimension Vista UCFP lot 11349BB; provide instructions for product replacement; and complete and return the attacted REQUEST FOR REPLACEMENT FORM via fax to (302) 631-8467. Siemens will replace any un-used inventory of this lot, at no charge.If you have any questions, please contact your Siemens Technical Solutions Center. your local Siemens Technical Support Representative or 800-441-9250. |  | 3395 | 3395 | Worldwide distribution: US (nationwide) including sates of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, ID, KY, KS, LA, MA, MD, MI, MN, MO, MS, MT, PA, NC, NE, NH, NJ, NY, OH, PA, SC, TN, TX, UT, VA, VT, WA, WI, WV, and PR; and country of: UK. | Terminated | August 27, 2014 | 616 | 3395 | Not_Computer
Z-0020-2014 | 66047 | No | CEO | Phosphorus (inorganic) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Clinical Chemistry Phosphorus Reagent |  |  | August 12, 2013 | 2013 | Abbott Laboratories, Inc | Reagent lots 64736UN12, 03991UN13, 51611UN13, 74805UN12, and 03992UN13 have incorrect expiration dates assigned by Supplier. | Error in labeling | The firm, Abbott Laboratories, sent a "PRODUCT CORRECTION" letter dated August 12, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to update existing inventory and manually track the implanted lots with the corrected expiration dates listed in the letter; retain a copy of letter and if you have forwarded any of the lots listed to another laboratory, provide a copy of the letter to them.If you or any of your healthcare providers you serve have any questions regarding this information, U.S. Customers please contact Customer Service at 1-877-4ABBOTT (422-2688). Customers outside the U.S., please contact your local area customer service. |  | 7213 | 7213 | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Albania, Algeria, Argentina, Austria, Australia, Azerbaijan, Bangladesh, Bahamas, Belgium, Bosnia &amp; Herzegovina, Brazil, Bulgaria, Burkina, Canada, Cayman Islands, Chile, China, Columbia, Croatia, Curacao, Cyprus, Czech Republic, Denmark, Dominican Republic, El Salvador, Egypt, Estonia, Ethiopia, Finland, France, Germany, Greece, Guatemala, Hong Kong, India, Iraq, Ireland, Israel, Italy, Ivory Coast, Jamaica, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malaysia, Mali, Mauritania ,Mexico, Moldova, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Qatar, Republic of Yemen, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Tajikistan, Taiwan, Tanzania, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, Turks and Caicos, United Arab Emirates, United Kingdom, Uruguay and Vietnam. | Terminated | April 06, 2015 | 602 | 7213 | Not_Computer
Z-0018-2014 | 66000 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Getinge 46-Series Medical Washer-Disinfectors, Listing # E208196 |  |  | July 30, 2013 | 2013 | Getinge Disinfection | Getinge Disinfection AB initiated a voluntary field safety correction for the Getinge 46-Series electrically-heated Washer Disinfectors due to possible overheating. | Nonconforming Material/Component | Getinge USA, Inc. sent a Device Correction Notices letter dated July 26, 2013 to all affected customers via US Postal Certified Mail within the U.S. and Canada. The letter identified the affected product, problem and actions to be taken. The letter indicate that all customers will be contacted by a Getinge representative for guidance if any machine within the serial number range of affected product is within their possession. If you should have any questions or concerns regarding this field correction, please call 1-800-475-9040, press "1" and enter extension 5036. |  | 143 | 143 units (142 domestically &amp; 1 unit internationally) | USA Nationwide in the states of AL, AR, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY and internationally in Canada. | Terminated | May 11, 2015 | 650 | 143 | Not_Computer
Z-0025-2014 | 66171 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific COGNIS CRT-D, Cardiac Resynchronization Therapy Defibrillator, models N118, N119. The following models are not available in the US: models N106, N107, P106, P107. Sterilized using ethylene oxide. Boston Scientific, 4100 Hamline Avenue North, St. Paul, MN 55112-5798 USA. |  |  | August 29, 2013 | 2013 | Boston Scientific CRM Corp | The performance of a low voltage capacitor in a subset of COGNIS CRT-Ds and TELIGEN ICDs manufactured prior to December 2009, may be compromised over time, causing increased current drain that can lead to premature battery depletion. | Component design/selection | An Important Medical Device Information letter dated August 2013 was sent to implanting and following physicians. The letter provided a description of the issue, describes clinical implications, provides the rate of occurrence, and gives recommendations for monitoring patients. |  | 38500 | 38,500 total worldwide | Worldwide Distribution-USA (nationwide) including Guam and Puerto Ricoand the countries of Argentina, Aruba, Australia, Austria, Belgium, Bermuda, Brazil, Canada, Chile, Hong Kong, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Finland, France, Gabon, Germany, Greece, Hungary, Iran, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Malaysia, Mexico, Netherlands, Netherlands Antilles, New Zealand, Norway, Panama, Peru, Poland, Portugal, Saudi Arabia, Singapore, Si Lanka, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad, UK, United Arab Republic, and Venezuela. | Terminated | April 10, 2014 | 224 | 38500 | Not_Computer
Z-0022-2014 | 66268 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Harvest Graft Delivery System; Model Number: GDP-10 |  |  | September 11, 2013 | 2013 | Harvest Technologies Corporation | Potential for leur connector to leak or have cracks. | Device Design | Harvest Technologies sent a notification letter dated September 10, 2013, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached form to confirm the number of kits received, used, and intending to return. They are asked to return the kits with the form. For questions contact Michelle Sullivan at 508-732-7500. |  | 730 | 730 cases (3 kits per case) | Worldwide Distribution - USA Nationwide: TX, NJ, MI, AZ, GA, LA, ME, MD, AR, MO, VA, MN, GA, MA, NY, PA, CO, CA and in the countries of Australia, Brazil, China, Singapore, Vietnam, Dominican Republic. | Terminated | February 10, 2014 | 152 | 730 | Not_Computer
Z-0021-2014 | 66254 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Advantx, Innova 2000, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, Innova 4100IQ. |  |  | August 16, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue involving missing screws in the overhead Video Monitor Suspension of the Advantx, Innova 2000, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, Innova 4100IQ Cardiovascular X-ray imaging systems. | Process control | Consignees were sent on 8/12/2013 a GE Healthcare "Urgent Medical device Correction" letter dated August 12, 2013. The letter was addressed to Hospital Administrators / Risk Mangers, Mangers of Radiology /Cardiology /Surgery Department and Radiologists /Cardiologists /Surgeons.The letter described the Safety Issue, safety Instructions, Affected Product details, Product Correction and Contact Information (US 800-437-1171, Japan 0120 055 919). |  | 2184 | 2184 (1063 US. 1121 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, and except ND, SD and VT, and the countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELARUS, BELGIUM, BOLIVIA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, COSTA ARICA, CROATIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY,GREECE, GUADELOUPE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, INDIA, INDONESIA, IRAN,IRELAND, ITALY, JAMAICA, JAPAN,JORDAN, KAZKHSTAN, REPUBLIC OF KOREA, KUWAIT, LEBANON, LIBYA ARAB JAMHIRIYA, LITHUANIA, MALAYSIA MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SLOVAKIA, SPAIN, SWEDEN, SYRIA, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UKRAINE, UNITED KINGDOM, and VENEZUELA. | Terminated | May 30, 2014 | 287 | 2184 | Not_Computer
Z-0015-2014 | 66294 | No | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Product Name: OSOM C. Difficile Toxin A/B Test |  |  | September 16, 2013 | 2013 | Sekisui Diagnostics LLC | Sekisuki Diagnostics has issued a Product Correction for certain lot numbers of the OSOM C. difficile Toxin A/B Test Kit, because it has been determined that there is potential for false negative results. | Other | Sekisui Diagnostic sent an Customer Recall Notification letter dated September 16, 2013, via email and/or FedEx. The letter identified the product the problem and the action needed to be taken by the customer. The customer notification letter informs the customers that Sekisui Diagnostics has issued a Product Correction for certain lots of the OSOM C. difficile Toxin A/B Test Kit (P/N 173). The customer notification lists the reason for recall, instructions to the customer, product description, required actions, and contact information. Customer Notification Letter lists the specific lot numbers being recalled for that catalog number. For each of these two catalog numbers, there is one letter directed to a distributor and one letter directed to an end user. Customer letters will be sent by email and/or FedEx. The letters have REQUIRED ACTIONS outlined in a box. If no response is received within 30 days, there will be directed follow-up communication. For further questions please call ( 858) 777-2600. |  | 391 | 391 kits | Worldwide Distribution - US (nationwide) and the countries of UK, Saudi Arabia, Dubai, Kuwait, Poland, Spain, India and Romania | Terminated | May 05, 2014 | 231 | 391 | Not_Computer
Z-0013-2014 | 66296 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT Big Bore Oncology, Brilliance CT Big Bore Radiology and Brilliance CT 16 (Air) |  |  | August 30, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | There are artifacts appearing on the Tracker images during threshold-triggered bolus scans with protocols using a Rotation Time of 0.4 seconds. | Software design | Philips sent an "Urgent - Medical Device Correction" letter dated August 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem If you need any further information or Support concerning this issue, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada, contact the Customer Care Solutions Center (1 -800-722-9377, Option 5: Enter Site ID or follow the prompts).This notice has been reported to the appropriate Regulatory Agency.Philips apologizes for any inconveniences caused by this problem |  | 258 | 258 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AR, AZ, CA, CO, CT, GA, HI, IL, IN, KS, MA, MD, MI, MO, MS, NC, NH, NJ, NV, NY, OH, OK, OR,TN, TX, UT, VA, VT and WV., and the countries of Australia, Austria, Bangladesh, Belgium, Brazil, Chile, China, Denmark, Egypt, France, Germany, Honduras, India, Indonesia, Italy, Kazakhstan, Latvia, Lebanon, Libyan Arab Jamahiriya, Mexico, Netherlands, Norway, Palestine, Philippines, Poland, Russia, Saudi Arabia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom and Vietnam. | Terminated | May 14, 2014 | 257 | 258 | Not_Computer
Z-0100-2014 | 65722 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | pocH-100i; sysmex; Automated Hematology Analyzer; 7.3" x 13.8" x 18.1" / 30.8 lbs.; Sysmex Corporation; 1-5-1 Wakinohama-Kaigandori, Chuo-ku, Kobe 651-0073, Japan; Sysmex America, Inc; One Nelson C. White Pkwy, Mundeline, IL 60060, U.S.A. |  |  | May 15, 2013 | 2013 | Sysmex America Inc | The data printed on the thermal printer may omit a digit or decimal point in the results or sample identification number. | Nonconforming Material/Component | The firm, Sysmex, sent a "Product Notification-FIELD CORRECTION" letter dated May 2013 to all customers in the U.S. and Canada on May 15, 2013 by U.S. Mail. The letter describes the product, problem and actions to be taken. The customers will be contacted by a Sysmex Service Representative within the following 90 days to schedule an appointment to replace the Printed Circuit Board (PCB) on the internal printer. The repair will be performed during an on-site service visit. The customers were instructed to, until the repair can be made, use the thermal printer for reporting results; to compare the sample identification, values and decimal points on the printout with the data on the LCD screen for accuracy.If you have any questions concerning this information, please email the Sysmex Technical Assistance Center (TAC). Simply click on email TAC button on the home page of the CRC, complete the form that appears and then click the submit to send your question to TAC. Sending an email to TAC is recommended for non-urgent requests.For urgent request, call the Technical Assistance Center at 1-888-879-7639. |  | 193 | 193 units | Worldwide distribution: US (Nationwide) including Puerto Rico and states of: AK, AL, AR, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, SD, TN, TX, UT, VA, VT, WA, WI, and WY; and country of Canada. | Open, Classified |  |  | 193 | Not_Computer
Z-0027-2014 | 66346 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | P21x/5-1 MHz transducer, Part Number P07698-23 |  |  | September 11, 2013 | 2013 | SonoSite, Inc. | FUJIFILM SonoSite Inc. remove P21x/5-1 Transducer because the firm discovered the Biopsy Bracket Sensor is not working as intended. | Error in labeling | SonoSite started calling their consignees on September 13, 2013. For customers who could not be reached via phone, a letter Urgent -Medical Device Removal will be sent on September 17, 2013. Affected customers will be offered a replacement transducer (P07698-24) free of charge. Consignees in the U.S. can call 1-877-657-8118, Monday-Friday, 6:00 a.m. - 5:00 p.m. (PST) or send E-mail to service@sonosite.com. Customers outside the US can call 425-951-1200. |  | 10 | 10 units | Worldwide Distribution - USA Nationwide in the states of Florida, Georgia, Oklahoma, Pennsylvania, Tennessee, Texas, West Virginia, and in the country Japan and Italy. | Terminated | December 11, 2013 | 91 | 10 | Not_Computer
Z-0029-2014 | 66105 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Model BP7507 OPTI LION E-Plus Cassettes 25 Cassettes with Sample Probes, OPTIMedical |  |  | August 09, 2013 | 2013 | OPTI Medical Systems, Inc | OPTI LION E-Plus Cassettes were labeled with an incorrect barcode. The incorrect barcode could perfect the performance of one or more parameters as the product ages. Also a Customer complaint from a distributor reported that a customer was getting QC failures. | Error in labeling | The firm, OPTIMedical, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 9, 2013 via email and fax to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use of the product; return any unused product to OPTI Medical System; distributors outside the U.S. may destroy unused product and certify the instructions on the enclosed Recall Response Card; contact your distributor or OPTI Medical System Technical Support at 1-800-490-6784 or technicalsupport@optimedical.com to obtain authorization number and shipping instructions; and complete and return the Recall Response Card to technical support via fax to 1-770-510-4447; mail to: OPTI Medical Technical Support, OPTI Medical Systems, Inc., 235 Hembree Park Drive, Roswell, GA 30076 or email: technicalsupport@optimedical.com Please contact OPTI Medical Systems, Technical Support department at 1-800-490-6784 if you have questions related to this notice. |  | 9025 | 361 boxes of 25 cassettes each | Worldwide distribution: US (nationwide) including states of: CA, MA, MN, and VA; and countries of: China, Guatemala, India, Iraq, South Korea, and Untied Arab Emirates. | Terminated | December 12, 2013 | 125 | 9025 | Not_Computer
Z-0030-2014 | 66310 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Perkin Elmer 226 Sample Collection Device, Serial numbers: 06132903 - 06198900. |  |  | September 05, 2013 | 2013 | Perkin Elmer Health Sciences, Inc. | Some devices may be contaminated on the back side of the device with glue in the blood application area. | Material/Component Contamination | The firm, Perkin Elmer, sent an "URGENT - MEDICAL DEVICE RECALL" letter dated September 5, 2013 via email to its consignees. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately notify all submitting facilities and personnel that have received the PerkinElmer 226 Sample Collection Device of the potential glue contamination and to distribute the enclosed record of actions form to all submitting facilities/personnel and encourage them to perform the inspection activities as listed in the letter; and complete and return the enclosed Recall Response Form return it to PerkinElmer Health Sciences via fax to: 864-299-8797 or email to: Rick.Haines@perkinelmer.com.Please contact PerkinElmer Health Sciences Greenville facility Quality Manager for further information 864-299-8787 (ext 104). |  | 65984 | 65,984 devices | US distribution in OH | Terminated | July 17, 2014 | 315 | 65984 | Not_Computer
Z-0028-2014 | 66325 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BBL (tm) MGIT (tm) Mycobacteria Growth Indicator Tube, part # 245122, packaged in 100 ct cartons and labeled in part |  |  | September 09, 2013 | 2013 | Becton Dickinson &amp; Co. | An in-vitro diagnostic reagent may be labeled with an incorrect barcode, which when scanned on associated diagnostic instrument may cause a false negative result. | Error in labeling | BD sent an Urgent Product Recall letter dated September 2013 to customers identified as microbiology lab managers. The letter identified the affected product, problem and actions to be taken. The letter advised that referenced lot numbers may contain an overlabel barcode that may begin to lift, exposing the original incorrect barcode pictured in the notification. Customers were instructed to complete the attached form and fax to Regulatory Compliance 410-316-4258. For further assistance regarding replacements contact BD Customer Service Department at 1-800-675-0908 and all other inquiries contact BD Technical Services Department at 1-800-638-8663. |  | 187000 | 187,000 tubes | USA Nationwide Disatribution in the states of: AL, AR, AZ, CA, CO, CT, FL, GA, IL, KY, MA, ME, MN, MO, NC, ND, NJ, NV, NY, OH, PA, TX, UT, VA, WA, WI | Terminated | October 29, 2013 | 50 | 187000 | Not_Computer
Z-0034-2014 | 64787 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity Cardio Imaging 5.0 is a software only Picture Archiving and Communication System (PACS). |  |  | October 02, 2012 | 2012 | GE Healthcare It | GE Healthcare is aware of a potential safety risk where an image labeling situation exists in the Image Viewer software of the Centricity Cardio Imaging product, version 5.0 SP1. Multiple studies may be displayed side by side by selecting "Compare Study" from the Study List, or by opening prior studies in the Patient Folder. In these cases, the Viewer differentiates the studies wityh a label (Curr | Software design | GE Healthcare IT mailed an IMPORTANT PRODUCT INFORMATION letter dated October 2, 2012 to the only consignee of this product. The letter notified the customer of the issue, provided temporary instructions on how to use safely use the software until a new software version can be installed, and advises that a software correction will be made available in the near future. Customers who have any questions can contact their local GEHC-IT service representative or call GE Healthcare's Remote Online Center (ROC) Customer Care Center at 1-800-437-1171. |  | 1 | 1 unit | USA Nationwide Distribution in the state of OK | Open, Classified |  |  | 1 | Computer
Z-0033-2014 | 66312 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Philips MR Systems: Asset 0.5T, Apollo 0.5T, Infinion 1.5T, Eclipse/Polaris 1.5T, Panorama 0.6T |  |  | August 07, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | There is a potential for water to collect in the vent pipe elbow in the magnet venting system. Water may freeze, blocking the venting system. | Device Design | Philips Healthcare sent an Urgent Medical Device Correction letter via certified mail on August 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information with all of their staff who need to be aware of the contents of this communication. Philips Healthcare will perform any necessary corrective actions prior to the customer's next cryogen refill. Customers with questions were instructed to contact their Philips field service representative: Technical Support Line 1-800-722-9377.For questions regarding this recall call 1-800-722-9377. |  | 17885 | Total 511 (Asset 0.5T,: 35 units Apollo 0.5T: 15 units, Infinion 1.5T: 116 units, Eclipse/Polaris 1.5T: 290 units, Panorama 0.6T: 55 units) | Worldwide Distribution - US (nationwide) including FL, CO, KY, MS, PA, NC, NJ, TX and Puerto Rico and internationally to Italy. | Terminated | March 26, 2014 | 231 | 17885 | Not_Computer
Z-0035-2014 | 66424 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX 5.2 DB Server (running Oracle 10.1.0.4.0), medical imaging Picture Archiving and Communication System (PACS). |  |  | September 19, 2013 | 2013 | AGFA Corp. | System downtime and/or slow performance may result if software in use is an old version of IMPAX and is not at current supported levels. | Software in the Use Environment | An "URGENT FIELD SAFETY NOTICE" letter was sent via email to the consignees on 9/19/2013, Acknowledgement, via FAX-Back or email, that the information was received and understood has been requested from the consignees. The letter specifically describes the safety alert for the device, informs consignees of the problem and recommends the upgrades for IMPAX and Oracle. |  | 4 | 4 | Worldwide distrbution: US(nationwide) distribution including NY and OK, and country of Canada. | Terminated | October 24, 2013 | 35 | 4 | Computer
Z-0031-2014 | 65952 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | UNITRAC Pneumatic Retraction Arm. |  |  | July 16, 2013 | 2013 | Aesculap, Inc. | Changes were implemented to the UNITRAC instructions for use to ensure safe use and potentially extend the life of the product. This included recommending the use of Sterile Cover JG901 and new Quick Release Coupling RT02OR to eliminate the need to sterilize the UNITRAC and a functional check was added prior to use to ensure optimal functionality of the UNITRAC system. | Process change control | A letter dated July 16, 2013 was sent requesting users to review the new instructions for use in their entirety, sign the attached Hospital Acknowledgment form and return to Aesculap. An Aesculap representative would be in contact to provide an in-service regarding the process for set-up of the Sterile Cover and required Functional Check. Discontinuation of use and repair of the device was recommended if the joints did not move freely or if the set position drifts during the Functional Check. |  | 217 | 217 | Nationwide Distribution. | Terminated | August 26, 2014 | 406 | 217 | Not_Computer
Z-0032-2014 | 66293 | No | LOM |  |  | Microbiology | PMA | FSE Image Upgrade Kit, Catalog # 92019C |  |  | August 20, 2013 | 2013 | Bio-Rad Laboratories | The FSE Image Upgrade kit (part 92019C) is recalled due to issue related to the export of data from the EVOLIS analyzer to Laboratory Information Systems (LIS) when using the MONOLISA Anti-HBs EIAQuantitative Determination in EVOLIS APF version 6.0. | Software change control | BioRad has sent an Urgent Product Correction letter dated August 19, 2013, to customers who had the EVOLIS FSE Upgrade Image Kit installed on the EVOLIS at their site by a Bio-Rad Laboratories Field Service Engineer (FSE).This issue impacts laboratory only if you are running the MONOLISA Anti-HBs EIA Quantitative Determination method with EVOLIS APF version 6.0 and TimeLiner Data Management Software. If you use this Quantitative Determination method with EVOLIS APF version 6.0, please contact Bio-Rad Laboratories for additional information or assistance.Bio-Rad Laboratories is currently developing and qualifying a software update CD consisting of the EVOLIS APF V6.0b and Timeliner Data Management Software to correct the issue. The software update CD will be installed by Bio-Rad Field Service Engineers at each of the customer sites wheninstalling the EVOLIS APF V6.0b.Customers should contact Bio-Rad Laboratories at 1-800-224-6723, Option2 and then Option 3 for any questions about this product correction. |  | 23 | 23 kits | US Distribution including the states of FL, IA, ME, MO, MS, NE, NY, TN, TX, and WA | Terminated | November 13, 2013 | 85 | 23 | Computer
Z-0049-2014 | 66304 | No | LZS |  |  | Ophthalmic | PMA | STAR Excimer Laser System, Model STAR S4 IR, Catalog Number: 0030-4864, Product is manufactured and distributed by AMO Manufacturing USA, LLC |  |  | September 16, 2013 | 2013 | Abbott Medical Optics, Inc. | This corrective action will address an identified failure of the lamp and illumination system within the laser unit caused by an unexpected failure of a thyristor component on the AC Power Distribution Printed Circuit Board (PCB). | Nonconforming Material/Component | The firm, Abbott Medical Optics, Inc., (AMO), sent an "ADVISORY NOTICE and a COVER LETTER FOR ADVISORY NOTICE" dated September 16, 2013 to all affected customers. The letters identified the product, problem, and actions to be taken by the customers. The customers were instructed to complete and return the Star Notification Acknowledgement form via fax to 1-714-247-4530 within 3 business days or mail to Abbott Medical Optics, 510 Cottonwood Drive, MS 347118, Milpitas, CA 95035, and share this advisory notice with your staff.If you have any questions related to operation of STAR Excimer Laser System or documentation requirements associated with this letter, please call AMO Technical Support at 1-800-511-0911 and you will be directed to an AMO Representative. |  | 834 | 834 units in total | Worldwide distribution: US (nationwide) including Puerto Rico and countries of: Asia Pacific, Canada, Europe, Africa, Middle East, Japan, and Latin America. | Terminated | November 12, 2014 | 422 | 834 | Not_Computer
Z-0051-2014 | 65969 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Synergetics 25 gauge Soft Tip Cannula, packaged in 6 units per box. Catalog number: 30.08.25. |  |  | July 30, 2013 | 2013 | Synergetics Inc | One lot of 25 gauge Soft-Tip Disposable Cannulas are co-mingled with 23 gauge cannulas; the product is labeled as 25 gauge cannulas. | Labeling Change Control | The firm sent recall notification letters to their consignees on 7/30/13. |  | 48 | 48/6-cannula boxes | Nationwide Distribution-USA including the states of MO, CO, KY, TN, NC, PA, IL CA, NY, IN, MI, LA, NJ and OH. | Terminated | April 07, 2014 | 251 | 48 | Not_Computer
Z-0044-2014 | 66317 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100NX Cassettes |  |  | September 18, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the three lots (13F049, 13F050, 13F051) STERRAD 100NX Cassettes because ASP has determined that the leak indicator on the cassettes is effective for only 12 months of the cassette's labeled 15-month shelf-life. | Under Investigation by firm | Advanced Sterilization Products (ASP) sent an Urgent Medical Device Recall letter on September 18, 2013, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to examine their inventory and return all affected cassettes using the enclosed prepaid UPS return lave to Stericycle. Customers were asked to count their inventory and record the data on the enclosed Business Reply Card and Packing Slip. Customers were asked to provide the notice to anyone in their facility that needs to be informed. Customers were also asked to maintain a copy of the notice with the affected product. Customers with questions or suspected problems were instructed to contact Stericycle directly at 1-877-257-7119.For questions regarding this recall call 949-453-6400. |  | 2995 | 2,995 units | Nationwide Distribution | Terminated | January 13, 2014 | 117 | 2995 | Not_Computer
Z-0043-2014 | 66328 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 200 Sterilizer (Single Door), Part Number: 10201. |  |  | September 18, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD 200 Sterilization System because it may emit an odor or smell into the surrounding environment that is not typical to normal operating conditions. | Under Investigation by firm | A customer notification letter was sent on September 18, 2013, to all customers who received the STERRAD 200 Sterilization System. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to indicate if they have received a STERRAD 200 System and to report any complaints or suspected problems to Advanced Sterilization Products (ASP) Professional Services directly at (888) 783-7723 and select option #2.An alternate design for the oil mist filter is being developed and qualification of new vacuum pump oil is in process. This action is to modify the PM interval to service the filter and oil more frequently. This will serve as an immediate/interim solution to reduce incidence of odor/smell complaints until the new oil mist filter design and vacuum pump oil are released.A Customer Letter will be sent to all STERRAD 200 customers to notify them of the issue and the corrective actions we are taking to resolve this issue. |  | 1305 | 1,305 units total (654 units in US , 651 OUS ) | Worldwide Distribution - USA (nationwide) and the countries of EA (Europe Middle East and Asia), APAC (Asia Pacific), Japan and Canada | Terminated | April 11, 2014 | 205 | 1305 | Not_Computer
Z-0045-2014 | 66400 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Merit Medical Performa Angiographic Catheter 5F Multipack. Part numbers 7560-43 and 7560-G3. |  |  | September 13, 2013 | 2013 | Merit Medical Systems, Inc. | 4 lots of Performa Cardiac Multipacks contain a JR4.0 catheter with a slight variation in tip shape. | Process control | Consignees were notified on 9/13/13 via Urgent Product Recall Notice, email, and product retrieval forms. Actions required of customers:1. Please immediately determine if any of the devices are within your facility, quarantinethem and discontinue use.2. Ensure that all personnel to whom the devices were distributed are made aware of this field action.3. Contact your Merit Sales Representative in order to arrange product return and replacement or credit.4. If you have any questions concerning this communication, please don't hesitate to contact Guido Sandulli at 1-801-826-4046. |  | 200 | 200 | US Distribution including the states of CO, VA, and MS. | Terminated | January 03, 2014 | 112 | 200 | Not_Computer
Z-0050-2014 | 66361 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Solea 1.0 Laser Surgical instrument for use in general and plastic surgery and dermatology. |  |  | September 12, 2013 | 2013 | Convergent Dental | Specific error condition could lead to unintended laser emission during laser activation. | Software design | The following Notice of Correction letter was sent out to all customers via email on 9/12/13, and followed up with a direct phone call to confirm receipt and review the contents of the letter: Dear Dr. [Customer name],This email is to inform you of a pending notice of correction regarding your Solea system, serial #[Customer serial number].We have identified a specific error condition that could lead to unintended laser emission during laser activation. The software defect that caused the error condition has been identified and tested.Please do not use your Solea system until further notice. Resolution of this issue will require the installation of an updated release of the Solea software. A representative from Convergent Dental will be in touch with you to schedule the upgrade.Convergent Dental is committed to product quality and patient safety. If you have questions about this issue, please contact us. Software fix and testing:New software version (Release 2.0.4.0) was created by the software development group. Convergent Dental has implemented two software changes to eliminate this problem. The first is a defect fix to correct the handling of fatal errors, specifically laser communication timeout errors, so that these errors are not able to be cleared. This will prevent the user from being able to re-enter the workflow after a fatal error. The second is a defect fix to not set the laser state to disconnected when a laser timeout error condition occurs. This will ensure that the command to stop laser emission is always sent to the laser control module anytime the laser cutting state is exited due to error condition or the normal workflow event of stepping off the foot pedal.The software development team ran full Unit and Integration Tests on the new software build. The software verification activities included execution of the full verification test plan (Doc #790-00001), which has been revised to include specific defect |  | 5 | 5 | US distribution including states of: MA, CT, NY. | Terminated | November 15, 2013 | 64 | 5 | Computer
Z-0054-2014 | 65742 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Licox Brain Tissue Oxygen Monitoring; 1 x Complete Brain Probe Kit |  |  | June 11, 2013 | 2013 | Integra LifeSciences Corp. | Integra LifeSciences Corporation has determined that there is a possibility that some lots of Licox CMP Brain Tissue Oxygen and Temperature Catheter Kits contain a hex wrench that does not loosen the set screw on the drill bit. | Nonconforming Material/Component | Integra LifeSciences Corporation sent a recall letter/return response form on 6/11/2013 by traceable courier service, traceable e-mail or facsimile. |  | 21 | 21 kits | Worldwide distribution: US (nationwide) and countries of: AR, BR, CH, CL,CR. DE, ES, FI, FR, GB, IL, IT, MX, PA, PL, PT, RU, TW, and ZA. | Terminated | February 18, 2014 | 252 | 1663 | Not_Computer
Z-0052-2014 | 65812 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Acuson SC2000 volume imaging ultrasound systems at software version 3.0. |  |  | July 10, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | Software issue results in inaccurate 3D image may result in ablation in wrong location. | Device Design | Siemens sent an Customer Safety Advisory Notification letter to all affected customers on July 10, 2013.The letter identified the product the problem and the action needed to be taken by the customerLetters provide steps to mitigate the risk and inform users that a software fix will be installed as soon as it is available. Siemens Service Engineers will contact all locations to arrange installation once fix is available. All personnel who use the ACUSON should be notified of the issue.Update September 17, 2013 : After discussion with CDRH, Siemens revised their letter to state that procedures using the ACUSON SC2000 system with BioSense Webster CARTO 3 equipped with CARTOSOUND Module and SoundSTAR Catheter should not be performed until the users system has been updated with the software upgrade that resolves the issue.For further questions please call (650) 694-5398. |  | 26 | 26 units | Worldwide Distribution - US (nationwide) including the states of IL, PA, AZ, OR, OH, NY, CO, TX and CA., and the countries of France, Germany, Japan, Portugal, Saudi Arabia, Taiwan and Turkey. | Terminated | November 25, 2013 | 138 | 26 | Computer
Z-0064-2014 | 66248 | No | MBB | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | SmartSetÂż Endurance GMV, Gentamicin, REF 3105-040, Medium Viscosity Antibiotic Bone Cement, Sterile. |  |  | September 16, 2013 | 2013 | DePuy Orthopaedics, Inc. | The SmartSetÂż GMV Endurance Gentamicin Bone Cement, product ref: number 3105-040 is intended for use in indications that have not been cleared by the US FDA. The intended use stated in the IFU for the SmartSetÂż GMV Endurance Gentamicin Bone Cement is for the fixation of prostheses to living bone in arthroplasty procedures of joints in which infection by gentamicin-sensitive organisms is a potentia | No Marketing Application | The firm DePuy Synthes, sent an "URGENT INFORMATION-MEDICAL DEVICE RECALL NOTICE" dated September 16, 2013 to its consignees via E-mail, hand delivery or fax depending on the most appropriate delivery method for each consignees. The Urgent recall notice described the recalled product, explained the reason for the recall, and directed consignees to cease further distribution, complete the attached Reconciliation form and quarantine and return all affected products through the DePuy Orthopaedics Sales consultant, Fax 1-787-287-7681 or Email to vbaez1@its.jnj.com. Consignees were directed to maintain a copy of the notice and forward the notice if product was further distributed. Clinical questions should be addressed to DePuy Orthopaedics, Inc's Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m. - 5 p.m. EDT). Recall questions or concerns should be addressed to the Quality &amp; Compliance Supervisor, at 1-787- 272-1900 x3114 or vbaez1@its.jnj.com (MF; 8 am- 5 pm EST). |  | 16040 | 16,040 units | Distributed in Puerto Rico and US Virgin Islands. | Terminated | February 24, 2015 | 526 | 16040 | Not_Computer
Z-0093-2014 | 65897 | No | GCC | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PIPELINE LS ACCESS SYSTEM, 15 Dilator, 5 mm, Insulated |  |  | July 31, 2013 | 2013 | DePuy Spine, Inc. | Dilators and Holding Clip are being used for a secondary function of neuromonitoring, for which it is not 510(k) cleared. | Under Investigation by firm | DePuy Synthes sent an Urgent Voluntary Product Recall Notification letter dated July 30, 2013, to all affected consignees via Fed Ex overnight delivery.Consignees were informed that DePuy Spine, Inc. is voluntarily recalling all lots of the PIPELINE LS ACCESS SYSTEM Dilators (1 st, 2nd, 3rd, and 4th) and Holding Clip. They were asked to review inventory located at their facility and quarantine any products listed in this recall notice. Consignees were instructed to contact their local DePuy Synthes Spine Sales Consultant to return any inventory located within their facility.Consignees with questions were instructed to call 508-828-3583. |  | 118 | 118 units | Nationwide Distribution including the states of AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI, WV, and WY. | Terminated | June 29, 2015 | 698 | 623 | Not_Computer
Z-0053-2014 | 66394 | No | DIF | Clinical toxicology control material. | Toxicology | Toxicology | 510(K) Exempt | Alere TriageÂż TOX Drug Screen Control 1 |  |  | September 23, 2013 | 2013 | Alere San Diego, Inc. | Alere San Diego, Inc. initiated a voluntary recall of Alere Triage TOX Drug Screen Control 1 (PN 94413, Lot C2681A) because the CODE CHIP that was provided with this lot causes the control to fail when run in the QC Sample Mode using Alere Triage TOX Drug Screeen products. The Alere Triage TOX Drug Screen test devices and patient results are not affected. | Under Investigation by firm | Alere San Diego, Inc. sent an Urgent Medical Device Notice dated September 23, 2013, to all affected customers via FedEx and outside US (OUS) consignees on September 24, 2013, via email. The customer notification letter titled "URGENT MEDICAL DEVICE NOTICE", informed customers of the recall, product description with part number and lot code (Alere TriageÂż TOX Drug Screen Control 1 PN 94413 Lot C2681A), reason for recall and impact (The purpose of the letter was to inform customers that the CODE CHIP" that was provided with this lot causes the control to fail when run in the QC Sample Mode using Alere TriageÂż TOX Drug Screen products. Customers were instructed to discontinue use of and discard the CODE CHIP contained within the Alere Triage TOX Drug Screen Control 1 PN 94413 Lot C2681A kit box. Customers were also instructed to provide a copy of the notice to their customers and instruct them to contact Alere for a replacement code chip. Customers were asked to complete and fax the enclosed Verification Form to confirm their receipt of the notice. Customers with questions were instructed to contact:(Alere San Diego, Inc.9975 Summers Ridge RoadSan Diego, CA 92121, U.S.A.Phone: 877 308 8287FAX: 858 805 8457E-mail: Responses.ts@alere.com).For questions regarding this recall call 858-805-2000, ext 3015. |  | 38 | 38 | Worldwide Distribution - USA (nationwide) and countries of: South Africa and Canada. | Terminated | September 26, 2014 | 368 | 38 | Not_Computer
Z-0065-2014 | 66121 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | VARISOURCE Breast Template System for high-dose rate Brachytherapy. |  |  | August 09, 2013 | 2013 | Varian Medical Systems, Inc. | The recall was initiated after Varian Medical became aware the Lexan templates of the Breast Bridge Template System could become bent and turn opaque after steam sterilization. | Nonconforming Material/Component | Urgent Medical Device Correction, Urgent Field Safety Notice were mailed by FedEx tracked delivery to all consignees on August 12, 2013. |  | 3 | 3 sets | Worldwide distribution: US (nationwide) distribution in CA and in the country of India. | Terminated | April 10, 2014 | 244 | 3 | Not_Computer
Z-0066-2014 | 66341 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Brivo OEC715/785/865 C-Arm, Models: 715 Prime, 785 Essential, and 865 Plus. |  |  | September 18, 2013 | 2013 | GE Healthcare Beijing , Hangwei Medical Systems, Co. LTD | GE is recalling certain lots of GE Brivo OEC715/785/865 C-Arm due to the potential for the wires in the high voltage cable bundle to fail, resulting in loss of system functionality. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated September 18,2013, to all affected customers to inform them of the issue and provide instructions for immediate mitigation of the issue. Customers were instructed to ensure that all potential users in their facility were made aware of this safety notification and the recommended actions. Customers with questions were instructed to call 1-800-874-7378 option 8.For questions regarding this recall call 801-295-4952. |  | 29 | 29 | Worldwide Distribution - USA (nationwide) and Internationally to China, India, Japan, Malaysia, Australia, Ireland, France, Germany, Spain, Hungary, Belgium, and Switzerland. | Terminated | January 27, 2014 | 131 | 29 | Not_Computer
Z-0082-2014 | 66336 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | MAQUET Getinge Group CARDIOHELP Base Unit REF 70104.7999 REF 70104.8012 Manufactured for: MAQUET Cardiopulmonary AG Made in Germany. |  |  | September 12, 2013 | 2013 | Maquet Medical Systems USA | It has come to the attention of MAQUET that in certain rare instances, the CARDIOHELP-i device may display an error message with audible alarm, indicating "battery defective." The alarm will occur-upon start up or when the main power supply is unplugged from the main power source. The human machine interface touch screen of the CARDIOHELP-i device may gradually become black due to the human machine interface (HMI) back light shutting down. | Software Manufacturing/Software Deployment | MAQUET issued their customers "Urgent-Medical Device Field Correction" notices/Customer Fax Response forms dated September 16, 2013 via FedEx on September 16, 2013. The notification informs the customer of the issue with the product; how to identified affected product; and the action needed to be taken by the customer. The field correction includes service on the affected product. Authorized MAQUET Service Technicians will service the affected product to correct the issue. Customers were requested to complete and fax back (or email a scanned copy) of the enclosed response form. Customers could contact their local field representative or the MAQUET Customer Service line at 1.888.627.8383 (press option 2 followed by option 2) Monday through Friday, 8am to 6pm EST.For questions regarding this recall call 973-709-7000. |  | 63 | 63 | Nationwide Distribution including AZ, FL, CO, TX, OH, VA, MN, WI, SC, IL, ND, UT, CA, OR, NY, IA, CT, SD, MI, HI, AL, MD, KS, MO | Terminated | June 07, 2017 | 1364 | 63 | Not_Computer
Z-0086-2014 | 65918 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | HemoCueÂż Urine Albumin Microcuvettes, 25 Microcuvettes HemoCue AB Angelholm, Sweden |  |  | July 16, 2013 | 2013 | HemoCue AB | HemoCue AB has become aware of a production related problem involving punctured HemoCueÂż Urine Albumin Microcuvettes single pack pouches. | Packaging | HemoCue sent an Urgent Field Safety Notice dated July 24, 2013. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed to :1. Quarantine any HemoCueÂż Urine Albumin (110608); lot numbers 1208848 to 1304847 you may have in your possession2. Return the Field Safety Notice Verification Form to: fax: 562-668-5794We apologize for the inconvenience this advisory has caused you. Please feel free to contact Technical Support at 800-426-7256 with any questions. |  | 517800 | 10,356 boxes/50 microcuvettes | Worldwide Distribution - Nationwide Distribution including the states of AL, AR, AZ, CA, CA ,CT, FL, GA, HE, ID, IL, IN, LA, MA, ME, MI, MN, MO, MS, NC, NH, NM, NY, OH, PA, SC,TX, UT, VA, VT, WA and WI., and the countries of Australia, Bahrain, Belarus, Canada, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Hong Kong, India, Indonesia, Italy, Kenya, Lesotho, Malawi, Mauritius, Mexico, Myanmar, Nigeria, Norway, Omen, Poland, Qatar, Romania, South Africa, South Africa, South Korea, Spain, Sweden, Taiwan, Tanzania, Ukraine, Yemen and Zambia. | Terminated | April 28, 2014 | 286 | 517800 | Not_Computer
Z-0083-2014 | 66322 | No | MIH |  |  | Cardiovascular | PMA | Valiant Thoracic Stent Graft with Captivia Delivery System, |  |  | September 19, 2012 | 2012 | Medtronic Vascular, Inc. | Firm received reports of difficulty in flushing the graft cover of the delivery system during device preparation as required by IFU. | Component design/selection | A Medical Device Information letter was sent to all affected customers on September 19, 2012. Additionally, all Medtronic Field Representatives were notified of the issue via a Field Memo |  | 1000 | ~1000 | International distribution:Austria, Belgium, Cyprus, Czech Republic, Denmark, Finland, Germany, Greece, Hungary, Iceland, Ireland, It lay, Latvia, Lithuania, Luxembourg, Netherlands, Norway, Poland, Portugal, Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey, United Kingdom. | Terminated | October 25, 2013 | 401 | 1000 | Not_Computer
Z-0067-2014 | 66205 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF PET/CT |  |  | August 16, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Issues with the Ingenuity TF PET/CT Software Version 4.0 could pose a risk for patients in the event that a patient rescan were to require reinjection of PET radiopharmaceutical. | Software design | Philips Healthcare sent an Field Safety Notice dated August 16. 3013, via certified mail to all US customers. Market Representative will distribute the letter internationally. The letter identified the product the problem and the action needed to be taken by the customer.Field Safety Notice 88200460 is intended to inform you about;" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct the problemAdditionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order.If you need any further information or support concerning these issues, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada contact the Customer Care Solutions Center (1-800.722.9377, option 5: Enter Site ID or follow the prompts). |  | 13 | 13 | Worldwide Distribution - US (nationwide) including the states of NY, OH and Pa, and the countries of Australia, Austria, Egypt, France, Germany, India and Netherlands. | Terminated | March 26, 2014 | 222 | 13 | Computer
Z-0081-2014 | 66483 | No | DTK | Cardiovascular intravascular filter. | Cardiovascular | Cardiovascular | 510(k) | Cordis OPTEASE Vena Cava Filter |  |  | September 27, 2013 | 2013 | Cordis Corporation | Cordis identified a printing error on one unit of the Optease Vena Cava Filter, in which the orientation arrow for the femoral approach was printed in the incorrect direction. The error resulted in the filter being implanted upside down, requiring and additional percutaneous procedure to retrieve the filter. | Error in labeling | Cordis sent an Urgent Medical Device Recall letter dated October 8, 2013, to all affected customers. The letter identified the product (Cordis OPTEASEÂż Vena Cava Filter Catalog Numbers 466F220A and 466F220B), the problem, and the action to be taken by the customer. Customers were instructed to:1) Read this Urgent Medical Device Recall letter.2) Immediately identify and set aside all product listed below in a manner thatensures the affected product will not be used.3) Review, complete, sign and return the enclosed Acknowledgement Formdirectly to Cordis at the fax number on the form.4) Return any affected product with reference to the RA# on the form. ContactCustomer Service (see Available Assistance section) if assistance is needed for return of the affected product. Credit will be provided or consignment levels reduced.5) Share this letter with others in your facility that need to be made aware of thisrecall.6) Contact any other facility to arrange the return of OPTEASEÂż if any productlisted below has been forwarded to them.7) Maintain awareness of this notice until all affected product has been returnedto Cordis.8) Keep a copy of this notice with the affected product.Please contact the special Customer Service number (800) 551-7683,Monday through Friday from 8:30 AM to 6:30 PM ET.o Return instructions are provided on the Acknowledgment Form.Questions related to the Acknowledgement Form" Please contact Cordis QA at (786) 313-8730.For questions regarding this recall call 1-800-551-7683. |  | 29395 | 29,395 units | Worldwide Distribution - US (nationwide) including Hawaii and Puerto Rico and states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and countries of: Canada and Panama. | Terminated | September 23, 2014 | 361 | 29395 | Not_Computer
Z-0074-2014 | 65799 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂż Femoral Component (CR-FLEX GSF POR FEM, C-L; CR-FLEX GSF POR FEM, C-R; CR-FLEX GSF POR FEM, C-L MINUS; CR-FLEX GSF POR FEM, C-R MINUS; CR-FLEX GSF POR FEM, D-L; CR-FLEX GSF POR FEM, D-R |  |  | August 30, 2013 | 2013 | Zimmer, Inc. | The low density polyethylene (LDPE) bag used to package implants adheres to the highly polished surface of the devices. | Package design/selection | August 30, 2013 Urgent Recall Notices were sent to Distributors, hospital risk managers and surgeons identifying the issue and directing consignees to make all affected personnel aware of the contents of the notification. The notice identified the products involved, the issue, possible adverse events associated with the issue and recommended that implants found with LDPE bag adhesion not use the product and return it to Zimmer sales representatives. Questions or concerns should be addressed to Zimmer customer call center at 1-877-946-2761 between the hours of 8 a.m. and 5 p.m. EST. |  | 11694 | 11694 | Worldwide Distribution: US (nationwide) and countries of: Argentina, Australia, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Dominican Republic, Ecuador, El Salvador, Germany, Guatemala, Honduras, Hong Kong, India, Israel, Jamaica, Japan, Korea, Malaysia, Mexico, New Zealand, Nicaragua, Panama, Peru, Singapore, Taiwan, Thailand, and Venezula. | Terminated | September 25, 2014 | 391 | 204220 | Not_Computer
Z-0042-2014 | 66433 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Brand: The ForeCYTE Breast Health Test, Mammary Aspiration Specimen Cytology Test (MASCT), MASCT System Kits, Clarity System Kits, MASCT Patient Sample Kits &amp; Clarity Patient Sample Kits Instrument, Biopsy. |  |  | September 19, 2013 | 2013 | Atossa Genetics, Inc. | Atossa Genetics is recalling ForeCYTE Breast Health Test, Mammary Aspiration Specimen Cytology Test (MASCT) due to concerns raised by the U.S. Food and Drug Administration (FDA) in a warning letter received by Atossa in February 2013. The FDA raised concerns about (1) the current Instructions for Use (IFU); (2) certain promotional claims used to market these devices; and (3) the need for FDA cl | No Marketing Application | Atossa Genetics Inc. issued a press release to announce a voluntary recall on October 4, 2013.The firm started sending out the URGENT - VOLUNTARY PRODUCT RECALL NOTIFICATION ForeCYTE Breast Health Test; Mammary Aspiration Specimen Cytology Test (MASCT) RESPONSE REQUIRED letter to consignees on October 7, 2013. The letter identified the affected product, problem and actions to be taken. The firm initiated a voluntary recall to remove the ForeCYTE Breast Health Test and the Mammary Aspiration Specimen Cytology Test (MASCT) device from the market. The MASCT device has not been cleared by the FDA for the screening or diagnosis of breast cancer. In addition, the ForeCYTE Breast Health Test has not been cleared or approved by the FDA for any indication. The ForeCYTE Breast Health Test and the MASCT device are not a replacement for screening mammograms, diagnostic imaging tests, or biopsies. Patients should follow the recommendations and instructions of their physician with respect to breast cancer screening and diagnosis.Distributors are instructed to notify downstream customers and physicians are instructed to notify patients who have been tested using the MASCT and/or ForeCYTE systems. Consignees should return the Response Form back to Atossa.Distributors and customers should call 800-219-4629 to obtain a Returned Materials Authorization (RMA) number to return affected products. Customers with questions about this voluntary recall or Atossa's products, or would like assistance with the recall documentation may contact Atossa at 1-888-219-4629.Any problems experienced with the use of this product may be reported to the FDA at http://www.fda.gov/Safety/MedWatch/HowToReport/default.htm,Online (form available to fax or mail), or call FDA 1-800-FDA-1088. |  | 58277 | 577 MASCT System kits, 101 Lab Kits, and 2557 Patients Kits (5 test per kit) | USA Nationwide Distribution in the states of: AL, CA, FL, GA, IL, IN, KY, MD, ME, MI, MO, NC, ND, NJ, NV, NY, OK, PA, SC, TN, TX, UT, VA, WA, Washington DC, and Puerto Rico. | Terminated | August 19, 2014 | 334 | 58277 | Not_Computer
Z-0089-2014 | 66472 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT is a gamma camera for Single Photon Emission Computed Tomography (SPECT), one unit per box. |  |  | September 03, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Philips received reports from the field that the Flat Panel Detector (FPD) does not always remain securely latched in the stowed position. | Device Design | Philips Healthcare sent an Urgent - Medical Device Correction letter (Field Safety Notice (FSN 88200462), dated August 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the notice with all members of their staff who need to be aware of the contents of the letter. Customers were asked to retain a copy with the equipment instruction for use. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. A Field Service Engineer will visit the site to install the corrected software. The firm will be tracking the progress.Customers with questions were instructed to contact their local Philips representataive or local Philips Healthcare office. For North America and Canada contact the Customers Care Solutions Center at 1-800-722-9377, option 5. |  | 371 | 371 units | Worldwide Distribution - USA including AZ, CA, CO, DC, DE, FL, GA, HI, IA, IL, KY, MA, MD, MN, MO, MS, MT, NC, NJ, ND, NE, NV, NY, OH, OK, OR, PA, SD, TX, VT, WA, and WI.Internationally to Canada, Algeria, Argentina, Australia, Austria, Belgium, Brazil, China, Czech Republic, Denmark, Finland, France, Gabon, Germany, India, Indonesia, Ireland, Italy, Japan, Korea Republic, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Terminated | August 18, 2014 | 349 | 371 | Computer
Z-0087-2014 | 66290 | [6 empty] | chromID Salmonella Agar, Reference No. 43621. |  |  | September 06, 2013 | 2013 | bioMerieux, Inc. | Complaints for Salmonella strains growing on this medium producing an uncharacteristic white colored colony instead of the expected pale pink to mauve colored colonies. This could potentially lead to Salmonella colonies not being detected. | Nonconforming Material/Component | Recall letters were issued to the US customers by Biomeriux, Durham, NC on 9/6/2013. |  | 90 | 90 | Nationwide Distribution. | Terminated | June 24, 2014 | 291 | 90 | Not_Computer
Z-0092-2014 | 66352 | No | DZJ | Bone cutting instrument and accessories. | Dental | Dental | 510(k) | REF 5100-088, TPS Micro Driver, For Use with Stryker Total Performance System |  |  | October 02, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | It was discovered that the silicone potting failed to properly cure after undergoing the prescribed curing possess. The potting material is used to insulate, stabilize and provide a moisture barrier to the TPS MicroDriver main Printed Circuit Board Assembly (PCBA). Inadequate cure of the silicone potting could potentially allow it to leak to the outside of the handpiece or cause an electrical ma | Nonconforming Material/Component | Stryker sent an Urgent Medical Device Correction Notification letter via overnight mail and interntion consignees via E-mail on October 2, 2013. The Urgent Notice identified the recalled product, the reason for the recall, potential adverse risks to health associated with the recall. Customers were instructed to:1. Immediately review this Recall Notification2. Locate, and quarantine, the units listed in this notification. Do not use the recalled TPS MicroDriver until it has been repaired. A list, of units shipped to your facility, is included on the attached Business Reply Form.3. Return the enclosed Business Reply Form to confirm receipt of this notification and identify how many affected units are currently in your inventory.4. If you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF.5. Fax the completed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762.6. Upon receipt of the Business Reply Form, Stryker will send you a pre-paid shipper to send your affected device(s) back to Stryker to be repaired. Unit will be repaired and returned to your facility7. Send back all affected devices using the pre-paid shipper provided to you by Stryker.8. If you require a loaner, please contact the Stryker Service Department at 888-308-1983.For questions regarding this recall call 269-389-4518. |  | 67 | 67 | Worldwide distribution: US (nationwide) and countries of: Australia, Canada, France, Germany, Italy, Japan, Spain, Switzerland, and United Kingdom. | Terminated | April 15, 2014 | 195 | 67 | Not_Computer
Z-0088-2014 | 66434 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Si Surgical System IS3000 with the da Vinci Surgeon Console which contains a high Resolution Stereoscopic Viewer. |  |  | October 17, 2011 | 2011 | Intuitive Surgical, Inc. | Improper restraints during transportation of the da Vinci system could cause the HRSV Monitor to become loose from its mount. | Use error | Advisory letters were sent to all affected customers on 10/17/11 informing them to contact Customer Service prior to moving the system outside of the hospital. |  | 1090 | 1,090 consoles | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Austria, Belgium, Canada, Chile, China, Denmark, Egypt, Finland, France, Greece, Germany, India, Indonesia, Israel, Ireland, Italy, Kuwait, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Qatar, Romania, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | October 28, 2013 | 742 | 1090 | Not_Computer
Z-0084-2014 | 66379 | No | FCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | PENTAX Endoscopic Accessories. OF-B194. Made in Japan. |  |  | September 16, 2013 | 2013 | Pentax Medical Company | PENTAX Medical has become aware of a product issue associated with the OF-B194 Gas/Water Feeding Valve. A manufacturing defect may cause improper stoppage of CO2 gas during an endoscopic procedure with the potential for serious hazard to and/or user. | Use error | PENTAX Medical sent an "Urgent Device Recall" notification letter with /Response forms dated September 16, 2013, via UPS Certified delivery to all affected customers. . The notification identified the affected product; informed the customers of the nature of the issue; advised of potential adverse patient consequences that may occur. The customer was asked to immediately examine their inventory and quarantine product subject to the recall. The customer was instructed to identify and notify their customers if they further distributed the product. The notification instructs users to return the OF-B194 Valves for replacement with corrected product. Customers were asked to complete the attached response form and fax back to Pentax Medical at 1.201.799.4063. An enclosed UPS Airway Bill is to be used to return recalled product. A contact phone number and email address is provided if the customer has questions (1.800.431.5880 xt2064).For questions regarding this recall call 800-431-5880.Pentax issued a press release on 10/31/2013.Pentax posted their press release on their website at www.pentax.com. |  | 76 | 76 units | Nationwide Distribution including Arizona, California, Florida, Idaho, Indiana, Kansas, Maryland, Missouri, Missippi, North Carolina, New York, Pennsylvania, Virginia, Wisconsin and Wyoming. | Terminated | July 09, 2014 | 296 | 76 | Not_Computer
Z-0101-2014 | 65920 | No | GKR | Automated hemoglobin system. | Hematology | Hematology | 510(k) | HemoCueÂż Hb 201 Microcuvettes, carton/4 boxes/25 pouches of Microcuvettes, HemoCue AB Angelholm, Sweden |  |  | July 25, 2013 | 2013 | HemoCue AB | HemoCue AB has become aware of a production related problem involving punctured HemoCue Hb 201 Microcuvettes single pack pouches. | Packaging | On 8/13/2013 the firm sent Field Safety Notices to their customers. |  | 169320 | 42,330 packages/4 boxes/25 microcuvetes | Worldwide distribution which includes the following states: AK, AL, AS, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY,MD, ME, MI, MN, MO, NC, ND, NE, NJ, NM, NY, OH, OK, OR, SC, SD, TN, TX, UT, VA, WA, WI, VT, WV &amp; WYand the countries of Armenia, Australia, Austria, Belgium, Cyprus, Czech Republic, Denmark, El Salvador Finland, France, Germany, Great Britain, Guatemala, Hong Kong, Israel, Ivory Coast, Netherlands, New Zealand, Norway, Peru, Poland, Russia, South Africa, Sweden, Switzerland and Taiwan. | Terminated | August 11, 2014 | 382 | 169320 | Not_Computer
Z-0098-2014 | 66362 | No | KWI | Elbow joint radial (hemi-elbow) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | Small Bone Innovations (SBi) rHead Radial Stem, rHead Recon Radial System, rHead Radial Stem Lateral Assembly, rHead Radial Stem 6mm Collar, rHead Recon Radial Stem 6mm Collar, rHead Radial Stem Lateral Assembly 6mm Collar, rHead Standard Extended Stem, rHead Radial Stem Plasma Coated, rHead Recon Radial Stem Plasma Coated, rHead Radial Stem 6mm Collar Plasma Coated, rHead Recon Radial Stem 6mm Collar Plasma Coated, uHead Ulnar Stem Standard, uHead Ulnar Stem 20mm Collar packaged in a sterile pouch inside a shrink-wrapped box. |  |  | September 06, 2013 | 2013 | Small Bone Innovations, Inc. | The recall was initiated due to an increased risk of implants breaching the sterile pouches inside the shrink-wrapped box. | Packaging change control | Email correspondence and a notification letter, dated September 6, 2013, were sent to users requesting the immediate return of affected devices. Another notification letter, dated December 20, 2013 was also sent for recall of the additional devices after the recall was expanded. |  | 2036 | 2036 devices | Worldwide Distribution-USA (nationwide) and the countries of Canada and Mexico. | Terminated | March 30, 2015 | 570 | 2036 | Not_Computer
Z-0099-2014 | 66238 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter CLEARLINK System Non-DEHP Three Lead Extension Set. 6.0" (15 cm), 1.0 mL, Luer Activated Valve, male Luer Lock Adapter, Sterile, nonpyrogenic, Baxter Healthcare Corporation, Deerfield IL 60015. |  |  | September 04, 2013 | 2013 | Baxter Healthcare Corp. | The firm is issuing a voluntary recall due to the potential for pinholes in the packaging. Pinholes are a breach in the sterile barrier and could lead to potential contamination. | Under Investigation by firm | Baxter sent an Urgent Product Recall letter dated September 9, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken by Customer/User:1. Locate and remove all affected product from your facility. If you distribute these products to other facilities or departments within your institution, please forward a copy of this communication to ensure that they locate and remove affected product from potential use. The product code and lot number can be found on the shipping carton label and individual product package.2. Acknowledge your receipt of this recall notification by completing the attached Customer Reply Form (Attachment 1) and return it to Baxter by either faxing it to 224-270-5457 or scanning and emailing it to fca@baxter.com. Returning the Customer Reply Form promptly will prevent you from receiving repeat notifications.3. The affected product should be returned for credit by contacting Baxter Healthcare Center for Service. The Center for Service can be reached at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time. Please have your Baxter eight-digit ship-to account number ready when calling.Please have the product code and lot number(s) available. |  | 3200 | 3,200 units | US Distribution including the states of OK, GA, CA, MN, TX and HI. | Terminated | April 19, 2017 | 1323 | 3200 | Not_Computer
Z-0108-2014 | 66166 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens RAPIDLabÂż 1265 system blood gas analyzer |  |  | September 17, 2012 | 2012 | Siemens Healthcare Diagnostics Inc | RAPIDPointÂż 400 Series, RAPIDPointÂż 500, orRAPIDLabÂż 1200 Series blood gas analyzer, users can inadvertently enter a wrong patient ID on the demographics screen, which ultimately could lead to a misidentification of the sample | Software Manufacturing/Software Deployment | Siemens Healthcare sent an Urgent Field Safety Notification letter dated September 2013 via Federal Express, to affected Siemens domestic customers. Effected Siemens customers outside the US was provided a copy of the Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each countries local regulations and procedures. Previously Siemens Healthcare had issued a Customer Notification on August 23, 2012. The letter identified the affected product, issue, information to mitigate the issue and actions to be taken. The firm will issue an upgrades software to users. For questions contact your local Customer Care Center - Technical Support Representative at 1-877-229-3711. |  | 2939 | 2939 units | Worldwide Distribution - USA Nationwide and in the countries of Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belgium, Bosnia Herzeg., Botswana, Brazil, Brunei, Bulgaria, Chile, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Fren.Polynesia, Germany, Greece, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Ivory Coast, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lithuania, Luxembourg, Macedonia, Malaysia, Mayotte, Netherlands, New Caledonia, New Zealand, Norway, P.R. China, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Russian Fed., Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Turkmenistan, U.A.E., United Kingdom, Vatikancity, Venezuela, Vietnam, | Terminated | April 01, 2015 | 926 | 11006 | Computer
Z-0117-2014 | 66243 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Brand Names: IntegraÂż Camino 110-4 Intracranial Pressure (ICP) Monitoring |  |  | September 10, 2013 | 2013 | Integra Limited | Integra LifeSciences has identified through an internal evaluation that in order to comply with US FDArequirements, certain CaminoÂż Intracranial Pressure Monitoring Kits require additional labelling to indicatethey are UNSAFE in an MR environment.Only the following CaminoÂż Intracranial Pressure Monitoring Kits are affected:110-4B, 110-4BC, 110-4G, 110-4H, 110-4HC, 110-4L, 110-4LC. | Under Investigation by firm | Integra sent an Urgent Medical Device Correction Field Safety Notice dted September 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and determine if they have any of the affected product. If so, customers were asked to ensure that the information contained in the letter be distributed to the appropriate personnel. Complete the attached form return it by email or fax as indicated on the form. Customers with questions were instructed to contact Customer Service at 1-855-532-1723.For questions regarding this recall call 1-800-997-4868, outside of the US 888-980-7742. |  | 79887 | 79,887 | Nationwide Distribution | Terminated | June 04, 2014 | 267 | 79887 | Not_Computer
Z-0111-2014 | 66524 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Single Site Curved Needed Driver |  |  | October 02, 2013 | 2013 | Intuitive Surgical, Inc. | A limited number of FDA-cleared instruments for da Vinci Single Site use, Bipolar Maryland and curved needle Driver, were shipped without the correct revision of the user manual which provided added instruction for use of these additional Single Site instruments including indication, warnings and precautions. | Labeling Change Control | The firm, Intuitive Surgical, sent an " Urgent Medical Device Correction" letter dated October 2, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure all affected personnel are fully informed of this notice; after receiving the updated manual, locate and discard all previous versions of the Single Site Instruments manual and replace with the new manual; complete and return the attached Acknowledgement Form via FAX to Intuitive Surgical, Inc., ATTN: REGULATORY COMPLIANCE, U.S. (408) 716-3040, or Scan and Email to: isi.compliance@intusurg.com.; Inform affected personnel when the correction has been completed; and Retain a copy of this Notice and the Acknowledgement Form for your files.If you have any questions, contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below:" North and South America: 800-876-1310 Option 3 (6 am to 5 pm PST)" Japan: 0120-56-5635 or 003-5575-1362 (9 am to 6 pm JST)" South Korea: 02-3271-3200 (9 am to 6 pm KSTJ)" Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 am to 6 pm CET). |  | 1118880 | 504 devices for a total of 2,220 affected clinical product. | Worldwide distribution: US (nationwide) and countries of:Australia, Belgium, France, Italy and United Kingdom. . | Terminated | September 04, 2014 | 337 | 1118880 | Not_Computer
Z-0119-2014 | 66059 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Xario Diagnostic Ultrasound System SSA-660A |  |  | August 05, 2013 | 2013 | Toshiba American Medical Systems Inc | The firm initiated this recall of certain Toshiba Diagnostic Ultrasound Systems because it has been found that under certain conditions, the temperature of the transducer head may increase in 2D+PW mode or 2D+CDI (POWER, ADF) +PW mode. When the system is used under both conditions listed below, the temperature of the transducer head connected to the system may increase in temperature and possibly | Device Design | Toshiba sent an Urgent Medical Device Correction letter dated August 5, 2013. along with a Reply Form via USPS return receipt mail. The letter identified the product, the problem, and the action to be taken by the customer. Until the software is corrected customers were instructed to change the preset settings as described in the letter. Customers were asked to complete and return the attached form and fax it to the toll free number. The form also can be sent via email to raffairs@tams.comCustomers with questions were instructed to call 800-421-1968 or contact their local Toshiba Representative at 800-521-1968.For questions regarding this recall call 714-730-5000. |  | 17 | 17 (1352 Total) | Nationwide Distribution including and US Territory Puerto Rico | Open, Classified |  |  | 17 | Computer
Z-0112-2014 | 66144 | No | FDS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fujinon/Fujifilm Ultrasonic Endoscope |  |  | July 31, 2013 | 2013 | Fujinon Inc | A limited number of endoscopes were shipped to FMSU-ESD with a manufacturing defect which included missing set screws used to more securely attach the Forceps Inlet (biopsy inlet port) to the endoscope. | Unknown/Undetermined by firm | The firm, FUJIFILM Medical Systems, USA, Inc. (FMSU-ESD) sent an "URGENT-MEDICAL DEVICE CORRECTION (UMDC)" and Customer Acknowledgement forms dated July 31, 2013 to their customers. The letter described the issue with the affected product; instructed customers on how to identify affected units; instructed customers to immediately cease use of the affected units; return affected units to FMSU-ESD for repairs and quality control inspection; and complete and return the attached "Customers Acknowledgement" form to FMSU-ESD, Attn: Regulatory Compliance, via fax to (973) 686-2616 or email: gwalljasper@fujifilm.com.If you have any questions regarding this notice, please do not hesitate to contact Director, Quality and Regulatory Compliance, at (973) 686-2636 or by email: gwalljasper@fujifilm.com. |  | 12 | 12 Units | Worldwide distribution: US (nationwide) including states of: NJ, PA and VA; and country of Canada. | Terminated | March 04, 2014 | 216 | 12 | Not_Computer
Z-0109-2014 | 66222 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Vented Spike Adapter Product Code 2C0471, Sterile, nopyrogenic. Manufactured by an affiliate of Baxter Healthcare Corporation, Clintec Nutrition Division, Deerfield IL 60015 USA. |  |  | August 30, 2013 | 2013 | Baxter Healthcare Corp. | Potential tears in the pouch of six lots of Vented Spike Adapter Product Code 2C0471 | Packaging | The firm, Baxter, sent an "Urgent Product Recall" letter dated September 10, 2013 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to locate and remove all affected product from your facility; If you distributed the products to other facilities or departments, forward them a copy of the letter; complete and return the attached Customer Reply Form via fax to 1-224-270-5457 or scan and email to fca@baxter.com; and contact Baxter Healthcare Center for Service at 1-888-229-0001 between hours of 7:00 am and 6:00pm Central Time to return affected product and receiving credit.For additional information contact the Center for One Baxter at 1-800-422-9837, Monday thru Friday during the hours of 8:00am to 5:00pm Central Time. |  | 49344 | 49,344 units | Nationwide distribution. | Terminated | April 19, 2017 | 1328 | 49344 | Not_Computer
Z-0116-2014 | 65471 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals, including radiologists, physicians, technologists, clinicians and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display, manipulate, archive, print and export images when connected with the Centricity PACS infrastructure. |  |  | June 11, 2013 | 2013 | GE Healthcare It | GE has recently become aware of potential issues due to out of context issues associated with specific workflows of the Centricity PACS RA1000 workstation. These issues are as follows: 1) In Centricity RIS-IC, and a RIS-Driven integration, the "close" button closes the exam in RIS-IC, but does not close the exam in Centricity PACS RA1000 Workstation. 2) Interrupted workflow could break the | Software design | URGENT MEDICAL DEVICE CORRECTION letters dated June 11, 2013 was mailed to all customers alerting them to these potential safety issues. The software was modified to correct these defects and a GE Healthcare service representative will contact each customer to make arrangements for installation of the new software version at no charge to the customer. Customers with concerns or questions about the market action can contact their local GE Healtchcare IT service representative or GE Healthcare's Remote Online Center - Customer Care Center at 1-800-437-1171. |  | 623 | 623 units | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV) and the countries of the UK, Taiwan, Switerzland, Sweden, Spain, Singapore, Saudi Arabia, Netherlands, Malaysia, Korea, Japan, Italy, Israel, India, Hong Kong, Germany, France, Denmark, China, Canada, Bermuda, Beligum and Australia. | Open, Classified |  |  | 623 | Computer
Z-0115-2014 | 66410 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Breeze C Scooter is a motorized electric scooter for elderly people. Brand name Breeze, model C. It is intended to provide increased mobility for elderly and/or disabled person and it is also an aid to independent living. |  |  | September 27, 2013 | 2013 | S.A.E. AFIKIM | Breeze C Scooter, motorized 3-wheeled electric scooter for elderly people, is recalled because it may become unresponsive during operation due to an electrical contact which may fault due to extensive wear, which potentially could result in the driver not being able to stop the scooter during operation except by turning off the ignition key. | Process control | Afikim Electric Vehicles sent an URGENT: MEDICAL DEVICE RECALL letter, dated September 27, 2013, to all affected customers and distributors. .Customers are advised to remove the ignition key and contact Afikim Electric Vehicles via telephone Monday through Friday at 1-800-809-3010 between the hours of 9AM and 5PM (EST). Consumers may also contact the company via e-mail at dov@afiscooters.com.Customers are instructed to return the Recall Response Form by fax it to 800-466-1171 or email to ofer@afiscooters.com.Consumers should report any adverse event or quality problems to FDA's MedWatch Adverse Event Reporting program either by online, fax, or regular mail. |  | 69 | 69 units | Nationwide Distribution: US including the states of California, Florida, Georgia, Illinois, Maine, Maryland, Massachusetts, Michigan, Minnesota, New Jersey, New York, Ohio, Oklahoma, Pennsylvania, Texas, Virginia, and Wisconsin. | Terminated | July 23, 2014 | 299 | 69 | Not_Computer
Z-0118-2014 | 64730 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS IW by GE Healthcare Dynamic Imaging Solutions is a device that receives medical images (including mammograms) and data from various imaging sources. Images and data can be stored, communicated, processed and displayed within the system or across computer networks at distributed locations. Lossy compressed mammographic images and digitized film screen images may only be interpreted using an FDA approved monitor that offers at least 5 M pixel resolution and meets other technical specifications reviewed and accepted by the FDA. Typical users of this system are trained professionals, including but not limited to physicians, radiologists, nurses, medical technicians an assistants. |  |  | June 15, 2012 | 2012 | Ge Healthcare It | Multiple issues were identified with the GE Centricity PACS-IW system that could pose a safety risk to patients. These issues include:1) With Auto-Fetch in use another patient's image may appear without user's intention. 2.a) In the case where dose administration time and image acquisition time fall on either side of midnight, the value will be under reported and close to zero for all pixe | Software design | GE Healthcare sent an IMPORTANT PRODUCT INFORMATION letter dated June 15, 2012 to direct accounts informing them of the Auto-Fetch functionality issue and advising them to discontinue the use of the Auto-Fetch feature and to change the system setting to not allow users privileges to use Auto-Fetch. Customers were requested to contact their Field Engineer or GE Service Representative with any questions or to call the GE Healthcare Remote Online Center Customer Care Center at 1-800-437-1171. GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 12, 2013 to all affected customers informing them of all three of the safety issues. The letter again adivsed customers to discontinue the use of the Auto-Fetch feature and to change the system setting to not allow users privileges to use Auto-Fetch. The letter further infomed customers that all three errors can be resolved through the installation of a new version of the software. Customers were requested to contact their Field Engineer or GE Service Representative with any questions or to call the GE Healthcare Remote Online Center Customer Care Center at 1-800-437-1171. |  | 847 | 847 units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, District of Columbia, Puerto Rico, and the U.S. Virgin Islands.Internationally to Austria, Azerbaijan, Bahrain, Belgium, Bosnia, Brazil, Canada, Columbia, Egypt, France, Germany, Ghana, Hungary, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Libya, Lithuania, Mexico, Netherlands, Nigeria, Palestine, Poland, Qatar, Russia, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, United Arab Emirates, United Kingdom, Uzbekistan, and Yemen.Center Recommended Depth - Consumers/User** | Open, Classified |  |  | 847 | Computer
Z-0114-2014 | 65997 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Synthes Application Instrument Sternal Zip Flex System |  |  | August 13, 2013 | 2013 | Synthes USA HQ, Inc. | When using the 1st generation Zip Fix Application Instrument it is possible to cut the implant while the tensioning trigger is being squeezed which could lead to loosening of the implant. | Labeling False and Misleading | Synthes sent a Urgent Medical Device Notification dated August 13, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action to be taken: This notification is to make you aware of the issue and the importance of completely releasing the tensioning trigger before cutting the implant. Please take the following action:"To ensure your receipt and understanding of this information, Please complete the verification section of this letter and return to Synthes by either of the following methods:" Fax: 1-888-352-9616 or" Scan/email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. |  | 319 | 319 | Worldwide Distribution - US (nationwide) and country of Mexico. | Open, Classified |  |  | 319 | Not_Computer
Z-0123-2014 | 66266 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | Globus Medical MicroFuse Putty, 2.5cc and 10cc. |  |  | August 28, 2013 | 2013 | Globus Medical, Inc. | The sterility of this product to a Sterility Assurance Level (SAL) of 10(-6) cannot be assured. | Process control | Globus sent an Urgent: Medical Device Recall Notification letter, dated August 28, 2013 to customers/users. The letter identified the affected units, issue, potential risk, product, and actions to be taken. Customers were requested to complete the attached response card and return of the affected product to Globus Medical. For questions and support contact Globus Medical by fax 1-610-300-1342 or email:recall@globusmedical.com. |  | 193 | 193 | USA Nationwide Distribution. | Terminated | November 14, 2014 | 443 | 193 | Not_Computer
Z-0126-2014 | 65469 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX-Revolution Mobile X-ray System. |  |  | June 06, 2013 | 2013 | Carestream Health, Inc. | Carestream Health Inc. is recalling the power plug on their DRX-Revolution Mobile X-Ray System due to the item generating sparks and becoming warm to touch. | Device Design | On June 6, 2013 Carestream Health Inc. initiated a recall for the power plug on their DRX-Revolution Mobile X-Ray System due to the item generating sparks and becoming warm to touch. Carestream Health Inc. will replace the power plug in all potentially impacted DRX-Revolution Mobile X-Ray Systems. Your Carestream service representative will contact you to answer any questions and schedule a visit to replace the plug of your Mobile X-Ray System as soon as possible. If you have any questions or concerns, or if you experience any issue with your plug prior to our visit, please contact the Carestream Customer Care Center at 1-800-328-2910, 8AM through 8pm Eastern Standard Time. |  | 180 | 180 units | Nationwide Distribution including the states of NY, TX, PA, OH, WI, CA, IL, CO, FL, IN, OK, NJ, GA, LA, AL, KS, AZ, MA, SC, MI, NC, IA, UT, VA, TN, and MN. | Open, Classified |  |  | 180 | Not_Computer
Z-0090-2014 | 65805 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | DRX Revolution Mobile X-Ray System, Model DRXR-1, Verona Street Rochester, NY 14608 |  |  | July 26, 2013 | 2013 | Carestream Health, Inc. | During an audit, Carestream discovered a DRX-1 System that captured an image which displayed an artifact. This issue was discovered in manufacturing. It occurs when the DRX product console software is updated to v5.6b software and the detector is not calibrated before use. | Software design | Carestream sent a Urgent-Medical Device Recall letter dated July 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Carestream Health Inc. recently discovered an image artifact issue due to a detector not being recalibrated with the 5.6B software upgrade. As our installation instructions did not require the user to recalibrate during an upgrade, there is a possibility your detector could do the same. When this occurs an image artifact could be evident on images captured in the Extended Exposure mode. These artifacts could potentially contribute to misdiagnosis or the necessity of patient re-exposure.We request that you take a moment to verify that your detector(s) are calibrated. At our next visit, Carestream will also verify that your detector was re-calibrated and, if not, re-calibrate the detector for you. Please note this only affects systems upgraded in the field.For questions or concerns, please contact the Carestream Customer Care Center at 1-800-328-2910, 8 am through 8 pm EST in the US. |  | 23 | 23 units installed total (11 domestically &amp; 12 internationally) | Worldwide Distribution - US (nationwide) including the states of New York, Massachusetts, Pennsylvania, Illinois, Texas, California and Louisiana., and the countries of Australia, Canada, Finland, Italy, Switzerland and United Kingdom. | Open, Classified |  |  | 23 | Computer
Z-0125-2014 | 66473 | No | LGW |  |  | Neurology | PMA | Medtronic Activa PC, model 37601, and Activa RC, model 37612, Multi-program Rechargeable Neurostimulator for Deep Brain Stimulation. |  |  | September 30, 2013 | 2013 | Medtronic Neuromodulation | Medtronic notified customers regarding over stimulation or stimulation in the wrong area related to select Medtronic neurostimulators. Upon a specific set of conditions, typically related to device recovery from an over discharge, there is a potential for over stimulation or stimulation directed to a lead electrode other than what was intended. | Software design | The firm, Medtronic, sent an "Urgent: Medical Device Correction" letter dated September 2013 entitled "Loss of Stimulation and Over Stimulation" to customers associated with the affected Medtronic neurostimulators. This letter described the product, problem and actions to be taken. The letter informed customers of the potential for these issues to occur and provided recommendations to significantly reduce the likelihood of the issues. The customers were instructed to avoid over discharge; if a rechargeable device becomes over discharged, contact Medtronic Patient Services at (1-800-510-6735) to resolve the POR and reduce the likelihood of over stimulation.If you have questions, please contact Medtronic Neurostimulation Technical Services at 1-800-707-0933 weekdays 7am-6pm CST. |  | 38294 | 38,294 | Worldwide distribution: US (Nationwide) including Puerto Rico and countries of: Algeria, Aruba, Austria, Belarus, Belgium, Brazil, Bulgaria, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Faroe Islands, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malta, Martinique, Mexico, Morocco, Netherlands, Netherlands Antilles, New Caledonia, Norway, Oman, Pakistan, Panama, Peru, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Terminated | November 20, 2015 | 781 | 102457 | Not_Computer
Z-0091-2014 | 66305 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences EMBOL-X Glide Protection System; Model Numbers: EXGF24D, EXGF24LLD, EXGF24MMD, EXGF24SSD, EXGF24XLD, EXGF24XSD, EXGFXS2D, EXGF24SS2D, EXGF24MM2D, EXGF24LL2D and EXGF24XL2D. |  |  | September 16, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling certain lots of EMBOL-X Glide Protection System due to deformation of the cannula tip. | Nonconforming Material/Component | The firm, Edwards Lifesciences, sent an "URGENT FIELD SAFETY NOTICE PRODUCT RECALL-ACTION REQUIRED" letter dated September 16, 2013 via Fed-Ex to its consignees/customers. Consignees were asked to review inventory for the cannulae and remove the cannulae from any kits received; quarantine affected material at site and return product to Edwards LIfesciences, Attn: Santosh Bhagat, 12050 Lone Peak Drive, Draper, UT 84020, Attention: RECALL, RGA #; call Edwards Customer Service at 800-424-3278 to obtain RGA# and replacement product (note: the filter is not affected and does not have to be returned with the cannulae); and complete and return the attached acknowledgment form via fax to Edwards Customer Service at 800-422-9329 within three days of receipt of this Field Safety Notice.If you have any questions, please call Edwards Customer Service at 800-424-3278 from the hours 6:00am - 4:30pm PST or contact your Edwards' sales representative concerning the recall. |  | 12034 | 12,034 | Worldwide distribution: US (nationwide) and country of: Canada. | Terminated | March 10, 2017 | 1271 | 12034 | Not_Computer
Z-0128-2014 | 66423 | No | EWL | Laryngeal prosthesis (Taub design). | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Blom-Singer Adjustable Tracheostoma Valve II |  |  | October 01, 2013 | 2013 | Helix Medical LLC | InHealth Technologies (IHT) is recalling the Blom-Singer Adjustable Tracheostoma Valve II (ATSV II) because the body component of the device is missing from the packaging. The ATSV II Valve product cannot be used as intended without the product body. | Packaging change control | A phone script and a customer notification letter are provided. The recall was initiated on 10/1/13 and customers will be notified that InHealth Technologies, a division of Helix Medical, LLC has initiated a recall of the following 3 lots of the Blom-Singer Adjustable Tracheostoma Valve (BE8025H and BE8025H-R3): 1057760, 1057919 and 1059222. The phone script provides the problem identified and the actions to be taken. Customers are instructed to complete the Customer response form and return by email to belinda.jackson@helixmedical.com, fax to (805) 566-5381 or by mail to InHealth Technologies, 1110 Mark Avenue, Carpinteria, CA 93013. Customers are also instructed to Contact UPS at 1-800-742-5877 and provide In Health Technologies UPS account no., 846718. UPS will come to your site, package the recalled product(s) and ship it to InHealth Technologies at InHealth Technologies expense. Customers are instructed to respond within 10 days. |  | 47 | 47 units | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, Croatia, Germany, Ireland, and Italy. | Terminated | January 21, 2014 | 112 | 47 | Not_Computer
Z-0129-2014 | 66196 | No | NDG | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ti-Magna Fx Washer |  |  | August 28, 2013 | 2013 | Zimmer, Inc. | Zimmer recently conducted a review of historic packaging validations completed. Based on this review, it was determined that a subset of product packaged at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. The number of samples required for packaging testing, to provide statisti | Packaging process control | The firm, Zimmer Inc., sent a letter entitled "Voluntary Recall of Ti-VFxII Condylar Tube Plates and Ti-Washers" dated September 12, 2013 to its customer. The letter described the product, problem and action to be taken. The customer was instructed to immediately discontinue use of the listed products. A Sales representatives from Zimmer K.K. or agencies will be in contact with you. Zimmer Inc. is planning to resume product manufacturing upon the completion of package validation review. Zimmer K.K. will contact you when a detailed schedule has been determined.If you have any questions, contact Zimmer K.K. Safety Office (9:00 to 17:00; closed weekends and public holidays) TEL: 0120-015-180. |  | 2632 | 2,632 devices | International distribution: Japan only. | Terminated | April 30, 2015 | 610 | 2632 | Not_Computer
Z-0132-2014 | 65586 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | DRX Revolution, Catalog #'s 1)1019397, 2) 1023415 (w/ installation warranty, US-only) and 3) 1023423 (w/ parts warranty, US-only). |  |  | June 19, 2013 | 2013 | Carestream Health Inc. | Carestream Health Inc. is conducting a recall for the Column end Cover present on the DRX Revolution Mobile X-Ray System due to the cover dislodging from the column. | Device Design | The firm, Carestream Health Inc., sent a "URGENT: Medical Device Recall" letter dated June 17, 2013 via courier service to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that there is no opportunity for interference between the End Cover and potential obstructions as the column is descending. Carestream will deploy a field service representative to apply a warning label to the column, and to provide an addendum to the Safety and Regulatory Information manual. Your Carestream service representative will contact you to answer any questions and schedule a visit as soon as possible. If you have any questions or concerns, please contact the Carestream Customer Care Center at 1-800-328-2910 from 8am through 8pm. |  | 331 | 331 units | Worldwide distribution: US (nationwide) including states of: AL, AR, AZ, CA, CO, FL, GA, IA, ID, IL, IN, IO, KS, LA, MA, MD, MI, MN, MO, NC, NY, OH, PA, SC, TN, TX, UT, VA, WI, and WV; and countries of: Australia, Brazil, Canada, Dubai, Germany, Japan, Kuwait, Malaysia, New Zealand, Saudi Arabia, Singapore, South Africa, Taiwan, and Thailand. | Open, Classified |  |  | 331 | Not_Computer
Z-0113-2014 | 66383 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | COULTER LH 750 |  |  | September 24, 2013 | 2013 | Beckman Coulter Inc. | The check valves may fail on the Coulter LH 750 Hematology Analyzer (PN 6605632 and A85570) and Coulter LH 780 Hematology Analyzer (PN 723585 and A90728). A failure may result in an air or liquid leak at the connection to the Vacuum Overflow Tank waste line or at the Backwash Tank drain line. | Component change control | Beckman Coulter sent a Urgent Product Correction letter dated September 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. . The customer notification letter dated September 12, 2013, titled "URGENT PRODUCT CORRECTION", provided product description with codes, the issue (reason for recall), impact, action, resolution, and contact information.Customers were instructed to:If you experience a "Backwash Tank Not Full" error message, review manual mode sample results since the last acceptable QC run. System safeguards that assist with detection of erroneous results include Startup, QC, system messages and flagging (e.g., R), delta checks, flagging from limit sets (H, L, a, c), definitive messages and decision rules that would trigger further review.Avoid physical contact with a possible leak by continuing to use personalprotective equipment as recommended in the Instructions for Use and by national regulations.If you have any questions concerning this notification, please contact our Customer Support Center: Via our website, http://www.beckmancoulter.com/customersupport/supportVia phone, call 800-526-7694 in the United States and Canada.Outside US and Canada, contact your local Beckman Coulter RepresentativeWe apologize for the inconvenience this issue has caused your laboratory. |  | 1697 | 1,697 units | Worldwide Distribution - Nationwide Distribution and the country of Canada.. | Terminated | October 09, 2014 | 380 | 1697 | Not_Computer
Z-0127-2014 | 66456 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Dekompressor Percutaneous Discectomy Probe. |  |  | November 04, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | Stryker has identified a potential for the auger and/or cannula to fracture when the Dekompressor is used contrary to the warnings in the Instructions For Use (IFU). Specifically, manually bending the cannula can weaken the device and may result in a fracture of the auger and/or cannula. This could result in the need for additional medical intervention to remove the broken piece, pain or loss of m | Device Design | On 11/4/13, an URGENT MEDICAL DEVICE Recall Notification was sent to all domestic consignees via mail and foreign by email. An advisory was sent to every physician that uses any of the products listed. Every physician thatuses any of these products must print their name and sign below which means you acknowledge that youve received a copy of the attached advisoryand that you understand the information regarding the proper usage of the Decompressor.Once the form has been signed by all physicians, the form mustbe completed and signed by the recall coordinator. Please fax to866-521-2762 or email to kara.spath@stryker.com. |  | 20385 | 20,385 total distributed since 2004 | Nationwide Distribution including the states of AR, AZ, CA, CO, FL, GA, IL, KS, MS, ND, and NE. | Terminated | April 29, 2014 | 176 | 20385 | Not_Computer
Z-0085-2014 | 66368 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Baylis Medical TorFlex Transseptal Guiding Sheath, Rx Only, non-pyrogenic, sterile. |  |  | September 23, 2013 | 2013 | Baylis Medical Corp * | The TorFlex Transseptal Guiding Sheath Kits are being recalled due to the possibility that the sheath could remove particulate of less than 5 mm from the dilator. The particulate could potentially migrate into a patients bloodstream. | Component design/selection | Baylis Medical sent an Urgent Device Recall Notification letter dated September 23, 2013, via overnight courier to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the response form provided and fax to 905-602-5671 ATTN: Quality Department to arrange for a free of charge replacement of the affected product. The form requires consignees to return the affected product. For questions contact Quality Assurance at 800-276-4416. |  | 370 | 370 devices | Worldwide Distribution - USA Nationwide in the states of : NE, NH, NC, FL, PA, VA, MI, CA, AL, and TN and the countries of China, Canada, Saudi Arabia, and Italy. | Terminated | March 31, 2015 | 554 | 370 | Not_Computer
Z-0133-2014 | 66504 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | The EkoSonic Endovascular System employs high frequency (2-2.5 MHz), low power (0.585 watt average power per transducer) ultrasound to facilitate the delivery of physician-specified fluids in the peripheral vasculature and to deliver solutions into the pulmonary arteries. |  |  | October 02, 2013 | 2013 | EKOS Corporation | The EkoSonic Control Unit is recalled because it had the incorrect power entry module installed. It is missing an AC filter to reduce the noise and electrical interference. | Mixed-up of materials/components | EKOS notified their customer by email on October 2, 2013, about the recall. EKOS has requested the customer to return of EkoSonic Control Unit PT3B-783, so that it can be replaced.CORRECTIVE ACTION:This single affected unit will be reworked with the correct power entry module and released into finished goods as a refurbished control unit.EKOS did not distributed any affected devices in the US.For any questions, please call EKOS Corporation at 425-415-3100. |  | 1 | one unit | Worldwide Distribution in Germany only. | Terminated | November 22, 2013 | 51 | 1 | Not_Computer
Z-0130-2014 | 66545 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS(R) II, NON-POROUS TIBIAL BASE, SIZE 5, LEFT, QTY: (1), REF 71420168. |  |  | October 07, 2013 | 2013 | Smith &amp; Nephew Inc | Swapped product. Non-porous tibial base size 5 L was mis-packaged as size 6 R and vice versa. | Process control | Smith &amp; Nephew sent an Urgent - Product Recall 1st Notification letter to all affected sales staff via email and FedEx on October 7, 2013. The hospital was notified by sales rep on October 9, 2013 and the notification letter was sent to the hospital via Fed Ex on October 11, 2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory, discontinue use, quarantine affected products immediately and return to Smith &amp; Nephew. Customers were asked to complete the attached Inventory Return Certification Form and fax to 901-566-7975. |  | 4 | 4 units | USA Nationwide Distribution in the states of: AZ, NC, IA, PA, OH, MD, and TN. | Open, Classified |  |  | 4 | Not_Computer
Z-0135-2014 | 66112 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Dispersion Aortic Perfusion Cannula with Duraflo Coating, Sterile, RX only For Single Use Only, |  |  | August 26, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling two lots of Duraflo Cannulas due to reported complaints involving cracks in the t-connector of the devices which were confirmed during product evaluation. | Process control | Edwards sent an Urgent Field Safety Notice Product Recall letter dated August 26, 2013 to all affected customer, via Fed-Ex. The product is sold directly to hospitals and distributors. The distributors involved will communicate with their customers directly. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their entire inventory, quarantine and return the affected products. An acknowledgement form is included for completion to assist in the assessment of the inventory. For questions contact Edwards Customer Service at 800-424-3278. |  | 140 | 140 units | Nationwide Distribution. | Open, Classified |  |  | 440 | Not_Computer
Z-0016-2014 | 65161 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens RAD Fluoro Uro Systems - Luminos dRF, Ysio and Uroskop Omnia with software versions VB10D, VB10F, VB10C and VB10E. |  |  | May 06, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Siemens has discovered through product monitoring a potential malfunction when using system operating with software versions VB10D, VB10F, VB10C and VB10. During a RAD examination the possibility exists that radiation may continue longer than expected and may result in unnecessary dose applied during an acquisition. Additionally, it is not possible for the user to stop the radiation manually; ho | Software design | Siemens initiated this correction on May 6, 2013 by sending Update InstructionXP005/13/S. Firm subsequently issued a letter on June 3, 2013 to all accounts.1. Siemens Field Modification of the Luminos dRF, Ysio &amp; Uroskop Omina running software versions VB10D, VB10F, VB10C and VB10E, which improves the signal transmission for the automatic exposure control, is available as of May 2013, and systems will be updated within one year.2. The revised software fix for models Luminos dRF, Ysio &amp; Uroskop Omina running software versions VB10D, VB10F, VB10C and VB10E, will be provided to all identified customers at no charge.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 workingdays of receipt of this letter in the manner specified in 21 CFR 1003.21 and1003.22. This office and the Food and Drug Administration (FDA) districtoffice coordinator noted below are to be included in the notification.2. Siemens Medical Solutions USA, Inc., will notify the FDA district office inPhiladelphia, Pennsylvania, with reports of progress regarding this fieldcorrection. |  | 23 | 23 | Nationwide Distribution | Terminated | November 06, 2014 | 549 | 23 | Computer
Z-0136-2014 | 66101 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Bayer brand A1CNow+ MULTI-TEST A1C SYSTEM TEST KITS (PART NUMBER 3021/08842610); Hemoglobin A1c Monitor; Glycosylated Hemoglobin Assay; Product is manufactured and distributed by |  |  | August 26, 2013 | 2013 | Bayer Healthcare Llc | Test results from these lots may produce a high-bias reading (by up to 20%) that is outside of limits defined for stability and product accuracy when product is stored at room temperature four (4) to five (5) months after release. | Pending | Bayer Healthcare sent an Urgent Medical Device Recall letter dated August 26, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory for the affected product and quarantine any affected product. Complete the attached Recall Response Form and fax to Bayer Customer Service. Customers were instructed to use the attached letter to notify affected next level customers of the recall. For questions regarding this recall call 408-524-2255, ext 290. |  | 138 | 138 Kits in US (Total 2452) | Worldwide Distribution - USA (nationwide) and Internationally to China, Argentina, Uruguay, Japan. | Terminated | January 14, 2014 | 141 | 138 | Not_Computer
Z-0036-2014 | 66335 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran Mini CT Scanner |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | A description of the defect in the product or the manner in which the product fails to comply with an applicable Federal standard; The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 ÂżÂż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evalu | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. |  | 320 | 320 | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1507 | 320 | Not_Computer
Z-0041-2014 | 66344 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT CT Scanner ID ENT00001 |  |  | December 14, 2009 | 2009 | Xoran Technologies, Inc. | Failure of the gantry to rotate. The complaint alleged that the scanner would emit radiation but would not rotate. | Radiation Control for Health and Safety Act | The service report in response to this complaint stated that the Head select relay board, chip and fan were replaced. The machine (ENTOOOO I) was subsequently tested. All tests were successful. |  | 320 | 320 | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1732 | 320 | Not_Computer
Z-0038-2014 | 66338 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran Minicat00202 |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evaluation of the hazards reasonably related to defect or the failure to comply with the Federal standard; An evaluation of the health ha | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  | 320 | 320 | Worldwide Distribution - US nationwide | Terminated | September 11, 2014 | 1507 | 320 | Not_Computer
Z-0040-2014 | 66340 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT00202 |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xorans specifications. | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  | 320 | 320 | Worldwide Distribution - Nationwide | Terminated | September 11, 2014 | 1507 | 320 | Not_Computer
Z-0037-2014 | 66337 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT CT Scanner |  |  | July 27, 2010 | 2010 | Xoran Technologies, Inc. | A description of the defect in the product or the manner in which the product fails to comply wvith an applicable Federal standard; The MiniCAT CT Scanners subject to this notification failed to meet Xoran's specifications for dose rate of 125 Âż 3 kVp. These 18 complaints received include dose rates both on the high and low end of Xoran's specifications. See ATTACHMENT 1 for details. (e) A n evalu | Radiation Control for Health and Safety Act | All of the MiniCAT CT Scanners subject to these complaints were serviced by Xoran technicians and the out-of-specifications devices were brought into specification. This was accomplished by either readjusting the settings or by installing a new component. Xoran's Research &amp; Development group is evaluating the current specification to see if the field specifications are appropriate and to ensure that proper testing protocols are communicated and are in place for use by 3 rd party testing organizations. In addition, Xoran has initiated CAPA # 18 0 that is looking at the source block with regard to evaluating the stability of the dose rating. |  | 320 | 320 | Worldwide Distribution - Nationwide | Terminated | September 11, 2014 | 1507 | 320 | Not_Computer
Z-0144-2014 | 66040 | No | HNJ | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | MLS Torpedo Mini-Light Ophthalmic single-port illumination. |  |  | June 27, 2013 | 2013 | Insight Instruments, Inc. | The product was registered and distributed as a Class I medical device and was subsequently determined to be a Class II medical device which requires FDA market clearance. | No Marketing Application | Insight Instruments, Inc. sent a letter dated April 8, 2013, via USPS priority mail to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue using and dispensing all affected product, perform a physical count of their inventory and complete the Return Response Form and fax to 772-219-9342. Customers were also instructed to call 800-255-8354 to arrange for return and prompt reimbursement.For questions regarding this recall call 727-219-9393. |  | 10212 | 10212 | Worldwide Distribution - USA (nationwide) including AZ, CA, FL, IN, MA, MD, TX, and WA and Internationally to Czech Republic, Germany, Netherlands, Spain, Switzerland, and Turkey. | Open, Classified |  |  | 10212 | Not_Computer
Z-0146-2014 | 66570 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Vision System Cart, a component of the |  |  | October 21, 2013 | 2013 | Intuitive Surgical, Inc. | Intuitive has identified that the metalized coating on the IS3000 HD lamp module may be incompatible with the illuminator control board, leading to an erroneous red Preventive Maintenance Message that shows up on monitors. | Nonconforming Material/Component | Intuitive Surgical sent an Urgent Medical Device Correction letter dated October 23, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to open the illuminator and replace any metal-finish lamp modules with plastic finish lamp modules, if customers do not have a replacement plastic-finish lamp module available, they may continue to use the metal-finish lamp module until they receive a replacement. If the message "illuminator lamp module error: Please reseat or replace lamp module" is displayed on the patient side monitor and the surgeon console, contact intuitive Surgical Customer Service. Customers were also instructed to examine any remaining or spare inventory to identify any metal-finish lamp modules and return them to Intuitive Surgical for credit. Customers were instructed to ensure all affected personnel are fully informed of this Urgent Medical Device Corrrection letter. Customers should complete the attached Acknowledgement Form and return it to Intuitive Surgical as instructed. Customers with questions should contact Intuitive Surgical Customer Service.For questions regarding this recall call 408-523-2244. |  | 684 | Standalone: 161 units; With Vision Starter Kit: 240 units; With IS3000 Vision Cart: 283 units | Worldwide Distribution - USA (nationwide) and Internationally tAustralia, Austria, Belgium, Brazil, *Bulgaria, Canada, Chile, China, CzechRepublic, Denmark, Finland, France, Germany, Great Britain, India, Israel, Italy, Mexico, Monaco, Netherlands, Norway, Pakistan, Panama, Qatar, Romania, Russian Fed., Saudi Arabia, Singapore, South Korea, Sweden, Switzerland, Taiwan, and Turkey. | Terminated | September 04, 2014 | 318 | 684 | Not_Computer
Z-0142-2014 | 66464 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | S5 Heart-lung machine, Stockert S5 System. Catalogue No: 48-30/40/50-00. |  |  | September 24, 2013 | 2013 | Sorin Group Deutschland GmbH | Sorin Group is recalling certain models of S5 Perfusion Systems due to malfunction of cardioplegia control. | Software Manufacturing/Software Deployment | Sorin Group sent an Urgent Field Safety Notice dated September 24, 2013, to all affected customers. The letter informed users of the potential for failing of automatic cardioplegia delivery function and notified them that they would be contacted by the Sorin Service Group to arrange an appointment to update the firmware on site. Customers were asked to complete the Response form to confirm they had received, read and understood the Field Notice. For questions regarding this recall call 303-467-6527. |  | 84 | 84 | Worldwide Distribution - USA including WA, NY, VA, TN, SC, TX, CA, CO, MA, WI, MD, AL, ND, and Puerto Rico and Internationally to Italy, Finland, Germany, Sweden, and Great Britain. | Terminated | July 10, 2014 | 289 | 84 | Not_Computer
Z-0147-2014 | 66382 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | UniCel DxH 800 and DxH 600 Coulter Cellular Analysis Systems |  |  | September 24, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall for UniCel DxH 800 Coulter Cellular Analysis System (PN 629029) and UniCel DxH 600 Coulter Cellular Analysis System (PN B23858) because the company's internal investigation indicates that specific lots of check valves used in several locations within the DxH 800 and DxH 600 systems may fail. A failure may result in an air or liquid leak at the VCS Waste Chamb | Under Investigation by firm | The firm, Beckman Coulter, sent an "URGENT PRODUCT CORRECTION" letter dated September 12, 2013, its customers. The letter includes the products descriptions and codes, issue, impact, action, resolution, instructions to the customers, and contact information. The customers were instructed to review the three samples analyzed prior to error message,if you experience a "Vent/overflow chamber detected overflow' error; to avoid physical contact with a possible leak by continuing to use personal protective equipment as recommended in the Instructions for Use and by national regulations; share this information with your staff and retain this notification as part of your Quality System documentation; if you have forwarded any of the affected products, provide them a copy of this letter; and to complete and return the enclosed RESPONSE FORM within 10 days via fax to: 786-639-7500 /7501/7502/7504; email: Regulatory.Notifications@beckmancoulter.com (note: please include "DxH system check valve letter dated 09/12/2013" in email subject field) and/or by mail to: Beckman Coulter, Inc., Attn: Regulatory Affairs Mail Code 31-B06, 11800 S.W. 147th Avenue, PO Box 169015, Miami Fl 36116-9015. .Your service representative will contact you to schedule an on-site visit for check valve inspection and replacement of any affected valves.If you have any questions concerning this notification, please contact our Customer Support Center: via our website, http://www.beckmancoulter.com/customersupport/support, via phone, call 800-526-7694 in the United States and Canada, and Outside US and Canada, contact your local Beckman Coulter Representaitive. |  | 1025 | 1,025 units | Worldwide distribution: US (nationwide) including Puerto Rico, and countries of: Australia, Belgium, Canada, China, Croatia, Czech Republic, France, Germany, Hong Kong, India, Ireland, Italy, Japan, Macao, Netherlands, New Zealand, Panama, Poland, Portugal Qatar, Russian Federation, Spain, Sweden, Switzerland, Taiwan, Thailand, Trindad and Tobago, Turkey, United Arab Emirates, and United Kingdom.. | Terminated | April 21, 2014 | 209 | 1025 | Not_Computer
Z-0145-2014 | 66457 | No | FJL | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Individual units labeled in part: REF: 4630022; for S(a)line resectoscopes; Vaporization electrode, sterite; 4mm telescope 30 25 / 12: only for continuous irrigation sheath 24,5 Fr.; LOT 426111; 2011/06; 2016/06; STERILE; RICHARD WOLF; |  |  | September 18, 2013 | 2013 | Richard Wolf Medical Instruments Corp. | Richard Wolf Medical Instrument Corporation recieved a complaint stating that a 4630024 vaporization electrode was packaged in a 4630022 pouch. | Process control | The firm, Richard Wolf, sent an "Urgent: Medical Device Recall" letter dated September 16, 2013, to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to inspect their inventory of vaporization electrodes from lot 426111, contact Richard Wolf Medical Instrument Corporation at 800-323-9653 x 344 or ext 225 or ext. 380 to inform Richard Wolf Medical Instrument Corporation if the customer has or does not have any devices from this lot, return the affected devices with the return authorization provided by Richard Wolf Medical Instrument Corporation for credit, and complete and return the attached Recall Response form via fax to: 847-913-0924; email: dclark@richardwolfusa.com; or mailed to: Richard Wolf Medical Instrument Corporation, 353 Corporate Woods Parkway, Vernon Hills, IL 60061-3110.Customers who require urgent replacement of the affected devices can contact their customer service representative for replacement.If you have any questions, call 800-323-9653 x 344. |  | 78 | 78 cartons (234 units) | Nationwide distribution: US including states of: CA, GA, IL, KY, LA, MI, MD, NY, OR, TX, VA and Puerto Rico. | Terminated | April 24, 2017 | 1314 | 78 | Not_Computer
Z-0139-2014 | 66526 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Patient Cart used in conjunction with the da Vinci S surgical and SI System. |  |  | November 30, 0012 | 12 | Intuitive Surgical, Inc. | Increased number of complaints regarding fluid entering the patient card base of the da Vinci S and SI systems. | Device Design | Intuitive Surgical Inc sent an Product Notification to all affected customers via email on November 30, 2012.E-mail notifications advised users of an addendum to the User Manual for Covering the Patient Cart column that provided a recommended method for covering the da Vinci patient cart column to prevent splashing on the cart during a procedure. The addendum was attached to the e-mail and is available online.If you have any questions, please contact Customer Service at 1-800-876-1310. Outside the United States call +41 21 821 2020 or +800.0821.2020 |  | 1754 | 1754 patient carts | Worldwide Distribution - USA including Puerto Rico and the countries of Australia, Austria, Argentina, Belgium, Bulgaria, Brazil, Canada, Chile, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, India, Indonesia, Ireland, Israel, Japan, Kuwait, Mexico, Netherlands, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Uruguay and Venezuela. | Terminated | November 04, 2013 | 730824 | 1754 | Not_Computer
Z-0138-2014 | 65782 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Stereotactic Circular Collimator - 3D Line stereotatic Hardware Accessories |  |  | July 23, 2013 | 2013 | Elekta, Inc. | Potential for clinical errors. | No Marketing Application | Elekta sent an Important Field Safety Notice dated 7/2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised of the issue and recommendation to resolve either by upgrading their system or to discontinue a clinical use of the device.For further questions please call ( 770 ) 300-9725. |  | 89 | 89 | Nationwide distribution: USA including the states of AZ, CA, FL, GA, MO, NJ, NM, NC, OR, TN and WV. | Terminated | November 18, 2014 | 483 | 89 | Not_Computer
Z-0155-2014 | 66587 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ORBIT PT. |  |  | July 24, 2013 | 2013 | Trumpf Medical Systems | An odor and smoke developed - The cause was determined to be a power supply unit sitting in the controller column. | Component design/selection | The user facility was contacted via phone in early August. They were contacted 9/17/2013 to schedule the replacement of the power pack because the new power pack became available. |  | 1 | One | US Distribution: NJ only. | Terminated | June 03, 2014 | 314 | 1 | Not_Computer
Z-0167-2014 | 66461 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Aquatec Ocean VIP, Aquatec Ocean Dual-VIP, Aquatec Ocean E-VIP mobile shower and toilet commode with tilting function. |  |  | August 31, 2013 | 2013 | Invacare Corporation | The snap-on fixtures of the backrest cushion on the Mobile Shower and Toilet Commode Aquatec Ocean VIP and Ocean Dual VIP do not withstand a required pull out force of 500 N, when the product is used in a reclined position and or tilted position. This may result in the snap-on fixture coming unlatched. | Nonconforming Material/Component | Invacare sent an Urgent Medical Device Recall Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to acknowledge their receipt and understanding of the urgent recall notice and indicate the number of affected product in their inventory on the enclosed Provider Response Card. Customers were instructed to examine their existing stock and quarantine any affected product and fax the completed Provider Response Card immediately to Invacare Corporation at 440-326-3544. Customers were responsible for ensuring that the recall was conducted to the customer level. Correction kits were provided to the customers with the communication. Customers were instructed to document the correction details on the enclosed Product Tracking Sheet and fax to Invacare Customer Service at 440-326-3544.For questions regarding this recall call 440-329-3544. |  | 538 | 538 (US only) 3317 (Total) | Worldwide Distribution - USA including PA and Internationally to Canada. | Terminated | September 15, 2014 | 380 | 538 | Not_Computer
Z-0148-2014 | 66339 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius LibertyÂż Cycler |  |  | October 14, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Fluid leaking may be observed inside the pump compartment of the Liberty Cycler and may result in dialysate becoming contaminated, potentially causing peritonitis | Device Design | Fresenius Medical Care contacted all affected consignees by telephone on 10/14/13 to inform them of the product safety alert. Follow-up Urgent Safety Alert letter was sent on 10/18/13 and 10/23/13 by Certified Mail, Return Receipt requested. The letter identified the affected product, problem and actions to be taken. Customers were instructed to report any events associated with this safety alert to Customer Service Team at 1-800-323-5188. |  | 16600 | 16,600 | USA Nationwide Distribution | Open, Classified |  |  | 16600 | Not_Computer
Z-0168-2014 | 66239 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | COBAS INTEGRA 800 Analyzer with software 9864.C2 |  |  | September 16, 2013 | 2013 | Roche Diagnostics Operations, Inc. | Roche COBAS INTEGRA 800 system, Software Version 9864.C2 does not perform necessary ISE Service Actions. Whenever the ISE rack is removed, the ISE counters in the software are reset to zero and the corresponding Service Actions are not conducted automatically by the system. Additionally, certain maintenance activities that have to be performed manually by the operator are not requested on the sof | Software Design Change | Roche sent an Urgent Medical Device Correction letter dated September 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Roche Diagnostics recommends tracking all ISE counters and manually requesting the corresponding Service Actions listed in the notification be monitored until updated software becomes available during Q4 2013. If the affected product was moved to another site, the notification is to be forwarded. consignees were asked to complete the attached faxback form and fax it to 1-877-266-0997 and file the Urgent Medical Device Correction (UMDC) for future reference. Question should be directed toRoche Support Network Customer Support Center 24 hours a day, seven days a week at 1-800-428-2336For questions regarding thisr ecall call 317-576-3911. |  | 527 | 527 units | Nationwide Distribution and Puerto Rico | Terminated | May 12, 2015 | 603 | 527 | Computer
Z-0151-2014 | 66465 | No | LTI |  |  | Gastroenterology/Urology | PMA | Adjustable Gastric Band, packaged alone in transparent PETG inner and outer blisters with Tyvek lids. |  |  | September 30, 2013 | 2013 | Ethicon Endo-Surgery Inc | During aging studies, the firm determined that the gastric belts were unable to pass the test requirements established for Force to Lock. | Unknown/Undetermined by firm | Ethicon sent an Urgent: Medical Device Recall letter dated September 30, 2013 to affected consignees via overnight, UPS to allow tracking of the receipt. The letter identified affected products, problem and actions to be taken. Each consignee is requested to respond to a provided Business Reply Form (BRF), indicating whether they have affected product, how much and whether they will/have returned the product or whether they no longer have any of the recall product. For questions contact your local sales representative or call our Customer Support Center 1-877-384-4266 option 6. |  | 1248 | Product Code: BD2XV - 1,248 units | Worldwide Distribution: USA (nationwide) including the states of: Alabama, Connecticut, Indiana, Michigan, North Carolina, South Carolina, Alaska, Delaware, Kansas, Minnesota, Ohio, Tennessee, Arizona, Florida, Kentucky, Mississipp,i Oklahoma, Texas, Arkansas, Georgia, Louisiana, Nebraska, Oregon, Virginia, California, Idaho, Maryland, Nevada, Pennsylvania, Washington, Colorado, Illinois, Massachusetts, New Jersey, Rhode Island, and West Virginia; and the countries of: Argentina, France, Malaysia, Spain, Australia, Georgia, Martinique, Sweden, Belgium, Germany, Mexico, Switzerland, Brazil, Greece, Netherland, Taiwan, Chile, Hungary, Panama, Tunisia, China, India, Philippines, United Arab Emirates, Czech Republic, Israel, Poland, United Kingdom, Egypt, Italy, Portugal, Venezuela, Estonia, Kenya, Russia, Finland, Kuwait, and Saudi Arabia. | Terminated | March 26, 2014 | 177 | 7438 | Not_Computer
Z-0153-2014 | 66253 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ISE REF Valve, Part No. MU7638 for the AU480, AU680, and AU5800 Chemistry Analyzer. |  |  | September 09, 2013 | 2013 | Beckman Coulter Inc. | The recall was initiated because Beckman Coulter has confirmed that some AU Clinical Chemistry Analyzer ISE units can experience premature ISE REF (Reference) valve failure. Premature valve failure can be detected by the presence of air bubbles in the ISE REF solution tubing on the ISE unit. If air bubbles are present, accuracy of the ISE test results may be affected. | Nonconforming Material/Component | A customer notification letter dated 9/9/13 was sent to customers who received the ISE REF Valve, Part Number MU3226 for the AU400, AU640, AU2700, AU5400 and the ISE REF Valve, Part Number MU7638 for the AU480, AU680, AU5800 Chemistry Analyzers. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Beckman Coulter Customer Support Center via http://www.beckmancoulter.com/customersupport/support, via phone at (800) 854-3633 in the US and Canada. Outside of the US and Canada, customers are instructed to contact their local Beckman Coulter Representative. |  | 4584 | 4584 units total (1173 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Canada, Vietnam, Zambia, and Zimbabwe. | Terminated | September 30, 2014 | 386 | 4584 | Not_Computer
Z-0166-2014 | 66425 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | OmniPod Insulin Management Product Demonstration Kit; Model Reference Number: (PDK-ZXP490) |  |  | September 27, 2013 | 2013 | Insulet Corporation | Alcohol Prep Pads contained within the kit are not properly labeled. | Labeling design | Insulet sent an Important Corrective Action Notification letter dated September 27, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers are asked to discontinue use and properly dispose of the affected product. For questions contact Customer Care at (866) 941-0294. |  | 5050 | 5,050 kits | Worldwide Distribution - USA Nationwide and the country of Israel. | Terminated | March 04, 2014 | 158 | 6769 | Not_Computer
Z-0164-2014 | 66351 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | PortexÂż Saddleblock Tray with Drugs  25g Quincke; 4795-20 |  |  | September 23, 2013 | 2013 | Smiths Medical ASD, Inc. | Visual particulate in the glass ampules of 5% Lidocaine HCL and 7.5% Dextrose Injection, USP, 2 mL Ampoule, NDC # 0409-4712-01, Hospira Lot Number 23-227-DK. These ampules are included in certain Portex Spinal Anaesthesia Trays. | Material/Component Contamination | Smiths Medical sent an Urgent Recall Notice dated September 23, 2013, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory for the affected product and remove from use. Complete the Confirmation Form and return by fax to 603-358-1017 or by email to spinal@smiths-medical.com. Upon receipt of the completed form, a customer service representative would contact them to arrange for exchange of their unused affected devices for credit or replacement. Customers were also instructed to forward the notice to all personnel who need to be aware within their organization. Customers with questions were instructed to contact Smiths Medical's Customer Service Department at 800-258-5361.For questions regarding this recall call 800-258-5361. |  | 222 | 222 | Nationwide Distribution including AK, AR, AZ, CA, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, RI, SC, TX, UT, VA, VT, WA, WI, WY. | Open, Classified |  |  | 7549 | Not_Computer
Z-0154-2014 | 66532 | Yes | NIQ |  |  | Cardiovascular | PMA | Resolute Integrity Zotarolimus-eluting Coronary Stent: |  |  | October 07, 2013 | 2013 | Medtronic Vascular, Inc. | A small number of units in lot number 0006573585 of the Resolute Integrity Zotarolimus-Eluting Coronary Sten System may contain a compliance chart incorrectly referencing a 2.25 mm diameter instead of the correct 3.0mm. | Process control | Urgent Medical Device Recall Letters were sent on October 7, 2013 via UPS 3-day service. |  | 201 | 201 units | Nationwide Distribution. | Terminated | February 11, 2014 | 127 | 201 | Not_Computer
Z-0174-2014 | 66597 | Yes | OAD |  |  | Cardiovascular | PMA | US Safire Duo Ablation Catheter MediGuide Enabled, REF A700240 and Cool Path Duo Ablation Catheter, MediGuide Enabled REF A700244, Rx only, Sterile EO. |  |  | October 11, 2013 | 2013 | St. Jude Medical | St. Jude Medical Cardiovascular and Ablation Technologies Division is performing a voluntary notification regarding selected batches of Coolpath Duo Ablation Catheter MediGuide Enabled products and Safire Duo Ablation Catheter MediGuide Enabled. These units were inadvertently shipped with the incorrect version of the Instructions for Use (IFU). | Error in labeling | The firm, St. Jude Medical (SJM), sent a "Urgent Medical Device Notice" dated 11 October 2013 to its consignees. The letter described the problem, the product involved in the recall and the actions to be taken. The consignees were instructed to discard the IFU received with the batches and replace it with the IFU received in conjunction with the notification letter. SJM has determined that it is acceptable to continue using these devices as long as the product is used in accordance with the IFU.If you have any questions, contact Vice President, Quality, at 651-756-5402. |  | 59 | 59 | US distribution: KS, MA, UT,TX, KY, OH. | Terminated | March 10, 2016 | 881 | 59 | Computer
Z-0141-2014 | 66138 | [6 empty] | Harmony Equipment Management System, Model #603X/80XX, STERIS Corporation, Mentor, OH |  |  | August 15, 2013 | 2013 | Steris Corporation | DrÂżger Medical GmbH has notified STERIS of the potential for a break in the drive screw contained in the motorized ceiling supply unit, a component of the Harmony EMS system. The ceiling supply unit controls the motorized adjustment of the height and rotation of the supply column. Should a break occur in the drive screw, there is potential for the adjustable arm to drop down, causing the EMS sys | Component design/selection | STERIS sent an Urgent Voluntary Field Correction Notice dated August 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.STERIS action- Drager Medical has provided STERIS instructions for inspection of the alignment of the mounting plate of the motor unit. If it is found that the mounting plate is bent, the drive screw within the motor unit must be replaced. STERIS will visit your facility to inspect the motorized ceiling supply unit of your Harmony EMS system(s) and if damage is observed, will replace the drive screw within the motor unit. For further information or if you have any questions regarding our visit to your facility, please contact STERIS Field Service Dispatch at 1-800-333-8828. |  | 127 | 127 units | Nationwide distribution: USA including the states of Alabama, Arkansas, Florida, Georgia, Illinois, Indiana, Kansas, Kentucky, Louisiana, Maryland, Massachusetts, Michigan, Minnesota, Missouri, Nebraska, New Jersey, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Texas, Virginia, Washington and Wisconsin. | Terminated | April 28, 2014 | 256 | 127 | Not_Computer
Z-0173-2014 | 66132 | No | CEW | Parathyroid hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Intact PTH Calibrator (6 pack); Catalog Number: SMN 10492387 |  |  | August 16, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | PTH Calibrator packaged with only high calibrator. | Packaging | Packaging error was brought to Siemens' attention through escalations received from (3) customers in the United States. Internal Investigation confirmed the issue was limited to the five (5) kits distributed to these customers. As the customers were aware of the issue and non affected product was provided to the customers to use during Siemens' investigation, no further communication was warranted. |  | 595 | 595 kits total (5 kits with packaging error) | Nationwide Distribution including MD, PA, and CT. | Terminated | February 13, 2014 | 181 | 595 | Not_Computer
Z-0170-2014 | 66408 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | HAMILTON-C2 Ventilator: |  |  | September 26, 2013 | 2013 | Hamilton Medical, Inc. | Manufacturer defined limits for improperly functioning loudspeakers of the HAMILTON C2 alarm system were exceeded. | Device Design | Hamilton Medical sent an Medical Device Safety Alert and Corrective Action letter dated September 27, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The corrective action does not require any action by the operator. In the case of an alarm situation as described in the letter, users should follow the instruction for use (Section 8 - responding to alarms in the operator manual) and have the ventilator serviced by an authorized service technician. An addendum to the Operators Manual as a result of the field correction was provided to customers with the Safety Alert Letter.Manufacture will develop and provision a test protocol for the service tech which test the proper functioning of the loudspeaker during the annual service inspection.We appreciate your support in this matter and sincerely regret any inconveniences that this action may cause you. We consider this action as necessary to ensure that our customers receive only safe and effective products with high quality.Further questions please call (775) 426-6331 Ext. 210 |  | 316 | 316 devices distributed in the US. | Nationwide distribution: USA including Puerto Rico and the states of IL, MT, IN, ND, AZ, MT, KY, NC, OH, MA, MO, VA, CA, NV, TX, WA, NY, GA, SC,MN, WY, PA, ND, UT, NE, CO and MI. | Open, Classified |  |  | 316 | Not_Computer
Z-0169-2014 | 66659 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | VitreaCoreÂż software. Versions subject to the limitation: |  |  | October 21, 2013 | 2013 | Vital Images, Inc. | Vital Images, Inc. has found an issue in VitreaCore software though internal testing. There is a potential for incorrect and possibly reversed image orientation for certain snapshots taken in VitreaCore software versions 6.0, 6.1, 6.2, 6.3 (except 6.3.1), 6.4 (except 6.4.4 and 6.4.5), as well as their upgrades, and 6.5. | Software design | Vita Images, Inc. sent an Urgent Medical Device Recall letter on October 21, 2013, to all affected customers via USPS First class mail. The letter described the Reason for Voluntary Recall (Nature of software issue, Potential scenarios which exhibit the issue), Risk to Health (How to confirm if the Snapshots are correct), Actions to be Taken by Customer (Recommended Next Steps which includes to complete and sign the enclosed Effectiveness Check Form and send to fieldnotices@vitalimages.com or fax to 1-952-540-3717), and Actions Being Taken by Vital Images (Vital Images Contact Information: If you have questions, direct Vital Images customers and customers of Toshiba America Medical Systems (TAMS) should contact Vital Images Customer Support at +1-800-208-3005, or by e-mail to support@vitalimages.com; customers through a distributor will be contacted by that distributor).For questions regarding this recall call 952-487-9574. |  | 598 | 598 | Nationwide Distribution - all states in continental USA including DC but not OK or RI.OUS: Australia, Austria, Belgium, Brazil, Canada, China, Denmark, Finland, France, Germany, Honduras, Hong Kong, Iceland, Ireland, Italy, Korea, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Qatar, r Saudi Arabia, Serbia, Singapore, Slovenia, South Africa,Spain , Sweden, Switzerland, Thailand, Turkey, Ukraine. | Terminated | January 16, 2015 | 452 | 598 | Computer
Z-0143-2014 | 66249 | No | OBK | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | Dual Luer Lock Cap |  |  | September 10, 2013 | 2013 | Baxter Healthcare Corp. | The firm is recalling lots 10043 and 10044 due to loose particulate matter found in the packaging. Particulate matter entering the fluid path from the Luer lock caps may result in embolic events. | Employee error | Baxter sent an Urgent Product Recall notification dated September 16, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised to :1. Locate and remove all affected product from your facility. If you distribute these products to other facilities or departments within your institution, please forward a copy of this communication to ensure that they locate and remove affected product from potential use.The product code and lot number can be found on the shipping carton label and individual product package.2. Acknowledge your receipt of this recall notification by completing the attached Customer Reply Form (Attachment 1) and returning it to Baxter by either faxing it to 1-224-270-5457 or scanning and emailing it to fca@baxter.com. Returning the Customer Reply Form promptly will prevent you from receiving repeat notifications.3. The affected product should be returned for credit by contacting Baxter Healthcare Center for Service. The Center for Service can be reached at 1-888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time. Please have your Baxter eight-digit ship-to account number ready when calling. Please have the product code and lot number(s) available.4.If you are a dealer, wholesaler,or distributor/reseller that distributed any product to other facilities, please forward this communication. Baxter distributed this product to customers between June 19, 2013 and August 20, 2013.The Center for One Baxter is available at 1-800-422-9837, Monday throughFriday during the hours of 8:00 am to 5:00 pm Central Time. |  | 628992 | 628,992 units | Worldwide Distribution - USA (nationwide) and the country of Canada. | Terminated | June 05, 2015 | 633 | 628992 | Not_Computer
Z-0171-2014 | 63859 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Handset component of the Handicare USA's Minivator Bath Bliss Reclining Bath Lifts (Bath Bliss), Part Number BP300-100/BP300-400RSUS (handset) - Model Numbers BBR0311USAB (blue) and BBR0311USSS (silver) |  |  | November 29, 2012 | 2012 | Handicare Accessibility | Firm became aware that two handsets in the affected batches became inoperable while the user was in the tub. | Device Design | Handicare sent an Urgent Medical Device Recall letter dated November 29, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately advise all of their customers affected by the recall letter to immediately remove all affected product from use. Upon receipt of affected product, replacement handsets would be sent to customers. Customers with questions were instructed to call 866-276-5438 or email Customerservice.usa@handicare.com. For questions regarding this recall call 610-266-5260. |  | 3301 | 3301 handsets | Worldwide distribution: USA (Nationwide) and country of: Canada. | Terminated | December 08, 2014 | 739 | 3301 | Not_Computer
Z-0178-2014 | 66331 | No | CJY | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Calcium Arsenazo, CALA-0250, For the Quantitative Determination of Total Calcium in Human Serum. |  |  | September 04, 2013 | 2013 | SEPPIM SAS | Calcium Arsenazo Reagent, reference CALA-0250, is recalled due to complaint of plasma specimens report high results than serum. | Device Design | ELITech Group of Bothell, WA sent an Urgent: Product Recall Notice dated September 9, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken:" Laboratories must not use plasma for the determination of total calcium when using this product. The product can still be used to determine total calcium with serum sample." ELITech Clinical Systems SAS has revised the package insert and kit box label for the CALCIUM ARSENAZO reagent reference CALA-0250. All mentions of the use of the plasma will be removed from the documentation. If you have inventory of CALCIUM ARSENAZO, reference CALA- 0250. All lots. in your laboratory, please discard the insert that you have and replace it with the insert enclosed with this notice." If you have transferred any of this product to another location, please forward a copy this bulletin to that location." All users receive CALCIUM ARSENAZO, CALA-0250, from distributors. To be certain everyone who is possibly affected by this change is contacted, we are sending this communication to all registered users. As confirmation of receipt of this notice, please return the attached receipt confirmation to Technical Support by fax (401-642-9003) within 48 hours of receipt.We apologize for the inconvenience that this situation can cause in your laboratories; we remain at your disposal for further information. If you have any questions, please contact Technical Support at 1-855-ELITECH (855-354-324)We thank you for your cooperation and the attention which you will give to this bulletin. |  | 95 | 47 kits in US and 48 kits in Sweden | Worldwide Distribution - US including the states of KY and UT and the country of Sweden. | Terminated | December 12, 2013 | 99 | 95 | Not_Computer
Z-0181-2014 | 66267 | No | DLZ | Phenobarbital test system. | Toxicology | Toxicology | 510(k) | Emit 2000 Phenobarbital Assay and Syva Emit 2000 Phenobarbital Assay |  |  | August 30, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Imprecision with certain lots | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Recall letter dated August 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease use of the affected product. Siemens will replace unused affected product at no charge. For replacements customers were instructed to incude this information on the Effectiveness Check form included with the letter. The Effectiveness Check form must be returned even if no replacement reagents are needed. The recall letter was to be retained with customers laboratory records and forwarded to those who may have received this product. Customers with questions were instructed to contact Siemens Technical Solutions or their Siemens technical support representative.For questions regarding this recall call 800-441-9250. |  | 4963 | 4963 | Worldwide Distribution - USA (nationwide) and Mexico. | Terminated | August 27, 2014 | 362 | 4963 | Not_Computer
Z-0176-2014 | 66191 | No | NGX | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | Pinook Mini Massager |  |  | May 23, 2013 | 2013 | Pinook-Usa | Failure to submit a premarket submission and gain approval of a medical device. | No Marketing Application | On October 18, 2013 the firm sent the following notification to their customers:This is to inform you of a product recall involving: Pinooks Mini-massager and Pinooks Micro Massager model number PT-11 See enclosed product label (for ease in identifying the product at retail/user level). This recall has been initiated due to failure to submit a premarket submission and gain approval.We began shipping this product on 11/27/2012. Immediately examine your inventory and quarantine product subject to recall. In addition, if you may have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.Pinook USA is removing the products (mini and micro massager) from the market until receiving FDA approval to market them again under a new 510(k).If you have stock on your end please return them to our office:901 C FL PKWY STE A6, Orlando FL, 32824.Before returning, please contact Sheva Ben-Shushan: accounting@pinookusa.com in order to receive an RMA number.Once you have RMA number, place that number on the package label.This recall should be carried out to the retail level. Your assistance is appreciated and necessary to prevent patient harm. If you have any questions, call Dvir Lev-Ran, telephone number: 407-854-0055 Ext 304. This recall is being made with the knowledge of the Food and Drug Administration. |  | 28994 | 28,994 units (Mini and Micro) | Worldwide Distribution-USA (nationwide) including Puerto Rico and Guam and the states of AZ, CA, CO, FL, IL, MA, MD, MO, MT, NC, ND, NJ, NV, NY, TX, VA, and WA and the country of Canada. | Open, Classified |  |  | 28994 | Not_Computer
Z-0175-2014 | 66475 | ph buffer | JCC | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | pH 7.0 Buffer Solution Pint and Quart Bottles, Part Number 02.0030 for Pints and 02.0031 for Quarts. Packaged in Plastic Bottles. Can be sold individually or in cartons with 6 bottles per box (for quart bottles only). |  |  | September 13, 2013 | 2013 | Mesa Laboratories, Inc. | Mesa Laboratories, Inc. is recalling certain lots of pH 7.0 Buffer Solution due to mold growth. | Material/Component Contamination | An email letter was sent to consignees on October 2, 2013 asking all customers to discard the product and a replacement product and tubing will be issued. |  | 58466 | 58,466 bottles | Worldwide Distribution-USA (nationwide) and the country of Canada. | Open, Classified |  |  | 58466 | Not_Computer
Z-0180-2014 | 66067 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | CARESTREAM DRX -Mobile Retrofit Kit, Catalog number: 1019397. Made in U.S.A. by Carestream Health, Inc. 150 Verona Street Rochester, NY 14608 |  |  | August 29, 2013 | 2013 | Carestream Health Inc. | Carestream Health Inc. has conducted a voluntary field action which includes the installation of a software patch on the following: DRX-Revolution Mobile X-Ray, DRX-Mobile and DRX Transportable Systems with DirectView Version 5.5 and 5.6 console software when used with DRX-1 Detector version 69 firmware. Carestream recently received reports of a non-diagnostic image on exams when a Normal Exposure | Software design | The firm, Carestream Health Inc., sent an "Urgent: Medical Device Recall" letter dated August 29, 2013 to its customers. The letter described the product, problem and actions taken. Carestream has developed a software patch that will allow Normal Exposure Mode and Extended Exposure Mode images to be taken in any order and will install the patch in all DRX-Revolution, DRC-Mobile and DRX-Transportable Systems with software version 5.5 or higher and Detector firmware version 69 or higher. No DRX Transportables were affected. Your Carestream service representative will install the patch at your next service call. If you have questions or concerns, or if you experience an issue prior to our visit, please contact the Carestream Customer Care Center at 1-800-328-2910, 8am through 8pm Eastern Standard Time. |  | 364 | 364 units | Nationwide distribution:AL, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MI, MO, NE, NH, NJ, NM, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WA, WI, and WV. | Open, Classified |  |  | 554 | Computer
Z-0191-2014 | 66145 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | DePuy Mitek FMS Intermediary Tubing with One-Way valve |  |  | August 29, 2013 | 2013 | DePuy Mitek, Inc., a Johnson &amp; Johnson Co. | Pillow valve included in the FMS Duo outflow tubing set and the FMS Solo intermediary tubing set may not perform as intended. may result in backflow of irrigation fluid into the One Day Set which may potentially lead to patient cross-contamination | Component change control | DePuy Mitek inititaed telephone notification on Ausgust 29, 2013 and sent an Urgent Voluntary Product Recall letter dated September 5, 2013 to US and OUS affected customers/affiliates. The customers were notified of the affected products, problem and actions to be taken. The customers were instructed to immediately check all inventories to locate and return affected product following the enclosed instructions. Stericycle is cooridianting the returns. If you have any questions or concerns in regards to this recall, please contact Stericycle directly at 1-866-737-1928. |  | 368 | 368 pk | Worldwide Distribution - USA (Nationwide) and in the countries of: Australia , Austria , Belgium , Canada , Czech Republic , Denmark , Finland , France, Germany , India, Ireland , Israel, Italy , Latvia , Netherlands , Norway , Nepal, Poland , Portugal , Spain , Sweden , Switzerland , United Kingdom. | Open, Classified |  |  | 29648 | Not_Computer
Z-0189-2014 | 66392 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMA SPECTRUM Infusion Pump with Master Drug Library component (SHARP LCD screen, 55074-2 revision B) |  |  | September 24, 2013 | 2013 | Baxter Healthcare Corp. | Baxter Healthcare Corporation is recalling the LCD display screen on the SIGMA SPECTRUM Infusion Pump with Master Drug Library for not meeting standards for withstanding an electrostatic discharge. | Under Investigation by firm | Baxter Healthcare Corporation sent an Urgent Product Recall letter dated September 24, 2013 to all affected customers via first class mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate the affected product, remove them from use and acknowledge your receipt of the recall notification by completing the provided Customer Reply Form and fax to 224-270-5467. Customers were asked to contact Baxter Healthcare Medina at 800-356-3454 (choose option 1) Monday through Friday during the hourse of 7:00 am and 7:00 pm Eastern Time to order replacement device(s). For technical questions contact Baxter Healthcare Medina at 800-356-3454, (choose option1) and for clinical questions contact Medical Information Services at Baxter 800-933-0303. |  | 346 | 346 units | Worldwide Distribution - USA Nationwide in the states of: AR, AZ, CA, CO, CT, FL, HI, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NM, NY, OH, OK, OR, PA, TN, UT, VA, WA and WV and the country Canada. | Open, Classified |  |  | 346 | Not_Computer
Z-0182-2014 | 66598 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Prostiva RF Therapy Hand Piece Kit, REF 8929, Rx only, sterilized using Ethylene Oxide. |  |  | October 21, 2013 | 2013 | Urologix, Inc. | A Hand Piece high deployment force exceeded the maximum specification. | Process change control | Consignees were sent a Urologix Urgent Medical Device Recall letter dated October 21, 2013. The letter described the problem and the device involved in the recall. Advised consignees to examine their inventory and remove the recalled product. Requested consignees to complete and fax the "Customer Acknowledgement Form" to Hope Przybilla at 763-475-1443 or call her for questions. The consignees local representatives will be personally assisting them with the return process and ensuring that replacement product is offered and received promptly. |  | 142 | 142 | Nationwide Distribution-including the states of AZ, CA, FL, GA, IL, KY, MA, MN, NC, NJ, NV, NY, OH, OR, TN, TX, and UT. | Terminated | April 21, 2014 | 182 | 142 | Not_Computer
Z-0188-2014 | 66726 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare elance Vital Signs Monitor, Model 93300, with Option S. It is a patient monitor with arrhythmia detection or alarms. |  |  | October 17, 2013 | 2013 | Spacelabs Healthcare, Llc | Spacelabs Healthcare elance Vital Signs Monitor with Option S, model 93330, is recalled because the monitor will not permit connection of any of the Spacelabs TruLink line of SpO2 sensors. | Nonconforming Material/Component | Spacelabs Healthcare sent the "URGENT -MEDICAL DEVICE CORRECTION Spacelabs Healthcare elance Vital Signs Monitor with Option S" letter, dated 01 November 2013, to the US consignee on November 1, 2013.Spacelabs Healthcare will send the customer letter (translated as necessary) to all international subsidiaries and distributors of record via email on November 6, 2013.USERS ARE ADVISED THAT:-They may not be able to monitor their patient's oxygen saturation because the monitor will not allow the TruLink SpO2 sensor to be connected. The clinician will be required to find an alternate method to monitor oxygen saturation.-They will be contacted at the earliest possible date to schedule a convenient time and Spacelabs will correct all of your facility's affected monitors at no cost.Customer in the US can call Spacelabs Healthcare, Inc. of Snoqualmie, Washington at 800-522-7025.If you have any questions about this corrective action program, please contactSpacelabs Healthcare at +86 512 87178304 for Technical Support. |  | 89 | 89 units total (1 in US and 88 International) | Worldwide Distribution-Distributed in Puerto Rico and the countries of France, Indonesia, Korea, Nicaragua, Philippines, and Taiwan. | Terminated | February 25, 2014 | 131 | 89 | Not_Computer
Z-0193-2014 | 66624 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns" Soft-FlowÂż Aortic Cannula without Suture Flange, angled tip, wire-reinforced with luer port, aortic cannula, 8.0 mm (24 Fr) OD with 3/8" connector, 14" (36 cm) long. |  |  | October 23, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | During an in-process inspection, Terumo Cardiovascular Systems (Terumo CVS) identified the presence of loose fiber particulate that exceeded finished product specifications on certain product lots of Sarns" Soft-FlowÂż Aortic Cannulae and Sarns" Venous Return Cannulae. | Employee error | On 10/23/13, Terumo Cardiovascular Systems sent an URGENT MEDICAL DEVICE RECALL LETTER to their consignees. The letter identified the reason for the device removal, identified the affected product and associated potential hazards, and provided recall instructions to their customers regarding affected product return. Questions regarding this recall are directed to CVS Customer Service M-F, 8am-6pm at 1-800-521-2818. |  | 90 | 90 units | Worldwide Distribution-USA including DC and the states of MO,TX,OK, LA, MA, CA, WI, and MI and the country of Canada. | Terminated | July 18, 2014 | 268 | 230 | Not_Computer
Z-0190-2014 | 66546 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | Sterrad 100NX, Product Code: 10104. |  |  | October 18, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the Sterrad 100NX Sterilizer some healthcare workers detected odors/smells emanating from the 100NX Sterilizers, and it may be related to normal breakdown of the oil used in the Sterrad 200 vacuum pumps. | Under Investigation by firm | A recall letter was sent to all customers on 10/18/13 who purchased the Sterrad100NX Sterilizer. The letter informed the customers of the problems identified and the actions to be taken. Customers are instructed to contact ASP Professional Services directly at (888) 783-7723 and select option #2. Customers are instructed to report any complaints or suspected problems with their Sterrad 100NX System to Stericycle at (877) 598-5705. |  | 2106 | 2106 units total (1224 units in US) | Worldwide Distribution-USA (nationwide) and the countries of EMEA, APAC, Japan, Canada. AE, AU, BE, BH, CA, CH, CL, CN, CO, DE, EC, ES, GB, GR, HK, HU, IL, IN, IT, JP, KR, KW, MX, MY, NX, PA, PH, PK, PL, PR, QA, RU, SA, SG, SI, TH, VE, VN, ZA, and Tripoli. | Terminated | January 16, 2014 | 90 | 2106 | Not_Computer
Z-0197-2014 | 66120 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMA SPECTRUM Infusion Pump with Master Drug Library. |  |  | June 10, 2013 | 2013 | Baxter Healthcare Corp. | Baxter Healthcare Corporation issued an Urgent Device Correction for the SIGMA SPECTRUM Infusion Pump with Master Drug Drug Library due to reports of upstream occlusion alarms and air-in-line alarms occurring during infusions. | Process design | On June 10, 2013 Baxter Healthcare Corporation issued an Urgent Device Correction for the SIGMA SPECTRUM Infusion Pump with Master Drug Library. The notices were dated June 5, 2013 and were distributed via first class mail. Until further information is provided, users should continue to follow the SIGMA SPECTRUM Infusion Pump Operator's Manual, which identifies a number of environmental and clinical usage factors that contribute to upstream occlusion and air-in-line alarms. For high-risk patients, consider having additional devices as backup. If the biomedical engineer finds that the sensor is not functioning properly, contact Baxter Healthcare at 800-356-3454 (choose option 1). For clinical questions, please contact Medical Information Services at Baxter at 800-933-0303, Monday through Friday 8:00am to 5:00 pm Central Time. |  | 273835 | 273,835 units | Worldwide Distribution-USA including Puerto Rico and Canada. | Terminated | April 20, 2017 | 1410 | 273835 | Not_Computer
Z-0186-2014 | 65739 | No | ONW | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Combiset Blood Line Hemodialysis use |  |  | August 29, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Updated Instructions for Use: Bloodline connections not tightened per instructions in the operator's manual may disconnect and lead to blood loss | Labeling design | Fresenius issued a Field Notification Letter dated 8/29/13 to Patient Clinics/ associated at home patients and provides additional detail instructions to ensure the connection to the transducer protector bloodline female luer interface is secure.This field advisory notice sent by Certified Mail, Return Receipt Requested. Accounts requested to return both the receipt and the completed reply form (enclosed) to confirm receipt of the notice.If you have any questions regarding this field advisory, please contact FMCNA Customer Service Team at 1-800-323-5188. Customer Service will be available Monday-Friday 7am-6pm CST. |  | 508029 | 508,029 cases | Worldwide Distribution-USA (nationwide) and the country of Canada. | Open, Classified |  |  | 540759 | Not_Computer
Z-0187-2014 | 66574 | Yes | OAD |  |  | Cardiovascular | PMA | Safire Duo Ablation Catheter MediGuide Enabled, REF A700240 and A700238, Rx only, Sterile EO. |  |  | October 11, 2013 | 2013 | St. Jude Medical | St. Jude Medical Cardiovascular and Ablation Technologies Division is performing a voluntary notification regarding selected batched of Safire Duo Ablation Catheter MediGuide Enabled. These units have a cable connector configuration that is different than intended. | Process change control | Consignees were sent a St. Jude Medical "Urgent Medical Device Notice" dated 11 October 2013. The letter was addressed to Cath lab manger / Risk Manger. The letter described the problem and the product involved in the recall. The letter stated that it is acceptable to continue using these devices as long as the product is used in accordance with the IFU. |  | 46 | 46 | Nationwide Distribution including the states of KY, OH, TX, MA, and NC. | Terminated | June 09, 2014 | 241 | 46 | Not_Computer
Z-0200-2014 | 65863 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Product Catalog (part number 20-0020) utilized with the Electric and Pneumatic Motor Systems including: MicroMax Motor Systems, XMax Motor Systems, EMax2 Motor Systems, and Emax2Plus Motor Systems, all models and serial numbers. |  |  | May 06, 2011 | 2011 | The Anspach Effort, Inc. | Anspach Effort, Inc. Palm Beach Gardens, FL initiated a correction of its medical Device labeling for its Pneumatic Systems Operating Manual, Cutting Burr Chart and Product Catalog due to a recent review of their regulations and compliance requirements that determined the claim that the device is capable of dissection of biometals, bioplastics and cement has been verified, but is is not part of th | Labeling Change Control | The firm, Anspach Effort, Inc., sent an "Urgent Medical Device Labeling Correction" letter dated June 27, 2011 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to do the following: 1) Destroy or delete any items in their inventory immediately and replace with the attached revised Operating Manual, Product Catalog and/or Cutting Bur Chart. 2) Complete and return the attached Customer Reply Form indicating disposition of the labeling materials and confirming their receipt of this letter, and fax to 1-800-327-6661. 3) If they distributed any of the products to other services or facilities, please forward this information as appropriate. For additional copies, contact Anspach Customer Support at (800) 327-6887. |  | 2020 | 2,020 | Worldwide Distribution: USA Nationwide and the countries of: Egypt, India, Germany, Turkey, Canada, Saudi Arabia, Greece, United Kingdom, Cayman Islands, Kazakhstan, Republic of Korea, Taiwan, Hong Kong, China, Lebanon, Argentina, Denmark, Belgium, Netherlands, Mexico, France, Australia, Panama, Italy, Guam, United Arab Emirates, Israel, Polan, Venezuela, Austria, Hungary, Lithuania, Norway, Malaysia, Colombia, Costa Rica, Ecuador, Jamaica, Peru, Thailand, Italy, Czech Republic, Cameroon, South Africa, Portugal, Finland, Sweden, Nigeria, Bahamas, Cambodia, Dominican Republic, Switzerland, Chile, Singapore, Spain, Syrian Arab Republic, Jordan, Japan. | Terminated | January 28, 2016 | 1728 | 2020 | Not_Computer
Z-0272-2014 | 66689 | No | LFD |  |  | Dental | 510(k) | Biotene Moisturizing Mouth Spray, with xylitol, For dry mouth symptom relief, 44.3ml/1.5 FL. OZ., Sugar Free Sweetened with xylitol and sorbitol, GlaxoSmithKline, Consumer Healthcare, L.P., Moon Township, PA 15108. |  |  | October 17, 2013 | 2013 | GlaxoSmithKline, LLC. | Label indicates, "Sweetened with xylitol and Sorbitol", however product does not contain sorbitol. | Labeling Change Control | GlaxoSmithKline sent an Urgent Medical Device Recall letter dated October 17, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all affected product. If product was further distributed customers were instructed notify their customers of the recall and the actions required. Customers were asked to complete and return the enclosed business replay card even if they have no affected product in their inventory. Customers should complete the return goods packing slip, enclose it with their return goods shipment and affix the return label to the package. Customers with questions should call Customers Service at 1-800-743-4014.For questions regarding this recall call 919-269-1194. |  | 483204 | 483,204 bottles | Nationwide Distribution . | Terminated | July 26, 2017 | 1378 | 483204 | Not_Computer
Z-0201-2014 | 66725 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Titan 3T MRI System, MRT-3010A/5. |  |  | October 28, 2013 | 2013 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems, Inc. is recalling due to a potential problem with Specific Absorption Rate (SAR)values when using the Toshiba Titan 3T System. When the system is used in combination with either the QD Head Coil (MJQH-142A) or the Large KneeSPEEDER (MJAJ-182A), SAR values lower than the actual SAR values may be displayed on the operatingmonitor. | Nonconforming Material/Component | Toshiba America Medical Systems, Inc. sent an "URGENT: MEDICAL DEVICE CORRECTION". letter dated October 28, 2013 via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS for retention. The letter listed Device, Serial Number, reason for recall, problem, corrective actions, request to customers, and contact information. |  | 18 | 18 units | USA Nationwide Distribution in the states of: CA, FL, GA, IA, NV, NY, LA, IN, PA, OR, TX, and WI. | Terminated | December 24, 2014 | 422 | 18 | Not_Computer
Z-0273-2014 | 66664 | No | EBW | Dental handpiece and accessories. | Dental | Dental | 510(k) | Osseocare Pro Console &amp; Set Osseocare Pro, Model 1600870-001, 1700470-001, and 1700471-001. |  |  | October 24, 2013 | 2013 | Nobel Biocare Usa Llc | The recall was initiated because there is a potential for malfunction of the OsseoCare Pro when used with the OseeoCare Pro iPad application software release 1.2.0.7. | Software Manufacturing/Software Deployment | Bien Air Dental send and Urgent Field Safety Notice dated October 24, 2013, to all affected customers. The letter notified customers of the recall, description of the problem, description of the product (OsseoCare Pro iPad app software release 1.2.0.7), action to be taken by user, an alternative solution, and contact details. The letter was also accompanied by an acknowlegement form. Customers with questions were instructed to call 1-800-433-2436.For questions regarding this recall call 714-282-4800. |  | 1082 | 1082 (391 US) | Worldwide Distribution - USA (nationwide) United States and Internationally to Austria, Belgium, Brazil, Canada, Croatia, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Hungary, India, Ireland, Italy, Japan, Latvia, Lithuania, Luxembourg, Mexico, Netherlands, Norway, Philippines, Poland, Portugal, Romania, Russia, South Africa, Spain, Sweden, Switzerland, Turkey, and the United Kingdom. | Terminated | April 10, 2014 | 168 | 1082 | Computer
Z-0270-2014 | 66550 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | AW VolumeShare 4 (ADVANTAGE WORKSTATION 4.3.), AW VolumeShare 5 (AW VOLUMESHARE 5 WITH ANGIOVIZ OPTION), AW Server. |  |  | September 19, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the image orientation within GSI viewer on AW products.The GSI Viewer application running on AW/Volume Share or on the AW Server when making a batch and rotating the batch group clockwise 180 degrees (Inferior to Superior) will display images in the Superior orientation. The orientation annotation on the image is correc | Software design | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated September 19, 2013 to its consignees/customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (800-437-1171 for USA and Canada or contact local GE Healthcare Applications Specialist). The customers were instructed to "Please ensure that all potential users in your facility are made aware of this safety notification and the action taken". GE Healthcare will provide a software update for all affected systems to address the issue. A GE Healthcare service representative will contact you to arrange for this correction. This activity will be performed at no cost to you.A second updated letter was sent on October 22, 2013 to correct the Affected Product Details communicated in the previous letter dated September 19, 2013. The new statement reads "All versions of GSI Viewer prior to version 2.00-OM on AW and prior to version 2.20- OB on AW Server are affected".If you have any questions, please contact us at 800-437-1171 (USA and Canada) or contact local GE Healthcare Applications Specialist. |  | 1458 | 1458 (336 US, 1122 OUS) | Worldwide distribution: US (nationwide) including states of: AL, AZ, CA, CT,DE, FL,GA, HI, IL, IN, IA, KS, MD,MA, MI,MN, MS, O. NV, NH, NJ, NY, NC, OH, OK, OR, PA, SD, TN, TX , UT, VA, WA, WI. OUS: ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, BYELO RUSSIAN, CANADA, CHINA, DENMARK, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONDURAS, HON KONG, HUNGARY, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KOREA, KUWAIT, LATVIA, LIBYA, LITHUANIA, LUXEMBOURG, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TURKEY, UKRAINE, and UNITED KINGDOM.. | Terminated | July 31, 2014 | 315 | 1458 | Computer
Z-0210-2014 | 66449 | No | HBL | Neurosurgical head holder (skull clamp). | Neurology | Neurology | 510(k) | Composite Series Skull Clamp, one clamp per tote. |  |  | August 29, 2013 | 2013 | Integra LifeSciences Corp. | An investigation of an adverse trend of complaints for the Skull Clamp index locking knob not easily unlocking or disengaging and malfunctions of the 80 lbs. torque knob was initiated. None of the complaints resulted in a report of patient injury, adverse health consequence, or a delay / prolongation of the surgical procedure. | Device Design | Integra sent an Urgent: Voluntary Medical Device Recall Notification dated August 29, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Your local Integra NeuroSpecialist will set up a time with you to discuss replacement options but in the interim, Integra is asking that you do the following:1. Review your inventory and determine if you have any MAYFIELDÂż Composite Series Skull Clamps (A3059). If so, stop using them immediately and remove them from service to prevent any inadvertent use of the Composite Series Skull Clamp.Either while the Integra NeuroSpecialist is on-site or independently;2. Complete the attached form. If you do not have affected product, check the box: I do not have Composite Series Skull Clamps.3. If you do have affected product, check the box: I do have Composite Series Skull Clamps. Record the total quantity and Serial No.(s) of all Skull Clamps in your inventory. The Serial No. is located on the clamp upright comprised of two letters (SS) followed by six numerals.4. Complete the other information and return by email or fax as indicated on the form, or give it to the Integra NeuroSpecialist who will return it for you.For assistance returning and exchanging these products or any other questions that you may have, please contact Integra Customer Service at 1-855-532-1723.We sincerely apologize for any inconvenience this may have caused and thank you for your cooperation. |  | 37 | 37 units | US Distribution including the states of AL, AR, CA, CT, FL, GA, IA, IL, MS, NY, OH, SC, TN, and TX. | Terminated | January 14, 2014 | 138 | 37 | Not_Computer
Z-0269-2014 | 66560 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips brand IntelliSpace PACS 4.4, Image Management System; The system is a software package used with general purpose computing hardware to acquire, store, distribute, process and display images and associated data throughout a clinical environment. Product is manufactured and distributed by Philips Healthcare Informatics, Inc. Foster City, CA |  |  | October 14, 2013 | 2013 | Philips Healthcare Informatics, Inc. | Philips Healthcare became aware of a software problem with the IntelliSpace PACS system, in that a timing issue has been identified that may result in a condition where the device does not store additional images received before processing is complete. The missing images may contain pathology which may result in a different diagnosis if they had been available at time of interpretation. | Software in the Use Environment | Philips sent an Urgent Field Safety Notice dated October 14, 2013, to all affected customers. The letter identified the product, problem, and actions to be taken by the customers.Please review the following with all radiology technicians who use lntelliSpaceBY CUSTOMER/USER PACS:" Assure you and your colleagues are aware of the problem." Verify that all images sent for each study are present on the iVaultbefore the images are purged from the modality.o Use the image count from the lower right corner of thethumbnails as seen in Enterprise only. Single images (such asCR, DX and MG) show no image count, so the user shouldcount those thumbnails without image counts as one. Add all thenumbers up to come up with the total number of viewableimages, stored to lntelliSpace." Compare number of images sent from the modality and those sent fromother 3rd party source to image count in IntelliSpace PACS." If the image count on the modality or 3 party source indicates there aremore images to the study than are present on the iVault, the studyshould be resent from the modality and/or 3 party source." Resending the studies will typically correct the problem but thetechnician should again verify the image count is correct.If you need any further information or support concerning this issue, pleaseINFORMATION AND contact your Philips representative or contact Care at isitecare@philips.com. Please call (650) 293-2300 for further questions. |  | 71 | 71 units | Worldwide Distribution - USA (nationwide) and country: Canada. | Terminated | June 11, 2014 | 240 | 71 | Computer
Z-0202-2014 | 66485 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes 8.5 MM Medullary Reamer |  |  | October 03, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received for intraoperative reamer head breakages which can allow for unretrieved fragments of non-implant grade material. | Device Design | Synthes sent an Urgent Notice: Voluntary Medical Device Recall Notification letters, dated October 3, 2013, to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory and immediately remove the affected lots from stock. For questions call 610-719-5450 or contact your Synthes Sales Consultant. |  | 2979 | 2979 | Worldwide Distribution: USA (Nationwide) and countries of: Canada and Mexico. | Open, Classified |  |  | 2979 | Not_Computer
Z-0271-2014 | 65792 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo.plaza image processing system |  |  | August 08, 2011 | 2011 | Siemens Medical Solutions USA, Inc | The firm became aware of an unintended behavior when using the syngo.plaza. The "Rename" functionality can cause an unintended patient merge if the patient is sent to another system. This issue affects customers who transfer patient data between two or more syngo.plaza or syngo Imaging XS systems. | Software design | Siemens sent a Customer Safety Advisory Notice dated August 3, 2011, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to forward the Customer Safety Advisory Notice to their customers. If necessary, customers should have the letter translated technically into the local language and copy it themselves. A letter was sent to direct accounts to inform them of the issu and provide instructions to avoid encountering the problem.For questions regarding this recall call 610-219-6300. |  | 19 | 19 | Nationwide Distribution including MO, MA, NY, WI, OH, NY, CA, FL, NC, MN, and WA. | Terminated | November 06, 2014 | 1186 | 19 | Not_Computer
Z-0219-2014 | 66306 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Carestream Vue PACS; Virtual Enterprise SW StandAlone |  |  | September 16, 2013 | 2013 | Carestream Health Inc. | Carestream Health, Inc. has issued a recall when using the CARESTREAM Client Workstation where under very specific conditions, some images from multi-slice studies such as CT and MR may not load into the viewer application due to a software problem. | Software design | Carestream sent an Urgent Medical Device Recall letter dated September 13, 2013, to all affected customers via Fed Ex on September 16, 2013, informing them that a software problem may have the potential to occur under certain conditions when using the CARESTREAM Client Workstation. When certain conditions are met, it appears that the child background process completes its tasks before the parent process and as a result the parent process considers its work done and terminates. In the event that a software problem does occur, the scan may not be fluent or the image numbers may not be sequential. A software patch is currently being deployed by Carestream Health. Customers were instructed that they would be contacted by a Carestream Service Engineer to coordinate the deployment. Customers with questions were instructed to call 1-866 737 2487 or they may call the Carestream Customer Care Center at 1-800-328-2910, 8AM through 8pm Eastern Standard Time.For questions regarding this recall call 1-800-328-2910. |  |  |  | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, FL, IA, IL, MA, MD, ME, MN, MO, NC, NH, NV, NY, OH, OK, PA, RI, TN, UT, VA WA, and WI and Internationally to Canada, Japan, and Kuwait. | Terminated | December 21, 2015 | 826 |  | Computer
Z-0285-2014 | 64263 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Plum A+ Infusion System is a cassette based multifunction infusion system that allows two lines in and one line out and can be used for standard, piggyback, or concurrent delivery. The Plum A+ Infusion System is designed to deliver parenteral, enteral, or epidural infusions over a broad range of infusion rates from multiple fluid container types. The Plum A+ Infusion System includes an infuser, disposable IV sets (hereafter called a set), optional accessories, and this operator's manual. |  |  | February 01, 2013 | 2013 | Hospira Inc. | Plum A+ infusers have the potential for the distal (occlusion) press sensor pin to break. The distal pressure sensor pin is a part of the overall subsystem that measures the pressure within the distal line of the administration set and indicates the presence of a full or partial distal occlusion. A broken distal pin can only be detected via a visual inspection of the distal pressure pin. A broken | Component design/selection | Hospira sent an Urgent Device Field Correction letter dated February 1, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter recommends that facilities immediately visually inspect their Plum a+ devices to determine if the distal pressure pin is broken or damaged via instructions included in the letter. If a broken or damaged distal pressure pin is observed, remove the device from service and contact Hospira at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM, CST) to report the issue. If the return of defective pumps significantly impairs a facility's ability to operate normally, Hospira will provide loaner pumps. The letter also encourages customers to insert cassettes into pumps in accordance with the instructions found in the operator's manual. Customers were asked to complete the attached reply form and return it via fax to the number on the form. Customers with questions were instructed to call 1-800-441-4100 or 1-800-241-4002, option 4.The letter also asked direct accounts that further distributed the pumps notify their customers of the URGENT DEVICE FIELD CORRECTION and ask them to contact Stericycle at 1-866-201-9068 to receive a reply form.. For questions regarding this recall call 800-441-4100, |  | 71 | 71 units | Worldwide Distribution - USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico and U.S. Virgin Islands; and countries of: Costa Rica and Dominican Republic. | Open, Classified |  |  | 231084 | Not_Computer
Z-0291-2014 | 66397 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay System, Part Number: 81600N |  |  | September 12, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling the Access 2 Immunoassay System because they determined that the fuses in the system may be of the wrong amperage. | Mixed-up of materials/components | A phone call was made on 9/12/13 to customer who received the Access 2 System. The phone script explains the problem identified of a possible wrong fuse and may result in failure of the fuse to operate properly. Customer was instructed that their Field Service Engineer, will call them back the following week to check on their system and replace the fuses if necessary. If you have any questions concerning this information, please call 952-393-8546. |  | 8 | 8 units total (1 unit in the US) | Worldwide distribution: US (nationwide) in the state of: TX and countries of: Colombia, Hong Kong and Korea. | Terminated | November 14, 2013 | 63 | 8 | Not_Computer
Z-0280-2014 | 66444 | No | JDZ | Cement mixer for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Acetabular Pressurizer; single use and six units per box. |  |  | November 01, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | On March 14, 2013 a Work Order was opened for 120 units (20 boxes) of the Acetabular Pressurizer (P/N 0206-522-000 Rev. H). On March 15, 2013, when the product was going to be placed in the pouch to be sealed, the manufacturing operator noticed that some assemblies were cracked. As a result of the injection pressure of the cement gun, a cracked tiplet may release from the Hub during cement applica | Nonconforming Material/Component | The firm, Stryker Instruments, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 11/01/2013, to all domestic consignees/customers and an email was sent to the International Subsidiaries and Sales Representatives. The letter described the product, problem and actions to be taken. The customers were instructed to do the following:1. Immediately review this Recall Notification2. Immediately check all stock areas and/or operating room storage and quarantine any affected product found.3. Complete the enclosed Business Reply Form to confirm receipt of this notification and identify how many, if any, affected items are currently in your inventory. Please complete and return the Business Reply Form even if youdon't have any affected product on hand.Note: Your signature on the BRF indicates that you received and understand this Notification and have followed the instructions in this Notification.4. If you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF.5. Fax the completed Business Reply Form to Stryker Instruments Regulatory Department, 866-521-2762.6. Upon receipt of the Business Reply Form, if you have recalled product on hand, Stryker will email a pre-paid shipper. This shipper can be used to return the recalled Acetabular Pressurizers to Stryker.7. Upon receipt of recalled Acetabular Pressurizers, credit will be applied to your account.For questions regarding this recall please contact Stryker Instruments: Monday-Friday 8am-5pm (EST) at 269-389-2921 or email: kellyjo.whipple@stryker.com. |  | 87 | 87 boxes (522 each) | Worldwide distribution: US (nationwide) including states of: TX, FL, IL, IA, ND, WI, VA, and NJ.; and countries of: CANADA, POLAND and SWITZERLAND. | Terminated | February 10, 2014 | 101 | 87 | Not_Computer
Z-0277-2014 | 66199 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VerSysÂż Cemented Revision Calcar (CRC) Hip System Femoral Stems |  |  | August 28, 2013 | 2013 | Zimmer, Inc. | Based on a review of historic packaging validations, it was determined that a subset of product packaged in one of the clean rooms at a single Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. The number of samples required for packaging testing (seal integrity and seal strength) to provid | Packaging process control | Removal of the subject devices from the field including distributor and hospital inventories. Distributors and Risk Managers and Surgeons at hospitals that received a direct shipment from Zimmer or its Distributors received a notification on 09/19/2013. |  | 12866 | 12866 units | Worldwide distribution: US (Nationwide) including Puerto Rico and Virgin Islands; and countries of: Australia, Algeria, Belgium, Bolivia, Canada, Canary Islands, Chile, Croatia, Finland, France, Germany, Greece, Guatemala, Honduras, India, Israel, Italy, Japan, Jordan, Lima-Peru, Malaysia, Morocco, Nicaragua, Portugal, Saudi Arabia, Seoul Korea, Serbia, Singapore, South Africa, Spain, Switzerland, Turkey, Utd.Arab.Emir., United Kingdom, and Venezuela. | Terminated | April 30, 2015 | 610 | 12866 | Not_Computer
Z-0281-2014 | 66600 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Ray Station Radiation Therapy Treatment Planning System, Model Numbers 3.0, 3.5 and 4.0 |  |  | October 23, 2013 | 2013 | RAYSEARCH LABORATORIES AB | Ray Search Laboratories found an issue with the dose calculation for Elekta motorized wedges in Ray Station 3.0, 3.5 and 4.0. The error presents when a wedge is used for a beam and the wedge angle differs from 60 degrees. | Software design | Medical Device Correction Letters dated 10/23/13 were e-mailed to the user accounts on 10/23/13. |  | 195 | 195 units domestically | US Distribution in the states of: California, Colorado, Delaware, Florida, Maryland, Missouri, New Jersey, New York, Ohio, Pennsylvania, Tennessee, and Virginia. | Open, Classified |  |  | 195 | Not_Computer
Z-0282-2014 | 66537 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Draeger Movita 603x/DVE 803x Ceiling Supply Unit. It is intended for use in operating theaters and intensive care |  |  | September 30, 2013 | 2013 | Draeger Medical, Inc. | Cases were encountered with the Draeger Movita 603x/DVE 803x Ceiling Supply Units where the drive screw in the motor unit broke and the lift arm fell down approximately 60 cm (23.6 inches). | Equipment maintenance | An "Urgent Medical Device Recall" notification letter was sent to users in September 2013. The issue was described and recommended actions were provided. Customers may contact (800) 543-5047 and select 1, then 2. |  | 139 | 139 | Worldwide distribution: US (Nationwide)and in the country of Canada. (note: the ceiling supply units are not classified as a medical device by Health Canada). | Terminated | September 03, 2014 | 338 | 139 | Not_Computer
Z-0279-2014 | 66506 | No | MHY |  |  | Neurology | PMA | Medtronic Activa PC, model 37601, and Activa RC, model 37612, Activa SC, models 37602 and 37603. Multi-program Rechargeable Neurostimulator for Deep Brain Stimulation. |  |  | September 30, 2013 | 2013 | Medtronic Neuromodulation | Medtronic notified customers regarding Loss of Stimulation of specific Implantable Neuro-Stimulators under the following specific conditions: Switching between groups with Multiple programs or creating program groups. Therapy can be restored with the Physician Recharge mode (RPM) of the Implantable Neurostimulator Recharger (INSR). | Software design | In September 2013 Medtronic issued a medical device correction letter titled "Loss of Stimulation and Over Stimulation" to customers associated with the affected Medtronic neurostimulators. This letter informed customers of the potential for these issues to occur and provided recommendations to significantly reduce the likelihood of the issues. Medtronic field representatives will be updating the model 8870 software application card to version AAR/01, which is contained in the physician programmer. A Physician Reply Form was asked to be signed and returned by fax or mail for acknowledging the receipt of the Urgent Medical Device Correction letter. When the Medtronic field rep updates the software application card of the physicains' programmer, they will complete a Rep Reply For: 8870 Card Update to AAR/01. |  | 62634 | 62,634 | Worldwide distribution: US (Nationwide) including Puerto Rico; and countries of: Algeria, Aruba, Austria, Belarus, Belgium, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Faroe Islands, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Hungary, Iceland, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malta, Martinique, Mexico, Morocco, Netherlands, Netherlands Antilles, New Caledonia, Norway, Oman, Pakistan, Panama, Peru, Poland, Portugal, Qatar, Reunion, Romania, Russian Federation, San Marino, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, and Venezuela. | Terminated | November 20, 2015 | 781 | 126797 | Computer
Z-0206-2014 | 64091 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+3 Infusion Pump System; |  |  | February 05, 2012 | 2012 | Hospira Inc. | Hospira has received reports of distal pressure sensor calibration drift on Plum A+ Infusers. The following error codes may be displayed during setup or infusion signifying possible calibration drift: E180/N180, E181/N181, E186/N186, E187/N187 or E346. Additionally, if the pressure sensor calibration has drifted, users may experience false distal occlusion alarms or distal occlusions that may not | Nonconforming Material/Component | Hospira sent an URGENT DEVICE FIELD CORRECTION letter dated February 5, 2013 was sent to all direct accounts notifying them of the issue. For device users who experience these issues, Hospira recommends the following actions be taken: 1) Remove the pump from service and perform a Performance Verification Test (PVT) Distal Occlusion Test to determine if the distal pressure sensor is performing correctly. 2) If the device does not pass the PVT, it may indicate that the distal pressure sensor has drifted. Perform the appropraite troubleshooting and repair activities defined by the user's facility, whihc may include returning the device to Hospira for further diagnosis and servicing. The remediation for this issue is to perform a recalibration of the distal pressure sensor. Hospira also intends to release in 2013 a recommendation of a yearly test to determine if the distal pressure sensor's calibration has drifted. Customers who suspect that their device(s) may be affected by this issue are urged to contact Hospira at 1-800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) to report the issue. Hospira will provide loaner devices if necessary. Customers are requested to complete and return the reply form included with the notification letter per the instructions on the reply form.For questions regarding this recall call 224-212-2000. |  | 214 | List No. 12348: 214 units | Worldwide Distribution - USA (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico and U.S. Virgin Islands; and countries of: Croatia, Dominican Republic, Turks and Caicos. | Open, Classified |  |  | 357778 | Not_Computer
Z-0172-2014 | 65795 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee systems |  |  | January 21, 2013 | 2013 | Siemens Medical Solutions USA, Inc | The firm became aware of a potential issue with Artis zee systems bearing model numbers 10094135, 10094137, and 10094141 equipped with a card collimator from a specific lot where a dislocation of the scatter radiation filter inside the card collimator may occur. | Device Design | Siemens sent an letter to all affected customers in January 2013, notifying them of the potential issue and urging customers to make appointments with Siemens to have software update AX 027/12/S realized in order to resolve the problem. The software update was made available from January 22, 2013, on. Customers were also asked to pass along this information if the device was no longer in their possession.The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 17 unit(s) will consist of the following:1. The firm sent out a customer notification letter in January 2013 notifying them of the issue and highlighting the approved procedures of the Corrective Action Plan.2. Siemens directed in the customer notification letter, for the end-user to make an appointment with the Customer Service organization to have safety update AX 027/12/S realized.3. As stated in the safety update AX 027/12/S and in the e-mail response the update to correct this issue will be provided to the customer at no expense.4. Siemens indicated in the customer notification letter that the AX 027/12/S update will be made available from January 22, 2013 and on.5. The AX 027/12/S update provides general information, work instructions and a completion certificate on how to correct the dislocated scatter radiation filter inside the collimator.You may proceed with implementation of the CAP (subject to the conditions noted above). For further questions please call (610) 219-6300. |  | 17 | 17 | US Distribution including the states of AZ, KY, PA, IA WV, FL, MO, OK, CA, AR, MT, NY, KS, and CT. | Terminated | October 29, 2014 | 646 | 17 | Computer
Z-0274-2014 | 65760 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | S-LIFT Extension Shims |  |  | August 28, 2013 | 2013 | SpineFrontier, Inc. | Fit variation among all lots of S-LIFT extension shims with corresponding retractor blades. The firm's investigation revealed that extension shims may not fit down all retractor blades. | Device Design | SpineFrontier sent an Urgent Voluntary Recall letter dated August 28, 2013, to all affected customers. The letter identified teh product, the problem, and the action to be taken by the customer. Customers were asked to initiate the immediate return of the parts noted and complete an acknowledgement form and return it to SpineFrontier via fax or email. Customers with questions were instructed to call 978-232-3990. For questions regarding this recall call 978-232-3990. |  | 67 | 67 | Nationwide Distribution including FL, TX, OK, MO, IL, OH, VA, OR, CO, KS, CT | Open, Classified |  |  | 67 | Not_Computer
Z-0276-2014 | 66450 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 3.5MM LOCKING SCREW, T15 HEXALOBE, SELF TAPPING, 34MM |  |  | September 26, 2013 | 2013 | OrthoPediatrics Corp | The firm identified, through complaints that the 3.5MM LOCKING SCREW packaging material was marked as a 3.5mm screw and the part number was correct but that the description indicated a 7.0mm screw. Upon investigation, it has been determined that a lot number 009U2 was packaged and mislabeled internally by OrthoPediatrics. | Labeling mix-ups | OrthoPediatrics sent an Urgent Medical Device Recall letter dated September 26, 2013, to all affected customers.The letter identified the product the problem and the actions needed to be taken by the customerThe sales representatives and distributors who have been identified via shipping records as having received one or more discrepant devices affected by this Recall were requested to quarantine the devices identified in this Recall by telephone on 18-September-2013. If you or any of your customers have any of these devices in the original packaging please take possession of them and quarantine them in preparation for returning them to OrthoPedlatrics. They are misbranded. If the Screw has already been removed from the original packaging and the original packaging has been discarded and the Screw placed in a sterilization tray, it need not be returned.For further questions please call ( 574) 268-6379. |  | 7 | 7 units shipped | Worldwide Distribution - US including Michigan and the countries of Ireland and Australia | Terminated | December 26, 2013 | 91 | 7 | Not_Computer
Z-0275-2014 | 66608 | No | HOZ | Ophthalmic sponge. | Ophthalmic | Ophthalmic | 510(K) Exempt | DeRoyal (R) Surgical Eye Spears, REF 30-049, Sterile EO, Manufacturer: DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849, MADE IN GUATEMALA |  |  | August 12, 2013 | 2013 | DeRoyal Industries Inc | One lot of eye spears was labeled as being x-ray detectable, but contained non x-ray detectable product. | Employee error | DeRoyal Industries sent an Urgent Voluntary Recall letter dated August 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease usage and return the product to:DeRoyal RecallsRA#: Eye Spears1755 Hwy 33 SouthNew Tazewell, TN 37825Ship via UPS Ground using Shipper #312780 Distributors were requested to notify their consignees of the same. The recall is extended to the user level. Customers with questions were instructed to call 1-800-251-9864.For questions regarding this recall call 865-362-2334. |  | 726 | 726 cases (7242 units) | Nationwide Distribution | Terminated | August 04, 2014 | 357 | 726 | Not_Computer
Z-0331-2014 | 66547 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ALTRUSÂż THERMAL TISSUE FUSION, Tissue Fusion Handpiece, 5 mm O.D., 16 cm Shaft, REF 60-9520-001, STERILE -- CONMED CORPORATION 525 FRENCH RD, UTICA, NY 13502 --- The ConMed AltrusÂż Thermal Tissue Fusion System is comprised of a dedicated energy source and disposable handpiece used to ligate (seal) and divide (cut) blood vessels and tissue bundles that fit into the jaws of the handpiece. The AltrusÂż System utilizes a thermal energy platform to achieve the desired clinical effect. The ConMed AltrusÂż Thermal Tissue Fusion handpiece is a single use device which is provided sterile. The device uses scissors to parallel jaw closure mechanism with one flat jaw and one crowned jaw in which the vessels/tissues are grasped and through which pressure and heat are applied. It is a multi-functional device capable of vessel sealing, grasping and dissecting during open general surgical procedures. The two jaw sizes allow for applications in a variety of clinical environments for access in confined areas as well as large open areas. |  |  | October 23, 2013 | 2013 | ConMed Corporation | The affected lots were exposed to a higher than acceptable level of particulate matter during manufacture and subsequently distributed. ConMed evaluated the potential risk that this issue could present to the patient and concluded the risk of injury to be negligible; however exposure could cause a minor, local, self correcting infection. | Material/Component Contamination | The firm, ConMed, sent an "URGENT: VOLUNTARY DEVICE RECALL" letter dated October 23, 2013 and Business Reply Forms to the domestic consignees/customers via USPS Priority Mail. Distribution of the recall letter to the foreign consignees will follow once the letter has been translated into the different foreign languages. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately stop the use of these devices; review your inventory for any of the devices listed; contact all of those organizations within your facility and any other facilities that you may have supplied or given these affected products to; complete and return the ATTACHMENT I Effectiveness Check Business Reply Form via fax to: ConMed Recall Coordinator at 315-624-3225.If you have any of the affected devices, please complete Attachment I and return it with the devices to: ConMed Corporation, 525 French Road, Utica, NY 13502Attn. Ed Kovac, Return via: FedEx Account # 487553646. Please indicate with your returned devices, if you are requesting credit or replacement of the devices. Please do not return used devices. If you do not have any devices to return, please complete Attachment I, indicating you have no devices and fax it to 315-624-3225, Attn: ConMed Recall Coordinator. For questions, please contact ConMed Recall Coordinator at 315-624-3237 or e-mail altrus@conmed.com. |  | 65 | 65 handpieces (Domestic: 41 devices; Foreign: 24 devices) | Worldwide distribution: US (Nationwide) and countries of: Canada, Germany and Italy. | Open, Classified |  |  | 212 | Not_Computer
Z-0335-2014 | 66601 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-Ray System, |  |  | October 11, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | If the couch, with the footrest extension attached, is positioned where the table travel is within the bore of the gantry, and the user begins to lower the couch, the footrest extension can contact the gantry cover or come in contact with the user due to the length of the accessory. | Device Design | Philips sent an Urgent - Medical Device Correction letter dated October. 11, 2013, to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue the use of the footrest extension. Philips also, stated they will replace the old footrest extension with the current one in use. For questions contact your local Philips representative or local Philips Healthcare office. |  | 9 | 9 units | Worldwide distribution: USA (nationwide) Distribution in the states of DC, GA, IL, IN, NY and OH and the countries: Germany, Israel and Netherlands. | Terminated | August 04, 2015 | 662 | 9 | Not_Computer
Z-0329-2014 | 65352 | No | FEQ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Olympus Endoscopic Flushing Pump, Model OPF-2. |  |  | May 20, 2013 | 2013 | Olympus America Inc. | There is a possibility that the pump head was not shipped in the locked position. If the pump head is in an unlocked position this may, under certain circumstances, cause the pump to stop and go into power stand-by mode. | Component design/selection | Notification letters were sent to customers on May 20, 2013. |  | 263 | 263 | Worldwide distribution: USA (nationwide) and the country of Canada. | Terminated |  |  | 263 | Not_Computer
Z-0333-2014 | 66635 | No | HJO | AC-powered slitlamp biomicroscope. | Ophthalmic | Ophthalmic | 510(k) | Zeiss brand IOLMaster, Version 1-5, Biomicroscope, Slit-Lamp, AC-Powered, |  |  | October 23, 2013 | 2013 | Carl Zeiss Meditec, Inc. | The user database delivered for download contained an erroneous value. Specifically the "aO" constant of the lens AMO Sensar AR40E is listed as "0.0" in the internal lens database of the IOLMaster. The correct value is -2.420. This could lead to the implantation of lens with a wrong diopter value. | Software in the Use Environment | The firm, Zeiss, sent an "Urgent Medical Device Correction" letter dated October 23, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customers. Customers were instructed to immediately update the Laser Interference Biometry (ULIB) lens constants, provided as an Attachment 1 and complete and return Attachment 2- Confirmation Form via mail to: Stericycle, Inc., 6026 Lakeside Blvd., Indianapolis, IN 46278; Attn: Zeiss IOLMaster.If you have any questions, call the following number for assistance: 1 (866) 891-1922. |  | 4907 | 4,907 units | Nationwide distribution. | Terminated | April 17, 2014 | 176 | 6108 | Not_Computer
Z-0330-2014 | 64806 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | March 12, 2013 | 2013 | GE Healthcare It | GE Healthcare has become aware of a potential safety issue associated with the use of GE Centricity Laboratory where panel reflex does not occur if DIP.Specimen is blank. Panel reflexing fails to occur when ad-hoc re-routing functionality is used and the following are true: 1) The ordered panel is a slave panel, and; 2) The panitem specimen group field of one or more of the ordered items on the ma | Software design | GE Healtchare sent an Urgent Medical Device Correction letter dated March 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter requested that customers review their workflows to determine whether this series of events could happen at their sites. Customers who self-identified this series of events as a possible workflow are recommended to contact their GE service representative for a software correction. These customers may also elect to identify master panels that may be re-routed ad hoc and then to ensure that the specimen group field is updated to NULL if currently set to blank.Customers with any questions about this issue are advised to contact their GE service representative or the GE Healthcare IT Help Desk: North America at tgg.nahelpdesk@med.ge.com.For questions regarding this recall call 847-277-4468. |  | 7 | 7 units | Worldwide Distribution - USA (nationwide) including GA and Internationally to Canada and the United Kingdom. | Open, Classified |  |  | 7 | Computer
Z-0264-2014 | 66567 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | ProViaÂż CROSSING GUIDEWIRE; Models: 15PROV300HS, 3PROV180HS, 3PROV180SS, 3PROV300HS, 3PROV300SS, 9PROV180HS, 9PROV180SS |  |  | October 21, 2013 | 2013 | Medtronic Vascular | Potential for PTFE (polytetrafluroethylene) coating to delaminate and detach from Guidewire. Medtronic guide wires are steerable guide wires that are used for the introduction and placement of diagnostics or interventional devices inthe coronary and peripheral vasculature, and may be used to reach and cross atarget lesion. Medtronic guide wires are not intended for cerebral vasculature.Medt | Nonconforming Material/Component | The firm, Medtronic, sent a "URGENT MEDICAL DEVICE RECALL" letter dated October 21, 2013 (in the US via UPS overnight service, 1-day delivery) to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately remove and quarantine all potentially affected guidewires that remain in inventory; return the affected product to Medtronic; if replacement product is needed, your Medtronic representative can assist you with identifying suitable replacement product, and to complete and return the Customer Confirmation Certificate confirming receipt of the notification via fax to Medtronic at 651-367-0612 to the attention of Customer Focused Quality (100% customer confirmations will be required).Should have any questions, please contact your Medtronic Representative or call 763-526-2513.Notifications to customers concerning the additional affected product began December 19, 2013 (in the US via UPS overnight service, 1-day delivery). Customers are requested to return a Confirmation Certificate confirmingtheir receipt of the notification. 100% customer confirmations will be required from the additional consignees. |  |  |  | Worldwide distribution: US (Nationwide) including states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, and WV; and countries of: Albania, Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Guyana, Honduras, Hong Kong,Hungary, Iceland, India, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malaysia, Mexico, Montenegro, Nepal, Netherlands, New Zealand, Nepal, Nicaragua, Norway, Oman, Poland, Portugal Puerto Rico, Qatar, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, United Kingdom, Uruguay, and Yemen. | Terminated | December 15, 2014 | 420 |  | Not_Computer
Z-0292-2014 | 65572 | No | LDP | Acetaminophen test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products ACET Slides, 90 slides/pack; catalog number 1037274 |  |  | June 12, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics is recalling Slide Dispense cartridges used with some VITROS Chemistry Product Slides due to slide dispense issues. | Equipment maintenance | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated June 12, 2013, to all affected customers via FedEx regarding Slide Dispense issues with various VITROS Chemistry Product Slides. VITROS ALT, CL-, GGT, LDH, K+, Na+ and TP slides use a different ABP, they are not affected by this issue. OCD provided instructions that may allow their customers to reposition the ABP if a jam occurs. OCD will provide you with a credit for any cartridge where the ABP becomes misaligned within the cartridge body in such a way that you are unable to reposition the ABP successfully to resolve the slide dispense issue. In addition to the Confirmation of Receipt form, OCD also enclosed three (3) copies of a Request for Credit form for you to complete when you require credit for affected cartridges. Complete and return the Confirmation of Receipt by June 21, 2013. If you experience slide dispense condition codes more than once a week, call our Customer Technical Support at 1-800-421-3311.For questions regarding this recall call 585-453-4224. |  | 16732 | 16732 units (12208 units domestically &amp; 4524 internationally) | Worldwide Distribution - USA (nationwide) including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY including Puerto Rico and Internationally to Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Panama, Singapore, Venezuela, England, Germany, France, Italy, and Spain. | Terminated | April 11, 2017 | 1399 | 16732 | Not_Computer
Z-0338-2014 | 66763 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Vascu Guard, Peripheral Vascular Patch with Apex Processing. |  |  | October 28, 2013 | 2013 | Synovis Surgical Innovations, Inc. | Synovis, a subsidiary of Baxter International Inc., is recalling certain shipments of VASCUGUARD and DURA-GUARD because Synovis, a subsidiary of Baxter International Inc., is recalling certain shipments of VASCUGUARD and DURA-GUARD because the shipping configurations in which the 6 fluid ounce units containing a specific revision of the Instructions for Use (IFU) shipped in a divided cell shipping | Packaging change control | A customer notification letter dated 10/28/13 was sent to all customers who purchased the Vascu Guard and the Dura Guard. The letter informs the customers of the problem identified and the actions to be taken. Customers with any questions regarding the recall are instructed to contact Customer Service unit at (800) 423-2090. If customers are distributors or if they have have distributed the recalled products to any other health care provider, they are instructed to notify their accounts and provide them with a copy of the recall notification. If customers would like Baxter to perform the recall task for them, then they are instructed to contact Stericycle at 1-866-912-8417. |  | 6014 | 6014 units | Nationwide Distribution. | Terminated | June 03, 2014 | 218 | 6460 | Not_Computer
Z-0355-2014 | 66558 | No | ONA | Stereotaxic instrument. | Neurology | General &amp; Plastic Surgery | 510(k) | Sterile Disposable Kit ( Contains Skin Tensioner ASY-25067; Three Mark Needle Assembly ASY-25094; Nose Cone Assembly ASY-25098; Dissection 4 mm PRT-25117, follicle Trap PRT 23999): |  |  | October 13, 2013 | 2013 | Restoration Robotics | Reports of discoloration (residue) on the pins of the skin tensioner have been received by several customers. Discoloration due to insufficient cleaning of stainless steel pins during passivation process. The residue hasbeen identified as a compound of iron and chlorine. | Process control | Restoration Robotics sent an Urgent Recall letters dated October 13, 2013 to all distributors via first class mail. Direct accounts were notified by phone as well as fax. Customers were informed of the affected product, problem and actions to be taken. Replacement products are being distributed to each site and will be delivered to each customer. For questions call 408-883-6760 |  | 515 | 515 US, 315 International (all packaging cnfigurations) | Worldwide Distribution - USA Nationwide in the countries of: Korea, Japan, Singapore, Thailand, United Kingdom, Canada, and Taiwan. | Terminated | December 10, 2013 | 58 | 515 | Not_Computer
Z-0357-2014 | 66439 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Trinica Anterior Lumbar Plate (ALP) system, part 07.01058.001. Non Sterile, Rx only. The Trinica Anterior Lumbar Plate System is a temporary supplemental fixation device consisting of a variety of shapes and sizes of plates and screws. The Trinica Anterior Lumbar Plate System is used as an implant for the correction and stabilization of the spine. This system provides temporary stabilization and augments the development of a solid spinal fusion. Additionally, this system provides the surgeon with the ability to supplement an interbody device with anterior plate fixation. The Trinica Anterior Lumbar Plate System components can be locked into a variety of configurations and each construct may be customized to individual cases. The plates are low profile and anatomically designed to provide optimal fit from either anterior or anterior-lateral approach. This system also features anti-migration locking caps to help secure the fixation screws. All Trinica Anterior Lumbar Plate System implant components are made from titanium alloy (Ti-6AI-4V). |  |  | October 09, 2013 | 2013 | Zimmer, Inc. | Zimmer Spine, Inc. is initiating a correction regarding the Trinica ALP Instrument Tray, part 07.01058.001 (all lots). The Ratchet Handles, part 07.00438.001 may not be effectively sterilized if left in their designated location(s) in the instrument tray. This is due to the current bracketing in the tray potentially interfering with complete steam penetration into the cannulated handle. | Device Design | Zimmer Spine sent a Medical Device Correction letter dated October 9, 2013 via FedEx A Medical Device Correction letter, dated October 09, 2013 via FedEx to distributors, Hospitals and Surgeons who have received and/or used the device. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification and ensure affected personnel were aware of its contents. Only sterilize the Ratchet Handles external to the instrument tray. Customers were to be notified when replacement trays were available. Customer with questions were instructed to call the Zimmer Customer Call Center at 866-774-6368.UPDATE: July 17, 2014 Zimmer issued another letter to consignees to inform them of a redesigned tray is available and will replace the current tray. A Zimmer representative will conduct the swap-out of the current Rev A tray for the new Rev B design tray. A Return Confirmation Form is asked to be emailed back to CorporateQuality.Postmarket@zimmer.com. For assistance or questions, contact Zimmer Spine Customer Service at 866-774-6368. |  | 94 | 94 Tray sets | Worldwide Distribution - USA (nationwide) and Internationally to Nationwide and OUS countries include: Australia, Austria, Belgium, Canada, France, Germany, Italy, New Zealand, Spain, Switzerland, and the UK. | Terminated | February 24, 2015 | 503 | 94 | Not_Computer
Z-0340-2014 | 66712 | No | NLH | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | Reprocessed Diagnostic Electrophysiology (EP) Catheters and Diagnostic EP Catheter Cables, Model Numbers: D6-DR-252-RT, F6-QF-252RT, 201101, 201104, 401442, 401860, D7-A20-131-RT, 401661, 401972, 401977, 401981, 401982, C6-MR/MST4-SA, CB3434CT. |  |  | October 18, 2013 | 2013 | Stryker Sustainability Solutions | Stryker Sustainability Solutions is recalling the Reprocessed Diagnostic Electrophysiology (EP) Catheters Cables due to a potential sterility breach in the package seal. | Process control | Stryker sent an Urgent Medical Device Removal Notification letter to affected customers. The letter informed customers of the affected product, problem identified and the actions to be taken. Customers are instructed to complete the "Recall Effectiveness Form" and fax to Moira Barton-Varty, Senior Director Regulatory Affairs at (480) 809-4957. Customers with any serious adverse events or product problems are instructed to call Stryker Sustainability Solutions at (888) 763-8803. |  | 46 | 46 units total | US Distribution: in the states of: Arizona, California, and Ohio. | Terminated | December 06, 2013 | 49 | 46 | Not_Computer
Z-0346-2014 | 66069 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA 2400 System Software V4.01. |  |  | September 03, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics issued a Urgent Medical Device Correction due to Ration Parameters Flagging Behavior for All Software Versions utilized on ADVIA 1200, 1650, 1800 or 2400 Chemistry Systems. | Software design | On September 3, 2013, Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice for the ADVIA 1200, 1650, 1800 and 2400 Chemistry Systems due to a software issue. The software issue has caused the Chemistry Systems to generate reports without error flags. Notices were sent to domestic customers via U.S. and outside the US the distribution is determined at the country level. The customers were instructed to contact Siemens Healthcare Diagnostics Technical Solution Center if they have technical questions regarding this issue at 877-229-3711 (prompt 18). |  | 876 | 876 units | Worldwide Distribution-USA (nationwide) and the countries of Sweden, Switzerland, Taiwan, Turkey, United Kingdom, and Vietnam. | Open, Classified |  |  | 3759 | Computer
Z-0337-2014 | 64805 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Centricity Perinatal and Centricity Intensive Care is intended to be used in clinical departments healthcare delivery systems such as a Labor and Delivery, Postpartum Maternal Care, newborn Nursery, critical care units including Neonatal Intensive Care Units (NICU), and may also be used in physicians offices and outpatient clinics. This product is primarily intended to serve the purpose of electronic documentation of clinical data and is designed to accept, transfer, display, calculate, store and manage clinical data. System capabilities provide the user with the ability to acquire data from the medical devices and to document, annotate, display, store, print, export and retrieve patient clinical practice information. This product is intended for professional use only. |  |  | March 11, 2013 | 2013 | GE Healthcare It | 1) Under rare network conditions, a Centricity Perinatal process may maintain an incorrect count of Patient slots in use which may lead to an inability to access the patient roster. The user is temporarily unable to review or edit form and chart information. This could result in a potential delay of care; however active fetal strips and permanent patient data storage are not affected. 2) When | Software design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 11, 2013, to all affected customers. The letters requested that direct accounts ensure that all potential users of the software were made aware of the potential problems and the recommended actions. Direct accounts were asked to return a customer response form and to return it to GE Healthcare via fax at 847-939-1522 or e-mail at QARAFMl1@ge.com.Version 6.94 of the software has been modified to correct these defects.GE Technical Support: 800-433-2009. |  | 1132 | 1,132 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of: Australia, Canada, Denmark, Germany, Hong Kong, Korea, South Korea, Thailand and United Kingdom. | Open, Classified |  |  | 1132 | Computer
Z-0351-2014 | 66727 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare qube Compact Patient Monitor, Model 91390. It is a patient monitor with arrhythmia detection or alarms. |  |  | October 17, 2013 | 2013 | Spacelabs Healthcare, Llc | Spacelabs Healthcare qube Compact Patient Monitor, Model 91390, may fail to power ON due to incorrect programmed microcontroller component. | Nonconforming Material/Component | Spacelabs Healthcare sent the "URGENT -MEDICAL DEVICE CORRECTION Spacelabs Healthcare qube Compact Monitor" letter, dated 01 November 2013, to the US consignees/customers on November 1, 2013. Spacelabs Healthcare will send the customer letter (translated as necessary) to all international subsidiaries and distributors of record via email on November 6, 2013. The letter described the product, problem and actions to be taken. The customers were instructed to advise staff of the potential for the monitor to fail to turn ON.Spacelabs Healthcare will contact their customers at the earliest possible date to schedule a convenient time for Spacelabs to correct all of your facility's affected monitors at no cost. If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 54 | 54 units total (42 units in the US and 12 outside US) | Worldwide distribution: US (nationwide) including states of: Georgia, Idaho, New York, Oklahoma, Tennessee, Texas, and Wyoming and in countries of: Bahrain, Canada, France, Great Britain, Poland, and The Netherlands. | Terminated | April 02, 2014 | 167 | 54 | Not_Computer
Z-0342-2014 | 66688 | No | JEB | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TruSystem Surgical Table (TS7000)  Leg Section One Part Light HU |  |  | October 28, 2013 | 2013 | Trumpf Medical Systems, Inc. | The coupler point (hood) of the Leg Section One Part Light may become detached from the leg section frame. | Packaging | TRUMPF sent an Urgent Safety Information letter dated October 28, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed to inspect Leg Section One Part Light for damage to the coupler point (hook) between each case. If any damage is noticed to the coupler, they were instructed to stop use immediately and contact TRUMPF Technical Service. If they drop or collide the Leg Section One Part Light into anything then stop use immediately and contact TRUMPF Technical Service. Review the User Instruction for proper use and care of the device. The leg section should be used only as intended as part of the TRUMPF TS7000 surgical table. The intended use is to support and position the patients legs during the pre-operative and post-operative stages of surgery. The leg section should be removed during surgery. TRUMPF will be exchanging the hooks, counter plates and screws of the coupler point to increase the area that the load is distributed over. This will allow the screws securing the hood to the leg frame to withstand external forces. TRUMPF anticipates parts to be available the first week of November. TRUMPF trained technicians will complete the TRUMPF Medical Systems, Inc. for the part exchange. Completion is anticipated by November 11, 2013.Further questions please call (843) 534-0606 |  | 15 | 15 components | US Distribution including the states of CA, IL, NJ, NY, NC, TX, WA and SC. | Terminated | September 16, 2014 | 323 | 15 | Not_Computer
Z-0348-2014 | 66773 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Candela Laser GentleLase Pro, Powered Laser Surgical Instrument for dermatological use - Model number: 9914-00-9015. |  |  | November 06, 2013 | 2013 | Candela Corporation | Unintended single pulse maybe emitted to handpiece prior to pressing of trigger or footswitch and cause injury. | Software design | Candela notiifed Distributors on 11/8/13 through e-mail, Certified return receipt request, or FedEx package containing distributor letter, customer letter for translation and customer lists. A software update to Version 2.0 will be issued to the user by a field service representative. |  | 1258 | 1258 units | Worldwide distribution: US (nationwide) and Internationally to: Azerbaijan, Bulgaria, Chile, Czech Republic, Greece, Israel, Kuwait, Qatar, Saudi Arabia, Serbia, Turkey, and the United Arab Emirates and to Subsidiaries in Australia, France, Germany, Italy, Japan, Portugal, Spain, and the United Kingdom. | Terminated | June 08, 2015 | 579 | 1258 | Computer
Z-0341-2014 | 66529 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Soft Tissue Retractor Small Extendible |  |  | September 24, 2013 | 2013 | Synthes USA HQ, Inc. | The potential exists for the locking screw nut to break on the Soft Tissue Retractor if it is overt-tightened. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 31, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory and immediately remove the affected lots from stock, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Complete the Verification Section at the letter by checking the appropriate box indicating affected product has been located and return the Verification Form with the product to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If customers do not have the identified product, they were instructed to complete the Offset blade facilitates easy preparation of the epipereosteal cavity for percutaneous plate insertion.e attached Verification Section and return to Synthes. |  | 374 | 374 | Worldwide Distribution - USA (Nationwide) and Canada. | Terminated | June 01, 2015 | 615 | 374 | Not_Computer
Z-0347-2014 | 65584 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Polyglycolic-Lactic Acid Absorbable Surgical Suture |  |  | December 13, 2012 | 2012 | Riverpoint Medical, LLC | Riverpoint Medical is recalling PGLA Surgical Suture (Vilet) because the label indicates the suture is undyed and it is actually a violet colored suture. | Employee error | Riverpoint sent an Urgent Medical Device Recall letter in December 2012 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to review their inventory and immediately quarantine the product in question, if found. In addition, if they have further distributed the recalled product, the consignees were advised to identify their customers and notify them of this product recall. All affected product is to be returned to Riverpoint Medical. Consignees were asked to call 503-517-8001 to return recalled product and request a credit or replacement.For questions regarding this recall call 503-517-8001. |  | 1932 | 1932 pieces | US Distribution in the states of: AZ, FL, and TX. | Terminated | November 21, 2013 | 343 | 1932 | Not_Computer
Z-0336-2014 | 66670 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Vascutek Gelsoft Vascular Graft (peripheral) |  |  | October 23, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Terumo Cardiovascular Systems (Terumo CVS) became aware of a customer complaint that a 15mm x 8mm Gelsoft graft was received which the customer believed was a 15mm x 6mm Gelsoft graft unit. The manufacturer confirmed a batch of product code 631508 was incorrectly labeled. | Error in labeling | Terumo CVS called affected consignees (via telephone)on 10/23/2013 of this device removal. Terumo CVS advised users to discontinue use and return all affected grafts in inventory. Terumo will replace or issue credit for returned grafts. Terumo provided consignees with the reason of the removal, potential hazard, and description of affected product. Terumo addressed any questions from the consignees. Consignees with questions were instructed to contact Terumo CVS Customer Service at 800-521-2818.For questions regarding this recall call 734-741-6173. |  | 9 | 9 | US Distribution in the states of: NC and CA. | Terminated | February 27, 2014 | 127 | 9 | Not_Computer
Z-0365-2014 | 66396 | No | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Chemistry Calibrator (CAL) used with the ADVIA Chemistry Direct Bilirubin_2 Reagent, Siemens Material Number (SMN) 10312279, REF 09784096, Part Number T03-1291-62; IVD |  |  | September 24, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics is conducting a field correction for Siemens Chemistry Calibrator Lots 090445B and 090445C on the ADVIA 1200, 1650, 1800, and 2400 Chemistry Systems. The System Specific Values (SSVs) for the Direct Bilirubin Assay (DBIL_2) have been reassigned for these lots. Siemens has determined that the DBIL_2 SSVs currently assigned to these lots may cause an average shift of | Under Investigation by firm | Siemens Healthcare Diagnostics sent an Urgent Medical Device Correction Letters dated September 24, 2013 and Response Forms to the US consignees via FedEx. Urgent Field Safety Notices were also distributed to customers outside the US. Customers are instructed to use the reassigned System Specific Values for the affected lots listed in the letter. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Recall letter within thirty (30) days. The questionnaire is to be faxed to the Technical Solutions Center at 302-631-7597. On September 30, 2013, Siemens sent a revised customer notifications to correct the Part Number. |  | 4243 | Domestic: 1191 units; Foreign: 3052 units | Worldwide Distribution - USA Nationwide in the states of: Argentina, Austria, Belgium, Bulgaria, Canada, Chile, China, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, France, French Polynesia, Germany, Guadeloupe, Guatemala, Great Britain, Greece, Korea, Indonesia, Israel, Italy, Kazakhstan, Latvia, Liechstenstein, Lithuania, Luxembourg, Macedonia, Martinique, Mexico, Netherlands, Norway, Peru, Poland, Portugal, Russian Federation, Serbia, Singapore, Slovenia, Slovakia, Spain, Sweden, Switzerland, Thailand, Taiwan, Tunisia, Turkey, United Arab Emirates, Uruguay, Vatican City, and Vietnam. | Open, Classified |  |  | 4243 | Not_Computer
Z-0362-2014 | 66609 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Fixed Offset Adapter, A-642100. |  |  | October 24, 2013 | 2013 | Ossur North America Inc | Ossur inititated a voluntary recall of the A-642100 Fixed Offset Adapter because it can develop fatigue cracks in normal use over an extended period of time and ultimately separate completely. | Device Design | Ossur sent an "Safety Notice Voluntary Product Recall" letter dated October 29, 2013, via regular mail. Follow-up telephones are also made. The letter identified the product the problem and the actions to be taken by the customer.Âżssur is committed to providing superior customer service and we will do our utmost to minimize any inconvenience this may cause your practice. To ensure patient safety we kindly ask you to take the following action:" Please identify the users of each of the A-642100 adapters, see attached list." Also, please check your inventory for any A-642100s you may have on hand but have not fitted on users. A follow-up phone call will be made a few days later to coordinate replacements. A new replacement component will be provided to the customer. For further questions please call (949) 360-3634. |  | 66 | 66 units | US Distribution including the states of AL, TX, NY, CA, VA, GA, WI, NC, ID and NJ. | Terminated | June 17, 2014 | 236 | 66 | Not_Computer
Z-0358-2014 | 66391 | No | LRJ | General purpose disinfectants. | General Hospital | General Hospital | 510(K) Exempt | Ready to Use Thymo-Cide, Germicidal Detergent, containing 0.051% thymol, packaged in 1-qt. bottles. The firm name on the label is Wexford Labs, Inc., Kirkwood, MO. |  |  | September 10, 2013 | 2013 | Wexford Labs Inc | Expired active ingredient was used in the manufacture of the product. | Nonconforming Material/Component | Wexford Labs, Inc. sent undated letters by e-mail to affected customers beginning September 10, 2013, and follow-up phone calls were initiated requesting return of the affected product. E-mail letters dated September 26, 2013, were issued requesting return of product and asking the customers to notify their downstream customers. Customers with questions were instructed to call 800-506-1146.For questions regarding this recall call 314-966-4134, |  | 7020 | 7,020/1-qt. bottles (1755 gallons) | Nationwide Distribution - USA including FL, SC, TX, MO, CA, OH, OK, MA, IN, and GA. | Terminated | April 23, 2014 | 225 | 7020 | Not_Computer
Z-0361-2014 | 66534 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Aesculap Arcadius XP L Trial Insertion Instrument (ME020R/ME020R-US) |  |  | September 20, 2013 | 2013 | Aesculap, Inc. | The trial insertion instrument may disassemble during use due to a potentially weak weld on then trial inserter shaft block area located near the thumb wheel. | Device Design | An Important Product Removal Notification dated September 20, 2013 was sent to customers. |  | 30 | 30 | Worldwide distribution: USA (nationwide) GA, TN, CA, WY, NV, TX, OK, OH, and IL; and internationally to Canada. | Terminated | December 16, 2014 | 452 | 30 | Not_Computer
Z-0360-2014 | 66720 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Salter Labs EtCO2 Sample line 15' (4.6 M) tube, male luer connectors. |  |  | November 01, 2013 | 2013 | SALTER LABS | It has been determined that three lots of model number 4515-15-25 CO2 Sample Line 15' tube, male Luer lock connectors may contain a 30' tube instead of the 15' length. | Labeling Change Control | Salter Labs called the five customers that received product from the three lots, then follow-up immediately by sending a recall notification letter directly to each customer. Each letter will identify all shipments of subject product that they received.All customers letters were sent certified mail, return receipt requested. |  | 91 | 91 cases (25 units per case) | Distributed in the states of MS, TX FL,, WI, MI, VA, NY, MO, NC, and MS. | Terminated | December 17, 2013 | 46 | 91 | Not_Computer
Z-0364-2014 | 66639 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Juggerknotless Drill Bit, REF 110003173, |  |  | October 10, 2013 | 2013 | Biomet, Inc. | Investigation identified that the design of the drill bit may make the suture anchor difficult to insert into the bone. The implant may skip across the hole or miss the hole during insertion. | Device Design | Biomet sent an Urgent Medical Device Recall Notice letter dated October 10, 2013 to all consignees. The letter identified the affected product, reason for recall, and actions to be taken. The letter stated customers will be contacted on an individual basis in an attempt to reconcile the affected product. Customers were directed to locate and remove implicated devices and fax a copy of the response form to 574-372-1683 prior to returning product. For questions call 574-372-1570. |  | 25 | 25 | USA Nationwide Distribution in the states of VA, CA, Ky, and IN | Terminated | July 16, 2014 | 279 | 25 | Not_Computer
Z-0363-2014 | 66137 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Model FL14E Rose |  |  | October 23, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | It was reported that the siderail may appear latched when it is not, due to lack of proper preventive maintenance. This might result in bruises and/or lacerations to a patient who is leaning against the siderails if the siderail releases. Additionally, the siderail may not effectively constrain a patient to prevent them from rolling or falling off the bed. | Equipment maintenance | Stryker Medical sent a Medical Device Notification letter dated October 23, 2013, to all affected consignees. The letter described the recalled product, potential injuries associated with the side rail sticking, as well as providing preventive measures to be taken to mitigate the risk. Consignees were directed to locate the beds listed in this notice, ensure that the siderails have been inspected and adjusted, as required, continue to routinely inspect the siderails and functionality using the enclosed, enhanced product maintenance information. Problems and concerns should be addressed to Stryker at (1-800-327-0770, option 2). Consignees were asked to Fax (269-488-8691) or e-mail (ProductFieldAction@stryker.com ) the enclosed return form and directed to forward the notice if product was further distributed and advise Stryker of the beds new location.For questions regarding this recall call 269-324-6609. |  | 24575 | 4126 units USA; 20449 inits Internationally | Worldwide Distribution - USA (nationwide) and Internationally to Canada; Japan, South America, Latin America and China. | Terminated | May 16, 2014 | 205 | 24575 | Not_Computer
Z-0359-2014 | 66300 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | BIOMET Spine, 5.5 TI 5.5x50mm Translation Screw Polaris 5.5 TI Spinal System. Rx Only. Non-Sterile. |  |  | August 22, 2013 | 2013 | Ebi, Llc | Polaris Translation Screw Failures were recently reported to Biomet Spine. These failures are generally described as dissociation of the screw head from the shaft of the Polaris Translation Screw. | Device Design | Urgent Product Information notices dated 8/22/2013 were mailed via FedEx to distributors on 8/23/2013. Follow-up emails with a copy of the UPI notice were sent to distributors on 8/23/2013. The UPI notice informed the customers of the issue with the product and how to identify the affected product. Customers were advised that the Polaris Translation Screw System and the SpF Implantable Spine Fusion Stimulator not be used in conjunction with any implants containing cobalt chrome. Firm contacts were provided in case customers had any questions. EBI, LLC had a conference call with their domestic distributors on 8/23/2013 informing them they are required to notify their customers. |  | 7092 | 7,092 units | Nationwide distribution including Puerto Rico. | Terminated | August 31, 2016 | 1105 | 7092 | Not_Computer
Z-0352-2014 | 66721 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | CAS I/II Absorbers that are parts of the BleaseSirius Anesthesia Workstation, the BleaseFocus Anesthesia Workstation with the following part numbers: 12200900; 12200901; 12200902; 12200903; 14200100, and service kits part number 050-0659-00 and 050-0901-00. |  |  | October 17, 2013 | 2013 | Del Mar Reynolds Medical, Ltd. | Spacelabs Healthcare Ltd. of United Kingdom recalls , and Service Kits, Part Number 050-0659-00 and 050-0901-00. The Bag to-Vent switch in CAS I/II Absorbers may fail due to loose fastening hardware in an absorber. | Process control | Spacelabs Healthcare sent an Urgent Medical Device Correction letter dated November 1, 2013, to all affected customers. Spacelabs Healthcare also sent a customer letter (translated as necessary) to all international subsidiaries and distributors of record via e-mail on November 6, 2013. The letter informed customers of a potential defect in CAS I/II Absorber products employed in the Spacelabs Healthcare BleaseSirius Anesthesia Workstation, and Service Kits, Part Number 050-0659-00 and 050-0901-00.The Bag-to-Vent switch in CAS I/II Absorbers may fail. In the worst case, this would result in a large leak from the absorber circuit and could impede the ability to provide ventilation.Customers were instructed to immediately advise their staff of the problem. Customers with questions were instructed to contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support.For questions regarding this recall call 800-522-7025. |  | 37 | 37 units (19 in the US and 18 International) | Worldwide Distribution - USA (nationwide) including Oklahoma and North Carolina and Internationally to Libyan Arab Jamahiriya, Finland, Jordan, Mexico, Papua New Guinea, and Taiwan. | Terminated | April 24, 2014 | 189 | 37 | Not_Computer
Z-0368-2014 | 66662 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ 7 Ultrasound System, Model 795200, 795201, 795202. |  |  | October 17, 2013 | 2013 | Philips Medical Systems North America Inc. | A problem has been detected with Philips Model EPIQ 7 Ultrasound System, software version 1.0 and 1.0.1. Using the International System Date Format may result in calculation errors in Gestational Age Gestational Age (GA) and Established Due Date (EDD). | Software design | Affected customers will be notified via the Urgent-Field Safety Notice (FSN79500286), Philips Model EPIQ 7 Ultrasound system, dated 2013 October. Customers are informed that using the International System Date Format in the Philips Model EPIQ 7 Ultrasound System, software version 1.0 and 1.0.1. may result in calculation errors in Gestational Age and Established Due Date (EDD). Philips will upgrade customer systems to software version 1.1, which corrects the issue, for a free of charge. |  | 306 | 306 units total (206 international; 100 in the US). | Worldwide Distribution-USA (nationwide) including the states of CT, HI, IL, IN, KY, MA, ME, MI, NC, OR, PA, TX, and WA, and the countries of Australia, Austria, Belgium, Chile, Denmark, Finland, France, Germany, India, Ireland, Republic of Korea, Latvia, Netherlands, New Caledonia, New Zealand, Norway, Poland, Portugal, Romania, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | July 16, 2014 | 272 | 306 | Computer
Z-0366-2014 | 66640 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Green Rusch - Lite Disposable Metal Laryngoscope Ridge Blade to visualize upper airway and aid in placement of tracheal tube. |  |  | October 15, 2013 | 2013 | Teleflex Medical | Labeling inconsistency | Other | Recall letters were sent to customers and distributors on 10/15/2013 via courier. The letter indicated that the consignee was to immediately discontinue use and quarantine any products with the affected catalog numbers. The letter also included a Recall Acknowledgment Form which was to be faxed back to the firm. |  | 6200 | 6,200 ea. | Nationwide distribution: AL, AZ, CA, CO, FL, GA, IL, IN, IA, KY, LA, MD, MA, MI, MN, MO, NE, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA, WA, WV and WI. | Terminated | November 04, 2015 | 750 | 6200 | Not_Computer
Z-0367-2014 | 66631 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Discovery NMCT 670, Discovery NM 630, Optima NMCT 640, &amp; Brivo NM 615. |  |  | October 25, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue related to a SPECT scan procedure used for pulmonary imaging.Contact between the Detector (GE Healthcare Discovery NM/CT 670, and/or Optima NM/CT 640, and/or DiscoveryNM630, and/or Brivo NM 615 all configurations) and the Patient's elbow may occur during SPECT protocols using 450 angular steps in scenarios when a Patients a | Employee error | GE Healthcare sent an "Urgent Medical Device Correction" letter dated September 25, 2013, to all affected customers. The letter identified the product the problem and the actions to be taken by the customer.We plan to correct all of the affected Discovery NMCT 670, Discovery NM 630, Optima NMCT 640, &amp; Brivo NM 615 systems that were distributedThe plan for public warning is to contact all customers with an affected Discovery NMCT 670, Discovery NM 630, Optima NMCT 640, &amp; Brivo NM 615 system by means of an Urgent Medical Device Correction letter. In the US this will be sent by a traceable means (either US Postal Service with receipt acknowledgement (registered mail) or equivalent courier service such as FEDEX). This letter will be sent to 3 titles within the affected accounts; Managers of Nuclear Medicine, Hospital Administrators / Risk Managers, Managers of Radiology/Cardiology. A GE Healthcare Field Engineer will perform the required software update on each affected system.Contact Information (GE Healthcare Call Center: USA or Canada 800-437-1171 and Other countries:( 262-896-2890). |  | 464 | 464 (143 US, 321 OUS) | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MI, ME, MD, MA, MN, MS,MO,MT, NV, NH, NJ, NY, NC, OH, OK, PA, PR, SC, SD, TN, TX, UT, VA, WV and WI., and the countries of ALGERIA, AUSTRALIA, BELGIUM, BRAZIL, BULGARIA, CANADA, CHINA, COLOMBIA, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, INDIA, IRELAND, ISRAEL, ITALY,. JAPAN, REPUBLIC OF KOREA, KUWAIT, MEXICO, NETHERLANDS, NEW ZEAL NAD, NORWAY, PAKISTAN, POLAND,QUATAR, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN,K THAILAND, TURKEY and UNITED KINGDOM. | Terminated | April 01, 2014 | 158 | 464 | Computer
Z-0369-2014 | 66552 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Titanium Trochanteric Fixation Nail (TFN)-Helical Blade |  |  | October 03, 2013 | 2013 | Synthes USA HQ, Inc. | A potential exists for non conforming Synthes Titanium Trochanteric Fixation Nail (TFN)-Helical Blades to be distributed. | Process control | Synthes sent a Field Safety Notice dated October 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the options provided in the Field Safety Notice to clarify and mitigate the potential risk. Customers were asked to prepare a list of the affected part and lot numbers and contact the Complaint Unit at 610-719-5596 to receive a Return Authorization Number. Complete and return the Verification Form with the product and Return Authorization number. Customers were also asked to forward the Field Safety Notice to anyone in their facility that needs to be informed. Customers were asked to maintain a copy of the Field Safety Notice.Customers with questions were asked to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 350195 | 350195 | Worldwide Distribution - USA (nationwide) and Internationally to Mexico and Canada. | Open, Classified |  |  | 350195 | Not_Computer
Z-0372-2014 | 66263 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | *** 1) SYMBIQ One-Channel Infuser; |  |  | August 28, 2013 | 2013 | Hospira Inc. | Hospira shipped the incorrect replaceable battery pack (Part No: HSP3566-E10) instead of the correct E09 battery pack (Part No. HSP3566-E09) for the Symbiq Single and Dual Channel Infusers. | Employee error | The firm, Hospira, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated August 28, 2013 to its customers. Hospira contracted with Stericycle, Inc., in Indianapolis, IN for customer notification, response, tracking, returns, effectiveness checks and destruction of returned product starting on August 28, 2013. The letter described the product, problem and actions to be taken. Customers were instructed to: 1) complete and return the attached reply form via fax to 1-866-470-3291 or e-mail to Hospira2921@stericycle.com, even if you do not currently have the affected batteries; 2) contact Stericycle at 1-877-779-7810 (M-F, 8am - 5pm ET) if they require additional response forms; and, 3) return any affected batteries upon receipt of the new (correct) replacement batteries to Stericycle. Customers are also instructed to forward the recall information to any accounts that they may have shipped the affected batteries to and have them contact Stericycle to recieve a reply form and return label.For further inquiries, please contact Hospira Advanced Knowledge Center at 1-800-241-4002, option 6, availabel 24 hours a day/7 days per week. |  | 59 | 59 batteries | US Distribution in states of: : KY, MN, ND, NJ, NY, PA, and SD. | Open, Classified |  |  | 59 | Not_Computer
Z-0370-2014 | 66056 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Medcare Products Max-500 Stand, Total Support, patient stand, model #400006. The responsible firm name listed in the Operations Manual is Medcare Products, Burnsville, MN. |  |  | October 30, 2013 | 2013 | Ergosafe Products LLC DBA Prism Medical Services USA | The lift might fail with the labeled 500-lb. load rating. | Device Design | The firm, Medcare Products, sent an "URGENT FIELD CORRECTIVE ACTIONS" letter dated October 21, 2013 via certified mail on October 30, 2013 to its customers. The letter described the product, problem and action to be taken. The customers were instructed to peel off the boom stickers and serial number label and replace with the new ones supplied; review the revised Manual and dispose of the original; and complete the action report form once the products have been relabeled. Please refer to your Operations Manual or visit our website at www.medcarelifts.com to view the manual online.If we can answer any questions that you may have regarding the testing standards or the relabeling of the products, please do not hesitate to contact us at 1-800-695-4479. |  | 97 | 97 units | Nationwide distribution. | Terminated | May 17, 2017 | 1295 | 330 | Not_Computer
Z-0373-2014 | 66459 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | GE Vivid E9 ultrasound system. GE Vivid E9 is a Track 3 diagnostic ultrasound system, which is primarily intended for cardiac imaging and analysis, but which also includes vascular and general radiology applications. The Vivid E9 incorporates a variety of electronic array transducers operating in linear, curved linear, sector/phased array or matrix array format, including two dedicated CW transducers and several real time 3D transducers. It consists of a mobile console with keyboard control panel; color LCD/TFT touch panel, LCD color video display and optional image storage and printing devices. It provides high performance ultrasound imaging and analysis and has comprehensive networking and DICOM capability. |  |  | October 11, 2013 | 2013 | GE Healthcare, LLC | GE became aware of a potential safety issue due to a system slow down and system lock up associated with the software of Vivid E9 ultrasound system. During a critical procedure the Vivid E9 Scanner may gradually become unresponsive and eventually lock up, with the result that the vivid E9 Scanner must be rebooted. | Component change control | An Urgent Medical Device Correction letter, dated 10/11/2013, was sent to 3 departments within the affected hospitals. The letter described the safety issue, and identified affected devices. The letter stated to reboot the scanner in an unresponsive /lock up condition. Also if the system is being used for an interventional procedure, the TEE probe should remain in the patient while system is rebooting. A GE healthcare service engineer will install a correction to affected devices. |  | 710 | 710 | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Argentina, Australia, Austria, Belarus, Belgium, Canada, Chile, Colombia, Croatia, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Indonesia, Italy, Japan, Korea (Republic of), Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Malaysia, Netherlands, Norway, Oman, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, and United Kingdom. | Terminated | January 23, 2015 | 469 | 710 | Computer
Z-0349-2014 | 64832 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Pump is a small and lightweight, single channel infusion device designed for use in the home, in the hospital, or anywhere electronic infusion is required. The GemStar Pump can be powered by AC mains adaptor, rechargeable battery pack, docking station, or two disposable AA alkaline batteries. When powered by batteries, The GemStar Pump is ideal for ambulatory patients. |  |  | March 15, 2013 | 2013 | Hospira Inc. | The proximal and distal pressure sensor calibration can drift resulting in the pump failing the Proximal or Distal Occlusion Operational Test, as described in the GemStar Technical Service Manual, or reporting one of the following errors during device setup or infusion: 1) Cassette Check - D; 2) Cassette Check - P; 3) Proximal Occlusion; 4) Distal Occlusion; 5) Pressure Calibration Error; 6) Bad Pressure Sensor Event; 7) Bad Pressure State; 8) Distal Pressure is Out of Range; 9) Proximal Sensor is Out of Range. A pump with this issue may, instead of reporting an error, not detected occlusions or issue false occlusion alarms, which will stop the infusion and invoke visual and audible warnings to the user. | Device Design | Hospira sent an Urgent Device Field Correction letter dated March 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter asks that customers immediately perform proximal and distal occlusion tests as defined i nthe GemStar Technical Service Manual. If the device fails either of the tests, remove it from clinical service and contact Hospira at 1-800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM) to report the issue and arrange for the return of the device for recalibration. The letter further reminds customers that the proximal and distal occlusion tests should be performed on an annual basis as a part of the yearly GemStar maintenance schedule. Customers were also requested to complete and return the attached reply form and return it to Hospira by fax to 1-888-714-5077 or by e-mail to Hospira7906@stericycle.com.Customers were directed to notify their consignees if they further distributed the product and confirm they have done so by return of the reply form to Stericycle. For questions regarding this recall call 800-441-4100. |  | 4132 | 1) List Number 13000: 2,162 pumps; 2) List Number 13100: 15 pumps; 3) List Number 13150: 1,955 pumps | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY. Internationally to Australia, New Zealand, China, Hong Kong, Taiwan, Japan, Korea, Philippines, Thailand, Malaysia, Canada, Austria, Bahrain, Belgium, Croatia, Denmark, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Iceland, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malta, Netherlands, Norway, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, UK, United Arab Emirates, Brazil, Chile and Colombia. | Terminated | March 07, 2017 | 1453 | 47972 | Not_Computer
Z-0376-2014 | 66359 | No | MZW | Dental cement. | Dental | Dental | 510(k) | ExciTE F DSC Soft touch Single Dose, Dual-curing Regular Ref. 50 |  |  | September 23, 2013 | 2013 | Ivoclar Vivadent, Inc. | Ivoclar Vivadent, Inc. is recalling a lot of ExciTE F DSC Dental Adhesive found to produce a blue green discoloration when used in combination with cements. | Material/Component Contamination | The firm, Ivoclar Vivadent, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated September 19, 2013 via first class mail to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to check your stock and segregate any material marked batch R59595; contact the Ivoclar Vivadent Customer Service at 800-533-6825 as soon as possible between the hours of 9:00am and 7:00pm EST Monday through Friday and a representative will arrange to have your stock replaced; both unopened and opened packages can be returned; and complete and return the Medical Device Recall Return Response Acknowledgment and Receipt Form.Please communicate locally with Ivoclar Vivadent Customer Service at 800-533-6825 for any questions. |  | 63 | 63 refill packs (50 single use applications per refill) | US Distribution: including states of PA, WA, WI, NY, LA, TX, NJ, CA, OR, IA, SC, IN, IL, and AZ. | Open, Classified |  |  | 63 | Not_Computer
Z-0375-2014 | 66657 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker |  |  | September 24, 2013 | 2013 | Stryker Spine | Multiple complaints have reported that the hex tip of the torque wrench fracturing during final tightening. | Process design | Stryker sent an Urgent Medical Device Removal Notification letter dated September 25, 2013, with product accountability forms to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory and hospital locations to identify any affected product, reconcile any Torque Wrenches with Catalog Number with Lot Numbers referenced in the letter. Customers were asked to fax a copy of the Customer Response Form to Regulatory Compliance at 855-632-9049 or you email it to Spine-RegulatoryActions@Stryker.com. Customers were asked to return the affected product using the enclosed mailing label to: RECALL RA 2013-105Stryker Spine59 Route 17 South,Allendale, New Jersey 07401branches/agencies, the field and surgeons via Fed Ex on starting on 9/24/2013.For questions regarding this recall call 201-760-8298.Stryker sent another Urgent Medical Device Removal Notification Letter dated 1/29/2014 which included additional lots and one product. |  | 651 | 220 US units; 431 OUS units | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Canada, Chile, France, Canada, Germany, Hong Kong, Ireland, Italy, Japan, Korea, Netherlands, New Zealand, Spain, Iraq, Egypt, and Turkey. | Open, Classified |  |  | 651 | Not_Computer
Z-0385-2014 | 66605 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Aviva Combo Blood Glucose Meter (as a part of the ACCU-CHEK Aviva Combo System)\ |  |  | October 25, 2013 | 2013 | Roche Diagnostics Operations, Inc. | There is a software synchronization issue with the ACCU-CHEK Combo system. In rare cases, when the Manual Pump option is chosen on the ACCU-CHEK Aviva Combo meter, there is a possibility of receiving an incorrect bolus advice recommendation that may cause a temporary under delivery of insulin. | Software design | The firm, Roche Diagnostics, sent an "URGENT MEDICAL DEVICE CORRECTION (UMDC)" letter dated October 29, 2013, via UPS to their consignees/customers. The letter described the product, problem and actions to be taken. The UMDC informed them of the software synchronization issue with the ACCU-CHEK Combo system that may cause a temporary under delivery of insulin. The customers were instructed to ensure they properly deliver the correction bolus when utilizing the Manual Pump option, as instructed in the Owners Booklet; and complete and return the Business Reply Card via fax to 1-877-282-0223. The Distributors were instructed to provide the UMDC to their direct customers. The HCPs were instructed to remind their patients to follow the instructions in the Accu-Chek Aviva Combo Meter Advanced Owners Booklet when utilizing Manual Pump option.Questions about the information contained in this UMDC are directed to ACCU-CHEK Pump Support, 24 hours a day, seven days a week at 1-800-688-4578. |  | 6595 | 6,595 units | Nationwide Distribution. | Terminated | January 14, 2015 | 446 | 6595 | Computer
Z-0379-2014 | 66663 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Product Brand Name(s): Alere Cholestech LDXÂż Lipid Profile"GLU Test Cassette, Model 10-991. |  |  | October 23, 2013 | 2013 | Alere San Diego, Inc. | Alere initiated a recall of Cholestech LDX Lipid Profile-GLU test cassettes, Catalog Number 10-991, Lot # 320167, becuase an incorrect lot number was printed on the outer 10 pack kit box label. The incorrect Lot # number printed was 321067. | Labeling Change Control | Alere initiated this recall by sending out notification letters to customers via fax, email, or direct mail. The letter titled "URGENT MEDICAL DEVICE NOTICE", dated October 23, 2013 is accompanied by a response form/Verification form. The notification informed customers of the reason for recall, product description with codes, required actions by the customers, and contact information. |  | 3621 | 3,621 kits | Worldwide Distribution: USA (nationwide) and Internationally to Canada. | Terminated | September 25, 2014 | 337 | 3621 | Not_Computer
Z-0377-2014 | 66561 | No | HCH | Aneurysm clip. | Neurology | Neurology | 510(k) | Mizuho Sugita Titanium Standard Aneurysm Clip II , |  |  | October 10, 2013 | 2013 | Mizuho America, Inc. | Specific item numbers form Lots 27-J and 29-J of Sugita T2 Aneurysm Clips. Mini Temporary Type, are discolored. | Process control | Mizuho sent a Product Recall Notice dated October 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to verify serial number before opening package and conduct a physical examination of the product for discolored Clips from the affected lots. Customers were also asked to fill out the Confirmation of User Receipt, and return to Misuho America. If customers should find affed clips they were instructed to quarantine them an call Mizuho America Customer Service at 800-699-2547 for return instructions. The Clips will be exchanged at no charge. Customers were also asked to ensure that all appropriate persons within the Hospital have been notified of the Recall Notice.For questions regarding this recall call 510-324-4500, ext 122. |  | 32 | 32 | Nationwide Distribution | Terminated | January 13, 2014 | 95 | 67 | Not_Computer
Z-0380-2014 | 66411 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | 2.4 and 2.7 MM Locking Screw SLF-TPNG with T8 Stardrive Recess 12 MM |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | Specific lots of the SS 2.4mm and SS 2.7mm Locking Screw (self-tapping) due to a mislabeling issue. The SS 2.7mm Screw was labeled as SS 2.4 mm Screw. | Labeling mix-ups | Synthes sent an URGENT NOTICE: MEDICAL DEVICE RECALL letter dated July 31, 2013, to all affected customer. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 if they have any of the affected product to obtain a Return Authorization Nuber, complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Fax the documents to Synthes at 1-866-792-6446 or email to FieldAction@synthes.com. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 98 | 98 | Nationwide Distribution including PA, CA, FL, MO, KS, IN, IL, WV, KY, LA, NM, SC, IA, TX, AZ, GA, NC, OK, NY, TN, MI, CO, and MT. | Open, Classified |  |  | 98 | Not_Computer
Z-0383-2014 | 66739 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MAS Urinalysis Control Level 1, Multi-analyte control (Assayed and Unassayed), Catalog/Part Numbers: 024222; Product is manufactured and distributed by Microgenics Corporation, a subsidiary of Thermo Fisher Scientific, Clinical Diagnostic Division, Fremont, CA |  |  | October 29, 2013 | 2013 | Thermofisher Scientive | Via one (1) consumer complaint and in-house testing has confirmed that the product does not meet the 30 day open vial stability specifications for Lot: UA1514011. | Under Investigation by firm | ThermoFisher Scientific, sent an Urgent Medical Device Field Correction letter dated October 29, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use, segregate and discard the affected product. Customers with questions were instructed to contact mas.controls@thermofisher.com. For question regarding this recall call 1-800-232-3342. |  | 914 | 914 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and internationally to Hong Kong. | Terminated | March 17, 2014 | 139 | 914 | Not_Computer
Z-0382-2014 | 66569 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Tri-Nex Healing Abutment, 0.50 mm x 5 mm height, Catalog No. HA-L-50-5. |  |  | September 13, 2013 | 2013 | Southern Implants, Inc | Southern Implants is recalling the Tri-Nex healing abutment because the package was incorrectly labeled and packaged as the Tri-Nex healing abutment, 05.0 mm x 5 mm height (HA-L-50-5). The package contains the Tri-Nex healing abutment 05.0 mm x 6.0 mm Flare x 5 mm Height (HA-L-50W-5). | Labeling design | Southern Implants, Inc. sent an Urgent Product Action letter dated September 11, 2013, to all affected customers. The letter identiied the product, the problem, and the action to be taken by the customer. Customers were instructed to fill out the customer response form and fax/email the completed form to Customer Support at (866) 903-9272 or customersupport@keystonedental.com.For questions regarding this recall call 866-903-9272. |  | 197 | 197 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Guatemala, Taiwan. | Terminated | January 21, 2014 | 130 | 197 | Not_Computer
Z-0384-2014 | 66415 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Matrix Threaded Persuader |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes became aware that during a procedure using the Matrix Threaded Persuader, the Reduction Insert malfunctioned causing pieces to break off. Once there is a breakage to the Reduction Insert it will not allow the instrument to function as required. | Labeling design | The firm, Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated July 31, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to review your inventory and immediately remove the affected lots from stock; If you do have the affected devices, contact Synthes at 1-800-479-6329 to obtain the RA# and return the product to with completed form to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132; complete and sign the Return Receipt Requested form via fax at 1-888-731-7955 or email to: FieldAction@synthes.com, even if you do not have the affected product.If you have questions, please call 610-719-5450. |  | 634 | 634 | Worldwide distribution: US (nationwide) and countries of Canada and Switzerland. | Terminated | August 31, 2015 | 761 | 634 | Not_Computer
Z-0381-2014 | 66792 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | REF 200068 Security (TM) Enclosed Carpal Tunnel System |  |  | October 25, 2013 | 2013 | Biomet, Inc. | Security Blades were manufactured oversized and potentially won't advance or may be difficult to advance through the mating Security Clip. | Nonconforming Material/Component | 10/25/2013 Urgent Medical Device Recall Notifications were sent to each Distributor and Risk/Recall Manager at each direct account identifying the recalled device, possible adverse health consequences and instructing consignees to immediately locate and remove devices from circulation, following instructions on the FAX back form and returning it within 3 business days to FAX 574-372-1683 and returning product to Biomet Return Goods, 56 East Bell Drive Warsaw In 46580. Questions should be directed to 574-372-1570 M-F, 8 a.m. to 5 p.m. |  | 1130 | 1130 | Worldwide Distribution: USA (nationwide) and Internationally to Japan. | Terminated | May 06, 2015 | 558 | 1130 | Not_Computer
Z-0388-2014 | 66420 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Tensioning Device |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | All lots of the Synthes Tensioning Device were recalled due to the potential for the reaction fork to come loose from the main body of the device and not allow the wire to release, nor the device to apply tension. | Device Design | Synthes sent an Urgent: Medical Device Recall letter dated July 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer."Please remove this product from your inventory and return it to Synthes. A replacement product will be sent to you when available. If you DO have any of the identified devices, please take the following steps:"Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided."Return the Verification Form (page 3 of this letter) with the product to:" Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indicating that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information."Return the document to Synthes by: Fax: 877-272-3517"Email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or email Fieldaction@synthes.comWe apologize for any inconvenience this may cause and thank you in advance for your attention and cooperation. |  | 369 | 369 | Nationwide Distribution. | Terminated | June 01, 2015 | 670 | 369 | Not_Computer
Z-0386-2014 | 66542 | No | LAS | Clinical toxicology control material. | Toxicology | Toxicology | 510(K) Exempt | Single Analyte Urine DAU Control (5 m L |  |  | October 03, 2013 | 2013 | Lin-Zhi International Inc | Product catalog # A68825 Norbuprenorphine 13 ng/mL Level 2 Controls gives a decreased concentration reading due to the degradation of the NBUP analyte in solution. | Under Investigation by firm | Beckman Coulter sent a Manufacturer Field Action Letter dated November 2013, to the one distributor. The letter identified the product the problem and the action needed to be taken by the customer.Please share this information with your laboratory staff and retain this notification as part of your laboratory Quality System documentation. If you have forwarded any of the affected product(s) listed above to another laboratory, please provide them a copy of this letter.If you have any questions regarding this notice, please contact our Customer Support CenterÂżÂż Via our website, http://www.beckmancoulter.com/customersupport/supportÂżÂż Via phone, call 1-800-854-3633 in the United StatesÂżÂż Outside the United States, contact your local Beckman Coulter Representative.We apologize for any inconvenience. |  | 60 | 60 items | US Distribution in CA. | Terminated | December 26, 2013 | 84 | 60 | Not_Computer
Z-0387-2014 | 66421 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Large Buttress Compression Nut For Blade Guide Sleeve TFN |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | The Large Buttress Compression Nut For Blade Guide Sleeve TFN was recalled because of complaints that the compression nut did not thread completely onto the blade guide sleeve and became stuck. | Process control | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 31, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action to be taken: Our records indicate you received this product. Synthesis requesting that the product be returned:If you DO have any of the identified devices, please take the following steps:"Call Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided."Return the Verification Form {page 3 of this letter) with the product to:"Credit/Returns, Synthes, .1101 Synthes Avenue, Monument, CO 80132.If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information."Return the documents to Synthes by." Fax: 1-888-731-7952 or" Scan/email: FieldAction@synthes.comIf you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. Thank you for your attention to this issue |  | 135 | 135 | US Distribution including the states of TX, PA, NY, GA, CO, IN, AI, AR, IL, MI, TN, CA, AZ, MS and VA. | Terminated | June 01, 2015 | 670 | 135 | Not_Computer
Z-0389-2014 | 66755 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Multi Parameter Patient Monitor (with Arrhythmia Detection and Alarms) |  |  | October 07, 2013 | 2013 | Mindray DS USA, Inc. dba Mindray North America | Mindray has identified an issue with V Series Monitor where the V Dock power pins connecting the V Series Monitor to the V Dock may lose spring tension, potentially causing the screen to temporarily go blank, lose communication with the central station monitor (when one is in use) or shut down. No patient related events have been reported to Mindray concerning the issue. | Environmental control | Mindray sent an Urgent Corrective Action letter dated October 7, 2013 to all of their customers. The letter identified the affected product, problem and actions to be taken. Customers should contact their Mindray Service Representative at 1-800-288-2121 Monday through Friday from 830 am to 530 pm to arrange for the replacement of the affected product. |  | 1839 | 1839 units | Worldwide Distribution - USA Nationwide | Terminated | April 27, 2015 | 567 | 1839 | Not_Computer
Z-0397-2014 | 66882 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | CSF-Unitized Shunt, Contoured Regular, Low Pressure, Catalog No. 46022; |  |  | November 18, 2013 | 2013 | Medtronic Neurosurgery | Medtronic Neurosurgery is recalling the CSF-Unitized Shunts, Delta Shunt Assembly Kit, and the Strata II Shunt Assembly Kit because they were found to have label discrepancies. | Other | Medtronic Neurosurgery sent an Urgent Medical Device Recall letter dated November 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to provide a copy of the letter to all those who need to be aware of the matter within their organization or to any organization where the potentially affected product may have been transferred. Customers with questions were instructed to call 805-571-8725. |  | 288 | 288 units total (147 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to US and worldwide: Austria, Belgium, Bosnia, Herzegovina, Canada, Croatia, Germany, Japan, Poland, Portugal, Taiwan, United Arab Emirates, United Kingdom. | Terminated | February 06, 2014 | 80 | 327 | Not_Computer
Z-0393-2014 | 66551 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Material Mix (Adjustable Cervical Distractor-Right, Matrix Distractor Rack, T-PAL Spacer Remover, Implant Inserter) |  |  | October 15, 2013 | 2013 | Synthes USA HQ, Inc. | One material batch used for production of the affected parts of the Synthes Material Mix were mixed in with an incorrect material with a lower than specified hardness and corrosion resistance. | Nonconforming Material/Component | SYNTHES sent an Urgent Notice: Medical Device Recall letter dated October 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number. Complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product was located. Customers were also asked to indicate the number of devices found and note the Return Authorization Number. Return the Verification Form with the product to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers should also return the Verification Form if no affected product was found in their inventory. Return all documents to Synthes by fax: 888-679-4875 or Scan/email: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 44 | approximately 44 | Worldwided Distribution: USA (nationwide) including IL, MO, TX, WA and Internationally to Canada and Switzerland. | Terminated | August 31, 2015 | 685 | 44 | Not_Computer
Z-0374-2014 | 66231 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | (Siemens Rad Fluoro Uro Systems) Ysio, Luminos Agile, Luminos dRF, and Uroskop Omnia systems with software version VC10 |  |  | September 03, 2013 | 2013 | Siemens Medical Solutions USA, Inc | A potential malfunction and hazard to patients exists when using the Ysio, Luminos Agile, Luminos dRF, and Uroskop Omnia systems with software version VC10, which may result in radiation continuing longer than required and overexposed images that are not of diagnostic quality. | Software design | Siemens Healthcare sent a Field Safety Notice dated September 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to forward the Field Safety Notice to any new users and provide the new owner information to Siemens Healthcare. Customers were asked to send back the attached form confirming receipt of the Field Safety Notice. Customers with questions should call 610-850-4634.For questions regarding this recall call 610-219-6300. |  | 357 | 357 | Worldwide Distribution - USA (nationwide) and internationally to Canada. | Terminated | February 24, 2015 | 539 | 357 | Computer
Z-0404-2014 | 66581 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 14.5 mm Outer Protection Sleeve for Suprapatellar |  |  | October 14, 2013 | 2013 | Synthes (USA) Products LLC | Complaints were received reporting that a product packaged as a 14.5 mm Outer Protection Sleeve for Suprapatellar contained a 12.0 mm Outer Protection Sleeve for Suprapatellar. | Labeling mix-ups | Synthes sent an Urgent Notice: Medical Device Recall letter dated October 14, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 if they have any of the affected product to obtain a Return Authorization Number, complete the Verification Section at the end of the letter and return the Verification Form with the affected product to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132If customers do not have the affected product they should complete the attached Verification Section and return all documents to Synthes by:Fax: 866-229-7778Scan/email: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 55 | 55 | Nationwide Distribution | Terminated | August 18, 2015 | 673 | 55 | Not_Computer
Z-0401-2014 | 66723 | No | JDP | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 4.5MM TI VA-LCP CRVD Condylar Plate/6H/159MM/RT-STER |  |  | May 29, 2013 | 2013 | Synthes USA HQ, Inc. | A plate was inadvertently released to a sales consultant that was restricted for sale. | Employee error | Synthes sent a letter dated November 7, 2013, to the implanting surgeon to follow up on a complaint received. The firm requested any additional information the surgeon might have about the use of the plate as well as a status on the patient following use. For questions the surgeon was instructed to call 800-620-7025.For questions regarding this recall call 610-719-5000. |  | 1 | 1 | US Distribution in Colorado. | Terminated | May 21, 2015 | 722 | 1 | Not_Computer
Z-0394-2014 | 66580 | No | HTN | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Spiked Washers 13.5/6.5 |  |  | October 03, 2013 | 2013 | Synthes USA HQ, Inc. | During a retrospective assessment, Spiked Washers 13.5/6.5 were discovered to be mislabeled as Spiked Washers 13.5/5.5. | Finished device change control | SYNTHES sent an Urgent Notice: Medical Device Recall letter dated October 3, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product located. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes Avenue, Monument, CO 80132Customers were also asked to return the Verification Form indicating they did not have the affected product also. All documents were to be returned to Synthes by:Fax 1-877-270-5855Scan/email: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450 or contact their Synthes Sales ConsultantFor questions regarding this recall call 610-719-5000.. |  | 3087 | 3087 | Nationwide Distribution | Terminated | December 18, 2015 | 806 | 3087 | Not_Computer
Z-0400-2014 | 66836 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | LANX Timberline MPF Lateral Modular Plate Fixation System Straight Split Tip Driver - Short ("Driver"). Model Number 8630-0107. |  |  | November 04, 2013 | 2013 | Lanx, Inc. | Lanx, Inc. is recalling various lots of Timberline MPF Straight Split Tip Driver - Short ("Driver") because the width of the most distal end from the Driver handle (the tip of the hex portion of the Driver) is oversized. | Process control | LANX sent an URGENT: Recall letter dated November 5, 2013, to all affected customers via Federal Express, Overnight delivery. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all affected product to the firm. Customers with questions were instructed to contact Field Inventory Management.For questions regarding this recall call 303-501-8402. |  | 25 | 25 units | Nationwide Distribution including AZ, PA, WI, NY, TX, PR, and TN. | Terminated | December 29, 2015 | 785 | 25 | Not_Computer
Z-0390-2014 | 66666 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.7mm and 5.0mm Dynamic Locking Screw (DLS) |  |  | July 22, 2013 | 2013 | Synthes USA HQ, Inc. | All lots of 3.7mm and 5.0mm Dynamic Locking Screw (DLS) were recalled due to complaints of breakages at the distal tip of the pin identified after successful healing and during planned removal of the device. | Labeling False and Misleading | Synthes sent an Urgent Notice: Medical Device Recall letter to Sales Reps and End Users on July 22, 2013. The letter identified the affect product, problem and actions to be taken. For questions call the Recall center at 610-719-5450. |  | 40248 | 40248 | Worldwide Distribution - USA (nationwide) and internationally to Canada. | Open, Classified |  |  | 40248 | Not_Computer
Z-0391-2014 | 66416 | No | HXP | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Matrix 5.5.mm Left and Right Coronal Bender |  |  | June 22, 2013 | 2013 | Synthes USA HQ, Inc. | A complaint was received of the tips of the Matrix 5.5.mm Left and Right Coronal Bender breaking while bending a spinal rod component of the Matrix System. | Component design/selection | Synthes sent an Urgent Notice: Medical Device Recall letter dated June 22, 2013 to users and sales reps. The letter identified the affected product, problem and actions to be taken. For questions call 610-719-5450. |  | 209 | 209 | Worldwide Distribution - USA (nationwide) including GA, FL, CA, UT, IA, AZ, OH and Internationally to Canada. | Terminated | September 10, 2015 | 810 | 209 | Not_Computer
Z-0403-2014 | 66815 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD NX, Product Code: 10033 |  |  | November 11, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD NX and STERRAD 100NX sterilization systems because ASP has determined that the bronchoscopes may contain a polyurethane material in the suction channel location that is not listed in the current claims for interior channels (lumens) for processing in STERRAD systems. | Labeling design | Advanced Sterilization Products sent an Urgent Medical Device Field Safety Notificaiton letter on November 11, 2013, to all affected customers. The letter informed customers of the problems identified and the actions to be taken. Customers were instructed to provide the notice to anyone in their facility that needs to be informed and post this notice where appropriate. Customers were instructed to contact ASP at 1-888-783-7723 for any suspected problems. ASP is collaborating with KARL STORZ Endoskope to determine the sterilization impact of the suction port lumen material. Customers are instructed to contact KARL STORZ Endoskope Technical Support at (800)421-0837, ext 5350 for alternative reprocessing instructions.For questions regarding this recall call 949-453-6400. |  | 4381 | Total 4,381 units | Nationwide Distribution | Terminated | January 22, 2014 | 72 | 4381 | Not_Computer
Z-0396-2014 | 63230 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Synthes Titanium Minimally Invasive Reduction Screw System Locking Cap, Part Number 04.631.502 |  |  | September 07, 2012 | 2012 | Synthes USA HQ, Inc. | This action is being initiated following a detailed investigation in response to reported complaints of post-operative locking cap loosening. Synthes has concluded that if the construct is mis-aligned during final tightening, the potential exists for locking caps to loosen post-operatively in certain situations. | Use error | Synthes sent an Urgent: Medical Device Product Removal letter on September 7, 2012, to all affected consignees. Consignees were requested to immediately cease using the product, examine their inventory and remove the parts from inventory. If they had the parts, they were further instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the verification section at the end of the letter and return the verification form. If they did not have the parts, they were requested to complete the verification section at the end of the letter and return the verification form. Consignees with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 745 | 745 | Nationwide Distribution including CA, CO, CT, FL, GA, ID, IL, IN, LA, MA, MD, MI, MN, MT, NC, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA, WA, and WI. | Terminated | December 18, 2015 | 1197 | 1502 | Not_Computer
Z-0392-2014 | 66389 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Baxter, VIAL-MATE Adaptor, 2B8071, Rx Only, STERILE R, Baxter Healthcare Corporation, Deerfield, IL 30015 USA |  |  | September 25, 2013 | 2013 | Baxter Healthcare Corp. | Incomplete foil seal on one lot of sterile product. | Packaging process control | Baxter Healthcare sent an Urgent Product Recall letter dated September 25, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to locate and remove all affected product from their facility, contact Baxter's Center for Service at 1-888-229-0001 to arrange for return and credit. Complete the attached Customer Reply Form and return it to Baxter by fax or email. Customers with questions were instructed to call 1-800-422-9837. For questions regarding this recall call 224-948-4770. |  | 38400 | 38,400 units | Nationwide Distribution including WV, KY, IA, AR, OH, IL, IN, MI, KS, TN, WI, MN, MO. | Terminated | July 22, 2014 | 300 | 38400 | Not_Computer
Z-0405-2014 | 66577 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Implant Holder for Synfix (TM)-LR |  |  | October 08, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received which described the SynFix LR implant holder breaking at the interface between the implant and the holder. | Unknown/Undetermined by firm | Synthes sent an Urgent Notice: Medical Device Notification letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to returny any of the affect devices, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section indicating the affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers were also instructed to return the Verification Section even if they don't have the affected product. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 548 | 548 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Open, Classified |  |  | 548 | Not_Computer
Z-0428-2014 | 66509 | No | ELZ | Preformed crown. | Dental | Dental | 510(K) Exempt | 3M ESPE, Unitek Stainless Steel Permanent Molar Crowns |  |  | June 05, 2013 | 2013 | 3M Company - Health Care Business | 3M is recalling specific lots of 3M ESPE Unitek Primary Stainless Steel Crowns - Series UR4 contain LL3 crowns which are incorrectly labeled as UR4 crowns. | Process control | 3M Company sent a Voluntary Recall letterdated June 4, 2013, to all affected consignees. The letter described the problem and the product involved in the recall. Advised consignees to review their inventory and quarantine the product and do not further distribute it. For replacement they should call Customer Service at 3M ESP Dental Products at 1-800-237-1650. If they have further distributed the product they requested they contact their customers to inform them of the recall. 3M prepared a letter to be used in their communications with their customers. The customers should contact 1-800-634-2249.For questions regarding this recall call 651-733-5718. |  | 1419 | 1419 (1325 US, 94 OUS) | Worldwide Distribution - USA including IN, TX, PA, FL, NV, SC, and TN and Internationally to Canada and Saudi Arabia. | Terminated | December 11, 2013 | 189 | 1419 | Not_Computer
Z-0415-2014 | 66864 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | SIGMAÂż LCSÂż HIGH PERFORMANCE" |  |  | November 18, 2013 | 2013 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics, Inc. is issuing a voluntary device correction of the HP M.B.T. Keel Punch Knee Instrument because of the potential for tabs on the connection end to fracture and be left in the patient. | Device Design | DePuy Orthopaedics sent an Urgent Information - Medical Device Correction letter to all affected customers on November 18, 2013. Immediately after initiation on November 18, the company halted initiation with our US distributors and international affiliates because the surgical technique referenced and images from the surgical technique were not correct. No surgeons received the November 18 surgeon letter. The recall was re-initiated on November 20, 2013. The communication to the surgeons who used the affected devices will be delivered by Distributors either by mail or in person and includes the surgeon letter.The device correction includes awareness, reiteration of surgical techniques and instrument care, and visual check of the device as outlined in the Field Notice and Surgeon letter. The devices are not being removed from the market, but the company is investigating potential design changes to the HP M.B.T. Keel Punch Knee Instrument to reduce the potential for the tabs to fracture. All DePuy United States Distributors and DePuy affiliate contacts for Canada, Latin America and International will be notified via email. The US Distributors will identify surgeons who use the affected instruments on Distributor Response Card #1. The US Distributor will notify surgeons either by mail or in person with the written communication. Once all device correction activities are complete, the US Distributor will complete and return Distributor Card #2 to DePuy Orthopaedics, Inc.The US distributors are the US consignees and are responsible for notifying all new HP M.B.T. Keel Punch Knee Instrument surgeon users of this device correction.All territories are required to ensure the HP M.B.T. Keel Punch Knee Instrument Device Correction is promptly initiated. Customers were asked to review the instructions included, along with the attached documents, whichcontain important information regarding notification of surgeons affected by this Device Correction.Customers |  | 58703880 | 29,235 units in total / 2008 through November 2013 | Worldwide Distribution - USA (nationwide) including AR,AZ,CA,CO,FL,GA,HI,IA,IL,IN,KY,LA,MA,MD,ME,MI,MN,MT,NC,NV,NY,OH,PA, TX,VA,WA,WI and Internationally to: AUSTRALIA AUSTRIA BELGIUM BRAZIL CANADA CZECH REPUBLIC DENMARK EGYPT FRANCE GERMANY GREECE HOLLAND HUNGARY INDIA IRELAND ISRAEL ITALY JAPAN KOREA MALAYSIA NEW ZEALAND NORWAY P R CHINA PAKISTAN PANAMA POLAND PORTUGAL PUERTO RICO RUSSIA SAUDI ARABIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SWEDEN SWITZERLAND TURKEY UAE UK | Terminated | February 24, 2015 | 463 | 58703880 | Not_Computer
Z-0410-2014 | 66414 | No | OAR | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | Vertebroplasty Needle Kit 10g Diamond Tip |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | Complaints were received of inconsistencies in the expiration dates shown on the packaging of the Vertebroplasty Needle Kit 10g Diamond Tip device. | Incorrect or no expiration date | Synthes sent an Urgent Notice: Medical Device Notification letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to returny any of the affect devices, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section indicating the affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers were also instructed to return the Verification Section even if they don't have the affected product. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 40 | 40 | Nationwide Distribution including IN, MT, HI, and FL. | Terminated | April 16, 2015 | 625 | 40 | Not_Computer
Z-0140-2014 | 66314 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens RAD Fluoro Uro system (Luminos dRF and Uroskop Omnia with software version VD10A/G |  |  | September 03, 2013 | 2013 | Siemens Medical Solutions USA, Inc | When using systems operating with software versions VD10A/G during a RAD examination, a malfunction can result in radiation continuing longer than required resulting in an overexposed image that is not of diagnostic quality, prompting repetition of the examination. | Software design | Siemens sent an Field Safety Notice dated September 3, 2013, to all affected customers.The letter identified the product the problem and the action to be taken by the customer. Action to be taken: 1. It is not possible for the user to prevent this malfunction. The exposure time with the use of automatic exposure control is limited to 500ms, which is the maximum duration of radiation in this mode. The exposure can be interrupted by the user at anytime by releasing the exposure button.2. There are two possibilities to solve this issue:a.) Switch off the system if the message "wait for FLC" in status bar does not disappear approximately. 30s after the OGP selection. The system gets ready after switching the system on and reselecting the Torno OGP.b.) Press Shift+F10. The pop up message has to be confirmed and you have to wait approximately.40s until! the "i" icon in the status bar stops blinking. After that a RAD OGP has to be selected and then the Torno OGP. Thereafter, the system gets ready.3. It is not possible for the user to prevent this malfunctionThe Center for Devices and Radiological Health (CDRH) understands that yourCorrective Action Plan (CAP) will involve 59 units and consist of the following elements:1. You have notified affected customers with Field Safety Notice XP032/13/S.2. You have prepared a field modification of the Luminos dRF and Uroskop Omniasystems with software version VD10A/G.3. Distribution of this field modification will began in September 2013 as UpdateInstruction XP015/13/S at no cost to the user.For further Customer Support call (610) 219-6300. |  | 59 | 59 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | March 09, 2015 | 552 | 59 | Computer
Z-0406-2014 | 66413 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 2.4 mm LC-DCP Plate 6 Holes/51mm |  |  | July 30, 2013 | 2013 | Synthes (USA) Products LLC | A complaint was reported regarding a screw being found in a package labeled for 2.4 mm LC-DCP Plate. | Labeling mix-ups | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory, immediately remove the affected product and return to Credit/Returns, Synthes. For questions call 610-719-5450. |  | 13 | 13 | Worldwide Distribution - USA (nationwide) including IN, NY, WA, NE, MD, TX, AL, MI, and Internationally to Canada. | Terminated | August 17, 2015 | 748 | 13 | Not_Computer
Z-0409-2014 | 66418 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 4.5 mm LCP Curved Condylar Plate |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | The Synthes 4.5 mm LCP Curved Condylar Plate was recalled because a Straight Condylar Plate is etched with a Curved Condylar Plate part number. | Process control | Synthes sent an Urgent Notice: Medical Device Notification letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to returny any of the affect devices, call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section indicating the affected product has been located, also indicate the number of devices found and note the Return Authorization Number. Return the Verification Form to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132Customers were also instructed to return the Verification Section even if they don't have the affected product. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 13 | 13 | Worldwide Distribution - USA including ID, FL, TN, NY, LA and Internationally to Canada. | Terminated | August 31, 2015 | 762 | 13 | Not_Computer
Z-0407-2014 | 66419 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Titanium Cannulated Humeral Nail |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | The packaging incorrectly depicts the Synthes Titanium Cannulated Humeral Nail as a 3 distal hole nail instead of a 2 distal locking hole nail. | Pending | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization Number, complete the Verification Section and return to Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132. If customers do not have any affected product they should still complete Verification Form and return to Synthes by fax to 1-888-731-7954 or Scan/email: FieldAction@synthes.com.Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-519-5000. |  | 14 | 14 | Worldwide Distribution - USA (nationwide) including SC, IA, NC, TX, KY, VA, AL, PA, and Internationally to Canada. | Terminated | September 10, 2015 | 772 | 14 | Not_Computer
Z-0411-2014 | 66654 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Shortened Plug Driver; Low Volume Specialty Device; Qty:1; Rx Only; Non Sterile; Biomet Spine, Parsippany, NJ 07054 |  |  | September 24, 2013 | 2013 | Ebi, Llc | Biomet received a complaint on 5/20/2013 reporting that the tip of the instrument sheared off during a removal case. | Nonconforming Material/Component | Biomet sent an Urgent Medical Device Recall Letters/Fax Back Response Form dated September 24, 2013, via Fed Ex. The customer is asked to examine their inventory and return the recalled product following the attached instructions. The Fax Back Response Form should be completed and returned to 973-887-1347 prior to the return of the product. Any questions please call (973) 299-9300 x 2162. |  | 2 | 2 units | US Distribution to CA | Terminated | May 15, 2015 | 598 | 2 | Not_Computer
Z-0426-2014 | 65985 | [6 empty] | Custom Ultrasonics System 83 Plus 2 and Plus 9 Endoscope Washer/Disinfector. |  |  | March 12, 2013 | 2013 | Custom Ultrasonics, Inc. | Certain System 83 Plus 2 and System 83 Plus 9 units running the Windows 7 operating system were not fully validated prior to distribution. | Software design | A notification letter, dated February 26, 2013, was sent to direct accounts to inform them of the recall and that technicians would visit facilities to replace the Windows 7 operating system with an appropriately validated Windows XP operating system. |  | 118 | 118 | Worldwide Distribution: USA (nationwide) and Internationally to: Canada, Slovakia Republic, Sydney and Peru. | Terminated | June 26, 2014 | 471 | 118 | Not_Computer
Z-0408-2014 | 66412 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Slipped Capital Femoral Epiphysis System |  |  | July 31, 2013 | 2013 | Synthes USA HQ, Inc. | There is a possibility for the Guide Wires of the Slipped Capital Femoral Epiphysis System to stick in the cannulation of the drill bit and the drill bits to break during surgery. | Device Design | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review their inventory, immediately remove the affected product and return to Credit/Returns, Synthes. For questions call 610-719-5450. |  | 166 | 166 | USA Nationwide Distribution | Terminated | June 01, 2015 | 670 | 166 | Not_Computer
Z-0425-2014 | 65356 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 13mm addEASE(TM) Binary Connector with 17 Ga. Needle for use with PAB(R) Containers, BC1300 |  |  | April 04, 2013 | 2013 | B. Braun Medical, Inc. | Possible solution leakage at the interface of the addEASE(TM) Binary Connector (Product Codes BC2000 and BC1300) and the PAB(R) bag. Based on extensive root cause investigation, it has been determined that the leakage is the result of a misaligned cannula (needle) on the addEASE(TM) Binary Connector, which pierces the PAB(R) bag stopper to access the solution. The misaligned cannula may result i | Equipment maintenance | B Braun sent an Urgent - Medical Device Removal Notification letter dated April 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by he customer. Customers were instructed to review the Product Recall Notification in its entirety and ensure that all users in their organization are informed about this voluntary product recall. Determine their current inventory of the affected lots within their distribution center. Contact their customers that have received the affected lots and inform them of this voluntary product recall. Combine both their inventory, and the inventory reported by their customers, and utilize their specific "Product Removal Acknowledgement" form to record the total number of individual units (within partial cases) and the number of full-unopened cases. If no inventory remaining enter zero (0). Return the completed Acknowledgement Form to BBMI Quality Assurance in the self addressed envelope or fax to 610-849-1197. Customers should call Customer Support Department at 800-227-2862 to arrange for eturn and replacement product or for questions.For questions regarding this recall call 800-227-2862. |  | 4000 | 4,000 units | Nationwide Distribution | Terminated | August 26, 2014 | 509 | 1750079 | Not_Computer
Z-0429-2014 | 66530 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Soft Tissue Retractor Small Extendible. |  |  | August 01, 2013 | 2013 | Synthes USA HQ, Inc. | The small extendible Soft Tissue Retractor was recalled due to locking nut (clamping sleeve) malfunction. | Use error | Synthes sent an Field Safety Notification letter dated August 1, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to take the following actions:Exchange the old surgical technique guide with the new guide version F provided with this notification.Review the revised surgical technique guide.Forward this Field Safety Notification to anyone in your facility that needs to be informed.If the technique guide has been forwarded to another facility, contact that facility.Maintain awareness of this notice until all technique guides have been exchanged.Maintain a copy of this notice.Please complete and return this page by either of the following methods: Fax: 877-567-9320 or Scan/email: FieldAction@synthes.comThank you for your attention to this issue. |  | 374 | 374 | Worldwide Distribution - USA (nationwide) and Internationally to Canada and Mexico. | Terminated | April 16, 2015 | 623 | 374 | Not_Computer
Z-0183-2014 | 66610 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical's InziiÂż 12/15mm Retrieval System. |  |  | October 21, 2013 | 2013 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall of the InziiÂż 12/15mm retrieval system. During shipment, the retrieval system packaging has the potential to become punctured with small holes, which could compromise the sterile barrier. The likelihood of this situation to occur and result in patient harm is highly unlikely; however, out of an abundance of caution for patient safety and a commitmen | Package design/selection | Customer notifications of the recall began the week of 10/21/2013 via UPS overnight with letters requesting that all affected products be returned to Applied. The notification letter is dated October 21, 2013 and titled "URGENT: MEDICAL DEVICE RECALL". The notification was accompanied by a Recall Notification Confirmation Form. The letter informed customers of the recall by providing the product description with codes (Model and Lot Numbers). reason for the recall, actions to take and contact information. |  | 29506 | 29,506 units | Worldwide Distribution: US (nationwide) and countries of: Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Lebanon, Malaysia, Mexico, New Zealand, Netherlands, Norway, Peru, Poland, Portugal, Saudi Arabia, Singapore, Spain, South Africa, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | September 15, 2014 | 329 | 29506 | Not_Computer
Z-0137-2014 | 66353 | No | JCA | Bleeding time device. | Hematology | Hematology | 510(K) Exempt | Surgicutt fully automated incision making device for bleeding time determination. ITC 8 Olsen Avenue, Edison, NJ 08820. 732.548.5700. www.itcmed.com. A subsidiary of Thoratec Corporation. |  |  | August 22, 2013 | 2013 | ITC-Nexus Dx | ITC has determined that the blade in their Surgicutt with Blotting Paper (Product Code: SUB50I; Lot: AK301M) devices may be loose. | Nonconforming Material/Component | ITC sent a Medical Device Recall letter dated September 20, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove/quarantine affected product in their inventory and to return the affected product to ITC. Customers with questions were instructed to call 1-800-631-5945.Questions regarding this recall call 732-548-7000, ext 4476. |  | 300 | 6 boxes (Each box contains 50 Surgicutt Devices plus 55 sheets of Blotting Paper). | Nationwide Distribution including Arizona andTexas.. | Terminated | June 09, 2014 | 291 | 300 | Not_Computer
Z-0430-2014 | 65408 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista(R) 500 and Dimension Vista(R) 1500 - Reloaded Vista Vials - limited to systems using software versions 3.4 and 3.5.1. |  |  | May 20, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | There is the potential for the vial volume of "in use" Vista calibrators, controls, or sample diluent to incorrectly revert to full volume when removed and reloaded. Dimension Vista(R) 500 and Dimension Vista(R) 1500 - Reloaded Vista Vials - limited to systems using software versions 3.4 and 3.5.1 | Software design | A notification letter was sent to consignees in May 2013. |  | 776 | 776 | Worldwide Distribution: USA (nationwide) and iinternationally to Canada. | Terminated | April 10, 2014 | 325 | 776 | Computer
Z-0431-2014 | 66705 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Easy Diagnost Eleva |  |  | October 30, 2013 | 2013 | Philips Medical Systems, Inc. | The delivered application software was an outdated version to the model of equipment. | Software design | Philips Medical Systems sent an Urgent - Field Safety Notice dated October 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Cusromers were informed that the delivered application software was an outdated version to the model of their expipment and that a Philips Field Engineer will be implementing an upgrade of the application software. Customers were instructed to review the information contained in the letter with all memebers of their staff who need to be aware of its contents. Customers with questions were instructed to contact their Philips representataive: Technical Support Line 1-800-722-9377.For questions regarding this recall call 1-800-722-9377. |  | 19 | 19 devices | Nationwide Distribution including MN, NH, DE, OR, MA, TX, OH, FL, CA, WA, and SC. | Terminated | January 09, 2017 | 1167 | 19 | Computer
Z-0438-2014 | 66724 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SIEMENS brand ARTISTE Linear Accelerator Linac systems with component: any photon unflat beam option (Multiple X feature) in combination with the IMRT or mARC option. A family of linear accelerator systems is to deliver X-ray photon and electron radiation for the therapeutic treatment of cancer. The ARTISTE/ONCOR systems are manufactured by: SIEMENS AG, MEDICAL SOLUTIONS, GERMANY and distributed SIEMENS Healthcare IM CR RO, CR Radiation Oncology, Martinez, CA |  |  | October 23, 2013 | 2013 | Siemens Medical Solutions USA, Inc | A safety risk regarding a potentially existing dark current radiation phenomenon on the Linear Accelerator (LINAC) in combination with IMRT or mARC (rotational IMRT) treatments using unflat (flattening filter free) beams. In case there are many long periods of field shape changes in a treatment plan, the dose due to dark current radiation may sum up to a significant value. | Device Design | Siemens sent an Urgent Medical Device Correction letter dated October 23, 2013 including the release of update instruction TH024/13/S to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. |  | 26 | 26 units | Worldwide Distribution: USA (nationwide) including states of: FL, NC, NY, PA, TN, UT, and WI; and countries of: Angola, Brazil, Canada, P.R. China, Germany, India, Japan, Republic Korea, Poland, and Thailand. | Terminated | January 22, 2014 | 91 | 29 | Not_Computer
Z-0439-2014 | 66764 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | C-Series: Clinac, Trilogy, Trilogy Tx. , Novalis high energy linear accelerators and UNIQUE single energy linear accelerator; Versions 7, 8 and 9. The UNIQUE is not sold in the US. |  |  | October 24, 2013 | 2013 | Varian Medical Systems, Inc. | This correction is to notify users that a solution to a previous correction has been developed and Varians reps will be contacting locations to schedule installation. | Device Design | Varian sent an Urgent Medical Device Correction letter, P/N CP-05645 Rev D dated October 24, 2013 to their affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated all affected sites will be contacted by Varian representatives to schedule an upgrade to their devices. Questions may be directed to 888-827-4265 for USA and Canada. Europe +41 41 749 8844 |  | 3809 | 3,809 devices | Worldwide Distribution - USA Nationwide and in the countries of Algeria, Angola, Argentina, Australia, Austria, Azerbaijan, Bangladesh, Belgium, Belorussia, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Georgia, Germany, Greece, Guadeloupe, Guam, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Uthuania, luxembourg, Macau, Macedonia, Malaysia, Martinique, Mauritania, Mexico, Moldova, Monaco, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Reunion, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, SriLanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Zimbabwe. | Terminated | January 13, 2016 | 811 | 3809 | Not_Computer
Z-0436-2014 | 66866 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Assembly, Patient Side Manipulator on da Vinci Surgical Systems: |  |  | November 11, 2013 | 2013 | Intuitive Surgical, Inc. | Reports of friction within certain instrument arms can interrupt smooth instrument motion. This can be felt by the surgeon as resistance in the movement of the master.In this situation, the instrument can stall momentarily and then suddenly catch-up to the master position if the surgeon pushes through the resistance. | Device Design | The firm, Intuitive Surgical, sent an "Urgent Medical Device Recall" letter dated November 11, 2013 all affected customers via traceable mail. The letter described the product, problem and actions to be taken. The customers were instructed to take the following actions: 1. Ensure all appropriate personnel are fully informed of this Notification. Forward this letter to all Surgeons and OR staff who perform da Vinci surgical procedures, Risk Manager, OR Director, Purchasing, and Biomedical Engineering staff. 2. Complete the attached Acknowledgement Form and return it to Intuitive Surgical via Fax to: Intuitive Surgical, Inc., ATTN: REGULATORY AFFAIRSU.S. +1 (408) 716-3040, or Scan and Email: isi.compliance@intusurg.com.3. An Intuitive Surgical Representative will contact your site to schedule the inspection of the instrument arms on your system(s) as indicated above. Please work with the Intuitive Surgical Representative to schedule this activity.4. Please inform appropriate personnel when the correction has been completed.5. Please retain a copy of this letter and the Acknowledgement Form for your files.If you need further information or support concerning this medical device correction, then please contact your Clinical Sales Representative or contact Intuitive Surgical Customer Service at the numbers listed below. Âż North and South America: 800-876-1310 Option 3 (6 am to 5pm PST) Âż Japan: 0120-56-5635 or 03-5575-1362 (9am to 6pm JST) Âż South Korea: 02-3271-3200 (9am to 6pm KSTJ) Âż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 am to 6 pm CET). |  | 1386 | 1,386 | Worldwide Distribution: US (nationwide) and to countries of: Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Denmark, France, Germany, India, Israel, Italy, Japan, S. Korea, Luxembourg, Mexico, Monaco, Netherlands, Norway, Romania. Russian Federation, Singapore, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | February 18, 2014 | 99 | 1386 | Not_Computer
Z-0435-2014 | 66758 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | VersaSafe Extension Set, Model No. 21000M-007 |  |  | November 08, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the VersaSafe Extension IntravascularAdministration Set because it may have excessive adhesivewhich will prevent the male spin collar to operate as intended. | Employee error | A customer letter and a distributor letter will be distributed on 11/8/13 to all customers who purchased the VersaSafe Extension Set, Model 21000M-07, Lot Number 13066550. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return to CareFusion the enclosed, pre-addressed and postage paid, Recall Response Card. |  | 3500 | 3500 units | Nationwide Distribution. | Terminated | April 17, 2014 | 160 | 3500 | Not_Computer
Z-0434-2014 | 66502 | No | FRG | Sterilization wrap. | General Hospital | General Hospital | 510(k) | STERITITE Containers: STERITITE Container Kit, 1/4 size, 3" high, product code: 13902; STERITITE Mini Mid Size, 3" high, product code: 13903; STERITITE Container Kit, 1/2 size, 4" high, product code: 13904; STERITITE Container Kit, 3/4 size, 4" high, product code: 13905; STERITITE Container Kit, Full size, 4" high, product code: 13906. |  |  | October 03, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the SteriTite containers because they do not have consistent data to show that sterility assurance level of 10(-6) is achieved during the half-cycle test runs on new, unprocessed containers. | Device Design | Advanced Sterilization Products (ASP) sent a Urgent Medical Device Field Safety Notification letter dated October 3, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers who received the SteriTite Containers.Customers were instructed to:Provide this notice to anyone in your facility that needs to be informed.Maintain a copy of this notice with the affected product.Maintain awareness of this notice.Customers are instructed to contact Stericycle at (888) 912-7092 for additional questions related to the recall letter. Customers are instructed to report any complaints or suspected problems related to the recall by contacting ASP at (888) STERRAD (888-783-7723). |  | 267 | 267 units total (252 units in US) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Vietnam, Thailand, Malaysia, China, Korea (South), Mexico, Spain, Singapore, Hong Kong and Portugal. | Terminated | January 13, 2014 | 102 | 267 | Not_Computer
Z-0433-2014 | 65147 | No | NSX |  |  | General Hospital | Contact ODE | syngo(R) Lab Data Manager - Version VA11B and VA12A Systems |  |  | May 10, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Under certain conditions the system may not perform as intended, causing the release of results to the Laboratory Information System that should have been held for manual review due to auto-verification rules or the delay/omission of result transmission to the LIS. | Software design | Siemens sent an Urgent Medical Device Correction letter in May 2013 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the letter in its entirety and ensure that all operators understand the information presented. Specific actions needed to be implemented by the customer were outlined in the letter. Customers were instructed to retain the letter with their laboratory records, and forward the letter to those who may have received the affected product. Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative.For questions regarding this recall call 800-441-9250. |  | 262 | 262 | Worldwide Distribution: USA (nationwide) and internationally to Canada. | Terminated | August 26, 2014 | 473 | 262 | Not_Computer
Z-0450-2014 | 66759 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set with Texium Close Male Luer, Model No. 24301-0007T |  |  | November 08, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set with Texium Closed Male Luer because it is not assembled with low-sorbing tubing. | Employee error | The firm, CareFusion, sent an "URGENT: Medical Device Recall Notification" letter dated November 8, 2013 to all customers who purchased the SmartSite Low Sorbing Infusion Set with Texium Closed Male Luer. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to immediately complete and return to CareFusion the enclosed, pre-addressed and postage paid, Recall Response Card and return all recalled product directly to the distributor form whom it was purchased. Customers with questions are instructed to contact their distributor directly or call the CareFusion Support Center at (888) 562-6018.If you have recall related questions or need support, please call CareFusion Support Center at 1-888-562-6018; hours: 7am to 4pm PST or Technical Support at 1-888-812-3229; hours 7am to 5pm PST. |  | 480 | 480 units | US Distribution in CA, WY, OH, UT, and MA. | Terminated | May 27, 2014 | 200 | 480 | Not_Computer
Z-0441-2014 | 66828 | No | FED | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Navigator HD l3F/15F x 36cm Ureteral Access Sheath Set Product No M0062502280, RDF Catalog No 250-228, Sterile. |  |  | November 05, 2013 | 2013 | Boston Scientific Corporation | The 13F/15F Navigator HD Ureteral Access SheathSet are being recalled after receipt of 5 reports where the tip of the dilator has separated fro the body of the dilator. | Component design/selection | Urgent recall notification letters were sent via Federal Express Priority mail on November 5, 2013. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it toBoston Scientific. The notification also requests the completion of a Reply Verification Tracking Form asa method of documenting the presence or absence of affected product in their inventory. Follow-up communications will be made to the accounts that do not respond to this inquiry. |  | 939 | 939 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Ireland, Denmark, Italy, France, Canada, Great Britain, Switzerland, South Africa, Germany, and the Netherlands. | Terminated | December 08, 2014 | 398 | 1662 | Not_Computer
Z-0443-2014 | 66632 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | Blanketrol III Data Export Software Version 1.10 |  |  | May 15, 2013 | 2013 | Cincinnati Sub-Zero Products Inc | During a customer inquiry regarding a beta version of the Bill Data Export Software, Model USB-127, it was discovered that the wrong software version was released to a customer. | Software Manufacturing/Software Deployment | Cincinnati Sub-Zero sent a Urgent Medical Device Notification letter dated July 15, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They instructed their customers to to do the following:1) Upon receipt of the corrected Blanketrol III Data Export Software disc, immediately collect and destroy all previous Blanketrol III Data Export Software disc from this lot # currently in possession and replace all downloaded Blanketrol III Data Export Software with the correct version provided by CSZ.2) After the Blanketrol Data Export Software Version 1.00 have been destroyed and downloaded versions replaced with version 1.10, complete and return the enclosed response form as soon as possible to acknowledge receipt of this notification and to inform CSZ that you have performed and completed the requested actions. Return the form by fax to (513) 772-9119 scan and e-mail the information to BIIIExportSoftware@cszinc.com, or mail to:Cincinnati Sub-Zero12011 Mosteller RoadCincinnati, OH 45241 U.S.A.For further questions please call 1-(800) 989-7373 or (513) 772-8810. |  | 12 | 12 Units | Worldwide Distribution - US Distribution including the following states AK, CO, FL, MA, MI, and NC. and Internationally to Canada, Germany and United Kingdom. | Terminated | May 13, 2014 | 363 | 12 | Computer
Z-0451-2014 | 66757 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Abbott i-STAT PT/INR cartridges Abbott Point of Care Inc., Princeton, NJ 08540 |  |  | October 28, 2013 | 2013 | Abbott Point Of Care Inc. | Abbott Point of Care Inc (APOC) has determined that i-Stat PT/INR cartridges have the potential to exhibit incorrectly elevated INR. | Nonconforming Material/Component | Abbott Point of Care Inc. (APOC) issued an Urgent Recall Notice on October 28, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of cartridges from PT/INR lots in this letter. All unused cartridges would be credited and are to be returned as instructed on the enclosed Business Reply Card. If customers provided cartridges from the lot number listed above to another institution they were instructed to provide a copy of this letter to them.Customers with questions were instructerd to contact Abbott Point of Care Technical at 1-800-366-8020, Option 1.For questions regarding this recall call 800-366-8020. |  | 914400 | 794,928 units US; 119,472 units OUS | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Brazil, Chile, China, Finland, France, Germany, Greece, Hong Kong, Ireland, Italy, South Korea, Kuwait, Mexico, Netherland, New Zealand, Nigeria, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, south Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, UAE, and United Kingdom, | Terminated | May 29, 2014 | 213 | 914400 | Not_Computer
Z-0448-2014 | 66333 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | GENESIS(R) II, PS HIGH FLEXION ARTICULAR INSERT, SIZE 5-6, 11 MM QTY: (1), REF 71421516, STERILE EO, Smith &amp; Nephew, Inc. Memphis, TN 38116 USA |  |  | August 27, 2013 | 2013 | Smith &amp; Nephew Inc | The Tyvek(TM) lid on one, or both, of the two trays that form adouble barrier around the insert may not have been adequately sealed. | Package design/selection | Smith &amp; Nephew sent an Urgent Recall Notification letter dated August 27, 2013, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. They were instructed to cease use and distribution of the product and quarantine for return. Customers were instructed to contact Smith &amp; Nephew via email at FieldActions@smith-nephew.com or fax 901-566-7975 to obtain a return authorization (RA) number. Customers with questions were also instructed use this email address. Customers with affected product were instructed to return it along with the completed inventory Return Certification Form to the address listed on the form. Non-responding consignees were notified again on September 27, 2013, by email.For questions regarding this recall call 901-396-2121. |  | 46 | 46 pieces | Worldwide Distribution - USA including Nevada and Internationally to United Arab Emirates, Italy, Spain, Belgium, Bulgaria, Poland, France, Japan, Netherlands, Australia, Turkey, Russian Federation, Germany and United Kingdom. | Open, Classified |  |  | 227 | Not_Computer
Z-0454-2014 | 64762 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | The GE Centricity Web; Software Version 3.0x, is intended for use under the direct supervision of a licensed healthcare practitioner. The Centricity Enterprise Web is an image and information distribution system for the clinical review of medical images and reports. The Centricity Enterprise Web is not intended for primary diagnosis. The Centricity Enterprise Web is available as a supplemental sub-system to Centricity PACS or as a stand alone Web-based image and information distribution system. The Centricity Enterprise Web provides both the server software and a client application, which utilizes off the shelf browser technology. Centricity Enterprise Web supports several interface protocols which allows the system to be integrated with any image or information systems such as a hospital information system (HIS), radiology information system (RIS), or electronic medical record (EMR), which supports one of the supported integration protocols. |  |  | April 19, 2012 | 2012 | Ge Healthcare It | GE Healthcare has become aware of an issue associated with rejected images using the WADO protocol within GE Centricity Web that may impact patient safety. The WADO (Web Access to DICOM Objects) interface allows for embedding DICOM images as JPEGs on a web page and is available as a licensed operation on Centricity Web. When rejected images are archived by Centricity PACS to Enterprise Archive, an | Software design | Urgent Medical Device Correction letters dated April 19, 2013 were sent to customers. The letter requests that customers ensure that their system is configured to disable the WADO interface using the included instructions until the software revision can be installed at their site. The letter further informs customers that updated software will be sent to them for installation.The software has been modified to correct these defects. The modified software will be released and installed in user sites through GE Healthcare FMI# 85206R. |  | 53 | 53 units | US Distribution incluidng states of: CA, FL, GA, IA, IL, LA, MO, NJ, NY, OK, TN, WA, and WI. | Open, Classified |  |  | 53 | Computer
Z-0455-2014 | 66538 | No | OVE | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Synthes Zero-P VA Implant 7 MM Height Parallel Sterile |  |  | September 30, 2013 | 2013 | Synthes USA HQ, Inc. | A potential non conformance was identified with the Zero-P VA plate where the screw could back out if the blocking mechanism is not visually confirmed to properly cover the head of the screw. | Process design | Synthes sent an email message dated March 23, 2012, to the sales representative who received the affected product. The email identified the product, the problem, and the action to be taken. The letter requested the sales rep examine inventory, stop using the product, and quarantine the affected product to ensure that it will not be used. Synthes instructed the customer not to remove any product from their account or return any product without further instruction. For questions they were instructed to call 610-719-5364.For questions regarding this recall call 610-719-5000. |  | 1 | 1 | Nationwide Distribution - only AZ. | Terminated | April 16, 2015 | 563 | 1 | Not_Computer
Z-0453-2014 | 66896 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Sunrise Brand Zippie Voyage, Chair with Casters, Early Intervention Device (Stroller), Model: EIZ14; Product is manufactured and distributed by Sunrise Medical (US) LLC., Fresno, CA. |  |  | November 11, 2013 | 2013 | Sunrise Medical (US) LLC | There is a remote possibility for the plastic latch to deform over the lifetime of use. When this happens there is the possibility for the seat to become dislodged from the base. And when not using the positioning strap, there is a risk that the occupant could slide out of the seat and sustain injury. | Nonconforming Material/Component | Urgent Medical Device Field correction letters were sent to Sunrise Suppliers/Dealers and Zippie Customers on November 2013 by first class mail. Foreign notification was supplied to suppliers and regulatory affairs director in Canada. |  | 345 | 316 units in US, 29 units international | Worldwide Distribution-USA (nationwide) and the countries of Canada, Australia, Hong Kong, and China. | Terminated | March 18, 2014 | 127 | 345 | Not_Computer
Z-0432-2014 | 66407 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TB-0545PC: THUNDERBEAT 5 mm , 45 em, Pistol Grip TB-0535PC: THUNDERBEAT 5 mm, 35 em, Pi stol Grip TB-0545IC: THUNDERBEAT 5 mm, 45 em, In line Grip TB-0535IC: THUNDERBEAT 5 mm, 35 em, Inline Grip TB-0520IC: THUNDERBEAT 5 mm, 20 em, In line Grip TB-051 OIC: THUNDERBEAT 5 mm, I 0 em, In line Grip |  |  | September 27, 2013 | 2013 | Olympus America Inc. | Complaints of damage to various models of Thunderbeat Hand Instruments during surgical procedures, including cracking, breaking, or deformation of the components of the probe tip or jaw. | Use error | Olympus sent an Urgent Medical Device Field Correction Action letter dated September 27, 2013, to all affected customers. The letter identified the product the problem and the actions needed to be taken by the customer.Please make the attached Procedural Tips and Techniques instructions available to all users of the THUNDERBEAT hand instruments. Please also maintain a copy with the Instruction Manual(s) fot"the device.OCA requires that you indicate on the enclosed questionnaire your acknowledgement of receipt of this corrective action and the contact information for your facility.This additional user instruction is also being added to the Instruction Manual(s) for the THUNDERBEAT hand instruments. If you would like a copy of the updated Instruction Manual(s), please contact OAI Customer Service at 1-800-848-9024.For further questions please call (484) 896-5688. |  | 8709 | 8709 | Worldwide Distribution - USA (nationwide) including Puerto Rico and Internationally to Canada. | Terminated | February 02, 2015 | 493 | 8709 | Not_Computer
Z-0452-2014 | 66417 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Titanium Trochanteric Fixation Nail (TFN)System |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | Two lots of Titanium Trochanteric Fixation Nail (TFN) System are missing the second bend in the nail. | Nonconforming Material/Component | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Synthes at 1-800-479-6329 if they have any of the affeccted product to obtain a Return Authorization Number, Complete the Verification Section of the letter by checking the appropriate box indicating the affected product has been located. Return the Verification Form with the product to:Credit/ReturnsSynthes1101 Synthes AvenueMonument, CO 80132.Customers should complete and return the Verification Frorm even if they do not have any of the affected product and return by fax to 1-888-386-2077 or Scan/email: FieldAction@synthes.com. Customers with questions were instructed to call 610-719-5450 or contact their Synthes Trauma Sales Consultant.For questions regarding this recall call 610-719-5000. |  | 17 | 17 | Nationwide Distribution including IL, CA, AZ, FL, TX, UT, VA, NY, NM, and MT. | Terminated | September 10, 2015 | 772 | 17 | Not_Computer
Z-0457-2014 | 66431 | No | GCX | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Product is labeled in part: LIST NO. 43449; RECEPTAL CANISTER; DO NOT FILL 900 ABOVE THIS LINE; Product of China; Hospira, Inc.,; Lake Forest, IL 60045 USA. |  |  | July 17, 2013 | 2013 | Hospira Inc. | Hospira has recieved reports that the vacuum needed to create suction with the 1L Receptal System cannot be achieved due to the canister and liner not properly seating. | Process control | An URGENT DEVICE RECALL letter dated July 17, 2013 was mailed to direct consignees. The letter requested the consignees: 1) quarantine any affected products, 2) complete and return the reply form included with the letter, 3) notify healthcare professionals within the consignee's organization of the recall, 4) return the affected products to Stericycle in accordance with the instructions in the letter. Consignees are further requested to forward a copy of the notification materials to any customers that they may have distributed the products to. |  | 7576 | 3,560 units distributed from June 2004 - December 2012 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CO, FL, GA, IL, IN, KS, MD, MI, MN, MO, MT, NC, ND, NE, NJ, NY, PA, SD, TX, UT, VA, and WA and countries of: Canada and EMEA. | Open, Classified |  |  | 7576 | Not_Computer
Z-0458-2014 | 66627 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon Stratafix Symmetric PDS Plus Knotless Tissue Control Device Ethicon LLC Guaynabo, Puerto Rico 00969. |  |  | September 25, 2013 | 2013 | Ethicon, Inc. | Ethicon Stratafix Symmetric PDS Plus Knotless Tissue Control Device has a small number of tab failures and fascial dehiscences in lower abdominal incisions. | Unknown/Undetermined by firm | Ethicon Inc. sent recall notifications/business reply forms on 9/25/2013 via UPS next day certified mail. |  | 37164 | 37,164 eaches | Worldwide Distribution-USA (nationwide) and countries of: Caribbean, Chile, Singapore and Malaysia. | Completed |  |  | 37164 | Not_Computer
Z-0461-2014 | 66809 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | NexGen Complete Knee Solution |  |  | November 19, 2013 | 2013 | Zimmer, Inc. | This is an expansion of the June 2013 recall of NexGen Micro components. Five (5) additional instruments have been identified that are associated only with the NexGen Micro components. There is the potential for these devices to be inadvertently utilized during surgery, resulting in a delay. The associated implants would not be available due to the previous recall. | Labeling design | Zimmer sent an Urgent Medical Device Recall Notice dated November 19, 2013, to all affected customers. The Notice outlined the products recalled; the reason for the recall, possible adverse events related to the recall and directed consignees to locate, quarantine and count all affected product recording the information on the Inventory Return Form acknowledging that recall directions were followed. Further questions should be directed to 1-877-946-2761. Affected product will be removed by a Zimmer Sales representative, Question should be directed to 1-877-946-2761 8a.m - 5 p.m. EST.For questions regarding this recall call 800-613-6131. |  | 638 | 638 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Canada, Denmark, Ecuador, Germany, India, Italy, Japan, Korea, Malaysia, Mexico, Netherlands, Peru, Singapore, South Africa, Spain, Switzerland, UAE, and UK. | Terminated | May 06, 2015 | 533 | 3027 | Not_Computer
Z-0466-2014 | 66948 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Vasoview Hemopro Endoscopic Vessel Harvesting System Maquet Cardiovascular LLC 45 Barbour Pond Drive Wayne, NJ 07470 |  |  | October 30, 2013 | 2013 | Maquet Cardiovascular, LLC | Maquet has received an increase in the number of complaints for "self-activated/remains activated/overheats" for the Hemopro VH-30000, "Hemopro 1" | Component design/selection | Maquet Inc. sent an Urgent - Medical Device Recall (Removal) Immediate Action Required Letter dated October 30, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please examine your inventory immediately to determine if you have any of the affected VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System (HEMOPRO 1) VH-3000 with the lot number(s), as specified above. The VH-3000 catalog number and lot number are located on the package label. Please remove the affected VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System and place in a secure location.NOTE: VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System HEMOPRO 1 lot numbers not listed are not affected by this recall and should not be removed. HEMOPRO 2 is not affected by this recall.Your MAQUET Cardiovascular Sales Representative will contact you shortly to assist with the completion of the enclosed Acknowledgement Form ( Fax Back Form), the return of any affected VASOVIEW HEMOPRO Endoscopic Vessel Harvesting System HEMOPRO 1 product and arrange for necessary replacement.We apologize for any inconvenience this may cause. If you have any questions, please contact your local MAQUET Cardiovascular Sales Representative or MAQUET Customer Service at 1-888-880-2874, Monday through Friday, between the hours of 6:00 a.m. and 5:00 p.m. PST.Thank you for your cooperation and immediate assistance. |  | 12259 | 12,259 units | Worldwide Distribution - USA (nationwide) and the countries of Brazil, Czech Republic, Denmark, France, Hong Kong, Italy, Middle East, Netherlands, Singapore, Switzerland, and the United Kingdom. | Terminated | April 24, 2015 | 541 | 12259 | Not_Computer
Z-0465-2014 | 66823 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Axillent Debakey Grasping Forceps, Part No. 33821DYD |  |  | November 06, 2013 | 2013 | Karl Storz Endoscopy America Inc | KARL STORZ is recalling the Axillent Debakey Grasping Forceps because the 33821WTD Wavy grasper was mislabeled as 33821DYD Axillent Debakey Grasping Forceps. | Labeling mix-ups | Karl Storz Endosocopy sent a Safety Alert letter dated November 6, 2013, to all customers who received the 33821DYD Axillent Debakey Grasping Forceps. The letter informed the customers of the problems identified and the actions to be taken. Customers were asked to set aside the mislabeled product and return the mislabeled instruments with the RMA (Return Material Authorization) number issued to them. Customers were instructed to fax or email completed recall response form to (424) 218-8519. Customers with any questions were instructed to call (800) 421-0837 ext 8289.For questions regarding this recall call 746-708-8440. |  | 7 | 7 units | Nationwide Distribution including states of: TX, PA, and NY. | Terminated | January 13, 2014 | 68 | 7 | Not_Computer
Z-0456-2014 | 66808 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iba COMPASS SW Version 3.0, Catalog Number CS10-100 (medical linear accelerator) |  |  | August 20, 2013 | 2013 | Iba Dosimetry Gmbh | The development partner of IBA informed that there is problem in a SW component used in COMPASS when converting Region of Interest (ROI) geometries to dose grid representation. Under certain conditions this problem generates too large ROI voxel volumes causing erroneous dose statistics and DVHs. | Software design | Iba sent an Field Safety Notice on August 20, 2013, by email and letter delivered by UPS, Please distribute this notice to all those who need to be aware within your area of responsibility where the affected software have been installed (e.g. colleagues, satellites).Please maintain awareness of this notice and resulting action until we have provided you with a new software version.We would kindly ask you to acknowledge the receipt of this Field Safety Notice. Please sign and send back the attached Confirmation of Receipt within four weeks after the receipt of this Field Safety Notice.We deeply apologize for the inconvenience. We are working on a new software version with corrected functionality. This should become available during next fall and we will ship it to you as soon as it is released, together with the installation and verification instructions.Should you have questions or require additional information, please contact the Customer Service Team at: Service Department IBA Dosimetry GmbH Bahnhofstrasse 5DE-90592 Schwarzenbruck, GermanyTel.:+49 9128 607 38Fax: + 49 9128 607 26Email: service@iba-group.com. |  | 33 | 33 units | Worldwide Distribution - US Distribution including the states of CT, MA, OH, CA, SC, NY and OK, and the countries of International distribution (from Germany) Poland, Germany, GB, The Netherlands, Spain, Turkey, Canada, Mexico, China, Hong Kong and Singapore. | Terminated | February 24, 2014 | 188 | 33 | Computer
Z-0464-2014 | 66830 | [6 empty] | Centurion Sterile # 84 Rubber Bands Reorder EB84, Caution: This product contains natural rubber latex which may cause allergic reactions. |  |  | October 28, 2013 | 2013 | Centurion Medical Products Corporation | Package labeling indicates both "latex free" and "contains natural rubber latex" . The rubber bands do contain natural rubber latex. This could cause a significant risk to users with latex allergies. | Labeling False and Misleading | Centurion sent a Urgent Recall Notification letter via Certified Mail October 31, 2013, return receipt to all affected customers. The affected Centurion Medical Products Corporation sales representatives were notified via email on October 28, 2013. Customers were instructed to destroy all implicated product and complete the accountability record included with the notice and fax to 517-546-3356. Customers were asked to forward a copy of the notice if product was further distributed. Additional notices will be mailed to non-responsive customers via Certified Mail ReturnReceipt, and will be documented in the recall file. For further questions please call (517) 546-5400. |  | 500 | 500 lots | US Distribution including the states of GA, LA and NY. | Terminated | January 29, 2014 | 93 | 500 | Not_Computer
Z-0467-2014 | 66761 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 1) Accolade II 127 Neck Angle Hip Stem, Size # 3 |  |  | October 25, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field indicating that an Accolade II Hip Stem Size #3 was identified in packaging with an Accolade II Hip Stem Size #5. | Packaging process control | Stryker sent an e-mail to their domestic locations on 10/11/2013. Notification letters and product accountability forms were sent via Fed Ex on 10/25/2013 with return receipt to branches/agencies and hospital risk management, chief of orthopaedics and surgeons. |  | 20 | 20 units | Nationwide Distribution including states: AL, GA, IL, and WA. | Terminated | March 31, 2015 | 522 | 20 | Not_Computer
Z-0469-2014 | 66596 | research use only/pathology | OTU |  |  | Pathology | Enforcement Discretion | Digital Pathology System (DPS) Software version 1.1 |  |  | October 01, 2013 | 2013 | Omnyx Llc | There is an issue associated with Digital Pathology System (DPS) Software, in which the annotation tools within the DPS version 1.1 Image Viewer will display incorrect measurements and area calculations for a region of interest on a scanned slide, under certain conditions. | Software design | Omnyx sent a Urgent Field Safety Notice letter dated October 1, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Please discontinue use of the measurement and area calculation features of the annotation tools (line, rectangle/square, ellipse/circle, and freeform) in the DPS Image Viewer at this time. If any of these tools have been used for measurements or area calculations on clinical cases, verify measurements and calculations on glass slides.Please confirm receipt of these instructions by completing the attached form and scanning/emailing it to FieldAction@ omnyx.com or faxing it to 001-412-894-2137Product Correction: An updated version of the DPS SW will be released and installed on your system as soon as it becomes available.Please note that the appropriate regulatory authorities have been informed of this field action.We apologize for any inconvenience that this may have caused you and are working diligently to address this issue.If you have any additional questions or concerns, please contact you GE Healthcare account representative. Further questions please call (412) 891-2100. |  | 2 | 2 | International Distribution to Canada and Switzerland only. | Terminated | April 16, 2014 | 197 | 2 | Computer
Z-0468-2014 | 66586 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Arrow Transradial Artery Access products |  |  | October 09, 2013 | 2013 | Arrow International Inc | There is a risk that the introducer needles packaged within the kits can be partially or fully occluded, which can cause delayed treatment or an adverse event when used on a patient. | Process control | Arrow International sent an Urgent Medical Device Recall Notification letter dated October 14, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product,to return product complete the enclosed Recall Acknowledgement Form and fax to 1-800-343-2935, ATTN: Customer Service. Customers were asked to completed the Recall Acknowledgement Form even if they have none of the affected product.Customers with questions were instructed to contact their local sales representative or Customer Service at 1-800-343-2935.For questions regarding this recall call 610-378-0131. |  | 294430 | 294430 | Worldwide Distribution - USA (Nationwide) and Internationally to Australia, Brazil, Canada, China, Ireland, New Zealand, Singapore, and Thailand. | Terminated | August 12, 2015 | 672 | 294430 | Not_Computer
Z-0427-2014 | 66674 | No | OPT | Near Infrared (NIR) Brain Hematoma Detector. | Ophthalmic | Ophthalmic | 510(k) | Cutera brand truSculpt infrared System (optional), Massager, Vacuum, radio frequency induced heat, Product is manufactured and distributed by Cutera, Inc., Brisbane, CA |  |  | October 16, 2013 | 2013 | Cutera, Inc. | Cutera received two adverse events associated with the edge of the electrode, when it was pressed against neck skin over the thyroid cartilage and due to the absence of adequate underlying conductive tissue, increased skin heating occurred. Both patients developed linear welts and blisters consistent with partial thickness thermal burns to the skin. | Employee error | Cutera sent an IMPORTANT SAFETY INFORMATION letter dated October 16, 2013, to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. It is important that you share this correspondence and the updated Operator Manual andvtreatment guidelines with all staff at your office that use the truSculpt, so they are aware of this new safety information. Please immediately remove and dispose of the current truSculpt Operator Manual and clinical guidelines and replace them with the new and revised truSculpt Operator Manual and clinical guidelines.Customers with questions were instructed to call 415-657-5500 or 415-657-5568. |  | 451 | 451 | Worldwide Distribution - USA (nationwide) and Internationally to Canada | Terminated | December 09, 2013 | 54 | 451 | Not_Computer
Z-0473-2014 | 66582 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes TomoFix Plates |  |  | October 09, 2013 | 2013 | Synthes USA HQ, Inc. | A labeling correction was initiated related to the Surgical Technique Guide J5606-C "TomoFix Osteotomy System" due to the potential for a surgical delay if the surgeon inserts the superior screws at a lesser angle (4 or 6 degrees) and is not aware that those angles are not present. | Labeling False and Misleading | Synthes sent an Urgent Field Safety Notification/Medical Device Labeling Correction letter dated October 9, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to exchange the old surgical technique guide version J5606-C with the new version J5606-D provided with this notification. Review the revised surgical technique guide (Pages 25 &amp; 26). Forward this Field Safety Notification to anyone in their facility that need to be informed.If the surgical technique guide has been forwarded to another facility, contact that facility.Maintain awareness of this notice until all surgical technique guides have been exchanged.Maintain a copy of this notice. Customers with questions were instructed to contact their SynthesSales Consultant. For questions regarding this recall call 610-719-5000. |  | 2912 | 2912 | Worldwide Distribution - USA (nationwide) Canada and Mexico | Terminated | December 18, 2015 | 800 | 2912 | Not_Computer
Z-0470-2014 | 66826 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | Coulter LH 750 Hematology Analyzer, Part No. 6605632, |  |  | November 08, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling the LH 750 and LH 780 Hematology Systems because the optical degradation may potentially cause a delay in reporting results for nucleated red blood cells (nRBC), white blood cell differential and/or reticulocytes. | Under Investigation by firm | Beckman Coulter sent an customer notification letter dated 11/8/13 to all customers who purchased the LH 750 and LH 780 Hematology Systems. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact Beckman Coulter's Customer Support Center at http://www.beckmancoulter.com/customer/support or call (800) 526-7694 in the US and Canada. Outside the US and Canada, customers are instructed to contact Beckman Coulter Representative.On 2/17/14, Beckman Coulter expanded the recall to inform Coulter LH 750 and Coulter LH 780 Hematology Analyzer customers of the issue related to LH Cleaner and Coulter Clenz. |  | 4155 | 4155 units total (1676 units in US) | Worldwide Distribtion - USA Nationwide including the countries of Algeria, Andorra, Angola, Australia, Austria, Bahrain, Bangladesh, Belgium, Bermuda, Bosnia and Herzegovina, Bulgaria, China, Colombia, Croatia, Czech Republic, Ecuador, Egypt, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Jordan, Korea, Republic of, Kuwait, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Macao, Malaysia, Mayotte, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, Virgin Islands, U.S., and Zimbabwe. | Terminated | December 25, 2014 | 412 | 4155 | Not_Computer
Z-0477-2014 | 66717 | No | HSN | Ankle joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Trabecular Metal Total Ankle Tibial Base Component Sizes 1-6 |  |  | November 12, 2013 | 2013 | Zimmer, Inc. | As part of routine manufacturing process monitoring by Zimmer, it was discovered that the Tibial Base and Talar Components from the Total Ankle system have the potential of containing elevated levels of manufacturing materials residue on the implants. The specific material was found to be mass media utilized during the manufacturing process to provide the uniform surface finish. This mass media | Manufacturing material removal | Zimmer sent an URGENT MEDICAL DEVICE RECALL notification letter dated November 12, 2013 to all consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to locate and remove the affected product from their inventory and notify their Zimmer representataive. The Zimmer representative will remove the recalled product from their facility. For patients that previously had the affected product implanted, consignees were instructed to continue their operative follow up routine. Consignees with questions were instructed to call 1-877-946-2761.For questions regarding this recall call 800-613-6131. |  | 50310 | 390 units (US) 129 units (outside of US) | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Canada, China, Finland, Germany, and Italy. | Terminated | October 15, 2014 | 337 | 50844 | Not_Computer
Z-0474-2014 | 66409 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes Trochanteric Fixation Nail |  |  | July 30, 2013 | 2013 | Synthes USA HQ, Inc. | Certain lots of Trochanteric Fixation Nail were distributed to the field with the incorrect labeling of the length and diameter of the nail on the packaging. | Employee error | Synthes sent an Urgent Notice: Medical Device Recall letter dated July 30, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following actions:Call Syntehs at 1-800-479-6329 to obtain a Return Authorization Number.Complete the Verification Section checking the appropriate box indicating affected product has been located.Return the Verfication Form with the product to:Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.If no affected product found customers were instructed to take the following steps:Complete the attached Verification Section at the end of the letter by checking the appropriate box indicating no affected product has been located.Return the documents to Synthes by fax to 610-251-9005 or Scan/email: FieldAction@synthes.com. Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 12 | 12 | Nationwide Distribution including WA, FL, LA, TX, GA, IN, OH, CA, MI, and PA. | Terminated | September 10, 2015 | 772 | 12 | Not_Computer
Z-0483-2014 | 66908 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Cancellous Bone Screw |  |  | October 01, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field indicating that a 5.5 mm Cancellous Bone Screw 40 mm was identified in packaging associated with a 6.5 mm Cancellous Bone Screw 25 mm. | Labeling mix-ups | Stryker sent notification letters/accountability forms via Fed Ex on October 1, 2013, with return receipt. to branches/agencies and hospital risk management, chief of orthopedics and surgeons.Domestic branch locations were notified by e-mail on 9/16/2013.The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to fax the Product Recall Acknowledgement Form to 855-251-3635, return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, New Jersey, 07430. Customers were instructed to attach the fluorescent organge PRODUCT REMEDIATION sticker to their return, indicating the Produc Remediation # RA 2013-145. Mark the outer box with the words "Product Recall."Customers with questions were instructed to call 201-972-2100.For questions regarding this recall call 201-831-5000. |  | 2232619272 | Catalog Number 2030-6525-1- 4 units; Catalog No: 2029-5540-1-7 units | Nationwide Distribution | Terminated | July 15, 2015 | 652 | 2232619272 | Not_Computer
Z-0478-2014 | 66874 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCOÂż EvolutionÂż steam sterilizer Transfer Carriage. |  |  | October 18, 2013 | 2013 | Steris Corporation | The firm discovered that unlatching can occur when overloaded transfer carriages are pulled or maneuvered incorrectly or with excessive force. | Device Design | Steris sent an Urgent Voluntary Field Correction Notice letter dated October 18, 2013 to their customers. The letter identified the affected product, problem, and actions to be taken. The letter informed customers that a Steris Service Representative will contact affected facility to arrange for Steris to upgrade the Transfer Carriage(s). For questions contact Steris Field Service Dispatch at 1-800-333-8828. |  | 936 | 936 units | Worldwide Distribution - USA Nationwide in the states of: AL, AZ, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MI, MO, MT, NE, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY and the District of Columbia including the countries of: Brazil, Canada, Chile, Guam, Mexico, Saudi Arabia, United Arab Emirates &amp; Venezuela. | Terminated | May 13, 2014 | 207 | 936 | Not_Computer
Z-0480-2014 | 66714 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGenÂż Complete Knee Solutions Stemmed Tibial Component Precoat |  |  | October 31, 2013 | 2013 | Zimmer, Inc. | Two complaints were received in which it was reported that the stem would not lock into the tibia. Evaluation of the returned device found that the stemmed portion of the tibial plate would not hold the stem as the connection was too loose. The taper diameter was found to be oversized by .010 . | Device Design | All distributors will be notified via electronic mail. Hospital risk managers and distributors with affected inventory will also be notified via courier. Surgeons that have implanted this manufacturing lot will be identified and notified by the distributor. Distributors will be sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory, as well as identifying and notifying surgeons that have implanted this manufacturing lot. Distributors will return on-hand affected products to Zimmer and ensure all of their products are accounted for using Attachment 1 provided in the letter. Distributors are to complete and return Attachment 2 indicating which surgeons have implanted the product and to which they have provided notice to. Hospitals that received direct shipments from Zimmer will be sent a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product and returning it to their Zimmer sales representative.Communications outside of the United States will occur approximately two weeks after the United States communications.Zimmer will conduct effectiveness checks in the following way:Distributors/Hospitalso 100% of the FedEx notifications will be accounted foro An inventory certification form will need to be returned from each distributor that certifies that all available affected product from their territory has been returned.o Additional notifications via email and FedEx will be sent to those who have not completed the required certification formo Accounts will be deemed unresponsive after 3 attempts.Further questions please call 800-613-6131. |  | 128 | 128 units | Worldwide Distribution - USA (nationwide) including the states of CA, IN, OH, GA, WI, IA, NE, NJ, PA, NY, MD, NC, AZ, NM, TX, CO, IL, AL, MI, WA, FL and KS, and the countries of China, Japan and Taiwan. | Terminated | April 30, 2015 | 546 | 128 | Not_Computer
Z-0482-2014 | 66834 | No | JOZ | Automated platelet aggregation system. | Hematology | Hematology | 510(k) | VerifyNow IIb/IIIa Test, 10-Test Kit, Catalog No. 85310, |  |  | November 14, 2013 | 2013 | Accumetrics Inc | Accumetrics is recalling the VerifyNow IIb/IIIa 10-Test because it contains incorrect information for the amount of time in which to perform the test following collection of the patient sample. The correct instruction is to perform the test within 15 minutes of collecting the patient sample. | Other | Accumetrics sent an Urgent Medical Device Correction letter dated November 14, 2013 to end users for the VerifyNow IIb/IIIa 10-Test. The letter informs the customers of the problems identified and the actions to be taken. The customers are instructed to call Accumetrics Customer Support at (800) 643-1640 option 2 or email at support@accumetrics.com. Customers are instructed to complete and return the enclosed Customer Account Tracking Form as soon as possible.For questions regarding this recall call 858-643-1600. |  | 69 | 69 units | Nationwide Distribution including NY, PA, DC, CT, IN, NJ, AR, TX, AZ,and NM. | Terminated | December 23, 2013 | 39 | 69 | Not_Computer
Z-0481-2014 | 66978 | No | NJM | Orthodontic plastic bracket. | Dental | Dental | 510(k) | Opal Orthodontics by Ultradent. Avex CX2 Ceramic Orthodontic Brackets |  |  | November 19, 2013 | 2013 | Ultradent Products, Inc. | Ultradent Products Inc. is recalling various lots of Avex CX2 Orthodontic Brackets because the design control process was not complete prior to distribution. | Process control | Ultradent Products notified consignees via email on November 19, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return all unused product. Sales representatives will pick up remaining unplaced brackets or returns will be sent back using pre-paid envelopes back to Ultradent recall coordinator for quarantine. Consignees were asked to complete evaluation forms to be returned by fax to 800-793-5206.For questions regarding this recall call 801-553-4420. |  | 217 | 217 units | Nationwide Distribution including UT, CA, WA, VA, OH, and TN. | Terminated | October 06, 2014 | 321 | 217 | Not_Computer
Z-0479-2014 | 66533 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Aesculap GNI161 Bipolar Foot Control |  |  | September 10, 2013 | 2013 | Aesculap, Inc. | Complaints were received for the GN161 Bipolar Foot Control reporting that the bipolar energy did not stop after release of the foot control. | Device Design | Aesculap sent an Important Recall Notification letter dated September 23, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return affected product and receive replacement product or credit. An Inventory Sheet is attached and must be completed. For questions call 610-984-9265 or 610-984-9414. |  | 85 | 85 | Worldwide Distribution - USA Nationwide and countries of: Austria, Germany, Japan, and Taiwan. | Terminated | October 20, 2014 | 405 | 85 | Not_Computer
Z-0484-2014 | 66790 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Helion S Exam Light or H300 |  |  | February 19, 2013 | 2013 | Trumpf Medical Systems, Inc. | Possible fatigue failure of the plastic joint may occur after an average use of seven years. | Nonconforming Material/Component | TRUMPF sent an Urgent Recall Notice dated February 19, 2013, to all US affected customers. The letter identified the product the problem and the action needed to be taken by the customers. The recall notice instructed customers to verify if they still have these light systems, and if they do have the light system verify if they have a metal or plastic ring at the joint of concern (per the manufacturer's Safety Notice). TRUMPF will directly work with affected customers to replace the affected arm system.For further questions please call ( 888 ) 474-9359. |  | 287 | 287 | US Distribution including the states of AR, CA, CO, KS, MA, MI, MO, NJ, NY, NC, OH, OK, PA, SD, TX and WA. | Open, Classified |  |  | 287 | Not_Computer
Z-0489-2014 | 66970 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | HGE Healthcare Patient Monitor B30, Procare Monitor B20, Procare Monitor B40, and Monitor B40 |  |  | October 11, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to ECG filter settings associated with the B20/B40 V1, B30 and B40 V2 Patient Monitors.The B20, B30 and B40 Patient Monitors set the ECG filter setting incorrectly if the monitor starts with impedance respiration set in the ON position. Instead of the user selected filter, an additional 0.5 Hz high pass filter is used. The | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated October 10, 2013, to all affected consignees. The letter described the safety Issue, Safety Instructions, Affected Product, Details, Product Correction and Contact Information (Technical Support at 1-800-558-7044).Consignees were asked to retain a copy of the instructions with the user manual.For questions regarding this recall call 262-513-4122. |  | 16061 | 16,061 (343 US, 15,718 OUS) | Worldwide Distribution - USA (nationwide) to CA, CT, FL, GA, MO, NC, NJ, PA, TX, WI and Internationally to:FRANCE, PORTUGAL, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BARBADOS, BELGIUM, BOLIVIA, BOSNIA 7 HERZEGOVINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, NEW ZEALAND, CHILE, CHINA, COLOMBIA, CONGO, COSTA RICA, COTE DLVOIRE, CROATIA, CYPRUS,. CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND, GABON, GERMANY, GREECE, GUATEMALA, HONDURAS, HUNGARIA, INDIA, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MEXICO, MOLDOVA, MOROCCO, MYAMMARA, NAMBIA, NAPAL, NETHERLANDS, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PHILIPPINES, POLAND, PORTUGAL ROMANIA, RUSSIA, SENEGAL SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SIR LANKA, SWEDEN, SWITZERLAND, THAILAND, THE NETHERLANDS, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES , UNITED KINGDOM, URUGUAY, UZEBEKISTAN, VENEZUELA, and VIETNAM. | Terminated | February 03, 2016 | 845 | 16061 | Not_Computer
Z-0472-2014 | 66638 | No | BSF | Carbon dioxide absorber. | Anesthesiology | Anesthesiology | 510(K) Exempt | Multi Absorber Original, Disposable |  |  | September 19, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to air leakage associated with the CO2 Multi Absorber. | Device Design | GE issued an Urgent Medical Device Correction letter dated September 19, 2013, to all affected customers via Fed Ex overnight mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to not use the affected product, isolate all affected product and return the attached form via fax to 800-535-7923. Once the information is received, customers would be contacted with an RMA and replacement information. If product was distributed further the notice should be forwarded to those customers. For questions customers should call 1-800-345-2700, option 2 followed by ooption 2. For foreign customers 1-800-932-0760, option 2.For questions regarding this recall call 262-513-4122. GE issued a press release on 12/30/2013. |  | 444115 | US: 59,721 boxes, 358,326 canisters; OUS: 3,724 boxes, 22,344 canisters | Worldwide Distribution - USA (nationwide) Puerto Rico and Internationally to Canada, Costa Rica, Mexico, and Venezuela. | Terminated | March 10, 2015 | 537 | 444115 | Not_Computer
Z-0512-2014 | 66888 | No | MJN | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Stryker brand SUPER COMPLIANT 7MM X 15MM, Catalog Number: |  |  | November 04, 2013 | 2013 | Stryker Neurovascular | Products were shipped labeled with "use by" dates that exceed currently available shelf-life data. All products remain within the earliest expiration date of April 2015. | Software Manufacturing/Software Deployment | Stryker Neurovascular sent an Urgent Medical Device Recall letter dated November 4, 2013, all affected users. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and segregate the affected products. Customers with questions were instructed to contact via NeuroComplaints@stryker.com. Or Fax the response form to 1-800-876-4355.For questions regarding this recall call 510-413-2500. |  |  |  | Worldwide Distribution - USa (nationwide) and Internationally to:Austria, Belgium, Czech Rep, Germany, France, Great Britain, Hungary, Italy, Lithuania, Netherlands, Norway, Poland, Slovakia, Spain, Turkey, Australia, Canada, Chile, Hong Kong, India, Korea, Malaysia, Russia and Japan | Terminated | August 27, 2014 | 296 |  | Not_Computer
Z-0493-2014 | 66967 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | G7 Acetabular Screw 6.5mmx35mm |  |  | November 26, 0013 | 13 | Biomet, Inc. | Investigation identified that the outer label, inner label, and patient label all have the incorrect diameter size of 10.5mm. The screw is actually a 6.5mm diameter and correctly corresponds to the items number on the label. | Labeling Change Control | Beginning November 25 Urgent recall electronic notices were sent by Biomet UK to all distributors with hospitals outside the USA notified by ground mail November 27,2013. The notification identified the product and reason for the recall, asking for immediate location and discontinued use of the product and return to Biomet UK, Ltd. Questions can be addressed to Biomet USA at 01 574-372-1570 M-F 8 am to 5 pm. |  | 532 | 532 in USA all under Biomet USA control | Worldwide Distribution - USA Nationwide and in the countries of Europe, Italy, UK, Australia, Japan. all product in USA is under Biomet control. | Terminated | September 29, 2014 | 730792 | 532 | Not_Computer
Z-0492-2014 | 66536 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Aesculap Columbus Revision CoCr 15mm distal and posterior spacers, sizes F5-F7 |  |  | September 19, 2013 | 2013 | Aesculap, Inc. | The Columbus Revision CoCr 15 mm distal and posterior implant spacers, sizes F5-F7, were not compatible. | Component change control | Domestic consignees were contacted by phone and a recall notification letters was also sent to these Users on/about September 26, 2013. |  | 24 | 24 | Distributed in FL and MD. | Terminated | December 22, 2014 | 459 | 24 | Not_Computer
Z-0518-2014 | 66859 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Kimberly-Clark Single Shot Epidural Pain Management Trays containing the BD LOR Syringe (BD Code 405198) |  |  | October 24, 2013 | 2013 | Kimberly-Clark Corporation | Kimberly-Clark Single Shot Epidural Pain Management Trays contain the BD 7 mL Epilor plastic Luer-Lok LOR Syringes under recall by Becton Dickinson (BD) because the product may stall or stick when traveling within the barrel of the syringe. | Component design/selection | Kimberly Clark Corporation sent a Product Advisory Notice letter to affected consignees. The letter identified the affected product, problem and actions to be taken. For questions call 770-587-8393 or via email |  | 389300 | 389,300 units | U.S. Nationwide Distribution | Open, Classified |  |  | 389300 | Not_Computer
Z-0516-2014 | 66894 | No | GGN | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(K) Exempt | HemosIL High Abnormal Control 3 UNASSAYED |  |  | November 15, 2013 | 2013 | Instrumentation Laboratory Co. | HemosIL Low, Abnormal, Normal Controls UNASSAYED (no-value assigned) lacks 510(k) approval | No Marketing Application | Instrumentation Laboratory issued Important Correction Notice via USPS to US customers dated November 2013 advising users in the US that unassayed controls are not acceptable for use in their unassayed (non-value assigned) format. IL will no longer sell or distribute controls without assayed values for the US market as of November 15, 2013, and are converting US customers to an assayed control format. Additional contact with customers via phone calls and a reminder email.An enclosed response tracking form provided to be returned. For technical questions, contact the IL Technical Support Center at 1-800-678-0710, Option # 6 (ACL Coagulation Products).For replacement product if value assignments are not available for your product lot, please contact your IL Customer Service Representative at 1-800-955-9525, Option # 6.For questions regarding this recall call 781-861-4467. |  | 84561 | 84561 kits | Nationwide Distribution | Terminated | March 25, 2014 | 130 | 171384 | Not_Computer
Z-0520-2014 | 66918 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Mega Needle Driver, 8MM assembly of the da Vinci S IS1200 and da Vinci Si, IS2000. |  |  | November 15, 2013 | 2013 | Intuitive Surgical, Inc. | Firm has become aware of the potential for detachment of the jaw insert for Large Needle Drivers and Mega Needle Drivers. | Software design (manufacturing process) | Intuitive Surgical sent a Field Safety Notice dated November 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure all appropriate personnel are fully informed of the Notification. Discontinue use of affected product, Use the attached instructions to identify and segregate affected product. Return affected product to Intuitive Surgical following the established Returned Materials Authorization (RMA) process. Complete the Acknowledgement Form and return to Intuitive Surgical. Customers with questions were instructed to contact Intuitive Surgical Customer Service. Customers were asked to retain a copy of this letter and the acknowledgement form for their files.For questions regarding this recall call 800-876-1310, option 3.Japan 0120-56-5635 or 03-5575-1362South Korea 02-3271-3200Europe, Middle East, Asia and Africa +800 0821 2020 +41 21 821 2020For questions regarding this recall call 408-523-2100.Urgent Medical Device Recall notification letters were sent to all affected sites on November 15, 2013. |  | 3540 | 15, 236 UPDATED 12/19/13 to 11,235 | Worldwide Distribution - USA (nationwide and Puerto Rico) and Internationally to Finland, Belgium, Netherlands, UK, Czech Republic, Sweden, Germany, Austria, Switzerland, Norway, Denmark, South Korea, France, Ireland, and Slovakia. | Terminated | September 04, 2014 | 293 | 95754 | Not_Computer
Z-0513-2014 | 66626 | No | LFW | Tobramycin test system. | Toxicology | Toxicology | 510(k) | Thermo Fisher Scientific/Microgenics brand CEDIA Tobramycin II Assay, Catalog Number 100018, Product is manufactured and distributed by |  |  | October 16, 2013 | 2013 | Microgenics Corp | Microgenics Corp., part of Thermo Fisher Scientific, is recalling CEDIA Tobramycin II Assay (Lot Number: 60169764) due to calibration stability problem relating to two complaints received. | Under Investigation by firm | Microgenics Corp., part of Thermo Fisher Scientific sent an Urgent Medical Device Correction letter dated Octobere 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and discard. Customers with questions were instructed to contact mgc@thermofisher.com. For question regarding this recall call 1-800-232-3342. |  | 370 | 370 units | Worldwide Distribution - USA including WI, UT, ND, MT, WA, CA, OK and Internationally to Australia, Canada, and Germany. | Terminated | February 03, 2014 | 110 | 370 | Not_Computer
Z-0519-2014 | 66521 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access Immunoassay Systems Reaction Vessels. |  |  | October 02, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is initiating a recall for the Access Immunoassay Systems Reaction Vessels, including the Acess/Access 2 Immunoassay Systems, SYNCRON LXi 725 Clinical System, and UniCel DxC 600i SYNCHRON Access Clinical System because Beckman Coulter has observed an intermittent increase in Access System "Wash Carousel Motion Errors" with specific lots of Access SystemReaction Vessels (RVs). I | Pending | Beckman Coulter sent an Urgent Product Correction (PCA) letter dated October 1, 2013, along with a customer response form to all affected customers on October 2, 2013. The customer letter informed customers of the recall with the issue, impact, action (If your Access Immunoassay System experiences Wash Carousel Motion Errors that cannot be corrected with normal troubleshooting measures as outlined in the Access 2 Reference Manual section 6.3, contact Beckman Coulter Customer Technical Support.If you do not have Access System RVs with the lot numbers listed above, or if your Access Immunoassay System is not experiencing Wash Carousel Motion Errors, no additional action is required.), resolution, and contact information, website: http://www.beckmancoulter.com. Customers in the United States and Canada with questions were instructed to call 1-800-854-3633, outside the United States and Canada, customers were instructed to contact their local Beckman Coulter representataive.For questions regarding this recall call 714-961-4941. |  | 16154 | 16,154 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Japan, India, Hong Kong, Guatemala, China, Brazil, Brunei Darussalam, Costa Rica, El Salvador, JordanMalaysia, Mexico, Panama, Philippines, South Africa, Taiwan, Thailand, Trinidad and Tobago, United Arab Emirates, Venezuela, and Viet Nam. | Terminated | October 23, 2014 | 386 | 16154 | Not_Computer
Z-0517-2014 | 65051 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension(R) TACR Flex(R) reagent cartridge (DF107) |  |  | April 25, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens has confirmed that the TACR method may demonstrate reduced on-board stability which may result in imprecise and inaccurate QC and patient results. | Component design/selection | Siemens sent an Urgent Medical Device Correction letter dated April 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A recall notification letter was sent to consignees in April 2013. Customers using the affected product should load Flex reagent cartridges on board immediately prior to use and remove the Flex after 8 hours. Do not load more Flexes than the lab will use in 8 hours and do not pre-hydrate TACR Flexes. Customers were asked to complete the attached Effectiveness Check form and fax it to the Siemens Technical Solutions Center at 302-631-8467. Customers were asked to forward the communication to anyone to whom may have distributed the product. Customers with questions were instructed to contact the Siemens Technical Solutions Center or their local Siemens technial support representative. For questions regarding this recall call 800-441-9250. |  | 13751 | 13,751 | Nationwide Distribution including AK, AL, AR, AZ, CA, DC, FL, GA, IA, IL, LA, MD, MI, MO, MS, NC, NH, NJ, NM, OH, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, and WI. | Terminated | November 12, 2014 | 566 | 13751 | Not_Computer
Z-0490-2014 | 66735 | [6 empty] | Nephros Dual Stage Ultra Filter (DSU) Brochure |  |  | October 28, 2013 | 2013 | Nephros Inc | Promotional materials for non-medical water filtration products were determined to promote claims which constitute marketing the product as a medical device. | Labeling False and Misleading | Nephros Inc., sent an Urgent Medical Device Recall letter dated October 29, 2013 to all affected customers via UPS 2nd Day Air. The letter identified the affected product, problem and actions to be taken. Customer are asked to provide a customer information list, immediately remove and discard all affected product labeling materials and any independent labeling materials they have created for the affected products. Customers are instructed to complete and sign the attached Response Form to acknowledge receipt, understanding and compliance with the instruction. For questions contact the Nephros representative at 201-343-5202 x 100.The letter was also posted at : http://www.nephros.com/download/news/DSU_Recall_Letters.pdf |  | 3367 | 3,367 untis with affected labeling | USA Nationwide Distribution | Terminated | July 06, 2016 | 982 | 3367 | Not_Computer
Z-0527-2014 | 66862 | No | ECQ | Dental hand instrument. | Dental | Dental | 510(K) Exempt | Broncho Flexible Fiberscope, models: 11001BN1, 11002BD1, 11004BC1, and 11009BC1. |  |  | November 13, 2013 | 2013 | Karl Storz Endoscopy America Inc | KARL STORZ Endoscopy America, Inc. is recalling the Flexible Broncho Fiberscope or a Flexible Intubation Fiberscope with a suction port because the material used in the suction channel portion was thermoplastic polyurethane (TPU), which is not listed in STERRAD systems claims for interior endoscope lumens. | Under Investigation by firm | The firm, KARL STORZ, sent a "A SAFETY ALERT-ATTENTION NEEDED" letter dated November 8, 2013 to all their customers on November 13, 2013. The letter describes the product, problem and actions to be taken.The letter informs the customers that the endoscopes incorporated a lumen material that is only approved for use as an external surface material and not included in the STERRAD NX and 100NX FDA-cleared claims for use in flexible endoscope lumens. The customers were instructed to follow the recommended processing methods included in the letter.Customers with questions about the letter are instructed to contact the undersigned and the Technical Services at (800) 421-0837, ext 5350. |  | 141 | 141 units | Nationwide Distribution. | Terminated | March 27, 2014 | 134 | 1657 | Not_Computer
Z-0525-2014 | 66496 | No | LTK | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | VITROS CA 125 II Calibrators |  |  | October 09, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics issued a Product Correction Notification for VITROS Immunodiagnostic Products CA 125 II Reagent Pack and Calibrators due to a positive shift in patient results. | Under Investigation by firm | The firm, Ortho Clinical Diagnostics, sent a "IMPORTANT PRODUCT CORRECTION NOTIFICATION" letter dated October 9, 2013 to their clients/customers informing them of a shift in results that may be observed when performing patient sample correlation tests between VITROS CA 125 II Calibrators and Reagent Packs. The letter described the product, problem and actions to be taken. The customers were informed that it is acceptable to continue using any remaining inventory until its expiration date providing that quality control results are within acceptable limits (Note: Lots 1220 and 1230 are expired). The customers were also instructed to consider the need to review patient results using the affected lot(s). If you suspect that a previously reported result may have been affected, provide this information to your Laboratory Medical Director and the requesting physician or health care provider so that appropriate actions may be taken for the patients; Complete and return the attached Confirmation of Receipt form no later than October 18, 2013 via FAX TO: Ortho Clinical Diagnostics at 1-888-557-3759 or 1-585-453-4110, and Forward the information in this notification if you have distributed this product outside of your facility.If you have any additional questions, please contact Customer Technical Services at 1-800-421-3311. |  | 601 | 601 units | Worldwide Distribution: US (nationwide) including states of: AK, CA, CT, CO, DC, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, VA, VT, WA, WI, WV, WY, and Puerto Rico, and countries of: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy &amp; Spain. | Open, Classified |  |  | 601 | Not_Computer
Z-0485-2014 | 66886 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle Blood Glucose Test Strips; |  |  | November 18, 2013 | 2013 | Abbott Diabetes Care, Inc. | Certain lots of FreeStyle and FreeStyle Lite Blood Glucose Test Strips produce erroneously low blood glucose results when using FreeStyle Blood Glucose Meters, FreeStyle Flash Blood Glucose Meters and the FreeStyle blood glucose meter built into the OmniPod system. Erroneously low results that are not recognized may pose significant risks to your health. | Other | Abbott Diabetes Care sent an Urgent Product Recall letter in November 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers who were affected were instructed to take the following steps:Call Abbott Diabetes Care Customers Service immediately at 1-888-736-9869 for replacement of the affected product. While waiting for replacement strips to arrive, use an alternate method to measure their blood glucose.If only test strips available to customers were from affected lots, customers should not stop testing their blood glucose. Customers were instructed to dispose of the test strips in accordance with their local and state disposal regulations. Customers with questions were instructed to call Abbott Diabetes Care Customers Service at 1-888-736-9869.For questions regarding this recall call 510-239-2775.Abbott issued a press release at the request of CDRH on november 27, 2013. |  | 141219 | 141,219 cartons for human use - all lot numbers | Worldwide Distribution - USA (nationwide and Puerto Rico) and Internationally to Australia, Austria, Belgium, Denmark, Egypt, Finland, France, Germany, Jamaica, Ireland, Israel, United Kingdom, Luxembourg, Netherlands, Norway, Sweden, Switzerland, UAE, and Canada.**Center Recommended Depth - Consumers/User** | Terminated | August 03, 2015 | 623 | 141219 | Not_Computer
Z-0533-2014 | 66885 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Integra Disposable, Convenience Kit Single Shot Epidural Tray Reorder Number 340269 Rx Only |  |  | October 28, 2013 | 2013 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences Corporation is recalling certain lots of Integra Kits for Epidural Anesthesia Administration as a result of a Product Advisory Notice from BD Medical for the LOR syringes indicating the plunger of the syringe may stall or stick when travelling within the syringe. | Component design/selection | BD Medical sent a Product Advisory Notice letter dated August 23, 2013 to Iaffected consignees by via traceable courier. The letter identified the affected product, problem and actions to be taken. Distributor have been asked to forward copies of the notification to their customers. For questions call BD Customer Service at 1-888-237-2762. |  | 359250 | 359,250 | USA Nationwide Distribution and one foreign consignee in Mursalat. | Terminated | January 29, 2014 | 93 | 359250 | Not_Computer
Z-0530-2014 | 65477 | No | JXB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | remel RapID Yeast Plus System, 20 tests, Ref: R8311007, IVD, containing a 15-ml bottle of RapID Yeast Plus Reagent A and a 10-ml bottle of RapID Yeast Plus Reagent B. |  |  | April 04, 2013 | 2013 | Remel Inc | The two reagents supplied in the RapID Yeast Plus System kit (Reagent A and Reagent B) are mislabeled. The labels were reversed. | Labeling mix-ups | Thermo Fisher Scientific sent a Medical Device Recall letter dated April 3, 2013 to affected customers. The letter identified the affected product, problem and actions to be taken. The recalling firm instructed customer to review their inventory for affected lot and discard the product. A Medical Device Recall Return Response was enclosed for the customer to indicate they understood the instructions and to report the amount of product they discarded. For questions call Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 150 | 150 kits | Worldwide Distribution - USA Nationwide in the states of AL, AZ, CA, CT, FL, ID, IN, KY, MA, MI, MS, NJ, OH, PA, TX, WI and the countries of Canada, UK, and Japan. | Terminated | March 12, 2014 | 342 | 150 | Not_Computer
Z-0524-2014 | 66992 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | InstaTrakÂż 3500 Plus, ENTrakTM Plus, InstaTrakÂż 3500, ENTrakTM Computer-Assisted Navigation Systems (USA), |  |  | November 26, 2013 | 2013 | GE OEC Medical Systems, Inc | GE OEC Medical Systems Inc is recalling InstaTrakÂż 3500 Plus, ENTrakTM Plus, InstaTrakÂż 3500, ENTrakTM Computer-Assisted Navigation Systems (USA) due to several potential safety issues related to the use of these products. | Software Design Change | GE Healthcare sent an An Urgent Medical Device Correction letter dated November 26, 2013, to all affected customers. The letter identified the product, the problem, and provided instructions for immediate mitigation of the issue. No product is being returned. The affected systems in the field will be corrected.Customers with questions were instructed to contact the service team at 800-874-7378.For questions regarding this recall call 801-536-4952. |  | 654 | 654 | Nationwide Distribution | Terminated | April 17, 2017 | 1238 | 654 | Not_Computer
Z-0523-2014 | 66964 | No | HTD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | outer box: Ref 906768, CurvTek Eye Needle 22 MM X-large Stainless Steel, Single Use Gamma Sterilized, Rx |  |  | November 26, 2013 | 2013 | Biomet, Inc. | Inner label of PN: 906768 Lot: 055450 CurvTek Eye Needle 22mm X-Large states incorrectly PN: 906760 7mm Medium. | Labeling mix-ups | Biomet sent an Urgent Medical Device Recall Notice dated November 26, 2013, to all affected customers. The notice identifies the product, problem, and actions to be taken by the customers. The letter included a Customer Response Form for customers to complete and return to the firm by FAX 574-372-1683. Customers with questions were instructed to call 574-372-1570.For questions regarding this recall call 574-372-1570. |  | 21 | 21 | Nationwide Distribution including MI, IA, IL, IN, and TN. | Terminated | July 07, 2015 | 588 | 21 | Not_Computer
Z-0528-2014 | 65734 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Test Discs, CTO425B Ciprofloxacin CIP 5 mcg, 50 discs per cartridge, 5 cartridges per box. The firm name on the box label is Oxoid Ltd., Basingstoke, UK. |  |  | November 21, 2012 | 2012 | Remel Inc | Boxes labeled as CIP5 may contain cartridges of AMP5. | Mixed-up of materials/components | Thermo Fisher sent a Field Safety Notice dated November 13, 2012, to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. Customers were instructed to check stocks of the affected product if necessary and contact Customer Service or their local Oxoid supplier who will issue a replacement. The notice should be passed on to all who need to be aware within their organization. Customers were also instructed to discard any product labeled as AMP5. For questions, customers were instructed to call +44 (0)1256 841144.discard any product labeled AMP5.For questions regarding this recall call 913-895-4185. |  | 300 | 60/5/50-disc cartridge boxes | Nationwide Distribution including AL, CA, CO, FL, IA, KS, LA, MA, MS, NE, NH, NJ, NY, OK, OR, RI, TN, TX, VA, WA WI, and Puerto Rico. | Terminated | March 12, 2014 | 476 | 300 | Not_Computer
Z-0487-2014 | 66736 | [6 empty] | Nephros SafeSpout |  |  | October 28, 2013 | 2013 | Nephros Inc | Promotional materials for non-medical water filtration products were determined to promote claims which constitute marketing the product as a medical device. | Device Design | Nephros sent an Urgent Medical Device Recall letter dated October 28, 2013, with a return response form to all affected customers via UPS 2nd Day Air. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to cease shipment of the affected product immediately. Obtain an RGA number to return the new and used filters for credit. In addition, customers were instructed to immediately remove and discard all copies of the labeling materials that pertain to the affected product. Remove all electronic copies of these materials from any publicly accessible locations (e.g. website). Customers with questions were instructed to contact their Nephros representative or call 201-343-5202, ext 100.For questions regarding this recall call 201-343-5202.The letter was also posted at:http://www.nephros.com/download/news/SafeSpout_SafeShower_Recall_Letters.pdfNephros sent a new recall letter dated 2/20/204 as they determined they have additional customers. |  | 3742 | 3742 units | Nationwide Distribution | Terminated | June 28, 2017 | 1339 | 8946 | Not_Computer
Z-0522-2014 | 65915 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | KayserBetten model IDA beds |  |  | April 11, 2013 | 2013 | mobility unlimited inc | Some joints on the doors of a limited number of IDA beds may loosen or separate over time which can cause the doors to drop outward and cause injury. | Device Design | Mobility Unlimited sent a notification letters dated April 12, 2013 to affected customers The letter identified the affected product, problem and actions to be taken. The letter informed customers that a Retrofit Kits along with illustrated instructions will be sent. For questions call 1-800-574-7880. |  | 33 | 33 beds | Worldwide Distribution - USA Nationwide in the states of AL, AZ, CA, FL, GA, IL, MN, MO, NY, PA, TN, TX, VA, and Canada. | Terminated | October 08, 2014 | 545 | 33 | Not_Computer
Z-0539-2014 | 66989 | No | LNX |  |  | General Hospital | 510(k) | ABACUS TPN Calculation Software, Baxa: 8300-0045 ABACUS CE (Calculator Edition), 8300-0046 ABACUS SE (Single-Workstation Edition), 8300-0047 ABACUS ME (Multi-Workstation Edition) |  |  | December 03, 2013 | 2013 | Baxter Corporation Englewood | Baxter Healthcare Corporation is issuing a voluntary correction for ABACUS Total parenteral Nutrition (TPN) Calculation Software versions 2.0 through 3.0 due to the possibility that a patient could receive another patient's TPN order as a result of data corruption. | Unknown/Undetermined by firm | Baxter sent an Urgent Device Correction letter dated December 4, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check the patient order against the final bag label as part of the TPN acceptance procedure to ensure that the patient's TPN order is correct. Customers will be contacted by Baxter to arrange for installation of a software upgrade specific to their compounder's configuation. Customers were asked to acknowledge the receipt of this notification by completing the attached Customer Reply Form and return it to Baxter by fax to 1-224-270-5457 or scanning and emailing it to fca@baxter.com. If the product was distributed to other facilities customers were instructed to notify end-user customers in accordance with their customary procedures. Customers with questions were instructed to call 1-800-678-2292.For questions regarding this recall call 800-422-9837. |  | 2211 | 2,211 | Nationwide Distribution | Terminated | October 16, 2015 | 682 | 2211 | Computer
Z-0535-2014 | 66944 | No | ELW | Impression material. | Dental | Dental | 510(k) | 3M ESPE Imprint II Regular Body , Catalog Number 9379, and Paradigm Regular Body VPS, Catalog Number 5315, Impression Material that are intended to reproduce the structure of a patients teeth. |  |  | November 21, 2013 | 2013 | 3M Company - Health Care Business | Specific lots of Imprint II and Paradigm impression material refill packs are being recalled because the moisture displacement feature in both is greatly reduced which may impact the quality of the dental impression. | Employee error | 3M ESPE sent an Voluntary Recall Notification Letter dated November 21, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers were instructed to:1) Review your inventory and identify any products from these lot numbers2) Quarantine products from these lots and do not further distribute3) If you have shipped any products from these lots to your customers, inform them of this voluntary recall using the attached Dear Doctor Letter and Inventory Checklist4) To return and replace product from these lots, contact 3M ESPE Dental ProductsCustomer Service at 1-800-237-1650If you have any questions, please contact the above number for assistance. Once again, we sincerely apologize for any inconvenience this may cause you. |  | 1977 | 1977 kits | Worldwide Distribution - USA (nationwide) and the countries of Canada, France, Germany, Hong Kong, Italy, Japan, Korea, Mexico, United Arab and Emirates. | Terminated | June 02, 2014 | 193 | 1977 | Not_Computer
Z-0536-2014 | 66963 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MagNA Pure 96 System |  |  | November 25, 0013 | 13 | Roche Diagnostics Operations, Inc. | This issue is only relevant for customers importing xml files into the MagNA Pure 96 IVD instrument software for sample identification. There is a potential for sample mismatch when using the MagNA Pure 96 IVD instrument (catalog number 06541089001). Importing an order file in xml format to the MagNA Pure 96 IVD instrument software may result in a wrong sequence of samples in the graphical overvi | Software design | Roche Diagnostics sent an URGENT Medical Device Correction Notice dated November 25, 2013, to all affected customers. The letter identifies the product, problem, and actions to be taken by the customers.Customers were advised to discontinue use of the Pipetting Report and to use the sample list file instead. The MagNA Pure 96 Instrument Operators Guide is being updated with an addendum addressing this issue. The letter included a Customer Response Form for customers to complete and return to the firm by FAX : 1-888-670-4135. Please contact the Roche Support Network Customer Support Center at 1-800-526-1247, 24 hours a day, seven days a week, if you have questions about the information contained in this UMDC. |  | 27 | 27 | US Distribution including the states of AL, AZ, CA, CT, GA, NY, NC, NJ, OH, TX, VA and WA. | Terminated | September 04, 2014 | 730768 | 27 | Computer
Z-0537-2014 | 65476 | No | JTO | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Streptex* kit, Ref R30950501 ZL50, Rapid latex test for Streptococcal grouping, IVD. The kit contains reagent groups A, B, C, D, F, G, a control, and an extraction enzyme. The firm name on the label is Remel Europe Ltd., Dartford, England. |  |  | May 02, 2013 | 2013 | Remel Inc | The product may give weak or false negative test results. | Process control | ThermoFisher sent a Medical Device Recall letter dated April 26, 2013, or May 2, 2013, to all affected customers via first class mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify any personnel who need to be notified of the potential for a false negative test result, review their inventory for affected product lots and discard the affected product. Customers were also instructed to return the attached Medical Device Recall Return Response acknowledging the disposal of the affected product. Customers with questions were instructed to contact the Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International).For questions regarding this recall call 913-888-0939. |  | 768 | 768 kits | Nationwide Distribution | Terminated | March 12, 2014 | 314 | 877 | Not_Computer
Z-0546-2014 | 66589 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | Continuous Peripheral Nerve Block Catheter Kit and Set |  |  | October 09, 2013 | 2013 | Arrow International Inc | A labeling inconsistency was discovered in which the lidstock states that the kit contains a stimulator cable with 1.5mm socket, however, the kit contains a stimulator cable with 1.5mm pin. | Labeling mix-ups | The firm, Arrow International, Inc., sent an "Urgent Medical Device Recall Notification" letter dated October 14, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue use and quarantine any of listed affected products/lots you have in stock; return affected products, customer service representative will contact you with RGA# and instructions for the return of the product, and complete enclosed Recall Acknowledgement form and fax to 1-800-343-2935, Attn: Customer Service, even if you have not affected product.If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-800-343-2935. |  | 121 | 121 | Worldwide distribution: US (nationwide) including states of: MA, LA, CA, GA, KS, TX, PA, NY, MO, WI, NJ, MS, CT, ID, AL, CO, and PR; and Singapore. | Terminated | September 21, 2015 | 712 | 121 | Not_Computer
Z-0541-2014 | 66363 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Dimension Vista Homocysteine Flex Reagent Cartridge |  |  | September 18, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Dimension Vista HCYS lot 13064MA shows erratic recovery of QC and patient samples. | Unknown/Undetermined by firm | Siemens sent a Urgent Medical Device Recall letter dated September 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Actions to be taken by Customer:Please discontinue use and discard any remaining inventory of Dimension VistaÂż HCYS FlexÂż reagent cartridge (Catalog# K7044) lot 13064MA.Please complete the attached form and fax it to (302) 631-8467 to facilitate your no charge replacement needs.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center- Technical Solutions or your local Siemens technical support representative. For further questions please call (800) 441-9250. |  | 302 | 302 | Worldwide Distribution - US Distribution including the states of OH, MD, CO, NJ, IA, NC, VA, WI and DC., and the countries of Austria, Denmark, France, Italy, Netherlands and Norway. | Terminated | October 01, 2014 | 378 | 302 | Not_Computer
Z-0540-2014 | 66987 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | MedfusionÂż 4000 Syringe Infusion Pumps |  |  | November 22, 2013 | 2013 | Smiths Medical ASD, Inc. | Smiths Medical is conducting a recall involving all MedfusionÂż 4000 Syringe Infusion Pumps.Smiths Medical has identified an issue with electrical shorting at the AC power receptacle. Shorting can occur when saline solution or other conductive fluids (e.g., potassium chloride) enter the AC power receptacle. When this happens, sparking and/ or flames, accompanied by smoke, may appear at the AC po | Device Design | Smiths Medical sent an "Urgent Medical Device Correction" letter dated November 22, 2013, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customers.All customers will be provided with instructions on inspection, cleaning, and prevention regarding this issue, for the continued safe and effective use of these pumps. Additionally, customers will be instructed to:1. Return damaged pumps for servicing or damaged AC power cords for replacement to Smiths Medical.2. Return the completed Confirmation Form to acknowledge receipt and understanding of the information provided in the Urgent Medical Device Correction letter, via fax or email, to Smiths Medical for processing."The Instructions for preventing this issue", "Mitigations", Conclusions", " Action to be taken by the User", and "Transmission of this Urgent Medical Device Correction". The consignees were also asked to complete the "Urgent Medical Device confirmation Form" and fax it to 1-800-237-8033 or by sending an electronic copy via e-mail to medfusion4000nov2013@smiths-medical.com.For questions regarding this information, consignees should contact their local Sales Representative, Smiths Medical's Customer Service Department at 1-800-258-5361, or via e-mail at medfusion4000nov2013@smiths-medical.com.UPDATE: Consignees were sent a follow-up letter dated 13 January 2014 from the one dated November 22, 2013. The letter was addressed to Risk / Safety Managers, Biomedical Professionals, Clinicians who oversee the use of Medfusion Pumps, Distributors, and other users of these devices. The "Urgent Medical Device Recall" letter included Advise on Action to be Taken by the User. This included customers to have their Pumps undergo a free correction. The correction consists of the installation of a new AC power Receptacle seal. Smith Medical provided options for completing the correction process which included: 1) Consignees were a |  | 9560 | 9560 (8601 US, 959 OUS) | Worldwide Distribution - USA (nationwide ) including the states of : AZ, NY, TN, MS, AR, WV, VA, VT, PA, DE, DC, OH, CO, NH, MI, CA, FL, CT, WA, MD, NM, RI, MA, TX, NC, IL, KY, WI and MI., and the countries of CANADA, SAUDI ARABIA and AUSTRALIA. | Terminated | March 25, 2015 | 488 | 9560 | Not_Computer
Z-0545-2014 | 66983 | No | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | SEAC Advanced Dental System, Model A591561 |  |  | November 05, 2013 | 2013 | ASI Medical, Inc. | ASI Medical is initiating a field correction for the SEAC Advanced Dental System because the miniature plastic air tank can potentially rupture. | Device Design | ASI Medical sent an Field Correction Action and Safety Notice letter dated November 5, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.REQUIRED FIELD CORRECTIONOur field correction will consist of sending you a warning label to be applied to the back service access panel of the cart and an ASME safety pop off valve that will open and exhaust if pressure should reach over 100 psi. This safety pop off valve is designed be easily installed in line between the air compressor and air tank in your system. We are requesting your cooperation in making this field correction. |  | 337 | 337 | Worldwide Distribution - USA (nationwide) including the countries of Austria, Aruba, Saudi Arabia, Canada, Germany, Spain and West Indies. | Terminated | August 11, 2014 | 279 | 337 | Not_Computer
Z-0544-2014 | 66698 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PedFuse Reset Screw; Model: 01-80115-55 |  |  | October 18, 2013 | 2013 | SpineFrontier, Inc. | During routine servicing, it was discovered that one 8mmx55mm Reset screw was assembled with an 8mm tulip head and a 9mm screw. The implant should haveconsisted of an 8mm tulip head and an 8mm screw. A recall was initiated to remove all remaining screws of the same lot from the field to investigate the nonconformity. | Software design (manufacturing process) | SpineFrontier sent an Voluntary Recall Notification dated November 5,2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.All affected implants are to be returned to SpineFrontier for inspection. SpineFrontier is requiring confirmation that you have received this notice and that you will initiate the immediate return of the implants listed above.IMMEDIATE RESPONSE REQUIRED.Please feel free to contact us at 978-232-3990 Monday through Friday8:30AM to 5:30PM EST with any questions |  | 35 | 35 Devices | US Distribution to VA and TX. | Open, Classified |  |  | 35 | Not_Computer
Z-0543-2014 | 66844 | No | MVN | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ClearSight(TM) Toric 1 day (ocufilcon D), 30 daily disposable contact lenses for astigmatism, STERILE, Rx only --- CooperVision Hamble, SO31 4RF, UK - Scottsville, NY 14546 USA - Juana Diaz, PR 00795 USA |  |  | November 20, 2013 | 2013 | CooperVision Inc. | Lots were labeled with incorrect expiration dates on secondary package (carton). The primary package (blister) is marked with the correct expiration date. | Incorrect or no expiration date | CooperVision sent an Urgent Medical Device Recall Letters dated November 20, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Our records indicate you received some of the affected product. Enclosed you will find a list of affected carton lot numbers, with the correct expiration date, that were shipped to your account." Please immediately examine your inventory," Stop any further distribution of carton(s) subject to the recall," Quarantine carton(s) with incorrect expiration date, and" CooperVision will make arrangements to remove affected, quarantine product.As soon as possible, please complete the attached response form and fax it to 1.866.484.1048 or scan and email to info@coopervision.com. If you have product to return please contact customer service for instructions and to schedule pick-up and removal of affected, quarantined product. Your account will be credited once the product is returned and processed.We appreciate your assistance and apologize for any inconvenience you or your patients experienced by this removal. If you have any further questions, please feel free to contact us at 800-341-2020. |  | 45 | 45 cartons | Worldwide Distribution - USA (nationwide) and the countries of Aruba, Canada, Colombia and Trinidad. | Terminated | July 22, 2015 | 609 | 198 | Not_Computer
Z-0550-2014 | 66842 | No | MRG | Vitamin D test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FastPack Free T4 Immunoassay |  |  | November 08, 2013 | 2013 | Qualigen Inc | Qualigen initiated this recall because the FastPack kit may cause lower than expected results due to a manufacturing defect. | Process change control | Qualigen, Inc., notified customers of the recall by calling customers via telephone (with a telephone script) on 11/08/2013. The telephone (script) included the reason for recall, product description with codes, instructions and contact information. The customers were instructed to examine your inventory and discontinue use of the kits immediately. Qualigen will send youreplacement kits and provide you with a shipping label to return all of the defective kits including any that were partially used.Feel free to contact System Support at 760.579.6900 if you have anyquestions or require assistance with returning your kits. |  | 6 | 6 | US Distribution to states of: AL, CA, KS, KY, IN, MA, NC, NY, NV, OK, SC, TN, TX, UT, VA, and WA. | Terminated | September 18, 2014 | 314 | 223 | Not_Computer
Z-0534-2014 | 66865 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8000 and 8015 with Alaris PCA module model 8120 and Auto ID barcode scanner model 8600. |  |  | December 05, 2013 | 2013 | CareFusion 303, Inc. | CareFusion is recalling the Alaris PC unit (model 8000 and 8015) because the Alaris PC unit will produce an audible alarm and the Alaris PC unit will display a System Error message. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated December 5, 2013, to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to contact CareFusion Customer Advocacy at (888) 812-3266, if they experienced a SYSTEM ERROR on the Alaris PC unit. CareFusion will contact all affected customers within 60 days to initiate the scheduling process for the remediation.Customers are instructed to complete and return the enclosed Customer Response Card to expedite the corrective action process. Customers with questions were instructed to contact CareFusion Support Center at 888-562-6018, or for Technical Support call 888-812-3229. |  | 4824 | 4824 total units (4242 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Netherlands, and Australia. | Completed |  |  | 4824 | Not_Computer
Z-0557-2014 | 66900 | No | LOM |  |  | Microbiology | PMA | VITROS Anti-HBs Calibrators &amp; VITROS Anti-HBs Reagent Packs |  |  | October 25, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics issued a Product Correction Notification for VITROS Immunodiagnostic Products Anti-HBs Calibrators and Anti-HBs Quantitative Reagent Packs due to false positive results. | Other | On October 25, 2013, Ortho Clinical Diagnostics issued product correction notifications to their US customers via US Postal Service Priority Mail and email notices to their foreign affiliates informing them of the recall. Ortho Clinical Diagnostics advised their customers to review previously reported results between &gt; 12 and &lt;30 mIU/mL and consult the Laboratory Medical Director. Customers were instructed to utilize alternative methods for anti-HBs testing. Customers are advised to seek alternative methods for anti-HBs testing. For additional information, please contact the OCD Customer Technical Service representative at 1-800-421-3311. |  | 1920 | 1,920 kits | Nationwide distribution - all states. | Open, Classified |  |  | 1920 | Not_Computer
Z-0554-2014 | 66873 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | PedFuse Reset Screw Inserters; PedFuse Screw Inserter 05-7 REspond Part# 11-80111 |  |  | November 20, 2013 | 2013 | SpineFrontier, Inc. | Screw inserters may not mate properly with hex interface of the screws. | Device Design | The firm, SpineFrontier, Inc., sent an "Urgent PedFuse Screw Inserter Voluntary Recall Notification" letter on November 20, 2013 to its consignees/customers. The letter described the product, problem and actions to be taken. The customers were instructed to initiate the immediate return of the recalled products and to complete a recall ACKNOWLEDGEMENT FORM via fax or email to: SpineFrontier, Inc., 500 Cummings Center, Suite 3500, Beverly MA 01915, Attn: Customer Service; Fax: 978-232-3991; email: customerservice@spinefrontier.com.Please feel free to contact at 978-232-3990 Monday thru Friday 8:30AM to 5:30PM Eastern with any questions. |  | 25 | 25 | US Distribution to states of:Texas, New York, Illinois, Ohio, Virginia, Washington D.C., and Florida.. | Open, Classified |  |  | 42 | Not_Computer
Z-0531-2014 | 66930 | No | LZP | Intraocular fluid. | Ophthalmic | Ophthalmic | PMA | Viscocel and Viscocel Plus. |  |  | November 25, 2013 | 2013 | C.L.R. Medicals International, Inc. | CLR Medicals International Inc., initiated a nationwide recall of Viscocel and Viscocel Plus, in response to FDA inspection findings of deficiencies in the Quality System Regulations for Medical Device Manufacturers, 21 Code of Federal Regulation (CFR) Part 820 and lack of Premarket Approval (PMA) for devices sold or distributed in domestic commerce. The Viscoelastic products, intended for human | No Marketing Application | CLR Medicals initiated a nationwide recall on November 25, 2013, of Viscocel and Viscocel Plus, by issuing a press release,in response to FDA inspectional findings of deficiencies in the Quality System regulations for Medical Device Manufacturers, 21 Code of Federal Regulation (CFR) Part 820 and lack of Premarket Approval (PMA) for devices sold or distributed in domestic commerce. The Viscoelastic products, intended for human and animal use, have been determined to be a public health concern due to distribution in the US without FDA approval, which evaluates the safety and efficacy of the product, and violations of the QSR regulations related to sterility, and biocompatibility that may result in the release of contaminated products. CLR Medicals International, Inc. sent a Medical Device Recall letter dated December 6, 2013, to all affected customers. The letter informed customers of the reason for recall, product description with manufacturing date range, possible health hazard, instructions on what to do with recalled product, and contact information for the firm and to report any adverse events to FDA MedWatch. The recall notification letter was accompanied by a response form. Customers with questions were instructed to call 909-398-1880.For questions regarding this recall call 909-398-1880. |  | 249592 | 249,592 units | Worldwide Distribution - USA (nationwide) and Internationally to Egypt, Mexico, and Canada. | Terminated | December 05, 2014 | 375 | 249592 | Not_Computer
Z-0558-2014 | 66971 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | HT Connect Peripheral Guide Wire |  |  | November 25, 2013 | 2013 | Abbott Vascular, Inc. | Abbott Vascular has initiated a recall of the HT Connect Peripheral Guide Wires due to a small number of devices exhibiting partial delamination of the PTFE coating. | Pending | Abbott Vascular sent an Urgent Device Recall letter dated November A recall communication letter dated November 25, 2013, to all affected customers. The letter informed customers of the problems identified and the actions to be taken. The letter informs the customers of the 9 part numbers and 9 description of the products. Abbott Vascular is instructing the customers that the returned products will be replaced with similar products pending availability. Abbott Vascular has attached a customer response form for customers along with the customer notification letter.For questions regarding this recall call 951-914-2400. |  | 54895 | 54,895 units total (5,356 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Australia, French Guyana, Macedonia, South Korea, Austria, Germany, Malaysia, Spain, Bahrain, Greece, Martinique, Sweden, Belgium, Hong Kong, Netherlands, Switzerland, Brazil, Indonesia, New Zealand, Taiwan, Bulgaria, Iran, Norway, Thailand, Canada, Iraq, Poland, Turkey, China, Ireland, Portugal,United Arab Emirates, Czech Republic, Italy, Saudi Arabia, United Kingdom, Denmark, Latvia, Singapore, Estonia, Libya, Slovakia, Finland, Lithuania, Slovenia, France, Luxembourg, South Africa. | Terminated | July 14, 2014 | 231 | 54895 | Not_Computer
Z-0552-2014 | 66956 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Plan Parameters workspace within |  |  | November 18, 2013 | 2013 | Varian Medical Systems, Inc. | An anomaly has been traced to a software deficiency within the Plan Parameters workspace. When a dose relevant parameter is changed and "do no clear MUs and Reference Point dose" is selected, the Monitor units may change. | Software design | Varian Medical Systems will provide all affected consignees with an Urgent Medical Device Correction Letter commencing on November 18, 2013 by federal express. . The Letter will also be distributed to the Varian Sales, Marketing, and Service organizations, informing them of the reported issue. |  | 6701 | 6701 | Worldwide Distribution. | Terminated | March 13, 2015 | 480 | 6701 | Computer
Z-0556-2014 | 67014 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Radiology 12.0, Picture Archiving Communication System (PACSA). |  |  | October 15, 2013 | 2013 | Mckesson Information Solutions LLC | When merging two patient records into one patient record using McKesson Radiology 12.0, the resulting patient record is missing the Contrast Allergy information for the second patient record. | Other | Consignees were notified via letter (Field Safety Notice) on 10/15/2013. |  | 6 | 6 units | MO, NY, TN, WV and France. | Terminated | June 03, 2014 | 231 | 6 | Not_Computer
Z-0551-2014 | 66949 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Integra Cranial Access Kit, Model Number HITHSP10, Rx Only. |  |  | November 26, 2013 | 2013 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | Integra LifeSciences is voluntarily recalling anyunexpired lots of HITHSP10 Cranial Access Kits due to identified labeling discrepancies. kit contains DuraPrep and Monocryl sutures whose labeling does not align with the intended use and labeling of the Cranial Access Kit. | Labeling mix-ups | Integra sent an Urgent Voluntary Medical Device Recall letter dated November 26, 2013, to all affected consignees. The letter identified the product, the problem, the action to be taken by the consignee and included a Recall Acknowledgement and Return form. Consignees were instructed to review their inventory and determine if they have any affected product, stop using the affected product immediately and remove from service. Consignees were instructed to complete the attached form even if they do not have the affected product in their inventory and return by email or fax as indicated on the form. When the form is received Customer Service will send an RMA number and directions to return the product. Customers with questions were instructed to contact Customer Service at 1-855-532-1723 or their Integra sales representative at 909-201-4099.For questions regarding this recall call 609-936-6822. |  | 731 | 731 | Nationwide Distribution | Terminated | April 09, 2014 | 134 | 731 | Not_Computer
Z-0559-2014 | 66535 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | B. Braun Introcan Safety IV Catheter PUR 24G, 0.7X19mm-US, catalog number 4251601-02. |  |  | September 27, 2013 | 2013 | B. Braun Medical, Inc. | One lot of Introcan Safety IV Catheters may not contain the anti-needlestick safety clip which is used for the prevention of needlestick injuries. | Labeling False and Misleading | B. Braun sent an Urgent Medical Device Recall Notification letter dated September 27, 2013 to all affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to utilize the attached "Product Removal Acknowledgement Form" to record the number of individual units in their inventory and fax the completed form to 610-848-1197. To ensure proper credit for affected product, customers were instructed to return the product to B. Braun Medical Inc. For questions contact Customer Support Department at 800-227-2862. |  | 55200 | 55,200 units | USA Nationwide Distribution including Puerto Rico. | Terminated | October 08, 2014 | 376 | 55200 | Not_Computer
Z-0560-2014 | 66934 | No | HTD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Edwards Lifesciences Ergonic Double Acting DeBakey Forceps with Narrow Straight Tip. Model Number ERGF01330. |  |  | November 20, 2013 | 2013 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling one lot of Ergonic instrument due to mislabeling. Ergonic scissors standard curved labeled as a Double Acting DeBakey forcep. Edwards control confirmed that the laser etching on the device was incorrect. | Labeling mix-ups | Edwards sent an " Urgent Field Safety Notice" to all affected customers dated November 20, 2013. The letter identified the product the problem and the action needed to be taken by the customers.Action to be taken by user:Our records show that you received one of these instruments. Our sales representative will be on-site to remove this instrument and provide a replacement. Please assist our sales representative in reviewing your entire inventory for the instrument listed in the recall letter. The form has been pre-populated with the inventory from our records. If you have additional inventory of this affected lot within your control, please add the quantity to the section on the form provided. Please remove the instrument from use, quarantine the device at your site and support the Edwards Sales Representative to return this product to Edwards. An acknowledgment form is included to assist you in the assessment of your inventory.While our sales representative is on-site, please complete the attached acknowledgment form and fax it back to Edwards Customer Service at 800.422.9329. The return of this form allows us to confirm that you have reviewed this notice and have taken appropriate action. Please contact Customer Service at 800.424.3278 if you have questions about the return and crediting process.The Customer Service organization can answer questions about tracking of the replacement instrument provided by the Edwards Sales Representative. If you have questions that have not been answered by this letter, please call Edwards Customer Service at 800.424.3278 from the hours of 6:00AM - 4:30PM PST . |  | 3 | 3 | Distributed to 3 customers in FL only. | Terminated | February 10, 2014 | 82 | 3 | Not_Computer
Z-0569-2014 | 66713 | No | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IMMULITE/IMMULITE 1000 Systems FOL Folic Acid, REF/Catalog Number LKFO1 (100 tests), Siemens Material Number (SMN) 10380902, IVD; for in vitro diagnostic use with the IMMULITE and IMMULITE 1000 Analyzers --- Siemens Healthcare Diagnostics Products Ltd. |  |  | November 06, 2013 | 2013 | Siemens Healthcare Diagnostics | potential for an accelerated decline in counts per second (CPS) for three IMMULITE/IMMULITE 1000 Folic Acid kit lots. This may lead to a failed adjustment before the labeled expiration date is reached. | Unknown/Undetermined by firm | Siemens sent an Urgent Medical Device Correction notices dated November 7 2013, via FedEx to the US customers. The letter identified the product the problem and the action needed to be taken by the customers.Customers are instructed that If the adjustment slope is less than or equal to 1.8 and quality control results are within the established ranges, the kit may be used to generate patient results. If the adjustment slope is greater than 1.8, customers are instructed to discontinue use and discard the kit lots. Ex-US distribution is determined at the country level. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within thirty (30) days. Field service personnel were sent a support document describing the issue and instructing them how to deal with customer questions.For further questions please call CUSTOMER CARE CENTER NUMBER: (302) 631-7597. |  | 1122 | Domestic: 442 kits; Foreign: 680 kits | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of Argentina, Bosnia-Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Cyprus, Denmark, Dominican Republic, Estonia, France, Germany, Great Britain, Greece, Guatemala, Hungary, India, Italy, Latvia, Mexico, Netherlands, Panama, Paraguay, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Tunisia, United Arab Emirates, Uzbekistan and Venezuela. | Terminated | March 23, 2017 | 1233 | 1122 | Not_Computer
Z-0561-2014 | 65131 | No | JMG | Ferritin immunological test system. | Immunology | Clinical Chemistry | 510(k) | Dimension(R) Chemistry Wash (RD701), Siemens Material Number (SMN) 10445052 |  |  | May 06, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics received complaints of Chemistry Wash causing QC and patient sample result shifts on CTN1, LTN1, FT4, TSH, PBNP and LPBN when using Chemistry Wash lots RD23031, RD23111, RD23241, RD23311, RD23391, and RD23461. These lots shipped between 12/12/2012 and 3/22/2013. Internal investigation has confirmed shifts can be encountered in two scenarios: 1 - When an affected l | Software design | Siemens sent an Urgent Medical Device Recall letter dated May 2013 to all affected customers. The letter identified the product and lots, the reason for the field action, the risk to health and the action(s) to be taken. Customers were instructed to discontinue use and discard inventory of the affected product. Siemens will replace any unused inventory of the affected lots at no charge. For questions contact your Siemens Technical Center or your local Siemens technical support representative. |  | 17719 | 17,719 | US Nationwide and OUS. | Terminated | August 27, 2014 | 478 | 17719 | Not_Computer
Z-0571-2014 | 66754 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Instrument Manager driver : Driver sysu100i that connects to the Sysmex UF-100 analyzer |  |  | November 06, 2013 | 2013 | Data Innovations, Inc. | When connected to the Sysmex UF-100, theInstrument Manager driver sysu100i driver software versions 7.00.0001, 7.00.0002, 7.00.0003, 7.00.0004, 7.00.0005, 8.00.0001, 8.00.0002, 8.00.0003,and 8.00.0004, and is configured with option "Strip Zeros from Result" set to "Yes" may result in sample error | Software design | Data Innovation notified consignees by Important Medical Device Recall Notification dated 11/5/13 to all list serv recipients of the problem and posted critical alert on Customer Web Portal. A certified letter was issued to impacted customers and consignees if no response was recieved from e-mail. |  | 24 | 24 systems | NationwideCanada and Britain | Terminated | December 05, 2014 | 394 | 24 | Computer
Z-0570-2014 | 67046 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator with CO2/EtCO2 Measurement |  |  | December 10, 2013 | 2013 | Philips Medical Systems, Inc. | MRx defibrillator displays a -?- for EtCO2 and does not display EtCO2 values when patient CO2 level falls below 7.6 mmHg | Software design | Philips Healthcare issued an Customer Information letter dated February 20, 2012, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer with the addendum to the Instructions for Use.If you have a ? in the parameter block and no CO2 waveform on the display, the waveform source is invalid. If invalid, check patient, confirm airway status and examine the FilterLine to make sure it is connected to the device and not kinked or pinched. For more information on the EtCO2 Option seeMonitoring Carbon Dioxide chapter in the HeartStart MRx Instructions for Use.Should you have any questions or concerns, please contact your local Philips representative at 1-800-722-9377. |  | 18529 | 18,529 units | US Nationwide and OUS. | Terminated | April 18, 2017 | 1225 | 18529 | Not_Computer
Z-0567-2014 | 66976 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Ziehm Vision R C-Arm Interventional Fluoroscopic X-Ray System. |  |  | November 21, 2013 | 2013 | Ziehm Imaging Inc | Engineering change in 2006 resulted in a minor non-conformity of their devices. Specifically the emergency switch was changed from a standard component to just an option on the devices. | Device Design | Ziehm Imaging sent an "URGENT RECALL NOTICE" letter dated November 22, 2013, to all affected customers. The letter identified the products the problem and the action needed to be taken by the customers. Customers were advised and instructed to:Your Model and Serial number can be found on the product label attached to the side of the mobile C-arm Stand. Our records indicate that your facility has one or more of the products listed above that are being used in your facility.Recall of the product was initiated because according to our manufacturing documentation, the referenced device installed at your facility was not equipped with an emergency stopping device for vertical travel of the c-arm imaging assembly. To fully comply with applicable standards the referenced device must be retrofitted with an emergency stopping device identification and label.The issue:According to our manufacturing documentation, the referenced device installed atyour facility was not equipped with an emergency stopping device.To fully comply with applicable standards the referenced device must be retrofitted with an emergency stopping device. The manufacturer Ziehm Imaging GmbH advised customers they may continue to operate the affected device with caution until the correction has been carried out on-site. During this time Ziehm Imaging GmbH advises the operator to be aware of the following safetymeasures which may resolve the possibility of potentially hazardous situations by:" Operating the power-off switch located either on the C-arm Stand top panel or on the monitor cart" Removing the power cable from the wall outlet or," Activate the push-buttons for reverse vertical movement and keeping the pushbutton pressed until patient /operator is no longer in the risk zone or the device stops Please alert all staff members working with the device to these potential hazardous situations, and familiarize them with all possibilities to resolve the hazardous situation |  | 103 | 103 devices | USA (Nationwide) Distribution including the states of AK, AL, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MT, MA, MD, MI, MN, MN, MO,MS, NC ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and Hawaii. | Terminated | June 23, 2014 | 214 | 386 | Not_Computer
Z-0579-2014 | 67003 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | OrthoPediatrics ACL Reconstruction System is a pediatric and small stature adult ACL reconstruction system containing Sleeve Installation Tool Set , Part 01-1008-0040/0047, 6mm; 6.5mm, 7mm, 7.5mm, 8mm, 8.5mm, 9mm, and 10mm. |  |  | November 07, 2013 | 2013 | OrthoPediatrics Corp | Part # 01-1008-0040/0047, ACL Sleeve Installation/ Insertion Tool, shipped November 4, 2013 - November 6, 2013 may be etched with the wrong size. | Labeling design | November 7, 2013 all consignees were notified of the Urgent Device recall via telephone calls and E-mail messages describing the product and reason for the recall,. The notice included directions to identify and quarantine the recalled devices returning product to OrthoPediatrics with completed return form to OrthoPediatrics via mail, email or courier service, or byphotographing with a smart phone and emailing to logistics@orthopediatrics.com or by answering therecall specific questionsin an email to mfox@orthopediatrics.com |  | 40 | 40 | Nationwide Distribution including the states of CA, KY, OH, RI, and TN. | Terminated | May 09, 2014 | 183 | 40 | Not_Computer
Z-0577-2014 | 66935 | [6 empty] | ProPlan CMF Guide(s) Model(s) |  |  | November 08, 2013 | 2013 | Materialise USA LLC | Patient Specific Guide contained the incorrect case report. The case report contains patient specific surgical plan information. | Labeling Change Control | Customer/ ordering physician was made aware of the issue11/08/2013. Revised accurate case report was immediately provided. |  | 1 | 1 | Distributed in Canada. | Terminated | January 07, 2014 | 60 | 1 | Not_Computer
Z-0573-2014 | 66914 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Indus Invue Screws: IM71058-XX: 04.2mm, SelfTapping, Tapered. |  |  | November 20, 2013 | 2013 | SpineFrontier, Inc. | Mismarked and unmarked screws | Packaging process control | On 11/20/2013, consignees were notified of the email via letter. Consignees were instructed to return the screws listed in the recall notification letter and return the Recall Acknowledgement form to SpineFrontier via fax or email. |  | 993 | 993 | Nationwide Distribution- including DC and the states of OH, TX, KS, MA, IN, CO, NY, AL, and MI. | Open, Classified |  |  | 1915 | Not_Computer
Z-0580-2014 | 66980 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | steroEOS Workstation with software versions: 1.3, 1.4 and 1.5 used in processing 3D X-ray images |  |  | November 29, 2013 | 2013 | EOS Imaging | Error copying information to the clinical file and to the report when two 3D models are opened simultaneously on the sterEOS workstation | Software design | EOS imaging issued a safety information letter dated November 29, 2013, to the US customer sites. The letter identifed the affected product, problem and actions to be taken. Users of the sterEOS workstation are asked not to simultaneously open several pairs of images in sterEOS 3D. The letter informed customers that EOS imaging will, at it's own expense, corect all systems affected by means of a software update. An EOS imaging representative will contact affected customers to make an appointment for this update once it is available. Users of the sterEOS workstation are asked not to simultaneously open several pairs of images in sterEOS 3D. For questions call +33 (0) 155 25 6060. |  | 36 | 36 units | USA Nationwide Distribution | Terminated | January 13, 2015 | 410 | 36 | Computer
Z-0572-2014 | 66916 | No | LIP | Treponema pallidum treponemal test reagents. | Microbiology | Microbiology | 510(k) | Captia Syphilis-G Test Kit |  |  | August 27, 2013 | 2013 | Clark Laboratories, Inc. (dba,Trinity Biotech USA) | Trinity Biotech recalled Trinity Biotech Captia Syphilis IgG kits due to false negative results. | Process design | Trinity Biotech issued an Urgent Product Information Request letter to their customer on August 27, 2013. The letter identified the affected product, problem and actions to be taken. The request was to determine if their customer did or did not backfill microtiter-strips during the analytical process. If a customer backfilled, they were advised to return the completed request form. On September 11, 2013 Trinity Biotech issued a notice to their customers advising them to re-evaluate any negative patient results for those who perform backfilling. For questions contact Trinity Biotech Help Desk at 800-325-3424, option 2. |  | 38 | 38 kits | USA Nationwide Distribution in the state of WA | Open, Classified |  |  | 38 | Not_Computer
Z-0581-2014 | 67042 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REFLECTION(R), MICROSTABLE(R), ACETABULAR LINER, 20o OVERHANG, 28 MM ID, 54-56 OD, QTY: (1), STERILE EO, REF 71742854, Smith &amp; Nephew, Inc. Memphis, TN 38116 USA. |  |  | November 25, 2013 | 2013 | Smith &amp; Nephew Inc | One batch of the Reflection(TM) Polyethylene Acetabular Liners may be out of tolerance causing the liner not to fully seat into the shell. | Under Investigation by firm | All affected sales staff was notified of problem and the recall via email on 11/25/2013. There were no US hospitals affected by this recall. |  | 38 | 38 units | Distributed in China and Spain. | Open, Classified |  |  | 38 | Not_Computer
Z-0593-2014 | 67052 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | June 07, 2013 | 2013 | Philips Medical Systems, Inc. | Error in the printing of the therapy energy setting labels on Philips HeartStart M3536A MRx Monitor/Defibrillators. The first energy setting is labeled 10 when it should be labeled 1-10 | Labeling Change Control | Philips sent a Customer Information Letter dated June 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.During an EI of the firm, the firm had not reported the correction to FDA.The Customer Information letter informs the customers what the behavior is and under what circumstances. The correction consisted of a replacement label that was provided free of charge to all units affected by this issue. Philips Healthcare representatives have been contact customers with affected devices to arrange for installation of the replacement label.To correct this issue, Philips will replace the label free of charge. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the replacement label. This issue does not affect current product shipments.If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377.Philips apologizes for any inconvenience caused by this problem. |  | 93 | 93 units | Worldwide Distribution - US Distribution including the states of IN, MS, TX and VA., and the countries of DENMARK, FRANCE, GERMANY, SPAIN and SWITZERLAND. | Terminated | March 22, 2017 | 1384 | 93 | Not_Computer
Z-0603-2014 | 67064 | No | MNJ | Cardiopulmonary bypass blood reservoir. | Cardiovascular | Cardiovascular | 510(k) | Terumo Cardiovascular Custom Cardiovascular Procedure Kits containing an Overpressure Safety Valve, LN130B, in the Arterial Recirculation Line |  |  | June 04, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Over-Pressure Safety Valves not cleared with an indication for use on the arterial side of the extracorporeal bypass circuit. | Component design/selection | Terumo Cardiovascular issued Customer notification letters on June 4, 2013. Accounts advised to: Review this notice and decide to continue to use or not use the Cardiovascular Procedure Kits as configured until a replacement kit with a one-way valve indicated for use on the arterial side of the perfusion circuit is available. If you have questions, contact Terumo CVS Customer Service: 800-521-2818. |  |  | 24,227. | NationwideCanada | Terminated | October 01, 2014 | 484 |  | Not_Computer
Z-0587-2014 | 66499 | No | FZS | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DIST LAT FEM LOCK PLT, LT |  |  | November 05, 2013 | 2013 | Zimmer, Inc. | Based on a review of historic packaging validations, it was determined that a subset of product packaged at a Zimmer facility required testing to ensure that previously packaged product from this facility met the established requirements for sterile barrier integrity. This testing determined that a packaging tray configuration which consists of an inner and outer sterile tray, resulted in failure | Process change control | On 11/5/2013, Urgent Medical Device Recall Notifications were sent to affected distributors, hospital Risk Managers, and doctors informing them of the failure. All distributors were notified via electronic mail. Hospital risk managers, surgeons and distributors with affected inventory within the expiration date of the device will also be notified via courier. The letter identifies the issue, health risks, and their responsibilities. Questions concerning the recall are directed to 1-877-946-2761Communications outside of the United States will occur approximately two weeks after the United States communications.Zimmer will conduct effectiveness checks in the following way:Distributors/Hospitals/Surgeons100% of the notifications will be tracked to ensure delivery of the notifications. An inventory certification form will need to be returned from each distributor that certifies that all available affected product from their territory has been returned.Additional notifications via email and FedEx will be sent to those who have not completed the required certification form or returned affected product.Accounts will be deemed unresponsive after 3 attempts. |  | 42 | 42, 065 units total | Worldwide Distribution. | Terminated | May 05, 2015 | 546 | 42 | Not_Computer
Z-0602-2014 | 67094 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Camino Advanced Monitor with Waveform Display Model CAM01 Integra Neurosciences, 311 Enterprise Dr Plainsboro, NJ 08536 USA |  |  | November 26, 2013 | 2013 | Integra LifeSciences Corp. | Integra became aware through a complaint from a single non-US customer that if the optional SYS/DIAS mode is used on CAM01 monitor that contains Mod 4 Firmware, the displayed values for the systolic and diastolic intracranial pressure (ICP) are reversed. But, if the CAM01 is used with an auxiliary bedside monitor, the bedside monitor displays these values correctly. | Device Design | Integra sent an Urgent Medical Device Correction letter and Correction Acknowledgement form dated November 26, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Our records indicate that you received one or more of the affected serial numbers that will require this firmware update. Integra asks that you do the following:1. Review your inventory and verify that you have Camino Advanced Monitors with a serialnumber formatted as MExxxxxx04. If so, you can continue to use the CAM01monitor(s), being aware that if you choose to use the optional SYS/DIAS mode, thevalues on the CAM01 display will be reversed.2. Complete the attached form. If you do not have the product with the affected serialnumber in service at your facility, check the box:I do not have any product with a serial number of the form MExxxxxx04.3. If you do have the affected serial number, check the box:I do have the product with a serial number of the form MExxxxxx04, and pleaseindicate the serial numbers you have.4. Complete the other information and return it by email or fax as indicated on the form.Complete the form and return it by email or fax as indicated on the form. Should you have any questions regarding these instructions, please contact Customer Service at 1-855-532-1723.We apologize for any inconvenience this may cause and thank you for your cooperation in this effort. |  | 276 | 275 CAM01 Monitors | Worldwide Distribution - USA ( Nationwide ) and the countries of : UK, Saudi Arabia, South Africa, Brunei, China, France, Poland, Turkey, Greece, Sweden, United Arab Emirates, Russia, Spain, Pakistan, Finland, Israel, Italy, Slovenia, Japan, Poland, Malaysia, Germany and Panama. | Terminated | February 13, 2015 | 444 | 276 | Not_Computer
Z-0599-2014 | 66895 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Asante Conset Infusion Set with 60cm tubing &amp; 9mm cannula packaged in 5-packs,singles, or inside a Combination Kit (FG-5229-4); |  |  | November 19, 2013 | 2013 | Asante Solutions, Inc. | Reports of the infusion tubing disconnecting from the connector, resulting in leakage of insulin from the infusion set and providing patient with less insulin than intended. | Device Design | Asante sent urgent medical device recall notifications letters to distributors and customers on November 19, 2013 by US mail, return receipt requested. |  | 1979 | 1979 kits | Nationwide Distribution. | Terminated | April 24, 2014 | 156 | 6977 | Not_Computer
Z-0583-2014 | 66909 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Ball Impactor Tip |  |  | September 30, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker has received complaints associated with cracks and/or fracture of the Ball Impactor Tip and Rim Impactor Tip Instruments. | Device Design | Stryker sent an Urgent Medical Device Correction letter dated October 1, 2013 to all affected customers via Fed Ex. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and fax the attached Product Correction Acknowledgement form to 855-251-3635. |  | 1024 | 1,024 units | USA Nationwide Distribution | Open, Classified |  |  | 2714 | Not_Computer
Z-0584-2014 | 67012 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Essenta DR, a Stationary x-ray system, X-ray system |  |  | December 05, 2013 | 2013 | Philips Medical Systems, Inc. | X-ray system C-arm casting may loosen and fall | Nonconforming Material/Component | Philips Medical Systems issued a Urgent Field Safety Notice-Warning Label and Inspection, dated November 21, 2013, advising of potential of C-arm detaching and requesting a daily inspection and a safety warning label . If you need any further information or support concerning this issue, please contact your local Philips representative. 1-800-722-9377. |  | 424 | 424 units | Worldwide Distribution - USA (Nationwide) including Puerto Rico and the countries of : Afghanistan, Argentina, Australia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Czech Republic, Egypt, France, Germany, Greece, Hungary, India, Indonesia Iran,Italy, Jordan, Kazakstan, Kenya, Liberia, Lithuania, Morocco, Netherlands, New Zealand, Nigeria, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Thailand, Turkey, Ukraine, United Arab, United Kingdom and Viet nam. | Open, Classified |  |  | 424 | Not_Computer
Z-0605-2014 | 67036 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | VERIFYÂż Dual Species Self-Contained Biological Indicator, Distributed by STERIS Corporation, Mentor, OH. |  |  | September 10, 2013 | 2013 | Steris Corporation | STERIS has identified that the population of the G. stearothermophilus spores that monitor STEAM sterilization cycles does not meet specifications for Lot # 131004 of the Verify Dual Species Self-Contained Biological Indicators. As a result the spore count on the affected product is lower than the labeled minimum and the product may incorrectly indicate that a 5-log spore kill was achieved. | Use error | On September 10, 2013, the firm sent Urgent Voluntary Recall Notices to their customers. |  | 1238 | 1,238 boxes | Worldwide Distribution-USA (nationwide) including the states of AK, Al, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY, and the countries of BAHRAIN, CANADA, COLUMBIA, GERMANY, ITALY, MALAYSIA, PHILIPPINES, ROMANIA, SPAIN, and THE NETHERLANDS. | Terminated | February 14, 2014 | 157 | 1238 | Not_Computer
Z-0604-2014 | 66981 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | STERRAD Cyclesure Biological Indicator (BI, Part No. 14324. |  |  | November 26, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD Cyclesure Biological Indicator (BI) because ASP has identified a microscopic intermittent and low frequency flaw in the outer vial of the STERRAD Cyclesure 24 Biological Indicator. | Nonconforming Material/Component | A customer notification letter was sent to all customers on 11/26/2013 who purchased the STERRAD CYCLESURE Biological Indicator (BI). The letter informs the customer of the problems identified and the actions to be taken. The customer notification letter includes a table with the products on recall which includes the manufacturing dates, lot numbers, expiration dates, and quantity. |  | 4048 | 4048 units total (4014 units in US) | Worldwide Distribution. | Terminated | March 05, 2014 | 99 | 4048 | Not_Computer
Z-0607-2014 | 66979 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | IntelliFill i.v. Pharmacy Compounding Device (PCD). |  |  | November 08, 2013 | 2013 | Baxter Healthcare Corp. | Following a retrospective review for Baxter Daytona Complaint CR-954, it was identified that the IntelliFill i.v. Pharmacy Compounding Device (PCD) (product code: IFIV) has the potential to continue to run and fill syringes with the HEPA filtration disabled. | Nonconforming Material/Component | A Service Bulletin was sent to all IntelliFill i. v. customers February 22, 20 I 0 by USPS mail.The Service Bulletin informs customers that IntelliFill i.v. operators are to be trained to listen for the running HEPA filtration fan motor and observe/feel for airflow coming from the automation deck.The field service engineer dispatched to the site replaced the damaged sensor on 11-Feb-2010 and a Service Bulletin was sent to all customers regarding the function and use of the HEPA filter on 22-Feb-2010.A new HEPA sensor was designed and installed into all current IntelliFill i.v. devices between 21-July-2011 to 9-Feb-2012. |  | 28 | 28 devices | Nationwide Distribution including the states of CA, CO, CT, FL, IL, LA, MA, MD, MI, NC, NE, NY, OH, OR, PA, TN, TX, and WV. | Terminated | March 23, 2015 | 500 | 28 | Not_Computer
Z-0606-2014 | 66902 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | RTC 25-C Inline Aerosol Tee Adapter. |  |  | November 13, 2013 | 2013 | Instrumentation Industries Inc | Some packages of the RTC 25-C may have been shipped to customers with incomplete Installation and Usage Directions due to a printing error. | Labeling Change Control | A recall notification letter, dated November 13, 2013, was sent to Users. |  | 220 | 220 | Nationwide Distribution. | Terminated | July 09, 2014 | 238 | 220 | Not_Computer
Z-0608-2014 | 66968 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CPT Hip System, Parts 00-8114-002-10 and 00-8011-004-00 |  |  | December 10, 2013 | 2013 | Zimmer, Inc. | Zimmer is initiating a lot specific recall of the CPT Hip Stem, Versys Build-Up Block, and Trabecular Metal Reverse Glenosphere medical devices due to reports of the products being packaged in the incorrect outer carton. Product labels and patient labels are correct. | Labeling mix-ups | Zimmer sent an Urgent Medical Device Recall letter dated December 10, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification and ensure affected personnel were aware of the contents. If affected product was found, it should be quarantines and their Zimmer representative notified. The Zimmer sales representative will remove the recalled product from their facility. Customers with questions were instructed to call 1-877-946-2761. . |  | 36 | 36 | Worldwide Distribution -USA including OH, FL, WI, MI NJ, SC, CA, NC, PA, OR, WA, TN, KY, OK, KS, AZ, TX, CO, NV and Internationally to Germany.There have been 71 distribution events. Note that this product is lot controlled and may be distributed multiple times if it was returned and returned to the shelf. 7 units have been returned, including three through the complaint process, which leaves 64 affected units in the field subject to removal. | Terminated | September 11, 2014 | 275 | 64 | Not_Computer
Z-0618-2014 | 66468 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | S3 Med/Surg Bed |  |  | November 26, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | Stryker Medical has identified a potential issue associated with Secure II and S3 Med/Surg beds built between April 2nd, 2005 and July 26th, 2011. The affected beds have the potential to experience damage to the Foot End Cover and/or Power Coil Cable due to reduced clearance between the Foot End Cover and Foot End Lift Header. | Device Design | Stryker Medical sent an Urgent Medical Device Recall Notification letter on November 26, 2013, to all affected customers. The notification included the products being recalled, the product description, product issue, potential hazards, and actions needed. Customers were instructed to locate and inspect the beds listed in the notice and ensure that no damage has occurred to the foot-end cover, power coil cable, or sensor coil cable. If the beds were not damaged customers may continue to use them. If damage occurred, customers wsere instructed to remove the beds from service and contact Stryker at 1-800-327-0770, Option 2, then Option 2 to arrange a service call. Continue to inspect the foot-end cover and coil cables as part of their routine preventive maintenance. Customers were instructed to fax the enclosed business reply form to 269-488-8691 or email productfieldaction@stryker.com. Customers were instructed to forward the letter if the product was loaned or sold to other customers. If any beds were disposed of and were no longer in use, customers should advised Stryker by providing them with the serial number. Questions or concerns are directed to Stryker Customer Service at 1-800-327-0770. Consignees are to fax (269-488-8691) or email (productfieldaction@stryker.com) the business reply form to confirm receipt of this notification.For questions regarding this recall call 269-324-6609. |  | 108769 | 108,769 Total | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Chile, China, Argentina, Brazil, Germany, Korea, Mexico, Singapore, UK, Europe, Middle East, Africa, and Latin America. | Terminated | May 29, 2015 | 549 | 108769 | Not_Computer
Z-0614-2014 | 67128 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Catheter, Continuous Flush |  |  | December 06, 2013 | 2013 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is conducting a recall for 3 lots of Maestro Microcatheters with a graphicdiscrepancy on the carton/box label end panel. The catheter graphic on the front of the cartonbox (correct) and the graphic on the end panel (incorrect) do not match. | Packaging process control | Merit Medical Systems sent an Urgent Product Recall Notice dated December 6, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use of the Maestro Microcatheters identified and isolate them in a quarantined area until the product is over-labeled by their Merit Sales Representative Customers were instructed to ensure that all personnel to whom the devices were distributed were made aware of the field action. Customers with questions were instructed to contact their sales representative or call 1-801-826-4062.For questions regarding this recall call 801-253-1600. |  | 59 | 59 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | January 27, 2014 | 52 | 59 | Not_Computer
Z-0613-2014 | 67035 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | April 10, 2013 | 2013 | Philips Medical Systems, Inc. | Device Operating on Battery Power May Shutdown without Warning if exposed to elevated levels of electromagnetic interference from other radio frequency (RF) energy sources and | Environmental control | Philips issued a Medical Device Correction letter April 2013 to all affected customers along with an IFU Addendum titled, HeartStart MRx Electromagnetic Compatibility. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that the condition could be avoided by operating the device on external AC or external DC power. Customers were instructed to follow the device instructions when using battery power to minimize or eliminate potential interference from other RF sources. Customers with questions were instructed to contact their local Philips representative.For questions regarding this recall call 978-687-1501. |  | 64972 | 64,972 | Worldwide Distribution - USA (nationwide) and Internationally to the following countries:AFGHANISTANALBANIAARGENTINAARUBAAUSTRALIAAUSTRIAAZERBAIJANBAHRAINBANGLADESHBELGIUMBOLIVIABOSNIA AND HERZEGOWINABRAZILBRUNEI DARUSSALAMBULGARIACAMBODIACANADACHILECHINACOLOMBIACOSTA RICACROATIA (local name: Hrvatska)CYPRUSCZECH REPUBLICDENMARKECUADOREGYPTEL SALVADORESTONIAETHIOPIAFINLANDFRANCEGERMANYGREECEGUADELOUPEHONG KONGICELANDINDIAINDONESIAIRAQIRELANDISRAELITALYJAPANJORDANKAZAKHSTANKENYAKOREA, REPUBLIC OFKUWAITLAOSLATVIALEBANONLIBYAN ARAB JAMAHIRIYALITHUANIALUXEMBOURGMACAUMACEDONIAMALAYSIAMALTAMEXICOMONGOLIAMOROCCONAMIBIANEPALNETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNICARAGUANIGERIANORWAYOMANPAKISTANPALESTINEPANAMAPAPUA NEW GUINEAPARAGUAYPERUPHILIPPINESPOLANDPORTUGALPUERTO RICOQATARREUNIONROMANIARUSSIAN FEDERATIONSAUDI ARABIASERBIASINGAPORESLOVENIASOUTH AFRICASPAINSRI LANKASWEDENSWITZERLANDTAIWAN, PROVINCE OF CHINATANZANIA, UNITED REPUBLIC OFTHAILANDTRINIDAD AND TOBAGOTUNISIATURKEYTURKMENISTANUGANDAUKRAINEUNITED ARAB EMIRATESUNITED KINGDOMUZBEKISTANVENEZUELAVIET NAMYEMEN | Terminated | April 13, 2017 | 1464 | 64972 | Not_Computer
Z-0616-2014 | 67002 | No | MLR | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | SYSTEM 1E Liquid Chemical Sterilant Processing System, Distributed by STERIS Corporation, Mentor, OH |  |  | November 21, 2013 | 2013 | Steris Corporation | Customers indicated fluctuations in environmental conditions sites and gradual degradation/deformation of check valve o-rings following frequent contact with S40 sterilant use solution result in nuisance aborts. | Device Design | STERIS sent an Urgent Voluntary Field Correction Notice dated November 21, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed that a STERIS Field Service Representative would contact their facility to schedule installation of the new hardware and updated software on their SYSTEM IE Liquid Chemical Sterilant Processing System(s). Customers were instructed that they may continue to operate the SYSTEM IE processor prior to the completion of the update. Customers with questions were instructed to contact STERIS Field Service Dispatch at 1-800-333-8828.For questions regarding this recall call 440-392-7519. |  | 6028 | 6,028 units | Nationwide Distribution - USA (nationwide) and the District of Columbia and Puerto Rico.Product was also shipped to the following countries: Afghanistan, Bahrain, Hong Kong, Italy, Saudi Arabia, Turkey &amp; United Kingdom. | Terminated | June 16, 2014 | 207 | 6028 | Computer
Z-0615-2014 | 65007 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Hand Switch - Electric Pen Drive, Part Number 05.001.012 |  |  | April 15, 2013 | 2013 | Synthes USA HQ, Inc. | Synthes received a complaint from a customer in which the Hand Switch for the Electric Pen Drive was not labeled correctly. The "on" and "lock" labels were interchanged. | Process design | Synthes sent an Urgent Notice: Medical Device Recall letter dated April 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for the affected product, remove from inventory, and return them to Synthes immediately for replacement. Customers were instructed to call Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number. Complete the Verification Section at the end of the letter by checking the appropriate box indicating affected product has been located. Return the Verification Form with the product to: Credit/Returns, Synthes1101 Synthes AvenueMonument, CO 80132If customers do not have the affected product they should still complete the Verifcation Section at the end of the letter by checking the appropriate box and faxing the document to 610-251-9005 or email to: FieldAction@synthes.comCustomers with questions were instructed to call 610-719-5450 or email FieldAction@synthes.com.For questions regarding this recall call 610-719-5000. |  | 290 | 290 | Worldwide Distribution - USA (nationwide) and Internationally to Mexico and Canada.. | Open, Classified |  |  | 290 | Not_Computer
Z-0620-2014 | 64807 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | March 12, 2013 | 2013 | GE Healthcare It | GE Healthcare is aware of a potential safety issue with the use of GE Centricity Laboratory Instrument Interface (IF) where free-text sent from IM is not being transcribed into free-text in iNET. The instrument may upload results to IF/iNET when order/request is not yet in IF/iNET. In such a situation, test level supporting free-text sent from the instrument is discarded (not captured). | Device Design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated March 12, 2013 to all affected customers. The letter notified customers of the problem, suggested that customers experiencing this problem should retransmit the result from the instrument once the order exists in iNET and verify that the item level free-text is present in iNET, and to contact their GE service representative to request or install the updated software. Customers with additional questions can contact the GE Healthcare IT Help Desk at 1-847-589-8496 or tgg.nahelpdesk@med.ge.com.For questions regarding this recall call 847-589-8496. |  | 5 | 5 units | Worldwide Distribution - US including MN and Internationally to India, Australia, Canada, and England. | Open, Classified |  |  | 5 | Computer
Z-0619-2014 | 67069 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx Gram-Negative QuickFISH BC. |  |  | December 11, 2013 | 2013 | AdvanDx, Inc. | Kits incompatible with certain fluorescence microscope light sources | Process design | AdvanDx issued an email notification along with a PDF file of the new package insert to each customer followed by written notification dated December 6, 2013 and a paper copy of the new package insert via Fed Ex courier. All customers received an acknowledgement form to fill out and return to AdvanDx so that the effectiveness of the replacement can be evaluated. All existing inventory of Gram-Negative QuickFISH BC will be repackaged with the new package inert Customers with questions were instructed to contact AdvanDx Technical Support at 1-866-376-0009 or techwupport@advandx.com.For questions regarding this recall call 781-376-0009. |  | 66 | 49 kits US; 17 kits OUS | Worldwide Distribution - USA including CA, IL, IN, OH, NJ, MI , MO, NY, WI. Internationally to Denmark. | Terminated | January 13, 2014 | 33 | 66 | Not_Computer
Z-0491-2014 | 66744 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX-Revolution Mobile X-Ray System -- Made in U.S.A. by Carestream Health Inc. 150 Verona Street, Rochester, NY . 14608 --- The DRX Revolution Mobile system is a mobile diagnostic x-ray system for generation of x-rays for examination of various anatomical regions. These products are mobile diagnostic x-ray systems designed to perform radiographic x-ray examinations on all pediatric and adult patients, in all patient treatment areas. |  |  | August 29, 2013 | 2013 | Carestream Health Inc. | It was discovered that Carestream received a report from a customer that 2 different patients were re-exposed due to the fact that the processed images were not able to be found in the system. | Radiation Control for Health and Safety Act | Carestream sent a letter to all affected customers informing them of the issue, how to detect the issue and actions to be taken by Carestream to correct the issue. The actions to be taken by Carestreamn to correct the issue will include the installation of a software patch to mitigate the effects of the observed failure and remove the possibility of a re-exposure in this failure mode. The correction will be implemented within the US to all affected devices at identified customer sites. For further questions call Carestream Customer Care at 1-800-328-2910, 8 am through 8 pm Eastern Standard Time. |  | 310 | 310 systerms domestically | Nationwide | Terminated | November 18, 2015 | 811 | 310 | Computer
Z-0623-2014 | 67050 | No | GET | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Aesculap Hi-Line XXS handpiece |  |  | December 09, 2013 | 2013 | Aesculap, Inc. | Certain lots of the Aesculap Hi Line XXS handpiece (part # GB790R) were recalled due to a possible failure of the handpiece. | Device Design | Aesculap sent an Important Correction &amp; Removal Notification letter dated December 9, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to quarantine all affected product and returned to AIC (USA). Customers were instructed that a Aesculap Sales Representative would contact them to schedule return of any affected inventory and replace their product. Customers were asked to complete the inventory sheet with the quantity being returned. If customers could not locate the product, they were to provide an explanation as to why the inventory would not be returned (discarded, etc.). Customers with questions were instructed to call 610-984-9414,For questions regarding this recall call 800-258-1946, ext 5067. |  | 14 | 14 | Worldwide Distribution - USA including TX, NY, PA, and TN. Internationally to Austria, Switzerland, Germany, Denmark, Spain, Finland, France, Hong Kong, Israel, Italy, Japan, Netherlands, Poland, Russian Federation, and Taiwan. | Terminated | December 02, 2014 | 358 | 14 | Not_Computer
Z-0624-2014 | 66273 | No | HXB | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PedFuse Bone Probe; Part: SI81008 spinal pedicle fixation. |  |  | September 11, 2013 | 2013 | SpineFrontier, Inc. | Bone Probes used to create a pathway in pedicle bone were manufactured and marked to incorrect specifications. | Process change control | On 9/13/2013, consignees were notified of the recall via email. They were asked to return the recalled product and complete and return the recall acknowledgement form via email or fax. |  | 13 | 13 devices | US Distribution: including states of: TX and VA. | Terminated | October 08, 2014 | 392 | 13 | Not_Computer
Z-0628-2014 | 67039 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 42-411571 QTY 1 CATALOG CAUTION. SEE INSTRUCTIONS FOR USE |  |  | November 26, 2013 | 2013 | Materialise USA LLC | Shipped femur guide contained a non-conformance which could possibly impact fit of the drill insert. | Nonconforming Material/Component | Materialise USA LLC notified the Field Representative on November 26, 2013 of the recall by phone followed by E-mail notification. The notification identified the device and reason for the recall. Notification directions included instructions to find and quarantine the device and return it Materialise. A replacement device was shipped November 27, 2013 receiving the defective device was called, followed by E-mail notification the Recall, which identified the device, problem, and instructed the distributor to find and quarantine product. For questions regarding this recall call 734-662-5057. |  | 1 | 1 | US Distribution: MO only. | Terminated | January 14, 2014 | 49 | 1 | Not_Computer
Z-0532-2014 | 66064 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Spectre Wireless Footswitch/Handswitch system Model 98/99, Model FDS, Model FDB. |  |  | August 22, 2013 | 2013 | Trans American Medical, Inc. | TransAmerican Medical Imaging is recalling certain lots of TMI Spectre Wireless Footswitch/Hand switch for GE OEC 9800/9900 and Philips FDS and FDB Cath Labs due to potential safety issue related to the use of this product producing uncommanded X-ray. | Component design/selection | TransAmerican Medical sent an Urgent Medical Device Correction letter dated August 26, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure that all potential users in their facility are made aware of the safety notification and the recommended actions. If the customers system does produce an uncommanded X-ray, turn the system off immediately. Remove the Spectre Wireless Footswitch and receiver and reinstall the standard OEM footswitch. Follow proper shutdown/system reboot instructions as indicated in the system operator's manual. Customers will be contacted to make arrangements for replacement of their TMi Spectre Wireless Footswitch at no charge. Customers will be provided a new TMI Spectre Wireless Footswitch and instructions on how to return the affected footswitch. Customers with questions were instructed to call 801-796-7335, ext 203. |  | 68 | 68 | US Distribution: including states of: WA and UT. | Terminated | December 29, 2015 | 859 | 68 | Not_Computer
Z-0626-2014 | 65409 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | HbA1c Flex(R) reagent cartridge (K3105A, Siemens Material Number 10470481) for the Dimension(R) Clinical Systems. |  |  | April 22, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Positive bias on the lots listed compared to alternate methods and a higher frequency of "Above Assay Range" flags with the lots listed. | Process control | Siemens sent an Urgent Medical Device Recall letter in April 2013 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discard any remaining inventory of the affected product. Siemens will replace any unused inventory at no charge. Customers were asked to complete the attached form and fax it to 302-631-8467 to indicate they have received the information. Customers were asked to retain the letter with their laboratory records and forward the letter to those who may have received the affected product. Customers with questions were instructed to contact their Siemens Technical Solutions Center or their local Siemens technical support representative.For questions regarding this recall call 1-800-441-9250. |  | 11006 | 11006 Total | Worldwide Distribution - USA (nationwide) and internationally to Canada and Mexico. | Terminated | February 12, 2014 | 296 | 11006 | Not_Computer
Z-0611-2014 | 67108 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 840 ventilator systems software part number 4-070212-85, revision AB-AG. |  |  | December 16, 2013 | 2013 | Covidien | Covidien is conducting a voluntary field corrective action on Puritan Bennett 840 ventilator systems with certain software revisions in response to customer reports of ventilator diagnostic code XB0069 in which the device stops mechanical ventilation due to a software error. | Software design | Covidien sent an Urgent Medical Device Voluntary Field Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to continue using their Puritan Bennett 840 ventilators until it is feasible to install the software update discussed in the letter. Covidien has developed a software solution to address this issue. Upon availability, there will be several ways to install the software update. To initiate the process for updating your software and to select the method that best suits your facility, please visit our software update management portal at www.PB840technicalupdate.com and follow the instructions to register your ventilators. If you need further assistance contact our Technical Support Department @ 1-800-255-6774, option 4, then option 1.For questions regarding this recall call 800-255-6774. |  | 56295 | 16,923 US, 39,372 OUS | Worldwide Distribution - USA (nationwide) and Internationally to Austria, Estonia, Belgium, Spain, Italy, Poland, Bulgaria, Finland, Portugal, Switzerland, France, Lithuania, Romania, Cyprus, Great Britain, Czech Republic, Greece, Latvia, Slovenia, Germany, Hungary, Slovak Republic, Ireland, Netherlands, Turkey, Croatia, Andorra, Albania, Algeria, Bahrain, Bosnia and Herzegovina, Egypt, Georgia, Iran, Iraq, Israel, Jordan, Kenya, Kuwait, Lebanon, Libya, Macedonia, Montenegro, Morocco, Russia, Saudi Arabia, Serbia, South Africa, Syria, Sudan, Tunisia, UAE, Uganda, Puerto Rico, Mexico, Columbia, Chile, Brazil, Argentina, Uruguay, Panama, Costa Rica, Peru, China, Korea, Japan, Taiwan, Hong Kong, Malaysia, Singapore, Sri Lanka, Australia, New Zealand. | Terminated | November 29, 2016 | 1079 | 56295 | Computer
Z-0627-2014 | 67093 | No | ODP | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | C-Plus" 10mm x 12mm Webless implants |  |  | October 07, 2013 | 2013 | RTI Surgical, Inc. | As the result of a recent internal review of regulatory documents, we have determined that the 10 mm x 12 mm Webless C-Plus implants were incorrectly documented as meeting the requirement for labeling as a partial Vertebral Body Replacement Device (VBR). While the use of this device as a partial VBR is unlikely due to the size of the device, use of this device as a partial VBR could result in a s | Labeling False and Misleading | RTI Surgical, Inc. sent an Important Notice letter dated October 8, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed remove fall affected product from their inventory and send back to Customers Service at 375 River Park Circle, Marquette, MI 49855 by October 16, 2013. Customers with questions were instructed to call 906-226-4489.For questions regarding this recall call 906-226-4812. |  | 1906 | 1906 | Nationwide Distribution - USA including FL, MO, NY, CA, MI,and IL. | Terminated | February 26, 2014 | 142 | 1906 | Not_Computer
Z-0622-2014 | 67025 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Integra Large Rickham -Style Reservoir |  |  | November 13, 2013 | 2013 | Integra LifeSciences Corp. | Integra became aware through an adverse complaint trend that Rickham Style Reservoirs with bent or broken connectors that were shipped in sleeve-type packages were found prior to use but after the sterile package was open. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall letter dated November 13, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory for the affected product and if located stop using immediately. Customers were asked to complete the attached form and return it to Integra as indicated and keep a copy of the form for their records. Once the form was received by Integra they will send an RMA number, directions to return the product, and input an order to replace the quantity they indicated on the form.Customers with questions were instructed to call 1-855-532-1723.For questions regarding this recall call 855-532-1723. |  | 78 | 78 packages | Worldwide Distribution - USA including FL, NE, WI, WA, CT, ND, AL, NY, NH, OR, OK, MN, and TX. Internationally to Belgium, Czech Republic, Germany, France, Great Britian, Italy, Poland, and Portugal. | Terminated | February 13, 2015 | 457 | 272 | Not_Computer
Z-0629-2014 | 66622 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius 2008 Series Hemodialysis Machines: |  |  | November 12, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | 2008 Series Hemodialysis Machines complaints of saline bags inappropriately filling during recirculation and setup | Under Investigation by firm | Fresenius Medical issued a Product Notification letter dated November 12, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers who encounter a saline bag that is inappropriately filling, should discontinue the setup and remove the machine from service until the drain line and drain-associated hydraulic components can be checed by a qualified service technician. Customers were also instructed to ensure that the hemodialysis machine is installed and operated in compliance with the Machine Specification section of their operator's manual, including the specification located in the Water/Drain Section. Customers with questions were instructed to call 1-800-227-2572.For questions regarding this recall call 800-662-1237. |  | 115142 | 111,504 units US and 3,638 Canada and Mexico- | Worldwide Distribution - USA (nationwide) and Internationally Mexico and Canada. | Open, Classified |  |  | 115142 | Not_Computer
Z-0632-2014 | 67131 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Anspach B-Blue-S 11.25 cm attachment and Anspach B-Green Standard Craniotome. |  |  | June 11, 2012 | 2012 | The Anspach Effort, Inc. | Anspach was contacted by a hospital to inform Anspach that a Blackmax motor had been used in a case where the patient was diagnosed with Creutzfeldt-Jacob Disease (CJD). | Unknown/Undetermined by firm | none |  | 2 | 1 B-Blue-S and 1 B-Green | US Distribution: Maryland only. | Terminated | April 18, 2016 | 1407 | 6 | Not_Computer
Z-0630-2014 | 67040 | No | KCY | Pneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Stryker Color Cuff, REF 5921-034-135 |  |  | December 06, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | The Color CuffÂż Sterile Disposable Tourniquet Cuffs with Quick-Connect Fitting may have an incorrect end-connector.A customer reported having received a luer connector instead of a quick connect fitting. An incorrect connector maylead to an inoperable system due to incompatibility with the Smart Pump. | Component change control | Distributor was notified by E-mail 11/26/2013, with follow-up UPS letters sent to Hospital Risk/ Materials Managers identifying the product, reason for the recall, and risk. The notice asked consignees to find and quarantine all implicated product, with hospitals requested to complete the Business Reply Form and FAX to 866-521-2762. Stryker will email a pre-paid shipper and credit all accounts. Question and concerns should be addressed to Kelly Jo Whipple 269-389-2921 or kellyjo.whipple@stryker.com |  | 5 | 5 boxes (50 units) | US Distribution: including states of: MI, OH, and TX. | Terminated | February 27, 2014 | 83 | 5 | Not_Computer
Z-0633-2014 | 66919 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | The Propaq LT Series monitors. Model numbers 802LTAN, 802LT0N, and 802LTRN. |  |  | November 21, 2013 | 2013 | Welch Allyn Protocol, Inc | Welch Allyn will update Propaq 802 Series Vital Signs Monitor, models are 802LTAN, 802LTON and 802LTR with software because the fault conditions relating to hardware or software failure or electrical noise caused by proximal defibrillation activity could, in rare instances, potentially interrupt device operations. | Device Design | Welch Allyn began sending out the Urgent Medical Device Correction letter, dated December 03, 2013, to consignees on December 10, 2013. The letter identified the product, the problem, and the action to be taken by the customer. Customer Actions: In the unlikely event that operation of your Propaq LT is interrupted and the display becomes blank or white, pressing the power (on/off) button will restart the device and quickly bring it back to full operational capacity.This notice should be passed on to all those who need to be aware within your organization, or to any organization where the potentially affected devices have been transferred.Welch Allyn is in the process of developing a software installation service tool that will enable installation of the software update by your biomed, IT staff or other technically qualified personnel.Upon release, the service tool will be made available to you via a download from Welch Allyn's website at www.welchallyn.com/PropaqltUpgrade or other media.In the meantime and on an on-going basis, the software update will be included if you schedule routine or other service with Welch Allyn or an authorized service provider.Customers with questions were instructed to call Welch Allyn Technical Support at 1-800-535-6663.For questions regarding this recall call 503-530-7935. |  | 6003 | 6003 units total (1715 units in the US and 4288 outside the US) | Worldwide Distribution - USA (nationwide) and Internationally to United Arab Emirates, Argentina, Australia, Belgium, Brazil, Canada, Switzerland, Chile, China, Columbia, Costa Rica, Czech Republic, Germany, Denmark, Ecuador, Spain, France, United Kingdom, Greece, Hong Kong,Taiwan, Ireland, Italy, Japan, Korea, Republic of, Kuwait, Lithuania, Luxemborg, Mexico, Malaysia, Mozambique, Namibia, Netherlands, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Sierra Leone, El Salvador, Thailand, Turkey, Uganda, Vietnam, Kosovo, and South Africa. | Terminated | February 17, 2015 | 453 | 6003 | Computer
Z-0649-2014 | 67150 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Catalys Precision Laser System (United States) - Catalys-U: |  |  | December 20, 2013 | 2013 | Optimedica Corporation | Potential for unexpected heating of the transformer mounting bolt located within the system chassis and the subsequent melting of a non-slip computer shelf mat, which may manifest an odor and visible smoke. | Under Investigation by firm | The firm, OptiMedica, issued an "ADVISORY NOTICE" dated December 20, 2013 to all customers with a Catalys Precision Laser System via Federal Express. OptiMedica also released a Field Service Bulletin (Document No: SB-00065) internally for the impacted product with instructions to implement the corrections through the OptiMedica service organization. The notice describes the product, problem and actions to be taken. The customers were instructed to complete the site header information on the attached fax-back form and return via fax to 1-408-890-4681 within the next three (3) business days or mail to OptiMedica at 1310 Moffette Park Drive, Sunnyvale, CA 94089.If you have any questions related to the operation of the Catalys Precision Laser System or the documentation requirements associated with this letter, please call OptiMedica Technical Support at 1-855-662-0911 and you will be directed to an OptiMedica representative who will be pleased to assist you. |  | 78 | 78 (39 in US, 39 internationally) | Worldwide Distribution: US (nationwide) including states of: TX, NC, OH, IA, UT, KY, ME, CA, NY, MA. NJ, IL, PA,, TN, FL, MO, GA, HI, SC, NE, MI, and GA and countries of: Dominican Republic, Germany, Switzerland, Australia, Spain, Japan, United Kingdom, Canada, France, The Netherlands, Austria, and Singapore. | Terminated | June 26, 2014 | 188 | 78 | Not_Computer
Z-0647-2014 | 67132 | No | CAW | Portable oxygen generator. | Anesthesiology | Anesthesiology | 510(k) | On Site Gas Systems POGS 33C Portable Oxygen Generation System |  |  | December 19, 2013 | 2013 | On Site Gas Systems Inc. | Heat exchanger hose becomes weak and may fail. | Device Design | On Site Gas Systems issued letter on 12/19/13 and stated replacement parts to retrofit the blue Federal Heater hose with stainless steel will be provided and the letter included the Technical Bulletin TB 12-19-13-1 . |  | 198 | 198 units | US Distribution: PA only. | Terminated | October 02, 2014 | 287 | 198 | Not_Computer
Z-0638-2014 | 67005 | No | GDZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sandel Healthcare Safety Devices Weighted Safety Scalpel #3 Disposable Safety Handle with #11 Blade |  |  | November 12, 2013 | 2013 | Ansell Sandel Medical Industries LLC | Select lots of SANDEL Weighted Safety Scalpels are packaged with the scalpel blade exposed beyond the safety shield which could compromise the integrity of the package and the sterility of the scalpel leading to injury. | Under Investigation by firm | The firm, Ansell, sent an "URGENT: VOLUNTARY RECALL OF ANSELL SANDEL WEIGHTED SAFET SCALPEL" letter dated November 11, 2013 to its customers and posted a website notification on November 11, 2013. The letter described the product, problem and actions to be taken. The customers were instructed: Do not ship SANDEL Weighted Safety Scalpel products 2210-L, 2210-N, 2211-L, 2211-N, 2215-L, 2215-N, or 2220-N to a hospital or end user with the listed lot numbers; quarantine SANDEL Weighted Safety Scalpel products in your inventory; and complete the attached response form and return it to attention: Ansell; via fax at 818-534-2510 or email sandel-cs@ansell.com. Additionally, if you have distributed this product to other accounts, please notify them of the voluntary recall and instruct them to return any outstanding stock to you as indicated in the attachment.If you have any questions or need additional information, please contact Customer Service by calling 866-764-3327 (6:00am-5:00pm PST/9:00am-8pm EST) or email sandel-cs@ansell.com |  | 96 | 96 cases (288 units) | Worldwide Distribution: US (nationwide) including states of: AL, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MN, MS, NC, NE, NY, OK, PA, PR, RI, TN, TX and WI; and country of: Canada. | Terminated | April 06, 2016 | 876 | 1521 | Not_Computer
Z-0642-2014 | 67034 | No | MEG | Piston syringe. | General Hospital | General Hospital | 510(k) | Jelco Smiths Medical Hypodermic Needle-ProÂż Fixed Needle Insulin Syringe, 1.0ml 29g x 1/2 in |  |  | July 10, 2013 | 2013 | Smiths Medical ASD, Inc. | Visible orange/ brown particulate on the needles of the 29g Hypodermic Needle-ProÂż Fixed Needle Insulin Syringes | Storage | The firm, Smiths Medical, sent an "URGENT FIELD SAFETY NOTICE" dated July 10, 2013 to its consignees/customers and distributors. The notice described the product, problem and actions to be taken. The Distributors were instructed to notify their customers. The consignees/customers were instructed to inspect your inventory for the suspect lots and remove all affected devices from use and complete and return the Urgent Field Safety Notice Confirmation Form by fax to 603-358-1017 or by email to FNS@smiths-medical.com within five (5) days of receipt of this notice. If you or your facility has distributed these affected products to other persons or facilities, please promptly forward the recipients of this Urgent Field Safety Notice. Upon receipt of the completed form, a customer service representative will contact you to arrange for exchange of your unused affected devices for credit or replacement.If you should have any questions regarding this information, please contact Smiths Medical's Customer Service Department at 800-258-5361 |  | 363015 | 363015 | Worldwide Distribution: US (nationwide) including states of: AZ, CA, FL, IL, LA, MO, MS, NH, NY, OH, PA, TN, TX, VA, WI, and WV; and countries of: Canada, Italy and Spain. | Open, Classified |  |  | 1125640 | Not_Computer
Z-0634-2014 | 67134 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Tracer IV wheelchairs, Heavy duty manual wheel chair, 24" seat width, 450 lbs. weight capacity. Distributed by Invacare, Elyria, OH 44035 |  |  | April 16, 2012 | 2012 | Invacare Corporation | The wheels on Tracer IV wheelchairs (Tracer IV 24" seat width, 450 lbs. weight capacity manual wheel chair) were warped and wobbled when rolled. | Component design/selection | The firm, Invacare, sent a memo dated December 4, 2013 to their customer. The memo described the product, problem, and actions to be taken . On February 8, 2012 representatives from the recalling firm visited the customer to investigate the complaint. The wheelchairs were returned and destroyed.If you have any question, please call 440-329-6875 or email: ssaigal@invacare.com. |  | 227 | 227 Wheelchairs | US Distribution: Michigan only. | Terminated | January 10, 2014 | 634 | 227 | Not_Computer
Z-0677-2014 | 67166 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 14" (36 cm ) Appx 0.50 ml, Smallbore Ext Set w/MicroClaveÂż Clear, |  |  | December 24, 2013 | 2013 | ICU Medical, Inc. | ICU Medical is recalling certain lots of the NanoClave Needlefree Connector product line because of a small percentage of devices that may leak. | Equipment maintenance | icumedical sent an Urgent Medical Device Recall letter dated December 24, 2013, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to return completed verification forms. Customers with questions were instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30AM and 4:00PM Pacific Time: (866) 829-9025 and select option 8, or email: customerservice@icumed.com.ICU Medical reported an update on January 6, 2014, an expansion of the recall to additional lots. Updated recall letter was sent out to all customers on additional lots on January 6, 2014 via UPS with tracking.For questions regarding this recall call 949-366-2183. |  | 100 | 100 units | Worldwide Distribution - USA including UT, TX, GA, WI, VT, NH, AL, OH, TN, IL, CA, KY,HI, NC, and Internationally to CA, BE, IN, KW, and AU. | Terminated | April 01, 2014 | 98 | 147975 | Not_Computer
Z-0651-2014 | 66603 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | The Vantage Digital Panoramic System delivers high-value features standard on every system: - Five pre-programmed exam settings, including bitewing, minimize exam set-up time and allow routine diagnostic exams to be performed extraorally for improved clinical efficiency and patient experience. - Progeny Clarity Enhanced Filter Set provides superior diagnostic imaging quality - Exclusive Focal Trough Alignment reminder assists in consistent patient positioning - Progeny Imaging software included for download on an unlimited number of operatory workstations with no additional costs - Exclusive VantageTrust remote support and training service offered at no charge for the first 60 days of use - Ceph upgradeable for expanded extraoral exam capabilities - The Progeny Vantage Panoramic System s proudly designed and manufactured in the USA; 2013 Midmark Corporation; 675 Heathrow Drive Lincolnshire, IL 60069; 847-415-9800 Toll-free 888-924-3800 |  |  | September 20, 2013 | 2013 | Midmark Corp dba Progeny Inc | An error was identified in the Real Time Controller (RTC) firmware versions 3.0, 3.1, 3.2, and 3.3 used in Vantage Panoramic X-Ray systems. The error may cause the column of the Vantage Panoramic X-Ray unit to continue moving vertically despite the release of the movement control button (i.e. failure to stop"). | Software design | An URGENT MEDICAL DEVICE FIELD CORRECTION letter dated September 20, 2013 was sent to all consignees. Letters sent to distributors include instructions on how to identify the defective units, to quarantine the machine(s), and to contact Midmark at 847-415-9739 or via- e-mail at abavuso@midmark.com. Midmark will either schedule an on-site field correction or arrange for the replacement of the affected device(s) depending on the customer's preference. Midmark further requested the distributors provide Midmark with the names, addresses, phone numbers, and e-mail addresses of any customers that purchased any defective units. Field corrections include the upgrading of RTC firmware to revision 3.4 on the defective units and will be performed as no additional charge by Midmark personnel. Letter sent to practitioners include instructions on how to identify the defective units, instructions on how to temporarily reduce the likelihood of the unintentional movement, and to contact Midmark at 847-415-9739 or via- e-mail at abavuso@midmark.com. Midmark will schedule an on-site field correction by upgrading the firmware to revision 3.4. |  | 58 | 58 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AK, AL, CA, CO, CT, FL, GA, IL, KS, MA, NC, NH, NJ, OH, OK, PA, SC, TN, TX, VA, and WI, and the countries of Australia, Canada, and Russia. | Open, Classified |  |  | 58 | Computer
Z-0650-2014 | 66398 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Nova StatStrip/StatSensor Hospital Blood Glucose Meter; Nova StatStrip Lactate Hospital Meter System; Nova StatSensor Creatinine Hospital Meter, |  |  | October 10, 2013 | 2013 | Nova Biomedical Corporation | Revised Battery Instructions for Removal and Expiration Date due to battery swelling and cell leakage | Component design/selection | On November 18, 2013, an URGENT Field Correction Letter was sent to all StatStrip Glucose Hospital meter, StatStrip Lactate Hospital Meter and StatSensor Creatinine Hospital Meter customers via email to the Point of Care Coordinator to alert users that lithium batteries that are utilized beyond their intended useful life may swell and/or crack posing a potential safety hazard. In some cases, the swelling may potentially cause the battery door on the meter to dislodge and render the meter inoperable. |  | 200000 | est. 200,000 | Nationwide Distribution. | Terminated | March 24, 2014 | 165 | 200000 | Not_Computer
Z-0648-2014 | 67161 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | SOLAR 8000 AND TRANSPORT PRO WITH PATIENT DATA MODULE (Marketed As GE Healthcare CARESCAPE Patient Data Module) |  |  | December 16, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue which affects all Patient Date Modules produced from April 2007 through June 2010, encompassing Fiscal Week 13, 2007 through 26, 2010 due to fluid ingress into the top cover assembly which can result in heat buildup. | Nonconforming Material/Component | Consignees were sent on 12/16/2013 a GE Healthcare "Urgent Medical Device Correction" letter dated December 5, 2013. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Contact Technical Support at 1-800-558-7044 or your local Service Representative). |  | 17444 | 17,444 (11,418 units US, 6,026 units OUS). | Worldwide Distribution: US (nationwide) including states of: AL, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, O, NE, NV,NH, NJ, NM, NY, NC, OH, OK, OR,PA PR, RI,SC, TN, TX, UT, VT, VA, WA, WI, and WY. Internationally to: ARGENTINA, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, REPUBLIC OF KOREA, KUWAIT, LEBANON,LITHUANIA, MALAYSIA, MAURITIUS, MOROCCO, NETHERLANDS, NEW ZEALAND, NORWAY, POLAND, QATAR , SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | April 19, 2016 | 855 | 17444 | Not_Computer
Z-0529-2014 | 66833 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | EOS System X- ray beam |  |  | November 07, 2013 | 2013 | EOS Imaging | EOS imaging discovered during production internal testing that the X-ray beam collimation set up by the operator at the user interface of the acquisition workstation is different from the actual collimation ofthe X-ray beam operated by the EOS system during the X-ray acquisition. | Device Design | CDRH approves the following conditions for EOS Imaging SA:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator notes below what are to be included in the notification. The draft letter provided on November 28, 2013, is deficient for the following reasons:a. It does not provide a statement that all repairs will be made withoutcharge (21 CFR 1003.21(3)).b. It does not provide instructions, with respect to the use of the product pending the correction of the defect (21 CFR 1003.21(1). Specifically, the letter should include directions for the user to determine if the collimation defect is currently present in their system, which would initiate immediate correction, rather than at the next semi-annual service.2. The customer notification letter will be sent in the manner required by 21CFR 1003.21(b).3. The repairs outlined are to be completed by July 1, 2014. All systems under service contracts (which includes semi-annual service) would be visited between receipt of this letter and July 1, 2014. All repairs, including those to systems which are not under service contracts, are also expected to be completed by July 1, 2014.EOS may proceed with implementation of the CAP (subject to the conditions noted above. The customer notification letter may be sent out immediately upon addressing the above deficiencies. However, it is suggested that you provide a copy of the revised letter to FDA for approval prior to issuance.For further questions please call (678) 564-5400. |  | 22 | 22 EOS systems are installed in US | Nationwide Distribution | Terminated | January 27, 2015 | 446 | 22 | Not_Computer
Z-0654-2014 | 66323 | No | HSZ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | ***REF BLACKMAX-N***Swivel/Angle Motor Assembly***Rx Only***The Anspach Effort, Inc. |  |  | April 25, 2013 | 2013 | The Anspach Effort, Inc. | The firm Anspach Effort of Palm Beach Gardens, FL initiated a voluntary removal of the Pneumatic Drill System, including the motor hand-piece, Auto Lube foot control, without a pressure relief valve (PRV) due to the hand piece's outer hose can rupture without warning if the air flow is blocked during use.. | Component design/selection | The firm, Anspach sent an "URGENT" Medical Devics Removal" letter dated April 24, 2013, to its customers. Customer Immediate Actions: 1. Screen your current inventory and return all Anspach Pneumatic Drill System without a pressure Relief valve (PRV) from your inventory. 2. If you have any product listed on the attached form to arrange for immediate return and replacements. 3. complete the attached Customer Reply form indicating the product being returned. 4. Upon receipt of a replacement product, equipment being returned must be received by Anspach within 10 business days of the replacement request. Should you have any queries, please contact Anspach Product Support at (800) 327-6887. |  | 2785 | 2,785 | Worldwide Distribution: US (nationwide) including Puerto Rico and states of: AL, AR, AZ, CA, CO, CT, DC, DE, GA, FL, IA, ID, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, VE, VT, WA, WI, WY, and WV. Internationally to: AT, AU, BR, BS, CA, CH, CL, CO, CZ, CR, DE, DK, DO, EG, ES, FI, FR, GB, GR, HK, HU, ID, IL, IN, IS, JP, KH, KR, LB, LT, MX, MY, NL, NO, NZ, PA, PE, PH, PL, PT, SA, SG, SY, TH, TR, TW, and ZA. | Terminated | July 20, 2015 | 816 | 4997 | Not_Computer
Z-0708-2014 | 67124 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Multiples Anspach Micro Curved Attachments (MCA) and Curved Burr Support S leves (Cutters and Attachments) |  |  | October 26, 2012 | 2012 | The Anspach Effort, Inc. | During a recent review of documentation Anspach determined that the current directions for use for the MCA cutting burrs and MCA Micro Curved Attachment do not provide clear details regarding the intended use specifically for the MCA burrs. The MCA burrs are intended for use in delicate bone in otologic procedures such as cochleostomy. The use of the device in dense bone removal may result in cu | Labeling False and Misleading | Anspach sent an Urgent Medical Removal letter dated October 26, 2012, to all affected customers. The letter stated that they were initiating a Medical Device removal regarding the MCA Cutting Burrs. During a recent review of documentation it was determined that the current directions for use for the MCA cutting burrs (item 18-0079) and MCA Micro Curved Attachment (item 18-0060) did not provide clear details regarding the intended use specifically for the MCA burrs. The MCA burrs are intended for use in delicate bone in otologic procedures such as cochleostomy. The use of the device in dense bone removal may result in cutter/ burr fracture. If a fracture occurs, contact Customer Support at (800) 327-6887 or email ProductSupportTeam@synthes.com. If the procedure was done successfully, no additional patient follow up is necessary.Customer Immediate Actions:1. Please remove and return any MCA items from your inventory immediately.2. Complete the attached reply form indicating your receipt of this letter. Return the completed form by fax or email to the number or email address provided on the form. If you distribute any of the products to other services or facilities, please forward this information as appropriate. Customers with questions were instructed to contact Anspach Customer Support at (800) 327-6887 or email ProductSupportTeam@synthes.com.For questions regarding this recall call 1-800-327-6887. |  | 14330400 | 6,824 imits (4,7324 in USA / 2,100 Foreign) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Netherlands, Great Britain, France, Sweeden, Italy, Spain, Japan, South Africa, Australia, Switzerland, | Terminated | June 23, 2014 | 605 | 14330400 | Not_Computer
Z-0652-2014 | 67087 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Arthrocare SPORTS MEDICINE, Knot Pusher, REF 25-3011, LOT #1049402, CON 1, 2013-10, NON-STERILE, Rx only |  |  | December 09, 2013 | 2013 | ArthroCare Medical Corporation | A sharp edge on the cannulated tip of the device may damage or cut a suture. | Nonconforming Material/Component | All consignees were provided with written notice via Urgent Recall Notification dated December 9, 2013 with delivery confirmation. In addition, the firm plans on notifying the EU, and the governments of France, Germany, Switzerland, and Spain. The firm will also notify the Swedish government because the firm's EU representative is located in Sweden. |  | 76 | 76 | Worldwide distribution: US (nationwide) including states of: CA, FL, GA, IL, IN, KY, MN, MO, NY, OH, TX, and WI. Internationally to: Canada, France, Germany, Spain, and Switzerland. | Terminated | April 08, 2015 | 485 | 76 | Not_Computer
Z-0653-2014 | 66870 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | Product Brand Name(s): |  |  | June 03, 2013 | 2013 | Alere San Diego, Inc. | Alere initiated this recall of certain lots of the Oral Fluid Drug Screen Device because the sponge may become dislodged from the Saliva Collector handle prior to or during the oral fluid sample collection process. If the sponge becomes dislodged during the specimen collection process, the donor may inadvertently swallow the sponge which could present a choking risk. | Other | Firm sent customers notification letter via combination of fax, email or direct mail. The notification letter dated June 3, 2013, titled "URGENT PRODUCT SAFETY NOTICE", informed customers of the recall, reason for recall, product description with lot codes, potential hazard statement, customer required actions, and contact information. |  | 280710 | 280,710 | US Distribution: including states of: AL, CA, PA, and VA. | Terminated | March 18, 2014 | 288 | 280710 | Not_Computer
Z-0717-2014 | 66030 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Leica BIOSYSTEMS; BOND Polymer Refine Red Detection; Catalog No: DS9390; Leica Biosystems Newcastle Ltd; Balliol Business Park West; Benton Lane; Newcastle Upon Tyne NE12 8EW; United Kingdom; www.LeicaBiosystems.com; Bond Polymer Refine Red Detection is a biotin-free, polymeric alkaline phosphate (AP)-linker antibody conjugate system for the detection of tissue-bound mouse and rabbit IgG and some mouse IgM primary antibodies. |  |  | July 31, 2013 | 2013 | Leica Biosystems Newcastle Ltd. | The affected lots of the product may contain a contaminant of fungus that can cause a loss of staining sensitivity and/or obstructions to fluid delivery to the slide. | Process control | Leica Biosystems sent an Urgent Field Safety Correction Notice dated July 31, 2013, to all affected customers. The notice requests that direct accounts either destroy any unused reagents or confirm that the product has been used and in no longer in stock. Direct accounts are requested to sign the notice and either fax or scan and e-mail it back to Leica Biosystems Newcastle Ltd. The notice further requests that direct accounts who further distributed the recalled products forward the notification to their customers.Customers with questions were instructed to contact their Leica representataive if a replacement is necessary, for any questions customers were instructed call 847-405-5413 or microsystems.com.For questions regarding this recall call 847-405-5413. |  | 2344 | 1) 2,304 units; 2) 37 units | Worldwide Distribution - USA including AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, and WV); and Internationally to Chile, Denmark, Finland, France, Germany, Ireland, Italy, Netherlands, Sweden, Switzerland, Poland, Romania, Singapore, South Korea, and the United Kingdom. | Terminated | August 14, 2017 | 1475 | 2344 | Not_Computer
Z-0716-2014 | 67142 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Locking Cortical Screw |  |  | December 19, 2013 | 2013 | Biomet, Inc. | 3.5mmx38mm Locking Cortical Screw, implants are being recalled since an investigation identifiedthat a 3.5mmx28mm Cortical Locking Screw is labeled as a 3.5mmx38mm Cortical Locking Screw. | Labeling mix-ups | URGENT MEDICAL DEVICE RECALL notices were sent to all consignees on December 19, 2013. Distributors who have affected hospitals in their regions were notified and provided with a copy of the notice that went to each hospital. Consignees are being contacted on an individual basis in an attempt to reconcile the product 100%. Fax back forms were provided to all consignees in order for Biomet to track returns. Non-responders will be contacted on an individual basis. |  | 52 | 52 | US Distribution including states of: AZ, CA, GA, IL, KY, MI, NC, NH, NY, OH, PA, SD, TX, UT, and VA. | Terminated | July 07, 2015 | 565 | 52 | Not_Computer
Z-0714-2014 | 67056 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Draeger Infinity Acute Care System Monitoring Solution |  |  | December 10, 2013 | 2013 | Draeger Medical Systems, Inc. | After a test in which a low SpO2 (yellow) medium grade alarm initiated in neo-natal mode and was subsequently paused, the high grade life threatening SpO2 alarm (red) did not activate when the SpO2 values were decreased to reach the red alarm limit. | Software design | An Urgent Medical Device Recall Letter, dated December 2013, was sent to Users. Once available Drager will update components of the IACS Monitoring Solution System. |  | 439 | 439 | US Distribution including the states of: MA, NH, IN, CA, TX, NY, OK, and IL. | Terminated | May 20, 2015 | 526 | 439 | Not_Computer
Z-0715-2014 | 67177 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Anspach MCA Cutting Burr, DFU Rx Only |  |  | March 29, 2012 | 2012 | The Anspach Effort, Inc. | Current directions for use for the MCA cutting burs (item 18-0001) do not provide enough details regarding the intended use specifically for the MCA burrs. | Error in labeling | Anspach sent an Urgent Medical Device Labeling Correction letter dated Marach 29,, 2012, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to include the item 18-0079 with their inventory immediately, complete the attached reply form indicating their receipt of the letter. Return the completed form by fax to the number or email provided on the form. If product was further distributed to other facilities, the correction letter should be forwarded to them. Customers with quesitons should call 1-800-327-6887 or email customer.support@synthes.com or ProductSupportTeam@Synthes.com.For questions regarding this recall call 561-627-1020. |  | 492 | 492 (321 US) | Worldwide Distribution - USA (nationwide) and Internationally to Great Britain, Australia, Switzerland, Sweden, France, Japan, Spain, Italy, Canada, and Netherlands. | Terminated | April 15, 2016 | 1478 | 492 | Not_Computer
Z-0709-2014 | 66869 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | King Systems |  |  | December 13, 0013 | 13 | King Systems Corp. | KLTD214 airways found inside KLTSD414 packages. The KLTSD414 airway has a lumen for use of a gastric suction tube; the KLTD214 airway does not include a lumen for a gastric suction tube. | Labeling Change Control | December 13, 2013 Urgent Recall Notices were sent to all Direct Accounts identifying the recalled item, risks associated with the incorrectly labeled product and asking customers to stop use, quarantine and destroy any recalled product on hand. Customers were asked to froward the notifcation if the product was further distributed and FAX the completed recall acknowledgement form to 317-776-5175 or email wnethercutt@kingsystems.com. Call the customer service representative (800.262.8673) to arrange for credit of your affected product(s). |  | 108 | 108 cases( 540 products) | Distributed in the states of CA, NY | Terminated | May 15, 2014 | 730638 | 108 | Not_Computer
Z-0712-2014 | 67139 | [6 empty] | Boston Scientific Extractor Pro XL Retrieval Balloon Catheter 12-15 mm |  |  | December 18, 2013 | 2013 | Boston Scientific Corporation | Extractor Pro XL Retrieval Balloon Catheters may have the incorrect skive hole location or balloon size | Process control | Boston Scientific issued an Urgent Medical Device Recall Removal letter dated December 18, 2013, to all affected customers via Federal Express Priority mail.3. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product from their inventory.For questions regarding this recall call 508-382-9555. |  | 34 | 34 units | Worldwide Distribution - USA (Nationwide Distribution and Puerto Rico) and Internationally. | Terminated | August 19, 2014 | 244 | 111 | Not_Computer
Z-0713-2014 | 67044 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | pO2 D999 membrane units. Distributed by Radiometer America, Inc., Westlake, OH 44145 |  |  | June 10, 2013 | 2013 | Radiometer America Inc | RADIOMETER became aware that the regarded lot pO2 units 942-042R0971 were produced with inner solution for pCO2 membrane units. This solution is a salt solution similar to the pO2 inner solution, but also contains glycerol which can be reduced at the cathode causing a suddenly rising zero current, causing the membrane unit to fail the following Cal2 calibration. | Mixed-up of materials/components | Radiometer sent a Product Notification letter dated June 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their inventory and remove any affected product, discard all affected product in their possession, and complete page 2 of the letter and return to their Radiometer representative. Radiometer will replace any quantity of affected product they have at their location. Customers with questions were instructed to contact Radiometer America Technicial Support at 1-800-736-0600, opt 4.For questions regarding this recall call 440-871-8900, ext 209. |  | 1254 | 209 boxes/6 membranes | Worldwide Distribution - US (nationwide) including states of: CO, CT, GA, MA, MO &amp; MS and country of: Canada. | Terminated | June 19, 2014 | 374 | 1254 | Not_Computer
Z-0725-2014 | 67136 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Optima XR220amx and Optima XR200amx with Digital Upgrade Mobile Radiographic Systems. Model numbers 5555000-5, 5555000-6, and 5421698. |  |  | November 07, 2013 | 2013 | GE Healthcare, LLC | There exists a software issue associated with patient selection from the worklist on the Optima XR200 with Digital Upgrade, and Optima XR220amx Mobile X-ray Systems. Intermittently when a large number of patient exams are retrieved from the HIS/RIS system, the exams when displayed on the mobile x-ray system are not properly displayed. If the user continues to attempt to select one of the exams w | Software design | GE sent an Urgent Medical Device Correction letter, dated 11/7/2013 to consignees. The letter described the issue, identified devices affected and stated that users should pay close attention to the patient exam/entry selected from the worklist prior to selecting the Start Exam button to ensure that the system did not update and select an incorrect exam. The user may also confirm during the exam that the appropriate patient information was selected via the display on the upper left portion. GE will fix the units with a software fix at no cost. |  | 1027 | 1027 | Worldwide Distribution-USA (nationwide) including Washington D.C., and Guam and the countries of ALBANIA, AUSTRALIA, BAHRAIN, BELGIUM, BRAZIL, CANADA, CHILE, CHINA , COLOMBIA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HUNGARY, INDIA, IRAQ, IRELAND, ITALY, JAPAN, JORDAN, KOREA, REPUBLIC OF, KUWAIT, MALAYSIA, MEXICO, NEW ZEALAND, NIGERIA, NORWAY, POLAND, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, THAILAND, TRINIDAD AND TOBAGO, TURKEY, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | May 06, 2014 | 180 | 1027 | Computer
Z-0718-2014 | 63071 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Anspach Single Use Sterile Bone Cutting Burs and Irrigation Clips. Rx Only |  |  | September 09, 2011 | 2011 | The Anspach Effort, Inc. | Recall was initiated due to the observation of separation between clear layers of the pouch during real time aging studies. As a result, it is possible to have a breach in the sterile barrier. | Packaging change control | Recalling firm notified direct accounts by letter on 09/09/11. Direct accounts were instructed to screen their inventory immediately for product packaged with the Tyvek Pouch/Tube Configuration, complete the reply form, and return the products to the firm. |  | 1860048 | 1,805,022 cutting burs and 55,026 irrigation clips | Worldwide Distribution - USA (nationwide) and the countries of Japan, Canada, Nigeria, Guam, China, Czech Republic, Thialand, Ecuador, Lithuania, Argentina, Turkey, Cayman Island, Jamaica, India, Jordan, Germany, Brazil, Mexico, Israel, Cameroon, Switzerland, South Africa, Greece, Finland, Taiwan, Hong Kong, France, Spain, Cyprus, Australia, United Kingdom, Sweden, Bahama, Korea, Malaysia, Iceland, Portugal, Polan, Austria, Belgium, Egypt, Singapore, Netherland, Arab Emirates, Kenya, Italy, Venezula, Columbia, Dominican Republic, Chile, Panama, Honduras, Costa Rica, Philippines. | Terminated | April 15, 2016 | 1680 | 1860048 | Not_Computer
Z-0719-2014 | 66907 | No | FGX | Manual gastroenterology-urology surgical instrument and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | TWISTER PLUS 22mm Rotatable |  |  | November 20, 2013 | 2013 | Boston Scientific Corporation | Sterility of device may be compromised due to lack of pouch seal integrity | Packaging process control | Boston Scientific issued Customer Notification letters to US accounts via Federal Express Priority mail on November 20, 2013. The notification instructs the account to check their inventory, remove any affected product from their inventory, and return it to Boston Scientific. The notification also requests the completion of a Reply Verification Tracking Form. |  |  | 0 eaches | Worldwide Distribution-USA (nationwide) and the countries of Canada, Austria, Australia, New Zealand, France, Spain, Portugal, Italy, Switzerland, Denmark, and Netherlands. | Terminated | February 24, 2015 | 461 | 5338 | Not_Computer
Z-0730-2014 | 66990 | No | MBB | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | PalacosÂż R+G Radioplaque Bone Cement 1 x 40 g Single (with Gentamicin), 40.8 g Methyl Acrylate Copolymer, 20 ml Methyl Methacrylate Monomer 0.5 g Gentamicin, Distributed by Zimmer Dover, OH 44522 |  |  | November 04, 2013 | 2013 | Zimmer, Inc. | The patient label associated with the product may be incorrect in that the lot number on the patient label may be different than the lot number on the outer carton label. | Other | Zimmer sent an Urgent Device Removal letter dated November 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification, idetntify the product and remove from use. If the affected product has been utilized, follow institutional protocol to determine if it is appropriate to place a copy of the notification letter into patient records for the affected product. Complete the attached Response Form and return it via fax to 800-871-7270. Customers were instructed to return a copy of the completed response form along with their returned product to ensure proper credit. Customers with questions were instructed to call 866-608-3941.For questions regarding this recall call 800-613-6131. |  | 1534 | 1,534 units | Nationwide Distribution including AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, NC, NJ, NV, NY, OH, OK, OR, PA, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY. | Terminated | October 08, 2014 | 338 | 1534 | Not_Computer
Z-0729-2014 | 67146 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Anspach Carbide Cutting Burrs &amp; Drums: |  |  | December 17, 2013 | 2013 | The Anspach Effort, Inc. | Some carbide cutting burr fractured during use. | Process design | The firm, Anspach, sent an "URGENT: MEDICAL DEVICE REMOVAL" letter dated December 19, 2013 to their costumers. The letter described the product, problem and actions to be taken. The customers were instructed to complete the attached reply form confirming your receipt of this letter and return the completed form by Fax to: 1-800-327-6661 or Email: customer.support@synthes.com. Reference RMA 200059216 when returning any devices. If you distribute(d any of the products to other services or facilities, please forward this information as appropriate . Should you have any queries please do not hesitate to contact Anspach Customer Support at 800-327-6887 or e-mail at customer.support@synthes.com. |  | 28562 | 28,562 products for the entire recall | Worldwide distribution: US (Nationwide) including state of Hawaii and Internationally to: Australia, Poland, Canada, South Africa, Israel, and Switzerland. | Open, Classified |  |  | 28562 | Not_Computer
Z-0737-2014 | 66788 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Atrium Ocean Water Seal Chest Drain ATS (Autotransfusion) Blood Recovery System with stopcock Blood Recovery System; Product Code: 2050-300. |  |  | November 19, 2013 | 2013 | Atrium Medical Corporation | Chest Drain tubing of of the ATS Blood recovery may leak or disconnect | Device Design | Atrium Medical issued notification letters to direct consignees by UPS next day on 19 November 2013, and indirect consignees (distributors customers) by UPS next day or USPS priority (for customers with a PO Box) on 27 November 2013. The Letter explain the issue and the steps that should be taken. The customer will be offered replacement with unaffected product or credit. The firm describes Procedures for Emergency Use. The firm identifies potential product substitutions. These products are the same drainage product without the auto-transfusion line.International Distributors. International distributors provided with the Recall Authorization' package and the customer letter.Accounts are requested to complete the Recall Reply Form acknowledging receipt of the notification.Contact Atrium Medical Customer Service at 1-800-370-7899, Monday through Friday between 9:00 am to 5:00 pm. |  | 6684 | 6,684 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Canada, Austria, Ireland, Romania, South Africa, Colombia, Lebanon Netherlands, Jordan, Russia, and Taiwan. | Terminated | June 11, 2015 | 569 | 71990 | Not_Computer
Z-0747-2014 | 67115 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Carescape Patient Data Module |  |  | November 27, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to intermittent failure of the spring loadedmounting latch of PDMs produced from April 2007 through December 2012. | Device Design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated November 26, 2013 to its consignees/customers. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The consignees/customers were instructed to please ensure that all potential users in your facility are made aware of this safety notification and the recommended actions. GE Healthcare will provide a new latch at no charge once it is available. GE will contact you to arrange for installation of the latch. If you have any questions regarding this notification, please contact Technical Support at 1-800-558-7044, or your local Service Representative. |  | 43033 | 43,033 (25,681 USA; 17,352 OUS) | Worldwide Distribution: US (nationwide) including states of: AL, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, IA, KS,KY, LA, ME, MD, MA, MI, MN, MS, NO, NE, NV, NJ, NM, NY, NC, OH, OK,OR, PA, RI, SC, TN, TX, UT,VT, VA, WA, WI, and WY. Internationally to: Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Czech Republic, Denmark, Egypt,England, Estonia, Finland,France, Germany,Greece, Guatemala, Hungary,India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Korea, Kuwait, Lebanon, Lithuania, Malaysia, Mauritius, Mexico, Morocco, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Poland, Portugal, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Republic of China, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 43033 | Not_Computer
Z-0746-2014 | 67043 | No | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | STERRAD 100S Sterilization System, Product Code: 10101 |  |  | December 16, 2013 | 2013 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the STERRAD NX, STERRAD 50, and STERRAD 100S Sterilization Systems because some healthcare workers detected odors or smells possibly emanating from STERRAD Systems. | Under Investigation by firm | Advanced Sterilization Products (ASP) sent an Urgent Medical Device Field Safety Notification dated December 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return the customer response forms. Customers with questions were also instructed to contact ASP Professional Services directly at (888) 783-7723 and select option number 2.For questions regarding this recall call 949-453-6400. |  | 2366 | 2366 units | Nationwide Distribution | Terminated | March 27, 2014 | 101 | 8183 | Not_Computer
Z-0733-2014 | 66202 | No | CJY | Calcium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Consult DiagnosticsÂż Eon Calcium Reagent For In-Vitro Diagnostic Use |  |  | August 29, 2013 | 2013 | Vital Diagnostics, Inc. | High positive calcium bias on plasma sample versus results on serum | Device Design | Vital Diagnostics issued a Recall Notice dated August 29, 2013, to all affected customers via certified mail. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use plasma for the determination of total calcium when using the affected product. The use of serum is acceptable, thus the product can still be used to determine total calcium with this sample type. Vital Diagnostics released a corection to the package insert for the Envoy Calcium Reagent. Customers with the affected product in their laboratory, were instructed to discard the insert packed with the product and replace it with the insert enclosed with this bulletin. If product was transferred to another location, the bulletin should be forwarded to that location. The package insert was revised indicating to only use Serum samples when running Calcium. Vital Diagnostics issued a revised (expanded) notice dated September 19, 2013, to include the Consult DiagnosticsÂż Eon Calcium Reagent (PN: 1429) requesting users to discontinue using the single vial Eon Calcium kit (PN: 1429). Customers with questions were instructed to call Vital Diagnostics Technical Support at 1-855-354-8324..For questions regarding this recall call 401-642-8400. |  | 213 | 213 units | Nationwide Distribution | Terminated | August 26, 2015 | 727 | 608 | Not_Computer
Z-0734-2014 | 67099 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Testing Disc, MEM10, REF: CT0774B, 10 mcg meropenem, Packed in packs containing 5 cartridges per pack, 50 discs per cartridge. Firm name on the label is Oxoid, Ltd., Basingstoke, United Kingdom. |  |  | November 21, 2013 | 2013 | Remel Inc | Individual discs in the lot may not be sufficiently impregnated with antibiotic, resulting in smaller than expected zone sizes. | Unknown/Undetermined by firm | Oxoid and Remel part of Thermo Fisher Scientific sent an Urgent Medical Device Recall letter dated November 20, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.The letter asked customers to destroy the affected product, if any was found on hand. The letter also asked if the customer further distributed the product that they inform their downstream customers. A response form was included with the letter and the letter instructed customers to complete and return the response form to indicate they had received the notification and that they had complied with it.If you have any questions please contact our Technical Service Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 225 | 45/5/50 - discs/cartridges/packs | Nationwide Distribution including the states of AL, AZ, CA, CO, FL, HI, IA, IL, KY, LA, MD, MI, MN, MO, MS, NC, NY, OH, SC, TN, TX, VA, WI, and WV | Terminated | December 11, 2014 | 385 | 225 | Not_Computer
Z-0731-2014 | 64724 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Centricity PACS System versions 3.X and higher and 4.X and higher |  |  | September 15, 2012 | 2012 | Ge Healthcare It | There is a potential loss of image(s) associated with the Centricity to Centricity (C2C) exam transfer module of Centricity PACS System versions 3.X and higher and 4.X and higher when transferring from the source to a destination server. When another process in the destination server attempts to access the same object or table, the transfer process of a particular image may be terminated. Once ter | Software design | GE Healthcare sent an Important Product Information letter dated September 15, 2012 to affected customers. The letter identified the affected product problem and actions to be taken. GE Healthcare will provide a patch to 3.2.X and 4.0 Systems to remediate the issue. For questions contact your local GEHC-IT service representative. |  | 748 | 748 units | Worldwide Distribution - US: Nationwide (AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WV, WY); FOREIGN: Australia, Austria, Belgium, Bermuda, China, Denmark, Egypt, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Korea, Kuwait, Malaysia, Malta, Netherlands, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, United Kingdom. | Open, Classified |  |  | 748 | Computer
Z-0748-2014 | 65366 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Merge Mammo is a multi-modality, vendor-neutral digital mammography workstation that enables imaging centers to use a single workstation to display and read images from different vendors and different acquisition devices, including digital mammography, ultrasound, MRI and CT. |  |  | August 26, 2013 | 2013 | GE Healthcare It | There is a potential safety issue with the 8.0.2 version software of the Merge (Cedara) Mammo viewer. The Merge Mammo viewer can load and display images not in synch with the selected or highlighted images in the thumbnail navigator. This incorrect image display issue could result in an inaccurate reading on the Merge Mammo workstation. This could result in a missed intervention or mis-diagnosis. | Software design | GE Healthcare sent an URGENT MEDICAL DEVICE CORRECTION letter dated August 26, 2013, to all affected customers. The letter includes important safety instructions for healthcare providers to refer to that can minimize the chance that an incorrect image will be used in the diagnosis or treatment of a patient. The letter further states that a GE Healthcare service representative will contact the customer to schedule a time to install a software update to version 8.0.2 that will correct the software defect.. Customers with any concerns about this defect or recall are encouraged to contact their local GEHC-IT service representative or the GE Remote Online Center (ROC) Customer Care Center at 1-800-437-1171.For questions regarding this recall call 800-437-1171. |  | 28 | 28 units | Worldwide Distribution - USA including AK, AR, CA, IA, ID, IL, MI, MS, MT, NJ, NV, NY, OR, PA, SC, TN, WI and Internationally to Israel, Mexico, and Spain. | Open, Classified |  |  | 28 | Computer
Z-0749-2014 | 63624 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Micro Curved Attachment (MCA) Cutting Burr, DFU 18-0001. All lots are included in recall. Part Number MCA |  |  | March 29, 2012 | 2012 | The Anspach Effort, Inc. | The Anspach Effort, Inc. is recalling all lots of the Anspach Micro Curved Attachment (MCA) Cutting Burrs and Dissection Tools. The recall was initiated after the indications for use were found to be broader than the validation of the MCA Cutting Burrs and its directions for use did not clearly state the correct intended use of the device. After a labeling correction, Anspach Effort, Inc. decide | Labeling Change Control | Firm is performing a Medical Device Recall by sending customers a recall letter and requesting the return of Anspach Micro Curved Attachments and Dissection Tools currently on hand. |  | 163 | 163 | Worldwide Distribution: US (nationwide) and Internationally to: AU, BE, CA, CH, ES, FR, GB, IT, JP, NL, SE, and ZA. | Terminated | March 11, 2015 | 1077 | 1412 | Not_Computer
Z-0773-2014 | 64740 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Imagecast PACS is an image management system intended to be used by trained professionals, including but not limited to physicians, nurses and medical technicians. The system is used with general purpose computing hardware to acquire, transmit, process and store images and data throughout a clinical environment. Data and images are acquired through DICOM compliant imaging devices and modalities. Lossy compressed mammographic images and digitized film screen images must not be reviewed for primary image and interpretations. Mammongraphic may only be interpreted using an FDA approved monitor that offers at least 5 Mpixel resolution and meets other technical specifications reviewed and accepted by the FDA. |  |  | March 04, 2011 | 2011 | Ge Healthcare It | In Imagecast PACS versions prior to 3.6.51.01, the estimated magnification factor (0018,1114) was not accounted for. These versions of Imagecast PACS look for values to be present in the following tags to calculate a measurement from pixels to millimeters (mm): (0020,0030) Pixel Spacing and, (0018,1164) Imager Pixel Spacing. Because the estimated magnification factor is not used, distance meas | Software design | Important Product Information Letters dated March 4, 2011 were sent to all customers that received the defective software. The letters advised customers to upgrade their PACS software version to Imagecast 3.6. The software upgrade can be requested from GE Support via the eService portal or by telephone at 1-800-437-1171. If the customer elects to not upgrade the software, instructions were provided for a new system calibration procedure. |  | 43 | 43 units | Nationwide Distribution: US including states of: AL, AZ, CT, FL, GA, IA, IN, LA, MA, MI, MN, MO, NC, NH, NJ, NY, OH, PA, TX, VA, WI, WV + District of Columbia. | Open, Classified |  |  | 43 | Computer
Z-0770-2014 | 67129 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Weck, Electrosurgical Coagulation Suction Tube 6 (15.2cm), 11 French, with Stylet, Teleflex Medical, Research Triangle, NC 2770. |  |  | December 19, 2013 | 2013 | Teleflex Medical | Sterility cannot be guaranteed. | Nonconforming Material/Component | Consignees were notified via letter on 12/19/2013. They were instructed to, immediately discontinue use and quarantine any products with the catalog numbers listed above. Consignees are to return all affected product to Teleflex Medical per the instructions on the Urgent Recall Notice. If they have no affected stock, they were to complete the enclosed Recall Acknowledgment Form and fax it to Teleflex Medical Customer Service. Distributors were instructed to communicate this recall to any of their customers who have received product. |  | 17450 | 17,450 ea. | Nationwide Distribution. | Open, Classified |  |  | 17450 | Not_Computer
Z-0772-2014 | 67098 | No | FGY | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BD Interlink threaded lock cannula |  |  | November 26, 2013 | 2013 | Becton Dickinson &amp; Company | BD Interlink Threaded Lock Cannula REF# 303369 and lot #3021392 is being recalled due to the potential that a small percentage of the lot may exhibit open seals which adversely impacts unit package integrity and product sterility. | Packaging | BD Diagnostics sent an Urgent: Voluntary Product Recall letters/Notice of Return dated November 26, 2013, to all affected customers. Customers were asked to identify and segregate stock of the affected number, inform the appropriate personnel in their organization to discontinue use of the affected product, return the afected product, following the instructions on the enclosed packing slip, BD would issue a credit for product returned. If product was further distributed customers were instructed to identify their customers and notify them at once of the product recall using the provided customer communication. Customers were asked to complete the attached notice and fax to 855-731-0157. Customers with questions were instructed to call 1-855-731-0157.For questions regarding this recall call 755-731-0157. |  | 268500 | 268,500 products | Worldwide Distribution - USA (nationwide) and Internationally to Canada and Japan. | Open, Classified |  |  | 268500 | Not_Computer
Z-0775-2014 | 67065 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | Estrogen Receptor AB-11 (Clone ID5) |  |  | December 06, 2013 | 2013 | Lab Vision Corporation | Reports that the product is not performing as expected. In house testing showed no staining on breast cancer tissues. | Under Investigation by firm | Lab Vision Corporation part of Thermo Fisher Scientific sent an Urgent Medical Device Recall letter dated December 6, 2013, to all affected customers. Revised letters will be sent on January 8, 2014. Changes include that the reagent is an Analyte Specific Reagent (ASR) and not an in vitro diagnostic device in the US.They also clarify that a false positive was reported by a customer and a false negative was confirmed during in house testing. Letters also correct the expiry date from 12/2013 to 03/2015. Data sheet was also updated to eliminate information stating that the presence of ER in breast tumors indicates an increased likelihood of response to estrogen therapy and also that the positive control must be determined by the investigator.If you have any questions, please contact Technical Service at 1-800-522-7270, option 2, option 7. |  | 5 | 5 | Worldwide Distribution: US (nationwide) including the states of TX and MA and internationally to: Turkey. | Terminated | February 25, 2014 | 81 | 5 | Not_Computer
Z-0776-2014 | 66298 | No | HTE | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Arena-L Trial Handle; Part: 11-81002 |  |  | September 24, 2013 | 2013 | SpineFrontier, Inc. | Handles and heads may not be compatible. | Other | The firm, SpineFrontier, sent an "URGENT Arena-L Trials Voluntary Recall Notification" letter dated September 24, 2013 sent to its consignee/customer. The letter described the product, problem and actions to be taken. The consignee/customer were instructed to confirm receipt of the notice and contact Customer Service to initiate the immediate return of the listed parts. In addition, consignee/customer were instructed to complete a recall ACKNOWLEDGEMENT FORM and return it to SpineFrontier via fax to: Attention: Customer Service at 978-232-3991 or email to: customerservice@spinefrontier.com.Please feel free to contact us at 978-232-3990 Monday through Friday 8:30AM to 5:30PM EST with any questions. |  | 23 | 23 devices in total | US Distribution: VA only. | Completed |  |  | 23 | Not_Computer
Z-0779-2014 | 67179 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Anspach MIA 16 (16cm Minimally Invasive Attachment), Rx Only |  |  | May 06, 2013 | 2013 | The Anspach Effort, Inc. | Product did not reach expected sterility assurance level | Unknown/Undetermined by firm | The firm, Aspach, sent an "URGENT: MEDICAL DEVICE LABELING CORRECTION" letter dated May 13, 2013 to its consignees/customers. The letter described the product, problem and actions to be taken. The consignees/ customers were instructed that Sterrad 100S system should not be used on the MIA 16 attachment (the information only applies to only Sterrad sterilization of this devices and not to any other attachments.), and complete and return the Return Receipt Requested Customer Reply form via fax to: 1-800-327-6661 or email: ProductSupportTeam@synthes.com. If you distribute any of the products to other services or facilities, please forward this information as appropriate.Should you have any queries, please do not hesitated to contact Anspach Product Support at (800) 327-6887. Hours of Operation are 8:30am-5:30pm Eastern Standard Time or email: ProductSupportTeam@synthes.com. |  | 183 | 183 | Nationwide Distribution: FL, IL, CA. IN, MI, CA, NJ, TX, NE, NC, OR, PA, TN, DC, UT, CO, MO, KY, MD, WV, NY, and MN. | Terminated | April 18, 2016 | 1078 | 183 | Not_Computer
Z-0782-2014 | 67123 | No | JEQ | Powered toothbrush. | Dental | Dental | 510(K) Exempt | UP &amp; UP kids' power toothbrush |  |  | December 18, 2013 | 2013 | Ranir Corporation | Ranir, LLC is recalling all lots of UP &amp; UP brand Kid's Power Toothbrush and TopCare brand Clifford Kids Power Toothbrush because the tufted disk may become detached from the power brush handle. If thetuft retention disk detaches from the handle, it is a potential choking hazard for children three years old and younger. | Device Design | Ranir, LLC phoned their two direct accounts notifying them of this recall on December 12, 2013. On December 18, 2013, Ranir, LLC sent an URGENT: MEDICAL DEVICE RECALL notification letter was sent to their direct accounts. The notification stated the affected product, issue, and instructions for responding to the notification, including instructions for return of the affected product. Questions are directed to Paula Bojsen (616)698-8880. |  | 601616 | 601,616 total | Distributed in MN and IL. | Terminated | August 20, 2015 | 610 | 601616 | Not_Computer
Z-0785-2014 | 67130 | No | IZJ | Cine or spot fluorographic x-ray camera. | Radiology | Radiology | 510(k) | ADAC ARGUS Gamma Camera Systems &amp; ADAC CIRRUS Gamma Camera Systems. Distributed by Philips Medical Systems, Cleveland, OH |  |  | December 04, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The recalling firm determined that visual inspections of the fork joints behind the detector that are conducted during six month preventive maintenance (PM) may not detect small cracks or deficiencies in the fork weldment. | Nonconforming Material/Component | On October 30, 2013 &amp; December 4, 2013, the firm sent Urgent Medical Device Recall letters to their customers. |  | 255 | 255 units | Worldwide Distribution: US (nationwide) including states of: AL, AR, CA, CO, CT, DC, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, ND, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TN, TX, VA, VT, WA, WI &amp; WV; and Internationally to: Austria, Belgium, Canada, Denmark, Germany, Great Britain, Italy, Netherlands, Norway, Pakistan, Russia, Singapore &amp; South Korea. | Terminated | November 06, 2015 | 702 | 255 | Not_Computer
Z-0784-2014 | 66846 | No | KRO | Knee joint femorotibial metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Segmental Systems; Polyethylene Inserts, Size C. |  |  | December 19, 2013 | 2013 | Zimmer, Inc. | After evaluation of 13 complaints(0.4% rate) for hyper-extension of the Zimmer Segmental System, Zimmer is updating the related surgical technique and instructions for use to provide additional instructions and patient conditions that may place excessive loading on the polyethylene insert. | Device Design | Depth of recall" Notifications were sent to customers, surgeons and risk managers, which use the affected product. Distributors were also informed of the notifications sent to surgeons and risk managers. The customers were identified by performing an invoice search and distribution history for the Polyethylene Inserts (00-5850-012-95 and 00-5850-013-95) since product launch. Recall Communication" All distributors were notified via electronic mail. Distributors that have received affected inventory were also notified via courier. Hospital risk managers and surgeons were also notified via courier.On 12/19/13, hospital risk managers and surgeons were provided with a notification identifying the issue and their responsibilities. These responsibilities include ensuring the affected personnel are aware of the contents of the notification. On 12/19/13, Distributors were sent a notification identifying the issue and to provide a copy of the notification sent to surgeons and risk managers. The distributors will also be instructed as needed to assist with identifying a comprehensive list of surgeons. " Communications outside of the United States will occur approximately two weeks after the United States communications.Effectiveness ChecksZimmer will conduct effectiveness checks in the following way:"Surgeons/Hospitalso100% of the FedEx notifications will be accounted for.oSecond and third notification attempts will be made.oFailed delivery attempts following the third attempt will result in the account being deemed as nonresponsive. |  | 5305 | 5305 distribution events | Worldwide Distribution. USA nationwide and Germany, with further distribution into EU countries. | Terminated | February 24, 2015 | 432 | 5305 | Not_Computer
Z-0790-2014 | 66933 | No | LLH | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BD GeneOhm Cdiff Assay, Part #441401, containing Lysis Kit sample buffer tubes, 200 tests/kit; labeled as ***GeneOhm Sciences Canada, Inc. 2555 boul. du Parc-Technologique Quebec, Qc, Canada, G1P 4S5*** |  |  | October 11, 2013 | 2013 | Becton Dickinson &amp; Co. | An incorrect reagent was packed with an in vitro diagnostic kit, which may cause an increase in unresolved results or false negative results for Clostridium difficile in patient samples. | Error in labeling | Becton Dickinson sent an Urgent Product Recall Letter dated October 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.They were advised that a portion of the BD GeneOhm Cdiff assay test kits from one lot, 08T13108, may contain incorrect bags of sample buffer tube. Customers were requested to return response forms indicating acknowledgement of the recall and quantity of remaining product that was discarded. Please fax to 410-316-4258. |  | 33 | 33 kits | US Distribution including the states of NY, IN, UT, MN, OH, MA, SD, CA, NJ, and PA. | Terminated | February 19, 2014 | 131 | 33 | Not_Computer
Z-0786-2014 | 67153 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | SIMVIEW NT and SIMVIEW NT with IDI: |  |  | December 11, 2013 | 2013 | Siemens Medical Solutions USA, Inc | In a very rare occurrence an intended movement of the SimView NT simulation system components could not be stopped with usual hand control or the collision detection functionality. | Device Design | Siemens sent an Urgent Field Safety Notice to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. What preventive measures can the user take?To stop the movement in this case, push the emergency button / stop button at control console or push the in-room emergency button. The emergency button/ stop button functionality is described in the User Manual.There is no preventive measure to avoid this situation. This letter is intended to bring this potential safety issue to your awareness and to ensure that your personnel are trained about the use of the emergency button / stop button inside the simulation room and on the control console. Always monitor any movement of the simulator components.The Update Instruction containing the Urgent Field Safety Notice will be delivered by the Siemens Customer Service Group, as per their procedureSiemens is working on a solution to mitigate the described situation. You will be informed as soon as the update distributing this solution is available.Please include this Field Safety Notice in your SimView NT Simulator System Owner Manual chapter Safety Advisory Letters where it should remain.In the interests of safety, we ask that you perform the described preventive measures and inform all affected personnel immediately.The relevant National Competent Authority has been informed of this Customer Safety Advisory Notice.We regret any inconvenience that this may cause, and we thank you in advance for your understanding. For further questions please call (925) 602-8157. |  | 65 | 65 active devices | Class II Recall - Worldwide Distribution - US Distribution including the states of California, Louisiana, Massachusetts, Pennsylvania, New York, Kentucky and Michigan., and the countries of Bulgaria, Germany, Belarus, Russian Fed., P.R. China, India, Republic Korea, Germany, Costa Rica, Argentina, Portugal, Greece,Bulgaria, Vietnam, Hungary, Italy, Palestinian Ter, Hungary, Jordan, EgyptPoland, Greece, Curacao, St Eus, Malaysia, Egypt, Malaysia, South AfricaTurkey, Thailand, Qatar and Iraq. | Terminated | March 31, 2014 | 110 | 65 | Not_Computer
Z-0787-2014 | 67126 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | PRIMUS, ONCOR, ARTISTE Linac systems with component 550 TxT Treatment Table. |  |  | December 05, 2013 | 2013 | Siemens Medical Solutions USA, Inc | Update to correct the possibility for patients or operators to have their fingers trapped or pinched between the table top and the bottom frame of the 550 TxT Treatment Table. Combined with this update instruction is a second complaint issue related to an incident in which a cable inside of the TxT" Treatment Table has been damaged due to contacting a moving part. The resulting short-circuit cau | Device Design | The firm, Siemens, sent a "Medical Device Correction" letter to affected users/customers starting on December 5, 2013 via Federal express. The letter described the product, problem and actions to be taken. Siemens has arranged for modification of the affected 500 TxT Treatment Table. A service representative will modify the mounts of the TT-A, TT-D, TT-M, or TT-s table tops ( Pleases note that this modification only applies to the Siemens provided table tops TT-A, TT-D, TT-M, and TT-S.). In addition, the Siemens service representative will modify the cable routing inside of the patient table base. The customers were instructed to include this Field Safety Notice in your Digital Linear Accelerator System Owner manual chapter 'Safety Advisory Letters' where it should remain.If you have any questions, please contact the Senior Manager, Regulatory Affairs at 1-925-293-5442 or email: christine.dunbar@siemens.com. |  | 305 | 305 active devices. | Worldwide Distribution: US (nationwide) and internationally to: Canada, Italy, P.R. China, Argentina, Pakistan, South Africa, Egypt, Iran, Vietnam, Indonesia, Brazil, Hungary, Bulgaria, India, Spain, Switzerland, Trinidad, Tobago, Italy, Algeria, Croatia, Turkey, Japan, Germany, Portugal, Poland, France, Sri Lanka, Russian Federation, France, Lebanon, Slovakia, Ukraine, Pakistan, and Korean Republic. | Terminated | January 30, 2015 | 421 | 305 | Not_Computer
Z-0780-2014 | 67078 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO RTP System |  |  | November 25, 2013 | 2013 | Elekta, Inc. | When the number of fractions is changed for a proton spot beam, the MU are recomputed and rounded if necessary, but the dose is not recalculated as it should be even after saving and re-opening the plan. | Software Design Change | An Important Safety Notice, LINXIO0017 Proton Sport Beams, Changes in Fractionation not Updating Dose, is in the process of being distributed to affected customers. The release date for this notice is November 25, 2013. The notice included a Confirmation of Receipt which was to be returned to Elekta.For further questions please call (404) 993-5886. |  | 13 | 13 | Worldwide Distribution - US Distribution including the states of IL, IN, LA, NJ, TX, WA and WI., and the countries of Belgium, Germany, Italy and Netherlands. | Terminated | May 13, 2014 | 169 | 13 | Not_Computer
Z-0796-2014 | 67047 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | DRX-Evolution with FF WS |  |  | November 25, 2013 | 2013 | Carestream Health, Inc. | Carestream Health Inc. has issued a recall for the Carestream DRX-Evolution (FF WS and CSJ WS) and the Kodak DirectView DR 7500 System due to the potential for unexpected device movement. | Nonconforming Material/Component | The firm, Carestream Health Inc, sent an "URGENT: Medical Device Recall" notice dated November 20, 2013 to their customers. The notice described the product, problem and actions to be taken. The customers were instructed that the Emergency Stop (E-Stop) button be used on the motorized FF WS and CSH WS Buckys for any observed uncontrolled movement as indicated in the user guide. If this movement is observed, please notify Carestream Health, and Carestream will schedule a site visit to resolve the problem. In addition, Carestream Health Inc. will apply a label(s) indicating the location of the E-stop button for customers using the FF WS. If you have any questions or concerns, please contact the Carestream Customer Care Center in the U.S. at 1-800-328-2910. |  | 675 | 675 units Total (145 units domestically &amp; 530 units internationally) | Worldwide Distribution: US (nationwide) including states of: CA, CO, FL, GA, IL,IN, MD, MI, MN, NC, ND, NJ, NY, OH, OK, PA, SC, SD, TN, TX, VA, WA, and WI; and internationally to: Australia Belgium, Canada, China, Finland, France, Germany, Hong Kong, Italy, Netherlands, New Zealand, Poland, Portugal, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 675 | Not_Computer
Z-0801-2014 | 67168 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-Ray System (Brilliance CT Big Bore Oncology &amp; Brilliance CT Big Bore Radiology) Distributed by Philips Healthcare System, Cleveland, OH. |  |  | December 04, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Machine Name in TumorLOC Basic Mode is obscured. When TumorLOC is in Basic Mode and Set Vertical Laser equals Yes, the Machine Name label and field are covered by the controls above it. The Machine Name cannot be viewed or changed. | Software design | On 12/4/2013 the firm sent Urgent Medical Device Corrections letters to their customers. |  | 32 | 32 units | Worldwide Distribution-USA (nationwide) including the states of AK, MD, NH, NV, OH, and OR, and the countries of Australia, China, France, India, Indonesia, Italy, Korea, Mexico, Poland, Romania, Russia, Spain, Switzerland, Taiwan, Thailand &amp; United Arab Emirates. | Terminated | July 01, 2015 | 574 | 32 | Not_Computer
Z-0795-2014 | 65184 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Antimicrobial Susceptibility Test Discs, Ertapenem 10 mcg., CT1761B, packaged in cartridges of 50 discs, 5 cartridges/pack. The firm name on the label is Oxoid Ltd., Basingstoke, Hampshire, England. |  |  | April 09, 2013 | 2013 | Remel Inc | Cartridges may contain individual discs which were not impregnated with antibiotic. | Process design | ThermoFisher Scientific sent a Medical Device Recall letter dated March 22, 2013, to affected customers. Letters dated April 9, 2013, were issued for the remaining three products. Customers were instructed to notify all personnel who need to be notified and review their inventory and discard any affected product. Customers were also instructed to return the attaced Medical Device Recall Response acknowleding the dispoasl of the affected product. Customers with questions were instructed to contact the Technical Services Department at 800-255-6730 or 913-888-0939 (International).For questions regarding this recall call 913-895-4185. |  | 30 | 30 packs | Nationwide Distribution: US including Puerto Rico. | Terminated | January 21, 2014 | 287 | 129 | Not_Computer
Z-0800-2014 | 67030 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens ADVIA 1200 and ADVIA 1650 Chemistry Systems. |  |  | December 17, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | An overflow flag //// is displayed in the result column for a test when the algorithm for the test cannot calculate a result or when the concentration cannot be determined because it is outside the absorbance limits, and it is possible for the user to misinterpret or edit the result. | Software design | The distribution of the Urgent Medical Device Correction letter scheduled to begin on December 17, 2013.Âżx Field service personnel will be notified of the Urgent Medical Device Correction Letter via the Siemens Knowledge Base, as well as information regarding placing labels on the customer system to remind the customer that ///// is not a reportable result.Âżx The Customers were instructed to contact Siemens Healthcare Diagnostics Technical SolutionCenter if they have technical questions regarding this issue.Âżx Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction. |  | 1563 | 1563 | Worldwide Distribution. USA nationwide and the country of Canada. | Terminated | August 26, 2014 | 252 | 1563 | Not_Computer
Z-0804-2014 | 67125 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Video Monitor Suspensions that may be on the following Fluoroscopic and Radiographic X-Ray systems: Precision 500D, Legacy, Prestilix, Monitrol, RFX, SFX, Tilt-C, Prestige VH, Prestige SI, and Prestige SII |  |  | November 08, 2013 | 2013 | GE Healthcare, LLC | Missing screws in the overhead Video Monitor Suspension. GE discovered that the 4 outer screws that hold the overhead Video Monitor Suspension to the carriage were missing since installation of the system. GE is not aware of any fall of the overhead monitor suspension. | Employee error | GE Healthcare Systems sent an " Urgent Medial Device Correction" letter dated November 7, 2013. The letter was addressed to Hospital Administrators / Risk Mangers, Managers of Radiology /Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (USA 800-437-1171, Japan 0120-055-919, For other countries, please contact your local GE Healthcare Service Representative). |  | 5651 | 5651 | Worldwide Distribution - US Nationwide including Washington D.C., Guam, and Puerto Rico. Outside the US to include: Afghanistan, Algeria, Argentina, Australia, Austria, Azerbaijan, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Brazil, Bulgaria, Burkina faso, Cameroon, Canada, Chile, China, Colombia, Costa Rica, Cote D'Ivoire, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Germany, Ghana, Greece, Guadeloupe, Guatemala, Guyana, Honduras, Hong Kong, Hungary, Iceland, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Korea (Republic of), Kuwait, Lebanon, Libyan Arab Jamahiriya, Macedonia, Malaysia, Mali, Malta, Martinique, Mexico, Morocco, Netherlands, New Caledonia, New Zealand, Niger, Norway, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, Yemen. | Terminated | November 03, 2014 | 360 | 5651 | Not_Computer
Z-0802-2014 | 66591 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Civco Belly Board, Reusable non-sterile carbon fiber belly board with cushion, REF 125012, packaged 1 per box. |  |  | October 25, 2013 | 2013 | Civco Medical Instruments Inc | Scales adhered to the product may be misaligned from the zero reference such that one side will not match the other. | Device Design | Civco Medical Instruments Inc sent a FIELD SAFETY NOTIFICATION letter dated October 28, 2013 via UPS to their customers. The notification identified the affected device, problem, and field safety corrective action. The letter describes how customers can check to see whether the scales on their device are out of alignment. The letter states CIVCO will rework the system to correct the difference. For questions contact sales representative at CIVCO Medical Solutions, Radiation Oncology at 800-842-8688 or 712-737-8688. |  | 138 | 138 belly boards | Worldwide Distribution - USA Nationwide in the states of AL, AR, AZ, CA, CO, CT, FL, GA, IL, KS, MA, MD, MI, MN, MT, NC, NJ, NY, OH, OK, PA, SC, TN, TX, VA, WA, WI, and WV. Foreign distribution was made to Brazil, Canada, Chile, Czech Republic, France, Germany, Iraq, Israel, Italy, Japan, Malaysia, Saudi Arabia, Slovenia, South Korea, Spain, Turkey, and United Arab Emirates. | Open, Classified |  |  | 140 | Not_Computer
Z-0799-2014 | 67178 | light, microscope | IBJ | Microscopes and accessories. | Pathology | Pathology | 510(K) Exempt | Compensator Slider Lambda 6x20 component (part number 473704-0000-000) utilized on all upright ZEISS microscopes (Axio Imager, Axio Scope. A1, Axio Lab. A1). |  |  | December 20, 2013 | 2013 | Zeiss, Carl Inc | Carl Zeiss Microscopy, LLC has recalled the Compensator Slider Lamda 6x20 component utilized on Standard Pathology Microscopes due to misalignment of the lambda plate. | Nonconforming Material/Component | Carl Zeiss, Inc. sent an Important Medical Device Corrective Action letter dated December 20, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised that their service provider would contact them to schedule a visit in which a factory trained service technician would install a new Compensator Slider Lambida 6x20. Customers should return the old Compensator Slider to the factory for testing. Customers with technical questions were instructed to call 914-681-7815. |  | 335 | 335 units Total (14 domestically &amp; 321 internationally) | Worldwide Distribution - Nationwide including NY, TX, NC, FL, ME, MA and MI; and Internationally to Arab Emirates, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Egypt, France, Germany, Hong Kong, India, Italy, Kosovo, Lithuania, Mexico, The Netherlands, New Zealand, Peru, Poland, Spain, and Taiwan. | Terminated | April 08, 2015 | 474 | 335 | Not_Computer
Z-0805-2014 | 67013 | No | LNX |  |  | General Hospital | 510(k) | ORCHESTRA/ORCHESTRA PLUS Programmer |  |  | November 20, 2013 | 2013 | Sorin Group Italia S.r.l. | Sorin has voluntarily issued a notification to physicians related to the overestimation of the residual longevity displayed by its programmer (ORCHESTRA or ORCHESTRA PLUS) during a follow-up exam of patients implanted with a REPLY or ESPRIT pacemaker. | Software design | On Nov 20 and Nov 25, 2013, the Dear DoctorLetter was sent by certified mail to 1,310 physicians who are following 13,003 patientsimplanted with a REPLU or ESPRIT pacemakers in the US. Sorin provided the following recommendations to the physicians: You should consider checking the battery impedance of the last follow-up exam. In casethe battery impedance is greater than or equal to 3.5kÂż, a follow-up visit must bescheduled within a maximum of 6 months from the last follow-up visit. When pacemaker operation is checked by the simple application of a magnet, a magnetrate less than 95 min-1 should trigger a follow-up exam in the pacemaker centre. As a general rule, a maximum of 6 month follow-up interval when the battery impedancebecomes greater than or equal to 3.5 kÂż. This recommendation should also be followedsubsequent to the installation of the new programmer software version. |  | 1718 | 1,718 | Nationwide. | Terminated | September 08, 2014 | 292 | 1718 | Computer
Z-0807-2014 | 67085 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | SOLAR 8000 AND TRANSPORT PRO WITH PATIENT DATA MODULE (Marketed As GE Healthcare CARESCAPE Patient Data Module) Product Usage: The Patient Data Module (PDM) is intended to provide uninterrupted acquisition of physiologic parameter data on adult, pediatric and neonatal patients during non-transport/bedside and transport patient care episodes. |  |  | December 05, 2013 | 2013 | GE Healthcare, LLC | Intermittent noise on ECG and/or Respiration waveforms associated with movement of the ECG cable connector on the DASH, PDM, and/or Tram Module. May result in reduced ECG, Respiration and arrhythmia detection performance. | Pending | GE Healthcare sent an Urgent Medical Device Correction letter dated December 5, 2013, to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers with questions were instructed to contact Technical Support at 1-800-558-7044. For questions regarding this recall call 262-513-4122. |  | 2516 | 2516 (2,321 US, 195 OUS) | Worldwide Distribution: US (nationwide) including DC and PR except SD, VT and WY; and Internationally to: ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BELGIUM, BRAZIL, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, EGYPT, GERMANY, UNITED KINGDOM, IRELAND, IRAQ, ISRAEL, JAMAICA, JAPAN, JORDAN, KOREA, KUWAIT, LEBANON, LITHUANIA, MARITIUS, MALDOVA, NEW ZEALAND, NIGERIA, PAKISTAN, PANAMA, PARAGUAY, PERU, POLAND, QATAR, RUSSIA, SLOVAKIA, SOUTH AFRICA, SWEDEN, SWITZERLAND, TUNISIA, UNITED ARAB EMIRATES, VENEZUELA, CHILE, CANADA, and BERMUDA. | Open, Classified |  |  | 51589 | Not_Computer
Z-0813-2014 | 67255 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | AMS-530 7 Microbore Trifurcated Extension sets |  |  | January 07, 2014 | 2014 | Churchill Medical Systems, Inc. | Defective component may cause extension sets to leak. | Nonconforming Material/Component | Vygon sent an Urgent Medical Device Recall letter dated January 7, 2014 to each customer via courier service and email. The letter identified the affected product, problem and actions to be taken. The recall letter instructs customers to: 1. Check your stock for the products included within the scope of the recall. Immediately cease use and distribution, and quarantine all affected product immediately. 2. Count and document your affected inventory on the Recall Acknowledgement and Inventory Return Form 3. Fax or email the form to Vygon including customer contact information. 4. Vygon customer service will contact the customer and issue a Return Merchandise Authorization number. 5. Return the product as well as the original signed Recall Acknowledgement and Inventory Return Form using the instructions from Customer Service involving product pick-up. 6. All inventory will be credited or replaced at no cost to the customer. Distributors were provided with instructions on how to handle product that was further distributed by them. For questions contact your local sales representative or Vygon's Customer Service Department at 1-800-473-5414, or by email at customerservice@vygonus.com. |  | 3850 | 3850 | US Nationwide Distribution in the states of: NY, VA, IL, MO, OR, FL, DE, OH | Terminated | November 03, 2014 | 300 | 3850 | Not_Computer
Z-0810-2014 | 66996 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Medrad Veris MR Monitor units |  |  | November 21, 2013 | 2013 | Medrad Mr Inc | The main board, P/N 301641, installed in some Medrad Veris MR Monitor units may be faulty and could lead to unexpected shutdown of the system while in use, resulting in the loss of vital signs information from the monitor. | Device Design | The firm, Bayer HealthCare, sent a "MEDICAL DEVICE FIELD CORRECTIVE ACTION" letter dated November 21, 2013 via certified mail to its users/customers. The letter described the product, problem and actions to be taken. The firm stated that, "You may choose to continue to use your Veris units but you MUST be vigilant about this potential issue. Should a shutdown occur, the system can be quickly restarted by pressing the POWER button. If thisproblem does occur, we ask that you report this immediately to Bayer Customer Support at 1-877-229-3767." The customers were instructed to verify the serial number on your Medrad Veris unit(s). If this problem does occur, report this immediately to Bayer Customer Support at 1-877-229-3767. The customers were also instructed to fill out and fax back the attached form. Bayer Service Representative will then be in touch to schedule an appointment to come and replace your main board, at no cost to you.Bayer also informed their customers that based on the maximum inventory that can be produced by the supplier, they anticipate that this Field Corrective Action will take up to six months.If you have questions, please contact our Customer Support team at 1-877-229-3767. |  | 922 | 922 units | Worldwide Distribution: US (nationwide) and Internationally to: Canada. | Terminated | March 30, 2015 | 494 | 922 | Not_Computer
Z-0816-2014 | 66875 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABL90 FLEX Analyzer Radiometer Medical ApS, Âżkandevej 21 |  |  | October 15, 2013 | 2013 | Radiometer America Inc | RADIOMETER become aware that if the ABL90 FLEX is tilted, this may potentially cause a negative bias on sodium (Na+) results. | Other | Radiometer sent an Urgent Medical Device Recall letter dated November 5, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check all of their ABL90 FLEX analyzers and ensure that the analyzer is placed on a horizontal platform/table. If the analyzer is placed on a roll stand, please ensure that the support plate is aligned horizontally and locked in this position. Customers were asked to complete and fax the accompanying acknowledgment form in order to confirm their receipt of this field safety notice. Customers with questions were instructed to contact Radiometer America Technical Support at 1-800-736-0600, option 4.For questions regarding this recall call 440-871-8900. |  | 3414 | Total 3414 (US 524, CA 31, ROW 2859) | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI &amp; WV and Internationally to Australia, Austria, Canada, China, Czech Republic, Denmark, Finland, France, Germany, Hungary, Italy, Japan, Netherlands, Norway, Poland, Russia, Singapore, South Africa, Spain, Sweden, Switzerland, Turkey &amp; United Kingdom. | Terminated | June 19, 2014 | 247 | 3414 | Not_Computer
Z-0788-2014 | 66493 | No | KGG |  |  | Neurology | PMA | TRUFILLÂż n-Butyl Cyanoacrylate (n-BCA) Liquid Embolic System; Two 1 g tubes of n-BCA, one 10 mL vial of Ethiodized Oil and one 1 g vial of Tantalum Powder |  |  | October 11, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Incorrect instructions for use (IFU). | Employee error | The firm, CODMAN Neuro, sent a "Medical Device Correction Notice" letter dated October 21, 2013 to its consignees/customers. The letter describes the product, problem and actions taken. The letter informed the consignees/customers of the incorrect statement in the IFU. The Corrected IFU was provided. Codman Neuro is in the process of updating the Instructions for Use to include the corrected information. Consignees/customers were instructed to complete and return the Correction Notice Acknowledgement Form and fax it to Codman Neuro at 508-977-6665.For additional information, please contact Codman Neuro representative at 1-866-491-0974. |  | 3860 | 3,860 | Worldwide Distribution: US (nationwide) including Puerto Rico; and internationally to: Costa Rica and Russia. | Terminated | July 28, 2015 | 655 | 11843 | Not_Computer
Z-0812-2014 | 67164 | No | CAH | Breathing circuit bacterial filter. | Anesthesiology | General Hospital | 510(k) | Vital Signs Anesthesia Circuits |  |  | December 26, 2013 | 2013 | Vital Signs Colorado Inc. | A circuit in use was occluded by a blue port cap that was inadvertently attached to the circuit as a result of an internal inspection error and will cause occlusion of gas flow. | Process control | GE Healthcare sent an Urgent Medical Device Correction letter on December 12, 2013 to all affected customers via certified letter to follow the safety instructions to:1) Please follow all safety instructions in accordance with the instructions for use (IFU).2) During installation, visually inspect the Vital Signs Breathing Circuits for any plugged ports. If 15mm or 22mm blue port caps are attached to any ports, please remove and discard.3) Prior to use, the circuits should be pressure tested for leaks and flow tested for obstructions in accordance with the ventilator manufacturers specifications. The long-term solution is to update the instructions for theassembly operator and modify the circuit air test process by replacing the cap with a plug tethered to the workstation so the plug cannot be left on the circuit.If you have any questions or concerns regarding this notification, please contact GEMSIT Customer Service at +1-800-588-7044 (Domestic) or Vital Signs Customer Service at +1-800-932-0760 (International). Hours ofOperation: 8:00 am EST to 6:00 pm EST.Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately per the contact information above. |  | 2332680 | 1,827,920 units US , 504,760 units OUS | Worldwide Distribution - USA (nationwide) including Puerto Rico; and Internationally to: Belgium, Canada, Chile, Mexico, Costa Rica, Kuwait, United Kingdom, Greece, Czech Republic, Latvia, Turkey, China, France, Italy, Brazil, Korea, Germany, Thailand, Egypt, Austria, Dominican Republic, India, Israel, Colombia, Japan, Netherlands, Venezula, Saudi Arabia, Argentina, Malaysia, Sweden, Lebanon, South Korea, Finland, United Arab Emirates, Taiwan, Panama, Uruguay, Singapore, Hong Kong, Ecuador, Peru, Oman, Spain, Lithuania, and South Africa. | Terminated | July 09, 2014 | 195 | 2332680 | Not_Computer
Z-0818-2014 | 67057 | No | MKW | Cyclosporine test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension Vista Cyclosporine Flex reagent cartridge |  |  | December 05, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Complaints were received regarding variability in recovery of QC and patient results with certain lots of the Dimension Vista CSA Flex reagent cartridges. | Nonconforming Material/Component | Siegmens sent an Urgent Medical Device Recall lletter dated December 2013 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use and discard any remaining inventory of the affected product. Customers were also instructed to complete the attached form and fax it to 302-631-8467. Customers were asked to retain the letter with their laboratory records, and forward the letter to any who have received the affected product. Customers with questions were instructed to contact Siemens Customer Care Center - Technical Solutions or their local Siemens technical support representative.For questions regarding this recall call 800-441-9250. |  | 1586 | 1586 | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | August 27, 2014 | 265 | 1586 | Not_Computer
Z-0817-2014 | 66474 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products COCM Reagent |  |  | October 04, 2013 | 2013 | Ortho-Clinical Diagnostics | Inaccurate quality control results. | Other | On October 4, 2013 Ortho Clinical Diagnostics issued a Product Correction Notification to customers via US Postal Service Priority Mail or FedEx regarding VITROS Chemistry Products COCM Reagent due to inaccurate quality control results. Ortho Clinical Diagnostics advised their customers to discontinue use and discard any recalled product. OCD customers with questions should be contact Customer Technical Services at 1-800-421-3311. |  | 3342 | 3,342 units | Worldwide Distribution. US nationwide (including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, IN, KY, LA, MD, MI, MO, MS, NC, NE, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, Puerto Rico), Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, and Spain. | Open, Classified |  |  | 3342 | Not_Computer
Z-0819-2014 | 67149 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Composix L/P with Echo PS 10" X 13" Product Code 0144113 |  |  | January 03, 2014 | 2014 | Davol, Inc., Subs. C. R. Bard, Inc. | Product labeling does not match product configuration. | Labeling Change Control | Davol, Inc., Subs. C. R. Bard, Inc.sent a customer notification letter on January 3, 2013, to all affected customers via Fedex to inform consignees that the Instructions for Use (IFU) and graphic contained onthe inside of the product carton bottom references two features added to the inflation tube that are not present on the product they received. Please contact the Davol Customer Service Department at 1-800-556-6275 or C. R. Bards Medical Services &amp; Support Department at 1-800-562-0027 if you have any questions. Consignees are asked to complete an effectiveness check form and return it via fax to 401-825-8753.For questions regarding this recall call 800-556-6275, |  | 19 | 19 | Nationwide Distribution including CA, PA, IL, VA, FL, GA, NY, TX, MO, NE, IN, NM, OH, IA. | Terminated | December 23, 2015 | 719 | 65 | Not_Computer
Z-0809-2014 | 67182 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Pulmonary Hemi-Artery SG |  |  | November 22, 2013 | 2013 | CryoLife, Inc. | Serological testing for the donor was performed with a blood sample that may have been hemodiluted due to the administration of normal saline to the donor prior to death. | Other | CryoLife sent an Urgent Human Tissue Recall letter dated November 22, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to contact CryoLife to arrange for the return of the affected product. If the affected product had been implanted or discarded, customers should provide information regarding the status of the affected product on the attached self-addressed postcard. Customers with questions were instructedto contact the Field Assurance Department at 800-438-8285.For questions regarding this recall call 770-419-3355. |  | 1 | One | US distribution in CA. | Terminated | February 20, 2014 | 90 | 1 | Not_Computer
Z-0815-2014 | 66991 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Titanium Single Lumen Low-Profile port, with Pre-Attached open-Ended Silicone 6.6Fr Catheter. Product Code 0602180 |  |  | December 06, 2013 | 2013 | Bard Access Systems | Bard Access Systems is recalling various Low Profile Port with 6.6 Fr Single Lumen Broviac Catheter because a small number of Silicone catheters in the may be undersized to 5.3 Fr instead of 6.6 Fr as stated in the labeling for the device. | Process control | BARD sent an Urgent Product Recall letter dated October 2013 to affected customers via certified letter. The letter identified the affected product, problem and actions to be taken. Customers were asked to check all inventory locations, remove and return any unused affected prodcut to Bard Access Systems through return goods authorization number RAP80498 for a credit or replacement, and to complete the Reply Form and Inventory Reconcilation Form. For questions contact BAS Customer Service at 1-800-290-1689. |  | 200 | 200 | Worldwide Distribution - US nationwide, Belgium, and Canada. | Terminated | August 05, 2014 | 242 | 1558 | Not_Computer
Z-0811-2014 | 64966 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | 1) The FEMTO LDV Z2 Femtosecond Surgical Laser is an ophthalmic surgical laser indicated for use in the creation of corneal incisions in patients undergoing LASIK surgery or other treatment requiring lamellar resection of the cornea at constant depth relative to the corneal surface. |  |  | April 16, 2013 | 2013 | Ziemer Usa Inc | A defect in the controller board can lead to the condenser to the video monitor becoming too warm. This could potentially cause the video monitor to shut off. | Component design/selection | A recall notification letter dated April 16, 2013 was mailed to direct accounts. The letter notified the direct accounts of the problem and advised them that a technician would be arriving at their facility during the week of June 2, 2013 to replace the defective controller board. The technician will also perform preventative maintenance on the device at the same time. Direct accounts are advised to contact Ziemer USA, Inc. at 618-462-9301 if they have any questions about this issue. |  | 5 | 5 units | Worldwide Distribution: US (nationwide) including states of: NJ and IL; and Internationally to: Finland, Japan, and Germany. | Open, Classified |  |  | 5 | Not_Computer
Z-0823-2014 | 66694 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens CentraLink" Data Management System V14.0.4, V14.0.5 or V14.0.8 |  |  | October 30, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Under extremely rare circumstances, a patient result that had been previously disabled may be released without proper review. (Disabled Test in CentraLink System or Aptio System May Become Enabled in CentraLink System.) | Software design | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated October 23, 2013 to its customers. The firm also sent a follow-up letter dated October 30, 2013. The letter describes the product, problem and actions to be taken. The customers were instructed Until a service visit is scheduled, do the following: Instruct the laboratory staff to recognize this sequence of events and Ensure that whenever the operator disables a method on an Aptio system or a CentraLinksystem and replenishes reagent for that method, the operator then confirms that the method remains disabled in the CentraLink system or disables it again. The operator must allow up to five minutes to verify that the test remains disabled. In addition, the customers were instructed tocomplete and return the FIELD CORRECTION EFFECTIVENESS CHECK response form via fax to Siemens Healthcare Diagnostics Customer Care Center at 800-441-9250, and to retain this letter with your laboratory records, and forward this letter to those who may have received this product. After a service visit is performed and an interim solution is implemented, the laboratory staff will be trained on this update to understand the minor but necessary changes in the workflow.If you have any questions, please contact your Siemens Customer Care Center at 800-441-9250 or your local Siemens representative. |  | 40 | 40 | US Distribution in the states of: MN, NJ, and NV. | Terminated | December 15, 2014 | 411 | 40 | Not_Computer
Z-0826-2014 | 67122 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Port Kit with Access Infusion Set under the following labels: 1) 8F/SL/Plastic Port/F, Product Number H965450140, Catalog Number 45-014; |  |  | November 25, 2013 | 2013 | Navilyst Medical, Inc. | Navilyst Medical has recalled a Port Kit with Access Infusion Set due to a possible breach of sterilization barrier. | Under Investigation by firm | Navilyst Medical, Inc., sent an Urgent Medical Device Recall - Immediate Action Required letter dated November 25, 2013 to their customers via Federal Express. The letter identified the affected product, problem and actions to be taken. Customers are instructed to segregate and return all affected devices to Navilyst Medical, Inc. Distribution Center Attn: QA Returns 10 Glens Falls Technical Park Glens Falls, NY 12801 and complete and return the Reply Verification Tracking Form provided in the recall notification via Fax to: 1-800-782-1357. For questions call 1-800-833-9973 between 8:30 am and 7:00 pm (EST). |  | 920 | 920 Total kits | US Nationwide Distribution in the states of CA, CT, DC, MA, NJ, NY, PA, MD, VA, WV, NC, GA, FL, TN, OH, MN, MT, IL, MO, NE, OK, TX, NM, and WA. | Open, Classified |  |  | 920 | Not_Computer
Z-0825-2014 | 67226 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 |  |  | December 22, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | The locking mechanism that is intended to hold the control panel in a fixed position can become loose and fail to lock the rotation of the control panel, making it unable to steer or control during transport. | Device Design | Customer Safety Advisory Letters were sent to all affected users on December 22, 2013. Letters advised of the issues and the actions they should take until a Siemens Field Engineer will visit and upgrade the control panel. |  | 271 | 271 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Canada, Ecuador, Finland, France, Germany, India, Ireland, Israel, Italy, Japan, Mexico, P.R. China, Philippines, Portugal, Republic Korea, Russian Fed. Saudi Arabia, Serbia, South Africa, Spain, Suriname, Sweden, Taiwan, Turkey, U.A.E., and United Kingdom. | Terminated | March 05, 2015 | 438 | 271 | Not_Computer
Z-0821-2014 | 67163 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 840 Series Ventilator PB840 Operators Manual Addendum PN 066009A 09/02 |  |  | January 25, 2013 | 2013 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is correcting labeling to clarify the operational life of the oxygen sensor (O2 sensor), PN 4-072214-00 used in the Puritan Bennett" 840, and PN, G-062010-00, used in Puritan Bennett" 740, and760 ventilator systems. The Operators Manual for PB840 states a nominal life of one year and the Addendum incorrectly states 2 years. | Labeling Change Control | The firm, Covidien, sent an "URGENT MEDICAL DEVICE VOLUNTARY FIELD CORRECTION" letter dated January 24, 2014 to its Consignees/customers via certified mail. The letter described the product, problem and actions to be taken. The customers were instructed to read and understand the enclosed information; communicate this information to those who need to be aware of this labeling update; for the Puritan Bennett 840 ventilator only: If you have addendum P/N 066009A 09/02, remove and discard the addendum and replace it with the new information in Attachment I; for the Puritan Bennett 740 and 760 ventilators only: Add the new information in Attachment I; Ensure all O2 sensors in use or in inventory conform to the instructions in this notification. to your current operators manual and service manual, and read and understand the warnings, cautions and notes related to safety, ventilator use and operation that are contained in the Preface of the Operators Manuals. are asked to remove incorrect labeling and replace with the revised labeling.Covidien enclosed a labeling update (Attachment I) which provides detailed information related to the life of oxygen sensors used with the Puritan Bennett 840, 740 and 760 ventilators. Covidien has validated another oxygen sensor, P/N 10097559, for use in the Puritan Bennett 840, 740 and 760 ventilators. Information about that sensor is also included in the update in Attachment I.Should you have any questions regarding this letter or to report any issues with the Puritan Bennett" 840, 740 or 760 ventilators contact our Technical Support Department @ 1-800-255-6774, option 4, then option 1, to ensure proper device reporting procedures are followed. |  | 70160 | 70,160 | Worldwide Distribution: US (nationwide) and Internationally to: Albania, Argentina, Australia, Bahrain, Belgium, Belize, Bermuda Bolivia, Bosnia and Herzegovina, Brazil, Cambodia, Canada, Chile, China, Colombia, Costa Rica, Croatia, Ecuador, Egypt, El Salvador, Finland, France, French Polynesia, Georgia, Germany, Greece, Guam, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Lebanon, Libya, Lithuania, Macedonia, The former Yugoslav, Malaysia, Mexico, Moldova, Republic of, Montenegro, Morocco, Netherlands, Nicaragua, Northern Mariana Islands, Oman, Panama, Paraguay, Peru, Philippines, Poland, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Switzerland, Syria, Taiwan, Thailand, Tunisia, Turkey, Uganda, Ukrainian SSR, United Arab Emirates (UAE) United Kingdom, Uruguay, Venezuela and Yugoslavia. | Terminated | November 24, 2014 | 668 | 78119 | Not_Computer
Z-0824-2014 | 67209 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound System; |  |  | December 22, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | The user interface assembly on the ACUSON SC2000 may become loose with the potential for the entire module to fall off of the ultrasound system. | Nonconforming Material/Component | Siemens sent an Customer Safety Advisory Notification letter on December 22, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Customers were instructed to:.If you observe looseness in the attachment of the user interface module to the column on the ultrasound system, immediately stop using your system and contact your Siemens service representative for repair.Your local Customer Service Engineer will contact you to make the necessary repairs on your system. You will be contacted to schedule an appointment. If you have any questions, please contact your local service support person for information regarding timelines and status.Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization will need to be aware of thls issue. As always, patient safety issues are a high priority, and we thus alert customers of product concerns which may occur undervarlous conditions. To date, no patient injury has been reported.This problem was discovered as part of our ongoing quality process. ÂżÂż ÂżÂżWe sincerely regret any inconvenience this condition may cause in your daily operations. For further questions please call (650) 694-5398. |  | 843 | 843 active units | Worldwide Distribution - US (nationwide) and the countries of Albania, Australia, Austria, Bangladesh, Belgium, Bosnia Herzeg, Brazil, Brunei, Canada, Columbia, Denmark, Ecuador, Egypt, Finland, France, Germany, Iceland, India, Israel, Italy, Japan, Kuwait, Malaysia, Mauritius, Mexico, Netherlands, Norway, Oman, P.R. China, Philippines, Poland, Qatar, Rep of Yemen, Republic Korea, Romania, Russian Fed. San Marino, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E. United Kingdom and Venezuela. | Terminated | May 06, 2015 | 500 | 843 | Not_Computer
Z-0183-2014 | 66610 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Applied Medical's InziiÂż 12/15mm Retrieval System. |  |  | October 21, 2013 | 2013 | Applied Medical Resources Corp | Applied Medical is conducting a voluntary recall of the InziiÂż 12/15mm retrieval system. During shipment, the retrieval system packaging has the potential to become punctured with small holes, which could compromise the sterile barrier. The likelihood of this situation to occur and result in patient harm is highly unlikely; however, out of an abundance of caution for patient safety and a commitmen | Package design/selection | Customer notifications of the recall began the week of 10/21/2013 via UPS overnight with letters requesting that all affected products be returned to Applied. The notification letter is dated October 21, 2013 and titled "URGENT: MEDICAL DEVICE RECALL". The notification was accompanied by a Recall Notification Confirmation Form. The letter informed customers of the recall by providing the product description with codes (Model and Lot Numbers). reason for the recall, actions to take and contact information. |  | 29506 | 29,506 units | Worldwide Distribution: US (nationwide) and countries of: Australia, Austria, Belgium, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Ecuador, Finland, France, Germany, Hong Kong, India, Ireland, Israel, Italy, Japan, Lebanon, Malaysia, Mexico, New Zealand, Netherlands, Norway, Peru, Poland, Portugal, Saudi Arabia, Singapore, Spain, South Africa, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | September 15, 2014 | 329 | 29506 | Not_Computer
Z-0830-2014 | 67176 | No | GFF | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Anspach REF: TURQ-434-4 Twist Drill for 1.5mm x 4mm screw. Rx only |  |  | March 31, 2011 | 2011 | The Anspach Effort, Inc. | Cutter was longer than was stated on the label | Pending | Anspach sent an Urgent Medical Device Product Removal letter dated March, 32, 2013, to all affeced customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inventory their batches and return the products for replacement. Customers needing assistance were instructed to contact Anspach at 800-327-6887.For questions regarding this recall, call 800-327-6887. |  | 29 | 29 cutting burs | Worldwide Distribution - USA including CA, CO, IN, KY, MD and TX and Internationally to Japan, Taiwan, United Kingdom, Sweden, and Germany. | Open, Classified |  |  | 628 | Not_Computer
Z-0844-2014 | 67188 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | KODAK DirectView DR 7500 Dual Detector System, MODEL DR 7500, Manufactured by Carestream Health, Inc. Made in U.S.A. |  |  | December 26, 2013 | 2013 | Carestream Health, Inc. | Carestream Health, Inc. has recalled DR 7500 Dual Detector System utilizing Version 5.X Software due to a possible patient image display error. | Software Manufacturing/Software Deployment | On December 26, 2013 Carestream Health, Inc. distributed Urgent Medical Device Recall notification letters dated December 19, 2013 to their customers via Fed Ex. Carestream Health, Inc. recommends that once the daily detector calibration process is initiated on the DR 7500 System running Version 5.X software it should not be interrupted. The calibration process should be completed only when patients are not being imaged. A Carestream Health Service representative will contact all customers to install a software modiciation to Version 5.X software which will not allow the cancellation of detector calibration. Customers will questions can call 1-800-328-2910 7 days per week on a 24 hour basis. |  | 164 | 164 units Total (42 domestically &amp; 122 internationally) | worldwide Distribution - US (Nationwide) including states of: IN, PA, MD, MI, NY, GA, MN, OK, VA, TX, WI, CA, IL, TN, WA, NC, SC, MA, NJ, and KY; and Internationally to: Canada, New Zealand, South Korea, Australia, Hong Kong, Saudia Arabia, Guinea, Kuwait, Israel, Oman, South Africa, China, UK, Sweden, Poland, Denmark, Germany, Finland, Czech Republic, Italy, France, Portugal and Spain. | Open, Classified |  |  | 164 | Computer
Z-0828-2014 | 67278 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Accuray CyberKnife Robotic Radiosurgery System; Accuray Incorporated Sunnyvale, CA. |  |  | January 13, 2014 | 2014 | Accuray Incorporated | Potential Safety issue with Synchrony Boom Arm Mounting Assembly - one complaint of mounting assembly detaching. | Device Design | Urgent Medical Device Correction notification letter were sent out beginning January 14, 2014 by trackable method. Letters describe potential problem and provide inspection instructions to be performed prior to use. |  | 44998 | 302 devices; 149 in US, 153 out of US. | Worldwide Distribution. - US (Nationwide) and Mexico; and Internationally to: Australia, Belgium, Canada, China, Colombia, Czech Republic, England, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Korea, Malaysia, Luxembourg, Netherlands, Poland, Russia, Saudia Arabia, Spain, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Kingdom, Venezuela, and Vietnam | Terminated | November 25, 2014 | 316 | 44998 | Not_Computer
Z-0832-2014 | 67306 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT &amp; Brilliance iCT SP Computed Tomography X-Ray Systems, Philips Healthcare, Cleveland, OH |  |  | December 09, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Patient images exhibited ring artifacts. | Software design | The firm, Philips Healthcare, sent an "URGENT - Medical Device Correction" letter dated December 9, 2013 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to do the following to avoid the occurrence of these specific ring artifacts:1. Disable Automatic Scan Time for all Exam Cards2. Leave Automatic Scan Time on during planning and prior to clicking Goto initiate the scan, click the acquisition series of the scan and disableAutomatic Scan Time, then proceed with the scan.3. Disable Edit Before Final Recon (Reconstruction) for all Exam CardsIf a ring artifact is present, clinical judgment should be used by thecustomer/user to determine if the images acquired in the dataset contain enough information to make a diagnosis.Philips will be releasing a Field Change Order (FCO) to correct the issue with software version release 4.1.1.19261. In addition, a Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, pleasecontact your local Philips representative or local Philips Healthcare office. ForNorth America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 7 | 7 units | Worldwide Distribution: US (nationwide) including states of: FL, IN, NY, OR and PA; and internationally to: Australia and Israel. | Terminated | May 29, 2014 | 171 | 7 | Computer
Z-0839-2014 | 66742 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | DeRoyal (R) Aquarius Hot/Cold Therapy Combo Unit, w/ Foot/Ankle Blanket NS, REF T656NS, Rx only, Manufacturer DeRoyal Industries, Inc. 200 DeBusk Lane, Powell, TN 37849. |  |  | September 17, 2013 | 2013 | DeRoyal Industries Inc | The unit displays an error code of hose kinking when no hose kink is present; the unit fails to monitor the temperature of the water in the blanket as intended and controlling it through the automatic temperature control. This failure could result in temperature injuries to the treatment site; and there have been two reports of electric shock. | Device Design | DeRoyal sent a "URGENT VOLUNTARY RECALL DEROYAL HOT/COLD THERAPY UNIT" letter dated September 17, 2013. The letter identified the product, problem, and actions to be taken by the customers. |  | 25 | 25 units | Nationwide Distribution. | Terminated | June 05, 2017 | 1357 | 1343 | Not_Computer
Z-0841-2014 | 66931 | [6 empty] | ZYM B Reagent (REF 70493) |  |  | November 08, 2013 | 2013 | Biomerieux France Chemin De L'Or | bioMÂżrieux identified a visual defect and activity issue on the ZYM B reagent (REF 70493); leading to a false negative results of some API biochemical tests using the ZYM B reagent. | Under Investigation by firm | bioMerieux sent an Urgent Product Removal Notice letter dated November 15, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to ensure the letter and guidance were distributed and received by all appropriate personnel in their organization, customers were instructed to discontinue use and destroy ZYM (B 70493), API NH (10400) API Listeria (10300) with ZYM B ampules having a visual defect.Customers still in possession of some kits of ZYM B (70493-Lots: 1) without visual defect that is used with API 20 Strep (20600), API Coryne (20900), API Staph (20500) and API ZYM (25200), have the option to:a) Use the kits of ZYM B (70493) still in your inventory with the following mandatory requirement for the API Strep and API Coryne: perform a daily quality control with the ATCCÂż strain described in the paragraph Reading and Interpretation of the package insert of each API strip prior to using the ZYM B ampules.b) Use the kits of ZYM B (70493) still in your inventory without any requirement with API Staph and API ZYM, orc) Discard all kits of ZYM B (70493) still in their inventory.If you are still in possession of some kits of API NH (10400-Lots: See Annex 1) without visual defect of ZYM B, you have the option to:a) Use the kits of API NH (10400) still in your inventory with the mandatory requirement: perform a daily quality control with the ATCCÂż strain 49917 Haemophilus paraphrophilus described in the paragraph Use of the reagents of the package insert of API NH prior to using the ZYM B ampules or,b) Discard all the kits of API NH (10400) still in your inventory.5. If you are still in possession of some kits of API Listeria (10300-Lots: See Annex 1) without visual defect of ZYM B, we ask you to use the API Listeria (10300) still in your inventory with the mandatory requirement: perform a daily quality control with the ATCCÂż strain 33090 Listeria innocua as |  | 39439 | 39439 kits (includes ZYM B reagents, API Listeria * API NH kits) | Worldwide Distribution - USA (nationwide) and Internationally to U.S. Nationwide; Algerie, Argentina, Armenia, Australia, Austria, Azerbaidjian, Bahrain, Bangladesh, Belarus, Belgium, Benin, Bolivia, Brazil, Brunei, Bulgaria, Burkina Faso, Cambodia, Cameroon, Canada, Chad, Chile, China, Colombia, Democratic Republic of Congo, Costa Rica, Cote d'Ivorie, Germany, Czech Republic, Dominican Republic, Dutch Antilles, Dutch Antilles, Ecuador, Egypt, Republica De El Salvador, Estonia, Germany, Ethiopia, France, Gabon, Georgia, Germany, Greece, Guatemala, Guinea, Herzegovina, Honduras, Hungary, India, Germany, Indonesia, Iraq, Israel, Italy, Jamaica, Japan, Iraq, Kazakhstan, Kenya, Korea, Kosovo, Kuwait, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Malaysia, Malaysia, Mali, Malta, Mexico, Romania, Mongolia, Morocco, Mozambique, Singapore, Netherlands, Managua, Niamey Niger, Nigeria, Nigeria, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russia, Uganda, Kingdom of Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sudan, Sweden, Switzerland, Syria, Tanzania, Italy, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, Uganda, Ukraine, UAE, United Kingdom, Uruguay, Venezuela, Vietnam, Vietnam, Yemen and Zimbabwe | Open, Classified |  |  | 39439 | Not_Computer
Z-0845-2014 | 67319 | No | NCX | Thermal regulating system. | Cardiovascular | Neurology | 510(k) | InnerCool RTx Endovascular System, Model Numbers: 861470 (120V), 861472 (240V). |  |  | January 13, 2014 | 2014 | Philips Healthcare | Phillips Healthcare initiated this action because the Main Control Board (MCB) may produce an inaccurate low patient temperature reading and low temperature alarm. | Nonconforming Material/Component | Philips Healthcare Recall Notice for this proactive performance correction was provided verbally to hospital biomedical engineers. Philips product support engineering scheduled and conducted the update on customer devices to the most current MCB configuration. |  | 50 | 50 | Worldwide Distribution - US (Nationwide) and Internationally to: Austria, France, Germany The Netherlands, Spain, Sweden, and Switzerland. | Terminated | February 13, 2014 | 31 | 50 | Not_Computer
Z-0831-2014 | 66197 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT and Brilliance iCT SP |  |  | July 16, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | An artifact that resembles thrombus may appear on the image. | Device Design | Philips sent an Field Safety Notice dated July 15, 2013, to all US customers via certified mail. Market representatives will distribute the letter internationally. The letter identified the product the problem and the action needed to be taken by the customers. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office.For North America and Canada, contact the Customer Care Solutions Center (1"800-722"9377,Option 5: Enter Site 10 or follow the prompts).This notice has been reported to the appropriate Regulatory AgencyPhilips apologizes for any inconveniences caused by this problem |  | 506 | 506 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states AL, AR, AZ, CA, DC, DE, FL, GA, IA, IL, IN, KY, MA, MD, MN, MO, MS, NC, NV, NY, OH, OK, OR, PA, SC, TX, VT, WA, and WV., and the countries of Argentina, Australia, Austria, Azerbaijan, Bangladesh, Belgium, Brazil, Chile, China, Czech Republic, Denmark, Egypt, France, Germany, Greece, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Lebanon, Netherlands, New Zealand, Norway, Oman, Panama, Philippines, Poland, Qatar, Romania, Russia, Saudi Arabia, Singapore, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Uzbekistan. | Terminated | September 19, 2014 | 430 | 506 | Not_Computer
Z-0840-2014 | 66917 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Optima CT520 and Optima CT540 Computed Tomography (CT) systems. |  |  | October 09, 2013 | 2013 | GE Healthcare, LLC | There is an issue with the Manual Film Composer feature on some CT products. There is an opportunity, while following a specific workflow, to create a film image with one patient's images and another patient's name in the footer. This is also an issue when within Manual Film Composer, if a color image is selected for printing, another patient's black and white image appears in the preview screen | Software design | Consignees were sent on 10/9/2013 a GE Healthcare "Urgent Medical Device Correction" letter dated October 9, 2013. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Directors and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Contact your local GE Healthcare Service Representative or the GE Healthcare Call Center : US or Canada : 800-437-1171 and Other countries: 262-896-2890). |  | 272 | 272 (29 US; 243 OUS) | Worldwide Distribution-USA (nationwide) including the states of AL, CO, FL, IN, KS, KY, MN, NC, OH, TX, and VA, and the countries of ALGERIA, AUSTRALIA, BELGIUM, BULGARIA, CANADA, CHINA, ABIDJAN COTEDIVOIRE EGYPT, FRANCE, GERMANY, GREAT BRITAIN, INDIA, IRAQ, ITALY, JAPAN, KENYA, LATVIA, LEBANON, MALAYSIA, MARTINIQUE,MOROCCO, NEW ZEALAND, NORWAY, PAKISTAN, POLAND, PORTUGAL ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWITZERLAND, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, and VIETNAM. | Terminated | June 19, 2014 | 253 | 272 | Not_Computer
Z-0833-2014 | 67171 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Symmetry and CODMAN brands, QUAD-LOCK Sterilization Container. Sterilization container systems manufactured and distributed under the following names: |  |  | December 05, 2013 | 2013 | Symmetry Medical/SSI | The Quad-Lock Sterilization Container System may not maintain an effective sterile barrier. | Process control | The firm notified their consignees of the recall by letter and email on 12/02/2013 instructing customers not to use the containers until further notice from Symmetry Surgical. |  | 2991 | 2,991 units | Worldwide Distribution - US (Nationwide) and PR; and Internationally to: Italy, Australia, Saudi Arabia, UK, China, Canada, Germany, New Zealand, Colombia, Korea, and Chile. | Terminated | April 13, 2015 | 494 | 2991 | Not_Computer
Z-0857-2014 | 67199 | Yes | DWA | Cardiopulmonary bypass pump speed control. | Cardiovascular | Cardiovascular | 510(k) | CP5 centrifugal pump system |  |  | December 04, 2013 | 2013 | Sorin Group Deutschland GmbH | Sorin is initiating a field correction on S5/C5 Heart-lung machine and CP5 centrifugal pump system due to reports of pump speed control knob failures resulting in no response. | Device Design | Sarin Group US distributed a Field Safety Notice dated December 5, 2013 to all affected customers and distributors. The letter informed users of the potential for shaft encoder failures and notified them that they would be contacted by the Sorin Service Group to arrange an appointment to repair the units onsite. Customers and distributors were also provided with a Response form to confirm they had received, read and understood the Field Notice. They were instructed to return the completed form to assist in monitoring the effectiveness of the communication. For questions call 303-467-6527 or 303-467-6503. |  | 75 | 75 | Worldwide Distribution - US (Nationwide) and PR; and Internationally to:BE, ES, FI, GB, NL, SE, AR, AT, AU, CA, CH, CN, CR, EG, HK, ID, JO, JP, KR, KW, LB, LY, MM, PL, RO, RS, RU, SG. TH, TN, TR, TW, UA, VN, and ZA. | Terminated | December 28, 2015 | 754 | 398 | Not_Computer
Z-0849-2014 | 67055 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Teflon Tube, sterile |  |  | November 07, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker became aware during laboratory testing that there is a potential that the seal integrity of the outer pouch (sterile barrier) may be compromised for certain lots of the product Teflon Tube, sterile. | Process design | The firm, Stryker, sent an "URGENT MEDICAL DEVICE REMOVAL" letter dated December 17, 2013 to its customers via e-mail. Preliminary notification letters and Product Accountability Forms were sent via FedEx on November 15, 2013 with return receipt to the Stryker locations. Notification and Product Accountability Forms were sent 12/18/2013 via Fed Ex with return receipt. Updated letters were sent via Fed Ex on 12/19/2013 with return receipt. The letter described the product, problem and actions to be taken. The customers were instructed to quarantine affected devices in local inventory as well as retrieve any devices located in hospitals; inform users of the Medical Device Removal and pass this to all those individuals who need to be aware within your organization; return all affected products available to your location to: Stryker Osteosynthesis, c/o Colleen O'Meara, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ, 07430, REF: PFA#2013-168 or contact Stryker customer service and refer to PFA #2013-166 for returning the product to us; complete and return the enclosed Business Reply Form and fax a copy to: 1-865-251-3635 or email a copy to Recall Coordinator at aminah.crawford@stryker.com, and keep a copy of the completed and executed Business Reply form for your records.If you have any further questions, contact Manager, Regulatory Compliance at 1-201-972-2100 or email: colleen.omeara@stryker.com. |  | 34 | 34 units | US Distribution in states of: GA, IL, KY, MI, MO, NJ, TN, UT and VA. | Terminated | March 31, 2015 | 509 | 34 | Not_Computer
Z-0846-2014 | 67213 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | Simplexa Flu A/B &amp; RSV Direct assay, Model MOL2650. |  |  | January 10, 2014 | 2014 | Focus Diagnostics Inc | Focus Diagnostics is initiating an urgent safety notice correction for Simplexa Flu A/B &amp; RSV Direct assay (MOL2650) because Focus Diagnostics has received some customer complaints of Simplexa Flu A/B &amp; RSV lots due to sporadic false positive RSV signals, which may result in a higher RSV false positive rate. | Pending | Focus Diagnostics sent an Urgent Safety Notice Correction letter dated January 10, 2014, via email and a hard copy letter by mail. The letter informed customers of the correction, product description with codes, product intended use, issue, risk to health, background, recommendation, actions by customers, and contact information. Customers were instructed to check their inventory for the affected product, if customers detect any positive RSV results, confirm the positive result by an alternate method. Customers were asked to complete the "URGENT: SAFETY NOTICE Acknowledgement and Receipt Form Response Required". Customers with questions were instructed to contact their Technical Services department at 800-838-4548, select option 3.For questions regarding this recall call 714-220-1900. |  | 1052 | 1,052 kits | Worldwide Distribution - US (Nationwide) including states of: NY, MA, NJ, OH, CA, HI, CO, MN, FL, PA, IA, WI, AZ, TX, TN, IL, SD, MO, NM, AR, WV, MI and Internationally to: Sweden, Israel, Canada, and Slovenia. . | Terminated | October 15, 2014 | 278 | 1052 | Not_Computer
Z-0850-2014 | 66832 | No | GXI | Radiofrequency lesion probe. | Neurology | Neurology | 510(k) | Kimberly-Clark RadiOpaque Radiofrequency Cannula, Gauge 22, Length 100 mm, Active Tip 10 mm, Distributed in the US by Kimberly-Clark,Product code PMF22-100-10 |  |  | October 29, 2013 | 2013 | Kimberly-Clark Corporation | Product may contain a cannula with an active tip length of only 5mm rather than 10mm as indicated on the label. | Employee error | The firm, Kimberly-Clark Corporation, sent an "URGENT: Voluntary Medical Device Recall Notification" letter dated October 29, 2013 to all direct customers and distributors via federal express priority overnight. The letter described the product, problem and actions to be taken. The customers were instructed to evaluate your Facility's inventory of Kimberly-Clark RadiOpaque Radiofrequency Cannula to determine if any of the two impacted product lots remain within your facility. If product is present, discontinue use of the impacted product and place in a quarantined location; complete and return the attached Product Recall Response Form-Customers via fax to: 1-920-225-4182 within five (5) business days of receipt of this letter (complete and return form even if either of the impacted product lots is NOT present in your facility). Once response form is received, a Kimberly-Clark representative will contact you regarding return and replacement of the product. If you require further assistance, please call Kimberly-Clark Customer Service at 800-KC-HELPS (800-524-3577). |  | 550 | 55 cases (10 units per case = 550 units) | US Distribution in states of: AZ, CA, DE, FL, LA, MD, MA, NJ, NY, OH, PA and TX. | Terminated | September 03, 2014 | 309 | 550 | Not_Computer
Z-0855-2014 | 66969 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Orthofix Firebird Spinal Fixation System Instrumentation Modular Screw Driver, |  |  | November 26, 2013 | 2013 | Orthofix, Inc. | Orthofix received 6 complaints which resulted in a reportable events due to extended surgical times greater than 30 minutes for the Modular Screw Driver (PN 52-1332). The complaints alleged that the screw driver's collet would malfunction resulting in the surgeon being unable to use the Modular Screw Driver to effectively place Modular Screws, which may result in a delay of surgery. | Nonconforming Material/Component | Orthofix sent an Urgent Medical Device Recall Notification letter dated November 26, 2013 by certified mail to distributors. The letter identified the affected product, problem and actions to be taken. The letter instructed distributors to immediately cease any further distribution, return affected products to Orthofix at no cost and complete the attached reply form. For questions contact your local Orthofix representative. |  | 443 | 443 | Worldwide Distribution - US Nationwide including (Puerto Rico) and the countries of: Spain, Germany, Australia, Peru, Mexico, Colombia, South Africa, and Israel. | Terminated | December 09, 2014 | 378 | 443 | Not_Computer
Z-0848-2014 | 67100 | No | BTI | Portable air compressor. | Anesthesiology | Anesthesiology | 510(K) Exempt | Compressor Mini |  |  | December 05, 2013 | 2013 | Maquet Cardiovascular Us Sales, Llc | Maquet Compressor Mini may deliver compressed air at a temperature higher than specified. | Material/Component Contamination | Maquet sent a product recall urgent-medical device field correction letter/customer fax response via Fed Ex on December 5, 2013, to their customers.Customers are advised of the issues with the device and provided instructions. The correction will include an update with replacement of the affected connector and if needed replacement of the thermoelectric cooler if it damaged. These operations will be performed by an authorized Maquet representative. Any questions can be directed to customer service at 1-888-627-8383 option 2 followed by option 3 Monday through Friday between the hours of 8 am and 6 pm EST.Customers should completed and return the enclosed confirmation form or scan a copy to field actions@maquet.com or fax (973) 807-9290. |  | 237 | 237 units | Worldwide Distribution US, Canada and Japan. | Terminated | February 05, 2015 | 427 | 237 | Not_Computer
Z-0847-2014 | 67207 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Tapered Screw-Vent Implant, MTX, 4.7mm x 11.5mm, 4.5mm Platform. Catalog Number TSVWB11, Lot 62470008. |  |  | January 06, 2014 | 2014 | Zimmer Dental Inc | Zimmer Dental is conducting a voluntary recall of a single lot of the Tapered Screw-Vent Implant, Catalog TSVWB11, Lot 62470008, because some of the packages of this lot may have the Cap Label state Âż3.7 x 10mm instead of Âż4.7 x 11.5mm. | Labeling Change Control | Zimmer Dental initiated the recall on 01/06/2013 by sending out recall notifications via FedEx to domestic customers and via electric mail to international customers.Customers are sent a letter identifying the issue, product description with codes, risk, instructions, and contact information. The letter's instructions include locating the affected product and return unused product to Zimmer Dental and notifying the amount of product that has been consumed. |  | 307 | 307 | Worldwide Distribution-USA (nationwide) and the countries of Thailand, Japan, Hong Kong, Morocco, and UAE. | Terminated | July 07, 2014 | 182 | 307 | Not_Computer
Z-0851-2014 | 66585 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | User Defined Method Flex Assignment/Siemens Dimension Vista 500 or Dimension Vista 1500 System with software version 3.5.1 or lower |  |  | October 07, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | When utilizing the Routine Inventory screen to enter a User Defined Method (EMPTY) Flex the system may assign the User Defined Method Flex to a different Flex that is currently in inventory on the system, and then use the incorrect Flex Cartridge to process the user defined method. | Software design | Siemens sent a Urgent Medical Device Correction letter dated October 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Please use the following actions if using Vista SW Versions 3.5.1 and below to prevent the issue from occurring:1. When assigning a User Defined (EMPTY) Flex, use the Advanced Inventory Screen per the Operators Guide.2. Verify that EMPTY Flex is no longer listed in the inventory before proceeding with any testing.3. If the EMPTY Flex is still listed in the inventory, contact your Siemens Customer Care Center - Technical Solutions for assistance.Please retain this letter with your laboratory records and forward this letter to those who may have received this product. Discuss the content of this letter with your Medical Director.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center - Technical Solutions at (800-441-9250) or your local Siemens technical support representative. |  | 2038 | 2038 | USA (nationwide) Distribution including Puerto Rico. | Terminated | August 27, 2014 | 324 | 2038 | Not_Computer
Z-0852-2014 | 67117 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | The ACUTE Innovations Modular RibLoc System (RibLoc U plus) is comprised titanium plates and screws that are used to stabilize rib fractures, fusions and osteotomies during the normal bone healing process. |  |  | October 30, 2013 | 2013 | ACUTE Innovations, LLC | The ACUTE Innovations Ribloc U plus Rib Fracture Plating System is recalled because it has the potential to malfunction during installation in a surgery. | Under Investigation by firm | The firm sent the Urgent Notice: Device Recall letter, dated 30 Oct. 2013, to their consignees.The recall notification letter requested consignees to return the product to Acute and a return authorization number provided.Consignees should contact Acute Innovations Customer Service at 866-623-4137 for any questions regarding this recall notification. |  | 176 | 176 units | US Distribution in states of Arizona, Georgia, Illinois, Indiana, Utah, and Wisconsin. | Terminated | March 13, 2014 | 134 | 176 | Not_Computer
Z-0854-2014 | 66982 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CARESCAPE Monitor B650 and CARESCAPE Monitor B850. |  |  | December 27, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of potential safety issues associated with the CARESCAPE Monitor B850 and CARESCAPE Monitor B650. These included 25 issues related to NIBP (Non-InvasiVentilation Alarms:ve Blood Pressure) , ECG, Central Monitor, General, Tram, and Bed to Bed Issues. | Software design | GE Healthcare sent an Urgent Medical Device Correction letter between December 27, 2013 and June 27, 2013, to all affected customers. The letter described the Safety Issues and Instructions, Affected Product Details, Product Correction and Contact Information. Customers with questions were instructed to contact Technical Support at 1-800-558-7044 or their local Service Representative.For questions regarding this recall call 800-558-7044. |  | 40239 | 40,239 (17,182 units US, 23,147 units OUS). | Nationwide Distribution including DC and PR, except DE, SD, ND, and VT. OUS: Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belgium, Bermuda, Bolivia,Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic ,Denmark, Ecuador, Egypt, El Salvador, Estonia ,Finland, France, Germany, Greece,Guam, Guatemala, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kuwait,Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Malta, Mexico, Morocco,Namibia, Nepal, Netherlands, New Zealand, Nicaragua, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia,Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Tanzania, United Republic of Tunisia, Turkey, United Arab Emirates, United Kingdom, Venezuela, Uruguay. | Open, Classified |  |  | 40239 | Not_Computer
Z-0853-2014 | 67325 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Innova IGS 630. Biplane Cardiovascular and Interventional Imaging System. |  |  | November 15, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with respect to IGS 630 Imaging Systems.While performing fluoroscopy on a biplane fluoroscopy unit, there is a potential for loss of the x-ray imaging function when the user changes the size of the lateral FOV (Field of View) and releases the biplane footswitch pedal simultaneously. This could lead to a system lock up requiring | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated November 15. 2013. The letter was addressed to Hospital Administrators / Risk Mangers, Mangers of Radiology / Cardiology / Surgery Department and Radiologists / Cardiologists / Surgeons. The letter described the Safety Issue, Safety Instructions, Affected Product Details, and Product Correction. The letter stated that GE Healthcare will correct all affected systems at no cost. For questions call (In US 800-437-1171; for other countries contact your local GE Healthcare Service Representative). |  | 28 | 28 (12 US, 16 OUS). | Worldwide Distribution - US Nationwide in the states of CA, CO, FL, NV, NJ, PA, TX and the countries of CZECH REPUBLIC, FRANCE, JAPAN, REPUBLIC OF KOREA, SAUDI ARABIA, and TURKEY. | Terminated | March 28, 2014 | 133 | 28 | Not_Computer
Z-0861-2014 | 65179 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Codman Certas Programmable Valve In Line Valve only with SiphonGuard Device |  |  | May 16, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Codman Certas Programmable Valves used for hydrocephalus may not operate properly | Device Design | Codman issued an Urgent Notice dated May 16, 2013, to all affected customers. The letter identifed the affected product, problem and actions to be taken. Mitigation instructions are provided and if not followed or are not successful, either issue could lead to symptoms associated with under- or over-drainage. Units not implanted are to be returned. A U.S Acknowledgement Form and fax are requested to be completed and faxed to: 1-888-943-4193, or e-mail the form to codman8894@stericycle.com. If you have any questions on concerns in regards to this recall, please contact your Codman Representative or Stericycle at 1-888-943-2395. |  | 1481 | 1481 units | Worldwide Distribution - US Nationwide, Australia, Austria, Belgium, Brazil, Switzerland, Czech Republic, Germany, Spain, Finland, France, Great Britain, Israel, Italy, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Sweden, Slovenia, Slovakia, Turkey, South Africa, and UK. | Terminated | June 02, 2015 | 747 | 5791 | Not_Computer
Z-0791-2014 | 66311 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT, Brilliance 64 and Brilliance Big Bore |  |  | March 07, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | The DoseRight feature suggest a mAs based on the measured patient size, a reference size and a reference mAs. When scanning large children, the suggested mAs may be higher than clinicians would expect. | Software design | Philips Medical Systems sent a Field Safety Notice letter dated February 25, 2011, to all affected customers.The letter identified the product the problem and the action needed to be taken by the customers.Your Corrective Action Plan (CAP) is approved and consists of:1. Notifying all affected users with Field Safety Notice 72800527/72800528 starting March 8, 2011. This Field Safety Notice has been reviewed and is approved.2. Following up with all responses by dispatching a Field Service engineer to each site to install the software update at no cost to the user.3. Tracking the responses from users by certified mail.For further questions please call (440) 483-7600. |  | 23 | 23 | Worldwide Distribution - US Distribution including the states of DC, IL, IN, MI, MO, NY, OH, OR, PA, TX and VT., and the countries of Australia, Belgium, China, Germany, India, Israel, Italy and The Netherlands. | Terminated | April 28, 2014 | 1148 | 23 | Computer
Z-0774-2014 | 67060 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Carestream DRX Evolution System (Standard Q) |  |  | August 20, 2013 | 2013 | Carestream Health Inc. | Carestream received a report of an exposure being initiated without a patient image captured. The patient was imaged 2 times before an image was captured. The field engineer who visited the site reported the system allowed an exposure without the assignment of a bucky or image receptor. | Radiation Control for Health and Safety Act | The Center for Devices and Radiological Health (CDRH) understands that the CAP (Corrective Action Plan) involving 4 units will consist of the following:a. Distribution of a letter to all affected customers informing them of the issue, how to prevent the issue including a reminder of instructions for proper x-ray system setup and operation found in the User Manual. The letter will also include actions to be taken by Carestream to correct the issue. This communication will be marked and distributed pursuant to the requirements of 21 CFR 1003.21.b. The next stage of correction will be the installation of a software patch to any potentially affected customer sites. The software patch has been developed to prevent the system from allowing an exposure when the bucky and detector are not selected. This patch has been verified and proven successful prior to rollout in manufacturing. Once receiving FDA approval on the corrective action plan, the correction will be implemented within the US to all affected devices at identified customer sites.c. The completion of the Corrective Action Plan, once approved, is expected to take approximately 3 months.d. Carestream will provide FDA with a report demonstrating effective closure of the activity.CDRH approves the draft of the customer notification letter and CAP subject to the following condition:Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.For further questions please call 585-627-6473. |  | 4 | 4 units | Nationwide Distribution including TX, IL, and GA. | Open, Classified |  |  | 4 | Computer
Z-0865-2014 | 64795 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Centricity Laboratory System is intended to be an information system designed to support the clinical and administrative activities associated with the provision and utilization of clinical laboratory services and facilities, e.g., the storing and delivering of analytical results. It is a specially designed data program application (software), which is supplied for installation in existing mainframe or decentralized computers or a computer network. This software system will not be the software directly used to run, steer, or control any specific laboratory analyzer or equipment. It is intended for use in laboratory facilities, including a central laboratory or in a multiple laboratory environment servicing satellite laboratories, reference laboratories, clinics, etc. Centricity Laboratory System is not intended for direct patient contact. |  |  | November 20, 2012 | 2012 | GE Healthcare It | GE Healthcare is aware of a potential safety issue associated with the use of GE Centricity Laboratory Core Lab calculated results in two scenarios. In Scenario 1: Calculated results greater than 6 digits plus a decimal are reported as 0 instead of TOO BIG when significant figures are not defined for the item. In Scenario 2: Calculated results greater than 6 digits plus a decimal are trunc | Software design | An URGENT MEDICAL DEVICE CORRECTION letter date March 12, 2013 was sent to customers. The letter requests that customers review all calculated results to determine if either of the scenarios could occur at their site. If a customer determines that either of the scenarios proposed in the letter could occur at their site, the customer can contact their GE service representative to request or install the software update. Customers with questions can contact the GE Healthcare IT Helpdesk at 1-847-589-8496 or tgg.nahelpdesk@med.ge.com. |  | 7 | 7 units | Worldwide Distribution - US (Nationwide) to states of: MN, GA, and CA; and International to: England. | Open, Classified |  |  | 7 | Computer
Z-0900-2014 | 67298 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | Atlantis PV, 8.5F, 15 MHz Peripheral Imaging Catheter: |  |  | January 06, 2014 | 2014 | Boston Scientific Corporation | Reports of physicians experiencing inability to pass a 0.035" guidewire through the guidewire lumen and exit port at the (proximal) Y-manifold of the imaging catheter. | Equipment maintenance | Boston Scientific sent an Urgent Medical Device Recall Removal letter dated January 14, 2014 to all affected customers by overnight mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete an Account Reply verification Tracking Form in order to received a Return Goods Authorization number for product return. For questions contact your local Sales Representative. |  | 1068 | 1068 single units | US Nationwide Distribution and Mexico | Terminated | October 29, 2014 | 296 | 1068 | Not_Computer
Z-0887-2014 | 67147 | No | HQB | Ophthalmic photocoagulator. | Ophthalmic | Ophthalmic | 510(k) | 27ga ONE-STEP (tm) Endo Illuminator, Synergetics, inc., length 7.5 ft, sterile / EO, 56.02.27P. Packed 12 units per box. Manufactured for Synergetics, Inc., O'Fallon, MO. |  |  | December 09, 2013 | 2013 | Synergetics Inc | Faulty seals on the outer pouch of various, double-pouched, single-use, sterile, ophthalmic devices could result in seal failure. | Use error | Domestic customers were sent recall notices by FedEx on December 9, 2013. Foreign customers were sent recall notices by e-mail between December 6 and December 13, 2013. The recall notice instructed customers to return the recalled product. |  | 96 | 96 units | Worldwide Distribution-USA (nationwide) and the countries of Argentina, Austria, Brazil, Canada, China, Egypt, France, Germany, Greece, Hong Kong, Italy, Japan, Mexico, Morocco, the Netherlands, Pakistan, South Korea, Spain, Sweden, Switzerland, and the United Kingdom. | Terminated | March 25, 2015 | 471 | 2855 | Not_Computer
Z-0870-2014 | 67245 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | BD BACTEC FX 40 instrument, catalog number 442296, manufactured by BD Diagnostic Systems, 7 Loveton Circle, Sparks, MD 21152. |  |  | December 09, 2013 | 2013 | Becton Dickinson &amp; Co. | Improperly functioning component of a diagnostic medical device may cause false negative results in clinical specimens. | Nonconforming Material/Component | BD Diagnostics initiated a field correction by phone on December 9, 2013, to affected customers and followed the calls with letters, e-mails, or fax. BD Diagnostics sent a secondary Urgent Field Corrective Action letter dated November 2013 to all affected customers. The recalling firm Secondary notifications were serial number specific and explained that some BD BACTEC FX instruments contain racks that may demonstrate variable indicator light intensities. Specifically, the Anonymous vial indicator lights which should appear yellow may instead appear more green in color. BD noted that affected instruments continue to function properly and will not report an incorrect result. The issue is limited to the yellow indicator lights only. Red and green indicator lights appear as intended. BD is developing software to correct this issue and anticipates that the new software will be available to begin instrument updates by February, 2014. The new software will not affect instrument if the indicator lights are functioning as intended. Until the software can be updated on your instrument, BD provided workflow instructions to ensure that the indicator light issue does not occur.For questions regarding this recall call 410-316-4054. |  | 2 | 2 units | Worldwide Distribution - USA (nationwide) and Internationally to IN, AU, CN, AR, BE, PE, KR, BR, CO, SG, TH, BN, HK, JP, and CA. | Terminated | August 20, 2014 | 254 | 200 | Computer
Z-0871-2014 | 67243 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | ARIATELE TELEMETRY TRANSMITTER, Model 96281, with SpO2 Option C. |  |  | December 19, 2013 | 2013 | Spacelabs Healthcare, Llc | The AriaTele Model 96281 with SpO2 monitoring Option C is recalled because the AriaTele (transmitter) display and the Telemetry Central Station display may show a SpO2 value when the sensor is not connected to the patient or in certain sensor failed conditions. | Software design | On January 6, 2014, Spacelabs sent an "URGENT - MEDICAL DEVICE CORRECTION Spacelabs Healthcare AriaTele Model 96281 with SpO2 Option C" letter, dated 03 January 2014 to US consignees.On January 9, 2014, Spacelabs e-mailed the "URGENT - MEDICAL DEVICE CORRECTION Spacelabs Healthcare AriaTele Model 96281 with SpO2 Option C" letter (translated as necessary), dated 03 January 2014, to foreign consignees.The letter described product, problem and actions to be taken. The Consignees where instructed to follow the "Advise on what action the clinician takes" on the letter.Spacelabs Healthcare will contact you at the earliest possible date to schedule a convenient tie for Spacelabs to correct all of your facility's affected transmitters (software update) at no cost.If you have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 72933 | 1057 units (988 in the US and 69 outside of the US). | Worldwide Distribution: US (nationwide) including states of: AZ, CA, FL, GA, IA, IL, LA, MA, ME, MI, MO, MS, NY, PA, SC, TN, TX, VA, and Puerto Rico; and Internationally to: Canada, Chile, and Panama. | Terminated | June 12, 2014 | 175 | 72933 | Computer
Z-0866-2014 | 67203 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | SET SCREW RETAINING DRIVER, REF 9010000820, QTY: 1 EA, Medtronic Sofamor Danek, USA, Inc. |  |  | January 03, 2014 | 2014 | Medtronic Sofamor Danek USA Inc | The retaining tabs component of the Set Screw Retaining Drive may be oversized, which will not allow proper mating between the Set Screw Retaining Drive and the set screw. | Reprocessing Controls | The recall notices were mailed directly to Risk Managers and Sales Representatives, with Recall Questionnaires and Confirmation forms included, on 01/03/2014. |  | 7 | 7 units | US Distribution in the states of CA, NC, TX, IN, NJ, and MI. | Terminated | May 16, 2014 | 133 | 7 | Not_Computer
Z-0901-2014 | 67206 | No | BWN | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Radial Assist RAD BOARD, used to support the |  |  | December 18, 2013 | 2013 | Merit Medical Systems, Inc. | Merit Medical Systems, Inc. is voluntarily recalling one lot (B507171) of RAD BOARD RB 100 devices. The affected devices are missing the main label which graphically depicts patient orientation on the board. The issue may allow the board to be oriented in the wrong direction where the semi-radiopaque material in the board (embedded Xenolite) could obscure patient anatomy during fluoroscopy. There | Labeling Change Control | Merit Medical Systems, Inc.sent a letter voluntarily recalling one lot (B507171) of RAD BOARD" RB100 devices. The affected devices are missing the main label which graphically depicts patient orientation on the board. The issue may allow the board to be oriented in the wrong direction where the semi-radiopaque material in the board (embedded Xenolite) could obscure patient anatomy during fluoroscopy. There have been no reports of patient harm or injury from Merit customers as a result of this issue. Merit has chosento remove this lot from the field by replacing the unlabeled RAD BOARDs. Merit requests that customers stop using these devices immediately until the product has been replaced by their Merit sales representative.Customers with questions were instructed to call 1-801-208-4344 or 1-801-208-4365.For questions regarding this recall call 801-208-4301. |  | 11 | 11 | Worldwide Distribution - USA including NJ, TX, WV, MN, GA, and Internationally to Mexico and Japan. | Terminated | February 13, 2014 | 57 | 11 | Not_Computer
Z-0867-2014 | 66055 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | T3F, Free T3 Assay, REF/Catalog Number LKF31, Siemens Material Number (SMN) 10381626; an IVD Immunoassay kit for use with the IMMULITE/IMMULITE 1000 Analyzers --- Shipping or unit package: 100, 200, 500 and 600 test kits --- Origin: UK Siemens Healthcare Diagnostics Products Ltd. Llanberis, Gwynedd, LL55 4EL UK. |  |  | August 27, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics confirmed customer complaints regarding an increase in the number of euthyroid patients (those with normal function of the thyroid) demonstrating values above the recommended normal range as published in the Instructions For Use for the IMMULITE/IMMULITE 1000 and/or IMMULITE 2000/IMMULITE 2000 XPi for several lots of Free T3 kits. A positive bias in quality control | Nonconforming Material/Component | Urgent Medical Device Recall Letters (dated 8/27/13) and Field Recall Effectiveness Check Forms were sent to the affected customers (starting on 8/27/13) informing them to discontinue use of and discard the IMMULITE Free T3 reagent kit lots. Customers were instructed to contact their local Siemens representative for assistance. Customers are requested to complete and return the Effectiveness Check Form included with the recall letter within thirty (30) days. |  | 2201 | 2201 kits in total | Worldwide Distribution-USA (nationwide) and the countries of Albania, Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bulgaria, Bosnia-Herzegovina, Brazil, Belarus, Canary Islands, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, France, Germany, Great Britain, Greece, Guatemala, Honduras, Hungary, Iceland, India, Indonesia, Ireland, Israel, Iran, Italy, Jordan, Republic of South Korea, Kuwait, Kazakhstan, Latvia, Lebanon, Libya, Lithuania, Macedonia, Mexico, Malaysia, Malta, Nepal, Netherlands, Norway, Panama, Peru, Philippines, Pakistan, Paraguay, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, Spain, Sri Lanka, South Africa, Switzerland, Turkey, Thailand, Taiwan, Uruguay, Usbekistan, Venezuela, Vietnam, and United Arab Emirates. | Open, Classified |  |  | 2201 | Not_Computer
Z-0905-2014 | 66835 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Medtronic Advanced Energy Aquamantys3 Pump Generator; Product Catalog Number: 10-1357 (Generator assembly), 40-404-1(as shipped, including accessories) |  |  | November 12, 2013 | 2013 | Medtronic Advanced Energy, LLC | The internal protection circuitry on the electrosurgical output that prevents a patient from exposure to DC voltage when the device is activated is being compromised during normal use. | Device Design | Medtronic sent an Urgent Product Removal Notification on November 12, 2013, via Next Day FEDEX. The communication advises users to immediately stop using the affected AQM3 electrosurgical generators and quarantine them until they can be returned to Medtronic Advanced Energy. Medtronic field personnel will collect them in order to remove them from service and return them to Medtronic Advanced Energy. Customers with questions were instructed to contact Customer Service at 866-777-9400.For questions regarding this recall call 866-777-9400. |  | 242 | 242 devices | Nationwide Distribution including CA, MO, PA, NJ, NY, OH, VA, TX, NC, DE, KS, TN, LA, IL, MI, SC, WI, and FL. | Terminated | April 03, 2014 | 142 | 242 | Not_Computer
Z-0904-2014 | 66912 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA N Control, Model No. A11A01653 |  |  | November 21, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling the ABX PENTRA N Control because it has been updated to include a revised Target Value assignment for Potassium. | Under Investigation by firm | Horiba sent an customer notification letter dated November 32, 2013, to all affected customers. The letter was sent via email to all customers who purchased the ABX PENTRA 400. HORIBA Medical is informing all ABX PENTRA N Control customers running the Potassium-E assay on the ABX PENTRA 400 with I.S.E. Module that the Target Value and Target Range for Potassium has been updated. The letter informs the customers of the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days by fax (949) 453-0600 Customers with questions are instructed to contact their local HORIBA Medical representative.For further questions please call (949) 453-0500 X 208 |  | 698 | 698 units total (80 units in US) | Worldwide Distribution - USA (nationwide) | Terminated | October 28, 2014 | 341 | 698 | Not_Computer
Z-0902-2014 | 67191 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Neptune 2 Rover Ultra Waste Management System. |  |  | December 30, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | The purpose of this correction is to upgrade the Neptune 2 Waste Management Systems in the field to the recently obtained 510(k) clearance specifications. | Device Design | On 12/30/13, Stryker sent an URGENT MEDICAL DEVICE UPDATE notification to consignees including government accounts and Sales Representatives, Stryker Canada and 3 in-house accounts and 1 account for lost rovers. Domestic and government consignees were sent the notification by mail FedEx overnight, and Sales Representatives and Stryker Canada were sent the notification by email. Notification described the affected product, instructions for the consignees, and upgrade information. Stryker also stated that a training kit and in-servicing will be provided to consignees. For questions regarding this recall, please contact Stryker Instruments: Monday-Friday 8am-5pm (EST)Neptune Customer Care Center 855-458-7441 or 269-389-2316strykerinstrumentsrecalls@stryker.com |  |  |  | Nationwide and Canada, Germany, Australia, New Zealand | Terminated | January 12, 2016 | 743 |  | Not_Computer
Z-0910-2014 | 67152 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | MooreMedical Supporting Health &amp; Care...MooreBrand Blood Glucose Monitor for Self-Testing &amp; Multi-Patient Use TRUEtrack. Manufactured by: NIPRO DIAGNOSTICS, Fort Lauderdale, FL 33309 U.S.A. |  |  | October 02, 2013 | 2013 | Nipro Diagnostics, Inc. | Manufactured with an incorrect factory set unit of measure. | Software design (manufacturing process) | Nipro Diagnostics sent an Urgent Medical Device Voluntary Recall letter dated December 31, 2013 to all affected customers. The customers were instructed to do the following: 1) Screen their inventory immediately for any of the serial numbers listed on the back of the products. 2) If they have any of the affected products, return the product by contacting Stericycle at 1-866-236-4518. 3) Complete and return the enclosed response form via fax at 1-888-358-7294. Customers with questions were instructed to call Stericycle at 1-866-236-4518.For questions regarding this recall call 954-677-9201. |  | 105 | 105 | Worldwide Distribution - USA (nationwide, HI and Puerto Rico and Internationally to Canada, European Union, Jamaica, Suriname, Trinidad and Tobago, Kazakhastan, Russian, Turkey, Australia, Bangladesh, Pakistan, Cambodia, Laos, Myanmar, Thailand, Vietnam, and India. | Terminated | January 28, 2016 | 848 | 606 | Not_Computer
Z-0907-2014 | 67138 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens SOMATOM Definition, Definition AS, Definition Flash, and Definition Edge systems. |  |  | November 15, 2013 | 2013 | Siemens Medical Solutions USA, Inc | During clinical operation, a malfunction may occur causing a system display freeze with SOMATOM Definition, Definition AS, and definition Edge while running software version syngo CT2012 B and syngo CT2013B. | Software design | A recall notification letter, dated November 15, 2013, was sent to Users. |  | 147 | 147 | Nationwide Distribution. | Terminated | March 10, 2015 | 480 | 147 | Computer
Z-0925-2014 | 67346 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Needleless Connector, Model No. 2000E |  |  | January 22, 2014 | 2014 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Needlefree Connector, model number 2000E, because of connection issues. The affected SmartSite connector lots may unintentionally disconnect from a female luer, may be difficult to disconnect from a female luer, or may fail to disconnect from a female luer once attached. Leakage may also be observed if the connector disconnects from the female luer during in | Under Investigation by firm | The firm, CareFusion, sent an "URGENT: Medical Device Recall Notification" letter dated January 22, 2014 to customers. The letter described the product, problem and actions to be taken. Customers were instructed to discontinue use of the listed lot numbers, contact CareFusion for a replacement of the affected lots, and if you have inventory of the listed products immediately complete and return to CareFusion the enclosed, pre-addressed and postage paid, Recall Response Card. All recalled product should be returned directly to the distributor from whom it was purchased. If you have any question, contact your distributor directly or call the CareFusion Support Center at numbers below.For recall related questions, customers are instructed to contact CareFusion Support center at (888) 562-6018. For adverse event reports, customers are instructed to contact Customer advocacy at (888) 812-3266 or email customerfeedback@carefusion.com. For technical questions regarding the Alaris System, customers are instructed to contact technical support at (888) 812-3229. |  | 3500000 | 3,500,000 units (2,136,200 units in US; 1,404,700 units foreign) | Worldwide Distribution: US (nationwide) including Puerto Rico; and internationally to: Australia, Canada and New Zealand. | Terminated | August 06, 2015 | 561 | 3500000 | Not_Computer
Z-0918-2014 | 67362 | Yes | NPT |  |  | Cardiovascular | PMA | Crimper Model 9100CR26 |  |  | January 20, 2014 | 2014 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling two lot numbers of the 26mm transcatheter heart valve crimper because they may not open sufficiently when the handle isrotated in the full open position. As a result, user may have difficulty fitting a 26mm Sapien transcatheter heart valve into the crimper orifice when preparing the valve for crimping. | Nonconforming Material/Component | Edwards Lifesciences issued an Urgent Product Notice to all customers. Letters were sent via Fed Ex on 01/20/14. The letter outlines steps to be taken to enable a valve to fit into the open orifice, and allow the valve to be crimped according to normal instructions. Customers were instructed to forward the Urgent Product Notification letter to necessary personnel within their organization. Customers were also instructed to acknowledge they have reviewed and understand the notification letter by signing and dating the form included with the letter. Customers with questions were instructed to contact their Edwards Clinical Specialist. |  | 422 | 422 | Nationwide Distribution | Terminated | March 24, 2014 | 63 | 422 | Not_Computer
Z-0908-2014 | 67242 | No | LWR |  |  | Cardiovascular | PMA | Product Name: Model 7305- Mosaic Aortic Obturator |  |  | January 09, 2014 | 2014 | Medtronic Cardiovascular Surgery-the Heart Valve Division | Medtronic, Inc. initiated a recall of Medtronic Mosaic Aortic Obturator Model 7305 used with the Mosaic Porcine Aortic Bioprosthesis, Model Number 305,because of higher-than-expected transvalvular gradients occurring post implant. | Device Design | Medtronic sent an Customer Notification letter dated January 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Medtronic recommends surgeons continue to remain diligent in selecting the appropriate size of Mosaic Aortic Bioprosthesis. Adoption of the updated Mosaic Aortic Bioprosthesis sizing chart (iEOA chart) and Mosaic Obturators/Sizers will facilitate optimal aortic valve size selection.Your Medtronic Field Representative will be contacting you about receiving a new Mosaic Aortic Obturator/Sizer set and sizing chart. If you would prefer to order the new Mosaic Aortic Obturator/Sizer set directly through Medtronic Customer Service, contact them at 1-800-848-9300.Medtronic will notify all applicable regulatory agencies, as required, about this matter. Please share this notification with others in your organization as appropriate. We appreciate your review of this notification and apologize for the inconvenience that it may cause. If you have any questions, please contact your Medtronic Field Representative or Lifeline Technical Services at 1-877-526-7890. |  | 8276 | 8276 units | Worldwide Distribution - USA (nationwide) | Terminated | May 30, 2014 | 141 | 8276 | Not_Computer
Z-0906-2014 | 67215 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | ***REF 186-01-08***InteGrip ACETABULAR AUGMENT***SMALL 8mm. Use with 48/50mm Shell.***EXACTECH, Gainesville, FL 32653-1630. |  |  | September 11, 2013 | 2013 | Exactech, Inc. | Exactech is recalling the InteGrip Acetabular Augments due to an out of range condition for an in-vitro biological evaluation standard. | Employee error | Exactech sent an "Important Product Recall Notice" dated September 11, 2013, to all affected customers. The letter identified the product and the action needed to be taken by the customers. Customers were instructed to cease distribution or use of the products. Extend this information to your accounts that may have this product in their possession. Verify if you have any of the subject InteGrip Acetabular Augments (catalog numbers 186-01-08 to 186-05-13) in the specified lots. Complete and fax back the attached form. Further questions, please call 1-800-392-2832. |  | 235 | 235 | USA Distribution including the states of : FL, VA, NY, OH, ME, TX, CO, and GA. | Terminated | August 04, 2014 | 327 | 235 | Not_Computer
Z-0911-2014 | 67358 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision 500D, Precision MPI &amp; Precision RXI X-ray imaging systems. |  |  | December 12, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue involving the Precision 500D, Precision MPI &amp; Precision RXI X-ray imaging systems overhead LCD Video Monitor Suspension (Single monitor &amp; Dual monitor).There is a potential for the locking nut to loosen out of the LCD Video Monitor Suspension yoke. This will not lead to an immediate fall of the suspended monitor since it will | Device Design | The firm, GE Healthcare, sent an "Urgent Medical Device Correction" letter dated December 11, 2013 to its consignees. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology/ Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The consignees were instructed to immediately contact your local GE Healthcare Service Representative if you observe any inappropriate performance or excessive movement of the monitor suspension yoke.GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction.If you have any questions or concerns regarding this notification, please call one of the following numbers: United States: 800-437-1171, Japan: 0120-055-919, for other countries contact your local GE Healthcare Service Representative. |  | 816 | 816 (722 US, 94 OUS). | Worldwide distribution: US (nationwide) including Guam and DC but excluding CT, DE, HI and RI; and internationally to: AUSTRALIA, BULGARIA, CANADA, CHILE, DOMINICAN REPUBLIC, GERMANY, HONDURAS, IRELAND, ITALY, KUWAIT, LEBANON, MALTA, NAMIBIA, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SLOVAKIA, SPAIN, SWITZERLAND, TAIWAN, TURKEY, UNITED KINGDOM, VENEZUELA, and YEMEN. | Terminated | September 11, 2014 | 273 | 816 | Not_Computer
Z-0922-2014 | 67378 | No | DZC | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | SuperElastic Regular force NiTi lower or upper Natural arches. |  |  | December 13, 2013 | 2013 | Highland Metals, Inc. | Labeling issue: Product labeled as "Upper" may contain "Lower" archwires with a red midline mark. Product labeled ad "Lower" may contain "Upper" archwires with a black midline mark. | Employee error | Highland Metals sent an Urgent Field Safety Notice with a response form to all affected consignees on December 13, 2013, via email.. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to pass the Field Safety Notice to all those who needed to be aware within their organization or to any organization where the potentially affected devices have been transferred. Consignees were instructed to complete the Acknowledgement of Receipt and Action form and fax to 1-408-271-2962 or email to quality@highlandmetals.com. For questions regarding this recall call 408-271-2955. |  | 9 | 9 packages of each lot number. | Worldwide Distribution - USA including IL, KY, MI, NY, CA, FL, LA and Internationally to Switzerland., Denmark, Greece, Australia, England, and Canada. | Terminated | February 04, 2014 | 53 | 46 | Not_Computer
Z-0914-2014 | 67225 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Stand-Alone Resuscitation Unit, Bag and Mask, part number M1226444, labeled in part ***GE Healthcare 8880 Gorman Road Laurel, MD 20723***www.gehealthcare.com***. |  |  | December 12, 2013 | 2013 | GE Healthcare | Medical device adjustment knob may be loose and affect the accuracy of blended Air/Oxygen gas mixture to patients. | Process control | GE Healthcare notified end users with a letter flagged as Urgent Medical Device Correction letter on 12/12/2013. The notification advised that GE Healthcare has become aware of a potential safety issue associated with the loosening of the blender knob shaft on all Resuscitation units. The firm advised that accounts ensure that all potential users in their facility are made aware of this safety notification and the recommended actions described in the notification:1. Perform blender accuracy check as outlined in Section 2.5-2.5.1of the Service Manual to verify that the O2 concentration set at21% is within 16-26% O2 and the concentration at 100% settingis within the recommended range (95-105% O2).2. If the blender accuracy test results do not meet specifications,remove the unit from patient use and contact a GE HealthcareService Representative at the number listed below.3. If the blender accuracy check passes in Step 1 above, you maycontinue to use your resuscitation system. The following precautionsare recommended to ensure blender accuracy is maintained betweensystem checks.Âż Use an independent oxygen analyzer, as determined by your hospitalpolicy, to check blended Air/O2 accuracy.Âż When delivering Oxygen, use a pulse oximeter to monitor your patient.Customers were requested to contact a GE Healthcare Service Representative at 1-800-345-2700 to arrange for correction if the device is found to be defective. |  | 15 | 15 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Uruguay, Venezuela, Albania, Algeria, Australia, Austria, Bahrain, Belarus, Belgium, Bolivia, Canada, Chile, China, Costa Rica, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Kazakhstan, Korea, Kuwait, Lithuania, Macedonia, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nigeria, Norway, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Thailand, Tunisia, Turkey, United Arab Emirates, and United Kingdom. | Terminated | June 01, 2015 | 536 | 12201 | Not_Computer
Z-0921-2014 | 67172 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | B. Braun AccuFlo Elastomeric Infusion Pump 200ML/H 100 ML |  |  | December 13, 2013 | 2013 | B. Braun Medical, Inc. | Internal testing of two lots of the AccuFlo CT-2000-100 Elastomeric Infusion Pump identified flow rates which are outside the specified infusion time and may result in overly rapid infusion of medications. | Process design | B. Braun Medical sent an Urgent Medical Device Recall letter dated December 13, 2013 to affected customer. The letter identified the product the problem and the action needed to be taken by the customer. The letter was sent to Progressive Medical Inc., the Importer and B. Braun's sole customer.WHAT ACTIONS SHOULD I TAKE?1. Further use and/or distribution of the affected product should be discontinued immediately. Please identify and quarantine any impacted materials in your possession.2. Review this recall notification in its entirety and ensure that all users within your organization and any consignees your organization may have are informed of this recall. If you are a distributor, please forward this recall notice to any customers who are in receipt of the impacted product.3. Determine your current inventory of the affected lots within your facility. Do not destroy any affected product.4. Utilizing the attached "Recall Acknowledgement" form record the number of full cases and individual units (for any loose units or partial cases) in your possession. If you have no inventory, please indicate that you have zero (0) inventory on the "Recall Acknowledgement" form.5. Return the completed "Recall Acknowledgement" form to B. Braun Medical Inc. Quality Assurance who will administer the recall on behalf of B. Braun Melsungen AG. The form must be e-mailedto PA_QualityAssurance.BBMUS.Service@bbraun.com or faxed to (610) 849-1197 within two (2) weeks of receipt even it the inventory in your possession is zero.6. Instructions for Credit, Replacement, Et Return:a.If you have purchased the product from a third-party and have any full cases, partial cases, or unused individual pieces in your possession please contact the seller to obtain instructions for handling these units.b.If you have purchased the product directly from B. Braun Medical Inc. and have any full cases, partial cases, or unused individual pieces of these affected products, ple |  | 30240 | 30, 240 units | US Distribution one importer/wholesale distributor in MO. | Terminated | September 03, 2014 | 264 | 30240 | Not_Computer
Z-0919-2014 | 67275 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX RIS QDOC 5.8 |  |  | January 08, 2014 | 2014 | AGFA Corp. | Patient name displayed (printed) on the Patient Report was the wrong patient name. | Unknown/Undetermined by firm | On January 8, 2014, an "URGENT FIELD SAFETY NOTICE" (UFSN) letter was emailed to consignees. The letter described the safety alert and mitigation, including directions to consignees as to how the disable Microsoft Word Background Printing. Acknowledgment via FAX-Back or email, that the information was received and understood was requested from the consignees. |  | 14 | 14 | CA, FL, GA, ME, OH, OR, SC, SD, TX and Canada | Terminated | February 28, 2014 | 51 | 14 | Not_Computer
Z-0924-2014 | 67133 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA Urine Control L/H. |  |  | December 18, 2013 | 2013 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is informing all ABX PENTRA Urine Control L/H customers running the Creatinine assayon ABX PENTRA Chemistry Analyzers that the Target Range for Creatinine has been updated. The ABX PENTRA Urine Control L/H lot specific annexes have been updated toinclude revised acceptable Target Ranges for Creatinine.No other analyte valuesare affected. | Pending | HORIBA Medical sent an Urgent Safety Notice dated December 18, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to update their Quality Control procedures to include the new Creatinine TargetRanges for the ABX PENTRA Urine Control L/H Reference #: A11A01674 and Lot #s: 1207001, 1303601, and 1300701. Customers were instructed to complete and return the enclosed response form to verify their understanding of the Field Safety Notice. Customers with questions were instructed to contact their local HORIBA Medical representative.For questions regarding this recall call 949-453-0500. |  | 531 | 531 | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | October 09, 2014 | 295 | 531 | Not_Computer
Z-0920-2014 | 66953 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Bicarbonate Jug Adapter to the 2008 Series Hemodialysis machine |  |  | December 17, 2013 | 2013 | Fresenius Medical Care Holdings, Inc. | Bicarbonate Jug Adapter distributed prior to510(k) approval | No Marketing Application | Fresenius notified customers by telephone and Formal notification on 12/17/13 via the Urgent Bicarbonate Jug Adapter Recall letter sent certified mail with signature confirmation and fax-back form.Customers were instructed to examine their stock immediately to determine whether they have any of the affected parts on hand. If customers have the affected parts, they are instructed to discontinue use and place all Bicarbonate Jug Adapters in a secure area for return to Fresenius Medical Care Renal Therapies Group, LLC (FMC-RTG). Customers instructed to contact their Fresenius Medical Care Technical Services Team for instructions on how to return the recalled product. |  | 703 | 703 units | Nationwide Distribution. | Open, Classified |  |  | 703 | Not_Computer
Z-0926-2014 | 67248 | No | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | The Rotoprone therapy System is an advanced patient care system for the treatment and prevention of pulmonary complications. The Rotoprone Therapy System provides up to 62 degrees in both prone and supine positions as well as Trendelenburg/Reverse Trendelenburg positioning. |  |  | September 11, 2013 | 2013 | ArjoHuntleigh Inc. | The lock pin can become stuck in the "in" position during the manual rotation, resulting in the inability to prone a patient | Device Design | ArjoHuntleigh Inc. sent a Technical Bulletin package to all affected service centers and an Urgent Field Safety Notice package to all affected customers. For rental units, a training plan was established for employees. For sold units, the notice contain a revised user manual and label explaining what to do if the lock pin becomes stuck in the "in" position. Customers that purchased the affected product were required to complete and return a Customer Response form to ArjoHuntleigh to confirm receipt and understanding of the field safety notice, and that the required actions were completed. For questions call 630-785-4885. |  | 231 | 231 units | US Nationwide and Middle East | Terminated | January 08, 2015 | 484 | 231 | Not_Computer
Z-0932-2014 | 67167 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Diagnostic imaging systems (GEMINI LXL Model #882412, TruFlight Select Model #882438, GEMINI TF PET/CT16 Model #882470, GEMINI TF PET/CT 64 Model #882471, GEMINI TF Ready Model #882473 &amp; GEMINI TF Big Bore Model #882476) |  |  | December 04, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The PET Reconstruction Server (PRS) database may randomly lock up after the low dose CT acquisition is complete and will not allow the PET acquisition to begin, resulting in an unexpected incomplete study. | Software design | On 12/4/2013, the firm sent Field Safety Noticies to their customers. |  | 386 | 386 units | Product was shipped to the following states: AL, AR, AZ, CA, CO, DC, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NJ, NV, NY, NY , OH, OR, PA, TN, TX, VA, VT, WA &amp; WI.Product was also shipped to the following countries: Australia, Belgium, Brazil, Bulgaria, Canada, China, Croatia, Cuba, Denmark, Egypt, France, Germany, Hong Kong, Hungary, India, Indonesia, Italy, Japan, Kuwait, Lebanon, Malta, Mexico, Netherlands, Pakistan, Philippines, Poland, Portugal, Republic of Korea, Russia, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, Ukraine, United Kingdom, Venezuela &amp; Vietnam. | Terminated | February 02, 2016 | 790 | 386 | Not_Computer
Z-0928-2014 | 67224 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO RTP System. |  |  | November 25, 2013 | 2013 | Elekta, Inc. | When bolus is present, the Effective Depth (with bolus) to the weight point returned on the Source Data Report is incorrect. | Software design | Corrective Action 1 - An important safety notice, LINXIO0018 - Incorrect Effective Depth for Bolused Beams is in the process of being distributed to affected customers. (Release date November 25, 2013). Corrective Action 2 - Develop a fix to the defect in an expeditious timeframe and make available in Patch Release 4.0.03 and Release 5.00.01. |  | 1424 | 1424 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AL, AK, AZ, CA, CO, CT, DE, FL, GA, IL, IN, KS, KY, LA, ME, MD, MI, MN, MS, MO, MT, NE, NV, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, QA, WA, WV, WI, and WY, and the countries of Albania, Algeria, Argentina, Austria, Australia, Bosnia, Bulgaria, Belgium, Bahrain, Brazil, Belarus, Botswana, Bangladesh, Canada, China, Columbia, Costa Rica, Croatia, Cuba, Cyprus, Chile, Czech Republic, Ecuador, Egypt, Estonia, France, Finland, Gabon, Germany, Great Britain, Greece, Guatemala, Hungary, India, Indonesia, Italy, Ireland, Israel, Iraq, Jordan, Japan, Kosovo, Kuwait, Lithuania, Libya, Latvia, Morocco, Malaysia, Malta, Mexico, Myanmar, Netherland, New Zealand, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Slovakia, South Africa, South Korea, Spain, Sudan, Syria, Taiwan, Thailand, Turkey, Turkmenistan, Ukraine, Venezuela, and Vietnam. | Terminated | December 02, 2015 | 737 | 1424 | Not_Computer
Z-0927-2014 | 66946 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CARESCAPE Monitor B850, a multi-parameter high acuity patient monitor intended for use in multiple areas within a professional healthcare facility. |  |  | December 20, 2013 | 2013 | GE Healthcare, LLC | There is a potential for communication loss associated with the F5-01 Frame when connected to CARESCAPE Patient Data Module (PDM) in the CARESCAPE Monitor B850. The PDM patient parameter acquisition module can sometimes spontaneously lose the communication with the host patient monitor when attached to the F5-01 Frame due to an issue with the PDM Ethernet communication hardware. The communicatio | Component change control | GE Healthcare sent an "Urgent Medical Device Correction" letter dated December 20, 2013 to all affected customers. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter addressed the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Call Technical Support at 1-800-558-7044 or your local Service Representative). |  | 12 | 12, 945 units (6,239 US, 6,706 OUS) | Worldwide Distribution - US Nationwide including Puerto Rico and the countries of: Australia, Austria, Azerbaijan, Bahrain, Belgium, BRAZIL, Canada, CHILE, COSTA RICA, Cyprus, Czech Republic, Denmark, ECUADOR, Egypt, Estonia, Finland, France, Germany, Greece, GUAM, India, INDONESIA, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Korea, Republic of, Kuwait, Lithuania, Malaysia, Malta, Mexico, NEPAL, Netherlands, New Zealand, Norway, Pakistan, PANAMA, PARAGUAY, PERU, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tanzania, United Republic of, Turkey, United Arab Emirates, United Kingdom, URUGUAY. | Open, Classified |  |  | 12 | Not_Computer
Z-0930-2014 | 67021 | No | DIP | Diphenylhydantoin test system. | Toxicology | Toxicology | 510(k) | ADVIA CentaurÂż Systems Phenytoin Calibrator N ( PHTN) 6-pack |  |  | December 18, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | ADVIA CentaurÂż Systems Phenytoin Calibrator N, lots CN58 and higher, change in recovery observed with the ADVIA Centaur Systems Phenytoin assay compared to recoveries using previous Calibrator N lots. | Under Investigation by firm | Siemens Healthcare Diagnostics Customer Notification notice was sent to affected Siemens Healthcare Diagnostic customers in the United States on December 18, 2013.. The notifications inform customers that Siemens Healthcare Diagnostics received customer feedback regarding the performance of the ADVIA CentaraÂż Systems Phenytoin assay, indicating that the assay was recovering higher relative to the performance of other assays in proficiency studies. Siemens internal investigation confirmed the positive bias for Calibrator N lot CN 57 in the ADVIA Centaur Systems Phenytoin assay. This lot has expired and is no longer in use. Expected performance has been restored starting with Calibrator N lot CN58. Siemens is notifying customers that have received Calibrator N lots CN58 and higher that there will be a change in recovery observed with the ADVIA Centaur Systems Phenytoin assay compared to recoveries using previous Calibrator N lots. Siemens is providing data to inform customers of the expected difference between Calibrator lots CN58 and higher versus prior Calibrator N lots.Customers were instructed to retain the letter with their laboratory records, and forward to those who may have received the affected product.Customers with questions were instructed to contact their Siemens Customers Care Center or their local Siemens technical support representative.For questions regarding this recall call 508-668-5000. |  | 1294 | 1294 (144 units US) | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia,Brazil,Canada, Columbia, China, Ecuador, Egypt, India, Japan, Korea, Mexico, Saudi Arabia,Singapore, Thailand, and Uruguay. | Open, Classified |  |  | 1294 | Not_Computer
Z-0931-2014 | 67326 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Trifurcated Extension Set, model no. ME 1224. |  |  | January 16, 2014 | 2014 | CareFusion 303, Inc. | CareFusion is recalling the Trifurcated Extension Set, models ME 1224 and MP 9230, because they may have a defect in the trifurcated component that could result in a component leak. | Nonconforming Material/Component | A customer notification letter and a distribution letter are dated 1/22/14 were sent to customers to inform them that CareFusion is recalling the Trifurcated Extension Set, models ME 1224 and MP 9230 may have a defect in the trifurcated component that could result in the component to leak. The letters inform the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return to CareFusion the enclose, pre-addressed and postage paid, recall response card. Customers are instructed that once CareFusion receives the recall response card, the CareFusion Support Center will provide instructions for return of the affected lot number and replacement of the affected lot. Customers are instructed to contact CareFusion Support Center at (888) 562-6018, Customer advocacy at (888) 812-3266, customerfeedback@carefusion.com, or technical support at (888) 812-3229. |  | 9550 | 9,550 units | Nationwide in US and Canada | Terminated | August 06, 2014 | 202 | 9550 | Not_Computer
Z-0933-2014 | 66861 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | The Gaymar Medi-Therm device provides a means of regulating patient temperature by supplying temperature-controlled water through a connector hose to accessory Gaymar Hyper/Hypothermia blanket(s)/body wrap(s). The blanket/body wrap provides an interface for heating or cooling the patient. Accessory YSI 400 series patient probe interfaces between the |  |  | December 27, 2013 | 2013 | Stryker Medical Division of Stryker Corporation | Stryker Medical has identified that on some units of the MTA7900, there is a lack of a water temperature display while the unit is in Auto mode. While in Auto mode, only the patient temperature is displayed along with the set point. The water display will be blank during Auto mode. If the user wants to see the water temperature displayed and switches out of Auto mode to Manual mode, then back to | Device Design | On Dec. 27, 2013, Stryker Medical notified their consignees, identifying the product, reason for the recall and associated risks. Consignees were asked . Locate the units listed attach enclosed warning tags to the units listed. For units that do not have the water temperature displayed, advise staff not to switch to manual mode in order to display the water temperature.Stryker will make arrangements to update the software so that the water temperature will be displayed in both Auto and Manual modes. Fax (269-488-8691) or e-mail (ProductFieldAction@stryker.com ) the enclosedacknowledgment to confirm receipt of this notification. Please forward a copy of this letter to the new users and advise us of their new location and let Stryker know if any units have been discontinued. Questions or concerns, should be addressed to Stryker Customer Service at (1-800-327-0770). Monday-Friday 8 a.m-5 p.m. (ET). |  | 1002 | 1002 units | nationwide and worldwide | Terminated | November 04, 2014 | 312 | 1002 | Computer
Z-0936-2014 | 67249 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | The Symbionic Leg is an integrated prosthetic leg for transfemoral and knee disarticulation amputees. |  |  | February 12, 2012 | 2012 | Ossur Americas, Inc. | OSSUR initiated this recall of the Symbionic Leg because it may have a possible defect in a small batch of electronics. | Nonconforming Material/Component | OSSUR contacted all affected customers via phone on February 12, 2013, informing them that a voluntary recall for this product has been implemented, along with the recommended instructions to return the product. A new replacement prosthetic will be provided to the customer.OSSUR has inspected and made appropriate modification to all inventory. Âżssur will also further strengthen the design and our manufacturing process to ensure an even more robust product.For further questions please call (949) 303-3239. |  | 17 | 17 units | US Distribution including the states of AL, DC, GA, NY, NJ, OK and PA. | Terminated | February 06, 2014 | 725 | 17 | Not_Computer
Z-0935-2014 | 67315 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | LOCATORÂż Implant Attachment System. |  |  | January 16, 2014 | 2014 | Zest Anchors LLC | Zest Anchors, LLC is voluntarily recalling specific lots of LOCATORÂż Implant Attachment Systems because it is currently provided with a label that indicates that the final torque applied to the abutment should be 30 N-cm. It was recently noted by Zest that in order to provide an optimal performance by increasing the factor of safety between the failure load of the abutment screw and the recommend | Unknown/Undetermined by firm | Zest Anchors, LLC initiated this recall by sending customers recall notification letter on 01/16/2014 via FedEx overnight.The customer notification letter, titled, "IMPORTANT: MEDICAL DEVICE VOLUNTARY RECALL NOTICE", informed customers of the recall, reason for voluntary recall, product information with part numbers and lot numbers, actions to be taken, and contact information. The customer notification letter was accompanied by a response form titled "MEDICAL DEVICE RECALL RETURN RESPONSE Acknowledgement &amp; Receipt Form". |  | 3130 | 3,130 | Distributed in Ohio and Switzerland. | Terminated | May 15, 2014 | 119 | 3130 | Not_Computer
Z-0937-2014 | 67237 | No | GFO | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | Hemochron Jr. Citrate APTT Cuvette |  |  | December 19, 2013 | 2013 | ITC-Nexus Dx | ITC has determined that some Citrate APTT cuvettes may recover higher than expected results in normal individuals. | Device Design | The firm, ITC The Point of Care, sent an "URGENT MEDICAL DEVICE RECALL" letter dated December 19, 2013 a long with an attached Health Professional Account Tracking Form via UPS to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to forward this communication to all those within your organization who need to be aware of this matter including any facility that you forwarded any of the affected product to; check all inventory to determine if you have any of the affected lots listed; if you have any affected lots, stop using them, remove from inventory and return to ITC; complete the attached Health Professional Account Tracking Form and return it within 10 days via fax to: 866-429-3132, email: RA@itcmed.com or mail to: ITC Technical Support, 20 Corporate Place South - Piscataway, NJ 08854 . It is important that you complete the form whether or not you have remaining inventory of these lots. In addition, contact ITC Technical Support immediately if you become aware of an unreported adverse event related to the use of affected J103C cuvettes.ITC will review your response form and will contact you to arrange for anyinventory to be returned to ITC and ITC will replace all unused, returned product.If you have questions, please contact ITC Technical Support: Telephone 800-579-2255 or 732-548-5700, Extension 4011 or E-mail Techsupport@itcmed.com. |  | 25965 | 25,965 curvettes (7,425 cuvettes US; 18,540 cuvettes OUS) | Worldwide Distribution: US (nationwide) and countries of: Austria, Belgium, Canada, Serbia, Germany, Estonia, Italy, Spain and Saudia Arabia. | Terminated | October 30, 2014 | 315 | 25965 | Not_Computer
Z-0934-2014 | 67302 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.5mm Locking Cortical Bone Screw, T15 Hexalobe, Self Tapping, 20mm |  |  | January 08, 2014 | 2014 | OrthoPediatrics Corp | OrthoPediatrics has initiated a Voluntary RECALL of Part # 00-0903-2620, 3.5MM Locking CorticalScrew, Lot # 009SV because the device within the package may be incorrect. These packages incorrectly contain a 3.5mm Non Locking Cortical Screw # 00-1050-3524. | Labeling mix-ups | Orthopediatrics Corporation has requested sales representatives and distributors who have been identified via shipping records as having the possibility of having received one or more discrepant devices affected by this Recall were requested to quarantine the devices identified in this Recall by telephone on January 8, 2014. The decision to Recall these devices was made and a formal written Recall Notice sent to the consignees via email on January 13, 2014.Notification will not be sent to hospitals at this time provided that all devices areretrieved or accounted for within 60 days of the initiation of recall. Hospitals will be notified on an as needed basis if OrthoPediatrics determines that all recalled devices have not been recovered or accounted for, and that contacting the hospitals involved will expedite the recall efforts. All discrepant devices subject to this voluntary Recall are planned to be returned from distribution to OrthoPediatrics Corp. in Warsaw, Indiana and maintained in quarantine. All remaining unshipped devices from these lots have been placed in quarantine at OrthoPediatrics Corp.All quarantined devices at OrthoPediatrics will be inspected and repackaged to correct the mislabeling per normal documented procedures. Objective evidence will be maintained of the rework of the devices.For questions regarding this recall call 574-268-6379. |  | 12 | 12 units | Worldwide Distribution - USA including OK, KY, FL, LA, OH, CO, MO, MN and Internationally to Ireland, Saudi Arabia, Italy, Turkey, South Africa, and Germany. | Terminated | April 23, 2014 | 105 | 12 | Not_Computer
Z-0938-2014 | 67350 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Malibu Polyaxial Screwdriver, 8.0 Plus, Model Number 91-2109; |  |  | January 22, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences has identified through an investigation of complaints that there may be the potential for the shaftof the Malibu Polyaxial Drivers to break during use. None of the complaints reported a patientinjury. However, some complaints did indicate the surgery was slightly prolonged while the additional driver providedin the kit was retrieved and used. | Manufacturing material removal | Integra LifeSciences initiated this recall by sending recall notification letter to US consignees, delivered by traceable courier service. Non US consignees notified of the recall by either traceable e-mail or facsimile. The letter dated, January 23, 2014, titled "URGENT: VOLUNTARY MEDICAL DEVICE RECALL IntegraÂż Malibu Polyaxial DriversCatalogue Numbers: 91-2109 &amp; 91- 2110", informed customers of the recall and provided product description with codes, picture of the product, advised customers to: Inspect their inventory of Malibu Polyaxial Drivers.- If you have product(s) with a lot number listed on the following page, STOP USING THEM IMMEDIATLY.- If you do not have Drivers with a lot number listed on the following page, you can continue to use them.Customers were also instructed that regardless if they have Drivers with the affected lots, to complete the attached form and return it to Integra as indicated. Customers were instructed to keep a copy of the form for their records., and contact information. Customers with questions were instructed to contact Customers Service at 1-866-942-8698.For questions regarding this recall call 609-275-0500. |  | 40 | 40 | Worldwide Distribution - USA including IA, FL, LA, IL, CA, TX, AZ, NC, MD, NM, VA, and Internationally to South Africa. | Terminated | December 17, 2014 | 329 | 40 | Not_Computer
Z-0939-2014 | 67246 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | ***REF 606-01-05***OPTEFORM Allograft Paste 5 cc, bone void filler of bony defects in dental and orthopedic applications.***STERILE Using Radiation. ***Rx Only***Distributor: EXACTECH 2320 NW 66th Court, Gainesville, FL 32653 USA***Manufacturer: RTI Biologicals, Inc. 11621 Research Circle, Alachua, FL 32615 USA. |  |  | December 06, 2013 | 2013 | RTI Surgical, Inc. | RTI Surgical, Inc. of Alachua, FL is recalling their OPTEFORM Allograft Paste, 5cc as it may contain a 10 cc fluid dispenser. The 10cc fluid dispenser does not include the correct fluid fill line reference for reconstitution of 5cc past allografts. | Other | RTI Surgical issued a: URGENT: Medical Device Recall Notification dated December 6, 2013 to their customer. The customer was instructed to quarantine the product immediately and return to RTI Surgical, Inc. for final disposition. For questions or concerns, call 386-418-8888. |  | 14 | 14 units | Distributed to Florida. | Terminated | September 15, 2014 | 283 | 14 | Not_Computer
Z-0940-2014 | 67027 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens ADVIA Chemistry Systems models 1200, 1650, 1800 and 2400 |  |  | November 06, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | The ADVIA Chemistry software systems using certain software versions allows the system to automatically run assay calibrations and quality control (QC). | Software design | Siemens sent an Urgent Medical Device Correct letter, dated November 6, 2013, to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to retain the recall notification letter with their laboratory records and forward to those who may have received the affected product. |  | 3759 | 3759 | Worldwide Distribution and USA Nationwide. | Terminated | August 10, 2015 | 642 | 3759 | Computer
Z-0941-2014 | 66467 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Track System with enGen Select v5.0 with autoverification enabled |  |  | October 08, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics issued an Urgent Product Correction Notification involving the enGen Select V5.0 &amp; enGen Custom v3.2.2 due to an anomaly that affects the autoverification feature and custom rules that depend upon data elements used by the referenced enGen Systems. | Software design | Ortho Clinical Diagnostics (OCD), issued an Urgent Product Correction Notification letter dated October 7, 2013 and on December 17, 2013 to their customers via Federal Express or via US Postal Service. The letter identified the affected product, problem and actions. The notices advise the customers to ensure that all required data elements are entered and in the correct format if using the autoverification or custom rules for the referenced enGen Laboratory Automation Systems. OCD trained service representatives will contact all affected customers. Customers with questions may call the Customer Technical Service at 1-800-421-3311. |  | 8 | 8 units | Worldwide Distribution - USA Nationwide in the states and countries of: AZ, FL, WV, TX, PA, DE, MI, OH, TN, IL, NC, Chile, India, Italy, Portugal, Spain, Sweden, UK | Terminated | April 11, 2017 | 1281 | 28 | Not_Computer
Z-0949-2014 | 65214 | No | LKM | Automated cell counter. | Hematology | Pathology | 510(k) | The UF-1000i is a fully automated urine particle analyzer intended for vitro diagnostic use in urinalysis. The instrument can perform screening for abnormal samples with a high degree of accuracy, so it increases automation and efficiency in laboratories. |  |  | March 31, 2013 | 2013 | Sysmex America, Inc. | The current Century Break patch did not contain the same setting for the Century Break year for both the Red Hat Linux and GUI files. Due to the mismatch in file settings, the software could not properly calculate a purge date for the database files based on the customer data retention setting in the application. This resulted in the database files meeting its maximum capacity and causing system p | Software design | A Product Notification letter dated July 2012, was sent to Siemens Healthcare Diagnostics. The letter suggested that an attached procedure should be followed for an server that requires the Century Break parameter change. If Siemens Healthcare Diagnostics has any questions about the procedure, the firm's representatives can contact the Sysmex Technical Assistance Center at 1-888-879-7639. A response form was also included with the letter requesting the Siemens Healthcare Diagnostics complete and return it. A second Product Notification letter dated September 2012, was sent to Siemens Healthcare Diagnostics. The letter contained largely the same instructions as the July 2012, letter. A third Product Notification letter dated March 2013, was sent to Siemens Healthcare Diagnostics. The letter include substantially similar instructions as the pervious letters; however, the procedure for changing the Century Break parameter was expanded. |  | 214 | 214 units | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and the states of AL, AR, CA, CO, CT, FL, GA, HI, IA, IS, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, SC, SD, TN, TX, VA, WA and WI and the country of Canada. | Open, Classified |  |  | 214 | Computer
Z-0968-2014 | 66498 | micro mixer | JRB | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products Mixing Cup Arrays utilized on VITROS 250 &amp; 350 Chemistry Systems |  |  | October 10, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics issued a product correction notification for VITROS Chemistry Mixing Cup Arrays utilized on VITROS 250 and VITROS 350 Chemistry Systems. Increase in Software error due to affected Mixing Cup Arrays. | Nonconforming Material/Component | Ortho Clinical Diagnostics sent an Important Product Correction Notification dated October 10, 2013 to all affected customer. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to complete and return the attached Confirmation of Receipt form. For questions contact the Customer Technical Services at 1-800-421-3311. |  | 306240 | 6380 units (each unit contains 48 arrays with 10 microcups) | Worldwide Distribution - US Nationwide and the countries of: Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, PR, Republic of Panama, Singapore, Venezuela, England, France, Germany, Italy, Spain | Open, Classified |  |  | 306240 | Computer
Z-0827-2014 | 67264 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Product trade name t:slim" Insulin Delivery System |  |  | January 10, 2014 | 2014 | Tandem Diabetes Care Inc | Tandem Diabetes CareÂż, Inc. announced that it is initiating a voluntary recall of specific lots of insulin cartridges that are used with the t:slimÂż Insulin Pump. The affected cartridges may be at risk for leaking. The cause of the recall was identified during Tandems internal product testing, and has not been associated with any complaints or adverse events reported by customers. However, a c | Equipment maintenance | Tandem Diabetes Care sent an Urgent Medical Device Voluntary Recall letter dated January 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed check the lot numbers of their cartridge supply against the llist of affected lot numbers. Do Not use any cartridges from the lots affected by this recall. If customers don't have useable cartridges, discontinue using the pump and revert to their backup plan until new supplies arrive. Customers were instructed to call Tandem Technical Support at 1-877-801-6901 to receive replacement cartridges at no charge. Customers were asked to return the enclosed return response form via email, fax, or mail. UPDATED: Tandem Diabetes Care expanded the recall om January 20, 2014. Tandem Diabetes CareÂż, Inc. (NASDAQ: TNDM), or Tandem, today announced that it is expanding a voluntary recall of specific lots of insulin cartridges that are used with the t:slimÂż Insulin Pump.For questions regarding this recall calll 858-366-6900, ext 6963. |  | 12 | 12, 807 boxes (10packs) | USA (nationwide) | Terminated | September 23, 2014 | 256 | 12 | Not_Computer
Z-0945-2014 | 65667 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MICROTEK(TM) MEDICAL BUR HOLE LATEX FREE PROBE COVER W/ GEL, featuring IsoSilk(R), 13cm x 244 cm, 5" x 96", REF PC3688EU, STERILE EO, QTY 20, Microtek Medical, Inc., Columbus, MS 39702 |  |  | July 02, 2013 | 2013 | Ecolab Inc | Certain lots of these Bur Hole Probe Covers did not undergo endotoxin (LAL) testing to the proper levels. | Process design | The recall was conducted to the consumer/user level. The firm notified their direct consignees of the problem and the recall by written letter mailed on July 2, 2013. The notice requested the consignees to quarantine any recalled product on hand, and complete and return the response form by July 31, 2013. The consignees who further distributed the recalled product were asked to notify their consignees. |  | 12546 | 12,546 individual units | Worldwide Distribution-USA (nationwide) and the countries of Austria, Germany, Finland, Ireland, France, Spain, Great Britain, Belgium, Netherlands, Switzerland, Jordan, United Kingdom, Saudi Arabia, Hungary, Latvia, Malta, Norway, Poland, Portugal, and Canada. | Terminated | May 14, 2014 | 316 | 29258 | Not_Computer
Z-0964-2014 | 67234 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Navigation Compatible Accolade Broach Handle |  |  | November 20, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received reports from the field of the impaction plate dissociating/fracturing from the main body of the broach handle. | Process design | Stryker Orthopaedics sent notification letters and acknowledgement forms via Fed Ex as follows:Branches-11/20/2013Hospital Risk Manager - 11/21/2013 |  | 211 | 211 units | IL MA, MN, ND, OR, &amp; PA | Open, Classified |  |  | 211 | Not_Computer
Z-0967-2014 | 67017 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | BrainLAB's VectorVision spine is intended for use as an intraoperative image-guided localization system for minimally invasive surgery. It links a freehand probe, tracked by a passive marker sensor system to virtual computer image space on a patient's preoperative image data that is processed by a VectorVision workstation. |  |  | November 15, 2013 | 2013 | Brainlab AG | Brainlab has determined that pins of thee Instrument Star Unit ML (Calibration with ICM4) with specific serial numbers might have been made of an incorrect material that is not suitable for the necessary disinfection and sterilization of the product. If corrosion occurs and is not detected by the user, and the device is used during surgery, corroded particles of the pin could, directly or indirect | Nonconforming Material/Component | BrainLab send a A FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 15, 2013, to all affected customers. The letter requested that customers verify that they are in possession of one or more of the affected devices, remove the affected devices from use, and that Brainlab will provide replacement parts for the customers. Customers with questions can contact Brainlab by telephone at 1-800-597-5911 or via e-mail at us.support@brainlab.com.For questions regarding this recall call 800-597-5911. |  | 25 | 25 units | Worldwide Distribution - USA including CA, MN, NC, NY, OH, TX and DC. Internationally to Australia, Chile, Germany, Hungary, Italy, Panama, Qatar, Russia, Thailand, United Arab Emirates. | Open, Classified |  |  | 25 | Not_Computer
Z-0966-2014 | 66995 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | Normative Data Template CD for use with the I-Portal devices. |  |  | November 21, 2013 | 2013 | Neuro Kinetics, Inc. | The data provided on the Normative Data Template CD for use with the I-Portal devices has not received clearance by FDA through the premarket 510(k) notification process. | No Marketing Application | A recall notification letter, dated November 2013, was sent to Users. A follow up letter regarding scheduling removal of the software is planned for December 2013. |  | 148 | 148 | Distributed USA (nationwide) and the country of Canada. | Terminated | July 20, 2015 | 606 | 148 | Computer
Z-0965-2014 | 66947 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | D-10 Dual Program, which contains Update Kit D-10 A1c program diskette. |  |  | November 22, 2013 | 2013 | Bio-Rad Laboratories, Inc. | D-10 A1c Dual Program Floppy was programmed incorrectly and cartridge injections counts are not being updated correctly when switching between the D-10 Dual Short Program and Extended Program on the D-10 Hemoglobin Testing System. | Software design | BIO-RAD sent an Urgent Medical Device Correction letter dated November 22, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to continue to use the affected update kit disk(s) included with the reorder pack(s), but manually track the number of injections on the cartridge, reviewing QC results closely throughout the life of the reorder pack or destroy the affected disk(s) included with these reorder pack(s) and replace with the new Update Kit disk(s) The new Update Kit disks will elminate the problem and may be used with all reorder packs. The replacement disks do not contain reagent or calibrator information; after installing teh new Update Kit disk, they ust manually update the component information using instructions provided with the disk.For questions regarding this recall call 510-741-6265. |  | 2906 | 2906 kits (318 Kits in US; 2588 kits internationally) | Worldwide Distribution - USA (nationwide) and Internationally to Latin America, Asia Pacific, Canada, Australia, New Zealand and India. | Terminated | March 20, 2014 | 118 | 2906 | Not_Computer
Z-0971-2014 | 67437 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, Coyote" MONORAIL" PTA Balloon Dilatation Catheter, UPN # H74939185201510, Catalogue # 39185-20151. |  |  | February 04, 2014 | 2014 | Boston Scientific Corporation | Boston Scientific is initiating this Recall Removal of one lot of the Coyote" OVER-THE-WIRE PTA Balloon Dilatation Catheter and one lot of the Coyote" MONORAIL" PTA Balloon Dilatation Catheter. Boston Scientific has become aware that the balloons of the devices from these lots may not deflate as expected during a procedure. To date, Boston Scientific has received three complaints for this issue. T | Device Design | Consignees were sent a Boston Scientific " Urgent Medical Device Recall Removal" letter on 2/4/2014. The letter was addressed to Materials Manager and Field Action Contact. The letter described the product involved in the recall and the problem. Advised consignees to segregate the affected lots and return to Boston Scientific in accordance to the enclosed recall instructions. The instructions also requested consignees to complete and return the Account Reply Verification Tracking Form (RVTF). |  | 35 | 35 for both products | US: AL, CA, FL, IL, IN, MI, MO, NY, OH, PA, TX.OUS: ITALY, GERMANY, FRANCE, GREAT BRITAIN, NETHERLANDS, WHITE RUSSIA, INDIA, JAPAN. | Terminated | August 04, 2014 | 181 | 35 | Not_Computer
Z-0979-2014 | 67305 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF Big Bore PET/CT, Computer Tomography System, Philips Healthcare System, Cleveland, OH |  |  | December 05, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | The TF Big Bore patient table experienced mechanical binding during horizontal table motion resulting in an automatic Emergency Stop (E-stop). An E-Stop will interrupt a current scan, resulting in an incomplete study. | Software design | ON 12/5/2013 the firm sent Medical Device Correction Letters to their customers. |  | 47 | 47 Units | Product was shipped to the following states: AR, CA, CO, FL, IL, LA, MA, MD, MI, MN, MO, NC, NE, NJ, NY, OH, PA, TX, VA, VT &amp; WI.Product was also shipped to the following countries: Belgium, Canada, China, Denmark, France, Germany, Italy, Japan, Netherlands, Saudi Arabia &amp; Venezuela. | Terminated | August 25, 2015 | 628 | 47 | Not_Computer
Z-0976-2014 | 67374 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging, Picture Archive Communication System (PACS). |  |  | December 23, 2013 | 2013 | Mckesson Information Solutions LLC | Images may be archived to an incorrect or invalid archive location resulting in incorrect images or no images being displayed for a study following retrieval from archive. | Software design | McKesson sent a field notice to customers with the potentially affected software versions on 12/23/2013. |  | 618 | 618 units | U.S. Nationwide (Puerto Rico), Australia, Canada, France, GU, IE, and The United Kingdom. | Terminated | April 25, 2017 | 1219 | 618 | Computer
Z-0977-2014 | 67267 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Medtronic GRAFTON (R) DBM, Flex 2.5cm x 10cm, REF T42100, Rx only, STERILE A, QTY: 1 EA, Medtronic Sofamor Danek, USA, Inc., 1800 Pyramid Place, Memphis, TN 38132 |  |  | January 06, 2014 | 2014 | Medtronic Sofamor Danek USA Inc | Product was labeled as being measured as 2.5 cm x 10 cm, but contained product that measured as 2.5 cm x 5 cm. | Packaging process control | The firm, Medtronic, sent an "URGENT-V voluntary Market Removal (Recall)" communication letter and questionnaire packet (via FedEx) dated January 3, 2014 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to remove the product from normal storage locations; Do not use the product; and as applicable, contact your Medtronic Sales Representative to coordinate the return and replenishment of affected product or call our Global Quality Department at 800-876-3133 ext 3197. In addition, the customers were instructed to disseminate this information to additional personnel within your facility as appropriate. A Medtronic Sales Representative will contact you to facilitate the return of any impacted Set Screw Retaining Drivers you may have in your possession.If you need assistance or if you have any related questions or concerns, please do not hesitate to contact your Medtronic Sales Representative or our Global Quality Department at 800-876-3133 extension 3197. |  | 30 | 30 units | US Distribution in the states of: MO, FL, MN, OH, MI, TX, and NJ. | Terminated | June 24, 2014 | 169 | 30 | Not_Computer
Z-0974-2014 | 65753 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 5 x 93 US Probe Cover CMS-6049, Sterile, Single Patient Use. Custom Medical Specialties, Inc., Pine Level, NC, Federal (USA) law restricts this device to the sale by or on the order of a physician. |  |  | July 12, 2013 | 2013 | Custom Medical Specialties, Inc. | No endotoxin (LAL) testing to proper level. | Under Investigation by firm | Microtek - Ecolab sent a Urgent Medical Device Recall Letter dated July 12, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you purchased this product CMS-6049 and this recall is dating back as far was 2008. Attached is a list affected lot number, however, all lots of the CMS-6049 are being recalled. Please inspect your inventory and return any product to us.To prevent any potential patient harm, discontinue the use of the probe cover.In order to advise the Food and Drug Administration about the effectiveness of this recall, you are requested to complete and return the enclosed questionnaire promptly fax or email it no later than July 29, 2013. It is necessary to report to us even if all products have been used.Any product still in your inventory is to be returned to us using a Return Authorization Number and our shipping number which will be given to you by calling (919) 202-8462Ext. 205. |  | 2610 | 2,610 | US Distribution including the states of CT, FL, IL, NJ, NY, PA, TN, VA and WI. | Terminated | April 08, 2014 | 270 | 2610 | Not_Computer
Z-0975-2014 | 67384 | No | KKY | Polytetrafluoroethylene with carbon fibers composite implant material. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker CMF Customized Implant Kit X-Large; 5444-0-401 X-Large Implant (2) |  |  | December 06, 2013 | 2013 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field that an incorrect patient-specific Simplex P CMF Customized Implant Kit was delivered to a hospital. | Labeling Change Control | Stryker Orthopaedics sent Urgent Medical Device Recall Notification letters/Product Recall Acknowledgement Forms dated 12/6/2013 to US branch/agency. surgeons, hospital risk managers, and sales representatives. |  | 1 | one unit | Distributed in PA. | Terminated | May 06, 2015 | 516 | 1 | Not_Computer
Z-0981-2014 | 67343 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare, PET Discovery 610, Discovery 710, Discovery 600 and Discovery 690. |  |  | November 20, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to scatter overcorrection associated with the List Mode Replay feature of your Discovery 600, 610, 710 and Discovery 690. | Software Manufacturing/Software Deployment | GE Healthcare sent an Urgent Medical Device Correction letter dated November 20, 2013, and December 4, 2013, to all affected consignees. The letters described the Safety Issue, Safety Instructions, Affected Products Details, Product Correction and Contact Information ( In the US Call 1-800-437-1171. For other countries contact your local GE Healthcare Service Representative). Customers were given the following instructions:To avoid potential issues, replay the list mode data prior to another acquisition.In the event another acquisition is made before replaying the list data, or following reboot, do the following steps:Âż Perform a PET only acquisition using the same tracer that is intended for the List Mode Replay.Âż Then, proceed to perform the List Mode Replay and subsequent reconstructions.Âż Make sure to perform a quality check on any images created after any replay of data.AffectedProduct DetailsPET Discovery 610 and PET Discovery 710, with software versions pet_coreload.44 and pet_mfk.44. The software version is visible on the homepage of the service desktop.To determine the software version currently on your system, do the following:1. On the right monitor upper left corner select the Service Icon.2. Verify that the Service window is under Home.3. Locate the SW Build Version listed in the center of the page.For questions regarding this recall call 800-437-1171. |  | 376 | 376 (101 US, 275 OUS) | Worldwide Distribution - USA including AL, AZ, AR, CA, CO, FL, GA, IL, IN, IA, KY, LA, MN, MO, NE, NJ, NM, NY, NC, OH, OK, PA, SD, TN, TX, UT, VA, WA, WV, WI and Internationally to:AUSTRALIA, AUSTRIA BELGIUM BOSNIA AND HERZEGOVINA, BRAZIL, CANADA, CHILE, COLOMBIA, CHINA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FRANCE, FINLAND, GEORGIA, GERMANY, HONG KONG, IRAN, ISRAEL INDIA,M ITALY, JAPAN, JORDAN, KUWAIT, LEBANON, MALAYSIA, MEXICO, MOROCCO, NEW ZEALAND, NORWAY, RUSSIA, REPUBLIC OF KOREA, POLAND, ROMANIA, SAUDI ARABIA, SPAIN, SWEDEN SWITZERLAND, TAIWAN, TUNISIA, THAILAND, TURKEY, UNITED KINGDOM. | Terminated | July 29, 2014 | 251 | 376 | Computer
Z-0983-2014 | 67387 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Invivo Corporation Expression Information Portal (Model IP5) is intended to be used as a secondary display control unit, where the device monitors the vital signs of a patient undergoing an MRI procedure from the MR control room, induction room, or recovery room and then outputs the measured patient data to a hospital information system (HIS). |  |  | December 04, 2013 | 2013 | Invivo Corporation | Invivo Corporation located at 12151 Research Parkway, Orlando, FL 32826 is recalling the Expression Information Portal (Model IP5). Upon power up, the IP5 may fail to produce audio. | Component design/selection | Philips Healthcare sent an "URGENT - MEDICAL DEVICE RECALL" letter dated December 4, 2013, to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to refer to the "Action to be Taken by Customer/User section on the attached pages. A Philips representative will contact the customer regarding their affected devices. All affected devices will have a firmware update installed to correct the problem free of charge. If you need any further information or support concerning this problem, please contact your local Philips representative: 877-468-4861 option 1, then option 2.Philips apologizes for any inconveniences caused by this problem. |  | 299 | 299 | Worldwide Distribution - USA (nationwide) and the countries of Australia, Austria, Denmark, France, Germany, Netherlands, Poland, Singapore, Sweden, Switzerland and United Kingdom. | Open, Classified |  |  | 299 | Not_Computer
Z-0982-2014 | 67370 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, Dash 3000, Dash 4000 and Dash 5000 Patient Monitors, Dash Port2 Docking Station. |  |  | November 27, 2013 | 2013 | GE Healthcare, LLC | GE has become aware of multiple issues affecting the Dash 3000/4000/5000 Patient Monitor and Dash Port 2 Docking Station:DASH: 1. Dash stops communicating on the Unity Network resulting in Loss of Monitoring at the CARESCAPE" CIC Pro Central Station*.2. Dash displays distorted ECG waveforms in combination monitoring mode. 3. Dash incorrectly becomes Time Master on the Unity NetworkÂż.4. | Software design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated November 27,2013. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Contact Technical Support at 1-800-558-7044 or your local Service Representative). |  | 178711 | 178,711 (92,123 US, 86,588 OUS) | Worldwide Distribution: USA (nationwide) including DC and PR; and Internationally to: ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BANGLADESH, BELARUS, BELGIUM, BOLIVIA, BOSNIA-HERZEGOVINA, BRAZIL, BRUNEI, BULGARIA, CAMBODIA, CAMEROON, CANDA, CARIBBEAN, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DJIBOUTI, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GABON, GEORGIA, GERMANY, GHANA, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, IVORY COAST, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIA, MACEDONIA, MALAYSIA, MALTA, MAURITIUS, MEXICO, MOLDOVIA, MOROCCO, NEPAL, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA MONTENEGRO, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, TAJIKISTAN, TANZANIA, THAILAND, TRINIDAD TOBAGO, TUNISIA, TURKEY, TURKMENISTAN, UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM, YEMEN, and ZIMBABWE. | Open, Classified |  |  | 178711 | Not_Computer
Z-0984-2014 | 66958 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay System Analyzer. |  |  | November 26, 2013 | 2013 | Beckman Coulter Inc. | The Access 2 Immunoassay System Analyzer and the UniCel DxI 600 Access may have been manufactured with non-conforming bushings. Non-conforming bushings have a remote probability of generating incorrect results. | Nonconforming Material/Component | Beckman Coulter initiated this recall by sending recall notifications to customers. An Urgent Product Correction letter sent to affected customers. The recall letter dated November 22, 2013 informed customer of the recall, products with description and codes, issue, impact, action, resolution, and contact information. |  | 3 | 3 units | Worldwide Distribution: US Distribution in South Dakota and the countries of Germany, Lebanon, and China. | Terminated | July 27, 2014 | 243 | 4 | Not_Computer
Z-0986-2014 | 67390 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes 3.2 mm Guide Wire 400mm. |  |  | January 27, 2014 | 2014 | Synthes, Inc. | Synthes 3.2 mm Guide Wire 400mm included an incorrect raw material listed on the label. | Labeling Change Control | A recall notification letter, dated January 27, 2013, was sent to End Users and sales consultants. |  | 1132096 | 1,132,096 | Worldwide Distribution-USA (nationwide) and the countries of Canada and Switzerland. | Terminated | December 18, 2015 | 690 | 1132096 | Not_Computer
Z-0987-2014 | 67236 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SAFSITEÂż Injection Sites and IV Sets manufactured with SAFSITEÂż Injection Sites. |  |  | January 08, 2014 | 2014 | B. Braun Medical, Inc. | Complaints were received reporting leakage on certain lots of SAFSITEÂż Injection Sites and IV Sets manufactured with SAFSITEÂż Injection Sites.The SAFSITEÂż Injection Site is intended to provide needle free access on IV and Extension Sets. | Process control | A recall notification letter, dated January 7, 2014, was sent to Users and distributors. |  | 2357944 | 2,357,944 | Nationwide Distribution. | Terminated | November 03, 2014 | 299 | 2357944 | Not_Computer
Z-0988-2014 | 67241 | No | LEZ |  |  | Ear Nose &amp; Throat | Enforcement Discretion | Siemens Prisma 2K Hearing Aid |  |  | July 01, 2005 | 2005 | Siemens Hearing Instruments, Inc | Siemens became aware of an issue where an infant or child could open the tamper resistant battery compartment of the Siemens Prisma 2K hearing aid under certain unintended use conditions. | Device Design | Siemens sent a notice letter to affected customers in 7/2005 and 12/2005. The letter identified the affected product, problem and actions to be taken. Siemens will install an additional security measure to affected product at no charge. For questions contact your Sales Representative at 800-766-4500. |  | 738 | 738 units | USA Nationwide Distribution | Terminated | July 06, 2017 | 4388 | 738 | Not_Computer
Z-0989-2014 | 67174 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Aesculap proSA Adjustment Disc Size Large (L) |  |  | December 23, 2013 | 2013 | Aesculap, Inc. | The accuracy is out of specification for the Aesculap proSA Adjustment Disc Size L. | Employee error | Aesculap sent an Important Correction and Removal Notification dated December 23, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers..Yom"Aesculap Sales Representative will contact you to schedule a time to remove the FV795T from Set FV792T. Please use FV790T in the absence of the FV795T to ensure the correct adjustment has been made. Please complete the attached Inventory Sheet which is necessary to comply with FDA regulations. When completing the inventory sheet, please fill in the quantity being returned. If you cannot locate the product, please provide an explanation as to why the inventory will not be returned (discarded, etc.).AIC (USA) appreciates your cooperation on this matter and apologizes for the inconvenience this may cause. Thank you for your patience and continued support of this product. Please call (610) 984-9265 or (610) 984-9414 with any questions. |  | 79 | 79 | Worldwide Distribution - US Distribution including the states of AZ, MD, MA, WI, MI, MN, DE, KS, and FL., and the countries of Australia, Germany, Great Britain, Norway and Japan. | Terminated | October 01, 2014 | 282 | 79 | Not_Computer
Z-0992-2014 | 67396 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Matrix Mandible Short Threaded Drill Guide intended for oral, maxillofacial surgery; trauma and reconstructive surgery, specifically for fractures of the subcondylar region of the mandible and fractures of the condylar basis region of the mandible. |  |  | January 27, 2014 | 2014 | Synthes, Inc. | One lot of Matrix Mandible Short Threaded Drill Guides (part No. 03.503.043) has a gray colored band instead of a green colored band as pictured in the Technique Guide. | Process control | A recall letter, dated January 27, 2014, was sent to end users and sales consultants. |  | 12 | 12 | US distribution to MI, TX, and CA. | Terminated | July 20, 2015 | 539 | 12 | Not_Computer
Z-0991-2014 | 67492 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | MORPH ACCESSPRO Steerable Introducer; |  |  | February 07, 2014 | 2014 | Biocardia, Inc. | Tyvek packaging may be worn, causing loss of package integrity. | Process control | Sales Representative called each affected hospital on February 7, 2014, to notify them of the recall. BioCardia has initiated a recall of Morph AccessPro due to a potential compromise of the sterile barrier. Use of this product may result in patient infection. Sales representatives contacted hospital sites identified in Attachment 3 by telephone. Script provided as follows: "BioCardia has initiated a recall of Morph Access Pro due to a potential compromise of the sterile barrier. Use of this product may result in patient infection. Our records indicate that we sold you. The recall is limited to these units in this lot sold to you. We would like to confirm if you still have the unit(s). If you do, we ask that you quarantine the unit(s) and our customer service department will contact you to coordinate the return of the unit(s)."For further questions please call (650) 226-0133. |  | 11 | 11 units | US Distribution including the states of FL, MA, WY. OH, MI and KS. | Terminated | April 03, 2014 | 55 | 11 | Not_Computer
Z-0993-2014 | 67076 | No | IPL | Standup wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Superstand stand-up wheelchair |  |  | May 16, 0013 | 13 | The Standing Company | Power Superstand Standing Wheel chair, flex shaft (the flexible wand which supports the joystick control module) may fracture if it is bent repeatedly and /or beyond 90 degrees. A newly designed Flexshaft which corrects this problem is available at no charge. | Component design/selection | The Standing Company sent an Urgent Medical Device Field Correction/Part Recall dated February 6, 2014, to all affected customers. The Urgent Recall notification identified the Flexshaft holding the control joystick on all Power Standing Wheelchairs were being replaced after investigation of complaints found the flexshaft could break because the inner spring of the shaft had failed, although the outer core remained intact. The notice explained how the failure would appear and that the break could happen without warning. The notice explained that the break would not cause the flexshaft to fall off the chair, or hit the ground and the chair could still be mobilized by having the user lean over and pick up the joystick. The notice warned customers not to over bend the flexshaft until replacement was performed. Consignees were asked to call 1-800-STANDING ( 1 800-782-6346) immediately if their wheelchair needed immediate flexshaft replacement. The notice stated that consignees would be contacted to set up a time and date to replace their current flexshaft, at no charge. Consignees were asked to complete and return the response form with customer contact information to Standing Wheelchair Company 5848 Dixie Hwy, Saginaw, MI 48601; Fax (989) 746-9185 or E-mail mail@thestandingcompany.com. Questions and Concerns should be addressed to Customer Service Manager Chris Moran, 1 800 782-6346.For questions regarding this recall call 989-746-9100. |  | 119 | approximately 119 | Nationwide Distribution including AL, CA, DE, IA, IL , IN, KS, KY, MI, MO, NJ, OH, OK, TX, WA, WI,and WV. | Terminated | June 25, 2014 | 730890 | 119 | Not_Computer
Z-0990-2014 | 67154 | Yes | FKB | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Orbit soft Infusion sets, 9 mm soft cannula, P10 (42") tubing; Single-Use Set |  |  | December 20, 2013 | 2013 | Perrigo Diabetes Care | Perrigo was notified by our supplier, Ypsomed AG, that they are recalling these batches due to being manufactured with a non-conforming female luer-lock connector. The luer-lock connectors are incompatible with standard male luer-lock connectors for use in connection to compatible infusion pumps. Perrigo became was notified of a quality complaint related to this issue on November 20, 2013 from a | Nonconforming Material/Component | Perrigo Diabetes Care sent a recall notice dated December 16, 2013, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to segregate the affected productt and request return of all undistributed/unused sets from their customer warehouses to Omni Manufacturing Services S.A. de C.V., Parque Industrial Queretaro, La Noria Lt. 13. Mz 3 No. 125, Santo Rosa Jauregui, MX-76220 Santiago de Queretaro. Customers were instructed to distribute the attached "Urgent Field Safety Notice" to all users or patients who have been shipped any products from the affected lots. Customers with questions were instructed to call Perrigo Diabetes Care at 1-888-775-1997. |  | 7924 | 7924 units | Nationwide Distribution including GA &amp; TX | Terminated | August 05, 2014 | 228 | 7924 | Not_Computer
Z-1045-2014 | 67505 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | CartoUnivu Module, Model: Carto 3 Software versions 3.2.2 and 3.2.3; Catalog No. KT-5400-124 |  |  | February 12, 2014 | 2014 | Biosense Webster, Inc. | Biosense Webster has initiated a recall of the CartoUnivu Module within the Carto 3 System (V3.2.2 and V3.2.3) when used with Siemens Axiom Artis VB Fluoroscopy systems in particular procedural configurations only. If the table rotation is not returned to zero point, there is a misalignment between the Carto 3 System map display and the fluoroscopic capture. | Other | Biosense Webster sent an Urgent Field Safety Notice dated February 12, 2014, to all affected customers to inform them that Biosense Webster has become aware of an issue with a particular configuration of the CartoUnivu Module withn the Carto 3 System (v3.2.2 and v3.2.3) when used with Siemens Axiom Artis VB Fluoroscopy systems in particular procedural configurations only. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions are instructed to contact their Biosense Webster sales representative or cal (866) 473-7823, Monday through Friday from 7am to 8pm EST. Customers with questions about the acknowledgement form and its return, are instructed to call (909) 839-8468.For questions regarding this recall call 909-839-8500. |  | 8 | 8 units total (5 units in US) | Worldwide Distribution - USA including CA, KY, NJ, MA, and UT and Internationally to Belgium, Austria, and Germany. | Terminated | May 23, 2014 | 100 | 8 | Not_Computer
Z-1063-2014 | 67446 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Gelzone Shoulder Sleeve, Catalog Number SLV634, Lot Number 852766. |  |  | January 29, 2014 | 2014 | Implantech Associates Inc | Implantech initiated the voluntary recall of Gelzone Shoulder Sleever, SLV-634, Lot Number 852766 because it may not satisfy customer requirements for the robustness of the Velcro closure. | Nonconforming Material/Component | Implantech sent a Recall Notification letters dated January 29, 2014 via mail and fax to the affected customers. The recall notification letter informed customers of the voluntary recall, product description with code information, reason for recall, contact information, and actions to be taken. The letter also contains a response section to be completed and returned via fax. For questions call 1-800-733-0833 x235. |  | 300 | 300 | Worldwide Distribution - USA (nationwide) in the states of: CA, AZ, IL, NC, CO and the country of: South Africa. | Terminated | March 12, 2015 | 407 | 300 | Not_Computer
Z-1058-2014 | 67090 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | External Implant RBT, D3 Threadform 4.0 X 10.5mm, REF 40105D3, STERILE R, Rx Only, BioHorizons(R), 2300 Riverchase Center, Birmingham, AL 35244 |  |  | December 10, 2013 | 2013 | BioHorizons Implant Systems Inc | The body length of the External Implant 40105D3 Lot 1302762 labeled as 10.5mm in length is in fact 12mm in length. | Mixed-up of materials/components | The firm, BioHorizons, sent an "Urgent Medical Device Recall" letter dated December 10, 2013 to its Consignees. The letter described the product, problem and actions to be taken. The customers were instructed to return unused affected implants and receive replacements free of charge; complete and return the enclosed Customer Reply Form via fax to (205) 484-2195 or scan to PDF and e-mail to implants@biohorizons.com; If you have affected implants, upon receipt of your completed Customer Reply Form, you will be contacted by BioHorizons Customer Care and provided a Return Authorization number. In addition, if you have already placed any affected implants without incident, no special patient action needs to be taken.If you have any questions regarding this information, please contact our Customer Care Department at (888) 246-8338, Monday through Friday, 7:30am to 6:00pm Central Time. |  | 119 | 119 units | Worldwide distribution: US (nationwide) to states of: LA, CA, NC, IL, PA, and NY; and Internationally to: Seoul, Korea; New Delhi, India; Taipei City, Taiwan; and Madrid, Spain. | Terminated | March 28, 2014 | 108 | 119 | Not_Computer
Z-0994-2014 | 67406 | [6 empty] | Integra Cusa Excel Rx Only |  |  | December 05, 2013 | 2013 | Integra LifeSciences Corp. | The blue COAG button on the CUSA Excel CEM nosecone could stick in the 'activated' position. | Material/Component Contamination | Integra sent an Urgent: Medical Device Correction Letter/Acknowledgement Return Form dated December 5, 2013 via traceable courier to their customers. The notification: advises customers of the nature of the issue and how it could affect CEM Nosecone use, advises customers what to do if the issue occurs and requests customers complete and return the Acknowledgment Return Form. The form should be faxed to rc@integralife.com or 1-609-275-9445.Any questions can be directed to customer service at 855-532-1723. |  | 17253 | 17253 boxes | Worldide Distribution - USA Nationwide | Open, Classified |  |  | 17253 | Not_Computer
Z-1005-2014 | 66500 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | VERSA-FX II KEYLESS TUBE, VERSA-FX II KEYLESS SHORT, VERSA-FX II SHORT TUBE PL, VERSA-FX II STD TUBE PLAT and VERSA-FX II SUPRACONDYLR |  |  | October 11, 2013 | 2013 | Zimmer, Inc. | Zimmer recently conducted a review of all historic packaging validations at its Zimmer Manufacturing B.V. (ZMBV) facility in Ponce, Puerto Rico. Based on this review, it was determined that the packaging operations conducted in the Building 1 packaging operation at this facility were not properly validated. As a result, Zimmer conducted seal integrity testing to evaluate if previously packaged product from this facility met the established requirements for sterile barrier integrity. The acceptance criteria established for this testing was 99% reliability with a 90% confidence interval. These packaging tray configurations met the acceptance criteria for sterile barrier integrity of both the inner and outer trays, however, Zimmer has decided to recall all manufacturing lots that have not expired and that were sterile packaged in the ZMBV facility Building 1 packaging operations. | Packaging | Distributors received a letter identifying the issue and their responsibilities dated November 4, 2013. These responsibilities include quarantining the affected fielded inventory in their possession, providing hospital risk managers with the recall notification letter and retrieving affected fielded inventory and direct sales from the hospitals. |  | 7044680 | 7,044,680 total | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, Canary Islands, China, Czech Republic, Denmark, Finland, France, Germany, Greece, India, Italy, Japan, Korea, Luxemburg, Netherlands, New Zealand, Norway, Portugal Reunion, Russian Federation, Singapore, Slovakia, Spain, Sweden, Switzerland, Turkey, Taiwan, Thailand, and United Kingdom. | Terminated | April 14, 2017 | 1281 | 7044680 | Not_Computer
Z-1061-2014 | 67230 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Synergy XVI |  |  | November 25, 2013 | 2013 | Elekta, Inc. | Potential collision risk when using XVI and external beam shaping devices. | Software design | Elekta sent an Important Field Safety Notice dated November 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Corrective Action #1 - Potential collision risk when using XVI and external beam shaping devices. (Only software version R4.2.1 from this notice are applicable to this recall). Recommendations to Users - Monitor the Actual column of the Table Move Assistant dialog box. If the Actual is blank, do not press the ASU button on the FKP. To correct the error - (1) Restart XVI. (2) Put the patient in their initial setup position. (3) Scan the patient again, and continue the usual workflow. Always monitor the patient when do RATM. When you do Volume View acquisitions with the Apex collimator attached, make sure that the stop angle of the Volume View parameters is in the lower half of the gantry rotation. This decreases the risk of a collision with the patient. Read and follow the advice in the Important Field Safety Notice (200 01 507 081) and acknowledge receipt to your Elekta representative within 30 days. Corrective Action #2 -Permanent Solution - XVI Risk control measure - changes to the software to prevent table movement if the start position has not been recorded correctly. Recommendations for changes to product are - (1) Table zero should default to NIHII not zero. (2) Check that the table zero is valid because calculating absolute table positions. (3) Check prescribed table position is not more than 2.5cm from actual table position when permitting ASU from outside room. Target Release XVI R.4.2.2 estimated for release April 2014.For further questions please call (770) 300-9725. |  | 312 | 312 | Worldwide Distribution: US (nationwide) including Puerto Rico and the states of: AL, AK, AZ, CA, CO, CT, FL, GA, IL, IA, KY, MD, MA, MI, MN, MS, MT, NE, NV, NJ, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, VT, VA, WA, WV, WI, DC, and Internationally to: Algeria, Australia, Austria, Belarus, Belgium, Brazil, Canada, Chile, China, Cyprus, Czech Republic, France, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Malaysia, Mexico, Pakistan, Portugal, Russia, Singapore, Slovakia, South Korea, Spain, Taiwan, Thailand, and United Kingdom. | Open, Classified |  |  | 312 | Computer
Z-0995-2014 | 67400 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Select 3D Arterial Cannulae. Vented, models 78618 (18 Fr.) and 78620 (20 Fr.). Non-Vented Cannulae, models 78718 (18 Fr.) and 78720 (20 Fr.). Sterilized using Ethylene Oxide. Non-pyrogenic. |  |  | January 30, 2014 | 2014 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic is recalling all models of its Select 3D and Select CAP Arterial Cannula product families due to reports of the cannula body splitting near the suture collar. | Under Investigation by firm | The firm, Medtronic, sent an "Urgent Medical Device Recall" letter dated January, 2014 via UPS on January 30, 2014 to their customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately identify and quarantine all listed product in your inventory; return all listed product to Medtronic and contact Customer Service at 800 854-3570 and reference RADAR #167826 to initiate a return and credit of product; complete and return the attached Customer Confirmation Certificate via fax to Medtronic at 651-367-0612 to the attention of Customer Focused Quality, and share this notification with others in your organization as appropriate.Select 3D and Select CAP Arterial Cannula are not available for purchase at this time. Please contact your Medtronic sales representative or Medtronic LifeLine Technical Services at 1 877-526-7890 or (+1 763) 526-7890 to assist with alternative cannula selection and any additional questions you may have related to this notification. |  | 19483 | 19483 total (7669 US, 11814 OUS) | Worldwide distribution: US (Nationwide) including Puerto Rico, and Internationally to: Algeria, Argentina, Austria, Belgium, Bosnia And Herzegovina, Brazil, Canada, Cyprus, El Salvador, France, Georgia, Germany, Greece, Hungary, India, Ireland, Italy, Japan, Jordan, Kuwait, Netherlands, Oman, Paraguay, Poland, Portugal, Russian Federation, Saudi Arabia, Slovakia, Spain, Sweden, Switzerland, Thailand, Turkey, United Kingdom, Venezuela, and Yemen. | Terminated | October 06, 2014 | 249 | 19483 | Not_Computer
Z-1059-2014 | 67231 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Synergy XVI |  |  | November 25, 2013 | 2013 | Elekta, Inc. | XVI can incorrectly calculate the target position of the treatment table. | Software design | Elekta sent an Important Field Safety Notice dated November 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Potential collision risk when using XVI and external beam shaping devices released November 25, 2013 (Only software versions R4.5, R4.5.1, R4.6 and R5.0 are applicable to this recall). (All software versions R4.6 in this notices are applicable to this recall). Recommendations to Users - Monitor the Actual column of the Table Move Assistant dialog box. If the Actual is blank, do not press the ASU button the FKP. To correct the error - (1) Restart XVI. (2) Put the patient in their initial setup position. (3) Scan the patient again, and continue the usual workflow. Always monitor the patient when you do RATM. When you do Volume View acquisitions with the Apex collimator attached, make sure that the stop angle of the Volume View parameters is in the lower half of the gantry rotation. This decreases the risk of a collision with the patient. Read and follow the advice int he Important Field Safety Notices and acknowledge receipt to your Elekta representative within 30 days. Corrective Action #2 - Permanent Solution - XVI Risk control measure - changes to the software to prevent table movement if the start position has not been recorded correctly. Recommendations for changes to product are - (1) Table zero should default to NIHII not zero. (2) Check that the table zero is valid because calculating absolute table positions. (3) Check prescribed table position is not more than 2.5cm from actual table position when permitting ASU from outside room. Users of software version 4.5, 4.5.1 and 4.6 will receive a mandatory upgrade to XVI 5.0. The estimated release date for the fix for the XVI 5.0 is December 2014.For further questions call ( 770 ) 300-9725. |  | 1028 | 1028 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AL, AK, AZ, AR, CA, CT, FL, GA, ID, IL, KY, LA, MD, MA, MI, MN, MO, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SD, TN, TX, VT, VA, WA, WV, WI, WY and DC., and the countries of Algeria, Australia, Austria, Bahrain, Belarus, Belgium, Brazil, Bulgaria, Canada, China, Columbia, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Iraq, Ireland, Israel, Italy, Japan, Jordan, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom and Venezuela | Open, Classified |  |  | 1028 | Computer
Z-1062-2014 | 67283 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Bausch &amp; Lomb AKREOS AO Micro Incision Lens |  |  | October 16, 2013 | 2013 | Bausch &amp; Lomb Surgical, Inc. | Lens was manufactured with incorrect raw material. | Nonconforming Material/Component | The firm, Bausch + Lomb, telephoned and sent an "URGENT - MEDICAL DEVICE RECALL" letter dated October 17, 2013 to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to: 1) Determine the disposition of the lenses; 2) Complete and provide the enclosed acknowledgement form to the sales respresentatives collecting the lenses, and 3) Return all unused products. If you have any questions, please contact Bausch + Lomb at (800) 338-2020. |  | 336 | 336 IOLs (283 IOLs in the US, 53 IOLs outside the US) | Worldwide Distribution: US (nationwide) and Internationally to: Great Britain, France, Spain, Portugal, Sweden, Russia and Guadeloupe. | Terminated | September 17, 2014 | 336 | 336 | Not_Computer
Z-0980-2014 | 66777 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Merit Medical's Custom Procedural Trays or Kits containing 1% Lidocaine HCl Injection, USP, 10 mg/mL, Lot 25-090-DK. Custom Procedural Tray or kit is a combination of legally marketed medical devices placed into one container for the convenience of the user and used within the limits of the product's intended uses. |  |  | September 30, 2013 | 2013 | Merit Medical Systems, Inc. | Custom Procedural Trays/Kits contain 1% Lidocaine HCl injection which were recalled because one confirmed customer report where visible particulate were found in the primary product container in the form of dark red/black particles. The particular matter may dislodged through the tissue and could possibly cause a pneumothorax or hemothorax. | Other | Merit Medical sent via FedEx an "URGENT PRODUCT RECALL NOTICE" dated September 30, 2013 to all affected customers. The letter identified the affected product, description of problem and recommended user actions to be taken. Consignees were given two options: 1) Based upon immediate patient need, they can keep the affected trays and discard the Lidocaine at the point of use (a) Merit will provide the customer with warning labels, along with replacement Lidocaine for each affected tray based on their current inventory or (b) The label will instruct the user to dispose of the Lidocaine at the point of use and replace it with the new lot. 2) They may return all affected trays remaining in inventory to Merit for credit or replacement. Consignees were requested to complete and sign the Customer Response Form and immediately fax a copy of the completed Customer Response Form to Joyce Braxton at 1-801-316-4852 and mail the completed original Customer Response Form using the provided postage-paid, self-addressed envelope to Merit. If you choose to return the trays, please return the trays as instructed on the Customer Response Form. If customers choose to keep the trays, Merit can send them labels and Lidocaine upon receipt of the Customer Response Form. If you have any questions, please call 1-804-416-1048 and/or 1-801-316-4822.For questions regarding this recall call 804-416-1030. |  | 20872 | 20, 872 trays | Nationwide Distribution including AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, WV. | Terminated | May 13, 2014 | 225 | 20872 | Not_Computer
Z-1057-2014 | 67261 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Driving Cap/Threaded |  |  | January 07, 2014 | 2014 | Synthes, Inc. | The Synthes Driving Cap/Threaded can bind in the IM Nail Radiolucent Insertion Handles when assembled becoming difficult to separate. | Process design | Synthes sent an Urgent Notice - Medical Device Recall Letter dated January 7 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Action to be taken: Our records indicate that your facility has the product(s) subject to this Recall. Synthes asks that you review your inventory and immediately remove the affected lots from stock.Please take the following actions:If you DO have any of the identified devices, please take the following steps:"Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization Number."Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, telephone number and signature in the spaces provided."Return the Verification Form (page-3 of this letter) with the product to:"Credit/Returns, Synthes, 1101 Synthes Avenue, Monument, CO 80132.If you DO NOT have the identified product, please take the following steps:"Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affected product has been located. :Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information."Return the documents to DePuy Synthes by:" Fax: 877-271-1473 or" Scan/email: FieldAction@synthes.comIf you have any questions,please call 610-719-5450. Thank you for your attention to this issue. |  | 143 | 143 | USA (nationwide) Distribution. | Open, Classified |  |  | 143 | Not_Computer
Z-1068-2014 | 66893 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Renasys EZ PLUS, a Canister Component of of Renasys EZ Negative Pressure Wound Therapy System |  |  | November 14, 2013 | 2013 | Smith &amp; Nephew Inc. | Modification of the bacterial overflow guard (filter) and related changes to product labeling. | Device Design | The correction strategy is as follows:" An Urgent Medical Device Correction Notice has been prepared to be distributed to all US consignees via Federal Express delivery for confirmation of the delivery." A copy of the Urgent Medical Device Correction Notice is to be posted on Smith &amp; Nephew's websites, URL corporate, http://www.smith-nephew.com/news-and-media/, and myrenasys.com, concurrent with the distribution of the Notice to consignees." US distribution of the original design of Canister Kits codes (66800912, 66800913, 66801066, 66800423 and 66800058) subject to this correction ceased November 1, 2013." The correction has been implemented on a change being affected basis. The modified design is being distributed to US customers effective November 1, 2013. No product will be removed from the market." Periodic progress reports to FDA, as required per 21 CFR 806 will be submitted monthly until completion of the correction actions.. |  | 6871375 | 6,871,375 canisters | Worldwide Distribution: USA (nationwide) including states of: AL, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KY, LA, MA, MI, MN, MO, MS, MT, NC, NH, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, VA, WA, WV, and Puerto Rico; and the countries of: Austria, Australia, Belgium, Switzerland, China, Cyprus, Germany, Denmark, Spain, Finland, France, UK, Greece, Israel, Italy, Malta, Netherlands, Norway, Paraguay, Russia, Saudi Arabia, Sweden, Thailand, Turkey, South Africa, Mexico, and Canada. | Terminated | August 04, 2014 | 263 | 13670515 | Not_Computer
Z-1054-2014 | 66080 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | STERILE SAMPLE BUILD; REF: SAMP0147; Packaged for Medline Industries, Inc., Mundelein, IL 60060; 1-800-Medline; www.medline.com |  |  | August 14, 2013 | 2013 | Medline Industries, Inc. | Guidewire .035x150 3MMJ TCFC, item number 88241, has a potential for the coating to flake off of the wire. The guidewire is used in a number of surgical convenience kits assembled and marketed by Medline Industries, Inc. | Nonconforming Material/Component | The firm, Medline Industries, Inc., sent a "RECALL IMMEDIATE ACTION REQUIRED" letter dated August 14, 2013, to its direct accounts/customers. The letter described the product, problem and actions to be taken. The customers were instructed to immediately check their stocks for inventory of the affected products, quarantine any affected products, return the product using the enclosed shipping slips for credit-If additional stickers are needed, please call 866-359-1704, and complete and return the enclosed FAX-BACK VERIFICATION FORM listing the quantity of affected product(s) on hand via fax to Medline at 866-767-1290.If you have any questions, please contact 866-359-1704. |  | 5 | 5 units | Nationwide Distribution: US to states of: AL, CA, FL, IL, LA, OH and VA. | Open, Classified |  |  | 2853 | Not_Computer
Z-1069-2014 | 67193 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | VirtuoSaphÂż Plus Endoscopic Vessel Harvesting System, Part number: VSP550, Sterile***Rx Only*** |  |  | December 12, 2013 | 2013 | Terumo Cardiovascular Systems Corporation | Sterility of medical devices intended for use in surgical procedures may be compromised. | Process control | Terumo CVS notified afected end users by phone on December 12, 2013 and advised that they had become aware that the sterile packaging barrier was breached in specific lots of VirtuoSaphÂż Plus Endoscopic Vessel Harvesting Systems. Customers were requested to return all unused affected product and verify that all users at their institution have been made aware of the risks associated with using the affected devices. For questions call 734-663-4145. |  | 26 | 26 units | USA Nationwide Distribution in the state of LA, NJ, and NY. | Terminated | July 28, 2015 | 593 | 26 | Not_Computer
Z-1066-2014 | 67404 | No | MBL | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Reflection Interfit Shell, 52 MM OD, 3 HOLE, REF 71336052, QTY 1, STERILE R, Smith &amp; Nephew, Inc. |  |  | January 23, 2014 | 2014 | Smith &amp; Nephew Inc | Inner spherical radius of the shell is undersized. | Under Investigation by firm | Smith &amp; Nephew Inc. sent an Urgent - Product Recall 1st Notification letter dated January 23, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Actions for Hospital Representatives:1. Please inspect your inventory and locate any unused devices from the above listed product and batch numbers, and quarantine them immediately.2. Complete the last two columns in the Inventory Return Certification Form on the following page, indicating the quantities that need to be returned, and include your phone and fax number in the spaces provided. If you do not have product to return, please place an x in the column No product to return .3. Please contact Smith &amp; Nephews Safety Affairs Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to obtain a return authorization (RA) number.4. Return any affected product to the address listed on the Inventory Return Certification Form. Please indicate the RA number on your return shipment.Actions for Smith &amp; Nephew Inventory Specialists1. Please inspect your inventory and locate any unused devices from the above listed product and batch numbers, cease distribution and use of those products, and quarantine them immediately.2. Carry out a physical count of all affected product in your territory and record this data on the Inventory Return Certification Form on the following page.3. If you do not have product to return, please place an x in the column No Product to Return , and then fill in the Acknowledgement of Responsibility at the bottom of the form.4. Fax a copy of the completed Inventory Return Certification Form to: +1-901-566-79755. If you have affected products, return the product, along with the completed Inventory Return Certification Form, to the address indicated |  | 10 | 10 units | Worldwide Distribution: US Distribution to NJ and FL; and the countries of: Colombia and Costa Rica. | Open, Classified |  |  | 10 | Not_Computer
Z-0963-2014 | 67266 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | ISIS HVT Tracheal Tube Cuffed with Subglottic Secretion Suction Port and Preloaded Stylet, 9.0, Rx only, |  |  | January 06, 2014 | 2014 | Teleflex Medical | The tracheal tube can kink during patient use. | Other | Teleflex Medical sent an Urgent Medical Device Recall Notification letter dated January 6, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to immediately discontinue use and quarantine any affected product. Return the product by completing the enclosed Recall Ackknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service to receive a Return Goods Authorization (RGA) number and instructions for the return of the affected product to Teleflex Medical. Consignees with no affected stock, should also complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881. Consignees with questions were instructed to call 1-866-246-6990.For questions regarding this recall call 919-572-8014. |  | 52489525 | 190,871 each (148, 275 US) | Worldwide Distribution - USA (nationwide) Puerto Rico and Internationally to Canada, Barbados, Buenos Aires, Bolivia, El Salvador, Mexico, Columbia, Chile, Costa Rica, Dominican Republic, Ecuador, Barbados, Belgium, South Korea, Thailand, New Zealand, Netherlands. | Terminated | June 17, 2016 | 893 | 52489525 | Not_Computer
Z-1077-2014 | 67159 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PFC SIGMA Cruciate Retaining (CR); Cemented Femoral devices; (Size 5/RT) |  |  | December 09, 2013 | 2013 | DePuy Orthopaedics, Inc. | Knee implant femoral devices were found to have anomalous microstructure. Porosity and the presence of blocky carbides, within the cobalt chrome part, can decrease the overall fatigue strength. This decrease in strength could cause implants to fracture under what would normally be acceptable loads. | Process design | The recall is extended to the DePuy Distributor, Hospital, and Surgeon levels. The two affected distributors were contacted by telephone. Written communication to the two hospitals and one surgeon who received the devices will be delivered by DePuy Orthopaedics, Inc. via email or regular mail. The one affected surgeon will also be contacted by telephone by the Medical Safety Officer and a metallurgist, who will discuss the surgeon communication with the surgeon.The sales representatives aid customers in the affected device returns, as needed. The devices will be returned through the normal DePuy Returns process, to attention of Returns and marking H13-30C on the outsideof the box. All (21) remaining devices have been verified as being returned and in quarantine.Effectiveness will be determined by product reconciliation, receipt of all hospital reconciliation forms, and receipt of international declaration from the International Affiliate. DePuy will follow-up with the affected hospitals until all hospital reconciliation forms are returned, and DePuy will follow-up with the International Affiliate until all international actions are complete and the international declaration is completed and return. |  | 2 | 2 | Worldwide Distribution-USA including the states of IA, CA, KY, and FL, and the countries of Finland, Germany, Sweden, and Czech Republic. | Terminated | June 10, 2014 | 183 | 2 | Not_Computer
Z-1078-2014 | 67401 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision 500D and Proteus XR/a equipped with Wireless Digital Radiography Option (WDR1). |  |  | January 20, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a software issue associated with patient selection from the worklist on your WDR1(Precision 500D and Proteus XR/a equipped with Wireless Digital Radiography Option) that may impact patient safety.Intermittently when a large number of patient exams are retrieved from the HIS/RIS system, the exams when displayed on the WDR1 are not properly displayed. I | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated January 20, 2014. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Mangers and Radiologists. The letter described the Safety Issue, Affected Product Details, Safety Instructions, Product Correction and Contact Information ( In USA call 800-437-1171, For other countries contact your local GE Healthcare Service Representative). |  | 359 | 359 (333 US, 26 OUS). | Worldwide Distribution - USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, DC, FL. GA, HI. IA, ID, IL, IN, LA, KY, ME, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OK, OR, PA, TN ,KY, TX, UT, VA, WA, WV, WI and WY., and the countries of ARGENTINA, AUSTRALIA, CANADA, FRANCE, IRELAND, ITALY, JORDAN, ROMANIA, RUSSIA, UNITED ARAB EMIRATES and UNITED KINGDOM. | Terminated | November 03, 2014 | 287 | 359 | Computer
Z-1073-2014 | 67405 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage PureFlow-B Solution; Model/Catalog Number: RFP-401. Product is packaged in a case containing two individual 5L bags. |  |  | January 24, 2014 | 2014 | NxStage Medical, Inc. | Mislabeled product. | Packaging process control | NxStage sent a Voluntary Medical Device Recall letter dated January 24, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They were asked to check all boxes in inventory for the recalled lot and segregate and quarantine them. In addition, they were asked to complete the recall reply form and return it by fax to NxStage Medical as well as to contact NxStage Customer service to arrange for return of all affected product and for replacement to be sent. Distributor was asked to notify all downstream customers and to provide NxStage Medical with a list of customers. Please complete and fax this form to 978-687-4810 even if you do not have any of the affected product on hand.If you have any questions or comments, please feel free to contact NxStage Customer Service at 1-866- NXSTAGE (1-866-697-8243). |  | 969 | 969 cases (2 bags per case) | US Distribution including the states of NY, TX, ID, AR, CA and AL | Terminated | August 29, 2014 | 217 | 969 | Not_Computer
Z-1075-2014 | 67066 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | pCO2 D788 Membranes. Distributed by Radiometer America, Inc., Westlake, OH. |  |  | July 05, 2013 | 2013 | Radiometer America Inc | RADIOMETER has become aware that some D788 pCO2 membranes can cause biased measurement results on patient results as well as QC results. | Under Investigation by firm | On 7/18/2013, the firm sent Product Notification letters to their customers. |  | 38890 | 38,890 units (7,860 units imported into the US) | Worldwide Distribution-USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY, and the countries of Bermuda and Canada. | Terminated | September 19, 2015 | 806 | 38890 | Not_Computer
Z-1076-2014 | 67235 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista Enzyme 1 Calibrator (ENZ 1 CAL) lots 3FD034 and 3HD012. |  |  | January 06, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Complaints were received regarding Quality Control (QC) shifts with Pseudocholinesterase (PCHE) and imprecision during PCHE calibration with the Dimension Vista Enzyme 1 Calibrator (ENZ 1 CAL). | Nonconforming Material/Component | A recall notification letter, dated January 2014, was sent to Users. |  | 3496 | 3496 | Nationwide Distribution. | Terminated | May 06, 2015 | 485 | 3496 | Not_Computer
Z-1074-2014 | 67307 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Single and Dual use Emission Computer Tomography gamma cameras (BrightView XCT, BrightView X &amp; BrightView SPECT) Philips Healthcare Systems, Cleveland, OH. |  |  | December 19, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | Three problems were reported to the firm:When positioning for a relative 180 degree non circular acquisition the second (bottom) detector was not positioned correctly causing an inappropriate rotation orbit calculation. When the acquisition was started and the detectors began to move into position, the incorrect calculation resulted in the detectors coming in contact with the patient.As | Software design | On 12/19/2013 the firm sent Field Safety Notices to their customers. |  | 1143 | 1,143 units | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, VT, WA, WI, WV &amp; WY, and the countries of Algeria, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Gabon, Germany, Greece, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom &amp; Viet Nam. | Terminated | January 12, 2015 | 389 | 1143 | Not_Computer
Z-1079-2014 | 66430 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trabecular Metal" Reverse Shoulder System |  |  | October 17, 2013 | 2013 | Zimmer, Inc. | As a result of complaints for disassociation of the TM Reverse Glenosphere from the Baseplate, Zimmer updated the related surgical technique and released a video to provide additional instructions to users on how to properly assemble the Glenosphere to the Baseplate. | Device Design | In November and December 2012 Sales staff were provided updated surgical technique and associated video to share with current customers, surgeons, and risk managers. October 17, 2013 Distributors ( via E-mail),and Risk Managers and Surgeons (via traceable courier) were sent written Urgent Recall notifications of these labeling changes. Current customers are identified through an invoice search and distribution history for the affected devices for the past 12 months, since 9/1/12. |  | 73002 | 73,002 units | Worldwide Distribution-USA (nationwide) and the countries of Germany, India, Malaysia, China, Singapore, Korea, Taiwan, and Brazil. | Terminated | September 12, 2014 | 330 | 73002 | Not_Computer
Z-1080-2014 | 67175 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Siemens Artis zee Ceiling Systems. |  |  | December 16, 2013 | 2013 | Siemens Medical Solutions USA, Inc | During the lifetime of Artis zee Ceiling systems, there is an increased wear due to potential sub-optimal cable guidance that can lead to cable damage. | Device Design | In December 2013, a recall notification letter was sent to Users. |  | 284 | 284 | Nationwide Distribution. | Terminated | March 09, 2015 | 448 | 284 | Not_Computer
Z-1090-2014 | 67429 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator; Model(s), Catalogue, or Code Number: M3535A, M3536A, M3536J, M3536MC |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | A component of the MRx Processor Board may be susceptible to damage from electrostatic discharge (ESD), which can disrupt ECG and SpO2 functionality leading to an inability to perform. | Device Design | The firm, Philips, sent an "URGENT - Medical Device Correction" letter dated February, 2014 with an attached Field Safety Notice to each known affected customer on 2/18/14. The letter describes the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The customers were instructed to continue to use your MRx prior to receiving the hardware upgrade; if a failure has been detected that may disable some ECG and SpO2 functionality during patient care, treat your patient per existing protocols and if the condition persists, take the device out of use and call for service from Philips authorized service personnel.Philips is initiating a correction to affected devices. A hardware upgrade will be provided free of charge for all units affected by this issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the hardware upgrade.If you have questions regarding this notification or need any further information or support, please contact your local Philips representative or call us at 1-800-722-9377. |  | 39128 | 39,128 | Worldwide Distribution: US (Nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, GU, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and Internationally to: Canada, AFGHANISTAN, ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELGIUM, BOSNIA AND HERZEGOWINA, BRAZIL, BRUNEI DARUSSALAM, BULGARIA, CHILE, CHINA, COLOMBIA, CROATIA (local name: Hrvatska), CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GUADELOUPE, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MACEDONIA, MALAYSIA, MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NETHERLANDS ANTILLES, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NORWAY, OMAN, PAKISTAN, PALESTINE, PANAMA, PAPUA NEW GUINEA, PERU, PHILIPPINES, POLAND, PORTUGAL, PUERTO RICO, QATAR, REUNION, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWEDEN, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, TRINIDAD AND TOBAGO, TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIET NAM, and YEMEN. | Open, Classified |  |  | 39128 | Not_Computer
Z-1091-2014 | 67342 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | The FOCUS is a microprocessor controlled intra-oral x-ray unit with a Load Matching Frequency (LMF) DC generator, which produces dental images on digital and film media. For the digital images the FOCUS is connected to the Sigma USB terminal which is connected to a PC, which reads the image data from a sensor and displays the image on the PC monitor. In order to make installation and service faster and easier the electronics is integrated into the horizontal arm. |  |  | January 14, 2014 | 2014 | Instrumentarium Dental, PaloDEx Group Oy | Instrumentarium Dental has delivered a number of FOCUS 3.0 Intraoral X-ray units in November 2013 with a potential flaw in the mounting hardware. A number of wall mount units were delivered with FOCUS X-ray machines that had not been properly welded at our supplier. As a result of the incomplete welding operation, the wall mount unit on the FOCUS X-ray unit may not properly support the unit as exp | Device Design | The firm, Instrumentarium, sent an "URGENT MEDICAL DEVICE RECALL" letter dated January 14, 2014 was sent to all PaloDEx direct accounts (including subsidiaries). The letter inested recipients to perform a field inspection on potentially defective FOCUS 3.0 X-ray units (units with serial numbers F21221 - F21641 and F21643). If a defective wall mount is discovered during the inspection, the FOCUS X-ray unit must be taken out of service and the wall mount replaced before continuing use of the device. Replacement wall mounts will be made available free of charge for the field repair. A Technical Bulletin including instructions on how to perform the inspection was included with the letter. Direct accounts were further instructed to return the enclosed Acknowledgement Letter confirming receipt of the URGENT MEDICAL DEVICE RECALL letter by mail to Instrumentarium Dental, Jussi Tanskanen, Nahkelantie 160, Fl-04300 Finland; or, via e-mail to Jussi.tanskanen@instrudental.com. The direct accounts are to complete and return the enclosed Confirmation Letter to document that all of the potentially affected units have been inspected and repaired if necessary to Instrumentarium Dental, Jussi Tanskanen, Nahkelantie 160, Fl-04300 Finland; or, via e-mail to Jussi.tanskanen@instrudental.com. A nearly identical URGENT MEDICAL DEVICE RECALL letter dated January 22, 2014 was sent to all US customers. The instructions were the same as noted above in the January 14, 2014 letter to PaloDEx direct accounts.If there are any questions, do not hesitate to contact us at the following: Marketing Director, General dentistryMika.anttila@instrudental.comTel. +358 40 768 3340For more information you can contact also:Director, Quality and Regulatory AffairsTarja.paivinen@instrudental.comTel. +358 40 719 0513After sales ManagerMika.Kyllonen@instrudental.comTel. + 358 40 080 5305 |  | 101 | 101 units (US Distribution) | Worldwide Distribution: US (Nationwide) including states of: AK, AZ, CA, CO, GA, IL, KY, MD, MT, NC, NV, OR, PA, TN, TX, UT, VA, WA,and WI; and Internationally to: Belgium Brazil, Chile, China, Czech Republic, France, Germany, Italy, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | May 22, 2015 | 493 | 101 | Not_Computer
Z-1089-2014 | 66660 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Azure Anterior Cervical Plate System, 2-Level Plate, 36 mm, Implant Grade Titanium and Nitinol, Orthofix Inc., RX Only |  |  | October 29, 2013 | 2013 | Orthofix, Inc | There is a possibility that the locking mechanism of the Azure Anterior Cervical Plate may fracture during or after implantation. | Component design/selection | Orthofix sent a Urgent Medical Device Recall Notification dated October 29, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Immediate Action:1. Immediately cease any further distribution or use of the AZURE" plates.2. Notify your surgeons and staff customers that Azure" plates cannot be used and are to be returned to Orthofix. at no cost to you, see Instructions below.3. Please complete the attached reply form indicating that you have received this letter.Product Return Instructions: To return the Azure plates, please follow these Instructions. 1. Check you inventory of Azura Implant Trays..2. Complete the attached form with the quantity you are returning.3. Contact your local Orthofix representative or Orthofix customer servicerepresentative at 888--298-5700 for return instructions and shipment information and replacement with a different Orthofix Anterior Cervical Plate System.. |  | 4172 | 4,172 units | Worldwide Distribution: USA (nationwide) and the country of Germany. | Terminated | December 09, 2014 | 406 | 4172 | Not_Computer
Z-1086-2014 | 67388 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Specialty Straight Static Inserter Assembly |  |  | December 18, 2013 | 2013 | Stryker Spine | Stryker has received complaints from customers relating to fracture along the weld that binds the handle to the shaft of the static inserter. | Device Design | Stryker sent Urgent Medical Device Removal Notification letter and to their customers via email on 12/18/2013. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine their inventory, return affected product using enclosed pre-paid mailing label, complete Customer Response Form and fax to Regulatory Compliance at 855-632-9049 or mail it to Spine-RegulatoryActions@Stryker.com. For questions call 201-760-8287, 201-760-8298 or 201-760-8041. |  | 21 | 21 units | US Nationwide Distribution | Terminated | August 25, 2014 | 250 | 21 | Not_Computer
Z-1087-2014 | 66774 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products Na+ Slides, manufactured by Ortho Clinical Diagnostics 513 Technology Blvd Rochester, NY 14652 |  |  | October 17, 2013 | 2013 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics has recalled VITROS Chemistry Products Na+ Slides and the Calibrator Kit 2 due to false positive serum and plasma sample results. | Under Investigation by firm | Ortho Clinical Diagnostics sent an Urgent Product Correction Notification dated October 17, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Ortho Clinical Diagnostics recalled VITROS Chemistry Products Na+ Slides and the Calibrator Kit 2 due to false positive serum and plasma sample. Please forward a copy of the notification letter if the products were further distributed. Complete and return the attached Confirmation of Receipt form by October 25, 2013. Customers with questions please call Customer Technical Service at 1-800-421-3311. |  | 548663 | 548,663 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY., and the countries of Australia, Brazil, Canada, Chile, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, Germany, France, Italy and Spain. | Open, Classified |  |  | 548663 | Not_Computer
Z-1072-2014 | 67004 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare - GiraffeÂż and PandaÂż iRes Infant Warmer |  |  | October 01, 2013 | 2013 | GE Healthcare, LLC | Oxygen and air wall inlet fittings on the back panel of the resuscitation systems were reversed during assembly. | Component design/selection | GE Healthcare sent an Urgent Medical Device Correction letter dated November 18, 2013, with a return receipt to all affected customers. The letter identified the products, the problem, and safety instructions to be taken by the customer along with the product correction. It is recommended that all potential users in your facility are made aware of this safety notification and the recommended actions. Safety issues described the description and affected product details, safety instructions, photos showing the back of the resuscitation system where the Air and Oxygen fittings are located and an appendix of the Affected Product Serial Number Definition and Configurations of Fittings. For questions or concerns regarding this notification, please call the phone number listed below: United States: 1-800-345-2700. |  | 172 | 12, 373 units total, 8,807 = US, 3,394 = OUS, 172 = unknown | Worldwide Distribution - USA Nationwide including the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL,GA. HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and the countries of Albania, Austria, Australia, Bahrain, Belarus, Belgium, Canada, China, Costa Rica, Cyprus, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Greece, Guatemala, Germany, Honduras, Hong Kong, Hungry, India, Indonesia, Ireland, Italy, Kazakhstan, Korea, Lithuania, Malaysia, Mexico, Netherlands, Norway, Peru, Poland, Romania, Russia, Saudi Arabia, South Africia, Spain, Sweden, Switzerland Turkey, United Arab, United Kingdom, and Venezuela. | Terminated | July 22, 2016 | 1025 | 172 | Not_Computer
Z-0978-2014 | 65309 | No | IXI | Radiation therapy beam-shaping block. | Radiology | Radiology | 510(k) | DMLC IV-ERGO |  |  | May 01, 2013 | 2013 | Elekta, Inc. | It is possible for an intermittent error to occur in the gantry sensor of the DMLC IV Systems | Component design/selection | Elekta sent an Important Field Safety Notice letter dated May 2013 to affected customers. The letter identified the scope, problem, clinical impact, solution, and safety reference. For question call +44 (0)1293 544422 |  | 35 | 35 units | US Nationwide Distribution in the states of CT, FL, GA, IA, MO, NY, ND, OK, RI, TX, VA, WA | Terminated | January 29, 2015 | 638 | 35 | Not_Computer
Z-1109-2014 | 67483 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL, Defibrillator/Monitor |  |  | February 14, 2014 | 2014 | Philips Medical Systems, Inc. | Philips HeartStart XL may experience a power board failure causing failure to defibrillate | Component design/selection | Philips Healthcare issued an Urgent Medical Device Correction notification/Field Safety Notice dated February 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed Philips Healthcare Representative will perform a hardware upgrade free of charge. Customers were instructed while awaiting the replacement power board for your device, you may continue to use the HeartStart XL. Identify a readily available backup defibrillator to use in the event the affected HeartStart XL fails to deliver therapy. If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377 |  | 5234 | 5234 units | Worldwide Distribution and USA Nationwide and the countries of AUSTRALIA, BANGLADESH, BELGIUM, BRAZIL, CANADA, CHINA, COLOMBIA, ETHIOPIA, FINLAND, FRANCE, GEORGIA, GERMANY, HUNGARY, INDIA, INDONESIA, IRELAND, ITALY, JAPAN, KAZAKHSTAN, KOREA, REPUBLIC OF, MALAYSIA, MEXICO, MONACO, NEPAL, NETHERLANDS, PORTUGAL, SINGAPORE, SOUTH AFRICA, SPAIN, SRI LANKA, SWITZERLAND, TAIWAN, PROVINCE OF CHINA, THAILAND, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, and VIET NAM, | Open, Classified |  |  | 5234 | Not_Computer
Z-1111-2014 | 67441 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobasÂż PCR Urine Kit 100 PKT IVD |  |  | January 06, 2014 | 2014 | Roche Molecular Systems, Inc. | Numerous complaints have been filed that leaking cobas Âż PCR media 4.3 mL IVD have been received from different lots of thecobas Âż PCR Urine kits, cobas Âż PCR Female Swab kits, and cobas Âż PCR media kits. | Device Design | Roche Molecular Diagnostics Inc. sent an Urgent Medical Device Correction letter/Fax-back form dated January 6, 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to follow the test-skpecific Instruction For Use for handling of cobas PCR media 4.3 ml IVD. The use of protective equipment should be worn as indicated in the labeling. Inspection each tube and/or blister packet for evidence of leakage prior to use. If leakage of the cobas PCR media 4.3 mL IVD is observed, do not use the tubes. Discard the affected tubes per local guidelines and contact Roche Support Network Customer Support Center at 1-800-526-1247. Complete the attached fax form and fax it to 1-855-695-8564. File this Urgent Medical Device Correction (UMDC) for future reference. Customers with questions were instructed to call 1-800-526-1247.For questions regarding this recall call 908-253-7200. |  | 84 | 84 kit batches | Nationwide Distribution | Terminated | August 07, 2014 | 213 | 84 | Not_Computer
Z-1097-2014 | 67250 | [6 empty] | The SYMBIONIC LEG is an integrated prosthetic leg for transfemoral and knee disarticulation amputees. |  |  | October 15, 2012 | 2012 | Ossur Americas, Inc. | OSSUR is recalling the Symbionic Leg because it is not recognizing the low battery warning signal from the Symbionic Leg before it powers down. | Other | Ossur Americas contacted all affected customers via phone on October 15, 2012 informing them that a software update for this product is required, along with the recommended instructions to return the product. A loaner prosthetic will be provided to the patient while the software update is being implemented. |  | 144 | 144 units | US Nationwide distribution. | Terminated | February 28, 2014 | 501 | 144 | Computer
Z-1107-2014 | 67328 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor |  |  | January 24, 2014 | 2014 | Philips Medical Systems, Inc. | HeartStart XL+ battery charge time to 100% capacity at 35ÂżC (95ÂżF) is slightly longer than the specified 3 hour duration as stated in the labeling | Component design/selection | Philips Healthcare issued a Customer Information Letter on January 24, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers with an IFU Addendum titled: HeartStart XL+ Instructions for Use Addendum. Philips has revised the labeling in the addendum to specify the battery charge time with device turned off and AC power connected. The revised specification is: With temperature at 25ÂżÂżC (77ÂżÂżF), less than 3 hours to 100% capacity; less than 2 hours to 80% capacity.Please contact your local Philips Representative at 1-800-722-9377. |  | 7157 | 7157 units | Worldwide Distribution - USA (nationwide) including Puerto Rico and the countries of ALGERIA ANGOLA ARGENTINA AUSTRALIA AUSTRIA AZERBAIJAN BAHRAIN BELGIUM BOLIVIA BULGARIA CANADA CHILE CHINA COLOMBIA COSTA RICA CROATIA (local name: Hrvatska) CYPRUS CZECH REPUBLIC DENMARK DOMINICAN REPUBLIC EGYPT EL SALVADOR ESTONIA ETHIOPIA FINLAND FRANCE GABON GERMANY GREECE GUATEMALA HONG KONG HUNGARY ICELAND INDIA INDONESIA IRAN (ISLAMIC REPUBLIC OF) IRAQ IRELAND ISRAEL ITALY JORDAN KENYA KOREA, REPUBLIC OF KUWAIT LATVIA LEBANON LIBERIA LIBYAN ARAB JAMAHIRIYA LITHUANIA MACAU MALAYSIA MALTA MAURITIUS MOROCCO MYANMAR (Burma) NAMIBIA NETHERLANDS NEW ZEALAND NORWAY OMAN PALESTINE PERU PHILIPPINES POLAND PORTUGAL,QATAR ROMANIA RUSSIAN FEDERATION SAUDI ARABIA SENEGAL SERBIA SEYCHELLES SINGAPORE SLOVAKIA (Slovak Republic) SLOVENIA SOUTH AFRICA SPAIN SRI LANKA SWEDEN SWITZERLAND SYRIAN ARAB REPUBLIC TAIWAN, PROVINCE OF CHINA TANZANIA, UNITED REPUBLIC OF THAILAND TRINIDAD AND TOBAGO TUNISIA TURKEY UNITED ARAB EMIRATES UNITED KINGDOM and VIET NAM. | Open, Classified |  |  | 7157 | Not_Computer
Z-1108-2014 | 67481 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | M3536A HeartStart MRx with Wireless Link Upgrade (861289) |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | Wireless Link may be unable to transmit data when configured for a specific data flow. | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter to all consignees on February 18, 2014. The Letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips is initiating a correction to affected devices. Philips will release a new version of Wireless Link software V2.0.0.3 sub1 that corrects the issue. The correction will be provided free of charge to all units affected by this issue. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the software upgrade.Customers with questons were instructed to call 1-800-722-9377.For questions regarding this recall call 800-722-9377. |  | 220 | 220 devices | Worldwide Distribution - USA including AZ, CA, FL, IL, IN, KS, KY, MA, MD, MI, MN, MO, NC, NY, OH, OK, OR, PA, SC, TN, TX,and VA and Internationally to Australia, Norway, and Spain. | Terminated | November 12, 2015 | 632 | 220 | Computer
Z-1092-2014 | 67067 | No | KWT | Shoulder joint metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | 00434811113, REF I4348-111-13 |  |  | January 10, 2014 | 2014 | Zimmer, Inc. | This single manufacturing lot has the potential for the bond strength between the Trabecular Metal material and substrate to be below the minimum required specification. Testing found bond strength approximately 80% of the minimum required by specification in one of the three zones of the Trabecular Metal. The other two zones exceeded the minimum specification. | Nonconforming Material/Component | Zimmer sent an Urgent Medical Device Recall dated January 10, 2014 to distributors via electronic mail. The hospital/surgeon that has implanted this manufacturing lot will be identified and notified by the distributor. The letter identified the affected product, problem and actions to be taken. Distributors were instructed to locate and remove the affected product in their territory, as well as identifying and notifying the hospital/surgeon that has implanted this manufacturing lot as applicable. The distributor will return on-hand affected products to Zimmer and ensure all of their products are accounted for using the return form provided in the letter. Distributors and hospital consignees are to complete and return the inventory return form and acknowledgement form indicating which hospital/surgeon has implanted the product and to which they have provided notice to. Email a copy of the completed return form to CorporateQuality.PostMarket@Zimmer.comCommunications outside of the United States will occur approximately two weeks after the United States communications.For questions call 1-877-946-2761 |  | 10 | 10 units | Distribution USA in the state of Illinois. | Terminated | September 11, 2014 | 244 | 10 | Not_Computer
Z-1094-2014 | 67375 | No | MNB |  |  | Obstetrics/Gynecology | PMA | NovaSure Impedance Controlled Endometrial Ablation System Radiofrequency Controller, Model Number: 71978-001 |  |  | January 24, 2014 | 2014 | Hologic, Inc | NovaSure Radiofrequency Control Units may not meet a requirement of the electrical standard, IEC 60601-1, | Component design/selection | Hologic sent an Important Medical Device Information letter to all affected consignees on January 24, 2014 via certified letter. The letter contains instructions for scheduling the return and replacement of the affected unit, provides contact information, and includes a postage-paid return card for the consignee to complete and return to Hologic. |  | 93 | 93 units | Worldwide Distribution - USA Nationwide and the countries Canada, France, Germany, and Sweden | Terminated | July 17, 2014 | 174 | 93 | Not_Computer
Z-1104-2014 | 67313 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Optiflux F200A High Flux Dialyzers; Product Number: 0500320A. |  |  | February 04, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Dialysate Port Leak During Priming of the Dialyzers | Under Investigation by firm | Fresenius Medical issued Notice to Clinics and Clinics with Patients Letters Dated - 1st Mailing - February 4, 2014 and Patients Letters Dated - February 5, 2014 - these letters contained the correct information.Corrected letters to Clinics and Clinics with Patients issued 2/14/14 via certified mail with return receipt and fax-back form . Users are advised the leak may be stopped by tightening the dialysate port cap more securely.Questions associated with this product notification to your Customer Service Team at 1-800-323-5188. |  | 5244 | 5,244 units | Worldwide distribution: USA (nationwide) and the country of: Mexico. | Open, Classified |  |  | 7741662 | Not_Computer
Z-1095-2014 | 67445 | No | MCX |  |  | Cardiovascular | PMA | CSI, Diamondback 360 Coronary Orbital Atherectomy System, Part Number 70058-02, Model Number DBEC-125 |  |  | January 27, 2014 | 2014 | Cardiovascular Systems, Inc. | Cardio vascular Systems Inc. has initiated a recall to request the immediate removal and return of Diamondback 360 Coronary Orbital Atherectomy Device (OAD).The device mistakenly contains a saline line that is used in our peripheral atherectomy device. This saline line was not part of the FDA approval for use in the coronary device. It therefore does not meet specification for this device. C | Process control | Cardiovascular Systems Inc. sent an "Urgent Medical Device Recall" letter dated January 27, 2014 to all affected customers. The letter described the affected product, recall description and actions to be taken. Customers were instructed to complete and return the attached Customer Acknowledgement Form and return the device i available. For further information they should contact CSI Customer Service at 651 Campus Drive, Saint Paul, MN 55112, Tele 877-274-0901, Fax 612-677-3355. |  | 318 | 318 | USA Nationwide Distribution in the states of AZ, CA, FL, MA, MN, MS, NC, NJ, NY, PA, TN, TX, VA. | Terminated | May 05, 2014 | 98 | 318 | Not_Computer
Z-1121-2014 | 67490 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Everest 20cc Inflation Device; Model Number and Catalog Number: AC2205P |  |  | February 06, 2014 | 2014 | Medtronic Vascular | Incorrect master label. | Labeling mix-ups | Medtronic sent an Urgent Medical Device Recall letter on February 6, 2014, via UPS 3-day delivery. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the affected product that remains in their inventory. Return the affected product to Medtronic: Contact Customer Service at 888-283-7868 and reference RADAR # 170276 to initiate a return and credit of product. Complete the attached Customers Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. If any affected product has been forwarded to another facility, customers were asked to notify that facility accordingly and facilitate the retrieval of the affected product. Customers with questions were instructed to contact their Medtronic representative.For questions regarding this recall call 888-283-7868. |  | 191 | 191 devices | Nationwide Distribution including TX, ME, IA, PA, OR, and NH. | Terminated | June 02, 2014 | 116 | 191 | Not_Computer
Z-1110-2014 | 67427 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator used |  |  | February 04, 2014 | 2014 | Philips Medical Systems, Inc. | HeartStart MRx Defibrillator Monitor therapy connection could experience accelerated wear causing risk for patients and/or caregivers | Device Design | Philips Medical North America issued on 2/4/14 the Urgent Medical Device Correction Notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of free of charge service support for the MRx therapy connection, including inspection of therapy connection, and installation of a stabilizing collar if not already present. If wear is detected, therapy components (including therapy pads/cables, CPR therapy cable, external paddles cable, and internal paddles cable/adapter) within their service life will be replaced.A Philips Healthcare representative will contact customers with affected devices to arrange for service. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to inspection and installation of upgrade hardware by a Philips Authorized Service Provider provided that upon receipt of this notification: Ensure you are performing ongoing therapy connection inspections on all of your MRx devices to detect wear. The HeartStart MRx Instructions for Use, 453564396411 Ed 1, 453564307761 Ed 2, or 453564174011, a previously circulated Therapy Connection Maintenance and Inspection When Used in Transport Addendum, Edition 6, describe how users can identify wear. If wear is detected, remove affected devices immediately from use and contact local Philips representative to arrange service or call 1-800-722-9377. |  | 58466 | 58466 units | Worldwide Distribution: USA (nationwide) and to countries of: AFGHANISTAN ALBANIA ARGENTINA ARUBA AUSTRALIA AUSTRIA AZERBAIJAN BAHRAIN BANGLADESH BELGIUM BOLIVIA BOSNIA AND HERZEGOWINA BRAZIL BRUNEI DARUSSALAM BULGARIA CAMBODIA CHILE CHINA COLOMBIA COSTA RICA CROATIA (local name: Hrvatska) CYPRUS CZECH REPUBLIC DENMARK ECUADOR EGYPT EL SALVADOR ESTONIA ETHIOPIA FINLAND France GERMANY GHANA GREECE GUADELOUPE HONG KONG ICELAND INDIA INDONESIA IRAQ IRELAND ISRAEL ITALY JAPAN JORDAN KAZAKHSTAN KENYA KOREA, REPUBLIC OF KUWAIT LAOSÂż LATVIA LEBANON LIBYAN ARAB JAMAHIRIYA LITHUANIA LUXEMBOURG MACEDONIA MALAYSIA MEXICO MONGOLIA MOROCCO MOZAMBIQUE MYANMAR (Burma)Âż NAMIBIA NEPAL NETHERLANDS NETHERLANDS ANTILLES NEW CALEDONIA NEW ZEALAND NICARAGUA NORWAY OMAN PAKISTAN PALESTINE PANAMA PAPUA NEW GUINEA PARAGUAY PERU PHILIPPINES POLAND PORTUGAL PUERTO RICO QATAR REUNION ROMANIA RUSSIAN FEDERATION SAUDI ARABIA SERBIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SRI LANKA SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA TANZANIA, UNITED REPUBLIC OF THAILAND TRINIDAD AND TOBAGO TUNISIA TURKEY TURKMENISTAN UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN VENEZUELA VIET NAM YEMEN | Open, Classified |  |  | 58466 | Not_Computer
Z-1120-2014 | 67438 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Defibrillator/Monitor M3538A Lithium Ion Batteries |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | Philips HeartStart MRx Monitor/Defibrillator, when operating on battery power only, may experience an unexpected shutdown if exposed to elevated levels of electromagnetic interference from RF energy source | Device Design | The firm, Philips Healthcare, sent an "URGENT- Medical Device Correction" letter dated February 2014 and issued a Field Safety Notice on February 18/ 2014 to its customers. The letter and field safety notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue. Philips will replace the affected batteries free of charge.Philips instructed customers to follow the Action to be Taken by Customer/User section of the Field Safety Notice. Fixed, portable, and mobile radio frequency communications equipment can affect the performance of medical equipment. See the HeartStart MRx Instructions for Use for the minimum recommended separation distance between RF communications equipment and the HeartStart MRx. The condition can be avoided by operating the device on AC or external DC power or by the following actions when operating on battery.If you need any further information or support concerning this issue, please contact your local Philips representative or call us at 1-800-722-9377. |  | 43792 | 43,792 batteries | Worldwide Distribution: US (nationwide) and Internationally to: ALBANIA, ARGENTINA, ARUBA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BANGLADESH, BELGUIM, BOLIVIA, BRAZIL, CANADA, CHINA, COLOMBIA, CONGO, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, EL SALVADOR, ESTONIA, ETHIOPIA, FINLAND, FRANCE, GERMANY, HONG KONG, ICELAND, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAKHSTAN, KENYA, KOREA REPUBLIC, KUWAIT, LATIVA, LEBANON, LIBYAN ARAB, LITHUANIA, MALAYSIA, MALTA, MEXICO, MONGOLIA, MOROCCO, MOZAMBIQUE, MYANMAR, NEPAL, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, ROMANIA, RUSSIAN FEDERAL, SAUDIA ARABIA, SERBIA, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN PROVIDENCE, THAILAND, TRINIDAND AND TABAGO, TUNISIA, TURKEY, TURKMENISTANIA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, UZBEKISTAN, VENEZUELA, and VIETNAM. | Open, Classified |  |  | 43792 | Not_Computer
Z-1117-2014 | 67377 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | XL CONVALESCENT RECLINER 5291/529S with Padded armrests with guard, Dual fold down side-tables, Removable side panels, Built-in head-rest. |  |  | September 06, 2013 | 2013 | Winco Mfg., LLC | A supplier manufacturing assembly defect with the right hand recline mechanism. Use of this product may result in the loss of functionality of the three (3) position lock bar when in a reclined position. | Process control | Winco sent a Recliner Field Correction Notification letter dated September 11, 2013 to all affected consignees. The letter identified the product, the problem and action to be taken by the consignees. The consignees were instructed to complete and return the enclosed field correction response form back to Winco so they can send replacement mechanism(s) along with instructions on how to exchange the mechanism(s). Customers with questions were instructed to contact Customer Care at 800-237-3377 or email QA-RA@wincomfg.com.For questions regarding this recall call 352-854-2929, ext 130. |  | 369 | 369 units (all models) | Worldwide Distribution - USA nationwide including Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Florida, Georgia, Idaho, Illinois, Indiana, Kanas, Kentucky, Louisiana, Massachusetts, Maryland, Maine, Michigan, Minnesota, Missouri, Mississippi, North Carolina, North Dakota, New Jersey, New Mexico, New York, Ohio, Oklahoma, Pennsylvania, South Carolina, South Dakota, Tennessee, Texas, Virginia, Vermont, Washington, Wisconsin, West Virginia, Puerto Rico and Internationally to the country of Canada. | Terminated | December 30, 2014 | 480 | 369 | Not_Computer
Z-1093-2014 | 67380 | No | NYQ |  |  | Immunology | PMA | Product Name: HER2 CISH pharmDx Kit |  |  | January 21, 2014 | 2014 | Dako North America Inc. | Dako has initiated a recall of HER2 CISH pharmDxTM Kit because weak red signals have been observed. | Device Design | Dako sent a Recall Notification letter dated January 21, 2014to affected customers via mail. The letter identified the affected product, description of the problem, actions to be taken by the user, communication, and Dako contact information. The letter was accompanied by a "Recall Form (Product Recall Form to customers)". For questions call 805-566-5464. |  | 49 | 49 | Worldwide Distribution - USA Nationwide and in the country of Chile. | Terminated | June 04, 2014 | 134 | 49 | Not_Computer
Z-1118-2014 | 67487 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator; Models: |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | ECG trunk cable and connector block of the MRx could be susceptible to accelerated wear, which could result in an interrupted ECG signal | Component design/selection | Philips Healthcare issued Urgent Medical Device Correction notification/Field Safety Notice on 2/18/14 informs customers of the issue and identifies what action Philips plans to take to remedy the issue.Philips will perform a hardware upgrade free of charge.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to receiving the hardware upgrade, while remaining aware that if the problem were to occur you could experience the following symptoms:Loss of demand mode pacing, nability to perform synchronized cardioversion with paddle andDisruption of ECG monitoring or INOP messages.Not every serial number within the listed range is affected by this Medical Device Correction. If you have questions about whether your MRx is affected by this issue, please contact Philips at 1-800-722-9377. |  | 49929 | 49929 units | Worldwide Distribution-USA (nationwide) and the countries of AFGHANISTAN ALBANIA ARGENTINA ARUBA AUSTRALIA AUSTRIA AZERBAIJAN BAHRAIN BANGLADESH BELGIUM BOLIVIA BOSNIA AND HERZEGOWINA BRAZIL BRUNEI DARUSSALAM BULGARIA CANADA CHILE CHINA COLOMBIA COSTA RICA CROATIA (local name: Hrvatska) CYPRUS CZECH REPUBLIC DENMARK ECUADOR EGYPT ESTONIA ETHIOPIA FINLAND FRANCE GERMANY GREECE GUADELOUPE HONG KONG ICELAND INDIA INDONESIA IRELAND ISRAEL ITALY JAPAN JORDAN KENYA KOREA, REPUBLIC OF KUWAIT LAOS LATVIA LEBANON LIBYAN ARAB JAMAHIRIYA LITHUANIA LUXEMBOURG MACAU MACEDONIA MALAYSIA MALTA MEXICO MOROCCO NAMIBIA NETHERLANDS NETHERLANDS ANTILLES NEW CALEDONIA NEW ZEALAND NICARAGUA NIGERIA NORWAY OMAN PAKISTAN PANAMA PAPUA NEW GUINEA PERU PHILIPPINES POLAND PORTUGAL PUERTO RICO QATAR REUNION ROMANIA RUSSIAN FEDERATION SAUDI ARABIA SINGAPORE SLOVENIA SOUTH AFRICA SPAIN SRI LANKA SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA TANZANIA, UNITED REPUBLIC OF THAILAND TRINIDAD AND TOBAGO TUNISIA TURKEY UGANDA UKRAINE UNITED ARAB EMIRATES UNITED KINGDOM UZBEKISTAN VENEZUELA VIET NAM YEMEN. | Open, Classified |  |  | 49929 | Not_Computer
Z-1112-2014 | 67415 | No | JSK | Supplement for culture media. | Microbiology | Microbiology | 510(K) Exempt | Oxoid Legionella BCYE Growth Supplement, For 500 ml of medium, SR0110C, packaged 10 vials/box, lyophilized. The firm name on the label is Oxoid Ltd., Basingstoke, Hants, England. |  |  | January 24, 2014 | 2014 | Remel Inc | The product may contain high levels of microbial contamination | Unknown/Undetermined by firm | Thermo Fisher sent an Urgent Medical Device Recall letter dated January 24, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. Please notify any personnel who need to be aware of the potential for false presumptive Legionella spp. identification or masking of the recovery/isolation of Legionella spp. Review your inventory for the affected product and discard it. You should review results reported using these lots and consider retesting and/or seek appropriate expert advice. Return the attached Medical Device Recall Return Response acknowledging receipt of the notice and disposal of the affected product.If you have further distributed these product lots, please contact those entities, advise them of the situation and provide them with a copy of this letter. You should insert your contact information, email, and fax numbers in the acknowledgement form and request that they return it to you.If you have any questions, please contact our Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 440 | 44/10/500-ml. vial boxes | US Distribution including Puerto Rico and the states of FL, IL, MA, ME, MI, and PA. | Terminated | July 23, 2014 | 180 | 440 | Not_Computer
Z-1096-2014 | 67424 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Invivo Corporation Expression IP5 Information Portal |  |  | October 03, 2013 | 2013 | Invivo Corporation | Invivo Corporation located at 12151 Research Parkway, Orlando, FL 32826 is recalling their Expression Information Portal (Model IP5). Three different IP5 devices lost audio during use. | Software design | Philips Healthcare sent an "URGENT - MEDICAL DEVICE RECALL" letter dated October 3, 2013 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Customers were instructed to refer to the "Action to be Taken by Customer/User section on the attached pages. A Philips representative will contact the customer regarding their affected devices. All affected devices will have a software update installed in order to correct the problem free of charge. Contact your local Philips representative for further information and support. |  | 271 | 271 | Worldwide Distribution - USA and the countries of Australia, Austria, Denmark, France, Germany, Netherlands, Poland, Singapore, Sweden, Switzerland, United Kingdom. | Open, Classified |  |  | 271 | Computer
Z-0973-2014 | 67363 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Medtronic O-arm Imaging System Mobile X-ray system. |  |  | July 25, 2013 | 2013 | Medtronic Navigation, Inc. | Review of on-system electronic product labeling and Instructions for comply with applicable Use in response to a customer inquiry uncovered gaps in compliance FDA standards with the requirements for information to be provided to users in 21CFR 1020.30(h) | Radiation Control for Health and Safety Act | Medtronic Navigation Inc. issued Notiifcation on 3/4/14 and included thee errata sheet to all affected customers updating the missing and incorrect information and ensure that labeling is now in compliance with 21 CFR Part 1020.30(h).This will be followed by a software update that contains the same information on the errata sheet, which serves to correct the electronic version of the Instructions for Use.The paper errata sheet will be sent via traceable means and the software upgrade will be performed by Medtronic Service either through remote access or by site visit of Medtronic service personnel.For further questions please call (978) 698-6000. |  | 449 | 449 in US | USA (nationwide) Distribution | Open, Classified |  |  | 449 | Computer
Z-1122-2014 | 67545 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Triton FP Infusion Pump, Part Number 400000, packaged in boxes. |  |  | February 17, 2014 | 2014 | WalkMed Infusion, LLC | WalkMed Infusion is initiaing a recall on their Triton FP Infusion Pump, Part Number 400000 due to the label on the right side of the pump containing an incorrectinstruction. | Labeling design | WalkMed Infusion sent an Important Medical Device Recall letter dated February 14, 2014, to all affected customers. The letter identified the product the problem and the action customers need to take.Our records Indicate that you have purchased product affected by this recall (see the attached list).Our sales representative will contact you to determine a plan for label replacement for your units. Replacement of the label will include a simple process of removing the old label (PN 300063), replacing it with a corrected label(PN 400033) and documenting completion of the replacement. This action can be completed by you,your customers,our sales representative, or by returning the infusion pump to WalkMed Infusion.For further questions please call (720) 351-4934 or (720) 351-4965.Product will be corrected in the field when possible and supported by the customer. Otherwise, WalkMed Infusion employees with travel to the customer and complete the correction activities. |  | 797 | 797 | US Distribution including the states of IL, MA, KS, and, NJ. | Terminated | October 06, 2014 | 231 | 797 | Not_Computer
Z-1123-2014 | 67385 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | AMO TECNIS CL Foldable Silicone IOL with OptiEdge Design, Model Z9002, diopter size: 16.0 and 21.0. |  |  | January 28, 2014 | 2014 | AMO Puerto Rico Manufacturing, Inc. | Certain Silicone TECNIS IOLs may be labeled with the incorrect diopter power. | Nonconforming Material/Component | Beginning on January 28, 2014, AMO began distribution of the Recall Notification Letter to the 72 AMO TECNIS IOL customer accounts within the United States who had received lenses from the 2 affected P/O's. 100% delivery of the recall communication to US customers was confirmed on January 31, 2014. Included with the Recall Notification Letter is a facsimile delivery confirmation which each customer is instructed to fax back to AMO. |  | 169 | 169 units | Worldwide Distribution-USA (nationwide) and the countries of Australia, Great Britain, Canada and Japan. | Terminated | April 19, 2017 | 1177 | 169 | Not_Computer
Z-1113-2014 | 66527 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-101A Aquilion 32/64 ONE Computed Tomography X-Ray System; TXS-101/R Aquilion RXL Computed Tomography X-Ray System; TSX-301A Aquilion One Computed Tomography X-Ray System; TSX-301B Aquilion Premium Computed Tomography X-Ray System; TSX-302A Aquilion Prime Computed Tomography X-Ray System. |  |  | October 01, 2013 | 2013 | Toshiba American Medical Systems Inc | Toshiba intitiated this recall because their investigation revealed that due to a problem of the X-ray output control software, the X-ray output conditions may not be set correctly. As a result, X-ray exposure under the scan conditions used for contrast study may not start. | Software design | Toshiba sent an Urgent Medical Device Correction letter on October 1, 2013, via USPS return receipt mail and accompanied by a response form. The recall notification letter provided the Device Model, serial number, SID, problem, corrective action, request to customers, and contact information (800) 421-1968 or (800) 521-1968. Toshiba designed and test an upgraded ROM for the SC2 PWB to prevent recurrence of the fault. Service technicians will replace the ROM on site at no cost to the system owners. All repairs are expected to be completed by April 2014 and progress reports will be provided on the effectiveness of the plan.For questions regarding this recall call 714-730-5000. |  | 338 | 338 units | Nationwide Distribution. | Terminated | March 03, 2015 | 518 | 338 | Computer
Z-1124-2014 | 67458 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MEDITECH Microbiology - Calculator/data processing module, for clinical use. |  |  | January 17, 2014 | 2014 | Medical Information Technology, Inc. | Potential for patient results being removed from LAB/EMR. | Device Design | Beginning on January 17, 2014 MEDITECH distributed a notification to MAGIC and Client Server customers who have the Microbiology product in TEST and/or LIVE only for the releases noted in this correction report. This notification was made via e-mailed task updates that can be printed by the customer. Task updates sent via e-mail are immediately transmitted to the customer.Instructions provided to customers: Task Description: CS LAB DTS 11696/MG LAB DTS 16703 - MIC: Inactive procedure prompt not displaying previous results. We are setting up this task to inform you of a potential patient safety issue that may result in patient results being removed from LAB/EMR. This task will continue to be updated to provide you with the status on this issue and details on the deployment of changes to address. This issue has been determined to impact both Client Server and MAGIC 5.6 and higher. Until the changes to address above are delivered, please refrain from inactivating procedure prompts in the MIC Procedure Dictionary. Although patient results are not permanently deleted from the record, because of the potential impact on a patient should this issue occur and go undetected, we felt it necessary to inform you of the issue. We strongly recommend delivery of the changesto address the issue.Âż We have engaged all necessary resources within MEDITECH to identify and resolve this issue as expediently as possible.Âż CS LAB DTS 11696 - MIC: Inactive procedure prompt not displaying previous resultsÂż MG LAB DTS 16703 - MIC: Inactive procedure prompt not displaying previous resultsÂż Inactivating Procedure Prompts in the MIC Procedure Dictionary causes data to be removed from existing specimens. These results are also removed from the EMR. However, once the prompts are re-activated, the results are again present on the specimens.Âż We strongly recommend delivery of this change to your system to alleviate this potential patient safety concern. Please confirm via this task th |  | 737 | 737 | Worldwide Distribution: US (nationwide) and to countries of: Canada, Ireland, and United Kingdom. | Open, Classified |  |  | 737 | Not_Computer
Z-1133-2014 | 67475 | No | KZP | Bracket adhesive resin and tooth conditioner. | Dental | Dental | 510(k) | BencoDental tartar &amp; stain remover ultrasonic cleaning solution |  |  | February 10, 2014 | 2014 | Heraeus Kulzer, LLC. | As the result of a complaint to Heraeus Kulzer's customer service department, Heraues Kulzer is recalling Benco Tartar and Stain Solution part #1010-451, lot #13121434 because it has incorrect labeling on the immediate container. While the outer carton correctly states "Benco Tartar and Stain", the immediate gallon container in which the solution is contained incorrectly states "Benco Temporary C | Labeling mix-ups | Heraeus Kulzer, notified their one consignee of the recall February 4, 2014 via phone and then by official electronic notification. On February 7, 2014, an Important Communication Product Recall letter was also mailed to Benco direct accounts. On 2/10/14, Heraeus Kulzer also called Benco customers notifying them of the recall. The notifications explained the issue, identified the affected product and actions to be taken. For questions call 574-293-5444 |  | 812 | 812 units (gallons) | USA Nationwide Distribution in the states AL, AR, AZ, CA, CO, FL, GA, IL, IN, KS, KY, LA, MI, MN, MO, MS, NC, ND, NE, NM, NV, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, WV | Terminated | June 18, 2014 | 128 | 812 | Not_Computer
Z-1129-2014 | 67421 | No | GEY | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Punch Hair Matic |  |  | February 18, 2013 | 2013 | MEDICAMAT S.A. | Medicamat is initiating a field correction due to misbranded marketing literature for the Punch Hair Matic SAFER. | Labeling Change Control | MEDICAMAT sent an Advisory Notice dated May 7, 2013, to the affected customer. The letter identified the product the problem and the action needed to be taken by the customer.The notification advises customers of the issue (misbranded brochure and website). The letter advises the distributor to delete the misbranded language from their communication, marketing, or other information provided to clients. They are also advised to ask their clients change their marketing if necessary and provided corrective labeling.For further questions please call +33146562404.. |  | 71 | 71 units | US Distribution including TX. | Terminated | March 10, 2014 | 385 | 71 | Not_Computer
Z-1127-2014 | 67349 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | Model Kodak 2100 Intraoral X-Ray System, Catalog Numbers 5303144, 8011579 -- Carestream Health, Inc. 150 Verona Street Rochester, NY 14608 -- Made in France -- TROPHY 77435 Marne La Vallee --- COMMON/USUAL NAME: Dental X-Ray Systems -- CLASSIFICATION NAME: Unit, X-Ray, Extraoral with timer --- There are four major components: x-ray tube, x-ray generator, collimator and image receptor. These systems are permanently installed diagnostic dental x-ray systems for generation of x-rays for examination of the teeth, jaw and oral structures. |  |  | January 24, 2014 | 2014 | Carestream Health Inc. | Carestream received problem reports relating to apparent early failure of Kodak 2X00 Intraoral X-ray device arms leading to arm breakage. | Component design/selection | Urgent Medical Device Recall Letters (dated 1/24/2014) were sent to the consignees on 1/24/2014 via Certified Mail, Return Receipt Requested. |  | 347 | Domestic: 347 units (total for both) | Worldwide distribution. USA (nationwide) and the country of Canada. | Terminated | June 18, 2015 | 510 | 347 | Not_Computer
Z-1132-2014 | 67444 | No | NPJ | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Restoris Multicompartmental Knee (MCK) System is an implant system designed to be used with MAKO's Robotic Arm Interactive Orthopedic System (RIO). It is composed of a unicompartmental implant system (Restoris MCK Uni) and a patellofemoral implant system (Restoris MCK PF). The Restoris MCK Uni is designed for use when load bearing ROM is expected to be less than or equal to 155 degrees. In Restoris MCK combination where multi-compartmental areas are being treated, the Restoris MCK components were designed with 3 mm of gap between the components to ensure that the components do not interfere. |  |  | January 09, 2014 | 2014 | Mako Surgical Corporation | The product may be mislabeled and could result in the incorrect implant being used. | Error in labeling | MAKO Surgical Corporation sent an Important Medical Device Recall Notice letter dated January 9, 2014 to all affected customers. The letter identified the product, the problem, and the actions to be taken by the customer. The letter informed the customers that their MAKOplasty Sales Specialist will coordinate with them the return of the product that will be replaced at no charge. Customers were asked to sign and return the Acknowledgement form to MAKO Surgical using one of the following methods: Email, Fax or Mail. For questions, customers were instructed to call (954) 628-1721. For questions regarding this recall call 954-628-0607. |  | 41 | 41 units | Nationwide distribution including AZ, FL, HI, OH, MI, NC NH, WA, TX, WI, OK, CA, RI, LA, ID, IL, IN, and TN. | Terminated | August 14, 2014 | 217 | 41 | Not_Computer
Z-1130-2014 | 67468 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | FMP X-alt Acetabular Liner |  |  | February 07, 2014 | 2014 | Djo Surgical | Packaging error -- two different types and sizes of acetabular liners used during hip replacement surgery were mistakenly packaged into each other's outer boxes after sterilization, resulting in the products being mislabeled. The units labeled as Model/Catalog Number 931-28-248 Lot/Serial Number 685F1034 on the outside and containing Model/Catalog Number 932-36-252, Lot# 728F1089 on the inside (1 | Packaging process control | DJO Global sent an Urgent Field Safety Notice dated February 7, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. All 10 devices with the outer label of 931-28-248, Lot#685F1034 were found in-house and were quarantined. Theagencies who have the affected product (devices labeled as932-36-252, Lot# 728F1089 containing 931-28-248, Lot#685F1034) were contacted to quarantine their devices. Of the affected devices, 2 of the lot of 10 have already been returned via product complaint.Customers were instructed to pass on the notice to all those who need to be aware within their organization or to any organization where the potentially affected devices have been transferred. Customers were instructed to place a replacement order and receive an RMA number to return all affected devices.. Customers with questions were instructed to call 512-834-6255. |  | 931282480 | 10 units of Model/Catalog Number 931-28-248 Lot/Serial Number(s) 685F1034 labeled asModel/Catalog Number 932-36-252, Lot# 728F1089 | Nationwide Distribution - USA including TX, LA, FL, CA, and AZ. | Terminated | October 01, 2014 | 236 | 931282480 | Not_Computer
Z-1125-2014 | 67498 | No | JDP | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes 5.0mm Variable Locking Screw |  |  | February 05, 2014 | 2014 | Synthes, Inc. | Labeling correction: The package insert in one 5.0mm Variable Locking Screw package was a GP0006 (Anterior Cervical Vertebrae Plate System) insert, which should have been a GP2848 (4.5mm VA-LCP Curved Condylar Plate System insert). | Employee error | Synthes sent an Urgent Field Safety Notification/Medical Device Labeling Correction letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Exchage the previously shipped package insert (GP0006) with the correct package insert (GP2848 provided with the notification).Review the package insert.Forward the Field Safety Notification to anyone in their facility that needs to be informed.If the package insert has been forwarded to another facility, contact that facility.Maintain awareness of this notice until all package inserts have been exchanged.Maintain a copy of the notice.Customers with questions were instructed to maintain a copy of the notice.For questions regarding this recall call 610-719-5000. |  | 45 | 45 | Nationwide Distribution including GA, CA, TX, MO, OH, PA, WA, AL, IN, AZ, MI, IA, TN, KY, NC, VA, NY, LA, IL, UT, and KS. | Terminated | August 31, 2015 | 572 | 45 | Not_Computer
Z-1131-2014 | 67435 | No | IOL | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | Accessory shorts for the AlterG Anti-Gravity treadmill. |  |  | January 28, 2014 | 2014 | AlterG, Incorporated | Unapproved material used by vendor in subset of shorts causing them to stretch more than usual and are uncomfortable for patient. | Component change control | The firm, AlterG, Inc, sent "Medical Device Recall Notification Anti-Gravity Short Exchange" letters dated February 6, 2014 to all customers beginning February 7, 2014 by Federal Express. The letter describes the product, problem and actions to be taken. The customers were instructed to review your inventory of shorts to identify any that are labeled ZIN1212 or ZIN0413 and return any shorts with these labeled production date using the return label provided.If you have any questions about this shorts exchange program or if there are any other ways we can be of assistance to you, please email us at support@alterg.com or call the support team at any time at 510-2705369. |  | 2221 | 2221 devices | Worldwide Distribution:US (nationwide) and to countries: Australia, Netherlands, United Kingdom, Spain, Germany, France, Denmark, Switzerland, Poland, Slovak Republic, Austria, and United Arab Emirates. | Terminated | May 13, 2014 | 105 | 2221 | Not_Computer
Z-1134-2014 | 67442 | No | HFX | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | DeRoyal ST. Circumcision Clamp, REF 32-1622, Size 1.3 cm, 1 Per Pack, Rx Only STERILE EO, distributed by DeRoyal Industries. |  |  | January 16, 2014 | 2014 | DeRoyal Industries Inc | Three lots of circumcision clamps were manufactured out of specification and as a result may not properly clamp at the surgical site completely. | Process control | DeRoyal initiated the recall by sending overnight letters to its direct accounts on 01/16/2014. If distributors inform DeRoyal of further distribution, DeRoyal will contact those consignees in the same manner. If necessary, further communications will be sent in the form of another mailing, phone calls, and emails to materials management personnel when possible. |  | 2504 | 2504 units | US Distribution to states of: NJ, AL, TX, WA, IL, OK, VA, FL, OR, and KS. | Terminated | October 13, 2015 | 635 | 2504 | Not_Computer
Z-1146-2014 | 67332 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA Multical. |  |  | January 20, 2014 | 2014 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is informing all ABX PENTRA MultiCal, N Control, and P Control customers running theCreatinine Rate Blank Method assay on the ABX PENTRA 400 and PC200 that the Target Values for Creatinine RateBlank Method have been updated. This issue is only applicable to Creatinine Rate Blank Method values reported on theABX PENTRA 400 and PC200 Chemistry Analyzers only. No other analyzers | Under Investigation by firm | HORIBA Medical initiated this recall by sending customers recall communication via FedEx week of January 20, 2014. The recall communication was a Field Safety Notice letter and Acknowledgement Form. The letter, dated 01/20/2014, titled "URGENT SAFETY NOTICE", informed customers of the recall by providing product description and codes, issue, impact, actiona/resolution, and contact information. The letter contained the following attachments: Field Safety Notice Acknowledgement Form and Updated Annexes with re-assigned values. |  | 2023 | 2,023 units | United States nationwide distribution. | Terminated | December 05, 2014 | 319 | 2023 | Not_Computer
Z-1137-2014 | 67454 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PRE1055: Prevena Dressing, 5 Pack |  |  | January 30, 2014 | 2014 | KCI USA, INC. | Prevena Incisional Management System Kit/Dressing Connector - the male connector on the tubing is difficult to connect to the female connector on the Prevena canister tubing | Nonconforming Material/Component | KCL sent an Urgent Voluntary Medical Device Correction letter dated February 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to locate, quarantine and return the affected device. For a replacement of the affected product and questions customers were instructed to contact KCL Customer Technical Support at 1-800-275-4524. A Product Reconciliation Form is attached to be completed and returned. |  | 20850 | 20,850 units total | USA Nationwide Distribution | Terminated | April 06, 2015 | 431 | 20850 | Not_Computer
Z-1144-2014 | 67436 | [6 empty] | Simplexa Flu AlB &amp; RSV Direct Assay Kits, intended for use as an aid in the differential diagnosis of influenza A, influenza B, and RSV viral infections in humans and is not intended to detect influenza C. Model Number MOL2650. |  |  | February 04, 2014 | 2014 | Focus Diagnostics Inc | Focus Diagnostics initiated the recall of the certainSimplexa Flu AlB &amp; RSV Direct assay kits due tothe potential of false positive Flu A, Flu B and RSV signals. | Under Investigation by firm | Focus Diagnostic initiated this recall on February 4, 2014 by sending recall notification letters to customers via email. The letter dated February 4, 2014, titled "URGENT: SAFETY NOTICE REMOVAL", informed customers of the recall (removal). The letter provided the following information: product description with code, reason for recall, the issue, the risk to health, recommendation, background, actions by customers, contact information, labeling, and response sheet. |  | 2175 | 2,175 kits | Worldwide Distribution: US (nationwide) and to countries of:Canada, Sweden, Belgium, Spain, Portugal, Slovenia, and Israel. | Terminated | November 01, 2014 | 270 | 2175 | Not_Computer
Z-1060-2014 | 66287 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | McKesson Anesthesia Care |  |  | March 15, 2013 | 2013 | MCKESSON TECHNOLOGIES INC. | There was an occurrence where the patient case data did not match the patient data when the case was recalled in the Anesthesia Care Record (ACR) in that it included data from another case. | Software design | McKesson sent a Clinical Alert dated March 15, 2013 to all affected customers. Additionally, phone calls were placed to each customer, followed up by an email to provide them written copies of the communication and Clinical Alert and obtaining acknowledgement that they have read and understand the issue and preventive actions to take. Customers with questions were instructed to contact McKesson Customer Support at 800-442-6767 (option 3).For questions regarding this recall call 404-338-3556. |  | 10 | 10 | Nationwide Distribution including CA, GA, IN, KY, MD, MI, OH, OK, and WA. | Terminated | May 01, 2014 | 412 | 10 | Not_Computer
Z-0972-2014 | 67227 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | TSX-301C Aquilion One (Vision Edition) Computed Tomography X-Ray System. |  |  | December 11, 2013 | 2013 | Toshiba American Medical Systems Inc | Toshiba America Medical Systems, Inc. initiated this recall due to the failure of the parts used in the X-ray high-voltage generator. In some cases, X-ray exposure cannot be performed. | Other | Toshiba America Medical Systems, Inc. initiated this recall by sending "Urgent Medical Device Correction," letter dated December 11, 2013, informed customers of the recall by providing, the device name, serial number, problem statement, corrective action statement, and contact information. A customer reply form accompanied the recall notification letter.Strategy:The majority of the letters are sent via USPS return receipt mail. The letter contains a return reply form that is to be faxed to TAMS for retention. After two weeks contact will be made with each non responding location and a copy of the letter will be faxed to that location and the fax receipt will serviceas proof of notification. Upon release of the new part it will be assigned to a TAMS employee with instructions for implementation at each site. Each incident is assigned a unique task number and each task is trackedthrough to completion. In the event that a customer refuses the installation of the corrected software/hardware the customer would be requested to sign a service case record stating that the corrected software/hardware was refused.CDRH approves the CAP subject to the following condition:Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. For further questions please call (714) 730-5000. |  | 8 | 8 units | US Distribution including the states of: MD, WV, OH, CA, MN, IN, NY and WI. | Terminated | October 16, 2014 | 309 | 8 | Computer
Z-1126-2014 | 67038 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | *** 1) List number 126970465 labeled in part: 12697-65 LifeShield LATEX-FREE HEMA; Y-TYPE BLOOD SET, Nonvented, 88 Inch with 210 Micron Filter Pump, Prepierced Injection Site and Secure Lock; Made in Costa Rica; Product Inquiries should be directed to Hospira, Inc., Lake Forest, IL 60045 USA; |  |  | October 25, 2013 | 2013 | Hospira Inc. | Several lots of the LifeShield Latex-Free HEMA Y-Type Blood Set IV Administration Set were assembled incorrectly. | Process design | The firm, Hospira, sent "URGENT MEDICAL DEVICE CORRECTION" letters dated October 25, 2013 to all direct accounts/customers.The letter described the product, problem and actions to be taken. The letters instructed customers to: discontinue use of the recalled product lots; search storage areas, supply carts, and other patient care areas for affected lots and to remove any discovered recalled products; quarantine all recalled products; complete and return the attached RESPONSE REQUIRED form via fax at 1-888-356-6430 or e-mail to: Hospira6107@stericycle.com; return affected products using the labels supplied with the letter; if you have not received a return label or reply form or require additional assistance; and, if the recalled products were further distributed, to forward the recall information to them and have them call Stericycle at 866-229-7775 (Monday - Friday, 8:00 AM - 5:00 PM EST) to receive a reply form. The letter further instructs customers to ensure that the facility's protocol(s) for administering fluids and blood products are followed and to report any backflow, inability or difficulty priming, or no flow issues to Hospira at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) or via e-mail at ProductComplaintsPP@hospira.com. For further inquiries, please contact Hospira using the information provided: Reports of adverse events or product complaints can be made by contacting Hospira Global Complaint Management at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) of via e-mail at ProductComplaintsPP@hospira.com. General medical inquiries can be made by contacting Hospira Medical Communications at 800-615-0187 (24 hours a day, 7 days a week) or via e-mail at medcom@hospira.com. This information was also posted to Hospira's Internet website. |  | 173689182 | *** 1) List Number 126970465: 21,860 units; *** 2) List Number 127200465 : 20,256 units; *** 3) List Number 046419701 : 26,900 units | Worldwide Distribution: US (nationwide) including states of: AL, AR, AZ, CA, FL, GA, IL, IN, KS, KY, LA, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OR, PA, TN, TX, VA, WI, WV and Puerto Rico; and countries of Canada, Espana, France, Ireland, and Slovenia. | Open, Classified |  |  | 173689182 | Not_Computer
Z-1081-2014 | 67217 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RT Desktop |  |  | November 25, 2013 | 2013 | Elekta, Inc. | During the treatment table, gantry, and collimator calibration procedures, it is possible to type incorrect values or invert a value. | Software design | Elekta sent an Important Field Safety Notice November 2013 to all affected customers. Corrective Action #1 - Release of Important Field Safety Notice 200 01 502 041 - "Recommended quality assurance after a calibration procedure" Released 25 November, 2013. Important Field Safety Notice is in the process of being distributed to customers advising them of the recommended Quality Assurance checks to avoid/detect the problem. The permanent solution to this problem is currently been developed. Corrective Action #2 - Permanent Solution - All users of Desktop Pro software will be informed of a mandatory upgrade to Integrity 1.2 upon its release which will prevent calibration from being inverted.For questions regarding this recall call 770-300-9725. |  | 489 | 489 | Worldwide Distribution - USA including AL, CA, CT, GA, FL, IL, IN, KS, KY, LA, MD, MA, MI, MN, MO, NJ, NY, OH, PA, TN, TX, VA, and Internationally to Argentina, Austria, Bangladesh, Belgium, Botswana, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cuba, Czech Republic, Egypt, Finland, France, Germany, Greece, Hungary, India, Indonesia, Iran, Israel, Italy, Japan, Lebanon, Mexico, Morocco, Netherland, Nigeria, Pakistan, Panama, Peru, Philippines, Portugal, Russia, Serbia, Solvenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, United Kingdom, Venezuela, Vietnam, and Yemen. | Terminated | August 13, 2015 | 626 | 489 | Computer
Z-1149-2014 | 67562 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Heparin I.V. Flush Syringe is available is pre-filled polypropylene syringes, 100 units/mL, 3 mL fill in a 12 mL syringe. |  |  | February 19, 2014 | 2014 | Medefil Incorporated | The bar code that is printed on the Master Carton label does not correspond to the product code fully displayed on the Master Carton label. The non-matching information can only be detected by using a bar code reader. The defective bar codes will be read as Product No. MIH-3335 ( Heparin I.V. Flush Syringe 100 units/mL; 5 mL fill in 12 mL Syringe) whereas this product is actually Product No. MIH-3 | Labeling Change Control | URGENT MEDICAL DEVICE CORRECTION letters dated February 19, 2014 were sent to all direct accounts. The letters included instructions to: quarantine the recalled products; and,complete and return the attached customer response form (via Fax to 630-681-9100 or via mail to Medefil Inc., 250 Windy Point Drive, Glendale Heights, IL 60139) so that the correct number of replacement labels can be sent from Medefil. |  | 1608 | 1,608 Master Cartons | USA (nationwide) including the states of AZ, CA, CO, FL, GA, IL, LA, MA, MD, MI, MN, MO, MS, NC, NJ, NY, OH, OK, PA, SC, TX, UT, and WI, and the country of Canada. | Open, Classified |  |  | 1608 | Not_Computer
Z-1084-2014 | 67184 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Agility. |  |  | November 25, 2013 | 2013 | Elekta, Inc. | During the treatment table, gantry, and collimator calibration procedures, it is possible to type incorrect values or invert a value. | Software design | Corrective Action #1 - Release of Important Field Safety Notice 200 01 502 043 - "Recommended quality assurance after a calibration procedure" Released 25 November, 2013. Corrective Action #2 - Permanent Solution - All users of Integrity 3.1 software will be informed or a mandatory upgrade to Integrity 3.2 upon its release which will prevent calibration from being inverted. Target Release for Integrity 3.2 will be available at the end of February 2014. |  | 300 | 300 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AL, AK, AZ, CA, CT, FL, GA, IL, KY, LA, MD, MA, MI, MO, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SD, TN, TX, VT, VA, WA, WV, WI, and WY, and the countries of Algeria, Argentina, Australia, Austria, Bahrain, Belarus, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Czech Republic, Cuba, Denmark, Ecuador, Estonia, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Nigeria, Norway, Pakistan , Panama, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sudan, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, and Vietnam. | Open, Classified |  |  | 889 | Computer
Z-1148-2014 | 66119 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Model Number/ Description: |  |  | March 25, 2011 | 2011 | GE Healthcare, LLC | When performing a secondary capture generated from GSI Viewer, the header information is correcton the images. However, when changing the ROI (region of interest) and zooming in, the DICOMheader does not change to reflect the new image, thus the line measurements may be inaccurate. | Software design | No consignee letter was sent out. FMI No. 25431 shows where the Field Engineer had to communicate with the hospital (Instead of them sending a letter).GE Healthcare has updated all of the affected Advantage Workstation systems affected by the GSI viewer issue. At this time, deployment of the recall is considered complete. As an effectiveness check, they have tracked all service requests for the field action to completion.These service requests record the repair of the affected units. Exceptions to completion of the service request have been recorded as well. Records of field action completion will be retained per their processes and procedures. |  | 283 | 283 (59 US, 224 OUS) | Worldwide Distribution: US (nationwide): AL, AZ, CA, CT, FL, IL, IN, IA, KY, MD, MA, MI, MS, NV, NH, NJ, NY, PA, WA, and WI; and countries of: AUSTRALIA, BELGIUM, BRAZIL, CANADA, CHINA, DENMARK, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN, KUWAIT, LATVIA, NEW ZEALAND, POLAND, SAUDI ARABIA, SINGAPORE, SOUTH AFRICA, SOUTH KOREA, SWEDEN,, SWITZERLAND, TAIWAN, and UNITED KINGDOM. | Terminated | March 13, 2014 | 1084 | 283 | Not_Computer
Z-1119-2014 | 67392 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Greatbatch Medical, Offset Cup Impactor. Part Numbers: P9485, T10243, T10281, T10287, T10484, T10491, T10507, T10600, T10604, T10753, T10861, T11209, T11340, T12121, T12230, T12288, T12305, T12391, T12699, T12767, T12861, T13344, T13480, T13642, T13722, T13999, T14155, T14384, T15311, T15752, T15822, T15949, T16117, T16427, T16604, T16611, T16661, T16829, T16934, T17062, T17203, T17238, T17321, T6318, T7821, T8042, T8043, T8044, T8087, T8088, T8093, T8095, T8177, T8184, T8277, T8310, T8333, T8468, T8487, T9109, T9196, T9205, T9316, T9348, T9360, T9556, T9557, T9747, T9894, T9954, T9955, T9999. |  |  | January 31, 2014 | 2014 | Greatbatch Medical | Greatbatch Medical has initiated a global Field Corrective Action for the Standard Offset Cup Impactor since the product was first released in 2004.Greatbatch has identified through internal Sterility Assurance Level validation testing that a sterility assurance level of 10-6 is not achieved on the Standard Offset Cup Impactor when it is subjected to the steam sterilization cycles and associated dry times that are specified in the current Instructions For Use (IFU) provided with the device. The Standard Offset Cup Impactor is provided non-sterile and must be sterilized prior to use in a surgery. Greatbatch has not received any reports of deaths, illnesses, injuries, or other adverse effects associated with this issue. | Process design | Greatbatch sent an Urgent Medical Device Field Corrective Action letter dated January 31, 2014, to all affected consignees. The letter was addressed to "Dear Valued Customer". The letter described the Reason for Field Corrective Action, Risk to Health, and Instructions to Customers. The instructions to customers included to review the enclosed Field Corrective Action Response Form, to examine their inventory and identify all Standard Offset Cup Impactors that are in their possession and immediately inform those responsible for sterilization to utilize the updated instructions. Also, advised consignees to immediately notify all customers of the Field Corrective Action and to complete and return the Field Corrective Action Response Form. For questions consignees were instructed to contact their local Greatbatch Medical representative at 1-763-951-8235 or e-mail them at FieldActionCenter@Greatbatch.com.For questions regarding this recall call 763-951-8235.UPDATED 3-28-2014: An updated Urgent Medical Device Correction letter was sent out to consignees 3/28/2014 to tell them of an additional validated sterilization parameter in addition to additional devices that are affected by the recall. A notice was asked to be posted where sterilization of the Standard Offset Cup Impactors occurs so personnel responsible for the sterilization of the devices would see the 2 validated cycles. A response form was asked to be returned to Greatbatch. |  | 4870 | 4870 worldwide (2482 US), EXPANDED: 7245 devices | Worldwide Distribution - USA including CA, IN, KS, MA, NJ, TN,TX,UT and Internationally to Australia, Austria, Belgium, China, France, Germany, Greece, Ireland, Italy, Japan, Luxembourg, Netherlands, Norway, South Korea, Spain, Sweden, Switzerland, Taiwan, United Kingdom. | Terminated | October 31, 2016 | 1004 | 4870 | Not_Computer
Z-1147-2014 | 67430 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Solar 8000M and Solar 8000i Patient Monitor software version 5.4, 5.5, or 5.6 with the Automatic-View-On- Alarm (AVOA) feature enabled. |  |  | February 13, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with the Solar 8000M/i Patient Monitor Automatic-View-On-Alarm (AVOA) feature.The Solar 8000M/i AVOA feature will stop functioning if the patient monitor is discharged from the CIC Central Station while a remote automatic bedside alarm view is currently being displayed on the Solar. Once this issue occurs, subsequent remote AV | Software design | GE Healthcare sent a "Urgent Medical Device Correction" letter dated February 12, 2014. The letter was sent to all affected customers. It identified the product the problem and the action needed to be taken by the customers. This issue affects all Solar 8000M/i Patient Monitors running software version 5.4, 5.5, or 5.6 with the Automatic-View-On-Alarm (AVOA) feature enabled.GE Healthcare will provide a correction at no charge once it is available. We will contact you to arrange for this correction.Please contact Technical Support at 1-800-558-7044 or your local Service Representative. |  | 10211 | 10,211 (7,549 units US, 2,662 units OUS) | Worldwide Distribution - USA (nationwide) including the states of AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, ID, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, NS, NO, NE, NH, NH, NM, NY, NC, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, WV, WI and WY., and the countries of : AUSTRALIA, CANADA, AUSTRIA, BANGLADESH, BELGIUM, BOSNIA AND HERZEGOVINA, BRAZIL, CHILE, CHINA, COSTA RICA, CYPRUS, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, INDIA, INDONESIA, JAPAN, MALAYSIA, MEXICO, NAMBIA, NETHERLANDS, NORWAY, POLAND, RUSSIA, SAUDI ARABIA, SINGAPORE, SOUTH KOREA, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY, UNITED KINGDOM, SERBIA AND YUGOSLAVIA. | Terminated | July 10, 2015 | 512 | 10211 | Computer
Z-1143-2014 | 67550 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Trilogy Ventilators |  |  | February 11, 2014 | 2014 | Phillips Respironics, Inc. | Internal testing revealed a potentially defective component on the power management board that could affect the function of certain Trilogy Ventilators resulting in loss of therapy to patients. | Component design/selection | Philips sent an Urgent - Medical Device Recall letter on February 11, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the information contained in the Notice with all members of their staff who need to be aware of the contents of the communication. For further information or support concerning the issue, customers were instructed to contact Philips Respironics at 877-387-3311.For questions regarding this recall call 724-733-0200. |  | 572 | 572 | Worldwide Distribution - USA (nationwide) and the countries of France, Great Britian, China, India, Italy, Korea, Kuwait, Nehterlands, and Singapore. | Terminated | April 17, 2015 | 430 | 572 | Not_Computer
Z-1082-2014 | 67223 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Desktop Pro TM linear accelerator control software intended to assist a licensed practitioner in the delivery of radiation to defined target volumes. |  |  | November 25, 2013 | 2013 | Elekta, Inc. | During the treatment table, gantry, and collimator calibration procedures, it is possible to type incorrect values or invert a value. | Software design | Corrective Action #1 - Release of Important Field Safety Notice 200 01 502 040 - "Recommended quality assurance after a calibration procedure" Released 25 November, 2013. All software versions (r7.01) detailed in this notice are affected by this recall. 200 01 501 041 - Recommended quality assurance after a calibration procedure Released November 25, 2013. Only software versions R6.0 and R6.1 are affected by this recall. Important Field Safety Notices are in the process of being distributed to customers advising them of the recommended Quality Assurance checks to avoid/detect the problem. The permanent solution to this problem is currently been developed. Corrective Action #2 - Permanent Solution - All users of Desktop Pro software will be informed of a mandatory upgrade to Integrity 1.2 upon its release which will prevent calibration from being inverted. |  | 1309 | 1309 | Worldwide Distribution: US (nationwide): AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IA, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, including Puerto Rico, DC; and countries of: Albania, Argentina, Australia, Austria, Bangladesh, Belarus, Belgium, Bosnia &amp; Herzego, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Gabon, Germany, Ghana, Greece, Hong Kong, Hungary, India, Ireland, Israel, Italy, Japan, Jordan, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Singapore Slovakia, South Africa, South Korea, Spain, Sudan, Suriname, Sweden, Switzerland, Syria, Taiwan, Thailand, Turkey, United Kingdom, Ukraine, Venezuela, and Vietnam. | Open, Classified |  |  | 1309 | Computer
Z-1151-2014 | 67641 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Hand Control 12FR Suction Coagulator with 10 ft cord and holster, Sterile, Disposable, Model Number 0043-25, packaged TyvekÂż polyester peel pouch, 25 pouches per box. |  |  | February 27, 2014 | 2014 | Megadyne Medical Products, Inc. | Megadyne Medical Products, Inc. is voluntarily recalling Hand Control 12Fr Suction Coagulator with 10ft cord and 132405 Holster, Sterile, Disposable, Lot 132405 due to the potential for a small number of these suction coagulators to self-activate immediately on connection to the electrosurgical generatorand to remain activated until it is disconnected. | Use error | Megadyne Medical Product sent an Urgent Medical Device Recall letter dated February 27, 2014, to all affected customers with follow-up phone calls and emails. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed toexamine their inventory immediately to determine if they have any of the affected product. If so, customers were instructed to remove the product from service and contact Megadyne Customer Service at 1-801-576-9669 or 1-800-747-6110 for return authorization and replacement or credit. If customers have distributed any the affected product, they were instructed to immediately contact their distribution points to have the product returned to Megadyne. Customers with questions were instructed to contact Customer Service at 1-800-747-6110 or 1-801-576-9669 or email Regulatory @Megadyne.com.For questions regarding this recall call 801-576-9669. |  | 494 | 494 | USA including FL, MN, LA, NJ, NM, PA, TX and Internationally to Canada. | Terminated | July 09, 2014 | 132 | 494 | Not_Computer
Z-1153-2014 | 67440 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Treatment packs used on the VisuMax Laser Keratome. |  |  | February 03, 2014 | 2014 | Carl Zeiss Meditec AG | Carl Zeiss Treatment pack, lot number M 130010, size M, may actually contain size S. | Mixed-up of materials/components | The firm, Zeiss, sent an "Important Field Safety Advisory Note" letter dated February 3, 2014 to its customer. Letter indicates that a phone call preceded the letter on January 23, 2014. The letter also describes the product, problem and actions to be taken. The customers were instructed to quarantine all treatment packs size "M", batch number M 130010; return the material to Carl Zeiss Meditec, Inc., 5160 Hacienda Dr, Dublin, CA 94568 and Zeiss will issue a credit for all material received, upon receipt of the returned materia, and to complete and return the Conformation Form with the envelop provided.If you have any questions regarding this field corrective action, please call 925-557-4832 or email at paul.shahan@zeiss.com. |  | 11 | 1 package containing 10 individual treatment packs | US Distribution: NV only. | Terminated | March 07, 2014 | 32 | 11 | Not_Computer
Z-1065-2014 | 65451 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | All Elekta digital accelerators with electrons and beam modulator, MLCi/MLCi2, Agility or Asymmetric Heads |  |  | May 28, 2013 | 2013 | Elekta, Inc. | Customers are resetting the default values outside recommended factory settings. | Device Design | Elekta sent an Important Field Safety Notice dated May 2013 to all affected customers. The Notice identified the affected product, problem, clinical impact and actions to be taken. Customers were instructed to adhere to the instructions and advice given in manuals and system dialogues. For question contact your local Elekta office. |  | 472 | 472 units | USA Nationwide Distribution in the states of: AL, AK, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VT, VA, WA, WV, WI, WY, DC and including Puerto Rico. | Open, Classified |  |  | 472 | Not_Computer
Z-1150-2014 | 67353 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Mobile X-ray System, Model: MUX-100, MUX-100H |  |  | November 15, 2013 | 2013 | Shimadzu Medical Systems | Shimadzu Corporation is recalling the Shimadzu Mobile X-ray Sytems (MUX-100 and MUX-100H) because the connecting metal plate which connects the main chassis and the column of MUX-100 may crack. | Nonconforming Material/Component | A Field Safety Notice and an Urgent Voluntary Medical Device Recall Notice were sent to customers to inform them that SHIMADZU CORPORATION has identified a potential issue in certain manufacturing lots of the Shimadzu Mobile X-ray Systems. Customers are informed that as a result of their investigation, it has been learned that the connecting work of very few column used in specific lots for the MUX series was not done correctly, and it may cause a crack. Therefore, if the machine is continuously used in that condition, there is the possibility that the column will start to lean. The notices informs the customers of the actions to be taken. Customers are instructed to contact Shimadzu Medical Systems National Technical Support at (800) 228-1429. |  | 56 | 56 units | USA Nationwide Distribution | Terminated | June 25, 2015 | 587 | 56 | Not_Computer
Z-1152-2014 | 67418 | No | NOT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(k) | HER2 Image Analysis |  |  | February 07, 2014 | 2014 | Leica Biosystems Imaging, Inc. | In an abundance of caution, Aperio (now known as Lecia Biosystems) is providing additional information to clarify that the use of Digital IHC Breast/ER/Dako Image Analysis, Digital IHC Breast/HER2 Digital Read, Digital IHC Breast/HER2/Dako Image Analysis, Digital IHC Breast/PR Digital Read and Digital IHC Breast/PR/Dako Image Analysis with the RUO products identified in that earlier notification i | Labeling False and Misleading | Aperio (now known as Leica Biosystems) sent an recall notification letter on February 7, 2014, via USPS mail. The letter dated January 31, 2014, informed customers of the reason for the notification, product description, provided additional information to clarify that the use of Digital IHC Breast/ER/Dako Image Analysis, Digital IHC Breast/HER2 Digital Read, Digital IHC Breast/HER2/Dako Image Analysis, Digital IHC Breast/PR Digital Read and Digital IHC Breast/PR/Dako Image Analysis with the RUO products identified in that earlier notification is also for Research Use Only and not for use in diagnostic procedures. The notice also included updated manuals for the image analysis algorthithms so that the intended use in the hardware and software manuals is consistent. An acknowledgement form and contact information was also provided.We sincerely apologize for any inconvenience this may have caused. If you have any questions, please do not hesitate to contact Aperio Technical Support at (866) 478-3999. |  | 78 | 78 | US Distribution including the states of FL, WA, PA, MS, TX, MN CA, NY, NJ, NC, CO and OH. | Terminated | April 15, 2014 | 67 | 78 | Computer
Z-1155-2014 | 67473 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | MRx may display a Red X in the Ready For Use (RFU) indicator during automated tests informing the user that a failure has been detected that may prevent the delivery of a shock and pacing | Component design/selection | Philips Healthcare issued and Urgent Medical Device Correction notification/Field Safety Notice on 02/18/14 to inform customers of the issue, Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:You can continue to use your MRx prior to receiving the hardware upgrade, provided that if you observe that the MRx is displaying a Red X in the Ready For Use (RFU) indicator\during automated tests, a failure has been detected that may prevent the delivery of a shock or pacing due to this issue. If this is observed you should turn the Therapy Knob to Monitor. An INOP message describing the failure is displayed. If needed, run an Operational Check for further information. If the condition persists, take the device out of use and call for service.Identify a readily available backup defibrillator to use in the event the affected HeartStart MRx fails to deliver therapy.Philips will perform a hardware upgrade (replace therapy boards) free of charge to be installed by a Philips Field Service Engineer. Contact your local Philips representative or call us at 1-800-722-9377. |  | 3541 | 3541 units | NationwideForeign:AUSTRIABAHRAINBELGIUMBRAZILCANADACHINAEGYPTESTONIAFINLANDFRANCEGERMANYGHANAICELANDINDIAINDONESIAIRELANDISRAELITALYJAPANJORDANKENYAKOREA, REPUBLIC OFLATVIALEBANONLIBYAN ARAB JAMAHIRIYALITHUANIAMALAYSIAMEXICOMOZAMBIQUEMYANMAR (Burma)NETHERLANDSNETHERLANDS ANTILLESNEW CALEDONIANEW ZEALANDNORWAYPAKISTANPERUPOLANDPORTUGALQATARRUSSIAN FEDERATIONSAUDI ARABIASINGAPORESOUTH AFRICASPAINSWEDENSWITZERLANDTAIWAN, PROVINCE OF CHINATHAILANDTURKMENISTANUNITED ARAB EMIRATESUNITED KINGDOMVIET NAM | Open, Classified |  |  | 3541 | Not_Computer
Z-1156-2014 | 67495 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panaroma Central Station |  |  | January 08, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | Mindray has identified two issues with the Panorama Central Station. Full Disclosure data may lag behind the Panorama system time clock when the Panorama CentralStation is in use with MindrayÂżs V-Series Monitor. Additionally, event indicators lag behind the Panorama system time clock when the Panorama Central Station is in use with a Panorama Telepack. | Software design | Mindray DS USA Inc. sent an Urgent Corrective Action letter dated January 8, 2014 via certified mail to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to contact Mindray Service Representative, at 1-800-288-2121 (Monday  Friday 8:30 am-5:30 pm), to arrange for the software update of the Panorama System(s) in their facility. This work will be performed at no cost. |  | 732 | 732 units in the US | Worldwide Distribution - USA Nationwide Distribution | Terminated | January 21, 2015 | 378 | 732 | Computer
Z-1154-2014 | 67470 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | RetCam 3, RetCam Portable and RetCam Shuttle; |  |  | January 24, 2014 | 2014 | Clarity Medical Systems Inc | Software anomaly for RetCam 3, RetCam Shuttle and RetCam Portable with software versions 6.0.x, 6.1. x and 6.2x where the patients age is calculated incorrectly. | Software Manufacturing/Software Deployment | Urgent Safety Alert letters were sent to Risk Manager and to NICU US consignees via FedEx on 1/24/2014.Template of letters to OUS distributors to be sent on distributors letterhead and sent to device recipients via electronic communication and./or hard copy on 1/24/2014. |  | 480 | 166 Ret Cam Shuttles, 288 RetCam 3 and 23 RetCam Portables. | Worldwide Distribution-USA (nationwide) including DC and the states of NM, LA, CA, MI, PR, GA, MD, MT, AL, MA, NY, IL, OH, TX, VA, MI, FL, MO, KY, TN, CT, NJ, PA, UT, SD, OK, MS, and NC and the countries of Algeria, Australia, Austria, Bangladesh, Belarus, The Netherlands, Brazil, Bulgaria, Chile, China, Columbia, Czech Republic, Denmark, Ecuador, Egypt, Finland, France,India, Indonesia, Iran, United Kingdom, Israel, Italy, Kuwait, Lebanon, Malaysia, Mexico, Sweden, Pakistan, Peru, Philippines, Poland, Portugal, Saudi Arabia, Taiwan, Thailand, Turkey, UAE, Argentina, Uzbekistan, Oman, Canada, Germany, and Switzerland. | Terminated | August 11, 2014 | 199 | 480 | Computer
Z-1166-2014 | 67308 | No | JOJ | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | VERIFYÂż SixCess 270FP Challenge Pack, STERIS Corporation, 5960 Heisley Road, Mentor, OH 44060, 440-354-2600. |  |  | December 19, 2013 | 2013 | Steris Corporation | STERIS has identified that the outer shipping package for one lot (Lot #024552A) of VERIFYÂż SixCess Challenge Packs incorrectly identifies the lot as Lot #024522A. | Labeling False and Misleading | The firm, Steris, sent an "URGENT VOLUNTARY RECALL NOTICE" dated December 20, 2013 to their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to verify if the outer packaging of your on-hand inventory contains lot number 024522A; destroy any remaining inventory of the affected product in your possession and contact STERIS Customer Service at 1-800-548-4873 to obtain no cost replacement product and if you do not have the affected product in your inventory, no further action is required.If you have any questions regarding this matter, please contact your STERIS Account Manager or STERIS Customer Service at 1-800-548-4873. |  | 1880 | 94 boxes/20 units | Worldwide Distribution: US (nationwide) including states of: IL, LA, ME, MI, NC, OH, OK, VA, WI &amp; WV; and to the country of: Canada.Product was also shipped to Canada. | Terminated | March 26, 2014 | 97 | 1880 | Not_Computer
Z-1164-2014 | 67658 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Boston Scientific, RunWay Guide Catheter, UPN H749389696640, Catalog Number 38969-664. |  |  | March 03, 2014 | 2014 | Boston Scientific Corporation | The sterile pouch of the devices from this lot may not be completely sealed and sterility may be compromised. The most common adverse health consequence expected from an unsealed pouch is a transient bacteremia, which may require antibiotic therapy. | Packaging | Consignees were sent a Boston Scientific "Urgent Medical Device Recall Removal " letter dated March 3, 2014. The letter described the problem and the product involved in the recall. The letter advised consignees to return the affected product and to complete and return the Account Reply Verification Tracking Form. The local Sales Representative can answer any questions that you may have regarding this Recall Removal. |  | 7 | 7 | TX and Russia | Terminated | June 09, 2014 | 98 | 7 | Not_Computer
Z-1163-2014 | 67257 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Innova 2100IQ, Innova 3100, Innova 3100IQ, Innova 4100, Innova 4100IQ Innova 2121IQ, Innova 3131IQ; Cardiovascular X-ray imaging systems (X-ray generator installed with Rotor V3 version and Heater V4 version boards). |  |  | December 11, 2013 | 2013 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue involving the innova systems. The Innova system can unexpectedly stop delivering X-rays after a power-on or a reset cycle which may result in the loss of real-time interventional imaging. No injuries have been reported due to this issue. | Software design | GE Healthcare sent an "Urgent Medical Device Correction " letter dated December 10, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter described the Safety Issue, The long term solution is to provide a software correction for all of the affected systems that were distributed. The plan for public warning is to contact all customers with an affected system by means of an Urgent Medical Device Correction letter within the US. In the US this will be sent by a traceable means (either US Postal Service with receipt acknowledgement (registered mail) or equivalent courier service such as FEDEX). This letter will be sent to the following titles within the affected accounts:GE Healthcare will correct all affected systems at no cost to you. A GE Healthcare service representative will contact you to arrange for this correction. If you have any questions or concerns regarding this notification, please call one of the following phone numbers: United States: 800 437 1171For other countries, please contact your local GE Healthcare Service Representative. |  | 98 | 257(159 US, 98 OUS) | Worldwide Distribution - USA (nationwide) including the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, ME, MD, MA, MI, MS, MO, NV, NJ, NY, NC, ND, OH, OK OR, PA, PR, RI, SC, TN, TX, UT, VA,WA,WY and WI., and the countries of ARGENTINA, AUSTRALIA, CANADA, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GERMANY, HUNGARY, INDIA, IRELAND, ITALY, JAPAN, JORDAN, REPUBLIC OF KOREA, MALAYSIA, MEXICO, NEW ZEALAND, NORWAY, PHILIPPINES, POLAND, RUSSIA, SAUDI ARABIA, SPAIN, SRI LANKA, TURKEY and UNITED KINGDOM. | Terminated | July 29, 2014 | 230 | 98 | Computer
Z-1161-2014 | 67477 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Smith &amp; Nephew RENASYS-G Sterile, Gauze Dressing Kit with Soft Port REF #'s: 66800961 (The RENASYS Soft Port Dressing Kits are used in conjunction with negative pressure wound therapy pumps) |  |  | January 09, 2014 | 2014 | Smith &amp; Nephew Inc. | Reports of wound fluid and /or blood that were not evacuated from beneath the wound dressing and the pump blockage alarm did not activate has been associated with inadequate Soft Port sealing at the wound site or improper positioning of the Soft Port orifice over the wound contact dressing drape. Inadequate sealing can result in air leaks and contribute to inefficient fluid evacuation from the wo | Packaging change control | A product bulletin titled "Soft Port Enhancements" dated March 14, 2013, was released for external distribution describing the changes to the Soft Port products, e.g. increasing the Soft Port orifices size and adding a second slit to the release paper covering the port adhesive, and stating the purpose of the modifications to the device. These product bulletin was provided to sales representatives for distribution to customers requesting information regarding the design enhancement made to the Soft Port product range. The bulletin was not distributed to all the customers.To ensure that the product modification and its purpose was appropriately communicated to consignees of affected product a Medical Device Correction Notice, dated 01/22/2014, retrospectively advising customers of the modification to the Soft Port products has been prepared and was sent to consignees of the affected products on the same date. |  | 7353 | 7,353 individuals units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, KY, LA, MA, MS, MT, NC, NJ, NV, NY, OH, OR, PA, TN, TX, VA, WA, WI, WV, HI and the countries Dubai, Curacao, Argentina, Australia, Chile, Germany, France, England, Hong Kong, Kuwait, Lebanon, Libya, Sri Lanka, Malaysia, New Zealand, Manila, Saudi Arabia, South Africa, Singapore, Turkey, Taipei and Vietnam. | Terminated | September 26, 2014 | 260 | 2423739 | Not_Computer
Z-1165-2014 | 67515 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | MAQUET PLEGIOX Cardioplegia Heat Exchanger |  |  | January 31, 2014 | 2014 | Maquet Cardiovascular Us Sales, Llc | The MAQUET post-marketing vigilance program has identified a discrepancy that thedesired cardioplegia temperature may not be obtained when the PLEGIOX CardioplegiaHeat Exchanger (PLEGIOX) is used in conjunction with a heater/cooler unit. | Device Design | Maquet Inc.sent an" Field Safety Notice" dated January 31, 2014, to all affected customers via Fed Ex. The letter identified the product the problem and the action needed to be taken by the customer.Any questions or additional information, can be obtained from the MAQUET representative, or MAQUET Customer Service at 1-888-627-8383 (press option 2, followed by option 3), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST.Please complete and return the enclosed confirmation form, acknowledging your receipt and understanding of this communication, by emailing a scanned copy to Fieldactions@maquet.com or via fax at (973) 807-9290. |  | 1424 | 1,424 units | Worldwide Distribution - USA (nationwide) and the countries of Brazil, Germany and Turkey | Terminated | April 24, 2015 | 448 | 1424 | Not_Computer
Z-1172-2014 | 67508 | No | JOZ | Automated platelet aggregation system. | Hematology | Hematology | 510(k) | VerifyNow IIb/IIIa Test, Catalog No. 85310, 10-Test Kit and Catalog No. 85011, 25-Test kit. |  |  | February 13, 2014 | 2014 | Accumetrics Inc | Accumetrics is recalling the VerifyNow IIb/IIIa test because it may result in the reporting of an erroneous low platelet aggregation unit (PAU) result. An erroneous low PAU result may cause a low percent inhibition calculation or low baseline PAU. | Nonconforming Material/Component | Accumetrics sent an Urgent Medical Device Recall letter dated February 13, 2014 to all affected customers. The recall letter informs the customers of the problems identified and the actions to be taken. Cusotomers are instructed to complete the attached Customer Account Tracking Form and return it within 10 days by fax, email or mail. Customers with questions are instructed call Accumetrics Customer Support at (800) 643-1640 option 2 or email at support@accumetrics.com. |  | 576 | 576 units total (524 units in US) | USA Nationwide Distribution | Terminated | August 23, 2014 | 191 | 576 | Not_Computer
Z-1173-2014 | 67528 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Device Name: ABX PENTRA ML |  |  | February 14, 2014 | 2014 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical initiated this recall of all software versions of ABX PENTRA ML (Model A11A06152) due to a software bug with the potential of transmitting incorrect results to the Laboratory Information System. | Software design | Horiba sent an "URGENT SAFETY NOTICE letter datedFebruary 13, 2014 via by mail to customers. The letter identified the affected product, issue, impact, actions to be taken. Customers were instructed to share this information with laboratory staff and retain this notification as part of their laboratorydocumentation. Customers were requested to complete and return the enclosed response form via one of the following methods:Fax: 949-453-0600Email: jim.knowles@horiba.comMail:HORIBA MedicalAttention: Jim Knowles, Recall Coordinator34 Bunsen DriveIrvine, CA 92618.For questions contact your local Horiba Medical representative. |  | 15 | 15 | USA Nationwide Distribution | Terminated | August 22, 2014 | 189 | 15 | Computer
Z-1175-2014 | 67157 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare CARESCAPE Monitor B650 and CARESCAPE Monitor B850. |  |  | February 28, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of potential safety issues associated with the CARESCAPE Monitor B850 and CARESCAPE Monitor B650. These issues include both NIBP (Non-Invasive Blood Pressure) and Bed to Bed Issues. | Software design | GE Healthcare sent an Urgent Field Safety Notice dated February 26, 2014, to all International consignees. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issues and Instructions, Affected Product Details Product Correction and Contact Information ( Contact Technical Support at 1-800-558-7044 or their local Service Representative).GE Healthcare determined that there were no US sites affected since all units were under GE control therefore, no customer letter was deployed in US.For questions regarding this recall call 847-277-4472. |  | 676 | 676 units OUS | Worldwide Distribution - USA (nationwide) and Internationally to AUSTRIA, BELGIUM, CZECH REPUBLIC, FINLAND, FRANCE GERMANY, GREECE, ITALY, LITHUANIA, POLAND, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM. | Terminated | April 29, 2015 | 425 | 676 | Not_Computer
Z-1176-2014 | 67566 | No | MTD | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | PENTARAY NAV High-Density Mapping Catheter and PENTARAY NAV ECO Catheter, Catalog No. D128201, D128202, D128204, D128205, D128207, D128208, D128210, D128211. |  |  | February 20, 2014 | 2014 | Biosense Webster, Inc. | Biosense Webster is recalling the PentaRay Nav Catheter because it has the potential to have partial tip separation from the catheter shaft.Biosense Webster is expanding the recall on 5/1/14 to include all lots of the PentaRay Nav Catheter and the PentaRay Nav ECO Catheter because they can lead to an insufficient bond of the distal tip to the catheter shaft. | Process design | The firm, Biosense Webster, sent an "URGENT FIELD SAFETY NOTICE MEDICAL DEVICE-VOLUNTARY FIELD RECALL (REMOVAL)" letter dated February 18, 2014 to all customers who purchased the PentaRay Nav Catheter. The letter describes the product, the problems identified and the actions to be taken. The customers were instructed to read the "Description of the Problem" section carefully; immediately identify and set aside all affected lots; maintain a copy of this letter with the affected PentaRay NAV Catheters until all units are returned to Biosense Webster; complete and return the Voluntary Field Removal Certification form via mail to: Biosense Webster, Inc., Attn: Field Action Coordinator, 15715 Arrow Highway, Irwindale, CA 91706, Fax to: 909-839-8514 or email to: rkazmi@its.jnj.com; arrange for return of all affected units you may have in your inventory; pass on this notice to anyone in your facility that needs to be informed, and if any of the affected product have been forwarded to another facility, contact that facility and arrange for the return.Customers with questions related to the information provided in the customer notification, are instructed to contact their Biosense Webster sales representative or call (866)473-7823, Monday through Friday from 7am to 8pm EST. Customers with questions related to the Acknowledgement form and its return, are instructed to contact the Field Action Coordinator at (909) 839-8468.Biosense Webster expanded the recall on 5/1/14 by sending a customer notification letter to all customers on 5/2/14 to include all product lots of the Pentaray Nav Catheter (Catalog No: D128201, D128202, D128204,D128205, D128207, D128208, D128210, D128211). The letter provides important information about the affected products and instructions on how the customers can return the affected products to Biosense Webster. |  | 15672 | 15,672 units total (12,804 units in US) | US Nationwide Distribution. | Terminated | October 24, 2014 | 246 | 15672 | Not_Computer
Z-1174-2014 | 67391 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes Small Notch Titanium Reconstructive Plate |  |  | January 27, 2014 | 2014 | Synthes, Inc. | Certain modules containing the Synthes Small Notch Titanium Reconstructive Plate set was distributed by Spine sales consultants although it was contraindicated for certain spine applications. | Labeling design | Synthes sent an Urgent Medical Medical Device Recall letter, dated January 27, 2013, to Users and Sales Consultants. The letter identified the affected product, reason for recall, potential hazard, and actions to be taken. Customers were instructed to complete the verification section. For questions call 610-719-5450 or contact your Synthes Sales Consultant. . |  | 18747 | 18747 | USA Nationwide Distribution | Open, Classified |  |  | 18747 | Not_Computer
Z-1170-2014 | 67183 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Contra-Angle Torque Driver Kit For Certain Internal Connection |  |  | July 22, 2013 | 2013 | Biomet 3i, LLC | During recent revised sterilization validation testing, the included surgical trays did not meet the Sterility Assurance Level required. | Unknown/Undetermined by firm | Biomet 3i sent an Urgent Medical Device Field Correction Notice dated July 2013 to all affected customer via email, facsimile, or postal mail. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to complete the attached response form and fax to 561-514-6316. For questions contact Biomet 3i customer service at 1-800-342-5454. |  | 250 | 250 | Worldwide Distribution - USA Nationwide in the states of AL, AZ, CA, CO, CT, DC, DE, FL, GA, IL, KS, KY, LA, MA, MD, MI, MN, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV, including Hawaii and Puerto Rico and the countries of : Argentina, Austria, Canada, China, Costa Rica, Germany, Dominican Republic, Spain, France, Great Britain, Hong Kong, Hungary, Israel, Japan, Korea, Mexico, Poland, Portugal, Paraguay, Sweden, Singapore, Taiwan, and Uruguay. | Terminated | April 21, 2015 | 638 | 1163 | Not_Computer
Z-1194-2014 | 67459 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart MRx Monitor/Defibrillator |  |  | February 18, 2014 | 2014 | Philips Medical Systems, Inc. | When the HeartStart MRx.is used with the Q-CPR Meter in defibrillation mode, the Q-CPR Meter may incorrectly display the Do Not Touch the Patient icon. | Software Design Change | The firm, Philips Healthcare, sent an "URGENT -Medical Device Correction/Field Safety Notice" letter dated February 2014 on February 18, 2014 to its domestic and foreign customers. The letter describes the product, problem and actions to be taken. The customers were instructed to follow the Action to be Taken by Customer/User section of the attached Field Safety Notice; to continue to use your HeartStart MRx without the Q-CPR Meter attached. If you choose to use the HeartStart MRx with the Q-CPR Meter attached and you observe that the Q CPR Meter is displaying the Do Not Touch the Patient icon at an unexpected time, follow the MRx Instructions for Use and remove the meter from the patient and continue CPR according to your organization's protocol.Philips plans to take to remedy the issue. Philips will provide a software upgrade free of charge.If you have questions regarding this notification or need any further information or support, please contact your local Philips representative or call at 1-800-722-9377. |  | 6583 | 6583 units | Worldwide distribution: US (nationwide) and to countries of: ARUBA AUSTRALIA AUSTRIA BELGIUM BRAZIL CANADA ECUADOR EGYPT FINLAND FRANCE GERMANY HONG KONG INDIA INDONESIA IRELAND ISRAEL ITALY KENYA KOREA, REPUBLIC OF KUWAIT LATVIA LIBYAN ARAB JAMAHIRIYA MALAYSIA MEXICO MOROCCO NETHERLANDS NORWAY OMAN PAKISTAN PHILIPPINES POLAND QATAR RUSSIAN FEDERATION SAUDI ARABIA SINGAPORE SOUTH AFRICA SPAIN SWEDEN SWITZERLAND TAIWAN, PROVINCE OF CHINA THAILAND TURKEY UKRAINE UNITED ARAB EMIRATES and UNITED KINGDOM. | Open, Classified |  |  | 6583 | Computer
Z-1179-2014 | 67521 | No | JJW | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cannabinoids DAU Calibrator Control Level 1 |  |  | February 06, 2014 | 2014 | Lin-Zhi International Inc | Product catalog #s 0007c Cannabinoid 37.5 ng/mL Control Level 1 and 0008c Cannabinoid62.5 Control Level 2 of the lots mentioned may give decreased concentration readings towardsthe end of their shelf life due to use of dated but unexpired secondary stock used to create theproducts. | Nonconforming Material/Component | Lin-Zhi International, sent an Recall letter dated February 6, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The Recall letter identified the product by catalog number and product description as well as lot numbers. The reason for recall is explained and that the issue was confirmed by Lin Zhi Quality Control. The 2 affected lots were made with unexpired but dated secondary stock that is the cause of the problem.Lin Zhi requests that customers inform laboratories and customers to properly dispose of any of the affected lots that are still in use or unopened. Lin Zhi should be notified of the amounts for credit. Questions should be directed to 408-732-3883. |  | 216 | 216 items | US Distribution including the states of SC, NC, LA, FL, TX,, AL and GA and Internationally to Germany. | Terminated | March 10, 2015 | 397 | 346 | Not_Computer
Z-1177-2014 | 67463 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | bioMerieux PREVI Isola System (General purpose,microbiology, diagnostic Medical Device), (REF 29500/29500R), bioMerieux, Hazelwood, MO. |  |  | February 10, 2014 | 2014 | Biomerieux Inc | The firm has determined the product may fail to dispense the sample to the agar plate resulting in a "failure to dispense" in conjunction with urine specimens. This could lead to a "false" negative growth result. | Software design | The firm sent a recall notification letter to all of their consignees on 2/10/2014. |  | 471 | 471 instruments | Distribution was made to AK, CA, CO, CT, DC, FL, GA, IL, IN, KY, LA, MA, MO, NC, NJ, NY, OH, OR, PA, RI, SC, TN, TX, VA and WI. | Terminated | March 16, 2017 | 1130 | 471 | Not_Computer
Z-1190-2014 | 66837 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | G1762: 3 mL Line Draw, Luer lock syringe (heparin concentration: 23.5 I.U. per mL) with Filter-ProÂż device, syringe tip cap, and 5 mL Luer lock syringe |  |  | November 22, 2013 | 2013 | Smiths Medical ASD, Inc. | issue with the plunger tip of the PortexÂż Arterial Blood Sampling Line Draw Syringes | Component change control | Smiths Medical sent an Urgent Field Safety Notice dated November 22, 2013, to all affected consignees via FedEx beginning on November 23, 2013. Distributors were instructed to notify their customers. Customers were instructed to inspect their inventory for the affected product and quarantine the affected syringes. Complete and return the attached Urgent Field Safety Notice Confirmation Form by fax to 603-358-1017 or by email to LineDraw@smiths-medical.com. Upon receipt of the completed form, a customer service representative will contact them to arrange for exchange of their unused affecte devices for credit or replacement. Customers were instructed to forward the notice to all personnel who need to be aware within their organization and to any organization where the potentially affected devices have been transferred. Customers were instructed to report any issues with the product to Smiths Medical's Global Complaint Department at 1-866-216-8806 or globalcomplaints@smiths-medical.com.For questions regarding this recall call 603-352-3812. |  |  |  | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV. Internationally to Australia, Austria, Belgium, Belize, Brazil, Canada, Czech Republic, Denmark, England, Finland, France, Germany, Greece, Hungary, Ireland, Israel, Italy, Japan, Netherlands, Norway, Poland, Scotland, Spain, and Sweden. | Open, Classified |  |  |  | Not_Computer
Z-1193-2014 | 67439 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products 25-OH Vitamin D Total Calibrators; REF/CATALOG # 684 4055, IVD --- COMMON/USUAL NAME: VITROS Vitamin D Calibrators -- Ortho-Clinical Diagnostics High Wycombe HP12 4DP, UK --- For use in the calibration of the VITROS ECi/ECiQ Immunodiagnostic Systems, the VITROS 3600 Immunodiagnostic Systems, and the VITROS 5600 Integrated System for the quantitative measurement of total 25-OH Vitamin D in human serum. |  |  | January 15, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics has determined that the VITROS Immunodiagnostic Products 25-OH Vitamin D Total Calibrators, Lots 0091 through 0235, used with VITROS Immunodiagnostic Products 25-OH Vitamin D Total Reagent Packs, may generate negatively biased results. As a result, users must immediately discontinue using all current lots of VITROS 25-OH Vitamin D Total Calibrators and Reagent Packs. | Component design/selection | On 1/22/2014, Customer Letters were sent via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees to inform them of the issue and advise them to discontinue use of the affected lot(s) and to inform their Laboratory Medical Director if they suspect that a previously reported result may have been affected. On 1/22/2014, Distributor Letters were sent via US Postal Service Priority Mail and/or FedEx overnight courier to US distributor consignees and instructed them to notify their consignees of the issue and actions. |  | 1554 | 1554 units domestically | Nationwide | Terminated | April 03, 2017 | 1174 | 1554 | Not_Computer
Z-1198-2014 | 67464 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Infinia and Infinia Hawkeye Nuclear Medicine Systems. For Dual-Head Variable Angle Gamma Camera, Hawkeye Option for Dual-Head Variable Angle Gamma Camera (this includes GE Quasar Nuclear Medicine Systems). Nuclear Medicine Systems Infinia devices Product Usage: The intended use of the Quasar system is to perform general Nuclear Medicine imaging procedures for detection of radioisotope tracer uptake in the patient body, using a variety of scanning modes supported by various acquisition types and optional imaging features designed to enhance image quality in Oncology, Cardiology, Neurology and other clinical diagnostic imaging applications. |  |  | February 07, 2014 | 2014 | GE Healthcare, LLC | Potential safety issue related to unintended radial detector motion, which may occur during patient setup or at the end of a patient scan. This failure mode has been observed on a single system, in which the detector moved to the home position and then slowly starting drifting inward toward the patient. This generated both a visual and audio alert condition on both the console and the gantry. N | Device Design | GE sent an Urgent Medical Device Correction letter, dated 2/7/2014 to medical facilities beginning 2/7/2014. The letter was addressed to Manager of Nuclear Medicine, Hospital Administrators/Risk Managers, and Managers of Radiology/Cardiology. The letter described the issue and provided safety instructions. For Infinia and Infinia Hawkeye, Millennium VG, VG Hawkeye and VariCam nuclear Medicine Systems, the letter stated that a software upgrade would correct the issue. For the Helix Nuclear Medicine Systems, GE stated that a software upgrade can not be applied to the older systems and that additional information will be provided at a later date. UPDATE ON THE HELIX DEVICES: An Urgent Medical Device Correction letter, dated 4/7/2014 was sent to customers who still have a GE Helix Nuclear Medicine System. The letter requested to immediately stop using the device, and reiterated the March 2008 letter in which GE stated that they would no longer be able to offer full service support. The main reason for this was due to unavailability of parts required for maintenance and repair of the device because outside suppliers cease to produce the older parts. GE offered assistance in removal and disposal of the Helix system. |  | 3062 | 3062 units (1182 US, 1880 OUS) | Worldwide Distribution-USA (nationwide) including Puerto Rico and Guam and the countries of Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, Georgia, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea (Republic of), Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Malta, Martinique, Mexico, Morocco, Namibia, New Zealand, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, and United Kingdom. | Terminated | November 03, 2014 | 269 | 3062 | Computer
Z-1201-2014 | 67516 | No | FIP | Water purification system for hemodialysis. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Mar Cor Purification CWP Reverse Osmosis System. Models 101, 102, 104, and 106. |  |  | June 27, 2012 | 2012 | Mar Cor Purification | There exists a potential situation with the crimp connection at the pump head assembly. The potential exists that the reverse osmosis pump head wiring could build resistance resulting in resistive heating of the surrounding components. If this occurs, the pump head assembly may experience melting of the plastic components and potentially create smoke/fire and shut down of the unit. | Component design/selection | Mar Cor Technicians started visiting the consignees to make the correction, in June 2012 Mar Cor called the consignee prior to the visit. An Urgent Medical Device correction letter was sent to customers beginning 2/27/2014. the letter identified the affected units, described the issue, stated that a correction has been, or will be made to the devices and asked for a response form be sent back to the firm. |  | 130 | 130 | nationwide. No distribution outside the US. | Terminated | August 22, 2014 | 786 | 130 | Not_Computer
Z-1195-2014 | 67447 | No | NQI | Immunomagnetic circulating cancer cell selection and enumeration system. | Immunology | Pathology | 510(k) | CELLTRACKS ANALYZER IIÂż (Product Code 9555) |  |  | January 08, 2014 | 2014 | Veridex, LLC | Janssen Diagnostics, LLC received complaints of camera failures on the CELLTRACKS ANALYZER IIÂż. | Device Design | Janssen Diagnostics, LLC sent an Important Product Correction Notification letter dated January 8, 2014 to all affected customers via FedEx overnight courier letter and Foreign Affiliate Consignees by email on January 8, 2014. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the Confirmation of Receipt Response form attached. For questions contact Customer Technical Service at 1-877-837-4339. |  | 19 | 19 units | Worldwide Distribution - USA Nationwide | Terminated | July 06, 2017 | 1275 | 19 | Not_Computer
Z-1197-2014 | 67625 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | XiO RPT System |  |  | February 17, 2014 | 2014 | Elekta, Inc. | Dose and Monitor Unit values are not computed correctly when Elekta Motorized Wedges are used with Dynamic Conformal arcs. | Software design | Elekta, Inc. issued an Important Safety Notice (LINXi000019) issued February 14, 29, 2014. The Safety Notice identified the product, the problem, and the action to be taken by the customer. All affected customers of the issue globally by February 28, 2014. Provide a fix to the defect in an expeditious timeframe. It is mandatory for all customers to apply the fix. Version 5.00.01 will be released for the fix in March 2014. Follow up with the affected customers for confirmation receipt of the LIN, and confirmation of installation of patch for the fix by November 30, 2014.For questions regarding this recall call 770-670-2422. |  | 1776 | 1776 | USA including AL, AK, AZ, CA, CO, CT, DE, DC, FL, GA, IL, IN, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NJ, NY, NC, ND, OH, OK, PA, RI, SC, TN, TX, VA, WA, WV, WI, WY, Puerto Rico and Internationally to Albania, Algeria, Argentina, Australia, Austria, Bahamas, Barbados, Bermuda, British Guiana, Boliva, Brazil, British West Indies, Burundi, Canada, Chile, China, Columbia, Costa Rica, Cuba, Curacao, Cyprus, Czechoslovakia, Dominica, Equator, Estonia, Egypt, Finland, France, Gambia, Germany, Great Britain, Georgia, Greece, Guatemala, Hong Kong, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Lativa, Lithunia, Madagascar, Malaysia, Malta, Mexico, Mongolia, Myanmar, Netherlands, New Zealand, Nigeria, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, San Salvador, Saudia Arabia, Serbia, Singapore, Sir Lanka, South Korea, Spain, Sudan Slovenia, Suriname, Tajikistan, Thailand, Turkey, Turkmenistan, Taiwan, Ukraine, Uruguay, Venezuela, Viet Nam, Zambia. | Terminated | July 21, 2015 | 519 | 1776 | Not_Computer
Z-1208-2014 | 67627 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | REF 623421 RS-ROM Noiles Rotating Hinge, Femoral Rotating Hinge Cemented (with Pin), Right , X-Small, Sterile R |  |  | February 24, 2014 | 2014 | DePuy Orthopaedics, Inc. | DePuy Orthopaedics has identified the potential for holes to develop in the inner and outer flexible pouches that form the sterile barrier for S-ROM Noiles Rotating Hinge Femur with Pin devices. The outer carton and shrink wrap are intact. A package redesign is underway to resolve this issue. S-ROM Noiles Rotating Hinge Femur with Pin devices will remain in US Distributors inventory during the pa | Packaging | DePuy Orthopaedics sent an Urgent Information - Recall Notice dated February 24, 2014, to all affected customers. The notice identified the product, reason for the recall and potential risks associated with using the recalled product. The notice includes response cards for the US Distributors to verify they have read, understood, and are adhering to the directions given in the field communication and for hospitals where consignment inventory is removed a hospital letter is available, as needed. No hospital signatures are required. Distributors are asked to send completed distributor cards to DePuy. Product related questions should be addressed to DePuy Orthopaedics, .Sales Consultant. Clinical questions should be addressed to DePuy Scientific Information Office 1-888-554-2482; recall questions to 574-371-4917. |  | 132 | 132 | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Australia, Austria, China, Czech Republic, Denmark, Dubai / UAE, Estonia, France, Germany, Hungary, Iberia, India, Ireland, Israel, Italy, Malaysia, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Russia, Singapore, Slovenia, Spain, South Africa, Sweden, Switzerland, Turkey, United Kingdom, and Vietnam. | Terminated | March 03, 2015 | 372 | 790 | Not_Computer
Z-1196-2014 | 67578 | No | KNH | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Gyrus ACMI Disposable Falope-Ring Band Applicator Kits. REF 006889-901 and REF 005280-901 (with 8 mm Disposable Trocan). Sterile EO. Single Use Only. |  |  | February 19, 2014 | 2014 | Gyrus Medical, Inc | Due to an anomaly in the packaging process for specific lots of Disposable Falope-Ring Band Kit, it is possible that a defective seal could be present that may allow a breach of the package's sterile barrier and may compromise the sterility of the product. The breach may or may not be easily seen. | Process design | Olympus sent an Urgent Medical Device Recall letter dated February 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Olympus requires you to take the following action:I. Immediately cease any further use of any affected product you have, remove it from your inventory and quarantine it until it is shipped back to us.2. Call your Olympus customer service representative at 1-888-524-7266 to obtain a Returned Goods Authorization so that you may return the product with no charge to you. Olympus will issue a credit t or replacement to your facility y for any returned product.3. Please note on the enclosed questionnaire that you have received this information.4. Fax the completed questionnaire to 484-896-7128 regardless of whether you have any affected inventory at your facility.ln addition, if you may have further distributed this product, please identify your customers, notify them at once of this product recall, and appropriately document your notification process. Your notification to your customers may be enhanced by including a copy of this recall notification letter. |  | 856 | 856 units | US Distribution. | Terminated | April 14, 2015 | 419 | 856 | Not_Computer
Z-1145-2014 | 67520 | No | LYB |  |  | Dental | 510(k) | Denlase: Diode Laser Therapy System and Penlase: |  |  | July 15, 2013 | 2013 | China Daheng Group Inc. | Failure to provide calibration procedures to purchasers, as required by 21 CFR 1040.11(a)(2). For example, the user manual did not contain calibration procedures for Denlase or Penlase Lasers | Radiation Control for Health and Safety Act | China Daheng Group, lnc.(CDH}, ensures tha they will execute seriously their plans for the repair of their electronics products Denlase: Diode Laser Therapy System and Penlase: Dental Laser Therapy System.China Daheng Group, lnc.(CDH}, will send their updated 238 pes of Diode Laser Therapy System operation manual and 200 pes of Dental soft Tissue Laser brochure to their Distributor. They will pay all of their expense including manufacture cost and freight cost.China Daheng Group, lnc.(CDH}, will remedy these defects through changing the manuals and brochures for their distributor, purchasers and any subsequent transferee. They will make new version manuals and brochures. Then they will send them to their customers and distributor after finishing inspection by their quality department according to the procedure Denlase-Y0029A for manual and brochure inspection.CDRH approves the CAP subject to the following condition:.Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification. |  | 238 | 238 | CA, KS, IL, PA, NJ, WI, WY, KY, FL and NYIndia, Indonesia, Turkey, Denmark, and France, | Terminated | April 17, 2015 | 641 | 238 | Not_Computer
Z-1064-2014 | 65450 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | SL &amp; SLi Series (Mod 15/18/20/25) Linear Accelerator |  |  | May 28, 2013 | 2013 | Elekta, Inc. | Customers are resetting the default values outside recommended factory settings. | Device Design | The firm, Elekta, sent an "Important Field Safety Notice (IFSN 200 01 103 073 - Electron Applicator Auto-Tracking Defaults)" released on May 28, 2013 and is in the process of being distributed to all customers. The Notice requests users to please adhere to the instructions and advice given in manuals and system dialogues, which explicitly state that a change to auto-tracking values could compromise compliance with safety standards and can cause increased unwanted radiation dose to the patient. All machines should be checked against the default settings attached. If the actual settings are greater than the default settings, then Elekta's recommendation is that the default settings should b re-applied and all relevant physics checks (e.g. beam uniformity, output factor) performed. If, however, auto-tracking values are increased beyond their factory-set defaults, then electron applicator radiation leakage measurements should also be made to ensue that the system remains compliant to IEC safety. All users of Desktop Pro and Integrity 1.1 software will receive a mandatory upgrade to Integrity 1.2 (released 28-Jan-14. The upgrade path to Integrity 1.2 differs depending on system configurations at each customers site, therefore Elekta has released three separate Important Field Safety Modifications in order to upgrade all customers successfully. Important Field Safety Notice 200 01 501 036  Important Field Safety Modifications 200 02 501 037  Applicable System Configuration Desktop Pro with Fast Tuning Magnetron, New Patient Support System; Important Field Safety Notice 200 01 501 038  Important Field Safety Modification 200 02 501 039  Applicable System Configuration Desktop Pro with Linac Control System MK3i, Fast Tuning Magnetron, New Patient Support System; Important Field Safety Notice 200 01 502 035  Important Field Safety Modification 200 02 502 036  Applicable System Configuration All Integrity 1.1 sites. All users of Integrity 3.1 will receive a man |  | 449 | 449 | Worldwide Distribution: US (Nationwide) including states of: AK, CA, CT, FL, IL, IN, IA, LA, MD, MA, MI, MN, NJ, NY, OH, OK, OR, PA, TN, TX, VA, WA, WV, and Puerto Rico; and countries of: Argentina, Australia, Austria, Belarus, Belgium, Bosnia &amp; Herzegovinia, Brazil, Canada, China, Czech Republic, Denmark, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Italy, Japan, Jordan, Libya, Mexico, Netherlands, New Zealand, Norway, Panama, Philippines, Poland, Portugal, Russia, Senegal, Serbia, Singapore, Slovenia, South Africa, Spain, Sudan, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Open, Classified |  |  | 449 | Computer
Z-1210-2014 | 67686 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | DeRoyal(R) LIMB HOLDERS, COTTON ADJUSTABLE, UNIVERSAL, REF M2017 |  |  | February 21, 2014 | 2014 | DeRoyal Industries Inc | Moldy smell in product | Packaging | DeRoyal initiated an Urgent Voluntary Recall dated February 21, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. DeRoyal has asked that products be destroyed, and that notice of destruction be sent back. For questions contact your DeRoyal Sales Representative or DeRoyal Customer Service at 1-800-251-9864. |  | 8962 | 8,962 pairs | Worldwide Distribution - USA Nationwide in the states of FL, NC, IA, AL, IL, TX and the countries of Argentina. | Terminated | May 08, 2015 | 441 | 8962 | Not_Computer
Z-1209-2014 | 67572 | No | HBT | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Reusable Clip Applier Manufactured by AtriCure, Inc., 6217 Centre Park Drive, West Chester, OH 45069. |  |  | January 23, 2014 | 2014 | Atricure Inc | The firm was notified by a customer of a cracked spring in the handle of the RCD1 device. | Nonconforming Material/Component | On 1/24/2014 the firm sent Urgent Recall Notification letters to their customers. |  | 46 | 46 units | Nationwide Distribution-including the states of AL, CA, CO, CT, DE, FL, KY, MD, MI, MN, MO, NC, NY, OH, OR, PA, TN, UT, VA &amp; WA. | Terminated | May 13, 2014 | 110 | 46 | Not_Computer
Z-1213-2014 | 67517 | No | FIP | Water purification system for hemodialysis. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Mar Cor Purification Millenium HX Portable Reverse Osmosis System. Model 3025975. |  |  | June 26, 2013 | 2013 | Mar Cor Purification | High inlet water pressure beyond specification and the solenoid valve does not meet application which could result in a delay in treatment. | Component design/selection | Mar Cor sent an Urgent Medical Device Correction letter dated March 5, 2014 to all affected customers. The letter identifies the affected devices, what the issue is and tells customers to contact Mar Cor to schedule correction. A response form was asked to be completed and returned. For questions contact Scott Nichols, Project Coordinator/Service Trainer at 1-800-633-3080. |  | 232 | 232 units | USA Nationwide Distribution including Puerto Rico. | Terminated | January 22, 2015 | 575 | 232 | Not_Computer
Z-1214-2014 | 67529 | No | DJR | Methadone test system. | Toxicology | Toxicology | 510(k) | Product Brand Name(s): Alere TriageÂż TOX+MTD Drug Screen, Catalog Number 94400, Lots W56013B and W56035B. |  |  | February 05, 2014 | 2014 | Alere San Diego, Inc. | Alere San Diego, Inc. is initiating a recall of the Alere TriageÂż TOX Drug Screen, Catalog Number 94400, Lots W56013B and W56035B because Alere's investigation has shown that these lots have a frequency of false negative results for PCP (Phencyclidine) that is higher than the frequency shown in the package insert. If you have questions about previously reported results. | Process control | Alere San Diego, Inc. sent out recall notification letters to customers via direct mail or email. The letter, dated February 3, 2014, titled "URGENT MEDICAL DEVICE RECALL", provided customers the following information: product description with codes, reason for recall, impact, contact information and customer/distributor required actions. A Customer/Distributor Verification Form was provided. Customers were instructed to discontinue use and discard the affected lots. Alternately, if you elect to continue to use the product customers must verify if it will not be used to report results for PCP. Customers were instructed to return the Customer Verification Form by fax to 858-805-8457 with the amount of product they have for credit. The information should be shared with their laboratory staff and the notification should be retained as part of their laboratory Quality System documentation. If the product has been forwarded customers were instructed to provide a copy of the letter to them. Customers with questions were instructed to contact Alere San Diego, Inc. at 877-308-8287 or email Triage.Support@alere.com.For questions regarding this recall call 858-805-2000, ext 3015. |  | 20130 | 305 kits (239 kits of W56013B, 66 kits of W56035B) | Worldwide Distribution - USA (nationwide) and Internationally to Taiwan. | Terminated | October 27, 2014 | 264 | 20130 | Not_Computer
Z-1215-2014 | 66787 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | Dental Abutment labeled in part: "4555 Riverside Drive Palm Beach Gardens, FL 33410 USA 1-800-342-5454***EDAZ , EDAZx and ILDExxx***BellaTek Zirconia Abutment Position Number: XX Abutment; Pilier; Abutment; Abutment; Pilastro; Pilar***Rx only***Non-sterile, Sterilize prior to use.***Biomet 3i Dental Iberica S.L.*** |  |  | August 26, 2013 | 2013 | Biomet 3i, LLC | Reports relating to Biomet 3i's BellaTek Zirconia Dental Abutment fractures were received. | Device Design | The firm, BIOMET 3i, sent an "URGENT: Medical Device Removal - BellaTek Zirconia Abutments" letter dated August 2013 via email, facsimile or postal mail to its customers. The letter described the product, problem and actions to be taken. All customers were instructed to check their inventory for the affected product; return the abutments to BIOMET 3i for replacement and full credit; complete and return the attached Recall Return Response Form, and to advise their customers and forward a copy of the Medical Device Removal Notice and the contained recommendations.BIOMET 3i has decided to initiate a design change to mitigate the fracture risk, and consequently, will cease distribution of BellaTek Zirconia Abutments unit l the corrective actions are completed.If you have any further questions, please contact Customer Service at 1-800-342-5454 for assistance. |  | 10454 | 10,454 | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, KS, KY, LA, MA, MD, MI, MN, MO, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, Hawaii and Puerto Rico; and countries of: Spain, Australia, Canada, Colombia, Japan, Uruguay, Mexico, Estonia, United Kingdom, Denmark, Germany, Spain, Italy, Poland, Ireland Netherlands, France, Sweden, Portugal, Australia, France, Belgium, Switzerland, Norway, Czech Republic, France, Holland, and Luxembourg. | Terminated | September 26, 2014 | 396 | 10454 | Not_Computer
Z-1216-2014 | 67559 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Dynamics Picture Archiving and Communication System (PACS). Model numbers 10091804, 10091805, 10091806, 10091807, 10091673 |  |  | January 27, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Potential patient data mixup. Improper error handling could allow for two different patients to be displayed in syngo Dynamics. Images shown for one patient and reporting system shown for a different patient, leading to the potential for misdiagnosis. | Software design | A recall notification letter, dated January 27, 2014, was sent to end Users. Customer Safety Advisory Letter SY016/14/S. Siemens is preparing a Service Patch modification of syngo Dynamics v9.5 which will be available in March 2014. |  | 50 | 50 | Nationwide distribution to MN, AR, MA, TX, NJ, WV, GA, CA, PA, OH, MI, WI, TN, NY, KS, CT, NY, MO, IL, IN, OR, IA, and WA. | Terminated | March 16, 2015 | 413 | 50 | Not_Computer
Z-1212-2014 | 67472 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | FreeStyle Flash Blood Glucose Monitors System. |  |  | February 19, 2014 | 2014 | Abbott Diabetes Care, Inc. | Abbott Diabetes Care has identified through internal testing and investigation that all non-applied voltage legacy meters have the potential to produce out of range control solution results and/or erroneously low blood glucose results when used in conjunction with Free Style test strips lots within expiry. | Material/Component Contamination | The firm, Abbott Diabetes Care, sent an "IMPORTANT: URGENT PRODUCT RECALL"- letter dated February 14, 2014 to all customers for whom they have a record of receiving 0mV meters on February 19, 2014 via trackable method. Abbott also issued Press Release on March 14, 2014. The letter describe the product, problem and actions to be taken. The customers were instructed to discontinue use of your FreeStyle Flash or FreeStyle Blood Glucose Meter and call Abbott Diabetes Care Customer Service at 1-888-345-5364 for replacement at no charge; upon receiving replacement meter; immediately dispose of your old meter in accordance with your local and state disposal regulations; and please read your test strip insert carefully to ensure that you are using the correct test strip type for your Blood Glucose Meter. The firm, OmniPod, sent a ACD letter dated February 7, 2014 and included Abbotts February 14, 2014 letter to its customers. The letter described the product, problem, and actions to be taken. The customers were informed that you can safely continue to use the OmniPod System for your insulin delivery needs. The customers were instructed to read all of the materials included in the Abbott communication carefully and make sure you understand how it affects you; to immediately discontinue use of FreeStyle test strips with the Freestyle meter built into the OmniPod system; call Abbott Diabetes Care Customer Service immediately at 1-877-584-5159 for replacement strips; while waiting on replacement strips, use an unaffected meter such as the FreeStyle Freedom blood glucose meter; you can always use an alternative blood glucose meter.The notification will be posted on the Abbott Diabetes Care website at 222.abbottdiabetescare.com.If you have any further questions, please call Abbott Diabetes Care Customer Service at 1-877-584-5159. For questions or concerns with the continued use of your OmniPod Insulin Management System contact Insulet Product Support at 1-800-591-34 |  | 25203 | 25,203 units | Worldwide Distribution: US (nationwide) and countries of: Austria, Canada, Germany, Israel, Netherlands, Norway, Sweden, Switzerland, United Kingdom. | Terminated | August 11, 2015 | 538 | 25208 | Not_Computer
Z-1217-2014 | 67486 | No | MNI | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Gibralt Spine System Polyaxial Screw |  |  | January 06, 2014 | 2014 | Exactech, Inc. | The products have been identified to contain springs which have been manufactured with type-304 stainless steel rather than the specified titanium alloy. | Mixed-up of materials/components | Exactech sent an "URGENT PRODUCT RECALL NOTICE" dated January 6, 2014, to all affected customers. The letter identified the product and action needed to taken by the customers. Customers were instructed to: 1) Immediately cease distribution. 2) Extend the information to accounts in possession of the product. 3) Verify if they have any of the subject Gibralt Screws from the specified lots. 4) Complete and fax back the attached form. 5) Return affected products. For further questions, please call ( 352) 327-4847. |  | 242 | 242 | USA Distribution including the states of FL, NY, OR, IL, PA, MO and NJ. | Terminated | August 04, 2014 | 210 | 242 | Not_Computer
Z-1221-2014 | 67598 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns Malleable Dual-stage Venous Return Cannulae, 34/46 Fr with 1/2" connector, 15" (38 cm) long |  |  | March 03, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | During in-process inspection, Terumo Cardiovascular Systems (Terumo CVS) identified exposed basket wire on certain product lots of Sarns Malleable Dual-stage Venous Return Cannulae and Sarns Dual-stage Venous Return Cannulae. | Device Design | Terumo sent an Urgent Medical Device Recall letter dated March 3, 2014, to all affected consignees. The letter included the reason for removal, potential hazards, and customer instructions. Terumo CVS advised users to stop using the affected devices and return any unused product on hand. Questions or concerns were directed to Terumo CVS Customer Service: 1-800-521-2818 M-F, 8am-6pm.For questions regarding this recall call 734-741-6173. |  |  |  | Worldwide Distribution - USA including CA, CT, FL, GA, IL, IN, LA, MI, MN, MO, NC, NY, PA, TN, TX, and Internationallly to United Arab Emirates (UAE), Taiwan, Thailand, Malaysia, Belgium, Japan, Mexico, and Canada. | Terminated | August 12, 2014 | 162 |  | Not_Computer
Z-1226-2014 | 67282 | No | CIG | Bilirubin (total or direct) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products White Reference Slides, White Correction Factor Slide (WCF), DT Slides, 25 Slides, Part Number J02315; Made in USA by Ortho- Clinical Diagnostics, Inc. 100 Indigo Creek Drive, Rochester, NY 14626 --- This product is a part/ tool used when performing the correction factors adjustment on the VITROS Chemistry Systems. It does not have a 510(k) number. |  |  | January 09, 2014 | 2014 | Ortho-Clinical Diagnostics | An incorrect assay value (0.3785) was printed on the product carton for the 460 nm wavelength (filter) that may cause a delay in the ability to obtain and report TBIL, Bu, Bc, and derived tests results. The correct assay value is 0.8658. | Error in labeling | Ortho Clinical Diagnostics sent an Urgent Distributor Product Correction Notification letter dated January 9, 2014, via FedEx overnight courier to all affected customers. The letter notified users of the issue and the correct 460 nm value. Foreign affitliates were notified on January 9, 2014 via e-mail. The letter advised users to inspect their current inventory of VITROS White Correction Slides to determine if Lot 9052-0045-6358 is present. If there is remaining inventory of VITROS White correction Slides, Lot 9052-0045-6358, the remaining inventory may be discarded or the user may continue to use the product with the updated 460nm value provided. If VITROS White Correction Slides, Lot 9052-0045-6358 have been previously used to perform the reflectometer correction factors adjustment , then this procedure should be repeated using the corrected 460 nm values or an alternate lot of VITROS White Correction Slides. Customers were instructed to complete and return the Confirmation of Receipt Form and identify all customers and any other branches who were shipped the affected lot from their facility and send a copy of the customer communication. Customers were also instructed to forward the information if they have distributed the affected product outside of their facility. Customers with questions were instructed to call Customer Technical Services.For questions regarding this recall call 908-218-8776. |  | 64 | 64 units worldwide | Worldwide Distribution - USA (nationwide) Puerto Rico, and Internationally to Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Terminated | July 07, 2015 | 544 | 64 | Not_Computer
Z-1225-2014 | 67622 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart FR3 Automated External Defibrillator (AED), models 861388 and 861389. |  |  | February 06, 2014 | 2014 | Philips Medical Systems North America Inc. | Philip's is recalling HeartStart FR3 Automated External Defibrillator (AED) because the Printed Circuit Board (PCB), which contains the circuitry that supplies energy to the device, could become separated from the electrical contact points that hold the board in place. In the event the PCB becomes separated from the electrical contact points, the AED may not be ready to deliver a shock in the event of an emergency. | Process control | Philips sent an Urgent Medical Device Recall Letter dated February 10, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If you are have an affected device, please contact Philips for a replacement AED at 1-800-263-3342 and select option 5. Live technical support is available Monday through Friday, 7:00AM to 5:00PM PST. You can also leave a message24 hours a day, seven days a week; a customer support representative will return your call promptly on the next business day.Customers who have an affected device will be supplied with a replacement AED device of the same model as the recalled device from Philips, free of charge. Please keep your AED in service until you receive a replacement if there are no alternative defibrillators readily available.Philips has asked Distributors to provide them with the following information within five calendar days: -The name, address, telephone number and, if available, email address of all customers who received affected devices from you or from any of your sub-distributors. -The serial numbers of any affected devices remaining in their possession, or in the possession of their sub-distributors.-Philips directed distributors with affected devices in inventory to not sell these products, but to note them as in their possession as part of their response to Philips. -Philips communicated in this letter that they are responsible for returning to Philips any units in their inventory or in the inventory of their sub-distributors, and Philips will provide instructions for returning the affected devices. |  | 251 | 251 units total (128 in USA and 123 untils outside USA) | Worldwide Distribution - US Distribution including the states of AL, CA, CO, WA DC, FL, GA, HI, KS, MD, ME, MI, MS, NC, NM, NY, PA, TX, VA, WA and WI., and the countries of Australia, Austria, Canada, France, Germany, India, Italy, Ireland, Japan, Malaysia, Netherlands, Norway, Switzerland, Taiwan and United Kingdom. | Terminated | June 29, 2017 | 1239 | 251 | Not_Computer
Z-1234-2014 | 67584 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MEDITECH Microbiology; Client Server Releases 5.6, 5.65, 5.66, 5.67, 6.0, 6.05, 6.07, 6.1, 6.13 and 6.14. |  |  | February 21, 2014 | 2014 | Medical Information Technology, Inc. | Potential for erroneous result reporting. | Device Design | Beginning on February 3, 2014 MEDITECH distributed a notification to Client Server customers who have the Microbiology product in TEST and/or LIVE only for the releases noted in this correction report. This notification was made via e-mailed task updates that can be printed by the customer. Task updates sent via e-mail are immediately transmitted to the customer. The issue will be corrected by CS LAB DTS 12279. |  | 438 | 438 | Distributed USA (nationwide) and Canada. | Terminated | November 12, 2015 | 629 | 438 | Not_Computer
Z-1233-2014 | 67194 | No | LYC | Bone grafting material. | Dental | Dental | 510(k) | ENDOBON XENOGRAFT GRANULES 0.5 ML (Bovine bone substitute) |  |  | October 03, 2013 | 2013 | Biomet 3i, LLC | Individual pieces of a batch of Biomet's Endobon Xenograft Granules 0.5 ML (Bovine Bone substitute) may have an incorrect label with the wrong batch number. | Error in labeling | Biomet 3i sent a Medical Advisory Notice dated October 3, 2013 via mail, facsimile or telephone to all affected customers. The notice identified the product, the problem, and the action to be taken by the customers. The customers were instructed to do the following: 1) Review their inventor for any incorrectly labeled products and return them to their local Biomet distributor. 2) Provide this information to persons in their organization that uses or order these products. 3) Ensure that a copy of this letter is provided to any other organization to which the products maybe have been transferred. Customer were also instructed to respond to the attached Response Form via fax at 561-514-6316 or email to: domesticcomplaints@biomet.com. For any questions or concerns, call Biomet 3i Customer Service at 1-800-342-5454.For questions regarding this recall call 561-776-6706. |  | 704 | 704 | Worldwide Distribution - USA including CA, DC, FL, LA, MA, MO, NC, NJ, NY, TX, VA, WA and Hawaii and Internatially to Austria, Belgium, Canada, Switzerland, Germany, Spain, France, Great Britain (UK), Ireland, Mexico, Poland, Portugal, and Singapore. | Terminated | December 11, 2014 | 434 | 704 | Not_Computer
Z-1231-2014 | 67630 | Yes | BTS | Bronchial tube. | Anesthesiology | Anesthesiology | 510(k) | Rusch Brochopart and Rusch Brochopart White |  |  | February 28, 2014 | 2014 | Teleflex Medical | The product may fail to achieve seal of right lung due to the cuff inflating to one side | Device Design | Teleflex sent a Urgent Medical Device Recall Notification dated February 28, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed:1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 63568 | 63,568 ea. | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and the states of AL, AZ, CA, CO, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MD, MA, MI, MN, MS, MO, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WV and WI,and the countries of United Arab Emirates, Austria, Australia, Azerbaijan, Belgium, Belgrade, Bahrain, Canada, Switzerland, Cameroon, China, Costa Rica, Cyprus, Czech Republic, Germany, Denmark, Dominican Republic, Algeria, Ecuador, Estonia, Egypt, Spain, Finland, France, Great Britain, Greece, Guyana, Croatia, Hong Kong, Hungary, Ireland, Israel, India, Iraq, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Latvia, Libya, Morocco, Serbia, Malaysia, Martinique, Nigeria, Netherlands, Norway, Oman, Panama, French Polynesia, Philippines, Poland, Portugal, Qatar, Reunion, Russia, San Salvador, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, San Marino, Thailand, Trinidad, Tunisia, Turkey, Venezuela, Yemen and Zambia | Open, Classified |  |  | 63568 | Not_Computer
Z-1232-2014 | 67612 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart Home and OnSite (HS1) automated external defibrillators (AEDs). |  |  | February 06, 2014 | 2014 | Philips Medical Systems North America Inc. | Philips is recalling HeartStart Home and OnSite (HS1) automated external defibrillators (AEDs) because one or both contact pins may be contaminated by residue from the soldering process. This contamination could prevent an adequate connection between the pads cartridge and the AED. If poor contact between the pads cartridge and AED occurs, the device may be unable to deliver therapy in an emergency. | Vendor change control | Philips began sending out the Urgent Medical Device Recall HeartStart Home and OnSite (HS1) letter, dated 2014 FEB 10, to consignees on 2/17/14. The letter identified the product, the problem, and the action to be taken by the customer. Customers are advised to contact Philips for a replacement AED at 1-800-263-3342 and select option 5. Live technical support is available Monday through Friday, 7:00AM to 5:00PM PST. Customers can also leave a message 24 hours a day, seven days a week. DISTRIBUTORS are instructed to provide Philips with the information provided in the letter. Customers who have an affected device will be supplied with a replacement AED from Philips, free of charge. If customers/distributors need any further information or support concerning this issue, please visit www.philips.com/HeartStartHS1PinAction. |  | 5075 | 5075 units total (2724 units in the US and 2351 units outside the US) | Worldwide Distribution - USA Nationwide | Terminated | March 23, 2017 | 1141 | 5075 | Not_Computer
Z-1256-2014 | 67503 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Natural-KneeÂż II Constrained Knee System |  |  | February 25, 2014 | 2014 | Zimmer, Inc. | A complaint was received for a discolored All Poly Tibia. Root cause analysis found that the discoloration was the result of free radical oxidation. Theoxidation process was advanced prematurely as a result of the machined component being stored in a transparent container exposed to artificiallight for six years prior to inert packaging and sterilization of the device. | Storage | Beginning 2/25/2014 all distributors were notified via electronic mail. Hospital risk managers and distributors, surgeons with affected inventory will also be notified via courier or direct mail. The Urgent Lot Specific Recall Notice identified the issue and responsibilities. Distributors were asked to locate and remove the affected product in their territory, as well as identifying hospitals and surgeons who have previously implanted the affected product. Distributors were asked to return on-hand affected products to Zimmer and ensure all of their products are accounted for using the form provided in the letter and email it back to corporatequality.postmarket@zimmer.com. Hospital risk managers and surgeons were asked to locate affected product and return it to their Zimmer sales representative. Questions or concerns should be addressed to customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST Communications outside of the United States will occur approximately two weeks after the United States communications.100 % effective checks will be done. |  | 3638 | 3638 units | Worldwide Distribution-USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, FL, IL, IN, KS, LA, ME, MI, MO, MS, NC, ND, NE, NJ, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA, and WI, and the countries of Canada, Australia, India, Austria, Japan, Czech Republic, Seoul Korea, France, Venezuela, Germany, Luxembourg, Spain, and Switzerland. | Terminated | March 24, 2015 | 392 | 3638 | Not_Computer
Z-1242-2014 | 67618 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Part 475620 Intramedullary Bone Saw CAM Assembly |  |  | February 26, 2014 | 2014 | Biomet, Inc. | Incorrect labeling on specific lots of the Intramedullary Bone Saw.The Saw Blade Assembly was identified as CAM Assembly. | Labeling False and Misleading | The firm, BIOMET, sent an "URGENT MEDICAL DEVICE SAFETY ALERT" letter dated February 26, 2014 via Fed Ex overnight to domestic consignees. Australia and GSCC were notified by email and also received a hard copy letter via mail. The letter described the product, problem and actions to be taken. The letter stated that this action requires the immediate location and discontinued use of the product and its return to Biomet. The Consignees were were instructed to immediately locate, discontinue use of the product and remove the identified device(s) listed; carefully follow the instructions on the FAX Back Response Form and reply within three (3) business days; complete and return Form via Fax to: 574-372-1683 prior to return of product, and use priority carrier for your shipment. If you have further distributed this product, you MUST notify hospital personnel of this action via enclosed "Dear Risk/Recall Manager Notice", and note that you are charged with the location and return of these products. Shipping address is Return Goods, 56 E. Bell Drive Warsaw, IN 46582. Questions related to the notice should be directed to 574-372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 1 | 1 | Worldwide Distribution: US (nationwide) to states of: NY, TX, MI, and M; to countries of: Australia and The Netherlands. | Terminated | July 07, 2015 | 496 | 17 | Not_Computer
Z-1235-2014 | 67416 | No | INI | Motorized three-wheeled vehicle. | Physical Medicine | Physical Medicine | 510(k) | Buzzaround XL scooters' |  |  | January 21, 2014 | 2014 | Golden Technologies, Inc. | Buzzaround XL scooters' front to rear lockup can become unintentionally disengaged due to possible improper fit of the front alignment cup to the rear alignment screw. | Device Design | Golden Technologies sent a recall notification letter, dated January 14, 2014 to providers/dealers. A second letter, dated January 21, 2014 was sent to these same customers. Based on a slow initial response, the firm intends to send out the correction kit with the inspection acknowledgement form as soon as possible. Based on the overall response from dealers, a third consumer letter was drafted (notification3) for mailing directly to known consumers, if necessary. For help contact a special technical service hotline at 1-800-624-6374 ext. 392. |  | 5044 | 5044 | Worldwide Distribution - USA Nationwide | Terminated | July 23, 2015 | 548 | 5044 | Not_Computer
Z-1245-2014 | 67592 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Optima CT580/Discovery CT590RT scanners running software version 13BW16.3. |  |  | February 14, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to a software issue associated with the Head Scan Protocols used on your Optima CT580 or Discovery CT590RT scanner.A potential hazardous situation may occur during a CT head scan causing a novel and rare artifact that may hide pathology. When using the head and large head SFOV, which applies the Iterative Bone Option (IBO) | Software design | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated February 14, 2014 to its Consignees. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Directors and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The consignees were instructed to ensure that all potential users in your facility are made aware of this safety notification and recommended actions. A GE Healthcare Service representative will perform the required software update on each affected system. This activity will be performed at no cost to consignees.Please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. You may also contact GE Healthcare Call Center: United States or Canada: 800-437-1171 and in other countries, please contact your local GE Healthcare Service Representative. |  | 33 | 33 (21 US, 12 OUS) | Worldwide Distribution: US (nationwide) to states of: CO, IN, MA, NJ, NC, OH, PA, TX, and VA; and countries of: Algeria,France, Germany, Poland, Russia, and Taiwan. | Terminated | April 29, 2015 | 439 | 33 | Computer
Z-1263-2014 | 67685 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Philips Model EPIQ 5 Ultrasound System. Intended for diagnostic ultrasound imaging and fluid flow analysis. |  |  | February 28, 2014 | 2014 | Philips Medical Systems North America Inc. | A problem has been detected with Philips Model EPIQ 5 and 7 Ultrasound System. The system will report incorrect Doppler velocities if you use High Pulse Repetition Frequency (HPRF) with multiple sample volume gates. While in High Pulse Repetition Frequency (HPRF) in Display Zoom Velocity Scale using manual spectral Doppler trace, the velocities in the velocity results table may be overstated by | Software design | The firm, Philips Healthcare Ultrasound, sent an "Medical Device Correction" letters dated 2014 Feb 28 via Certified Mail via the USPS by March 5, 2014. The letters described the product, problem and actions to be taken. The customers were instructed to not use HPRF in PW Doppler mode, to avoid the noted issues. Philips Field Service Engineer will be contacting them to schedule software upgrade to the systems free of charge.If you need any further information or support concerning this issue, please contact the Quality Specialist II at 978-659-4519 or your local Philips representative. |  | 262 | 489 total =227 units in US and 262 units Internationally (no Canadian consignees) | Worldwide Distribution: US (nationwide) including states of: AL, AK, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NM, NJ, NM, NY, OH, OK, PA, SC, TN, TX, VA, WA, and WV; and countries of: Australia, Austria, Belgium, Canada, Chili, Ecuador, Egypt, Finland, France, Germany, Greece, India, Ireland, Italy, Japan, Korea Rep, Malaysia, Netherlands, New Zealand, Norway, Peru, Philippines, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey and United Kingdom. | Open, Classified |  |  | 262 | Computer
Z-1265-2014 | 67702 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Halogen Lamp component in RetCam 3 System - |  |  | March 07, 2014 | 2014 | Clarity Medical Systems Inc | One lot of Halogen lamps have the potential to fail sooner than the expected life of the lamp which could result in damage limited to the internal electrical components inside the RetCam 3 computer module. | Nonconforming Material/Component | The firm, Clarity, sent a "Field Bulletin" dated March 12, 2014 to customers - risk managers and distributors by e-mail on March 12, 2014. They plan to follow this with a second letter on or about March 28, 2014 which will include a second bulletin, replacement bulbs and installation instructions. Clarity stated in the letter to its customers "In the meantime, you may continue to use your RetCAM 3."If you have any question of concerns, please contact Clarity at +1 925 621 322 or email to: service@retcam.com |  | 234 | 234 lamps | Worldwide Distribution: US (nationwide) to states of: TN, MA, CA, TX, MD,FL, MS, IL, and PA; and to countries of: France, Thailand, Egypt, Mexico, Kuwait, India, Austria,Japan, Sweden, Peru, United Kingdom, Philippines,The Netherlands, China,, UAE,Israel, Germany, Italy. | Terminated | August 11, 2014 | 157 | 234 | Not_Computer
Z-1238-2014 | 67588 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | SLC 2-way standard Lubricious Coated |  |  | February 07, 2014 | 2014 | Convatec Inc. | The lumen of the Medline Sterile Foley Catheter was found to be blocked by excess lubricious coating material. | Device Design | Convatec sent Urgent: Voluntary Medical Device Recall letters dated 2/7/2014 via UPS and email. Foreign customers were provided the notice via DHL. The letter identified the affected product, problem and actions to be taken. Customers are asked to perform a count of recalled product in inventory. All consignees should be contacted and be instructed to return product based on their procedures. Customers were instructed to complete and return the enclosed respoonse form. |  | 5540 | 5,540 catheters | Worldwide Distribution and in the state of TX | Terminated | December 18, 2014 | 314 | 100580 | Not_Computer
Z-1246-2014 | 67590 | No | JTN | Antimicrobial susceptibility test disc. | Microbiology | Microbiology | 510(k) | Oxoid Vancomycin 30 mcg, each cartridge contains 50 susceptibility discs, CT0058B. The firm name on the label is Oxoid Ltd., Basingstoke, Hampshire, England. |  |  | February 18, 2014 | 2014 | Remel Inc | Some of the discs may not be impregnated with the antibiotic. | Process control | ThermoFisher sent an Urgent Medical Device Recall letter dated February 18, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify any personnel who need to be notified of the potential for a false indication of resistance to Vancomycin. Customers were also instructed to review their inventory for the affected product and discard it. Customers should review results reported using this lot and consider retesting and/or seek appropriate expert advice. Return the attached Medical Device Recall Return Response acknowledging receipt of the notice and disposal of the affected product. If product was further distributed, customers should contact those entities, advise them of the situation and provide them with a copy of the letter. Customers with questions were instructed to contact the Technical Services Department at 800-255-6730 (US) and 913-888-0939 (International).For questions regarding this recal lcall 913-895-4185. |  | 200 | 40/5/50-disc cartridge packages | Nationwide Distribution | Terminated | June 25, 2014 | 127 | 200 | Not_Computer
Z-1247-2014 | 67775 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | 2.5MM Aggressive Blade |  |  | March 18, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the single use 2.5MM Aggressive Blade because the line on the affected part (P/N 5290-928-000, Lot #13280CG2) was a straight line without a number scale or logo. The line required for this part number includes a number scale that goes from 1cm  7cm. | Employee error | Stryker Instruments sent an Recall Notification dated March 18, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed if you have further distributed this product, please forward this letter and the attached Business Reply Form (BRF)to all affected locations. Please indicate each location on the BRF. A FedEx shipping label will be emailed to each account upon receipt of the signed Business Reply Form. This label will allow the account to return the product to be destroyed.A Business Reply Form will be included in each mailing. Return of this form will serve as confirmation that the account has received the notification and executed the recall. Follow up phone calls, emails, and faxes will be sent to non-responding accounts as necessary.Product destruction will be handled at Stryker Instruments, Kalamazoo. Labeling will be defaced and the Aggressive Blades will be disposed of in the trash compactor.For questions regarding this recall please contact Stryker Instruments:Monday-Friday 8am-5pm (EST) 269-389-2921 |  | 90 | 90 units | US Distribution including the states of NV, CA, WA, TN, OR, SD, ID, IN and CO., and Internationally to Malaysia | Terminated | May 15, 2014 | 58 | 90 | Not_Computer
Z-1268-2014 | 67599 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | AMS URIC ACID For the In vitro quantitative determination of Uric Acid in serum |  |  | February 11, 2014 | 2014 | AMS Diagnostics, LLC | Marketing the devices outside 510(k) requirements | No Marketing Application | AMS Diagnostics sent a Recall Notification letter dated February 11, 2014 to all affected customers. The letter instructed the customers to: 1) Review their inventory for the lots listed and if they currently have them, to discard and notify the firm for a replacement. 2) If they don't have the lots listed, please check the box, sign, date and return the letter via fax at (954) 217-0046 or email to dj@amsdiagnostics.com. Customers with questions were instructed to call 954-217-0040 ext. 302.For questions regarding this recall call 954-217-0040. |  | 13 | 13 kits | Nationwide Distribution including FL, GA, MD, CA, NC, MA, SC, WY, and TX. | Open, Classified |  |  | 296 | Not_Computer
Z-1270-2014 | 67607 | No | DJG | Opiate test system. | Toxicology | Toxicology | 510(k) | Single Analyte DAU Calibrator (5 mL) |  |  | February 18, 2014 | 2014 | Lin-Zhi International Inc | Customer reported that the Synchron NBUP cutoff calibrator (10 ng/mL) could not be qualified through quality control testing. The cutoff calibrator was recovering at a concentration close to the NBUP Control Level 1 (7 ng/mL). | Packaging | Lin-Zhi International sent a Recall letter dated February 18, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Letters acknowledge that Lin Zhi was notified by the customer that the product could not be qualified and that this was confirmed by Lin-Zhi internal testing.Letters request that the distributor notify all customers of the situation and ask them to destroy any remaining product from the affected lots. the root cause is identified and letters also state that a corrective action has been implemented to prevent future problems.Questions may be directed to Lin Zhi International at 408-732-3856. |  | 7200 | 120 units, 60 of each lot number | US Distribution including CA. | Terminated | April 25, 2014 | 66 | 7200 | Not_Computer
Z-229-2014 | 67639 | No | DSQ |  |  | Cardiovascular | PMA | HeartMate II LVAD Pump and Pock Controller Kit; |  |  | March 04, 2014 | 2014 | Thoratec Corporation | Discovery of serious injuries and deaths associated with the process of changing from a primary System controller to their back-up System controller in patients using the Pocket System controller model. | Device Design | Thoratec sent an Urgent Medical Device Correction letter dated March 4, 2014 to all affected customers. The letter was sent via e-mail. All distributors outside of the USA and Canada to whom affected product has been distributed have been notified of this field correction notice and have been provided copies of the Customer Letter to forward to their customers. The letter identified the affected product , problem and actions to be taken. Customers were requested to complete and sign the attached Acknowledgement Form and return it to Thoratec via fax (925-847-8571) or e-mail a scanned copy tocustomer.service@thoratec.com. For questions call 1-800-528-2577. |  | 36 | 36 | Worldwide distribution: US (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA, and WI; and countries of: Austria, Belgium, Czech Republic, France, Germany, Ireland, Israel, Italy, Lebanon, Lithuania, Poland, Saudi Arabia, Singapore, Slovenia, Sweden, Switzerland, The Netherlands, Turkey, and United Kingdom. | Terminated | August 27, 2015 | 541 | 218533 | Not_Computer
Z-1266-2014 | 67576 | No | IMP | Wheelchair component. | Physical Medicine | Physical Medicine | 510(K) Exempt | Otto Bock Aerial and Aerial Active Wheelchair Cushions. Material numbers: SK757 476C00=SK701, 476C00=SK702, 476C00=SK703, 476C00=SK704, 476C00=SK705, 476C00=SK706, 476C00=SK707, The following material numbers were not distributed within the US: 476C00=SK751, 476C00=SK752, 476C00=SK753, 476C00=SK754, 476C00=SK755, 476C00=SK756, 476C00= SK757. |  |  | January 31, 2014 | 2014 | Bock,Otto,Orthopedic Ind,Inc | Otto Bock is recalling all Aerial and Aerial Active Wheelchair Cushions from specific material numbers. During internal testing, it was discovered that there is a small risk of flammability to the cushion after it has been washed at high temperatures. The cushions should either be returned to Otto Bock or destroyed. | Unknown/Undetermined by firm | An Urgent Voluntary Device Recall letter was sent to customers beginning 1/31/2014. The letter identified the affected material numbers, described the reason for recall, and asked that product be returned or destroyed. A response Questionnaire was asked to be completed and returned. |  | 464 | 464 units (US 113, OUS 351) | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Bosnia and Herzegovina, Brazil, China, Columbia, Croatia, Denmark, France, Germany, Greece, Hong Kong, Hungary, Iceland, Japan, Mexico, Netherlands, Poland, Romania, Russia, South Africa, Spain, Sweden, Switzerland, Thailand, and United Kingdom. | Open, Classified |  |  | 464 | Not_Computer
Z-1271-2014 | 67657 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MEVION S250, System, Radiation Therapy, Charged-Particle, Medical used for Proton Radiation Therapy |  |  | March 05, 2014 | 2014 | Mevion Medical Systems, Inc. | Geometric positioning error can occur after a rotational correction has been applied to one or more treatment fields | Software design | Mervion issued an Important Safety Notice letter on March 5, 2014 by visit to sites to the single customer using the product to prevent the use of multiple isocenters treatment modalities. The letter identified the affected product, problem and actions to be taken. For questions call 978-540-1555. |  | 1 | 1 unit | USA Nationwide Distribution in the state of MO | Terminated | April 03, 2014 | 29 | 1 | Not_Computer
Z-1272-2014 | 67556 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis Systems |  |  | February 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Certain units of the AXIOM Artis Systems are using a flat detector cooling unit with inappropriate connectors that may possibly leak, causing a system failure and possible drop out of the gantry system. | Device Design | Siemens sent a recall notification letter, dated February 7, 2014, to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions call +49 9191 18 0. |  | 37 | 37 | USA Nationwide Distribution in the states of WI, KY, OH, NJ, NC, FL, PA, MA, IA, TX, NY, IL, MI, AZ, CA, and LA. | Terminated | December 30, 2014 | 326 | 37 | Not_Computer
Z-1276-2014 | 67710 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | WeckÂż Vista", Universal Laparoscopic Port, Size 5/10mmx100mm, |  |  | March 11, 2014 | 2014 | Teleflex Medical | Complaints were received of leakage of insufflation gas through the device, which can cause difficulty or delay in the completion of laparoscopic surgical procedures. | Nonconforming Material/Component | Teleflex sent an Urgent Medical Device Recall Notification letter dated March 11, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return all affected product to Teleflex Medical per the instructions provided in the letter. Customers were asked to complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 55521 | 55,521 ea. in total | Worldwide Distribution - USA Nationwide including Puerto Rica and the countries of: Australia, Brazil, China, Canada, Malaysia, Korea, Ireland Singapore and Vietnam. | Terminated | August 15, 2017 | 1253 | 55521 | Not_Computer
Z-1274-2014 | 67263 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Drill Template and the Saw Guide for the Ulna Osteotomy System. |  |  | January 07, 2014 | 2014 | Synthes, Inc. | The Drill Template and the Saw Guide for the Ulna Osteotomy System may exhibit jamming, bending, or breaking of the screw connecting the Saw Guide to the Drill Template or jamming, bending, or breaking of the K-wire during fixation of the Drill Template. | Component design/selection | A recall notification letter, dated January 7, 2013, was sent to sales consultants and Users. |  | 343 | 343 | Nationwide Distribution-including the states of AR, AK, VA, IN, PA, OH, CT, TX, MI, WI, NC, CO, IL, and MA. | Terminated | September 10, 2015 | 611 | 343 | Not_Computer
Z-1275-2014 | 67615 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DICOM Export Format for Sum Dose with Velocity v3.0.0. |  |  | July 25, 2013 | 2013 | Velocity Medical Solutions, LLC | An anomaly was detected such that, under certain conditions, a summed dose created by Velocity replaces the calculated dose from an approved plan in Eclipse when importing the dose from Velocity. | Software design | Velocity Medical Solutions sent a Urgent Field Safety Notice dated July 25th, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Velocity advises customers creating Dose volumes in Velocity to refrain from exporting these dose files to Eclipse until the customers install has been upgraded to Velocity 3.0.1, which will be available August 5th, 2013.A new version of Velocity that corrects this error, Velocity v3.0.1, will be available on August 5th, 2013 for upgrade at no cost. Please contact your Velocity Medical representative to obtain this upgrade.Service representative will upgrade the application at your earliest convenience.Please advise the appropriate personnel working in your medical department of the content of this letter.We sincerely apologize for any inconvenience and thank you in advance for your co-operation.Contact Information. Helpdesk: (404) 920-1966 or (855) 857-0802 and press 2. |  | 54 | 54 installs | Worldwide Distribution - US (nationwide) including the states of AL, CA, CO, CT, DC, FL, GA, IL, IN, LA, MA, MD, MI, MS, NC, NY, PA, SC, TN, TX, WI, and the countries of Belarus, Italy and Switzerland. | Terminated | September 11, 2014 | 413 | 54 | Not_Computer
Z-1273-2014 | 67638 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare G2 Clinical Access Mislabeled CDs for PN: 063-1829-11 Rev A and Update Kit, PN: 040-1565-10. |  |  | February 14, 2014 | 2014 | Spacelabs Healthcare, Llc | Spacelabs has voluntarily recalled G2 Clinical Access Software CD because the programs installed on the Clinical Access CD are incorrect and it will prevent the customer frominstalling Clinical Access. | Labeling Change Control | Spacelabs Healthcare began sending out the Medical Device Correction letter, dated 28 February 2014, to domestic customers on 28 February 2014 and to international customers by 7 March 2014. The letter identified the product, the product, the problem, and the action to be taken by the customer. Customers are also directed to call 1-800-522-7025 and select 2 for Technical Support.Customers are advised to dispose the CD, PN: 063-1829-11 Rev A previously supplied and replace it with the replacement CD (Rev B) provided. |  | 36 | 36 units total: 17 units were distributed in the US and 19 units distributed outside the US. | Worldwide Distribution- USA (nationwide) including the states of Alabama, Florida, Georgia, Louisiana, Missouri, New York, South Carolina, Tennessee, Texas and Wyoming, and the countries of Australia, Canada, Czechoslovakia, China, Colombia, France, Great Britain, Italy, Netherlands, Poland, Saudi Arabia, and Turkey. | Terminated | May 20, 2014 | 95 | 36 | Computer
Z-1293-2014 | 67733 | No | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Sutures, Non-absorbable, Synthetic, Polypropylene, Sterile, Rx only, |  |  | March 11, 2014 | 2014 | Teleflex Medical | The products are being recalled because they did not meet minimum needle attachment strength requirements. | Employee error | Teleflex sent an Urgent Medical Device Recall Notification letter dated March 11, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to return all affected product to Teleflex Medical per the instructions provided in the letter. Customers were asked to complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. For questions contact your local sales representative or Customer Service at 1-866-246-6990 |  | 32271 | Total 32,271 ea. | Worldwide Distribution - U.S. Nationwide and the countries of Canada, Germany, and Ireland | Terminated | August 15, 2017 | 1253 | 32271 | Not_Computer
Z-1297-2014 | 67672 | No | MIH |  |  | Cardiovascular | PMA | Ovation Prime Abdominal Stent Graft System; |  |  | February 27, 2014 | 2014 | Trivascular, Inc | Communication to European physician users of the Ovation Prime Abdominal Stent Graft system to notify them of changes to Instructions for Use. | Device Design | Trivascular sent an Important Medical Device Information letter dated February 27, 2014, to all European physicians who have scheduled cases through March 31, 2014 or until IFU;s are translated and available for European physicians. The letter provided notification of an update that will appear in a future revision of the product Instructions for Use (IFU). The IFU will be available in their country when the required translations and product registration/notification have been completed. This IFU does not require rework or return of the affected product. Customers with questions were instructed to contact their local TriVascular representative.For questions regarding this recall call 707-543-8805. |  | 45 | 45 | Worldwide Distribution including Austria, Belgium, Czech Republic, France, Germany, Greece, Hungary, Italy, Luxembourg, Norway, Poland, Spain, and Sweden. | Terminated | April 11, 2014 | 43 | 45 | Not_Computer
Z-1299-2014 | 67624 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | D-10" Hemoglobin Testing System, |  |  | February 27, 2014 | 2014 | Bio-Rad Laboratories, Inc. | On a rare occasion the D-10 software might generate an extra result by duplicating the result form the previous sample and the Sample ID/Injection number from the next sample. | Nonconforming Material/Component | Bio-Rad Laboratories, Inc. sent an Urgent Medical Device correction letter dated February 27, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. A Customer medical Device Correction Response Form must be collected from each customer to ensure that they have received this communication. Forms should be faxed to the Bio-Rad CSD Regulatory Affairs Department at 510-741-3954 along with the completed Subsidiary Medical Device Correction Response Form.For questions regarding this recall call 570-741-6114. |  | 483 | 483 instruments in US. OUS to be determined. | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Japan, China, India, SriLanka, Argentina, Paraguay, Uruguay, Chile, Peru ,Bolivia ,Mexico ,Panama, ,Costa Rica ,Ecuador, Venezuela, Colombia, Brazil, Jamaica, Austria, Switzerland, France, Germany, Greece, Portugal, Spain, Italy, UK, Belgium, Sweden Norway, Finland ,Israel/Palestine, Tunisia, Algeria, Egypt, ,Morocco, Jordan, Yemen, Iraq, Kuwait, Qatar, Oman, Lybia, Lebanon ,UAE, Saudi Arabia, Pakistan, Afghanistan, Turkey, South Africa, Chad, Kenya, Benin, Burkina Faso, Cameroon, Gabon, Senegal, Ghana, Mali, Ivory Coast, Nigeria, Djibouti, Czech Republic, Slovakia, Slovenia, Serbia, Hungary, Romania, Poland, Russia, Ukraine, Kazakhstan, Bulgaria, Belarus, Georgia, Mauritius, Maldives, Vietnam, Philippines, Korea, Thailand, Hong Kong , Indonesia, Taiwan, Malaysia, Singapore, New Zealand, and Australia. | Terminated | June 04, 2014 | 97 | 483 | Computer
Z-1300-2014 | 67567 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | PowerPort Slim Implantable Port, indicated for patient therapies requiring repeated access to the vascular system. Product Codes: 1716000, 1716070 |  |  | February 27, 2014 | 2014 | Bard Access Systems | Labeling discrepancy.The label states that the implantable port contains silicone suture plugs, however approximately 23% of the units contained in the affected lots have a PowerPort Slim titanium implantable port without silicone suture plugs. | Packaging | Notified by Urgent Customer Notification Letter date February 13, 2014, customers were asked to immediately examine their inventory and identify any product subject to this notification. In addition, if they have further distributed this product, please identify customers and notify them at once of this notification. Customers may include a copy of this letter in your notification to customers. Customers may choose to use the involved product with the understanding that the kit may contain a PowerPortÂż Slim implantable port without suture plugs, or may return the applicable product for replacement or credit. |  | 777 | 777 | US Nationwide distribution including WA, KY, IL, and IA. | Terminated | September 05, 2014 | 190 | 777 | Not_Computer
Z-1202-2014 | 67016 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Artis Q and Artis Q biplane systems |  |  | December 05, 2013 | 2013 | Siemens Medical Solutions USA, Inc | There exists the possibility that during clinical operation, a breakdown of the x-ray tube could result in loss of x-ray imaging for the Artis Q and Artis Q biplane systems with a flat detector as 40 HDR (large 30x40). | Device Design | Siemens sent an Important Customer Safety Notice letter dated November 22, 2013, to affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were instructed to discontinue use until the affected product has been successfully replaced by a Siemen's representative.at no cost. |  | 19 | 19 | USA Nationwide Distribution in the states of TN, MA, CA, IL, OH, FL, UT, WA, GA, IA, NY, MO, IN, and CO. | Terminated | November 06, 2014 | 336 | 19 | Not_Computer
Z-1302-2014 | 67712 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Morph AccessPro Steerable Introducer; |  |  | March 03, 2014 | 2014 | Biocardia, Inc. | Firm discovered the potential for particulates in the inner lumen of Morph Access Pro Steerable Introducer, which could result in am embolic event. | Other | The firm, BioCardia, Inc., sent an "URGENT-VOLUNTARY MEDICAL DEVICE RECALL" letter dated March 3, 2014 its customers, hospital's Director of Risk Management, via USPS mail and priority mail to Cardiac Catheterization Laboratory. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your inventory; quarantine the product; return the product to BioCardia and contact Customer Service at 650-226-0148 to coordinate product return; in the event you no longer have this device as a result of further distribution, identify the purchaser and notify them at once of this recall; and complete and return the enclosed MEDICAL DEVICE RECALL RETURN RESPONSE Acknowledgement and Receipt Form via FAX to 650-631-3731 or mail to: BioCardia, Inc., 125 Shoreway Road, Suite B, San Carlos, CA 94070 as soon as possible.If you have any questions, call the Vice President, Quality Assurance at 650-226-0143. |  | 2541 | ~2541 units - both catalog numbers. | Worldwide Distribution: US (nationwide) and country of: Israel. | Terminated | June 19, 2014 | 108 | 2541 | Not_Computer
Z-1305-2014 | 67646 | No | NTE | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | FIBERNET Embolic Protection System: |  |  | March 03, 2014 | 2014 | Medtronic Vascular, Inc. | Medtronic is recalling a specific subset of lot numbers of the FIBERNET Embolic Protection System (EPS) where the yellow polyimide tube on the FIBERNET EPS was not removed from the FIBERNET EPS during the manufacturing process and subsequently may remain on the FIBERNET EPS. | Manufacturing material removal | Medtronic sent an Urgent Device Recall letter dated March 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records show that your facility has received one or more of these FIBERNET EPS, as shown in the attachment to this notification. Consequently, Medtronic requests you immediately take the following actions:1. Immediately remove and quarantine all potentially affected product that remains in your inventory.2. Return these units to Medtronic. Contact Customer Service at 1-888- 283-7868 and reference RADAR # 174028 to initiate a return and credit of product. A Medtronic representative can assist in facilitating the return and replacement of this product as necessary.3. Complete the attached Customer Confirmation Certificate and email to RS.CFQFCA@Medtronic.com, or fax to Medtronic at 651-367-0612 to the attention of Customer Focused Quality.Medtronic has taken the necessary steps to prevent any future shipment of the potentially affected product. Regulatory agencies will be notified about this recall as applicable.Please share this notification with others in your organization as appropriate. lf any FIBERNET EPS within the scope of this recall have been forwarded to another facility, please notify that facility accordingly and facilitate the retrieval of the affected product.We apologize for the inconvenience this may cause you; please be assured that patient safety and product quality remain our primary concern. Should you have any questions, please contact your Medtronic representative or Medtronic LifeLine Technical Services at 1 (877) 526-7890 or (+1 763) 526-7890. |  | 50 | 50 Total all sizes: 18 in US, 32 Internationally | Worldwide Distribution - US including the states of IL, PA, MA, VA. RI and LA., and the countries of Germany, Italy, Poland, Romania, Saudi Arabia, Kosovo, Slovakia and Spain. | Terminated | June 03, 2014 | 92 | 50 | Not_Computer
Z-1306-2014 | 67513 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | i-STAT G3+ Cartridge |  |  | January 01, 2014 | 2014 | Abbott Point Of Care Inc. | Abbott Point of Care has determined that some individual patient results generated with thei-STAT G3+ cartridge lot N13183 have the potential to exhibit incorrectly elevated PCO2 anddepressed pH results. | Process control | APOC sent an Urgent Recall Notice dated January 2014 via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the affected product and provide a copy of this letter to any institution that was forwarded the affected product. All unused cartridges will be credited and are to be returned as instructed on the enclosed Business Reply Card. To ensure continuity of supply at their facility, customers were instructed to follow their standard process to place an order or call Customer Sedrvice at 1-800-323-9100, Option #2, then Option #9). Customers with questions were instructed to contact Abbott Point of Care Technical Support at 1-800-366-8020, Option 1 of their Abbott Point of Care representative.For questions regarding this recall call 800-323-9100. |  | 48 | 48, 450 US, 1400 OUS (49,850 catridges) | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | October 20, 2014 | 292 | 48 | Not_Computer
Z-1307-2014 | 67504 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Radiation Therapy Treatment Planning System, models 2.5, 3.0, 3.5 and 4.0; a software system designed for treatment planning and analysis of radiation therapy. |  |  | February 06, 2014 | 2014 | RAYSEARCH LABORATORIES AB | Potential for dose errors due to software program errors. | Software Design Change | RaySearch Laboratories sent an Field Safety Notice, Medical Device Correction letter dated February 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.ACTIONS TO BE TAKEN BY THE USERA workaround that eliminates the described error is to select the discard rotations option in the Image Registration Module before creating a rigid registration intended to be used for hybrid deformable registration. Another option is to use the ROI based deformable registration when applicable, as this functionality is not affected by the error.SOLUTIONThis issue will be resolved in a patch release of RayStation, scheduled for market release February 28, 2014. Release schedule is subject to regulatory approval for some markets. In the meantime, this field safety notice is being distributed to all customers.CONTACT INFORMATIONFor further information, please contactFor the Americas:Customer Support, RaySearch Americas, at +1 877 778 3849 For Europe, Asia and rest of the world:RaySearch Laboratories AB, at +46 8 5450 6130 Thank you for your cooperation, and we apologize for any inconvenience |  | 151 | 151 units Total (88 units domestically &amp; 63 units internationally) | Worldwide Distribution - US including the states of NY, OH, MD, CT, CO, TN, IL, AZ, MS, NJ, PA, DE, FL, CA and WA., and the countries of New Zealand, Australia, Republic of Korea, China, Italy, Germany, Norway, Belgium, Switzerland, United Kingdom and Spain. | Completed |  |  | 151 | Computer
Z-1298-2014 | 67502 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Draeger PS500 Power Supply Unit of the Infinity ACS Workstation Critical Care (Evita Infinity V500), Evita V500 and /or Babylog VN500 Ventilators, used for ventilation during short term transport of patients within a hospital. |  |  | February 07, 2014 | 2014 | Draeger Medical Systems, Inc. | The battery capacity of the optional PS500 Power Supply Unit of the Infinity ACS Workstation Critical Care (Evita Infinity V500) did not last as long as expected. | Component design/selection | Drager sent an Urgent Medical Device Recall letter dated February 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to use an affected device for patient transport only if absolutely necessary; not to rely on the battery charge status indicator, always supervise the patient and the ventilator during transport; ensure that they have a manual breathing bag available for manual ventilation as recommended in the instructions for use. If the power fail alarm occurs during transport, immediately provide manual ventilation and connect ventilator to a wall power source to resume ventilation. DragerService representative will schedule the implementation of a permanent solution as quickly as possible. The letter should be distributed to users of the Evita V500 and /or Babylog VN500 Ventilators to make them aware of this device behavior and risk mitigation measures. Customers with questions were instructed to call 1-800-543-5047. For questions regarding the operation and/or servicing of their Draer ventilator contact Technical Support at 1-800-543-5047. For questions regarding this recall call 215-721-5400. |  | 1544 | 1544 | Nationwide Distribution | Terminated | August 17, 2015 | 556 | 1544 | Not_Computer
Z-1310-2014 | 67628 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | TDWorkstation version 11.01.A or higher. TDWorkstation Evolution. |  |  | February 17, 2014 | 2014 | Technidata S.A. | If a conflict occurs between a Demography Query and a Result Transmission by the instrument, the TDWorkstation could send an incorrect patient demography to an instrument. | Software design | An email dated 2/17/14 was sent to their customers to inform them about concerns only if their site uses TDWorkstation and TD-Workstation Evolution products that meet all the listed criterias. The criterias are listed in the customer notification letter and it informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the reply form and return it by fax (514) 270-2770 with in 15 days from receiving the customer notification form. Customers that do not respond the reply form, Technidata will consider those customers do not wish to install the correction. |  | 85 | 85 units total (1 unit in US) | Distributed in the state of PA. | Terminated | March 31, 2014 | 42 | 85 | Not_Computer
Z-1309-2014 | 67558 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis Zee Angiographic x-ray systems. |  |  | February 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc | An inappropriate cable connection was assembled for certain Artis zee systems. which may cause a system failure and possibly lead to a restricted maneuverability of the gantry system. | Device Design | A recall notification letter dated February 7, 2014, was sent to End Users. |  | 9 | 9 | MN, WI, PA, MI, AL, OH, NC, AR, and NJ. | Terminated | March 16, 2015 | 402 | 9 | Not_Computer
Z-1308-2014 | 67511 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RaySearch Ray Station, Radiation Therapy Treatment Planning System, Stand-alone software |  |  | February 06, 2014 | 2014 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories has recalled RaySearch Ray Station 4.0 software due to clinical dose calculation errors during radiation therapy. | Software Design Change | On February 6, 2014, RaySearch Laboratories distributed Urgent Medical Device Correction notices to their customers via email regarding the affected product. To determine if the version you are using is affected, bring up the About box in the RayStation application and check if the number reported is "4.0.2.9". If so, the notice applies to your version. Customers are advised to inform all planning staff and all users about the instructions listed on the notice. Customers are advised to inspect their system and identify all installed units with the referenced software version numbers listed on the notice. Corrective labeling is distributed to the customers in the form of the safety notification attached to this report and the software problem will be resolved in RayStation patch 4.0.3 scheduled for market release end of February 2014 and in all future releases. For further information and return of the reply form, please contact: Freddie Cardel, Director of Customer Support, RaySearch Americas, (877) 778-3849 or freddie.cardel@raysearchlabs.com. |  | 19 | 19 units (16 units domestically &amp; 3 units internationally) | OH, MD, NJ, WA, IL, MA, FL, Spain, United Kingdom, Italy, | Open, Classified |  |  | 19 | Computer
Z-1311-2014 | 67749 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Zoll Medical X Series, PropaqMD, PropaqM Defibrillator/Pacemaker/Monitor With System Software Version 02.10.02.00 or Higher, automated external defibrillator. |  |  | March 05, 2014 | 2014 | ZOLL Medical Corporation | Defibrillator/Pacemaker/Monitor May Latch in a Continuous Device Reset Loop and cannot be used | Software design | Zoll Medical issued a Notification letter daetd March 5, 2014, to all affected consignees. The notification letter provides information on how to avoid the issue and instructs customers to contact ZOLL to get affected devices updated. Customers will be given the option to have a software upgrade kit sent to them to update their devices or return devices to ZOLL Service department in Chelmsford Massachusetts to have the device upgraded with the latest software release (Software version 02.16.04.00). Customers with questions were instructed to contact the 24/7 technical support numbers 1 (800) 348 9011 or +1 (978) 421 9460 to assist users with any aspect of this notice. |  | 3169 | 3,169 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, AE, Afganistian, Australia, Belgium, Czech Republic , China, Chile, Croatia, Denmark, France, Great Britain, Hungary, Israel, Italy, Japan, Korea, Kuwait, Russia, Saudi Arabia, Singapore, South Africa,Thailand, and Turkey. | Open, Classified |  |  | 3169 | Computer
Z-1312-2014 | 67623 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Pinnacle3 Radiation Therapy Planning (RTP) System, software version 9.0. |  |  | February 11, 2014 | 2014 | Philips Medical Systems (Cleveland), Inc. | Philips Medical Systems have recently determined that a software nonconformance can cause incorrect beam geometry. This issue affects Pinnacle3 Radiation Therapy Planning (RTP) System, software version 9.0 manufactured from 09/29/09 to 12/01/2013. | Software design | Philips Healthcare sent an Urgent Field Safety Notice dated February 10, 2014, to all affected consignees. The letter included the Problem Description, Hazard Involved, How to Identify Affected Products, Action to be taken by Customer /User, Actions Planned by Philips and Further Information and Support. Consignees were instructed to review the following information with all members of their staff who need to be aware of the contents of this communication including clinics within their institution that are remotely connected to their Pinnacle Server. Customers were questions were instructed to contact their Philips representative. For customers in North America please contact Customer Care Solutions Center at 1-800-722-9377. For all other countries their local Philips Healthcare Office should be contacted.For questions regarding this recall call 1-800-722-9377. |  | 1016 | 1016 | Worldwide Distribution - USA including US: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV and Internationally to AUSTRALIA, AUSTRIA. BELGIUM, CANDA, CHINA, DENMARK, FRANCE, GERMANY, HONG KONG, HUNGARY, INDIA, IRELAND, ITALY, JAPAN, LUXEMBOURG, MALAYSIA, MEXICO, MONTENEGRO, NETHERLANDS, NEW ZEALAND, OMAN, RUSSIA, SOUTH ARICA, SOUTH KOREA. SPAIN,. SWITZERLAND, TAIWAN, THAILAND, TURKEY, EMIRATES, UNITED KINGDOM.. | Terminated | November 21, 2014 | 283 | 1016 | Computer
Z-1313-2014 | 67582 | No | MEF | Teething ring. | Dental | Dental | 510(K) Exempt | Infantino Squeeze &amp; Teethe Monkey, "Go Gaga Squeeze &amp; Teether Coco the Monkey". Non-Fluid Filled Teething Ring. |  |  | January 06, 2014 | 2014 | Infantino LLC | Infantino initiated this voluntary recall of all lots of Infantino Squeeze &amp; Teethe Monkey, "Go Gaga Squeeze &amp; Teethe Coco the Monkey", Model Number 206-647, because the tail of the monkey may pose a choking hazard to young children. | Other | Infantino, LLC initiated this recall by notifying its customer on January 6, 2014 and Consumer Product Safety Commission (CPSC) posted the press release on 02/20/2014. CPSC posted the press release on the CPSC website. The CPSC recall press release dated February 20, 2014 informed consumers should stop using this product unless otherwise instructed and it is illegal to resell or attempt to resell a recalled consumer product. The CPSC press release contains a Recall Summary with the name of product, hazard, remedy, consumer contact information "Infantino toll-free at (888) 808-3111 between 8 a.m. and 4 p.m. PT Monday through Friday or online at www.infantino.com and click on Recall Information on the home page.", Recall Details on the number of units, description, incidents/injuries, remedy, sold exclusively at Target stores nationwide, Importer, Manufactured in China, and CPSC Consumer Information Hotline and Media Contact. The CPSC press release also contained photos of the product.Infantino, LLC also posted on their company website information on this recall. |  | 144000 | 144,000 units | US nationwide. | Terminated | April 23, 2015 | 472 | 144000 | Not_Computer
Z-1314-2014 | 67647 | No | DYJ | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | AOA Distalizer Lock Nut Screw, found in products listed as: Distal Jet, Spring Jet, Mesial Jet, and Uprighter Jet. Removable functional orthodontic appliances that are to modify the growth of the jaws in a prescribed growth pattern. |  |  | February 24, 2014 | 2014 | Allesee Orthodontic Appliances | The Lock Nut component of the Jet appliances may be missing the screw collar, which serves as a mechanical stop to prevent over-tightening in appliance assemby. Without this collar, an over-tightened lock nut may stop the piston assembly from moving freely, subsequently preventing intended movement of the teeth. | Nonconforming Material/Component | An advisory letter dated January 10, 2014 was sent to customers. The letter titled "ATTENTION: FIELD CORRECTIVE ACTION", informed customers the reason for recall, product description and devices affected dates of manufacturer, potential hazard, a photo comparison of a lock nut with screw collar and lock nut with no screw collar, instructions to the customers, and contact information. The letter was accompanied by an acknowledgement form titled "DISTALIZER JET LOCK NUT ACKNOWLEDGEMENT FORM".A follow up letter dated February 24, 2014 on March 7, 2014 to customers. This letter addresses the issue of the replacement screws as a follow up to the initial January 10, 2014 letter. The follow up letter, titled "ATTENTION: MEDICAL DEVICE RECALL" is dated Februray 24, 2014, refrences the January 10, 2014 letter that AOA sent, attaches the letter for reference, and informs customers that "AOA is now offering the option to obtain a Lock Nut with Screw Collar to replace within devices determined to be affected by this recall". The follow up letter informs customers of the reason for recall, the description and manufacturing dates of product affected, instructions to follow, and contact information. The customers were instructed to 1) continue to monitor the patient during the time of treatment; or 2) remove the device from the patient's mouth and return it to AOA for evaluation/replacement; or 3) request a replacement Lock Nut Screw with collar by contacting our Customer Support Department at 1-800-262-5221, and complete and return the attached acknowledgement form titled "DISTALIZER JET LOCK NUT RECALL ACKNOWLEDGEMENT FORM" via fax to 262-886-6879.If you have any questions, contact the Customer Support Department at 1-800-262-5221. |  | 411 | 411 | Worldwide Distribution: US (nationwide) and the country of : Canada. | Terminated | August 23, 2014 | 180 | 411 | Not_Computer
Z-1315-2014 | 67549 | [6 empty] | Omni Micro-electrode/reference electrode for cobas b221 analyzer, Model/Catalog/Part Number: |  |  | February 28, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Reference electrode used beyond the guaranteed in-use 52 week lifetime,may leak and potentially cause erroneous pH and/or sodium results. Other Ion parameters are not affected. | Device Design | Roche Diagnostics Operations sent an Urgent Medical Device Correction letter dated February 27, 2014 via UPS Ground (receipt required) to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the enclosed fax form and fax it to 1-888-345-0253. For questions contact Roche Customer Business Support Center at 1-800-428-5076. |  |  |  | Worldwide Distribution - USA nationwide including PR and the country of Guam. | Open, Classified |  |  |  | Not_Computer
Z-1321-2014 | 67861 | No | EJF | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Titanium Orthos Brackets, an orthodontic bracket. 022 +9/+9/0, Part Number 449-6212, Lot Number 121359169. A fixed attachment on a tooth which holds an archwire during orthodontic treatment. |  |  | March 25, 2014 | 2014 | Ormco/Sybronendo | Ormco Corporation inititated recall of U2R Titanium Orthos (0.22+9/+9/0), Part Number 449-6212, Lot Number 121359169, because the color ID dot was marked on the wrong tie-wing of the brackets. | Other | Recall notification letters sent to all affected consignees via 1st class mail on March 24, 2014. The consignees are instructed to complete the Return Form and have the option to return any affected product in their inventory to Ormco Corporation.The letter is dated "March 24, 2014" and titled: "URGENT: MEDICAL DEVICE RECALL". The letter informed customers of the recall, reason for recall, product description with part number and lot number, instructions to follow, and contact information. The letter was accompanied by a "Recall Return/Acknowledgement Form". |  | 150 | 150 units | US and Canada. | Terminated | October 27, 2014 | 216 | 150 | Not_Computer
Z-1327-2014 | 67819 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 5/10/12 mm x 100mm Weck Vista Universal Cone Open Access, Catalog No. 405933. |  |  | March 11, 2014 | 2014 | Teleflex Medical | Complaints of leakage of insufflation gas through the device. | Nonconforming Material/Component | Letters were issued on March 11, 2014 to customers informing them of the recall. The letter asked the consignee to immediately discontinue use and quarantine any products on hand. The letter included a Recall Acknowledgement Form which is to be returned. A letter was also issued to distributors on March 11, 2014, requesting them to conduct a sub-recall. |  | 55521 | 55,521 total | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of AL, AZ, AR, CA, CO, CT, FL, GA, IL, IN, KS, KY, MD, MA, MI, MN, MS, MO, NH, NJ, NY, NC, ND, OH, OR, PA, SC, TN, UT, VA, WA, WV, WI and the countries of Australia, Brazil, Canada, China, Israel, Korea, Malaysia, Singapore, and Vietnam. | Terminated | January 28, 2016 | 688 | 55521 | Not_Computer
Z-1317-2014 | 67734 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Citation TMZF HA 132 degrees Neck Angle, V40 Hip Stem - Right Size #3: catalog # 6265-5113 , Left Size #4: catalog #6265-5104 , and Right Size #4: catalog #6265-5114; Stryker Howmedica Osteonics Hip Implants. |  |  | February 03, 2014 | 2014 | Stryker Howmedica Osteonics Corp. | Stryker received a report from the field indicating that a citation TMZF HA 132 degrees neck angle V40 Hip Stem - Left Size # 4 was identified in a packaging associated with a citation TMZF HA 132 degrees Neck Angle V40 Hip Stem - Right Size #3. | Error in labeling | Stryker sent an Urgent Medical Device Recall Notification letters and Product Accountability Forms to Hospital Risk Management, Chief of Orthopaedics, and Surgeons via Fed Ex with return receipt on February 26, 2014. The letter identified the affected product, problem and actions to be taken. Customer were instructed to return the affected product to the attention of Regulatory Compliance. Please attach the fluorescent orange PRODUCT REMEDIATION sticker to your return. For questions call 201-972-2100. |  | 48 | 48 units | USA Nationwide Distribution | Terminated | December 02, 2014 | 302 | 48 | Not_Computer
Z-1318-2014 | 67716 | No | FLF | Ethylene oxide gas sterilizer. | General Hospital | General Hospital | 510(k) | 3M" Steri-Vac" Gas Sterilizer/Aerators, Model #'s 4XL,5XL, 8XL |  |  | March 17, 2014 | 2014 | 3M Company - Health Care Business | Recently, 3M received a single report of a potential exposure to Ethylene Oxide (EO) at a hospital installation involving 3M" Steri-Vac" Gas Sterilizers. The cause was traced to a rare situation involving a failure of the EO cartridge puncture assembly. This assembly was only in production during the period November 10, 2006 through December 14, 2007 and in service during the period February 7, 20 | Device Design | 3M sent an "Urgent Medical Device Field Correction" letter dated March 6, 2014. The letter was sent to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. You are being notified because our records indicate you received a unit that contains the cartridge puncture assembly that is the subject of this field correction. 3M will contact you to schedule an appointment for a field correction on any unit in the range of serial numbers provided in the table. 3M's goal is to service the affected units with minimal disruption to your facility. Units manufactured outside of the dates November 10, 2006 to December 14, 2007 are not affected. .For Further Information contact 3M Health Care Service Line at 1-800-228-3957 or your local 3M representative. |  | 419 | 73 US, 346 OUS | Worldwide Distribution - USA (nationwide) Distribution including the states of AL, AZ, CA,CO, FL, GA, HI, IL, IN, IA, KS, KY, MD, MA, MI, MN, MS, NE, NJ, NY, NC, ND, OH, OK,OA, SC, TX, WA, WV and WI. and the countries of COLOMBIA, VENEZUELA, UAE, THAILAND, GERMANY, CHINA, KOREA, CANADA, JAPAN, RUSSIA, INDIA, POLAND, PERU, AUSTRIA, ISRAEL, INDONESIA, SPAIN, CHILE, MALAYSIA, SWITZERLAND, EL SALVADOR, ECUADOR, LEBANON, MEXICO, ARGENTINA, BULGARIA, TAIWAN, TURKEY,ROMANIA, GUATEMALA and NEW ZEALAND. | Terminated | January 22, 2015 | 311 | 419 | Not_Computer
Z-1316-2014 | 67586 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE 2000; IMMULITE 2000XPi; VersaCell |  |  | February 03, 2014 | 2014 | Siemens Medical Solutions Diagnostics | There is a potential for incorrect results to be reported from processing of a sample tube. | Device Design | Siemens sent an Urgent Medical Device Correction letter dated February 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The customers was told how to correct the issue if it has occurred, and how to prevent the issue from occurring. The completed form should be faxed to (302) 631-7597. Any questions can be directed to Siemens Technical Support Representative. Siemens plans to implement a correction at a later time. |  | 1006 | 1006 units | Nationwide and Foreign Distribution. | Open, Classified |  |  | 1006 | Not_Computer
Z-1322-2014 | 67574 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Integrity 1.1. |  |  | March 28, 2013 | 2013 | Elekta, Inc. | The problem is that the "static tolerances" from the calibration files (loaded database) are used for beam delivery instead of the machine calculated values. | Software design | An Important Field Safety Notice (IFSN) has been released and distributed to all customers. Integrity 1.1 Incorrect Static Tolerances for Linac Geometric Axes, 200 01 502 032. The notice explains the issue and that the user should apply the static Tolerance Updater tool to check tolerances and adjust as required. Important Field Safety Modification Digital Accelerator FCO Ref: 200 02 502 031 Integrity R1.1 Static Tolerance Updater Tool was released 03-2013. Elekta has released Integrity version 1.2 which contains the permanent fix for this issue. The permanent solution fixes the error that made the digital accelerator protection functions use the value loaded from the calibration database. The solution now ensures the protection functions use the values the digital accelerator calculates. Important Field Safety Notice 200 01 502 035 has been released to inform customers the solution is available. Important Field Safety Modification 200 02 502 036 will track the installation of the software by Elekta service engineers. |  | 498 | 498 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CT, FL, GA, KY, LA, MD, MO, NM, NY, NC, OH, OK, PA, TN, TX, and WA and the countries of Argentina, Australia, Austria, Belarus, Belgium, Brazil, Canada, China, Colombia, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Estonia, France, Germany, Greece, Honduras, Hong Kong, India, Indonesia, Iraq, Ireland, Italy, Japan, Jordan, Malaysia, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Peru, Poland, Romania, Russia, Saudi Arabia, Senegal, Singapore, Slovakia, South Korea, Spain, Sudan, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | July 26, 2017 | 1581 | 498 | Computer
Z-1331-2014 | 67741 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens Medical Solutions, ACUSON P300 Ultrasound system with P300 PA 230E cardiac probe (Esaote product 7348). |  |  | March 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | ACUSON P300 ultrasound systems using the PA230 transducer with a user-defined preset which includes the 100% setting for output power could cause the transducer surface temperature to exceed the maximum allowable limit. | Device Design | Siemens sent an Customer Safety Advisory Notifications on March 7, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.A new version of software resolving this error will be provided free of charge. Your local Customer Service Engineer will contact you to schedule an appointment to make the necessary repairs on your system. If you have any questions, please contact your local service support person for information regarding timelines and status. Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization who need to be aware of this issue. As always, patient safety issues are a high priority. To date, no patient injury has been reported. This problem was discovered as part of our ongoing quality process.For further questions please call (650) 694-5398.We sincerely regret any inconvenience this condition may cause in your daily operations. |  | 675 | 675 | Class II Recall - Worldwide Distribution - US Distribution and the countries of Australia, Austria, Belgium, Bhutan, Bolivia, Botswana, Bulgaria, Canada, Chile, Colombia, Croatia, Dem. Rep. Congo, Denmark, Ecuador, France, Germany, Greece, India, Indonesia, Ireland, Israel, Italy, Ivory Coast, Japan, Kuwait, Malaysia, Mauritius, Mongolia, Morocco, Mozambique, Myanmar, Namibia, Netherlands, Norway, Peru, Philippines, Poland, Portugal, Republic Korea, Romania, Saudi Arabia, Serbia, Seychelles, Singapore, South Africa, Spain, Switzerland, Taiwan, Trinidad,Tobago, Turkey, U.A.E. and United Kingdom. | Terminated | November 12, 2015 | 615 | 675 | Computer
Z-1320-2014 | 67844 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | Merit Custom Syringe Kit, Convenience Kit, I.R. Embolization Pack, K02-01010A, Sterile EO. |  |  | March 13, 2014 | 2014 | Merit Medical Systems, Inc. | The products are labeled as sterile but were not sterilized. | Release of Material/Component prior to receiving test results | Merit Medical sent an Urgent Product Recall Notice dated March 14, 2014 to all customers and sales representatives via email. The letter identified the affected product, problem and actions to be taken. Customers were provided a Product Retrieval Form, product identification information, instruction to immediately quarantine any devices and discontinue use, ensure all personnel to whom devices were distributed are made aware of this field action, and instructions to contact their Merit representative to arrange product return and replacement. For questions call 1-801-316-4822. |  | 15 | 15 | Worldwide Distribution - USA Nationwide in the state of WI and countries of Thailand and Hong Kong. | Terminated | July 07, 2014 | 116 | 95 | Not_Computer
Z-1329-2014 | 67736 | No | FTM | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Proceed Surgical Mesh, Product Codes PCDB1, PCDT1, and PCDJ1. |  |  | February 19, 2014 | 2014 | Ethicon, Inc. | The possibility exists that an incomplete seal on the packaging compromised the sterility of the device and introduced the potential for delamination. | Under Investigation by firm | Ethicon sent an Urgent Medical Device Recall letter dated February 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are asked to examine their inventory and determine if they have any of the affected product. If any product has been distributed, those hospitals should be notified. The business reply form should be completed and faxed or emailed to Stericycle at 1-888-641-9732 or ethicon5798@stericycle.com. For clinical or product support contact a customer representative or call the Customer Support Center at 1-877-ETHICON to speak with a registered nurse or contact Stericycle at 1-888-345-0524 and reference event #5798. Stericycle is handling returns. |  | 223 | 223 units | US Distribution including the states of FL, NJ, MO, IL, PA, NC. TX, TN, WA, ID, UT, WV, MS. CA, AL, MA. GA, LA and OK. | Open, Classified |  |  | 223 | Not_Computer
Z-1335-2014 | 67758 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 5X250 MM Curved Cannula, Arm 2 |  |  | March 13, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive Surgical has become aware of the potential for certain lots of the Single-Site 5mm Curved Cannulae to become damaged during use. The second part of this action is to update Single-Site labeling. These labeling changes instruct users how to inspect cannulae prior to use and to appropriately handle cannulae during use. | Device Design | Intuitive sent an Urgent Medical Device letter dated March 25, 2014, to all users of cannulas that have the potential to be damaged during use as well as a second letter to all single site cannula users that provided clarity to instruction for use for inspections and handling of the product.Certain lots manufactured prior to August of 2012, and users were asked to use Attachment A to help identify affected lots numbers. Phase 1 - Cannulas should then be inspected with information provided in Attachment B. Any damaged devices should be quarantined and returned to intuitive. If inspection does not find damage, the device may continue to be used. Any other cannulae and any non damaged product should be used according to Attachment C's instructions for use.Phase 2 - Once replacement cannulae are available, the undamaged cannulae should be returned.The Acknowledgement form should be completed and returned to Intuitive. Intuitive will contact consignees when replacement cannulae are available.Questions should be directed to Customer Service at 800-876-1310.Those consignees that use the 5 mm canulae but do not have affected product were provided with a letter that only includes the Single Site Cannula inspection and Use addendum. |  | 5217 | 5217 total cannulae | Worldwide Distribution - USA and the countries of Australia, Belgium, Brazil, Chile, China, Columbia, Denmark, France, Germany, India, Italy, Lebanon, Monaco, Netherlands, Norway, Romania, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | September 03, 2014 | 174 | 5217 | Not_Computer
Z-1330-2014 | 67849 | No | DYW | Orthodontic plastic bracket. | Dental | Dental | 510(k) | Spirit MB Bracket, UR Central +14T +5A.018, Part Number 494-0110, Lot Number 081351013. |  |  | March 25, 2014 | 2014 | Ormco/Sybronendo | Ormco Corporation initiated this recall of Spirit MB Brackets, Part Number 494-0110, Lot Number 081351013, because the ID dot, which designates the bracket as distal/gingival, is misplaced. The misidentification may lead an orthodontist to place the bracket on the wrong tooth, leading to movement 10 degrees in the wrong direction. | Use error | Ormco Corporation initiated this recall by sending recall letters to customers on March 25, 2014 via 1st class mail. The letter identified the affected product, problem and actions to be taken. The consignees are instructed to complete the "Return Form" and return any affected product in their inventory. The letter was written in Japanese language and accompanied by a "Return Form". |  | 150 | 150 units | Internationally Distributed in the country of Japan. | Terminated | April 18, 2014 | 24 | 150 | Not_Computer
Z-1336-2014 | 67597 | No | HGX | Powered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Playtex Nurser Deluxe Double Electric Breast Pump, Model X06578AO, powered by a 12V DC power supply, Model SY-12120, which is included with the package. |  |  | March 18, 2014 | 2014 | Energizer Personal Care | Some of the power adapters outer casings may become loose and separate, resulting in a potential for electric shock. | Component change control | Energizer sent and Voluntary Retail Recall Notice letter dated March 18, 2014 to retail customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to stop sale and return the affected product. The letter provides instruction to customers to return the affected product. For questions or concerns, contact Energizer Personal Care at 1-888-207-1492 or visit www.playtexproducts.com.On 9/9/2014 the firm expanded the recall to one additional production lot to include: Suspect date codes for Breast Pumps: P12315-XXXX to P13205-XXXX. Suspect adaptor date codes: 1238 to 1324Revised Press issued 9/9/2014. |  | 35803 | 35,803 | USA Nationwide Distribution in the states of: AL, AZ, CA, CO, CT, DE, FL, GA, IA, IL, IN, KS, KY, MD, MI, MN, MO, NC, NJ, NV, NY, OH, ON, OR, PA, SC, TN, TX, VA, VT, WA, WICanada - one consignee, Energizer (recalling firm's) company store | Completed |  |  | 35803 | Not_Computer
Z-1337-2014 | 67665 | [6 empty] | Syphilis Test Strip (Serum) |  |  | March 04, 2014 | 2014 | LumiQuick Diagnostics Inc. | Products shipped to US locations that do not have approved applications for PMA , IDE or 510(k) clearance. | No Marketing Application | Urgent Device Recall notification letters were sent to all US consignees on March 4, 2014 by e-mail and US mail as well as a phone call. The letter identified the products and actions to be taken by the customers. |  |  | ALL | Distributed in the states of CA, FL, NJ, TX, and NY. | Terminated | May 09, 2014 | 66 |  | Not_Computer
Z-1417-2014 | 67757 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 7mm Drain Kit, Flat Full Perforated, Saber 100ml Bulb and 3C System Sterile (10/case) |  |  | March 20, 2014 | 2014 | Aspen Surgical Products, Inc. | The recall has been initiated due to a potential compromise of the sterile barrier of the Surgidyne Wound Drainage and Surgidyne Bulb Evacuator &amp; Collection System medical devices. | Packaging process control | Aspen Surgical sent an Urgent Medical Device Recall letter dated March 20, 2014, to all affected customers and will be carried out at the distributor and retailer level, which can include doctor's offices, hospitals, and clinics. The distributor was instructed to contact their local FDA District office, conduct a sub-recall and to forward this Recall Notification and Response form to their retailers.A written notification to the consignees was sent by certified mail.Aspen has given the customers the option to either 1.) Return the product 2.) Destroy the product with a Certificate of Destruction to be provided to Aspen. Any returned product to Aspen will be put on MRB and dispositioned by Quality. The returned product will not be reprocessed.In order to determine effectiveness of the recall, any consignee that does not return the response form and/or product, will be contacted within three weeks of receipt of the Recall letter. Non-responders will be contacted a total of three (3) times in order to verify the effectiveness of the recall activity.Customers with questions were instructed to call 888-364-7004, ext 177 or 1-888-364-7004, ext. 156.For questions regarding this recall call 616-698-7100. |  |  |  | Worldwide Distribution - USA including LA, MA, NH, TX, FL, OH, NC, PA, CA, UT, TN, MI, CO, WY, OK, WA, MN, PA, IN, VA, IA, IL, MD, CT, DE, SC, NY, AR, MA, AZ, NJ, MO, GA, KS, WV, NE, NC, Puerto Rico, and Internationally to Canada, UAE, England, Italy, Belgium, Costa Rica, and Ecuador. | Terminated | November 25, 2014 | 250 |  | Not_Computer
Z-1437-2014 | 67691 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | GLOBALÂż Anchor Peg Glenoid Plus |  |  | March 28, 2014 | 2014 | DePuy Orthopaedics, Inc. | Urgent Device Correction notifications are being sent to USA distributors with follow-up to surgeons concerning updated surgical techniques stressing the possibility of the shoulder pin guide breaking and potentially being left in the patient. | Device Design | Beginning 3/28/12014 DePuy United States Distributors and DePuy affiliate contacts for Canada and International will be notified via email. The US Distributors will be responsible for adhering to the directions outlined in the field communication, including identifying the Surgeon users and notifying them either by mail or in person with the Surgeon Letter and providing access to revised surgical techniques. The notice emphasized technical points regarding the use of the GLOBAL APG+ 2.5mm Breakaway Guide Pin (singleuse shoulder instrument) that have been added to the surgical techniques. Risks were explained and surgical techniques revised were identified and available through the distributor. Distributors were asked to identify surgeons using the device, filling out the response form. For clinical questions from surgeons, please contact DePuy Orthopaedics, Inc.s Scientific InformationOffice at 18885542482 (MF; 8 a.m.  5 p.m. EDT).For product related questions, please contact your local DePuy Orthopaedics, Inc., Sales Consultant." For questions about recall information provided: Kim Earle, Recall Coordinator, 5743714917 (MF; 8 a.m.  5 p.m. EDT) |  | 21675 | 21,675 | USA Nationwide Distribution including Canada | Terminated | December 04, 2014 | 251 | 21675 | Not_Computer
Z-1436-2014 | 67724 | No | JPK | Hematology quality control mixture. | Hematology | Hematology | 510(K) Exempt | Boule Con-Diff Tri-Level |  |  | February 24, 2013 | 2013 | Clinical Diagnostic Solutions | Elevated MCV results on the 1311-682, 1311-683, and 1311-684 lots. | Under Investigation by firm | Clinical Diagnostic Solutions sent a Notification letter dated February 26, 2014 to affected customers. The letter identified the affected product, problem and the actions to be taken. For questions contact Technical Support at 1-800-453-3328. |  | 3540 | 3,540 | Worldwide Distribution: USA Nationwide and the countries of IVORY COAST, MEXICO, and Philippines. | Terminated | August 07, 2014 | 529 | 3540 | Not_Computer
Z-1439-2014 | 67857 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor |  |  | March 31, 2014 | 2014 | Philips Medical Systems, Inc. | A component on the HeartStart XL+ Therapy Board could malfunction potentially affecting the ability to deliver therapy | Reprocessing Controls | Philips Healthcare issued an Urgent-Medical Device Correction letter dated March 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer/user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem.A component on the HeartStart XL+ Therapy Board could malfunction potentially affecting the ability to deliver therapy.Please see the attached Field Safety Notice that provides information on how to identify affected devices and instructions on actions to be taken. Please follow the ACTION TO BE TAKEN BY CUSTOMER / USER section of the notice. To correct this issue Philips will provide a replacement therapy board to customers with affected units free of charge.This voluntary correction has been reported to the appropriate regulatory agencies.If you need any further information or support concerning this issue, please contact your local Philips representative at Philips representative contact details to be completed by the KM / country.Philips apologizes for any inconveniences caused by this problem.Further questions please call (978) 687-1501. |  | 19 | 19 units | Worldwide Distribution - US Distribution including the states of IN and WA., and the countries of : CHINA FRANCE GERMANY ITALY MALAYSIA PANAMA SAUDI ARABIA TURKEY and UNITED KINGDOM. | Terminated | October 21, 2015 | 569 | 19 | Not_Computer
Z-1438-2014 | 67862 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Philips HeartStart XL+ Defibrillator/Monitor with system |  |  | April 02, 2014 | 2014 | Philips Medical Systems, Inc. | Software communication failure may occur on the HeartStart XL+ locking the user out of clinical mode and possibly causing a delay in therapy. | Software design | Philips Healthcare sent an Urgent Medical Device Correction letter dated March 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer/user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem.A software communication failure may occur on the HeartStart XL+ locking the user out of clinical mode and possibly causing a delay in therapy.Please see the attached Field Safety Notice that provides information on how to identify affected devices and instructions on actions to be taken. Please follow the ACTION TO BE TAKEN BY CUSTOMER / USER section of the notice. To correct this issue Philips will provide a software upgrade to customers with affected units free of charge.This voluntary correction has been reported to the appropriate regulatory agency.If you need any further information or support concerning this issue, please contact your local Philips representative at 1-800-722-9377.Philips apologizes for any inconveniences caused by this problem. |  | 9872 | 9872 units | Worldwide Distribution - USA (nationwide) including the countries of: ARGENTINA AUSTRALIA AUSTRIA BELGIUM CANADA CHINA CZECH REPUBLIC DENMARK FRANCE GERMANY ITALY LATVIA NETHERLANDS NEW ZEALAND NORWAY PHILIPPINES PORTUGAL SOUTH KOREA SPAIN SWITZERLAND THAILAND TURKEY and UNITED KINGDOM. | Open, Classified |  |  | 9872 | Computer
Z-1441-2014 | 67752 | No | DSA | Patient transducer and electrode cable (including connector). | Cardiovascular | Cardiovascular | 510(k) | FAS-LOC 6 Foot Disposable Extension Cable with Safe Connect/Screw Down. |  |  | March 05, 2014 | 2014 | Remington Medical Inc. | Remington Medical discovered an error on the FL-601-97 IFU; part number IFU-RM-0154. Specifically, the polarity for the black chuck in the picture is wrong. It shows (+), but should show (-). | Labeling Change Control | Remington Medical Inc sent an Field Safety Notice dated March 5, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please be aware that the typographical error described above is included in the FL-601-97 Instructions for Use (RM-0154 Rev 05) provided with the lots that you have received as identified below. RMI has corrected the typographical error in the revised FL-601-97 Instructions for Use (RM-0154 Rev 5.1)which is attached to this communication. Please incorporate the information contained in the attachment with the existing FL-601-97 labeling (Instructions for Use) provided with the product. If you have distributed these products further, please forward copies of this notice and the revised Instructions for Use to anyrecipients as appropriate.If you have any questions regarding this communication, please contact RMIs sales representatives or customer service team at 800-989-0057, Monday - Friday, between the hours of 8 am - 4pm EST.Patient safety is the highest priority for Remington Medical, Inc. We appreciate your assistance with this matter and regret any inconvenience this may cause. |  | 41499 | 41,499 cables | USA (nationwide) Distribution including Canada | Open, Classified |  |  | 41499 | Not_Computer
Z-1444-2014 | 67714 | No | LGZ | Infusion pump. | General Hospital | General Hospital | 510(k) | Smiths Medical HOTLINEÂż Fluid Warming Set L-80. |  |  | March 17, 2014 | 2014 | Smiths Medical ASD, Inc. | Injection port of the Set may leak blood/ IV fluids from the septum of the injection port or the septum may become dislodged | Nonconforming Material/Component | Smiths Medical issued Urgent Field Safety Notice Recall on 3/17/14 via mail service. Distributors will be instructed to notify their customers. The Field Safety Notice will also include an Urgent Recall Notice Confirmation Form (Confirmation Form) that consignees will be instructed to send back to Smiths Medical for carrying out the action and tracking effectiveness.Accounts are to inspect inventory and return recalled product. If you should have any questions regarding this information, please contact Shirley Duggan,Quality System Engineer at 1-781-763-9335 or by E-Mail Shirley.duggan@smiths-medical.com. |  | 870 | 870 units | Worldwide Distribution - USA Nationwide in the states of: AU, CA, CL, CN, CZ, GT, ID, IN, IT, JP, LY, NZ, QA, SG, SK, TR, TW. | Terminated | June 09, 2015 | 449 | 394080 | Not_Computer
Z-1445-2014 | 67755 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Alcon LenSx Ophthalmic Laser, for use in cataract surgery for the creation of corneal cuts/incisions, anterior capsulotomy and laser phacofragmentation. |  |  | March 14, 2014 | 2014 | Alcon LenSx, Inc. | Alcon LenSx received reports of unexpected downward motion of the gantry. | Device Design | Alcon LenSx Inc advised customers of a Medical Device Correction being initiated for specific serial numbers of the LenSx Laser System for use in cataract surgery for the creation of corneal cuts/incisions, anterior capsulotomy and laser phacofragmentation.Alcon sent customer notification letters via mail during the week of March 17, 2014. The letter titled "Medical Device Correction", advised customers of the correction and provided: the product's description with serial number(s), description of the potential condition, advise on action to be taken by the user, transmission of the notice, and a contact reference person. The letter also contained a section for customers to sign and return the page of the Notice to Alcon. |  | 75 | 75 | Worldwide Distribution. | Terminated | October 16, 2014 | 216 | 75 | Not_Computer
Z-1448-2014 | 67675 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Pega Medical Inc., Fassier-Duval IM Telescopic System, male component (solid shaft), Part number: M040-SS-110; Used in Implant Catalogue number: FD-040(LS)-SS. orthopedic intramedullary rod. |  |  | February 07, 2014 | 2014 | Pega Medical Inc. | The Male Components in this lot are made of material with lower strength than manufacturing specifications. | Nonconforming Material/Component | Pega Medical sent an Urgent Field Safety Notice-Recall letter dated February 7, 2014, to their customers via email. The letter identified the affected product, the problem, and the action to be taken by the customer. Customers were instructed to identify the affected devicee, verify if it was used in a patient and place the affected device in quarantine immediately if not used. If not used, return the affected device to Pega Medical Inc. to the address provided. Pega's customer notification stated "If used, it is recommended to continue the normal patient follow up.Customers with questions should call (450) 688-5144 ext. 242 or email egarcia@pegamedical.com |  | 15 | Total of 5 units (2 units in the US and 3 units in Germany) | Worldwide Distribution - USA Nationwide in the states of Missouri, Texas, and the country of Germany. | Terminated | April 23, 2014 | 75 | 15 | Not_Computer
Z-1449-2014 | 67762 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Cobas c8000 Modular Analyzer Series Cobas c502, Part Number 5964067001, a clinical chemistry analyzer intended for the in vitro quantitative / qualitative determination of analytes in body fluids. |  |  | March 13, 2014 | 2014 | Roche Diagnostics Operations, Inc. | The Initial Cassette Volume Check (ICVC) is not activated on the cobas c 502 module of the cobas 8000 modular analyzer series when manually filled cobas c packs are loaded. This issue occurs with tests that use cobas c pack MULTI or empty pre-labeled cobas c packs. The ICVC feature allows the reagent probe to dive into the reagent cassette and confirm the reagent level in the bottles. This ensures | Software Manufacturing/Software Deployment | March 17, 2014 all consignees will be notified by Urgent Medical Device Correction sent via UPS Ground (receipt required). The notice identified the issue, root cause, affected assays with clinical significance. Consignees were instructed to follow the steps listed in the Urgent Medical Device Correction to correctly activate the ICVC and verify the correction. If your system is not connected to the Internet through the cobas link, and you do not have the option to run the affected reagents on an alternate system (i.e., cobas c 311 or cobas c 501 analyzer), please contact the Roche Support Network Customer Support Center at 1-800-428-2336 to have an update scheduled via a Roche Field Representative. If your system is connected to the Internet through the cobas link, follow the steps below:" Remove all reagent cobas c packs on board your cobas c 502 module for the assays listed on pages 1-2." Dispose of these reagent cobas c packs according to your local waste management ordinances." Load new cobas c packs for the affected assays, and visually verify the ICVC is performed." Complete the attached fax form (5326-00-0314) and fax it to 1-888-276-6166." Please consult with the physician or pathologist at your facility to determine specific clinical implications for your patients." File this Urgent Medical Device Correction (UMDC) for future reference. |  | 76 | 76 units | USA Nationwide Distribution including San Juan PR | Terminated |  |  | 76 | Not_Computer
Z-1451-2014 | 67876 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Sit/stand Rolling Walker, SKU/Item Number 66838. Intended for use by those with difficulty walking and weighing 225 lbs. or less. |  |  | March 27, 2014 | 2014 | Central Purchasing LLC DBA Harbor Freight Tools | Harbor Freight Tools intiaited a recall of Sit/Stand Rolling Walker (SKU/Item Number 66838, sold June 2009 to March 2014) because, over time, there is the possibility that a walker wheel may come off this product following repeated use without tightening the wheel bolt to the frame, presenting a potential hazard for the user. | Under Investigation by firm | Recall communication to web site customers who purchased the product, notifying them of recall and availability of full refund for product.Store posting notifying retail customers of the recall and availability of fullrefund for product (with a similar posting provided on the HFT web site). Consumers are informed of the recall, product info. with photo, item number, reason for recall, hazard, instructed to stop using these walkers immediately, and contact information provided telephone number, and e-mail address. Consumer letters also provided with a Merchandise Return Label.On 04/02/14, an updated consumer letter issued. |  | 43621 | 43,621 | Nationwide Distribution. | Terminated | August 05, 2015 | 496 | 43621 | Not_Computer
Z-1447-2014 | 67821 | orthosis, limb brace | IQI | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | K-Lite Knee 2 Stays with Visco, Model No. 122015. |  |  | March 14, 2014 | 2014 | Breg Inc | Breg is recalling the Economy Elastic Knee Sleeve Open and K-Lite Knee 2 Stays with Visco because they contain latex but were not labeled with the required caution label. | Process change control | Breg representatives informed customers of the recall by telephone on March 14, 2014 and then a customer notification letter was sent to customers via email onMarch 17, 2014. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the return response form. Customers with questions are instructed to contact Breg Customer Care at (800) 321-0607. |  | 2 | 2 units | USA Nationwide Worldwide in the states of: TX, KY, IN, and FL. | Terminated | April 14, 2014 | 31 | 22 | Not_Computer
Z-1440-2014 | 67677 | Yes | BTL | Powered emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Draeger Oxylog 3000 plus ventilators, emergency and transport ventilator. Part number: 5704833 |  |  | March 05, 2014 | 2014 | Draeger Medical Systems, Inc. | The potentiometers (knobs) for setting the ventilation parameters may have contact interrupted and the device discontinues ventilation, gives an audible alarm and displays the error message Poti unplugged. | Process control | Drager sent an Urgent Medical Device Recall letter dated February 2014 to affected customers. The letter identified the affected device, problem and actions to be taken. The notification directed customers to perform manual ventilation in the event of a failure. Drager plans to repair the affected devices free of charge. If you have any questions regarding the Recall letter please call 1-800-543-5047 (press 1 at the prompt, then 2, then 32349). For questions regarding the operation and/or servicing of your DrÂżger ventilator please contact DrÂżgerService Technical Support at 1-800-543-5047 (press 4 at the prompt). |  | 3 | 3 | USA Nationwide Distribution in the states of WI and NY. | Terminated | October 31, 2014 | 240 | 3 | Not_Computer
Z-1450-2014 | 67804 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | 400003 Medcare Lift with Scale |  |  | May 10, 2013 | 2013 | ErgoSafe Product LLC dba Prism Medical Services | MedCare Products has determined that a number of their Floor Lifts manufactured from 9/1/2010 to 11/7/2012 do not meet MedCare's high standards. The original design includes a pin that over time and years of use may stop performing as designed and may break. | Nonconforming Material/Component | Medcare sent a Notification letter dated May 10th, 2013 to affected customers. The letter described the affected product, problem, actions to be taken and provided a replacement assembly. It instructed consignees to replace the old assembly with the new one included in the package. For questions they can contact Todd Perrine in their Technical Support Department at 952-300-2101. |  | 514 | 514 | USA Nationwide Distribution in the states of: CA, CO, IL, IN, IA, KS, KY, LA, ME, MD, MI, MN, MS, NO, NE,NH,NY, ND,OH, OK, OR, PA , SD, TX, WI. | Terminated | November 20, 2014 | 559 | 514 | Not_Computer
Z-1452-2014 | 67808 | No | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Reusable latex breathing bags with multiple part numbers and sold as part of test lung part number and reusable anaesthesia sets. |  |  | March 25, 2014 | 2014 | Draeger Medical, Inc. | Reusable latex breathing bags with various part numbers and sold as part of test lung and reusable anaesthesia sets were recalled when Draeger became aware of an incident where a user suffered an acute allergic reaction after using a DrÂżger latex test lung. The user did not suffer permanent injury. The breathing bag that belongs to the test lung and other latex breathing bags are labelled NR .A | Labeling design | A recall notification letter, dated March 2014, was sent to wholesale distributors and end users. |  | 433 | 433 | USA Nationwide Distribution in the states of: MI, SC, CO, OH, NE, WI, NC, SD, AZ, TX, MA, WA, PA, KY, MN, GA, OK, NY, IN, VA, UT, MT, ME,FL, CA, MD,MO, IA, TN, MD, WV, IL,NJ, CT, and AR. | Terminated | October 08, 2014 | 197 | 433 | Not_Computer
Z-1455-2014 | 67797 | Yes | FJK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Gambro Cartridge Blood Set, blood transport system for hemodialysis, Model Number(s): 101025 (003410510) - Cartridge set, STND PRM LN and 103401 (003414500) - Cartridge set, PRM_LN_INJ_PT. The Gambro Cartridge Sets are single use sterile tubing sets intended to provide extracorporeal blood transport circuit for hemodialysis treatments for Gambro Phoenix and COBE Centrisystem 3 (and 3+) Dialysis Delivery Systems. |  |  | December 09, 2013 | 2013 | Gambro Renal Products, Incorporated | Occluded heparin tubing events preventing anticoagulation dosing on the Gambro Cartridge Blood Sets. | Process control | On 12/9/13 and 3/12/14, a field safety notice sent via UPS overnight delivery informing customers of issue and if any customer wants to return product, a hold at distribution centers and manufacturer quarantine areas. Upon reconciliation of returned product a destruction disposition with certification. |  | 4597875 | 306,525 sets (20435 boxes of 15 sets) | Worldwide Distribution-USA (nationwide) and the countries of Canada, Mexico, and Colombia. | Terminated | December 22, 2015 | 743 | 4597875 | Not_Computer
Z-1442-2014 | 67781 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | SpaceLabs Arkon Anesthesia Delivery System with Version 2.0 Software installed. This system includes a ventilator. |  |  | March 10, 2014 | 2014 | Del Mar Reynolds Medical, Ltd. | The System may go into a controlled failed state if a series of parameters on the user interface are activated in a specific sequence. | Software in the Use Environment | Spacelabs Healthcare sent an Urgent Medical Device Correction letter, dated 11 March 2014, to consignees. The letter identified the product, the problem, and the actions to be taken by the customer. The letter stated that Spacelabs Field Service personnel will be contacting consignees to schedule a convenient time to install, at no cost, a software upgrade.Consignees with questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 16 | 16 units distributed in the US | USA Nationwide Distribution in the states of North Carolina and South Carolina | Terminated | August 20, 2014 | 163 | 16 | Computer
Z-1456-2014 | 67499 | No | GXL | Skullplate screwdriver. | Neurology | Neurology | 510(K) Exempt | Synthes CMF Battery Powered Driver, Part no. 305.840. Used to place screws in the cranialfacial bones, and to drill and/or tap holes to prepare cranialfacial bones for screw insertion. |  |  | January 30, 2014 | 2014 | Synthes, Inc. | The CMF Battery Powered Driver graphic case contains outlines of the Battery Powered Driver, taps, screwdriver blades, and a screw holding sleeve. However, the device is not intended to be used in conjunction with a tap for tapping holes in a patient's skull, upper and/or lower jaw. Therefore some outlines in the graphic case suggest a potential use of this device that is considered off-label. | Labeling design | Synthes sent an Urgent Notice: Field Safety Notification letter dated January 30, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please take the following actions;Please see the Battery Powered Driver Users Manual for on-label procedures. You can access an electronic version of the Battery Powered Driver Users Manual by accessing the following : linlc http://www.synthes.com/pages/default.aspx. Click on North America Tab at the top of the screen, and then type Battery Powered Driver Users Manual in the search bar to download the Users Manual Forward this Field Safety Notification to anyone in your facility that needs to be informed. If the graphic case has been forwarded to another facility, contact that facility. " Maintain a copy of this notice.If you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. Thank you for your attention to this issue. |  | 20081 | 20081 | Nationwide Distribution | Open, Classified |  |  | 20081 | Not_Computer
Z-2250-2012 | 62856 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Newport HT70 and HT70 Plus Ventilators, Model Number: HT70 and HT70 Plus. |  |  | August 03, 2012 | 2012 | Newport Medical Instruments Inc | May emit a continuous high priority alarm and the ventilator may stop ventilating, due to a component failure on the control board. | Nonconforming Material/Component | Newport Medical sent an Urgent Field Correction Notice letter dated August 3, 2012 to all affected customers, via DHL and FedEx courier services. The letter identified the affected product, reason for recall and actions to be taken. Customers were instructed to immediately stop using the affected ventilator and retrieve the replacement kit provided free of charge. In addition, the letter states that if the product was further distributed, to contact their customers and notify them of the recall. Customers were asked to complete and return the enclosed Replacement form as replacements are completed. For questions contact Newport Medical Technical Support Department at 800-451-3111, extension 282, for U.S. customers; 714-427-5811, extension 500, for international customers.". |  | 72 | 72 | Worldwide distribution: USA (nationwide) and countries of: Brazil, Chile, Egypt, France, Hong Kong, India, Japan, Norway, Philippines and South Africa. | Terminated | April 14, 2014 | 619 | 72 | Not_Computer
Z-1457-2014 | 67560 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The Synthes Hohmann Retractor, 43 mm Long Shank, Angled 235 mm |  |  | February 11, 2014 | 2014 | Synthes, Inc. | The Synthes Hohmann Retractor was mis-etched on the product and package as part number 399.24 instead of part number 399.22. | Error in labeling | Synthes sent an Urgent Notice: Medical Device Recall letter dated February 11, 2014, to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to Complete the attached Verification Section form. For questions call 610-719-5450. |  | 20 | 20 | Worldwide Distribution - USA Nationwide Distribution in the states of: TN, MD, NY, IN, NJ, WA, VA,and the country of Canada. | Terminated | September 10, 2015 | 576 | 20 | Not_Computer
Z-1458-2014 | 67609 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | BIOMET 3i Endosseous Dental Implant. |  |  | August 29, 2013 | 2013 | Biomet 3i, LLC | Due to residual machining fluid identified by discoloration (darker in appearance) on the external dental implant surface. | Process control | Biomet 3i sent an Urgent Medical Device Notice dated September 9, 2013, to all affected customers. The notice identified the product, problem and actions to be taken. Each customer was instructed to check their respective inventory for the affected product and to return any unused quantities along with the attached Recall Response Form. Biomet 3i will issue replacement product at no charge. For questions or concerns, call Biomet 3i at 1-800-342-5454. |  | 34744 | 34,744 | Worldwide Distribution - USA and countries of United Arab Emirates, Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, Chile, Colombia, Costa Rica, Germany, Dominican Republic, Spain, France, Great Britain (UK), Greece, Hong Kong, Ireland, Israel, India, Italy, Japan, Korea (South), Lebanon, Liechtenstein, Lithuania, Mexico, Netherlands, Poland, Portugal, Paraguay, Saudi Arabia, Sweden, Singapore, Turkey, Taiwan, Uruguay, Vietnam and South Africa. | Terminated | October 17, 2014 | 414 | 34744 | Not_Computer
Z-1464-2014 | 67763 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity Core 128, Computed Tomography X-ray systems, Philips Medical Systems (Cleveland), Inc. |  |  | March 06, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | A customer reported that after upgrading to software version 3.5.5 from 3.5.4, reconstructions would intermittently not complete. | Software in the Use Environment | .The firm, Phillips, sent an " URGENT ÂżÂż Medical Device Correction" letter dated2013 March 06, to its costumers . The letter disc ribes the product, problem and actions to be taken. The costumers were instructed to Verify that all planned image reconstructions are completed, and Perform off-line reconstructions if expected images are missing. The firm will release A Field Change Order (FCO) to correct the issue with software version release 3.5.7.and A Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter Site ID or follow the prompts). |  | 116 | 116 units | Worldwide Distribution:US ( Nationwide) Including stages of MI, OR, PA, TX &amp; WV; and Countries of: Australia, Bangladesh, Belgium, Brazil, Bulgaria, China, France, Germany, India, Indonesia, Ireland, Japan, Nicaragua, Oman, Palestine, Russia, South Korea, Spain, Switzerland, Thailand, Turkey &amp; United Kingdom. | Terminated | November 05, 2014 | 244 | 116 | Computer
Z-1459-2014 | 67774 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Extension Set, Model No. 20029E, intravascular administration set. |  |  | March 19, 2014 | 2014 | CareFusion 303, Inc. | Disconnections and leakages at the connection of the smallbore tubing and the SmartSite Y-port. | Nonconforming Material/Component | CareFusion sent an Urgent Medical Device Recall Notification letter dated March 2014 to all customers. The letter informs the customers of the problems identified, potential risk and the actions to be taken. Customers are instructed to complete and return to CareFusion the enclosed recall response card. Customers with recall related questions are instructed to contact CareFusion Support Center at (888) 562-6018. Customers with adverse event reports are instructed to contact Customer Advocacy at (888) 812-3266 or email customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris System are instructed to contact Technical Support at (888) 812-3299. |  | 7700 | 7,700 units | USA Nationwide Distribution in the states of: AR, AZ, CA, CO, FL, ID, MD, MI, NC, NV, OK, OR, PA, TX, VA, WI | Terminated | June 06, 2015 | 444 | 7700 | Not_Computer
Z-1460-2014 | 67780 | No | IZZ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | AXIOM Vertix MD Trauma systems |  |  | March 05, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential issue and possible hazard to patients when using the AXIOM Vertix MD Trauma systems. In rare cases, steel ropes inside the lift column of the system can be defective without triggering the safety lock, which can result in the U-arm dropping down unexpectedly during movement in vertical direction, potentially causing serious injury. | Component design/selection | Siemens sent a Safety Advisory Notice dated March 5, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised as a first check it is strongly recommended for the users to check whether metallic dust or rubbed off parts of metal are visible underneath the lifting column or around the system. If this is the case, it is strongly recommended to immediately stop using the Vertix MD Trauma system and call the local Siemens service. To avoid any risk until the implementation of the modification mentioned below, it is furthermore strongly recommended to perform up/down movements of the lifting column not directly above the patient, but complete the vertical movement beside the patient and then move the system horizontally above the patient.We appreciate your understanding and cooperation with this Safety Advisory Notice and ask you to immediately instruct your personnel accordingly. Please ensure that this Safety Advisory Notice is placed in the system's instructions for use until the update has been installed. If you have sold or otherwise disposed of this equipment and it is no under your control, we kindly ask that you forward this Safety Advisory Notice to the new user of the equipment. Please also inform us about the new owner of the equipment.We apologize for any inconvenience this may cause. Further questions please call (610) 219-6300. |  | 2 | 2 | US Distribution including MO and OH. | Terminated | March 16, 2015 | 376 | 2 | Not_Computer
Z-1453-2014 | 67893 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Pipeline Embolization Device (PED). Used endovascular treatment of adults with intracranial aneurysms in the internal carotid artery. |  |  | April 01, 2014 | 2014 | Micro Therapeutics Inc, Dba Ev3 Neurovascular | Potential for the PTFE (polytetrafluoroethylene) coating to delaminate and detach from the delivery wire. | Process change control | A customer notification letter was sent on 4/1/14 to all customers who purchased the Pipeline Embolization Device and the Alligator Retrieval Device. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the customer response form and fax the completed form to Covidien at (763) 591-3519. Customers with questions are instructed to contact Dr. Mark A. Turco at (202) 310-5120.A press release was sent on 4/11/14 to Business Wire Newsroom to inform the public about the recall from Covidien. The recall is regarding certain lots of its Pipeline Embolization Device and Alligator Retrieval Device where the polytetrafluoroethylene (PTFE) coating applied to the delivery wire could delaminate adn detach from the devices. Consumers are instructed to report any adverse reactions or quality problems experienced with the use of the product to the FDA's MedWatch Adverse Event Reporting program either online, by regular mail or by fax. Consumers with may also contact Covidien Customer Service at (800) 716-6700 between 7am to 7pm (central) or email at customerserviceus@covidien.com. |  | 32 | 32 units total (23 units in US) | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Germany, France, UK, Turkey, Argentina, Canada, Colombia, Mexico, , Austria, Belgium, Denmark, Hungary, Italy, Netherlands, Norway, Poland, Portugal, Russia, Slovakia, Slovenia, Spain, Switzerland, and UAE. | Terminated | October 31, 2014 | 213 | 32 | Not_Computer
Z-1462-2014 | 67867 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | The OEC 9900 Elite mobile fluoroscopy system (image intensified fluoroscopic x-ray) is designed to provide fluoroscopic and spot-film images of the patient during diagnostic surgical and interventional procedure. |  |  | March 31, 2014 | 2014 | GE OEC Medical Systems, Inc | GE Healthcare has identified a potential safety issue related to the screws that attach the workstation monitor spring arm to the extension arm on certain OEC 9900 systems, causing the spring arm and monitor to fall forward. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated March 31, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Use caution when positioning or extending the workstation articulating arm. In addition, refrain from positioning staff or patients directly underneath or in front of a workstation monitor.Your local GE Healthcare Field Engineer will contact you to coordinate the replacement of the screws, and ensure that your system is operating safely and meets all specifications. These activities will be performed at no cost to you.Should you have any questions or concerns regarding these corrections, please do not hesitate to contact the service team for further information at 800-874-7378 option 8.Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately.UPDATE: On June 20, 2014. GE Healthcare Surgery began distributing an Urgent Medical Device Correction Letter to customers with affected units. The letter makes clear that this field correction applies to all user sites with the listed products, regardless of whether service work was performed on the monitor extension arm since 3/31/2014. |  | 453 | 453 units | Worldwide Distribution - USA, and the countries of Australia, Argentina, Canada, China, Cambodia, Columbia, Japan, Ecuador, Japan, Korea, Hong Kong, India, Israel, Libya, Mexico, Nigeria, Peru, Russia, Thailand, Taiwan, Saudi Arabia and Singapore. | Terminated | October 26, 2015 | 574 | 453 | Not_Computer
Z-1461-2014 | 67837 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT 16, Brilliance CT 64, Brilliance CT Big Bore Oncology, Brilliance CT Big Bore Radiology, Brilliance iCT, Ingenuity CT, Ingenuity Core, Ingenuity Core 128 and Ingenuity Flex scanners are Computed Tomography X-ray systems. |  |  | March 27, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Block assembly-lock stop subframe of service latch was not manufactured to engineering specifications from the supplier, causing premature fracture of the threaded rod on the patient support service latch. As a result, the table top can become free floating due to the disengaged service latch in the table tops subframe. | Nonconforming Material/Component | A Field Safety Notice was sent to all US consignees via certified mail. Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule a time to implement the Field Change Order. |  | 135 | 135 | Nationwide Distribution - USA including the states of CA, FL, IL, IN, KY, MD,SC, and TX. | Terminated | February 02, 2016 | 677 | 135 | Not_Computer
Z-1463-2014 | 67835 | No | ONG | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | FRAXEL DUAL 1550/1927 Laser System, (FRAXEL SR 1500 DUAL Laser System), for dermatological procedures. |  |  | March 20, 2014 | 2014 | Solta Medical, Inc. | Obsolete versions of the Laser Treatment Settings card for the Fraxel DUAL 1550/1927 Laser system may have contained incorrect indication for use for treatment of Melasma. | Labeling Change Control | Solta Medical sent an Correction Important Product letter dated March 20, 2014 to affected customer. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the enclosed response form. |  | 1077 | 1077 units | Worldwide distribution | Terminated | April 27, 2015 | 403 | 1077 | Computer
Z-1465-2014 | 67925 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | Terumo Cardiovascular Procedure Kit- X-Coated Pump and Table Pack with FX25RE |  |  | March 28, 2014 | 2014 | Terumo Cardiovascular Systems Corp | Incorrect customer notification label applied to the Tyvek lid of this specific lot of TerumoÂż Cardiovascular Procedure Kits | Labeling mix-ups | Terumo CVS contacted the affected user on 3/28/14 by phone via phone script to discontinue use and return all affected units in inventory. Terumo CVS will replace or issue credit for returned product. |  | 37 | 37 units | CA | Terminated | December 12, 2014 | 259 | 37 | Not_Computer
Z-1466-2014 | 67983 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Fem-Flex II Arterial and femoral Cannulae, arterial cannula sizes 16, 18 and 20 French, models: FEMII016A, FEMII016AS, FEMII018A, FEMII018AS, FEMII020A, and FEMII020AS. |  |  | April 09, 2014 | 2014 | Edwards Lifesciences, LLC | Potential for separation of dilator tips. | Nonconforming Material/Component | Customers were notified via letter on 4/4/14 to return affected products. |  | 13115 | 13,115 | Nationwide, including foreign, govt/VA/Military consignees. | Open, Classified |  |  | 42000 | Not_Computer
Z-1469-2014 | 67569 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Brainlab, Spine &amp; Trauma 3D 2.0, Navigation Software. An intraoperative image-guided localization system to enable minimally invasive surgery. Catalog Number: 22264 |  |  | November 18, 2013 | 2013 | Brainlab AG | The "fluoro match registration" function of the Navigation Software Spine &amp; Trauma 3D 2.0 allows the user to intraoperatively match e.g., CT data sets to the current patient anatomy visible on fluoroscopic images, used by the navigation software during spinal surgeries to display the position and orientation of instruments. For this registration function in combination with a digitally integrated | Software design | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 18, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter included instructions for the modification of C-arm configurations: 1) If possible, Use analog image transfer between C-arm and Brainlab Navigation System and to consult with the local Brainlab Customer Support Representative for this modification, and 2) If possible, adjust the image size of the C-arm to 8 bit and to follow the instructions provided with the C-arm and if required to consult the C-arm manufacturer. For measures when performing a Fluoro match registration, the following instructions are provided to reduce the probability that the algorithm results in an inaccurate match: 1) perform the preoperative CT scans according to the descriptions in the Brainlab scan protocol, 2) set the threshold so that a smooth bone surface is displayed and the quality of the chosen bone threshold influences the accuracy of the match result (refer to section "Bone Threshold" in the Software User Guide"), 3) use the Radiolucent Spine Reference Clamp instead of the Spine Reference X-Clamp, 4) make sure that no additional objects (e.g. wires, retractors, instruments, etc.) are within the view of the C-arm. Additional reminders in the letter include: 1) always verify registration accuracy by holding the pointer or instrument tip to at least three anatomical landmarks and verifying their position in the software, 2) verify that the registration is at the correct level on the patient and data set, and 3) accuracy must be checked on the treated bone surface. Brainlab will provide a software update with the issue solved to affected customers tentatively by January 2014. Customers with questions can contact the Customer Hotline at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 133 | 133 units | Worldwide Distribution - US Distribution including the states of AZ, CA,GA, IL, IN, LA, MA, MI, NY, NC, OH, PA, TN, TX and WI and the countries of Australia, Austria, Belgium, Brazil, Chile, Czech Republic, France, Germany, Italy, Japan, Malaysia, Netherlands, Norway, Philippines, Poland, Russian Federation, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey and United Kingdom. | Open, Classified |  |  | 133 | Computer
Z-1468-2014 | 67793 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | CODMAN CERTAS Programmable Valves, central nervous system shunt, models as follows: |  |  | December 13, 2013 | 2013 | Codman &amp; Shurtleff, Inc. | Ensuring all customers have the tools to assist in verifying the setting of the valve on x-ray images and the location and orientation of implanted CODMAN CERTAS Programmable Valves. | Process control | Codman Neuro issued notification packages dated 12/13/13 to Domestic Hospitals and international accounts. Specific letters and associated inclusion of the additional aids, include A) a valve setting verification guide with instructions for use with x-ray images, and B) a plastic valve cut out overlay). A response reply form was included to be completed and returned to confirm effectiveness of the notification. |  | 9445 | 9,445 units | Nationwide Distribution. | Terminated | April 21, 2016 | 860 | 9445 | Not_Computer
Z-1470-2014 | 67788 | No | KLT | Neuroleptic drugs radioreceptor assay test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products CRBM Slides, REF 889 2382, IVD -- Ortho-Clinical Diagnostics, Inc. For in vitro diagnostic use only. VITROS Chemistry Products CRBM Slides quantitatively measure carbamazepine (CRBM) concentration in serum and plasma using VITROS 250/350/950/5,1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System. |  |  | March 05, 2014 | 2014 | Ortho-Clinical Diagnostics | The firm identified a potential for biased carbamazepine (CRBM) results to be generated when using VITROS CRBM Slides, Lot 3920-0080-8403. | Under Investigation by firm | The firm, Ortho Clinical Diagnostics, sent an "URGENT PRODUCT CORRECTION NOTIFICATION" letter (Ref. CL14-086; dated March 5, 2014) via FedEx Overnight Courier and/or US Postal Service Priority Mail (for PO Boxes only) to its US consignees (OCD Direct, Drop-Ship and US Federal Government). In addition, OCD proactively communicated with the 3 affected customers via phone to inform them of the potential issue and expedite product replacement. The letter described the product, problem and actions to be taken. The customers were instructed to immediately discontinue using and discard all remaining inventory of VITROS CRBM Slides, Lot 3920-0800-8403; review previously reported results using this lot; discuss any concerns regarding previously reported results with their Laboratory Medical Director or with the requesting physician; complete the attached Confirmation of Receipt form no later than March 14, 2014; and forward the information in this notification if you have distributed this product outside of your facility. If you have any additional questions, please contact Customer Technical Services at 1-800-421-3311. |  | 20 | 20 units | US Distribution to states of: IN, MD, and TN. | Open, Classified |  |  | 20 | Not_Computer
Z-1475-2014 | 67820 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | ZYM B Reagent (REF 70493). |  |  | March 06, 2014 | 2014 | BioMerieux SA | bioMerieux identified a visual defect and activity issue on the ZYM B reagent (REF 70493) and FB reagent (REF 70562). The activity issue lead to a false negative result of some Rapid ID 32 strips and API biochemical tests using the ZYM B reagent or FB reagent | Other | Letters were issued to customers on March 10, 2014. The letter contained instructions to be used if they still were using the product. The letter also included a report form which was to be returned. |  | 33853 | 33,853 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico, and the states of AL, AK, AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WV, WA, WI, and WY and the country of Algeria, Angola, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Benin, Bolivia, Bosnia, Brazil, Brunel, Bulgaria, Burkina Faso, Cambodia, Cameroon, Canada, Central African Republic, Chile, China, Colombia, Costa Rica, Croatia, Czech Republic, Dutch Antilles, Ecuador, Egypt, El Salvador, Estonia, Erythrea, France, Gabon, Georgia, Germany, Ghana, Greece, Guatemala, Haiti, Herzegovina, Honduras, Hungary, India, Indonesia, Iraq, Israel, Italy, Japan, Jamaica, Jordan, Kazakhstan, Kenya, Korea, Kosovo, Kuwait, Latvia, Lebanon, Lithuania, Macedonia, Malaysia, Mali, Mexico, Moldavia, Morocco, Mozambique, Myanmar, Netherlands, Nigeria, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovenia, South Africa, Spain, Sweden, Thailand, Trinidad &amp; Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Vietnam. | Terminated | March 14, 2017 | 1104 | 33853 | Not_Computer
Z-1473-2014 | 67777 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Oxoid Legionella Latex Test, DR0800M, containing 50 tests per box. The box contains a variety of components, one of which is Oxoid L. pneumophila Group 2-14 Latex Reagent, DR0802. The responsible firm on the label is Oxoid Ltd., Basingstoke, Hants, England. |  |  | March 17, 2014 | 2014 | Remel Inc | A reagent contained within the product may return false negative results. | Under Investigation by firm | The recalling firm initiated a recall by issuing letters dated 3/17/14 via regular mail to their customers. |  | 72 | 72/50-test boxes | Distribution was made to CA, FL, IL, MA, ME, NJ, PA, TN, TX, VA, and VT. There was no foreign/military/government distribution. | Terminated | December 11, 2014 | 269 | 74 | Not_Computer
Z-1472-2014 | 67803 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Fassier-Duval IM Telescopic System. Orthopedic intramedullary rod implant. Catalogue number FD-040(L)-SS. |  |  | March 05, 2014 | 2014 | Pega Medical Inc. | The outer package label contained the incorrect length of the device. There is a label on the outer package and an additional label on the product contained in the package. | Error in labeling | There is an error on the outside label of the implant. The length was mistaken on the label. It was indicated 245mm instead of 338mm. It has to be clarified that the Fassier-Duval implant is provided with two labels, one that is outside the tube used for packaging and one inside the tube. |  | 48 | 48 | Worldwide distribution, USA nationwide, Greece, United Arab Emirates, Peru, Canada, Israel, United Kingdom, France, Germany, Colombia, Mexico, Singapore, Lithuania, Denmark, and Spain. | Terminated | May 19, 2014 | 75 | 48 | Not_Computer
Z-1481-2014 | 67535 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RaySearch Ray Station Radiation Therapy Treatment Planning System, Stand-alone software. Versions 2.0, 2.5, 3.0, 3.5 and 4.0. |  |  | February 14, 2014 | 2014 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories has recalled "RaySearch Ray Station 2.0, 2.5, 3.0, 3.5 and 4.0" software due to dose miscalculations for bolus/external/fixation support structures on CT-image. | Software design | RaySearch Laboratories sent a Field Safety Notice, Medical Device Correction letter dated February 14, 2014 to their customers via email. The letter identified the affected product, problem and actions to be taken by the user. Customers with questions may contact Freddie Cardel, Director of Customer Support at 1-877-778-3849 or via email at freddie.cardel@raysearchlabs.com. |  | 251 | 251 units (148 units domestically &amp; 103 units internationally) | Worldwide Distribution - USA Nationwide in the states of: AZ, CA, CT, CO, DE, FL, IL, NJ, NY, MA, MD, MS, OH, PA, TN, VA and WA. and the countries of Australia, New Zealand, Belgium, New South Wales, Republic of Korea, Italy, Germany, China, France, Switzerland, Netherlands, Spain, Austria, United Kingdom and Norway. | Open, Classified |  |  | 251 | Computer
Z-1480-2014 | 67913 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery MR450, Discovery MR750, Optima MR450w, Optima MR 450w GEM, Discovery MR750w, and Discovery MR750w GEM products. Magnetic Resonance Diagnostic Device. |  |  | April 02, 2014 | 2014 | GE Healthcare, LLC | Potential safety issue with gradient coil electromechanical connections associated with GE MR Products.Gradient coil electromechanical connections may loosen causing an increase in resistance. The increased resistance can lead to increasing heat during scanning which, in turn, can cause material in the Rear Endbell enclosure to overheat, generate smoke in the scan room, and cause some localize | Process change control | Consignees were sent on 4/2/2014 a GE Healthcare "Urgent Medical Device Correction" letter dated April 2, 2014. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Mangers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product correction and Contact Information (Fore questions, contact your local service representative ). |  | 1546 | 1,546 (704 US, 842 OUS). | Worldwide Distribution-US (nationwide) except the states of ND and RI, including DC, GU, PR, and the countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA,AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BRAZIL, CANADA, CAYMAN ISLANDS, CHILE, CHINA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, HONG KONG, HUNGARY, INDIA,INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, KOREA, KUWAIT, LEBANON, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, POLAND, PUERTO RICO, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, UZBEKISTAN, VENEZUELA. | Terminated | April 08, 2015 | 371 | 1546 | Not_Computer
Z-1479-2014 | 67722 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | LifeShield LATEX-FREE HemoSet 100 mL Burette I.V. PlumSet, Nonvented, 105 inch with Prepierced Reseals, OPTION-LOK and Inline Dual Channel Cassette. |  |  | February 25, 2014 | 2014 | Hospira Inc. | Hospira identified an incorrect lower lid (set component) had been supplied and was being used during the manufacturing process of list number 11241-03, a HemoSet containing a burette and blood filter, intended for infusion of blood/blood products. In a gravity delivery, the correct lower lid dispenses 15 drops per mL and the incorrect lower lid found dispenses 10 drops per mL. | Mixed-up of materials/components | Hospira sent a URGENT MEDICAL DEVICE RECALL letter dated February 25, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. 1) check their inventory and immediately quarantine any affected product; 2) complete the attached reply form and return it to the fax number or e-mail address on the form; 3) inform potential users within the direct account's organization of this product of the recall; 4) return affected products to Stericycle using labels provided with the notification; and 5) follow the instructions on the return label for returning products. Additional labels and guidance on returning products can be obtained by calling Stericycle at 1-888-240-4282 (Monday - Friday, 8:00 AM - 5:00 PM ET). If the recalled products were further distributed, the recall notification documents should be forwarded to those accounts. Those accounts can contact Stericycle at 1-888-240-4282 (Monday - Friday, 8:00 AM - 5:00 PM ET) to obtain a reply form and guidance on the return process. |  | 3893 | 3,893 sets | US Distribution including the states of FL, IA,IN, KS, KY, ME, MO, MT, PA, TN, TX, UT, VA and WA. | Open, Classified |  |  | 3893 | Not_Computer
Z-1485-2014 | 68041 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | Merit Laureate Hydrophilic Guide Wire, Catalog No. LWSTFS35260EX, Straight Tip Stiff Shaft, 0.035" (0.89mm), 260 cm (102"). |  |  | April 14, 2014 | 2014 | Merit Medical Systems, Inc. | Discrepancy between the carton and unit labeling for the Merit Laureate Hydrophilic 0.035" 260 cm Guide Wire, catalog number LWSTFS35260EX, which has a straight tip and a stiff shaft. This lot was partially mixed with another lot of Laureate0.035" 260 cm Guide Wire, catalog number LWSTDA35260EX, which has an angled tip, standard shaft and Chinese unit labeling. | Process control | Customers were notified on 4/14/14 via certified letter with instructions to identify product, discontinue use, and return products. |  | 173 | 173 | Worldwide Distribution: USA Nationwide and China. No military/govt/VA consignees. | Terminated | July 17, 2014 | 94 | 359 | Not_Computer
Z-1487-2014 | 67959 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | iLab Ultrasound Imaging System; iLab 3.0 (Polaris 1.0) software. Used for ultrasound examinations of intravascular pathology. |  |  | April 03, 2014 | 2014 | Boston Scientific Corporation | Specific iLab Ultrasound Imaging Systems using newest 3.0 (Polaris 1.0) software change the pullback speed from the default setting of 0.5 mm/sec to 1.0 mm/sec, the Longview distance measurement may display an inaccurate value that is double the actual distance. | Nonconforming Material/Component | Boston Scientific sent an Urgent Medical Device Field Correction letter to all affected customers via hand delivery by BSC sales representatives. The letter identifed the affected product, problem and actions to be taken. Customers were asked to immediately post the Field Correction information in a visible location near the system(s) to ensure it is easily accessible to all users of the device. For question contact your local Sales Representative. |  | 9 | 9 units | USA Nationwide Distribution in the states of NY, NC, DC, FL, PA | Terminated | March 18, 2015 | 349 | 9 | Computer
Z-1471-2014 | 67466 | No | DSQ |  |  | Cardiovascular | PMA | Heartware HVAD Pump Implant Kit (Heartware Ventricular Assist Device) Catalog Numbers: 1100, 1101, 1102, 1103, 1104, and 1205. For use as a bridge-to-cardiac transplantation in patients who are at risk of death from refractory end-stage left ventricular heart failure. The HeartWare VAS is designed for in-hospital and out-of-hospital settings, including transportation via fix-wing aircraft or helicopter. |  |  | December 06, 2013 | 2013 | HeartWare Inc | The housing to the Pump's driveline connector became partially or fully separated from the front portion of the driveline connector. | Process control | HeartWare, Inc. sent an Urgent Medical Device Correction letter to all their hospitals and doctors on December 12, 2013 instructing that at each routine clinic visit, they should inspect the patient's driveline connector as described in the attached Technical Bulletin. For already implanted patients, arrange a follow up visit at the earliest convenience to check patient's driveline connector. They were asked to sign and return the Acknowledgement Form to HeartWare. For questions, contact your HeartWare Representative or HeartWare Customer Service at 1-877-367-4823. |  | 4586 | 4,586 | Worldwide Distribution - USA Nationwide in the states of MN, IL, FL, WI, MO, MA, NC, CA, GA, AL, CO, TX, OH, NY, VA, UT, KY, MD, AZ, SC, IN, PA, DE, NJ, LA, OR, WA, MI, IA, TN and DC and countries: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, France, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Singapore, Slovakia, South Africa, Spain, Switzerland, Turkey and United Kingdom. | Terminated | November 03, 2016 | 1063 | 4586 | Not_Computer
Z-1488-2014 | 67905 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | GE Vivid E9 Ultrasound System, a general-purpose ultrasound system, specialized for use in cardiac imaging. |  |  | March 27, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue where the probe power surveillance settings for 4 probes (C1-5-D, C2-9-D, IC5-9-D and M5Sc-D) are incorrect. This may result in a situation where probes can overheat. | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter dated March 27, 2014. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology / Managers of Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (Contact your local GE Healthcare service representative). |  | 447 | 447 (81US, 366 OUS) | USA Nationwide Distribution in the states of: CA, CO EL, GA, IN, KS, KY , LA, MA, MI, MO, MT, NY, NM, NV, US: OH, TX, VA, WI.OUS: ARABIC EMIRATES, AUSTRALIA, AUSTRIA, BELGIUM, COLOMBIA, CZECH REPUBLIC, DENMARK, FINLAND, FRANCE, GERMANY, GREAT BRITAIN, INDIA,IRELAND, ITALY, JAPAN, KOREA, NETHERLAND, NORWAY, POLAND, PORTUGAL, RUSSIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN. | Terminated | January 14, 2015 | 293 | 447 | Not_Computer
Z-1489-2014 | 68034 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | ev3 EverCross 0.035" OTW PTA Dilatation Catheter, 6mm x 120mm x 135cm, REF: AB35W06120135. and ev3 EverCross 0.035" OTW PTA Dilatation Catheter, 6mm x 40mm x, 80cm REF: AB35W06040080. Peripheral Dilatation Catheter |  |  | April 09, 2014 | 2014 | Ev3, Inc. | Covidien is recalling 2 production lots of EverCross 0.035 PTA Balloon Catheter because of a product labeling error. There is a discrepancy in the PTA Balloon length and the delivery catheter length printed on the product labeling. | Labeling Change Control | Consignees were sent a Covidien "Urgent Product Recall" letter dated April 09, 2014. The letter described the problem and the product involved in the recall. The letter described the required actions which included to immediately quarantine and discontinue use of the affected devices and to return affected product. Requested consignees to complete the Verification Form and return it to their Covidien Sales Representative. For questions they can contact Covidien Customer Service at 1-800-716-6000. |  | 39 | 39 units (31 OUS, 8 US) | Worldwide Distribution-USA including the states of WV, TX, OK, IN, NY, TX, and NC, and the countries of AUSTRALIA, ISRAEL, TURKEY, UNITED KINGDOM, ITALY, FRANCE, and CHINA. | Terminated | October 06, 2014 | 180 | 39 | Not_Computer
Z-1490-2014 | 67509 | No | JFY | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | VITROS Chemistry Products CDM PROM (Calibration Data Module) No. 0168, Catalog Number/REF 199 9077, used in conjunction with: VITROS Chemistry Products CRSC DT Slides, Generation (GEN) 82, REF 6802721, and |  |  | January 29, 2014 | 2014 | Ortho-Clinical Diagnostics | Customers were unable to calibrate, or use previously calibrated, VITROS DT CRSC slides (REF/Product Code 6802721), GEN 82, on the VITROS DT II System, when using Calibration Data Module (CDM) Rev. 168 (REF/Product Code 199 9077), as CDM 168 does not contain calibration parameters for DT CRSC. | Process control | Ortho Clinical Diagnostics sent an Urgent Product Correction Notification Letters dated January 29, 2014, via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees to inform them of the issue. Customers were advised to Inspect their inventory to determine if they had VITROS CRSC DT Slides, GEN 82. - To use their current inventory of GEN 82 until the replacement order arrives providing that they have CDM PROM 0167 installed on their system and calibrate using VITROS DT Calibrator Kit, Lot 42. - If product replacement is required, discard their remaining inventory of VITROS CRSC DT Slides, GEN 82 upon receipt of their replacement order. - Complete and return the Confirmation of Receipt Form by February 7, 2014, by fax to 1-888-557-3759 or 1-585-453-04110, or email to: ConfirmationAdmin@its.jnj.com. Indicate on the form if they have any inventory of the affected lot that will require replacement. - Forward this information if they have distributed this product outside of your facility. Distributors were also instructed to identify all customers and any other branches who were shipped VITROS CRSC DT Slides, GEN 82 from their facility and send a copy of the Urgent Product Correction Notification Letter to inform them of the issue. Foreign affiliates were informed by e-mail on January 22, 2014, of the issue and instructed to notify their consignees of the issue and required actions. |  | 5056 | US: 284 units, Ex-US: 4772 units | Worldwide Distribution - US Distribution including the states of AR, FL, IA, IL, KY, MI, MN, ND, NY and VA, and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 5056 | Not_Computer
Z-1491-2014 | 64791 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | The CEP 8 SpectrumGreen (SG) ASR Probe Kit, 20ul (list 06J37-018) is an Analyte Specific Reagent (ASR) containing a single vial of CEP 8 spectrum green probe and a Certificate of Analysis (COA). Chromosome enumeration probes (CEP) are chromosome-specific FISH probes that hybridize to highly repetitive human satellite DNA sequences, usually located near centromeres. CEP signals enable the identification and enumeration of human chromosomes in interphase and metaphase cells from fresh and archived samples. The label for CEP 8 SG ASR Probe Kit indicates it is an Analyte Specific Reagent, and analytical and performance characteristics are not established. |  |  | March 28, 2013 | 2013 | Abbott Molecular | 20 kits of CEP 8 SG ASR Probe (list 06J37-018, lot 440023), contained Vysis CEP 4 (Alpha Satellite) SG Probe, 20u, (part 32-112004 lot 438814), instead of the expected CEP 8 probe. | Packaging change control | Customers were sent Field Correction Notices dated 3/28/13 instructing them to inspect their CEP 8 SpectrumGreen Probe kits for the presence of CEP4 probe vials. If any CEP4 vials are discovered, customers should contact their local Abbott Molecular Technical Support Represtative for coordination of reagent replacement with the correct CEP8 vial. If customers have any specific questions they can contact Abbott Molecular at 1-800-553-7042, Option 2. |  | 20 | 20 kits | Nationwide (CA, CO, CT, FL, HI, IL, IN, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NM, NV, NY, OH, OR, PA, TN, TX, VA, WA, WI); Canada | Terminated | March 08, 2016 | 1076 | 20 | Not_Computer
Z-1505-2014 | 67568 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | MULTIX FUSION. |  |  | February 13, 2014 | 2014 | Siemens Medical Solutions USA, Inc | It was observed that the corrugated hose holder above the 3D rail on the Multix Fusion system became loose. According to the Installation Instructions the corrugated hose holder shall be afixed with transversal carriage, with Loctite 243 and Torque 10nm during system installation. Siemens initiated the update instruction to verify and secure the hose holder with Loctite 243. A PCB D802 Rev. 00 b | Process change control | Affected customers were initially notified with a recall letter on/about 2/12/14 and were sent Update Instructions to affected customers. |  | 12 | 12 | Distributed in the states of AK, PA, KY, OH, WI, MI, FL, and NY. | Terminated | March 16, 2015 | 396 | 12 | Not_Computer
Z-1497-2014 | 67895 | No | GAQ | Stainless steel suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Monofilament and Multifilament Non-absorbable Steel Suture; Product Codes: 0520197-2, 29-727, 29-7270M4, 29-7281M4, E29-5553M4, E29-7270M4, E29-7272M4, E29-727A, EP4013P, EP4110N, X-2235, X29-7271M2, X-5667M8 and X-6730M5. |  |  | March 11, 2014 | 2014 | Teleflex Medical | The products are being recalled because they did not meet minimum diameter requirements. | Other | Teleflex sent an Urgent Medical Device Recall Notification letter dated March 11, 2014 to all affected constomers. The letter identified the affected product, problem and actions to be taken. Customers were instructed if they have affected stock immediately discontinue use, quarantine, and return all affected product to Teleflex Medical per the instructions provided in the letter. Customers were asked to complete the enclosed Recall Acknowledgement Form and Fax it to 1-508-964-6078, Attn: Customer Service. For questions contact your local sales representative or Customer Service at 1-508-964-6042. |  | 3816 | 3,816 ea | Worldwide Distribution - CA, CO, FL, IN, KY, MA, NM, PA, Canada, Ireland and Germany. | Terminated | December 17, 2015 | 646 | 3816 | Not_Computer
Z-1494-2014 | 67831 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrow Pressure Injectable PICC Powered by Arrow VPS Stylet Catalogue number CDC-35052-VPS, peripherally inserted central catheter. |  |  | March 24, 2014 | 2014 | Arrow International Inc | Kits are packaged with the incorrect catheter. The kits contain a 40cm catheter and should contain a 50cm catheter. | Labeling mix-ups | A recall notification letter, dated March 24, 2014 was sent to end users." Urgent Medical Device Recall Notification " |  | 789 | 789 | Nationwide Distribution- Including the states of Az, Tn, Fl, Ga, Ca | Terminated | September 21, 2015 | 546 | 789 | Not_Computer
Z-1493-2014 | 67746 | No | FYD | Air-handling apparatus for a surgical operating room. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Stryker PenAdapt"; Rx Only, Sterile. |  |  | March 26, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | During packaging verification testing, a failure occurred due to cuts in the packaging potentially causing a breach in the sterile barrier. Additional testing identified microscopic aluminum metal shavings and foreign particles estimated in size to be between 0.01inches x 0.015 inches to 0.03 inches x 0.04 inches that were introduced during the manufacturing process and were sterilized within the | Material/Component Contamination | On March 26, 2014, Stryker Instruments sent URGENT MEDICAL DEVICE RECALL NOTIFICATION to their consignees via mail and email. Notification included product description, reason for recall, risk to health, and actions to be taken by the Customer/User. For questions regarding this recall please contact Stryker Instruments:Monday-Friday 8am-5pm (EST)Angela Ragainis/Kelly Jo Whipple269-389-4354 / 269-389-2921angela.ragainis@stryker.com / kellyjo.whipple@stryker.comCustomers will be asked to sign the Business Reply Form stating they have read the Notification and followed all instructions within. Customers will return affected product to Stericycle for destruction using a return label provided by Stryker or Stericycle. Upon receipt of affected product a credit will be issued to customer account.In addition, facilities are being asked to research their records for patient issues that may have been associated with the use of a recalled device, such as infection and/or inflammation. |  | 35310 | 35,310 unexpired units | Worldwide Distribution-USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV,and WY, and the countries of Austria, Canada, Sweden, England, UAE, Netherlands, Spain, and Hong Kong. | Terminated | December 11, 2014 | 260 | 35310 | Not_Computer
Z-1501-2014 | 67976 | [6 empty] | Hot Ice, 1 fl. oz., Meta Labs LLC, UPC 80705-2 |  |  | February 03, 2014 | 2014 | Altasource LLC dba Meta Labs LLC | Firm is not registered as a drug manufacturer. | No Marketing Application | Meta Labs notified customers of the Voluntary Recall Notification letter on February 4, 2014. The letter requested that the products be returned. For questions call 770-645-0602. |  | 155 | 155 | USA Nationwide in the states of AZ, AR, CA, CO, FL, GA, ID, IL,KY, LA, MI, NE, NV, NY, NC, OH, PA, TN, TX, UT, and VA. | Open, Classified |  |  | 274 | Not_Computer
Z-1496-2014 | 67795 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | TRABECULAR METAL MODULAR ACETABULAR SYSTEM, Shell with Cluster Holes, Porous 60mm. Catalog number 00-6202-060-22; Hip implant component. |  |  | April 01, 2014 | 2014 | Zimmer, Inc. | The shell was missing the locking ring. | Nonconforming Material/Component | The firm, Zimmer, sent an "URGENT MEDICAL DEVICE RECALL-LOT SPECIFIC" notification dated April 2014 to their consignees. The notification included product description, product issue, health risks, and consignee responsibilities (including directions on how to return the affected product). The consignees were instructed to locate all products from lot 62490588 and quarantine them immediately; carry out a physical count of all affected product; complete and return the Inventory Return Certification Form via email to: corporatequality.postmarket@zimmer.com; return the recalled product, and notify Zimmer of any hospitals that you have further distributed the affected product and of surgeons that have implanted the product. Consignees with questions pertaining to the recall are directed to call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 17 | 17 units | Worldwide Distribution: US (nationwide) in states of: US: GA, WI, NJ, CA, NY, NC, PA, FL, CT, ME, TX, and IL; and countries of: Australia and Canada. | Terminated | September 11, 2014 | 163 | 17 | Not_Computer
Z-1503-2014 | 67859 | No | MBQ |  |  | Physical Medicine | Contact ODE | Digital Diagnost, stationary radiographic system; Models: 712020/22/25/26/50/52/54/55/57/60/62/80/81/82/85 |  |  | March 17, 2014 | 2014 | Philips Medical Systems, Inc. | The footboard of Patient support for stitching can be fixed in upper position. When the fix clamp (hook) is not correctly adjusted into its correct position, the footboard is not securely held. | Under Investigation by firm | On March 17, 2014, consignees were notified of the recall via letter. The letter explains what the issue is, actions firms can take to avoid or minimize the occurrence of the issue, and the actions planned by Philips to correct the issue.For further information or support concerning this issue, please contact your local Philips Representative. 1-800-722-9377 |  | 985 | 985 devices | Nationwide: US: AL, AR, AZ, CA, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, MA, MD, ME, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, WY / Worldwide including Canada &amp; Mexico | Open, Classified |  |  | 985 | Not_Computer
Z-1492-2014 | 67678 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235 Proton Therapy System Legacy Software version 7.2.7 and higher, models 110, 112, 113, and 116. |  |  | December 04, 2013 | 2013 | Ion Beam Applications S.A. | Internal testing of the Proteus 235 Proton Therapy System revealed incorrect dose rate calculation in case of beam line option other than Zero (0). | Software design | Ion Beam Applications ( IBA) sent a recall notification letter dated on/about December 5, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Short term action: A user notice (see annex 1) has been sent to impacted sites where the affected legacy PTS software version was delivered. This user notice explains the workaround required to be put in place in order to fully mitigate the risk and the 1BA action plan to permanently fix this issue.Corrective action:A patch on the PTS Legacy SW was requested and in implementation. In the same time, the solution was put in main code of PTS legacy software version 7 to be integrated in the next release.The customer representative signing underneath confirms that the user notice above is communicated to the appropriate employees.We apologize in advance for any inconvenience this may cause and we would like to thank you for your cooperation.The IBA Site Staff or the Operations manager is available to provide you with additional information and/or guidelines if necessary.Further questions please call : 215-972-7777 |  | 4 | 4 | US Distribution including the states of OK, IL, NJ, and WA. | Terminated | June 12, 2014 | 190 | 4 | Computer
Z-1507-2014 | 67834 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrowgard Blue Plus Antimicrobial Multi-lumen Central Venous Catheter, catalogue number AK-45854-CDC. |  |  | July 17, 2013 | 2013 | Arrow International Inc | Arrow is recalling the Arrowgard Blue Plus Antimicrobial Multi-lumen Central Venous Catheter due to a labeling inconsistency. One portion ofthe lidstock states that the kit contains 8.5FR x 16CM catheter, while the rest of the lidstockcorrectly identifies the catheter as 8.5FR x 20CM | Labeling False and Misleading | Arrow International Inc, ("Arrow") sent a Urgent Medical Device Recall Notification dated July 17, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. 1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog and lot number listed above. 2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to : 1-866-804-9881. Attn: Customer Service. This will allow us to document the amount of product you have on hand for return Goods Authorization (RGA) number and will provide instructions for the return of product to Arrow International.3.If you have no affected stock please complete the enclosed Recall Acknowledgement Form and fax it to : 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.For further questions please contact Arrow Customer Service at 1-800-343-2935. |  | 170 | 170 | US Distribution including the states of IA, OR, WV, MI, AZ, IL, PA, GA, DE, NC, LA, and VA. | Terminated | June 02, 2015 | 685 | 170 | Not_Computer
Z-1502-2014 | 67896 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | On-Board Imager - Exact Arm Elbow Motor Shaft (kVD, kVS, and MV) with ''Type 01" elbow motors. |  |  | March 20, 2014 | 2014 | Varian Medical Systems, Inc. | Varian has identified a failure with the shaft of the Type 01 elbow motor for Exact arms used in the On-Board Imager (OBI) device due to belt over-tensioning. | Equipment maintenance | Varian sent an Urgent Medical Device Correction letters March 20, 2014 to affected custoers via trackable mail. The letter was also distributed to Varian Sales, Marketing and Service organizations to inform them of the issues. The letter identified the affected product, description of problem, details, and recommended user actions to be taken. A Varian Customer Support Service representative will arrange a site visit to check the bel tension of the Type 01 elbow motors. If over tensioned belts are detected, the elbow motor will be replace with a Type O2 elbow motor. Questions may be directed to 888-827-4265. |  | 2622 | 2622 | Worldwide Distribution. | Terminated | May 01, 2015 | 407 | 2622 | Not_Computer
Z-1495-2014 | 67868 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips MDC PACS R2.X / IntelliSpace PACS DCX R3.1 software, software |  |  | April 02, 2014 | 2014 | Philips Medical Systems, Inc. | Philips MDC PACS Release R2.x and Philips IntelliSpace PACS DCX R3.x, may have incorrect density measurement on Enhanced CT/MR examinations | Software design | Philips Healthcare sent a Urgent Medical Device Correction Notification/Field Safety Notice dated March 2014, to inform customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips will replace the software free of charge.Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:Until you receive a software upgrade from Philips, the user must configure the MRI and CT devices to send regular CT or MR DICOM image objects (not enhanced) to PACS.. This issue does not affect images that are not enhanced image objects.Further questions please call (978) 687-1501. |  | 373 | 373 units | Worldwide Distribution - US Distribution including Puerto Rico and the states of CA, FL, IL, NJ, OK and WA., and the countries of : Argentina Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Bosnia, Brazil, Bulgaria, Chile, China, Egypt, Ethiopia, France, Germany, Ghana, Greece, Guinea, India, Iraq, Israel, Italy, Latvia, Libya, Lithuania, Malawi, Malta, Mexico, Mozambique, Nepal, New Zealand, Papua, New Guinea, Paraguay, Philippines, Poland, Portugal, Romania, Russia, Slovakia, Slovenia, Solomon Islands, South Africa, Spain, Thailand, Turkey, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 373 | Computer
Z-1504-2014 | 68027 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | Siemens ACUSON X700 Ultrasound Systems with software version 1.0.00, 1.0.01, 1.0.02, 1.1.00 and 1.1.02. ultrasound imaging system. |  |  | April 02, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | When fourSight 4D ultrasound imaging, 3-Scrape real time 3D imaging, or Anatomical M-mode are used during an Obstetric or Gynecology study, data from the previous patient could replace the current patient's date in the report. | Software Design Change | Siemens sent a Customer Safety Advisory Notification letter on April 2, 2014 to all affected customers via certified mail. The letter identified the affected product, problem and what actions to take until software is upgraded. For questions call 1-888-826-9702. |  | 490 | 490 | Worldwide Distribution - US Nationwide and the countries of Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia Herzeg., Brunei, Chile, Colombia, Croatia, Ecuador, Egypt, France, Germany, Greece, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kenya, Kuwait, Latvia, Libya, Liechtenstein, Malaysia, Mauritius, Montenegro, Namibia, Nepal, Nigeria, Norway, Palestinian Ter, Philippines, Poland, Portugal, Qatar, Rep. of Yemen, Republic Korea, Romania, Saudi Arabia, Slovakia, South Africa, Spain, Switzerland, Taiwan, Turkey, Turkmenistan, U.A.E., United Kingdom. | Terminated | November 12, 2015 | 589 | 490 | Computer
Z-1506-2014 | 67648 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Sengstaken Tube and Sengstaken-Blakemore Tube. Teleflex product codes 204800, 204802, and 204803, a gastrointestinal tube and accessories, consisting of flexible or semi-rigid tubing used for instilling fluids into, withdrawing fluids from, splinting, or suppressing bleeding of the alimentary tract. |  |  | March 04, 2014 | 2014 | Teleflex Medical | Sterility of the product cannot be guaranteed. | Other | Letters were issued on 3/4/2014 to consignees asking them to immediately discontinue use and quarantine any products with the recalled catalog numbers. Distributors were also asked to conduct a sub-recall. The recall letters included a Recall Acknowledgment Form which was to be faxed back to Teleflex. |  | 117801 | 117,801 ea. | Worldwide Distribution-USA (nationwide) including the states of AL, AK, AR, CA, CO, FL, GA, ID, IL, IA, KY, LA, MD, MA, MI, MN, MS, MO, NE, NH, NJ, NM, NY, NC, OH, PA, SC, TN, TX, UT, VT, VA, WA, WV, and WI, and the countries of Algeria, Andorra, Armenia, Argentina, Austria, Australia, Azerbaijan, Belarus, Bermuda, Bahamas, Boliva, Brazil, Canada, Chile, China, Costa Rica, Cyprus, Czechoslovakia, Germany Guatemala, Denmark, Estonia, Egypt, Finland, France, Great Britain, Georgia, Guadeloupe, Greece, Guyana, Curacao, Hungary, Israel, India, Italy, Jordan, Japan, Kuwait, Kazakhstan, Luxembourg, Latvia, Libya, Madagascar, Macedonia, Martinique, Malta, Mauritius, Mexico, Nicaragua, Nigeria, Netherlands, Norway, Oman, Peru, French Polynesia, Poland, Portugal, San Marino, Singapore, Spain, Suriname, San Salvador, Thailand, Tunisia, Turkey, Uruguay, Uzbekistan, Wallis &amp; Futuna Islands, Zambia. | Terminated | January 28, 2016 | 695 | 117801 | Not_Computer
Z-1508-2014 | 67818 | No | DIO | Cocaine and cocaine metabolite test system. | Toxicology | Toxicology | 510(k) | ToxCup Drug Screen Cup, Catalogue Numbers: HT15, HT15A, PT11, PT11A, PT15A, PT23A, a screening device for the detection of drugs and drug metabolites in human urine. For in vitro diagnostic use only. |  |  | March 26, 2014 | 2014 | Branan Medical Corporation | May be labeled with incorrect result interpretation graphics on the ToxCup lid label. | Labeling Change Control | The firm, Branan Medical Corp (BMC), sent an "URGENT MEDICAL DEVICE RECALL" letter dated March 26, 2014 to all the customers who received the ToxCup Drug Screen Cup. The letter describes the product, problems identified and the actions to be taken. Customers are instructed to discontinue use and/or sale of product; quarantine all units of these product lots; destroy unused product(s) in compliance with your laboratory and state requirements; share this information with your staff and retain notification as part of your QS documentation; if distributed to your customers, provide a copy of this letter to them; complete and return the enclosed Customer Verification Form via Fax to: Branan Medical Corporation at 949-598-7167 or by e-mail at: branan.complaints@brananmedical.com within 10 days. Should you have any questions about the information contained in this notification, contact Branan Medical Corporation at (949) 598-7166 or email branan.complaints@brananmedical.com. |  | 9800 | 9,800 units | Worldwide Distribution: US (nationwide) and country of Canada. | Terminated | July 23, 2014 | 119 | 9800 | Not_Computer
Z-1511-2014 | 68032 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | GE Healthcare, CardiacVX [CardiacVX (1.1.0 Ext. d) and CardiacVX Flow (1.1.0 Ext. d) analysis software running on Advantage Workstation]. Model #'s 5434389-2 and 5434384-2. |  |  | April 07, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to higher than actual flow values reported on CardiacVX and CardiacVX Flow running on Advantage Workstation. | Software design | GE Healthcare sent an Urgent Medical Device Correction" letter dated August 21,2013, to all affected consignees. The letter was addressed to Hospital Administrators /Risk Managers, Radiology Department Managers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. Customers were instructed until the issue is corrected on their system with the affected CardiacVX or CardiacVX Flow software for phase-contrast protocols customers use the fast CINE phase-contrast sequence with the magnitude weighting user CV option disabled. For directions on how to update user CV, please refer to the MR operator manual. If customers choose to perform imaging using magnitude weighting, do not utilize CardiacVX or CardiacVX flow for flow value calculations.with the affected product. Customers with questions were instructed to contact their local service representative.For questions regarding this recall call 262-513-4122. |  | 101 | 101 (1 US , 100 OUS) | Worldwide Distribution - USA including IN and Internationally to AUSTRALIA, CANADA, CHILE, FRANCE, ITALY, JAPAN, KOREA, KUWAIT, NEW ZEALAND, PANAMA, POLAND, SAUDI ARABIA, SPAIN, SWEDEN, SWITZERLAND, UNITED KINGDOM, | Terminated | May 29, 2014 | 52 | 101 | Computer
Z-1484-2014 | 67523 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SIGMA SPECTRUM Volumetric Infusion Pump with Master Drug Library |  |  | February 07, 2014 | 2014 | Baxter Healthcare Corp. | Baxter Healthcare Corporation has issued an Urgent Device Correction for the SIGMA SPECTRUM Volumetric Infusion Pump with Master Drug Library due to repeated System Error 322 occurrences. | Software design | On February 7, 2014, Baxter Healthcare Corporation distributed Urgent Device Correction notices and Customer Reply Forms to their customers via first-class mail. Affected customers are instructed to contact Baxter Healthcare Corporation at 1-800-356-3454 (choose option 1) Monday through Friday during the hours of 7:00 am to 7:00 pm Eastern Time. Baxter technicians would determine if an adjustment or repair is necessary. Customers are asked to complete the attached Customer Reply Form and return it to Baxter either by faxing to 1-224-270-5457 or e-mailing a scanned copy to fca@baxter.com. |  | 277096 | 277,096 units (revised) | Worldwide Distribution - USA Nationwide including PR and the country of Canada. | Terminated | May 25, 2017 | 1203 | 277096 | Not_Computer
Z-1509-2014 | 67972 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Monodek Violet Synthetic Absorbable Surgical Sutures, MF 0 TC43/HR26 48 |  |  | March 11, 2014 | 2014 | Teleflex Medical | Product does not meet minimum knot tensile strength requirements. | Under Investigation by firm | Consignees were notified by an Urgent Medical Device Recall Notification letter, dated 3/11/2014. The letter identified the affected product and the reason for recall. Customers were instructed to immediately discontinue use of and quarantine any affected product in stock. The affected product is to be returned; and the Recall Acknowledgement Form should be completed and faxed to the number provided regardless of whether customers have affected product in stock. Questions should be directed to a local sales rep or Customer Service at 1-866-246-6990. |  | 3072 | 3,072 ea. | Worldwide Distribution -- USA, including the state of MA, and the country of Germany | Open, Classified |  |  | 3072 | Not_Computer
Z-1516-2014 | 67885 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | The Acumed Tension Band Pin is a fixation pin which is .0625" in diameter and is available in lengths between 35mm and 90mm. It is to be used in conjunction with commonly used orthopedic wire. |  |  | March 26, 2014 | 2014 | Acumed LLC | Certain lots of Acumed Tension Band Pins are being recalled because the implantable product was manufactured with a grade of stainless steel that is not within specifications. | Component change control | Acumed sent an Urgent Medical Device Recall letter dated March 26, 2014 to all affected customers. The letter identified the affected products, problem and actions to be taken. Customers are instructed to immediately stop using and/or distributing all lots identified in the letter. Return affected products as instructed in the letter. For questions contact Acumed Business Service via email at BusinessService@acumed.net or via phone at 877-627-9957. |  | 6128 | 2544 unit in the US and 3584 units outside the US | US Nationwide Distribution - in the states of CA, CO, FL, GA, LA, MI, MN, MO, OH, OR, PA, PR, TX, UT, VA, WA and WY.Product was distributed outside the US to: Australia, China, Figi, France, Great Britain, Italy, Japan, Korea, Maylasia, Sweden, South Africa, Spain. | Terminated | February 04, 2015 | 315 | 6563 | Not_Computer
Z-1513-2014 | 67719 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 6.0 is a patient positioning and monitoring system. |  |  | February 24, 2014 | 2014 | Brainlab AG | When using multiple radiation treatment targets within a single plan, the ExacTrac v.6.0.3 might position the patient to an unintended radiation treatment point. As a result, there is a possibility that the radiation treatment dose may be delivered to the unintended target position. | Software design | A FIELD SAFETY NOTICE / PRODUCT INFORMATION letters, dated February 24, 2014, were sent to all direct accounts. The letter identified the affected product, the reason for the recall, and included instructions for direct accounts to take while a permanent fix can be deployed. The letter also advised that a software update will be made available starting the end of April 2014. Customers with questions or concerns can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 41 | 41 systems (13 systems in US; 28 systems to foreign countries) | Worldwide Distribution -- USA, including the states of CA, IN, LA, MA, OH, TX, and WI, and the countries of Australia, China, France, Germany, Japan, Netherlands, and the United Kingdom. | Open, Classified |  |  | 41 | Computer
Z-1514-2014 | 68088 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | HexaPOD evo RT Couchtop. |  |  | April 17, 2014 | 2014 | Elekta, Inc. | If screws are not properly tightened with a torque wrench they may become loose over time. | Component design/selection | Corrective Action #1 - Important Field Safety Notice 618 01 302 002 was issued April 2014. The Notice informs end users about the problem and the potential clinical impact. It also give a recommendation how the risk can be mitigated - inspect the screws and if loose contact Elekta Service to tighten the screws with 20 Nm. The notice also advises the customer that a modification kit will be released that introduces an improved screw solution. The solution will be implemented on all affected products when available. Corrective Action #2 Permanent Solution - Replace Current Screws - An improved screw has been designed and will be available for Elekta Field Service organization to replace on all affected products at the end of April. The new design solution provides a very long used thread depth, exceeding the standard recommended thread-depths by far. This design will prevent the screws from loosening. |  | 266 | 266 | Worldwide Distribution-USA including DC and the states of AZ, CA, CT, FL, GA, IL, HI, KY, MD, MA, MI, MS, MO, NJ, NY, NH, NC, ND, OH, OK, OR, PA, SC, TX, VA, WA, WI, and WY, and the countries of Algeria, Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Czech Republic, Denmark, Finland, France, Germany, Greece, Hong Kong, India, Israel, Italy, Japan, Kazakhstan, Mexico, Netherlands, New Zealand, Peru, Poland, Portugal, Russia, Saudi Arabia, South Korea, Span, Sweden, Switzerland, Taiwan, Thailand, Turkey, and United Kingdom. | Terminated | January 29, 2015 | 287 | 266 | Not_Computer
Z-1518-2014 | 67210 | No | JHR | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | IMMULITE 2000 /IMMULITE 2000 XPi Systems TIE Total IgE, REF L2KIE2, SMN 10380873, and L2KIE6, SMN 10380872, IVD --- Siemens Healthcare Diagnostics Products Ltd. |  |  | December 20, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens confirmed an overall average positive bias of 23% against the WHO 2nd IRP 75/502 with the IMMULITE Systems Total IgE assays. | Under Investigation by firm | Urgent Medical Device Correction Letters (dated 12/20/13) and effectiveness check response forms were sent to customers on 12/20/13 via FedEx to inform them of the issue. Customers were requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within thirty (30) days. Field service personnel were sent a support document describing the issue and instructing them how to deal with customer questions. |  | 18489 | 18,489 units | Worldwide distribution. | Open, Classified |  |  | 24442 | Not_Computer
Z-1510-2014 | 67901 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Nuclear gamma cameras (ADAC Vertex Plus, ADAC Solus, ADAC Vertex V60 &amp; ADAC Vertex Classic) Philips Medical System, Cleveland, OH 44143. |  |  | March 10, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Slippage of the radius drive belt in the relative 180 degree orientation of the superior positioned detector head will allow the detector head to drift down and potentially come in direct contact with the patient. | Device Design | On 03/10/2014, the firm sent Urgent - Medical Device Correction letters to their customers. The letter identified the affected product, the reason for recall, the hazard involved, and the actions that should be taken by customers. Customers are to cease use of the identified procedures until implementation of the Field Change Order. Additionally, customers should sign and return the attached Customer Response Form within 10 days of receipt, acknowledging that they have read and understood the Field Safety Notice and Addendum. The Field Safety Notice must be placed in customers' User Documentation. Customers are to contact their local Philips Healthcare representative or office, if they need further information or support. |  | 1043 | 1,043 cameras | Worldwide Distribution -- USA, and the countries of Australia, Austria, Belgium, Brazil, Canada, China, Denmark, Egypt, France, Germany, Hong Kong, Ireland, Italy, Korea, Lebanon, Malaysia, Netherlands, Norway, Pakistan, Portugal, Russia, Singapore, Spain, Sweden, Thailand, &amp; United Kingdom. | Terminated | August 04, 2015 | 512 | 1043 | Not_Computer
Z-1521-2014 | 67786 | applicator, absorbent tipped, sterile | KXG | Absorbent tipped applicator. | General Hospital | General Hospital | 510(K) Exempt | Fisherbrand Sterile Swabs, Calcium Alginate Fiber Tipped Aluminum Applicator Swab, .025 and .035 in Diameter, all lots, 100 per pack |  |  | February 05, 2014 | 2014 | Fisher Scientific Co | A customer reported the presence of an organism in a lot of Calcium Alginate Fiber Tipped Wood Applicator Swab. | Process control | Fisher Scientific sent an Urgent: Medical Device Recall/Return letter dated March 18, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. ACTIONS TO BE TAKEN BY THE USER OR DISTRIBUTOR:Fisher Scientific requests that you discontinue use and distribution of all products referenced above. We also request that you return all product (opened and unopened) as soon as possible. A credit will be issued for returned product.User: Fisher Scientific requests that you complete the Acknowledgment Form enclosed with this letter, even if you do not have the above mentioned product(s) in your inventory, within three (3) working days.Distributor: Our records indicate that you may have purchased one or more of the above products for re-sale to your customers. You should complete the attached Acknowledgment Form for inventory you have received and/or which is still in stock. Fisher Scientific requests that you complete and return the Acknowledgment Form enclosed with this letter, even if you do not have the above mentioned catalog numbers in your inventory, within three (3) working days. In addition, please contact your affected customer base to advise them of the expanded recall, by providing them with a copy of this letter. You should insert your contact information, email and fax numbers in the Acknowledgment Form and request that your customer return the Acknowledgment Form to you, and you notify us of the information conveyed in such Acknowledgment Form(s).ACTION BY FISHER SCIENTIFIC:Fisher Scientific has no inventory of the affected products as it was discontinued in May 2013. However, we do have alternative products that we can recommend as replacements. Please contact customer service for information on product alternatives.We appreciate your immediate attention to this expanded voluntary recall. By returning the attached Acknowledgment Form you wil |  | 7430 | 7430 | Nationwide Distribution | Terminated | August 10, 2015 | 551 | 7430 | Not_Computer
Z-1520-2014 | 67883 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Chemistry Products FS Diluent Pack 3, REF 680 1754, 3 x 45 mL, (D1)/15 mL (D2), IVD --- Ortho-Clinical Diagnostics, Inc. Rochester, NY, 14626 |  |  | March 26, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) confirmed the potential for lower than expected C-Reactive Protein (CRP) results only to be generated when using VITROS Chemistry Products Specialty Diluent (Lot F3168) or VITROS Chemistry Products FS Diluent Pack 3 (Lot 01-3266). The VITROS Instructions for Use states that if the if a C-reactive protein concentration exceeds the systems measuring (reportable or dynamic) range, dilute the sample with VITROS Specialty Diluent or a patient sample containing a low concentration of CRP. An initial threefold dilution is recommended. Internal testing determined that results using VITROS CRP Slides have the potential for approximately 51% bias using a 3x dilution factor. | Other | On 3/26/2014, Urgent Product Correction Notification Letters were sent via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them to discontinue use of the affected lot(s), discard the affected lots and to inform their Laboratory Medical Director if they suspect that a previously reported result may have been affected. Replacement product will be sent for the product discarded. J&J; Foreign affiliate consignees were notified by e-mail informing them of the issue on 3/26/2014. |  | 1042 | 1042 distributed (220 went to warehouse in France) | Nationwide, Puerto Rico, and foreign distribution to Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Terminated | March 27, 2017 | 1097 | 2698 | Not_Computer
Z-1522-2014 | 67902 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra CUSA Excel Disposable Wrench. Catalog Number C5601. The CUSA Excel+ Sterile 23 kHz Torque Wrench is an accessory used with the CUSA Excel+ System ultrasonic surgical instrument. The torque wrench is used to properly secure the CUSA tip to the CUSA handpiece in the sterile field. |  |  | March 10, 2014 | 2014 | Integra LifeSciences Corp. | Some 23 kHz Torque Wrenches were incorrectly assembled and contained a 36 kHz Hex Wrench Insert instead of the 23 kHz Hex Wrench Insert. | Under Investigation by firm | Integra sent an Urgent: Voluntary Medical Device Recall letter dated March 10, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The written notifications were delivered to US consignees by traceable courier service and non-US consignees by traceable email or facsimile. Integra is asking that you do the following:1. Review your unused inventory of 23 kHz CUSA Sterile Torque Wrenches and determine if they are on the affected lot number list on the next page and if they are still in the box.2. Complete the attached form.- If you do not have an affected lot number on the list (see page 2), YOU CAN CONTINUE TO USE THE WRENCHES. Check the box I do not have any affected products on the list- If you do have product(s) on the affected lot number list, AND THEY ARE STILL IN THE BOX, determine if the box also has a green dot sticker as in the photo below. If so, YOU CAN CONTINUE TO USE THE WRENCHES. Check the box I do not have any affected products on the list on page 2.3. Complete the other information on the form if you do or do not have affected products, return the form by email or fax as indicated on the form. Keep a copy of the form for your records.When your form is received, if you have noted you have affected products Customer Service will contact you and provide an RMA number and directions to return the product, as well as input an order to replace the quantity you indicated on the form.Should you have any questions regarding these instructions, please contact Customer Service at 1-855-532-1723.We apologize for any inconvenience this may cause and thank you for your cooperation in this effort. |  | 23904 | 3,984 devices packaged as 6 devices / box (664 boxes) | Worldwide Distribution - USA (nationwide) including the countries of Australia, Canada, and Switzerland. | Terminated | February 13, 2015 | 340 | 23904 | Not_Computer
Z-1525-2014 | 66841 | No | MHY |  |  | Neurology | PMA | Medtronic Activa RC (37612), Activa PC (37601), Activa SC (37602), Activa SC (37603), Activa PC+S (37604), Multi-program Rechargeable Neurostimulator for Deep Brain Stimulation for movement disorders. |  |  | February 27, 2014 | 2014 | Medtronic Neuromodulation | Medtronic neurostimulators (used with pain and deep brain stimulation therapy) have a cycling feature that, when enabled, allows stimulation output to be toggled on and off based on the programmed settings. The use of this feature is determined by the physician and set using the clinician programmer. Current labeling indicates the use of cycling improves device longevity and recharge interval | Labeling False and Misleading | Medtronic sent a Medical Device Correction letter dated February 2014 to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to do the following:If you have patients with cycling enabled or wish to use cycling, use the updated calculations and graphs available online at professional.medtronic.com/cycling to assess the impact of cycling on device longevity. Medtronic will update the product labeling in the second half of 2014." Remind your patients to continue to check their battery status. Instructions on how to check battery status can be found in the patient programmer or recharger system manual." Pain Stimulation Trialing Only: The cycling feature should not be enabled on the External Neurostimulator (ENS) when estimating device longevity because the resulting longevity estimate for non-rechargeable devices may not provide accurate information.Consignees were also instructed to acknowledge receipt of the notice by completing the attached reply form and returning it to Medtronic using the contact details on the reply form. For questions consignees can contact Medtronic Neuromodulation Technical services at 1-800-707-0933 weekdays 7am - 6pm CST.For questions regarding this recall call 800-707-0933. |  | 244207 | 170,286 US, 73,921 OUS for all devices | Nationwide Distribution - all states including DC and Puerto Rico. | Terminated | November 07, 2014 | 253 | 244207 | Not_Computer
Z-1531-2014 | 67211 | No | CIZ | Androstenedione test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE/IMMULITE 1000 Systems Androstenedione (Test Code AND), REF LKAO1, SMN 10381158, IVD, 100 test kit size --- Siemens Healthcare Diagnostics Products Ltd. Llanberis, Gwynedd, LL55 4EL, UK |  |  | December 20, 2013 | 2013 | Siemens Healthcare Diagnostics | Siemens Healthcare Diagnostics confirmed over-recovery of samples with concentrations greater than 5.5 ng/mL with the IMMULITE/IMMULITE 1000 and IMMULITE 2000/IMMULITE 2000 XPi Androstenedione assays (LKAO1, L2KAO2). Siemens has determined that these lots are linear up to a concentration of 5.5 ng/mL, instead of 10 ng/mL as stated in the Instructions for Use (IFU). Quality Control materials at con | Under Investigation by firm | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated December 19, 2013, to its customers informing them that Siemens Healthcare Diagnostics has confirmed over-recovery of samples with concentrations greater than 5.5 ng/mL with the IMMULITE/IMMULITE 1000 and IMMULITE 2000/IMMULITE 2000 XPi Androstenedione assays (LKAO1, L2KAO2). The distribution of the customer communication began on 12/20/13 via FedEx (US). The letter described the product, problem and actions to be taken. Customers were instructed to take the following actions: - Please discuss this communication with your medical director. - Siemens recommends the use of Quality Control materials with at least two levels having concentrations less than 5.5 ng/mL. - Patient values of &gt; 5.5 ng/mL should be verified using an alternate method. - and complete and return the Effectiveness Check questionnaire attached to the Urgent Field Safety Notice within 30 days and fax to: CUSTOMER CARE CENTER at 302-631-7597.Field service personnel were sent a Service Knowledge Base (SKB) describing the issue and instructing them how to deal with customer questions.If you have any questions, please call 1-914-524-3074 or contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 892 | 892 | Worldwide distribution: US (nationwide) and countries of: Argentina, Australia, Austria, Belgium, Brazil, Canada, China (People's Republic of), Czech.Rebuplik, Denmark, Estonia, France, Germany, Great Britain, Italy, New Zealand, Portugal, Spain, Switzerland, Turkey and Vatican City. | Open, Classified |  |  | 8633 | Not_Computer
Z-1530-2014 | 67897 | No | CIT | Aspartate amino transferase (AST/SGOT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alfa Wassermann Diagnostic Technologies, LLC. ACE AST Reagent. |  |  | March 05, 2014 | 2014 | Alfa Wassermann, Inc. | Extremely low absorbance readings. | Process control | Alfa Wassermann Diagnostic Technologies, LLC issued an "Important: Advisory Notice/"Product Response Form" dated March 5, 2014 and sent on March 6, 2014 to their customers for ACEÂż ALT Reagent Product reorder number SA1046 Lot F3604. They issued a follow-up "Important: Advisory Notice/"Product Response Form" dated March 21, 2014 and sent on March 26, 2014 to their customers for ACEÂż AST Reagent Product reorder numbers SA1047, Lot F3609, and RX1047, Lot F3610. The notification identified the product(s), problem, and actions to be taken by the customers. Customers were instructed to destroy any affected product and complete and submit the enclosed Product Response Form. For questions or concerns, contact the Alfa Wassermann Diagnostic Technologies, LLC Customer Solution Center, Toll Free: (866) 419-ALFA (2532). |  | 53 | SA1047 lot F3609: 146; RX1047 lot F3610: 53 | US Nationwide Distribution | Terminated | August 13, 2015 | 526 | 784 | Not_Computer
Z-1527-2014 | 67950 | No | ERL | Ear, nose, and throat electric or pneumatic surgical drill. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | PolypVac 3.5mm Microdebrider; Catalog Number FG-000016; contents: 1 Microdebrider, 1 Stylet, 1 filter. Sterile. Used for the excision of polyps in the nasal passageways and in sinuses for which access has been previously established. |  |  | April 07, 2014 | 2014 | Laurimed LLC | The reciprocation mechanism in the PolypVac Microdebrider may turn off prematurely due to a component failure in the device handle. | Component design/selection | All direct accounts were notified on April 8, 2014 with an Urgent: Medical Device Recall of PolypVac Microdebrider letter, sent by first class mail. The letter identified the affected product, the reason for the recall, and stated that the firm is in the process of implementing corrections. The firm is requesting that all unused lots be returned to them. Replacement devices will sent to customers. In addition, customers are to complete the attached Lot Information sheet. Questions should be directed to bdubois@laurimed.com or 650-587-5296. |  | 58 | 58 | Distributed in the states of CA, VT, NC, DO, TX, FL, PA, OR, AZ, IL, ID, and VA. | Terminated | May 14, 2014 | 37 | 58 | Not_Computer
Z-1482-2014 | 67591 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Abbott Acclaim Infusion Pump |  |  | October 31, 2013 | 2013 | Hospira Inc. | Hospira has received customer reports of broken door assemblies on the Hospira/Abbott Acclaim Encore infusion pumps. If the door assembly breaks, it may prevent the door from closing properly and unrestricted flow may occur. If the door cannot be closed, the pump cannot be used which canresult in a delay in therapy. | Component design/selection | URGENT MEDICAL DEVICE CORRECTION letters, dated October 31, 2013, were sent to all direct accounts. The letters identified the affected product, as well as the reason for recall, risk to health, and instructions for device users. Users are to inspect each Hospira/Abbott Acclaim Encore infusion pump for door handle cracks prior to programming a therapy by following the provided instructions.The letter also included instructions for recipients to ensure that all potential users in the facility are made aware of the safety notification and the recommended actions; and complete the attached reply form and return it to the fax number or e-mail address on the form. Customers are to contact Stericycle at 866-891-0586 (Monday - Friday, 8:00 AM - 5:00 PM ET) to obtain additional reply forms if needed. If the devices were further distributed, those customers should be notified and asked to contact Stericycle at 866-891-0586 (Monday - Friday, 8:00 AM - 5:00 PM ET) to receive a reply form. The letter reminds customers that Hospira is retiring the Acclaim Encore infusion pump and will no longer support them as of December 31, 2014. |  | 13218 | 13,218 infusion pumps total | Nationwide Distribution -- AK, AL, AR, AZ, CA, CO, CR, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY. | Terminated | August 16, 2017 | 1385 | 13218 | Not_Computer
Z-1534-2014 | 67973 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CAPIO, Monodeck, Violet Monofilament Polydioxanone Suture, 1 x 48 inches (122cm) Absorbable Surgical Suture, Rx Only, Teleflex Medical. Product Code: 833-137. |  |  | March 11, 2014 | 2014 | Teleflex Medical | The product did not meet minimum and/or average minimum Teleflex resorption strength requirements. | Other | Teleflex Medical sent an Urgent Medical Device Recall Notification dated March 11, 2014. The letter identified the affected product, problem and actions to be taken. Custmers were instructed to return all affected product to Teleflex Medical per the instructions on the Urgent Recall Notice. Customers were asked to complete the enclosed Recall Acknowledgement Form and fax to 1-866-804-9881, Attn: Customer Service. For questions contact your local sales representative or Customer Service at 1-866-246-6990 |  | 7380 | 7,380 ea (total) | Worldwide Distribution - US Nationwide in the states of CA, CO, GA, LA, IL, MA, MI, MN, MO, NC, and in the countries of Ireland and Singapore. | Terminated | March 29, 2016 | 749 | 7380 | Not_Computer
Z-1533-2014 | 68141 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare Discovery MR450, Discovery MR750, Optima MR450w, Optima MR450w equipped with the GEM option, Discovery MR750w, and Discovery MR750w equipped with the GEM option products. |  |  | April 23, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue with certain GE MRI Products. Patient egress from the system can be delayed or impeded if the cradle release handle and/or cradle release block mechanism malfunctions or becomes misaligned. This could make the safe removal of the patient from the system difficult. | Process design | Consignees were sent on 4/23/2014 a GE Healthcare Ref # 60863 "Urgent Medical Device Correction" letter dated April 18, 2014. The letter was addressed to Hospital Administrators / Risk Mangers, Radiology Department Mangers and Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. If customers encounter the stated problem with the device, they should contact their GE Healthcare Service Representative for an inspection and necessary repairs. GE will correct all affected devices at no cost to the customers. For questions customers should contact their local service representative. |  | 1548 | 1,548 units (704 US, 844 OUS). | Worldwide Distribution -- USA, except ND, RI and WY, and including DC, Guam, and Puerto Rico; and, the countries of ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA,AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BRAZIL, CANADA, CAYMAN ISLANDS, CHINA, CROATIA, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, HONG KONG, HUNGARY, INDIA, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, KOREA, KUWAIT, LEBANON, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, POLAND, QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, T UREY, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, UZBEKISTAN, and VENEZUELA. | Terminated | January 21, 2015 | 273 | 1548 | Not_Computer
Z-1512-2014 | 67802 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar Docking Station, used with GemStar infusion pump, in combination with other GemStar accessories such as the GemStar Lock Box and patient bolus cord. The docking station is also used for hosting the GemStar pump for general infusions. List Numbers: 13075-XX-XX. |  |  | February 18, 2014 | 2014 | Hospira Inc. | There are two situations that may occur when using the GemStar Docking Station, List Number 13075-XX-XX, in conjunction with the GemStar infusion pump: 1) when the Docking Station is used in conjunction with a GemStar Phase 3 pump (List 13000-XX, 13100-XX, or 13150-XX) the potential exists for the GemStar Phase 3 pump to fail to power up while connected to the Docking Station, and 2) when either a GemStar Phase 3 (List 13000-XX, 13100-XX, or 13150-XX)or GemStar Phase 4 pump (List 13086-XX, 13087-XX, or 13088-XX) is used in conjunction with both a Docking Station and an External Battery Pack accessory (List 13073-XX) there exists a possibility that the GemStar pump will display error code 11/003 and give an audible alarm indicating excessive input voltage from the external sources. If the GemStar pump detects what is perceived to be more than 3.6 Volts measured on the external voltage input, the pump will stop the infusion and alarm both with an audible sound as well as a visual alarm. | Device Design | Hospira sent an URGENT MEDICAL DEVICE CORRECTION letter dated February 18, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. The letter included the following instructions: 1) To avoid failure to power up, turn the pump on first, before connecting the pump with the docking station. This will prevent the failure to power up. 2) To mitigate the potential for an 11/003 error code, remove the External Battery Pack accessory (List 13073) from the Docking Station in conjunction with an External Battery Pack accessory (List 13073), this practice should not continue. Please contact Hospira to discuss an appropriate alternative. Additional instructions included: 1) Inform potential users within the recipient's organization of the notification. 2) Should the facility experience a failure to power up or an 11/003 error code, report the issue to Hospira by calling 1-800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CT) or e-mail ProductComplaintsPP@hospira.com. 3) Compelte the attached reply form and return it to the fax number or e-mail address on the firm, even if there is no affected product in inventory. 4) If th products were further distributed, notify the accounts that received the products identified in the notification and ask them to contact Steriycle at 866-792-5451 (Monday - Friday, 8:00 AM - 5:00 PM ET) to receive a reply form.For questions regarding this recall call 800-441-4100. |  | 26765 | 26,765 docking stations total (6866 in US) | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Guam, Puerto Rico, and Internationally to Australia, Austria, Bahrain, Belgium, Canada, Chile, Egypt, France, Germany, Greece, Hong Kong, Italy, Jordan, Kuwait, Lebanon, Libya, Luxemburg, Netherlands, New Zealand, Oman, Portugal, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Turkey, United Arab Emirates, and United Kingdom. | Terminated | January 23, 2017 | 1070 | 26765 | Not_Computer
Z-1536-2014 | 67948 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | Anti EpCam PE - Analyte Specific Reagent. |  |  | April 04, 2014 | 2014 | BD Biosciences, Systems &amp; Reagents | Two lots of Anti EpCAM PE-Analyte Specific Reagent are mislabeled - the actual antibody present in the vial is CD9 PE. | Under Investigation by firm | Customer product recall notification letters were sent to all consignees on April 4, 2014 by certified mail. |  | 120 | 12 for lot 3295974; 10 of lot 3259822 | Worldwide Distribution-USA including the states of CA, TX, FL, MO, NY, MI, MO, and AZ and Seoul, Korea. | Terminated | September 28, 2015 | 542 | 120 | Not_Computer
Z-1538-2014 | 68093 | No | DWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Joseph Lamelas Atrial Lift System, Model Number MI-ALS-001. A single surgical instrument or tool for retraction of the atrial wall tissue in cardiac surgery. |  |  | March 19, 2014 | 2014 | Biorep Technologies | The support arm attached with an adhesive to the body of the blade is detaching while attempting to engage the visor component. | Nonconforming Material/Component | Biorep Technologies sent a letter dated March 19, 2014, to its affected customer. It was advised to due to the high risk, according to our risk categories shown in Table 2, at this time, it is recommended that product be retrieved from its points of use.Any further questions please call (305) 687-8074. |  | 80 | 80 boxes (480 units) | US Distribution to a distributor in MN. | Terminated | August 07, 2014 | 141 | 80 | Not_Computer
Z-1543-2014 | 67871 | No | FNK | Hydraulic adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | RetractaBed , Clinical Contour, Med - Mizer Inc..... Model SS 801, 802, 803, 803A: 80 " bed with 1, 2, or 3, motors and autocontour |  |  | April 18, 2014 | 2014 | Med-Mizer, Inc. | All configurations of the SS Retractabed, Clinical Contour, made prior to June 7, 2008 may have had insufficient welding on the foot cross tube. This tube is responsible for holding the two support rails together when using the foot side rails and keeps them from separating from the frame. If the weld on the tube breaks, it can lead to bed collapse or the rollers coming out of the channel on t | Process change control | On 5/1/14, Med-Mizer sent URGENT: MEDICAL DEVICE RECALL notification to their consignees. Notification included description of the affected product, reason for recall, risk to health, instructions on how to recognize how the device may fail, and instructions for actions to be completed by their consignees. Contact Information: 1-877-867-7365, sales@med-mizer.com, 8am-5pm EST. |  | 2980 | 2980 | USA (nationwide) | Open, Classified |  |  | 3172 | Not_Computer
Z-1528-2014 | 67947 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | INOmax DSIR, model 10007, Nitric Oxide Delivery System, with Software version 3.0.0. The INOmax DS delivery system delivers INOMAX (nitric oxide for inhalation) therapy gas and provides continuous monitoring of inspired O2, NO2, and NO, and a comprehensive alarm system. The primary targeted clinical setting is the Neonatal Intensive Care Unit (NICU) and secondary targeted clinical setting is the transport of neonates. |  |  | March 27, 2014 | 2014 | INO Therapeutics (dba Ikaria) | An anomaly impacts the monitoring system of INOmax DSIR Plus (software 3.0.0). When the integrated pneumatic backup switch is activated with low or no gas flow through the injector module, an intended NO cell failure alarm will occur. In addition, an unintended NO2 cell failure alarm may also occur. The NO/NO2 cell failure alarms are caused by an average NO/NO2 circuit concentration greater tha | Software Design Change | Ikaria sent an Urgent Field Corrective Action letter via FedEx on March 27, 2014. The letter identified the issue, and requested that all INOmax DSIR with software version 3.0.0 be returned to Ikaria. Customers with questions were instructed to contact Ikaria Customer Care at 1-877-566-9466.For questions regarding this recall call 877-566-9466. |  | 22 | 22 | Nationwide Distribution including DC and VA. | Terminated | November 21, 2014 | 239 | 22 | Computer
Z-1544-2014 | 68045 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | 12-Instrument Sterilization Tray; |  |  | April 16, 2014 | 2014 | Stryker Endoscopy | Gravity sterilization parameters provided in the 12 Instrument Sterilization reprocessing guide did not fully consider all worst-case scenarios during validation testing. | Reprocessing Controls | Stryker sent an Urgent Device Corrections letters dated April 18, 2014 to all US customers via certified mail. Stryker Sales representative and Stryker international distributors were notified with the letter and response card by e-mail. International sites are responsible for notifying customers. The letter identified the affected product, problem, potential hazards and actions to be taken. Customers were instructed to complete and sign the enclosed Acknowledgement of Receipt form and scan and e-mail to endorecall@stryker.com. For questions call 1-800-624-4422 option 3 or via email at endorecall@stryker.com. |  | 2591 | 2591 trays | Worldwide distribution | Terminated | August 18, 2014 | 124 | 2591 | Not_Computer
Z-1540-2014 | 67910 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Triage BNP Test for the Beckman Coulter Access Family of Immunoassay Systems, Model Number 98200, Lot Numbers: 329599, 331265. The Alere Triage BNP test is intended to be used as an aid in the diagnosis of congestive heart failure (heart failure), as an aid in the assessment of severity of congestive heart failure, for the risk stratification of patients with acute coronary syndromes, and for the risk stratification of patients with heart failure. |  |  | March 20, 2014 | 2014 | Alere San Diego, Inc. | These lots may not recover within range when tested using certain commercially available controls. | Under Investigation by firm | Alere San Diego, Inc. initiated this recall by sending out notification letters via email and/or regular mail on March 20, 2014. The letter titled "URGENT MEDICAL DEVICE CORRECTION", dated "March 20, 2014, informed customers of the reason for recall, and potential impact, product description with lots numbers, customer/distributor required action, and contact information. The letter was accompanied by a "Customer/Distributor Verification Form". Customers were instructed that they may continue to use reagent lots 329599 and 331265." If they are using an alternate commercial control with Reagent Pack Lots 329599 or 331265, they can receive the Alere TriageÂż BNP QC Controls PN 98201 by requesting them on the Verification Form."Customers were instructed to share this information with their laboratory staff, including their laboratorys medical director and retain this notification as part of their laboratory Quality System documentation.For questions regarding this recall call 858-805-2000, ext 3015." If they have forwarded the product listed above to another laboratory, please provide a copy of this letter to them." Complete and FAX the enclosed Verification Form within 10 days to confirm your receipt of this notice. Only one Verification Form is required per facility.Customers with questions were instructed to call 877-308-8287 or email Triage.Support@alere.com. |  | 12592 | 12,592 kits | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Canada, China, EU, Mexico, Russia, Singapore, South Africa, South Korea, Taiwan and United Arab Emirates. | Terminated | January 23, 2015 | 309 | 12592 | Not_Computer
Z-1537-2014 | 68016 | No | JPD | Hemoglobin A[bdi2] assay. | Hematology | Hematology | 510(k) | Variant II beta-thalassemia Short Program Model Numbers 270-2103 (250 tests) and 270-2154 (500 tests), Bio-Rad Laboratories, Inc. in vitro diagnostic. The Variant II Hemoglobin testing system uses the principals of high performance liquid chromatography (HPLC) for the separation and determination of normal and abnormal hemoglobin. |  |  | April 08, 2014 | 2014 | Bio-Rad Laboratories, Inc. | There is an issue with the buffers which impacts the quantification of the Hemoglobin F peak. Calibration and Quality control are likely to fail if the affected lots are used. | Nonconforming Material/Component | BIO-RAD Laboratories sent an URGENT: MEDICAL DEVICE LETTER in the US on April 8. 2014 and worldwide with letters sent to customers and subsidiaries on April 14, 2014. Subsidiaries were asked to translate the letter and distribute to all affected customersCustormer were asked to complete a CUSTOMER MEDICAL DEVICE CORRECTION RESPONSE FORM and fax it to the Bio-Rad CSD Regulatory Affairs Department at the number provided to confirm that you havereceived this important information |  | 457 | 457 kits | Worldwide Distrubution ; USA INCLUDING CA, NY, GA, NC, AL, NJ, CT, WA, NY, LA, Il, MA and Internationally to China, Australia, Canada, France, Great Britain, Hong Kong, Thailand, Singapore,and Brazil. | Terminated | May 30, 2014 | 52 | 457 | Not_Computer
Z-1545-2014 | 67930 | No | KQO | Automated urinalysis system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | CLINITEK Status+ Analyzer (portable urine chemistry analyzer). Catalog number 10490946. |  |  | February 26, 2013 | 2013 | Siemens Healthcare Diagnostics Inc | Device with Sample Interference Notes (SIN) enabled not cleared for US marketing. | Reprocessing Controls | On February 26, 2013, a Siemens Field Engineer from Siemens Healthcare visited customer site and disabled the Sample Interference Notes feature on the instrument. |  | 1 | 1 unit | US Distribution in the state of: CA. | Terminated | May 06, 2014 | 434 | 1 | Not_Computer
Z-1552-2014 | 68074 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | TX1 Tissue Removal System. TX1 Procedure Pack P/N 554-1003-001 (outer box), TX1 MicroTip P/N 554-1002-001 (inner tray). For use with the TX System Console as an ultrasonic surgical aspirator of soft tissue. |  |  | April 14, 2014 | 2014 | Tenex Health Inc | The sterile barrier in the packaging may be compromised due to cracks in the thermoform tray that may occur during transport of the product in extreme cold weather conditions, in conjunction with mechanical impact. | Packaging | Tenex Health sent an Urgent Medical Device Recall notification letter dated April 15, 2014, to all consigness. The letter informed consignees of the recall, manufacturer, product name, product code, reason for recall, recall carried out to the user level, instructions to customers, contact information, label examples, and response form was provided. Consignees were instructed to immediately examine their inventory per the enclosed instructions, if they identify any of the affected product within their inventory, they should segregate the product and immediately return only MicroTips with damaged trays to the address below, using the Tenex Health, Inc. FEDEX account number and referencing the Returned Goods Authorization (RGA) number.RGA#Tenex Health, Inc.26902 Vista TerraceLake Forest, CA 92630Customers were asked to complete and return the enclosed response form as soon as possible.Customers with questions should contact Tenex at (949) 454-7500, ext. 209or email tanaka@tenexhealth.com.For questions regarding this recall call (949) 454-7500. |  | 9927 | 9,927 units | Worldwide Distribution - USA (Nationwide) and Internationally to Panama and Singapore. | Terminated | July 16, 2014 | 93 | 9927 | Not_Computer
Z-1562-2014 | 67986 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | AlloFuse DBM Putty 5cc, Catalog No. 90038005 |  |  | March 25, 2014 | 2014 | AlloSource, Inc. | The donor was hemodiluted. | Process control | AlloSource sent letter via USPS Certified Mail Receipt to all affected customers dated March 25, 2014, and March 26, 2014. The letter identified the product the problem and the action needed to be taken by the customer. AlloSource requested that the distributor consignee notify subsequent consignees to determine disposition and to request return of unused product. Returned inventory will be quarantined physically and electronically upon receipt. Product initially quarantined and product returned will be destroyed following established procedures.Further questions please call (720) 873-4733 |  | 29 | 29 | Distribution US nationwide, including Michigan and a distributor in Colorado. | Terminated | August 11, 2014 | 139 | 29 | Not_Computer
Z-1554-2014 | 68095 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Bausch &amp; Lomb, Biotrue, Multi-purpose contact lens solution, 10 fl.oz., Sterile. |  |  | April 21, 2014 | 2014 | Bausch &amp; Lomb Inc- Greenville Solutions Plant | Out of specification results for one of the two disinfectants. | Nonconforming Material/Component | Consignees were notified via registered mail starting 04/21/2014. |  | 256440 | 256,440 bottles | US Nationwide. | Terminated | January 09, 2015 | 263 | 256440 | Not_Computer
Z-1549-2014 | 68067 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | PandaÂż Freestanding Infant Warmers, integrated with Nellcor SpO2. |  |  | April 02, 2014 | 2014 | GE Healthcare | Medical device software may be defective and may cause user-set alarm limits to result in false positive or false negative Oxygen saturation and pulse rate alarm notifications to the caregiver. | Software design | GE Healthcare notified Biomedical/Clinical Engineering managers, Labor &amp; Delivery/ NICU Nurse managers, Respiratory Therapy Managers and Directors of Risk Management by letter 04/02/2014. End users were advised that they may continue to use the system provided they follow the GE recommended actions:1) If in use with a patient, turn off the Nellcor SpO2 function in the control panel and switch to an approved alternate form of SpO2 measurement.2) Prior to using the affected GE warmer with the Nellcor SpO2 function, or resuming the use of this function with any patient, contact your hospital Bio-Med or qualified technician to check the Nellcor SpO2 alarm functionality using either an adult Nellcor SpO2 sensor or a Nellcor SpO2 simulator following the applicable instructions in the Appendix 3) If the test activates a Low Pulse Rate alarm, the user-adjusted Nellcor SpO2 alarm limits are functioning properly and the warmer can be placed back in clinical service.End users were also advised that alarm functionality must be re-checked following every service event requiring access to the internal electronics and during annual preventative maintenance, until revised software is installed.4) If the test does NOT activate a Low Pulse Rate alarm, remove the warmer from clinical service and contact your local GE Healthcare Service Representative. |  | 21 | 21 units | Worldwide Distribution. | Open, Classified |  |  | 1110 | Computer
Z-1563-2014 | 68114 | No | FIP | Water purification system for hemodialysis. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Mar Cor Purification, Central Water Platform (CWP), 102, 104 and 106. |  |  | April 24, 2014 | 2014 | Mar Cor Purification | Mar Cor Purification manufactured Central Water Plant (CWP) reverse osmosis system, models 102, 104 and 106. This correction has been initiated due to a potential situation with a specific lot of the water inlet solenoid valve. The potential exists that the water inlet solenoid valve could cease to function. If this occurs, CWP will stop producing reverse osmosis water. | Nonconforming Material/Component | Consignees were sent on 4/24/2014 a Mar Cor Purification "Urgent Medical Device Correction" letter. The letter described the problem and the product involved in the recall. Advised consignees to examine their CWP unit to determine if it is subject to this correction. If your unit is involved in the recall inform all staff of the potential issue. Contact Scott Nicholas, Project Coordinator / Service Trainer at 1-800-633-3038 to schedule replacement of the water inlet solenoid valve. This correction is to be carried to the user level. Consignees were requested to complete and return the "CWP Correction Notice Response Form". For questions they should contact Scott Nichols, Project Coordinator / Service Trainer at 1-800-633-3080. |  | 197 | 197 | US Nationwide Distribution - AL, AZ, AR, CA, CO, FL, GA, ID, IN, IA, KS, KY, LA, MA, MI, MN, MS, MO, MT, NH, NJ, NM, NY,NC, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WV, WI. | Terminated | November 20, 2015 | 575 | 197 | Not_Computer
Z-1555-2014 | 67711 | No | GAR | Nonabsorbable polyamide surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Nonabsorbable Surgical Suture, U.S.P., Sterile - Do Not Resterilize, Deknatel. |  |  | March 11, 2014 | 2014 | Teleflex Medical | Product does not meet minimum needle attachment strength requirements; thus, the needle can become detached from the suture during use. | Employee error | Urgent Medical Device Recall Notification letters were issued on March 11, 2014 to consignees. The letter requested that customers discontinue use and quarantine any product on hand. The attached Recall Acknowledgment Form should be completed and returned to Teleflex regardless of whether affected product is on hand or not. Once received a customer service representative will contact customers with instructions on how to return the affected product. Questions should be directed to Customer Service at 1-866-246-6990. |  | 32271 | 32,271 ea. | Worldwide Distribution -- USA, including the states of AL, CA, FL, GA, IL, IN, KS, LA, ME, MA, MI, MN, MS, MO, NV, NJ, NM, NY, OH, PA, RI, SC, TN, TX, UT, VA, WV, and WI; and countries of Canada, Germany, and Ireland. | Terminated | August 14, 2017 | 1252 | 32271 | Not_Computer
Z-1565-2014 | 63145 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | SpineFrontier S-Lift Lock Shaft Inserter, an instrument used with S-LIFT Intervertebral Body Fusion Device in spinal fusion surgery. Product Usage: The S-Lift Lock Shaft SI50001 mates with SI50002, Implant Inserter, to aid in securing the implant to the whole implant inserter assembly, SI50000. |  |  | December 23, 2011 | 2011 | SpineFrontier, Inc. | It has been verified through field use that there are two possible breakages that may occur when using SI50001 Revision D Lock Shaft Inserter, possibly resulting in instrument fragments entering the sterile field:Impact Cap: The impact cap on the Lock Shaft Inserter can shear off under high-impact force during insertion and removal.Pin: The pin used to guide the Lock Shaft Inserter into the ma | Device Design | Spine Frontier sent an Urgent Notification dated December 23, 2011, consignees. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the attached acknowledgement form. For questions call 866-914-7717 |  | 30 | 30 devices | US Nationwide - OR, FL, MA, TX, CO, OK, IL, and MO | Terminated | May 20, 2014 | 879 | 30 | Not_Computer
Z-1564-2014 | 67839 | No | CGN | Folic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens ADVIA Centaur Folate |  |  | March 25, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Serum controls may go out of range low with the ADVIA Centaur ÂżSystems Folate assay and a shift in whole blood patient may result with kit lots ending in 219 and 223 | Nonconforming Material/Component | Siemens issued An Urgent Medical Device Recall (UMDR) to US customers on March 25, 2014. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on March 24,2014. These notices inform the customer of reported controls failures with Folate ReadyPackÂż kit lots ending in 223. In addition the notices inform customers of a shift when moving from Folate ReadyPackÂż kit lots ending in 219 and kit lots ending in 222, 224, 225, 226 and 227. Customerswere instructed to discontinue use of ADVIA CentaurÂż Systems Folate kit lots ending in 219 and 223. ADVIA CentaurÂż systems Folate kit lots ending in 222, 224, 25, 226 and 227 are performing within specification and may continue to be used for folate interpretation. The notices should be reviewed with the facility's laboratory or medical director. Fax-back forms will be used to confirm customers received the notification. |  | 49123 | 49123 units | Worldwide Distribution-USA (nationwide) and the countries of Austria Belgium Bulgaria Cyprus Czech Republic Denmark Estonia Finland Finland France Fren.Polynesia Germany Great Britain Greece Guadeloupe Hungary Iceland Ireland Israel Italy Latvia Lithuania Luxembourg Malta Martinique Netherlands Norway Poland Portugal Qatar Reunion Romania Russian Fed. Serbia Slovakia Slovenia Spain Sweden Switzerland Turkey Unit.Arab Emir. | Terminated | April 30, 2015 | 401 | 49123 | Not_Computer
Z-1569-2014 | 67927 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Shimadzu Digital Radiography X-Ray System, Catalog No. DAR-8000f. |  |  | April 03, 2014 | 2014 | Shimadzu Medical Systems | Shimadzu Corporation is recalling the Shimadzu Digital X-Ray System because they have received a report of a lost image under certain circumstances. | Software design | A Service Bulletin and a customer notification letter were sent on 4/3/2014 for all the customers who purchased the Shimadzu Digital X-Ray System (Catalog No. DAR-8000f). The notification informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns about the device correction, are instructed to contact their local Shimadzu representative or contact Shimadzu National Technical Support at (800) 228-1429. |  | 685 | 685 units total (14 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Japan, China, Germany, Australia, Thailand, Sri Lanka, South Africa, UAE, Saudi Arabia, Korea, and Taiwan. | Terminated | June 25, 2015 | 448 | 685 | Not_Computer
Z-1568-2014 | 68031 | No | LKK |  |  | General Hospital | PMA | Refill Kit 8551/856X (includes Models 8561, 8562 and 8564) for use with Medtronic Implantable Programmable Infusion Pumps. |  |  | April 04, 2014 | 2014 | Medtronic Neuromodulation | Some refill kits in the affected lots may contain an extension tubing set that is occluded and healthcare professionals may be unable to aspirate or fill the SynchroMed pumps. There is no problem with the Lioresal Intrathecal (baclofen injection) contained in the affected drug refill kits. There have been no patient injuries beyond additional needle sticks associated with this issue. | Process control | Medtronic sent a "Medical Device Removal" letter dated April 2014 hand delivered, e-mail or faxed to affected customers. The letter described the problem and the product affected by the recall. The letter also contained the Action Required, Additional Information (Contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7 am - 6pm Central Time) and requested consignees to complete and fax or e-mail the "Medical Device Removal Reply Form". |  | 13784 | 13,784 | Worldwide Distribution - USA Nationwide in the states of NM, ND, RI, and VT and the countries of Australia, Singapore, Thailand, Central and Eastern Europe, Western Europe, Middle East, &amp; Panama. | Terminated | November 05, 2014 | 215 | 13784 | Not_Computer
Z-1546-2014 | 68038 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Alere INRatio2 PT/INR Professional Test Strips, Model Number: 99008G2. The Alere INRatio2 PT/INR Monitoring System (Professional Use), consisting of the Monitor and test strip, is used for quantitative determination of international normalized ratio (INR) in fresh, capillary whole blood to monitor the effect of warfarin therapy on clotting time by health care professionals. (in vitro diagnostic use). |  |  | April 16, 2014 | 2014 | Alere San Diego, Inc. | Test Strips may report an inaccurately low INR result. Several patients had a therapeutic or near-therapeutic INR with the Alere INRatio2 PT/INR Professional Test Strip but a significantly higher INR (outside of therapeutic range) when performed by a central laboratory. | Under Investigation by firm | Alere sent an Urgent: Medical Device Recall letter dated April 16, 2014 via fax, email or direct mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to discontinue use, return remaining inventory to Alere, and complete the provided verification form to acknowledge receipt of the recall notice. For questions contact Alere Technical Service by phone at 844-292-5373 or by E-mail at INRatio.Notification@alere.com. UPDATED: On 05/06/2014, Alere issued a press release on this recall to the consumer/public level. The press release, dated May 6, 2014, announced Alere Inc.s voluntary Class I recall in the U.S. of the Alere INRatio2 PT/INR Professional Test Strips (PN 99008G2). The press release also stated the following information: reason for recall, any adverse event reports, reason for the adverse event reports , root cause not yet determined, notified customers to immediately STOP using the product, product description with codes, contact information, FDA MedWatch information and about Alere and About the Alere INRation2 PT/INR Professional Monitoring System. |  | 99795 | Estimated 99,795 Total units | US Nationwide Distribution | Terminated | July 19, 2017 | 1190 | 99795 | Not_Computer
Z-1570-2014 | 67720 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMedÂż II Implantable Drug Infusion Pump, Model 8637-20, 8637-40. |  |  | February 26, 2014 | 2014 | Medtronic Neuromodulation | This recall provides important new information regarding overinfusion associated with the Medtronic SynchroMed II Implantable Pump. Overinfusion can result in a life-threatening overdose andcan also result in drug withdrawal due to premature emptying of the pump. Due to the low reported rate of occurrence of this issue and the inability to predict which pumps may be at risk,Medtronic is not re | Under Investigation by firm | Medtronic sent a "Urgent Medical Device Correction" letter dated March 2014. The letter was sent to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter provided the Explanation of the Issue, Scope and Severity, Recommendations, and Important Guidelines.Customer visits were started by Medtronic field Representatives on February 26th, 2014.Medtronic is communicating this information to the appropriate regulatory agencies globally, including the U.S. Food and Drug Administration. We are committed to continuing to improve our product performance and services to enable you to manage your patients in a safe and effective manner. If you have questions, please contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7am- 6pm CST. |  | 195198 | 195,198 pumps (146,435 US, 48,763 OUS) | Worldwide Distribution - All states in USA. OUS: List not provided at this time. | Open, Classified |  |  | 195198 | Not_Computer
Z-1572-2014 | 67921 | No | LKK |  |  | General Hospital | PMA | MedStream Programmable Infusion Pump, 40 mL, Product Code: 91-4201 |  |  | November 10, 2015 | 2015 | Codman &amp; Shurtleff, Inc. | A small number of specific pumps may trigger a low battery alarm at the beginning of the implanted life. | Device Design | Codman Neuro sent an Urgent Field Safety Notice to all affected customers on April 18, 2014. The letter identified the product the problem and the action needed to be taken by the customer. The letter states the reason for recall, information about managing patients with affected pumps, and returning affected non-implanted pumps.You are receiving this voluntary recall notification because our records indicate that you are the recipient of one or more MedStream Pumps affected by this recall. Please do not implant any identified products included in this recall. Please complete the enclosed Acknowledgement Form and fax the completed form to: 001-508-977-6665. It is important that we receive this acknowledgement form, even if you have no affected product in your inventory.You may also e-mail the form to us at Codmanproductcomplaints@dpyus.jnj.comWe request that you immediately check all inventory to determine if you have affected product. Please contact your Codman Neuro sales consultant to assist you in the return process. Please use the included instructions to report your inventory status and return affected product. Note this recall applies only to |  | 34 | 34 | Worldwide Distribution including the countries of Belgium Germany, Spain, United Kingdom, Ireland, Italy, Russia and Turkey. | Terminated | November 10, 2015 | 0 | 124 | Not_Computer
Z-1574-2014 | 68109 | No | HTT | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ANSPACH Cutting bur, 3 mm Diamond Ball, Standard Length. REF L-3SD Use with LONG Attachments. Used for cutting and shaping bone, including the spine and cranium. |  |  | January 08, 2014 | 2014 | The Anspach Effort, Inc. | Anspach product code L-3SD contained product code S-3SD which is shorter in length than the labeled product. | Labeling mix-ups | Anspach sent an "URGENT: MEDICAL DEVICE REMOVAL" letter dated January 13, 2014, to all affected customers. The letter described the product, the problem and actions to be taken. If they have distributed any of the products to other services or facilities, customers were instructed to please forward this information as appropriate. Customers with questions you contact Anspach Customer Support at (800) 327-6887, option 3 or email customer.support@synthes.com.. Customers were asked to complete the attached reply form confirm receipt of the letter, return the completed form by fax to 800-327-6661.For questions regarding this recall call 561-494-3706. |  | 30 | 30 cutting burs affected | Nationwide Distribution including FL, TX, NC, IL, CO, and PA. | Terminated | January 20, 2015 | 377 | 30 | Not_Computer
Z-1576-2014 | 68143 | No | EIA | Dental operative unit and accessories. | Dental | Dental | 510(k) | Enaly 1000 BT-12 Ozone Generator, Model # 1000BT-12 |  |  | April 23, 2014 | 2014 | William C. Domb, D.M.D., A Professional Corporation | William Domb is recalling the Enaly 1000 BT-12 Ozone Generator because it is not approved or cleared by the FDA for medical use. | No Marketing Application | William Domb, DMD sent an Urgent Medical Device Recall letter on April 23, 2014, to all affected customers who purchased the Enaly 1000BT-12 Ozone Generator. The letter informed the customers of the problems identified, risk to health, and the actions to be taken. Customers were instructed to discontinue use of the Enaly 1000 BT-12 Ozone Generator and return their units to the following address:William Domb, DMD190 N Mountain AvenueUpland, CA 91786Customers were also instructed to complete the Medical Device Field Action Return Response form. For questions customers should contact William Domb at (909) 949-2968 or by email at bill@smilestudio.com from Monday-Friday at 8am-5pm Pacific Time.For questions regarding this recall call (909) 949-2968. |  | 16 | 16 units | Worldwide Distribution - USA including IN, HI, WA, TN, CA, TX, FL, NC, and Internationally to Canada, Portugal, and Ireland. | Terminated | January 08, 2015 | 260 | 16 | Not_Computer
Z-1573-2014 | 67970 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SCC SA HIS version 4.0.7.0 when used with SoftMic version 4.0.7.0. |  |  | January 29, 2014 | 2014 | SCC Soft Computer | When an isolate is resulted without a SNOMED code, isolate information in the downstream system may be incomplete or missing. As a result, there is the potential for the delay or omission of patient treatment updates. | Software design | On 2/10/2014 SCC Soft Computer used a proprietary task management system to send a Correction Communication to the consignee with the affected software. The letter identifies the product, problem and corrective actions. |  | 1 | 1 | Distributed in IL. | Terminated | November 03, 2014 | 278 | 1 | Computer
Z-1581-2014 | 68145 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | DISP 9730950 STRL SPHERE, 1 PK INSERT |  |  | April 24, 2014 | 2014 | Medtronic Navigation, Inc. | Medtronic Navigation is recalling certain Medtronic single-use, disposable devices because there was not sufficient control to assure that returned, unused devices were appropriately examined for product and/or package damage to ensure sterility. | Nonconforming Material/Component | Domestic Urgent Medical Device Safety Notice letters were shipped on April 24 and 25, 2014. The letters indicated that a Medtronic representative will contact the hospital and arrange for a visit to collect affected products and arrange for no charge replacements or credit. Questions regarding this recall should be directed to Medtronic Navigation Technical Services at 1-800-595-9709. |  | 26674 | 26,674 total | Nationwide Distribution | Terminated | December 29, 2014 | 249 | 26674 | Not_Computer
Z-1577-2014 | 68099 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Recalled product is a component of the Newport Medical Instruments e360 |  |  | April 18, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Reports of the audible alarm not sounding when the ventilator is in an alarm state. The ventilator, as well as the visual and remote alarms, continue to operate normally. | Component design/selection | Newport Medical Instruments, Inc. (Newport), a wholly-owned subsidiary of Covidien is conducting a voluntary field corrective action via letters to be sent during the week of April 28, 2014.The letter titled "URGENT MEDICAL DEVICE FIELD SAFETY CORRECTIVE ACTION" informed customers of the recall by providing: reason for recall, product description, information on affected serial numbers regardless of software versions, issue identified, recommended actions to help mitigate risk to the patient, actions to eliminate the occurrence of this issue, important safety reminders, and contact information. Customers are to contact Technical Support Department to receive a software update kit, and then follow the instructions provided once the kit has been received. The attached verification and completion form should be completed and faxed. |  | 5765 | 5,765 | Worldwide Distribution -- USA, including the states of Arizona, California, Colorado, Connecticut, Florida, Georgia, Kentucky, Massachusetts, Maryland, Missouri, Nevada, Pennsylvania, Rhode Island, Texas, Virginia, and Washington; and, the countries of Algeria, Bangladesh, Belgium, Bolivia, Brazil, Bulgaria, Chile, China, Colombia, Dominican Republic, Ecuador, Egypt, England, France, Germany, Greece, Honduras, Hong Kong, India, Indonesia, Ireland, Israel, Japan, Kazakhstan, Kenya, Lebanon, Liechtenstein, Lithuania, Mexico, Morocco, Namibia, Nepal, Netherlands, Nicaragua, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Republic of Armenia, Republic of KOS, Russia, Saudi Arabia, Singapore, Slovenia, S. Africa, Sri Lanka, Taiwan, Thailand, Tunisia, Turkey, UAE, Uruguay, Vietnam, and Yemen. | Terminated | March 01, 2016 | 683 | 5765 | Computer
Z-1575-2014 | 68015 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | NavioPFSTM System, Orthopedic Sterotaxic Instrument, model #'s NPFS-02000, NPFS-02010, NPFS 02020. |  |  | April 11, 2014 | 2014 | Blue Belt Technologies MN | Blue Belt Technologies is conducting a voluntary recall for their Navio PFS system which is used during orthopedic knee surgery. One of the electronic modules inside the computer cart assembly may have a defect in one of its components that could result in a latent failure. | Software design | Beginning April 3, 2014 consignees were hand delivered a Blue Belt Technologies, Inc. "Urgent Medical Device Recall" letter. The letter was addressed to Safety Officer / Regulatory Compliance Officer. The letter included the Description of the Issue, Description of Safety Concerns, Affected Product, Action to be taken by user, Transmission of this field safety notice and Contacts at Blue Belt Technologies. For questions in the USA contact Customer Support Blue Belt Technologies 2900 Northwest Blvd, Suite 40, Plymouth, MN 55441, USA, 1-763-452-4922 (Direct) lbrink@bluebelttech.com. In the European Union, Ross Lethbridge, European Product Manager, Blue Belt Technologies, Office number 207, 3000 Aviator Way. Manchester Business Park, Manchester, UK M22 5TG, Mobile: +44 7787 415069, rlethbridge@bluebelttech.com. |  | 11 | 11 | USA: CA, PA, OH, TX.OUS: BELGIUM, SCOTLAND, UNITED KINGDOM. | Terminated | June 10, 2014 | 60 | 11 | Not_Computer
Z-1603-2014 | 68017 | No | DWL | Medical support stocking. | General Hospital | General Hospital | 510(K) Exempt | DROP LOK" Knee Brace***LATEX FREE" |  |  | March 19, 2014 | 2014 | Zimmer, Inc. | During a transfer of products from a recently shutdown facility, the firm discovered raw material labeled as latex free actually contained latex. | Labeling mix-ups | Zimmer sent an Urgent: Device Removal Letters dated April 8, 2014 to their customers. The letter identified the affected product, probllem and actions to be taken. Customers were instructed to review this notification, iIdentify and quarantine the affected product to prevent further distribution or use, complete the attached Response Form and return it via email. orporateQuality_PostMarket@zimmer.com. Return affected product to: Zimmer SurgicalAttn: QA/RA Dept.- Recall200 West Ohio AvenueDover, Ohio 44622 USAPlease include a copy of the Response Form with the shipment.For returns outside the US, please email Rhonda.duncan@zimmer.com obtain an IRA (international return authorization) number. The IRA request should include the part number(s) being returned and the quantity. Please write the associated IRA number on the outside of the box.4. Zimmer will credit your account for returned Drop-Lok" Knee Braces, Cartilage Knee Braces, Hinged Knee Supports, or Neoprene Tennis Elbow Supports. Please return a copy of the completed response form along with your returned product to ensure proper credit.Important: Please distribute this notification to all personnel within your organization who need to be aware. If you have further transferred affected product(s), please provide the customer's information on the Business Response Form to Zimmer.For questions call 330-354-0989 |  | 131 | 131 units | Worldwide Distribution - US Nationwide in the states of: AK, AZ, CA, FL, GA, IA, IL, IN, LA, MD, MI, MO NC, NE, NV, NY, OH, OK, OR, PA, RI, SD, TX, UT, VA, VT, WA, WI &amp; WV and countries of: AUSTRALIA, CANADA, GERMANY, IRAQ, ITALY, SAUDI ARABIA, TUNISIA &amp; UNITES ARAB EMIRATES. | Terminated | April 14, 2015 | 391 | 3548 | Not_Computer
Z-1608-2014 | 68026 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | Ultrasound needles, 22G (0.7 mm) x 50mm, packaged individually, 10 needles/box, item number REF USC050-22. Needles for peripheral nerve blocking under ultrasound guidance. |  |  | March 19, 2014 | 2014 | RM Temena GmbH | Ultrasound needles may be defective and leak at the needle hub. | Process design | Temena sent an Urgent Recall letter dated March 24, 2014, to the affected customers. The letter identified the product the problem and the action needed to be taken by the customer.IMPORTANT INSTRUCTIONS:1. Ensure that the devices questions will not be applied to humans.2. Opened packages shall not be used any more.3. Please send us back the request of products (including opened packages)4. Inform your distribution partners / further resellers about this recall to achieve a 100% recall of all unused products.5.Please add a copy of our invoice / respectively of our delivery notes, to carry out a cost reimbursement or receive compensation delivery or a credit note.6. Please use attached form to answer Temena SAS.Further questions please call: +33(0)1.30.86.05.30We sincerely apologize for the inconvenience this may have caused by this product recall we thank you in advance for your support |  | 7370 | 7370 needles | US Distribution including the states of MD and IL. | Completed |  |  | 7370 | Not_Computer
Z-1566-2014 | 67884 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | ABACUS TPN Calculation Software, Product Codes: 8300-0045 ABACUS CE (Calculator Edition), 8300-0046 ABACUS SE (Single-Workstation Edition), 8300-0047 ABACUS ME (Multi-Workstation Edition), 8300-0156 ABACUS 3.0 CE, 8300-0157 ABACUS 3.0 SE, 8300-0158 ABACUS 3.0 ME, 8300-0166 ABACUS 3.1 DE, 8300-0167 ABACUS 3.1 CE, 8300-0168 ABACUS 3.1 SE, 8300-0169 ABACUS 3.1 ME. ABACUS is a Windows - based order entry TPN software application for comprehensive calculations and label printing. The primary use of Abacus is for Total Parenteral Nutrition (TPN) order calculations. |  |  | April 02, 2014 | 2014 | Baxter Corporation Englewood | 1. ABACUS v3.1 may calculate quantities of electrolytes that are double the expected values during the creation of TPN orders.2. ABACUS v3.1 may automatically add additional sterile water to a formula equal to the volume of a premixresulting in an over dilution.3. All software versions of ABACUS software display the calcium phosphate curve points for Premasolincorrectly.4. All software v | Software design | Baxter sent an Urgent Device Correction letter dated April 21, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. . 1. Follow safe compounding practice, which includes 1) using of the Summary button to verify the order against the calculated amounts prior to completing the order, 2) verifying that the ordered ingredients and quantities displayed in the software and printed on the Bag label and the Solution Formula label match the PN prescription prior to preparation, and 2) using a filter for administration of a PN bag. Use of the ABACUS Software must comply with your facility's policies, procedures, and training, including verification that all information entered into the ABACUS Software is accurate and clinically appropriate.2. To determine if any of these Issues are present in your current ABACUSconfiguration, please note the following immediate actions should be taken:"Double Salts: Baxter contacted all facilities installed by Baxter with ABACUS v3.1 to correct the Double Salts Issue via telephone between April 4, and April 16, 2014.To avoid the Double Salts Issue, a pharmacy installed with ABACUS v3.1 should 1) ensure the "Include non-electrolyte ion contributions in balance" checkbox is checked for all salt-based templates, and 2) compare the ordered amount of salts to the delivered amount on the Ions tab in ABACUS."Sterile Water Added to TPN with Pre-Mix: Customer should ensure that all premix not intended for use on the compounder are marked as "Pre-Mix PN Solution" type in the Formulary. Customer should then check the template, under the "Infusion" tab, to ensure the "Infusion Mode" is set to "Pre-Mix (Cyclic)"."Premasol CaP04 Curve Data Points: Review Attachment 1 for information relevant to the use of calcium phosphate curves for Premasol 3% Dex 25 |  | 2034 | 2034 | Worldwide Distribution - USA nationwide, Canada, Colombia, Dominican Republic, Panama, Austria, Bosnia Herzegovina, Bulgaria, Bahrain, Estonia, France, Germany, Georgia, Hungary, Ireland, Israel, Jordan, Kuwait, Latvia, Morocco, Oman, Qatar, Saudi Arabia, Slovenia, Slovakia, South Africa, Switzerland, United Arab Emirates, United Kingdom, Hong Kong, and Taiwan. | Terminated | January 20, 2016 | 658 | 2034 | Computer
Z-1611-2014 | 67932 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | DPM 6 Monitor in use with a Multi Parameter Module (MPM). Patient monitor used for monitoring, displaying, reviewing, storing, and alarming of multiple physiological parameters including ECG, arrhythmia detection, ST segment analysis, heart rate (HR), respiration rate (RESP), temperature, pulse oxygen saturation (SpO2), pulse rate (PR), non-invasive blood pressure (NIBP), invasive blood pressure (IBP), carbon dioxide (CO2), anesthetic gas (AG), impedance cardiograph (ICG), bispectral index (BIS), cardiac output (C.O.) and respiration mechanics (RM) |  |  | March 05, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | The MPM included a cardiac ST/Arrhythmia analysis feature which is not cleared for marketing. | No Marketing Application | Mindray DS Usa Inc. contacted their sole customer via phone on March 5, 2014. The affected products will receive software upgrade. |  | 4 | 4 units | US Nationwide Distribution in the state of OH | Terminated | April 27, 2015 | 418 | 4 | Computer
Z-1614-2014 | 68161 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Coaxial Micro-Introducer Kit. |  |  | April 28, 2014 | 2014 | Greatbatch Medical | Greatbatch identified that the documentation does not support the five-year shelf life of the Bard Access Division Coaxial Micro-Introducer Set. | Incorrect or no expiration date | Greatbatch sent an "Urgent Medical Device Recall" letter dated April 25, 2014 to affected customers via email and Fedex. The letter described the Reason for Recall, Risk to Health, and Instructions to Customers. Requested consignees to immediately place affected devices in quarantine and to complete and return the Field Correction Action Response Form. For questions they should contact their local Greatbatch representative or 1-763-951-8235, or e-mail FieldActionCenter@Greatbatch.com. |  | 6642 | 6,642 sets (66,420 devices) | US Nationwide Distribution in the state of UT. | Terminated | May 24, 2016 | 757 | 6642 | Not_Computer
Z-1616-2014 | 67961 | No | NBW | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ACCU-CHEK Compact Plus Test Strips |  |  | April 14, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Roche Diabetes Care has become aware the ACCU-CHEK Compact Plus test strips may produce erroneously low blood glucose readings in patients undergoing Ceftriaxone therapy. The interference is not described in the product labeling. | Device Design | On 4/14/14, Roche Diagnostics Operations issued URGENT MEDICAL DEVICE CORRECTION notification to consignees which included description of the affected product and issue, clinical significance, and actions required. Consignees are directed to call the ACCU-CHEKÂż Customer Care Service Center at 1-800-858-8072 for any questions about the information contained in the notification. |  | 590136 | 590,136 boxes | Nationwide Distribution. | Terminated | September 18, 2014 | 157 | 590136 | Not_Computer
Z-1612-2014 | 66775 | No | BYZ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Mucus Specimen Trap 40 cc, Specimen Trap 40 cc with additional transport cap; Product number DYND44140 (pack of 50), Product number DYND44140H (individual). Mucus Specimen Trap 80 cc, Specimen Trap 80 cc w/ additional transport cap; Product number DYND44180 (pack of 50). Used to collect mucus specimens during suction of fluid from the oral cavity, nose-throat area and/or bronchi of a patient. |  |  | November 04, 2013 | 2013 | Medline Industries Inc | There is a potential that the specified mucus specimen traps are not sterile. Use of the affected products may result in the contamination of patient specimens. | Process design | Medline sent a Recall Notification Letters dated November 4, 2013 to all affected customers.The letters included instructions for customers to: check their stock of the affected products; complete and return the enclosed response form even if none of the recalled products are in inventory; Medline will provide customers with a Return Goods Authorization upon receipt of the response form and credit will be issued upon receipt of the returned recalled products; and, notify any customers who may have received the recalled products. Customers who have any questions can contact Medline Industries at 866-359-1704. |  | 1610 | 1,610 mucus specimen traps | US Nationwide Distribution in the states of CA, MA, MD, MI, MN, NH, NJ, NY, OH, PA, TX, and UT. | Completed |  |  | 1610 | Not_Computer
Z-1553-2014 | 67603 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur XP Immunoassay system, an automated in vitro diagnostic analyzer. Catalog # 078-A011-03 - ADVIA Centaur XP, Catalog # 078-A011R03 - ADVIA Centaur XP refurb. |  |  | February 18, 2014 | 2014 | Siemens Healthcare Diagnostics | Instrument Cover Gas Spring failures. Over time, the gas spring may lose its effectiveness and fail to support the cover in its full or partially open position. This may lead to the cover falling. | Device Design | On February 18, 2014 Siemens Healthcare Diagnostics distributed Urgent Medical Device Correction notices to their customers via courier service (Urgent Medical Device Correction/). Upon receipt of the communication, customers are asked to follow "Actions to be Taken by the Customer". If the cover is unable to stay partially raised, immediately contact your Siemens technical support representative to arrange for a service. Complete and return the Field Correction Effectiveness Check attached to this letter within thirty (30) days by faxing to (302) 631-7597. Customers with questions should contact their Siemens technical support representative. |  | 5900 | 5,900 units total (1,679 domestically &amp; 4,221 internationally) | Nationwide and internationally | Open, Classified |  |  | 5900 | Not_Computer
Z-1613-2014 | 67960 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo Dynamics v9.5 system. A Picture Archiving and Communication System (PACS) intended for acceptance, transfer, display, storage, archive and manipulation of digital medical images, including quantification and report generation. |  |  | April 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Systems with the affected serial number are experiencing a software error which may lead to incorrect patient demographics display. | Software design | Siemens sent a Recall Notification letter dated March 26, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Siemens is preparing a Service Patch modification that would be performed to resolve the software issue. |  | 50 | 50 | US Nationwide Distribution in the states of AR, CA, CT, GA, IA, IL, KS, MA, MI, MN, NJ, NM, NY, OH, OR, PA, TX, TN, TX, WA, WI, and WV. | Terminated | December 15, 2014 | 256 | 50 | Computer
Z-1610-2014 | 68029 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PS-C Femoral Nonporous Rt Sz 6, Model Number 161-1604 Component of the Balanced Knee System and Balanced Knee Revision System, for total knee arthroplasty procedures. |  |  | April 01, 2014 | 2014 | Ortho Development Corporation | A femoral component containing pegs was found in a box for the pegless version. | Process control | Ortho notified customers via email and phone calls on April 1, 2014. Customers were asked to return product in inventory. |  | 5 | 5 | US Nationwide Distribution in the states of NV, and AZ. | Terminated | July 14, 2014 | 104 | 5 | Computer
Z-1619-2014 | 68042 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Revision K Actuator Test Boards sold as replacement parts for use with Fresenius 2008-Series Hemodialysis Machines (2008K, 2008K2, 2008K@home and 2008T). |  |  | April 25, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | 2008-Series Hemodialysis Machines with Revision K Actuator Test Boards may fail to detect a disconnected Ultrafiltration (UF) Pump | Device Design | Fresenius Medical issued Urgent Recall letter issued on 4/25/14 by certified mail with signature confirmation and fax-back form. Customers instructed to examine their inventory to determine whether they have any of the affected 2008T machines or if they have installed any of the affected replacement boards. If customers have the affected machines or replacement boards, they will be instructed to contact FMC-RTG Technical Service to have the Actuator Test Boards replaced. Contact the FMCNA Technical Service Team at 1-800-227-2572. |  | 39 | 39 units | USA (nationwide) and the country of Canada.* | Open, Classified |  |  | 118 | Not_Computer
Z-1615-2014 | 68144 | No | KWI | Elbow joint radial (hemi-elbow) polymer prosthesis. | Orthopedic | Orthopedic | 510(k) | ExploR 7x26mm Modular Radial Stem, Catalog number 11-210062. Elbow prosthesis orthopedic implant. |  |  | April 22, 2014 | 2014 | Biomet, Inc. | This lot of the ExploR Modular Radial Stem may be missing the wedge ramp threads. If the wedge ramp threads are missing from the stem, the likelihood of the set screw loosening from the radial stem intraoperatively is increased. If the set screw loosens, a revision surgery may be necessary. | Process control | Biomet sent an URGENT MEDICAL DEVICE SAFETY ALERT dated April 22, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 8 | 8 devices | US Nationwide Distribution in the states of NY, MD, NC, SC, MI, OR, and NV. | Terminated | January 13, 2015 | 266 | 8 | Not_Computer
Z-1617-2014 | 68208 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | McKesson Radiology-PACS. |  |  | April 17, 2014 | 2014 | Mckesson Medical Immaging | High priority alert icon was not displayed in the Study List when there are no Stat studies present. However, the High priority alert icon is visible when studies marked as Stat priority exist in the Study List. | Software Design Change | On April 17, 2014, McKesson Medical Imaging Company initiated a process for applying a software solution to all McKesson Radiology customer systems that may be affected to prevent future occurrences of this issue. The systems with the potentially affected software versions will be corrected without removing the systems from their location. |  | 45 | 45 | Worldwide Distribution-USA (nationwide) including the states of AL, CA, GA, ID, IL, IA, MD, MS, MO, NE, NV, NH, NY, NC, PA, TN, TX, VT, WA, WV, WY, and the countries of Canada and France. | Terminated | February 02, 2016 | 656 | 45 | Computer
Z-1567-2014 | 68066 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris Pump model 8100 with software version 9.1.18 and software upgrade kits with "Delay Until" Option and "Multidose" Feature |  |  | April 23, 2014 | 2014 | CareFusion 303, Inc. | CareFusion is recalling the Alaris Pump model 8100 version 9.1.18 because it may have a software issue that results in situation where the pump module will not properly delay an infusion when the "Delay Until" option or "Multidose" feature is used. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated April 23, 2014, to all affected customers. The letter informed the customers of the problems identified and the actions to be taken. Customers were informed that if the infusion starts earlier or later than intended and is not immediately detected and stopped by the clinician this could result in serious injury or death. Customers are informed although no adverse events or deaths have been reported there is a potential for this risk. If customers experience the issue while using the Alaris Pump module then they are instructed to contact CareFusion Customer Advocacy at (888) 812-3266. Customers are instructed to promptly complete and return the enclosed Customer Response Card to expedite the corrective action process.For questions regarding this recall call 858-617-4000. |  | 57219 | 57,219 units total (55,756 units in US) | Worldwide Distribution-USA (nationwide) and the countries of Canada, AE, AU, and KW. | Terminated | March 01, 2017 | 1043 | 57219 | Computer
Z-1607-2014 | 67684 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare HVAD Pump Implant Kit (Heartware Ventricular Assist Device) - HeartWare Battery |  |  | April 16, 2014 | 2014 | HeartWare Inc | Patients and caregivers are instructed that if a battery does not provide two hours of support or behaves unusually, that battery should no longer be used and should be replaced. | Under Investigation by firm | HeartWare sent an Urgent Medical Device Correction letter dated April 30, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to familiarize themselves with the letter and review their Patient Manual about proper power management. Discuss with their VAD Coordinator any questions they may have about the HeartWare System or the letter.Always keep two (2) power sources connected to your controller; never leave your controller connected to only one (1) power source except while briefly switching to another power source." Review the recommended practices for power management provided later in this letter. If a battery shows any of the abnormal battery behaviors described, stop using that battery and contact your VAD Coordinator immediately to replace it.Customers were instructed to confirm that they have received and understood the information. Sign and return the enclosed Confirmation form as soon as they can to HeartWare in the prepaid self-addressed envelope or by fax to 1-305-364-2665. For questions customers were instructed to contact their doctor or VAD coordinator.For questions regarding this recall call 305-364-1575. |  | 32349 | 32,349 | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, KY, LA, MA, MD, MI, MN, MO, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia and Internationally to Canada, Dominican Republic, Australia, Belgium, Italy, Germany, New Zeland, South Afica, Japan, Venezuela, Czech Republic, France, Switzerland, Brazil, UK, Saui Arabia, Taiwan, Slovakia, Luxemburg, Hong Kong, Lebano, India, Chile, Oldenburg, Israel, Netherlands, Croatia, Egypt, Turkey, Malaysia, Norway, Greece, Spain, Lithuania, Korea, Singapore, Poland, and Argentina. | Open, Classified |  |  | 32349 | Not_Computer
Z-1620-2014 | 68030 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | RAPIDFILL Syringe Strip, For use with RAPIDFILL Automated Syringe Filler (ASF): H93890200 RAPIDFILL Syringe Strip, White. H93890221 RAPIDFILL Syringe Strip, Yellow. H93890222 RAPIDFILL Syringe Strip, Orange. H93890223 RAPIDFILL Syringe Strip, Red. H93890225 RAPIDFILL Syringe Strip, Blue. H93890227 RAPIDFILL Syringe Strip, Salmon. H93890229 RAPIDFILL Syringe Strip, Violet. H93890231 RAPIDFILL Syringe Strip, Gray. H93890232 RAPIDFILL Syringe Strip, Green. Sterile 10 ml syringes with twist-off luer sealing tip caps banded together with label material. |  |  | April 16, 2014 | 2014 | Baxter Corporation Englewood | Possible breach in sterile barrier for inner and outer bags of theproduct. | Process control | Baxter sent an Urgent Product Recall letter dated April 30, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Baxter is requesting that you take the following actions:1. Locate and remove all affected product from your facility (the product code and lot number can be found on the individual product package or shipping carton).2. To return the affected product, receive credit and order replacement, please contact Baxter Healthcare Center for Service at 1-888-229-0001, Monday through Friday, between the hours of 7:00am and 6:00pm Central Time. Please have your Baxter eight-digit ship-to account number ready when calling.3. Complete the enclosed customer reply form, and return it to Baxter by either fax or scanned e-mail.4. If you are a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, please conduct a recall with your end-user customers in accordance with your customary procedures.If you have questions regarding the content of this communication, please call The Center for One Baxter at 1-800-422-9837 during the hours of8:00am to 5:00 pm Central Time, Monday through Friday. |  | 774600 | 774, 600 | Worldwide Distribution - US Distribution including the states of AL, AZ, CT, CA, FL, GA, OH, OR, IL, KS, MA, MI, MN, MO, NC, PA, HI, TN, VA, NV and WI and the country of Canada. | Terminated | March 23, 2015 | 341 | 774600 | Not_Computer
Z-1621-2014 | 67964 | No | OZT |  | Cardiovascular | Cardiovascular | 510(k) | InterValve V8 Balloon Aortic Valvuloplasty Catheter, model 172212C110, 192412C110, 212612C110, 232812C110. Used in Cardiovascular procedures for Balloon Aortic Valvuloplasty. |  |  | March 14, 2014 | 2014 | InterValve Inc | Leak issues were discovered through accelerated age testing. | Nonconforming Material/Component | Since this recall involved only 2 units at one site, the devices were picked up by the sales representative. There was no written communication regarding this.The site was verbally informed that there was a problem with the device that manifests itself during prepping and the devices need to be retrieved for retest. |  | 2 | 2 | Distribution to TN only | Terminated | June 16, 2014 | 94 | 2 | Not_Computer
Z-1622-2014 | 68107 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | syngo Imaging XS |  |  | April 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc | A complaint reported that the archiving method at one customer site was changed from implicit to explicit, and the cleanup procedure enabled by the explicit archiving method may have deleted data from archive itself. | Use error | Siemens sent a Customer Safety Advisory Notice dated April 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Our service organization already checked all syngo Imaging sites and did not find any other occurrence of the wrongly configured link.Additionally our service organization will include the check for such wrongly configured links into the existing proactive site monitoring procedures.We appreciate your understanding and cooperation with this safety advisory and ask you to immediately instruct your personnel accordingly. Please include this safety advisory notice in your operator's manual. In the interests of safety, we ask that you perform the above preventive measures and inform all affected personnel immediately.If you have sold this device/equipment and it is no longer in your possession, we kindly ask that you forward this safety notice to the new owner of this device/equipment. Please inform us about the new owner of the device/equipment.We regret any inconvenience that this may cause, and we thank you in advance for your understanding. For further questions please call (610) 219-6300.. |  | 64 | 64 | Nationwide Distribution | Terminated | December 15, 2014 | 256 | 64 | Not_Computer
Z-1629-2014 | 68021 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Anspach Irrigation Clip 40 |  |  | December 03, 2013 | 2013 | The Anspach Effort, Inc. | The Irrigation Clips were assembled with the incorrect (smaller) clips. | Mixed-up of materials/components | The firm, Anspach, sent an URGENT: Medical Device Removal letter dated December 3, 2013 to its customer. The letter described the product, problem and actions to be taken. Should you have any queries, please do not hesitate to contact Aspach Customer Support at (800) 327-6887. |  | 150 | 150 | US Nationwide Distribution in the state of Mississippi. | Terminated | April 18, 2016 | 867 | 150 | Not_Computer
Z-1623-2014 | 68129 | No | LJH | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | LATEX-FREE; EXTENSION SET; For Bladder Irrigation or Urinary Drainage, Compatible with T-U-R FLOW-POUCH RESERVOIR, list 6542, 71 inch (180 cm) Nominal Length; List No. 4693-01; single use; Hospira |  |  | March 25, 2014 | 2014 | Hospira Inc. | The Tyvek cover may not be completely sealed to the rest of the package allowing openings at the edge of the sterile package for the following products: List No. 04693-01, Latex Free Extension Set; List No. 06543-01, Nonvented Y-Set for Transurethral Resection; List No. 06544-01 Cystoscopy / Irrigation Set; List No. 06599-01, Large Bore Y-Irrigation Set (Latex-Free); and, List No. 15239-01, 4-Lea | Process design | Hospira sent an URGENT MEDICAL DEVICE RECALL letter dated March 25, 2014 to all affected customers. The letter included instructions for the direct accounts to: 1) check their inventories and quarantine any of the affected products; 2) complete and return the attached reply form to the fax number or e-mail address on the form; 3) inform potential users of these products of this recall; 4) return the affected products to Stericycle using the labels provided with the letter (additional labels or reply forms can be obtained by calling Stericycle at 877-377-5128, Monday - Friday, 8:00 AM - 5:00 PM EST); and, 5) if the products were further distributed, notify those customers and request that they contact Stericycle at 877-377-5128 (Monday - Friday, 8:00 AM - 5:00 PM EST). An updated URGENT MEDICAL DEVICE RECALL letter dated March 28, 2014 was sent to all affected customers. The letter contained instructions similar to those found in the March 25, 2014 letter. |  | 42292 | 42,292 sets (US only) | Worldwide Distribution - US: Nationwide in the states of (AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, UT, WA, WI, WV, WY) including District of Columbia and Puerto Rico; *** the countries of Barbados, Canada, Costa Rica, Dominican Republic, Hong Kong, Kuwait, Libya, Saudi Arabia, United Arab Emirates | Open, Classified |  |  | 2962342 | Not_Computer
Z-1628-2014 | 67923 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Lithium Ion Battery M4605A and M4607A for use with IntelliVue Patient Monitors. The monitor is used for monitoring and recording of, and to generate alarms for, multiple physiological parameters. |  |  | April 04, 2014 | 2014 | Philips Medical Systems, Inc. | The risk of battery failure increases with age, when a battery remains in use longer than 3 years after date of manufacture or 500 charge-discharge cycles. Such failure can result in overheating that in rare cases can cause the battery to ignite or explode. | Unknown/Undetermined by firm | Philips sent an Urgent: Medical Device Correction letter dated April 4, 2014 to all affected customers. The letter identified the affected product, problem, action to be taken and instructions for Use Addendum. Customers were instructed to review this information with all staff members who are involved in the operation and battery management of the IntelliVue patient Monitors and need to be aware of the contents of this communication. The addendum should be stored with the Monitor Instructions for Use. Please contact your local Philips representative or call us at 1-800-722-9377 with questions or concerns about this correction. |  | 188520 | 188,520 devices | Worldwide Distribution - Including US Nationwide, Canada, Czech Republic, Denmark, Netherlands, Estonia, Finland, France, Germany, Greece, Hungary, Italy, Japan, Kazakhstan, Norway, Poland, Romania, Russia, China, Slovakia, Sweden, and Turkey. | Open, Classified |  |  | 188520 | Not_Computer
Z-1631-2014 | 67924 | No | OSI | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | ECATS E 850 Bariatric Bed, True Air Technologies, Inc. |  |  | March 18, 2014 | 2014 | SCM True Air Technologies LLC | The firm is recalling the products due to lack of design controls, lack of a Device Master Record, and failure to approve specific vendors. | Nonconforming Material/Component | SCM True Air Technologies LLC sent an Urgent Medical Device Recall letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to remove the affected product from use and return to True Air Technologies immediately. The units should be decontaminated and bagged with a certification of disinfection before return to True Air Technologies. True Air will provide replacement units to replace the returned units. Customers were instructed to refer to the Medical Device Recall Return Response Form attached for acknowledgement of receipt and contact information for this voluntary recall. Customers with questions should call 800-682-7163. |  | 48 | 48 beds | Nationwide Distribution including GA, KY, and MI. | Completed |  |  | 155 | Not_Computer
Z-1633-2014 | 67833 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Arrowgard Blue and Arrowgard Blue Plus Pressure Injectable Central Venous Catheter, a multiple-lumen catheter that permits venous access to central circulation. Catalog number ASK-45703-PHF. |  |  | July 17, 2013 | 2013 | Arrow International Inc | The labeling states that the kit contains Introducer Needle: Echogenic 18 Ga. x 2.5 (6.35cm) XTW and 5mL ArroW Raulerson Spring-Wire Introduction Syringe; however, the kit contains, and should state, Introducer Needle: 18 Ga. x 2.5 (6.35cm) XTW and 5mL ArroW Raulerson Spring-Wire Introduction Syringe. As a result of this labeling inconsistency, if the clinician were to place the needle under x-r | Labeling False and Misleading | Arrow sent an Urgent Medical Device Recall Notification letter dated July 17, 2013, was sent to the customer requesting they discontinue use and quarantine any remaining affected product for return to Arrow.Our records indicate that you have received products that are subject to this voluntary recall. We are now asking that you take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the product and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-800-343-2935,Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. Your sales representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions and assistance for the return of product to Arrow International.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-800-343-2935, Attn: Customer Service. This will allow us to document your receipt of this letter. |  | 4285 | 4285 | US Distribution only to MI. | Terminated | March 30, 2015 | 621 | 4285 | Not_Computer
Z-1634-2014 | 68240 | No | DRY | Cardiopulmonary bypass on-line blood gas monitor. | Cardiovascular | Cardiovascular | 510(k) | CDI Blood Parameter Monitoring System; Catalog Numbers or Part Numbers - 145891, 500AVHCT; Serial Numbers - BPM SN B015552 (installed into CDI 500 Monitor SN 3535) |  |  | May 09, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | There is a potential for inaccuracies in blood gas readings/measurements. There is a potential risk of respiratory alkalosis or cerebral vasoconstriction with low paCO2. If the Hemoglobin and Hematocrit measurements are not accurate, there is a potential risk of unnecesary blood transfusion with resulting risk of infection or inflammatory reaction. BPM shipped without meeting manufacturing accep | Release of Material/Component prior to receiving test results | On 5/9/14, TCVS placed a phone call to the one affected customer regarding the recall. The phone call explained the reason for the recall, identification of affected product, potential hazard, and instructions for return of affected product. |  | 1 | 1 | one customer in IA | Terminated | July 11, 2014 | 63 | 1 | Not_Computer
Z-1638-2014 | 68224 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | MAMMOMAT Inspiration. |  |  | April 25, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential and possible hazard to the user when using the MAMMOMAT Inspiration PC monitor at the control desk, in that the holder of the PC monitor can break causing an unstable monitor to fall causing possible serious injury. | Device Design | A safety advisory notice, dated April 25, 2014, was sent to direct accounts to notify them of the issue and provide instructions to avoid the potential risk of injury. |  | 5 | 5 | Distributed in the states of GA, FL, NY, NV, and MA. | Terminated | December 15, 2014 | 234 | 5 | Not_Computer
Z-1635-2014 | 68012 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microline Renew Fenestrated Control Tip Grasper, Single Patient Use, Product Number: 3222 |  |  | April 14, 2014 | 2014 | Microline Surgical | Grasper jaw may break when force is applied to the jaw | Nonconforming Material/Component | Microline Surgical sent an Recall Notice dated April 14, 2014, via overnight mail to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Immediately examine your inventory and quarantine product subject to this recall. In addition, if you have further distributed this product, please identify your customers and notify them at once of this product recall. Your notification to your customers may be enhanced by including a copy of this recall notification letter.Please contact Microline Customer Service ( 978) 922-9810 for a RGA number to return all unused product. This recall should be carried out to the hospital level. Please complete and return the enclosed response form as soon as possible. |  | 540 | 540 units | Worldwide Distribution - US including the states of TX, IL, TN and NY, and the country of Chile. | Terminated | September 02, 2015 | 506 | 540 | Not_Computer
Z-1636-2014 | 68043 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TRIGEN INTERTAN NAIL: |  |  | April 04, 2014 | 2014 | Smith &amp; Nephew Inc | A manufacturing error resulted in the set screw being inserted too deep in the affected nails. | Process control | The firm initiated the recall via an "Urgent - Product Recall 1st Notification - Urgent" letter and e-mail on 04/04/2014, identifying the product, explaining the problem, and the potential risk associated with using the product. Customers are to inspect their inventory, and locate and quarantine any affected product. Additionally, the Inventory Return Certification form should be completed and returned along with any affected product identified, once a Return Authorization number has been obtained from Smith &amp; Nephew's Safety Affairs Department. |  | 1253 | 1,253 units total | Worldwide Distribution -- USA, France, Germany, Sweden, Switzerland, UK, Canada, Australia, Japan, China, India, Dubai, South Africa, South Korea, Mexico, and Venezuela. | Terminated | July 01, 2017 | 1184 | 1253 | Not_Computer
Z-1539-2014 | 68011 | No | BZO | Ventilator tubing. | Anesthesiology | Anesthesiology | 510(K) Exempt | Infant Nasal CPAP (continuous positive airway pressure) Prong, Nasal. |  |  | April 15, 2014 | 2014 | Fisher &amp; Paykel Healthcare, Ltd. | The Infant Nasal CPAP Prong has the potential to detach from the Nasal Tubing and therapy is likely to be interrupted. | Device Design | Fisher &amp; Paykel Healthcare sent an Urgent Medical Device Recall letter on April 11, 2014, to all affected customers. The letter informed them of the problems identified and the actions to be taken. Customers were instructed to complete Section A 'Inspection of Stock' on the Urgent Medical Device Recall Response Form and return it to their FPH Representative. Destroy the Affected Prongs by cutting them in half. Customers were also instructed to contact their FRH Representative for replacement product. For questions regarding this recall call 949-453-4000, ext 1423. |  | 62063 | 62,063 units total (14,420 units USA) | Worldwide Distribution - USA (nationwide) and Internationally to:ARGENTINAAUSTRALIABANGLADESHBELGIUMCANADACHILECHINAFRANCEGERMANYGREAT BRITAINGREECEHUNGARYINDIAINDONESIAIRAQISRAELITALYJAPANLIBYAMALAYSIANEPALNETHERLANDSNEW ZEALANDPERUPOLANDPORTUGALQATARROMANIARUSSIARWANDASAUDI ARABIASOUTH AFRICASPAINSRI LANKASWITZERLANDTAIWANTHAILANDTRINIDAD &amp; TOBAGOTURKEYUNITED ARAB EMIRATES | Terminated | September 30, 2014 | 168 | 62063 | Not_Computer
Z-1642-2014 | 67958 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | MXR-35 Diagnostic X-ray High Voltage Generator manufactured by DRGEM Corporation. |  |  | June 29, 2012 | 2012 | Fujifilm Medical System USA, Inc. | The MXR-35 X-ray Generator's labeling did not include the certification statement as required by 21 CFR Part 10 10.2. | Radiation Control for Health and Safety Act | Fujifilm will replace the violative labels with replacement labels provided by DRGEM, to bring the X-ray generator into compliance. You stated that you expect for all nine installed units to be corrected by the end of May 2014. The seven units in the Stamford, CT warehouse will be corrected when the generators are delivered to customer sites. Under 21 CFR 1003.31(c)(2), this exemption is granted subject to the following conditions:1. You must correct the missing certification statement for all units shipped into the US within the timeline you specified, by the end of May 2014.2. You must ensure that all future units that will be delivered and installed in the US will be in compliance with the labeling requirement, per 21 CFR 1010.2. |  | 14 | 14 | Worldwide Distribution - US Distribution including the states of TX and KY and Internationally to China. | Open, Classified |  |  | 14 | Not_Computer
Z-1649-2014 | 67858 | bandage, elastic | FQM | Elastic bandage. | General Hospital | General Hospital | 510(K) Exempt | Esmark Elastic Bandage (Sterile), 4" x 9'. Item DYNJ05116A, Individually wrapped, 20 bandages per case. Packaged by Medline Industries Inc. This product is used as an elastic bandage to support and compress a part of a patient's body. It is also used as a tourniquet to restrict blood flow to a part of a patient's body. |  |  | March 20, 2014 | 2014 | Medline Industries Inc | This lot did not go through the correct sterilization procedures. This product may potentially be non-sterile. | Employee error | Medline Industries issued an Immediate Action Required letter dated March 20, 2014 to all affected customers. The letters included instructions to: 1) immediately check inventories for the recalled products and quarantine the recalled products; 2) complete and return the enclosed destruction form listing the quantity of destroyed product (credit will only be issued if the completed form is received); and 3) If the customer is a distributor, promptly notify the distributor's customers that may have received the recalled products about this recall. Direct accounts and their customers that have any questions can contact Medline Industries at 866-359-1704. |  | 2860 | 2,860 bandages | Nationwide Distribution including AZ, CA, IA, IL, IN, KY, LA, MI, MS, NJ, NY, OH, PA, TX, and WA. | Open, Classified |  |  | 2860 | Not_Computer
Z-1647-2014 | 67787 | No | MKW | Cyclosporine test system. | Clinical Chemistry | Toxicology | 510(k) | Dimension VistaÂż CSA and CSAE |  |  | March 19, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | A complaint was received of Dimension VistaÂż CSA and CSAE under-recovery of cyclosporine when patient samples are run from the Small Sample Container. The under-recovery that has been observed is 18 to 44%. In March 2014, Siemens recalled due to the complaint received of Dimension VistaÂż CSA and CSAE reporting under-recovery of cyclosporine when patient samples are run from theSmall Sample C | Device Design | Siemens sent an Urgent Medical Device Correction letter dated March 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use the SSC with patient samples when running CSE or CSAE.Siemens also recommended a look back at previously generated CSA or CSAE results if the patient samples have been run from the SSC and that the content of this letter should be discussed with the Medical Director. Customers were asked to complete the attached form and fax it to (302) 631-8467.For questions regarding this recall call 800-441-9250. |  | 4175 | 4175 | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Belgium, Switzerland, Germany, Spain, France, Great Britain, Italy, Saudi Arabia, Australia, and Japan. | Terminated | August 10, 2015 | 509 | 4175 | Not_Computer
Z-1643-2014 | 67996 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Polyflux Revaclear, Capillary Dialyzer, Model Number 110633; |  |  | April 16, 2014 | 2014 | Gambro Renal Products, Incorporated | The firm received several complaints for internal blood leaks involving various lot numbers produced in late 2013. | Storage | The recall was initiated by an Urgent: Medical Device Recall letter delivered by UPS Overnight Delivery on April 16, 2014. The letter identified the affected product and reason for the recall. The firm is asking customers to remove and quarantine any remaining affected product from their inventory for collection and replacement. Customers are to also complete and return the Customer Reply Form to the Gambro Regulatory Affairs Department. Customers with affected product are to contact Gambro Customer Support or their Gambro representative to return the product. Questions should be directed to Customer Support at 1-800-651-2623. |  | 2461824 | 2,461,824 units total | Worldwide Distribution -- USA, including the territory of Puerto Rico, and the countries of Bermuda and Australia. | Terminated | December 08, 2015 | 601 | 2461824 | Not_Computer
Z-1648-2014 | 67643 | No | KLI | Total thyroxine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS TT4 Calibrator |  |  | February 20, 2014 | 2014 | Ortho-Clinical Diagnostics | Total T4 Calibrators and Reagent Packs may have calibration failures or low Quality Control results. | Other | Ortho Clinical Diagnostics sent an Urgent Product Correction Notifications dated February 19, 2014, to all affected customers via Fed Ex and/or US Postal Service Priority Mail. The letter identified the product the problem and the action needed to be taken by the customer.Customers are advised to immediately discontinue using and discard all remaining affected product. Customers should review their Quality Control results that were generated using the affected lot. Customers should consult their Laboratory Medical Director if they suspect that patient results were affected. Complete the confirmation of receipt form and return by March 4, 2014. Forward a copy of the Product Correction Notice if the product was distributed outside of your facility. Fax to :1.888.557.3759 or 1.585.453.4110 |  | 583 | 583 units Total (357 units domestically &amp; 226 units internationally) | Worldwide Distribution - USA (nationwide) and the countries of Australia, Brazil, Canada, China, Colombia, India, Japan, Mexico, Republic of Panama, Singapore, Venezuela, England, Germany, Italy and Spain. | Terminated | April 19, 2017 | 1154 | 583 | Not_Computer
Z-1646-2014 | 68262 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | Philips XRE, the dictation editor system included in the Philips XIRIS 8.1.50 |  |  | May 05, 2014 | 2014 | Philips Healthcare Informatics, Inc. | When using XIRIS Rio version (8.1.50.0) in conjuncture with iSite PACS a synchronization error can occur and if not noticed, information for the incorrect patient may be displayed | Software design | Philips Healthcare sent an Urgent Field Safety Notice dated May 5, 2014. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were advised of the synchronization problem and request that a copy of this notification be kept with the equipment instructions for use. The document goes on to outline the affected product, the problems and conditions when it can occur. A photo of the label for easy identity the actions to be taken are provided. Philips is informing users of the issue via the notification and will deploy a software correction to all customers when available.Questions may be directed to Philips reps or Care at 1-866-946-4701 or OneEMSProd@philips.com |  | 5 | 5 units | Worldwide Distribution - US including the states of FL and TX and the countries Canada and South Africa. | Terminated | April 28, 2015 | 358 | 5 | Computer
Z-1523-2014 | 67815 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Datascope System 98/98xt, CS 100/CS100i and CS 300 Intra-Aortic Balloon Pump. The Intra-Aortic Balloon Pump (IABP) is an electromechanical system used to inflate and deflate intra-aortic balloons in cardiovascular procedures. |  |  | March 21, 2014 | 2014 | Maquet Datascope Corp - Cardiac Assist Division | Potential mechanical failure of the fan assembly associated with the power supply. A fan assembly failure could result in the power supply overheating and cause the IABP to shut down without warning. | Nonconforming Material/Component | Maquet Inc. sent an Urgent Medical Device Field Correction letter dated March 21, 2014, to all affected customers. The letter identified the affected product, problem, and actions to be taken.. Customers were instructed that their Service Representatiave would contact them to schedule the replacement of the fan assembly at no cost to the customer. Upon completion customers were requested to sign a service repair order to verify satisfactory completion of the work. Customers with questions within Puerto Rico, were instructed to contact Technical Support in Mexico at +52-55-9000-8970, ext 104.For questions regarding this recall call 973-709-7652.A field correction letter was sent out on April 15, 2014 to additional consignees that did not receive the March 15, 2014 letter.Maquet issued press on May 9, 2014. |  | 11881 | 11,881 units (US 5183 and OUS 6698) | Worldwide Distribution - USA (nationwide) and approximately 104 countries. | Open, Classified |  |  | 11881 | Not_Computer
Z-1645-2014 | 67957 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 550 TxT Treatment Table of the Digital Linear Accelerator (LINAC) |  |  | April 01, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Table may lose calibration during patient treatment, which may cause the possibility of mistreatment. | Software design | Siemens sent an Medical Device Correction, Field Safety Notice of Improvement to End users in March 2014, informing them that a Siemens representative would install a software update.By installation of this software update the issues were solved by the following actions:1. A plausibility check was introduced for comparing the two independent encoder systems of each table axis.2. The time between two encoder comparisons was reduced to 50ms and an error code was introduced.Please include this Field Safety Notice in your Digital Linear Accelerator System Owner Manual chapter "Safety Advisory Letters" where it should remain.The relevant National Competent Authority has been informed of this update.We regret any inconvenience that this may cause, and we thank you in advance for your understanding. Further questions please call (610) 219-6300. |  | 120 | 120 | USA ( nationwide) Distribution. | Terminated | December 11, 2014 | 254 | 120 | Computer
Z-1640-2014 | 67620 | No | KOI | Electrical peripheral nerve stimulator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, NeuroMuscular Transmission Module, E-NMT-01 |  |  | April 09, 2014 | 2014 | GE Healthcare, LLC | When the E-NMT-01 module is used in conjunction with the ElectroSensor, the Neuromuscular Transmission (NMT) values may indicate a deeper level of muscle relaxation than the actual level of muscle relaxation. In the clinical situation visual movements of the hand are seen after TOF (Train of Four) stimulation, but the patient monitor shows no counts, or counts are not corresponding to the actual a | Software design | GE Healthcare sent an "Urgent Medical Device Correction" letter (Ref 36099) dated April 9, 2014 to affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information (contact Technical Support at 1-800-558-7044 or your local Service Representative).UPDATE: On May 11, 2015 a second GE Healthcare letter FMI 36099-2 was sent to OUS consignees. The updated report was for additional affected modules due to FRU installation. |  | 1794 | 1,794 (471 units US, 1,323 units OUS) | US Nationwide Distribution in the states of: AR, CA, CO,DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MI, MN, MO, NC, NV, NY, OH, OK, TN, TX, UT, WA, WI, WY.OUS: ARGENTINA. AUSTRALIA, AUSTRIA, BELGIUM, BOLIVIA, CANADA, CHILE, COLOMBIA, COSTA RICA, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, HONG KONG, HUNGRY, INDIA, IRA, ISRAEL, ITALY, JAPAN, KAZAKTSTAN, KUWAIT, LATVIA, LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, OMAN, PANAMA, POLAND, QATAR, ROMANIA, RUSSIAN, FÂżDÂżRATION, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SPAIN, SWEDEN, SWITZERLAND, TURKEY, TAIWAN, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA. | Terminated | September 28, 2015 | 537 | 1794 | Not_Computer
Z-1639-2014 | 67922 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Naturalyte Liquid Bicarbonate Concentrate, 6.4 Liter Bottle |  |  | April 10, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Naturalyte Liquid Bicarbonate maybe contaminated | Under Investigation by firm | Fresenius Medical Care North America issued an Urgent Medical Device Recall Letter dated April 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their stock immediately and determine whether they have any of the affected parts on hand. If customers have the affected product, they were instructed to discontinue use and place all Naturalyte Liquid Bicarbonate Concentrate in a secure area for return to Fresenius Medical Care Renal Therapies Group, LLC (FMC-RTG). Customers will be instructed to contact their Fresenius Medical Care Technical Services Team for instructions on how to return the recalled product. Firm will also conduct 100% telephone calls. If you have any additional questions, please contact your FMCNA Customer Service Team at 1-800-323-5188.Fresenius Medical Care sent an Urgent Expanded Medical Device Recall letter on May 1, 2014 for 9 additional lots. Customers were instructed to complete and return the attached FAX BACK FORM to confirm that they have received the notice and to indicate whether or not they have the affected product in their possession.For questions regarding this recall call 800-662-1237. Firm issued Press on 5/21/14. |  | 672784 | 672,784 units | Worldwide Distribution - USA (nationwide) and Canada. | Open, Classified |  |  | 672784 | Not_Computer
Z-1650-2014 | 68094 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | SW-2008 Snuggle Warm Sterile Cardiac Connective Warming Blanket; used with the Equator Convective Warmer, are intended for thermal regulation of a patient's temperture and are designed to prevent hypothermia and/ or reduce |  |  | April 22, 2014 | 2014 | Smiths Medical ASD, Inc. | The hose could separate from the blanket. | Process change control | On April 24, 2014, consignees were notified of the recall via letter. They were instructed to do the following: Inspect inventory and segregate any unused Blankets listed in the notification, complete and return the attached Confirmation Form by fax or e-mail within 7 days of receipt of this notice, and destroy the product. |  | 2120 | 2120 devices | Worldwide Distribution Including the States of CA, IL, KY, MS, NC, TN, TX, WA, WI and the Countries of BH, DE, NO, RU | Terminated | February 27, 2015 | 311 | 2120 | Not_Computer
Z-1651-2014 | 68223 | No | LBH | Cavity varnish. | Dental | Dental | 510(k) | Gluma Desensitizer Power Gel |  |  | April 23, 2014 | 2014 | Heraeus Kulzer, LLC. | The firm discovered that the Gluma Desensitizer Power Gel can be inadvertently extruded from the application syringe all at once when it is supposed to extrude on the desired area a small amount at atime. If extruded all at once, product could reach the dental patients mucosa, face or eyes or the dental professional could be exposed. This would cause irritation, burning or blistering. | Nonconforming Material/Component | Heraeus Kulzer sent a Voluntary Recall letter dated April 22, 2014 to all affected doctors and supply centers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to:1. Inspect your inventory, both in storage and in use, and identify the lot number on your kits and syringes. Refer to attached images.2. If you identify any Gluma Desensitizer PowerGel with the aforementioned kit or syringe lot numbers, immediately cease usage of this product and call Technical Services at 1 (574) 299-54093. Please advise if you forwarded this product to any other end users. We will need their contact information inClusive of name, address and phone number.For questions regarding this notification, please call our Technical Service team at (574) 299-5409. |  | 1278 | 5,112 syringes or 1,278 kits (1 Kit=4 syringes) | US Nationwide Distribution in the states of GA, TX, IL, CA, WA, WI, HI, MO, FL, NJ, TN, PA, NY, NC, OH, IN, CT, KY, AZ, NV, IA, | Terminated | November 05, 2014 | 196 | 1278 | Not_Computer
Z-1652-2014 | 67956 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Linear Accelerator (LINAC) models. |  |  | April 01, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens Radiation Oncology became aware that customers may be using the Siemens Linear Accelerator in combination with stereotactic accessories which have not been validated as being compatible with Siemens LINAC models. | Labeling design | Siemen sent an Urgent Medical Device Correction letter to all affected customers. The letter identified the affected product, problem and actions to be taken. For questions contact your local Elekta representative. |  | 9 | 9 | US Nationwide Distribution in the states of AL, ME, PA, IL, and CA. | Terminated | March 16, 2015 | 349 | 9 | Not_Computer
Z-1655-2014 | 67953 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur Systems TSH3 Ultra 2500 tests |  |  | April 10, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Samples containing fluorescein may show interference with the ADVIA CentaurÂż Systems TSH3 Ultra, Vitamin D and the ADVIA Centaur BRAHMS Procalcitonin assays. | Device Design | Siemens Healthcare Urgent Medical Device Correction (UMDC) was mailed to all affected Siemens Healthcare Diagnostics customers in the United States on April 11, 2014 for customer receipt on April 14, 2014. This notice informed the customers of fluorescein interference with the ADVIA Centaur Systems TSH3-Ultra and Vitamin D Total assays. This notice instructs customers that the TSH3-Ultra and Vitamin D assays can continue to be used for testing any sample that does not contain fluorescein. However, testing blood drawn from a patient, without allowing clearance of fluorescein is not recommended. This notice should be reviewed with the facility's laboratory or medical director.B. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on April 10, 2014. This notice informed the customer of fluorescein interference with the ADVIA Centaur Systems TSH3-Ultra , Vitamin D Total and ADVIA Centaur BRAHMS Procalcitonin assays. This notice instructs customers that the TSH3-Ultra, Vitamin D andProcalcitonin assays can continue to be used for testing any sample that does not contain fluorescein. However, testing blood drawn from a patient, without allowing clearance of fluorescein is not recommended. This notice should be reviewed with the facilitys laboratory or medical director. (Note: The ADVIA Centaur BRAHMS Procalcitonin assay is not sold in the US.) |  | 8052 | 8,052 kits | Worldwide Distribution-USA (nationwide) and the countries of Afghanistan Algeria Armenia Austria Bahrain Belgium Bulgaria Burkina Faso Canary Islands Croatia Cyprus Czech Republic Denmark Estonia Finland France Germany Great Britain Greece Hungary Iceland Iran Iraq Ireland Israel Italy Jordan Kazakhstan Latvia Lebanon Lithuania Luxembourg Macedonia Malta Netherlands Norway Oman Pakistan Poland Portugal Qatar Romania Russian Fed. Saudi Arabia Serbia Slovakia Slovenia Spain Sweden Switzerland Turkey Unit.Arab Emir. Vatican Yemen. | Terminated |  |  | 273253 | Not_Computer
Z-1660-2014 | 67963 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | LIFEPAK 1000 defibrillator. Affected Part Numbers: 320371500XXX |  |  | May 09, 2014 | 2014 | Physio-Control, Inc. | Physio-Control has become aware of incidents where customers have attempted to use theirLIFEPAK 1000 defibrillator and the device has shut down unexpectedly due to a very low battery.A defibrillator in this scenario has the potential to fail to deliver a shock, with the potential resultthat therapy is not delivered and a patient is not resuscitated. A software malfunction in the LIFEPAK 100 | Software Manufacturing/Software Deployment | On May 9, 2014 Physio-Control sent an Urgent Medical Device Recall Letter (dated May 2014) to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. The customer notification letter will provide new instructions on how to check the functional and power readiness of the device and advise the customer on the actions to take based on the symbol displayed.Customers are asked to record the readiness indicator(s) shown on each device and fax the confirmation sheet back to Physio-Control at 1-866-448-9567. Customers are also directed to contact Physio-Control at 1-800-442-1142 to obtain replacement batteries, if needed. Customers are provided a Device Readiness Guide for use in interpreting the Readiness Display on the device, and are directed to refer to the Operating Instructions which were provided upon purchase.In addition to notifying Physio-Control of any potential quality problems or adverse reactions or events experienced associated to the use of these products, Physio-Control asks customers to report directly to the FDA through the MedWatch Adverse Reporting Program online, by regular mail or by fax.Physio-Control reminds customers that it is critically important in understanding what the device and battery indicators mean on yourdefibrillator and what actions you need to take as a result. At any time the battery charge can be verified by following the instruction provided on page 2-5 of the Operating Instructions. It is also important that you always carry a spare fully-charged battery, as stated in the Operating Instructions.Customers are directed to call Physio-Control at 1-800-442-1142, 6:00 a.m. to 4:00p.m. (Pacific), Monday - Friday for any further questions.*********************************************************************************************On May 18, 2015, Physio Control sent the Urgent Medical Correction-Action Required LIFEPAK 1000 DEFIBRILLATOR to U |  | 96611 | 96,611 units total (40,974 in the US and 55,637 outside the US) | Worldwide Distribution-USA (nationwide) and the countries of ALBANIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BARBADOS, BELGIUM, BOLIVIA, BOSNIA AND HERZOGOVENIA, BRAZIL, BRUNEI DARUSSALAM, CANADA, CANARY ISLANDS, CAYMAN ISLANDS, CHILE, China, COLOMBIA, COSTA RICA, CROATIA, Curacao, CYPRUS, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EspaÂża, FAROE ISLANDS, FINLAND, FRANCE, FRENCH GUIANA, FRENCH POLYNESIA, GABON, GEORGIA, GERMANY, GREECE, GUAM, GUATEMALA, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, ISLAMIC REPUBLIC OF, IRAQ, IRELAND, ISRAEL, ITALY, Japan, JORDAN, KAZAKHSTAN, KENYA, KOREA, REPUBLIC OF, KUWAIT, LATVIA, LEBANON, LIBYA, LITHUANIA, MACAU, MACEDONIA, THE FORMER YUGOSL, MALAYSIA, MALDIVES, MALTA, MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NORWAY, OMAN, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, REUNION, ROMANIA, RUSSIA FEDERATION, SAN MARINO, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SRI LANKA, Suriname, South America, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, TURKMENISTAN, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, and VENEZUELA. | Open, Classified |  |  | 96611 | Computer
Z-1661-2014 | 68165 | Yes | LWP |  |  | Cardiovascular | PMA | Medtronic Temporary Pacemaker, Model 5348. Single Chamber External Pulse Generator (EPG) Temporary Pacemaker (AAI/VVI). |  |  | April 29, 2014 | 2014 | Medtronic Inc. Cardiac Rhythm Disease Management | Events related to a pacing rate outside of the intended setting, including events of sudden increased pacing rate up to the maximum setting of 180 pulses per minute (ppm), have been reported. This is caused by the development of high resistance on internal electrical connector contacts due to oxidation over time. | Device Design | Consignees were sent a Medtronic "Urgent Medical Device Correction" letter dated April 2014. The letter was addressed to Risk Manger, Health Care Professional. The letter included the Issue Description, Root Cause, Potentially Affected 5348 EPGs within the Serial Number Ranges, Malfunction Indications, and Recommended Actions. For questions related to EPG service policy or purchasing a replacement EPG, customers should contact your Medtronic representative. For technical questions, please call Instruments Technical Service at (1-800)638-1991. Medtronic asks that customers complete and sign the attached Customer Confirmation Certificate and fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality or e-mail to RS.CFOFCA@Medtrnic.com. |  | 16337 | 16,337 (8,989 US; 7,348 OUS) | Worldwide Distribution -- USA, including the territories of Guam and Puerto Rico; and, the countries of Albania, Algeria, Australia, Austria, Bahamas, Bangladesh, Belarus, Belgium, Benin, Bolivia, Bosnia And Herzegovina, Brazil, Bulgaria, Burkina Faso, Chile, Colombia, Costa Rica, Cote D'Ivoire, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Faroe Islands, Finland, France, Georgia, Germany, Greece, Guadeloupe, Honduras, Hong Kong, Hungary, India, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Luxembourg, Macau, Malaysia, Malta, Martinique, Mauritius, Mexico, Morocco, Nepal, Netherlands, Netherlands Antilles, New Caledonia, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Saudi Arabia, Senegal, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, and Yemen. | Terminated | October 08, 2014 | 162 | 16337 | Not_Computer
Z-1662-2014 | 68211 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IMPAX CV Pediatric Echo Reporting, IMPAX CV Echo Measurement Import via Optical Character Recognition (OCR) Service IMPAX. For the creation of digital structured reports for cardiovascular diagnostic studies such as echocardiography. |  |  | April 30, 2014 | 2014 | AGFA Corp. | Failure to effectively install and test a hotfix to correct the known software defect within the IMPAX CV 7.8 SU3-OCR Service. | Software design | AGFA Healthcare sent an URGENT FIELD SAFETY NOTICE LETTER dated April 30, 2014, to the consignee describing the problem that had occurred, the actions being taken by Agfa and the actions Agfa has required the site to perform. As part of the correction Agfa requires the consignee to review the affected reports to confirm that the data is correct. If errors are found the consignee will need to correct the data. The consignee was asked to acknowledge, via FAX-Back to 864-421-1664 or email, that the information was received and understood by the consignee. For questions the consignee was instructed to contact their local Agfa HealthCare organization at 877-777-2432.For questions regarding this recall call 864-421-1754. |  | 1 | One | Distribution to SC only | Terminated | September 18, 2014 | 141 | 1 | Computer
Z-1663-2014 | 68221 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM Aera magnetic resonance system. |  |  | April 24, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens discovered a processing error for the MAGNETOM Aera magnetic resonance system that may have resulted in some users not receiving Operator Manuals. | Labeling design | A customer safety advisory notice, dated April 24, 2014 was sent to notify direct accounts of the delivery of user manuals. |  | 5 | 5 | Distributed in the states of MN, CA, CO, TX, and IA. | Terminated | January 08, 2015 | 259 | 5 | Not_Computer
Z-1664-2014 | 68281 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Visitec Anterior Chamber Cannula, .40x22mm (27Gx7/8in) |  |  | May 06, 2014 | 2014 | Beaver-Visitec International Inc. | Contains a sharp tip needle when the product should contain a blunt tip cannula | Packaging process control | Beaver Visitec (BVI) sent a Product Recall Notice dated May 6, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately return any un-opened product of the affected lot so that BVI can replace the product or issue credit to their company. Customers were asked to fax or email the customer response form to 866-906-4304 or claimsUS@beaver-visitec.com as acknowledgement of the letter. Customers with questions should call 781-906-7986.For questions regarding this recall call 781-906-7950. |  | 9970 | 9970 units | Nationwide Distribution including AL, CT, IN, NC, NY, TX, and VA. | Terminated | September 05, 2014 | 122 | 9970 | Not_Computer
Z-1666-2014 | 68040 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | Sprotte Spinal Tray / Sprotte Spinal Needle 24 gauge with Introducer |  |  | April 03, 2014 | 2014 | B. Braun Medical, Inc. | B.Braun Medical, Inc. (BBMI) has received reports of 24 gauge Sprotte Needles that are missing the accompanying introducer needles. | Nonconforming Material/Component | B/Braun sent an Urgent Medical Device Recall Notification letter dated April 3, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the Device Recall Notification in its entirety and ensure that all users in their organization and other concerned persons are informed about this voluntary product recall. The letter should also be forwarded to their customers. Customers were instructed to determine their current inventory of the affected lots within their facility, DO NOT DESTROY ANY AFFECTED PRODUCT. Further use of the affected product should be discontinued immediately and quarantined. Customers should use the attached Product Removal Acknowledgement form to record the total number of inidividual units. If cusotmers have no remaining inventory they should check the box indicating zero inventory. Return the completed Product Removal Acknowledgement form to BBMI Quality Assurance department by fax to 610-849-1197 or email to PA_QualityAssurance.BBMUS_Service@bbraun.com within 2 weeks of receipt, even if the inventory is zero (0). If customers have any affected product they should call BBMI Customer Support Department at 800-227-2862 to arrange for return and replacement product. Customers with questions were instructed to call the Clinical and Technical Support Department at 1-800-854-6851.For questions regarding this recall call 610-596-2870. |  | 7400 | 7400 | Worldwide Distribution - USA (nationwide) and to Canada. | Terminated | December 02, 2014 | 243 | 7400 | Not_Computer
Z-1665-2014 | 68092 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professional, including radiologists, physicians, technologists, clinicians, and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display, manipulate, archive, print, and export images when connected wit the Centricity PACS infrastructure. |  |  | April 03, 2014 | 2014 | GE Healthcare | It has been identified that in a reject image workflow where multiple studies are simultaneously opened on the Centricity PACS RA1000 workstation and images are rejected from the active displayed study, images for the previously displayed study may also be rejected. The issue can occur with a previously displayed study that remains open in the background. If one or more, but not all, images in | Software design | An URGENT MEDICAL DEVICE CORRECTION letter dated April 3, 2014 was sent to all direct accounts. The letter included instruction to customers to use until the permanent fix for this issue is deployed: 1) Check to see if the DelRejImg task is active and running. a) Log into CA Tool. b) From the Navigation menu, select System Monitor. c) From the pull down menu, select Dbase Maintenance. d) Scroll to the DelRejImg task (this list is alphabetical). e) Review the Task Status. Possible statuses are IDLE or SUSPEND f) If the task status is set to SUSPEND The task is not running. g) If the task is set to IDLE, move to Step 2. 2) Check the Minimum Rejected Image Deletion Time. a) From the Navigation menu, select Sys. Configuration. b) From the pull down menu, select PACS Configuration. c) Scroll to Minimum Rejected Image Deletion Time (this list is alphabetical). d) Review the Value field. e) If the value field is a negative integer (such as the default setting of -1), the Delete Rejected Image process is NOT running and there is no patient safety hazard. If the value field is a positive integer (such as 1 or higher), the Delete Rejected Image process IS running and there is a potential patient safety hazard. 3) A GE Healthcare Service Representative will remove the DelRejImg task from the task table in the Centricity PACS database. This will prevent the DelRejImg task from running, and prevent any rejected images from being deleted. This feature is being removed entirely from the 3.2 release family of Centricity PACS and will be reinstated in a future 4.X workstation release of Centricity PACS. 4) Perform the reject image workflow with only one patient study open at a time. A future release of Centricity PACS will address the unintentional rejected image issue. Until a correction for the unintentional rejected image issue is available, a GE Healthcare service representative will remove the DelRejImg task table in the Centricity PACS database. This will remove the re |  | 297 | 297 units | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, and WV, and the countries of Australia, Canada, Chile, France, Germany, India, Ireland, Israel, Italy, Japan, Korea (South), Kuwait, Lebanon, Netherlands, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | June 05, 2015 | 428 | 297 | Not_Computer
Z-1667-2014 | 68195 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound Systems with software version 3.5 (VA35); ultrasound imaging system. |  |  | April 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Software issues with ACUSON SC2000 ultrasound with software version 3/5 (VA35) - may incorrectly display the color flow data in the wrong position and when Doppler Frequency is adjusted, the system does not restart the pulsed wave Doppler or continuous wave Doppler waveform. | Software design | Siemens sent an Customer Safety Advisory Notification letter on April 23, 2014, to all affected customers. Distributors in foreign countries will implement delivery of notification to their customers." Use the factory default presets with software version 3.5 instead of user-defined presets restored from software version 2.0or 3.0. Review all patient studies on exams performed after software revision 3.5 was installed on your system. If you are unsure when this software was installed on your system, please contact your local Customer Service representative." After adjusting the Doppler frequency, freeze and unfreeze the system. This action refreshes the display of the Doppler spectrum resulting in accurate measurements. Additional measurements performed after system freeze results in correct measurement values.Your local Customer Service Engineer will contact you to schedule an appointment to replace your existing version of software, free of charge. If you have any questions, please contact your local service support person forinformation regarding timelines and status.Until the new software has been installed on your system, please share this information with all personnel within your organization who need to be aware of this issue. As always, patient safety issues are a high priority, and we alert customers of product concerns which may occur under various conditions. To date, no patient injury has been reported as a result of these issues.We sincerely regret any inconvenience this condition may cause in your daily operations. For further questions call (650) 694-5398. |  | 257 | 257 units | Worldwide Distribution - USA including the states of AL, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, KS, LS, MD, MI, MN, MO, NH, NY, NC, OH, OR, PA, SC, SD, TN, TX, UT, VA, WA and WI., and the countries of Australia, Austria, Dominican Rep, France, Germany, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Mexico, New Zealand, Norway, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed, Saudi Arabia, Sweden, Switzerland, Turkey, U.A.E. and United Kingdom. | Terminated | November 12, 2015 | 568 | 257 | Computer
Z-1668-2014 | 68197 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound Systems with software version 3.5 (VA35) and using Stress Echo Feature. Ultrasound imaging system. |  |  | April 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Potential malfunction due to a software issue for ACUSON SC2000 ultrasound systems using software version 3.5 (VA35) when using the stress echo feature which could cause the need for the study to be repeated. | Software design | Siemens sent a Consumer Safety Advisory Notification letter on April 23, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to discontinue use of the stress echo feature with software version .5. For studies acquired with systolic-only captures: review all patient studies on exams performed after software revision 3.5 was installed on your system. If you are unsure when this software was installed on your system,please contact your local Customer Service representative. Foreign distributors will implement distribution of the field correction to their customers.For questions regarding this recall call 650-694-5398. |  | 177 | 177 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belarus, Dominican Rep., France, Germany, India, Indonesia, Israel, Japan, Mexico, New Zealand, Norway, Philippines, Poland, Republic of Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Sweden, Switzerland, Turkey, U.A.E., Ukraine, and United Kingdom. | Terminated | November 12, 2015 | 568 | 177 | Computer
Z-1671-2014 | 67527 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Intact PTH Calibrators consists of 6 bottles (4.0 mL each). Calibrator A contains Bis Tris Propane buffer with protein (bovine) stabilizer. Calibrators B-F contain PTH (synthetic peptide) in Bis Tris Propane buffer with protein (bovine) stabilizer. Preservatives: sodium azide and ProClin 300. |  |  | February 12, 2014 | 2014 | Abbott Laboratories | Abbott has confirmed that a performance shift in the ARCHITECT Intact PTH assay has the potential to generate falsely elevated results on patient samples. Results generated with impacted lots may demonstrate a positive shift relative to those generated with previous calibrator and/or control lots. The issue may also impact your established ARCHITECT Intact PTH reference ranges. Abbott ARCHITECT In | Under Investigation by firm | Product Information letters dated August 23, 2013 were sent to all direct accounts alerting the direct accounts to a positive bias observed in test results using some of the affected kits. The letter advised the direct accounts to take this information into account when interpreting patient results. Customers with questions were advised to contact Customer Service at 877-4ABBOTT (877-422-2688) (24 hours a day, 7 days a week) and to refer to Product Information Letter PI23AUG2013. Product Recall letters dated February 12, 2014 were sent to all direct accounts. The letter included instructions to: 1) immediately discontinue the use of, and destroy, any remaining inventory of ARCHITECT Intact PTH assay according to each customer's laboratory practices; 2) to continue Intact PTH testing at your laboratory, Abbott recommends that customers identify an alternate method for testing patient samples; 3) if further guidance is needed, please contact the customers's Abbott sales representative; 4) review the letter with the treating physician or healthcare provider to ensure they are aware of the issue and can determine whether re-evaluation of the patient is necessary, and that Abbott included a physician letter that you can use to notify your ordering physicians of this issue; 5) If the customer forwarded any product to other laboratories, please inform them of this Product Recall and provide a copy of the recall letter, customer reply, and physician letter to them; and, 6) complete and return the Customer Reply. Customers with questions were advised to contact Customer Service at 877-4ABBOTT (877-422-2688) (24 hours a day, 7 days a week). |  | 7780 | 7,780 calibrator kits | Worldwide Distribution - US Nationwide including District of Columbia and Puerto Rico; Algeria, Andorra, Angola, Argentina, Armenia, Austria, Azerbaijan, Bahamas, Bangladesh, Barbados, Belarus, Belgium, Bosnia &amp; Herzegovina, Brazil, Brunei, Bulgaria, Cameroon, Canada, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Ethiopia, Finland, France, Gaza and Jericho, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, ITALY, Jamaica, Japan, Jordan, Kenya, Korea (South Korea), Kuwait, Kyrgyzstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malaysia, Maldives, Mali, Mauritania, Mexico, Montenegro, Morocco, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestine, Panama, PanamÂż, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, San Marino, Saudi, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, SPAIN, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Ukraine, Uruguay, Venezuela, Vietnam, and Yemen | Terminated | May 22, 2015 | 464 | 17243 | Not_Computer
Z-1669-2014 | 68190 | No | GFG | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes ProDisc-C 2.0mm Milling Bit, Synthes 2.0mm Milling Bit HXC, Synthes 2.0mm Milling Bit with 2.35mm x 33.5mm Shaft, and Synthes 2.0mm Milling Bit with Flat Coupling. Spinal surgical instrument: The Synthes ProDisc-C 2.0mm Milling Bits are used as a part of the milling option for keel preparation of the vertebral bodies. The milling bits are coupled to a power tool and interface with the milling guide, which is connected to a trial implant for the keel preparation step. |  |  | April 29, 2014 | 2014 | Synthes, Inc. | Certain lots of the ProDisc-C 2.0mm Milling Bits have an increased potential to break during use. | Process control | DePuy Synthes sent a Recall letter, dated April 29, 2014, to affected customers. The letter identified the affected customers, problem and actions to be taken. The letter instructed customers to discontinue use and distribution of the device and to return affected devices to Synthes. For questions call 610-719-5450. |  | 6690 | 6690 | US Nationwide Distribution | Open, Classified |  |  | 6690 | Not_Computer
Z-1674-2014 | 68150 | Yes | MUQ |  |  | Cardiovascular | PMA | BG3515-5 US 5mL BioGlue Surgical Adhesive. |  |  | April 25, 2014 | 2014 | CryoLife, Inc. | Serum albumin component monomer failed to meet internally established end of shelf-life specification. | Process control | All consignees were notified of the voluntary actions by issuance of Field Correction/Removal Notice on April 24, 2014. This notices states which lot numbers are to be returned, and which lot numbers can be relabeled with a corrected expiration date. |  | 6190 | 6190 units | Distributed Nationwide including Puerto Rico and the states of AL, AK,AZ, AR, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NV, NH, NJ, NM , NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WV, and WI. | Terminated | November 10, 2014 | 199 | 6190 | Not_Computer
Z-1675-2014 | 68159 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Licox Test Kit REF BC10R Rx Only |  |  | April 16, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences has identified through a small number of complaints that some lot numbers of Licox Test Kits (Catalog Number BC10R) may provide erroneous "calibration test failure" message when they are used with the Licox PtO2 Monitor (LCSO2, or may provide an erroneous out of range result with the Licox CMP Brain Tissue Oxygen and Temperature Monitor (AC31). | Under Investigation by firm | Integra sent an "Urgent Voluntary Medical Device Recall/Return Response form notifications dated April 16, 2014 to US customers via traceable courrier service and to non-US customers via traceable email or facsimile. The customer notification advises them of the nature of the issue and the potential for adverse patient consequences; advises them to review their inventory and if they have BC10R Test Kit Resistors with an affected lot number to immediately stop using it; requests they identify and report to Integra if they do or do not have any affected BC10R affected lot numbers; and advises them how to return any affected products in their possession. Customers with questions were instructed to call 1-855-532-1723. |  | 49 | 49 test probes | Worldwide Distribution - USA including Illinois, Florida, Maryland, Massachusetts, North Carolina, Pennsylvania, South Dakota and Internationally to Australia, Austria, Belgium, Chile, Finland, Germany, Great Britain, Greece, Israel, Italy and Russia. | Terminated | October 24, 2014 | 191 | 49 | Not_Computer
Z-1673-2014 | 68293 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medrad Continuum MR Infusion system- non-wireless system |  |  | April 03, 2014 | 2014 | Bayer Corp | There is a potential safety risk associated with potentially damaged locking pins within the bracket of non-wireless versions of Continuum systems shipped prior to March 2008. | Component design/selection | Bayer sent an Urgent Medical Device Recall and Removal Letters dated April 3, 2014, and April 30, 2014, were sent to direct accounts to notify them of the bracket issue and to provide additional instructions for testing and/or returning the brackets.As a critical medical needs site currently using Continuum with a recalled bracket you have two options:1. Immediately return your Continuum(s) for refund.2. Take actions to ensure continued safe use of your system associated with the bracket recall. You will still be eligible for a refund when you return the system no later than June 30, 2015.When you are ready to return your Continuum(s), please refer to:Âż Enclosure 2: Disassembly Instructions.Âż Enclosure 3: Continuum Return Form.We sincerely regret the inconvenience and impact of the recalls and subsequent removal of Continuum from the market. Should you have questions regarding the pro-rated refund program, please call 877-229-3767. |  | 1347 | 1347 | Worldwide Distribution - USA (nationwide) including the countries of Australia, Brazil, Belgium, Canada,Germany, Singapore, South Africa, Poland, Mexico and UAE. | Terminated | April 16, 2015 | 378 | 1347 | Computer
Z-1696-2014 | 68124 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | BardÂż Ventralight" ST Mesh with Echo PS" 8"x10" Reorder Number: 5955810 |  |  | April 24, 2014 | 2014 | Davol, Inc., Subs. C. R. Bard, Inc. | Pouch holding sterile inflation assembly and inflation adapter may be open or have a weak seal. | Packaging process control | Davol, Inc., Subs. C. R. Bard, Inc.sent an Urgent Medical Device Recall letter dated April 28, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed to examine their inventory and quarantine product subject to the recall, contact Davol Customer Service at 1-800-556-6275 or C.R. Bard's Medical Services and Support Department at 1-800-562-0027 for instructions on how to return the product to Davol, Consignees were also instructed to complete and return the recall effectiveness check form, and notify their customers if the product was further distributed.For questions regarding this recall call 1-800-556-6275. |  |  |  | Worldwide Distribution - USA (nationwide) and Internationally to AUSTRIA, BELGIUM, CANADA, CYPRUS, ESTONIA, FINLAND, FRANCE,GERMANY, GREECE, IRELAND, ITALY, LUXEMBOURG, NETHERLANDS, NORWAY, POLAND, PORTUGAL, SOUTH AFRICA, SPAIN, SWEDEN, SWITZERLAND, and UNITED KINGDOM | Terminated | December 23, 2015 | 608 |  | Not_Computer
Z-1681-2014 | 68006 | No | MQP | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | Synthes XRL Medium Trial - Central Body, an instrument used as part of the Synthes XRL System, a vertebral body replacement device. For spinal surgery. The Synthes XRL device is intended to be used with the Synthes supplemental internal fixation systems (e.g., USS, including MATRIX, Pangea, and TSLP). |  |  | April 08, 2014 | 2014 | Synthes, Inc. | Two lots of the Synthes XRL Medium Trial-Central Body device were incorrectly etched with the incorrect dimensions for height. | Process control | Synthes sent an Urgent Notice: Medical Device Recall letter dated April 8, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action Required:Synthes has record that your facility has the product(s) subject to this recall. Synthes asks that you review your inventory and immediately remove the affected lots from stock.If you DO have any of the identified devices, please take the following steps: Contact your local Synthes Spine Sales Consultant to return any inventory located within your facility.Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found. Please include your name, title, telephone number and signature in the spaces provided.Return the Verification Form (page 4 of this letter) to: Fax: 877-270-9246 orScan/email: Synthes3041@stericycle.comIf you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affected product has been located. Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information.Return the documents to Synthes by: Fax: 877-270-9246 or Scan/email: Synthes3041@stericycle.comNote: If the Verification Form is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 4 of the notification.If you have any questions, please call 610-719-5450 or contact your Synthes Sales Consultant. |  | 3 | 3 | US Distribution including the states of FL, NJ and MN. | Terminated | September 10, 2015 | 520 | 3 | Not_Computer
Z-1678-2014 | 68179 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | IntelliSpace Portal (ISP), (ISP DX/HX/EX, Model number 881001; ISP IX, Model number 881030; ISP LX SPECT, Model number 881035), Philips Medical Systems. A medical software system for diagnostic imaging data viewing. |  |  | March 11, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | During SPECT reconstructions using Attenuation Correction and Scatter Correction, no Scatter Correction is being applied in the AutoSPECT Pro application on IntelliSpace Portal. In addition, resolution recovery is not applied correctly in SPECT reconstructions using Astonish in AutoSPECT Pro on Intellispace Portal. | Software design | On 03/10/2014 the firm sent Urgent Medical Device Correction Letters to their customers. |  | 157 | 157 units | Worldwide distribution, including US nationwide, Argentina, Australia, Austria, Canada, Cayman Islands, Chile, Denmark, Finland, France, Germany, India, Iraq, Israel, Italy, Latvia, Malaysia, Mexico, Netherlands, Panama, Peru, Poland, Saudi Arabia, Singapore, Slovakia, Spain, Switzerland, Turkey and United Kingdom. | Terminated | August 04, 2015 | 511 | 157 | Not_Computer
Z-1677-2014 | 68263 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen (TM) Laboratory Automation System configured with Data Innovations, LLC Instrument Manager (TM) (IM) v8.06 --- For in vitro quantitative measurement of a variety of analytes of clinical interest --- |  |  | March 20, 2014 | 2014 | Ortho-Clinical Diagnostics | Software Anomaly: Ortho Clinical Diagnostics (OCD) has become aware that if a sample requires follow-up action, such as dilution or retesting due to a mechanical issue, it is possible for that sample to be routed to the sample storage drawer (Storage) prior to the completion of the necessary pending action. A delay in patient test result(s) reporting could occur as operator is not alerted that results have not been produced and that further actions are required. | Software design | US consignees were sent Urgent Product Correction Letters (Cl2014-102, dated 3/20/2014) on 3/20/2014 by FedEx or USPS (for PO Box addresses only) overnight mail, informing them of the issue and instructing them to ensure that all required data elements are entered and in the correct format when using the auto verification feature or custom rules. J&J; Foreign affiliate consignees were notified by e-mail on 3/20/2014 informing them of the issue. |  | 81 | Domestic: 20 units, Foreign: 61 units | Worldwide Distribution-USA (nationwide) including Puerto Rico, and the countries of Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Terminated | April 11, 2017 | 1118 | 81 | Computer
Z-1699-2014 | 68080 | No | LXB | Hearing Aid. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Cochlear Baha Soft tissue gauge 6mm. |  |  | April 09, 2014 | 2014 | Cochlear Americas Inc. | Cochlear America is recalling Cochlear Baha Soft Tissue Gauge due to possible manufacturing imperfection in the smooth surface of the metal that could potentially leave residual metal on the patient's skin. | Process control | Customers were notified via letter on 4/17/14. This communication included an explanation of the recall and instructions to quarantine and return products. |  | 288 | 288 | Nationwide Distribution. | Terminated | October 06, 2014 | 180 | 288 | Not_Computer
Z-1698-2014 | 68065 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Novocastra Liquid Mouse Monoclonal Antibody Glutathione S-Transferase pi (NCL-L-GSTpi-438), a liquid tissue culture supernatant. For in vitro diagnostic use. |  |  | February 26, 2014 | 2014 | Leica Microsystems, Inc. | Novocastra Liquid Mouse Monoclonal AntibodyGlutathione S-Transferase pi (Product Code: NCL-L-GSTpi-438) does not function as intended up to the expiry date on the product labeling. This was detected through an ongoing stability program. There is a linkage between the age of the product and staining intensity. | Incorrect or no expiration date | Leica Biosystems sent a Medical Device Recall Notification letter dated February 26, 2014 to all direct accounts. The letter advised that clinicians may want to review results obtained from the affected products. The letter also included instructions for customers to: 1) appropriately destroy any unused or partially used NCL-L-GSTpi-438 (lots 6018777, 6023260,6026297, 6026980) reagent or confirm that they are not in stock; 2) complete and return the attached Medical Device Recall Notification Acknowledgement Form to 847-236-3747; and, 3) contact the local Leica representative immediately if replacement is required. Customers with questions or concerns about this recall can contact Leica via e-mail at LMGRA@leica-microsystems.com.For questions regarding this recall call 847-405-5413. |  | 33 | 33 units | Nationwide Distribution including AZ, CA, FL, and NY. | Terminated | August 14, 2017 | 1265 | 33 | Not_Computer
Z-1697-2014 | 68284 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Integra Newdeal SURFIX Fixed Angle Locking Screw 3.5mm x 22mm (catalog number 285322S) and 3.5mm x 30mm (catalog number 285330S). Screws are provided in different diameters and lengths and are used to anchor various sizes and shapes of bone fixation plates when treating bone fractures or reconstruction of the ankle, foot or hand. |  |  | March 28, 2014 | 2014 | Integra LifeSciences Corporation | New Deal has identified through a complaint report that some 3.5 mm Diameter x 22 mm Length Surfix screws in Lot no. F626 (Catalogue No. 285322S) have been packaged and labeled as 3.5mm Diameter x 30 mm Length Surfix screws (Catalogue No. 285330S). | Package design/selection | Integra issued an Urgent Voluntary Medical Device Recall letter dated March 28, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory for affected product, immediately stopo using or distributing them. Customers were asked to complete the Acknowledgement Form and return by email or fax as indicated on the form. When the form is received customer service will contact customers to provide an RMA number and directions to return the affected products. Customers with questions were instructed to call 1-855-532-1723.For questions regarding this recall call 609-275-2700. |  | 417 | 417 units | Worldwide Distribution - USA including AR, CA, NC, TX and Internationally to Austria, Chile, Denmark, Finland, France, Germany, Greece, Ireland, Italy ,Lebanon, Portugal, Russia, South Africa, Spain, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | April 18, 2016 | 752 | 417 | Not_Computer
Z-1680-2014 | 68305 | No | OEI | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Wound Drain - Label reads in part "Elutia ***Hubless Coated Silicone Surgical Wound Drain***" |  |  | May 15, 2014 | 2014 | Bacterin International, Inc. | The recalling firm reported there was inadvertent omission of quarterly (Q4 - 2013 and Q1 - 2014) dose audits/verifications on the affected products. All of the affected products were appropriately sterilized | Use error | Bacterin sent e-mail and FedEx letter with return receipt request to it's single consignee on May 15, 2014. Consignee is asked to separate the recalled product in a secure location for return to Bacterin International, Inc. Consignee is asked to complete tracking/verification form included in the letter/e-mail and call 1-888-886-9354 from 8:00AM to 5:00PM (Mountain Time) or send an e-mail to CS@Bacterin.com to obtain a Return Authorization (RA) number. |  | 390 | 390 | Distributed to one foreign consignee in Lebanon. | Terminated | June 02, 2014 | 18 | 390 | Not_Computer
Z-1676-2014 | 68091 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Aesculap Inc.(AIC) Miethke Ventricular Catheter with Deflector/Miethke Shunt System |  |  | April 21, 2014 | 2014 | Aesculap, Inc. | AIC (USA) received information regarding 3 complaints in which the deflector did not move freely on the ventricular catheter. | Device Design | Aesculap sent an Important Correction &amp; Removal recall letter dated April 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. An Aesculap Sales Representative will contact you to schedule return of any affected inventory and replace your product. Please complete the attached Inventory Sheet which is necessary to comply with FDA regulations. When completing the inventory sheet, please be sure to fill in the quantitybeing returned. If you cannot locate the product, please provide an explanation as to why the inventory will not be returned (discarded, etc.).AIC (USA) appreciates your cooperation on this matter and apologizes for the inconvenience this may cause. Thank you for your patience and continued support of this product. Please contact (610) 984-9414 with any questions. |  | 48 | 48 | US Distribution including the states of PA, KY, NY, RI, SC, KS, TX, CT, TN and WA. | Terminated | October 20, 2014 | 182 | 48 | Not_Computer
Z-1702-2014 | 67453 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | LifeCare 5000 Pump (Plum 1.6) infusion pump, indicated for volumetric infusion of critical drugs and fluids by intravenous (I.V.) administration. |  |  | January 24, 2014 | 2014 | Hospira Inc. | There is the potential for the door roller assembly on the Plum LifeCare 5000 Series and Plum XL families of infusers to break. In the event of a broken door roller pin, the door cannot appropriately lock the cassette in the right position. Depending on the conditions of a broken door roller pin, a number of events can occur. The door cannot be closed or a cassette/door alarm or a false occlusion | Component design/selection | Hospira sent an URGENT MEDICAL DEVICE RECALL letter dated January 24, 2014, to all affected customers via traceable UPS. The letters included instructions for recipients to: inspect the door roller assemblies; if a door roller or pin appears loose, broken or missing, remove the device from clinical service until the door assembly is replaced with a corrected door assembly and contact Hospira at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) to report the issue. All devices requiring repair will be serviced and replaced with the existing design door and/or pin. Hospira is in the process of designing a metal door roller pin to replace the plastic on currently available on the Plum XL and once that design is complete, any door roller/pin serviced on a Plum XL will be replaced with the new metal design. Recipients are further instructed to: inform potential users in their organization of the recall notification; complete the attached reply form and return it to the fax number or e-mail address on the form; and if the pumps were further distributed, to notify the accounts that received the devices and ask them to contact Stericycle at 877-597-9581 (Monday - Friday, 8:00 AM - 5:00 PM EST) to receive a reply form. The URGENT MEDICAL DEVICE RECALL letter also reports that Hospira is in the process of retiring the Plum LifeCare 5000 and Plum XL families of infusion devices as noted in the IMPORTANT PLUM INFUSION PUMP DISCONTINUATION AND TRANSITION NOTICE issued May 1, 2013. As of June 30, 2015, Hospira will consider the products within the Plum LifeCare 5000 and Plum XL Infusion Systems families retired and will no longer support them.For questions regarding this recall call 224-212-2000. |  | 132889 | List Number 02507: 132,889 pumps | Worldwide Distribution - USA nationwide including District of Columbia, Puerto Rico, and US Virgin Islands; Australia, Argentina, Bahrain, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, Colombia, Croatia, Czech Republic, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, Israel, Italy, Jordan, Kuwait, Luxembourg, Mexico, Oman, Peru, Poland, Portugal, Qatar, Russia, San Marino, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Syria, Taiwan, Turkey, United Arab Emirates, United Kingdom, and Yemen. | Open, Classified |  |  | 337666 | Not_Computer
Z-1705-2014 | 68357 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Ultraview SL Command Module, Model 91496. |  |  | May 08, 2014 | 2014 | Spacelabs Healthcare Inc | A software anomaly causes temperature alarm limits to change to limits which are not physiologically possible and result in nuisance false positive alarms. | Software Manufacturing/Software Deployment | Spacelabs Healthcare identified a software anomaly in the Ultraview SL" Command Module, Model 91496. During an internal audit, Spacelabs identified a reportable recall that was not reported. We have a product called the Spacelabs Medical Ultraview SL" Command Module, Model 91496. A software error was discovered at a customer site which was not a safety issue, but in the interest ofcustomer service it was decided to downgrade all their modules with an older software version. |  | 64 | 64 devices | US Distribution to GA. | Terminated | June 03, 2014 | 26 | 64 | Computer
Z-1709-2014 | 68266 | [6 empty] | CARINAsim MAN-1236 Rev 2 and CARINAiso MAN-1237 Rev 2 laser software. LAP isocenter marking control software. Laser systems for patient alignment. |  |  | October 07, 2013 | 2013 | LAP of America Laser Applications, LLC | Potential for patient to be marked incorrectly. Customers currently using CARINAsim or CARINAiso systems can encounter two safety relevant errors in the operating instructions (CARINAsim MAN-1236 Rev 2 / CARINAiso MAN-1237 Rev 2) | Software design | On 10/16/2013 Customers were sent an Urgent field safety notice for Corrective action concerning the operating instructions of CARINAsim and CARINAiso. The letter identified the affected product, as well as the issues involved. Users were to review their current operating steps and validate their operating sequence. Users should have received a modified software version with software update in February 2014. All users should have been notified of this letter and a copy should be maintained in the users' records. Questions should be directed to Thomas Gaudet on 561-416-9250 at the Service Center. |  | 280 | 280 Manuals | Worldwide Distribution -- USA, including the states of AK, AZ, CA, CT, DC, FL, GA, IA, ID, IL, KS, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WY and the territory of Puerto Rico; and, the countries of Brazil, Canada, Chile, Colombia, Cuba, Honduras, and Mexico. | Open, Classified |  |  | 280 | Computer
Z-1708-2014 | 68153 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MonoswiftÂż PGCL [poly (glycolide-co-caprolactone)] Absorbable Surgical Suture; USP size 3/0 PGCL; part number L936. Product is labeled sterile. The product is sealed inside an inner foil pouch with an outside TyvekÂż pouch. It is an undyed, monofilament synthetic absorbable suture with needle, size 3/0 (2.0 metric), PS-1 Needle |  |  | April 23, 2014 | 2014 | C P Medical Inc. | CP Medical, Inc., is recalling one lot of Monoswift PGCL [poly (glycolide-co-caprolactone)] Absorbable Sutures for failing to meet the USP minimum average value for knot pull (tensile) strength. | Employee error | CP Medical, Inc., sent a Recall Notice letter dated April 28, 2014 to affected customers. The letter identified the affect product, problem and actions to be taken. The letter instructed customers to immediately discontinue selling or using the recalled product and return all existing stock to CP Medical. They also request their customers to contact downstream customers and request that all their existing stock be returned. CP Medical, Inc. asks all customers to provide them with information as to the quantity of recalled product they are able to return and for all stock they are not able to recover, customers are asked to provide quantity information and location of the product for accountability purposes. The firm requests that all customers contact their customer service representative at 1-800-950-2763 for a Return Goods Authorization (RGA) # and shipping account number. All shipping costs, including those from customers, will be paid for by CP Medical. Replacement product (at no cost) will be provided. |  | 528 | US distribution:4 boxes/ 12 sutures per box (48 total); Outside US: 44 boxes/12 sutures per box. (528 total). | Worldwide Distribution - US Nationwide | Terminated | September 12, 2014 | 142 | 528 | Not_Computer
Z-1707-2014 | 67926 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 5.5, an Image Processing System used for patient positioning for radiation therapy or radiosurgery. |  |  | April 15, 2014 | 2014 | Brainlab AG | When using a specific plan update workflow in combination with a non-Brainlab treatment planning system (TPS), there is a possibility in ExacTrac 5.5.5 to generate an unintended blend of two initially separate plans of the same patient. | Software design | Brainlab sent a "FIELD SAFETY NOTICE / PRODUCT NOTIFICATION" letter dated March 13, 2014 was sent to all affected customers. The letter identifies the affected product, problem, and actions to be taken. The letter includes specific instructions to users that are intended to reduce the likelihood of similar events occurring when using ExacTrac software version 5.5.5. The letter further states that Brainlab will be providing a software update to all customers. Brainlab will contact all customers to arrange for the software update within the following 6 months. Customers with any questions about this issue can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 40 | 40 systems (worldwide) | Worldwide Distribution - US Nationwide in the states of AZ, AR, CA, FL, KY, MA, NE, NY, TN, TX and the countries of Argentina, Australia, Brazil, Finland, India, Japan, Netherlands, Oman, Russia, and Taiwan | Open, Classified |  |  | 40 | Computer
Z-1700-2014 | 68298 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Hologic Selenia Full Field Digital Mammography System: SEL-00002, SEL-00013, SEL-00030, SEL-00039. |  |  | May 13, 2014 | 2014 | Hologic, Inc. | Selenia Full Field Digital Mammography System defective relay switch, if there is a need to use the emergency off switch, the system may not turn off which could result in patient injury | Component change control | Hologic notified customers by letter dated 5/13/14 via certified mail or Federal Express to inform customers of the problem and provide acknowledgement of the notification and confirmation that their systems have been correctedA Dealer notification dated 5/13/14 sent by certified mail or Federal Express to inform Dealers of the need to notify their customers of the problem and to correct the identified Selenia system using Hologic Technical Bulletin TB-00806. On 5/16/14 the firm expanded the recall to an additional units Mammography System M-IV (2 units) and Selenia Mammography Systems (4) by letter dated 5/14/14.If there are any questions regarding this field correction, please call Hologic Customer Support at 877 371-4372. |  | 11 | 11 units | Distributed in Puerto Rico, and the states of CA, NY, and TX, and the country of Republic of Korea. | Terminated | August 18, 2014 | 97 | 14 | Not_Computer
Z-1706-2014 | 68219 | No | KPY | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | Kenex Radiation Shield and surgical lamps installed with GE Healthcare Interventional, Surgery, CT and X-Ray Imaging systems. |  |  | May 02, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to the improper installation of a Kenex ceiling suspended radiation shield and surgical lamp. This issue is not related to the design of the Kenex product. The affected products are limited to installations completed or contracted by GE Healthcare.There has been a reported incident of a radiation shield falling from the ceil | Equipment maintenance | GE Healthcare sent an "Urgent Medical Device Correction" letter # 12209 dated May 2, 2014. The letter was addressed to Director of Biomedical Engineering, Director of Radiology and Chief of Cardiology. The letter described the Safety Issue, Safety Instructions, Affected product Details, Product Correction and Contact Information. For questions they should call 1-800-345-2700. For other countries, they should contact their local GE Healthcare Service Representative. . |  | 3221 | 3221 (108 US, 3113 OUS). | Worldwide Distribution - US Nationwide, ALBANIA, ALGERIA, ANGOLA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BOLIVIA, BOSNIA &amp; HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, MACEDONIA, MALAYSIA, MALTA, MARTINIQUE, MAURITIUS, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PHILIPPINES, POLAND, PORTUGAL, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SENEGAL, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE, UNITED KINGDOM, VENEZUELA, VIETNAM, and YEMEN. | Terminated | February 25, 2015 | 299 | 3221 | Not_Computer
Z-1718-2014 | 68358 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Patient Side Cart assembly, 4-arm, for use with Intuitive Surgical, Inc. da Vinci Surgical System IS 4000. The patient-side cart is positioned at the operating room table and contains four arms that are positioned over the target patient anatomy. The endoscope attaches onto any arm. The patient-side assistant attaches/detaches the endoscope and instrument intra-operatively. |  |  | May 09, 2014 | 2014 | Intuitive Surgical, Inc. | Torque wrenches used to tighten certain PSC bolts were out of calibration. This could lead to over- or under-tightening of bolts that may result in system faults, mechanical sound (e.g. scraping, squeaking, rattling), cracks or vibration in the system. | Nonconforming Material/Component | Intuitive sent an Urgent Medical Device Correction letter dated May 13, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. All personnel who use the device should be notified of the issue and the acknowledgement form completed and returned.Intuitive reps will visit each affect sites to install a replacement PSC and ensure that it is properly integrated with the Da Vinci.For questions contact Intuitive Surgical at 800-876-1310 Option 3 (Monday - Friday 6 am to 5 pm PST). |  | 2 | 2 | Medical facilities in TX and FL | Terminated | June 04, 2014 | 26 | 2 | Not_Computer
Z-1710-2014 | 68090 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Liquid Mouse Monoclonal Antibody CD15 (NCL-L-CD15), a liquid tissue culture supernatant containing 15 mM sodium azide as a preservative. For in vitro diagnostic use. |  |  | January 10, 2014 | 2014 | Leica Microsystems, Inc. | Product does not function as intended up to the expiry date on the product labeling. There is a link between the age of the product and staining intensity. | Incorrect or no expiration date | Leica sent an Urgent Field Safety Correction Notice letter dated January 10, 2014 was sent ot all affected customers. The notice included instructions for recipients to: 1) destroy any unused or partially used NCL-L-CD15 (lots 6015073 and 6017936) reagent or confirm that the products are no longer in stock; 2) indicate on the attached Field Safety Correction Notice Acknowledgement Fom which action was taken; and, 3) contact the local Leica represetnative immediately if a replacement is required. Customers with any concerns about this notice can contact Leica via e-mail at LMGRA@leica-microsystems.com. |  | 145 | 145 units | Worldwide Distribution - US Nationwide in the states of CA, CT FL, IL, MO, MS, NC, NY, PA, and the countries of Argentina, Australia, Brazil, Colombia, Italy, Malaysia, New Zealand, Nigeria, Poland, Romania, Russia, Saudi Arabia, Serbia, South Africa, Sweden, Taiwan, Tunisia, Turkey, and United Kingdom | Open, Classified |  |  | 145 | Not_Computer
Z-1711-2014 | 68110 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Luminos dRFsystem |  |  | April 11, 2014 | 2014 | Siemens Medical Solutions USA, Inc | It was determined that under rare environmental conditions (i.e., extremely dry air in combination with a highly isolated floor) electrostatic discharges (ESD) &gt;8 kV may trigger an unintended movement of the AXIOM Luminos dRFsystem which may lead to an emergency situation and pose danger to the patients, operating personnel or to the unit. | Device Design | Siemens sent an Safety Advisory Notice dated April 11, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.What steps can the user take to avoid the potential risk of this issue?Press one of the red emergency STOP buttons immediately in case of an unintended movement. The system needs to be restarted afterwards to be fully functional again.How will the issue finally be resolved?Siemens is preparing a modification of the Luminos dRF system control console that will resolve this potential malfunction. The field modification will be available by end of May 2014.We appreciate your understanding and cooperation with this safety advisory notice and ask you to immediately instruct your personnel accordingly. Please ensure that this safety advisory notice is placed in the system's instructions for use until the update has been installed.If you have sold this device/equipment and it is no longer in your possession, we kindly ask that you forward this safety notice to the new owner of this device/equipment. Please inform us about the new owner of the device/equipment.For further questions please call (610) 219-6300. |  | 37 | 37 | Class II Recall - US Distribution including the states of IL,VA, TN, PA, TX, MN, GA, MI, IN, OH, FL, CA, OR, NY, ME, MO, OK, CO and NJ. | Terminated | December 15, 2014 | 248 | 37 | Not_Computer
Z-1714-2014 | 68363 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | Toshiba Infinix X-Ray Interventional System INFX-8000V/UU. An X-Ray system that is capable of radiographic and fluoroscopic studies used in an interventional setting. |  |  | May 12, 2014 | 2014 | Toshiba American Medical Systems Inc | Although the Flat Panel Detector (FPD) touch sensor is activated during use, the error message is displayed and the audible signal continues sounding during the table movement of longitudinal tilt, the movement of longitudinal tilt does not stop. | Other | Toshiba America Medical Systems, Inc. initiated this recall by sending recall notifications via USPS return receipt mail.The recall notification letter, dated May 12, 2014, titled "URGENT: MEDICAL DEVICE CORRECTION", informed customers of the following: device information, serial number, SID, reason for reason, purpose of the notification, corrective action, request to customers, contact information, and customer reply form.Customers with questions regarding this letter were instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968.For questions regarding this recall call 714-730-5000. |  | 2 | 2 | Nationwide Distribution including CA and IN | Terminated | October 02, 2014 | 143 | 2 | Not_Computer
Z-1719-2014 | 68193 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MACROLYTE Premie Dispersive Electrode, REF 440-2400, CONMED CORPORATION. The device is applied to the patient during electrosurgical procedures to provide a path for RF current to leave the patient and return to the generator. |  |  | May 08, 2014 | 2014 | ConMed Corporation | ConMed Corporation has received complaints regarding sparking, no output and burning at the cord set connection. The complaints have been confirmed by ConMed. | Nonconforming Material/Component | ConMed sent an Urgent Medical Device Recall letter dated May 8, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Corrective Action: All returned devices will be received into the facility and the lot number and quantity recorded on to an electronic spreadsheet. The returned product will then be stored in a secure quarantined area. At this time, management has not determined as to how the devices will be dispositioned. A protocol(s) will be written to address the destruction of the returned devices. This will be submitted to NYK-DO for approval prior to any action being taken with the returned devices. CAPA 153136 was opened. Fixtures were revised to allow for placement of the foil/gel substrate from the bottom of the foam. This is to maintain consistent dimensions on the product and not allow for float/movement of the substrate from being offset.For further questions please call (315) 624-3024 or fax to (315) 624-3030. |  | 56880 | Domestic - 12,600 units, Foreign - 44,280 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Belgium, Brazil, Canada, Ecuador, El Salvador, Finland, France, Great Britain, Honduras, India, Indonesia, Israel, Italy, Japan, Latvia, Lebanon, Malaysia, Mexico, Panama, Philippines, Portugal, Saudi Arabia, South Africa, Spain, Thailand, Trinidad, United Arab Emirates and Uruguay. | Terminated | April 11, 2017 | 1069 | 56880 | Not_Computer
Z-1720-2014 | 68039 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | SBi RingFix Olive Wires |  |  | April 07, 2014 | 2014 | Small Bone Innovations, Inc. | Small Bone Innovations (SBI) is recalling non-sterile RingFIX Olive Wires labeled as item #451-3004 Drill tip 1.8 mm diameter 400mm olive k-wire packaged with item #451-3005 Bayonet tip 1.8mm diameter 400mm olive k-wire inside. | Packaging process control | SBi sent an Urgent Medical Device Recall dated April 7, 2014, to affected customers (sales consultants, distributors, and end users). The letter identified the affected product, problem and actions to be taken. For questions call 215-428-1791 ext 248. |  | 44 | 44 | US Nationwide Distribution in the states of IN, NC, WA, NH, and CT. | Terminated | October 20, 2014 | 196 | 44 | Not_Computer
Z-1713-2014 | 68073 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | ArrowÂż Percutaneous Sheath Introducer Kits |  |  | April 14, 2014 | 2014 | Arrow International Inc | Arrow is recalling the ArrowÂż Percutaneous Sheath Introducer Kits due to dilator tip damage noted during internal inspection. Dilator tip damage has the potential to result in vesseldamage. | Process control | Teleflex/Arrow International sent an Urgent Medical Device Recall Notification letter dated April 17, 2014, to distributors and end users. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product and to complete the enclosed Recall Acknowledgement Form and fax it to 1-800-343-2935 to return product. Customers would be contacted by a customer service representative for a Return Goods Authorization (RGA) Numbers and would be provided instructions for the return of the product. Customers should complete the form whether they have affected product or not. Customers with questions were instructed to call 1-800-343-2935.For questions regarding this recall call 610-378-0131. |  | 15424 | 15424 | Worldwide Distribution - USA (nationwide) and Canada. | Open, Classified |  |  | 15424 | Not_Computer
Z-1722-2014 | 68245 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems Computed Tomography X-ray systems (Brilliance CT series, iCT, iCT SP, Ingenuity CT, Ingenuity Core, Ingenuity Core128, Ingenuity Flex, and the MX8000 Dual v. EXP scanners) |  |  | April 01, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | A customer reported that the table top had become free floating. A Field Service Engineer (FSE) confirmed that the service latch had become disengaged. | Component design/selection | Philips Healthcare sent an Urgent Medical Device Correction letter dated April 1, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed If they experience a horizontal, free-floating couch motion, to contact their Field Service Engineer immediately. Customers were instructed to retain a copy of this Field Safety Notice with the equipment Instructions forUSER Use (IFU). Customers with questions or needing further information were instructed to contact their local Philips representataive or local Philips Healthcare office. For North America, contact the Customer Care Solutions Center at 1-800-722-9377, Option 5.For questions regarding this recall call 440-483-3032. |  | 6839 | 6, 839 units | Worldwide Distribution: US nationwide including PR and DC, Afghanistan, Albania, Algeria, Argentina, Australia, Austria, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bhutan, Bolivia, Bosnia, Brazil, Brunei Darussalam, Bulgaria, Canada, Cambodia, Cape Verde, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Dubai, Dominican Republic, Ecuador, Egypt, Ethiopia, Finland, France, French Polynesia, Gabon, Georgia, Germany, Ghana, Greece, Greenland, Guadeloupe, Herzegovina, Hong Kong, Hungary, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kenya, Kuwait, Kyrgyzstan, LAO People's Dem, Rep., Latvia, Lebanon, Libya, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malawi, Malaysia, Malta, Martinique, Mauritius, Mexico, Monaco, Mongolia, Morocco, Mozambique, Myanmar, Nahar, Nepal, Netherlands, New Caledonia, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Turkmenistan, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Viet Nam &amp; Yemen. | Open, Classified |  |  | 6839 | Not_Computer
Z-1723-2014 | 68295 | No | LKK |  |  | General Hospital | PMA | Medtronic Model 8780 Ascenda Intrathecal Catheter. |  |  | May 02, 2014 | 2014 | Medtronic Neuromodulation | Medtronic has determined that an Ascenda catheter kit may have had endotoxin in excess of USP limits. | Release of Material/Component prior to receiving test results | There was only one consignee affected by the recall. The consignee was contacted by phone on 5/2/14 and the information was provided verbally. General Manager &amp; Vice President, Pain Stimulation/Targeted Drug Delivery spoke via phone with the physician managing the patient. The following information was provided:"Medtronic identified the potential for endotoxin levels exceeding USP requirements on product recently implanted (Model 8780, serial number N457062003)."There may be the potential for serious patient injury due to endotoxin within the first hours after implant."Medtronic recommends observing the patient for symptoms of meningitis. Also to be aware that CSF would not test with a positive culture because endotoxin induced meningitis results in an aseptic meningitis.The physician expressed confidence in the management of the patient and stated he was calling the nursing care facility to increase monitoring of the patient. Subsequent information was received regarding the patient indicating there were no signs of an endotoxin reaction. |  | 1 | 1 | Distributed in the state of CT. | Terminated | October 07, 2014 | 158 | 1 | Not_Computer
Z-1721-2014 | 67967 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Siemens SOMATOM Emotion 16, Emotion 6, and SOMATOM Spirit Computed Tomography Systems |  |  | March 21, 2014 | 2014 | Siemens Medical Solutions USA, Inc | The rubber damper may degrade and break, causing loosening of the motor mounting plates, which could lead to a series of events that could place the operator and patient at risk of injury. | Component design/selection | Siemens sent a Customer Safety Advisory Notice/ Field Safety Notice letter dated March 21, 2014, to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that the system modification would be implemented by Siemens service specialists in April 2014 to correct the problem. |  | 345 | 345 | US Nationwide Distribution in the states of IA, TX, FL, NE, NY, OH, NC, KS, IL, CA, OK including PR. | Terminated | March 16, 2015 | 360 | 345 | Not_Computer
Z-1717-2014 | 68312 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens Ysio Systems |  |  | May 09, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Excessive usage of cleaning liquid on Siemens Ysio Systems may cause infiltration of the generator console, which may lead to a short circuit causing an unintended release of radiation. | Use error | Siemens sent an Safety Advisory Notice letter dated May 9, 2014, to affected customer. The letter identified the affected product, problem and actions to be taken. For question call +49 (913) 84-9908. |  | 433 | 433 | US Nationwide | Terminated | August 18, 2015 | 466 | 433 | Not_Computer
Z-1712-2014 | 68326 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PATH(R) THREAD CUP ADAPTOR, REF 2007-0013, Wright Medical Technology, Inc. Orthopaedic surgical instrument. |  |  | April 30, 2014 | 2014 | Microport Orthopedics INC. | Difficulty removing the cup adaptor from the impaction handle during surgery | Device Design | Firm has initiated the recall by email on 04/28/2014 to all distributor's offices, including international, then followed with written communication on 05/02/2014 via Fed Ex P1 delivery to all United States consignees. Follow-up to non-responding consignees will commence 30 days after initiation. |  | 140 | 140 units | Nationwide Distribution. | Terminated | December 01, 2014 | 215 | 140 | Not_Computer
Z-1716-2014 | 68261 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | The NAV3i Platform Power Box supplies the electrical components with power (e.g. Stryker Nav3 Graphite Camera, SPC-3 Computer, IO-Tablet (touch screen) and the monitor). |  |  | May 13, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | The Stryker NAV3i Platform Power Boxes produced between December 9, 2013 and March 7, 2014 may have been assembled without electrical insulation of the soldering joints | Process change control | On May 13, 2014, Stryker Instruments sent URGENT MEDICAL DEVICE RECALL notification to their consignees. Notification included the description of the affected product, reason for the recall, risk to health, and actions to be taken by the Customer/User. Once Stryker receives the response from the customer, a Stryker Representative will contact the customer to set up a time to perform the upgrade to the affected power box. Old Power Boxes will be returned to Stryker and placed in an industrial compactor/recycler for destruction. Questions regarding the recall are directed to Kelly Jo Whipple M-F 8am-5pm @ 269-389-2921 or kellyjo.whipple@stryker.com |  | 20 | 20 | Distributed in the states of AR, CA, IN, LA, MA, MD, MI, MT, NV, and TX. | Terminated | September 15, 2014 | 125 | 20 | Not_Computer
Z-1728-2014 | 68431 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging |  |  | May 29, 2014 | 2014 | Mckesson Medical Immaging | Firm received a complaint that an unreported study performed that day could not be opened from the archive. | Software design | McKesson sent an Urgent Field Safety Notice dated May 20, 2014, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were informed that the potentially affected software versions will be corrected without removing the systems from their location. Customers were instructed to distribute the notice to all those within their organization who need to be aware and alert all those affilliated with their organization that are affected by this field safety notice. All customers receiving the Field Safety Notice should contact McKesson Support as soon as possible at 1-800-663-2533 to arrange installation of the software update. Customers with questions should contact the Customer Support Department and ask to speak to their Support Manager. |  | 659 | 659 | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OR, PA, RI, SC, TN, TX, VT, VA, WA, WV, WI, WY, and Puerto Rico. Internationally to Australia, Canada, France, Great Britain,and Guam. | Terminated | March 31, 2017 | 1037 | 659 | Computer
Z-1729-2014 | 68237 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | HLD Systems Model 610 Washer/High Level Disinfector. |  |  | April 24, 2014 | 2014 | Cenorin, LLC | Cenorin LLC is voluntarily recalling HLD Systems Model 610 Washer/High Level Disinfector due to a potential risk of the system not injecting the proper amount of cleaning fluid during the wash cycle and not achieving the desired level of debris reduction. | Equipment maintenance | Cenorin sent a letter on April 24, 2014 informing their customers of the potential failure of the pump. This letter also explains the customer's option of upgrading their entire HLD System to the new cleaning fluid delivery system or to continue purchasing a six month PM kit that has a replacement pump added. Customers are instructed to complete and return an enclosed response form within 15 days and to email Jamie Wright at jwright@cenorin.com or call 800-426-1042, extension 224 for questions. |  | 25 | 25 HLD systems | Distribution USA Nationwide | Terminated | June 30, 2015 | 432 | 25 | Not_Computer
Z-1730-2014 | 68142 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Maquet PowerLED 700. |  |  | April 16, 2014 | 2014 | Maquet Medical Systems USA | Maquet has identified an issue with the Maquet PowerLED 700 (certain serial numbers). One or several plastic fixation tabs can break, which can cause the ambient light module to partially detach. | Other | Maquet issued "Urgent-Medical Device Field Correction" notices, dated April 16, 2014 to affected customers via Federal Express. The notice identified the product, related issue and the action needed to be taken by the customer. The corrective action will include inspection of the fixation tabs of the Ambient Light Module adn replacement of the Ambient Light Module if tabs are found to be defective. The correction will be performed by an authorized MAQUET Service Technician. MAQUET Service Team representatives will schedule on-site service with customers. If customers have any questions or require additional information, the notice instructs them to contact their local MAQUET representative, or MAQUETCustomer Service at 1-888-627-8383 (option 3), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST. |  | 1169 | United States 1169 units | US Nationwide Distribution. | Terminated | April 06, 2016 | 721 | 1169 | Not_Computer
Z-1727-2014 | 68230 | No | GHH | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | Alere TriageÂż D-Dimer Test |  |  | May 02, 2014 | 2014 | Alere San Diego, Inc. | Alere initiated this recall because a limited number of Alere TriageÂż D-Dimer devices from PN 98100 Lot W53884B were incorrectly manufactured. As a result, when testing patient sample these affected devices will either generate an error code (e.g. E4 error), or they may generate an incorrect result. | Under Investigation by firm | Alere sent an Urgent Medical Device Recall letter dated May 2, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers are instructed to stop using all product from the affected lot, lot W53884B, and to discard any unused product. The customers were also instructed to complete and return the verification form indicatingany credit or replacement product requested. They were also instructed to contact their primary consignees and advise them of the notice.Should you have any questions about the information contained in this notification, please call (877) 308-8287. and /or email:Triage.Support@alere.com |  | 373 | 373 kits | Nationwide Distribution only. | Terminated | February 27, 2015 | 301 | 373 | Not_Computer
Z-1726-2014 | 66274 | No | OUE | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Immunology | 510(K) Exempt | Leica Ceiling Mount MS-1F &amp; MC-1F for use with the M500 N operating microscope. Article Number 10446657 utilizes a supply voltage of 110 V. Article Number 10446658 utilizes a supply voltage of 220 V. |  |  | September 11, 2013 | 2013 | Leica Microsystems, Inc. | The ceiling mount supplier encountered an event in which the drive screw in the motor unit, which lifts and lowers the lower lift arm, broke. As a result, the lift arm fell down approx. 60 cm (23.6 inches). | Nonconforming Material/Component | A Medical Device Correction letter, dated September 11, 2013, was sent to all direct accounts. The letters advised customers that a Leica Microsystems representative will contact them and make arrangements to conduct an inspection of the ceiling mount(s) and to provide additional preventative maintenance and inspection protocols for the devices. Customers are asked to complete and return the enclosed Acknowledgement Form and send it back to Leica Microsystems within 10 days of receiving the Medical Device Correction letter. Customers with questions should contact their Leica Microsystems representative. |  | 103 | 103 total | Worldwide Distribution -- USA, including the states of AZ, IL, MA, MO, NC, OR, TX, VA, and WI; and, the countries of Australia, Austria, Chile, China, Colombia, Czech Republic, Greece, Israel, Italy, Japan, Malaysia, Netherlands, Norway, Oman, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, and United Kingdom. | Terminated | April 24, 2017 | 1321 | 103 | Not_Computer
Z-1733-2014 | 68076 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare Diagnostics Vista Air Compressor. |  |  | April 14, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | A small number of Vista Air Compressors were miswired, which if installed and powered up, could cause a fire, or become an electrical shock hazard. | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Correction letter dated April 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.Actions to be Taken by the Customer1. If an Air Compressor has already been delivered to your laboratory, or isdelivered in the future for installation by your Siemens Healthcare CustomerService Engineer, please do not make any attempt to install the Air Compressorby yourself.2. Please notify your local Siemens Healthcare Diagnostics Customer Care Center when an Air Compressor is at your site so arrangements may be made for installation by a Siemens Healthcare Diagnostics Customer Service Engineer. Please discuss the content of this letter with your Medical Director.Please complete and return the Field Correction Effectiveness Check Form attached to this letter. Please retain this letter with your laboratory records and forward this letter to those who may have received this product.We apologize for the inconvenience this situation has caused. If you have anyquestions, please contact Siemens Customer Care Center - Technical Solutions at (800-441-9250) or your local Siemens technical support representative. |  | 4 | 4 | Worldwide Distribution: US Distribution to states of: IL, OH and TN; and country of: Italy. | Terminated | December 15, 2014 | 245 | 4 | Not_Computer
Z-1731-2014 | 68267 | No | ETD | Tympanostomy tube. | Ear Nose &amp; Throat | Ear Nose &amp; Throat | 510(k) | Armstrong Grommet with wire Otological Ventilation Tube |  |  | May 12, 2014 | 2014 | Summit Medical, Inc. | One lot of ventilation tubes have mixed tubes with wire attached to the slanted flange and ventilation tubes with wire attached to the straight flange. This poses no health hazard to the user or patient. | Employee error | Consignees was sent a Summit Medical "Medical Device Recall" letter, dated May 9, 2014. The letter described the problem and the product involved in the recall. Requested consignees to examine their inventory and to quarantine the product. Consignees need to immediately examine their inventory quarantine any affected product. Also, they are to complete and return the "Recall Response Form." For questions they can contact Nicole Dove at (651) 789-3921 or ndove@summitmedicalusa.com. |  | 9 | 9 boxes | Worldwide Distribution -- USA, including CA, and GERMANY. | Terminated | June 19, 2014 | 38 | 9 | Not_Computer
Z-1738-2014 | 68270 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Laboratory Automation Systems (enGen) |  |  | April 10, 2014 | 2014 | Ortho-Clinical Diagnostics | Software anomaly: potential for delay in the reporting of patient sample test results when using enGen with TCAutomation Software v2.6 , v3.2 and v3.5 Configured with Bypass Modules for ADVIA Centaur, Abbott Architect, VITROS and/or Tosoh AIA Systems. | Software design | Urgent Product Correction Letters (CL2014-114, dated 4/10/2014) were sent to all US customers and Foreign affiliates starting on 4/10/2014. The letter informed customers of the issue and provided required actions to be taken: Routinely monitor your Bypass module(s) -- If you detect a stoppage due to the software timing issue, follow the steps in the How to Detect the Bypass Module Time Issue and Release the Samples section on Page 2 of this notification. -- Complete the Confirmation of Receipt Form and return by 4/23/2014. -- Place this notification near each system or with your enGen (TM) Laboratory Automation System user documentation. |  | 108 | Domestic - 31 units; Foreign - 77 units | Worldwide Distribution: US Nationwide including Puerto Rico, and countries of: Australia, Bermuda, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela | Open, Classified |  |  | 108 | Computer
Z-1736-2014 | 68271 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System and VITROS 5,1 FS Refurbished, Product Codes/Catalog Numbers 6801375, 6801890 for in vitro diagnostic use. |  |  | April 10, 2014 | 2014 | Ortho-Clinical Diagnostics | Potential for biased results when a Multiple Window (MW) code is associated with results generated by the VITROS 4600 Chemistry System, the VITROS 5600 Integrated System and the VITROS 5,1 Chemistry System. MW codes may be associated with results from any of the following VITROS Chemistry Products using multiple point (multipoint) rate measurements. The MicroSlide assays that could be affected | Software design | Important Product Correction Letters (CL2014-116, dated 4/10/2014) were sent on 4/10/2014 via FedEx overnight courier or ORTHO PLUS e-Communications to all customers with VITROS 5,1 FS, 5,1 FS Refurbished, 4600,and/or 5600 Systems. They were informed of the issue and advised to: -- Review VITROS ALT, ALTJ, ALKP, ALKPJ, AST, ASTJ, CHE, CHEJ, CK, CKJ, GGT, GGTJ, LDH and LDHJ results generated on VITROS 4600/5600 and 5,1 FS Systems. Do not report any results associated with an MW code. -- Perform repeat testing in singlicate on all samples with results having an MW code. Only results without an MW code are acceptable. -- Review their current Laboratory Information System (LIS) configuration to determine if changes are required in order to restrict reporting of results with an MW code. Customers with questions should contact Customer Technical Services at 1-800-421-3311. --- Foreign affiliates were informed of the issue by e-mail on 4/10/2014 and instructed to notify their consignees of the issue and actions. |  | 1978 | Domestic - 597 units, Foreign - 1381 units | Worldwide distribution including US Nationwide, Puerto Rico, Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 4029 | Not_Computer
Z-1732-2014 | 68350 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RUSCH Easy Cath Coude Kit, 10 FR, Rx Only, Sterile, Distributed by: Teleflex Medical, Research Triangle Park, NC 27708. |  |  | May 21, 2014 | 2014 | Teleflex Medical | Product was packaged with a straight catheter rather than the correct coude (curved) catheter. | Mixed-up of materials/components | Teleflex Medical sent an Urgent Medical Device Recall Notification dated May 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately discontinue use and quarantine any of the affected product in stock. They were to complete and return the enclosed Recall Acknowledgement Form via fax to Teleflex Medical. Upon receipt, they will be contacted by a Customer Service Representative with instructions for the return of the product to Teleflex Medical.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 1000 | 1000 eaches | US Distribution including the states of CA, FL, OH and OK. | Terminated | November 23, 2015 | 551 | 1000 | Not_Computer
Z-1737-2014 | 68309 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Ultrasonic Transducer Kits |  |  | May 14, 2014 | 2014 | Beckman Coulter Inc. | Beckman Coulter is initiating a field action because some Ultrasonic Transducer Kits werepotentially assembled with incorrect mounting screws. | Mixed-up of materials/components | Beckman Coulter sent an "URGENT MEDICAL DEVICE RECALL" letter dated May 8, 2014 via first class mail. The letter identified the reason for recall, product information, issue, impact, action, resolution, contact information, and customer response form. Customers are advised to review historical quality control (QC) data and System Check results, and determine whether a review of patient test results is warranted.For questions regarding this letter, please contact our Customer Support Center:From our website: http://www.beckmancoulter.com/customersupport/support, By phone: call 1-800-854-3633 in the United States and Canada, and Outside the United States and Canada, contact your local Beckman Coulter representative |  | 82 | 82 units | Worldwide Distribution and USA Nationwide - and the countries of Australia, Canada, China, France, Hong Kong, Italy, Japan, Mexico, Portugal, Spain, Switzerland, Taiwan,and Turkey. | Terminated | December 25, 2014 | 225 | 82 | Not_Computer
Z-1740-2014 | 67998 | No | EBZ | Ultraviolet activator for polymerization. | Dental | Dental | 510(k) | bluephase style Light probe Pin-point 6&gt;2 mm black REF # 636 241. The accessory is a pin point light probe, an accessory to a dental curing light |  |  | April 14, 2014 | 2014 | Ivoclar Vivadent, Inc. | A stock check revealed broken glass fibers between the connector and the fiber bundle in a large number of these light probes. | Other | lvoclar Vivadent, Inc. sells the affected medical device through Dental Dealer partners (initial consignees) who in turn sell directly to the dentist's office. In this case, lvoclar Vivadent, Inc. handled direct drop shipment of the device to end user dental offices in 80% of the cases. Where lvoclar Vivadent hasdirect shipment information available, dentist offices will be notified directly.The Dental Dealer partners will be notified of this recall and asked to cooperate and provide name and address listings for all dental office customers where they have shipped the Pin Point probe. The dealer will also be given a customer communication to send directly to the dentists purchasing the device. |  | 241 | 241 units | Worldwide Distribution: US Nationwide and the country of Canada. | Terminated | November 05, 2015 | 570 | 241 | Not_Computer
Z-1739-2014 | 68135 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NaturaLyteÂż Liquid Bicarbonate Concentrate, 6.4 liter bottle to be used with a hemodialysis unit; Part Number: 08-4000-LB. |  |  | May 01, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Container cap may contain foreign material- 0.60% - 0.80% Manganese | Equipment maintenance | Fresenius Medical notified customer by letter dated 5/1/14 to examine stock immediately and instructed to discontinue use and place all Naturalyte Liquid Bicarbonate Concentrate in a secure area for return to FMC-RTG. The customer instructed to contact the Fresenius Medical Care Customer Service Team for instructions on how to return the recalled product. |  | 12 | 12 bottles | Distributed in the state of MO. | Open, Classified |  |  | 12 | Not_Computer
Z-1741-2014 | 68269 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen (TM) Laboratory Automation System Configured with TCAutomation" Software v3.5 -- For in vitro quantitative measurement of a variety of analytes of clinical interest --- |  |  | April 17, 2014 | 2014 | Ortho-Clinical Diagnostics | Software anomaly: potential buffer timeout issues when using enGen (TM) Laboratory Automation Systems Configured with TCAutomation (TM) Software v3.5. | Software Design Change | Ortho Clinical Diagnostics sent an Important Product Correction Notification dated April 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.On April 17, 2014, foreign affiliates were informed by e-mail of the issue and instructed to notify their customers of the issue and actions.The letters instruct customers: to ensure that samples prematurely routed to Check are detected and processed appropriately, please monitor completion of all pending tests following standard laboratory practice. Customers were informed of the issue and advised to: -- Routinely monitor pending result reports from your Laboratory Information System (LIS) to determine if results are pending for samples. -- Investigate to determine the next action identified for a sample if the sample is routed to the Check location prematurely. If you have any additional questions, please call Customer Technical Services at 1-800-421-3311. |  | 9 | Domestic: 6 units, Foreign: 3 units | Worldwide Distribution - US including the states of Arizona, Delaware, Ohio, Tennessee, and North Carolina; and the countries of Mexico, England and Italy. | Terminated | April 11, 2017 | 1090 | 9 | Computer
Z-1742-2014 | 68313 | No | NEC | Liquid bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | BioDerm Liberty 3.0 External Male Catheter Specialty Type with integral collection chamber (Component: SKIN-PREP Protective wipes). |  |  | May 31, 2011 | 2011 | BioDerm, Inc. | On 06/08/2011, BioDerm, Inc. issued a voluntary recall of a component skin wipes for the BioDerm Liberty 3.0 External male Catheter, BioDerm Liberty Pouch Clean &amp; Dry External Male Catheter and BioDerm Safe n' Dry. This component (wipes) were initially manufacturer by Smith &amp; Nephew and recalled due to a potential bacterial contamination. | Material/Component Contamination | Notice was sent to customers by letter dated June 8, 2011 and by phone. |  | 26040 | 26,040 units. | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, FL, NH, NJ, NY, PA, TX, and WI, and the countries of Canada and Netherlands. | Terminated | June 20, 2014 | 1116 | 26040 | Not_Computer
Z-1679-2014 | 68356 | No | DWJ | Thermal regulating system. | Cardiovascular | Cardiovascular | 510(k) | 115 V Blanketrol II, Model 222S Hyper-Hypothermia System . Used to lower or to raise a patients temperature and /or maintain a desired patient temperature through conductive heat transfer. Part number 86165. |  |  | January 08, 2014 | 2014 | Cincinnati Sub-Zero Products Inc | The incorrect voltage component was placed into the device during manufacturing. This may make the device unable to maintain patient temperature during cooling mode. The device will continue to indicate that the unit is cooling because the compressor will continue to run, however water in the reservoir will begin to return to ambient temperature. The actual water temperature will be displayed but | Mixed-up of materials/components | The firm initiated a telephone notification on 01/08/2014, followed by a written notification which was sent to the customers via letter or email on 01/31/2014.URGENT Medical Device Field Action (January 31, 2014)REF: 115V Blanketrol II Model 22S Hyper-Hypothermia System(1) Discontinue Use and contact CSZ(2)Upon receipt of new 115V (p/n 91113) solenoid coil(s), immediately remove and destroy/discard all the affected 220V (p/n 3600) solenoid coil(s)(3)After the affected 220V (p/n 36007) solenoid coil(s) have been removed, discard and replaced, please complete and return the enclosed response form as soon as possible to acknowledge receipt of this notification and to imform CSZ that you have performed and completed the requested actions. Return the form by fax |  | 7 | 7 units | The product was distributed to one customer in MI, and internationally (Taiwan) | Terminated | June 06, 2014 | 149 | 7 | Not_Computer
Z-1748-2014 | 68081 | formulations, paraffin, all | KEO | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Tissue-Tek PARAFORM PROC/EMB/MEDIUM, FORMULA 3 - 8x1Kg. |  |  | April 16, 2014 | 2014 | Sakura Finetek USA Inc | This Lot has variations in melting temperature. | Packaging process control | Customer notifications sent via e-mail letter.The letter titled "PRODUCT RECALL" dated 04/17/2014 informed customers (of Lot 263922), informed customers of the recall by providing the following information: product name, code, lot number, reason for recall, requested actions, contact information and an "FIELD CORRECTION EFFECTIVENESS CHECK" form. |  | 115 | 115 | Nationwide Distribution. | Terminated | September 02, 2014 | 139 | 115 | Not_Computer
Z-1746-2014 | 68349 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | IntelliVue MX40 Wearable Patient Monitor in use with the Philips IntelliVue Information Center iX (PIIC iX); 865350, 865351, 865352 |  |  | May 22, 2014 | 2014 | Philips Medical Systems, Inc. | IntelliVue MX40 (part numbers 865350, 865352) with software revisions B.00 or B.01 may not resume alarming when the alarms are paused and the pause timer ends. | Software design | On May 22, 2014, each affected customer was sent an Urgent Medical Device Correction notification/Field Safety Notice. The letter explains the problem, explains the action to be taken by the consumer, and the action planned by Philips. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of a free of charge software upgrade. A Philips Healthcare representative will contact customers with affected devices to arrange for service. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:The following two items must be configured when using MX40 with the IntelliVue Information Center iX: 1.Utilize the default state of Pause Alarms at Yellow only, which disables the ability to Pause/Suspend all Red level alarmsUtilize the default state of Remote Suspend Alarms as Off, which disables the use of the Pause Alarms feature at the MX40 device |  | 42600 | 42,600 | United States: NationwideForeign Countries:Germany, Austria, Finland, Sweden, Australia, Qatar, Canada, Netherlands, New Zealand, Portugal, United Kingdom, Spain, Belgium, France, Switzerland, Norway, Czech Republic, Poland, Italy, Singapore, Denmark, Iceland, India, Saudi Arabia, Ireland, Oman, Bahrain, Japan, Israel, Malaysia, South Africa, Turkey, Thailand, Latvia, Bulgaria, Indonesia, Slovenia, United Arab Emirates, Hong Kong, Aruba, Chile, Lebanon, Korea, Kuwait, Taiwan, Argentina, Hungary, Luxembourg, Mexico, Gabon | Open, Classified |  |  | 42600 | Computer
Z-1747-2014 | 68279 | No | OCY | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Boston Scientific, NovaGold High Performance Guidewire, Rx only, Sterilized using ethylene oxide, UPN M00552010 and M00552000, Manufactured by NeoMetrics Inc, 2605 Fernbrook Lane Suite J, Plymouth, MN 55447. |  |  | March 31, 2014 | 2014 | NEO METRICS, INC. | Neo Metrics, Inc. is recalling NovaGold Guidewire because the tip may uncoil and / detach more frequently than anticipated. There have been no reports of illness or injury. | Nonconforming Material/Component | NeoMetrics sent a letter dated March 31, 2014, to Boston Scientific. On March 24, 2014, Boston Scientific made NeoMetrics aware that Boston Scientific received 3 complaints from customers that the tip of the NovaGold guidewire detached and/or became uncoiled.Because the tips are breaking and/or uncoiling at a rate that was not anticipated, it has been decided that these units should be removed from the field.Therefore, NeoMetrics does not object to Boston Scientific executing this product removal in accordance with Section 4.4 of the distribution agreement.For further questions please call (763) 559-4440. |  | 2512 | 1256 boxes ( 2 wires/box) | Worldwide Distribution - US and the countries of Austria, Belgium, Germany, Spain, France, Great Britain, Italy, Netherlands and Sweden. | Terminated | January 16, 2015 | 291 | 2512 | Not_Computer
Z-1762-2014 | 68331 | No | HBE | Powered simple cranial drills, burrs, trephines, and their accessories. | Neurology | Neurology | 510(k) | Stryker 3.5mm Zyphr Elite Round Fluted Bur |  |  | May 22, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | There may be missing notches, partial notches or no notches at all in some Cutting Accessories that are used with the Elite" Attachments. | Nonconforming Material/Component | Stryker Instruments sent an URGENT MEDICAL DEVICE RECALL notification letter dated May 22, 2014, to the consignees. Recall notification included affected product description and identification, reason for recall, risk to health, and actions to be taken by the customers. Customers will be instructed to review the notification, quarantine unused product and to return a signed Business Reply Form. They will be advised that a Stryker Representative will be sent to their facility to inspect any affected product on hand. Once product is inspected by the Stryker Representative, they can return conforming product back to stock, for normal use. Non-conforming product will be removed from the facility by a Stryker Representative following the inspection process. The Stryker Representative will send the product to the manufacturer for destruction. For questions regarding this recall please contact Stryker Instruments:Kara Spath269-389-4518 kara.spath@stryker.com |  | 39 | 39 | Worldwide Distribution - US (nationwide) | Terminated | May 06, 2015 | 349 | 13739 | Not_Computer
Z-1782-2014 | 68317 | No | LKF | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical TAMPA CATHETER 5 French 33 cm. |  |  | May 13, 2014 | 2014 | CooperSurgical, Inc. | Sterility of the device may be compromised due to unsealed pouch | Process control | Cooper Surgical Inc notified consignees by letter dated 5/13/14 sent via Federal Express with confirmed delivery receipt. Consignees are requested to return for refund or exchange. If you have any further questions contact the firm at 203.601.5200. |  | 1180 | 1180 units | Distributed USA (nationwide) including the states of CO, CT, MO, VT, NY, OH, NJ, FL, NC, VA, AL, MA, PA, and CA, and the country of Canada. | Terminated | May 28, 2015 | 380 | 1180 | Not_Computer
Z-1784-2014 | 64273 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis EMPIRA RX PTCA Dilatation Catheter Catalog # 85R30300S For cardiovascular use. |  |  | February 05, 2013 | 2013 | Cordis Corporation | Dilatation catheters could exhibit radial versus axial tears should they burst during inflation. | Process design | An Urgent Medical Device Recall letter and Acknowledgement form was sent overnight to mulitple contacts in each account February 7, 2013. A representative will follow-up as necassary to facility obtaining signature, faxing the acknowledgment form to Cordis, collecting and returning units. |  | 80 | 80 units | Worldwide distribution: US states: AZ, FL, IL, LA, MA, NJ, OR, SC, and TX. Armenia, Austria, Belgium, Colombia, Czech Republic, France, Hungary, Iran, Israel, India, Italy, Kuwait, Latvia, Lebanon, Luxembourg, Mexico, Morocco, Poland, Portugal, Russia, Saudi Arabia, South Africa, Spain, Switzerland, and United Arab Emirates. | Terminated | June 23, 2014 | 503 | 80 | Not_Computer
Z-1783-2014 | 64507 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IntellVue Info Center iX 866023; PIIC iX Upgrade 866024; IntellVue Server iX Upgrade 866026; PIIC Classic Upgrade 866117 Patient Physiological Monitor. |  |  | March 05, 2013 | 2013 | Philips Healthcare Inc. | If a customer creates customized trend scales in the trend review tile and the iX or primary server reboots for any reason, the iX(s) will enter a reboot loop. If one iX reboots, the issue will impact only those patients monitored on that device. If the primary server reboots, all patients across all iX devices will be impacted. | Software design | On 3/5/13, Philips sent each known affected customer an Urgent Medical Device Correction notification/Field Safety Notice to inform them that Philips has received a customer report that the PIC iX (release A.01.XX) central stations that are connected to a database server (DBS) may not be able to complete a restart/reboot resulting in the loss of central monitoring. In order to prevent this from occurring, customers were advised to not modify/customize the default trend scales for reviewing stored patient data in the retrospective review applications. Recall letter also states that a Philips Healthcare representative will contact customers with affected devices to arrange for the installation of updated PIIX iX software resolving this issue on affected units. Philips will conduct these updates for all affected devices at no charge. |  | 899 | 899 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | May 17, 2017 | 1534 | 899 | Computer
Z-1781-2014 | 68111 | No | HTZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Matrix Mandible Short Cut Plate Cutter |  |  | April 22, 2014 | 2014 | Synthes, Inc. | It was discovered internally that the face of the Synthes Matrix Mandible Short Cut Plate Cutter has the potential for discoloration/corroded material in the affected lot. | Device Design | Synthes Inc. sent an Urgent Medical Device Recall Notice dated April 22, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If you DO have any of the identified devices, please take the following steps:Contact DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number . Complete the Verification Section at the end of this letter by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found. Please include your name, title, telephone number and signature in the spaces provided.Return the Verification Section (page 3 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Return a copy of the Verification Section (page 3 of this letter) to DePuy Synthes by: Fax: 866-324-3731 or Scan/email: Synthes7726@stericycle.comIf you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section at the end of this letter by checking the appropriate box indication that no affecteded product has been located. Please include your name, title, telephone number and signature in the spaces provided this return documentation acknowledges your receipt of medical device removal information.Return the Verification Section (page 3 of this letter) to DePuy Synthes by:Fax: 866-324-3731 or Scan/email: Synthes7726@stericycle.comNote: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes sales consultant. |  | 14 | 14 | US Distribution including the states of CA, CT, FL, GA, NY, SD and TX. | Terminated | September 10, 2015 | 506 | 14 | Not_Computer
Z-1641-2014 | 67991 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Portex Low Dead Space Connector with Sideport, 3.5mm |  |  | April 10, 2014 | 2014 | Smiths Medical ASD, Inc. | One lot of 3.0mm sized Connectors, Product Reorder No. 10035-05PS, Lot No. 2553426 were placed in packages labeled as size 3.5mm. | Packaging process control | Smiths Medical sent an Urgent Medical Device Recall Notice dated April 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and quarantine all affected products and to complete the attached Response Form and return it to Smith Medical within 7 days of receipt of the notice. Customers with questions were instructed to contact Smiths Medical's Customer Service Department at +1-800-258-5361. |  | 100 | 100 devices | US distribution in the state of North Carolina. | Terminated | November 29, 2016 | 964 | 100 | Not_Computer
Z-1786-2014 | 68332 | No | GBZ | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Taut Operative Cholangiogram Catheter, 4.5FR (1.5mm) x 76cm |  |  | May 16, 2014 | 2014 | Teleflex Medical | Several complaints were received for the metal support tube missing from the cholangiogram catheter. | Employee error | Teleflex sent an Urgent Medical Device Recall Notification dated May 16, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were Instructed:1. Immediately discontinue distribution and quarantine any products with the catalog numbers and lot numbers listed above.2. Using the provided customer letter and Recall Acknowledgement Form templates, communicate this recall to any of your customers who have received product included within the scope of the recall.3. Have the customers return any affected product to you, together with a completed Recall Acknowledgement Form, for consolidation and return to Teleflex Medical. In the event that an alternative approach is needed, contact Teleflex Medical Customer Service for more information at 1-866-246-6990.4. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1- 866-804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.5. If you and your customers have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letterTeleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 450 | 450 | Worldwide Distribution - USA including the states of AL, AZ, AR, CA, CO, FL, IL, LA, ME, MI, MO, NY, OH, OR, TN, TX, VT and WI., and the country of Belgium. | Terminated | November 04, 2015 | 537 | 450 | Not_Computer
Z-1789-2014 | 68226 | Yes | NPF | Apnea monitor. | Anesthesiology | Anesthesiology | 510(k) | Philips/Children's Medical Ventures Smart Monitor 2 Professional Series infant apnea monitor. |  |  | April 29, 2014 | 2014 | Childrens Medical Ventures | The battery pack wire harness is improperly assembled rendering the device inoperable. The wiring issue causes an error code to be displayed and results in the device sounding an alarm repeatedly at power up. The presence of the error code and the devices alarm prevents the device from functioning until the issue is addressed. | Component design/selection | Philips sent an Urgent Medical Device Correction letter dated April 29, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This Customer Information is intended to inform you about:1.what the problem is and under what circumstances it can occur2.the actions that you as a customer can take to minimize the effect of the problem3.the actions planned by Philips to correct the problem.If you need any further information or support concerning this issue, please contact your local Philips representative: For further questions please call (412) 380-8881.Philips apologizes for any inconveniences caused by this problem. |  | 112 | 112 | Worldwide Distribution - USA including the states of MI, WI, SC, FL, TX, KS, MA, AZ, UT, WV, IN, MN, NV and VA., and the countries of Australia, Belgium and Canada. | Terminated | April 06, 2015 | 342 | 112 | Not_Computer
Z-1785-2014 | 68292 | No | KNT | Gastrointestinal tube and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Convatec Flexi Seal Control Fecal Management System Rx Only |  |  | April 29, 2014 | 2014 | Convatec Inc. | Convatec Inc. is recalling Flexi-seal FMS Control Kit since it does not have a 510K. In addition, an internal assessment of product performance, including a review of customerfeedback, has confirmed that Flexi-Seal Control is not meeting expectations or those of customers. Specifically, the Auto-Valve feature that is unique to the Flexi-Seal Controlhas not been consistently performing relative | Device Design | Convatec Inc. sent an Urgent Voluntary Medical Device Recall letter dated April 29, 2014 via UPS. The letter identified the affected product, problem and actions to be taken. Customers are asked to immediately stop distributing and quarantine all recalled lots. Affected products must be returned to Distributor. Contact your Distributor for shipping instruction. Customers should completed the attached return response form and return it to the address on the response form. Questions can be directed to Gregory Kowzun at (336) 547-3710. Convatec issued a press release on 7/3/2014. |  | 97887 | 73,021 units US; 24,866 units OUS | Worldwide Distribution - USA Nationwide | Terminated | August 28, 2015 | 486 | 97887 | Not_Computer
Z-1787-2014 | 65483 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Beekley Spots Light Image O-Spots for Mammography |  |  | June 19, 2013 | 2013 | Beekley Corporation | Light Image O-Spots Dispenser mislabeled as Soft n' Stretchy Light Image S-Spots | Error in labeling | Beekley sent an Urgent Medical Device Recall letters dated June 19, 2013 via FedEx with signature and tracking confirmation. The letter identified the affected product, problem and actions to be taken. Firm requests users to examine inventory, remove and return recalled product. For questions contact Chelsea Fuller, Regulatory Manager at 1-800-233-5539 x 491 |  | 81 | 81 boxes | Worldwide Distribution - USA in the states of CA, FL, MA, NC, NY, and PA. | Terminated | June 16, 2014 | 362 | 81 | Not_Computer
Z-1788-2014 | 68426 | No | EJJ | Dental amalgam, mercury, and amalgam alloy. | Dental | Dental | 510(k) | Tytin Regular Set, Double Spill, 50 Capsules, Part Number 29948, Lot Number 3-1294. |  |  | May 29, 2014 | 2014 | Kerr Corporation | Kerr Corporation is recalling one (1) lot of Tytin Regular Set ( Part Number 29948, Lot Number 3-1294), because it may set faster than specified in the Directions for Use. | Unknown/Undetermined by firm | Kerr Corporation initiated this recall by sending the recall notification letter accompanied by a return form, to customers via USPS 1st class mail. The consignees are instructed to complete the return form and return any affected product to Kerr Corporation.The recall notification letter, dated May 29, 2014, titled "URGENT: MEDICAL DEVICE RECALL", informed customers of the recall by providing the following information: product description with part number and lot number, distribution date range, reason for recall, risk statement, instructions to customers and/or Kerr distributor on what to do with the recalled product, FDA MEDWATCH reporting system info, and company contact information, and instructed customers to complete the return form titled "Acknowledgement and Recall Return Form". |  | 2030 | 2,030 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, Mexico, Australia, Belgium, Czech Republic, Germany, Spain, France, Great Britain, French Guiana, Italy, Portugal, and Saudi Arabia. | Terminated | October 27, 2014 | 151 | 2030 | Not_Computer
Z-1793-2014 | 68229 | No | OIE | Specimen transport and storage container. | Pathology | Clinical Chemistry | Enforcement Discretion | APTIMA Urine Collection Kit; Catalog Number |  |  | April 17, 2014 | 2014 | Gen-Probe Inc | Incorrect labeling was used on each pouch of the affected lots of APTIMA Urine Specimen Collection Kits. The outer box of the kits are identified correctly as "APTIMA Urine Collection Kits"; however, the inner pouches are incorrectly labeled as "APTIMA Unisex Swab Collection Kit for Endocervical and Male Urethral Swab Specimens." | Labeling mix-ups | Hologic sent a Recall Notice letter dated April 14, 2014 via overnight mail (FedEx) to affected customers. The letter identified the affected product, problem and actions to be taken. The notification letter requests customers to: 1. segregate the entire contents of the affected kit lots2. reconcile and destroy all partial and/or full kits remaining in inventory and to document; the amount destroyed on the Customer Response Form provided with the Recall Notice; 3. if the urine collection kits were further distributed to other collection sites, then to notify them of this issue; and 4. complete and return the Customer Response Form to Technical Support (contact information is provided in the notification letter).For additional information, questions, or concerns, contact Hologic Molecular (Technical) Support.Toll Free: 888-484-4747Direct: 858-410-8511Fax: 858-410-8250Email: molecularsupport@hologic.com |  | 4740 | 4,740 units | Worldwide Distribution - US Nationwide and the countries of Canada (CA), Denmark (DK), France (FR), Germany (DE), and the United States (US), Belgium (BE), United Kingdom (GB). | Terminated | August 26, 2014 | 131 | 4740 | Not_Computer
Z-1792-2014 | 68294 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | cobas 8000 analyzer series Operators Manual Roche cobas |  |  | May 21, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Under certain conditions loose ProCell/CleanCell aspiration tube filters may cause: consecutive discrepant results; sipper alarm; foam on reservoirs and system reagent alarms. | Labeling design | Roche sent an Urgent Medical Device Correction (UMDC) letter dated May 20, 2014 via UPS Grount (receipt required) to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to: 1. Follow the instructions in the Precautions Regarding Handling the Aspiration Tube Filter section in this UMDC; 2. Discontinue performing ProCell/Clean Cell aspiration tube filter check as this maintenance item willbe performed by your Roche FSR as a part of service maintenance; 3. If you have distributed the products listed within this UMDC to other sites or facilities, please provide them with a copy of the UMDC; and 4. Complete the faxback form enclosed with this UMDC and fax it to 1-888-410-6620. For questions contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 if you have questions about the information contained in this UMDC. |  | 149 | 149 | US Nationwide Distribution | Terminated | June 04, 2015 | 379 | 1758 | Not_Computer
Z-1801-2014 | 68216 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | EZPass Suture Retriever, Nitinol, Single Pack, Part number 10007380. Orthopedic Manual Surgical Instrument, Sterile, single use arthroscopic instrument utilized in conjunction with a nitinol wire to aid in passing suture through soft tissue. |  |  | April 15, 2014 | 2014 | Biomet, Inc. | The nitinol wire or monofilament may buckle within the EZPass Suture Passer and become kinked making the wire/monofilament difficult to use. | Device Design | On 2/28/14, customers received a letter stating these devices have been recalled following an investigation which identified that the nitinol wire or monofilament may buckle within the EzPass Suture Passer and become kinked making the wire/monofilament difficult to use. The letter also stated, there is no adverse health outcome expected for the patient and this action requires the immediate location and discontinued use of the product and its return to Biomet. |  | 100 | 100 units | Nationwide Distribution. | Terminated | July 07, 2015 | 448 | 202 | Not_Computer
Z-1724-2014 | 68255 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Synergy XVI |  |  | April 23, 2014 | 2014 | Elekta, Inc. | XVI can incorrectly calculate the target position of the treatment table. | Software design | Elekta is responsible for the correction of all products with this defect.1. Notifications regarding the possibility of this software error and subsequent table location error and potential hazards have been sent out. The notifications include steps to take to reduce the possibility of the occurrence of an injury.2. A software correction has been released for some versions of the XVI software. All versions of software with this error will be corrected by the end of July 2014.3. The software correction will enable verification that values have been assigned to table position parameters; failure to pass verification will inhibit the system. The software correction will also detect internal communication errors (incomplete registrations) and inhibit the system when those errors are detected.4. The update will be mandatory and will be made at no cost to the users.CDRH approves this CAP. You may proceed with implementation of the CAP.For further questions please call (770) 670-2422. |  | 380 | 380 US | Worldwide Distribution - USA including the states of AZ, CT, IA, OH, MD, MI, MN, NV, TX and VA., and the countries of India, Ireland, Malta, Netherlands, South Africa and UK. | Terminated | July 17, 2017 | 1181 | 380 | Computer
Z-1800-2014 | 68355 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ, an oncology information system used to manage workflows for treatment planning and delivery. It supports information flow among healthcare facility personnel and can be used wherever radiotherapy and/or chemotherapy are prescribed. |  |  | May 23, 2014 | 2014 | Elekta, Inc. | When using CMA, if the user hovers the mouse pointer exactly above the direction dropdowns and then uses the mouse thumbwheel or touchpad scroll bar, the direction selection might change when the intention was only to move the mouse to a different area of the screen. | Software design | Elekta sent an Important Safety Notice on May 23, 2014, to all affected customers. The notice identified the product the problem and the action needed to be taken by the customer."Notify the affected customers of the issue - To be completed 06/22/2014"Provide a fix to the defect in an expeditious timeframe. It is mandatory for all customers to apply the fix - Version 2.50.05 SP8 &amp; Version 2.60 - To be released 05/16/2014"Follow up with the affected customers for: -Confirmation receipt of the LIN, AND -Confirmation of installation of patch for the fix- To be completed 11/30/2014. For further questions please call (404) 993-5886. |  | 723 | 723 | Worldwide Distribution - USA (nationwide) including Puerto Rico, and the countries of Albania, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Gabon, Germany, Ghana, Greece, Honduras, Hungary, India, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Norway, Pakistan, Panama, Peru, Poland, Portugal, Romania, Russia, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sudan, Suriname, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, Ukraine and Venezuela. | Open, Classified |  |  | 723 | Not_Computer
Z-1798-2014 | 68373 | Yes | OAD |  |  | Cardiovascular | PMA | IntellaTip MiFi XP Temperature Ablation Catheter; 10m x 8F (2.67mm) Large Curve. |  |  | May 15, 2014 | 2014 | Boston Scientific Corporation | Some units of Intella Tip MiFi XP Temperature Ablation Catheters were not manufactured according to specification. | Use error | Customers were informed of the recall via overnight letter sent on May 15, 2014. |  | 968 | 968 units total all models | Nationwide Distribution. | Terminated | June 25, 2015 | 406 | 968 | Not_Computer
Z-1805-2014 | 68286 | No | DKB | Clinical toxicology calibrator. | Toxicology | Toxicology | 510(K) Exempt | VITROS (R) Chemistry Products FS Calibrator 1, Product Code/ Catalog Number 6801873 -- For in vitro diagnostic use only --- VITROS Chemistry Products FS Calibrator 1 is used in conjunction with VITROS Chemistry Products Calibrator Kits to calibrate the VITROS 5,1 FS and 4600 Chemistry Systems and the VITROS 5600 Integrated System. |  |  | May 12, 2014 | 2014 | Ortho-Clinical Diagnostics | Calibrator values assigned for the calibrator kit lots were released using a reference method that was undergoing validation for a new lot of European Reference Material (ERM). The validation process was not complete prior to calibrator value assignment for VITROS (R) Chemistry Products hsCRP Reagent (Product Code 6801739), Generation (GEN) 28. | Process change control | Customer Product Correction Letters (CL2014-143, dated 5/12/14) were sent on 5/12/2014 via FedEx overnight courier and/or US Postal Service Priority Mail (for PO Boxes only) and /or ORTHO PLUS e-Communications to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them to discontinue use of the affected calibrator lot(s), and inform their Laboratory Medical Director if they suspect that a previously reported result may have been affected. Distributor letter (DL2014-143) was also sent on 5/12/2014 via US Postal Service Priority Mail and/or FedEx overnight courier to US Distributor consignees and instructed to notify their consignees of the issue and the required actions. Foreign affiliates were pre-notified via e-mail on 5/08/2014 that the communication of this issue and required actions is imminent and were sent the finalized recall communication on 5/12/2014. |  | 825 | Domestic - 464 units; Foreign - 361 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Terminated | March 23, 2017 | 1046 | 1898 | Not_Computer
Z-1810-2014 | 67740 | No | HQG | Prescription spectacle lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | Safety glasses made with "Rage" frames manufactured between February 20, 2014 and February 25, 2014. The color of the frames of the recalled glasses were either white or black. |  |  | February 26, 2014 | 2014 | Wal Mart Stores, Inc | Fifteen pairs of prescription, safety glasses did not meet required, industrial lens thickness; nor did these glasses have the OSHA-required manufacturer markings. | Employee error | Wal-Mart sent an Urgent Medical Device Recall letter dated March 12, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken by the Customer/User:Do not use the eyewear in any situation where safety-impact lenses are required. Replacement eyewear has been manufactured and available for the patient to pick up at the vision center where the original eyewear was purchased.Type of Action by the Company:Walmart contacted all patients on February 27 and 28. Patients were notified that their eyewear was not made at safety thickness, and that the patient should not wear the eyewear in an environment where safety thickness lenses are needed. Walmart manufactured replacement orders for all affected patients.Should you have any questions regarding this notice please call (479)277-0720 |  | 15 | 15 safety glasses | US Distribution including the states of AR, AL, CA, FL, KS, LA, MN, NJ, NV,OH, TX, and SC. | Terminated | January 08, 2015 | 316 | 15 | Not_Computer
Z-1806-2014 | 68162 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation software version 4.0.3.4 |  |  | April 25, 2014 | 2014 | RAYSEARCH LABORATORIES AB | This notice concerns a problem with the computation of ROI voxel volumes for cases where the CT has variable slice spacing. The dose grid volumetric representation of an ROI can be wrong. This affects all dose-volume properties for such ROIs, including DVH, dose statistics, clinical goals and constraints or objective functions. Furthermore, if material override ROIs are used, or if the external is not represented by contours, the bug can trigger an error in the 3D dose distribution. | Software design | On April 25, 2014, RaySearch emailed all consignees of the affected product a copy of the recall notice as well as the response forms. Back up notification methods include phone calls and site visits. |  | 6125 | 25 sites/245 units (US), 25 sites/124 units (foreign) | Worldwide Distribution-USA (nationwide) including the states of AZ, CA, FL, IL, MA, MD, MO, MS, NJ, NV, NY, OH, PA, TN, VA, and WA, and the countries of Australia, Belgium, Germany, France, Italy, New Zealand, Republic of Korea, Switzerland, Spain, and United Kingdom. | Terminated | June 26, 2017 | 1158 | 6125 | Not_Computer
Z-1811-2014 | 68526 | No | HXX | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Biomet Spine, Cypher MIS Screw Inserter, Catalog number 14-501669, instrument for Spinal Screw Fixation System. |  |  | May 15, 2014 | 2014 | Ebi, Llc | The Cypher MIS Screw Inserter may exhibit an increased rate of instrument tip failure. | Nonconforming Material/Component | Biomet sent an "Urgent Medical Device Recall Notice"/Fax Back Response Form dated May 15, 2014, via Fed Ex to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory for the affected product and if located to remove the product from circulation; follow instructions on the attached Fax Back Response Form prior to return of the product; and return the product using the Fed Ex label included with the notification. Customers were instructed, if if they further distributed the product, they must notify their customers of the recall via the enclosed "Dear O.R. Manger" notice. If the customers had any questions they were asked to please contact Biomet, Monday through Friday (8 am to 5 pm) at 973.299.9300 xt. 2322. |  | 32 | 32 units (30 domestic and 2 international) | Worldwide Distribution - USA (nationwide) and the country The Netherlands. | Open, Classified |  |  | 32 | Not_Computer
Z-1809-2014 | 68454 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Model Number(s): PB840 Ventilator (4-840120DIUU-XX). The product is a component of the Puritan Bennett 840 Ventilator: The Backlight Inverter Printed Circuit Board (BLI PCBA) for the 9.4 Display part number 4-079056-00. |  |  | June 04, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | In the case of a loss of GUI display due to a Backlight Inverter PCBA failure, the ventilator continues to provide uninterrupted ventilatory support at the programmed settings for the patient. However, there is a loss of display and thus there is a necessity to move the patient to another ventilator. | Component design/selection | Covidien will be issuing the Product Notification letter in 2014. Customers are requested to forward the letter to anyone they may have transferred or assigned for use one of the 9.4 display configuration ventilators. Service engineers will be replacing the BLI PCBA with a new design containing the 1.0 ampere fuse and then destroying the old PCBA. Ventilators replaced with newer models under the commercial program will be returned to Covidien for disposal.Should you have any questions regarding this letter or to report any issues with the Puritan Bennett" 840 ventilator contact our Technical Support Department at 1-800-255-6774, option 4, then option 1, to ensure proper device reporting procedures are followed. Thank you for your prompt attention to this matter. |  | 7371 | 7371 | Worldwide Distribution - USA (nationwide) including the states: AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY and District of Columbia., and the countries of Canada and Japan. | Open, Classified |  |  | 7371 | Not_Computer
Z-1807-2014 | 68430 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iGUIDE System |  |  | May 29, 2014 | 2014 | Elekta, Inc. | A bug in the software prevents the iGUIDE software from logging off. | Software design | Elekta sent an Important Field Safety Notice 618 01 303 024 dated May 29, 2014 to affected customers. The letter identified the affected product, problem, potential clinical impact and actions to be taken. It also gives a recommendation how the risk can be mitigated - Users should only use the iGUIDE software when the key is in the on position and the Position Guard is active. Users are advised to restart the workstation if any error occurs to ensure the software is running properly. The notice also advises the customer that a patch will be released that corrects the behavior of the iGUIDE software. This solution will be implemented on all affected products when available. For questions contact your local Elekta office. |  | 224 | 224 | Worldwide Distribution - USA (nationwide) in the states of AZ, CA, CT, HI, IL, IA, KY, MD, MA, MI, MO, NH, NJ, NC, ND, OH, OK, OR, PA, TX, VA, WA, WI, and the countries of Algeria, Australia, Austria, Brazil, Canada, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Italy, Japan, Kazakhstan, Mexico, Netherlands, New Zealand, Peru, Poland, Russia, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, United Kingdom and Vietnam | Open, Classified |  |  | 224 | Computer
Z-1808-2014 | 68100 | No | CCL | Oxygen gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Compact Airway Modules E-CO, E-COV, E-COVX, E-CAiO, E-CAiOV and E-CAiOVX. The Datex-Ohmeda S/5MT Compact Airway Module, E-CAiOVX family is used for monitoring hospital patients respiration, ventilation, and gas exchange status. |  |  | May 23, 2014 | 2014 | GE Healthcare, LLC | Potential loss of airway gas measurement in the Compact Airway Gas Modules. Loosening of the gas sampling pump hex screw in the module pump unit may lead to pump failure that may cause loss of respiratory airway gas parameter monitoring. If this issue occurs, the monitor will show an alarm note on monitor screen of either "Low gas sample flow" or "Check sample gas out". | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter dated May 23, 1014. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions they should contact Technical support at 1-800-558-7044 or your local Service Representative. |  | 4961 | 4,961 (1,946 units US, 3,015 units OUS) | Worldwide Distribution - US (nationwide): All states in continental USA including DC except DE, ME, ND, RI &amp; VT. and the countries of: ARGENTINA, AUSTRALIA, AUSTRIA, BAHAMAS, BAHRAIN, BELGIUM, BOSNIA AND HERZEGOVINA, BRAZIL, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT ESTONIA, FINLAND, FRANCE, GERMANY, GREECE, GHANA, HONG KONG, HUNGARY, INDIA, IRAQ, IRELAND, ISRAEL, ITALY, JAPAN, JORDAN, KAZAHKSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT LEBANON, LITHUANIA, MACEDONIA, MAURITIUS,MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY,, PANAMA, PERU, POLAND, PORTUGAL, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, VIET NAM, and URUGUAY. | Terminated | October 21, 2015 | 516 | 4961 | Not_Computer
Z-1825-2014 | 68535 | No | ISW | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | RUSH81 Prosthetic Foot |  |  | June 02, 2014 | 2014 | Ability Dynamics LLC | Ability Dynamics is recalling RUSH81 Prosthetic Foot because the existing bolts may be bottomed out in the threaded hole and this could lead to bolt failure. | Device Design | Ability Dynamics sent an Urgent Medical Device Correction Notice letter dated June 2, 2014 to all affected customers (Prosthetist) who purchased the RUSH Prosthetic Foot. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to fill out the Bracket Bolt Replacement Form and return to Ability Dynamics. Customers with questions are instructed to contact Rich Sainz, National Clinical Manager, 2082 E. University Dr. Tempe AZ 85281 (480) 717-1412, rsainz@abilitydynamics.com. |  | 175 | 175 units | Worldwide Distribution - US (nationwide) and in the countries of Australia and Canada. | Terminated | September 22, 2014 | 112 | 175 | Not_Computer
Z-1824-2014 | 68374 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Spacelabs Healthcare Ventilator Flexport Interface, Model 90436A-07 |  |  | May 09, 2014 | 2014 | Spacelabs Healthcare Inc | Spacelabs Healthcare is voluntarily recalling the Hamilton Galileo Ventilator Flexport, Model 90436A-07, where the monitored Minute Volumes (Vmin) has been reported at one time to reach ten times the actual value on the bedside monitor. | Component change control | Spacelabs changed their strategy for updating the software and sent the updated letter Urgent-Medical Device Correction- Hamilton Galileo Ventilator FlexportModel 90436A-07, dated 10 June 2015, to their consignees. The firm is going to send consignees new replacement flexport(s) with upgraded software and requested consignees to return old affected units via Return Goods Authorization (RGA) to Spacelabs and the firm will destroy affected units.Consignees are instructed to take the following actions:1. Please replace the affected devices and test for correct operation.2. Return your old device(s) using the RGA number and prepaid FedX shipping label.Use the Reference code: |  | 154 | 46 units distributed within the US; 108 units distributed outside the US. | Worldwide Distribution-USA (nationwide) including the states of AZ, IN, LA, GA, SC, FL, and NY, and 18 customers outside the US. | Terminated | November 16, 2015 | 556 | 154 | Computer
Z-1827-2014 | 68311 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Philips MR systems: Trolley Variable Height IRF (for Ingenia 1.5T &amp; 3.0T products) |  |  | May 15, 2014 | 2014 | Philips Medical Systems, Inc. | Ingenia customers have experienced clamping of the foot under the central column of the Height Adjustable Flex Trak trolley and the Trolley Variable Height IRF. | Device Design | Philips sent an Urgent Field Safety Notice (FSN78100414) dated May 2014 via certified mail to US and Canadian consignees. Philips Key Market representatives will distribute the letter internationally. The letter identified the affected product, explains the problem, actions that should be taken in order to prevent risks for users, and the actions planned by Philips to correct the problem. On May 31, 2014, Philips released a Field Change Order that requires service engineers to replace the front and rear covers of the Ingenia MR FlexTrak Trolley with new covers with protection bumper that prevents foot clamping. For further information and support concerning this issue, consignees should contact their local Philips representative at 1-800-722-9377. |  | 5 | 5 | Worldwide Distribution - US (nationwide), Canada, Mexico Australia, Austria, Germany, Belarus, Bulgaria, Brazil, China, Denmark, Estonia, Finland, France, Germany, Greece, Italy, Jordan, Korea, Republic of Korea, Netherlands, New Zealand, Norway, Oman, Pakistan, Russian Federation, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, and Vietnam. | Terminated | August 07, 2015 | 449 | 425 | Not_Computer
Z-1837-2014 | 68310 | No | PBI | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | G7 Acetabular System Positioning Guide Rod, orthopedic surgical instrument for hip prosthesis PN 110003458. |  |  | May 07, 2014 | 2014 | Biomet, Inc. | Guide rod may fracture during impaction of the acetabular shell and be retained in the inserter or the patient and could result in a delay in surgery. | Use error | The firm issued the following notice on May 7, 2014: URGENT MEDICAL DEVICE RECALL NOTICE:This notification is to inform you of an Urgent Medical Device Recall initiated by Biomet Orthopedics ("Biomet") which involves Part Numbers: 110003458 G7 Positioning Guide Rod and 110003500 G7 Positioning Guide Post. These instruments have been consigned and/or invoiced to your account. Biomet has initiated this action following an investigation which identified that positioning guide rod may fracture during impaction of the acetabula shell. If the positioning guide rod fractures, then the positioning guide post may become stuck to the insertion handle. There is no adverse health outcome expected for the patient.This action requires the immediate location and discontinued use of the product and its return to Biomet.The following actions are REQUIRED:Immediately locate and remove the identified device(s) listed below from circulation.Carefully follow the instructions on the enclosed "FAX Back Response Form .Fax a copy of the Response Form to 574-372-1683 prior to return of product.Use priority carrier for your shipment.If you have further distributed this product, you MUST notify hospital personnel of this action via the enclosed "Dear Risk/Recall Manager" notice. This letter MUST be given to hospital personnel responsible for receiving recall notices. However, you are charged with the location and return of these products. |  | 1047 | 1047 total | Worldwide Distribution-USA (nationwide) and the countries of Canada, Netherlands, New Zealand, and Japan. | Terminated | October 30, 2015 | 541 | 1047 | Not_Computer
Z-1834-2014 | 65274 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | NIPRO Safe Touch Safety Scalp Vein Set |  |  | May 04, 2013 | 2013 | Nipro Medical Corporation | There is a possibility of a crack in the CP luer connector. The position of the crack is at the connection site of the CP luer connector and the tubing. The crack could cause or contribute to leakage of fluids. | Component design/selection | Nipro sent a Recall Notification letter dated April 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This recall does extend to the customer level. Please notify all customers of the recall and collect all affected product that is in their possession. When all products have been received from your customers call tags and RMAs will be issued in order for you to return all products to Nipro Medical warehouse in Memphis TN. You will then be reimbursed for all products received.Please place all products you have in stock on hold. Use the attached faxback to respond to this notification, with product codes, lots and quantities you have in all of your distribution facilities. One form should be completed for each distribution location. Return forms as soon as possible to Nipro Medical, fax: 305.592.4621. Any questions or concerns can be addressed by phone at 305.599.7174 ext. 249. |  | 176000 | 176,000 pieces | Worldwide Distribution - USA and the countries of Canada, Bolivia, Chile, Uruguay, Belize, Dominica, Honduras, Jamaica, Paraguay and Venezuela. | Terminated | March 02, 2015 | 667 | 176000 | Not_Computer
Z-1833-2014 | 68072 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Synthes External Fixation Systems (Small, Medium, Large and DO), bone fixation appliance components. |  |  | April 10, 2014 | 2014 | Synthes, Inc. | Labeling changes have been made related to MR (Magnetic Resonance imaging) conditions as a result of changes in required protocols to designate a product MR Safe, MR Conditional, or MR Unsafe. | Labeling Change Control | Synthes sent an Urgent Notice dated April 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:All Synthes External Fixation devices should be treated as MR Conditional."Synthes asks that you review the information contained in this Labeling Notification and complete the Verification Section located on Page 3."Discard outdated revisions of the Technique Guides noted in the table below."Update your records with updated Labeling Infonnation."Forward this Medical Device Labeling Update to anyone in your facility that needs to be infonned, especially those personnel that conduct MR testing."If the Verification Fonn is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual in pg. 3."Updated product literature can be located on the Synthes website athttp://www.syntheskyo.com/trauma kyo/home/home.htm."Please see the attached inseii for current conditions for use in the MR environment."Maintain a copy of this notice.Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 5114052 | 5,114,052 | Worldwide Distribution - USA (nationwide) and Canada. | Open, Classified |  |  | 5114052 | Not_Computer
Z-1832-2014 | 68320 | equipment, laboratory, general purpose, labeled or promoted for a specific medical use | LXG | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VersaCell X3 Solution and VersaCell X3 Expansion Pack, a robotic laboratory sample-management system. It connects up to 3 instruments and provides a user interface that routes samples, reduces manual tasks and consolidates results. |  |  | April 25, 2014 | 2014 | Siemens Medical Solutions Diagnostics | The keyboard and monitor may not be properly secured to the monitor arm. | Device Design | Siemen's Healthcare Diagnostics sent an Urgent Medical Device Correction letter dated April 25, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return teh Field Correction Effectiveness Check form attached to the letter within 30 days by fax to 302-631-7597.Inspect the monitor, keyboard and bracket assembly. If the screw and washer assembly is secure as shown in Figure 1, the monitor and keyboard assembly is safe to use. After making adjustments or repositioning the monitor, keyboard and bracket assembly, inspect the screw adn washer to ensure it is ssecure as shown in Figure 1.If there is a gap between the underside of the monitor support bracket and cylindricalmounting feature as shown in Figure 2, contact your Siemens technical supportrepresentative to schedule a service visit as soon as possible. Front load and manage samples directly on the connected instruments until your Siemens technical support representative addresses the issue.Customers were instructed to retain the letter with their laboratory records, and forward the letter to those who may have received the affected product. Customers with questions were instructed to contact their Siemens Customer Care Center or their local Siemens technical support representative.For questions regarding this recall call 877-229-3711. |  | 19 | 19 units | Worldwide Distribution - USA including Iowa and New York and Internationally to Canada, Chile, France, India, Italy, Malaysia, Peru, Korea, and the United Kingdom. | Terminated | April 16, 2015 | 356 | 19 | Not_Computer
Z-1844-2014 | 68444 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 840 Ventilator (4-840120DIUU-XX) |  |  | June 03, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is conducting a medical device field correction for specific Puritan Bennett 840 Ventilators due to customer reports of loss of graphical user interface (GUI) display information while the ventilator continues to provide breath support. | Nonconforming Material/Component | Covidien sent an Urgent Medical Device Voluntary Field Correction letter dated 6/13/2014, to all affected customers. The letter identified the product the problem and and the action needed to be taken by the customer. Customers are requested to forward the recall letter to anyone they may have transferred or assigned for use one of the potentially affected ventilators. Service professionals will be identifying the affected PSUs and replacing them. The removed PSUs will be destroyed.To verify whether your devices are affected by this field safety corrective action and to initiate the process to receive the actions mentioned above, please complete the attached verification form and fax it to the Covidien contact indicated on the form. If you need further assistance, contact our Technical Support Department @ 1-800-255-6774, option 4, then option 1. |  | 104 | 104 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY., and the countries of Argentina, Australia, Belgium, Brazil, Canada, Chile, Colombia, Costa Rica, Egypt, Germany, Guatemala, Israel, Japan, Kazakhstan, Kenya, Lebanon, Libya, Macedonia, Mexico, Montenegro, Morocco, Panama, Peru, Russia, Saudi Arabia, Singapore, Switzerland, Taiwan, Tunisia, Turkey, Yugoslavia (Serbia &amp; Montenegro) Including GOVT/VA/Military | Open, Classified |  |  | 104 | Not_Computer
Z-1839-2014 | 67610 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Altrus Thermal Tissue Fusion Handpieces under the following specifications: 1) 10mm x 16 cm Catalog Number: 60-9510-001, 2) 10 mm x 23 cm Catalog Number: 60-9511-001, 3) 10 mm x 36 cm Catalog Number: 60-9512-001, 4) 5 mm x 16 cm Catalog Number 60-9520-001, 5) 5 mm x 23 cm Catalog Number 60-9521-001 and 6) 5 mm x 36 cm Catalog Number 60-9522-001. Products are packaged within cartons, 6 units per carton. |  |  | February 20, 2014 | 2014 | ConMed Corporation | ConMed Corporation initiated an Urgent Medical Device Correction for Altrus Thermal Tissue Fusion Handpieces due to risk of burns. | Device Design | On February 20, 2014 and February 28, 2014 ConMed Corporation distributed Urgent Medical Device Correction notices to their domestic and foreign customers via FedEx Priority. The notice was generated by ConMed for information purposes only. Customers are to ensure that the jaws of the energy vessel sealing device does not come in contact with unintended tissue, surfaces, or objects until the jaws have cooled sufficiently. Customers who have further distributed this product, please forward a copy of the correction notice to the recipient as appropriate. Customers will questions should contact ConMed Regulatory Affairs at (315) 624-3237 or by email at altrus@conmed.com. |  | 12653 | 12,653 units (10,769 domestically &amp; 1,884 internationally) | Worldwide Distribution. | Terminated | April 07, 2017 | 1142 | 12653 | Not_Computer
Z-1831-2014 | 67058 | detergent | JCB | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Coulter DxH Cleaner 10L, Part No. 628023. |  |  | November 08, 2013 | 2013 | Beckman Coulter Inc. | Beckman Coulter is recalling the UniCel DxH800 and UniCel DxH600 Coulter Analysis Systems because the optical degradation may potentially cause a delay in reporting results for nucleated red blood cells (nRBC), automated white blood cell differential and the reticulocyte. | Under Investigation by firm | An Urgent Product Correction letter, dated 11/08/2013, was sent to all the customers who purchased the UniCel DxH 800 and UniCel DxH 600 Coulter Cellular Analysis Systems. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Beckman Coulter's Customer Support Center via website, http://www.beckmancoulter.com/customersupport/support, via phone at (800) 526-7694 in US and Canada. For outside of the US and Canada, customers are instructed to contact their local Beckman Coulter Representative.Beckman Coulter sent an update on 2/14/14, to inform customers that the recall was expanded to include specific lots of DxH Cleaner. |  | 12924 | 12,924 units total (9,733 units in US) | Worldwide Distribution -- US, Canada, Australia, China, Belgium, Brunei Darssalam, Colombia, Croatia, Czech Republic, France, Germany, Gibraltar, Greece, Hong Kong, Hungary, India, Ireland, Italy, Japan, Korea, Kuwait, Macao, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Oman, Philippines, Poland, Panama, Portugal, Puerto Rico, Qatar, Romania, Russia Federation, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Turkey, United Arab Emirates, United Kingdom, and Vietnam. | Terminated | December 25, 2014 | 412 | 12924 | Not_Computer
Z-1843-2014 | 68352 | No | PAM | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | FirmArray Blood Culture Identification (BCID) Panel, Model 2.0 |  |  | May 27, 2014 | 2014 | BioFire Diagnostics, Inc. | BioFire has identified an increased risk of false positive results when the FilmArray Blood Culture Identification (BCID) Panel is used with bioMÂżrieux BacT/ALERT Standard Anaerobic (SN) blood culture bottles. False positive results have been observed for Pseudomonas aeruginosa and Enterococcus. | Labeling design | BioFire Diagnostics sent an Urgent Recall/Correction letter on May 28, 2014. Customers were instructed If the BCID Panel is used to test BacT/ALERT SN bottles, positive results for Pseudomonas aeruginosa and Enterococcus should be confirmed by another method prior to reporting the test results.Customer with questions were instructed contact customer support at Support@biofiredx.com or via telephone by dialing 1-800-735-6544 and selecting option 5 for customer support and then option 1 for FilmArray.For questions regarding this recall call 801-736-6354, ext 360. |  | 1084 | 1084 | Worldwide Distribution - USA (nationwide) European Union, Hong Kong, and VA. | Terminated | November 07, 2014 | 164 | 1084 | Not_Computer
Z-1842-2014 | 68339 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | VerSysÂż Hip System, Beaded Fullcoat Stems, Zimmer, Inc., Warsaw, IN 46581 |  |  | May 19, 2014 | 2014 | Zimmer, Inc. | During routine inspection, beaded hip stem forgings received from supplier failed to meet ZES 2A-102. Processing of Beaded CoCrMo Alloy Hip Stem, material requirements per ATS #14-04818 and ATS #14-05549, respectively. | Nonconforming Material/Component | On 6/2/2014 the firm sent "Urgent Medical Device Recall" letters to their customers. On 8/13/14, Zimmer contacted a distributor via phone pertaining to one additional lot (00-7843-011-06; lot 62547211) which had been included in the recall. Formal recall notices were issued to the distributor via email on 8/15/14. |  | 147 | 147 units | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CO, CT, FL, GA, IL, IN, MA, MI, MN, NC, ND, NJ, NY, OH, OK, PA, TN, TX, VA &amp; WI, and the countries of Canada, Colombia, Dominican Republic, Japan, Korea (S. Korea), Peru, Taiwan &amp; Thailand. | Terminated | March 24, 2015 | 309 | 147 | Not_Computer
Z-1840-2014 | 68379 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens LANTIS System ; LANTIS Commander, 10 Users ; LANTIS 6.1 Commander ; LANTIS 8.3 Commander |  |  | May 06, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential risk when using LANTIS OIS System client software with operating systems for which it has not been validated and released, resulting in incorrectly stored data, such as treatment records, and LANTIS database corruption. | Use error | Siemens sent an Urgent Medical Device Correction Field Safety Notice dated was sent to end users on/about May 6, 2014, to notify them of the safety issue and provide instructions to avoid potential risks. Customers were instructed to make sure all computers with the LANTIS client software installed are running on an operating system that was validated by Siemens for opration with LANTIS, e.g. Microsoft Windows XP Service Pack 3 (SP3). Customers were asked to include the Field Safety Notice in their System Owner Manual chapter "Safety Advisory Letters" where it should remain and that all affected personnel affected be informed.For questions regarding this recall call 1-610-219-6300. |  | 108 | 108 | Nationwide Distribution | Terminated | December 15, 2014 | 223 | 108 | Computer
Z-1841-2014 | 68362 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Integra UCR 3mm Hex, Large Axial Driver. |  |  | May 21, 2014 | 2014 | SeaSpine Inc | Integra has identified through an investigation of complaints that there may be the potential for the tip of asingle lot number of the UCR 3mm Hex, Large Axial Driver, 91-1123 to twist or become rounded duringuse. | Under Investigation by firm | Integra sent an Urgent Voluntary Medical Device Recall letter dated May 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed:1. If you have distributed the affected product, identify those customers and forward the attached Integra Medical Device Recall Notice to them.2. Inspect your inventory of UCR 3mm Hex, Large Axial Drivers (Catalogue Number: 91- 1123). If you have product with lot number MO16090B, STOP USING IMMEDIATELY. If you have product of any other lot number, you can continue to use it. The event has been isolated to this single lot.3. We also request that regardless if you have a UCR Driver with the affected lot, you complete the attached form and return it to us as soon as possible as indicated. We recommend you keep a copy of the form for your records.4. When your form is received, Customer Service will contact you and provide an RMA number, directions to return the product, and input an order to replace the quantity you indicate on the form.Should you have any questions regarding these instructions, please contact Customer Service at 1-866- 942-8698.We apologize for any inconvenience this may cause and thank you for your cooperation in this effort.:. |  | 7 | 7 | US Distribution including the states of : AL, CA, OH, NM and TN. | Terminated | December 15, 2014 | 208 | 7 | Not_Computer
Z-1845-2014 | 66315 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens RAD Fluoro Uro System (Ysio, Luminos dRF, and Uroskop Omnia) with software versions VB10G. |  |  | August 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc | When using systems operating with software versions VD10A/G during a RAD examination, a malfunction can result in the image appearing very dark, which may prompt unnecessary repetition of the examination. | Software design | Siemens sent a notification letter, dated August 28, 2013, to Affected customers. The letter identified the affected product, problem and actions to be taken. The letter provided work-around instructions and announced that a field modification was prepared to resolve the issue. |  | 76 | 76 | USA Nationwide Distribution in the states of OH, NJ, MA, CO, NY, CA, MD, OR, KY, PA, IL, VA, TX, MN, KS, GA, and WY. | Terminated | February 24, 2015 | 545 | 76 | Computer
Z-1846-2014 | 68384 | No | LHN | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ProBeam Proton Therapy System (HGA and HFB) |  |  | May 09, 2014 | 2014 | VARIAN MEDICAL SYSTEMS PARTICLE THERAPY GMBH | Anomaly with the ProBeam System where under certain conditions, the Treatment Control and Monitoring application could fail to send treatment history records to the ARIA database. | Under Investigation by firm | Varian sen an Urgent Device Medical Correction/Urgennt Field Safety Notice letter dated May 9, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. For questions cotact Varian Service Support. at debg-PT-service-support@varian.com. |  | 1 | 1 | US Distribution in tthe state of CA | Terminated | November 12, 2014 | 187 | 1 | Not_Computer
Z-1847-2014 | 68256 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | Abacus 3.1, Pharmacy Calculator |  |  | May 28, 2014 | 2014 | Baxter Corporation Englewood | Baxter Corporation is initiating a field correction following an upgrade to ABACUS Software v3.1, when Clinimix 4.25110 is used in an order, the Nutritional Summary and any labels containing energy information may display 1,000 times less protein than is actually present in the solution. | Software design | Baxter sent an Urgent Device Correction letter dated May 27, 2014 via certified mail to affected customers. The letter identified the affected product, problem and actios to be taken. Customers were instructed to: 1. Review and distribute this communication to pharmacy staff and compounding staff that use ABACUS software. This notice is intended only for customers that have upgraded to version 3.1 of the software.2. Please acknowledge receipt of this notification by completing the enclosed customer reply form and return it to Baxter by either faxing it to 1-224-270-5457 or scanning and e-mailing it to fca@baxter.com.3. If you have distributed ABACUS to other facilities, please notify the end users in accordance with your customary procedures. For questions call Baxter Techenical Support line at 1-800-678-2292. |  | 42 | 42 | US Nationwide Distribution and the state of VA. | Open, Classified |  |  | 42 | Computer
Z-1745-2014 | 67879 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT Big Bore Oncology Whole Body Computed Tomography X-Ray System &amp; Brilliance CT Big Bore Radiology Whole Body Computed Tomography X-Ray System. |  |  | December 19, 2013 | 2013 | Philips Medical Systems (Cleveland) Inc | During internal testing, the Brilliance CT Big Bore was found to be out of tolerance for radio frequency emissions. Specifically, at the 48MHZ frequency, the testing indicated the Brilliance CT Big Bore was 3.5dB uV/meter higher than the applicable IEC 60601-1-2 standard specification. | Nonconforming Material/Component | Philips Medical System sent an Urgent Medical Device Corrections letters dated December 10, 2013 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter informed customers that a field service Engineer will visit the facility to perform an Electro-Magnetic Interference (EMI) noise reduction up0date on each system. For information or support concerning contact your local Philips representative:For North America and Canada, contact the Customer Care Solutions Center (1-800-722-9377, OptionS: Enter site ID or follow prompts). |  | 454 | 454 Units | Worldwide Distribution - US (nationwide) in the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI &amp; WV.and countries of Algeria, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, China, Czech Republic, Denmark, Ecuador, France, French Polynesia, Gabon, Germany, India, Indonesia, Israel, Italy, Jordan, Libyan Arab Jamahiriya, Lithuania, Luxembourg, Malaysia, Martinique, Mexico, Morocco, Netherlands, New Zealand, Oman, Palestine, Philippines, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syrian Arab Republic, Thailand, Ukraine &amp; United Kingdom. | Terminated | August 25, 2015 | 614 | 454 | Not_Computer
Z-1848-2014 | 68456 | No | FCN | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | SteriGear Urinary Drain Bag with Fig Leaf Cover, 2000 ml, REF 10270, The Fig Leaf, SteriGear 362 S. University Avenue, provo, UT 84601, 1-800-398-3259, SteriGear.com, Sterile, EO |  |  | January 11, 2011 | 2011 | SteriGear, LLC | SteriGear notifed customers of circumstances in which The Fig Leaf urinary drain bag valve can close creating a negative pressure that can suck the non-return valve closed. | Device Design | SteriGear, LLC sent a special note to all affected customers between January 12, 2011 and September 28, 2012. The note identified the product, the problem, and the action to be taken by the customer. Customers were instructed on how to reposition the inlet tube. Customers with questions were instructed to contact Director of Product Development at 801-607-3211. No correction or removal was performed.For questions regarding this recall call 801-607-3209. |  | 110560 | 110,560 | Nationwide Distribution | Terminated | January 02, 2015 | 1452 | 110560 | Not_Computer
Z-1849-2014 | 68466 | No | CAN | Pressure regulator. | Anesthesiology | Anesthesiology | 510(K) Exempt | Model RG3000 and RG4000 Series Valve Integrated Pressure Regulator (VIPR) |  |  | April 07, 2014 | 2014 | Western / Scott Fetzer Company | Separation events have occurred, including the VIPR separating from the Aluminum cylinder. | Device Design | All initial consignees of subject product were notified by letter dated June 5, 2014 via U.S. mail. Initial consignees were directed to forward copies of the recall notification letters to their contacts in the event the subject product has been further distributed. |  | 118325 | 118,325 units | Worldwide Distribution-USA (nationwide) including Puerto Rico and the states of Alaska, Arkansas, Arizona, California, Colorado, Connecticut, Delaware, Florida, Georgia, Hawaii, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Michigan, Minnesota, Missouri, Mississippi, Montana, North Carolina, North Dakota, Nebraska, New Hampshire, New Jersey, New Mexico, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, Wisconsin, and West Virginia and the countries of Canada,China, Costa Rica, France, Honduras, India, Ireland, Italy, Japan, Korea (South), Malaysia, Panama, Spain and Taiwan. | Open, Classified |  |  | 118325 | Not_Computer
Z-1852-2014 | 68491 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare Qube Compact Monitor, Model 91390. |  |  | June 06, 2014 | 2014 | Spacelabs Healthcare Inc | Spacelabs Healthcare is recalling Qube Compact Monitor, Model 91390, due to a potential chance of losing network connection after switching from hardwired Ethernet connection to using the Qube Docking Station without powering down the device. The bedside monitor and all of its alarms will continue to function normally. No one has been injured as a result of this issue. | Component design/selection | Spacelabs Healthcare Snoqualmie, WA sent an Urgent Medical Device Correction letter, dated June 6, 2014, to all consignees. The letter identified the product, the problem, and the actions to be taken by the customer. The letter instructed customers using docking station ready qube monitors with qube docking stations immediately advise their staff of this situation; and, to follow the steps provided, if they continue to use a docking station ready qube monitor with a docking station.Spacelabs Healthcare will contact customers at the earliest possible date to schedule a convenient time for Spacelabs to correct customer's affected qube monitors at no cost. Questions should be directed to Spacelabs Healthcare Technical Support at 1-800-522-7025. |  | 4842 | 3383 in the USA; 1459 outside the USA | Worldwide Distribution -- USA, including the states of AL, AZ, CA, CO, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, ME, MI, MN, MO, MS, MT, NC, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, WA, WI, and WY; and, the countries of AFG,ARE, ARG, AUS, BHR, BOL, CAN, CHE, CHL, CHN, COL, CRI, DEU, DOM, ECU, FRA, GBR, IND, ITA, JOR, KWT, LBY, MYS, NLD, OMN, PAN, POL, PRT, ROU, SAU, SGP, THA, TTO, TUN, TUR, TWN, and VNM. | Terminated | February 17, 2015 | 256 | 4842 | Not_Computer
Z-1850-2014 | 68128 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Vertical Spring Arm for Single or Dual Monitor Mount (Harmony LL 500 and Harmony LL 700 surgical lighting systems), Steris Corporation. Model numbers YG19051 and YG19262. |  |  | April 04, 2014 | 2014 | Steris Corporation | Potential for the welded seam in affected monitor spring arms to crack and/or break. | Device Design | On April 11, 2014 the firm sent Urgent Voluntary Field Correction Notices to their customers. |  | 68 | 68 units | Distribution to US (states: MA, NY, OH, PA &amp; TX) and Canada. | Terminated | October 15, 2014 | 194 | 68 | Not_Computer
Z-1851-2014 | 68387 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | RUSCH, LaserTube (Rubber), Laser Resistant Tracheal Tube, cuffed, Teleflex Medical, Willy-Rusch, GmbH. |  |  | May 23, 2014 | 2014 | Teleflex Medical | Although the Natural Rubber Latex caution symbol is clearlydisplayed on the product label, the accompanying caution statement "Caution: This product contains Natural Rubber Latex which may cause Allergic Reactions" was not printed on the product label. | Error in labeling | Teleflex Medical sent a Urgent Medical Device Recall Notification dated June 6, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you have received products that are subject to this recall. We are now notifying our customers to take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the catalog numbers listed above from the lot numbers listed on the attachment.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866- 804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 1658 | 1658 ea | Nationwide Distribution and the country of Colombia. | Terminated | January 13, 2016 | 600 | 1658 | Not_Computer
Z-1853-2014 | 68485 | refractor, manual, non-powered, including phoropter | HKN | Manual refractor. | Ophthalmic | Ophthalmic | 510(K) Exempt | RT-5100 Refractor and RT-3100 Refractor; |  |  | June 09, 2014 | 2014 | Nidek Inc | Near Point Chart Arm of RT 5100 and RT-3100 Refractors may lower spontaneously with potential for patient injury. | Use error | Recall/Field Correction letters were sent to doctors on June 9, 2014 with a preventive maintenance instruction to be performed on the device by user. Also included are the necessary tools to perform actions. Customers are to complete and return the Recall Response Card. Questions are to be directed to Customer Service.Contact Information: Customer Service: 1-800-891-1772 E-mail: RT-5100 update@noritsu.com |  | 167 | ~4,285 units RT-5100; 167 units of RT-3100 | US distribution only. | Terminated | October 30, 2015 | 508 | 167 | Not_Computer
Z-1864-2014 | 68588 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Ingenia 3.0T, magnetic resonance imaging system. |  |  | June 13, 2014 | 2014 | Philips Medical Systems, Inc. | For some identified Ingenia systems shipped between December 2013 and March 2014, cover set rear magnetic M10 nuts were erroneously supplied instead of A4 stainless steel nuts. | Nonconforming Material/Component | On June 14, 2014, Field Safety Notice (FSN 78100432), was be sent to all US and Canadian consignees via certified mail. Philips Key Market representatives will distribute the letter internationally.The FSN instructs consignees to not let hospital technical staff or any person other than Philips service engineers dismount the rear covers before the system is inspected and/or corrected by Philips. Philips has alerted installation and service engineers to this hazard. A Field Change Order with reference FCO781 00432 is being released directing Philips field service engineers to inspect the Ingenia rear cover set nuts and replace them if needed with non-magnetic stainless steel nuts. For further information or support concerning this issue, consignees should contact their local Philips representative at 1-800-722-9377. |  | 18 | 18 | Worldwide Distribution-USA (nationwide) including the states of AL, AZ, CA, CO, GA, ME, MI, NY, OH, PA, TX, and WA, and the countries of Mexico, Argentina, Australia, Belarus, Belgium, Brazil, Chile, China, France, Germany, Hong Kong, India, Indonesia, Ireland, Italy, Japan, South Korea, Laos, New Zealand, Norway, Palau, Poland, Portugal, Qatar, South Africa, Sweden, Switzerland, Taiwan, Tunisia, Turkey, Turkey, United Kingdom, Venezuela, and Vietnam. | Terminated | January 27, 2016 | 593 | 96 | Not_Computer
Z-1854-2014 | 68354 | Yes | FKR | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NxStage Dialysate Sack (SAK) with Concentrate |  |  | May 14, 2014 | 2014 | NxStage Medical, Inc. | SAK Dialysate Concentrate contain aluminum levels which exceed internal product specification | Material/Component Contamination | NxStage issued Initial recall letter dated 5/14/14 to dialysis centers. An error was noted and an updated letter dated 5/15/14 was resent to dialysis centers. The updated 5/15/14 letter was emailed to all current active PureFlow SL patients who had provided an email address followed by UPS next business day letter sent to all current active PureFlow SL patients. Accounts are requested to examine inventory and remove product. Contact NxStage Customer Service to arrange for return of all affected product and forreplacement product to be sent. If you have any questions or comments, please feel free to contact NxStage Customer Service at 1-866-NXSTAGE (1-866-697-8243). |  | 13474 | 13474 cases | Worldwide Distribution - USA (nationwide) and Internationally to France, Netherlands, Spain, Sweden, and the UK. | Terminated | March 13, 2017 | 1034 | 69334 | Not_Computer
Z-1866-2014 | 68259 | [6 empty] | LAP LASER PICTOR 3D Multi Color Laser System |  |  | March 21, 2014 | 2014 | LAP of America Laser Applications, LLC | While projecting yellow lines it was notice that the yellow projection has the potential to splits into two separate red and green projections. In this case the difference between the red and green projection was up to 2.5mm. | Software design | LAP of America Laser sent an Urgent safety Information FFST 2014-2-3 via UPS Next Day Air to the customer. |  | 1 | 1 system | North Carolina | Open, Classified |  |  | 1 | Not_Computer
Z-1867-2014 | 63486 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens StreamLAB(R) Analytical Workcell; Catalog number 746002.901. |  |  | October 11, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc. | Firm has received reports from customers that the small pins at the bottom of the center door panel of the Input/Output Module, which protrude approximately 5/8 inch (16 mm), have contributed to trip and fall incidents. Customers are advised that when working in the area of the Input/Output Module to be aware of the two pins that protrude and use caution. | Device Design | The firm initiated their recall/correction of this product on October 11, 2012 by sending an Urgent Customer Notification letter to all StreamLAB(R) Automation Solutions customers. The letter informed customers that the firm received reports from customers that the small pins at the bottom of the center door panel of the Input/Output Module have contributed to two trip and fall incidents. Customers were advised that when working in the area of the Input/Output Module to be aware of the two pins that protrude and use caution. Also, StreamLAB operators were to be notified of the potential safety issue and the notice was to be posted on or near the system. Additionally, customers were to complete and return the attached Field Correction Effectiveness Check form. |  | 317 | 317 | Nationwide Distribution -- AL, AR, AZ, CA, CO, CT, FL, GA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, WI, and WV. | Terminated | August 26, 2014 | 684 | 317 | Not_Computer
Z-1869-2014 | 68386 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ |  |  | May 29, 2014 | 2014 | Elekta, Inc. | MOSAIQ does not re-calculate the agent volume when the user changes the ordering dose (to other than what was calculated) prior to saving the order. | Software design | Elekta issued an Important Safety Notice dated May 29, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Corrective action that shall be taken:Notify the affected customers of the issue - To be completed 06/22/2014Follow up with the affected customers for:Confirmation receipt of the LIN, ANDConfirmation of installation of patch for the fix- To be completed 11/30/2014The notice includes a confirmation which is to be returned.For further questions call (770) 670-2422. |  | 101 | 101 | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AL, AK, AZ, CA, CO, CT, DE, FL, GA, HI, IL, IN, IA, KY, LA, ME, MD, MA, MI, MO, MT, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TX, VT, VA and WA,. and the countries of Australia, Canada, New Zealand and United Kingdom. | Open, Classified |  |  | 101 | Not_Computer
Z-1868-2014 | 68330 | No | MBP | Resorbable calcium salt bone void filler device. | Orthopedic | Physical Medicine | 510(k) | Accell Evo3C Demineralized Bone Matrix, 5 cc, Model Number: 02-6000-050 , Lot Number: 132160. |  |  | May 19, 2014 | 2014 | Integra LifeSciences Corp. | This lot of Accell Evo3C Demineralized Bone Matrix, 5cc was incorrectly packaged in boxes that describe the product as Accell Evo3 Demineralized Bone Matrix. Evo3C is contraindicated for treatment of vertebral compression and not indicated for use in the posterolateral spine. | Packaging | Integra sent an Urgent Voluntary Medical Device Recall dated May 19, 2014 to customers via mail (traceable courier service). The recall notification letter identified affected product, reason for recall and actions to be taken Customers were instructed to complete the "RECALL ACKNOWLEDGEMENT AND RETURN FORM". and return to:FCA1@integralife.com or FAX to 1-609-275-9445. For questions contact Customer Service at 1-800-550-7155. |  | 61 | 61 | US Nationwide Distribution | Terminated | September 05, 2014 | 109 | 61 | Not_Computer
Z-1873-2014 | 68343 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Glucose Hexokinase Liquid Reagent For the in vitro quantitative determination of Glucose in serum. |  |  | April 07, 2014 | 2014 | JAS Diagnostics Inc. | JAS Diagnostics, Inc. of Miami Lakes, FL is recalling their Glucose Hexokinase reagent due to certain lots not performing according to their specifications. | Mixed-up of materials/components | JAS Diagnostics, Inc. sent an Important Product Recall Information letter dated April 7, 2014 to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to: 1) Screen their inventory for any of the lots listed in the letter and discontinue use immediately. 2) Contact the Customer Service Department at 305-748-2762 or 305-748-2763 to obtain replacement product. |  | 407 | 407 | Worldwide Distribution - US including Puerto Rico and the states of CA, FL, TX, PA, GA, KY, MI, NY, OK, MN, TN, VA, NC, NV, NJ and WV, and the countries of Barbados, Ghana, Trinidad, Tobago, Canada and Haiti. | Terminated | August 21, 2014 | 136 | 407 | Not_Computer
Z-1812-2014 | 68419 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | HydroFinity Hydrophilic guidewire; Nitinol Guidewire with Hydrophilic Coating, Sterile. |  |  | May 28, 2014 | 2014 | Nitinol Devices and Components, Inc. | Reports of outer polymer jacket to the core wire being damaged during use on the HydroFinity Guidewire. | Device Design | Covidien sent an Urgent Product Recall letter dated May 30, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return any unused devices to Covidien, their account would be credited. Customers were also instructed to::1) Stop using product listed in this letter immediately2) Segregate this product from other inventory3) Fill out the reply (verification) form at the end of this letterÂż If you do not have any product identified in this letter, please fax oremail the completed form to Covidien at (877) 523-9109 orcovidien8770@stericycle.com.Âż If you do have product, your sales representative will assist you incompleting the verification form and arranging for return of theproduct.Âż Please fax the completed form to Covidien (877) 523-9109 or email tocovidien8770@stericycle.com.4) Your sales representative will be available to answer any questions regarding this recall and assist you in completing the verification form, returning product and addressing any account credits.This action is being conducted with the knowledge of the United States FDA and other regulatory authorities. Âż Customers with questions were instructed to call (202) 310-5120.For questions regarding this recall call 202-310-5120. _ |  | 8485 | 8485 in US, 9595 - ROW - total, all models | Worldwide Distribution - USA (nationwide) and Internationally to Italy, France, Norway, Sweden, Germany, Belgium, United Kingdom and Canada. | Terminated | January 13, 2015 | 230 | 8485 | Not_Computer
Z-1870-2014 | 68531 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | IS4000 da Vinci Xi - Endoscope accessory; |  |  | June 11, 2014 | 2014 | Intuitive Surgical, Inc. | In some procedures when using the endoscope in conjunction with the IS4000 (da Vinci Xi) System, the surgical view has an orange-yellow tint which can make it hard to differentiate between tissue | Device Design | Medical Device Recall notifications were sent to the 4 affected sites on June 11, 2014. Letters describe the problem and the potential health risk. The two model numbers of the endoscope are listed and the statement that ISI will perform a software upgrade to the da Vinci System and replace the endoscopes.Letters request that the following actions take place:- Ensure that all appropriate personnel are fully informed of the information provided in the notification letter.- Complete and return the acknowledgement form using instructions provided.- Retain a copy of The letter for your records.Intuitive representatives will schedule a site visit to provide the re-calibrated endoscope, update the software and ensure that it is properly integrated with the da Vinci Xi Surgical System.Questions or concerns may be directed to Intuitive Customer Service at 800-876-1310. |  | 20 | 20 units | Distributed in the states of FL, TX, and NY. | Terminated | July 02, 2014 | 21 | 39 | Computer
Z-1872-2014 | 68405 | No | HRX | Arthroscope. | Orthopedic | Orthopedic | 510(k) | Shaver Handpiece Sterilization Tray. Model number 272-700-000. For sterilization of Stryker Endoscopy arthroscopic shaver. |  |  | May 19, 2014 | 2014 | Stryker Endoscopy | STERRAD 100S parameters provided in the Shaver Handpiece reprocessing guide (P10299 Revision A) did not fully consider all worst-case scenarios during validation testing. | Nonconforming Material/Component | Device Correction letters were sent on May 23, 2014 to all users by certified mail. Internation sites will receive notification and be responsible for contacting customers. |  | 1992 | 1992 trays | Worldwide Distribution. | Terminated | May 07, 2015 | 353 | 1992 | Not_Computer
Z-1876-2014 | 64247 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ANSPACH***REF 98-0020***Modified Arthroscopic Knotless Suture Anchor Deployment Tool with Suture Puller***Rx Only*** |  |  | November 28, 2012 | 2012 | The Anspach Effort, Inc. | The Anspach Effort, Inc. in Palm Beach Gardens, FL is recalling their Anspach Custom Devices. The units were not manufactured consistently with Quality System requirements. | Process design | The Anspach Effort, Inc. sent an Urgent Medical Device Recall letter dated December 21, 2012 to all affected customers. The letter identified the affected products, problems and actions to be taken. Customers were instructed to remove and return the affected products and complete the attached reply form and return to The Anspach Effort Inc. For questions contact Anspach Product Support at (800) 327-6887.For questions regarding this recall call 561-494-3706. |  | 1 | 1 | Worldwide Distribution - USA including AL, CA, FL, IA, IL, KY, MN, MS, NC, OH, OK, PA, TN, TX.and Internationally to Canada, Japan, Hungary, Korea and Sweden.. | Terminated | August 07, 2014 | 617 | 788 | Not_Computer
Z-1928-2014 | 68469 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Alere Triage BNP Calibrators for the Beckman Coulter Access Family of Immunoassay Systems, Model No. 98202 |  |  | May 30, 2014 | 2014 | Alere San Diego, Inc. | Alere San Diego is recalling Alere Triage BNP Calibrators for the Beckman Coulter Access Family of Immunoassay Systems because the bottle labeled as Calibrator Level 5 actually contains Calibrator Level 4. This results in a calibration failure when attempting to calibrate using a calibrator kit which contains the affected Level 5 Calibrator. | Error in labeling | Alere sent an Urgent Medical Device Recall letter dated May 30, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The primary consignees were instructed to block all stock of the affected lot in their warehouses; discontinue use and return all stock on hand, complete the provided verification form to document receipt of the notice and quantity to be returned, contact their primary consignees and advise them of the requireddiscontinuation and return of the affected lot. For customers (or their sub-distributors) who have inventory on hand, Alere San Diego, Inc. will provide replacement product to the customer for the quantity of units that they must return.Customers are instructed to complete the attached reply form with the amount of unused product that will be returned for replacement. Customers with questions are instructed to contact Alere San Diego at 9975 Summers Ridge Rd, San Diego, CA 92121, (877) 308-8287, triage.support@alere.com |  | 3121 | 3,121 kits | Worldwide Distribution - Nationwide and the countries of Australia, Canada, China, EU, Mexico, Russia, Singapore, South Africa, South Korea, Taiwan and United Arab Emirates. | Terminated | February 11, 2015 | 257 | 3121 | Not_Computer
Z-1927-2014 | 68537 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Healthcare ADVIA Acid/Base Reagent 1 and 2 Kit containing ADVIA Centaur Acid Reagent bottles, lot number: 454778 |  |  | June 02, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Swollen acid bottles due to a trace metal contamination in the Acid Reagent bottles | Material/Component Contamination | Siemens Healthcare Diagnostics issued Urgent Medical Device Recall (UMDR) to customers in the United States on June 02, 2014 and a Urgent Field Safety Notice (UFSN) issued outside the United States on June 02, 2014.These notices inform the customer of the potential for swollen acid bottles due to confirmed metal contamination. Customers with alternative lots in inventory were instructed to transition to the alternative lot upon receipt of the communication. Customers without an alternative lot of acid in inventory were informed that Siemens proactively shipped an alternative lot to their laboratory and they should transition to that lot upon its receipt. Siemens also informed customers without an alternative lot that they may continue to run and report patient results using the affected lot of acid.Questions contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 3839 | 3,839 kits | Worldwide Distribution - US Nationwide and the countries of Canada, Argentina, China, Colombia, Korea, Malyasia, New Zealand, Mexico, Guatemala, and Singapore. | Terminated | February 05, 2015 | 248 | 3839 | Not_Computer
Z-1930-2014 | 67011 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | MAS Omni Immune PRO Controls |  |  | December 02, 2013 | 2013 | Microgenics Corporation | The constituent, Inhibin A, was removed from MAS Omni IMMUNE and MAS Omni IMMUNE PRO controls; however the package insert was not updated to reflect the change. | Labeling Change Control | Urgent Medical Device Correction letters were sent on December 2, 2013 to all affected customers and distributors via trackable mail. |  | 1393 | 1393 | Worldwide Distribution-USA (nationwide) and the countries of Australia, China, Germany, Hong Kong and Malaysia. | Terminated | July 08, 2014 | 218 | 4578 | Not_Computer
Z-1874-2014 | 68447 | No | IPF | Powered muscle stimulator. | Physical Medicine | Physical Medicine | 510(k) | IF 8000; 1/F Interferential Electrical Stimulation Device; NMES (neuromuscular electrical stimulation) mode. |  |  | January 25, 2009 | 2009 | Zynex Medical, Inc. | Zynex notified customers that a malfunction in the IF 8000 can result in the device ceasing to operate, and in some instances,overheating and becoming hot to the touch. | Component design/selection | Consignees were notified via letter on 6/25/2009. The letter identified the reason for the recall, and stated that customers could contact customer service at 1-866-940-7030 if they encountered any difficulties with the affected product. Otherwise, they could continue using the product as they have in the past. |  | 1461 | 1,461 | Nationwide Distribution | Terminated | July 02, 2015 | 2349 | 1461 | Not_Computer
Z-1926-2014 | 66364 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension LOCI TNI Troponin I Calibrator RC 621. |  |  | September 17, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | Siemens Dimension LOCI TNI Troponin I calibrator lot 3DD007 may demonstrate an average upward shift of 24% in QC and patient results following calibration. | Process control | Siemens sent an Urgent Medical Device Recall letter dated September 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised immediate discontinuation of use and discard of the affected device.Please complete the attached Effectiveness Check form and fax it to the Siemens Customer Care Center at (302) 631-8467 to indicate that you have received this information. Also, please indicate any replacement product required.Please contact Siemens Customer Care Center- Technical Solutions at (800) 441 9250 if you are aware of any adverse event that occurred as a consequence of this issue, or for immediate product needs.Please retain this letter with your laboratory records, and forward this Jetter to those who may have received this product.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 1625 | 1625 | Distribution: US Nationwide including Puerto Rico | Terminated | August 27, 2014 | 344 | 1625 | Not_Computer
Z-1925-2014 | 68418 | No | JDC | Elbow joint metal/polymer constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | DISCOVERY ELBOW prosthesis, 4 MM X 100 MM RIGHT HUMERAL FLANGED I BOND COAT Tl 6AL 4V ALLOY. |  |  | May 15, 2014 | 2014 | Biomet, Inc. | Surface finish is different than specified. The implant has 30 grit blast on the exterior of the flange when the print calls for the glass bead blast. The patient may experience pain requiring surgical intervention. | Nonconforming Material/Component | Biomet sent an Urgent Medical Device Recall Notice letter dated May 28, 2014 to their customer. The letter identified the affected product, problem and actions to be taken. The letter instructed physician to monitor the patient for any pain that may require surgical intervention. For questions call (574) 372-1570. |  | 1 | 1 | US Nationwide Distribution in the state of California | Terminated | March 16, 2015 | 305 | 1 | Not_Computer
Z-1942-2014 | 68482 | No | HSX | Knee joint femorotibial metal/polymer non-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stride Femoral Component Size 2, PFSI-00053, Right Lateral /Left Medial. Unicondylar Knee prosthesis. |  |  | June 04, 2014 | 2014 | Blue Belt Technologies MN | Mislabeled. Packages labeled PFSI-00053, Right Lateral/Left Medial contains PFSI-00054, Right Medial/Left lateral implants. Packages labeled PFSI-00054, Stride Right Medial/Left Lateral contain PFSI-00053, Left Medial/Right Lateral. | Labeling mix-ups | Firm initiated recall on 6/4/2014 by Blue Belt Technologies representatives hand delivering an Urgent Medeical Device Recall letter. The letter was addressed to the attention of Safety Officer/Regulatory Compliance Officer. The letter described the issue of the labeling mix-up, asked for devices to be quarantined, and stated that a Blue Belt Technologies Clinical Specialist will be contacting the consignee or return and replacement. A Receipt Acknowledgement form was asked to be returned by fax or mail. |  | 2 | 2 | US Nationwide Distribution in the states of CA, PA, and TX | Terminated | August 04, 2014 | 61 | 2 | Not_Computer
Z-1934-2014 | 68457 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen Complete Knee Solution, MIS Total Knee Procedure, Stemmed Tibial Component, Precoat, Fixed Bearing, Size 4. Part of the NexGen systems of semi-constrained, non-linked condylar knee prosthesis. |  |  | May 19, 2014 | 2014 | Zimmer, Inc. | Zimmer is initiating a voluntary recall of specific identified lots due to the potential that the threads may be out of specification. | Equipment maintenance | Zimmer sent an Urgent Medical Device Recall letter June 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the notification and ensure affected personnel are aware of the contents. Locate and quarantine any affected product and notify their Zimmer sales representative. Their Zimmer sales representative will remove the recalled product from their facility. For patients that previously had this product implanted, it is recommended that they continue their normal operative follow up routine. Customers with questions were instructed to call 1-877-946-2761. |  | 41180 | 41,180 total | Worldwide Distribution - USA (nationwide) and Internationally to Australia, China, Hong Kong, India, Japan, Korea, Malaysia, Nicaragua, Singapore, and Taiwan. | Terminated | October 29, 2015 | 528 | 41180 | Not_Computer
Z-1920-2014 | 68503 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Medtronic Duet External Drainage and Monitoring System, Interlink Injection Sites, Catalog No. 46913. |  |  | June 09, 2014 | 2014 | Medtronic Neurosurgery | Medtronic Neurosurgery is recalling the Medtronic Duet External Drainage and Monitoring Systems because the patient line tubing in the Duet External Drainage and Monitoring Systems may become disconnected from the patient line stopcock. | Device Design | An Urgent Medical Device Recall letter, dated 6/9/14, was sent to all their customers who purchased the Medtronic Duet External Drainage and Monitoring Systems. The letter informs the customers of the problem identified and the actions to be taken. Customers are instructed to cease use of the affected product, return any used units, account for any product used and verify that connections are in the condition as stated in the Instructions for Use for product currently in use. Additionally, customers are to complete the Customer Product Accountability Form. Customers with questions about the recall letter are instructed to contact Bob Shokoohi at (805) 571-8725. |  | 3026 | 3,026 units | Worldwide Distribution -- USA, Australia, Belgium, Canada, Croatia, Cyprus, Denmark, Egypt, Finland, France, Greece, Iran, Israel, Italy, Kuwait, Netherlands, Norway, Poland, Puerto Rico, Romania, Russian Federation, Saudi Arabia, Singapore, South Africa, Switzerland, Turkey, and United Kingdom. | Terminated | January 14, 2015 | 219 | 10752 | Not_Computer
Z-1941-2014 | 68396 | No | OMN | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | illumipro-10 Incubator/Reader, Meridian Bioscience, Inc. |  |  | April 22, 2014 | 2014 | Meridian Bioscience Inc | Illumipro-10 instruments may have Block B chambers operating at an incorrect temperature as a result of incorrect temperature calibration. | Employee error | Meridian sent an Urgent Medical Device Recall letter dated April 25, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately return their unit to Meridian Bioscience. Meridian's Technical Support Staff will call them directly to provide them with a return authorization number and instructions for decontaminating, packaging and shipping the unit. For additional information customers were instructed to call 800-343-3858 (North America or 513-271-3700 (Worldwide). |  | 35 | 35 instruments | Worldwide Distribution - USA including AL, AZ, CA, CO, FL, IA, IN, LA, MA, ME, MI, MN, NJ, NY, OH, PA, and WV. Internationally to Australia. | Terminated | June 30, 2014 | 69 | 35 | Not_Computer
Z-1947-2014 | 68335 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Alcon CUSTOM-PAK containing Anterior Chamber Cannula 27G Component Part Number 585006 contained in Alcon Assembly Numbers 9795-37, 11233-14, and 13334-04. |  |  | May 16, 2014 | 2014 | Alcon Research, Ltd. | Alcon CUSTOM-PAK which is supposed to contain a 27G Anterior Chamber Cannula (27G x 7/8 in) actually contains a 27G Sharp Tip Needle | Employee error | Alcon sent an Urgert Market Withdrawal letter dated May 14, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory to determine if they have any affected product, place the affected product into quarantine to ensure it is not used, return the attached Response Form via fax to 817-302-4337 or email market.action@Alcon.com.Customers with questions were instructed to call 1-800-862-5266. |  | 684 | 684 total (525 units in US) | Nationwide Distribution including North Carolina and Maine | Terminated | December 09, 2014 | 207 | 684 | Not_Computer
Z-1945-2014 | 68640 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter, MiniCap with Povidone-Iodine solution, 5C4466P. For dialysis. |  |  | July 08, 2013 | 2013 | Baxter Healthcare Corp. | the white printed paper layer on the pouch packagingmay not be fully adhered to the pouch foil layer. The upper paper layer may serve as a protective barrier and help support overall packaging. | Packaging change control | An Urgent Product Recall notification was mailed via U.S.P.S. first-class mail to facilities, (dialysis centers and distributors) on July 8, 2013 and mailed by first-class mail to home patients on July 11, 2013. |  | 20940 | 20,940 units | Distributed in the states of AR, TN MO, KS, TX, MS, and AL. | Terminated | June 30, 2014 | 357 | 20940 | Not_Computer
Z-1835-2014 | 67713 | No | IXY | Wall-mounted radiographic cassette holder. | Radiology | Radiology | 510(K) Exempt | The Claymount Harmony is a grid cabinet designed to accept a digital x-ray imaging receptor and stationary grid. It contains signaling electronics to indicate the presence of the grid and detector as well as detector orientation. It may be mounted to a wall stand or incorporated in a radiographic table. |  |  | March 15, 2014 | 2014 | Claymount Assembli | Harmony1 models of the Harmony Bucky is incorrectly labeled in that the label does not list the manufacturer address, the date of manufacture, or a statement of compliance with 21 CFR Subchapter J. | Radiation Control for Health and Safety Act | CDRH approves the CAP subject to the following conditions for Claymount Assemblies BV:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. All forward production of Harmony devices must be properly tested, labeled, and reported when imported into the United States. For further questions please call (630) 271-9729. |  | 125 | 125 units shipped to United States | Worldwide Distribution - US including the states of CA, CO, FL, IL, NC, NJ, OH and PA and the countries of Canada, France, Italy, Netherlands, South Africa, Switzerland, Turkey and United Kingdom. | Open, Classified |  |  | 125 | Not_Computer
Z-1944-2014 | 68394 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | syngo Lab Data Manager, SMN 10800057, 10803189, and 10804573. |  |  | May 20, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Certain configuration parameters may cause the reference ranges to not hold results as expected and release them to the Laboratory Information System (LIS) that should have been held for manual review. | Device Design | An urgent medical device correction notice, dated May 20, 2014, was sent to end users to provide instructions to avoid potential risks associated with the software defect, until changes can be made by a Siemens technical representative. |  | 336 | 336 | Worldwide Distribution. US Nationwide, Canada, Australia, Belgium, Czech Republic, Germany, Denmark, Finland, France, United Kingdom, Italy, Netherlands, Portugal, Slovakia, Singapore, Japan, and Spain. | Terminated | August 10, 2015 | 447 | 336 | Computer
Z-1946-2014 | 68592 | No | MQK | Assisted reproduction labware. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Name: Pasteur Pipets, used by the embryologist in the IVF Laboratory. |  |  | June 20, 2014 | 2014 | ORIGIO, Inc | Identification of packaging breach of the nylon pouch sterilebarrier for the 3-pack configuration | Package design/selection | Field Safety Corrective Action / Urgent Product Recall letter dated June, 2014. |  | 134086680 | Total number of Boxes Distributed 32,760. US 4,093 / Canada 1,156 / EU &amp; Other 27,511 | Worldwide Distribution: US Nationwide, Canada, Australia, Barbados, Chile, China, Colombia, Denmark, Ecuador, France, Germany, Great Britain, Greece, India, Italy, Japan, Jordan, Malaysia, Netherlands, Peru, Russia, Singapore, South Africa, Spain, Taiwan, Turkey, and United Arab Emirates. | Terminated | July 20, 2015 | 395 | 134086680 | Not_Computer
Z-1836-2014 | 67715 | No | IXY | Wall-mounted radiographic cassette holder. | Radiology | Radiology | 510(K) Exempt | The Claymount Harmony is a grid cabinet designed to accept a digital x-ray imaging receptor and stationary grid. It contains signaling electronics to indicate the presence of the grid and detector as well as detector orientation. It may be mounted to a wall stand or incorporated in a radiographic table. It is intended to be used in general radiographic applications as a component of a radiographic system. |  |  | March 15, 2014 | 2014 | Claymount Assembli | Claymount Assemblies BV is recalling the Harmony2 model of Harmony Bucky because 1) the devices are incorrectly labeled in that the label does not list the manufacturer address, the date of manufacture, or a statement of compliance with 21 CFR Subchapter J, and 2) the device's insulation values do not comply with IEC 60601-1 in all cases. | Radiation Control for Health and Safety Act | MEDICAL DEVICE RECALL letters dated March 15, 2014 were sent to all direct accounts. The letters includes instructions to: 1) complete and return the Acknowledgement and Receipt Form accompanying the letter; 2) replace the Harmony2 Bucky(s) with new Harmony3 Bucky(s) that the direct account will be receiving in the near future; 3) if the direct account is a distributor and/or installer and it elects to have their customer(s) perform the replacement, the customer(s) should be provided with a copy of the letter, the replacement Harmony3 Bucky(s), and the replacement notification card, or the direct account can provide the names and addresses of their customers in which case Claymount will contact the customers directly; and, 4) once each unit is replaced, the replacement completion notification card should be mailed or faxed to the Claymount USA office and the Harmony2 Bucky(s) should be returned to "Claymount Americas Corp., ATTN: Alice Packard, 2545 Curtiss St., Downers Grove, IL 60515-4059" or faxed to 630-271-9995.For questions customers were instructed to call the following numbers:USA (630) 271-9729EUR +31 315 659 150 |  | 12 | 12 devices | Worldwide Distribution - USA including NC, OH, CO, CA, NJ, IL, OH, PA, FL and Internationally to Canada, China, France, Italy, Netherlands, South Africa, Switzerland, Turkey, and United Kingdom. | Open, Classified |  |  | 12 | Not_Computer
Z-1961-2014 | 68498 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | MIS Rod 2, a component of the NewPort System. |  |  | June 09, 2014 | 2014 | Integra LifeSciences Corp. | Integra Lifesciences is recalling NewPort System MIS Rod, 2 because it may be up to 30 degrees offset from specification requirements causing misalignment with the inserter. This could potentially result in the lordosis of the rod not aligning with the handle when the inserter is used in accordance with the current NewPort MIS System Surgical Technique. | Under Investigation by firm | A customer notification letter was sent on 6/9/14 to their customers and distributors who purchased the Integra NewPort System MIS Rod, 2 to inform them there may be a potential that the hex on some NewPort System MIS Rod, 2 are up to 30 degrees offset from specification requirements causing misalignment with the inserter. Customers are instructed to complete and return the customer response form and return by email or fax as indicated on the form. Customers with questions are instructed to contact Customer Service at (866) 942-8698. |  | 94 | 94 units | Worldwide Distribution. US including states of CA, FL, IL, MD, MO, TN, TX, and countries of Australia, Czech Republic and Spain. | Terminated | November 06, 2014 | 150 | 94 | Not_Computer
Z-1959-2014 | 68398 | Yes | CAC | Autotransfusion apparatus. | Anesthesiology | Anesthesiology | 510(k) | Atrium Express Dry Seal Chest Drain ATS Blood Recovery, Catalog Number: 4050-100N. |  |  | April 28, 2014 | 2014 | Atrium Medical Corporation | Product that was previously recalled because the Chest Drain tubing of the ATS Blood recovery may leak or disconnect, was released for distribution in error. | Nonconforming Material/Component | Smith &amp; Nephew notified Distributors on April 28, 2014 by telephone. A follow up e-mail and a recall notification letter were issued by UPS on 6/27/14. Distributors were requested to notify customers who received the lot. A reply acknowledgement form was to be completed and returned. Return Goods Authorizations for the return of the goods were issued to customers. |  | 195 | 195 Cases; Total 1170 units | Distributed in the states of AK, CA, MD, MN, MO, NC, NJ, OH, and UT. | Terminated | March 17, 2016 | 689 | 195 | Not_Computer
Z-1963-2014 | 68506 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | EXACTAMIX 1200 Compounder, model numbers 1200-DY, 1200-DX, 1200-DYR and 1200-DXR; EM1200 DY Display, Replacement EM 1200 DY Display. EXACTAMIX 2400 Compounder model numbers 2400-DY, 2400-DX, 2400-DYR and 2400-DYX. Baxa Corporation. An automated pumping system that compounds multiple sterile ingredients into a finished solution in a single patient IV administration bag. |  |  | June 10, 2014 | 2014 | Baxter Healthcare Corp. | If the universal ingredient (UI) in an active configuration is changed using the Configuration Editor, a flush of the outlet pump tube will not be initiated by the software. It could result in the original UI remaining in the tube which could be delivered into the next bag. | Software design | An Urgent Device Correction communication via 1st class mail was mailed on July 14, 2014, informing customers of a potential issue if the universal ingredient in an active configuration is changed using the Configuration Editor rather than following the normally prescribed and supported methods as outlined in the Operator Manual. |  | 1758 | 1758 | US Nationwide | Open, Classified |  |  | 1758 | Computer
Z-1958-2014 | 68347 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | Simplexa" Flu A/B &amp; RSV Direct, Model Number: MOL2650. |  |  | May 15, 2014 | 2014 | Focus Diagnostics Inc | Focus Diagnostics is providing an urgent safety notice for a correction to the labeling for Simplexa" Flu A/B &amp; RSV Direct (MOL2650). Focus Diagnostics received some customer complaints of Simplexa Flu A/B &amp; RSV Direct assays with sporadic false signals, which may result in a higher false results rate due to a potential software spectral matrix and Direct Amplification Disc storage. | Device Design | Focus Diagnostics sent an Urgent Safety Notice Correction letter dated May 15, 2014, to all affected customers. The notification, informed customers of the product, the problem, and the action to be taken by the customer. To decrease the likelihood of obtaining any false results caused by either manually entering a spectral matrix or storing the DAD improperly after each use, please follow the recommended warnings and precautions provided in the revised and attached Simplexa Flu A/B &amp; RSV Direct (PI.MOL2650 Rev. D) package insert.1. The spectral matrix must be installed in each 3M Integrated Cycler and should not be changed unless an updated QR code for the instrument is provided by Focus Diagnostics. The spectral matrix is unique to each 3M Integrated Cycler. The spectral matrix was provided with the 3M Integrated Cycler instrument on the cover of the 3M Integrated Cycler Hardware Manual. If the matrix label will not scan or cannot be found, contact Focus Diagnostics. The contact information is on the last page of this document. 2. Not installing or changing the spectral matrix can result in false results.3. After each use store DAD discs flat with the numbered foil side up.Customers were instructed to acknowledge that they have received this notification by signing the enclosed acknowledgement form and email the form to DxTS@focusdx.com or fax back to Focus Diagnostics Technical Services at 562-240-6526 within 10 business days. The Notice of Correction and Acknowledgement Form was sent to customers on 05/15/2014 via email and USPS mail. Additional customers were notified on 05/29/2014 via email and USPS mail.For questions regarding this recall call 714-220-1900. |  | 614 | 614 kits | Nationwide Distribution | Terminated | December 31, 2014 | 230 | 614 | Computer
Z-1949-2014 | 68393 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | DeVilbiss Model DV5x Series CPAP units. For use in treating Obstructive Sleep Apnea (OSA) by means of application of positive air pressure. Used in home and clinical environments. |  |  | May 13, 2014 | 2014 | DeVilbiss Healthcare LLC | A limited number of DV5x Series CPAP units were manufactured and shipped without the serial number programmed into the unit, causing the unit serial number on LCD display to appear as "RUN-TEST". | Software Manufacturing/Software Deployment | DeVilbiss Healthcare sent an Urgent: Medical Device Correction notice dated May 13, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. What to do if you have an affected DV5x Series CPAP unit:1. Verify your inventory of DV5x Series CPAPs and isolate any units that are on the included serial number list, these units will need to have the serial number corrected.2. Please contact DeVilbiss Healthcare Customer Service to either request a Sales Representative to correct the issue on site or return the unit back to DeVilbiss to be corrected.3. If you have already distributed any DV5x Series CPAPs included in this serial number list to end users, a patient communication has been included with this notification - a. Insert your contact information at the bottom of the "User Medical Device Correction Letter" included in this notice.b. Forward a copy of the "User Medical Device Correction Letter'' to each user towhom you have provided an affected DV5x Series CPAP unit.If you have one or more of the affected DV5x Series CPAP unit in your possession (as identified in the serial number list) and have any questions regarding the information provided in this communication, please contact DeVilbiss Healthcare Customer Service at 1-800- 338-1988 from 8:00AM-5:00PM EST. [Note: If you have already spoken with a DeVilbiss Healthcare Representative about this issue, there is no need to call unless you have further questions.] |  | 1809 | 1809 | Worldwide Distribution - USA (nationwide) and the countries of Colombia, Honduras, Iran, Thailand and Uruguay. | Terminated | December 02, 2014 | 203 | 1809 | Not_Computer
Z-1940-2014 | 68333 | freezer | JRM | General purpose laboratory equipment labeled or promoted for a specific medical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Stirling Ultracold Portable Shuttle, ULT-25. |  |  | March 20, 2014 | 2014 | Global Cooling Inc | The external power supply input is overheating at the connection to the freezer. | Use error | Consignees received an Urgent Medical Device Recall letter, dated March 20, 2014, a Recall Response letter, a User Alert, and a label to attach to the freezer power supply input. The letter identified the reason for the recall and how to recognize that the device may fail. Customers are to discontinue use of the device if there are any signs of a loose connection, corrosion, discoloration, or melting; and take the necessary actions identified, including contacting Global Cooling Customer Service at service@stirlingultracold.com. The Recall Return Response was to be completed and returned. |  | 413 | 413 units | Worldwide Distribution -- USA, including the states of AK, AR, AZ, CA, CO, CT, FL, GA, GU, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MN, MO, MT, NC, NE, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, and WA; and, the countries of Australia, Brazil, Canada, China, India, Indonesia, Japan, Korea, Kuwait, Mexico, Netherlands, Pakistan, Singapore, Switzerland, Taiwan, Thailand, United Arab Emirates, and United Kingdom. | Terminated | July 01, 2015 | 468 | 413 | Not_Computer
Z-1956-2014 | 68564 | No | OVN | Absorbable surgical gut suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | McKesson Medi-Pak, Laceration Trays with Cotton O.R. Towels, Sterile, Single Use, Latex Free, 20 Each per Case, Rx only, Distributed by: McKesson Medical Surgical, Inc., Richmond, VA 23228 |  |  | June 12, 2014 | 2014 | Medical Action Industries Inc | Some of the sterile kits may have incomplete seals which might compromise the sterility of the kits. | Nonconforming Material/Component | Consignees were notified via E-Mails, Certified Letters and Telephone on 06/12/2014. |  | 474 | 474 cs (9,480 ea) total for all kits involved | Distributed in the states of AR, AZ, FL, GA, IA, IL, KY, MA, MN, NC, NY, OH, TN, TX, and VA. | Terminated | February 03, 2015 | 236 | 474 | Not_Computer
Z-1962-2014 | 68529 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON S Family Ultrasound Systems at software versions VC25D, VC30A and VC30B; (S2000, S3000 and S1000 systems) Siemens Medical Solutions USA, Inc. Ultrasound imaging systems. |  |  | June 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | There is a potential measurement error with the Auto-Stats (auto statistics) measurement tool on the ACUSON S Family ultrasound system. | Software design | Siemens sent a Customer Safety Advisory Notification letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following: To avoid the potential risk of this issue, do not use the Auto-Stats tool; use alternate measurement tools. Only use the Velocity tool on frozen images.Customer Service Engineer will contact the customer to schedule an appointment to replace their existing version of software, free of charge. Customers with questions should contact their local service support person forinformation regarding timelines and status.Customers were instructed to share this information with all personnel withintheir organization who need to be aware of this issue.For questions regarding this recall call 650-694-5135. |  | 790 | 790 units | Worldwide Distribution - USA including CA, MO, VA, PA, WV, and Internationally to Angola, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Brazil, Canada, Chile, Colombia, Czech Republic, Denmark, Dominican Rep., Ecuador, Egypt, France, Germany, Greece, India, Iraq, Israel, Italy, Japan, Kazakhstan, Malaysia, Mexico, Morocco, Netherlands, Norway, P.R. China, Philippines, Poland, Portugal, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, U.A.E., Ukraine, United Kingdom. | Open, Classified |  |  | 790 | Computer
Z-1966-2014 | 68472 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Vision Troubleshooting Guide for the da Vinci Standard Surgical System; da Vinci S System and da Vinci Si System. |  |  | May 21, 2014 | 2014 | Intuitive Surgical, Inc. | Multiple updates to User Manuals, Instructions for Use (IFU), Quick Reference Guides (QRG), and Instrument Release Kits (IRK) for use with the da Vinci Surgical System, instruments and accessories. Removal of the Emergency Grip Release Wrench, part number 710142. | Device Design | Intuitive Surgical sent a Field Safety Notice dated June 16, 2014, to all consignees affected by changes in labeling, IFU User manuals and QRG's. Consignees were instructed to take the following actions:Ensure all appropriate personnel are fully informed of this Notification. Forward this letter to your Risk Manager, OR Director, Purchasing Manager, Biomedical Engineering staff and members of your medical staff who perform da Vinci surgery procedures.2. Use Attachment C to identify the labeling that has been updated.3. No manuals or products should be returned. Discard old versions of the User Manuals, IFUs, and QRGs for which you have received updated versions.4. Discard all versions of the Emergency Grip Release Wrench (Part Number 710142). Order additional IRKs from Intuitive Surgical as needed.5. Complete and return the attached Acknowledgement Form to acknowledge that you have received the updated Labeling, IRKs, in-service and have discarded all previous versions of Labeling and the emergency grip release wrench. Please return the completed Acknowledgment Form within two days of the in-service.6. Retain a copy of this Notice and the Acknowledgement Form for your records.If you need further information or support concerning this recall, please contact your Intuitive Surgical Representative or Intuitive Surgical Customer Service at the numbers listed below:North America: 800-876-1310 Option 3 (6 am to 5pm PST)For questions regarding this recall call 408-523-2602. |  | 2014 | All distributed prior to June 2014 | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Great Britain, Greece, Guadeloupe, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Malaysia, Mexico, Monaco, Netherlands, New Zealand, Norway, Pakistan, Philippines, Indonesia, Portugal, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovak Republic, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, Uruguay, Venezuela, and Vietnam. | Terminated | May 05, 2015 | 349 | 2014 | Not_Computer
Z-1977-2014 | 68490 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACS Canto and associated Fluidics Cart |  |  | June 06, 2014 | 2014 | BD Biosciences, Systems &amp; Reagents | Laser cooling fans and internal fan of the associated fluidics cart are not operational. | Nonconforming Material/Component | BD sent an Important Product Recall notification letters dated June 2014 to affected US customers by trackable mail. International affiliates were notified on June 6, 2014. The letter identified the affected product, problem and the actions to be taken. A Field Change notification has been established and a BD field Service Engineer will coordinate with each account to perform the necessary corrections. Questions or concerns may be directed to BD Customer Support at BD Biosciences, prompt 2. Out of US customers are to contact local BD rep with questions. |  | 35 | 35 installed devices | Worldwide Distribution - US Nationwide in the states of CA, OH, RI, GA, FL, MD, NY, WV, NJ, Ma, CT, IA, NM, MN and the countries of China and Belgium. | Terminated | January 27, 2015 | 235 | 35 | Not_Computer
Z-1976-2014 | 68519 | No | HPP | Operating headlamp. | Ophthalmic | Ophthalmic | 510(K) Exempt | IntegraÂż LED Battery Charger, Single Bay. |  |  | June 11, 2014 | 2014 | Integra LifeSciences Corp. | LED battery chargers may prematurely fail and will not charge the LED battery as intended | Manufacturing material removal | Integra initiated a voluntary recall on June 11, 2014 to consignees that have been shipped affected product lots of IntegraÂż LED Battery Chargers, Catalogue No. 90523.The written notifications were delivered to US consignees by traceable courier service and non-US consignees by traceable email or facsimile. The consignee notification: Advises them of the nature of the issue and the potential for an adverse patientconsequence. Advises them to review their inventory and how to determine if they have affected battery chargers. Advises them to immediately stop the use or distribution of any affected product lot numbers they may have. Requests they identify and report to Integra if they do or do not have any unexpired affected product lots.Advises them replacements battery chargers will be shipped for any affected products they affirm are in their possession and how to return their affectedproduct. Questions regarding these instructions, please contact Integra Customer Service at 1-800-431-1123 or your Integra sales |  | 1250 | 1250 battery chargers | Worldwide Distribution - US Nationwide and the countries of Canada, Australia, Arab Emirates, Denmark, Germany, Great Britain, India, Ireland, Israel, Japan, Kuwait, Lebanon, Malta., Netherlands, Panama, Poland, Saudi Arabia, Singapore, South Africa, Spain, Sweden, and Tunisia. | Terminated | December 19, 2014 | 191 | 1250 | Not_Computer
Z-1978-2014 | 68500 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Orthopaedics Triathlon Tritanium Patella Inserter, . NON-STERILE |  |  | May 01, 2014 | 2014 | Stryker Howmedica Osteonics Corp. | Tritanium Patella Inserter Instrument fracture during implantation. . | Process control | Stryker Orthopaedics issued an "Urgent Medical Device Recall Notification" letters/Notification Acknowledgement forms dated May 7, 2014 to their customers via Fed-Ex on May 7, 2014. The notification informs the customers of the issue(s) with the affected product; potential hazards; and how to identify and handle any affected product. Customers were instructed to return any affected product to Stryker Orthopaedics and fill out and return the enclosed Notification Acknowledgement form. For questions call (1-201-972-2100). |  | 22 | 22 Units | US Nationwide Distribution - in the states of Arizona, Florida, Kentucky, Louisiana and New Jersey. | Terminated | June 09, 2017 | 1135 | 22 | Not_Computer
Z-1982-2014 | 68644 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView XCT: product code 882482, BrightView X- upgraded with the XCT Flat Panel Detector (FPD): product code 882454. Gamma camera for Single Photon Emission Computed Tomography (SPECT) Medical Device for imaging. |  |  | June 23, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Flat Panel Display (FPD) failed to remain securely locked in the deployed position. | Device Design | Philips sent an Field Safety Notice dated June 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This Field Safety Notice (FSN) 88200490 and Addendum is intended to inform you about the following:" what the problem is and under what circumstances it can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct the problem.If you need any further information or support concerning this issue, please contact your local Philips Healthcare representative:For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377: follow the prompts).thcare is informing the consignee of this issue through a Field Safety Notice 88200490. The US consignees will receive this notice via certified mail and Philips Healthcare Key Market representatives will distribute the FSN internationally. Additionally, a Field Service Engineer will contact each site to schedule a time to conduct the appropriate field safety correction through Field Change Order (FCO) 88200490 that will address the above identified issue. |  | 407 | 407 | Worldwide Distribution - USA including the states of : AZ, CA, CO, DC, FL, GA, IL, MA, MD, MN, MO, MS, MT, NC, NJ, NY,OH, OK, OR, PA, TX, VT, WA,WI. Government Facilities: AZ, CA, CO, FL, HI, IA, KY, MA, MN, MO, NC, ND, NE, NV, OH, OK, SD, TX, and WV. and the countries of : Algeria, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Czech Republic, Denmark, Finland, France, Germany, Greece, Indonesia, Ireland, Israel, Italy, Japan, Lebanon, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Poland, Portugal, Russia, Saudi Arabia, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, UAE and UK. | Terminated | September 24, 2015 | 458 | 407 | Not_Computer
Z-1991-2014 | 68520 | No | MMI | Creatine phosphokinase/creatine kinase or isoenzymes test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA CentaurÂż Systems TnI-Ultra assay; Catalog Number 02789602 (100 tests), SMN 10317708 and 02790309 (500 tests), SMN 10317790 |  |  | May 30, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | The solid phase reagent in some of the ADVIA CentaurÂż TnI -Ultra ReadyPacksÂż, lot 084, is darker in appearance, creating a potential for incorrect control and patient results. | Nonconforming Material/Component | An Urgent medical Device Correction (UMDC) was mailed, via FedEx, to all affected Siemens Healthcare Diagnostics customers in the United States on May 30, 2014 for customer receipt on June 2, 2014. All international customers were provided with a copy of the Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each country's local regulations and procedures. Each customer is asked to complete a fax-back form. This notice informed the customers of the potential for incorrect control and patient results due to solid phase which is darker in appearance. The notice instructs customers to place one TnI-Ultra ReadyPack on the system at a time and to calibrate and run controls on each ReadyPack to ensure accurate control and patient values. Patient TnI-Ultra results produced with this assay are acceptable if they follow acceptable calibration with valid quality control results. Upon receipt of kit lots ending in 086 and higher, calibration of each ReadyPack is not required. This notice should be reviewed with the facility's laboratory or medical director. |  | 34283 | 34,283 kits | Worldwide Distribution-USA (nationwide) including PR and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY, and the countries of Canada, Mexico, Argentina, Austria, Australia. Bahrain, Belgium, Brazil, Bulgaria, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, French Polynesia, Great Britain, Greece, Guadeloupe, Hong Kong, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lebanon, Libya, Macedonia, Malaysia, ,Malta, Martinique, Netherlands, New Zealand, Norway, Pakistan, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Republic of Korea, Romania, Russian Federation, Serbia, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, Sri Lanka, South Africa, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, Uruguay, and Vietnam. | Terminated | February 26, 2016 | 637 | 34283 | Not_Computer
Z-1984-2014 | 68383 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes B37 Replacement Screws. M3.5 Screw for insertion handle guide/aiming arm, part numbers 3.113.025 and 03.113.026. Surgical instrument for use with Low Bend Medial Distal Tibia Plate. Orthopedic. |  |  | May 15, 2014 | 2014 | Synthes, Inc. | All lots of B37 Replacement Screws are being recalled due to potential mislabeling where the package may not contain the screws identified on the label. | Labeling design | An urgent medical device recall notice, dated May 29, 2014, was sent to end users and sales consultants to advise them of the issue and provide instructions for returning affected devices. |  | 44 | 44 | Distributed US nationwide, in DC and the states of KY, PA, IL, WA, AZ, CA, OK, FL, and KS. | Terminated | December 18, 2015 | 582 | 44 | Not_Computer
Z-1990-2014 | 68524 | No | LRG | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan WalkAway-40 plus Instrument and MicroScan Walkaway-96 plus instrument access door hinge. |  |  | June 02, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Springs contained in the access door hinge assembly on the Walk Away Plus instruments are becoming detached causing premature failure of the hinge assembly and a potential safety issue for the customers. | Component design/selection | Urgent Medical Device Recall notification letter was sent to all affected US customers on June 2nd by Federal Express.OUS notifications were also sent to regional Siemens organizations who distributed in-country notifications and relevant notifications. |  | 1027 | WA40: 560 US, 131 OUS, 9 already repaired; WA96: 81 US, 189 OUS, 57 repaired. | Worldwide Distribution-USA (nationwide) and the countries of Algeria Argentina Brazil Canada China Chile Columbia France Indonesia Libya Georgia Germany India Italy Japan Kazakhstan Lithuania Oman Pakistan Peru Poland Portugal Republic of Korea Romania Russia Saudi Arabia Spain Turkey. | Terminated | December 23, 2015 | 569 | 1027 | Not_Computer
Z-1987-2014 | 68495 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Microline Surgical ReNew Laproscopic Fenestrated Grasper Forceps Tips (Reusable) |  |  | May 12, 2014 | 2014 | Microline Surgical, Inc. | Potential for the Fenestrated Grasper Tip jaw to break when force is applied to the jaw | Nonconforming Material/Component | Microline Surgical notified Distributors, Hospitals (Customers/End Users) and Suppliers via letter dated May 15, 2014 with labeling, pictures, and Recall Response Forms via UPS Overnight Air, email, and/or fax. Product is requested to be returned and a RMA number will be issued. Contact 978-867-1775 for questions. |  | 6290 | 6290 units | Worldwide Distribution-USA including the states of CA, HI, ID, MA, MI, MS, PA, TN, UT, and TX. and the countries of Belgium, France, Ireland, Germany, Indonesia, Korea, Norway, Saudi Arabia, Switzerland, Turkey, and UK. | Terminated | August 13, 2015 | 458 | 6440 | Not_Computer
Z-1985-2014 | 68382 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Siemens Mobilett XP Digital Mobile X-ray System |  |  | May 09, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential fire hazard of certain Lithium-Ion batteries in the computers used in the original manufacturing or as spare parts between October 2010 and April 2011 for the Siemens Mobilett XP Digital mobile X-ray System, which might occur at any time with higher probability while the system is being charged and connected to the main power supply. | Component design/selection | Siemens sent a Safety Advisory Notice letter dated May 9, 2014, to affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated that a Siemens service technician would be visiting their site to identify and service affected devices. |  | 10 | 10 | US Nationwide Distributions in the states of MO, GA, FL, NJ, CA, PA and including NM. | Terminated | December 15, 2014 | 220 | 10 | Not_Computer
Z-1979-2014 | 68459 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | CNS-6200 Series Central Nurse Station and accessories. Model Number: CNS-6201 |  |  | June 03, 2014 | 2014 | Nihon Kohden America Inc | Software Version 02-26, when used with the Central Nursing Station (CNS) 6201, (PU-621 RA) may unexpectedly and without warning reboot, resulting in a period of approximately 3 minutes during which time the system does not display patient data and does not emit alarms based on that data should they be appropriate. | Software design | Nihon Kohden sent an Urgent Medical Device Safety Notification and Field Correction letter dated June 3, 2014 to affected customers. Recall notification letter, dated June 3, 2014 was hand-delivered to site by Nihon personnel. The letter provided the following information: product description with codes and software version, reason for recall, field correction information and actions to be taken, MedWatch reporting information, and contact information for recall. |  | 430 | 430 | Worldwide Distribution - US in the state of Oklahoma and in the countries of Bolivia, Chile, Germany, Japan, Korea, Taiwan, Yemen, andSouth Africa. | Terminated | April 28, 2015 | 329 | 430 | Computer
Z-1989-2014 | 68528 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Revolution XR/d. |  |  | June 04, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue due to a collimator installation error during a service maintenance activity involving the Proteus XR/a and Revolution XR/d X-ray imaging systems. . There was a reported incident of a collimator fall which caused a serious patient injury. | Equipment maintenance | GE Healthcare sent an Urgent Medical Device Correction letter dated June 4, 2014, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions in regard to this recall the consignees in the United States were advised to call 800-437-1171. For other countries contact your local GE Healthcare Service Representative.For questions regarding this recall call 800-437-1171. |  | 1053 | 1053 | Worldwide Distribution - USA including PR. DC, GUAM, all states except VT. Internationally to the following countries:ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BENIN, BOSNIA AND HERZEGOVINA, BOTSWANA, BRAZIL, BULGARIA, BURKINA FASO, CANADA, CHILE, CHINA, COLOMBIA, CZECH REPUBLIC, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GERMANY, GREECE, GUATEMALA, HONDURAS, HONG KONG, HUNGRY, ICELAND, INDONESIA, IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMHIRIYA, LITHUANIA, MALAYSIA, MALTA, MARTINIQUE, MEXICO, MOROCCO, NETHERLANDS, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, PAKISTAN, PANAMA, PHILIPPINES, POLAND, PORTUGAL, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, TAIWAN, TANZANIA, THAILAND, TUNISIA, TURKEY, UGANDA, UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM YEMEN. | Terminated | May 06, 2015 | 336 | 4513 | Not_Computer
Z-1960-2014 | 65562 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur HCY Diluent - 10mL Bottle Catalog number 09877493 (124533), SMN 10310370, and Ancillary Pack Catalog number 03302138 (124581), SMN 10378677. Used on the ADVIA Centaur, ADVIA Centaur XP and ADVIA Centaur CP platforms. For in vitro diagnostic use in the quantitative determination of total homocysteine (HCY) in serum or plasma. |  |  | June 25, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc | ADVIA Centaur Homocysteine dilution recovery on the diluent identified the percent recovery for the 1:10 patient sample dilutions lower than reported in the Instructions for Use (IFU). | Device Design | Siemens Healthcare sent an Urgent Field Safety Notice that was sent to all affected Siemens Healthcare Diagnostics customers in the United States on June 25, 2013. The Urgent Field Safety Notice was sent out for distribution to all customers outside the United States on June 24, 2013. These notices inform the customer that the 1:10 diluted patient recovery is lower than reported in the Instructions for Use (IFU). Customers should discontinue use of both manual and on board 1:10 dilutions of samples with the ADVIA Centaur HCY assay. Customers can continue to use a manual or on board 1 :2 dilution for samples with results outside of the assay range. Siemens has confirmed theperformance at the 1 :2 dilution level meets the IFU ranges.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Technical Solutions Center or your local Siemens technical support representative.For further questions please call (508) 668-5000. |  | 51 | 51 units | Worldwide Distribution - USA (nationwide) and the countries of : Austria, Belgium, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Finland, France, Germany, Great Britain, Greece, Hungary, Ireland, Italy Kazakhstan, Latvia, Liechtenstein, Luxembourg, Netherlands, Norway, Poland, Portugal, Romania, Saudi Arabia, Serbia, Slovakia, Spain, Sweden, Switzerland ,Turkey, Unit.Arab Emir and Vatican | Open, Classified |  |  | 51 | Not_Computer
Z-1980-2014 | 68234 | No | HKI | Ophthalmic camera. | Ophthalmic | Ophthalmic | 510(k) | Non-Mydriatic Auto Fundus Camera. Model AFC-330. An ophthalmic camera for use in capturing images of the retina and anterior segment of the eye. |  |  | June 16, 2014 | 2014 | Nidek Inc | Image taken by AFC-330 has a white spot which may affect diagnosis or evaluation of image. | Device Design | Field Correction notifications were sent to all affected users on June 16, 2014 by trackable method. |  | 185 | 185 cameras | Nationwide Distribution. | Terminated | March 03, 2015 | 260 | 185 | Not_Computer
Z-1983-2014 | 68496 | No | HSZ | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Anspach Pneumatic and Electric Motor Systems. |  |  | May 08, 2014 | 2014 | The Anspach Effort, Inc. | Several attachments and handpieces failed to meet the expected sterility assurance level (SAL) when sterilized by the method that is specified in the current Directions for Use (DFU) provided with the device. | Process control | Customers were notified 05/08/2014 by a Return Receipt Requested letter from Anspach with an URGENT: Medical Device Labeling Correction announcement, REGARDING: STERRAD 100S Hydrogen Peroxide Gas Plasma Technology no longer recommended for sterilization of Anspach products. The letter identified the problem and actions to be taken by the customer. Customers were to complete and return the attached reply form. If the products were distributed further, then those customers should have been notified as well. Customers were to contact Anspach Customer Support at 1-800-327-6887 if they want copies of the revised labeling. |  | 115988 | 115,988 units | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX,UT, VA, VT, WA, WI, WV, WY, and Hawaii, and the territories of Puerto Rico, US Virgin Island and Guam; and, the countries of Argentina, Australia, Austria, Belgium, Brazil, Cameroon, Canada, Chile, Costa Rica, Czech Republic, Dominican Republic, Egypt, Finland, France, Germany, Grand Cayman, Hong Kong, India, Israel, Jamaica, Japan, Korea, Lithuania, Malaysia, Mexico, Nassau, Netherlands, Nigeria, Palestine, Panama, Poland, Portugal, Saudi Arabia, Singapore, South Africa, Spain, Sweden, Switzerland, Syria, Thailand, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | January 29, 2016 | 631 | 115988 | Not_Computer
Z-1948-2014 | 68242 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Manual Resuscitator Bag Series: AF1000, AF2000, AF3000, AF4000, AF5000, BT2000, BT3000, BT4000, BT5000, BVM700, CPRM1000, CPRM2000, CPRM3000, PRO-1900, PRO-1000, PRO-2000, SC7000, SC8000, SC8020, SC8120, SC9000, SS3200, VN2000, VN3000, VN4000, VN5000, VT1000 , VN2102, VN2002. Emergency manual resuscitator for the Pediatric Population. |  |  | May 16, 2014 | 2014 | Ventlab LLC | The pop-off valve in the defective devices remain open and a squeeze of the bag may not generate enough force to force the duck bill valve open and therefore no air will be delivered to the victim. | Process change control | The firm, Ventlab, issued a press dated May 14, 2014, stating Ventlab, LLC has initiated a voluntary medical device removal of a limited number of Ventlab" Resuscitator Bags. The bag series and lot numbers were included in a table as well as manufacture dates. In addition, Ventlab, LLC sent an "URGENT: MEDICAL DEVICE PRODUCT REMOVAL IMMEDIATE ACTION REQUIRED" letter dated May 19, 2014 to its distributors and customers. The letter described the product, problem and actions to be taken. The customers were instructed to review and examine their inventory for affected product; stop using them and immediately contact Ventlab, LLC for further instructions on the return of these products; and if product was further distributed, identify their customers/consignees and notify them of this product removal. The customers were also instructed to complete and return the attached Return Response Form via fax at 1-800-400-8820, Attn: Quality Department, or email to: PFA@ventlab.com as soon as possible and inform all affected personnel of this removal.Ventlab will send new corrected replacement resuscitation bags to customers once you return the affected product.If you have any questions regarding this action, please call Ventlab at 1-800-237-5481 between the hours of 8:30 AM to 5:00 PM (EST) Monday through Friday, or 1-844-635-5326 or via e-mail at PFA@ventlab.com. |  | 251384 | 251,384 | US Nationwide Distribution in the states of OH, NY, TX, TN, FL, CT, and VA, | Terminated | November 17, 2015 | 550 | 251384 | Not_Computer
Z-1981-2014 | 68489 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Langston Dual Lumen Catheters, 6Fr, REF 5550 and 5540. The Vascular Solutions Langston Dual Lumen Pressure Monitoring catheter is used as a pressure measurement catheter and for delivery of contrast media during angiographic studies. |  |  | May 27, 2014 | 2014 | Vascular Solutions, Inc. | The inner catheter of some Langston V2 Dual Lumen Catheters have separated from the device hub during contrast pressure injections. This may require an intervention to retrieve the separated piece from the patient to prevent injury. | Component design/selection | Vascular Solutions sent an Urgent Medical Device Removal letter dated May 23, 2014, to all affected consignees.The letter described the problem and the product affected by the recall. Advised consignees to remove affected devices and place them in a secure area. Requested consignees to fill and return the "VSI Account Inventory Form" to VSI's Customer Service Department by email customerservice@vasc.com or fax 763-656-4251. VSI will replace all returned devices. Customers with questions were advised to contact their local Vascular Solutions Account Manager.For questions regarding this recall call 763-656-4300. |  | 8580 | 8,580 total (US 5714) | Worldwide Distribution - USA nationwide including DC and Puerto Rico, AUSTRIA, CANADA, GERMANY, ISRAEL, ITALY, NETHERLANDS, SPAIN, SWEDEN SWITZERLAND, and UNITED KINGDOM. | Terminated | January 30, 2015 | 248 | 8580 | Not_Computer
Z-1992-2014 | 68488 | No | MGQ |  |  | General &amp; Plastic Surgery | 510(k) | Amerigel Wound Dressing 1 oz. tube |  |  | March 21, 2014 | 2014 | Kova Laboratories, Inc. | Product may not have been manufactured pursuant to a validated process protocol and CGMP. | Process control | Kova Laboratories contacted customers with an URGENT: Medical Device Recall message by phone and e-mail 3/21/2014 and 4/21/2014, of the affected products to be returned. |  | 122396 | 122,396 / 1oz. tubes | Distributed in the states of FL and NJ. | Open, Classified |  |  | 122399 | Not_Computer
Z-1994-2014 | 68593 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Patient Side Cart Battery Box used in conjunction with the da Vinci A surgical system. |  |  | April 27, 2012 | 2012 | Intuitive Surgical, Inc. | Patient Side Cart battery boxes may overheat in rare charging conditions, causing battery bulging. | Device Design | No notification was sent. Issue is being addressed during routine maintenance or through complaint process. |  | 2606 | 2606 | Worldwide Distribution. | Terminated | July 07, 2014 | 801 | 2606 | Not_Computer
Z-1995-2014 | 68580 | No | LDS | Medical washer. | General Hospital | General Hospital | 510(K) Exempt | HLD System, Model 610, Medical Device Cleaning and High Level Disinfection Washer/Pasteurizer. |  |  | January 24, 2013 | 2013 | Cenorin, LLC | The temperature sensor/control system in the HLD Systems Model 610 may provide a false temperature reading. This caused the water to remain at ambient temperature throughout the pasteurization cycle rather than using hot water. | Device Design | Cenorin sent a Notification letter dated December 18, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter stated: Cenorin has developed an upgrade to mitigate this issue as well as provide additional process control and system monitoring benefits. Cenorin included a Return Response form with the recall letter to be completed and return. For questions contact HLD Systems Sales Representative for questions. |  | 25 | 25 HLD systems | Worldwide Distribution - US (nationwide): CA, DE, FL, MA, MI, NY, PA, TX, UT, WA, and WI and the countries of Uruguay and Taiwan. | Terminated | February 10, 2015 | 747 | 25 | Not_Computer
Z-2002-2014 | 68415 | No | FRC | Sterilization process indicator. | General Hospital | General Hospital | 510(k) | 3M Attest Auto-reader. Model numbers 390, 390G,and 490. |  |  | June 09, 2014 | 2014 | 3M Company - Health Care Business | Labeling on units shipped prior to May 16th, 2014 did not contain the statement "This product contains dry natural rubber" as required by the United States Food and Drug Administration (FDA) for medical devices. The non-slip pads at the bottom of the unit contain dry natural rubber. One incident of an allergic reaction in a sensitized individual has been reported. | Labeling Change Control | Consignees were sent on 6/9/2014 a 3M "Urgent Medical Devices Correction" letter dated May 21, 2014. The letter described the problem and the product involved in the recall. The letter described what action is 3M taking and what action was required of the consignees. Consignees were sent an adhesive label and a response card. They requested consignees to adhere the label to a portion of the plastic covering each unit in a location where it is visible to the end-user. After that, they are requested to fill out and return the response card. |  | 5343 | 5343 (2621 USA, 2722 OUS) | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the countries of AUSTRALIA, CANADA, CHILE , COLOMBIA, COSTA RICA, SOUTH AFRICA, BRAZIL, GERMANY, ECUADOR, DUBAI, JAPAN, HONG KONG , CHINA, KOREA, MALAYSIA, PANAMA, SINGAPORE, VATICAN, and THAILAND. | Terminated | June 03, 2015 | 359 | 5343 | Not_Computer
Z-2017-2014 | 68493 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Humidifier Adaptor, 040 Shelf Pak, French, Teleflex Medical, respiratory gas humidifier adaptor. |  |  | June 05, 2014 | 2014 | Teleflex Medical | The seals of the adaptor packaging may be creased which may potentially affect packaging integrity. | Packaging change control | Consignees were notified by letter on 06/06/2014. |  | 7758650 | 7,758,650 ea. (total for all types listed) | Worldwide Distribution-USA (nationwide) and the countries of Belgium, Canada, China, Germany, Great Britain, France, India, Malaysia, Mexico, Philippines, Singapore, Thailand and Zambia. | Terminated | July 02, 2015 | 392 | 7758650 | Not_Computer
Z-1998-2014 | 68512 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | 07-150-07 MIZUHO, MAL DISP DOPPLER PROBE BX 4, REF 138200 |  |  | May 28, 2014 | 2014 | Vascular Technology, Inc. | Report of the product penetrating the sterile barrier; thiscould render the product unsterile. | Process design | Vascular Technology sent an Urgent Medical Device Recall letter dated May 28, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This recall has been initiated due to a report of the product penetrating the sterile barrier; this could render the product unsterile. Use of this product may lead to infection in the patient.Please immediately examine your inventory and return all product subject to recall to Vascular Technology at the above mentioned address. In addition, if you have further distributed this product, please identify your customers and notify them at once of this product recall.Your notification to your customer must include instructions telling your customer to return the recalled product if unused or to account for the recalled product if it has been used.For further questions please call (603) 594-9700. |  | 164 | 164 | Worldwide Distribution - US including California and the countries of Australia, United Kingdom and Taiwan. | Terminated | February 11, 2016 | 624 | 370 | Not_Computer
Z-1996-2014 | 68552 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Henora REF: 70-4026-09V1 300psi I.V. Catheter Extension Set |  |  | March 28, 2014 | 2014 | RGI Medical Manufacturing, Inc. | Problem with low adhesion of the tubing to luer connection which may lead to a separation of the extension set while in use. | Process control | RGI Medical Manufacturing, Inc. phoned their direct account: Baxter Healthcare on February 28, 2014 and sent an Urgent: Device Recall letter dated March 21, 2014. The letter identified the affected product, problem and the actions to be taken. Contact 800-747-6870 to obtain instruction for returning the product to RGI |  | 19200 | 19,200 units. | US Nationwide Distribution in the states of AL, CA, DE, GA, IL, KY, MA, MD, ME, MI, MN, NJ, NM, NV, NY, OH, TN, TX, WA, and WI. | Open, Classified |  |  | 19200 | Not_Computer
Z-1997-2014 | 68243 | No | CCX | Patient position support. | Anesthesiology | Anesthesiology | 510(K) Exempt | Devon Supine, Lithotomy and Trendelenburg Positioning Kit; Product Code: 31151090 |  |  | May 08, 2014 | 2014 | Covidien LLC | Straps may separate from the foam pad. | Process control | Covidien Sales Representatives notified customers in person by beginning May 8, 2014. The Sales Representative will provided a copy of the Recall letter (attachment 4), and collect any remaining product at the site. The notification and recovery of the product is expected to be complete no later than May 14, 2014. |  | 264 | 264 devices | US Nationwide Distribution in the states of NJ, CA, TX, IL, MO, PA, OH, IN, MI, MN, FL, IN, DC, PA, VA, MS, ME, NV, AR, KY, NY, KS, SC | Terminated | September 04, 2015 | 484 | 264 | Not_Computer
Z-2018-2014 | 68244 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | TerumoÂż Advanced Perfusion System 1 |  |  | June 17, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | An evaluation of the APS1 Operator Manual found that existing instructions lack clarity and accuracy related to specific items. | Labeling Change Control | On June 17, 2014, TCVS sent URGENT MEDICAL DEVICE RECALL notification to consignees. Notification included reason for recall, potential hazard, correction, affected population, and customer instructions. Questions concerning the recall are directed to Terumo CVS Customer Service @ 800-521-2818 (M-F, 8am-6pm) |  | 1713 | 1713 units | Worldwide Distribution-USA (nationwide) including DC and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of AUSTRALIA, BELGIUM, CANADA, CHILE, Hong Kong, Indonesia, Japan, Malaysia, Mexico, Philippines, Singapore, South Korea, Taiwan, Thailand, UNITED ARAB EMIRATES (UAE), and Vietnam. | Terminated | September 22, 2014 | 97 | 1713 | Not_Computer
Z-2020-2014 | 68553 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Intuitive Surgical Inc. Large Clip Applier instruments used in conjunction with da Vinci S and da Vinci SI Surgical Systems |  |  | June 13, 2014 | 2014 | Intuitive Surgical, Inc. | Manufacturing variation in the grip assembly of the Large Clip applier which can lead to the instrument jamming during insertion when used with the 8 mm cannula or with certain reducer accessories.. | Process control | Urgent Medical Device Correction letters were sent on June 18, 2014 to all affected sites via trackable mail. |  | 847734 | 5290 for part 420230-04; 1964 for part 420240-06 | Worldwide Distribution. | Terminated | December 01, 2014 | 171 | 847734 | Not_Computer
Z-2025-2014 | 68608 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set, Model Codes: 10014855A, 10015861A, 10061661A |  |  | June 17, 2014 | 2014 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion Set because they were assembled with a drip chamber that consist of a 15 micron filter. The Directions For Use did not indicate that the drip chamber has a 15 micron filter and it does not provide the instruction that the set should not be used with blood administration due to the potential for occluding the 15 micron filter in the drip c | Labeling Change Control | A recall letter is dated 6/25/14 and will be sent out to customers who purchased the SmartSite Low Sorbing Infusion Set because they were assembled with a drip chamber that consist of a 15 micron filter. The Directions For Use did not indicate that the drip chamber has a 15 micron filter and it does not provide the instruction that the set should not be used with blood administration due to the potential for occluding the 15 micron filter in the drip chamber. The letter informs the customers of the problems identified, potential risk, and actions to be taken. Customers are instructed to promptly complete and return the enclosed mandatory Customer Response Card to acknowledge the receipt of this communication and to expedite the corrective action process. |  | 5600 | 5,600 units | Distributed in the states of FL, NM, WA, WI, and TX. | Open, Classified |  |  | 5600 | Not_Computer
Z-2026-2014 | 68678 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The STERIS 4085 General Surgical Table |  |  | June 27, 2014 | 2014 | Steris Corporation | STERIS engineering analysis has determined the D1 pressure switch in the hydraulic column manifold does not adequately detect pressure level. If one or more of the table floor lock feet become unlocked due to the loss of hydraulic pressure, there is a potential for the table hand control to incorrectly indicate that the floor lock feet are locked. | Device Design | STERIS sent an Urgent Voluntary Field Correction Noticel letter dated June 27, 2014 to affected Customers.via FedEx with tracking numbers for delivery confirmation. The letter identified the description of the product, description of the problem and the actions to be taken. For questions contact Steris Field Service Dispatch at 1-800-333-8828. |  | 3689 | 3689 | Worldwide Distribution - US (nationwide in the states of AL, AK, AZ, AR, CA,CO, CT, DE, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS MO, MT, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VT, VA, WA, WV, WI, WY in the countries of Canada, India, Vietnam, China, India, Korea, Mexico, costa Rica, Brazil, Taiwan, Australia, New Zealand, El Salvador, Indonesia, Hong Kong, Singapore, Nepal. | Terminated | September 16, 2015 | 446 | 3689 | Not_Computer
Z-2024-2014 | 68607 | No | GAR | Nonabsorbable polyamide surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Dafilon Suture Black 10/0 (0.2) 15 cm DRM4 NS, non-sterile, non-absorbable polyamide Surgical Suture, additional label affixed on the box states the product is sterile, packaged in a cassette holding 12 sutures |  |  | June 13, 2014 | 2014 | Aesculap, Inc. | Aesculap Inc. (AIC (USA)) has initiated a voluntary recall of Dafilon Black 10/0 (0.2) 15 cm DRm4 non-sterile non-absorbable polyamide Surgical Suture due to a labeling issue. Box labels do state non-sterile; however there is an additional label that is affixed on the box that states the product is sterile. | Labeling design | A correction and removal notification, dated June 13, 2014, was sent to end users, sales reps, and distributors which described the product, problem, and actions to be taken. |  | 11 | 11 | Distributed in the states of IL, NY, GA, and RI. | Terminated | October 07, 2014 | 116 | 11 | Not_Computer
Z-2022-2014 | 68546 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | STOPCOCK EXT. SET W/ 3-WAY STOPCOCK WITH INJ. SITE AND SPINLOCK CONNECTOR, 22 inches and 37 inches |  |  | May 29, 2014 | 2014 | B. Braun Medical, Inc. | Complaints were received that reported the stopcock of some Stopcock Extension Sets are assembled in the reverse position, resulting in the stopcock being closed when set in the open position and vise versa. | Process control | An urgent medical device correction notification, dated June 3, 2014, describing the product, problem, and actions to be taken, was sent to end users and distributors. |  | 10835 | 10,835 | USA (nationwide) including the states of PA, IL, NC, FL, IL, CA, CO, NE, LA, NC, LA, VA, TN, NE, TX, and KS. | Terminated | December 10, 2014 | 195 | 10835 | Not_Computer
Z-2023-2014 | 68399 | No | INN | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Aquatec Fixed Backrest - accessory for Mobile Shower and Toilet commode Aquatec Ocean, Ocean VIP and Ocean E-VIP family. |  |  | June 25, 2014 | 2014 | Invacare Corporation | The AquatecÂż Fixed Backrest withstands a load of only 250 newtons at uneven load. This may result in the Fixed Backrest becoming detached from the frame. | Device Design | The recall notification will be sent to the two consignees informing them of the recall and to notify their customers. The letter will include a provider letter, customer letter, provider response card and impacted units listing. |  | 465 | 465 units | Distributed in PA and Canada. | Terminated | April 15, 2015 | 294 | 465 | Not_Computer
Z-2034-2014 | 68390 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer M/L Taper Reduced Neck Standard &amp; Extended Offset, HIP JOINT PROSTHESIS, Sizes:4, 5, 6, 7.5, 9, 10, 11 &amp; 12.5. |  |  | May 23, 2014 | 2014 | Zimmer, Inc. | Black residue observed in the sterile pouch of the M/L Taper and Kinectiv stems. | Packaging | Zimmer sent an Medical Device Recall Letter on June 4, 2014, to their Distributors. On June 9, 2014, Zimmer sent Medical Device Recall Letters to the physicians. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed:1. Locate all products from the affected lots and quarantine them immediately, if not previously done so per the May 23, 2014 Notice of Suspension of Loaner Pool. 2. Confirm the physical count of all affected product in loaner pool in your territory and complete the Confirmation of Loaner Pool Inventory and Return Certification Form (Attachment 2), and email a completed copy of Attachment 2 to corporatequality.postmarket@zimmer.com.3. Return the recalled product along with the completed Confirmation of Loaner Pool Inventory and Return Certification Form (Attachment 2). 4. Identify additional customers that may have used this product in the past 12 months by completing the spreadsheet template provided to you via email or confirm that Zimmer has identified all current customers in your territory.5. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 852 | 852 units | Worldwide Distribution - USA (nationwide) and the countries of Australia, Canada, China, Europe, India, Japan, Korea and Singapore. | Terminated | March 06, 2015 | 287 | 2654 | Not_Computer
Z-2027-2014 | 68658 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Hospira MedNet Medication Management Suite (MMS). |  |  | June 06, 2014 | 2014 | Hospira Inc. | Hospira MedNet 5.5, 5.8.1, and 5.8.2 contains software defect where the dosing units of "nanog/kg/min" and "milliUnits/min" are not sent to a Plum A+ device (Version 13.40, 13.41, and 13.4.2) via Auto Programming when those dosing units are not properly established in the customized drug library. A dosing unit mismatch of "nanog/kg/min" or "milliUnits/min" on an unmatched medication will produce a | Under Investigation by firm | Hospira sent an URGENT MEDICAL DEVICE CORRECTION letter dated June 6, 2014 was sent to all affected customers. The letter included instructions for customers to: notify all potential users in their organization of the recall, and to complete and return the attached reply form. The letter further notified customers that they will be contacted by a member of Hospira's Clinical Services Team to make arrangements to review the drug library and implement the appropriate mitigation. Hospira will further work with the direct account to ensure Auto Programming installations receive Pharmacy consulting to support the implementation of the mitigation and that a correction for this issue will be addressed in a future version of the software. Reports of adverse events or product complaints can be made by contacting Hospira Global Complaint Management at 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) or via e-mail at ProductComplaintsPP@hospira.com. Additional information or technical assistance can be obtained by contacting the Hospira Advanced Knowledge Center at 800-241-4002, option 4 (24 hours a day, 7 days a week). |  | 13 | 13 units | US Nationwide in the states of FL, LA, PA, WA, WY and the country of Hong Kong | Open, Classified |  |  | 13 | Computer
Z-2039-2014 | 68646 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | Essential NEO, Audio/Visual Integration Unit |  |  | June 25, 2014 | 2014 | Karl Storz Endoscopy | KARL STORZ Endoscopy-America is recalling the Essential NEO because the connected power cable could potentially become overheated and result in smoke being generated and the cessation of function of the peripheral that the cable is connected to. | Nonconforming Material/Component | Storz sent an Safety Alert letter dated June 25, 2014, to all customers who purchased the Essential NEO. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the attached Acknowledgement/Response Form with contact information and return to KARL STORZ. Customers with questions are instructed to contact KARL STORZ' Technical Support Department at (800) 421-0837, option 3, then option 4. |  | 968 | 968 units | Nationwide Distribution including the states of: AL, AZ, CA, CO, CT, District of Columbia, DE, FL, GA, HI, IA, IL, IN, KY, LA, MX, MS, MD, MI, MA, MO, MN. NM, NV, NJ, NY, NC, NE, OK, OH, PA, SC, SD, TX, UT, VA, WI and WA., and Internationally to Canada. | Terminated | January 13, 2015 | 202 | 968 | Not_Computer
Z-2032-2014 | 68547 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Zimmer EDI PERIARTICULAR PLATING SYSTEM CANCELLOUS BONE SCREW, FULLY THREADED, HEX HEAD 4MM DIAMETER, 14 MM LENGTH, and EDI 4.0 MM CANCELLOUS SCREW, FULLY THREADED, 20 MM LENGTH. Temporary internal fixation devices, designed to stabilize fractures during the normal healing process. |  |  | June 12, 2014 | 2014 | Zimmer Manufacturing B.V. | Two lots of screws were commingled. Screws with etching and machining for lot number 62629432 were packaged in bags labeled for lot number 62628759; and screws with etching and machining for lot number 62628759 were packaged in bags labeled for lot number 62629432. The affected product was distributed between 3/6/14 and 4/21/14. | Process design | Zimmer sent an Urgent Medical Device Recall Lot specific letter dated June 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed:1. Please locate, quarantine, and return the affected product (part number 00-4840-020-00 lot number 62629432 and part number 00-2347-022-14 lot number 62628759) from all inventories.2. Please note that any hospitals that received direct shipments of this product from Zimmer will be notified via Fed Ex.3. Please review the list of hospital and surgeons to be provided to your territory via email.4. Provide the name, address, and phone number of any additional hospitals or surgeons that may have received or used the affected product, or confirm that the list provided is complete.5. Complete and return Attachment 1.6. Ensure that a copy of the Risk Managers and Surgeons letter is left with the facilities from which product is removed by providing them with a copy of the letter, if they are not identified as a facility receiving a letter directly from Zimmer.7. If after reviewing this notification you have questions or concerns, please contact Zimmer at 1-877-946-2761. |  | 496 | 496 | Worldwide Distribution - USA (nationwide) and the countries of India, Japan, China, Australia and Canada. | Terminated | June 24, 2015 | 377 | 496 | Not_Computer
Z-2036-2014 | 68696 | No | MEC | Medical washer-disinfector. | General Hospital | General Hospital | 510(K) Exempt | Reliance 1227 Cart and Utensil Washer/Disinfector |  |  | June 30, 2014 | 2014 | Steris Corporation | STERIS has identified that deformation of the washers door handle pin may occur when the door handle is rotated with sufficient force. Should deformation and/or breakage of the door handle pin occur there is a potential for the washer door to become inoperable. | Device Design | Customers will be notified via onsite visits by STERIS Field Service Technicians to each affected Customer site. These visits will be documented and monitored until completion. |  | 37 | 37 | US Nationwide Distribution - in the states of AL, AR, CA, DE, FL, GA, IL, IA, LA, MA, MI, NJ, NY, OH, OK, OR, SC, TX, VA. and in the country of Canada. | Terminated | July 01, 2015 | 366 | 37 | Not_Computer
Z-2040-2014 | 68591 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ColorSafe IV lines |  |  | August 17, 2011 | 2011 | Codan Us Corporation | CODAN US Corporation is recalling the ColorSafe IV (intravenous) product lines because they were manufactured and marketed prior to FDA approval. | Other | CODAN sent a recall letter dated August 17, 2011, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return any of the affected product. If they sold the product they should provide Codan the name and address of the customer. so that they may be contacted to retrieve the shipped products. Customers were instructed to contact CODAN Customers Service for a Returned Goods Authorization. Customers with questions were instructed to call 714-430-1303. Customers were instructed to not send used devices.For questions regarding this recall call 714-545-2111. |  | 7250 | 7250 units | Worldwide Distribution - US including TX, CT, NY and Internationally to Belgium. | Terminated | July 10, 2014 | 1058 | 7250 | Not_Computer
Z-2031-2014 | 68200 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar infusion pump, List numbers 13000 and 13100, GemStar Infusion Pump (pain management) List number 13150. Used for intravenous, arterial, subcutaneous, short-term epidural infusion and parenteral administration of general I.V. fluids, medications, nutritional fluids, and blood/blood products. |  |  | January 31, 2014 | 2014 | Hospira Inc. | The connection between the beeper subassembly and the pump may fail. The GemStar infusion pump will identify this failure during the "self-test" while powering up which will result in a Beeper Error ("code 10/001/000"). This Beeper Error (code 10/001/000) is a service alarm that places the pump in an inoperable mode and requires service before it can be returned to service. | Component design/selection | Hospira sent an URGENT MEDICAL DEVICE RECALL letter dated January 31, 2014 was sent to all affected customers. The letter included instructions for customers to: 1) inform potential users of these devices of the notification, 2) if the beeper error occurs during power up, remove the device from service immediately, 3) report any beeper errors to Hospira by calling 800-441-4100 (Monday - Friday, 8:00 AM - 5:00 PM CST) or e-mail ProductComplaintsPP@hospira.com, 4) complete the attached reply form and return it to the fax number of e-mail address on the form (even if there is no product in inventory), and 5) if the products were further distributed, notify those customers that received the products and ask them to contact Stericycle at 855-827-6578 (Monday - Friday, 8:00 AM - 5:00 PM EST) to receive a reply form. A subsequent letter was sent on February 18, 2014 that included a correction to the description of the error code generated when a beeper wire failure occurs. That letter included identical instructions as those found in the January 31, 2014 letter for recipients to follow. |  | 36881 | 36,881 infusion pumps | Worldwide Distribution - US: Nationwide in the state of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY including District Of Columbia, Guam, Puerto Rico; and the countries of Austria, Australia, Bahrain, Belgium, Canada, Chile, France, Germany, Greece, Guatemala, Hong Kong, Italy, Jordan, Kuwait, Lebanon, Libya, Luxembourg, Malaysia, Mexico, Netherlands, Oman, Portugal, Saudi Arabia, Slovenia, South Africa, Spain, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela | Open, Classified |  |  | 36881 | Not_Computer
Z-2038-2014 | 68424 | No | CAH | Breathing circuit bacterial filter. | Anesthesiology | General Hospital | 510(k) | PALL Medical Breathing Circuit Filter, REF BB50T and BB50T-BULK , Rx ONLY Single use, bi-directional bacterial/viral removal filter. |  |  | June 04, 2014 | 2014 | Pall Corporation | Possible leak from the filter allowing a small of amount of air to be released from the junction between the filter end cap (flat side of the filter) and the filter housing connection. | Other | Urgent Product Recall Letters (dated 6/03/2014) and Response Forms were sent to customers via fax and e-mail starting on 6/04/14. The letters instruct the direct accounts to examine their inventory and place the affect lots on hold in quarantine. For questions, or if a patient injury associated with this device has been reported to you, please call Karen Peterson directly at 516.801.9267 (office) or 516.644.6171 (cell). The letter for the bulk product also instructs customers to notify their customers of the recall and advise them not use the affected lots if the product had been distributed. |  | 24650 | 24,650 units domestically | Distributed in the states of MN, MO, OH, MA, NY, NJ, WY, AZ, CA, IL, LA, IN, CO, NC, and TX. | Terminated | April 01, 2015 | 301 | 24650 | Not_Computer
Z-2037-2014 | 68643 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GEMINI TF PET/CT16: 882470, GEMINI TF PET/CT 64: 882471, GEMINI TF Big Bore: 882476. A diagnostic imaging system for fixed or mobile installations that combines Positron Emission Tomography (PET) and X-ray Computed Tomography (CT) systems. |  |  | June 18, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | When beginning a procedure to start the gantry and CT Host, a Close Estop message may appear to the operator, but disappears before action can be taken. If this occurs it results in an inability to close E-Stop and the system is not operational for clinical use. This problem occurs randomly. | Software design | Philips Healthcare is informing the consignee of this issue through a Field Safety Notice (FSN) 88200487-488. The US consignees will receive this notice via certified mail and Philips Key Market representatives will distribute the FSN internationally. Additionally, a Field Service Engineer will contact each site to schedule a time to install the software update. |  | 33 | 33 | Distributed in the states of CA, CO, IL, OK, OR and WI. | Terminated | September 16, 2015 | 455 | 33 | Computer
Z-2029-2014 | 68024 | No | KFM | Nonroller-type blood pump. | Cardiovascular | Cardiovascular | 510(k) | Sarns Disposable Centrifugal Pump with X-Coating, 164275X, is a sterile, single use device. |  |  | April 14, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | During set-up and priming of the bypass circuit, leaks were detected on the centrifugal pumphead. | Device Design | Urgent Medical Device Recall letters, dated April 11, 2014, were sent to customers. The letters identified the affected product, as well as the reason for the correction, the potential hazard, the correction, the affected population, and instructions for customers. Customers are to review the notice and ensure that all users have received a copy. Additionally, customers are to confirm receipt of the communication by returning the attached Customer Response Form as instructed; and, return all effected centrifugal pumps. Terumo CVS will issue a Returned Goods Authorization upon receipt of the Response Form. Questions should be directed to Terumo CVS Customer Service at 800-521-2818. |  |  |  | Nationwide Distribution | Open, Classified |  |  |  | Not_Computer
Z-2041-2014 | 68468 | No | JSS | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | BBL CrystaLEAD Enteric/Nonfermentor (E/NF) Identification (ID) System, catalog number 245000. in vitro diagnostic for the identification of aerobic gram-negative bacteria that belong to the family Enterobacteriaceae as well as some of the more frequently isolated glucose fermenting and nonfermenting gram-negative bacilli. |  |  | May 21, 2014 | 2014 | Becton Dickinson &amp; Co. | A portion of the BBL Crystal Enteric/Nonfermentor (E/NF) Identification kits from lot numbers 3224281 and 3290379 are labeled with incorrect expiration dates. A portion of these kits contain Crystal Inoculum fluid bottles that will expire prior to the date that was printed on the kit carton. The expiration date printed on the individual Crystal Inoculum fluid bottles is correct. | Labeling mix-ups | BD Diagnostic sent an Urgent Product Recall letter dated May 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed:Please discard any outer kit boxes that are labeled with the incorrect expiration date. All components of the kit will perform as expected with the expiration dates on each component.Our records indicate you may have been shipped a portion of the above-referenced product between August 23, 2013, and the present day.Please complete the attached form whether or not you have any inventory remaining so that we may acknowledge your receipt of this notification. It is not necessary to contact BD by telephone. Simply complete and fax the form to: Attention: Regulatory Compliance, Fax: 410-316-4258.If further assistance is needed, please contact the BD Customer Service Department at 1-800-675-0908. For all other inquiries, please contact the BD Technical Services Department at 1-800-638-8663.Please accept our apology for any inconvenience this may cause. BD is committed to providing you with the highest quality products. Thank you for your continued support. : -"/ 06 . ;/ i )!.U |  | 18.21 | 18.210 | Worldwide Distribution - US (nationwide), Australia, Belgium, Canada, Colombia, China, Indonesia, India, Japan, Korea, Mexico, Malaysia, Peru, Philippines, Singapore, Thailand, and Taiwan. | Terminated | August 12, 2016 | 814 | 18.21 | Not_Computer
Z-2043-2014 | 68632 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia E camera system. Emission Computed Tomography System used to detect or image the distribution of radionuclides in the body or organ, using the following techniques: planar imaging, whole body imaging, tomographic imaging for isotopes with energies up to 588 keV. Nuclear medicine diagnostic device. |  |  | April 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Siemens Medical Solutions USA, Inc. has received reports indicating the collimator cart rear casters may become loose. | Device Design | A Customer Safety Advisory Notice dated April 23, 2014 was sent to every direct account. The letter included instructions for direct accounts to discontinue use of the collimator cart and notify the local service representative if they notice or suspect tha one or more casters are loose (which would be indicated by shaking while the cart is in motion). Otherwise, a service representative will be contacting the direct account within the following 3 months to schedule a retention bracket installation. Direct accounts with questions can contact Siemens Medical Solutions USA, Inc. at 800-888-7436. |  | 531 | 531 systems | Worldwide Distribution- US: Nationwide (AK ; AL ; AR ; AZ ; CA ; CO ; CT ; DE ; FL ; GA ; HI ; IA ; ID ; IL ; IN ; KS ; KY ; LA ; MA ; MD ; ME ; MI ; MN ; MO ; MS ; MT ; NC ; ND ; NE ; NH ; NJ ; NM ; NV ; NY ; OH ; OK ; OR ; PA ; RI ; SC ; SD ; TN ; TX ; UT ; VA ; VT ; WA ; WI ; WV ; WY) + District of Columbia + Puerto Rico; FOREIGN: Albania ; Algeria ; Angola ; Argentina ; Australia ; Austria ; Azerbaijan ; Bahrain ; Bangladesh ; Belarus ; Belgium ; Botswana ; Brazil ; Bulgaria ; Canada ; Chile ; China ; Colombia ; Costa Rica ; Croatia ; Cuba ; Czech Republic ; Dominican Rep. ; Ecuador ; Egypt ; El Salvador ; Finland ; France ; Georgia ; Germany ; Ghana ; Greece ; Hong Kong ; India ; Indonesia ; Iran ; Iraq ; Ireland ; Israel ; Italy ; Jamaica ; Japan ; Jordan ; Kazakhstan ; Kuwait ; Macedonia ; Malaysia ; Mexico ; Morocco ; Myanmar ; Netherlands ; New Zealand ; Niger ; Nigeria ; Norway ; Oman ; Pakistan ; Palestinian Territory ; Peru ; Philippines ; Poland ; Portugal ; Qatar ; Romania ; Russia ; Saudi Arabia ; Serbia ; Singapore ; Slovakia ; Slovenia ; South Africa ; South Korea ; Spain ; Sri Lanka ; Sweden ; Switzerland ; Syria ; Taiwan ; Thailand ; Trinidad and Tobago ; Tunisia ; Turkey ; Ukraine ; United Arab Emirates ; United Kingdom ; Uruguay ; Vietnam ; Yemen. | Open, Classified |  |  | 2989 | Not_Computer
Z-2044-2014 | 68570 | No | EFB | Dental handpiece and accessories. | Dental | Dental | 510(k) | Lynx TM20 Torquemaster Low Speed Handpiece MTI Dental Products |  |  | May 07, 2012 | 2012 | MTI Precision Products LLC. | Sharp edge on the body/housing of the Lynx TM20 TorqueMaster Low Speed Handpiece. | Software design (manufacturing process) | MTI notified their customers via e-mail on 5/712 and 5/8/12. MTI sent end-users recall letters on or about 5/9/12. |  | 134 | 134 devices | US Nationwide Distribution in the states of CA, FL, IL, KY, MA, MD NJ, NY, and RI. | Terminated | March 12, 2015 | 1039 | 134 | Not_Computer
Z-2050-2014 | 68584 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Zimmer Dental Tapered Screw-Vent Implant, Catalog TSVT6B13, Lot 62284006 |  |  | June 25, 2014 | 2014 | Zimmer, Inc. | Zimmer Dental is conducting a voluntary recall of a single lot of the Tapered Screw-Vent Implant, Catalog Number TSVT6B13, Lot number 62284006, due to a potential crack in the cap of the outer vial. | Under Investigation by firm | Recall notification letters sent to customers on 06/25/14.The recall letter, titled "URGENT DEVICE RECALL NOTICE", delivered to customers via a courier, with instructions for returning any unused affected product. The recall letter was accompanied with a response form and provided the following information to customers: product information with codes, reason for recall, risks, responsibilities, MedWatch Reporting and contacting information. |  | 129 | 129 units | Nationwide Distribution. | Terminated | October 30, 2014 | 127 | 129 | Not_Computer
Z-2051-2014 | 68581 | No | BTY | Predictive pulmonary-function value calculator. | Anesthesiology | Anesthesiology | 510(k) | Shape HF Cardiopulmonary System. Made up of Shape HF System Analyzer (PN 0001-9001) and Disposable Patient Interface (DPI), (PN 0004-4001). A pulmonary gas exchange testing system used during cardiopulmonary exercise testing where collection and review of gas exchange variables are indicated. The system will provide physiological data to physicians to aid in patient assessment. |  |  | May 28, 2014 | 2014 | Shape Medical Systems, Inc | Shape Medical has initiated a correction due to a mandatory software upgrade for the Shape HF Cardiopulmonary Testing System prior to use of the impacted DPI lot numbers. Use of the incorrect software version with the impacted DPIs could result in erroneous testing results potentially leading to incorrect diagnosis and incorrect treatment. | Under Investigation by firm | Shape Medical Systems sent an Urgent Medical Device Correction letter dated June 17, 2014, to all affected customers. Customers were informed that new orders of the Disposable Patient interface devices (DPI) will require a software upgrade for the Shape HF Cardiopulmonary Testing System prior to using the new DPIs. Impacted DPIs will be shipped with Instructions to conduct the software upgrade and Technical Service will contact customers shipped the impacted DPIs to confirm receipt and assist with the software upgrade. A Response Form was requested to be returned. Customers with questions were instructed to call Shape Technical Service at 1-888-906-6266 or 651-621-2990. |  | 800 | 800 | Nationwide Distribution | Terminated | May 22, 2015 | 359 | 800 | Computer
Z-2047-2014 | 68565 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | SUTUREFIX Ultra S with 1 #2 Ultrabraid (blue) |  |  | June 02, 2014 | 2014 | Smith &amp; Nephew, Inc. Endoscopy Division | Sterility of device is compromised due to breach in sterile barrier | Package design/selection | Smith &amp; Nephew issued an Urgent - Product Recall Notification letter dated June 2, 2014 via Federal Express to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to inspect their inventory and locate all affected product and quarantine them immediately.Complete the Return Certification Form .Contact the Smith &amp; Nephew Returns Group at 800-343-5717 (option 3) or send an email to ProductRecoverv@smith-nephew.com. They will provide customers with instructions for returning their product and receiving a replacement.For questions regardin g this recall call 978-749-1073. |  | 907 | 907 units | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, Denmark, France, Germany, Great Britain, Italy, Netherlands, Norway, Portugal, Spain, and Sweden. | Open, Classified |  |  | 3151 | Not_Computer
Z-2055-2014 | 67813 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | BAM (SHAPE) ARM, LEFT, an articulating, arm bracket that is attached to the service head of boom systems which are used for supporting and positioning equipment in operating rooms and patient care areas. The product comes as component of kit part number 0682000976, and the Arm part number in the kit is P17087. The kit, when assembled, supports a flat-panel video monitor that is attached to one of two types of Stryker Communications boom systems, EDS or FLEXiS. |  |  | March 17, 2014 | 2014 | Stryker Communications | During the use of the SHAPE Arm there is the potential for the monitor mount to come off of the SHAPE Arm if the lock washer and screw are not present. | Process control | Direct accounts in the US were notified by letter dated March 21, 2014, and asked to report any malfunctions to Stryker Communications immediately. In addition, Stryker Representatives will visit all US accounts to inspect all the product units on hand. If needed, Representatives will check that black, screw/lock washers supplied by Stryker are secured in place and will replace any that are not the typed furnished by Stryker. All units will be updated with a yellow, warning sticker on the monitor bracket which reads: "Failure to properly install the screw and lock washer below can result in the attached component/equipment falling." Stryker Communications' foreign distributors will oversee notification, inspection, and corrections of their accounts and report corrections and other findings to Stryker Communications. |  | 953 | 953 SHAPE Arms of all types | Worldwide Distribution. US nationwide including Puerto Rico, China, Philippines, Mexico, Korea, Japan, Colombia, Chile, Canada, Brazil, and Australia | Terminated | May 15, 2015 | 424 | 953 | Not_Computer
Z-2057-2014 | 68583 | No | MDB | Microbial growth monitor. | Microbiology | Microbiology | 510(k) | Thermo Scientific VersaTREK REDOX 1, Ref 7102-44, 80ml with Stir Bar (O2 aerobic), 50 bottles/box, IVD. The firm name on the label is Remel Inc., Lenexa, KS. |  |  | June 11, 2014 | 2014 | Remel Inc | Bottles did not provide positive signals within the time frame of the quality control specification for one of two strains of Streptococcus pneumoniae. | Employee error | The recalling firm made phone calls to their customers beginning 6/11/14 with follow-up letters dated 6/12/14 being issued via regular mail requesting the affected product be destroyed. Additional customers were later identified and additional phone calls were made on 6/16/14 and recall letters dated 6/16/14 were issued via regular mail. |  | 76 | 76/50-bottle boxes | Distribution was made to CO, IL, IN, NE, NJ, PA, SD, and a government facility located in VA. | Terminated | August 28, 2014 | 78 | 76 | Not_Computer
Z-2058-2014 | 68595 | No | CIZ | Androstenedione test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Coat-A-Count Direct Androstenedione, SMN 10381049, Catalog Number TKAN1. Siemens Healthcare Diagnostics |  |  | June 09, 2014 | 2014 | Siemens Healthcare Diagnostics Inc. | The assay is demonstrating a higher frequency of results greater than 10 ng/mL compared to alternate platforms. | Under Investigation by firm | A recall letter dated 6/9/14 was sent to all customers who purchased the Coat-A-Count Direct Androstenedione because the assay is demonstrating a higher frequency of results compared to alternate platforms. The letter informs the customers of the problems identified, risks to health, and actions to be taken. Customers are instructed to complete and return the Field Correction Effectiveness Check attached to the recall letter within 30 days. |  | 1164 | 1,164 units | Worldwide Distribution. US nationwide including CA, MI, NC, NE, NY, OH, PA, TX, and VA; Argentina, Austria, Belgium, Bolivia, Canada, Chile, Colombia, Finland, Germany, Great Britain, Greece, India, Ireland, Italy, Japan, Korea, Netherlands, Paraguay, Peru, South Africa, Spain, Sweden, Switzerland, and Uruguay | Terminated | July 24, 2015 | 410 | 1164 | Not_Computer
Z-2067-2014 | 68781 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Excella II Standard Pedicle Screw, Catalog No. E2S47525 and E2S47535, 4.75 mm x 25 mm. For spinal fixation. Packaged in heat-sealed plastic bag. Also provided in reusable steam sterilization tray as implant set. |  |  | July 07, 2014 | 2014 | Innovasis, Inc | Innovasis is recalling the Excella II Standard Pedicle Screw due to mislabeling of the size and catalog number. | Process control | Distributors were notified by e-mail on 7/7/14, with attached copies of the written letter and forms to be printed, signed and returned to Innovasis. Distributors have been instructed to return all inventory in stock. |  | 32 | 32 | Distribution US nationwide, including TX, CA, KS, MN, and OK. Foreign distribution to Japan. | Terminated | August 11, 2014 | 35 | 32 | Not_Computer
Z-2070-2014 | 68573 | No | LHI | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Excelsior Medical Sterile Triple Lead Tubing Set and Sterile Heavy Duty Triple Lead Tubing Set. To be used with the Pharm-Assist Dispensing Pump. Used for drug reconstitution and/or fluid transfer |  |  | June 05, 2014 | 2014 | Excelsior Medical Corp | There is a potential for a leak. A molding defect in the connector that joins the three leads to the main flow line may allow fluid to escape or air to enter the fluid path during transfer of fluids from one container to another. | Nonconforming Material/Component | Excelsior Medical issued "Urgent Medical Device Recall" notifications/"Customer Reply Forms" dated June 5, 2014 to affected customers via certified mail (return receipt request). The notice informed customers of the issue with the affected product; how to identify affected product; instructed the customers to discontinue use of tubing sets from the affected lots and return the attached form (Customer Reply Form) via fax to Excelsior to make arrangements for return of product. If product was further distributed by the customer the notification instructs them to identify the customer(s) that it was supplied to and notify them at once of the product recall. A phone number was provided in case customers had questions regarding the recall (1.800.487.4276 8:30 am - 5:00 pm EST) as well as an email address (recall12@excelsiormedical.com). |  | 331500 | 4250 tubing sets (PA03 - 78 cases; BX03 - 92 cases (each case contains 25 tubing sets)) | Worldwide Distribution - USA (nationwide) and Canada. | Terminated | August 19, 2015 | 440 | 331500 | Not_Computer
Z-2069-2014 | 68480 | No | EJK | Calcium hydroxide cavity liner. | Dental | Dental | 510(k) | Calcium Hydroxide Powder, USP |  |  | June 05, 2014 | 2014 | Axcentria Pharmaceuticals, LLC | One lot of Calcium Hydroxide Powder is being recalled because it was manufactured and distributed without an approved 510k, and the current label does not include instructions for use. | No Marketing Application | Axcentria Pharmaceuticals sent an urgent medical device recall notice, dated June 5, 2014, was sent to customers and described the product, problem, and actions to be taken. Customers were asked to fill out and return the form included.For questions customers were instructed to call 215-453-5055, ext 317. |  | 528 | 528 | US distribution to NY and CA | Open, Classified |  |  | 528 | Not_Computer
Z-2060-2014 | 68578 | No | GKP | Coagulation instrument. | Hematology | Hematology | 510(k) | ACL TOP (Base), automated coagulation laboratory instrument. |  |  | June 03, 2014 | 2014 | Instrumentation Laboratory Co. | Potential for sample misidentification. | Device Design | Consignees were notified via phone starting on June 3, 2014. On June 4, a USPS mailing was sent to all US/Canadian customers. On June 4, 2014, email notifications were sent to the IL Affiliates and International Distributors/Dealers (through the International Group) requesting them to translate the Field Safety Notice and provide notification to their customer base and secondary distributors.The Urgent Product Correction letter informs consignees of the issue and provides workaround instructions. Consignees are asked to complete a Mandatory Response Tracking Record and return it via Fax (781)861-4207 or email: ra-usa@ilww.com. For technical questions, consignees should contact the IL Technical Support Center at 1-800-678-0710. |  | 1074 | 1074 | Worldwide Distribution-USA (nationwide) including the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Canada, Mexico, Argentina, Australia, Austria, Belarus, Belgium, Bolivia, Brazil, Brunei, Chile, China, Costa Rica, Croatia, Czech Republic, Denmark, Dominican Republic, El Salvador, Finland, France, Germany, Gibraltar, Great Britain, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Korea, Kuwait, Libya, Lithuania, Luxembourg, Malaysia, Mexico, Netherlands, New Zealand, Norway, Panama, Peru, Philippines, Poland, Portugal, Quatar, Romania, Russia, Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, UAE, Uruguay, Venezuela, Vietnam. | Terminated | June 11, 2015 | 373 | 6647 | Not_Computer
Z-2059-2014 | 68582 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | V-Series Patient Monitoring System, Manufactured for Mindray DS USA, Inc. Used for monitoring of human physiological parameters. The V-Series Monitor also has the capability of performing IV Drug and Hemodynamic Calculations and interfacing with network devices. |  |  | May 19, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | ECG interpretation feature issues. False measurements reported when the algorithm cannot determine P wave or QRS axis. The Median Complex Report displays the representative beat 25% larger than actual signal input. When a patient date of birth is imported from an external source, the ECG interpretation algorithm uses the default age of 50, and interpretive statements may be incorrect. | Software design | Mindray issued "Urgent: V-Series Monitoring System Corrective Action" notices dated May 19, 2014, to customers that received the affected product via certified mail with return receipt verification. The notification informs the customer of the issues with the affected product and how to identify the affected product. The notification recommends that customers discontinue use of the 12-lead ECG interpretation feature until a software correction is available (anticipate August 2014). A point of contact is provided to customers in case they have questions.For further questions please call 1-800-288-2121. |  | 113 | 113 units (57 United States; 56 International) | Worldwide Distribution - US including the states of Louisiana, New Jersey, Montana, South Carolina and Wyoming. | Open, Classified |  |  | 113 | Computer
Z-2071-2014 | 68353 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens RAPIDCommÂż Data Management System. |  |  | May 22, 2014 | 2014 | Siemens Healthcare Diagnostics Inc | RAPIDCommÂż 5.0 Screen Layout will display incorrect test names | Software design | Siemens Healthcare issued an Urgent Field Safety Notice by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on May 22, 2014, both in the United States and elsewhere for communication with affected customers. This notice informs customers of the issue with the RAPIDCommÂż Data Management System (Version 5.0.A software correction will be issued to address the problem.Please retain this letter with your laboratory records. Complete and return the Field Correction Effectiveness Check attached to this letter within 30 days.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative.For further questions please call (781) 269-3000. |  | 6 | 6 systems | Worldwide Distribution - US including the states of OR and PA., and the countries of Denmark and UK. | Terminated | May 26, 2015 | 369 | 6 | Computer
Z-2073-2014 | 68674 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LP,Surgical Fibers, Bare Laser Fiber 600 micron, DBLF-60-2, Sterile EO. |  |  | June 20, 2014 | 2014 | Laser Peripherals LLC | Laser Peripherals LLC is recalling DBLF-60-2 Bare Laser due to a mislabeling that indicated a sterility expiration date as '2016-16' as opposed to the intended date of '2016-06'. | Labeling False and Misleading | Laser Peripherals contacted all consignees via telephone on June 20, 2014, in regard to the product being recalled by following the firm's Standard Operating Procedures. The SOP's included Instructions to immediately stop the use of the recalled product and provided other communications related to the recall.For questions regarding this recall call 612-280-1318. |  | 25 | 25 | Nationwide Distribution to IL, IN, and MD | Terminated | October 22, 2014 | 124 | 25 | Not_Computer
Z-2072-2014 | 68742 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | Integra Camino 110-4 Intracranial Pressure (ICP) Monitoring Catheters, |  |  | July 09, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences is recalling the Camino Intracranial Pressure Monitoring Kits because of revised labelling to indicate they are unsafe in an MR environment. | No Marketing Application | A customer notification letter dated 7/9/14 was sent to customers to inform them that Integra LifeSciences has recently revised the labeling of certain Camino Intracranial Pressure Monitoring Kits to indicate they are UNSAFE in an MR environment. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the form attached the letter and return it by email or fax. Customers with questions are instructed to contact Customer Service at (855) 532-1723. |  | 28752 | 28,752 units | Algeria, Austria, Belgium, Bulgaria, Cyprus, Egypt, Finland, France, Germany, Great Britain, Greece, Hungary, Greece, Hungary, Ireland, Israel, Italy, Jordan, Lebanon, Lithuania, Luxemborg, Netherlands, Poland, Portugal, Romania, Saudi Arabia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tunisia, Turkey, United Arab Emirates, Argentina, Brazil, Brunei, Chile, China, India, Japan, Korea, Mexico. | Terminated | October 14, 2014 | 97 | 28752 | Not_Computer
Z-2077-2014 | 68611 | No | CBR | Nitrous oxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Spacelabs Healthcare Ultraview SL Multigas Module, Model 92518 (Software V1.00.00-09). Provides a means of monitoring a variety of gas concentrations and alert clinical personnel when the concentration of anesthetic agent(s), oxygen, carbon dioxide or nitrous oxide moves outside of user-defined limits. |  |  | June 18, 2014 | 2014 | Spacelabs Healthcare Inc | With this software version, the Minimum Alveolar Concentration (MAC) value displayed is extremely high (20.1) and physiologically invalid. | Process control | Spacelabs sent an Urgent Medical Device Correction letter dated June 20, 20/14 to all affected customers. The letter identified the affected product, problem and the actions to be taken. Customers are informed that Spacelabs Field Service personnel will be contacting their facility to schedule a convenient time to confirm, at no cost, a software update is needed to resolve this issue. Customers with questions are instructed to contact SpacelabsHealthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 8 | 8 units total (3 units in US) | Worldwide Distribution - US in the state of Georgia and in the countries of : Malaysia, Paraguay, and France. | Terminated | December 08, 2014 | 173 | 8 | Computer
Z-2078-2014 | 68609 | No | EFA | Dental handpiece and accessories. | Dental | Dental | 510(k) | Titan T 5K Motor |  |  | June 13, 2014 | 2014 | Dental EZ Group Star Dental Division | The housing of 5K motor was laser marked incorrectly as a 20K motor.The recall was expanded in September 2014 to include additional lot #2251181. | Process control | Star Dental sent a Medical Device Recall letter, dated June 13, 2014 to affected customers. This letter identified the affected product, problem, and actions to be taken. For question call 866-383-4636 extension 4350.The recall was expanded in September 2014 to include additional lot #2251181. A letter, dated September 2014 was sent to customers who were impacted by the recall expansion. |  | 45 | 45 | US Distribution in the states of : IN, TN, WA, CA, UT, MN, and NY. | Terminated | February 24, 2015 | 256 | 45 | Not_Computer
Z-2075-2014 | 68395 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens syngo RT Dosimetrist 2.7 system |  |  | May 15, 2014 | 2014 | Siemens Medical Solutions USA, Inc | An update to the Virtual Simulation software of the syngo RT Dosimetrist is necessary to address a safety issue. In special scenarios within the Structure Operations function the system did not store the selected name from the list of predefined names. After reloading of the data, this structure name had been changed by the system. Additionally, a performance issue was addressed where import of DI | Software design | Siemens sent a Medical Device Correction Field Safety Notice of Improvement letter to affected customers via USPS Certified Mail. The letter described the affected product, problem, and actions to be taken. |  | 7 | 7 | US Nationwide Distribution in the states of NH, TN, OH, VA, and TX. | Terminated | December 15, 2014 | 214 | 7 | Computer
Z-2074-2014 | 68590 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter MiniCap with Povidone-Iodine Solution, 5C4466P, Baxter Healthcare Corporation. For dialysis. |  |  | June 17, 2014 | 2014 | Baxter Healthcare Corp. | Inadequate iodine and packaging related defects. | Packaging process control | The firm notified their direct consignees on 06/17/2014 by letter requesting customers to locate and remove all affected product from their facility, and contact Baxter Healthcare Center for Service to arrange for return and credit. |  | 585360 | 585,360 units | Distributed USA (nationwide) and Canada. | Terminated | April 03, 2015 | 290 | 585360 | Not_Computer
Z-2076-2014 | 68619 | No | JDQ | Bone fixation cerclage. | Orthopedic | Orthopedic | 510(k) | Tritium Sternal Plate System, Screw, 02.7mm X 12mm, 12 Pack, Sterile. |  |  | June 11, 2014 | 2014 | PIONEER SURGICAL TECHNOLOGY, INC. | Lack of Sterility Assurance | Employee error | On 6/11/14 and 6/12/14 the firm contacted their two consignees and surgeon via phone and explained that the product did not go through the sterilization process. The returns authorization number and shipping labels were discussed for the unused unit. On 6/18/14 and 6/20/14 an URGENT: MEDICAL DEVICE VOLUNTARY RECALL NOTIFICATION was sent to the customers as well as the one surgeon. The letter explained the sterility concerns and that the productis being recalled. Contact Dan Nelson, Manager of Quality, at 906-226-4489 if you have any questions or concerns. |  | 5 | 5 | US Nationwide Distribution in the states of TX and FL | Terminated | October 22, 2014 | 133 | 5 | Not_Computer
Z-2068-2014 | 68344 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | DATEX-OHMEDA S/5 SINGLE-WIDTH AIRWAY MODULE, E-MINIC AND ACCESSORIES. |  |  | June 11, 2014 | 2014 | GE Healthcare, LLC | Failure of the CO2 detector in Single-width Airway and Extension modules, which may cause a slow continuous decrease of measured Et-/FiCO2 values. Incorrect Et CO2/Fi CO2 value may impair clinical decision making for both mechanically and spontaneously ventilated patients because of incorrect low CO2 values. | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated June 11, 2014, to all affected consignees. The letter described the Safety Issue, Safety Instructions, Affected Product Details, and Product Correction and Contact Information. For questions customers were instructed to contact Technical Support at 1-800-558-7044 or their local Service Representative. |  | 2948 | 2,948 (761 units US, 2,187 units OUS) | Worldwide Distribution - USA nationwide, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BRAZIL, CANADA, CHILE, CHINA, Croatia, CZECH REPUBLIC, DENMARK, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, GREECE, GUATEMALA, HONG KONG, INDIA, IRAQ, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, REPUBLIC OF KOREA, MALAYSIA, NETHERLANDS, NEW ZEALAND, OMAN, PERU, POLAND, ROMANIA, RUSSIAN FEDERATION, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, SWITZERLAND, SYRIAN ARAB REPUBLIC, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VENEZUELA, LEBANON, REPUBLIC OF MALDOVIA, COSTA RICA, GABON, INDONESIA, JAMAICA, NIGERIA, NORWAY, NEPAL, PAKISTAN, PANAMA, PHILIPPINES, SENEGAL, SOUTH AFRICA, THAILAND, TUNISIA, and CAMBODIA. | Terminated | March 29, 2016 | 657 | 2948 | Not_Computer
Z-2079-2014 | 68437 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | RAPIDPoint 500 Measurement Clinical Chemistry System |  |  | May 23, 2014 | 2014 | Siemens Healthcare Diagnostics Inc | RAPIDPoint 500 Measurement Cartridges may have an error code one or more electrolyte parameters (e.g., Na+, Ca++, K+, Cl-) | Process design | Siemens Healthcare Diagnostics issued an Urgent Field Safety Notice by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on May 27, 2014 both in the United States and elsewhere for communication with affected customers. This notice informs customers of the issue with the RAPIDPoint 500 Measurement cartridges. Questions contact: Siemens Customer Care Center or your local Siemens technical support representative at 877-229-3711. What steps can be taken to mitigate the issue:" Install a new RAPIDPoint 500 Measurement Cartridge.or" Install a 405 Measurement Cartridge - The RAPIDPoint 405 cartridge is available in four sizes:Since the Lactate parameter is not available on the RAPIDPoint 405 cartridge, however, you will need to arrange to get your lactate results from other instruments in your central laboratory.or" Use an alternative system to obtain your electrolyte values . |  | 20507 | 20,507 cartridges | NationwideForeign:U.A.E.AlbaniaAustriaAustraliaBosnia Herzeg.BelgiumBulgariaBahrainBrazilCanadaSwitzerlandP.R. ChinaColombiaCzech RepublicGermanyDenmarkAlgeriaEstoniaEgyptSpainFinlandFranceUnited KingdomGeorgiaGreeceHong KongCroatiaHungaryIndonesiaIrelandIsraelIndiaItalyJapanRepublic KoreaKuwaitLithuaniaLuxembourgLatviaMalaysiaMexicoNetherlandsNorwayNew ZealandPeruPolandPortugalQatarRomaniaSerbiaRussian Fed.Saudi ArabiaSwedenSingaporeSloveniaTurkeyTaiwan | Terminated | June 10, 2015 | 383 | 20507 | Not_Computer
Z-2080-2014 | 68624 | No | LJH | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Disposable StrykeFlow Tip, Stryker Disposable Suction/I irrigator Tip; Sterile, Rx only, Latex Free. Stryker Suction Irrigation system, used to provide a clean and clear operative site through the evacuation and aspiration of irrigant solution. for endoscopic use. |  |  | June 20, 2014 | 2014 | Stryker Endoscopy | Internal review found that six lots of Disposable StrykeFlow Tips were shipped with a potentially deformed blister pouch which could impact sterility of the product. | Other | Stryker sent an Urgent Device Removal letter via certified mail on June 20, 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to return any unused product, inform individuals within their organization who need to aware of the device removal. Customers with questions were instructed to call 1-800-624-4422.For questions regarding this recall call 408-754-2664. |  | 496 | 496 boxes (6 tips per box) | Worldwide Distribution - USA (nationwide) and Internationally tAustralia, Canada, France, Greece, Netherlands, United kingdom and South Korea. | Terminated | October 15, 2014 | 117 | 496 | Not_Computer
Z-2087-2014 | 68442 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Guide Blocks for the 2 Column Plate 6 Hole Head (Volar Distal Radius Plate System). Orthopedic surgical instrument. |  |  | May 29, 2014 | 2014 | Synthes, Inc. | The Guide Blocks for the 2 Column Plate 6 Hole Head (Volar Distal Radius Plate System) were found to be labeled incorrectly. Part no. 03.111.600 (Right) was found in the package labeled part no. 03.111.601 (Left). The laser etching denoting the orientation on the part is correct. | Process control | An urgent field correction notification, dated May 29, 2014, was sent to end users and sales consultants which described the product, problem, and actions to be taken. |  | 180 | 180 | Distribution US Nationwide and Bermuda. | Open, Classified |  |  | 180 | Not_Computer
Z-2088-2014 | 68659 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Hemochron Jr. Whole Blood Microcoagulation System APTT Cuvette. For in vitro diagnostic coagulation test. Catalog number J103. |  |  | June 20, 2014 | 2014 | ITC-Nexus Dx | Test results may be erroneously elevated or report an out-of-range high (OOR-H) error code in non heparinized blood samples. | Device Design | ITC issued an Urgent Medical Device Recall Letters/Response Forms dated June 20, 2014, to End-users and US Distributors were notified. ITC asks that this communication be forwarded to all those within your organization who need to be aware of this matter. If any affected products have been forwarded to another facility, a copy of this communication should be forwarded to them immediately.Customers should check their inventory for the affected lot. The affected lot should not be used and quarantined until return to ITC. The attached form should be completed and returned to ITC within 10 days. ITC will review the form and contact the customer to coordinate any returns. For questions, please contact Technical Support: Telephone 732-548-5700, Extension 4011 E-mail: Techsupport@itcmed.comNormal business hours are Monday through Friday, 8:00 AM to 6:30 PM, Eastern Time.For urgent matters, Technical Support is available 24 hours per day, 7 days a week. |  | 3150 | US- 315 cuvettes; OUS - 2835 cuvettes | Worldwide Distribution - US including the states of Georgia, Virginia and New York., and the countries of Canada, France, Italy, Liechtenstein and Romania. | Terminated | December 11, 2014 | 174 | 3150 | Not_Computer
Z-2085-2014 | 68569 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B850, B650 or B450. Patient Monitor. |  |  | June 13, 2014 | 2014 | GE Healthcare, LLC | Partial loss of monitoring with the CARESCAPE Monitor B850, B650 or B450. When connecting the PDM (Patient Data Module) to the CARESCAPE B850, B650 or B450 monitor with software version 2 the PDM may not start communicating to the monitor and loss of the PDM parameters could occur. | Software design | Consignees were sent on 6/13/2014 a GE Healthcare "Urgent Medical Device Correction" letter #36103, 36104 and 36105 dated June 6, 2014. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Products Details, Product Correction and Contact Information. Consignees were advised to contact Technical Support at 800-558-7044 in regard to questions on this recall. |  | 41992 | 41,992 (20,509 US; 21,483 OUS) | Worldwide Distribution: US nationwide, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SERBIA, SAUDI ARABIA, RUSSIAN FEDERATION, ROMANIA, REUNION, QATAR, PORTUGAL, POLAND, PHILIPPINEs, PARAGUAY, PANAMA, NORWAY, NEW ZEALAND, NETHERLANDS, MYANMAR, MOROCCO, MEXICO, MALAYSIA, LITHUANIA, LEBANON, LATVIA, JORDAN, JAPAN, JAMAICA, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, ICELAND, HUNGARY, GREECE, GIBRALTAR, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT, ECUADOR, DENMARK, CZECH REPUBLIC, CYPRUS, CROATIA, CHILE, BRUNEI, BERMUDA, BELGIUM, BANGLADESH, AUSTRIA, AUSTRALIA. | Open, Classified |  |  | 41992 | Computer
Z-2081-2014 | 68567 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | AGFA Digital Radiography X-Ray system DX-D 100 |  |  | June 06, 2014 | 2014 | AGFA Corp. | When liquid comes in contact with the DX-D 100 touch screen, the device may incorrectly recognize this as user input altering device settings. | Use error | Agfa Healthcare sent an "URGENT FIELD SAFETY NOTICE" dated June 6, 2014, to the US and Canadian customers. The letter described the safety alert and mitigation. Acknowledgment, via FAX-Back or email, that the information was received and understood was requested from the customers.Although medical staff are required to disinfect hands on a frequent basis, users must ensure that their hands are dry before using the touch screen of the DX-0100 because liquid residue may activate the action buttons on the touch screen. Customers were instructed:"Do not operate touch screen monitor with wet hands!"Do not let liquids come in contact with the touch screen while the DXD 100 is powered on!"Always double check your parameter settings prior to exposing the patient. We would like to remind you of this specific warning statement in the DX-0100 User Manual.Please complete the feedback form as soon as possible and return it to us by June 16, 2014.Should the above information not apply to your facility or should the device have been transferred to another organization, please be so kind as to indicate this on the attached feedback form and pass this Urgent Field Safety Notice to the organization where the device has been transferred.We thank you for your careful attention to this issue and your continued support.If you have any questions about this matter, please contact your local Agfa Healthcare organization: 1-877-777-2432, prompt 1 and reference PR1406020002 |  | 216 | 216 units | USA (nationwide) including the states of AL, CA, FL, GA, IL, IN, IA, KY, LA, MA, MI, MN, MS, MO, NE, NJ, NY, NC, OH, SC, TN, TX, VA, WV, WI, WY and Internationally to Canada. | Open, Classified |  |  | 216 | Not_Computer
Z-2084-2014 | 68440 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | RECLAIM DISTAL TAPERED Distal Stem, various sizes. Taper Protector Sleeve. RECLAIM Revision Hip System Orthopedic Manual Surgical Instrument, used to ensure the locking taper surface of the Distal Stem is protected during proximal preparation. |  |  | June 03, 2014 | 2014 | DePuy Orthopaedics, Inc. | The product can be difficult to remove from the Distal Stem both out of the package andafter proximal reaming. | Device Design | DePuy Synthes sent an Urgent Information - Device Correction Notice letter dated June 24, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:1. Immediate Remedy: Discard the RECLAIMÂż Taper Sleeve Protector instrument before stem implantation.2. Changes to the literature and IFU: Documents are revised as follows:" RECLAIMÂż Revision Hip Systems Surgical Technique # DSUS/JRC/0614/0200  pages 11, 12, 13, 16,30, 34, 35 (figure numbering from pages 11 to 27 were also revised): Removed information regardingtaper sleeve protector." RECLAIMÂż Revision Hip Systems Surgical Technique FAQ # DSUS/JRC/0614/0202  pages 1 and 2:Removed information regarding taper sleeve protector." RECLAIMÂż Revision Hip Systems IFU #090200799 REV E  Page 5: Verbiage will be revised andimplemented as soon as possible, depending on regulatory approvals.3. Change: The company consulted with design surgeons and developed a strategy to eliminate the taper sleeve protector from the system. As a result of removing the taper sleeve protector, the company isevaluating whether there is a need for a protective component to be used during the distributionCustomers with questions were instructed to call 574-371-4917.For questions regarding this recall call 574-371-4577. |  | 3623 | 3,623 | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI &amp; WV and Internationally to AUSTRALIA, AUSTRIA, CANADA, CHILE, CHINA, CZECH REPUBLIC, DENMARK, FRANCE, GERMANY, IBERIA, IRELAND, ISRAEL, ITALY, NEW ZEALAND, NORWAY, POLAND, SINGAPORE, SLOVENIA, SOUTH AFRICA, SPAIN, SWITZERLAND &amp; UK. | Terminated | May 17, 2017 | 1079 | 14904 | Not_Computer
Z-2086-2014 | 68341 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NOVATION Femoral Stem, Tapered Extended offset, plasma- coated, press fit. Catalog number 160-01-13 |  |  | March 07, 2014 | 2014 | Exactech, Inc. | The femoral stem is lacking specified plasma coating. | Component design/selection | Exactech sent an Important Product Recall Notice dated March 10, 2014, to all affected customers. The notice identified the product the problem and the action to be taken by the customer. Description of Issue: The femoral stem is lacking specified plasma spray coating.Clinical Impact: Without plasma spray coating, the stem may not remain stable and will not have the possibility for biological fixation.In order to comply with applicable regulations and Exactech policies customers were instructed to :" Immediately cease distribution or use of these products." Extend this information to your accounts that may have this product in their possession." Determine if they have any of the affected devices " Fax back the attached form. Complete the attached fax back form and return it to Exactech.Customers were instructed to contact their Exactech inventory representative within 5 business days to confirm quantities at their location. For any inventory restocking questions related to this issue, please call 1-800-392-2832.For questions regarding this recall call 800-392-2832. |  | 10 | 10 | Worldwide Distribution including Spain, Greece and Japan | Terminated | August 04, 2014 | 150 | 10 | Not_Computer
Z-2089-2014 | 68795 | No | MAV | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | Perouse Flamingo, Model 0218TA and 0218TB intended for use during cardiovascular procedure to create, maintain and monitor pressure in the balloon catheter. |  |  | May 10, 2012 | 2012 | Perouse Medical | Perouse Medical in France initiated a voluntary recall of certain batches of Perouse FLAMINGO Inflations Devices due to a potential water-tightness default. The potential default could cause an air entry when the plunger is positioned around 20 mL. | Process control | Perouse Medical initiated this recall by sending an e-mail notification to consignees on May 10, 2012. A formal written recall notification letter was sent to consignees on May 16, 2012. The recall letter, written in English, on company's letterhead, dated May 16, 2012, titled "VOLUNTARY BATCH RECALL Of FLAMINGO Inflations Devices", notified the consignee of the following: product information with codes and batch codes, product description, description of the potential default, reason for the recall action, instructions to the consignee, contact information and assistance. Response form "INVENTORY FORM". |  | 2496 | 2,496 units | US. | Terminated | September 09, 2014 | 852 | 2496 | Not_Computer
Z-2090-2014 | 68718 | No | DRM | External cardiac compressor. | Cardiovascular | Cardiovascular | 510(k) | Zoll Circulation LifeBand Single Pack (accessory), and Nickel-metal hydride (NIMH) Battery. |  |  | July 16, 2013 | 2013 | Zoll Circulation, Inc. | Typographical error in the Japanese legal label of the AutoPulse Resuscitation System Model 100. The label is specific to the Japanese version of the product. | Labeling Change Control | Zoll Japan handled the recall. |  | 201 | 161 Nickel Metal batteries, 40 LifeBand | Distributed in Japan. | Terminated | July 22, 2014 | 371 | 201 | Computer
Z-2021-2014 | 68196 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVOCATE Redi-Code+ BMB-BA006A Blood Glucose test strips |  |  | May 05, 2014 | 2014 | DIABETIC SUPPLY OF SUNCOAST, INC. | "Advocate Redi-Code + BMB-BA006A Blood Glucose test strips" are inappropriately labeled and marketed as "Advocate Redi-Code + Blood Glucose test strips". | Labeling False and Misleading | Diabetic Supply of Suncoast, Inc. sent an Urgent Voluntary Medical Device Recall letter dated June 6, 2014, to all affected consignees. The firm is asking the distributors to examine their inventory immediately and quarantine any product subject to recall. In addition, if they may have further distributed this product, to please identify their customers and notify them at once of this product recall. Customers who have the BMB-BA006A Advocate Redi-Code+ test strips can contact Diabetic Supply of Suncoast, Inc. directly for further information to determine if the products are being used in the proper manner or if they need to be returned to Suncoast for replacement. Test strips will be replaced with relabeled (with new labels) boxes of test strips to avoid any chance of confusion. Consumers should stop using the above mentioned product and contact Suncoast to find out if this recall pertains to the products they have in their possession and if so how to have them replaced for the relabeled product manufactured by BroadMaster Bio-Tech. Customers with questions should call Diabetic Supply of Suncoast, Inc. at 561-296-0488.For questions regarding this recall call 561-254-1007. |  | 1852800 | 1,852,800 cases (50 strips ea) | Worldwide Distribution - USA (nationwide) and Internationally to Venezuela, Ecuador, Canada &amp; Vietnam | Terminated | June 13, 2017 | 1135 | 1852800 | Not_Computer
Z-2091-2014 | 68622 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | MK5 NX Controller--Wheelchair controller |  |  | June 24, 2014 | 2014 | Invacare Corporation | M41 Powered Wheelchairs received through Invacare Service Parts (ISP) have controllers that were programmed in a way this is not aligned to standard factory settings. The controller during different settings such as max forward speed, max forward acceleration, max reverse speed and max turning speed can result in momentary loss of control. | Process control | The controllers will be reporogrammed in the field upon installationInvacare Sent out a Technical Services Bulletin (Oct. 15, 2012) . Please call 800-832-4707, Monday-Friday, 8:00 am- 7:00 pm EST, for assistance. If you are aware of a consumer complaint regarding this controller, please contact the Invacare Customer Service Complaints Department at 877-374-8919, Monday-Friday, 8:00am- 7:00pm or cscomplaints@invacare.com |  | 819 | 819 | Nationwide Distriution: CA, SC, IL, UT, TN, KS, HI, NJ, LA, TX,FL, IN, AL, AZ, MN, WI, MN, MI, AR,MA, GA, MS, NJ, NM, OK, ME, PA, NY, NC, OK, MO, WA, IA, OR,KY, MD, NE, ID, VA, WA, ND, MT, OH, NH, CT. Also shipped to Canada | Terminated | October 15, 2014 | 113 | 819 | Not_Computer
Z-2092-2014 | 68667 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Solanas 2.1mm FIXED DEPTH DRILL, 10mm, Part No. 63928-10. Component of the SOLANAS III Posterior Stabilization System, a spinal fixation system intended to improve stability of the cervical and thoracolumbar area of the spine (C1-T3). |  |  | June 26, 2014 | 2014 | Alphatec Spine, Inc. | Alphatec Spine is recalling the 2.1mm Fixed Depth Drill 10mm because they received a report that the instrument malfunctioned and golden stop slid from its fixed position during use. | Unknown/Undetermined by firm | Alphatec Spine, Inc. sent an Urgent Medical Device Notification letter dated June 20, 2014, June 26, 2014, to all affected customers. Alphatec Spine received a report of the 10mm fixed depth drill from this lot which malfunctioned and golden stop slid from its fixed position during use. The letter informs the customers of the problems identifed and the actions to be taken. Customers are instructed to confirm that they received the notification letter by completing the customer response form. Completed forms are to be sent to Alphatec Spine at 5818 El Camino Rd, Carlsbad, CA 92008. Customers may also return forms by fax at (760) 431-0289. Customers with questions were instructed to contact Alphatec Spine's Customer Service Department on how to return the affected product.For questions regarding this recall call 760-431-9286. |  | 14 | 14 units total (12 units in US) | Nationwide Distribution including CT, NC, AL, WV, IL, AR, and NY. | Terminated | September 03, 2014 | 69 | 14 | Not_Computer
Z-2111-2014 | 68769 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Siemens syngo Imaging XS, version VA70A or higher |  |  | June 30, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens has become aware of the potential malfunction when using syngo Imaging XS, version VA70A or higher. In some cases a miscalculation may occur for the grey scale values in the functions Region of Interest , Pixel Lens , Edge enhancement and Histograms . This potential malfunction may occur for images with a pixel depth greater than 12 bit. All values calculated for images with a pixel d | Software design | Siemens sent a Customer letter dated June 30th, to all affected customers to inform them of a new software version VA70B_0114 that will be introduced to all syngo Imaging XS customers with version VA70A or higher to solve a potential issue regarding miscalculations done for the grey scale values in the functions "Region of Interest", "Pixel Lens", "Edge Enhancement", and "Histograms".For questions regarding this recall call 610-219-6300. |  | 141 | 141 | Nationwide Distribution | Terminated | February 24, 2015 | 239 | 141 | Computer
Z-2110-2014 | 68756 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | QLAB 10.1 Core Module PN 453561728001. QLAB Quantification Software is a software application package designed to view and quantify image data acquired on Philips Medical Systems ultrasound products. |  |  | July 09, 2014 | 2014 | Philips Ultrasound, Inc. | The QLAB a2DQ and aCMQ features have a defect that may use incorrect measurement values under specific conditions. | Software change control | Philips Medical Systems notified customers of the voluntary recall by sending a certified URGENT-Medical Device Correction letter, dated 2014 JUL 09, to all US and Canadian consignees. Philips representatives will distribute the letter internationally. Follow-up letters to wholesale customers will be provided and wholesalers are requested to forward copies of recall letters and response forms to their customers. Philips Medical Systems will also have Customer Reply Cards returned for tracking purposes. Customers are directed to call local Philips representative or Philips Customer Service at 1-866-767-7822. |  | 59 | 59 units total (12 units in the US and 47 units outside the US) | Distributed in the states of CA, IA, KY, MD, NY, PA, UT, VA, WA and WI.Also distributed in Australia, Belgium, Canada, India, Italy, Japan, South Korea, Netherlands, New Zealand, Spain, Sweden, and Switzerland. | Terminated | November 06, 2015 | 485 | 59 | Not_Computer
Z-2097-2014 | 68404 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Sarns Malleable Venous Return Cannulae 24, 28, 30, 32, 34, 36 &amp; 40 Fr single stage with 3/8" &amp; 1/2 flare with or without Xcoating |  |  | May 23, 2014 | 2014 | Terumo Cardiovascular Systems Corporation | During an in-process inspection, Terumo Cardiovascular Systems (Terumo CVS) identified the presence ofloose fiber particulate that exceeded finished product specifications | Process control | Terumo sent an Urgent Medical Device Recall letters dated June 27, 2014 to their customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to 1. Review this Medical Device Recall notice, 2. Assure that all users in your institution have received notice of this issue and remove any affected units from use, 3. Fill out the Product Tracking Sheet (enclosed), indicating the number of affected units at your institution and fax to 1-877-907-7509 or email to terumo4967@stericycle.com. 4. Complete the Packing Slip (enclosed) and follow the packing instructions to return the affected product, and 5. Terumo CVS will issue credit once the affected units are returned. For questions call Terumo CVS c/o Stericycle 1-866-629-6181. |  | 32880 | 32,880 units | Worldwide Distribution - US (nationwide) in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV &amp; WY and the countries of : Australia, Belgium, Canada, Chile, Japan, Malaysia, Mexico, Singapore, Taiwan, Thailand, United Arab Emirates (UAE) &amp; United Kingdom. | Terminated | February 24, 2015 | 277 | 376285 | Not_Computer
Z-2093-2014 | 68701 | No | BYG | Oxygen mask. | Anesthesiology | Anesthesiology | 510(K) Exempt | Oxygen Mask, Model No. 8120-7 |  |  | June 25, 2014 | 2014 | Salter Labs | The case label correctly states partial rebreather but the insert label incorrectly states non-rebreather. | Error in labeling | Salter Labs sent a Medical Device Recall letter dated June 25, 2014 to all affected customers. The letter identified the affected product, problems and actions to be taken. Customers are instructed to complete the attached Return Response Form and return it to Salter as soon as possible. Customers are instructed to send completed form to Salter by either Fax (800) 628-4690, attn: Sarah Velasquez, email to SVelasquez@us.salterlabs.com, or mail to Salter Labs, attn: Sarah Velasquez, 100 W. Sycamore Rd. Arvin, CA 93203. |  | 584 | 584 units | Worldwide Distribution - US (nationwide) and the countries of Brazil, United Kingdom, Indonesia, Morocco, Mexico, Norway, Venezuela, and South Africa. | Terminated | September 10, 2014 | 77 | 584 | Not_Computer
Z-2107-2014 | 68731 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCI, Breathing Circuit, Adult, Single Limb w/Heated Wire, [Respiratory Gas Humidifier, (Direct Patient Interface)]. The Hudson Heated Wire Ventilator Circuit is intended as a conduit for respiratory gas between a patient and a ventilator, and includes heated wires for use with a Concha Column Humidifier. |  |  | July 09, 2014 | 2014 | Teleflex Medical | The temperature probe does not properly connect to the temperature port. | Nonconforming Material/Component | Teleflex Medical sent an Urgent Medical Device Recall Notification dated July 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you have received products that are subject to this recall. We are now notifying our customers to take the following actions:1. If you have affected stock, immediately discontinue use and quarantine any products with the product and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866- 804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 11679 | 11,679 units | Worldwide Distribution - US including the states of CA, FL, KS, KY, MA, MI, MO, NC, NE, NJ, NY, OH, PA, TN, TX and VA, and the countries of Argentina, Chile, Colombia, and the Philippines | Open, Classified |  |  | 11679 | Not_Computer
Z-2115-2014 | 68703 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Pad-Pak and Pedi-Pak accessories for the HeartSine Samaritan PAD 300P and 350P device (SAM 300P and 350P). |  |  | June 12, 2014 | 2014 | Heartsine Technologies, Limited | A small number of sealed foil pouches containing the electrodes were found to be difficult to open, resulting in potential delay in therapy. | Packaging process control | HeartSine notified distributors in the United States by e-mail on June 12, 2014 and by follow-up letter on June 13, 2014. Distributors were given a letter to use in communications with customers on June 17, 2014. Customers were instructed to return affected product to there distributors. Distributors were asked to return the affected product to HeartSine. In order to effect an immediate return, replacement Pad-Paks and Pedi-Paks have already been delivered to all respective distributors. A seconde letter, dated June 30th, was sent to end users to clarify where on the label to find lot specific information.For question contact Quality Manager at +44 (0)28 90 93 94 19. |  | 3245 | 3245 | Worldwide Distribution - US (nationwide) and in Canada, Mexico and Latin America. | Terminated | April 16, 2015 | 308 | 3245 | Not_Computer
Z-2114-2014 | 68378 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B650. A multi-parameter patient monitor intended for use in multiple areas and intrahospital transport within a professional healthcare facility. |  |  | July 07, 2014 | 2014 | GE Healthcare, LLC | Possible sound loss associated with the CARESCAPE Monitor B650. When an external display is connected to the CARESCAPE Monitor B650 using a USB extender cable with an active USB hub, the system may experience a sound loss including alarm sounds from the patient monitor. This may occur if the external display is powered off and on during normal operation of the monitor. | Software design | GE Healthcare "Urgent Medical Device Correction" letter #36102 dated July 7, 2014 to affected customers. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions they can contact Technical Support at 1-800-558-7044 or your local Service Representative. |  | 21799 | 21,799 (7,635 units US, 14,164 units OUS) | Worldwide distribution. US nationwide including Puerto Rico, ZAMBIA, YEMEN, VIETNAM, VENEZUELA, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, THAILAND, UNITED REPUBLIC OF TANZANIA, TAJIKISTAN, TAIWAN, SYRIA, SWITZERLAND, SWEDEN, SRILANKA, SPAIN, SOUTH AFRICA, SLOVENIA, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIAN FÂżDÂżRATION, ROMANIA, QATAR, PORTUGAL, POLAND, PHILIPPINES, PERU, PARAGUAY, PANAMA, PAKISTAN, OMAN, NORWAY, NICARAGUA, NEW ZEALAND, NETHERLAND, NAMIBIA, MYANMAR, MOROCCO, MEXICO, MAURITIUS, MALAYSIA, LITHUANIA, LIBYAN ARAB JAMAHIRIYA, LEBANON, LATVIA, KUWAIT, REPUBLIC OF KOREA, KAZAKHSTAN, JORDAN, Italy, JAPAN, JAMAICA, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, ICELAND, HUNGARY, HONG KONG, HONDURAS, GUATEMALA, GREECE, GERMANY, GEORGIA, FRANCE, FINLAND, ESTONIA, EL SALVADOR, Egypt, ECUADOR, DENMARK, CZECH REPUBLIC, CYPRUS, CROATIA, COSTA RICA, COLOMBIA, CHINA, CHILE, CANADA, BULGARIA, BRUNEI, BRAZIL, BOLIVIA, BERMUDA, BELGIUM, BANGLADESH, BAHRAIN, BAHAMAS, AZERBAIJAN, AUSTRIA, AUSTRALIA, ARGENTINA. | Open, Classified |  |  | 21799 | Not_Computer
Z-2113-2014 | 68709 | No | MVC |  |  | Pathology | PMA | InSite HER-2 Detection Kit, an immunohistochemistry assay for in vitro diagnostic use. |  |  | July 01, 2014 | 2014 | Biogenex Laboratories, Inc. | Product was distributed without 510(k) approval. | No Marketing Application | Each customer contacted by phone - neither is in possession of the product. |  | 4 | 4 kits | Brazil and Ecuador. No US distribution | Terminated | July 24, 2014 | 23 | 4 | Not_Computer
Z-2116-2014 | 68586 | No | FDS | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EG-3630U Ultrasound Video Gastroscope |  |  | June 12, 2014 | 2014 | Pentax Medical Company | Aspiration needles used in combination with Ultrasound Gastroscopes can derail from the accessory elevator at the distal end of the endoscope. This may result in patient injury if the Instruction for Use is not followed carefully by the user. | Device Design | Pentax Medical sent an Urgent Medical Device Field Safety Corrective Action dated June 12, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Pentax also posted their notification on their website. For questions regarding this action please feel free to contact Paul Silva, RecallCoordinator atTel: 800-431-5880 ext 2064 (8:30am  5:00pm Mon-Fri, EST)Fax: 201-799-4063Email : paul.silva@pentaxmedical.com |  | 531 | 531 units | US Nationwide Distribution | Terminated | November 17, 2015 | 523 | 531 | Not_Computer
Z-2118-2014 | 68615 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic Nexdrive Micropositioning Drive. Models Ml-1000 and Ml-2000. |  |  | June 11, 2014 | 2014 | Medtronic Neuromodulation | Potential for misalignment of the Z-stage scale. Using one of these devices for a procedure could result in the microelectrode being inserted to an incorrect target depth. | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Recall letter dated June 2014 on June 11, 2014, to all affected customers. The letter described the problem and the product involved in the recall. The letter described the "Action Requested" which included to stop using and return the recalled device and to complete and return the Reply Form. Customers with questions were instructed to contact Medtronic Neuromodulation Technical Services at 1-800-707-0933.EXPANDED 10/02/2014: Medtronic sent an Urgent Medical Device Recall letter beginning 10/02/2014. The letter explained that this letter is an update to the June 2014 action, to include 4 additional lot numbers. The letter explained the reason for recall identified all lots recalled and asked that devices stopped being used. A response form provided instructions on how to return product and the form was asked to be returned within 10 days. Questions should be directed to Medtronic Neuromodulation Technical Service at 1-800-707-0933, weekdays 7 am - 6 pm, Central time. |  | 3000 | 300 EXPANDED 10/02/2014 130 devices | Worldwide Distribution - USA including CA, CO, FL, IL, KS, KY, LA, MN, MT, NC, NE, OR, TX, VA, and UT, and Internationally to Canada, Spain, and United Kingdom. | Terminated | January 22, 2015 | 225 | 3000 | Not_Computer
Z-2117-2014 | 68681 | Yes | CBP | Nonrebreathing valve. | Anesthesiology | Anesthesiology | 510(k) | CPR Shield With One Way Valve and Barrier Filter, Reorder No. 4921 --- Dynarex label, 100 per case --- Manufactured for: Dynarex Corporation - Made in China. |  |  | July 02, 2014 | 2014 | Dynarex Corporation | "THIS SIDE UP" is on the wrong side of the product. This error poses a potential health hazard if the wrong labeling is followed by caregivers. | Employee error | Voluntary Device Recall letters (dated 7/01/14) and recall verification forms were sent to customers via US Postal First Class Mail on 7/02/14. The letters instructed customers to check their inventory for the specific lots, discontinue distributing them, and quarantine them immediately. Customers were to promptly contact Dynarex Credit &amp; Return Representative Avi Celnick at 1-845-365-8200 ext. 6644 to arrange for product return and obtain an RGA #. For questions, please call Dynarex's toll free number 1-888-396-2739. |  | 250 | Domestic: 208 cases; Foreign: 42 cases | Worldwide Distribution -- USA, Canada, and Mexico. | Terminated | April 20, 2017 | 1023 | 250 | Not_Computer
Z-2119-2014 | 68757 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | Emerald Fiberoptic Blade, Macintosh 4, Teleflex Medical. A rigid laryngoscope is a device used to examine and visualize a patient's upper airway and aid placement of a tracheal tube. |  |  | July 09, 2014 | 2014 | Teleflex Medical | Packaging Error: Customers have reported receiving incorrect product. The box labeling indicates the correct product MAC 4 but the contents are incorrect, Miller 4 blades. | Other | Teleflex Medical sent an Urgent Medical Device Recall Notification dated July 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to:1. If you have affected stock, immediately discontinue use and quarantine any products with the product and lot numbers listed above.2. To return product, complete the enclosed Recall Acknowledgement Form and fax it to 1-866- 804-9881, Attn: Customer Service. This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA) Number and will provide instructions for the return of product to Teleflex Medical.3. If you have no affected stock, please complete the enclosed Recall Acknowledgment Form and fax it to 1-866-804-9881, Attn: Customer Service. This will allow us to document your receipt of this letter.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 120 | 120 units | US Distribution including the states of FL, GA, KY, MD, MS, MT, PA, SC and TX. | Terminated | November 09, 2015 | 488 | 120 | Not_Computer
Z-2122-2014 | 68470 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | InTouch Critical Care Bed, Stryker Medical. Model Numbers 2130, 2131, 2140, 2141, 2151, 2152, 2155, and 2156. |  |  | June 04, 2014 | 2014 | Stryker Medical Division of Stryker Corporation | Stryker has received complaints from the field alleging instances where the In Touch beds are hoisted or lifted from the floor and the casters have fallen off. The investigations into these casters have shown that in these cases the caster has broken on the top of the caster stem where the brake rod inserts into the caster. If more than one caster on a unit is damaged, brake functions can be affec | Device Design | On 06/30/1014 the firm sent "Urgent Medical Device Notification" letters to their customers. |  | 11632 | 11,632 units | Worldwide Distribution -USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI &amp; WV, and the countries of Argentina, Brazil, Canada, Chile, China, EMEA, Greece, Italy, Japan, Korea, Latin America, Mexico, Poland, Singapore, South Africa, Taiwan, and UK. | Terminated | March 14, 2016 | 649 | 11632 | Not_Computer
Z-2121-2014 | 68799 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Horizon Medical Imaging |  |  | May 05, 2014 | 2014 | Mckesson Medical Immaging | The firm has identified a software issue which may cause a discrepancy between the index of images expected by the application, and the actual content of the folder storing the images. | Software design | McKesson sent an Urgent Field Safety Notice dated May 5, 2014, to all affected customers informing them of the issue and advising them of the immediate actions to take to reduce the risk to the patient. These actions include - Users ensure that all expected images are display in the Image Viewer and PACS Administrators should engage McKesson Support for all issues with studies. McKesson is also providing a software solution to HMI systems that may be potentially affected by the problem to prevent future recurrence. Customers were instructed to contact McKesson Support as soon as possible at 1-800-663-2533 to acknowledge that they have read and understood the Field Safety Notice, and to arrange installation of the software update. Customers with questions were instructed to contact Customer Support and ask to speak to their Support Manager.For questions regarding this recall call 404-338-3556. |  | 659 | 659 units | Worldwide Distribution - USA including AL, AK, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, MT, NE, NV, NH, NM, NJ, NY, NC, ND, OH, OR, PA, RI, SC, TN, TX, VT, VA, WA, WV, WI, WY,and France, Canada, Puerto Rico, Guam, Great Britain, Ireland, Australia. | Terminated | April 13, 2017 | 1074 | 659 | Computer
Z-2123-2014 | 68706 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | PHAROS Excimer Laser, Model No. EX-308 |  |  | June 30, 2014 | 2014 | Ra Medical Systems Inc | RA Medical Systems is recalling the PHAROS EX-308 Excimer Laser because there is a possibility that the laser may calibrate with the iris closed. | Error in labeling | Ra Medical Systems sent a customer letter dated June 30, 2014, to who purchased the PHAROS EX-308 Excimer Laser Phototherapy System Hand Piece. The letter informed the customers of the problems identified and the actions to be taken. Customers were also provided with the Instructions for Use (IFU) for the PHAROS EX-308 Excimer Laser Phototherapy System Hand Piece and a replacement iris indication label for the handpiece. |  | 435 | 435 units | Worldwide Distribution - USA (Nationwide) and Internationally to: Korea, Kuwait, and Saudi Arabia. | Terminated | September 09, 2014 | 71 | 435 | Not_Computer
Z-2112-2014 | 68556 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | IPM Wound Gel Rx Wound Dressing |  |  | March 25, 2014 | 2014 | Wellspring Pharmaceutical Corporation | Potential for microbial contamination of the product. | Nonconforming Material/Component | WellSpring Pharmaceutical Corporation notified customers on March 25, 2014 via e-mail. The notification identified the affected product, the problem, and directed consignees to return all products. |  | 12540 | 3,135 cartons of 4/10g tubes | Worldwide Distribution -- USA, including the states of Alabama, Arkansas, Georgia, Illinois, Kentucky, Louisiana, Missouri, Mississippi, New Jersey, South Carolina, Tennessee, Texas, and Virginia; and the country of Canada. | Terminated | September 08, 2014 | 167 | 12540 | Not_Computer
Z-2124-2014 | 68744 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | DeRoyal KNEE ARTHROSCOPY PACK, REF 89-6192.04, 1 Per Pack, Distributed by: DeRoyal Industries, Inc.; Manufacturer: DeRoyal Industries, Inc. |  |  | June 04, 2014 | 2014 | DeRoyal Industries Inc | The firm distributed surgical kits which contained Irrigation Sets which were subsequently recalled by Hospira. | Packaging change control | The recall was initiated by letter delivered via UPS on 6/4/2014. The firm requested that the consignee contact Stericycle to arrange for return and replacement of the product. A second notice was sent on 7/22/2014 to those accounts who did not respond to the initial notice. Distributors were requested to notify their end-users or provide a list of end-users to DeRoyal for direct notification. Effectiveness checks will be sent via email for those customers who have provided an email contact and via UPS for those customers who have not responded. |  | 128 | 128 units | Distributed in the states of MI, GA, OH, IL, MO, and IN. | Terminated | May 05, 2015 | 335 | 358 | Not_Computer
Z-2126-2014 | 68789 | No | DPW | Cardiovascular blood flowmeter. | Cardiovascular | Cardiovascular | 510(k) | FlowCOUPLER. An implantable device that is used to detect blood flow in vessels. Used in the end-to-end anastomosis of veins and arteries normally encountered in microsurgical and vascular reconstructive procedures. The FlowCOUPLER System is intended to detect blood flow and confirm vessel patency intra-operatively and post-operatively at the anastomotic site. Baxter Healthcare. |  |  | July 11, 2014 | 2014 | Baxter Healthcare Corporation | Baxter Healthcare is recalling the FlowCOUPLER Device because there is a possibility that some units may not maintain electrical continuity during handling intra-operatively and post-operatively, resulting in a loss of Doppler signal. This may result in the inability to monitor anastomotic patency, with the possibility that blood flow may not be appropriately monitored. | Process control | Synovis Micro Companies Alliance, Inc. (a subsidiary of Baxter International Inc.) sent an Urgent Medical Device Recall letter on July 11, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter informs the customers that the recall is being conducted as a precautionary measure due to the possibility that some units may not maintain electrical continuity during handling intra-operatively and post-operatively, resulting in a loss of Doppler signal. This may result in the inability to monitor anastomotic patency, with the possibility that blood flow may not be appropriately monitored. The recall letter informs the customers of the safety risk and and actions to be taken. Customers with questions about the recall are instructed to contact the Synovis MCA Customer Service unit at (800) 510-3318 or (205) 941-0111. |  | 1315 | 1,315 units | Worldwide Distribution - USA (nationwide) and the countries of United Kingdom, Germany, Sweden, Spain, Israel and Denmark. | Terminated | May 27, 2015 | 320 | 1315 | Not_Computer
Z-2128-2014 | 68794 | No | KLE | Resin tooth bonding agent. | Dental | Dental | 510(k) | Henry Schein Natural Elegance SE Bond Self-Etch Light-Cure |  |  | July 16, 2014 | 2014 | Novocol, Inc. | Septodont is recalling Self-Etch bond because the material was incorrectly packaged. | Process design | An email was sent on 7/16/14 to the two distributors requesting the return of all units of Lot C4211-3 in stock. |  | 60 | 60 | Product was sent to two wholsalers/distributors in PA and NY. No product was distributed outside of the US or to government/VA consignees. | Terminated | December 31, 2014 | 168 | 84 | Not_Computer
Z-2130-2014 | 68697 | No | CCW | Rigid laryngoscope. | Anesthesiology | Anesthesiology | 510(K) Exempt | RUSCH, TruLite, One-Piece Disposable Laryngoscope Handle Blade Combination, Non-Sterile, Rx Only, Distributed by: Teleflex Medical. |  |  | June 30, 2014 | 2014 | Teleflex Medical | There is a possibility the handle of the laryngoscope may heat up and melt the end cap. | Device Design | Consignees were notified via letter on/about 06/30/2014. |  | 6246 | 6,246 units | U.S. Nationwide | Terminated | March 21, 2017 | 995 | 6246 | Not_Computer
Z-2127-2014 | 68445 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Small Bone Innovations 3.0/4.0mm AutoFIX" Cannulated Compression Screw System Sets, Implants, and Instruments |  |  | May 19, 2014 | 2014 | Small Bone Innovations | Small Bone Innovations (SBi) is recalling all lots of 3.0/4.0mm AutoFIX" Cannulated Compression Screw System Sets, Implants, and Instruments to update the IFU. | Process control | Small Bone Innovations sent an Urgent Medical Device Recall letter, dated May 19, 2014, to all affected customers. The letter that described the product problem, and the action to be taken by the customer. was sent to sales reps, stocking customers, and distributors. Customers were instructed that a copy of the IFU should be stored with each unpackaged AutoFIX 3.0/4.0mm Compression Screw Sets, Implants, or Instruments in their current inventory. Customers were instructed to return all packaged affected product which would be replaced by the firm with an updated IFU. A Federal Express return label was included to facilitate the return shipment. Customers were instructed to complete, sign, and return with the product the Return Material Authorization (RMA) document.For questions regarding this recall call 215-428-1791. |  | 506 | 506 | Nationwide Distribution including TN, OK, AZ, PA, TX, MO, NY, NC, CA, FL, WA, AL, IN, NH, SD, and OH. | Terminated | October 08, 2014 | 142 | 506 | Not_Computer
Z-2134-2014 | 68737 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Patient Data Manager 2.0 (Content manager 2.0, Patient Browser 4.0, DICOM Viewer 2.0) Brainlab Digital Lightbox. A system for the display of medical images. Model numbers 25100-05 and 25100-06. The software can transfer images to and from picture archiving and communication systems (PACS), file servers, or removable storage media. |  |  | June 27, 2014 | 2014 | Brainlab AG | The following defects have been identified for data loaded with Patient Data Manager version 2.0 (including subversions 2.0.0, 2.0.1 and 2.0.2): 1) When loading non-square pixel image data with PDM 2.0, the data set is not correctly transferred and displayed in the corresponding Brainlab navigation or planning software. When a data set is loaded, the pixels of the data set are sorted in order to | Software design | FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letters dated June 27, 2014 were sent to all direct accounts. The letters instruct product users to not upload image data sets with non-square pixels and to verify with their radiology department the modalities of the different imaging devices (e.g. CT, MRT, PET, ultrasound) from which data may be loaded onto a Brainlab device. Brain lab will provide a software update with this issue solved to affected customers with a tentative planned time line of availability being July 2014. Customers with questions can contact Brainlab by telephone at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 154 | 154 systems | Worldwide Distribution-USA (nationwide) including the states of AL, AK, AZ, CA, FL, GA, IA, IL, KS, MD, MO, MS, NC, NY, OH, PA, SD, TX, VA, VT, WA, and WI, and the countries of Australia, Austria, Belgium, Brazil, Chile, China, Ecuador, France, Germany, Japan, Netherlands, New Zealand, Norway, Qatar, Russia, Spain, Switzerland, Taiwan, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 154 | Computer
Z-2131-2014 | 68657 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Zimmer Periarticular Locking Plate System Distal Radial Volar Plate Lateral Column (1st generation stainless steel only). Orthopedic fixation plate. For temporary internal fixation and stabilization of osteotomies and fractures. |  |  | June 10, 2014 | 2014 | Zimmer, Inc. | A complaint was received in which it was reported that an 8-hole plate from lot 62276042 had an enlarged hole that is unable to lock the screw, which resulted in the screw head pulling through plate hole. | Device Design | Zimmer sent an Urgent Medical Device Recall letter dated June 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers.They instructed their distributors to do the following: 1. Please locate, quarantine, and return the affected product (part number 00-2358-014-04; lot 62276042) from all inventories.2. Please note that any hospitals that received direct shipments of this product from Zimmer will be sent a copy of the Risk Manager and Surgeon letter directly. Additionally, Risk Managers and Surgeon that were identified via invoice data will also receive a letter directly. 3. Please review the list of hospital and surgeons to be provided to your territory via email.4. Provide the name, address, and phone number of any additional hospitals or surgeons that may have received or used the affected product, or confirm that the list provided is complete.5. Complete and return Attachment 1.6. Ensure that a copy of the Risk Managers and Surgeons letter is left with the facilities from which product is removed by providing them with a copy of the letter, if they are not identified as a facility receiving a letter directly from Zimmer.7.If after reviewing this notification you have questions or concerns, please call 1-877-946-2761.They instructed the Surgeons and Risk Managers to do the following: 1. Review the notification and ensure affected personnel are aware of the contents.2. Inspect inventory for part number 00-2358-014-04; lot number 62276042. Quarantine the device and contact your Zimmer sales representative to remove and return the affected device. 3. If after reviewing this notification you have questions concerning this notification, please call 1-877-946-2761. |  | 28 | 28 units | Worldwide Distribution - US including the states of CA and OH., and the countries of BRAZIL, CHINA, KOREA and TAIWAN. | Terminated | January 09, 2015 | 213 | 28 | Not_Computer
Z-2133-2014 | 68400 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | DimensionÂż Integrated Chemistry Systems intended to measure a variety of analytes in human body fluids. |  |  | May 27, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | There is a misassembly issue with thephotometer installed on DimensionÂż Integrated Chemistry Systems instruments. Through internal investigation bySiemens Healthcare Diagnostics has confirmed that a limited number of Dimension photometers were misassembled. These photometers are performing within specifications, but may lead to a shortened life of the source lamp. | Employee error | An urgent field safety notice, dated May 2014, was sent to users to inform them of the issue and that a service representative would replace the misassembled photometer. |  | 15 | 15 | FLLAARIAOHOHWIILLAKSMT | Terminated | December 15, 2014 | 202 | 15 | Not_Computer
Z-2132-2014 | 68766 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | MC-500 Multicolor Laser Photocoagulator; Distributed by NIDEK Inc. Manufactured by NIDEK CO., Ltd, Japan. For use in ophthalmic and retinal procedures. |  |  | June 09, 2014 | 2014 | Nidek Inc | Laser aperture label was not affixed to device prior to shipment in the US. | Error in labeling | Nidek sent a Recall/Field Correction letter dated June 9, 2014, to all affected customers. The letter advised customers of the issues and informed customers that Nidek Representatives will visit each location to affix the proper labeling to the device. Customers with questions were instructed to contact Nidek Customer Service at 1-800-223-9044 M-F 8 am to 5 pm PST.For questions regarding this recall call 510-353-7785. |  | 29 | 29 units (26 already re-labeled) | Nationwide Distribution including IL NE NY, NH, CD, MI, TX, CA, PA, FL, KS, UT, MN, WI,. MO. and OR. | Terminated | November 24, 2014 | 168 | 29 | Not_Computer
Z-2143-2014 | 68647 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | Ultra-speed Dental Film, Catalog Number/REF 834 8658, KODAK and Carestream DENTAL labels --- Carestream Health Inc. |  |  | July 02, 2014 | 2014 | Carestream Health Inc. | The Indicator Dot on the film packet's outer printed paper labeling was incorrectly positioned relative to the location of the raised dot on the dental intraoral film. This may cause confusion for the dental practitioner. If the processed film is placed on the view box by referencing the film's raised dot, the image orientation will be rotated 180 degrees. | Package design/selection | Recall notices were mailed to US Customers via overnight carrier on 07/02/2014. Customers are asked to complete an inventory of goods to be returned which will be ultimately destroyed by recalling company. If you have any questions or concerns, please contact the Carestream Customer Care Center in the U.S. at 1-800-933-8031. |  | 1633 | 1633 boxes | Worldwide Distribution. | Open, Classified |  |  | 1633 | Not_Computer
Z-2138-2014 | 68774 | No | NXX | Staphylococcus aureus serological reagents. | Microbiology | Microbiology | 510(K) Exempt | AdvanDx Quick FISH Fixation Kit for BC ( Blood Cultures), 50 tests. In vitro diagnostic. The QuickFISH Fixation Kits are components that are sold separately for the Staphylococcus QuickFISH BC test |  |  | July 04, 2014 | 2014 | Advandx Inc | Positive result in the Internal negative control well for the Staphylococcus QuickFISH BC test, failing QC | Nonconforming Material/Component | AdvanDx notified customers by e-mail on July 4, 2014 of the problem and advised to discontinue use of the QuickFISH Fixation Kits with Lot#04914. Replacement kits were immediately shipped to all of the affected customers. Questions contact local AdvanDx representative or Technical Support 1-866-376-0009, fax 1-781-376-0111 or e-mail: techsupport@advandx.com.An acknowledgement of Receipt and Destruction of Materials requested to be completed. |  | 10 | 10 kits | Distributed in the states of AZ, ID, FL, MO, OH, MI, NJ, and VA. | Terminated | September 30, 2014 | 88 | 10 | Not_Computer
Z-2136-2014 | 68863 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Serrated Soft Tissue Shields. Orthopedic surgical instruments, part of the Silex Sacroiliac Fusion System. The soft tissue shields are surgical instrument guides in the Silex Sacroiliac Joint Fusion System for spinal surgery. |  |  | July 11, 2014 | 2014 | X Spine Systems Inc | Revision B of the tissue shields of the device was 7 mm longer than the previous design. | Employee error | X-Spine systems contacted both customers via telephone and instructed to return the affected product back to X-Spine.All product has been recovered. The distributors have already returned the recalled product. No product was utilized in any surgical procedure.Immediate Action Taken:X-spine inventory was reviewed and two pieces of each size were recovered from finished inventory. The X-spine sales person returned the product. The distributor has returned the remaining product. The devices were recovered prior to use in any surgical procedure.For further questions please call (937) 847-8410 |  | 5 | A total of 5-pieces of each part number distributed | US Distribution including the states of FL and MO. | Terminated | August 14, 2014 | 34 | 5 | Not_Computer
Z-2139-2014 | 68790 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 3600 Immunodiagnostic System --- For use in the in vitro quantitative, semi-quantitative and qualitative measurement of a variety of analytes of clinical interest, using VITROS Immunodiagnostic Products Reagents. The VITROS Software is a part of the VITROS 3600, 4600, &amp; 5600 Systems. The VITROS 3600, 4600, &amp; 5600 Systems share the same software base. VITROS Sample Programming consists of Sample Identification (SID), test requests, and patient demographics. Patient demographics include patient name, address, date of birth, gender, physician, and patient identification (PID). |  |  | July 07, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics has identified an anomaly with VITROS System Software Version 3.0 and below. Internal testing confirmed that under very specific conditions, incorrect patient demographics were associated with a sample identification (ID) number. | Software in the Use Environment | On 7/07/2014, Customer Letter (Ref. CL2014-191) was sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them of the required actions. On 7/07/2014, Distributor Letter (Ref. DL2014-191 was sent via US Postal Service Priority Mail and/or FedEx overnight courier to US Distributor consignees instructing them to notify their consignees of the issue and required actions. Foreign affiliates were informed by e-mail on 7/07/2014 of the issue and instructed them to notify their consignees of the issue and required actions. For questions, please contact Customer Technical Services at 1-800-421-3311. |  | 678 | Domestic: 142, Foreign: 536 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 5568 | Computer
Z-2137-2014 | 68406 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | KODAK DirectView DR 9000 System, Catalog Numbers 1966688 (US and Canada) and 8161937 (outside the US and Canada) --- A permanently installed diagnostic x-ray system composed of 3 main components: an operator console, U-Arm assembly with x-ray tube including a collimator, and a stationary generator. |  |  | May 30, 2014 | 2014 | Carestream Health Inc. | A control system failure may cause the U-Arm to move downward unexpectedly when the Radiology Technologist is positioning the equipment using the "Up" or "Down" control buttons. Downward movement occurs at three times the normal rate of speed and will continue until the device control button is released, the Emergency Stop button is activated or the equipment Bump Sensor contacts an object. | Nonconforming Material/Component | Urgent Medical Device Recall Letters (dated 5/29/2014) were sent to the consignees on 5/30/2014. Carestream Health recommends that the DR 9000 U-arm be moved before positioning the patient under the U-arm for a decubitus exam. This will remove any potential safety hazard posed to the patient as a result of any unexpected movement. For questions or concerns, please contact the Carestream Customer Care Center in the US at 1-800-328-2910, available 7 days per week on a 24 hour basis; Outside of the US, please call your local Service support number. |  | 146 | Domestic: 100 units; Foreign: 46 units | Worldwide Distribution-USA (nationwide) and the countries of Canada, China, Italy, Germany, Greece, Kuwait, Spain, South Africa, and United Arab Emirates. | Terminated | December 17, 2014 | 201 | 146 | Not_Computer
Z-2145-2014 | 68745 | No | HET | Gynecologic laparoscope and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Instructions for Use for the blue endo MOREsolution Electromechanical Laparoscopic Morcellator System, Model 55-000011, consisting of the control unit, hand-grip, flexible shaft, trocars/obturators, sleeves, valves, and blades. |  |  | July 16, 2014 | 2014 | Market-Tiers Inc | Some pages of the Instructions for Use were omitted, including pages regarding the Indications for Use, Warning, and Contraindications. | Labeling Change Control | The recalling firm issued a letter dated 7/16/14 via Federal Express requesting the customer destroy the incorrect IFU and replace it with the corrected version. |  | 100 | 100 units | Nationwide Distribution. | Terminated | May 11, 2015 | 299 | 100 | Not_Computer
Z-2144-2014 | 68862 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Tapered-Screw Vent Implant, HA, 4.7mmD x 8mmL x 4.5mm, Part No. TSVWH8 Dental implant component. |  |  | July 17, 2014 | 2014 | Zimmer, Inc. | Zimmer Dental is recalling the Tapered Screw-Vent Implant because the inner vial state "03.7 x 8mm" instead of "04.7 x 8 mm." | Process design | A recall letter dated 7/22/14 was sent to all customers who purchased the Tapered Screw-Vent Implant to inform them of the recall by Zimmer Dental. The letter informs the customers of the problems identified, risks involved, and the actions to be taken. Customers with questions are instructed to contact Zimmer Dental Customer Service at (800) 854-7019, 7-5pm, Mon-Fri. Customers are instructed to complete the customer response form and e-mail a copy to corporatequality.postmarket@zimmer.com or fax to (574) 372-4265. |  | 291 | 291 units | Worldwide Distribution-USA (nationwide) and the countries of Holland, Italy, and Japan. | Terminated | February 06, 2015 | 204 | 291 | Not_Computer
Z-2146-2014 | 68741 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac is intended to be used to place patients at an accurately defined point within the treatment beam of a medical accelerator for stereotactic radiosurgery or radiotherapy procedures, in order to treat lesions, tumors and conditions anywhere in the body when radiation treatment is indicated. ExacTrac may also be used to monitor the patient position during the treatment. |  |  | June 24, 2014 | 2014 | Brainlab AG | When using multiple isocenters (radiation treatment targets) within a single plan, in certain workflow conditions ExacTrac v.6.0.4 might move the patient to an unintended isocenter position, despite displaying the green "OK" icon. If this anomaly occurs and is not detected by the user, the radiation treatment dose at the linear accelerator may be delivered to the unintended target position. | Software design | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated June 24, 2014, to all affected customers. The letter included instructions for customers to: 1) Avoid using plans with multiple isocenters in ExacTrac 6.0.4 whenever possible, and 2) If plans with multiple isocenters cannot be avoided and such a plan is used in ExacTrac v.6.0.4, then always select "Yes" in the dialog displayed to "keep the previous correction shift" when switching between isocenters and always verify all final treatment positions. Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact you to schedule the update within the next 6 months. Customers with questions can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 15 | 15 systems | Worldwide Distribution - USA including CA, CT, and IL; Internationally to France, Germany, Japan, Netherlands, Russia, and Slovenia. | Open, Classified |  |  | 15 | Computer
Z-2147-2014 | 68700 | Yes | LWS |  |  | Cardiovascular | PMA | Boston Scientific ENDOTAK RELIANCEÂż SG, transvenous defibrillation lead, Models 0180 &amp; 0292, Sterile EO. |  |  | June 19, 2014 | 2014 | Boston Scientific CRM Corp | A review of manufacturing test records for Boston Scientific ENDOTAK RELIANCE implantable leads revealed a suspected test data recording error. Specifically, some test results were recorded as "failed" without any other indication of failure. There are no reported injuries from the devices. | Software Manufacturing/Software Deployment | Sales representatives hand delivered a Boston Scientific "Medical Device Retrieval" Letter dated June 19, 2014 or June 20, 2014 when they were retrieving the devices from the hospital shelves. The letter was addressed to Hospital Administrator. The letter described the problem, product being recalled and the retrieval of the device by Boston scientific sales representative. For further questions they can contact their local sales representative or Boston Scientific International Technical Services.Physicians were contacted via telephone starting on July 6, 2014. A Boston Scientific letter dated July 8, 2014 followed the telephone conversation. For further information physicians can contact their local Boston Scientific representative or Technical Services at 1-800-227-3422. |  | 21 | 21 | Worldwide Distribution -US (nationwide) in the states of OH, MN and country of France | Terminated | December 03, 2014 | 167 | 21 | Not_Computer
Z-2152-2014 | 68749 | No | MTF |  |  | Immunology | PMA | VIDAS TPSA, in vitro diagnostic, for use as an automated quantitative measurement of prostate specific antigen levels in human serum or plasma. |  |  | October 14, 2013 | 2013 | BioMerieux SA | Product is not approved or authorized for distribution in the US. | Packaging process control | A letter was issued on October 14, 2013 to consignees requesting the consignee to check their stock and destroy all kits of the VIDAS TPSA ref. 30428. The letter also requested that the consignee notify their accounts of the recall. An acknowledgment form was also included with the letter which is to be returned to BioMerieux. |  | 1902 | 1902 | FL, LA, MI, MS, NV, OH, OK, TN, TX | Completed |  |  | 1902 | Not_Computer
Z-2154-2014 | 68746 | No | FCM | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | RUSCH Foley Urinary Catheter Insertion Tray with 10cc syringe, BZK Swabs, Sterile, Rx only, Teleflex Medical Product Code: 76730. |  |  | July 10, 2014 | 2014 | Teleflex Medical | Labeling error: The manufacturing date is the same as the expiration date. The manufacturing date is correct, however the expiration date incorrectly indicates that the product is expired. | Process control | Consignees were notified via letter on/about 07/10/2014. |  | 720 | 720 | Nationwide distribution in CA, CT, FL, MA, MO, NJ, OH, OR and TX. | Terminated | November 25, 2015 | 503 | 720 | Not_Computer
Z-2151-2014 | 67640 | dam, rubber | EIE | Rubber dam and accessories. | Dental | Dental | 510(K) Exempt | Latex Dental Dams labeled as Non-Latex Dental Dams |  |  | December 11, 2013 | 2013 | Coltene Whaledent Inc | ColtÂżne/Whaledent Inc. received a complaint about the dental dams being the incorrect color. During an internal investigation, the firm found that Latex dental dams had been packaged in Non-Latex packaging. | Packaging change control | The firm issued a press release, called the distributor and sent a certified letter. The firm obtained the distributor's distribution list and sent recall notifications to all of these consignees.The firm included envelopes and return/reply cards, via response form mailed, certified mailing with return receipt. |  | 480 | 32 boxes, 15 individual dams per box | Distributed in Memphis, TN. | Open, Classified |  |  | 480 | Not_Computer
Z-2153-2014 | 68686 | [6 empty] | Panthera-Puncher 9 . |  |  | July 01, 2014 | 2014 | PerkinElmer Health Sciences, Inc. | Risk of incorrect results from samples prepared with Panthera-Puncher 9 | Process design | Perkin Elmer sent an Urgent Medical Device Recall letter dated July 1, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The Recall Letter delivered to the end customer by email facilitated through our Perkin Elmer Technical Support group located in Akron Ohio.Technical Support will be asked to make sure the customers receive the Recall Letter and implement the recommended precautions: "Removing the plates from the Panthera-Puncher 9 when not actively punching and processing samples in the instrument."Opening the plate track hatch if plates must reside in the Panthera-Puncher 9 while the instrument is idle or in sleep mode. This will interrupt power to the ionizer and eliminate any ionizer impact on the plate wells.For further information, please contact your local Perkin Eimer representative or GSTechnicaiSupport@PerkinEimer.com. For further information please call (781) 663-6900. |  | 144 | 144 units | Worldwide Distribution - USA (nationwide) and the countries of : Canada, Australia, Brazil, China, Denmark, Great Britain, Italy, France, Mexico, Qatar and Korea. | Terminated | October 21, 2015 | 477 | 144 | Not_Computer
Z-2160-2014 | 68971 | No | DRS | Extravascular blood pressure transducer. | Cardiovascular | Cardiovascular | 510(k) | Trifurcated Transpac IV Monitoring Kit w/Safeset Reservoir, 03 ml Squeeze Flush and 6 Needleless Valves, Item No. 011-46106-22 |  |  | July 18, 2014 | 2014 | ICU Medical, Inc. | ICU has become aware of an issue with a component where the connection on an arterial monitoring kit may crack and cause a fluid leak. | Nonconforming Material/Component | ICU Medical, Inc. sent an email dated July 18, 2014, to their customers, The email inform the customer of the problem and the action to be taken. The email stated that the US Sales Representative would remove the recalled product on July 22, 2014. If product was further distribution customers should notify their customers. Customer was informed of the RGA to have the recalled product returned.For questions regarding this recall call 801-264-1732. |  | 60 | 60 units | Worldwide Distribution - USA including Oregon and Internationally to Australia. | Terminated | September 25, 2014 | 69 | 3570 | Not_Computer
Z-2156-2014 | 68807 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Sonopet Ultrasonic Surgical System Console, Ultrasonic Aspirator, Stryker Medtech. |  |  | June 24, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | The firm received complaints regarding insufficient irrigation to the tip of the Sonopet system. Lack of irrigation can potentially lead to excessive heat at the tip. | Labeling design | The firm sent Medical Device Upgrade Notification letters dated July 11, 2014 to their customers. |  | 2460 | 2,460 units | Nationwide Distribution-including DC and the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI &amp; WV. | Terminated | October 27, 2015 | 490 | 2460 | Not_Computer
Z-2166-2014 | 68670 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | PerkinElmer Specimen Gate Screening Center software, version 1.4 and 1.6; Product Code: 5002-0500 |  |  | June 27, 2014 | 2014 | PerkinElmer Health Sciences, Inc. | Potential for demographic information and test result to be mismatch. | Software design | PerkinElmer sent an Urgent Medical Device Recall letter on July 1, 2014, to all affected customers. The letter informed customers of the recall and provided instructions. Customers were asked to complete the enclosed response form and fax it to 1-330-825-8520. For additional information, customers were instructed to contact their local PerkinElmer representative or email specimen.gate.support@perkinelmer.com.Customers with questions were instructed to call 1-800-321-9632. option 1. |  | 13 | 13 devices | Worldwide Distribution - USA including MA, GA, FL, NV, and Internationally to Canada, Ireland, Italy, Saudi Arabia, Belgium, and UK | Terminated | April 06, 2015 | 283 | 13 | Not_Computer
Z-2169-2014 | 68617 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | 5 mL Reaction Vessels (a consumable of the Abbott m2000sp Automated Fluid Handling System). The Abbott m2000 RealTime System is intended for use in performing nucleic acid amplification testing by polymerase chain reaction in clinical laboratories. The system is comprised of the m2000rt and the m2000sp instruments. |  |  | June 19, 2014 | 2014 | Abbott Molecular | Abbott Molecular identified that some of the 5 ml Reaction Vessels (RVs) contained in lots 56944001, 57026001, 57059001, and 57673001 (original vendor lot 2198077) may have an insufficient rim. An RV with an insufficient rim may not be held securely in the 1 mL Subsystem Carrier. This may cause the RV to be seated incorrectly in the carrier or drop through it. | Employee error | Abbott sent a Field Correction Recall / Urgent Field Safety Notice dated July 3, 2014, to all affected customers. June 19, 2014 were sent to all customers. The letters instructed customers to: 1) examine their inventories, 2) remove any of the affected lots, and 3) contact their Abbott Molecular Representativeat 1-800-553-7042 to coordinate replacement.Customers were also instructed to review this information with laboratory personnel and retain the communication for future reference.For questions regarding this recall call 224-361-7274 |  | 1688000 | 1,688,000 reaction vessels | Worldwide Distribution - USA (nationwide) and Internationally to:AUSTRALIAAUSTRIABELGIUMBENINCAMEROONCHINAETHIOPIAFINLANDFRANCEGERMANYINDIAINDONESIAIRELANDISRAELITALYKENYAMALAYSIAMALAWIMALINETHERLANDSNEW ZEALANDNORWAYPAKISTANPHILIPPINESPOLANDPORTUGALRUSSIASENEGALSINGAPORESOUTH AFRICASOUTH KOREASPAINSWEDENSWITZERLANDTAIWANTHAILANDTURKEYUKRAINEUNITED KINGDOMVIETNAM | Terminated | February 23, 2016 | 614 | 1688000 | Not_Computer
Z-2170-2014 | 68177 | No | PEN | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | Verigene Gram-Negative Blood Culture Nucleic Acid Test (BC-GN), performed using the sample-to-result Verigene System, is a qualitative multiplexed in vitro diagnostic test for the simultaneous detection and identification of selected gram-negative bacteria and resistance markers. BC-GN is performed directly on blood culture media using blood culture bottles identified as positive by a continuous monitoring blood culture system and which contain gram-negative bacteria as determined by gram stain. BC-GN is indicated for use in conjunction with other clinical and laboratory findings to aid in the diagnosis of bacterial bloodstream infections; however, is not used to monitor these infections. Sub-culturing of positive blood cultures is necessary to recover organisms for antimicrobial susceptibility testing (AST), for identification of organisms not detected by BC-GN, to detect mixed infections that may not be detected by BC-GN, for association of antimicrobial resistance marker genes to a specific organisms, or for epidemiological typing. The BC-GN test is performed on the Verigene System platform, which is a fully automated, bench-top, molecular diagnostics workstation consisting of Verigene Reader and a bank of up to 32 Verigene Processor SP units. The System enables the detection of bacterial DNA from blood culture, unformed stool, or nasopharyngeal swab, depending on the test, through automated nucleic acid extraction and hybridization. |  |  | April 17, 2014 | 2014 | Nanosphere, Inc. | Nanosphere has determined that a portion of BC-GN Test Cartridge Lot number 040214021A has an elevated risk for No Call results and False Negative results. The BC-GN Test Cartridge (20-006-021) is distributed as part of the Verigene BC-GN Test Kit (20-005-021), which is labeled as "IVD". Therefore, Nanosphere has initiated a voluntary recall to retrieve the following BC-GN Test Cartridges: Test Ca | Under Investigation by firm | All customers were initially notified of the recall by telephone on April 17, 2014. The customers were instructed to quarantine and return all affected product to Nanosphere. Nanosphere sent an Voluntary Product Recall Notification letter dated April 17, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letters included instructions: 1) on how to return the recalled products, and 2) to review any test results that were generated while using the recalled products to ensure that there has been no negative impact on patient care. Customers with questions can contact Nanosphere at 888-837-4436, option 2 or via e-mail at Technicalsupport@nanosphere.us.. |  | 580 | 580 test cartridges | US Distribution including the states of AZ, CA, FL, MN, MT, NE, NJ, NY, TX and WI | Terminated | March 14, 2016 | 697 | 580 | Not_Computer
Z-2168-2014 | 68708 | No | OZN | Clostridium difficile toxin gene amplification assay. | Microbiology | Microbiology | 510(k) | Nanosphere Verigene |  |  | October 16, 2013 | 2013 | Nanosphere, Inc. | Nanosphere has received report of an increased rate of "No Call - INT CTL 2" results for the Internal Control 2 (INT CTL 2) when negative samples are tested with certain lots of CDF Amplification Trays (20-011-022), which are contained in the CDF Amplification Reagent Kit (20-012-022); specifically for lots 061913022D and 071813022E. | Nonconforming Material/Component | Nanosphere sent an Field Notification Letter dated October 16, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letters included instructions for customers to: 1) discontinue use of the affected product, 2) discard any of the affected product, and 3) complete the attached form and return it to Nanosphere. Customers with questions can contact Nanosphere at 847-400-9000 or 888-837-4436 x 2. |  | 280 | 280 sets of trays | US Distribution including the states of CA and WA | Terminated | May 22, 2015 | 583 | 280 | Not_Computer
Z-2171-2014 | 68814 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist Stapler System User Manual Addendum and Wall Chart; |  |  | June 02, 2014 | 2014 | Intuitive Surgical, Inc. | Changes were made to the Stapler EndoWrist Stapler 45 System Instruments and Accessories User Manual Addendum (550986-03) and the EndoWrist Stapler 45 Wall Chart (PN 551524-01) in which the pH range specification for the enzymatic cleaner used throughout the cleaning process was modified from pH-neutral to pH-neutral to mild alkaline (mild alkaline is defined as pHd11). Mild Alkaline enzymatic | Labeling Change Control | Intuitive Surgical representatives hand delivered corrected manuals to the two affected sites.. |  | 2 | 2 | Distribution to Florida and South Carolina | Terminated | August 08, 2014 | 67 | 2 | Not_Computer
Z-2135-2014 | 68596 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | Enhancement Medical expression. model # EXP-150. Injectable Gel 1.5cc Syringe with hyaluronic acid concentration of 20 mg/ml. |  |  | August 07, 2013 | 2013 | Enhancement Medical, LLC | Manufacturer made a change in the production process that resulted in a change in final gel weight. RECALL EXPANDED 7/8/2014 Firm expanded their recall to include all lots of product. | Device Design | Enhancement Medical, LLC called all Independent Sales Representatives and then provided each representative a list of consignees with impacted product and instructed the Representatives to disclose low gel weight issue and potential for common treatment responses. Representatives were instructed to exchange the affected product. EXPANDED Enhancement Medical sent an Urgent Medical Device Recall letter beginning July 8, 2014 to all consignees whom received product. The letter identified all lots as being recalled, the reason for the recall, recommended that Expression not be injected subcutaneously, and asked for product to no longer be used and be returned to the recalling firm. An acknowledgement Form was asked to be mailed, e-mailed or faxed back to the firm.Customers with questions were instructed to call 414-918-4280. |  | 17875 | 17,875 | Worldwide Distribution - USA (nationwide including Puerto Rico) and Internationally to Canada. | Open, Classified |  |  | 17875 | Not_Computer
Z-2172-2014 | 68838 | No | LKK |  |  | General Hospital | PMA | Medtronic Ascenda Intrathecal Catheter models 8780 and 8781, and Ascenda Intrathecal Catheter Pump Segment Revision Kit, model 8784. |  |  | July 11, 2014 | 2014 | Medtronic Neuromodulation | Medtronic is removing specific lots of the Ascenda Intrathecal Catheters and Revision Kits, which are used with the implantable SynchroMed drug infusion pump. This recall is being conducted due to a single component of the catheter, the retainer ring, not meeting specification criteria. There is a possibility of unintentional disconnection of the catheter from the pump, or difficulty in disconne | Nonconforming Material/Component | Medtronic sent an Urgent Medical Device Removal letter dated July 11, 2014, to all affected customers. The letter described the issue, identified affected product, provided patient management recommendations for patients implanted with an affected device, asked for Risk Management to quaratine unused inventory within the hospital, and to contact Medtronic Customer Service at 1-888-638-7627 to facilitate return of devices. |  | 1200 | 1200 (551 US, 649 OUS) | Worldwide Distribution - USA (nationwide) including the states of AL,AR,AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, Ml, MN, MO, MS, MT, NC, NE, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI and WV., and the countries of Austria, Belgium, Cyprus, Denmark, France, Germany, Iraq, Italy, Lebanon, Netherlands, Panama, Poland, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, United Kingdom and Uruguay. | Terminated | July 17, 2015 | 371 | 1200 | Not_Computer
Z-2167-2014 | 68648 | No | JOP | Ultrasonic transducer. | Cardiovascular | Cardiovascular | 510(k) | Esaote PA230E Transducer |  |  | June 26, 2014 | 2014 | Biosound Esaote, Inc. | Using the PA230E transducer with a user-defined setting with the output power to 100% (Max) in 2D-TEI mode focusing at maximum depth, could cause the transducer surface temperature to exceed the maximum allowable limit. | Software design | On 06/26/2014, Esaote North America sent an advisory letter via FEDEX (with signature required) to each customer or consignee affected by this voluntary recall. The packet contained the advisory letter, including actions to be taken by the user, information about a pending software upgrade, and a business reply card to confirm they received the letter and have the equipment noted on the advisory letter. Customers with questions were instructed to call 800-428-7378.For questions regarding this recall call 1-317-813-6005 |  | 1301 | 1,301 units | Nationwide-Product was shipped to the following states: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, UT, VA, WA, WI &amp; WV. International**Product was also shipped to consignees in Canada | Terminated | July 17, 2015 | 386 | 1301 | Computer
Z-2173-2014 | 68633 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Implantable Insulin Pump, Model No. MMT-2007D |  |  | June 16, 2014 | 2014 | Medtronic MiniMed Inc. | Medtronic MiniMed is recalling the Implantable Insulin Pump because there have been reports of a pump malfunction resulting in the disconnection of the delivery of insulin prior to the end of the 7 year useful life of the device. | Under Investigation by firm | An Urgent Field Safety Notice Letter dated 6/16/14 was sent to all customers who purchased the Medtronic Implantable Insulin Pump. The letter informs the customers of the problems identified and the actions to be taken. Customers are informed that Medtronic MiniMed is working to make new unaffected pumps available as soon as possible. Customers are instructed to contact Elisabeth Andrieu at + 33 6 89 71 9056 or Dr. Francine Kaufman at +1 818.576.5331. |  | 764 | 764 units | New Jersey and Worldwide: Belgium, France, Netherland, and Sweden | Terminated | September 25, 2014 | 101 | 764 | Not_Computer
Z-2175-2014 | 68752 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | DX-D 600 - DXD Imaging Package |  |  | July 02, 2014 | 2014 | AGFA Healthcare Corp. | Some customers reported that while using the DX-D600 in room or exam preparation, the Overhead Tube Crane (OTC) showed increased velocity and uncontrolled longitudinal movement. | Device Design | AGFA HealthCare sent an URGENT FIELD SAFETY NOTICE letter dated July 2, 2014 to affected customers. The letter described the affected product, problem and actions to be taken. Customers were instructed to complete and return the Acknowledgment form, via FAX-Back or email. A Mandatory Service Bulletin DIS012.14E/No. 32, was released June 30, 2014, to implement the mandatory upgrade of full automatic DX-D600 systems to version 3.6 to prevent unexpected system movements. This Mandatory Service Bulletin was issued to Agfa DR Specialists to perform the service. A 2nd release will be implemented to upgrade semi-automatic systems. For questions contact your local AGFA HealthCare organization 1-877-777-2432. |  | 17 | 17 units | Worldwide Distribution - US Natinwide in the state of GA, IN, KS, KY, MA, NJ, SC, VA, DC and the country of Canada | Terminated | June 04, 2015 | 337 | 17 | Not_Computer
Z-2176-2014 | 68798 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Agfa Digital Radiography X-Ray System DX-D 100 |  |  | May 13, 2014 | 2014 | AGFA Healthcare Corp. | Unit was mislabeled with a factory label showing 32 KW output power. | Labeling mix-ups | Only one customer was affected no customer letter was issued field correction was conducted.. |  | 1 | 1 | US in the state of WV | Terminated | August 13, 2014 | 92 | 1 | Not_Computer
Z-2177-2014 | 68727 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Kodak DirectView DR 7500 Diagnostic X-Ray System |  |  | July 09, 2014 | 2014 | Carestream Health, Inc. | An on-site evaluation of the equipment parts that attach the Beta Assembly (x-ray tube and collimator) to the Over Head Tube Crane (OTC) Telescope was completed by Carestream Health (CSH) Service Engineering. Inspection of these parts revealed that the tab on the lock washer was not bent into the lock nut recess and because the Beta Assembly detached from the OTC telescope, the lock nut was not sufficiently tightened. | Device Design | Carestream will notify all DR 7500 customers of the issue and schedule a Field Engineer to inspect and repair the system if required. Specific instructions (Service Modification) for the inspection and repair will followed by the Carestream Field Engineer in order to complete the correction. Documented evidence of the Field Engineer visit will be noted in the Carestream service engineering record for each affected system. Carestream estimates that the plan of correction will be completed within 2-4 weeks |  | 589 | 589 systems | Worldwide Distribution - US Nationwide and the countries of (United Kingdom, United Arab Emirates, Turkey, Taiwan, Switzerland, Sweden, Spain, South Africa, Portugal, Poland, Oman, New Zealand, Netherlands, Kuwait, Italy, Israel, Hong Kong, Germany, France, Finland, Denmark, Dubai, Czech Republic, China, Korea, Belgium, Australia, and Amsterdam. | Terminated | May 12, 2017 | 1038 | 589 | Not_Computer
Z-2179-2014 | 68888 | No | MDS |  |  | Clinical Chemistry | PMA | Medtronic MiniMed Guardian Monitor, Model No. CSS7100, CSS7100K |  |  | July 21, 2014 | 2014 | Medtronic MiniMed Inc. | Medtronic MiniMed is recalling the Guardian REAL-Time monitors because of a manufacturing issue that has the potential to compromise the water resistance of the affected Guardian monitors. | Device Design | Medical Device Recall letters dated July 21, 2014 andJuly 22, 2014 were sent to distributors, patients, and healthcare providers who purchased the Guardian REAL time monitors. The letters inform the customers of the problems identified and the actions to be taken. Distributors with questions are instructed to email rs.nrdlrinbox@medtronic.com. If healthcare professionals want to speak to Medtronic's medical officers then they are instructed to call (818) 576-4211, ext. 64511. Customers and healthcare professionals' patients are instructed to call Guardian Recall Hotline at (800) 646-4633, ext. 21090, 8am-6pm, CST. |  | 86 | 86 units | Worldwide:Distribution - US Nationwide and the countries of Croatia, Denmark, Finland, Germany, Netherlands, Norway, Sweden, Switzerland. | Terminated | October 08, 2014 | 79 | 86 | Not_Computer
Z-2178-2014 | 68844 | No | IWZ | Radiographic film. | Radiology | Radiology | 510(K) Exempt | Radiomat M+ NIF 14 x 17; Radiomat SG 14 x 17 NIF; RX-B 8 x 10 in./IBF-004355; Radiomat B+ NIF 14 x 17; Radiomat B+ NIF 8 x 10; Radiomat B+ NIF 9 1/2 x 9 1/2; CX Ultra UV-G Plus NIF 11 x 14; Radiomat G Plus 14 x 17 NIF |  |  | July 16, 2014 | 2014 | AGFA Healthcare Corp. | Some sheets of the medical screen film are fogged. | Process design | On July 16, 2014, an "URGENT FIELD SAFETY NOTICE" letter was sent to the US and Canadian dealers (consignees). The letter described the safety alert and the request to quarantine affected product. Acknowledgment, via FAX-Back or email, that the information was received and understood was requested from the dealers. A customer Urgent Field Safety Notice was also provided to those dealers who had shipped affected product to end users so that it could be sent to their customers. |  | 2617 | 2617 | Worldwide Distribution and US Nationwide in the states of AL, AZ, CA, FL, IL, LA, MI, MN, MS, NE, NH, NJ, NC and country of Canada | Terminated | December 09, 2014 | 146 | 2617 | Not_Computer
Z-2181-2014 | 68803 | No | PDH | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Luma Wrap Phototherapy Swaddler is designed to provide centered and comfortable boundaries as a benefit for restive babies in the Neonatal Intensive Care Unit. |  |  | July 11, 2014 | 2014 | Beevers Manufacturing &amp; Supply, Inc. | Beevers Manufacturing &amp; Supply Inc. is recalling Luma Wrap neonatal phototherapy blankets because the product was inadvertently misclassified and marketed without a cleared 510(k) submission from the US FDA. | No Marketing Application | BMS is contacting their customers and requesting that distributors report the amount of affected product they have in stock, prior to returning affected product. BMS is also asking its distributors to forward the recall notification to its customers to whom affected product was sold. BMS expects those customers to also report to BMS the amount of affected product they have. |  | 22080 | 240 total (112/ lot 66812; 92/ lot 66813; 17/ lot 30660; and 19/ lot 67879). | Product was distributed to US, Germany, Austria, France, Italy, Spain, Australia. | Terminated | July 17, 2015 | 371 | 22080 | Not_Computer
Z-2180-2014 | 68634 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Syngo RT Therapist and Syngo RT Oncologist intended use of the SIEMENS branded Syngo RT Therapist and Syngo RT Oncologist software applications used with SIEMENS family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer. |  |  | June 17, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens became aware of a potential issue with the syngo RT Therapist 4.3.SP1 and syngo RT Oncologist 4.3SP1 software applications running on Siemens brand LINAC systems that may experience sporadic system crashes that will cause the inability to unload patient data. | Software design | In June, 2014, an urgent medical device correction field safety notice of improvement was sent to end users informing them of the product, problem, and action to be taken to avoid the potential safety issue. |  | 56 | 56 | UT, WI, NY | Terminated | December 15, 2014 | 181 | 56 | Computer
Z-2174-2014 | 68507 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Linear Accelerators of type ARTISTE, ONCOR, and PRIMUS with Automatic Sequenced Cancer Treatment Delivery Option. |  |  | May 28, 2014 | 2014 | Siemens Medical Solutions USA, Inc | A safety risk exists with automatically sequenced treatment technique using the SIMTEC Auto Field Option where automatic movements of the gantry and the tabletop during an auto-sequenced treatment have the potential to lead to a collision with the patient. | Use error | An Urgent Medical Device Correction Field Safety Notice Letter was sent to end users on/about June 2014 which described the product, problem, and actions to be taken to avoid encountering the safety risk to patients and users. |  | 160 | 160 | Nationwide and ROW, including Canada and Mexico. | Terminated | July 23, 2015 | 421 | 160 | Not_Computer
Z-2182-2014 | 68679 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Integra NewPort MIS System |  |  | June 27, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences is recalling the Integra NewPort MIS System because there is a possibility that components from the demonstration set may have been mixed with the NewPort MIS pedicle screw system components used in surgery. | Under Investigation by firm | The firm, Integra, sent an "URGENT MEDICAL DEVICE CORRECTION FIELD SAFETY NOTICE" dated June 27, to its customers. The letter informs the customers that Integra Lifesciences is recalling the Integra NewPort MIS System because there is a possibility that the components from the demonstration set may have been mixed with the NewPort MIS pedicle screw system components used in the surgery. The letter also informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the attached response form and return it by email or fax as indicated on the form. Customers with questions regarding the recall letter instructions are instructed to contact the Marketing Manager at (760) 207-7050. |  | 4 | 4 units | US Distribution in TX only. | Terminated | September 16, 2014 | 81 | 4 | Not_Computer
Z-2183-2014 | 68890 | No | NIG | Tumor-associated antigen immunological test system. | Immunology | Immunology | 510(k) | VITROSÂż Immunodiagnostic Products CA 19-9 Calibrators and Reagent Packs, Lot 1320 |  |  | July 16, 2014 | 2014 | ORTHO-CLINICAL DIAGNOSTICS | Customers may receive positively biased results when using VITROSÂż Immunodiagnostic Products CA 19-9 Total Reagent Packs, Lot 1320 and VITROSÂż CA 19-9 Calibrators, Lot 1320 on the VITROS ECi/ECiQ Immunodiagnostic Systems, the VITROS 3600 Immunodiagnostic System and the VITROS 5600 Integrated System. There is an unexpected observed change in CA 19-9 Calibrators, Lot 1320 calibrator values. | Other | Ortho Clinical Diagnostics sent an Urgent Product Correction Notification July 2014 to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to Immediately discontinue using and discard Lots 1320 of VITROS CA 19-9 Calibrators and Reagent Packs.Review previously reported VITROS 25-OH Vitamin D Total results using Lot 1320. Review quality control results used to verify calibration events. Results lower than the target mean may indicate potentially biased assay results.Discuss any concerns regarding previously reported results with the Laboratory Medical Director or with the requesting physician.Complete and return the attached Confirmation of Receipt form. OCD will:credit their account for all VITROSÂż CA 19-9 Calibrators and Reagent Packs that were discarded or ship repalcment product uspon request.Customers were instructed to forward the information in this notification if they have distributed this product outside of their facility.Customers with questions were instructed to contact their Customer Technical Service representative or contact the Customer Technical Service representative at 1-800-421-3311. |  | 475 | 475 | Worldwide Distribution - USA (nationwide) and Internationally to France and Japan | Open, Classified |  |  | 475 | Not_Computer
Z-2184-2014 | 68837 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Codman Certas - In Line Valve Only; Product Code: 82-8800 |  |  | July 03, 2014 | 2014 | Codman &amp; Shurtleff, Inc. | To clarify the CODMAN CERTAS Valve virtual off (setting 8) pressure specification outlined in the product's Instructions for Use (IFU). | Device Design | Codman Neuro sent an Urgent Medical Device Safety Notification dated July 17, 2014, to clinicians who use the device. The notification clarifies the CODMAN CERTAS Valve virtual off (setting 8) pressure specification outlined in the product's Instructions for Use (IFU). Consignees are also asked to complete the enclosed Acknowledgement Form and fax the completed form to 1-888-239-1305. For questions or concerns regarding this notification, consignees are instructed to contact their local Codman Neuro Representative or Scientific and Medical Affairs at SciMedAffairs@its.jnj.com or (866) 685-7325.For questions regarding this recall call 508-828-2726. |  | 9498 | 9,498 devices total | Worldwide Distribution - USA including AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MT, NC, ND, NE, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, WY; and Internationally to: Canada, Australia, Austria, Belgium, Brazil, Czech Republic, Denmark, Egypt, Finland, France, Germany, Great Britain, Greece, Hungary, Ireland, Israel, Italy, Kuwait, Lebanon, Luxembourg, Netherlands, Norway, Portugal, Russia, Saudi Arabia, Slovak Republic, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey. | Open, Classified |  |  | 9498 | Not_Computer
Z-2120-2014 | 68302 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GYC-1000 Green Laser Photocoagulator system |  |  | May 16, 2014 | 2014 | Nidek Inc | Laser Aperture label was not applied to certain GYC-1000 Green Laser Photocoagulators devices in the field as per FDA regulation, | Radiation Control for Health and Safety Act | Nidek sent an Medical Device Correction Notifications dated May 16, 2014, informing doctors of the lack of the Laser Aperture Label for GYC-1000 Green Laser Photocoagulator System. Letters included a Aperture label and instruction/illustration on where to affix it to the device. A response form is to be completed and returned indicating that they read the correction information and have affixed the label on the GYC-1000 device as instructed. Questions may be directed to Nidek Customer service at 1-800-223-9044 Monday - Friday from 8 am to 5 pm PST.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.You may proceed or continue with implementation of the CAP (subject to the conditions noted above). |  | 567 | 567 units | USA (nationwide) Distribution | Terminated | October 16, 2014 | 153 | 567 | Not_Computer
Z-2214-2014 | 68840 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROSÂż Chemistry Products CDM PROM Calibration Data Module Rev. No. 0168 and 0169 (Product Code 1999077) are used in conjunction with VITROS Chemistry Products NH3 DT Slides (Product Code 1532589) GEN 59 and VITROSÂż Chemistry Products DT Calibrator Kit lot 44 (Product Code 1957927) and tested on the VITROS DT60/DT60 II Chemistry System: IVD, Made in USA by Ortho-Clinical Diagnostics, Inc. Rochester, NY |  |  | July 15, 2014 | 2014 | Ortho-Clinical Diagnostics | Inability to calibrate or use VITROS NH3 DT Slides, GEN 59 with VITROS Chemistry Products DT Calibrator Kit, Lot 44 when CDM PROM 0168 or 0169 is installed on VITROS DT60/DT60 II Chemistry Systems. | Process control | Ortho Clinical Diagnostics sent an Urgent Product Correction Notification letter dated July 15, 2014,by Federal Express overnight to all affected customers. All US end-user consignees who had been shipped VITROS NH3 DT slides, GEN 59, Product Code 1532589 were informed of the issue. Customers were instructed to: a) use VITROS DT NH3 slides, GEN 59 with Calibration Data Module, Rev. 0168 or 0169 with VITROS DT Calibrator, Lot 42 if available; or b) If VITROS DT Calibrator, lot 42 is not available, the user is instructed to request CDM, Rev. 0170 to be used with VITROS NH3 DT slides, GEN 59 and VITROS DT Calibrator, Lot 44.Customers were instructed to forward this notice if the product was further distributed outside of their facility.OCD has manufactured a new CDM, Rev. 170, containing the necessary data/calibration mathematics for VITROS DT Calibrator, Lot 44 to be used to calibrate VITROS NH3 DT slides, GEN 59.Customers with questions were instructed to call Customers Technical Services at 1-800-421-3311.For questions regarding this recall call 908-218-8776. |  | 5906 | 5906 total: 240 (US); 5666 (OUS) | Worldwide Distribution - USA including MA, TX, KY, IL, MS, OH and Puerto Rico; Internationally to: Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 5906 | Not_Computer
Z-2211-2014 | 68763 | No | LGX | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General Hospital | 510(K) Exempt | Medical Positioning Rapid Response HUT table, Powered Medical Examination table, Model # 1219. |  |  | July 10, 2014 | 2014 | Medical Positioning Inc | Excessive weight on the table will cause tables with the affected component to drop into the Trendelenburg position more rapidly than normal. The product was manufactured with a component that can cause the device to malfunction. | Component design/selection | Medical Positioning Inc. sent Potential Product Hazard Bulletin letter dated July 8, 2014 to affected customers, via USPS certified return receipt mail. The letter identified the affected product, problem and actions to be taken. Customers were informed a Product Support Specialist will contact them to make arrangements for a Field Service Technician to perform a test designed to identify if the table component affects the table mechanics. If the table is identified as having the affected component, the component will be replaced. |  | 1 | 1 bed | Worldwide Distribution - US Nationwide in the states of CA, GA, FL, IL, IN, KS, KY, MD, MN, MO, MS, NC, NY, OH, TN, TX, VA and WI. | Terminated | January 03, 2017 | 908 | 40 | Not_Computer
Z-2199-2014 | 68788 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Prisma M100 PRE pump infusion set. Product number 107791. Each set is packaged in a plastic pouch, 4 pouches per carton box |  |  | July 18, 2014 | 2014 | Gambro Renal Products, Incorporated | Gambro initiated a field action on the Prisma and Prismaflex sets due to breakages of connectors resulting in leakage of liquid or blood. | Device Design | Gambro sent an Urgent Field Safety Notice dated July 18, 2014 to all end-users. A Customer Reply Form was attached to the notice for customers to complete and return. |  | 401451 | 401,451 | Worldwide Distribution. | Terminated | February 05, 2015 | 202 | 401451 | Not_Computer
Z-2212-2014 | 68651 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Device Brand Name: Distal Radius Fracture Repair System |  |  | June 23, 2014 | 2014 | Biomet, Inc. | The firm received complaints indicating that the drill guide becomes stuck in the jig. The drill guide can be placed in the jig about halfway before it becomes stuck. | Nonconforming Material/Component | On June 23, 2014 Biomet issued recall notification to their consignees via letter. Notification included affected product issue and instructions for return were provided. Distributors who had affected hospitals in their regions were notified and provided with a copy of the notice that was sent to the hospital. Consignees are being contacted on an individual basis in an attempt to reconcile the product 100%. A response form was provided to all consignees in order for Biomet to track returns. Non-responders will be contacted on an individual basis. |  | 37 | 37 | The product was distributed domestically and to internationally. Domestically to CA, PA, IA, CO, NY, MD, NC,AL, AZ, IL, GA, VA.Netherlands, Argentina, | Terminated | July 14, 2015 | 386 | 37 | Not_Computer
Z-2215-2014 | 68654 | No | JDG | Hip joint femoral (hemi-hip) metallic cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Moore Hip Prosthesis Long Fenestrated Stem 38 mm (1-1/2 Inch) Head Diameter 203 mm Length. |  |  | July 01, 2014 | 2014 | Zimmer, Inc. | Zimmer is initiating a voluntary recall of 20 production lots of Moore Hip Prosthesis Long Fenestrated Stems due to packaging design verification test failures. Specifically, 9 of 16 test samples failed simulated distribution and shipping testing when the Moore hip stem breached either the inner or outer package tray. Zimmer has not received any complaints of damaged package trays for these 20 lot | Packaging process control | On July 1, 2014 Zimmer, Inc. issued recall notification to their consignees via letter. This notice was issued to affected distributors and hospitals with instructions for returning the affected product. Distributors were sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory, as well as identifying hospitals and surgeons who have previously implanted the affected product and providing related contact information (name, address, and phone number). Distributors are to return on-hand affected products to Zimmer and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers and surgeons were provided with a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product and returning it to their Zimmer sales representative. |  | 68 | 68 | Worldwide Distribution - US (nationwide) in the states of IN, WI, MN, OR, and TX, and one foreign consignee in Canada. | Terminated | January 09, 2015 | 192 | 68 | Not_Computer
Z-2213-2014 | 68845 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Navigation Spine &amp; Trauma 3D Version 2.0 and 2.1 |  |  | May 21, 2014 | 2014 | Brainlab AG | Brainlab Navigation Software Spine &amp; Trauma 3D 2.0/2.1 offers automatic registration of intraoperatively acquired CT image data sets. To enable automatic registration, the software requires the gantry position of the scanner. The gantry position can either be entered manually or submitted automatically from the CT scanner. If automatic gantry communication is available, the so-called &quot; | Software in the Use Environment | The firm, Brainlab AG, sent "FIELD SAFETY NOTICE / PRODUCT NOTIFICATION" letters dated May 21, 2014 to all consignees. The letters included instructions for consignees to: 1) not use Automatic Registration of iCT Data in combination with Siemens SOMATOM Definition AS CT scan (with Software Somaris SOM 7 or higher and Automatic Gantry Communication (available for Spine &amp; Trauma 3D 2.0 and 2.1 only); 2) always make sure to perform careful accuracy verification according to the instructions in the Software User Guide; 3) Brainlab will perform corresponding service visits to change the gantry communication from automatic to manual to enable a safe and effective usage of this device combination, and 4) advise the appropriate personnel working in your department of the content of the letter.Customers with any questions can contact Brainlab via telephone at 800-597-5911 or e-mail at us.support@brainlab.com. |  | 17 | 17 systems | Worldwide distribution: US (nationwide) including states of: NC, NY, OH, WI;and countries of: Austria, China, Germany, Italy, Japan, Malaysia, and Saudi Arabia. | Open, Classified |  |  | 17 | Computer
Z-2155-2014 | 68806 | No | MCW | Intraluminal artery stripper. | Cardiovascular | Cardiovascular | 510(k) | CSI Cardiovascular Systems, Inc., Diamondback 360 Peripheral Orbital Atherectomy System, Model Number DBP-125MICRO145, Part Number 7-10003. |  |  | May 23, 2014 | 2014 | Cardiovascular Systems, Inc. | CSI has initiated a recall on the Diamondback 360 Peripheral Orbital Atherectomy Device because it may contain a saline sheath that may experience cracking, fracture, and release particulate during use. | Process control | Cardiovascular System, Inc. sent consignees an Urgent Medical Device Recall letter dated May 23, 2014. The letter described the Affected Product, Recall Description, Instructions which included to remove the affected product and return it to CSI and to complete and return the Customer Acknowledgement Form. For further Information they customers were instructed to contact Customer Service, Cardiovascular System, Inc., 877-274-0901.For questions regarding this recall call 877-274-0901. |  | 48 | 48 | Nationwide Distribution including AZ, AR, CA, CT, FL, IL, IA, MD, MI, NY, NC, PA, TN,and TX. | Terminated | October 01, 2014 | 131 | 48 | Not_Computer
Z-2236-2014 | 68831 | No | PEN | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | Verigene BC-GN Test Kit (Catalog number 20-005-021) includes: 1) 20 Verigene BC-GN Test Cartridges. Each Test Cartridge comes preloaded with all required reaction solutions, including wash solutions, oligonucleotide probe solution and signal amplification solutions required to generate a test result. BC-GN is indicated for use in conjunction with other clinical and laboratory findings to aid in the diagnosis of bacterial bloodstream infections; however, is not used to monitor these infections. |  |  | July 09, 2014 | 2014 | Nanosphere, Inc. | Nanosphere Inc. has recently determined through four customer complaints that a small number of Verigene Gram-Negative Blood Culture Nucleic Acid Test (BC-GN) Extraction Trays in Lot 042914021B have low levels or are missing Magnetic Microparticles (MMPs) required for DNA Extraction. | Under Investigation by firm | Nanosphere sent an Urgent Product Recall letter dated July 9, 2014 to all affected customers. All customers were initially notified of the recall by telephone on July 9, 2014. The customers were instructed to quarantine and return all affected product to Nanosphere. The letter included instructions: 1) on how to return the recalled products, and 2) to review any test results that were generated while using the recalled products to ensure that there has been no negative impact on patient care. Customers with questions were instructed to contact Nanosphere at 888-837-4436, option 2 or via e-mail at Technicalsupport@nanosphere.us. |  | 1045 | 1045 extraction trays (840 affected by recall) | Nationwide Distribution including CA, CO, IA, MT, OH, OK, OR, TX, and UT | Terminated | September 01, 2015 | 419 | 1045 | Not_Computer
Z-2233-2014 | 68892 | No | GKZ | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD Multi Check CD4 Low Control Catalog No. 340916 is intended as a complete process control for immunophenotyping by flow cytometry. |  |  | July 18, 2014 | 2014 | BD Biosciences, Systems &amp; Reagents | The outer box label contains the incorrect distribution date: Box states 2017-09-02 ; should be 2014-09-02 | Labeling mix-ups | Global hold initiated on July 18, 2014. Customer letters sent by certified mail on July 18, 2014 to advised users of the problem. |  | 85 | 85 | IN, MA, TX, VA, Ca, KY, MD, PA, GA, and BD locations in Singapore and Belgium | Terminated | May 19, 2015 | 305 | 85 | Not_Computer
Z-2232-2014 | 68747 | No | DAP | Fibrinogen/fibrin degradation products assay. | Hematology | Hematology | 510(k) | VIDAS D-Dimer Exclusion II, an automated quantitative test for use on instruments of the VIDAS family for the immunoenzymatic determination of fibrin degradation products in human plasma using the Enzyme Linked Fluorescent Assay technique. Also indicated for use in conjunction with clinical pretest probability assessment model to exclude DVT and PE disease. |  |  | April 17, 2013 | 2013 | BioMerieux SA | Customer complaints about not repeatable results below the detection limit (&lt;45 ng/mL). Therefore false negatives are being reported for results around the medical decision point of 500 ng/mL that initially tested at &lt;45 ng/mL. | Under Investigation by firm | A recall letter was issued on 4/17/2013 to US customers via FED Ex mailing. The letter included a fax back acknowledgment form. A recall notice was also sent via email to bioMerieux Subsidiaries and international distributors on 4/11/2013. |  | 2160 | 2160 units | AL, AR, CA, CO, CT, DE, FL, GA, ID, IL, IN, IA, KS, KY, LA, ME, MA, MI, MN, MO, MT, NE, NH, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, VT, VA, WA, WV, WI, WY | Terminated | October 08, 2014 | 539 | 2160 | Not_Computer
Z-2222-2014 | 68656 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Double Offset Rasp Handle, Right Hand |  |  | July 23, 2014 | 2014 | Zimmer, Inc. | The firm received complaints for failure of the handle to remain assembled to the rasp during use due to various failure modes, including fracture of the leaf spring. | Device Design | On July 23, 2014, Zimmer issued recall notification to their consignees. Distributors were notified via electronic mail and courier. Recall notification included a description of the issue, affected product, possible risks to patient, and instructions for responding to the recall. Questions or concerns regarding the recall are directed to Zimmer customer call center @ 1-800-348-2759 M-F, 8am-8pm ET. |  | 412 | 412 units | Worldwide Distribution-USA (nationwide) including the states of AK, AL, AZ, CA, CO, FL, GA, IN, KS, KY, MI, MN, NC, NJ, NY, OH, OK, OR, PA, SD, TN, TX, UT, and VA, and the countries of Australia, France, Germany, Italy, Switzerland, and Canada. | Terminated | September 15, 2015 | 419 | 412 | Not_Computer
Z-2237-2014 | 68953 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IntelliSpace Portal (ISP) DX/HX/EX, IntelliSpace Portal (ISP) IX, and IntelliSpace Portal (ISP) LX SPECT |  |  | July 27, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | The following software issues have been identified in the affected products.Problem 1: When reopening a bookmark generated from processing a MUGA (Multi-Gated Acquisition) scan within the NM Cardiac Application, the ejection fraction (EF) displayed should be the same as the ejection fraction (EF) originally displayed when the bookmark was first created. In some instances, the ejection fraction ( | Software design | Philips Healthcare issued an Urgent Medical Device Correction letter dated July 28, 2014, via certified mail to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Philips Key Market representative would distribute the FSN internationally. Additionally, a Field Service Engineer would contact each site to schedule a time to install the software update.Customers with questions were instructed to contact their local Philips representataive or local Philips Healthcare office. For North America and Canada customers were instructed to contact the Customer Care Solutions Center at 1-800-722-9377, option 1.For questions regarding this recall call 978-659-7703. |  | 1099 | 1099 Units | Worldwide Distribution - USA including AK, AL, AR, AZ, CA, CO, FL, GA, IL, IN, KS, LA, MA, MD, MI, MN, MO, MS, NC, ND, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, VT, WA, WI, WV, WY,Washington D.C. and Internationally to the following countries: Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Cayman Islands, Chile, China, Colombia, Congo, Costa Rica, Czech Republic, Denmark, Ecuador, Egypt, Ethiopia, Finland, France, Gabon, Germany, Greece, Guatemala, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kenya, Latvia, Lebanon, Libyan Arab Jamahiriya, Lithuania, Malaysia, Mauritius, Mexico, Mongolia, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Nicaragua, Norway, Oman, Pakistan, Palestine, Panama, Peru, Philippines, Poland, Portugal, Puerto Rico, Qatar, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Viet Nam, and Yemen. | Terminated | February 02, 2016 | 555 | 1099 | Computer
Z-2234-2014 | 68802 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Vital Sync" VPMP &amp; IM, software, V2.4. (DVD), packaged on a CD or other electronically controlled device. CDs are packaged in clam shell jeweled cases within a carton box. |  |  | July 17, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | Covidien is conducting a field correction of the Vital Sync Virtual Patient Monitoring Platform and Informatics Manager due to a mismatch between the alarm priority on the medical devices and the Vital Sync system. | Software design | The firm, Covidien, sent an "URGENT FIELD CORRECTIVE ACTION NOTICE" letter dated 7/16/14 to its customers on 7/17/14 via FedEx. The letter informs users that a representative will be in contact to schedule a software update.The letter also described the product, problem and actions to be taken. The customers were instructed to disregard the audible and visual alarm priority indicators and only use the alarm description text on the patient tile and on the patient detail page for alarm priority information, until the mismatch is corrected. The Vital Sync" Virtual Patient Monitoring Platform and Informatics Manager or Vital Sync" Remote Monitoring System may still be used until the software update is available.If you have questions about this letter or to report any issues with the Vital Sync" Virtual Patient Monitoring Platform and Informatics Manager or Vital Sync" Remote Monitoring System, contact our Technical Support Department at (800)635-5267, option 3, then option 1. |  | 3 | 3 | US distribution to states of: LA, IL, MS, CO, IN, GA, CA and TX. | Terminated | October 21, 2014 | 96 | 12 | Computer
Z-2238-2014 | 68848 | No | JMG | Ferritin immunological test system. | Immunology | Clinical Chemistry | 510(k) | Siemens Dimension Chemistry Wash (RD701) is an accessory of the flex cartridge system used on the Siemens Dimension RxL and Xpand assays. The wash buffer is used to clean the probe sensor between uses and prepares it for taking the next reading/measurement. |  |  | July 15, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received complaints for the Dimension Chemistry Wash (RD701) of QC and patient sample result accuracy shifts on Troponin I, CTNI (RF421C), LTNI (RF521C); Free T4, FT4 (RF410); Thyroid Stimulating Hormone, TSH (RF412); and N-terminal pro-brain natriuretic peptide, PBNP (RF423A) and LPBN (RF523A) when using Dimension&reg; Chemistry Wash lot RD41041.Internal inv | Material/Component Contamination | Siemens sent an Urgent Medical Device Recall letter dated July 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken by the Customer:1. Immediately discontinue use and discard all inventory of Chemistry Wash lotRD41041. If this lot was never in use on the instrument, no further action is needed. If this lot has been in use, continue with steps 2 through 5.2. Place an alternate lot of Chemistry Wash on the Dimension system.3. Prime the new bottle of Chemistry Wash 30 times.4. Recalibrate the CTNI, LTNI, TSH, FT4, PBNP and/or LPBN assays using the alternate lot of Chemistry Wash on board.5. For the listed assays, process Quality Control every 8 hours.Actions to be taken by Siemens:Your Customer Service Engineer will be making a visit to decontaminate the system if you have used lot RD41041. This process will take approximately two hours. Once decontamination has been completed, process and evaluate quality control. Resume routine QC testing frequency according to your standard laboratory practice.Siemens is shipping replacement product at no charge for all orders filled with lot RD41041. If you have questions about your replacement shipment, please contact Customer Service at (888) 588-3916.Please complete the attached form and fax it to (302) 631-8467 to indicate that you have received this information.Please retain this letter with your laboratory records. Please forward this letter to anyone to whom you may have provided this product.We apologize for the inconvenience this situation has caused. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. |  | 3200 | 3200 | USA (nationwide) Distribution and Internationally to Canada. | Terminated | August 26, 2015 | 407 | 3200 | Not_Computer
Z-2242-2014 | 68931 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard 360 Femoral Boss Reamer, |  |  | July 01, 2014 | 2014 | Biomet, Inc. | The Vanguard 360 Femoral Boss Reamers were manufactured oversized to the print. The instrument may not fit into the mating component. | Process change control | Urgent Medical Device Recall Notices dated July 15, 2014 were mailed with verification of delivery via FedEx.Response forms were provided to all consignees in order forBiomet to track returns. The letters informed the consignees to return the devices to Biomet and conduct a sub-recall sending the recall letter to all customers. |  | 461 | 461 | California, IA, FL, MD, UT, SC, AL, IL, CO, MA, KS, GA, WA, NV, AR, NH, OK,Pennsylvania, New York, Texas, Virginia, New Jersey,North Carolina, Kentucky, Michigan, Ohio, Wisconsin, Louisiana,Arizona, Indiana, South Dakota, Missouri &amp; Montana. In addition to Korea, B.V., Japan, Puerto Rico, UK, Chile, New Zealand, Australia, &amp; Canada | Terminated | October 30, 2015 | 486 | 461 | Not_Computer
Z-2240-2014 | 68960 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Treatment System, Model: Hi-Art, Catalog/Part Number: H-0000-0003, software versions 2.0.0 and 2.0.1 (Hi-ArtÂż 5.0.0 and 5.0.1). |  |  | July 16, 2014 | 2014 | TomoTherapy Incorporated | Accuray is voluntarily recalling TomoTherapy H Series software versions 2.0.0 and 2.0.1 (Hi-ArtÂż 5.0.0 and 5.0.1). Accuray has identified potential safety issues (anomalies) with these software versions. | Software design | Accuray sent an "Urgent Device Correction" letter dated July 16, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customers. These anomalies will be corrected with a new software version.You will be contacted by Accuray Customer Support to schedule a software upgrade to your system.If you have any questions or concerns regarding this issue, please do not hesitate to contact Accuray Customer Support at +1.877.668.8667 (USA) or +1.408.716.4700 (non-USA) or customersupport@accuray.com. |  | 168 | 168 (75 US; 93 OUS) | Worldwide Distribution - Nationwide including the states of AZ, AR, CA, FL, GA, IL, IN, KS, KY, MD, MN, NM, NY,NC, OK, OR, SD, TN, TX, UT, WA, WY, WI and WV., and the countries of AUSTRALIA, BELGIUM, CANADA, CHINA, ENGLAND, FRANCE, GERMANY, HONG KONG, ITALY, JAPAN, KOREA, MEXICO, NETHERLANDS, PHILIPPINES, RUSSIA, SPAIN, TAIWAN, TURKEY, MARTINIQUE, ROMANIA, SRI LANKA and SWITZERLAND. | Terminated | July 21, 2015 | 370 | 168 | Computer
Z-2239-2014 | 68948 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | TomoTherapy Treatment System, Model: Hi-Art, Catalog/Part Number: H-0000-0003, software versions 1.2.0, 1.2.1 and 1.2.2 (Hi-ArtÂż 4.2.0, 4.2.1 and 4.2.2). |  |  | July 17, 2014 | 2014 | TomoTherapy Incorporated | Accuray is voluntarily recalling TomoTherapy H Series software versions 1.2.0, 1.2.1 and 1.2.2 (Hi-ArtÂż 4.2.0, 4.2.1 and 4.2.2). Accuray has identified potential safety issues (anomalies) with these software versions. | Software design | Consignees were sent an Accuray "Urgent Device Correction" letter dated July 16, 2014. The letter was addressed to TomoTherapy System Medical Physicist. The letter described the Purpose of this letter, Description of the Potential Problems, Safety Instructions and software upgrade information. For questions or concerns they can contact Accuray Customer Support at 1-877-668-8667 (USA) or 1-408-716-4700(non-USA) or customersupport@accuray.com |  | 322 | 322 | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of US:; AL, AK, AZ, CA, CO, CT, FL, GA, HI, IL IN, IA, KS,KY, LA, MD, MA, MI, MO, MT, NE, NV, NJ, NM, NY, NC,ND, OH,OK, OR, PA, SC, SD, TX, UT, VA, WV, and WI, and the countries of AUSTRALIA, BELGIUM CANDA, CHINA, EL SALVADOR, ENGLAND, FRANCE, GERMANY, HONG KONG, INDIA, ITALY, JAPAN, KOREA, MALAYSIA, MEXICO, NETHERLANDS, PHILIPPINE, POLAND, RUSSIA, SAUDI ARABIA, SINGAPORE, SPAIN, SWEDEN, TAIWAN, THAILAND, TURKEY, and UNITED ARAB EMIRATES. | Terminated | September 08, 2015 | 418 | 322 | Computer
Z-2241-2014 | 68527 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Apex Arc Hip Stem, Size 2, HA Coated; Product Code: HC-21002-HA |  |  | May 09, 2014 | 2014 | Omnilife Science Inc. | Staining on hip implant. | Manufacturing material removal | All product was consigned to US sales Agents until implanted or returned to OMNI. All US Agents holding the inventory were notified on May9, 2014 by phone and follow up email that the product is being recalled. Customers have been instructed to immediately discontinue use and return the product to OMNIlife science. |  | 13 | 13 devices | US: CA, OK, VA, FL, MA, NY, TX | Terminated | August 19, 2014 | 102 | 13 | Not_Computer
Z-2244-2014 | 68819 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Magna-Fx Cannulated Screw Fixation System |  |  | July 22, 2014 | 2014 | Zimmer, Inc. | Zimmer is recalling various trauma screws that failed a leak test due to damage of the inner cavity only. | Package design/selection | On 7/22/14, Zimmer sent recall notification to the consignees explaining the issue, identifying affected product, providing the health risks, and instructing the consignees what to do for the recall. Questions or concerns, please contact Zimmer at 1-877-946-2761. |  | 47682 | 47,682 units total | National Distribution: AK, AL, AZ, CA, CT, DC, FL, GA, US VIRGIN ISLANDS, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, SC, SD, TN, TX, VA, WA, WI, WV, WYForeign Distribution: Canada, Mexico, ARGENTINA, AUSTRALIA, BELGIUM,BRAZIL, CHINA, EL SALVADOR, ENGLAND, GERMANY, INDIA, ITALY, JAPAN, JORDAN, KOREA, NEW ZEALAND, SAUDI ARABIA, SINGAPORE, SPAIN, SWITZERLAND, TAIWAN, UAE | Terminated | June 18, 2015 | 331 | 47682 | Not_Computer
Z-2252-2014 | 68979 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips IntelliSpace PACS 4.4 image management systems performs digital image processing, measurement and communication and storage. |  |  | July 24, 2014 | 2014 | Philips Healthcare Informatics, Inc. | A problem was detected in the Philips IntelliSpace PAC 4.4 software where if patient information is changed in the DICOM header and the study is exported before the image is updated in the PACS, the unchanged data may export causing the potential for diagnosis on incomplete data. | Software design | Urgent Medical Device Correction was sent to all affected sites on July 24, 2014 by certified mail. |  | 109 | 109 units | US and Australia, Belgium, Canada, Egypt, France, Germany, Netherlands, Oman, Saudi Arabia, Spain, Switzerland, United Arab Emirates, United Kingdom. | Terminated | February 11, 2016 | 567 | 109 | Computer
Z-2258-2014 | 68792 | No | JDS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | TRIGEN(R) INTERTAN(TM) 10MM X 18 CM NAIL, 125o, REF 71675201, QTY (1), STERILE R, Smith &amp; Nephew, Memphis, Inc., TN 38116 |  |  | July 02, 2014 | 2014 | Smith &amp; Nephew, Inc. | There was an error in manufacturing which resulted in the proximal lag screw hole being positioned incorrectly in several batches of TRIGEN(R); INTERTAN(TM); NAILS. | Under Investigation by firm | Smith &amp; Nephew Issued an Urgent Product Recall dated July 2, 2014 to all affected customers via E-mail and/or FedEx . The notification identified the effective product, problem and actions to be taken. Customers were instructed to: Inspect inventory and locate any unused affected devices, cease distribution and use of those products, quarantine them immediately, complete the last two columns in the Inventory Return Certification Form, and return any affected product to the address provided. For questions call 901-396-2121 |  | 601 | 601 units | Worldwide Distribution - US Nationwide and the countries of AUSTRALIA, CANADA, CHINA, INDIA, UAE, JAPAN, SOUTH AFRICA, SOUTH KOREA, UNITED KINGDOM. AUSTRIA, NORWAY, POLAND, PORTUGAL | Open, Classified |  |  | 601 | Not_Computer
Z-2253-2014 | 68898 | No | KWY | Hip joint femoral (hemi-hip) metal/polymer cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Pivot Bipolar Cup 22x42/43-51/53mm, REF 133-YYYY |  |  | July 11, 2014 | 2014 | Ortho Development Corporation | Ortho Development is recalling various Pivot Bipolar Cup, Patella, PS Tibial Insert, UC Tibila Insert, and CK Tibil Insert because the foil pouches did not properly seal, allowing air to seep back into the pouch after sealing which increases the risk that the material properties of the implant would be substantially degraded due to oxidation. | Packaging process control | Ortho Development notified all affected consignees/distributors of the recall via email on July 11, 2014 which included a list of the affected product (part numbers, lot numbers and quantities) under their control. An RMA number and return shipping label were also included with the email. Emails were followed up with phone calls to verify receipt of the notification. |  | 188 | 188 | Worldwide Distribution - US Nationwide and the country of Japan. No military/govt/VA consignees. | Terminated | October 06, 2014 | 87 | 227 | Not_Computer
Z-2264-2014 | 67814 | No | PDZ | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | All brochures printed before November 20, 2013 for the HyperBlue 1530 medical laser, including the product brochure and the online information regarding the product at www.hyperionMed.com before December 3, 2013. |  |  | November 20, 2013 | 2013 | Hyperion Medical | Brochures and web site information for the HyperBlue 1530 diode laser promoted features which had not been cleared in the product's premarket notification to the FDA. The web site and brochures had claims that it could use 940 nm, 808 nm, and 532 nm wavelengths as single wavelengths or in combination with other wavelengths, and a claim that the product could be used with a 15 mm hand piece. | Labeling False and Misleading | Hyperion Medical notified all sales representatives by private courier letter to stop distributing the the sales brochure and return or destroy existing stocks of the brochure. Sales representatives were instructed not to demonstrate the 15 mm hand piece with the HyberBlue 1530 and to remove from any of their individual internet web sites any information about multiple wavelength use or the 15 mm handpiece accessory. The letter to the sales representatives included delivery confirmation. The information on the firm's web site for this product regarding multiple wavelengths and the 15 mm handpiece was removed.For further information please call (214) 914-3133. |  | 3346 | 3,346 brochures. Quantity not applicable to web site. | US Distribution including the states of PA, MD, UT, TX, OH, GA, MO, FL, CO, and WA. | Terminated | May 15, 2015 | 541 | 3346 | Not_Computer
Z-2263-2014 | 68870 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | WaterLase iPlus Dental Laser System, Part No. 7200854-XX |  |  | June 12, 2014 | 2014 | Biolase Technology Inc | Biolase is recalling the WaterLase iPlus Dental Laser System because of a discrepancy present in the user interface software (version 1.7). If an error occurs in the procedure screen, the air setting will change to 1% and OFF. | Under Investigation by firm | An Urgent Medical Device Correction letter dated 6/12/2014 was sent to all customers who purchased the WaterLase iPlus Dental Laser System. The letter informs the customers of the problems identified and the actions to be taken. Customers are informed that a Field Service Engineer can install the software update. Customers with are instructed to contact Customer Service at (800) 321-6717. |  | 117 | 117 units | Nationwide in US and worldwide to: Canada, UK, Russia, Korea, Hong Kong, Libya, Iran, Iraq, Kuwait, Singapore, Armenia, Egypt, Pakistan, Sweden, New Zealand, Poland, Germany, Italy, UAE. | Terminated | October 06, 2014 | 116 | 117 | Computer
Z-2259-2014 | 68935 | No | MHY |  |  | Neurology | PMA | Medtronic, Activa PC, Model 37601, Method of Sterilization: Ethylene Oxide, Single Use Only, Rx Only. |  |  | July 14, 2014 | 2014 | Medtronic Neuromodulation | Medtronic is recalling six Activa PC (model 37601) Implantable Neurostimulators due to the potential for a damaged electrical component during manufacturing. | Process control | The firm, Medtronic, notified their Consignees on 07/14/2014 via telephone of the recall. Medtronic representative used telephone script to convey the information. The script was directed to Risk Management or Inventory Management. The caller was to inform consignees of the problem and product being recalled. Advised consignees to quarantine the product and provided the Medtronic Device Removal Reply Form to the consignees via e-mail or Fax. The Reply Form included contact information which was to call 1-800-633-8766 in case they needed to contact a Medtronic representative. The completed form is to be faxed back to 1-800-897-3899 or e-mail a PDF to neuro.quality@medtronic.com.If you have any questions, call 763-526-1294. |  | 6 | 6 | Distributed in the states of: MA, NC, OH, and TX. | Terminated | October 06, 2014 | 84 | 6 | Not_Computer
Z-2266-2014 | 68952 | No | DSY | Vascular graft prosthesis. | Cardiovascular | Cardiovascular | 510(k) | Gelweave (TM) Vascular Prostheses |  |  | July 15, 2014 | 2014 | Vascutek, Ltd. | The cautery that is provided with the Gelweave graft in the box has a shorter shelf life indicated on its label, versus that of the graft (4 years and 5 years, respectively). | Incorrect or no expiration date | Vascutek sent an Urgent Field Safety Notice dated July 15, 2014, to all affect customers. The letter identified the product the problem and the action needed to be taken by the customer. User Instructions:A. Vascutek advises that you ensure suitable equipment is available prior to implantation to trim the graft if required.B. For product received with a cautery in the interim, Vascutek request that you check the shelf life of each product before every operation. If you find that the cautery is out of date, remove and discard the cautery. If you require an appropriate cautery supplied as a separate item, please contact your local sales representative for a referral.C.Vascutek are taking this opportunity to reinforce the cautions from the Gelweave TM Vascular Prostheses IFU regarding soaking of the graft in saline to prevent focal burning if a cautery is to be used. This statement is already included in the IFU and now appears in bold, red font to ensure appropriate awareness.This notice needs to be passed on to all persons who need to be aware within your organization or to any organization where the devices have been transferred or distributed.Please maintain awareness of this Field Safety Notice while all actions are taken in your organization and until these products are provided without the cautery.Please return the User Return Slip by e-mail or fax to the Distributor's address given on page 3.This action by Vascutek Ltd. is being taken with the knowledge of the National Competent AuthorityÂż Medicines and Healthcare Products Regulatory Agency (MHRA).If you have any further questions or comments, please do not hesitate to contact us at tcvs.recall@terumomedical.com. Return Completed Form immediately fax to (734) 741-6149. If required, your Vascutek Ltd. representative can discuss and provide more information on the use of cautery with Vascutek products. For further questions call (800) 262-3304 ext. 6056 |  | 63162 | 63,162 | Nationwide Distribution including the states of : AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI and WV. | Terminated | March 19, 2015 | 247 | 63162 | Not_Computer
Z-2260-2014 | 68980 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Uni*Fuse INFUSION SYSTEM WITH COOPER WIRE, 5F x 90 cm x 10 cm Infusion Pattern, Catalog No./REF 12401806, Product No. H787124018065, STERILE -- AngioDynamics, 603 Queensbury Avenue, Queensbury, NY 12804 |  |  | August 06, 2014 | 2014 | Angiodynamics, Inc. | AngioDynamics became aware that Uni*Fuse Infusion Systems with Cooper Wire have the potential to contain an incorrect sized guidewire (too long). | Component design/selection | AngioDynamics sent an URGENT MEDICAL DEVICE RECALL Letter dated August 6, 2014, to all affected customers via Federal Express. Consignees were instructed to segregate and return all affected devices to AngioDynamics,Inc. and complete and return the Reply Verification Tracking Form, provided in the recall notification. Those customers with the affected product were also instructed to call AngioDynamics Customer Service at 1-800-772-6446.For questions regarding this recall call 718-742-4571. |  | 97 | Domestic: 70 units; Foreign: 27 units | Worldwide Distribution - USA (nationwide) and Internationally to Austria, Australia, Ireland, Malaysia, Sweden, and Switzerland. | Terminated | March 27, 2017 | 964 | 227 | Not_Computer
Z-2268-2014 | 68735 | No | DRT | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Blood pressure (BP) circuit of the ANSAR ANX-3.0 Autonomic Monitor system. |  |  | July 01, 2014 | 2014 | Ansar Group, Inc. Dba Ansar Medical Technologies, Inc | The blood pressure (BP) circuit occasionally fails to inflate the blood pressure cuff and the BP circuit fails to activate, resulting in no BP measured. | Unknown/Undetermined by firm | Recall notification letters sent to customers on July 8-9, 2014 by registered mail. |  | 351 | 351 | Nationwide Distribution. | Terminated | August 10, 2015 | 405 | 351 | Not_Computer
Z-2270-2014 | 68956 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips IP5 Information Portal |  |  | June 30, 2014 | 2014 | Invivo Corporation | The wireless function of the IP5 radio could be interrupted causing loss of wireless communication between the IP5 and host monitoring system. | Process design | The firm, Philips Healthcare, sent an "URGENT - Field Safety Notice" dated June 30, 2014 to its customers. The notice describes the product, problem and actions to be taken. The notice also provides the customers with instructions for use while waiting for the update to the IP5. A Philips representative will contact you regarding your affected device.If you need any further information or support concerning this issue, please contact your local Philips representative: 877-468-4861 option 1, then option 2. |  | 573 | 573 devices. | Worldwide distribution: US (nationwide) including states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, NE, NH, NJ, MN, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, WV, Hawaii, Alaska, and District of Columbia; and countries of: Australia, Austria, Denmark, France, Germany, Ireland, Netherlands, Norway, Poland, Saudi Arabia, Singapore, Slovenia, Spain, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 573 | Not_Computer
Z-2267-2014 | 68719 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | IntelePACS software versions 3-8-1 to 4-9-1, |  |  | July 03, 2014 | 2014 | Intelerad Medical Systems | In specific situation a partially truncated file may be written to the IntelePACS server. This partially truncated file still can be viewed by InteleViewer and InteleViewer will report to the user that the file in question is complete. This situation only occurs if a file is partially truncated. | Software design | Intelerad sent an Urgent Recall for Product Correction letter to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were informed that a new software release will be installed on the system. Afterwards, an audit will be run on the the entire PACS to identify all image data that may have been truncated. Customers were requested to complete and return the attached Client Response Form. |  | 230 | 230 | Worldwide Distribution - USA Nationwide and the countries of Scotland, Ireland, United Kingdom, Canada, Australia, and New Zealand. | Terminated | January 13, 2016 | 559 | 230 | Computer
Z-2269-2014 | 68768 | No | KNH | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Gyrus ACMI Falope-Ring Dilator, item 000878-4. The Falope-Ring Dilator is a conical plastic device that is a component of the Falope-Ring Band Contraceptive Tubal Occlusion Device procedure pack (FRB-30). The Falope-Ring Band Contraceptive Tubal Occlusion Device (TOD) consists of a cylindrical radiopaque silicone band that occludes the lumen of the Fallopian tube when placed over a looped segment of the tube. The FRB-30 (used with the reusable applicators) consists of one reusable Falope-Ring Guide, an instruction manual and 30 procedure kits. Each kit is provided sterile and includes one dilator and two Falope Bands. |  |  | July 08, 2014 | 2014 | Gyrus Medical, Inc | Falope-Ring Dilator for Reusable Falope-Ring Band Applicators have been sold as a stand-alone non-sterile component. The dilator is intended to assist placement of the Falope-Ring band on the reusable Falope applicator. | Packaging process control | Olympus sent a "Medical Device Recall" letter dated July 7, 2014 to affected customers. The letter was addressed to Surgery Department &amp; Risk Management Department. The letter described the product being recalled, problem and actions to be take. Consignees were advised to immediately cease further use of any affected product and to remove and discard it. Customers were to fax the completed questionnaire to 484-896-7128. For questions contact Laura Storms at 484-896-5688 or at laura.storms@olympus.com |  | 507 | 507 units | Worldwide - US Nationwide Distribution in the states of AL, AR, CA, CO, CT, DE, FL, GA, IA, IL, IN, LA, MD, MI, ,MN, MS, NH, HY, OH,OR, PA, TN TX, UT, VA, WA, WI,and WV. and the country of CANADA. | Terminated | May 07, 2015 | 303 | 507 | Not_Computer
Z-2265-2014 | 68738 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Synthes 3.0 MM TI Cannulated Screw Instrument and Implant Set |  |  | June 30, 2014 | 2014 | Synthes, Inc. | A screw set was returned from the field for destruction after being exposed to flood water, which was erroneously reprocessed and returned to the field. Initial investigative testing indicated the components within the set may be cytotoxic. | Employee error | An Urgent Field Safety Notification, dated July 21, 2014, was sent to surgeons explaining the product, problem, and action to be taken with regard to routine post-operative follow-up on patients exposed to the affected product. Questions or concerns should be directed to Leonie Rynn at 610-719-5408 or lrynn1@its.jnj.com. |  | 4 | 4 | Distributed in the states of WI and MI. | Terminated | December 18, 2015 | 536 | 4 | Not_Computer
Z-2295-2014 | 68973 | No | GAX | Nonpneumatic tourniquet. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Port Access Kit with MicroClave Clear Connector, Item No. B79104 |  |  | August 05, 2014 | 2014 | ICU Medical, Inc. | ICU Medical is recalling the BD PosiFlush Sterile Field Saline Flush Syringe because some unit packages may exhibit open seals which impacts package sterility and potentially product sterility. | Other | An Urgent recall letter dated 8/5/2014 was sent to all customers who purchased the BD PosiFlush SF Saline Flush Syringe 10ml. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions are instructed to contact ICU Customer Service Monday through Friday between 8:30AM to 4:00PM PST at (866) 829-9025 and select option 8, or email customerservice@icumed.com. Customers are instructed to complete the attached Medical Device Recall Response Form and fax to (801) 264-1755 or email to recall@icumed.com. |  | 400 | 400 units | Nationwide Distribution - USA including AL and MD. | Terminated | February 26, 2015 | 205 | 1000 | Not_Computer
Z-2288-2014 | 68653 | No | FSM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PERSONA (TM) The Personalized Knee System Tibial General Instrument Tray, non-sterile |  |  | July 01, 2014 | 2014 | Zimmer, Inc. | The firm discovered that the current Personal Tibial/General Fixed Bracket Trays may not have the specific geometry to accommodate the tibial drill and stop guide with proper position/orientation. The drill bracket and drill guide are not the correct geometry to allow for a proper fit in the Persona Fixed Case Kit which could lead to improper sterilization of the tibial drill and stop guide. | Device Design | On July 1, 2014, Zimmer, Inc. issued recall notification to their consignees via e-mail and/or letter. Notification included affected product issue and instructions for returning the affected product. All distributors were notified via electronic mail. Hospital risk managers, as well as distributors with affected inventory were also notified via courier. Distributors were sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory. Distributors are responsible for returning on-hand affected products to Zimmer and ensuring all of their products are accounted for using the form provided in the letter. Hospital risk managers were provided with a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product, quarantining the product, and returning it to their Zimmer sales representative. |  | 190 | 190 units | Worldwide Distribution - USA including AZ, CA, CO, FL, IL, IN, KS, MI, MN, NC, NJ, NY, OH, OR, PA, TN, TX &amp; VA and Internationally to AUSTRALIA, FRANCE, INDIA, ITALY, and GERMANY. | Terminated | June 09, 2015 | 343 | 190 | Not_Computer
Z-2192-2014 | 68443 | protector, skin pressure | FMP | Skin pressure protectors. | General Hospital | General Hospital | 510(K) Exempt | Gel-E Donut, Squishon 2 gel pillows |  |  | May 28, 2014 | 2014 | Childrens Medical Ventures | The gel-filled Gel-E Donut and Squishon product line has received a number of complaints about visible mold. The mold detected was determined to be Cladosporium and Penicillium Fungi, which are commonly found molds. | Material/Component Contamination | Childrens Medical Ventures sent an Urgent Medical Device Recall letter dated May 28, 2014, to affected customers via UPS. The letter identified the affected product, problem, and actions to be taken. Customers were instructed to review entire inventory and inspect for signs of visible mold, do not open the packaging to do the inspection, the mold should be seen through the transparent outer packing, complete and return the reply form provided on the last page of the letter, dispose of any product containing mold per your facilitys environmental guidelines, do not return any product to your Distributor or to Philips/Childrens Medical Ventures.and credit for disposed product will be given. If you need any further information or support concerning this problem, please contact your Distributor at the following phone number: |  | 332649 | 332,649 | Worldwide Distribution - USA (nationwide) including the country of Canada. | Terminated | November 29, 2016 | 916 | 332649 | Not_Computer
Z-2148-2014 | 68071 | No | MQN | Bone plate. | Dental | Dental | 510(k) | DePuy Synthes Craniomaxillofacial Distraction System (AB Distractor Bodies and BC Distractor Bodies) |  |  | April 16, 2014 | 2014 | Synthes, Inc. | DePuy Synthes is initiating a recall of certain lots of the Craniomaxillofacial Distraction System (AB Distractor Bodies and BC Distractor Bodies) because they may reverse post-operatively. | Device Design | The firm, DePuy Synthes, sent "Urgent Notice: Medical Device Recall" letters dated April 16, 2014 to affected customers. The letter identified the affected product, reason for recall, potential patient impact, mitigation steps for patients with distractors currently implanted and actions to be taken. Customers were instructed to review inventory, immediately remove the affected lots from stock, complete verification section of the letter and return the affected product. For questions call 610-719-5450 or contact your DePuy Synthes Sales consultant. |  | 1499 | 1499 | US Nationwide Distribution | Terminated | September 09, 2015 | 511 | 1499 | Not_Computer
Z-2291-2014 | 68966 | No | LZE | Nonpowered goniometer. | Orthopedic | Orthopedic | 510(K) Exempt | GE Healthcare, Discovery XR656P/N 5374989/5397837 is intended to generate digital radiographic images of the skull, spinal column, chest, abdomen, extremities and other body parts in patients of all ages. |  |  | July 25, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue involving the nylon hooks which support the FlashPad" detector on the Discovery XR656 wall stand systems.There have been reported incidents of detector falls, but no injuries have been reported. | Nonconforming Material/Component | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter dated July 25, 2014. The letter was addressed to Hospital Administrators / Risk Mangers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. For questions in the USA call 800-437-1171. For other countries contact your local GE Healthcare Representative. |  | 334 | 334 (162 US, 172 OUS). | USA: AL, CA, CO, FL, GA, HI, IL, IN, KS, LA, MD,MI, MN, MO, MT, NJ, NM, NY, NC, OH, OR, PA, SC, TN, TX, UT, VT, VA, WA, WV, WI.OUS: ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BERMUDA, CANADA, CHILE CHINA, DENMARK, FINLAND, FRANCE, GERMANY, HUNGARY, INDIA, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, MOROCCO, NEW ZEALAND, NORWAY, PAKISTAN, RUSSIA, SAUDI ARABIA, SLOVENIA, SPAIN, SWEDEN,THAILAND, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY. | Terminated | February 20, 2015 | 210 | 334 | Not_Computer
Z-2298-2014 | 68817 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona (TM) The Personalized Knee System |  |  | June 10, 2013 | 2013 | Zimmer, Inc. | Zimmer had received complaints that reported breakage of the Persona Knee Trial Articular Surface Provisionals (TASPs). This is a retrospective recall which covers Zimmer's corrective actions taken earlier this year concerning the TASPs. | Device Design | On February 10, 2014, Zimmer sent letter to surgeons who had filed complaints regarding the affected TASP fractures. The letter provided the surgeons with Zimmer's investigational findings of the issue. Surgeons were directed to send questions or any additional informtion to Zimmer.PER@Zimmer.com. The Zimmer Weekly Wrap-Up, dated February 14, 2014, was issued electronically to sales force. They were notified to engage in mandatory training to improve the usage of the TASPs and prevent potential breakage. Zimmer has taken additional corrective measures by initiating another recall in July 2014 whereby they notify the distributors as well as all users of the enhanced surgical technique. |  | 70986 | 70,986 distribution events | Worldwide Distribution-USA (nationwide) including DC and Puerto Rico and the states of AK, AL, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, MS, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SD, TN, TX, UT, VA, WA, WI, WY, and the countries of Australia, Canada, Germany, India, Japan, Korea, and Singapore. | Terminated | September 11, 2014 | 458 | 70986 | Not_Computer
Z-2535-2014 | 68983 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Philips IntelliSpace ECG Management System, with software option C61 needed to import ECGs from the Quinton Stress ECG System. |  |  | August 07, 2014 | 2014 | Philips Medical Systems, Inc. | Philips IntelliSpace ECG (IECG) Management Systems that are importing Stress ECG records from a Quinton Stress ECG system have the potential for the patient record to contain multiple and therefore inaccurate patient identifiers such as: patient name, patient ID (PID), Medical record number (MRN). This creates the potential for misdiagnosis and incorrect therapy. | Software design | Philips Healthcare sent an Urgent Medical Device Correction Notification / Field Safety notice dated August 7, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Field Safety Notice is intended to inform you about:" what the issue is and under what conditions it can occur" the actions that should be taken by the customer/user in order to prevent risks for patients" the corrective action planned by Philips to address the following issue1)Previously Imported Stress Records:Institutions that have imported Quinton Stress ECG records into IECG should review each previously imported record to make sure that multiple patient ID data is not present. Institutions that imported stress records at the end of each stress exam or did not batch process the records should not find any records with multiple Patient IDs in the record. However, we do recommend that you perform the review to confirm that multiple Patient IDs do not exist.2) Continued Use: Software is now available to address the issue described in this letter. If you have not already been contacted by Philips please call to arrange for immediate installation of the software upgrade.The correction will consist of free of charge software upgrade to A.01.06.. A Philips Healthcare representative will contact customers with affected devices to arrange for service. Further information or support concerning this issue, please contact your local Philips representative or call us at 1-800-722-9377 |  | 11 | 11 units | Worldwide Distribution - US including the states of CA, GA, NC, NH, NY and OR., and the countries of Canada and Spain. | Terminated | October 21, 2015 | 440 | 11 | Computer
Z-2290-2014 | 68753 | [6 empty] | PENTAX Upper/Lower G.I. Endoscope OF-B130 and OF-B194 Gas/Water Feeding Valve |  |  | July 17, 2014 | 2014 | Pentax Medical Company | The Instructions for Use for the Pentax OF-B130 and OF-B194 Gas/Water Valve IFU if not followed carefully may cause a potential risk to health. | Labeling design | Pentax Medical sent an Urgent Medical Device Recall Field Safety Corrective Action letter dated July 16, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.If the Instruction for Use are not followed carefully by the user there is a potential risk to health for the patient, such as a perforation.Customers are asked to complete the attached response form and fax to (201) 799-4063. Customers should see the attached revised Instructions for Use Z632 R09. If you have any questions regarding this action, please feel free to contact your Sales Representative at : 800-431-5880 ext 2064 |  | 1206 | 98 units US; 1108 units OUS | Worldwide Distribution - US Distribution including the states of CA, CO, IN, KS and ID., and the countries of Asia, Canada, Europe and Japan. | Terminated | August 20, 2015 | 399 | 1206 | Not_Computer
Z-2542-2014 | 68730 | No | JCF | Lymphocyte separation medium. | Hematology | Hematology | 510(K) Exempt | BD Vacutainer CPT Cell Preparation Tube with Sodium Heparin, 8 mL, 16 x 125 mm, Catalog #REF 362753, Lot #4063409 with expiry 03/2015, intended for the collection of whole blood and the separation of mononuclear cells. |  |  | July 03, 2014 | 2014 | Becton Dickinson &amp; Company | BD has initiated the recall of its Vacutainer CPT Cell Preparation Tube due to poor separation, resulting in a reduced number of recovered mononuclear cells. | Under Investigation by firm | BD issued an "Urgent Voluntary Product Recall" Letter dated July 3, 2014 and a "Notice of Return" Form to all affected customers. |  | 296160 | 296,160 tubes | US and Foreign | Open, Classified |  |  | 296160 | Not_Computer
Z-2543-2014 | 68801 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Zimmer PERSONA Distal Valgus Alignment Guide, Nonsterile. orthopedic surgical instrument. |  |  | July 30, 2014 | 2014 | Zimmer, Inc. | Persona Distal Valgus Alignment guide collet locks are seizing in the 'unlocked' position. | Device Design | An Urgent Medical Device Correction letter dated, July 31, 2014, was issued to consignees. The letter identified the affected product, the problem, the risks, and responsibilities. Risk managers and surgeons are to review the notification and the updated surgical technique provided in Attachment 1; as well as, ensure that all affected personnel are aware of the contents. Questions or concerns should be directed to Zimmer customer call center at 1-877-946-2761. |  | 2248 | 2,248 units | Worldwide Distribution -- United States, including Puerto Rico; and, to the countries of Australia, Austria, Japan, Belgium, Italy, Switzerland, India, Germany, Spain, France, Israel, Netherlands, Saudi Arabia, South Africa &amp; Canada. | Terminated | March 06, 2015 | 219 | 2248 | Not_Computer
Z-2544-2014 | 68277 | No | OYC |  |  | Clinical Chemistry | PMA | Medtronic MiniMed Paradigm Insulin Infusion Pumps, |  |  | March 13, 2014 | 2014 | Medtronic MiniMed Inc. | Medtronic is recalling the Medtronic MiniMed Paradigm Insulin Infusion Pumps because they have received reports that users have accidentally programmed the pump to deliver the maximum bolus amount. | Device Design | The firm, Medtronic, sent an "IMPORTANT MEDICAL DEVICE SAFETY INFORMATION" letter dated March 13, 2014 to its customers. The letter described the product, problem and actions to be taken. The letter informs the customers that because accidental button pressing errors may occur it is important that customers always confirm the insulin dose flashing on the display is correct before pressing ACT to start delivery. The letter informs the customers of what safety settings will minimize the risk of an accidental programming error and what actions to be taken. Should you have any questions about this letter or concerns with your keypad, please do not hesitate to contact the 24 hour helpline at (866) 222-7304 for a replacement. |  | 559374 | 559,374 units total (444,374 units in US) | Worldwide Distribution-Nationwide in US | Terminated | April 07, 2015 | 390 | 559374 | Not_Computer
Z-2553-2014 | 68854 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension EXL Integrated Chemistry System in vitro diagnostic devices intended to measure a variety of analytes in human body fluids. All Serial #s of material #s: 10636929, 10636928, 10486890, 10472176, 10639337, and 10639338. |  |  | July 16, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | The Reagent Lid hinge may lose its effectiveness and slowly shift downward during maintenance procedures, potentially resulting in physical contact and injury to the operator. | Device Design | Siemens notified end users of a potential safety issue with the Reagent Lid hinge on the Dimension EXL instrument via an Urgent Medical Device Correction letter, dated July 2014. The letter indicated Customer Service Engineers will inspect and adjust the tension on the reagent lid hinges during every visit and replace hinges which no longer properly adjust. Users were urged to use caution when the Reagent Lid is open and notify Siemens the Reagent Lid is shifting down. |  | 2757 | 2757 | Nationwide and internationally, including Canada and Mexico | Terminated | December 16, 2014 | 153 | 2757 | Not_Computer
Z-2550-2014 | 68818 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Trauma Tubes and Plates: |  |  | July 22, 2014 | 2014 | Zimmer, Inc. | Zimmer is recalling sterile-packaged implants and instruments packaged in six different package configurations due to packaging design verification test failures. Specifically, multiple test samples fromeach of the six configurations failed simulated distribution and shipping testing when the devices either breached inner or outer package trays or seal integrity was lost. | Package design/selection | Zimmer sent an URGENT MEDICAL DEVICE RECALL letter dated July 2014 to all affected customers. The letter identified the product, problem, and actions to be taken by the customers. Please call the Customer Call Center at 1-800-348-2759 for questions or concerns involving this notice. |  | 972125 | 972,125 units | Worldwide Distribution. | Terminated | September 15, 2015 | 420 | 972125 | Not_Computer
Z-2554-2014 | 68961 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | CURBELL MEDICAL CareSense Wireless Advanced LCD Bed and Chair Monitor, MODEL # CSM-BC400; --- COMMON/USUAL NAME: Bed Patient Monitor; CLASSIFICATION NAME: Bed Patient Monitor --- Monitor is connected to a sensor pad that monitors the patient if they attempt to arise from a bed, chair, floor mat, or toilet seat. The monitor will sound an alarm. |  |  | August 05, 2014 | 2014 | Curbell Medical, Inc. | The firm became aware of a potential problem that was initiated by a customer complaint. After consultation with the manufacturer, it was discovered that a resistor was incorrectly placed within the circuit board on the monitor. This change to the resistor was a planned change to address a product improvement (improve battery drain). However, this modification rendered the monitor's alarm via t | Component design/selection | Urgent Medical Device Recall Letters (dated 8/01/14) and Response Forms were sent to the consignees via e-mail (w/read receipt) on 8/05/14. |  | 323 | Domestic: 112 units; Foreign: 211 units | Nationwide, Canada, Ireland, Italy | Open, Classified |  |  | 323 | Not_Computer
Z-2281-2014 | 68346 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | AMO TECNIS CL Foldable Silicone IOL with OptiEdge Design, Model Z9002. Indicated for primary implantation for the visual correction of aphakia in adult patients. Tecnis CL Mono 13 mm with the following diopter sizes: 19.5-20 |  |  | May 13, 2014 | 2014 | AMO Puerto Rico Manufacturing, Inc. | Units may be labeled with the incorrect diopter power. | Error in labeling | AMO sent an Urgent Medical Device Recall dated May 13, 2014 to the affected customers. The letter identified the affected product, problem and actions to be taken. This recall only affects the specific Silicone Tecnis IOLs Model Z9002. Implanted IOLs with appropriate post-op refractive outcomes are not affected by this recall. Please underatake the following actions:1. Stop using and remove from your inventory the affected product.2.Complete and return the included Facsimile Form as AMO. If you have inventory to return, please complete the Facsimile Form, and returned to AMO. For questions contact AMO Customer Service Representative 877-266-4543. |  | 77055 | 77,055 IOLs in total | Worldwide Distribution - Nationwide (USA and Puerto Rico), and foreign: Latin America, Asia Pacific, Canada, Europe, Africa and Middle East, and Japan. | Open, Classified |  |  | 77055 | Not_Computer
Z-2545-2014 | 68934 | No | KRR | Knee joint patellofemoral polymer/metal semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Gender Solution Patello-Femoral Joint (PFJ) System Patello-Femoral Trochlea Component, Precoat, Size 4, Left. Knee prosthesis. |  |  | July 30, 2014 | 2014 | Zimmer, Inc. | Internal review found that five (5) PFJ implant forgings were identified as have forged laps during receiving inspection process for raw material. These five forgings were not segregated and were subsequently issued to final manufacturing under Part 00-5926-014-01; lot 62455332. A forged lap is a defect that forms when metal folds over itself during forging. | Nonconforming Material/Component | Zimmer sent an URGENT MEDICAL DEVICE RECALL letter dated July 2014 to affected distributors and hospital accounts. The letter identified the affected product, problem and the actions to be taken. The letter instructed customers to 1. Review the notification and ensure affected personnel are aware of the contents, 2. Locate and quarantine any affected product listed above and notify your Zimmer sales representative, and 3. Your Zimmer sales representative will remove the recalled product from your facility. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 15 | 15 units | Worldwide Distribution - US (nationwide) in the states of HI, CA, GA, MI, and IL and the countries of TAIWAN, BELGIUM, AUSTRALIA, FRANCE, JAPAN, and UK. | Terminated | January 20, 2015 | 174 | 15 | Not_Computer
Z-2243-2014 | 68914 | No | MMX | Embolectomy catheter. | Cardiovascular | Cardiovascular | 510(k) | CloverSnare 4-Loop Vascular Retrieval Snare. |  |  | July 17, 2014 | 2014 | Cook Inc. | This remedial action is the result of six product complaints associated with separation of the snare from the distal tip of the wire. | Device Design | Cook Medical sent an Urgent Medical Device Recall letter dated July 17, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review the attached list of affected products and lot numbers shipped to their account, and quarantine any affected product that remains unused. Immediately collect and return all unused affected products to Cook Medical as soon as possible for credit. Complete the attached Recalled Product Reply Form and return to Cook Medical either with the product or separately. Customers were asked to report any Adverse Event to Cook Medical Customer Relations at 800-457-4500 or 1-812-339-2235. |  | 696 | 696 devices total distribution (671 nationwide) | Worldwide Distribution - USA including AK, AL, AZ, CA, CT, DC, FL, GA, HI, IA, ID, IL, IN, KY, LA, MD, ME, MI, MN, MO, MS, NC, NE, NH, NJ, NV, NY, OH, OK, PA, SC, TX, UT, VA, WA, WI, WV and Internationally to:Canada, Austria, Belgium, Denmark, Germany, Great Britain, Ireland, Italy, Spain, Sweden, and Switzerland. | Terminated | June 16, 2015 | 334 | 696 | Not_Computer
Z-2555-2014 | 68986 | No | OJR | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | The Thermedx Fluid Management System is designed to provide irrigation fluids to irrigate, distend, and debride while warming the fluid to assist in the prevention of peri-operative hypothermia. The fluid is warmed via an internal infrared mechanism. Additionally the system will hold suction canisters for the collection of used irrigant. |  |  | February 10, 2014 | 2014 | Thermedx LLC | A vendor changed the manufacturing process of a component used in the canister ring that may affect the performance of the ring. The deficit display may be higher than the actual deficit when the fluid collection canister is fully loaded. | Component design/selection | Thermedx sent an Urgent Product Correction letter dated March 15, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were notified that the devices were in need of correction, and that it would be performed at their facilities by Thermedx. Customers with questions were instructed to call a Thermedx technical service representative at 1-888-542-9276.For questions regarding this recall call 440-542-0883. |  | 21 | 21 | Nationwide Distribution including HI, MS, NY, OH, PA, VA, WA, WV, and IL. | Terminated | August 04, 2015 | 540 | 21 | Not_Computer
Z-2559-2014 | 68710 | No | NTI | Drug metabolizing enzyme genotyping system. | Toxicology | Toxicology | 510(k) | Verigene CYP2C19 Nucleic Acid Test (CYP2C19 Test). Each CYP2C19 kit consists of 20 Verigene CYP2C19 Nucleic Acid Test Cartridges and 20 Verigene CYP2C19 Nucleic Acid Extraction Trays (with Tip Holder Assemblies). Clinical laboratory test. |  |  | March 18, 2014 | 2014 | Nanosphere, Inc. | Nanosphere has recieved reports related to an unacceptable increase in initial no-call rate when performing the Verigene CYP2C19 Nucleic Acid Test (2C19), which has been confirmed through internal testing. | Under Investigation by firm | Nanosphere sent an Field Notification Letter dated March 18, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letters included instructions for customers to: 1) discontinue use of the affected products, 2) discard the affected product, and 3) complete and return the attached for to obtain free-of-charge replacement kits. Customers with any questions about this recall should contact their Nanosphere technical support representative at 1 888 837-4436.. |  | 78 | 78 kits | Worldwide Distribution - US including the states of AL, CA, MD, ND, NJ, OK and TX., and internationally to Bulgaria, China and New Zealand. | Terminated | July 02, 2015 | 471 | 78 | Not_Computer
Z-2536-2014 | 68739 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | 4001PS Neonatal/Paed.Intubation Kit 2.5mm Murphy, LDS 15mm Conn |  |  | July 07, 2014 | 2014 | Smiths Medical ASD, Inc. | The tube fastener on the ET Tube Holder may compress the tracheal tube when used which may cause difficulty when locking the ET Tube Holder in place or difficulty when passing asuction catheter. | Process change control | Smiths Medical sent an Urgent Medical Device Recall Notice dated July 11, 2014, to all affected customers. The letter informed customers of the recall and instructed them to inspect their inventory and segregate any unused affected products and to complete and return the Recall Confirmation Form, by Fax 1-800-621-2590 or email to endotrachealtubeholders@smiths-medical.com.Customers with questions were instructed to call Smiths Medical's Customer Service Department at 1-800-258-5361.For questions regarading this recall call 866-216-8806. |  | 135518 | 135,518 (110,246 nationwide) | Worldwide Distribution - USA including AK, AR, AZ, CA, CO, CT, DC, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV and Internationally to CA, CY, GB, IE, IT, JE, KW, MT, PL, RU, SA, ZA. | Open, Classified |  |  | 135518 | Not_Computer
Z-2558-2014 | 68951 | No | KQT | Gastroenterology-urology evacuator. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | CollectEVAC, Sterile device with tubing and adaptor for use with ACMI and STORZ/Wolf resectoscopes. Part number 610-954. Rx only. Sterile EO. |  |  | August 01, 2014 | 2014 | Civco Medical Instruments Inc | A pinhole leak identified in the packaging of the recalled product potentially could compromise the sterility of the product. | Packaging | Civco Medical Instrucments contacted customers by phone and sent an Urgent Medical Device Recall letter dated August 1, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were requested to return all the affected product. Customers were instructed to contact Customer Service at 319-248-6520 with any concerns or questions. |  | 52 | 52 kits | Worldwide Distribution - USA including AZ, IL, MA, MD, MI, MN, MO, MS, LA, NC, NJ, NY, SD, TN, VA, WI, and Internationally to Brazil. | Terminated | January 28, 2015 | 180 | 52 | Not_Computer
Z-2557-2014 | 68834 | No | LOX | Percutaneous Transluminal Coronary Angioplasty (PTCA) Catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis EMPIRA RX PTCA Dilatation Catheter |  |  | June 23, 2014 | 2014 | Cordis Corporation | The balloon may exhibit resistance to the removal of the balloon cover. Weakening of the affected product in the region of the proximal balloon seal may occur when excessive force is applied to remove the balloon cover. This may lead to balloon burst, shaft separation or deflation difficulties, which could result in procedural delay, vessel trauma or surgery. | Environmental control | Cordis Corp sent Urgent Medical Device Correction Letters dated June 23, 2014, to all affected customers. Those in the Foreign Countries were sent via E-mail dated June 23, 2014. The letter identified the product the problem and the action needed to be taken by the customer.Cordis and Creganna-Tactx are not removing the products, but rather providing additional labeling (customer letter) advising the user not to use the product if unusual resistance is felt during removal of the balloon cover. Cordis and Creganna-Tactx explained in the letter to the consignees that is is known that balloons with hydrophilic coating are susceptible to expansion of the coating at extremely elevated humidity, therefore, it is important for the User to follow the current labeling, to store the product in a cool, dark and dry place.Action requested on your part:"Pass on this letter to anyone who needs to be informed in your facility and in any facility where potentially affected devices may have been transferred."Maintain awareness of this communication until the information has been incorporated into the Cordis EMPIRA" NC RX and EMPIRA" RX PTCA dilatation catheter labeling."Retain this letter with affected product."Report all balloon cover removal difficulties through the standard complaint process."For questions related to the issue please contact Cordis Endovascular Medical information Line 877-DEVICE5 (877-338-4235), Monday through Friday from 9:00AM to 5:00 PM EDT."For questions related to the Acknowledgment Form and its return, please contact Cordis QA at (786) 313-8730. |  | 250000 | 250,000 units. | Worldwide Distribution - USA including Puerto Rico and the states of AR, AZ, CA, CO, CT, DC, FL, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WV, WY and Hawaii., and the countries of Argentina, Brazil, Canada, Chile, Colombia, Mexico, Panama and Venezuela. | Open, Classified |  |  | 250000 | Not_Computer
Z-2302-2014 | 68536 | No | OEQ | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | Enforcement Discretion | Customed Angiodrape Pack |  |  | May 20, 2014 | 2014 | Customed, Inc | Customed has determined that there is the possibility that packaging integrity may be compromised on different catalog numbers (multiple lots). This could result in an injury to the patient due to product contamination or loss of sterility condition. | Package design/selection | Customed Inc. sent a Recall Notification letter dated June 3, 2014 via email and followed by letter to all affected customers. The letter identified the affected products, problem and actions to be taken. The letter informed customers if the affected products are in their inventory to: 1. Remove products from sale (quarantine them) and cease distribution. 2. Complete and return the enclosed acknowledgement form to Customed Inc .Recall Coordinator. Affected products should be returned to Customed Inc for credit.Customers with questions were instructed to contact Customed Inc. Recall Coordinator at 1-787-801-0100 Ext. 7540 |  | 39936 | 26 lots, 1536 units (multiple units per lot) | Worldwide Distribution - US (nationwide) in the state of Florida including Puerto Rico (USA); &amp; USVI (St John) and the countries of Tortola, Trinidad, Dominican Republic, Costa Rica and Mexico. | Terminated | August 14, 2017 | 1182 | 15299586 | Not_Computer
Z-2560-2014 | 69050 | No | MBI | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | ACL Sleeve Installation Tool Sizes: 6mm, 6.5mm, 7.0mm, 7.5mm, 8.0mm, 8.5mm, 9.0mm, 10.0mm |  |  | July 25, 2014 | 2014 | OrthoPediatrics Corp | Weld fractures on the Sleeve Installation Tool for ACL reconstruction system. | Nonconforming Material/Component | OrthoPediatrics sent a Notification letter dated July 25, 2014 to all affected customers. The letter Identified the affected product, problem and actions to be taken. The letter informed customers to take immediate possession of all affected product within their control and at any of their health care facilities and immediately return the recalled device(s) to OrthoPediatrics Corp. This notice and the attached Reply Form should be forwarded if the affected product was further distributed. . |  | 56 | 56 units | US Nationwide Distribution. | Terminated | May 19, 2015 | 298 | 56 | Not_Computer
Z-2562-2014 | 68650 | No | KNC | Obstetric table and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | AffinityÂż Four Birthing Bed, |  |  | June 26, 2014 | 2014 | Hill-Rom, Inc. | The attachment latch mechanism can become bent depending on user handling. Latch mechanism bending could lead to improper engagement of the foot section and this could result in injury to the user due to a fall. | Nonconforming Material/Component | Hill-rom issued an Urgent Field Safety Notice letter dated June 26, 2014 to affected customers. The Notification identified the affected product issue and instructions for product inspection. The letter states that the firm is working on making changes to the Lift-Off foot section that will mitigate risk even if damage was to occur. If you have any questions concerning the notice or requested procedure, please contact Hill-Rom Technical Support at 800-445-3720. |  | 13000 | 13,000 units | Worldwide Distribution - US Nationwide and the following countries: Argentina , Aruba, Australia, Austria, Bahamas, Bahrain, Belgium, Belize, Bermuda, Brazil, Brunei, Bulgaria, Cayman Islands, Chile, China, Colombia, Costa Rica, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Dominican Republic, Egypt, Equatorial Guinea, Finland, France, Germany, Greece, Guam, Guyana, Hong Kong, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kenya, Kuwait, Lebanon, Libya, Malaysia, Mauritius, Mexico, Morocco, Netherlands, New Zealand, Nigeria, Norway, Oman, Pakistan, Palestine, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Turks and Caicos Islands, United Arab Emirates, United Kingdom, US Virgin Islands, Venezuela and Vietnam. | Terminated | July 14, 2017 | 1114 | 13000 | Not_Computer
Z-2561-2014 | 68965 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens EasyLink Informatics System V4.0/V5 |  |  | August 13, 2013 | 2013 | Siemens Healthcare Diagnostics, Inc. | May not perform as intended under certain conditions, causing the release of results to the Laboratory Information System (LIS) that should have been held for manual review due to auto-verification rules or the delay/omission of resulttransmission to the LIS. | Software design | Siemens issued an Urgent Medical Device Correction (UMDC) letter on August 13, 2013 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to review this letter in its entirety and ensure that all operators understand the information presented. The issues described are specific to the features listed in the letter.. If the laboratory is not using a particular feature, then they are not affected by the issues described for that feature. For questions contact your Siemens Customer Care Center or your local Siemens technical supportrepresentative. |  | 1347 | 1347 | US Nationwide Distribution | Terminated | December 02, 2014 | 476 | 1347 | Not_Computer
Z-2563-2014 | 68912 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Roche cobas IT 1000 |  |  | July 31, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Potential that test results would be assigned to the wrong Patient ID if cobas IT 1000 receives an ADT ADMIT, TRANSFER or MERGE event message that contains a genuine Patient ID (e.g. 999999), which is the same as what has been configured as the Generic Patient ID (999999 by default). | Software design | Roche sent an Urgent Medical Device Correction letter dated July 31, 2014 via UPS Ground (receipt required) to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to follow the steps provided in the letter to prevent assigning patient results to an incorrect generic patient ID. Please contact the Roche Support Network Customer Support Center at 1-800-440-3638, prompt 5, 24 hours a day, seven days a week if you have questions about the information contained in this UMDC. |  | 148 | 148 Units | US Nationwide Distribution in the states of: AK, AL, AR, AZ, CA, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MD, ME, MI, MO, NC, NE, NH, NJ, NY, OH, OR, PA, PR, SC, TN, TX, VA, VT, WA, WI, and WV. | Terminated | June 16, 2015 | 320 | 148 | Not_Computer
Z-2149-2014 | 68665 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | SIEMENS ADVIA Centaur CP Immunoassay System BNP kit; SMN 10309044 - 100 test |  |  | June 23, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Complaints of failed calibration. | Other | Siemens sent an Urgent Medical Device Correction (UMDC) letter to all affected Siemens Healthcare Diagnostics customers in the United States on June 23, 2014 via Federal Express. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on June 23, 2014 via hard copy, email, and/or fax per each country's local regulations. These notices inform the customer of the potential for low calibrator deviation failures when using ADVIA CentaurÂż Systems BNP kit lots ending in 172 on the ADVIA CentaurÂż CP due to an error on the master curve card. Siemens informs customers if a successful calibration is achieved, the results obtained are valid. If customers are unable to obtain a successful calibration on the ADVIA CentaurÂż CP, they should switch to an alternative lot of BNP assay. The notices should be reviewed with the facility's laboratory or medical director.Fax-back forms will be used to confirm that the customers received the UMDC. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Medical Device Correction and completed the required actions. Customers with questions were instructed to contact their Siemens Customer Care Center or their local Siemens technical support representative.For questions regarding this recall call 877-229-3711. |  | 1191 | 1,191 kits | Worldwide Distribution - USA (nationwide including Puerto Rico) and Internationally to Canada, Albania, Argentina, Australia, Bosnia Herzeg., Brazil, Chile, Colombia, Egypt, Georgia, Hong Kong, India, Indonesia, Iran, Iraq, Israel, Japan, Kazakhstan, Lebanon, Malaysia, New Caledonia, Oman, P.R. China, Pakistan, Philippines, Qatar, Rep. of Yemen, Republic Korea, Russian Fed., Saudi Arabia, Serbia, Singapore, South Africa, Taiwan, Thailand, U.A.E., Ukraine, Venezuela, Vietnam, Austria, Bulgaria, Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Greece,Hungary,Iceland, Ireland, Italy, Latvia, Lithuania, Netherlands, Norway, Poland, Portugal, Romania,Slovakia, Slovenia, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 31, 2016 | 647 | 2152 | Not_Computer
Z-2564-2014 | 68758 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Hitachi Echelon Oval MRI system is a diagnostic imaging device (one unit per package) and is intended to provide the physician with physiological and clinical information, obtained non-invasively and without the use of ionizing radiation. |  |  | August 18, 2014 | 2014 | Hitachi Medical Systems America Inc | The Gradient Coil was found to have a failure mode that allowed it to overheat and become a burn hazard. | Device Design | Letters will be sent to customers. Hatachi Serivce will be sent to each site to exchange the Gradient Coil Assembly. |  | 39 | 39 systems | Worldwide Distribution: US (nationwide) including states of: CA, DE, FL, ID, KS, KY, MD, MS, NY, OH, and WY; countries of: Japan, Brazil, France, Germany, and Kyrgyzstan. | Terminated | September 14, 2015 | 392 | 39 | Not_Computer
Z-2566-2014 | 67547 | No | LZD | Total temporomandibular joint prosthesis. | Dental | Dental | PMA | Biomet Microfixation Temporomandibular Joint (TMJ) Patient Matched Left Fossa Part# CP751506 |  |  | September 20, 2013 | 2013 | Biomet Microfixation, LLC | Product mix occurred prior to final package and labeling process. | Mixed-up of materials/components | Biomet Microfixation, Inc. sent an Urgent Medical Device Field Notice dated September 20, 2013 to its customer. The customer was instructed to: 1) Immedicately quarantine the device. 2) Discontinue use and return to Biomet Microfixation. 3) If the product was further distributed, the letter must be forwarded to the firm's customers. 4) A replacement device will be shipped priority upon completion. Questions related to this notice should be directed to 904-741-9448 or 904-239-6671.For questions regarding this recall call 904-741-4400. |  | 1 | 1 | Worldwide Distribution to Denmark only. | Terminated | October 03, 2014 | 378 | 1 | Not_Computer
Z-2565-2014 | 68970 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | Medsolutions4U Multi-Function Electrodes to be used with PhysioControl Lifepak Defibrillators, models 9, 10C, 11, 12, 15, 20, 500 and are indicated for Defibrillation, Cardioversion, Pacing, and Monitoring |  |  | March 12, 2013 | 2013 | Bio-Detek, Inc. | Medsoultions4U Multi-Function Lifepak Electrodes, manufactured with an incorrect wire connector which could not interface with the defibrillators | Nonconforming Material/Component | Bio-Detek issued an Urgent Field Safety Notice to Medsolutions4 U on 3/12/13, informing them of the problem and actions that should be taken to remove the affected products from service. |  | 30 | 30 cases (300 electordes) | Worldwide Distribution in Germany and Slovenia. | Terminated | June 17, 2015 | 827 | 30 | Not_Computer
Z-2573-2014 | 69033 | No | KWP | Spinal interlaminal fixation orthosis. | Orthopedic | Orthopedic | 510(k) | STREAMLINE CT CERVICO-THORACIC SYSTEM; SET SCREW |  |  | September 10, 2012 | 2012 | PIONEER SURGICAL TECHNOLOGY, INC. | Some set screw thread profiles were found not to be within specifications. | Nonconforming Material/Component | The recall took place at the distributor/sales representative level. Communication was done through phone and email notifications on 9/10/2012.All affected set screws were to be traded out through a scheduled process. This was coordinated through Pioneer Surgical's customer service. Product was scrapped and no reconditioning took place. All affected set screws were swapped out. |  | 11 | 11 sets | Distributed in the states of NC, NH, NY, CA, MI, IL, and FL. | Terminated | September 30, 2014 | 750 | 11 | Not_Computer
Z-2574-2014 | 68984 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | FOCAL Sim Workstation, Radiation Treatment Planning System. |  |  | August 18, 2014 | 2014 | Elekta, Inc. | When DICOM is exporting an arc plan (VMAT, Dynamic Conformal of 3D Static Arcs) and using the Composite Field Sequencing (CFS) feature, the DICOM exported plan will not match the approved treatment plan. | Software design | Elekta, Inc. issued an Urgent Important Field Safety Notice letter (IFSN 382-01-MON-001) to all affected customers.. Confirmation of receipt is to be completed by October 31, 2014. The problem will be resolved in a patch to Focal 4.80. Development is underway with a release date of September 2014 with confirmation of installation of the patch for the fix no later than 6 months from release of the patch. An updated notice from the manufacturer is also being sent to affected customers (FCO: 382-01-MON-001 VID:2.0). For questions contact your local Elekta office. |  | 554 | 554 units | Worldwide Distribution - US (nationwide)in the states of AK, CA, FL, MI, MN, MO, NJ, NY, NC, OH, PA, TN, VA,WA and Puerto Rico, Algeria, Antigua &amp; Barbuda, Australia, Austria, Bahrain, Belarus, Brazil, Bulgaria, Canada, Chile, China, Colombia, Cyprus, Czech republic, Ecuador, Egypt, Estonia, France, Germany, Greece, India, Iraq, Israel, Italy, Japan, Kazakhstan, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Pakistan Portugal, Romania, Russia, Saudi Arabia, Singapore, Solvenia, South Africa, South Korea, Spain, Switzerland, Thailand, Turkey, and United Kingdom | Terminated | August 08, 2017 | 1086 | 554 | Not_Computer
Z-2568-2014 | 69072 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | LINEAR 7.5 FR. 40cc IAB Catheter |  |  | August 11, 2014 | 2014 | Datascope Corporation | During label reconciliation a duplicate serial number label was detected to be printed for the LINEAR 7.5 FR 40 cc Intra-Aortic Balloon Catheter. | Labeling False and Misleading | Maquet Inc.is sued an Urgent Medical Device Recall Letter/ Fax Back Form dated August 11, 2014. Customers are asked to quarantine the affected product and upon return they will be provided with an exchanged unaffected product. The attached form should be completed an returned to Maquet Inc. Questions can be directed to Maquet Technical Support at (888) 627-8383 - Press option "3" followed by option "1" Monday through Friday between the hours of 8 am and 5 pm (EST). |  | 1 | one unit | US Distribution to TX. | Terminated | February 05, 2015 | 178 | 1 | Not_Computer
Z-2571-2014 | 69138 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | V5Ms Transesophageal transducers; |  |  | August 13, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Reports of deterioration of material covering the articulating section of the V5Ms transesophageal transducer. | Under Investigation by firm | Siemens sent a Customer Safety Advisory notification to all affected customers on August 13, 2014. The notification identified the product, the problem, and the action to be taken by the customer. Customers were advised that if they notice any damage or wear on their transducer, or if the transducer fails to pass the leakage current test, immediately discontinue use of the transducer and contact their local Customer Service Engineer. If they have anyquestions, customers should contact their local service support person for information regarding timelines and status.Customers were instructed to share this information with all personnel within their organization who need to be aware of this issue.For questions regarding this recall call 650-417-8185. |  | 321 | 321 units | Worldwide Distribution - USA (nationwide) and Internationally tAustralia, Austria, Belgium, Bolivia, Brazil, Canada, China, Columbia, France, Germany, Great Britain, India, Ireland, Italy, Japan, Mexico, Netherlands, Norway, Peru, Poland, Portugal, Russian Federation, Saudi Arabia, Slovak Republic, South Africa, South Korea, Spain, Sweden, Taiwan, Thailand,and United Arab Emirates. | Terminated | March 06, 2015 | 205 | 321 | Not_Computer
Z-2572-2014 | 69087 | Yes | LWS |  |  | Cardiovascular | PMA | Ellipse Implantable Cardioverter Defibrillators (ICDs), Model No. CD1275, CD1311, CD2275, CD2311, CD1411, CD2411 |  |  | August 19, 2014 | 2014 | St Jude Medical Cardiac Rhythm Management Division | An anomaly may result in delayed, partial, or no delivery of high voltage therapy. Internal damage to the high voltage capacitors, which may result in extended charge time, occurring during capacitor maintenance or charging for high voltage therapy. It may present as a vibratory patient notification and upon interrogation, a programmer or Merlin.net alert. | Component design/selection | A customer letter dated 8/19/14 was sent to physicians with the affected Ellipse ICDs on 8/20/14. The letter identifies the problems and the actions to be taken.Customers with any questions or concerns are instructed to contact their local St. Jude Medical representative or St. Jude Medicals Technical Services Department. |  | 51492 | 51,492 units (27,396 units in US) | Worldwide Distribution: US (nationwide), Canada, Australia, New Zealand, Japan, EU, Middle East, Latin America, and Asia-Pacific. | Terminated | October 08, 2015 | 415 | 51492 | Not_Computer
Z-2592-2014 | 68993 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Split Cath III Trays, 16F x 28CM |  |  | August 04, 2014 | 2014 | Medical Components, Inc dba MedComp | The Split Cath III Trays were incorrectly labeled as 16F x 32CM Split Cath III instead of 28CM. | Labeling False and Misleading | Medcomp sent a Recall Notification letter, dated August 4, 2014, to all affected customers. The letter identified the affected product, problem, and action to be taken. Customers were requested to respond and return all affected product. For questions call contact Susan Smith at ssmith@medcompnet.com. |  | 110 | 110 | USA Nationwide Distribution | Terminated | December 10, 2014 | 128 | 110 | Not_Computer
Z-2591-2014 | 69022 | No | MRD |  |  | General &amp; Plastic Surgery | 510(k) | EN MENTOR Round High Profile Single Use Saline Breast Implant Sizer Single Use Saline Breast Implant Sizer |  |  | July 02, 2014 | 2014 | Mentor Texas, LP. | The device labeling is stamped with 330cc and the device is actually 380cc device. | Error in labeling | Mentor sent an Important Customer Notification letter dated July 17, 2014 to all affected customers and follow-up with a visit by the local sales representative. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately identify all unsused product from lot list provided and maintain a copy of this notice with the affected product. For question contact your local Sales Representative or Mentor Customer Service at 1-800-235-5731. |  | 34 | 34 units | U.S Nationwide Distribution in the states of Michigan, Maryland, Texas, New Jersey, Indiana, Nevada, Florida, California, New York, Tennessee, Pennsylvania, and Colorado. | Terminated | April 06, 2015 | 278 | 34 | Not_Computer
Z-2581-2014 | 68985 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco |  |  | August 18, 2014 | 2014 | Elekta, Inc. | Incorrect Delivery of Composite VMAT Fields - When DICOM is exporting an arc plan (VMAT, Dynamic Conformal or 3D Static Arcs) and using the Composite Field Sequencing (CFS) feature, the DICOM exported plan will not match the approved treatment plan. | Software design | The firm, Elekta, issued an "URGENT IMPORTANT FIELD SAFETY NOTICE (IFSN 382-01-MON-001)" to all affected customers on August 18, 2014. The notice describes the product, problems and actions taken. Confirmation of receipt is to be completed by October 31, 2014. The problem will be resolved in patches to the following releases - Monaco 3.30 and Monaco 5.00. Development is underway with a release date of September 2014 with confirmation of installation of the patch no later than 6 months from release of the patch. An updated notice from the manufacturer is also being sent to affected customers (FCO: 382-01-MON-001 VID:2.0). If you have any queries about this Notice, please contact your local Elekta office or call 770-300-9725 or email: Elekta_Global_PMS@elekta.com. |  | 928 | 928 units | Worldwide distribution: US (nationwide) including states of: AL, AK, AZ, CA, CO, CT, FL, IL, IN, MD, MA, MI, MN, MO, MT, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SD, TN, TX, VA, WA, WI, Puerto Rico; and countries of: Algeria, Angola, Antigua &amp; Barbuda, Australia, Austria, Bahrain, Belarus, Belgium, Brazil, Canada, Chile, China, Colombia, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, India, Iraq, Ireland, Israel, Italy, Japan, Kazakhstan, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Namibia, Netherlands, New Zealand, Norway, Pakistan, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sir Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, and Ukraine. | Open, Classified |  |  | 928 | Not_Computer
Z-2588-2014 | 68939 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | GALT VALVED TEARAWAY INTRODUCER SET GALT VTI RX ONLY UNITS 7 STERILE EO |  |  | July 25, 2014 | 2014 | Galt Medical Corporation | The adhesive used in the product could dislodge when objects are inserted through the valve which could pose an embolic risk to the patient. | Nonconforming Material/Component | Galt Medical sent an Urgent Medical Device Recall letters to all affected customers. The letter identified the affected product, reason for recall and actions to be taken. Customers were asked to identify and segregate the recalled lot (s) that is in their possession, complete and return the enclosed Recall Reply form, and return the recalled product to Galt Medical Corp. For questions contact your local sales representative or Recall Coordinator at 214-778-1306. |  | 11155 | 11,155 units in total | Worldwide Distribution - US (nationwide) Canada, Germany, the Netherlands, Ireland, Switzerland, France, Italy, and Greece. | Terminated | February 26, 2016 | 581 | 11155 | Not_Computer
Z-2576-2014 | 68878 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | PERSONA The Personalized Knee System Constrained Posterior Stabilized Tibial Articular Surface Provisional TOP with the following: L 3-5 CD; L 6-9 CD; L 3-5 EF; L 6-9 EF; L 10-11 EF; L 6-9 GH; L 10-12 GH; L 10-12 J; R 3-5 CD; R 6-9 CD; R 3-5 EF; R 6-9 EF; R 10-11 EF; R 6-9 GH; R 10-12 GH; and R 10-12 J. |  |  | August 07, 2014 | 2014 | Zimmer, Inc. | Reports of breakage of the Persona Knee Trial Articular Surface Provisionals (TASPs). The Persona primary knee surgical technique has been updated. This field action correction is to ensure that all distributors and users of this instrument are aware of the enhanced surgical techniques. | Device Design | On 8/7/2014, Zimmer notified all distributors via electronic mail. Distributors that have received affected inventory are also notified via courier. Hospital risk managers and surgeons are also notified via courier.Hospital risk managers and surgeons were provided with a notification identifying the issue and their responsibilities. These responsibilities include ensuring the affected personnel are aware of the contents of the notification. Distributors were sent a notification identifying the issue and to provide a copy of the notification sent to surgeons and risk managers. The distributors will also be instructed as needed to assist with identifying a comprehensive list of surgeons. |  | 77737 | 77,737 units total | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY; and countries of: AUSTRALIA, INDIA, JAPAN, KOREA, SINGAPORE, AUSTRIA, BELGIUM, FRANCE, GERMANY, ISRAEL, ITALY, LUXEMBOURG, NETHERLANDS, SAUDI ARABIA, SOUTH AFRICA, SPAIN, SWITZERLAND, UNITED KINGDOM, and UNITED ARAB EMIRATES. | Terminated | January 27, 2015 | 173 | 77737 | Not_Computer
Z-2590-2014 | 69021 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | iTRACK250A Canaloplasty Microcatheter Kit; Catalogue number: iT-250A; Rx only, sterile. |  |  | August 06, 2014 | 2014 | Ellex iScience, Inc. | One lot may not have been properly sealed, resulting in a non-sterile device. | Under Investigation by firm | Urgent Medical Device Recall letters were sent to all customers on August 6, 2014 by trackable mail. Letters advised on the potential lack of sterility of iTRACK250A, Canaloplasty Microcatheter Kit, lot 1404-03 due to improperly sealed pouches. Lack of sterility may pose the risk of infection to the patient. Customers are identify and segregate any affected product in their inventory. The affected product should be returned following the instructions provided. Replacement product will be provided at no charge. Questions should be directed to Customer Care at 1-888-846-4724, M-F 8 am to 4:30 Pm Eastern Time. |  | 148 | 148 devices | Worldwide Distribution -- US (nationwide) and countries of: Hong Kong and Switzerland. | Terminated | October 28, 2014 | 83 | 148 | Not_Computer
Z-2586-2014 | 69125 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | ARIATELE TELEMETRY TRANSMITTER, Model 96281, with SpO2 Option C. Physiological patient monitor. Used in a hospital environment. |  |  | July 02, 2014 | 2014 | Spacelabs Healthcare Inc | The transmitter display and the Telemetry Central Station display may show a SpO2 value when the sensor is not connected to the patient or in certain sensor failed conditions.This is an expansion of recall Z-0871-2014 because additional serial numbers were identified as being part of this recall action as part of the firm's retrospective analysis of all recalls since 2011. | Software design | Spacelabs notified both consignees via email of the expanded recall July 2, 2014. A f/u letter was sent by overnight delivery to the single new consignee the same day. Spacelabs sent an announcement of the expanded recall to both consignees identified in the expanded recall August 25, 2014. A field service technician from Spacelabs will visit each consignee and upgrade the software as necessary.If you have any questions about this corrective action program, please contact SpaceLabs Healthcare at 1-800- 522-7025, and select 2 for Technical Support. |  | 8 | 8 units | US Distribution to TX. | Terminated | October 16, 2014 | 106 | 8 | Computer
Z-2583-2014 | 68655 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Natural-Knee System Patella Bushings. Orthopedic surgical instrument. Part Number: 6290-00-702. |  |  | July 10, 2014 | 2014 | Zimmer, Inc. | Zimmer received a trend of complaints indicating corrosion of product. | Device Design | Zimmer issued an Urgent Medical Device Recall-Lot Specific notification via e-mail/letter dated July 19, 2014, to all affected customers. The notification identified the product the problem and the action needed to be taken by the customer. All distributors were notified via electronic mail. Hospital risk managers and surgeons, as well as distributors with affected inventory were notified via courier. Distributors were sent a letter identifying the issue and their responsibilities. These responsibilities include locating and removing the affected product in their territory as well as identifying hospitals and surgeons who have previously used the affected product. Distributors are to return on-hand affected products to Zimmer, and ensure all of their products are accounted for using the form provided in the letter. Hospital risk managers and surgeons were provided with a letter identifying the issue and their responsibilities. These responsibilities include locating the affected product, quarantining the product, and returning it to their Zimmer sales representative. If you have questions or concerns regarding this recall please call the customer call center at 1-877-946-2761. |  | 8638 | 8638 distribution events | Worldwide Distribution - US Distribution and the countries of : Argentina, Australia, Brazil, France, Germany, India, Japan, Spain, Switzerland, and Venezuela. | Terminated | August 20, 2015 | 406 | 8638 | Not_Computer
Z-2570-2014 | 68754 | No | LKF | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical H/S CATHETER (PART # 61-5007) |  |  | July 03, 2014 | 2014 | CooperSurgical, Inc. | Unsealed pouches. | Packaging process control | Customer notification letter was sent on July 1, 2014 via UPS with confirmed delivery receipt. The letter identified the product, the problem, and the action needed to be taken by the customer. At CooperSurgical's expense, arrangements were made to replace any of the affected products that customers may have on hand. For questions contact James Keller, Vice President, Regulatory Affairs/Quality Assurance at 203-601-5200. |  | 1400 | 1,400 units | Worldwide Distribution -- USA, in the states of CA, TX, IL, MI, MS, ID, NY, FL, MD, PA, OR, GA, CO, WA, NJ, HI, and NC, including DC; and the countries of the countries of Spain, England, Holland, and United Kingdom. | Terminated | May 28, 2015 | 329 | 1400 | Not_Computer
Z-2595-2014 | 69066 | No | LKB |  |  | General Hospital | Enforcement Discretion | DualCap IV Pole Strips Disinfectant Caps for Luer Access Valves, REF 500-LB, Sterile. |  |  | August 05, 2014 | 2014 | Catheter Connections, Inc. | Catheter Connections is recalling DualCap IV Pole Strips disinfectant caps for luer access valves due to potential improper packaging cap seal which may render the sealed cap non-sterile. | Packaging | The firm, Catheter Connections, sent an "Urgent: Medical Device Field Action" letter dated August 5, 2014 to its consignees/customers. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to identify any customers who received this product having lot no. 1406509; make all potential users in your facility are made aware of this notification and the recommended actions; forward the attached URGENT MEDICAL DEVICE CORRECTION LETTER to customers immediately via fax or email; search your inventory and instruct all users to check each cap coming from the specific affected lot; quarantine and discard any defective product; and complete and return the attached URGENT MEDICAL DEVICE FIELD ACTION REPLY FORM via fax to 1-888-862-2693 or email to: info@cathconn.com.Catheter Connections does not require the return of devices having the affected lot number.If you have questions, please contact Catheter Connections at 1-801-906-0820 (M-F, 8:00 am - 5:00 pm MST) and ask for Quality Department. |  | 24910 | 24,910 units | US Distribution to states of: CT, GA, MD, and PA.. | Terminated | October 08, 2014 | 64 | 24910 | Not_Computer
Z-2594-2014 | 68915 | No | NHA | Endosseous dental implant abutment. | Dental | Dental | 510(k) | BIOMET 3i Certain Bellatek Encode Healing Abutment |  |  | June 25, 2014 | 2014 | Biomet 3i, LLC | Identifiers on the label are correct but the label has the incorrect color coding. | Mixed-up of materials/components | Biomet 3i sent an Urgent - Medical Device Removal letter / and by phone call on and dated June 25, 2014, to all affected customers. The letter identified the product the problem and the action to be taken by the customer.If you have any Certain BellaTek Encode Healing Abutment(s) of the affected lot in your practice, please return the abutment(s) to BIOMET 3i in exchange for Cettain BellaTek Encode Healing Abutments, IEHA458 with a proper color stripe. Alternatively, you may return the Certain BellaTek Healing Abutment(s) for full credit.We also ask that you confirm receipt of this notification by completing the attached Recall Return Response Form, indicating your preference for replacement or credit and/ or notifying us if the product was used.Thank you for your support - we are committed to maintaining high standards of quality for all our products and apologize for any inconvenience this may have caused you.If you have any further questions, please contact Customer Service at 1-800-342-5454 for assistance. |  | 6 | 6 units | US Distribution to NH and TX. | Terminated | September 23, 2014 | 90 | 6 | Not_Computer
Z-2597-2014 | 68982 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | RADIANCE v. 2.4X, 2.5X and 2.6X |  |  | July 28, 2014 | 2014 | Radiometer America Inc |  | Software design | Radiometer sent an Product Notification letter dated July 31, 2014 to affected customers via FedEx. The letter identified the affected product, problem and actions to be taken. The letter informed customers that local RADIOMETER IT specialist will contact customers in order to schedule the upgrade to v. 2.74 or 3.5.2. All RADIANCE versions below 2.7X will be upgraded to v. 2.74 or 3.5.2.For questions regarding this letter or the Recall, please contact Radiometer America Technical Support at 1-800-736-0600 opt.4. |  |  |  | Worldwide Distribution - US (nationwide) in the states of AZ, CA, MO, MS, TX and UT. and the countries of: Austrialia, Austria, France, UK, Italy, Japan, Spain, India, and the Netherlands. | Terminated | October 15, 2014 | 79 |  | Not_Computer
Z-2598-2014 | 69119 | No | NRY | Percutaneous catheter. | Cardiovascular | Neurology | 510(k) | Merci Retriever |  |  | June 25, 2014 | 2014 | Concentric Medical Inc | Merci Retrievers were shipped to US customers past their expiration date, even though they were labeled with the correct expiration date. | Process control | Concentric sent an Urgent Medical Device Recall Notification dated June 25, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Letters were sent to Risk Manager- Director of Material and requested that the subject devices be quarantined and that Concentric be notified if the product has been distributed to other organizations. If so Concentric should be notified so that the additional customers can be notified.The response form should be completed and returned to Neuro Complaints @Stryuker.com. upon receipt of the response form, a representative will contact the firm to arrange for return and/or credit for the devices. Fax it to 1-866-876-4355. For further questions please call ( 650) 810-1707 Ext. 1716. |  | 9 | 9 devices | US Distribution including the states of PA, CA, TX, NM, and OH. | Terminated | October 29, 2014 | 126 | 9 | Not_Computer
Z-2599-2014 | 69058 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | CLAW PLATE 20mm INTERAXIS CHARLOTTE F&A; SYSTEM, REF 40S20020, 1 EACH, Rx ONLY, NON STERILE |  |  | July 30, 2014 | 2014 | Wright Medical Technology, Inc. | Plates manufactured from an incorrect raw material. | Device Design | Wright Medical sent an Urgent Medical Device Recall letter dated July 25, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for the affected product and immediately quarantine all affected product and return to Wright Medical as soon as possible. For return instructions customers were instructed to contact Customer Service at 800-238-7117. Customers were also asked to complete the attached response form, confirming receipt of the notice and the status of the product(s) on hand and return by fax to 901-867-7401.Customers with questions were instructed to call 901-290-5662.For questions regarding this recall call 800-238-7117. |  |  |  | Worldwide Distribution - USA (Nationwide) and Internationally to Canada Australia, Great Britain, Germany, Spain, Netherlands, France. | Open, Classified |  |  |  | Not_Computer
Z-2609-2014 | 68821 | No | DOP | Mass spectrometer for clinical use. | Clinical Chemistry | Toxicology | 510(K) Exempt | AB SCIEX API 3200MD" LC/MS/MS System with software: |  |  | July 15, 2014 | 2014 | Ab Sciex | Software- AB SCIEX API 3200MD LC/MS/MS SYSTEM with MultiQuantMD 3.0 software may provide incorrect data for patient diagnosis | Software design | AB SIEX issued letter on 7/15/14, notifying customers to implement this field correction, a temporary fix to be implemented immediately.Additionally, a software update is in development and is expected to be available in approximately 8 weeks. Upon availability, AB SCIEX will send a DVD with instructions on how to install the new software update. At that time, outlined temporary actions will no longer be necessary.If you have any questions regarding this recall, potential hazards or actions to be taken please feel free to contact AB SCI EX at +1 289 982 2531.Please confirm receipt of this letter by signing and faxing back the attached Acknowledgement Form. |  | 6 | 6 units | FL, NC, OK, PA, TX | Terminated | October 01, 2014 | 78 | 6 | Computer
Z-2610-2014 | 69109 | No | GAG | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | PROXIMATE Skin Stapler. PROXIMATE RH Rotating-Head Skin Stapler (Model PRR35), PROXIMATE PX Fixed-Head Skin Stapler (Model PXR35), Bulk/Non-Sterile PROXIMATE RH Rotating-Head Skin Stapler (Model CK 117), and Bulk/Non-Sterile PROXIMATE PX Fixed-Head Skin Stapler (Model CK 158). Skin stapler for use in a single patient, which is designed to apply rectangular stainless steel staples for routine wound closure. |  |  | August 12, 2014 | 2014 | Ethicon Endo-Surgery Inc | The device may have a non-conforming component that may cause the device to fire an unformed staple. | Nonconforming Material/Component | The firm, Ethicon Endo-Surgery, sent an "URGENT: MEDICAL DEVICE RECALL" letter dated August 25, 2014 to its customers via overnight, UPS to allow tracking of the receipt. The letter described the product, problem and actions to be taken. The customers were instructed to immediately examine your inventory to determine if you have recalled product on hand; remove the recalled product and communicate the issue to relevent staff; complete and return Business Reply Form (BRF) confirming receipt of this notice within three (3) business days, indicating whether they have affected product, how much and whether they will/have returned the product or whether they no longer have any of the recall product. The BRF may be sent to Stericycle by fax at 1-855-328-6659 or by email at ethicon3022@stericycle.com. Product replacement is available for customers who return unopened affected product.If you need clinical or product support, please contact your local sales representative or call our Customer Support Center at 1-877-ETHICON (1-877-384-4266). If you need additional shipping labels or a communications package, contact Stericycle at 1-855-328-0398 and reference Event 3022. |  | 42798 | 42,798 | Worldwide Distribution: US (nationwide) including states of: AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY; and countries of: Armenia, Belgium, Germany, Denmark, Finland, France, UK, Italy, Japan, South Korea, Netherlands, Norway, Russia, Sweden, Singapore, Taiwan, Tanzania, and South Africa. | Terminated | July 10, 2015 | 332 | 42798 | Not_Computer
Z-2617-2014 | 68877 | No | CCL | Oxygen gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | CARESCAPE Respiratory Modules E-sCO, E-sCOV, E-sCAiO, E-sCAiOV, E-sCAiOE, E-sCAiOVE; the Airway Gas Option N-CAiO and respective service exchange units. |  |  | August 08, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue due to a failure of an O2 sensor component associated with the CARESCAPE respiratory modules, the Airway Gas Option and their respective service exchange units. | Nonconforming Material/Component | Consignees were sent on 8/8/2014 a GE Healthcare "Urgent Medical Device Correction" letter Ref# 36106 dated August 8, 2014. The letter was addressed to Healthcare Administrator / Risk Manger, Chief of Nursing and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions you can contact Technical Support at 1-800-558-7044 or your local Service Representative. |  | 748 | 748 (307 units US, 441 units OUS) | US: CA, CO, FL, GA, IN, KY, LA, MD,MI, NY,NC,OH, PA, SC, TN, TX, VA,WI.OUS: UNITED KINGDOM, TURKEY, THAILAND, SWITZERLAND, SWEDEN, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, PHILIPPINES, PANAMA, PAKISTAN, NORWAY, NICARAGUA, NEW ZEALAND, NETHERLANDS, MEXICO, LEBANON, KUWAIT, REPUBLIC OF KOREA, JAPAN, ITALY, INDONESIA, INDIA, GERMANY, FRANCE, FINLAND, EGYPT, CZECH REPUBLIC, COLOMBIA, CHILE, BULGARIA, BELGIUM, AZERBAIJAN, AUSTRALIA. | Terminated | September 08, 2015 | 396 | 748 | Not_Computer
Z-2596-2014 | 68805 | cover, mattress (medical purposes) | FMW | Mattress cover for medical purposes. | General Hospital | General Hospital | 510(K) Exempt | Prime Care Âż Transcend Mattress, Prime Care Âż Transcend Perimeter Mattress, Prime Care Âż Transcend Advanced Mattress, Prime Care Âż Transcend Advanced Perimeter Mattress, Prime Plus Âż Bariatric Mattress and Prime Plus Âż Bariatric Perimeter Mattress, individual units packaged seperately. |  |  | August 11, 2014 | 2014 | Primus Medical LLC | The firm found that the top cover of the mattresses were delaminating. | Component design/selection | The firm sent out Urgent: Voluntary Field Correction/Removal letters, dated August 11, 2014, to all consignees with instructions for correcting the issue. Customers were to immediately examine their inventory and quarantine any affected product. Customers are to complete and return the response form provided. Questions can be directed to 1-877-638-2776. |  | 1557 | 1,557 units | Worldwide Distribution -- USA, including the states of TX, OH, PA, KS, FL, NV, IL, WA, MN, NJ, NC, UT, NY, CA, MD, CO. NE VA, and AZ; and, the country of Canada. | Terminated | July 19, 2017 | 1073 | 1557 | Not_Computer
Z-2614-2014 | 68843 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | ***Baxter***BURETROL Solution Set***105" (2.7 m) |  |  | July 14, 2014 | 2014 | Baxter Healthcare Corp | Baxter Healthcare Corporation is voluntarily issuing a recall for specific lots of Buretrol Solution set product codes due to complaints for separation between the burette chamber and the drip chamber. Affected product codes are: 2C7564, 2C8864 and 2H8864. | Nonconforming Material/Component | Action to be taken if product was purchased directly from Baxter:Baxter is requesting that you take the following actions: 1. Locate and remove all affected product from your facility (the product code and lot number can be found on the individual product package or shipping carton).2. Contact Baxter Healthcare Center for Service to arrange for return and credit. The Center for Service can be reached at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time. Please have your Baxter eight-digit ship-to account number ready when calling.3. Complete the enclosed customer reply form, and return it to Baxter by either fax Or scanned e-mail.4. If you are a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, please conduct a recall with your end-user customers in accordance with your customary procedures. Baxter distributed this product to customers between September 27, 2013, and April 25, 2014.Action to be taken if the products were purshased from a distributor or re-seller.1. Locate and remove all affected product from your facility (the product code and lot number can be found on the individual product package or shipping carton).2. Contact Baxter Healthcare Center for Service to arrange for return and credit. The Center for Service can be reached at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time. Please have your ship-to account number ready when calling.3. Follow your supplier's reply and recall process. Please do not return the customer reply form to Baxter. |  | 107232 | 107232 | Nationwide Distribution. | Terminated | March 15, 2016 | 610 | 352612 | Not_Computer
Z-2611-2014 | 68959 | No | PFF | High throughput genomic sequence analyzer for clinical use. | Immunology | Toxicology | 510(K) Exempt | Illumina Worklist Manager (IWM) (software v1.0.15), a component of Illumina MiSeqDx Platform. |  |  | July 28, 2014 | 2014 | Illumina Inc | A software component of the MiSeqDx instrument, called the Illumina Worklist Manager(IWM) v1.0.15, fails to perform as intended | Device Design | Illumina notified affected customers via phone, email and Product Quality Notification letter dated July 28, 2014 to describe the software anomaly. The letter identified the affected product, problem and actions to be taken. For questions contact your local support team or Illumina Technical Support at techsupport@illumina.com. |  | 35 | 35 | Worldwide Distribution - US (nationwide) and the countries of: France, Germany, Italy, Norway, UK and United Arab Emirates. | Terminated | November 12, 2014 | 107 | 35 | Computer
Z-2616-2014 | 69036 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Streamline MIS Spinal Fixation System Fixed Rod Holder |  |  | June 14, 2013 | 2013 | PIONEER SURGICAL TECHNOLOGY, INC. | The affected fixed rod holders did not allow the instrument to lock down on the spinal rod. | Device Design | Pioneer Surgical sent a Product Field Action letter on July 7, 2013, to all affected customers. The letter described the affected product and issue, and provided instructions for return of the recalled product. All 9 affected instruments were returned to Pioneer Surgical, reconditioned and sent back to the field. Customers with questions were instructed to call 906-226-4489.For questions regarding this recall call 906-226-4812. |  | 9 | 9 | Nationwide Distribution including TX, CA, IL, and FL. | Terminated | September 30, 2014 | 473 | 9 | Not_Computer
Z-2618-2014 | 68991 | No | NQX | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | The BD MAX MRSA Assay, catalogue #442953. |  |  | July 23, 2014 | 2014 | Becton Dickinson &amp; Co. | The performance of the assay no longer conforms to this statement in the Analytical Specificity section of the Package Insert:One hundred-eleven (111) out of 111 MSSA strains tested at extremely high concentrations (&gt; 106 CFU/swab), produced negative results with the BD MAX MRSA Assay.One of the samples tested with each of the CAP surveys MRS5-A 2014 and MRS5-B 2014 yielded false positive results. | Under Investigation by firm | Becton Dickinson sent an Field Correction Notification initiated on July 23, 2014. US customers were contacted via letters sent by UPS. The letter identified the product the problem and the action needed to be taken by the customer.The acknowledgement form included with the follow up letter that is faxed back to BD will be reconciled with the customer list. BD will contact 10% of non-respondents to verify receipt of the customer letter. Any product returned to BD will be quarantined and discarded.Nonetheless, you may observe some false positive results for MSSA strains when using the above referenced product. Please contact the BD Technical Service Department at 1-800-638-8663 if you have observed an increase in MSSA false positive results or have failed the above referenced CAP surveys. |  | 585048 | 585,048 tests | US Distribution including the states of : CA, FL, HI, IL, IN, MD, MI, MN, MT, NC, ND, NJ, NM, NV, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA and WI. | Terminated | May 09, 2016 | 656 | 585048 | Not_Computer
Z-2292-2014 | 68785 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Siemens Luminos dRF, Ysio or Uroskop Omnia systems. |  |  | July 07, 2014 | 2014 | Siemens Medical Solutions USA, Inc | It was discovered that during a RAD examination using Siemens Luminos dRF, Ysio or Uroskop Omnia systems with software version VB10C to VB10F and automatic exposure control, a highly unlikely malfunction can result in longer radiation exposure than required. This may also result in an overexposed image that is not of diagnostic quality. As a result the taken examination needs to be repeated. Thi | Device Design | Siemens had informed all affected customers about this issue via Customer Safety Advisory Notice XP023/13/S (dated June 3, 2013) reported under C&R; # 2240869-05/16/13-0013-C. A corresponding software update to fix the issue was originally planned to be rolled out by May 2014. The originally scheduled roll out date could not be met and to ensure that this new software version works accurately Siemens incorporated a comprehensive monitoring phase. Siemens has now scheduled the roll out for the software version VB10J to be completed by end of October 2014. A letter, dated July 7, 2014, with important safety information was sent to end users to notify them that the new software version VB10J will be introduced to solve the potential malfunction of the generator. |  | 192 | 192 | Nationwide Distribution. | Terminated | February 24, 2015 | 232 | 192 | Computer
Z-2623-2014 | 69089 | No | DYE | Replacement heart valve. | Cardiovascular | Cardiovascular | PMA | Hancock Apical Left Ventricle Connector, Model No. H174Axx, Catalog No. 0174ACP012. |  |  | August 15, 2014 | 2014 | Medtronic Cardiovascular Surgery-the Heart Valve Division | Medtronic has decided to discontinue production and distribution of this low-volume product. | Component design/selection | An Urgent Medical Device Recall letter, dated August, 2014, was sent to customers on 8/15/14 for specific serial numbers of Hancock Apical Left Ventricle Connector, model H174Axx, devices. This recall is the result of a decision by Medtronic to discontinue sales of this low volume product, in lieu of fulfilling new regulatory requirements to maintain its distribution in the United States. Customers are instructed to quarantine and return all listed product in their inventory to Medtronic and contact customer service at (800) 854-3570. Customers are instructed to complete the attached customer confirmation certificate and e-mail it to RS.CFQFCA@Medtronic.com or fax to Medtronic at (651) 367-0612 to the attention of Customer Focused Quality. |  | 2 | 2 units | Distributed in the states of MD and TX. | Terminated | November 06, 2014 | 83 | 2 | Not_Computer
Z-2625-2014 | 69043 | No | LGM | Patient care reverse isolation chamber. | General Hospital | General Hospital | 510(k) | Life Force Chamber |  |  | June 27, 2014 | 2014 | Life Force of Tampa, LLC | The Life Force chamber is marketed and promoted as a medical device, intended for use in the diagnosis of disease or other conditions or in the cure, mitigation, treatment, or prevention of disease, or is intended to affect the structure or function of the body. Life Force of Tampa does not have an approved application for premarket approval in effect, or an approved application for an investigat | Labeling False and Misleading | Life Force of Tampa LLC sent an URGENT: MEDICAL DEVICE RECALL letter dated June 27, 2014 to all affected customers. The letter instructed the customers to: 1) Immediately remove and discard all copies of the labeling material. 2) Remove all electronic copies of the materials from any publicly accessible locations (e.g. website). 3) Complete and return the Acknowledgment form to Life Force of Tampa. For questions or concerns contact Douglas G. Cornell at (813) 961-4655. |  | 10 | 10 | Nationwide Distribution in the states of AZ, FL, MD, NV, OK, PA, and TX. | Terminated | July 30, 2015 | 398 | 10 | Not_Computer
Z-2626-2014 | 68908 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Stryker Aria Implant Inserter, REF 48758500, |  |  | July 07, 2014 | 2014 | Stryker Spine | Stryker has initiated the recall of its Aria Standard and Specialty Implant Inserters due to complaints of product fracturing. | Device Design | Stryker Spine sent an "Urgent Medical Device Removal Notification/Customer Response" form dated July 7, 2014 to all their affected customers via Fedex. Customers were instructed to examine their inventory and hospital locations to identify the product. Fax a copy of the Customer Response Form to 855-632-9049 or email it to Spine-RegulatoryAction@Stryker.com.For questions regarding this recall call 201-760-8298. |  | 64 | 64 units | Worldwide Distribution - USA including OR, AZ, TN, MA, MD, GA, FL, NY, IA, IN, TX, MI, PA, UT, CA, MN, MO, OH, KS, OK, MT, FL, AL, CO, TN, and WI and Internationally to Australia. | Open, Classified |  |  | 121 | Not_Computer
Z-2624-2014 | 68954 | No | KCT | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Integra Container Lid, ÂżÂż Size, Red; Container Lid, ÂżÂż Size, Blue; Container Lid, ÂżÂż Size, Red; Container Lid, ÂżÂż Size, |  |  | July 30, 2014 | 2014 | Integra York PA, Inc | Integra LifeSciences has identified through an internal investigation that some colored lids for Integra containersystems were not yet cleared by the FDA for sale and were inadvertently distributed. | Employee error | Integra sent an Urgent Voluntary Medical Device Recall letter dated July 30, 2014, to the affected customer. The letter identified the product, the problem, and the action to be taken by the customer. The customer was instricted if they have distributed the affected product, they should identify those customers adn forward a copy of the Integra Medical Device Recall Notice to them. The customer was also instructed to complete the attached form and return by email FCA@integralife.come or FAX to 1-609-275-9445. For questions the customer was asked to contact Customer Service at 1-855-532-1723.For questions regarding this recall call 717-840-9335. |  | 84 | 84 | US only to one repacker/relabeler and three medical facilities. The relabeler was identified as: KLS Martin, 11201 St. Johns Industrial Parkway S, Jacksonville, FL 32246 (904) 641-7746 | Terminated | March 10, 2015 | 223 | 84 | Not_Computer
Z-2635-2014 | 69038 | No | LEJ | Sunlamp products and ultraviolet lamps intended for use in sunlamp products. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SUN DASH RADIUS 252 PN 8000809705. |  |  | July 25, 2014 | 2014 | JK Products &amp; Services, Inc | The integrity of the wires inside main electrical chord become compromised from excessive flexing (during the door open/close process). | Device Design | The owners of the affected device will be notified via certified letter. After the recalling firm confirms ownership, a technician will be dispatched to install a door clamp and door stop, and inspect the condition of the main electrical chord. |  | 1948 | 1948 units | Nationwide Distribution. | Open, Classified |  |  | 1948 | Not_Computer
Z-2628-2014 | 69062 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | BD BBL Vancomycin Screen Agar Catalog number 222204 |  |  | July 28, 2014 | 2014 | Becton Dickinson &amp; Co. | These lots of Agar may have exhibited breakthrough growth of vancomycin susceptible Enterococcus species control organisms. Agar that allows susceptible enterococci to grow appears as a falsely resistant culture. If a lab does not QC, the error would be identified as falsely resistant enterococci would have further AST workup. There could be a delay in the identification of false-resistance | Under Investigation by firm | BD Diagnostic System sent an Urgent Product Recall letter dated July 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate you may have been shipped the above-referenced lot numbers of product between March 5, 2014 and June 5, 2014. Please discontinue use of these lot numbers and discard any remaining packages. If further assistance is needed, please contact BD Technical Services Department at 1-800-638-8663.The acknowledgement form included with the follow up letter that is faxed back to BD will be reconciled with the customer list. BD will contact 10% of non-respondents to verify receipt of the customer letter. BD has asked distributors to send their customer list to BD for follow up. |  | 28540 | 28540 | Worldwide Distribution: US (Nationwide) including the states of AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI and WV; and Internationally to the countries of Belgium, Canada, Colombia, India, Mexico, Singapore and Taiwan. | Terminated | January 07, 2016 | 528 | 28540 | Not_Computer
Z-2637-2014 | 69086 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ProVue Instrument Driver for JResultNet (ProVue.jrm) v. 1.5.3 and below |  |  | August 15, 2014 | 2014 | Data Innovations, LLC | Potential issues with the ProVue driver have been identified which may cause data to be misassociated with patient results under a given set of circumstances. | Software design | Data Innovations LLC sent an Important Medical Device Information letter on August 15, 2014, to consignees for the JResultNet using ProVue Instrument Driver JResultNet (provue.jrm) v. 1.5.3 and below. Potential issues with the ProVue driver have been identified which may cause data to be mis-associated with patient results under a given set of circumstances. The user on the ProVue transmits the same results twice back to back by clicking the Send to Host button twice. The ProVue sends two ASTM messages back to back before ending transmission (sending an EOT character) The solution was to disconnect JResultNet from the ProVue instrument and apply the patch to driver provue.jrm when it is available. If you have further questions regarding this notice, please feel free to contact us at northamerica-support@datainnovations.com or 239-931-6004 Toll-Free (from the US and Canada): 1-800-332-0499. |  | 84 | 84 | Nationwide Distribution including AL, AZ, CA, CT, DC, FL, GA, IN, LA, MI, MO, NC, NH, NJ, NY, OH, OK, OR, PA, TN, TX, UT, VA, and WA. | Terminated | March 04, 2015 | 201 | 84 | Not_Computer
Z-2640-2014 | 69083 | No | MGQ |  |  | General &amp; Plastic Surgery | 510(k) | Amerigel Post-Op Surgical Kits . Kit includes: Amerigel Wound Dressing, Amerigel Wound Wash, 22x2 Gauze Pads, and 1x3 Fabric Bandages |  |  | July 11, 2014 | 2014 | Amerx Health Care Corp. | Not approved labeling claims (antimicrobial and autolytic debridement) | Labeling design | On July 21, 2014, Amerx Health Care Corporation telephoned all affected customers advising them of the recalled products. During the call the customers were instructed to: 1) Discontinue dispensing the lot(s) and promptly return via parcel post to the Clearwater facility. 2) If you dispose of the product on your own, provide proof of disposal. 3) You will be reimbursed by check or credit memo for returned or disposed goods and postage. |  | 116633 | 116,633 units | Worldwide Distribution-USA (nationwide) including the District of Columbia, Virgin Islands, Puerto Rico, and Guam, and the states of AK, AL, AZ, CA, CO, CT, FL, GA, HI, IL, IN, KS, KY, LA, MA, ME, MD, MI, MT, MO, NC, NH, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WY, and WV, and the countries of Canada, England, Africa, and Australia. | Terminated | July 22, 2015 | 376 | 522101 | Not_Computer
Z-2630-2014 | 68585 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | BD Vacutainer EDTA Tubes (LAVENDER TOP TUBE 2.0ML) |  |  | June 23, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Products were held at improper storage temperature | Storage | Fresenius Medical North America contacted customers via telephone on 6/23/14 by Customer Service and follow-up with formal letter notification, Urgent Recall by certified mail with signature confirmation and fax-back form. Customers instructed to examine their inventory to determine whether they have any of the affected Naturalyte Liquid Bicarbonate or BD Collection Tubes. If customers have the affected product, they will be instructed to contact FMC-RTG Customer Service to have the product replaced. On July 10, 2014 Fresenius issued a clarification letter to users to state only products from the identified lots that were delivered by RTG LLC to the specified facilities on May 5, 2014 are affected by this recall as the products were exposure to higher temperatures than recommended.If you have any additional questions, please contact your FMCNA Customer Service Team at 1-800-323-5188. |  | 26 | 26 boxes | Distributed in the states of LA, OK, and TX. | Open, Classified |  |  | 55 | Not_Computer
Z-2636-2014 | 69049 | No | DQO | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Volcano TrakBack II Disposable Pullback Device; |  |  | August 13, 2014 | 2014 | Volcano Corporation | Internal testing had identified a potential breach to the sterile barrier that may impact the sterility of the product. | Packaging | Volcano sent an Urgent Product Recall Notification letter dated August 13, 2014, to all customers who received the device in the last 25 months. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease use of the affected product and return unused devices to the Volcano distribution center. Customers were asked to prepare the attached inventory log and then contact Volcano Customer Service at 800-228-4728, option 4 to arrange for return of any affected product in their inventory. A credit would be issued for any returned product. Customers with questions were asked to call 800-228-4728.For questions regarding this recall call 978-439-3586. |  | 8251 | 8251 total devices (3777 US) | Worldwide Distribution - USA (nationwide) and Internationally to Korea, Thailand, Canada, Australia, Chili, Mexico, Argentina. | Terminated | June 30, 2015 | 321 | 8251 | Not_Computer
Z-2642-2014 | 68997 | No | JCX | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Vac-Assist Suction Units (SUA01 Suction Aspirator Unit w/800 cc Canister). |  |  | August 03, 2014 | 2014 | EMG Technology Co., Ltd. | suction canister lids may have occluded ports. | Nonconforming Material/Component | A recall letter, dated August 6, 2014, was sent via email which identified the product, problem, and actions to be taken. |  | 1452 | 1452 | US Distribution in IL. | Terminated | August 25, 2015 | 387 | 1452 | Not_Computer
Z-2567-2014 | 68627 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | C-Series High Energy Linear Accelerator - Models: Novalis Tx, Trilogy, Trilogy Tx, Clinac iX, Clinac 2100 C/D, Clinac 2300C/D, Clinac 21EX, Clinac 23EX with versions 7, 8, 9 and 6 MV configuration and maximum dose rate of at least 6000 MU/min only. |  |  | June 17, 2014 | 2014 | Varian Medical Systems, Inc. | Trend in reports of unexpected decrease in beam output in C-series High Energy Linear Accelerators for 6MV photon treatment mode. | Radiation Control for Health and Safety Act | Urgent Medical Device correction letters were sent on June 16, 2017 to all affected users by certified mail. |  | 3416 | 3416 | Worldwide Distribution-USA (nationwide) and the countries of Algeria, Argentina, Australia, Austria, Azerbaijan, Bangladesh, Byelarus, Belgium, Bosniak, Brazil;, Bulgaria, Burma, Canada, Chile, China, Columbia, Costa Rica, Czech Republic, Denmark, Dominican Republic, Ecuador, Estonia, Finland, France, Germany, hong Kong, Hungary, Iceland, India, Indonesia, Iraq, Ireland, Israel, Italy, Japan, Luxembourg, Macau, Macedonia, Malaysia, Mauritania, Mexico, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Romania, Russia, Saudi Arabia,, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Venezuela. | Open, Classified |  |  | 3416 | Not_Computer
Z-2644-2014 | 69032 | No | NHT | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(k) | JBAIDS Instrument System Software version 3.5.0.72. |  |  | August 13, 2014 | 2014 | Biofire Defense | BioFire Defense, LLC is initiating a field action for JBAIDS Software version 3.5.0.72 because the wrong version software may have been installed. | Software Manufacturing/Software Deployment | Consignees were notified on 8/13/14 via letter and telephone and were provider with instructions on how to install the correct software. |  | 43 | 43 Units w/ laptops | Distributed USA (nationwide) and Guam. | Terminated | October 24, 2016 | 803 | 43 | Computer
Z-2643-2014 | 68871 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry A1c_3 CAL (Calibrator) for use on the ADVIA Chemistry Systems, REF/Catalog Number/Siemens Material Number (SMN) 10491408, IVD |  |  | August 06, 2014 | 2014 | Siemens Healthcare Diagnostics | The ADVIA Chemistry Systems A1c_3 method may exhibit a percent bias of -9% to +11% for hemoglobin (Hb)A1c samples when using A1c_3 Calibrator lots 3HD044 and 3LD068. Depending on quality control limits, this issue may not have been detected. NOTE: The percent bias range is a percentage of the HbA1c result. It is NOT absolute HbA1c units. | Under Investigation by firm | Siemens sent an Urgent Field Safety Notice dated August 2014 to all affected consignees on August 6, 2014 via Federal Express. The letter instructed consignees to discontinue use and discard calibrator lots 3HD044 and 3LD068. Customers were requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within thirty (30) days. Distribution of the Urgent Field Safety Notices to foreign consignees was determined at the country level.For questions regarding this recall call 914-524-2955. |  | 1358 | Domestic: 1122 kits; Foreign: 236 kits | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Austria, Brazil, Canada, China, Colombia, Croatia, Cyprus, Denmark, Estonia, Finland, France, Germany, Greece, Guatemala, India, Indonesia, Istanbul, Italy, Lithuania, Mexico, Paraguay, Peru, Poland, Romania, Russian Federation, Serbia, Singapore, Slovakia, South Korea, Spain, Switzerland, Thailand, Uruguay, Vietnam, and Zambia. | Open, Classified |  |  | 1358 | Not_Computer
Z-2653-2014 | 68909 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | VERO Linear Accelerator System, Model Number MHI-TM2000; Common Name: Medical Linear Accelerator. |  |  | July 05, 2014 | 2014 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | Software Anomaly: If a user changes the calendar setting from Workday to Holiday or vice versa, that would erroneously change the status of treatment/fractions, which are completed and to become Treated (completed) status on the exact day when such change is made, to Untreated status. Furthermore, such treatment would be cloned and mistakenly added to the schedule as Untreated treatment | Software design | A Preliminary Customer Information Letter (dated 7/4/14) was emailed to customers on 7/05/14 from Mitsubishi's sales and service representative in Germany. The letter informed customers of the potential safety issue associated with the VERO Linear Accelerator and the recommended actions to be taken. A formal URGENT FIELD SAFETY NOTICE Letter (dated 7/15/14) was emailed to customers on 8/07/14 from Mitsubishi's sales and service representative in Germany. The letter instructed customers to ensure that all potential users in the facility are made aware of the safety notification and the recommended actions to be taken. |  | 3 | 3 units in use in the US | Distributed in the states of NY, FL, and TX. | Terminated | April 01, 2015 | 270 | 3 | Computer
Z-2648-2014 | 69136 | No | GDF | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Bard DuaLok Breast Lesion Localization Wire, |  |  | August 26, 2014 | 2014 | Bard Peripheral Vascular Inc | Bard Peripheral is recalling the Bard DuaLok Breast Lesion Localization Wire because it may be at risk of having portions of the wire protruding from the packaging, thus representing a breach of sterile barrier. | Other | A customer notification letter was sent on 8/27/14 to all customers who purchased the Bard DuaLok Breast Lesion Localization Wire. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to call Bard Peripheral's Recall Coordinator at (800) 321-4254 option#2 Ex 2727. Customers are instructed to fax the completed recall and effectiveness check form to BPV at (800) 994-6772. |  | 177760 | 177,760 units | Worldwide Distribution-USA (nationwide), Australia, Canada, and European, Latin American and Asian Pacific countries. | Terminated | March 06, 2015 | 192 | 177760 | Not_Computer
Z-2652-2014 | 68113 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Tenor is a mobile passive lift, intended to be used on horizontal surfaces for lifting and transfer in hospitals, nursing homes or other healthcare facilities. Transfers with the Tenor are made indoors in the patient's room, in communal areas or in a bathroom. Tenor is meant to be used with the patient type who: sits in a wheelchair, has no capability to support themselves at all, cannot stand unsupported and is not able to bear weight (not even partially), is dependent on caregiver in most situations, is a passive resident, might be almost completely bed ridden, is often stiff and has contracted joints, and/or is totally dependent. The Tenor has a safe working load of 320 KG (705 lbs). |  |  | June 10, 2014 | 2014 | ARJOHUNTLEIGH POLSKA SP. ZO.O. | ArjoHuntleigh received reports where the Tenor Lift without Scale (spreader bar) bolt broke during use. There is a remote risk concerning the breakage of the spreader bar bolt connection, part # TEN.018. During use, the bolt is subjected to mechanical forces, which may lead to breakage. | Device Design | ArjoHuntleigh sent an URGENT - FIELD SAFETY NOTICE letter dated June 10, 2014, to all affected customers. The letter included instructions for direct accounts to: 1) ensure that caregivers and users of the Tenor patient transfer lift referenced above are made aware of the field safety notice; 2) ensure that a copy of the field safety notice is placed in the Tenor Lift Instructions for Usemanual for future reference and training; 3) complete and sign the enclosed Customer Response Form and return the form to the ArjoHuntleighRegulatory Department; 4) and to be advised that an ArjoHuntleigh Service Technician will contact the facility representative listed on the Customer Response Form to schedule service to replace your spreader bar, free of charge. Customers with questions about this recall can contact ArjoHuntleigh at 800-323-1245 x 57985 or via e-mail at AHNA-RA@arjohuntleigh.com. |  | 382 | 382 lifts total worldwide (160 US) | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Bahrain, Canada, France, Germany, India, Ireland, Israel, Italy, Mexico, Netherlands, New Zealand, Poland, Romania, Saudi Arabia, South Africa, Spain, Sweden, United Kingdom, and Uruguay. | Open, Classified |  |  | 382 | Not_Computer
Z-2650-2014 | 69070 | No | KFF | Heparin assay. | Hematology | Hematology | 510(k) | Stago IVD, Liquid Anti-Xa, NDC 3607450003116. |  |  | July 18, 2014 | 2014 | Diagnostica Stago, Inc. | Diagnostica Stago Inc. received some customer complaints regarding a misreading of Reagent 1 (substrate) barcode on STA analyzers. | Under Investigation by firm | Diagnostica Stago sent a "Quality Information" letter/Customer Acknowledgement form to its customers. The letter identified the product, problem, and actions to be taken by the customers. For additional information and further support contact the Diagnostica Stago Hotline at 800-725-0607. |  | 514 | 514 kits | Distributed USA (nationwide) and the country of Canada. | Terminated | January 16, 2015 | 182 | 514 | Not_Computer
Z-2655-2014 | 69054 | No | MEB | Infusion pump. | General Hospital | General Hospital | 510(k) | B.Braun Accuflo 100ML/H 100ML Elastomeric Infusion Device. |  |  | August 06, 2014 | 2014 | B. Braun Medical, Inc. | The primary packaging label (inner pouch label) indicates that the package contents are Catalog Item: CT-1000-250, while the secondary packaging label (outer carton label) indicates the contents are CT-1000-100. The actual contents are a CT-1000-100. | Labeling Change Control | B. Braun Medical Inc. provided the recall notification received from B. Braun Melsungen AG to Progressive Medical Inc. who is the initial importer of the product in the U.S. market. The notification was sent via FedEx and e-mail on/about August 6, 2014 and identified the product, problem, and action to be taken. The notification indicated that if the affected units were further distributed that any distributors are required to forward the letter on to their consignees. |  | 8640 | 8,640 units | Distributed in the state of Missouri. | Terminated | January 29, 2015 | 176 | 8640 | Not_Computer
Z-2651-2014 | 69112 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | Coulter DxH Diluent, Catalog No. 628017 |  |  | August 25, 2014 | 2014 | Beckman Coulter Inc. | Beckman Coulter is recalling Coulter DxH Diluent because it may be associated with elevated platelet (PLT) background counts between 3-7 x 10(3) cells/uL. This will result in failed Daily Checks on the UniCel DxH 800 and DxH 600 Coulter Cellular Analysis Systems. | Under Investigation by firm | Beckman Coulter sent an Urgent Medical Device Recall letter dated August 25, 2014, to all customers that purchased the Coulter DxH Diluent. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to complete and return the enclosed response form within 10 days. Customers with questions were instructed to contact Beckman Coulter's Customer Support Center at http://www.beckmancoulter.com, or call (800) 526-7694 in the US and Canada. |  | 46050 | 46,050 units total (40,078 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Chile, Mexico, Panama, and Thailand.. | Terminated | February 25, 2015 | 184 | 46050 | Not_Computer
Z-2649-2014 | 69067 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Luminos dRF system with SW VD10 or on Luminos dRF Max |  |  | August 12, 2014 | 2014 | Siemens Medical Solutions USA, Inc | It was determined that a potential malfunction may occur on the Luminos dRF system with SW VD10 or on Luminos dRF Max when the table side control is exposed to fluids. Fluid may infiltrate the tableside control panel and result in malfunction. In the very unlikely case of fluid infiltration a permanent short circuit within the control panel of the Luminos dRF / Luminos dRF Max will automatically d | Equipment maintenance | Siemens sent a Customer Safety Advisory Notice XP045/14/S, dated August 12, 2014, to direct accounts, which identified the product, problem, and action to be taken to avoid potential risk to patients and users. Customers were instructed to avoid extensive exposure of fluids to the tableside control panel. When cleaning the system, customers were instructed not to appoly excessive fluids, only use damp towels, in accordance with the product manual. Customers were asked to forward the safety notice to any of their customers and to inform Siemens of the identify of the device's new owner where possible.For questions regarding this recall call 610-219-6300. |  | 33 | 33 | Nationwide Distribution including IL VA TN PA TX IN MN NM GA MI FL CA NY ME OK MO NJ and AZ. | Terminated | May 21, 2015 | 282 | 33 | Not_Computer
Z-2645-2014 | 69117 | No | OLI | Low level laser system for aesthetic use | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ultraslim. These include all products labeled as UltraSlim, Luxurian, BioSlim Lipo, Cellulize, Slim 'n Trim, UltraRed, UltraBlue, Slim-PRO, Dazzle Light, UltraSlim Cold Light, SlenderRay, and U-Slim It, as well as those products with private labelling branding. |  |  | May 06, 2014 | 2014 | Advanced Photonic Systems, LLC | Marketing without marketing clearance or approval. | No Marketing Application | Advanced Photonic Systems sent an Urgent Medical Device Recall letter dated July 12, 2014 via UPS to all those who are known by APS to have possession, custody, or control of the affected device. Customers were instructed to return the recalled product at the sole expense of APS. Each 90 days for the following year, APS plans to provide a progress report to the FDA that identifies which recipients have complied. For more information , customers may write to Advanced Photonic Systems, LLC at the address provided or call 1-800-210-5060 (9am to 5pm, Monday through Friday).For questions regarding this recall call 800-210-5060. |  | 235 | 235 Total | Worldwide Distribution - US- AL, AZ, CA, CO, FL, GA, ID, IL, IN, LA, MA, MD, MI, MN, MO, NV, NJ, NY, NC, OH, PA, TX, UT, and VA.Internationally to Canada, France, and Switzerland | Terminated | June 24, 2015 | 414 | 235 | Not_Computer
Z-2654-2014 | 69126 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | 4D Integrated Treatment Console (4DITC), versions 10 and 11. |  |  | August 18, 2014 | 2014 | Varian Medical Systems, Inc. | Due to user error, the 4DITC can improperly allow users to clear a safety interlock imposed by the 4DITC on behalf of Varian's Bar Code Conical Verification system from the BrainLab's ExacTract console when it is used with BCCV. | Software design | Varian issued an Urgent Medical Device Correction Notification letter dated August 19, 2014, to all affected sites. The letter informed users of the specific product and versions numbers as well as the issue. A list of third party systems that perform verification and communicate authorizations via ADI are included in Table 1 of the letter. The issue is only now to occur with the ExacTract positioning system in combination with a C-Series HE linac BCCV.User recommended actions are provided for users using auxiliary devices requiring ADI interface and for users performing SRS/SBRT.Varian has developed a fix for the issue and customer service representatives will be contacting all affected sites to schedule installation.A copy of the correction letter is to be kept with the most current labeling and all personnel who work in the radiotherapy department should be made aware of the content of the letter.The recall Return response form is to be completed and returned.Questions may be directed to 1-888-827-4265 for US and Canada. Contact information is also provided for other areas of the world. |  | 512 | 512 affected devices | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Brazil, Canada, Chili, China, Columbia, Czech Republic, France, Georgia, Germany, hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Latvia, Lithuania, Macau, Malaysia, Mexico, Netherland,s Portugal, Qutar, Russia, Singapore, Slovenia, South Korea, Spain, Sri Lanka, Switzerland, Taiwan, Thailand, Ukraine and United Kingdom. | Terminated | July 30, 2015 | 346 | 512 | Not_Computer
Z-2656-2014 | 69057 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Zenith Slide W-I-D-E Long Term Care Bed, Model(s) APS98674, APS9867406, APS986741, APS986742, ZG58674, ZG78674, ZZ78674, ZZ98674, and ZZ98774. |  |  | July 21, 2014 | 2014 | GF Health Products, Inc./d.b.a.Basic American Metal Products | The bed head deck may fail to elevate as expected during normal operation. When the non-operation occurs, the head deck remains in the lowered horizontal / flat position and will no longer function as intended. | Nonconforming Material/Component | Consignees were sent a Basic American Medical Products "Product Field Correction Notice" dated July 21, 2014. The letter provided the background information on the recall and the " Required Action" to be taken which includes the replacement of the head deck. Consignees need to contact 1-888-974-4347 to scheduled the replacement order. The replaced head decks should be disposed and disposal documented on the enclosed Letter of Disposal Agreement. The completed letter needs to be returned to Basic American Medical Products. For additional questions they can contact the Field Correction Coordinator at 1-888-974-4347 or electronically: fieldcorrection@grahamfield.com. |  | 1941 | 1941 | Worldwide Distribution: USA (nationwide) including the states of AL, AZ, AR, CA, CO, CT, FL, GA, IL, IN, IA, KY, LA, ME, MD, MA,MI, MN, MS, MO,MT, NE, NH, NJ, NM, NY,M NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, LT, VA, WA, WV, and WI and PUERTO RICO and the country of CANADA. | Terminated | August 11, 2015 | 386 | 1941 | Not_Computer
Z-2658-2014 | 69145 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ZOLL Pro-padz Sterile Adult Multi-Function Electrodes |  |  | August 26, 2014 | 2014 | Bio-Detek, Inc. | Sterility of device is not assured | Process control | BioDetek sent an Urgent Device Correction letter dated August 26, 2014, to all affected customers via UPS overnight delivery. There has also been notifications addressed to the 4 International accounts on August 27, 2014. Customers were requested to return the suspect product(s) in their possession and complete the Device Correction Response ReturnForm provided to each customer. Customer inventory will be replaced as soon as replacement electrodes are processed. Customers with questions can contact the technical support numbers at 1 (800) 348-9011 or +1 (978) 421-9460 24/7.For questions regarding this recall call 401-729-1400. |  | 198 | 198 cases | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Canada, and Germany. | Completed |  |  | 310 | Not_Computer
Z-2657-2014 | 69113 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TRx Vertical Positioner, Classification Name: Wheelchair Component - Power Positioning System. |  |  | August 28, 2014 | 2014 | Motion Concepts | Motion Concepts identified that there is a possibility of knee support failure occurring through fatigue of two attachment bolts in the Vertical Position System (TRx Vertical Positioner). | Device Design | URGENT RECALL NOTIFICATION Letters (dated 8/26/2014) were sent to customers on 8/26/14 via Federal Express. Motion Concepts is requesting that all customers that have purchased a Motion Concepts TRx VerticalPositioner system in between June 2011 - May 2014 to make arrangements for replacement of the knee support mount in accordance with the instructions enclosed in the letter. For questions, please contact us at 905-695-0134. |  | 59 | US: 59 units | US Distribution in states of: AR, CA, CT, DE, IA, ID, IL, IN, MI, MD, MO, NE, NY, OH, PA, TN, TX, and WA. | Terminated | June 03, 2015 | 279 | 59 | Not_Computer
Z-2668-2014 | 68874 | No | BRZ | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Pedi-Pak Pedi-Syringe Filter 60 mL Becton Dickinson/Monoject Syringe |  |  | July 17, 2014 | 2014 | Genesis BPS, LLC. | Genesis BPS is recalling numerous devices since they did not notify the agency of its intent to introduce material changes. | No Marketing Application | Genesis BPS sent an Urgent Medical Device Recall letter with a Recall Letter/Confirmation Form on July 17, 2014 to all affected customers. Customers were asked to immediately examine their inventory and quarantine product subject to recall. In addition, if there is further distribution of the product, customers should be identified and be notified of the recall. Product use should be discontinued and disposed of according to applicable federal regulations and internal protocols.The attached confirmation form should be completed and returned to Genesis by either fax or (201) 708-1104 or e-mail (info@genesisbps.com). Questions can be directed to (201) 708-1400 ext 2004 (9:00 and to 5:00 pm EST).For questions regarding this recall call 201-488-1174. |  | 7272 | 7272 units | Nationwide Distribution &amp; Switzerland | Terminated | March 04, 2016 | 596 | 19440 | Not_Computer
Z-2663-2014 | 69074 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur XP Immunoassay System, Catalog Numbers/Siemens Material Numbers (SMN) 10285219, 10316507, 10317207, 10317284, 10319668, 10320757, 10323213, 10324519, 10327135, 10327836, 10328940, 10329339, |  |  | August 13, 2014 | 2014 | Siemens Healthcare Diagnostics | The firm is conducting a field correction for the ADVIA Centaur and ADVIA Centaur XP Immunoassay systems, regarding manually loaded sample racks that are improperly loaded or pushed too far on the sample entry queue during normal operation of the ADVIA Centaur and ADVIA Centaur XP Immunoassay systems. Pushing the sample racks may cause misreads of Sample IDs (SID) if the sample rack loading instructions are not followed exactly as described in the ADVIA Centaur or ADVIA Centaur XP Operators Guide. | Software design (manufacturing process) | The firm, Siemens, sent "Urgent Medical Device Correction" Letters (dated 8/13/2014) to their consignees/customers informing them of the issue and actions to be taken. Field service personnel were sent a Support Bulletin describing the issue and instructing them how to deal with customer questions. The UFSN must be used by customers as an addendum to the system operating instructions until further notice. The consignees/customers were also instructed to review this letter with your Medical Director; complete and return the attached FIELD CORRECTION EFFECTIVENESS CHECK within 30 days via fax to the Customer Care Center at (302) 631-7597; retain the letter with your laboratory records and forward the letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center, your local Siemens technical support representative and/or call 914-524-2868. |  | 10515 | 10,515 (Domestic: 2771; Foreign: 7744) in total | Worldwide Distribution: US (nationwide) and countries of: Argentina, Austria, Australia, Bahrain, Bangladesh, Belgium, Belgrade, Brazil, Botswana, Canada, Chile, China, Colombia, Croatia, Czech Republic, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Georgia, Great Britain, Hungary, Indonesia, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kazakstan, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Myanmar, New Caledonia, Netherlands, Norway, New Zealand, Pakistan, French Polynesia, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Slovakia, South Africa, Spain, Switzerland, Thailand, Turkey, United Arab Emirates, Vatican, Venezuela, Vietnam, Yemen, and Zimbabwe. | Terminated | May 02, 2017 | 993 | 10515 | Not_Computer
Z-2664-2014 | 69045 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance 64 and Ingenuity CT are Computed Tomography X-Ray Systems intended to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data taken at different angles and planes. These devices may include signal analysis and display equipment, patient, and equipment supports, components and accessories. |  |  | September 01, 2011 | 2011 | Philips Medical Systems (Cleveland) Inc | During a Field Test (a customer external evaluation period), the customer detected a problem in the CIRS 4.0 Beta 2 software. When using the CIRS 4.0 Beta 2 software, during the recon during ready stage of reconstruction, images may be overlapped with, or superimposed on other images. Philips has disabled this feature at all sites evaluating 4.0 Beta 2 software. Philips will issue a software up | Software design | The firm, Philips, sent a Field Safety Notice (FSN 72800534) entitled "URGENT-Medical Device Correction" dated September 1, 2011 via mail on September 6, 2011 to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to review the following information with all members of your staff who need to be aware of the contents of this communication and retain a copy with the equipment Instructions for Use. A Field Service Engineer will visit each site to install the corrected software.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, option 5: Enter Site ID or follow the prompts). |  | 6 | 6 | Worldwide Distribution: US distribution in OH and countries of: Australia, China, Belgium, Israel, and Sweden. | Terminated | July 01, 2015 | 1399 | 6 | Computer
Z-2661-2014 | 69182 | No | KOD | Urological catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Rusch, Soft Rubber (Nelaton and Nelaton Robinson), Bladder Catheter, Sterile. |  |  | September 04, 2014 | 2014 | Teleflex Medical | Misbranding: Although the product labeling identifies the catheters as &quot;soft rubber,&quot; the Natural Rubber Latex caution statement was not printed on the product label. | Labeling Change Control | Consignees were notified via letter on/about 09/04/2014. Subsequently, consignees were notified of the expanded recall on 01/13/2015. |  | 193363 | 193,363 ea. + 1,140 units (total 194,503 units) | Worldwide Distribution-USA (nationwide) and the countries of Austria, Belgium, Bahrain, Belgrad, Canada, Croatia, Cypress, Czeck Republic, Denmark, Finland, Germany, Great Britain, Hungary, Ireland, Italy, Libya, Netherlands, Norway, Romania, Russia, Saudi Arabia, Slavonia, Slovakia, Spain, Switzerland, Turkey, Portugal and Zambia. | Open, Classified |  |  | 193363 | Not_Computer
Z-2666-2014 | 69051 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Brilliance CT Big Bore is a whole-body Computed Tomography X-ray system intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data from the same axial plane taken at different angles. Model Numbers 728244 and 728243. |  |  | August 31, 2012 | 2012 | Philips Medical Systems (Cleveland) Inc | Philips Healthcare received a complaint stating the system opens e-stop while sitting idle causing all movements and scan to stop. The problem has only occurred one time when the scan was out of Ready for Scan or scanning mode. | Software design | A Customer Letter (FSN 72800580), Attachment #1, will be sent to all US and Canadian consignees via certified mail. Philips Key Market representatives will distribute the letter internationally. Additionally, a Field Service Engineer will contact each site to schedule time to implement the Field Change Order. |  | 257 | 257 | Nationwide Distribution incuding states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, MA, MD, MI, MN, MO, MS, NC, ND, NH, NJ, NM, NY, OH, OK, OR, PA, TN, TX, VA, WA, and WI. | Terminated | July 01, 2015 | 1034 | 257 | Not_Computer
Z-2665-2014 | 69129 | [6 empty] | Siemens CentraLink Data Management System, software version 14. |  |  | August 19, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has identified a task scheduler issue with the CentraLink&quot; Data Management System, software version 14. The issue may cause the software to stop executing commands, including uploading validated results to the LIS. The issue is related to an internal software timer that overflows after 24 days. Automated rules triggered by task schedulers or by actions in the CentraLink user interface may also be affected. As each customer configuration is unique, in addition to the delay in uploading of validated results, the specific consequences of discontinued command execution will vary. | Software design | An Urgent Medical Device Correction (UMDC #10819364) letter was issued on August 19, 2014 to all affected customers. The letter instructed customers as follows:Siemens Healthcare Diagnostics is conducting a field correction for the CentraLink" Data Management System. To prevent the issue from occurring, log out of user sessions and restart the CentraLink services and translators as soon as it is convenient, then periodically restart going forward. Any restart frequency that occurs more often than every 24 days is acceptable and will prevent the issue from occurring. Refer to the Additional Information section for detailed instructions. |  | 858 | 858 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Canada, Chile, China, Colombia, Denmark, Egypt, Finland, France, Germany, Greece, Hungary, India, Ireland, Israel, Italy, Lebanon, Lithuania, Luxemburg, Malaysia, Mexico, Netherland, New Zealand Norway, Pakistan, Portugal, Peru, Qatar, Russia, Saudi Arabia, Singapore Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand United Arab Emirates, United Kingdom, Uruguay, Zambia. | Terminated | May 12, 2017 | 997 | 858 | Computer
Z-2619-2014 | 69088 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | STAT 2 Flow Controller; 80" (203 cm) Appx 14.8 mL, 20 Drop Primary Set |  |  | August 19, 2014 | 2014 | ICU Medical, Inc. | ICU is recalling the ConMed Stat 2 flow controller because it may deliver a significantly higher flow rate than was intended. | Nonconforming Material/Component | ICU Medical sent an Urgent Medical Device Recall letter dated August 19, 2014 to all affected customers. The letter identified the affected product, problems identified and the actions to be taken. Customers were instructed to complete the recall response form and fax to ICU Medical at (801) 264-1755 or by email: recall@icumed.com. Customers with questions or require assistance relating to this recall were instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address:customerservice@icumed.com.For questions regarding this recall call 801-264-1732. |  | 3750 | 3750 units | Worldwide Distribution - US (nationwide) in the state of WA, OR and the country of South Africa. | Terminated | November 18, 2014 | 91 | 6250 | Not_Computer
Z-2593-2014 | 69020 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CareFusion EnVe and ReVel Ventilators. |  |  | August 08, 2014 | 2014 | CareFusion 203, Inc. | CareFusion has identified a potential for damage to the power connector on the EnVe and ReVel ventilators due to misconnection by the operator. The input port on the ventilators must have the adapters correctly aligned to function properly. Incorrect alignment may cause the pins in the connector to bend, causing the PTV AC power adapter and/or PTV Auto Lighter Power Cord to be shortened, and pre | Component design/selection | CareFusion sent a Product Recall Notice dated August 8, 2014, to all affected customers. The letter listed the product involved in the recall, described the Issue, Potential Risk , Actions to be Taken by Carefusion, Action to be Taken by the Customer and Other information. Customers were not required to return the devices. Customers were asked to perform the following actions:Once the kit(s) has/have been received, these kits must be fitted and a confirmation must be returned to CareFusion for each kit fitted, identifying the ventilator serial numbers and/or number of AC Power Adapters, Auto Lighter Power Cords to which the kit has been fitted. The confirmation card will be in each kit.To avoid possible delays in providing ventilation and for future compatibility, customers must update their entire inventory. Warranty is void if vent-side LEMO adapter is not properly used in accordance with its instructions and ventilator is damaged due to misuse.Contact Information: CareFusion Support Center, SupportCenter@CareFusion.com Phone 1-888-562-6018, Fax 1-858-617-4851 (Monday through Friday 7AM to 4PM Pacific); Customer Service &amp; Technical Support, ACVentCare@CareFusion.com, Phone: 1-800-754-1914; FAX CS: 1-763-398-8403 (24 hours and 7 days a week between the hours of CST 8:00 am - 5:00pm. After hours and on weekends expect a one hour call back.For questions regarding this recall call 888-562-6018. |  | 1729 | 1729 | Nationwide Distribution including AL, AZ, CA, CO, CT, FL, GA, IL, IN, KS, KY, LA, MA, MI, MN, MS,M0, NE, NV,NH,NH, NY NMC, OH, OR,PA, RID, SC, TN, TX, UT VT, VA, WA, and WI. | Terminated | October 07, 2015 | 425 | 1729 | Not_Computer
Z-2671-2014 | 68974 | No | MZO |  |  | Microbiology | PMA | VITROS Immunodiagnostic Products Anti-HCV Reagent Pack for in vitro qualitative detection of immunoglobulin G antibody to hepatitis C virus (anti-HCV) in human serum and plasma. REF 680 2454, IVD -- Ortho-Clinical Diagnostics High Wycombe UK. Note that one sales unit has 5 Reagent Packs contained within it. Each Reagent Pack can test for 100 samples (patient, Quality Control, etc.). |  |  | August 06, 2014 | 2014 | ORTHO-CLINICAL DIAGNOSTICS | VITROS Immunodiagnostic Products Anti-HCV kit lot 9090 has been observed producing sporadic lower than expected VITROS Anti-HCV test results due to an approximate signal loss of up to 66.2% which can result in unexpected negative test results being obtained for reactive sample fluids (quality control and patient samples). | Packaging process control | All US consignees (users) who were shipped the affected lot VITROS HCV Reagent Pack were sent communications via FedEx and/or US Priority Mail (PO Boxes only) on 8/06/14, and advised of the issue and to immediately discontinue using all remaining inventory of VITROS Anti-HCV Reagent Packs Lot 9090. For additional information, please contact OCD's Customer Technical Service representatives at 1-800-421-3311. |  | 463 | 463 sales units | NY, NJ, CT, MA, PA, VA, MD, FL and GA | Terminated | April 14, 2017 | 982 | 463 | Not_Computer
Z-2670-2014 | 68989 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP BD PosiFlush(TM) SF Saline Flush Syringe 10 mL REF#306553 used in Centurion Convenience Kits. Recalled product is attached to the outside of the Convenience Kits. |  |  | July 25, 2014 | 2014 | Centurion Medical Products Corporation | Unit packages may exhibit open seals which impacts package integrity and potentially product sterility. | Packaging | Centurion sent an Urgent Product Recall Notice dated July 25, 2014, to all affected customers via Certified Mail. The notification discussed the affected product issue and provided instructions for initiating product return. Customers were instructed to identify and remove all inventory of these kits that they may have within their possession. Customers were asked to forward a copy of the notice to any customers to whom this product may have been further distributed. Customers were also asked to complete the enclosed accountability record, whether or not they have the product, and fax or email at their earliest convenience to 517-546-3356 or lcarpenter@centurionmp.com.Customers with questions were instructed to call (517) 546-5400, Ext 1135.For questions regarding this recall call 517-546-5400, ext 1156. |  | 21388 | 21388 | The product was distributed domestically to the following states: UT, WI, TX, NE, AZ, CA, IA, MD, NY, OK, PA, TN, TX, VA, NC, IL, MA, OR, WA, IN, SD. | Terminated | November 25, 2014 | 123 | 21388 | Not_Computer
Z-2672-2014 | 67912 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | EZ-IO 45mm (40kg &amp; up) Needle Set, part number 9079-VC-005; lot code 3606455 |  |  | March 31, 2014 | 2014 | Vidacare Corporation | The voluntary recall is due to the needle set not meeting its specifications. | Material/Component Contamination | Vidacare sent an Urgent Medical Device Recall Notification to all of its customers on March 31, 2014.The letter identified the product the problem and the action to be taken by the customers. The firm asked its customers to return the affected lot for a replacement.Customers were instructed:1. If you have affected stock, immediately discontinue use and quarantine any product with the part number and lot number as identified above2 To return product. complete the enclosed Medical Device Recall Acknowledgment and Return Response Form and fax it to 1 866-431-6194, Attn "Customer Service" This will allow us to document the amount of product you have on hand for return. A customer service representative will contact you with a Return Goods Authorization (RGA)Number and will provide instructions for the return of product to Vidacare Corporation and to facilitate replacement product if one has not already been provided3. If you have already used the affected stock. please complete the enclosed Medical Device Recall Acknowledgement and Return Response Form and fax it to 1-866-431- 6194, Attn" Customer Service". This will allow us to document your receipt of the letter.Vidacare Corporation is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-479-8500. |  | 172 | 172 | Worldwide Distribution-US (nationwide) including the states of LA, NY, CA, CT, WA, AK, NJ, TX, GA, MI, MO, MD, FL, KY, VA, OH, UT, SC, NC, IN, AZ, MN, IL, TN, OK, KS, NE, MT, MS, AR, ID, WV and PA., and Internationally to Canada | Terminated | January 30, 2015 | 305 | 172 | Not_Computer
Z-2678-2014 | 69102 | No | JAX | Pneumotachometer. | Anesthesiology | Anesthesiology | 510(K) Exempt | No packaging. The device is a computed tomography scanner. |  |  | January 09, 2007 | 2007 | Philips Medical Systems (Cleveland) Inc | When the Extended Display Firled ov View reconstruction is used, the CT numbers and anatomy in teh extended region are not accurate. | Labeling design | Philips sent a Product Safety Notification letter dated December 21, 2006 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to distribute the Product Safety Notification to all affected users and keep a copy of the notification with the Instructions for Use until further notice. Customers with questions were instructed to contact Customer Care Center at 1-800-722-9377, option 5 - Diagnostic Imaging, option 1 - CT.For questions regarding this recall call 440-483-7600. |  | 61 | 61 | Worldwide Distribution - US including MO, TX, MD, AR, VA, KY, RI, NJ, MA, NY, LA, PA, VT, SC, CA, OH, FL, MN, MI and Internationally to Netherlands, Canada, Finland, Austria, France, Malaysia, Denmark, Norway, UK, Australia, Belgium, New Zealand. | Terminated | November 06, 2015 | 3223 | 61 | Not_Computer
Z-2569-2014 | 68704 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee and Artis zeego systems. |  |  | June 05, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential problem with Artis zee and Artis zeego systems running software version VC21B and being used in conjunction with the Large Display, in that under certain circumstances, the release of radiation can become blocked unnecessarily. | Software design | Siemens sent an "Important Customer Safety Notice" dated August 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please forward this Safety information to other organizations that could be affected by this measure. If the device has been sold and therefore no longer in your possession, please forward this safety notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible.Please observe this Customer Safety Notice and comply with the corresponding measures until the update has been fully completed. For further questions please call (610) 219-6300. |  | 206 | 206 | Nationwide Distribution | Terminated | April 21, 2015 | 320 | 206 | Computer
Z-2674-2014 | 69178 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | DePuy Synthes TI Vectra Plates: |  |  | September 02, 2014 | 2014 | Synthes, Inc. | DePuy Synthes is initiating a voluntary medical device recall of certain lots of the TI Vectra Plates, which is a part of the DePuy Synthes Vectra Anterior Cervical Plate Family of Systems. The Vectra Systems are intended for anterior screw fixation to the cervical spine (C2-C7). It was discovered in certain lots, that the clips are potentially missing from the TI Vectra Plates resulting in the i | Employee error | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated September 2, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please take the following actions:If you DO have any of the identified devices, please take the following steps: Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number. Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided. Return the Verification Section (page 3 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Send a copy of the completed Verification Section to DePuy Synthes by: Fax: (610) 430-7083 or Scan/email: Fieldaction@synthes.comIf you DO NOT have the identified product, please take the following steps: Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical information. Return the documents to DePuy Synthes by: Fax: (610) 430-7083 or Scan/email: Fieldaction@sythes.com Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450, or contact your DePuy Synthes sales consultant. |  | 16 | 16 | US Distribution including the states of IL, SC, VA, TX, MA, NC, LA, CA, AL, CO, NE, OR and KS., and Internationally to Canada. | Terminated | September 02, 2015 | 365 | 16 | Not_Computer
Z-2677-2014 | 67900 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Circular Lock Screws, part # G826, lot # 11406. The Lock Screws are part of the Genesys Spine TiLock Pedicle Screw System. |  |  | April 02, 2014 | 2014 | Genesys Orthopedics Systems, LLC | The recall is being initiated because MK Precision (the Circular Lock Screw supplier) notified Genesys Spine that a portion (less than 20%) of the Circular Lock Screws in lot 11406 were manufactured using Stainless Steel in the center Post instead of the specified titanium alloy. | Nonconforming Material/Component | Genesys Spine sent an Urgent Medical Device Recall letter dated April 2, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately examine their inventory and quarantine any affected product. Customers were instructed to notify their customers (end users) of the recall and request that they return any affected product to Genesys Spine by completing the attached Recall Tracking Form to Genesys Spine, 1250 South Capital of Texas Hwy, Building Three, Suite 600, Austin, Texas 78746.Customers with questions were instructed to call 512-381-7070. |  | 642 | 642 | Nationwide Distribution including NV, CA, NY, MO, TX, FL, NJ, GA, TN. | Terminated | May 15, 2015 | 408 | 642 | Not_Computer
Z-2673-2014 | 69047 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | BD PosiFlush SF Saline Flush Syringe 0.9 Sodium Chloride Injection, USP 10 mL REF 306553. |  |  | July 22, 2014 | 2014 | Becton Dickinson &amp; Company | BD received some reports of open seals found on the BD Posiflush SF Flush Syringe 10 mL (Ref #306553). | Packaging | BD sent recall letters/recall response cards dated 7/22/2014 via e-mail or UPS 2nd day mail. Customers are asked to examine their inventory and quarantine any affected product subject to recall. Customers should complete the Recall Response Card and fax it back to BD at 1-201-847-6990. Any questions can be directed to 1-888-237-2762 option 3. |  | 3088320 | 3,088,320 units | Nationwide Distribution. | Terminated | April 21, 2015 | 273 | 3088320 | Not_Computer
Z-2676-2014 | 68721 | No | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The synthetic non-absorbable surgical sutures CARDIOFLON Evolution are composed of polyester [poly(ethylene terephtalate)] braids, coated, with or without PLEDGETS of different sizes. |  |  | June 25, 2014 | 2014 | Peters Surgical | A potential splitting of the surface treatment of the thread ends on certain sutures has recently been identified. | Process control | The sole US distributor (MedAlliance Inc.) was notified of the recall via e-mail on July 16, 2014. IMPORTANT AND URGENT - BATCH RECALL PROCEDURE letters dated June 25, 2014 under the Peters Surgical header were sent to all MedAlliance Inc. consignees. The letters included instructions for consignees to contact MedAlliance Inc. by phone (630-933-8661) or by e-mail (agoos@medalliancesolutions.com) for information on what to do with the recalled products. Consignees were also notified that they can contact Peters Surgical directly at qualite@peters-surgical.com for answers to any additional questions. MedAlliance Inc. sent VOLUNTARY PRODUCT RECALL/WITHDRAWAL LETTER[s] dated June 25, 2014 to all consignees. The letters included instructions for consignees to: 1) discontinue use of and segregate recalled product in a secure location for return to MedAlliance Inc., and 2) to complete and return the enclosed RETURNED GOOD AUTHORIZATION (RGA) FORM to MedAlliance Inc. to arrange for the return of the recalled products. Consignees with any questions about the recall can contact Ms. Alaina Goos, Senior Product Manager, by phone at 630-933-8661 x 104 or via e-mail at agoos@medalliancesolutions.com. |  | 22190 | 22,190 units | Worldwide Distribution- US: Nationwide (CA, FL, GA, IL, IN, LA, MD, MI, MO, MS, NC, NY, OH, OR, PA, SC, TX, VA, WA); and countries: ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BANGLADESH, BELGIUM, CAMBODIA, CHINA, CYPRUS, CZECH REPUBLIC, DOMINICAN REPUBLIC, EGYPT, FRANCE, GEORGIA, GERMANY, HUNGARY, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, LEBANON, LIBYA, MALAYSIA, MAURITIUS, MOLDOVA, MOROCCO, MYANMAR, NETHERLANDS, NEW ZEALAND, NORTH KOREA, PAKISTAN, PERU, POLAND, ROMANIA, RUSSIA, SAUDI ARABIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SRI LANKA, SWITZERLAND, TAIWAN, THAILAND, UNITED KINGDOM, VENEZUELA, VIETNAM, YEMEN | Open, Classified |  |  | 27166 | Not_Computer
Z-2679-2014 | 67727 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Mini Cannulated Titanium Headed and Headless Screw Set and 2.5, 3.0, and 4.0 mm countersink/depth gauges in the Mini Screw Set/System |  |  | April 25, 2013 | 2013 | Instratek, Incorporated | Faded and wrong markings on bone screw Countersink/Depth Gauge instruments in Mini Screw Set. | Nonconforming Material/Component | The firm plans to notify their customers of the recall via email, telephone and visit. The firm then plans to scrap and discard any affected product. |  | 96 | 96 | Nationwide: Maine, Oregon, Florida, Texas, Virginia, Arizona, Nebraska, California, Ohio, Illinois, Oklahoma, Washington, North Carolina, Arkansas, Tennessee, New Mexico, Connecticut, Colorado | Terminated | January 26, 2016 | 1006 | 96 | Not_Computer
Z-2691-2014 | 69140 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Justright Surgical, JustRight 5mm Reload, REFJR-REL25-2.0, Sterile, EO, Rx Only |  |  | August 27, 2014 | 2014 | Justright Surgical, LLC | JustRight Surgical is recalling specific lots of JustRight 5mm Stapler and Reload products because they may deliver poor distal staple formation, past the end of the knife cut. | Process design | Consignees were notified via letter on August 27, 2014 in addition, the sales rep will hand deliver the Recall Notification on 8/28/14. Acknowledgement Form will be provided and returned products will be sent to STD Med where it will be held in quarantine. |  | 39 | 3 cases containing 36 units | Distributed in the states of Colorado, California and New York. | Terminated | October 06, 2014 | 40 | 60 | Not_Computer
Z-2686-2014 | 69158 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Naturalyte Liquid Bicarbonate Concentrate |  |  | June 23, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Product was held at temperature above the labeled recommended storage temperature | Storage | Fresenius Medical North America contacted customers via telephone on 6/23/14 by Customer Service and follow-up with formal letter notification, Urgent Recall by certified mail with signature confirmation and faxback form Customers instructed to examine their inventory to determine whether they have any of the affected Naturalyte Liquid Bicarbonate. If customers have the affected product, they are instructed to contact FMC-RTG Customer Service to have the product replaced.A revised/clarification letter dated 7/10/14 issued to state only products from the identified lots that were delivered by RTG LLC to the specified facilities on May 5, 2014 are affected by this recall as products were exposed to temperatures higher than their recommended storage temperature during transportation on May 5, 2014. If you have any additional questions, please contact your FMCNA Customer Service Team at 1-800-323-5188. |  | 72 | 72 cases | Distributed in Texas. | Open, Classified |  |  | 72 | Not_Computer
Z-2680-2014 | 69148 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | The Ysio Systems with software version VC10 |  |  | August 27, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens discovered that an unlikely error may occur on the Ysio system with fixed detector in the wall stand. This error may result in line artifacts in the image. If these artifacts appear in the region of interest, the examination may need to be repeated. | Under Investigation by firm | The firm, Siemens, sent a "Safety Advisory Notice" dated August 27, 2014, to end users/customers. The notice identified the product, problem, and action to be taken. The customers were instructed to follow the instructions in the notice; promptly notify and instruct all the staff at your organization who have to be aware of this problem; forward this safety information to other organizations affected by this measure; and if the device has been sold and therefore no longer in your possession, forward to new owner.Siemens prepared a field modification of the Ysio systems with a hardware upgrade. This field modification will be available in September 2014.If you have any questions, please contact Regulatory Specialist at 610-219-4334 or Customer support at 610-219-6300. |  | 62 | 62 | Nationwide Distribution. | Terminated | March 30, 2015 | 215 | 62 | Not_Computer
Z-2689-2014 | 68955 | No | BZE | Breathing system heater. | Anesthesiology | Anesthesiology | 510(k) | CareFusion AirLife Heated Infant Breathing Circuit |  |  | July 31, 2014 | 2014 | Carefusion 2200 Inc | The CareFusion AirLife Heated Infant Breathing Circuit is being recalled due to a regulatory compliance risk involving a material change. The changes to the gas pathway may potentially impact the safety of the device and quality of the gas condensates to the patient. | No Marketing Application | CareFusion sent an URGENT: PRODUCT RECALL letters dated July 31, 2014 to all affected customers. The letters instructed customers to: 1) perform a 100% physical inventory to verify if any of the affected product codes are in stock; 2) complete and return the attached Recall Response Form within 15 days; 3) notify any customers/parties that the recalled products were distributed to of the recall; 4) distributors should have all their customers return the products and a completed Recall Response Form to the distributor for credit; and, 5) customers should destroy the recalled products according to their disposal procedure. If a customers procedure(s) does not allow for the destruction of the recalled products, the customers was advised to contact CareFusion AirLife Customer Support at 800-323-9088 (Monday - Friday, 8:00 AM - 5:00 PM CST) for product return options. Customers will receive credit for all destroyed products. Direct accounts with questions about this recall should contact CareFusion AirLife Customer Support at 800-323-9088 (Monday - Friday, 8:00 AM - 5:00 PM CST) for additional information. |  | 318280 | Approximately 318,280 devices | Worldwide Distribution - US Nationwide in the states of AZ, CA, FL, IL, KY, LA, MI, MO, NE, NY, OH, OK, PA, RI, TN, TX, UT including Puerto Rico and the countries of Canada, China, Kuwait, Mexico, Saudi Arabia, and Switzerland. | Terminated | February 08, 2016 | 557 | 318280 | Not_Computer
Z-2681-2014 | 69034 | No | PEN | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | The Verigene Gram Negative Blood Culture Nucleic Acid Test (BC-GN), performed using the sample-to-result Verigene System, is a qualitative multiplexed in vitro diagnostic test for the simultaneous detection and identification of selected gram-negative bacteria and resistance markers. BC-GN is performed directly on blood culture media using blood culture bottles identified as positive by a continuous monitoring blood culture system and which contain gram negative bacteria as determined by gram stain. |  |  | October 10, 2013 | 2013 | Nanosphere, Inc. | Nanosphere has received several reports relating to an increased rate of Processing Errors during procedures performed on the Processor SP caused by Tips in Tip Holder Assemblies for the following products and lots of Extraction Trays. | Nonconforming Material/Component | Nanosphere sent a Field Notification Letter dated October 10, 2013, to all affected customers. The letter identified the product, the problem and the action to be taken by the customer. Instructions were included in the letter for customers to determine whether or not the extraction tray lots should be returned. If the extraction tray lots are to be returned, customers were to complete and return the attached form and replacement tip holder assemblies will be supplied at no cost. Customers could then dispose of the recalled unused tip holder assemblies. Customers with questions were instructed to contact their Nanosphere technical support representative.For questions regarding this recall call 1-888-837-4436, ext 3. |  | 400 | 400 units | Worldwide Distribution - USA including AZ, CA, FL, GA, IA, IL, IN, KY, MA, MD, MN, MO, MT, NC, NE, NJ, NY, OH, OR, PA, TX, UT, VA, WA, WI, OH and Internationally to Bulgaria, Germany, Italy, Japan, South Korea, and Spain. | Terminated | June 16, 2015 | 614 | 5920 | Not_Computer
Z-2687-2014 | 69107 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | UniCel 600/800 Access Immunoassay Systems are in vitro diagnostic devices used for the quantitative, semi-quantitative, or qualitative determination of various analytes concentrations found in human body fluids. The cables and kits are used in UniCel DxI 600, 800 instruments, and are distributed to Beckman and Non-Beckman Service Engineers for replacement of worn and/or damaged parts during field services calls. |  |  | August 21, 2014 | 2014 | Beckman Coulter Inc. | Beckman Coulter is recalling the Sample Probe Cable because nonconforming sample probe cables may been installed on newly manufactured and/or recently serviced UniCel DxI 600 and 800 Access Immunoassay Systems, UniCel DxC 660i, 680i, and 860i, and UniCel DxC 880i Synchron Access Clinical Systems instruments. Use of a nonconforming sample probe cable can casuse a delay in reporting results, but wi | Under Investigation by firm | A customer notification letter dated 8/21/14 was sent to all customers who purchased the Sample Probe Cable for use with the UniCel DxI Family of Immunoassay Systems. The letter informs the customers of the problems identified and the actions to be taken. Customer are instructed to complete and return the enclosed response form within 10 days. Customers are instructed to call (800) 854-3633 in US and Canada. |  | 236 | 236 units total (70 units in US) | Nationwide in US and worldwide: Canada, China, Colombia, Croatia, Czech Republic, France, Germany, India, Italy, Japan, Korea, Kuwait, Malaysia, Mexico, Philippines, Puerto Rico, Russia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, United Kingdom, United States, Vietnam. | Terminated | April 30, 2015 | 252 | 236 | Not_Computer
Z-2688-2014 | 69137 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | RENASYS EZ Negative Pressure Wound Therapy (NPWT) System (RENASYS EZ) for patients who would benefit from a suction device particulary as the device may promote wound healing. |  |  | August 13, 2014 | 2014 | Smith &amp; Nephew, Inc. | Modification were made to the RENASYS EZ NPWT Systems without a 510(k) or PreMarket Approval notification. | No Marketing Application | All RENASYS EZ Products will be removed from the field and customers will be transitioned to other product. On August 13, 2014 the firm sent an "Urgent- Medical Device Removal RENASYS EZ" notification via FedEx overnight asking that the use of the RENASYS EZ should be immediately discontinued and the device should be returned to Smith &amp; Nephew. If you have questions, please reach out to your Customer Care representative or call 1-800-876-1261 Monday through Friday, 8:00 AM to 6:00 PM, Eastern Time. |  | 2142 | 2,142 units | Worldwide Distribution- US (nationwide) including MA, MN, NY and Puerto Rico; Canada, Mexico, Dubai, Australia, Chile, Colombia, Germany, France, England, Spain, Italy, Norway, Hong Kong, Shanghai, New Zealand, South Africa, Singapore, Japan, Thailand and the Caribbean Islands. | Terminated | August 06, 2015 | 358 | 2142 | Not_Computer
Z-2587-2014 | 68618 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | AXIOM Artis or Artis zee system |  |  | June 05, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There is a potential problem with the AXIOM Artis or Artis zee system in connection with the hand switch operated at the patient table which may result in uncontrolled release of an x-ray image. | Device Design | Siemens sent an Important Safety Information letter dated June 5, 2014, to affected customers. The letter described the affected product, problem, and actions to be taken. Customers were instructed to observe this safety notice and comply with the corresponding measures until the update has been fully completed. |  | 1067 | 1067 | US Nationwide Distribution. | Terminated | April 06, 2015 | 305 | 1067 | Not_Computer
Z-2692-2014 | 69105 | No | JGJ | Magnesium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ABX PENTRA Magnesium RTU Reagent, Part No. A11A01646 |  |  | August 11, 2014 | 2014 | Horiba Instruments, Inc dba Horiba Medical | Horiba Medical is recalling the ABX PENTRA Magnesium RTU Reagent because the on-board stability claims has been updated. Horiba has updated the ABX PENTRA Magnesium RTU package inserts by removing the 7-day on-boardstability claim and replacing it with a 1-day on-board stability claim. | Other | Horiba sent an Urgent Safety Notice letter dated August 11, 2014 to all customers running the Magnesium RTU assay on the ABX PENTRA 400 Chemistry analyzers that the On-Board Stability has been updated. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. For questions call 888-903-5001, ext 4266. |  | 359 | 359 units | US Nationwide Distribution. | Terminated | November 25, 2014 | 106 | 359 | Not_Computer
Z-2694-2014 | 68487 | No | HMT | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | eye-pak 7407 Tray Support Cover REF 8065740745 |  |  | June 03, 2014 | 2014 | Alcon Research, Ltd. | The peel pouches of the Tray Support Covers may be insufficiently sealed so that sterility cannot be assured. | Packaging process control | Alcon Research sent an Urgent Medical Device Recall letter dated June 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please follow these important steps:1) Review your inventory of Tray Support Covers to determine if you have any units from the affected Lots in stock.2) Place the product aside to ensure it is not used.3) Return the attached Response Form via fax or email to Alcon.Please return the Response Form even if you do not have any inventory from the affected lots.4) Alcon will contact you to arrange the return of your inventory.If you have any further questions, concerns, or need assistance with obtaining replacement product, please contact Alcon Customer Service at 1-800-862-5266. |  | 2462 | 2462 | Worldwide Distribution - Nationwide Distribution and the countries of Australia, Belgium, and Canada. | Terminated | July 27, 2015 | 419 | 2462 | Not_Computer
Z-2693-2014 | 68594 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Minstrel (with scale) |  |  | July 28, 2014 | 2014 | ARJOHUNTLEIGH POLSKA Sp. z.o.o. | ArjoHuntleigh received three reports where the bolt under the scale unscrewed from the scale. There is a remote risk concerning the bolt unscrewing from the scale. When the scale and bolt parts are not correctly serviced, or when they receive a sharp blow during or outside of use, it is possible the bolt holding the spreader bar to the scale could unscrew during subsequent use. As a result, the sp | Device Design | ARJOHUNTLEIGH sent an URGENT - FIELD SAFETY NOTICE letter dated July 28, 2014 to all customers. The letter identified the affected product, problem and actions to be taken. The letters provided customers with two options with how to address this recall: 1) because the devices are close to their 10-year lifetime, customers can return the devices for a trade-in discount that will be applied to the purchase of a new ArjoHuntleigh patient transfer lift; or 2) an ArjoHuntleigh Service Technician will install a new scale with added locking features to the existing unit free of charge. Customers were further instructed to: 1) ensure that all caregivers and users of the devices are given a copy of the letter; 2) place a copy of this letter with the device's instruction for use; and, 3) to complete and return the enclosed Customer Service Response form. For questions or assistance completing the Customer Response Form, please contact ArjoHuntleigh at 1-800-323-1245, ext. 57985 during office hours; Monday to Friday, 8:00am to 5:00pm (CST) or via email at AHNA-RA@arjohuntleigh.com. |  | 18 | 18 Lifts | Worldwide Distribution - US Nationwide in the state of MA and the countries of Australia, Germany, and Italy. | Open, Classified |  |  | 18 | Not_Computer
Z-2695-2014 | 69035 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Fuze Manual Tilt Wheelchair. |  |  | June 12, 2012 | 2012 | Pdg Product Design Group Inc | The hex bolt head might shear and cause the axle plate to detach from the lower frame of the wheelchair and could collapse the wheels while in use. | Component design/selection | PDG Product Design group sent all Canadian customers via an e-mail dated August 9, 2012. The letter identified the product the problem and the action needed to be taken by the customer. Immediately call PDG Customer Service @ 1-888-858-4422. Advise the Customer Service Representative that you have received this letter. Our Customer Service Representative will be happy to organize for you to receive a brand new chair.In the United States, we initially notified our Comfort Company who is our only United States Sales Representative. Comfort Company then notified U.S. Consignees via a letter which is provided for your reference. Also a letter was sent to our biggest customer in the USA, which is Cair Medical. It was sent on December 12, 2012 (PDG Letter to Cair Medical Fuze recall) to explain and outline how PDG handled the Fuze hex bolts breaking. A copy of this letter is also provided for your reference.Our Hong Kong customers were informed thru e-mail. Non- respondents were re- contacted by either e-mail or telephone.Effectiveness checks were done thru telephone. Defective product on hand at our facility has been scrapped and destroyed. Defective product that was removed from wheelchairs in the field was destroyed at the time the bolts were replaced. |  | 153 | 153 | Worldwide Distribution- US Distribution including the states of : AZ, CA, GA, KY, MI, MN, MO, MS, NC, NE, NY, OH, PA, SC, TX, VA, and WA., and Internationally to Canada and Hong Kong. | Terminated | July 23, 2015 | 1136 | 153 | Not_Computer
Z-2696-2014 | 68521 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | UltraStream Chronic Hemodialysis Catheter Standard Kit; Model Number 32101524 (24 cm), 32101528 (28cm), 32101532 (32 cm), and 32101536 (36 cm) and 32102540 (40 cm). The Rex Medical UltraStream Chronic Hemodialysis Catheter is a chronic hemodialysis catheter that can achieve high flow rates at low arterial pressure. |  |  | June 05, 2014 | 2014 | Argon Medical Devices, Inc | UltraStream Chronic Hemodialysis Catheter may develop cracks on the red (arterial) luer during attachment or manipulation. | Component design/selection |  |  | 3783 | 3,783 units | Nationwide Distribution. | Terminated | April 08, 2015 | 307 | 3783 | Not_Computer
Z-2707-2014 | 69223 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Medical Patient Monitor is a Physiological, Patient monitor With Arrhythmia Detection or Alarms. The Models are 91367, 91369, 91370, 91387, and 91393 are sold with Perioperative Mode (option D), which includes Start case/End case functionality, Upgrade Kit 040-1548-00. |  |  | August 26, 2014 | 2014 | Spacelabs Healthcare Inc | The clinician did not realize the monitor was in END CASE mode and was expecting audible alarms. Visual alarms continue to work normally.The firm performed a retrospective analysis of all recalls since 2011 and discovered 222 additional serial numbers were missed and were not included in the previous recall under RES 59456 (recall number Z-3032-2011). | Environmental control | Spacelabs sent an "URGENT - MEDICAL DEVICE CORRECTION (Expanded)" letter dated 5 September 2014 to all consignees in the US. This letter identified recalled model 91367, 91369, 91370, 91387, and 91393, Upgrade Kit 040-1548-00, Equipped with the Perioperative Option (-D), problem, and actions to be taken.On September 12, 2014, a recall notification letter (translated asnecessary) will be emailed to all international subsidiaries and distributors of record.The notification letter stated that the Operating Room staff may not be aware that alarm tones are off if the device has been left in the END CASE state. If the customers make use of the Perioperative START CASE / END CASE feature, they should advise the clinicians to familiarize themselves with this feature as described in the Perioperative section of the Operations Manual. In particular,they should know how to confirm the state of the alarm tones prior to starting a case.The firm's Field Service Engineers will contact customers to schedule a time to come out to update affected monitors at no cost.Customers have any questions about this corrective action program, please contact Spacelabs Healthcare at 1-800-522-7025. |  | 222 | 82 units distributed in US and 140 units distributed outside US | Worldwide Distribution- US (nationwide) including the states of: FL, GA, IA, KY, LA, MA, MI, MN, MO, MS, NJ, NY, OH, OK, PA, SC, TN, TX, WA, Guam and Puerto Rico and to the following countries: Argentina, Australia, Canada, China, Czech Republic, Finland, France, Germany, Greece, India, Israel, Italy, Netherlands, New Zealand, Panama, Paraguay, Poland, Singapore, Spain, Sri Lanka, Switzerland, Tunisia, United Kingdom, and Uruguay. | Terminated | June 26, 2015 | 304 | 222 | Not_Computer
Z-2706-2014 | 69116 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Vapotherm Precision Flow,Disposable Patient Circuit |  |  | August 13, 2014 | 2014 | Vapotherm, Inc. | Water may leak into the center gas lumen of the Patient Circuit | Process control | Vapotherm issued an Initial Notification on August 13, 2014. The letter was updated on August 20, 2014 to include two additional lots. The recall letter notified users that this defect is easily detectable during normal use and theprobability of serious injury or death is remote. The letter was sent to raise awareness of the recent increase in incidence rate of this issue and reminded users of the requirement in the operating manual that requires that the cannula be connected to the patient only after the set temperature is reached. The letter also advised patient care personnel not to use, or to discontinue use of any Disposable Patient Circuit if water is observed leaking into the center lumen or from the cannula. They were instructed to replace the device and report the defect to Vapotherm immediately for further investigation.Should you have any questions please contact our technical support team atEmail: ts@vtherm.com(US) 855.557.8276(INTL) 1-603.658.5121 |  | 207 | 207 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, France, Swizerland, Korea, Sweden, Germany , Belgium, Spain, Mexico, Scotland, UAE, Australia, Israel, Hong Kong, Ireland, kuwait, and Italy. | Terminated | October 02, 2015 | 415 | 48329 | Not_Computer
Z-2708-2014 | 69191 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Dako Autostainer Link 48 with software version Dako Link 4.0.3, Model No. AS480 |  |  | August 29, 2014 | 2014 | Dako North America Inc. | Dako is recalling the Dako Link 4.0.3 software on the Dako Autostainer Link 48 because Dako has determined that an upgrade to this version of the software may cause premature rejection of reagent bottles or inadequate aspiration of reagent if the instrument has also had a preventative maintenance or waste pump replacement, and the baseplate is not completely aligned. | Software change control | Dako North America sent an Urgent Medical Device Correction letter to all customers who were upgraded to the Dako Link 4.0.3 software that is used on the Dako Autostainer Link 48 and who also had preventative maintenance or a waste pump replacement. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed that Dako will contact customers to set up an appointment for a service visit in order to ensure that the instruments are running safely or reliably. Customers with questions are instructed to contact Dako. |  | 191 | 191 units total (57 units in US) | Worldwide Distribution - US Nationwide in the states of AT, AU, BE, CA, CN, DE, DK, ES, Fl, FR, GB, IE, IT, NL, NO, SE, US, DZ, GR, IS, OM, SG, UY. | Terminated | March 03, 2015 | 186 | 191 | Computer
Z-2709-2014 | 69123 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance CT System 6/10/16/16P/40 configuration. No packaging, sold as single unit. |  |  | December 27, 2005 | 2005 | Philips Medical Systems (Cleveland) Inc | The system has a gap that can appear between the upper and lower tilt stand covers. The fingers may be squeezed if a person was to place them in this region while the gantry was tilting from an angle to zero tilt angle. | Device Design | Customers will be contacted by a field service technician with the field correction order. The firm will install the new labeling to the device. |  | 364 | 364 | Worldwide Distribution -- AZ, CO, GA, IA, IL, IN, KS, KY, MD, MI, MN, MS, NC, NY, OH, OK, TX, VT, and WI; and, countries of Singapore, New Zealand, India, Australia, Korea, Japan, Korea, China, India, Taiwan, Australia, Turkey, France, South Africa, Spain, Italy, Austria, Belgium, Switzerland, Israel, Netherlands, England, and Brazil. | Terminated | November 06, 2015 | 3601 | 364 | Not_Computer
Z-2719-2014 | 69025 | No | GZB | Implanted spinal cord stimulator for pain relief. | Neurology | Neurology | 510(k) | Percutaneous Trial Lead Kit Model 3086 (8 ch percutaneous lead). |  |  | July 29, 2014 | 2014 | St. Jude Medical, Inc. | The Directions for Use (DFU) manual contains information for devices that are not approved in the US. The correct US approved versions of the manuals were sent to the customers to replace the incorrect international manuals. | Packaging process control | A field action was initiated with customer letters, dated July 29, 2014, and attached replacement DFUs. The letter identified the affected product and the reason for the recall. The recalling firm requested the consignees refer to the replacement DFU for guidance while using the device, and to discard the international manuals. Questions regarding the letter should be directed to customers' St. Jude Neuromodulation Representative or call the Technical Services Department at 1-800-727-7846. |  | 4 | 4 units | Distributed in the states of Texas and Illinois. | Terminated | August 21, 2015 | 388 | 6 | Computer
Z-2711-2014 | 69235 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Ingenuity CT, Ingenuity Core, and Ingenuity Core128 scanners are Computed Tomography X-ray systems intended to produce cross-sectional images of the body by computer reconstruction of X-ray transmission data taken at different angles and planes. This device may include signal analysis and display equipment, patient and equipment supports, components and accessories |  |  | September 11, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Ingenuity CT system can mislabel the exposure that results in an image in the incorrect order that it was taken by the user. | Software design | Philips sent an Urgent - Field Safety Notice Medical Device Correction letter dated September 5, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please review the following information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication. Please retain a copy with the equipment Instruction for Use.Additionally, a Field Service Engineer will contact each site to schedule a time to install the software update.If you need any further information or support concerning this issue, please contact your local Philips representative or local Philips Healthcare office. For North America, contact the Customer Care Solutions Center(1.800.722-9377 and follow the prompts). |  | 79 | 79 | Worldwide Distribution - US including the states of : AZ, CA, CT, FL, GA,MA, MS,OH, PA, TX, and WY., and the countries of : Australia, Denmark, Finland, France, Germany, India, Ireland, Japan, Malaysia, Netherlands, New Zealand, Norway, Russia, Slovakia, South Africa, South Korea, Sweden, Switzerland and United Kingdom. | Open, Classified |  |  | 79 | Computer
Z-2717-2014 | 68329 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CelluTome Epidermal Harvesting System Part number CT-H25 or Part Number CT-H50 |  |  | May 19, 2014 | 2014 | KCI USA, Inc. | Blade can shift during shipping which makes the product ineffective. | Device Design | KCI sent an Urgent - Voluntary Medical Device Field Notification letter dated June 6, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm is sending Field Notifications to affected customers instructing them to complete appropriate actions.Option I: Contact your local KCI Representative or KCI Customer Technical Support at 1-800-275-4524 and select Option 3 to arrange for return and replacement of the affected harvester(s).Option 2: If you prefer, you may re-set the blade after removal from the packaging by gently tapping the curved end of the harvester against the palm of your hand, as shown in Figure 3, until the blade is no longer visible in the harvester top plate, as shown in Figure 4. |  | 1331 | 1,331 units | US nationwide Distribution | Terminated | September 18, 2015 | 487 | 1331 | Not_Computer
Z-2712-2014 | 68932 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Symbia T Series consists of the T16, T6, T2, and T. The Symbia T integrates state-of-the-art SPECT and high quality spiral CT to provide precise attenuation correction and anatomical mapping. The system can also be used as a complete clinical SPECT system for general purpose, WB and SPECT applications. |  |  | July 28, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Siemens Medical Solutions USA, Inc. is performing a field correction because we recently detected a rubber isolator with unfavorable aging properties. The function of the rubber isolator, which is situated within the rotate motor mounting assembly, may degrade over time and may impact motor support. The event does not pose a hazard by itself but could lead to a rupture of the drive belt which in t | Component design/selection | Siemens sent an Undated Customer Safety Advisory Notice letter to all direct accounts (customers). The letters included instructions for customers to: 1) continue to use the system(s) while waiting for the field correction to be scheduled and performed unless the system exhibits abnormal noise or vibration in which case the system(s) should not be used and the local Siemens Medical Solutions USA service engineer should be contacted; 2) ensure that the letter is placed in the Systems Owner's Manual and share it with support staff; and, 3) if the product was sold to another party, forward the letter to that party. A local service engineer will contact the customer to schedule the servicing of the affected systems by December 2014 which will consist of replacing the mounting plate. If customers have any questions about this issue, customer can contact their local service representative by calling 800-888-7436. |  | 1039 | 1039 systems | Worldwide Distribution - USA including AL, AK, AR, AZ, CA, CO, CT, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, NC, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VA, VT, WAM WI, WY, District of Columbia, Puerto Rico and Internationally to Algeria,Argentina,Australia,Austria,Bangladesh,Belgium,Brazil,Bulgaria,Canada,Chile,China, Colombia,Croatia,Czech Republic,Denmark,Ecuador,Egypt,Finland,France,Germany,Hong Kong,Hungary,India,Indonesia,Iran,Ireland,Israel,Italy,Japan,Kazakhstan,Korea, Kuwait,Latvia,Libya,Lithuania,Malaysia,Mexico,Morocco,Myanmar,Netherlands,New Zealand,Norway,Oman,Pakistan,Peru,Philippines,Poland,Portugal,Qatar,Romania,Russia,Saudi Arabia,Singapore,Slovakia,Slovenia,South Africa,Spain,Sweden,Switzerland,Taiwan,Thailand,Tunisia,Turkey,United Arab Emirates,Ukraine,United Kingdom,Vietnam | Open, Classified |  |  | 1087 | Not_Computer
Z-2715-2014 | 69084 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur Immunoassay System (including refurbished units), |  |  | August 19, 2014 | 2014 | Siemens Healthcare Diagnostics | The firm has identified an issue with patient demographic information sent to the LIS from the ADVIA Centaur/ADVIA Centaur XP Immunoassay systems. Siemens has confirmed that under extremely rare circumstances patient demographic data from the previous order received from the LIS is merged with the next order. This issue can occur when the LIS data buffer on the ADVIA Centaur system becomes full | Software design | Siemens sent an Urgent Medical Device Correction Letter dated August 19, 2014, via Fed Ex to US customers and via email to regional countries for implementation outside the US (OUS). The letter identified the product the problem and the action needed to be taken by the customer.The letters inform the customers of the issue and specific actions to be taken if the instrument is interfaced to an LIS system that transmits patient demographics with each order. Field service personnel were sent a support bulletin describing the issue and instructing them how to deal with customer questions. Please complete and return the Field Correction Effectiveness Check attached to this letter within 30 days. If you have received any complaints of illness or adverse events associated with the products listed in Table 1, immediately contact your local Siemens Customer Care Center or your local Siemens technical support representative.Please fax this completed form to the Customer Care Center at (302) 631-7597. If you have any questions, contact your local Siemens technical support representative |  | 10515 | Domestic: 2771 units; Foreign: 7744 units | Worldwide Distribution including the United States., and the countries of Argentina, Austria, Australia, Bahrain, Bangladesh, Belgium, Belgrade, Brazil, Botswana, Canada, Chile, China, Colombia, Croatia, Czech Republic, Cyprus, Denmark, Egypt, Estonia, Finland, France, Germany, Georgia, Great Britain, Hungary, Indonesia, India, Ireland, Israel, Italy, Japan, Jordan, Korea, Kazakhstan, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malaysia, Malta, Martinique, Mexico, Myanmar, New Caledonia, Netherlands, Norway, New Zealand, Pakistan, French Polynesia, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovenia, Slovakia, South Africa, Spain, Switzerland, Thailand, Turkey, United Arab Emirates, Vatican, Venezuela, Vietnam, Yemen and Zimbabwe. | Open, Classified |  |  | 10515 | Not_Computer
Z-2714-2014 | 69194 | No | CAI | Anesthesia breathing circuit. | Anesthesiology | Anesthesiology | 510(K) Exempt | Hudson RCIÂż Pediatric Anesthesia Breathing Circuits, Single Patient Use, Rx Only, |  |  | September 08, 2014 | 2014 | Teleflex Medical | The pediatric breathing circuits can crack prior to and during use. | Device Design | Teleflex Medical sent an Urgent Medical Device Recall dated September 8, 2014, to all affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to immediately discontinue use, quarantine any products with the product codes and lot numbers provided in the notification, return any affected product, complete and return the enclosed Recall Acknowledgement Form and fax it to the attention of Customer Service. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 27176 | 27,176 units | Worldwide Distribution - US Nationwide in the states of AZ, CA, FL, IL, KY, OH, NJ, NY, TX, LA, OR, SD, WA, including Puerto Rico, and the countries of Belgium, Canada, Chile, China, Colombia, Korea, Mexico, Peru, Venezuela and Zaire. | Terminated | August 16, 2016 | 708 | 27176 | Not_Computer
Z-2720-2014 | 69063 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Dimension VistaÂż CHEM 1 CAL, lot 3GM081. |  |  | August 18, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has received complaints for low recovery of Magnesium (MG) Quality Control (QC) and patient samples following calibration of MG with the Dimension Vista&reg; CHEM 1 CAL, lot 3GM081. Investigation by Siemens confirms that QC and patient samples under-recovers MG by 0.25 mg/dL [0.10 mmol/L] across the assay range following calibration with 3GM081. | Process design | Siemens sent an Urgent Medical Device Dorrection letter, dated August 2014, to all affected consignees. The letter identified the product, the problem, and the actions to be taken. Consignees were instructed if they still have an active calibration of MG with this lot of CHEM 1 CAL, they should immediately calibrate MG with an alternate lot of CHEM 1 CAL. Consigness were instructed to retain the letter with their laboratory records, and to forward the letter to those who may have received the affected product. Customers with questions were instructed to contact their Siemens Customer Care Center - Techhnical Solutions or their local Siemens technical support representative. |  | 9769 | 9769 Total (6000 US) | Worldwide Distribution - USA (Nationwide) and Internationally to AustraliaAustriaBelgiumCanadaCzech RepublicFinlandFranceGermanyGreat BritainItalyJapanNetherlandsNew ZealandNorwayPortugalSlovak RepublicSloveniaSouth KoreaSpainSwitzerland | Terminated | February 02, 2015 | 168 | 9769 | Not_Computer
Z-2737-2014 | 69039 | No | GFC | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Persona The Personalized Knee System 2.5 mm Female Screw 48 mm Length Sterile Qty-2. Sterilized using irradiation, single use. |  |  | September 09, 2014 | 2014 | Zimmer, Inc. | The Persona 48mm X 2.5mm Female Screw is being recalled due to potential use of the screw in cortical bone, and corresponding stripping of the 2.5mm Persona Hex Driver. | Device Design | Zimmer sent an URGENT MEDICAL DEVICE RECALL notifications dated September 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Review the notification and the updated portion of the Persona surgical technique provided in Attachment 1, and ensure affected personnel are aware of the contents.2. Locate and quarantine any affected product listed above and notify your Zimmer sales representative.3. Your Zimmer sales representative will remove the recalled product from your facility.4. If after reviewing this notification you have further questions or concerns please call the customer call center at1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 16380 | 16,380 units | Nationwide Distribution including the states of : HI, CA, UT, NY, OH, IN, AL, FL, GA, WI, IA, MN, NE, SD, NJ, PA, ND, IL, MO, VA, MD, NC, MI, OR, WA, KS, KY, TN, WV, OK, CT, MA, ME, AZ, NM, TX, LA, ID, CO, MT, UT, NV and AK., and Internationally to Singapore, Japan, Germany and Canada. | Terminated | May 05, 2015 | 238 | 16380 | Not_Computer
Z-2729-2014 | 68725 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Reusable laparoscopic instruments are packaged as non-sterile. Cleaning and sterilization of devices must occur prior to use. |  |  | May 30, 2014 | 2014 | Instrumed International, Inc. | Additional caution statements have been added to the Instructions For Use for monopolar electrodes. 1) To avoid tissue carbonation, the operation voltage of the HF generator must not exceed 650 peak voltage (Vp) for all monopolar electrodes. and 2) For all monopolar electrodes, the useful life of this product is less than or equal to 50 cycles and less than or equal to 2 years. | Labeling design | Instrumed sent an URGENT: MEDICAL DEVICE CORRECTION letters dated May 30, 2014 to all consignees. The letters advised the consignees of the new warning statements on the product Instructions For Use and an instruction to complete and return the attached Response Request Form to Instrumed International. Consignees were further instructed to extend the recall to the retail level. For questions contact Instrumed Quality Manager at 847-908-6119. |  | 15 | 15 instruments | Distribution US Nationwide in the states of: AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY. | Open, Classified |  |  | 4504 | Not_Computer
Z-2736-2014 | 68014 | No | PEN | Multiplex nucleic acid assay for identification of microorganisms and resistance markers from positive blood cultures. | Microbiology | Microbiology | 510(k) | Verigene BC-GN Kit (Catalog number 20-005-021); 1) 20 Verigene BC-GN Test Cartridges; are qualitative multiplexed in vitro diagnostic tests for simultaneous detection and identification of selected gram-positive (BC-GP) and gram-negative (BC-GN) bacteria and resistance markers. |  |  | June 17, 2013 | 2013 | Nanosphere, Inc. | There is a specific Extraction Tray lot containing Tips that may slightly increase the occurrence of Process SP Processing Error when used due to the Processor SP experiencing difficulty releasing the Tip to the final location in the Tip Assembly. The difficult releasing the Tip can result in a Processing Error and occasionally the Tip to be dropped into the Processor SP. | Process change control | Undated recall notification letters were sent to all direct accounts (customers) June 17 - 19, 2013. The letters included instructions for customers to: 1) check their inventory for the recalled products; 2) set any recalled products aside for return; and, 3) contact Nanosphere to obtain replacement tip holders. Customers with questions about this recall can contact Nanosphere via e-mail at technicalsupport@nanosphere.us or by calling 888-837-4436. |  | 20 | 20 extraction trays | Worldwide Distribution:US (Nationwide) including states of: AZ, CA, FL, IL, KY, MD, MN, MT, OH, PA, TN, TX, VA, WA, WI; and countries of: Belgium, China, and Saudi Arabia. | Terminated | March 14, 2016 | 1001 | 1680 | Not_Computer
Z-2733-2014 | 69139 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Array LaserLink, Manufactured by Lumenis, |  |  | August 15, 2014 | 2014 | Lumenis, Inc. | Lumenis initiated a field-correction for the Array Laser Link&quot;, GA-0006700 (SN XXYYZZ) because the system may project an erroneous pattern display on the retina, which is different than the desired pattern. | Software design | Lumenis sent a Safety Advisory Notice dated July 15, 2014, to all affected consignees. The letter identified the product, the problem, and the action to be taken by the consignee. Consignees were instructed that Lumenis would contact consignees to schedule a product update, to be completed at a convenient time. For questions consignees should call 801-656-2549.For questions regarding this recall call 801-656-2690. |  | 16 | 16 | Worldwide Distribution - USA (nationwide) and Internationally to JAPAN, CHINA, and CANADA. | Terminated | October 21, 2014 | 67 | 16 | Not_Computer
Z-2734-2014 | 69053 | Yes | CBF | Hyperbaric chamber. | Anesthesiology | Anesthesiology | 510(k) | The Rejuvenator Chamber manufactured by Baro-Therapies, Inc. |  |  | August 08, 2014 | 2014 | Baro-Therapies, Inc | The firm is distributing the Rejuvenator device without an approved 510(k). | No Marketing Application | Baro-Therapies issued a Correction Notice letter on August 11, 2014, to all affected customers. The letter stated they have suspended marketing of the device and plans to shut down their website until they receive clearance from the FDA. For questions or concerns contact the firm at (850) 654-0051. |  | 10 | 10 | US Nationwide Distribution in the states of FL, GA, MS, TN and CO. | Terminated | August 06, 2015 | 363 | 10 | Not_Computer
Z-2731-2014 | 69213 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | ASSEMBLY, PATIENT SIDE CART, PS4000 used with the da Vinci Xi Surgical System. |  |  | September 11, 2014 | 2014 | Intuitive Surgical, Inc. | Correction due to the detection of a motor sensor failure in the Patient Cart instrument arms of the da Vinci Xi (IS4000) Surgical Systems after a mid-procedure restart. | Other | Intuitive sent an Urgent Medical Device Correction letter dated September 11, 2014 to all consignees via trackable mail system. Letters identified the issue and provides information for users should a error code be displayed. For questions contact your Clinical Sales Representative or Intuitive Surgical Customer Service |  | 75 | 75 devices | Worldwide Distribution - US nationwide and the countries of: Belgium, Germany, China(Hong Kong), India, Israel, Italy, and the United Kingdom. | Terminated | October 23, 2014 | 42 | 75 | Not_Computer
Z-2732-2014 | 69132 | No | LTI |  |  | Gastroenterology/Urology | PMA | LAP-BAND AP SYSTEM AP Large with Access Port, Allergan, LAP-BAND AP SYSTEM Access Port I. Product was distributed for both the AP Large and AP Standard sizes. |  |  | August 15, 2014 | 2014 | Apollo Endosurgery Inc | Product was distributed past its expiration date. | Software Manufacturing/Software Deployment | Firm called its consignees and advised them to either return or destroy the devices which had not already been implanted. |  | 10 | 10 | Distributed in the states of California, Delaware, Indiana, Missouri, New Jersey, New York, and Oregon. | Terminated | February 01, 2016 | 535 | 10 | Not_Computer
Z-2738-2014 | 69122 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | SPECTRUM Pump, Model No. 35700BAX. |  |  | September 03, 2014 | 2014 | Baxter Healthcare Corp. | One Service Technician may not have correctly serviced specific Sigma Spectrum Infusion Pumps according to established procedures during the time period of 5/5/2014 through 6/3/2014. | Equipment maintenance | Initial notification was initiated via telephone call by the Baxter Medina Service Center on 8/28/14 to all affected Sigma SPECTRUM Infusion Pump consignees. URGENT DEVICE CORRECTION Letters (dated 9/03/2014) were sent to the consignees via USPS first class mail on 9/03/14 formally informing them of the recall. The letters identified the affected product, the description of the issue, the hazard involved, as well as, the actions to be taken by customer. Customers are being instructed to immediately remove the pump from use and return the pump to Baxter for inspection. The firm will provide replacements. Customers are to contact Baxter Healthcare Medina at 800-356-3454 for technical questions regarding the letter. |  | 58 | USA: 56 units, Canada: 2 units | Worldwide Distribution -- USA and Canada. | Terminated | March 16, 2017 | 925 | 58 | Not_Computer
Z-0001-2015 | 69192 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Expression MR200 MRI Patient Monitoring System Model 866120 |  |  | September 08, 2014 | 2014 | Invivo Corporation | During production, excess metal (aluminum) debris was found stuck to the cling plastic that covers the device's pole assembly. The metallic debris could affect system performance due to the potential for a failure to power up or loss of power. This could result in the loss of monitoring and a delay in diagnosis of the patient. | Material/Component Contamination | The firm, Philips Healthcare, sent an "URGENT- Medical Device Recall" letter dated June 17, 2014, to its customers informing them of the product, problem and actions taken. The firm informed the customers that there are no additional actions required on their part. The Philips/Invivo Account Manager will contact them in order to arrange for replacement of the device. If you need any further information or support concerning this issue, contact your local Philips/Invivo representative: 877-468-4861 option 1, then option 2. |  | 10 | 10 units | Worldwide Distribution: US distribution in state of: CA and countries of: Poland, Singapore, and Sweden. | Terminated | January 09, 2017 | 854 | 10 | Not_Computer
Z-0002-2015 | 69258 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Siemens Ysio Max system |  |  | September 02, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Two malfunctions posing potential risk to patients were identified with Siemens Ysio Max, AXIOM Luminos dRF Max, Uroskop Omnia Max, and Luminos Agile Max systems with a specific set of serial numbers that will cause lost images due to an automatic Radiology Information System (RIS) worklist update, thus causing examinations to be repeated. Image label may also be displayed incorrectly after images | Software design | Siemens sent a Safety Advisory Notice, dated September 2, 2014, to all affected customers. The Notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to disable the automatic RIS worklist update and perform this procedure manually. Customers should disable the Measurement Field Display and Collimator Display by setting the display time for both setting to 0 sec (if editable). Customers should forward the safety notice to the new owner and inform Siemens of the new owner where possible. Customers with questions should call 610-219-2119 or 610-219-4834.For questions regarding this recall call 610-219-6300. |  | 10 | 10 | Nationwide Distribution - US including UT, MO, CO, MA, MN, MS, OH, SC, MI, and NY. | Terminated | February 24, 2015 | 175 | 18 | Not_Computer
Z-0007-2015 | 69094 | No | DIP | Diphenylhydantoin test system. | Toxicology | Toxicology | 510(k) | VITROS Chemistry Products PHYT (Phenytoin) Slides, REF/Product Code 829 8671, IVD. |  |  | August 20, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics identified a potential for biased results to be generated when using VITROS PHYT Slides, Lot 2613-0150-4913. Some positively biased results were identified during testing of a different VITROS PHYT Slide lot; the affected slides were not released for distribution. Since Lot 2613-0150-4913 was manufactured in the same event, there is the potential for slides from this lot to also be affected. | Software Manufacturing/Software Deployment | On 8/20/14, an URGENT PRODUCT CORRECTION NOTIFICATION Letter (CL2014-222, dated 8/20/14) was sent via FedEx overnight courier and/or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them to discontinue use of the affected lot(s) and to inform their Laboratory Medical Director if they suspect that a previously reported result may have been affected. One foreign affiliate was notified on 8/20/14. One distributor was notified on 8/20/14. The customers were advised to: - Immediately discontinue using and discard VITROS Chemistry Products PHYT slide Lot 2613-0150-4913. Review VITROS PHYT results obtained using Lot 2613-0150-4913 and discuss any concerns they may have regarding previously reported results with their Laboratory Medical Director or with the requesting physician. Complete the attached Confirmation of Receipt form no later than 8/28/14. Forward the information in this notification if they have distributed this product outside of their facility facility. |  | 349 | Domestic: 339 sales units; Foreign: 10 sales units | Distributed USA (nationwide) and in Singapore. | Terminated | April 11, 2017 | 965 | 349 | Not_Computer
Z-0006-2015 | 69024 | No | CIT | Aspartate amino transferase (AST/SGOT) test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ENVOY 500 AST Reagent Kit, reference Model/Catalog 55255, for Envoy 500. |  |  | August 11, 2014 | 2014 | ELITech Clinical Systems SAS | Some users of ENVOY 500 AST Reagent Kit, reference 55255, for Envoy 500 systems are observing that Quality Control (QC) fails before the product expiry is reached. | Process control | ELITech sent an Urgent Product Recall Notice letter dated August 12, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. To All Users,***Effective immediately, please discontinue using this product.***Actions to be taken:"Laboratories must not use this product and must destroy all remaining inventory of the product."If you have transferred any of this product to another location, please forward a copy of this bulletin to that location."All users receive ENVOY 500 AST REAGENT KIT, reference 55255, from distributors. To be certain everyone who is possibly affected by this is contacted, we are sending this communication to all registered users.As confirmation of receipt of this notice, please return the attached confirmation receipt to Technical Support by fax (401-642-9003) within 48 hours of receipt.On August 20, 2014:, customers were contacted with an interim solution.On September 3, 2014, the firm sent Technical Bulletin 5056 and Distributor Notice D057/Customer Notice C061 to all Envoy users notifying that the recalled Envoy AST Reagent (ref 55255) was being discontinued; the old Envoy AST Reagent (ref. 55250) was being reinstated.On September 23, 2014, the firm sent the URGENT: PRODUCT RECALL NOTICE "UPDATE", dated 19 September 2014, to all Envoy users stating that old (original) Envoy 500 AST Reagent Kit (ref 55250) will replace the recalled Envoy 500 AST Reagent Kit (ref 55255).As confirmation of receipt of this notice, please return the attached confirmation receipt to Technical Support by fax (401-642-9003) within 48 hours of receipt.Customers with questions about this recall, please call the Technical Support at 1-855-354-8324 (1-855-ELITECH). |  | 878 | 878 kits | Nationwide Distribution | Terminated | December 11, 2014 | 122 | 878 | Not_Computer
Z-0009-2015 | 69064 | No | EHD | Extraoral source x-ray system. | Dental | Radiology | 510(k) | eXpertDC intraoral x-ray, models: eXpert DC, NA, 75 Inch reach; eXpert DC, NA, 65 Inch reach; eXpert DC, CN, 55 Inch reach; eXpert DC, WW, 75 Inch reach; eXpert DC, WW, 65 Inch reach; eXpert DC, WW, 55 Inch reach; OEM Arm Assembly |  |  | August 12, 2014 | 2014 | Gendex Corp | Two (2) failures of the spring link components within the articulated arm assembly. The spring links broke during normal assembly process for the articulated arm during first cycle of the arm, prior to the 14lb tube head (full load) being installed onto the assembled articulated arm. | Nonconforming Material/Component | Gendex sent a Recall Notification Letter, dated August 12, 2014 and a second Communication Notification letter on September 22, 2014 to all affected customers. The letter identified the affected product, problem, and actions to be taken. For questions contact Gendex Customer Service at 1-800-323-8029. |  | 651 | 651 | Worldwide Distribution - US Nationwide and the countries of Europe, Canada, Australia, China, Russia, Taiwan, Singapore, and Thailand. | Terminated | March 24, 2017 | 955 | 651 | Not_Computer
Z-0010-2015 | 69322 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Precision 500D, Classical R&F; System. |  |  | September 05, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has become aware of a potential safety issue involving missing screws used for mounting the overhead Video Monitor to the tray assembly on a single monitor suspension. This issue affects the Precision 500D X-ray imaging systems. | Employee error | Consignees were sent on 9/5/2014 a GE Healthcare "Urgent Medical Device Letter" GE Ref:10893, dated September 5, 2014.The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology / Cardiology &amp; Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions in USA call 800-437-1171, for other countries contact your local GE Healthcare Service Representative. |  | 204 | 204 (199 US, 5 OUS). | Worldwide Distribution - US Nationwide - AL, AZ, AR, CA, CO, CT, FL, GA, HI, ID, IL, IN, IA, KS,KY, LA, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, OH, OR, PA, SC, TN, TX, UT, VA, WA, WV, WI and in the countries CANADA, GERMANY, UNITED KINGDOM. | Terminated | February 20, 2015 | 168 | 204 | Not_Computer
Z-0023-2015 | 69249 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Biomedix SELEC-3 I.V. Administration Set. Individually sealed into a TyvekÂż and LDPE/PET pouch. Secondary packaging includes 48 pouched units per box/case. STERILE |  |  | August 14, 2014 | 2014 | Biomedix, Inc. | Customer contacted Biomedix of complaint and returned the failed set along with 86 units of unopened stock (96 units originally shipped). Biomedix performed extensive investigation and found insufficient application of glue. Recall was initiated due to the rate of the failure. | Process control | Biomedix sent an PRODUCT RECALL NOTIFICATION letter dated August 14, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Please contact us with the quantity remaining in inventory or available to be returned to our facility. All units returned to Biomedix will be replaced with the same SELEC-3 item or a similar design at no charge.Please call (1-800-627-2765 ) for further questions. |  | 2928 | 2,928 | US Distribution including the states of OH, NC, and TX | Terminated | November 13, 2014 | 91 | 2928 | Not_Computer
Z-0011-2015 | 69092 | No | INM | Mechanical chair. | Physical Medicine | Physical Medicine | 510(K) Exempt | Nemschooff, Model Numbers - Serenity II:790-12, Serenity III: 791-12, Serenity III Exam: 791-12E, and Serenity IIS: 792-12. Designed to provide patient comfort during patient treatment sessions. The chair may include a manual control reclining mechanism. Customer configurable models are also available. The Serenity III Exam 791-12E was discontinued 11/1/2010. |  |  | January 18, 2011 | 2011 | Nemschoff Chairs, Inc. | Nemschoff Chairs, Inc. is recalling various model of the Serenity Treatment Chair because of potential elevated temperatures in the seat and back lumbar areas of the chair. | Device Design | The corrections have been performed via heat service kits being installed "on demand" to address heating element malfunctions starting on 1/18/2011. Consignee letter to all customers is pending HHE assessment. |  | 1868 | 1868 | Worldwide Distribution: USA (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, IL ,IN, KS, KY, MA, MD, MI, MN, MO, MN, MT, NC, ND,NE, BH, NJ NM, NV, NY, OH, OK, OR, PA, PE, SC, SD, TN, TX, UT, VA, WA, WI, WV, WY, and the country of CANADA. | Terminated | July 24, 2015 | 1648 | 1868 | Not_Computer
Z-0019-2015 | 67029 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | ArjoHuntleigh Contoura C1000 series beds are designed to accommodate patients weighing up to 990 lbs. (450 kg). Model C1000 is the basic bed. Model C1080 includes a weighing system and a power outlet for an ArjoHuntleigh approved mattress pump. |  |  | March 17, 2014 | 2014 | Arjo Hospital Equipment AB | There is a low but existing entrapment risk for a patient being treated on the Contoura 1000 and Contoura 1080 bariatric beds if they have a BMI &lt; 40. | Labeling design | ArjoHuntleigh sent an URGENT FIELD SAFETY NOTICE - IMMEDIATE ATTENTION REQUIRED letters dated March 17, 2014 to all affected cusomers. The letters included instructions for customers to: 1) ensure that the Contoura Bed Field Safety Notice brought to the attention of all caregivers and clinical staff and they are aware that the Contoura 1000 or 1080 medical bed is intended to be used as a bariatric medical bed, designed for the patients with BMI greater than 40; 2) ensure that all users of the Contoura 1000/1080 bed, applicable staff and caregivers are trained to ensure awareness of the revised Instructions for Use Manual and new bed label; 3) discard all previous versions of Contoura 1000/1080 Bed Instructions for Use Manuals and ensure that a copy of this Contoura Bed Field Safety Notice is placed in the enclosed revised Contoura 1000/1080 bed Instructions for Use Manual (746-173-AM-3 - 10/2013); 4) where applicable, include this requirement in subsequent training of applicable staff and caregivers; and, 5) complete and sign the enclosed Customer Response Form and return this form to the ArjoHuntleigh Regulatory Department. An ArjoHuntleigh Service Technician will contact the designated facility representative listed on the Customer Response Form to schedule service to apply the two labels shown above, free of charge. For question contact ArjoHuntleigh at 1-800-323-1245, ext. 57985. |  | 1695 | 1,695 beds (worldwide) | Worldwide Distribution - US: Nationwide in the states of (AZ, CA, CO, CT, GA, ID, IL, MA, MD, MN, NC, NE, OH, PA, RI, SD, TX, VA, WI); and the countries of Australia, Austria, Bahrain, Belgium, Canada, Denmark, France, Germany, Greece, Iceland, India, Ireland, Israel, Italy, Lebanon, Mexico, Netherlands, New Zealand, Norway, Saudi Arabia, Singapore, Spain, Sweden, Switzerland, United Arab Emirates | Open, Classified |  |  | 1695 | Not_Computer
Z-0021-2015 | 69280 | No | HTO | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Compress Face Reamer |  |  | September 17, 2014 | 2014 | Biomet, Inc. | Investigation determined that units supplied were missing the 4x21 degree helix angle. | Process control | On September 17, 2014, "URGENT MEDICAL DEVICE RECALL NOTICE" was sent to the consignee Biomet Global Supply Chain Center (GSCC) by email and FedEx. GSCC notified all of the customers within Europe, the Middle East, and Asia that received the affected product. The recall notification included a description of the reason for recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call (574) 372-1570. |  | 4 | 4 | International Distribution in the countries of Netherlands, Europe, the Middle East, and Asia. | Terminated | July 14, 2015 | 300 | 4 | Not_Computer
Z-0024-2015 | 69149 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension VistaÂż Mass CKMB Isoenzyme Calibrator (MMB CAL) an in vitro diagnostic product for the calibration of Creatine Kinase MB Isoenzyme (MMB) method on the Dimension VistaÂż System. |  |  | August 28, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has determined that Dimension Vista&reg; Mass CKMB Isoenzyme Calibrator (MMB CAL) shows drift during routine stability monitoring that exceeds our acceptance criteria for drift over the shelf life of this product. Siemens has observed drift from 3% to 12% at MMB concentrations across the assay range. Depending on quality control limits, this drift may not have been d | Process design | An Urgent Medical Device Recall notice, dated August 2014, was provided to all Dimension VistaÂż MMB CAL customers who received the affected lots to notify them of the issue, the potential risk to health, and actions to be taken by customers.Customers were instructed to discard lots 3JD045, 4BD064, 4CD012, and 4DD065, and recalibrate with lot 4FD073. |  | 6562 | 6562 | Worldwided distribution: US (nationwide) and countries of: Australia, Austria, Belgium, Canada, Denmark, Germany, Finland, France, Italy, Japan, Netherlands, New Zealand, Portugal, Saudi Arabia, Slovenia, South Korea, Spain, and Switzerland. | Terminated | March 30, 2015 | 214 | 6562 | Not_Computer
Z-0025-2015 | 69085 | No | LJS | Percutaneous, implanted, long-term intravascular catheter. | General Hospital | General Hospital | 510(k) | BioFlo PICC with ENDEXO Technology Maximal Barrier Nursing Kit, 3F-55 cm, UPN Product No. H965750231, REF/Catalog No. 75-023, STERILE --- |  |  | August 25, 2014 | 2014 | Navilyst Medical, Inc | NMI was notified by the manufacturer of the 10 ml BD PosiFlush SF Saline Flush Syringes, that it had initiated a recall on specific lots of the syringes as some syringe unit packages exhibited open seals which may impact package integrity and potentially affect product sterility. Based on the notification, NMI is recalling specific PICC Kits that contain the affected 10 ml BD PosiFlush SF Saline | Nonconforming Material/Component | BD Medical sent an Urgent Voluntary Medical Device Recall letter dated July 22, 2014, to all affected customers via Federal Express on August 25, 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to segregate and quarantine all affected product. Immediately discontinue shipment of the affected product. Complete the Recall Response Card form and fax it back to BD at 1-201-847-6990. If the affected product was further distributed cusotmers should identify their cusotmers and notify them immediately of the product recall. Customers with questions were instructed to contact BD Customer Service at 1-888-237-2762, option 3.For questions regarding this recall call 518-742-4571. |  | 20 | 20 units/pouches | Nationwide Distribution | Open, Classified |  |  | 825 | Not_Computer
Z-0038-2015 | 69104 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | RECLAIM Distal Reamer Extension |  |  | September 03, 2014 | 2014 | DePuy Orthopaedics, Inc. | Specific lots are being recalled due to the potential for the tabs to break, leaving fragments in the patient. | Device Design | On February 25, 2013, DePuy Orthopaedics, Inc. issued a device correction for the specific lots of the RECLAIM Distal Reamer Extension impacted by this recall. The letter stated the reason for the recall and swap out. A subsequent letter outlined a formal swap-out process of the affected reamer extensions as part of the device correction strategy. Newly designed devices are now available. Questions about device swap-out information may be directed to Kim Earle, Recall Coordinator, 574-371-4917; and Surgeon questions may be directed to DePuy's Scientific Information Office at 1-888-554-2482. |  | 875 | 875 units | Worldwide Distribution -- USA, including the states of ME, MD, PA, MA, FL, LA, IL, IA, OH, MN, TX, MT, WA, CA, HI, NV, AZ, KY, WI, CO, GA, MI, NC, VA, NY, TN, and IN; and, the countries of Canada, Australia, Austria, Belgium, Czech Republic, United Kingdom, France, Germany, Ireland, Israel, New Zealand, Norway, Poland, Slovenia, South Africa, Switzerland, and Chile. | Terminated | November 09, 2015 | 432 | 875 | Not_Computer
Z-0039-2015 | 69146 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | RENASYS EZ/ RENASYS EZ Plus 250 mL canister with Solidifier |  |  | July 18, 2014 | 2014 | Smith &amp; Nephew, Inc. | RENASYS EZ Canisters exhibiting visible evidence of deformation of the inlet port on the canister lid. | Device Design | Smith &amp; Nephew sent an Urgent Medical Device Correction Notice dated July 15, 2014 to all US customers via Federal Express delivery. The letter identified the affected product, problem and actions to be taken. The firm instructed customers to discard the affected product and replace with a new one from an unaffected lot. For questions contact 727-399-3785. |  | 8650 | 19,499 = (10,849 (Item 66800058) &amp; 8,650 (Item 66800913)) | Worldwide Distribution - US Nationwide in the states of FL, NC, CA, MO, MA, TX, NY, OH, CO, VA, IL, NJ, TN, NV, MN including Puerto Rico and countries ofCanada, Mexico, Dubai, Argentina, Australia, Colombia, Germany, Hong Kong, Jordon, Malaysia, Thailand, Tunisia, New Zealand, South Africa, Shanghai, Brazil, Saudi Arabia and Lebanon. | Terminated | August 05, 2015 | 383 | 55661 | Not_Computer
Z-0049-2015 | 69404 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | StarMedTec LightTrailÂż Reusable Fibers, 365 um; Material/Part Number: 6453 |  |  | September 30, 2014 | 2014 | Boston Scientific Corporation | Product is not cleared for use with lasers other than the Auriga XL system | No Marketing Application | Boston Scientific sent an " Urgent Medical Device Recall Removal Immediate Action Required" letter dated September 30, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The notification instructs the account to check their inventory, remove any affected product from their inventory, and contact Boston Scientific for further instructions on returns. The notification also requests the completion of a Reply Verification Tracking Form as a method of documenting the presence or absence of affected product in their inventory.We ask that you complete and return the Reply Verification Tracking form, included with this letter, according to the instructions on page 3, even if you no longer have inventory of the recalled lots. If you identify any of the affected lots within your inventory, please segregate it immediately and you will be contacted by Boston Scientific with additional instructions after you return your form. You will receive credit, as appropriate, for any affected product in your inventory. Fax to : Field Action Center at 1-866-213-1806.For further questions call (508) 382-9555 |  | 5 | 5 units | US Distribution including the states of NY and TX. | Terminated | November 26, 2014 | 57 | 45 | Not_Computer
Z-0042-2015 | 69377 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | McKesson Cardiology" Hemo intended for complete physiological/hemodynamic monitoring, clinical data acquisition, medical image and data processing, and analytical assessment. |  |  | August 27, 2014 | 2014 | Mckesson Information Solutions LLC | Software Error: The McKesson Cardiology Hemo calculation section incorrectly converts the Hemoglobin value before it is utilized in the applicable formula calculations. | Software design (manufacturing process) | On August 27, 2014, McKesson initiated a correction via a Field Safety Notice to all affected domestic customers. McKesson Israel will be responsible for contacting foreign consignees. |  | 272 | 272 devices | Worldwide distribution: US (nationwide) including states of: AR, CA, FL, DE, GA, HI, IL, IN, KS, LA, MA, MD, MN, MO, MS, MT, NJ, NV, OK, PA, SC, TNTX, , UT, WA; and countries of: Australia, Canada, Israel and the United Kingdom. | Terminated | May 18, 2016 | 630 | 272 | Computer
Z-0061-2015 | 69052 | No | MDM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Charnley PN Retract &amp; Handle, REF 962004000, Non-Sterile-DuPuy International Ltd. |  |  | August 27, 2014 | 2014 | DePuy Orthopaedics, Inc. | The small extraction peg of the Charnley Pin Retractor and Handle set is breaking from the pin. | Mixed-up of materials/components | The firm, Depuy Synthes, sent an "URGENT INFORMATION - RECALL NOTICE" dated August 27, 2014 to their consignees. The recall notification included a description of the reason for recall, affected product, clinical implications, and instructions for responding to the formal recall notification. The customers were instructed to cease using the affected components immediately, quarantine all affected components in a manner that ensures the affected components are not used; return affected devices immediately to your DePuy; if any affected product has been forwarded to another facility, contact that facility immediately to communicate recall and arrange return; maintain awareness and copy of this notice, and complete and return the Reconciliation forms via fax to 574-371-4939 or emailed to kearle@its.jnj.com. For questions about the device recall information provided, please contact Recall Coordinator, at 574-371-4917 (M-F; 8 a.m. - 5 p.m. EDT) or email Kearle@its.jnj.com.For product-related questions, please contact your local DePuy Orthopaedics, Inc. Sales Consultant. For clinical questions from hospitals/user facilities, please contact DePuy Orthopaedics, Inc. Scientific Information Office at 1-888-554-2482 (M-F; 8 a.m. - 5 p.m. EDT). |  | 65 | 65 units | Worldwide Distribution: US (Nationwide) including states of: AR, CA, CO, CT, IL, MA, NC, OH, PA, TN, TX, UT, VA, WA, WI; and Internationally to: AUSTRALIA, CHILE, UK, INDIA, IRELAND, JAPAN, NEW ZEALAND, RUSSIA, SINGAPORE, and SAUDI ARABIA. | Terminated | November 17, 2015 | 447 | 65 | Not_Computer
Z-0018-2015 | 69276 | fixative, formalin-containing | LDY | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Cardinal Health Prefilled Buffered 10% Formalin. |  |  | September 17, 2014 | 2014 | Richard-Allan Scientific Company | The affected lots could have NBF concentrations that are lower or higher than the desired specifications. | Mixed-up of materials/components | The firm issued an "URGENT MEDICAL DEVICE RECALL" letter to all affected customers. The notification identified the affected products, reason for voluntary recall, risk to health, actions to be taken, and instructions to complete the attached Acknowledgement Form. For questions, contact Sarah Rickert at 269-544-5628 or sarah.rickert@thermofisher.com. |  | 235 | 235 cases | Worldwide Distribution-USA (nationwide) including the states of PA, IL, CA, TX, GA, WA, MA, KY, OH, NY, VA, NC, IN, NJ, MN, UT, HI, FL, KS, TN, MO, MI, MD, LA, NM, DE, CO, WV, OR, WI, AZ, SC, IA, and NE and the country of Canada. | Completed |  |  | 1621 | Not_Computer
Z-0022-2015 | 69197 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare Integris Allura 15 -12 Fluoroscopy x ray system |  |  | August 08, 2014 | 2014 | Philips Medical Systems, Inc. | Philips had discovered through a field service testing where a device failed to comply with a Federal standard. | Radiation Control for Health and Safety Act | Philips Healthcare discovered through field service testing, that the device was in violation. Firm service rep. visited site on 7/01/14 and corrected unit.The field service engineer immediately reprogrammed the generator for the X-ray tube installed in the system and re-adapted the x-ray tube to bring the system back into compliance.At this time the Failure to comply has been corrected and the customer has been notified. This correction is to be free of charge to the customer.For further questions please call (978) 687-1501. |  | 1 | one system | US distribution to FL | Terminated | December 03, 2014 | 117 | 1 | Not_Computer
Z-0060-2015 | 69207 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare Carescape Patient Data Module |  |  | September 19, 2014 | 2014 | GE Healthcare, LLC | GE Healthcare has recently become aware of a potential safety issue to the ECG calculations following a disconnect/reconnect cycle with the Patient Data Module, when used with the Carescape Bx50 monitors.If the PDM is quickly disconnected and reconnected (a less than 10 second cycle) from a Carescape Bx50 monitor following a Patient Discharge or during active monitoring, the ECG waveform an | Software design | Consignees were sent on 9/19/2014 a GE Healthcare "Urgent Medical Device Correction" letter Ref# 336107 dated September 18, 2014.The letter was addressed to Director of Biomedical Engineering, Chief of Nursing and Healthcare Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions , Affected Product Details, Product Correction and Contact Information. For questions they can contact Technical Support at 1-800-558-70444 or your local Service Representative. |  | 1691517372 | 68,452 (39,505 units US, 24,711 units OUS, 4,236 units unidentified/unknown) | Worldwide Distribution: US (nationwide) including DC, PR, Guantanamo except DE, HI, ND, SD &amp; WY; and countries of: SOUTH AFRICA, KOREA, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TUNISIA, TRINIDAD TOBAGO, THAILAND, TANZANIA, TAIWAN, SWITZERLAND, SWEDEN, SPAIN, SLOVAKIA, SINGAPORE, SAUDI ARABIA, RUSSIA, QATAR, BRASIL, PORTUGAL, POLAND, PHILIPPINES, PERU, PANAMA, PARAGUAY, OMAN, NEW ZEALAND, NORWAY, NETHERLAND, NEPAL, MYANMAR, MEXICO, MOROCCO, MALDOVA, MAURITIUS, MALAYSIA, LITHUANIA, LEBANON, BAHRAIN, KUWAIT, REPUBLIC OF KOREA, JAPAN, JORDAN, JAMAICA, ITALY, ISRAEL, IRELAND, INDONESIA, INDIA, ICELAND, HUNGARY, HONG KONG, GUATEMALA, GREECE, GERMANY, FRANCE, FINLAND, ESTONIA, EGYPT,ECUADOR, DENMARK, CECH REPUBLIC, COSTA RICA, COLOMBIA,, CHINA, CANADA, BULGARIA, BELGIUM BANGLADESH, AUSTRIA, AUSTRALIA, and ARGENTINA. | Open, Classified |  |  | 1691517372 | Not_Computer
Z-0059-2015 | 69337 | No | OWQ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Biosense Webster's Reprocessed SOUNDSTARÂż eco Diagnostic Ultrasound Catheters, Catalog Nos: M-5723-115, Model Number BIO10438577. |  |  | September 24, 2014 | 2014 | STERILMED, INC. | Sterilmed, a division of Johnson &amp; Johnson Medical NV/SA (Sterilmed), has become aware of an issue affecting certain lots of Reprocessed SOUNDSTAR&reg; eco Diagnostic Ultrasound Catheters for which Sterilmed is the reprocessor and Biosense Webster is the distributor. Sterilmed determined that the units of these lots were incorrectly processed and as a result when the affected catheter is conne | Process control | Sterilmed sent an "Field Notification" letter dated September 22, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions requested on your part:1. Read the Description of the Problem section below carefully.2.Immediately identify and set aside all units from the affected lots of product listed above in a manner that ensures the affected product will not be used.3. Maintain a copy of this letter with the affected product.4. Review, complete, sign and return the attached Field Removal Certification Form in accordance with the instructions listed on the form.5. Arrange for return of any impacted units from the provided list of affected Reprocessed SOUNDSTARÂż eco Diagnostic Ultrasound Catheters that you may have in your inventory per the instructions on the Field Removal Certification Form.6. Pass on this notice to anyone in your facility that needs to be informed.7. If any of the affected Reprocessed SOUNDSTARÂż eco Diagnostic Ultrasound Catheters have been forwarded to another facility, contact that facility and arrange for the return.8. Maintain awareness of this notice until all affected product with lot numbers listed above has been returned to Biosense Webster.For questions related to this issue and product return please contact your Biosense Webster sales representative. For further questions please call (763) 488-3400. |  | 22 | 22 | US Distribution including the states of WA and CA. | Terminated | August 11, 2015 | 321 | 22 | Not_Computer
Z-0057-2015 | 69360 | No | HWQ | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Green Mamba Suture Passer. |  |  | September 26, 2014 | 2014 | Biomet, Inc. | An investigation identified that high level friction may exist between the needle and Mamba suture passer instrument, causing the needle to break during use. | Nonconforming Material/Component | On 9/26/2014, "URGENT MEDICAL DEVICE RECALL NOTICE" notifications were sent via courier to the affected distributors with instructions for returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, contact information, and instructions for responding to the formal recall notification. |  | 12 | 12 units | Nationwide Distribution-including the states of FL, NY, TX, VA, NV, MI, CA, IN, SD, GA, KY, NC, and AR. | Terminated | July 14, 2015 | 291 | 191 | Not_Computer
Z-0052-2015 | 69409 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | ZOLL Pro-padz Sterile Multi-Function Electrodes (MFE) |  |  | September 16, 2014 | 2014 | Bio-Detek, Inc. | Product sterility is not assured | Material/Component Contamination | Bio-Detek notiifed customers of the expanded recall on September 16, 2014 for Lot 1914 and September 22, 2014 for Lots 1614, 1814, 2114 via UPS overnight delivery. Customers were requested to return the suspect product(s) in their possession. A Device Correction Response Return Form was provided to each Customer. For questions call 1-800-348-9011 or 1-978-421-9460. |  | 237 | 237 cases (6/case) | Worldwide Distribution - US Nationwide and the countries of: Australia, Canada, Belgium, France, and Germany. | Completed |  |  | 489 | Not_Computer
Z-0063-2015 | 69270 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist One Vessel Sealer used in conjunction with the da Vinci Si Surgical System IS3000. |  |  | September 17, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive Surgical is initiating a voluntary correction relating to the labeling and software associated with the EndoWrist. One Vessel Sealer for the da Vinci Si (IS3000) Surgical System. | Device Design | Urgent medical device correction letter was sent to all consignees on 9/22/14 by trackable mail system.Letters identified the issue and provides information for users should a error code be displayed. |  | 1483 | 1,483 devices | Worldwide Distribution-USA (nationwide) and the countries of Australia, Austria, Belgium, China, Chile, Colombia, Czech Republic, Denmark, Finland, France, Germany, Greece, Italy, Japan, Netherlands, Norway, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | June 01, 2015 | 257 | 1483 | Computer
Z-0062-2015 | 69294 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | MaxLock Extreme Mod-Foot Concave Reamer; Individually packaged in a 3 x 4 bag and heat sealed with label on the outside. |  |  | September 19, 2014 | 2014 | Orthohelix Surgical Designs Inc | Incorrect sizing was being laser etched on the MaxLock Extreme Mod-Foot Concave Reamer | Device Design | An Urgent - Medical Device Recall customer notification, dated September 19, 2014, was sent out via email and FEDEx to all distributors with instructions on returning the product. The letter identified the affected product and the reason for the recall. Customers are to immediately cease distribution of the affected product, and complete and return the questionnaire. The letter also states that the recalled product will be replaced. Questions are to be directed to the Customer Service Department 1-888-494-7950. |  | 38 | 38 | USA Distribution -- in the states of WI, NM, TX, WA, MO, OH, UT, TN, MI, VA, NC, VA, and PA. | Terminated | October 31, 2016 | 773 | 38 | Not_Computer
Z-0064-2015 | 69173 | No | CHL | Blood gases (PCO2, PO2) and blood pH test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Omni Micro-electrode/reference electrode for cobas b221 analyzer. |  |  | September 12, 2014 | 2014 | Roche Diagnostics Operations, Inc. | The default references for normal values are inconsistent between cobas b 221 and other blood gas instruments. In addition, the reference ranges are different when comparing with corresponding ranges listed in the Instructions for Use for the analyzer and the reference literature source. | Labeling Change Control | The Urgent Medical Device Correction will be sent via UPS Ground (receipt required) to the consignee list. The Urgent Medical Device Correction will be sent via UPS Ground (receipt required) to the consignee list.Consignees are instructed as follows:" Check the current set of reference values on your cobas b 221 blood gas system and, if appropriate, change the reference ranges accordingly." If you have not established your own reference ranges, please follow the instructions in the Workaround section of the UMDC or contact the Roche Support Network Customer Support Center. |  | 492 | 492 devicees | Nationwide Distribution including Puerto Rico and Guam, and the states of AK, AL, AZ, CA, CO, CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, ME, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NY, OH, OK, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA and WI. | Open, Classified |  |  | 492 | Not_Computer
Z-0066-2015 | 68987 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow Select Kits (ASK) and Maximal Barrier Precautions Kits (CDC) with BD Medical Posiflush 10mL Saline Flush Syringe. |  |  | August 06, 2014 | 2014 | Arrow International Inc | The component supplier recalled their 0.9% Sodium Chloride Injection USP BD Posiflush SF Saline Flush Syringe 10mL because some unit packages may exhibit open seals which impact package integrity and potentially product sterility. | Packaging | Arrow International sent an Urgent Medical Device Recall Notification letter, dated August 6, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately discontinue use and quarantine any affected product. Customers were instructed to remove the BD pre-filled saline syringe from the exterior of the packaging. To return the BD syringes, customers were instructed to complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507, Attn: Customer Service. A customer service representative will contact the customer with a Return Good Authorization (RGA) Number and will provide instructions for the return of producet to Arrow International. All customers were asked to complete the enclosed Recall Acknowledgement Form and fax it to 1-855-419-8507. Attn: Customer Service. Customers with questions were instructed to contact their local sales representative or Customer Service at 1-866-246-6990. |  | 71255 | 71255 | Nationwide Distribution | Open, Classified |  |  | 71255 | Not_Computer
Z-0065-2015 | 69065 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Stryker T2 Tibia System Nail Insertion Sleeve, Elastic, Sterile, a) REF 1806-1406S, SPI 08-11 and b) REF 1806-1407S, SPI 08-13. |  |  | July 24, 2014 | 2014 | Stryker Howmedica Osteonics Corp. | Stryker Orthopaedics received a report indicating some of the Nail Insertion Sleeves had a different inner diameter than described on the package label. | Mixed-up of materials/components | Stryker Orthopaedics sent an "Urgent Medical Device Removal" letter dated July 23, 2014 and Business Reply Form to their Branches/Agencies via e-mail on July 23, 2014 and their Hospital Risk Management was sent Notification Letters and Product Accountability Forms via FedEx (with return receipt) by July 24, 2014. The recall letter informed the customer of the problem. The customer was asked to return all affected products available to Stryker Osteosynthesis, c/o Paul Jahnke, Stryker Orthopaedics, 325 Corporate Drive, Mahwah, NJ 07430, REF: PFA #2014-101 or Contact Stryker's customer service and refer to PFA #2014-101 for returning the product to Stryker. Customers were asked to complete and sign the enclosed Business Reply Form and fax a copy to: 1 (855) 251-3635 or email a copy to Aminah Crawford, Recall Coordinator at aminah.crawford@stryker.com. |  | 484 | 484 units | US Nationwide Distribution in the states of CA, FL, IN, MD, NC, NH, NY, OH, PA and WI. | Terminated | June 08, 2017 | 1050 | 484 | Not_Computer
Z-0069-2015 | 69179 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Expression MR200 MRI Patient Monitoring System |  |  | September 05, 2014 | 2014 | Invivo Corporation | Failure to produce the Non-Invasive Blood Pressure (NIBP) Measurement. | Process control | Philips Medical Systems sent an Urgent - Medical Device Correction letter dated July 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Our records indicate that you have an affected device. The following page provides additional instructions and actions that will be taken to address this problem.If you need any further information or support concerning this issue, please contact your local Philips representative: 877-468-4861 option 1, then option 2. |  | 24 | 24 units. | Worldwide Distribution - US including the states of CA, GA, MI, NJ, NV, and TX., and the countries of: Austria, Finland, Germany, Poland, Singapore, Sweden, Switzerland and United Kingdom. | Open, Classified |  |  | 24 | Not_Computer
Z-0070-2015 | 69368 | No | OMP | Powered suction pump. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | V.A.C. GranuFoam Silver Large Dressing is barrier to bacterial penetration and may help reduce infection. Featuring SensaT.R.A.C. Technology REF M8275099/10 STERILE Single Use only Latex Free Rx Only. |  |  | September 18, 2014 | 2014 | KCI USA, Inc. | A portion of the V.A.C. GranuFoam Silver Dressing M8275099/10 lot 50775977 was incorrectly manufactured with an incorrect expiry date on the sterile pouch. The correct expiry date of 2015-11 is on the case label, however the pouch label identifies the expiry date of 2016-11. | Error in labeling | KCI will send a customer letter to each consignee notifying them of the issue, and to return the affected units. |  | 1090 | 1090 pouches (109 boxes) | The affected product was distributed to Mexico, Turkey, and United Arab Emirates only. The affected product was not distributed to the US. | Terminated | July 24, 2015 | 309 | 1090 | Not_Computer
Z-0045-2015 | 69257 | Yes | DRO | External transcutaneous cardiac pacemaker (noninvasive). | Cardiovascular | Cardiovascular | 510(k) | Covidien Kendall Adult Multi-Function Defibrillation Electrodes |  |  | September 18, 2014 | 2014 | Covidien LLC | Covidien Medi-Trace Cadence and Kendall Defibrillation Electrodes Not Compatible for use with Philips model FR3 or FRx AEDs | Device Design | Covidien notified customers by letter via Federal Express on 9/18/14. The letter requests that they immediately verify that they are not stocking Covidien electrodes for use with Philips FR3 or FRx AED units. A customer reply acknowledgement form is provided to allow Covidien to track the effectiveness of the notification. Covidien is adding a label to all levels of packaging that alerts the user that the product is not compatible with the FR3 and FRx AED units. Distribution partners notified to address product in the distribution chain that is available to be relabeled. For questions customers should call (800) 962-9888, option 8, then extension 2500.For questions regarding this recall call 508-261-8000. |  | 31590 | 31,590 pairs | Worldwide Distribtion-USA (nationwide) including Puerto Rico and the countries of Canada Greece Israel Belgium Finland France Germany Great Britain Ireland Italy Norway Poland Portugal Spain Sweden Switzerland Brazil Chile Philippines Taiwan Australia New Zealand. | Terminated | May 24, 2017 | 979 | 644460 | Not_Computer
Z-0068-2015 | 69295 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | eFilm Workstation 4.0 and 4.0.1. eFilm Workstation is one of the components of a PACS (Picture Archiving and Communications System). |  |  | August 28, 2014 | 2014 | Merge Healthcare, Inc. | There is an issue related to eFilm Workstation versions 4.0 and 4.0.1 when having multiple studies open and utilizing the thumbnail panel to select multiple series from multiple studies may lead to the selection of an incorrect study. | Device Design | Merge sent an Important Product Information Notice letter, dated August 22, 2014 to consignees. The letter stated the issue, identified affected product, provided instructions, and told of a product update expected to be available by September 15, 2014. Customers with questions should contact Customer Service at 877-741-5369 |  | 2630 | 2,630 | Worldwide Distribution - US Nationwide and the countries of : Argentina, Australia, Canada, Chile, France, Germany, Hungary, Iceland, Italy, Lithuania, Mexico, New Zealand, Romania, Russia, South Africa, Sweden, and Yemen. | Terminated | June 01, 2017 | 1008 | 2630 | Computer
Z-0067-2015 | 69357 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | AC Powered adjustable bed frame. |  |  | September 15, 2014 | 2014 | Primus Medical LLC | Firm received several complaints that the head deck actuator bracket had separated from the bed frame. | Device Design | Primus Medical will send a letter notifying all customers with PCB900 beds at their facility of the potential head bracket separation. |  | 1726 | 1726 | Worldwide Distribution-USA (nationwide) including the states of CA, DE, FL, GA, IL, IN, KS, KY, MA, MN,MS, NC, NM, NV, NY, OH, PA, TN, TX, and VA, and the country of Canada. | Terminated | August 10, 2017 | 1060 | 1726 | Not_Computer
Z-0074-2015 | 69283 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist Stapler 45 instrument for the da Vinci Si System |  |  | September 18, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive is issuing an urgent product notice to all users of the EndoWrist Stapler 45 to discontinue use due to reports of inability to remove the stapler from tissue. | Device Design | Intuitive Surgical sent an Urgent Product Notice letter dated September 19, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to:-Immediately stop using the EndoWrist Stapler 45 instrument.-Quarantine the EndoWrist Stapler 45 instruments within your hospital. Keep the instruments quarantined until Intuitive Surgical sends you a subsequent communication with instructions regarding the quarantined product. -Ensure all affected personnel are fully informed of this notice. Forward this notice to your Risk Manager, OR Director, Purchasing Manager, Biomedical Engineering staff, and members of your medical staff who perform da Vinci Si surgery procedures using the EndoWrist Stapler 45 instrument.-Please retain a copy of this notice for your records with your da Vinci Si User Manual.If Âżyou Âżneed Âżfurther Âżinformation Âżor Âżsupport Âżconcerning Âżthis ÂżProduct ÂżNotice, Âżplease contact your Clinical Sales Representative or Intuitive Surgical Customer Service at 800-876-1310 Option Âż3 Âż(6 Âżam Âżto Âż5 Âżpm ÂżPST). Âż |  | 1034 | 1,034 Devices | US Nationwide Distribution. | Terminated | March 03, 2015 | 166 | 1034 | Not_Computer
Z-0076-2015 | 69290 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Multi-Lumen Central Venous Catheterization |  |  | September 18, 2014 | 2014 | Arrow International Inc | Arrow International, Inc. has initiated a voluntary recall for product code AK-12703, lot number RF3039028 because the label on the outer corrugate is labeled with the incorrect expiration date. All kits packaged within the corrugate are correctly labeled. There is no risk to the patient as the kits within the corrugate are labeled with the correctexpiration date. | Labeling False and Misleading | Arrow International sent an Medical Device Advisory Notifications letter dated September 18, 2014, sent to direct accounts to notify them about the product, problem, and actions to be taken.Arrow International, Inc. ( Arrow ) is notifying customers that the label on the outer corrugate for the above mentioned product code and lot number is labeled with the incorrect expiration date. All kits packaged within thecorrugate are correctly labeled. There is no risk to the patient as the kits within the corrugate are labeled with the correct expiration date.Arrow International, Inc. is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 65 | 65 | US Distribution including the states of MD, NC , LA , IN , OH, KY, CA , PA , NY and NV. | Terminated | December 02, 2014 | 75 | 65 | Not_Computer
Z-0077-2015 | 69234 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 4600 Chemistry System (VITROS 5,1 FS System family member), System Product Code 6802445. |  |  | September 08, 2014 | 2014 | Ortho-Clinical Diagnostics | Ortho Clinical Diagnostics (OCD) identified an anomaly with Software Version 3.0 and below on the VITROS 4600 Chemistry Systems and VITROS Integrated Systems. Internal testing confirmed that when using calibrator barcode labels supplied with VITROS Chemistry Products Calibrator Kits 1, 2, 3, 4, 6, or 9, an unexpected assay calibration may occur if assay targets are unassigned (i.e., hidden). | Software design | On 9/08/14, URGENT PRODUCT CORRECTION NOTIFICATION Letter (CL2014-232, dated 9/8/14) and Confirmation of Receipt Form were sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government only) to inform them of the issue and advise them of the required actions. Foreign affiliates were informed of the issue via e-mail on 9/8/14 and instructed to notify their consignees of the issue and the required actions. |  | 232 | USA: 97; Foreign: 232 | Worldwide Distribution-USA (nationwide) including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 1177 | Computer
Z-0075-2015 | 69340 | No | HWT | Template for clinical use. | Orthopedic | Orthopedic | 510(K) Exempt | Discovery Elbow Humeral Trial |  |  | September 24, 2014 | 2014 | Biomet, Inc. | An investigation identified that the product was incorrectly manufactured as a left humeral trial but is etched as a right humeral trial. | Process control | Biomet sent an "URGENT MEDICAL DEVICE RECALL NOTICE" letters dated September 24, 2014, via FedEx overnight to affected customers. The notification included a description of the reason for recall, instructions for responding to the formal recall notification, and affected product, actions to be taken. For questions call (574) 372-1570. |  | 6 | 6 | Worldwide Distribution - US Nationwide in the state of SD, OH, CT and the country of: The Netherlands and India. | Terminated | July 14, 2015 | 293 | 6 | Not_Computer
Z-0071-2015 | 69333 | No | DKZ | Amphetamine test system. | Toxicology | Toxicology | 510(k) | ArcPoint Labs 5 Panel Dip Drug Screen, Part No. APD-5M. |  |  | September 23, 2014 | 2014 | Ameditech Inc | Ameditech is recalling the ArcPoint Labs Dip Drug Screen Dip Card because the positive interpretation illustration was incorrect on the cap labels as well as on the card cap graphic in the package insert. | Labeling Change Control | A customer notification letter dated 9/23/14 was sent to all customers who purchased the ArcPoint Labs Dip Drug Screen Dip Card. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the Verification Form. Customers with questions are instructed to contact Ameditech at (858) 805-8714 or email at amd.complaint@alere.com. |  | 400 | 400 kits | Distributed in AL. | Terminated | February 27, 2015 | 157 | 800 | Not_Computer
Z-0079-2015 | 63897 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Medium 2-Hooks Spreader Bar for use with Loop Slings, sizes small, medium and large on the Maxi Sky 2 Ceiling Lift; Model 700-19425. |  |  | January 21, 2013 | 2013 | Arjo, Inc. dba ArjoHuntleigh | The upper portion of the hooks on the spreader bar are too flexible for holding the loops of the sling. The loop of the sling could potentially come out of the hooks if positioned in the upper portion of the hooks. | Device Design | URGENT FIELD CORRECTION RECALL letters dated January 21, 2013 were sent to all direct accounts (customers). The letters identified the potential safety related issue; included instructions on how to use the Medium Two Hooks Spreader Bars, and to complete and return the enclosed Customer Response form. Once the Customer Response form has been returned customers will be contacted by an ArjoHuntleigh service representative to make arrangements for the replacement of the existing spreader bars with spreader bars of a different design. Customers with any questions about this recall can contact ArjoHuntleigh by telephone at 800-323-1245, extension 57985, or by e-mail at usa.quality@arjohuntleigh.com. |  | 363 | 363 spreader bars | USA Distribution - AR, GA, MI, PA, and the District of Columbia. | Terminated | October 27, 2015 | 1009 | 363 | Not_Computer
Z-0081-2015 | 69228 | No | JIL | Urinary glucose (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | External power supply adaptors that accompany Clinitek Status+ |  |  | September 19, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | The external power supply for the Clintek Status analyzer, which is provided separately, is damaged. It can result in an electric shock to the User. | Process control | Siemens sent an Urgent Field Safety Notice dated September 2014, to all by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on September 19, 2014 both in the United States and elsewhere for communication with affected customers. This notice informs customers of the issue with the power supply and the potentially affected products that may contain a damaged power supply. Consignees are asked to complete a Field Correction Effectiveness Check form and return in to Siemens Healthcare Diagnostics via Fax to (312) 275-7795. All affected domestic customers were sent a hard copy of the Urgent Field Safety Notice via Federal Express on 9/23/14. All affected customers outside the US will be provided a copy of an Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each countries local regulations and procedures.Customers with questions were instructed to contact their local Siemens technical support representative.For questions regarding this recall call 877-229-3711. |  | 4200 | approx. 4200 | Worldwide Distribution - USA (nationwide) including Puerto Rico and Internationally Canada, Mexico, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Kuwait, Latvia, Malaysia, Mexico, Morocco, Netherlands, New Zealand, Norway, Pakistan, Poland, Portugal, Qatar, Republic Korea, Romania, Russian Fed., Saudi Arabia, Senegal, Singapore, Slovakia, Slovenia, South Africa, Sweden, Switzerland, Tadjikistan, Taiwan, Turkey, Turkmenistan, Unit.Arab Emir., United Kingdom, Vietnam. | Terminated | July 20, 2015 | 304 | 4200 | Not_Computer
Z-0080-2015 | 69253 | Yes | KDJ | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 15 L Cycler Drainage Bag |  |  | October 17, 2013 | 2013 | Baxter Healthcare Corp | The large tube clamp used to close the drain tube on the 15 L Cycler Drainage Bag may not fully close the drain tube and may result in leakage. | Component change control | Baxter sent an Urgent Product Recall letter dated May 30, 2014 to all affected customers. The letter identified the affected product, problem and actions to be taken. The recalling firm is instructing their customers to remove all affected products from their facility and quarantine the product subject to the recall. Customers are instructed to return the affected product to Baxter and receive a credit or replacement. For questions contact the Center for One Baxter at 1-800-422-9837. |  | 3085080 | 3,085,080 | Worldwide Distribution - US Nationwide (including Puerto Rico) and countries of, Canada, Australia, New Zealand, Thailand, Malaysia, Singapore, Korea, Japan, Philippines, Columbia, Ecuador, Austria, Bahran, Belgium, Bosnia &amp; Herzegovina, Bulgaria, Czech Republic, Denmark, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Hungary, Ireland, Isreal, Italy, Jordan, Kazakhstan, Latvia, Luxembourg, Malta, Norway, Oman, Poland, Portugal, Qatar, Russia, Saudi Arabia, Slovakia, South Africa, Spain (including Canary Islands), Sweden, Switzerland, The Netherlands, Turkey and United Arad Emirates. | Terminated | January 04, 2016 | 809 | 3085080 | Not_Computer
Z-0086-2015 | 69371 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Siemens Mammomat Inspiration mammography systems. |  |  | September 18, 2014 | 2014 | Siemens Medical Solutions USA, Inc | The stereo biopsy devices for Mammomat Inspiration mammography systems might have integrated a safety switch which causes a failure of the functionality. The pin implemented in the safety switch may not put enough pressure on the safety circuit to prohibit movement. The needle positioning device may move even with the safety switch being set sideward. If this occurs with the needle being already i | Process control | A safety advisory notice, dated September 18, 2014, was sent to end users that identified the product, problem, and actions to be taken. |  | 3 | 3 | US Distribution in states of: NY, VA, and MA. | Terminated | February 24, 2015 | 159 | 3 | Not_Computer
Z-0088-2015 | 69291 | Yes | LWP |  |  | Cardiovascular | PMA | Boston Scientific TELIGEN, Implantable Cardiac Defibrillator models E102, E110, the following models are not available in the US: models E103, E111, F102, F103, F110, F111. Sterilized using ethylene oxide. |  |  | September 17, 2014 | 2014 | Boston Scientific CRM Corp | Boston Scientific identified a subset of COGNISTM CRT-Ds and TELIGENTM ICDs that had experienced an increased rate of premature battery depletion due to compromised performance of a low voltage (LV) capacitor. Boston Scientific updated their August 2013 Physician communication to include additional devices that may experience premature battery depletion due to compromised performance of a low volt | Component design/selection | Boston Scientific sent an Important Medical Device Information letter to physicians beginning September 17, 2014. The letter provided a summary of the issue, described clinical implications and expected device performance, and provided patient monitoring recommendations. Boston Scientific is introducing upgraded program software (Model 2868, version 3.04) that enhances the effectiveness of the Safety Architecture tools. Questions should be addressed to Boston Scientific Representative or Technical Services 1-800-227-3422 or tech.services@bsci.com. |  | 31144 | 31,144 devices total | Worldwide Distribution - US Nationwide including American Samoa, Guam, Puerto Rico, and U.S. - Virgin Islands. OUS countries to include: Argentina, Aruba, Australia, Austria, Bahamas, Barbados, Belgium, Bermuda, Brazil, Canada, Cayman Islands, Chile, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Finland, France, French Polynesia, Germany, Greece, Guadeloupe, Haiti, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Korea, Kuwait, Latvia, Lebanon, Luxembourg, Malaysia, Mexico, Monaco, Mozambique, Netherlands, Netherlands Antilles, New Caledonia, New Zealand, North Yemen, Norway, Panama, Peru, Philippines, Poland, Portugal, Qatar, Reunion, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syrian Arab Republic, Taiwan R O C, Thailand, Trinidad and Tobago, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela. | Terminated | February 20, 2015 | 156 | 31144 | Computer
Z-0092-2015 | 69432 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | DePuy Synthes Radiolucent Retractor, part number 387.580, lot number 3620715 |  |  | October 02, 2014 | 2014 | Synthes, Inc. | Particle residue on the instrument from adhesive tape which was used to bind the device during transport. There may be a risk of tissue reaction if a patient is exposed to latex particles from the adhesive tape, or a systemic reaction such as anaphylaxis in patients with latex allergy. | Material/Component Contamination | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated October 2, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Please take the following actions:If you DO have any of the identified devices, please take the following steps:Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number.Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided.Send a copy of the completed Verification Section to DePuy Synthes by:Fax: (610) 430-7083 or " Scan/email: Fieldaction@synthes.comNote: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you DO NOT have the identified product, please take the following steps:Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information.Return the documents to DePuy Synthes by Fax: (610) 430-7083 orScan/email: Fieldaction@sythes.comNote: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 3 of the notification.If you have any questions, please call 610-719-5450, or contact your DePuy Synthes sales consultant. Thank you for your attention to this issue. |  | 11 | 11 | US Distribution including the states of PA, NC, TN, NY, WI, WA and MI. | Terminated | December 18, 2015 | 442 | 11 | Not_Computer
Z-0089-2015 | 69262 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Codman EDS 3 Drainage System with Ventricular Catheter (External Drainage System) |  |  | September 15, 2014 | 2014 | Codman &amp; Shurtleff, Inc. | tubing within the system that drains CSF may leak or disconnect from the joints and may result in over- or under-drainage of CSF from the ventricular system or introduction of air into the ventricular system (pneumocephalus) | Device Design | Codman Neuro sent an Urgent Notice dated September 12, 2014, to all affected customers via overnight mailing and OUS customers notified via email the week of September 15, 2014.Accounts were requested to isolate all inventory of the affected product and return to Stericycle at the following address:Stericycle 2670 Executive Drive, Suite A Attn: Event #8894 Indianapolis, IN 46241For patients currently being managed with the EDS 3 System, the system should be replaced immediately. In the case where no substitute drainage system is immediately available, the EDS 3 System may continue to be used until an alternative product can be obtained.If you have any questions or concerns regarding this notification, please contact your local Codman Neuro Representative or Scientific and Medical Affairs at SciMedAffairs@its.jnj.com or (866) 685-7325.For questions regarding this recall call 508-880-8000. |  | 66120 | 66,120 units | Worldwide Distribution - USA (nationwide) and Internationally to:United Arab EmiratesArgentinaAustriaAustraliaArubaBelgiumBahrainBoliviaBrazilBahamasCanadaSwitzerlandChilePeople's Republic of ChinaColumbiaCosta RicaCyprusCzech RepublicGermanyAlgeriaEstoniaEgyptSpainFinlandFranceGabonGreat BritainFrench GuianaGreeceGuatemalaHong KongCroatiaHungaryIrelandIsraelIndiaIranItalyJordanKenyaKuwaitKazakhstanLebanonLithuaniaLuxembourgLatviaMauritiusMexicoMalaysiaMozambiqueNew CaledoniaNigeriaNetherlandsNorwayNew ZealandOmanPanamaPeruPolandPalestinePortugalReunionRomaniaSerbiaRussian FederationRwandaSaudi ArabiaSwedenSingaporeSloveniaSlovakiaTunisiaTurkeyTrinadad And TobagoTaiwanUkraineUruguayVenezuelaMayotteSouth Africa | Terminated | August 11, 2017 | 1061 | 160823 | Not_Computer
Z-0094-2015 | 69301 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Continuum Acetabular System Trabecular Metal Shell with Multi Holes, Porous 68mm, Prosthesis, Hip, Semi-Constrained, Metal/Polymer |  |  | October 01, 2014 | 2014 | Zimmer, Inc. | The affected products are missing polar boss threads. | Equipment maintenance | Zimmer sent an "URGENT MEDICAL DEVICE RECALL- LOT SPECIFIC" notifications dated October 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Your Responsibilities1. Review the notification and ensure affected personnel are aware of the contents.2. If you find any product from the affected lot , quarantine the product and notify your Zimmer sales representative.3. Your Zimmer sales representative will remove the recalled product from your facility.4. For patients that previously had this product implanted, it is recommended that you continue your normal post operative follow up routine.5. If after reviewing this notification you have further questions or concerns please call the customer call center at1-800-348-2759 between 8:00 am and 5:00pm EST. |  | 19 | 19 | Worldwide Distribution: US (nationwide) in the states of : IA, CA, NC, AL, WA, FL, TN, MA, AZ, TX, UT, and NV; and Internationally to Japan. | Terminated | March 02, 2015 | 152 | 19 | Not_Computer
Z-0093-2015 | 69380 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | McKesson Cardiology Hemo. |  |  | June 25, 2014 | 2014 | Mckesson Information Solutions LLC | McKesson Cardiology Hemo front-end parameter boxes must undergo annualmaintenance testing to establish the proper functioning of the parameter boxes,and to identify in a timely manner whether a parameter box is out of calibration.McKesson has identified that its communication to customers regarding therequired maintenance tests was insufficient. While each unit has an attachedsticker i | Labeling design | Consignees were notified via Field Safety Notice on 06/25/2014. |  | 625 | 625 devices | Worldwide Distribution-USA (nationwide) and the countries of Australia, Botswana, Canada, Germany, New Zealand, Netherlands, Pakistan, Portugal, South Africa, Turkey and The United Kingdom. | Open, Classified |  |  | 625 | Not_Computer
Z-0098-2015 | 69319 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | DeRoyal(R) SafeLiner Suction Canister Kit, SEMI-RIGID, 1000cc, REF 71-6518, NON-STERILE, Rx Only, |  |  | August 22, 2014 | 2014 | DeRoyal Industries Inc | These lots of suction canisters and lids have an insufficient seal under low or intermittent suction condition. | Process control | DeRoyal sent an Urgent Voluntary Recall letter dated August 22, 2014, to all affected consignees by overnight letter or by email. DeRoyal was to follow up with non-responding accounts before October 1, 2014. Customers were instructed to check their inventory for the affected product. If they have affected product, customers were instructed to destory both the liner and the lid and dispose of it according to their facility's procedures. Customers were instructed to return the included Notice of Destruction to recalls@deroyal.com even if they no longer have affected product. Customers with questions were instructed to contact their DeRoyal Sales Representative or DeRoyal Customer Service at 1-800-251-9864.For questions regarding this recall call 865-362-2334. |  | 5112 | 5112 units | Nationwide Distribution including MI, OH, IN, FL, WV, VA, MI, NC, and NY | Terminated | April 15, 2015 | 236 | 19012 | Not_Computer
Z-0105-2015 | 69348 | No | HGM | Perinatal monitoring system and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Philips Avalon Monitors with software revision J.30.58: |  |  | October 03, 2014 | 2014 | Philips Medical Systems, Inc. | Philips Intellivue and Avalon Fetal Monitors in time-synchronized automatic/sequence mode, the NBP automatic measurement series is stopped | Software Design Change | Philips Healthcare issued on 10/3/14 the Urgent Medical Device Correction notification/Field Safety Notice. The letter informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.The correction will consist of a software upgrade free of charge.A Philips Healthcare representative will contact customers with affected devices to arrange for service. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice:During the interim period until the SW is upgraded please make sure, if the NBP repetition interval needs to be increased (e.g. from 5 to 15 minutes), to restart the automatic NBP measurement.Please review this information with all members of your staff who need to be aware of the contents of this communication. It is important to understand the implications of this communication. |  | 118 | 118 units | Worldwide Distribution-USA (nationwide) and the countries of AUSTRALIA AUSTRIA BAHRAIN BELGIUM CHINA COLOMBIA CZECH REPUBLIC DENMARK FINLAND FRANCE GERMANY HONG KONG INDIA IRELAND ISRAEL ITALY JAPAN KAZAKHSTAN NETHERLANDS NEW ZEALAND NICARAGUA NIGERIA NORWAY OMAN POLAND ROMANIA RUSSIA SINGAPORE SLOVENIA SOUTH AFRICA SOUTH KOREA SPAIN SWEDEN SWITZERLAND TAIWAN THAILAND TURKEY UKRAINE and UNITED KINGDOM. | Open, Classified |  |  | 7705 | Computer
Z-0106-2015 | 69436 | No | GZL | Depth electrode. | Neurology | Neurology | 510(k) | Box Label: 3 mm Cueva Electrode Kit, Sterile EO, Rx Only, REF. CNE1-3KDINX . |  |  | October 08, 2014 | 2014 | Ad-Tech Medical Instrument Corporation | The recall has been initiated due to concerns that the applicator wand may malfunction potentially causing the inability to monitor the VII cranial nerve. . | Process control | Consignees were sent on 10/8/2014 an AD-TECH "Medical Device Recall" letter dated October 6, 2014. The letter described the Reason for the Voluntary Recall, Risk to Health, Action to be taken by the Customer, Product and Distribution Information and Other Information. Requested consignees to return the "Acknowledgement and Receipt Form" to their Ad-Tech Clinical Specialists Fax 262-634-5668, Telephone 262-634-1555, customersupport@adtechmedical.com. For additional information they can use the same number. |  | 13 | 13 kits (26 electrodes, 26 applicator wands, 26 leadwires). | Worldwide Distribution: US (nationwide) in the states of: CO, FL, LA, NJ, OH, WA, and WI; and internationally to: Finland. | Terminated | April 08, 2015 | 182 | 13 | Not_Computer
Z-0099-2015 | 66541 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | CRN Crono Syringe CRN20 with 19G x 1 "Steel Needle" |  |  | October 18, 2013 | 2013 | Intra Pump Infusion Systems | Product has the potential to have a clear debris on the tip of the needle. | Device Design | Cane S.p.A. Medical Technology sent an Urgent: Correction and Removal Action Required letter dated October 21, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken:a. Stop using the 19G x 11/2" 1.10 x 40mm steelneedle included in the CRONO Syringe blister pack. B. Safely discard the 19G x 11/2" 1.10 x 40mm steel needle into a Sharps container.c. Use a TRANSFER DEVICE (e.g.,Needleless adapter with female-to-female luer connector) or alternative device to transfer medications into the CRONO syringe.d.For intravenous infusions,we are recommending that a filter also be used during infusions.e. Complete the required 'Acknowledgment and Receipt Form'and fax back to our Distributor;Intra Pump Infusion Systems at 630-845-2768.Contact Information for our US Distributor: INTRA PUMP INFUSION SYSTEMS " Monday through Friday 8:00am to 5:00pm CST,TollFree:1-866-211-7867 " Direct:630-845-7867 " Fax:630-845-2768 www.lntrapump.com |  | 25440 | 25,440 pieces | Worldwide Distribution: US (nationwide) in states of CA, PA, MA, IL, OH and NY; and Internationally to Canada. | Terminated | February 01, 2016 | 836 | 25440 | Not_Computer
Z-0008-2015 | 69313 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Esprit V1000 and V200 Ventilators, Model No. V1000 and V200; and the 3rd Generation Power Supply Repair Part (RP)Kits |  |  | September 17, 2014 | 2014 | Respironics California Inc | A failure of a specific component may prevent the ventilators from operating on AC power or transitioning back to AC power when operating on battery. Additionally, if a battery is not present or is depleted the ventilators will not operate. This failure will cause the loss of ventilator support, which may result in hypercarbia or hypoxemia. | Unknown/Undetermined by firm | Philips Healthcare sent an Urgent - Field Safety Notice dated September 17, 2014, to all customers who purchased the Esprit V1000 and V200 Ventilators to inform them that Respironics is recalling Esprit V1000 and V200 Ventilators installed with 3rd Generation Power Supplies and also 3rd Generation Power Supply Repair Part Kits. The letter identifies the affected devices, informs the customers of the problems, hazard involved, and the actions to be taken. Customers were reminded to be advised of the warnings provided. A Philips Respironics representative will contact customers to schedule a no-cost replacement for further information or support.Customers with questions were instructed to contact Philips Healthcare at 1-800-722-9377.For questions regarding this recall call 760-918-7300. |  | 2206 | 2,206 units total (665 units in US) | Worldwide Distribution - USA, Argentina, China, Egypt, Ethiopia, Honduras, Hong Kong, India, Indonesia, Iraq, Israel, Italy, Jordan, Libyan Arab Yamahiriya, Malaysia, Mali, Mexico, Myanmar (Burma). Nepal, Nigeria. Pakistan, Peru, Philippines, Serbia, Singapore, South Africa, Sri Lanka, United Republic of Tanzania, Thailand, Vietnam, Zimbabwe, Republic of Korea, and Netherlands. | Open, Classified |  |  | 2206 | Not_Computer
Z-0103-2015 | 69292 | No | LOM |  |  | Microbiology | PMA | ADVIA CentaurÂżÂż Systems HBc Total (HBcT) Kit. For In Vitro Diagnostic Use. |  |  | September 15, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Increased imprecision and positive bias is observed with patient samples and negative controls and can result in an increase in initial reactive results. | Unknown/Undetermined by firm | Siemens Healthcare issued An Urgent Medical Device Recall (UMDR) sent to all affected Siemens Healthcare Diagnostics customers in the United States on September 15, 2014. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on September 15, 2014. These notices inform the customer of the imprecision and positive bias and a potential increase in initial reactive results associated with ADVIA Centaur Systems HBc Total Lot 100064. Customers were instructed to discontinue use of the affected lots and dispose of the product according to local regulations or procedures. The notices should be reviewed with the facility's laboratory or medical director.Each customer will be required to return a completed confirmation fax-back to Siemens. |  | 5434 | 5434 | Worldwide Distribution-USA (nationwide) and the countries of Austria Bahrain Belgium Croatia Cyprus Czech Republic Denmark Estonia Finland France Germany Great Britain Greece Hungary Iraq Ireland Italy Jordan Latvia Lebanon Lithuania Luxembourg Netherlands Norway Poland Portugal Romania Slovakia Spain Sweden Switzerland Turkey Unit.Arab Emir and Yemen. | Terminated | December 28, 2015 | 469 | 5434 | Not_Computer
Z-0107-2015 | 69260 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mindray V21, Size: 515 mm X 335 mm X 685 mm, N.W.: 8 kg, G.W.: 11 kg, Qty:1 |  |  | July 31, 2014 | 2014 | Mindray DS USA, Inc. dba Mindray North America | Mindray DS USA Inc. initiated a voluntary field corrective action for V-Series Patient Monitor because there is an issue with the V-Series Drug Calculator function. | Software design | Mindray DS USA, Inc. sent a Recall Correction Action Letter to their customers via certified mail with return receipt on July 31, 2014. |  | 2247 | 2,247 units | Worldwide Distribution: US (nationwide) and internationally to: Austalia, Canada, Columbia, Germany, India, Indonesia, Ireland, Israel, Kuwait, Malaysia, New Zealand, Pakistan, Russian Federation, Saudi Arabia, South Africa, Sri Lanka, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, and Venezula. | Open, Classified |  |  | 2247 | Not_Computer
Z-0100-2015 | 69226 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | GYRUS ACMI, REF 3345 -Everest BiCOAG MOLly Forceps 5mm/45cm |  |  | August 27, 2014 | 2014 | Gyrus Medical, Inc | Due to an anomaly in the packaging process, it is possible that a defective seal could be present that may allow a breach of the package's sterile barrier and may compromise the sterility of the product. | Packaging process control | Olympus sent a Medical Device Recall letter dated August 26, 2014, via FEDEX to all affected consignees. The letter described the issue and the product involved in the recall. It also described the action that the consignee needs to take. Consignees were instructed to cease further use of the device, quarantine it and return it to the manufacturer.Consigness were instructed to complete and return by fax (484-896-7128) the questionnaire enclosed. If the product was further distributed consignees should notify their customers of the product recall. For additional information they can call 484-896-5688.For questions regarding this recall call 484-896-5688. |  | 5560 | 5520 US, 40 OUS | Worldwide Distribution - USA including US: AL, AZ, CA, FL, GA, HI, IL, IN, IA, KY, MA,MI,MS, NJ, NY, NC, OH, PA, SC, TN, UT, and WA.Internationally to ITALY, SPAIN, UNITED KINGDOM, and CHINA. | Terminated | September 15, 2015 | 384 | 5560 | Not_Computer
Z-0101-2015 | 69242 | No | BSB | Blood and plasma warming device. | Hematology | Hematology | 510(k) | Ultra Battery 1 Model: TA-UB1 NSN: 6515-01-592-5803 |  |  | July 25, 2014 | 2014 | Estill Medical Technologies, Inc | Estill has discovered that a fully discharged Ultra Battery 1 (TA-UB1) will reach 96.9% of its capacity during a 60 minute charging period. If you have a fully discharged Ultra Battery 1 (TA-UB1), in order to achieve a 100% of charge capacity you should charge it for a period of 90 minutes. In addition, the charging voltage range of 10-36V specified on the Ultra Battery 1 (TA-UB1) label and in t | Device Design | Thermal Angel sent a Notice of Corrections letter dated July 25, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to communicate the contents of the notice to all employees, consignees, consultants, or other parties affected by this correction. The firm also updated the labeling and instructions for use. For additional information cusotmers were instructed to contact Medical Technologies at (877)354-0286.The recalling firm issued a Notice of Corrections to its distributors and customers.For questions regarding this recall call 214-561-6001. |  | 6886 | 6,886 | Worldwide Distribution - USA (nationwide) and Internationally to Serbia. | Terminated | August 21, 2015 | 392 | 6886 | Not_Computer
Z-0111-2015 | 69420 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Symmetric Total Knee Femoral Components |  |  | September 30, 2014 | 2014 | Signal Medical Corporation | Packaging of a device was compromised | Packaging | Signal Medical sent a Notice letter dated September 30, 2014 to customers. Customers were instructed to carefully inspect devices to ensure package integrity before using. If packaging is compromised it should not be used. Return the affected device and packaging to Signal Medical Corporation. Replacement devices with new packaging configuration is shipped to each location with the old packaging configuration in the event that the packaging is found to be compromised. For questions contact Signal Medical Corp. at 1-800-246-6324. |  | 41 | 41 | Worldwide Distribution - US (Nationwide) in the state of: MO and internationally to: Panama. | Terminated | March 14, 2016 | 531 | 41 | Not_Computer
Z-0115-2015 | 66778 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia E is a multi-purpose SPECT system ideal for hospitals and outpatient centers. |  |  | October 28, 2013 | 2013 | Siemens Medical Solutions USA, Inc. | Reports of loose bolts in the detector support structure of Symbia and E.CAM systems. | Process control | Undated "Customer Safety Advisory Notice" letters were sent to all direct accounts (customers). The letters stated the problem, and included instructions for customers. Customers: 1) can continue to use the devices while waiting for an inspection to be scheduled and performed; and, 2) their local service engineer will inspect the detector pivot connections and make any necessary repairs on site. Customers with any questions can contact their local service representative or Siemens Molecular Imaging at 800-888-7436. |  | 562 | 562 devices | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and, the District of Columbia, Guam, and Puerto Rico; and, the countries of Albania, Algeria, Angola, Argentina, Austria, Australia, Azerbaijan, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia and Herzegovina, Botswana, Brazil, Bulgaria, Cameroon, Canada, Chile, China, Columbia, Congo, Costa Rica, Cuba, Croatia, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Honduras, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lebanon, Luxembourg, Macedonia, Madagascar, Malaysia, Mexico, Morocco, Myanmar, Nepal, Netherlands, New Zealand, Niger, Nigeria, Norway, Oman, Pakistan, Palestine, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Sri Lanka, Sweden, Switzerland, Syria, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Vietnam, and Yemen. | Open, Classified |  |  | 5424 | Not_Computer
Z-0114-2015 | 69256 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | ASK-05501-AFH1, Lot Number 23F13K0969 Anesthesia Conduction Catheter, Epidural Anesthesia Kit |  |  | September 15, 2014 | 2014 | Arrow International Inc | Arrow International, Inc. is notifying customers that lot 23F13K0969 of product code ASK-05501-AFH1 contains an incorrect component. The lidstock indicates that the kit contains one Syringe, 10mL, Luer-Slip Loss of Resistance, Glass but contains one Syringe, 10mL, Luer-Lock Loss of Resistance, Glass. | Labeling mix-ups | Arrow International sent a Medical Device Advisory Notification dated September 15, 2014, to affected customers notifying them that lot 23F13K0969 of product code ASK-05501-AFH1 contains an incorrect component, and requesting they provide this Advisory Notice to all those who need to be aware of it within your organization. Questions were directed to local sales representative or Customer Service at 1-866-246-6990. |  | 300 | 300 | US Distribution in Utah only. | Terminated | December 18, 2015 | 459 | 300 | Not_Computer
Z-0109-2015 | 69418 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Microseal Augmented (20 degrees) Liner, acetabular liner component made from UHMW polyethylene (ASTM F648). Sizes 36x52x20; and 36x50x20. |  |  | September 23, 2014 | 2014 | Signal Medical Corporation | Incorrect dimension on four liners was detected | Process design | Signal Medical Corporation verbally notified the only consignee via telephone. Please call Signal Medical Corp. at 1-800-246-6324 with any questions. |  | 12 | 12 devices | US Distribution in MO only. | Terminated | January 20, 2015 | 119 | 12 | Not_Computer
Z-0113-2015 | 69204 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Bionector Needleless connector. Bionector is a multi-purpose catheter accessory; a closed needle-less system permitting blood sampling, intermit ant injection and continuous infusion of fluids or medications to patients in a hospital or homecare situation. |  |  | August 28, 2014 | 2014 | Vygon Corporation | Vygon has uncovered potential problems with membrane elasticity in the 896.038, Bionector Needleless connector and is issuing a voluntary recall of the 896.038, Bionector Needleless connector. | Component design/selection | Vygon sent an Urgent Medical Device Recall letter dated August 28, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their stock for the affected product and immediately cease distribution, and quarantine all affected product immediately. Count and document the affected inventory on the "Recall Acknowledgement and Inventory Return Form. Customers were asked to fax or email the form to Vygon including customer contact information. Vygon customer service will contact the customer and issue a Return Merchandise Authorization number. Customers were instructed to return the product as well as the original signed Recall Acknowledgement and Inventory Return Form using the instructions from Customer Service involving product pick-up. All inventory will be credited at no cost to the customer.For questions regarding this recall call 800-473-5414. |  | 42250 | 42250 | Worldwide Distribution - US including CA and Internationally to CA. | Terminated | March 30, 2015 | 214 | 42250 | Not_Computer
Z-0110-2015 | 69467 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Mortara Surveyor S12 and Surveyor S19 Patient Monitors. |  |  | October 07, 2014 | 2014 | Mortara Instrument, Inc | Martara Instrument, Inc. has recently became aware of a potential safety hazard involving our Surveyor rand S12 S19 Patient Monitors. A hazardous situation may be created when mounting these monitors using the 75mm VESA mounting hole pattern provided in the battery cover of these monitors. If the mounting screws provided by Mortara are not used and if the mounting screws used are longer than the p | Device Design | Mortar Instruments, Inc. sent an "Urgent Medical Device Correction" letter dated October 3, 2014. The letter was addressed to Health Care Administrator / Risk Manger and Director of Biomedical / Clinical Engineering.The letter described the problem and the product involved in the recall. Customers were advised to contact Mortara Technical Support at 1-888-667-8272 or TechSupport @mortara.com to schedule a service representative to update their Surveyor S12 and Surveyor S19 units or for additional questions they can contact their local Service Representative. The update will include replacement of the battery cover and replacement of the battery pack if mounting screw damage is evident. International consignees received similar dated letter except it provided instructions detailing the correction steps and requested customers to complete and return the Response Acknowledgement and Receipt Form. For questions they can contact Technical Support at +1.414.354.1600 or TechSupport@mortara.com, or contact your local Service Representative. |  | 624 | 624 | Worldwide Distribution: US (Nationwide) in the states of: AZ, AR, CA, DC, FL, ID, IL IN, KY, LA, ME, MD, MA, MI, MN, MO, MT, NV, NJ, NY, NC, OK PA, SC, TN, TX, VA, and WI; and internationally to: AUSTRALIA, BANGLADESH, CANADA, CHILE, CROATIA, EGYPT, GERMANY, ITALY, JORDAN, LEBANON, MALAYSIA, MOROCCO, OMAN, PHILIPPINES, POLAND, ROMANIA, SAUDI ARABIA, THAILAND, THE NETHERLANDS, UNITED KINGDOM. | Terminated | February 13, 2015 | 129 | 624 | Not_Computer
Z-0108-2015 | 69499 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | St Jude Medical, 10F Fast-Cath Trio" Hemostasis Introducer, REF # 406308, Sterile EO. |  |  | October 14, 2014 | 2014 | St Jude Medical | St. Jude Medical is conducting a voluntary recall of SMJ Batch # 4641580 of the Fast-Cath Trio" Hemostasis Introducer, Catalog number 406308. The inner diameter of the 10F dilator is too small, such that the guidewire cannot advance through the device which may result in a procedural delay while another dilator of the desired size is obtained. | Process control | St. Jude Medical sent an Urgent Medical Device Recall Notice dated October 14, 2014, to all affected consignees. The letter described the problem and the product affected by the recall. Advised consignees to discontinue us of the device. A St. Jude Medical sales representative were to contact the consignee in the near future to facilitate removal and reconciliation of the recalled device. Customers with questions were instructed to contact their St. Jude Medical sales Representative.For questions regarding this recall call 651-756-2295. |  | 220 | 220 | Nationwide Distribution including FL, IN, ME, MA, NY, NY, OH, RI,TX, and WV. | Terminated | March 25, 2015 | 162 | 220 | Not_Computer
Z-0118-2015 | 69236 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | The Aeroneb Professional Nebulizer System is a portable medical device for multiple patient use that is intended to aerosolize physician-prescribed solutions for inhalation to patients on and off ventilation or other positive pressure breathing assistance. The Aeroneb Professional Nebulizer System is suitable for use in adult and pediatric patients. |  |  | September 02, 2014 | 2014 | Aerogen Ltd. | Aerogen has received a small number of complaints in relation to the main power adapter manufactured by FRIWO GmbH and used in the Aeroneb Solo and Pro Nebuliser systems. The complaints received identified that the power adapter cover had broken and dislodged from the body of the adapter prior to use. | Nonconforming Material/Component | URGENT FIELD SAFETY NOTICE letters dated September 2, 2014 were sent to all direct accounts (customers). The letters included instructions for customers to: 1) compare their power adapter date codes against those found in the letter as being subject to this recall; 2) until new power adapters are received, customers may continue to use the recalled power adapters if the customer inspects each power adapter prior to use and finds that there is no evidence of damage or separation of the housing and unplug the adapter after each use; and, 3) contact the Aerogen or appropriate Aerogen sales partner immediately so that the power adapter can be replaced free of charge. The letter further requests that the recall notification be forwarded to any parties that may have received the recalled power adapters. Customers with questions can contact Aerogen at 866-423-76436. |  | 2421 | 2,421 power adapters | Worldwide Distribution-USA (nationwide) including the states of CA, GA, IL, MS, NC, and TX, and the countries of Australia, Belgium, Canada, China, Czech Republic, Denmark, Egypt, Finland, France, Germany, Hungary, India, Ireland, Israel, Italy, Kuwait, Libya, Malaysia, Netherlands, Norway, Poland, Qatar, Russia, South Korea, Singapore, Sweden, Switzerland, Thailand, Turkey, United Arab Emirates, United Kingdom, and Vietnam. | Terminated | April 25, 2016 | 601 | 4003 | Not_Computer
Z-0123-2015 | 69354 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | The Persona Tibial Articular Surface Inserter is a sterilizable instrument intended for |  |  | October 07, 2014 | 2014 | Zimmer, Inc. | Affected product did not undergo heat treatment hardening step leading to the potential for deformation of the tip. | Process control | Zimmer sent an " URGENT MEDICAL DEVICE RECALL - LOT SPECIFIC " notifications dated September 2014, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 618 | 618 units | Worldwide Distribution - USA (nationwide) including the states of NY, OR, NC, OH, TX, FL, CO, IN, TN, PA, CA, KS, MO, WI, VA, MN, NJ, UT, MA, GA, MI, OK, AZ, MS, IL and AL., and the countries of Korea, Singapore, Canada and Japan | Terminated | September 23, 2015 | 351 | 618 | Not_Computer
Z-0122-2015 | 69321 | No | CGR | Cortisol (hydrocortisone and hydroxycorticosterone) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | IMMULITE /IMMULITE 1000 Systems COR Cortisol, REF/Catalog Number LKC01, SMN 10381388, IVD. |  |  | September 19, 2014 | 2014 | Siemens Healthcare Diagnostics | Customer complaints were received for positive bias reported on the controls used on the IMMULITE /IMMULITE 1000 on the Cortisol Assay reagent lots IMMULITE /IMMULITE 1000 (LKCO1) Lots 381,382 and 383 . Subsequently, the firm confirmed a positive bias on the IMMULITE/IMMULITE 1000 (LKCO1) Lots 381, 382 and 383. | Nonconforming Material/Component | US Customers were sent an Urgent Medical Device Recall Letter (UMDR2014-09-19, dated 9/19/14) via FedEx on 9/19/14. Outside US Customers were sent an Urgent Field Safety Notice Letter. The letters inform the consignees of the recall of the IMMULITE/IMMULITE 1000 Cortisol (LKCO1) assay, kit lots 381, 382 and 383. Customers were asked to take the following actions: - Please review this letter with your Medical Director. - Discontinue use of and discard the IMMULITE/IMMULITE 1000 Cortisol kit lots listed in the letter. - Review your inventory of these products to determine your laboratorys replacement needs and to provide information to Siemens for reporting to the authorities. --- Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Field Safety Notice letter within thirty (30) days. |  | 815 | US: 772 units, Canada: 43 units | Worldwide Distribution: US (nationwide) and internationally to Canada. | Terminated | November 18, 2015 | 425 | 815 | Not_Computer
Z-0117-2015 | 69379 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Straumann Bone Level Tapered Implant 04.1 mm RC, SLActiveÂż 8mm, RoxolidÂż, Loxim |  |  | September 19, 2014 | 2014 | Straumann USA, LLC | Assembled with Narrow Connect (NC) transfer piece instead of Regular Connect (RC) | Packaging process control | Strauman notified consignees of this Field Corrective Action by letter dated September 19, 2014 and delivered via hand by Straumann Representative. Accounts are requested to return inventory. Contact (978) 747-2514 for assistance. |  | 6 | 6 units | US Distribution in states of: FL and MA. | Terminated | March 13, 2015 | 175 | 6 | Not_Computer
Z-0129-2015 | 69449 | No | DBF | Ferritin immunological test system. | Immunology | Immunology | 510(k) | ADVIA Centaur Calibrator C-Used with the ADVIA Centaur Ferritin assay on the ADVIA Centaur CP system only. (In-Vitro Diagnostic) |  |  | October 09, 2012 | 2012 | Siemens Healthcare Diagnostics, Inc | Ferritin Analytical Sensitivity and Calibrator C Realignment with kit lots ending in 22 and 23 used with the ADVIA Centaur Ferritin assay on the ADVIA Centaur CP system | Nonconforming Material/Component | Siemens issued an Urgent Field Safety Notice via e-mail to all US customers affected on October 9, 2012. This notice informs the customer of the Field Correction and the actions that need to be taken. The Urgent Field Safety Notice instructs customers that upon receipt of Calibrator C Lots ending in 25 and higher (Cal C lots CC25LA and higher and Cal C lots CC25HA and higher), discontinue use of the Calibrator C lots ending in 22 and 23 with the ADVIA Centaur Ferritin assay when used on the ADVIA Centaur CP system. A confirmation fax-back form is included to ensure that all customers worldwidehave been notified and have complied with the mandatory action. Further questions please call (508) 668-5000 |  | 2440 | 2440 kits | Worldwide Distribution - Nationwide Distribution and the countries of Argentina, Australia, Brazil, Canada, China, Columbia, Egypt, Hong Kong, India, Israel, Indoenesia, India, Japan, Korea, Mexico, Malaysia, Philippines, Saudi Arabia, Singapore, Taipei, Uruguay, Venezuela, Viet Nam and South Africa. | Terminated | June 18, 2015 | 982 | 10937 | Not_Computer
Z-0124-2015 | 69417 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Spirotome" Soft-Tissue Biopsy Needle Set. |  |  | September 26, 2014 | 2014 | Cook Medical Incorporated | MEDINVENTS, the manufacturer of the Spirotome&quot; Soft-Tissue Biopsy Needle Set distributed by Cook Medical, has initiated a voluntary recall of these products. The recall hasbeen initiated due to customer complaints that the protective caps were dislodged in the pouch. Furtherinvestigation disclosed potential dislodging of protective caps on the needle tips on other distributedproducts. | Package design/selection | On 9/26/2014, URGENT: MEDICAL DEVICE RECALL notifications were sent to the affected distributors and institutions Risk Managers/Recall Administration with instructions for returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call Customer Relations at 1-800-457-4500 or 1-812-339-2235. |  | 1239 | 1239 units | Worldwide Distribution-USA (nationwide) including the states of Arizona, California, Colorado, Hawaii, Iowa, Illinois, Massachusetts, Maryland, Michigan, Missouri, New Mexico, Pennsylvania, Texas, Virginia, Vermont, and Wisconsin, and the countries of AT; BE; CH; CY; CZ; DE; DK; ES; FI; FR; GB; HU; IC; IE; IS; IT; LB; LT; LU; NL; NO; PT; QA; SE; TR; and ZA. | Terminated | February 02, 2016 | 494 | 1239 | Not_Computer
Z-0135-2015 | 68201 | No | JJX | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | HbA1c Calibrators LN 4P52-02 consists of: 4P52-02 Hemoglobin A1c 1 x 1.6 mL and 1 x 1.6 mL. A1c Calibrators (lyophilized) contain hemoglobin and glycated hemoglobin from human whole blood. Prior to lyophilization, the calibrator matrix is an MES-buffered solution. Preservative: Ofloxacin. |  |  | April 30, 2014 | 2014 | Abbott Laboratories | HbA1c Calibrators LN 4P52-02, lot 45063UQ12 may contain a value sheet for a different HbA1c calibrator lot. If calibrator lot 45063UQ12 is used in the calibration of the HbA1c assay using values from an incorrect calibrator value sheet then all results, including quality control and patient, may be falsely elevated. | Under Investigation by firm | Product Correction letters dated April 30, 2014 were sent to all customers. The letters included isntructions for customers to: 1) verify that the lot numbr listed on each calibrator bottle agrees with the lot number printed on the value sheet; 2) If HbA1c Calibrators lot 45063UQ12 has already been utilized in your laboratory, verify the concentrations configured for this calibrator lot match the values in the letter; 3) If these do not match, use the values in the letter when calibrating the HbA1c assay using calibrator lot 45063UQ12; and, 4) follow your laboratory protocol regarding the need for reviewing previously reported patient results if incorrect calibrator values were used. |  | 183 | 183 calibrator kits | Worldwide Distribution in the countries of Austria, Canada, Cyprus, Denmark, Finland, France, Germany, Greece, Hong Kong, Ireland, Italy, Kuwait, Malaysia, Norway, Poland, Romania, Russia, Saudi Arabia, Thailand, Turkey, and United Arab Emirates. | Open, Classified |  |  | 183 | Not_Computer
Z-0132-2015 | 69101 | No | MWS | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | 510(K) Exempt | Acrobat V Vacuum Off-Pump System |  |  | August 13, 2014 | 2014 | Maquet Cardiovascular, LLC | An internal finding identified the presence of pinholes/cuts in some of the Tyvek pouches of the vacuum tubing sets included as an accessory to Acrobat V Vacuum Off-Pump System, Acrobat SUV Vacuum Off-Pump System, Axius XPOSE 3 Access Device and XPOSE 4 Access Device. | Package design/selection | Maquet Inc. issued Recall letters/Acknowledgement forms dated via Fed-Ex on 8/13/2014. |  | 16532 | 16,532 units (19,270 units international) | Worldwide Distribution: US (nationwide) and countries of: ARGENTINA, AUSTRIA, BANGLADESH, BELARUS, BRAZIL, BRUNEI, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CZECH REPUBLIC, DENMARK, ECUADOR, GERMANY, FINLAND, FRANCE, GREECE, HONG KONG, INDICA, INDONESIA, ISRAEL, ITALY, JAPAN, MEXICO, MALAYSIA, NETHERLANDS, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAK REPUBLIC, SOUTH KOREA, SPAIN, SWITZERLAND, TAIWAN, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM AND VIETNAM. | Terminated | March 30, 2016 | 595 | 16532 | Not_Computer
Z-0138-2015 | 69279 | No | NAD | Occlusive wound dressing. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Biatain Super Non-Adhesive, Catalog number 46300, 10 pcs, REF 0463001400. Country of Origin:HU. |  |  | September 08, 2014 | 2014 | Coloplast Manufacturing US, LLC | Coloplast is conducting a recall for one lot (4201931) of the Biatain Super Non-Adhesive wound dressing which was mislabeled. | Labeling mix-ups | Coloplast sent an "Urgent Voluntary Recall Notification" dated September 9, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If they have further distributed the product they need to contact their customers immediately. For questions they can contact Coloplast Customer Care Representative at 800-533-0464.A second follow up "Urgent Voluntary Recall Notification" letter dated September 23, 2014 was also sent to consignees. |  | 6 | 6 | US Distribution to MO and RI. | Terminated | January 21, 2015 | 135 | 6 | Not_Computer
Z-0137-2015 | 69296 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Animas Vibe Insulin Infusion Pump and System. |  |  | October 25, 2012 | 2012 | Animas Corporation | Animas determined that keypad wear over time was contributing to an increased complaint rate associated with damaged keypads on the Vibe Insulin Infusion Pump and System. Animas has implemented an improvement to the keypad related to the material used to reduce susceptibility to damage and increase durability. | Device Design | Animas initiated the notification to Distributors by letter, dated July 2012, and Patients and Health Care Professionals were notified by letter in September 2012. The delay in notifying patients and health care professionals was driven by the need to have sufficient inventory of replacement product on hand in order to meet anticipated demand before sending the letters. |  | 1938 | 1938 | Worldwide Distribution in the countries of France, Italy, Sweden, and UK.. | Terminated | June 29, 2015 | 977 | 1938 | Not_Computer
Z-0095-2015 | 68047 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The GemStar Power Supply is an accessory for the GemStar Infusion Pump. It is a two piece assembled switching power supply. The Power Supply converts alternating current (AC) line input voltage of 120 or 240 V to an output to the device of 3.3 Volts DC. |  |  | February 03, 2014 | 2014 | Hospira Inc. | The GemStar Infusion pump may not receive direct current (DC) power from the power supply. In one instance, smoke was observed and the GemStar pump was operating on battery power while connected to the 3 volt DC power supply. | Process control | An URGENT MEDICAL DEVICE CORRECTION letter dated February 3, 2014 was sent to all direct accounts. The letter included specific instructions for how to handle and use the power supplies so as to minimize the chance of any device failures. There were additional instructions for customers to: 1) Inform potential users in their organization of the product of this notification; 2) Complete the attached reply form and return it to the fax number or e-mail address on the form (even if you do not have any affected product in inventory); 3) If the products were further distributed, their customers should be notified of the recall and asked to contact Stericycle at 888-641-9735 to receive a reply form; and, 4) if a potential user believes that their power supply is damaged for any reason, it should be removed from service and the user should contact Hospira Global Complaints at 800-441-4100 to report the concern. The letter further notes that there is no need to return fully functional power supplies (unless they malfunction). |  | 978 | 978 power supplies | Worldwide Distribution -- USA, including the states of AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY; and, the countries of Austria, Belgium, Brazil, Columbia, France, Germany, Greece, Netherlands, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 07, 2017 | 1128 | 13112 | Not_Computer
Z-0134-2015 | 69160 | No | LPL | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | 510(k) | ProClear Toric contact lenses are made from a material containing 62% water and 38% omafilcon B, a random copolymer of 2-hydroxyethylmethacrylate and methacrylic acid. The lenses have a tint (VAT Blue 6) which is added to make the lens more visible for handling. |  |  | September 17, 2014 | 2014 | Cooper Vision Caribbean Corp. | Routine quality monitoring system has identified that a limitednumber of lots were manufactured with the cylinder axis outside the allowable specification. | Other | The firm sent out a notification letter with the following instructions:Please take the following actions immediately:" Examine your inventory for any affected lot numberso Use the enclosed report to inspect your inventory for affected product that is still in yourWarehouse." Stop any further distribution of the specified lots of Preclear T oric lenses." Quarantine any Preclear Toric lenses with the specified lot numbers." Within the next five (5) business days, please complete the attached Recall Response Form andfax it to 866-484-1048 or scan and email to info@coopervision.com.o If you have indicated that you have product to return, you will be contacted by aCooperVision Representative with instructions for pick-up and removal of affected,quarantined product. Your account will be credited once the product is returned andprocessed.This action is being taken in full cooperation with the Food and Drug Administration (FDA). The recall onlyaffects Proclear T oric lenses with the lot numbers specified on the enclosed list. No other Preclear T oricproducts- or any other CooperVision products - are affected. |  | 3355194 | 13639 lenses/246 product lots | Worldwide Distribution: US (nationwide) and internationally to: Canada, Latin America, EU, Asia Pacific, and Africa/Central Eurasia countries. | Terminated | June 10, 2015 | 266 | 3355194 | Not_Computer
Z-0126-2015 | 69347 | No | LYL | Soft (hydrophilic) contact lens care products. | Ophthalmic | Ophthalmic | 510(k) | Equate Moisture Last Multi-Purpose Solution, 2 pack, 12 fl. oz. (355 mL), Distributed by Walmart Stores, Inc. |  |  | September 25, 2014 | 2014 | Bausch &amp; Lomb Inc- Greenville Solutions Plant | Product lot did not meet the specification for its disinfectants at the time of manufacture. | Other | Consignees were notified via letter starting 09/25/2014. |  | 18792 | 18,792 bottles (9,396 twin packs) | US Nationwide Distribution in states of: AL, AR, AZ, CA, CO, DE, FL, GA, IA, IL, IN, KS, LA, MS, NC, NH, NY, OH, OR, PA, TN, TX, UT, VA, and WI. | Terminated | June 12, 2015 | 260 | 18792 | Not_Computer
Z-0127-2015 | 69334 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Dimension VistaÂż HbA1c |  |  | August 14, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | In March, Siemens Healthcare Diagnostics issued an Urgent Medical Device Correction for decreased recovery of patient samples with the Dimension Vista&reg; CSA &amp; CSAE Flex&reg; reagent cartridges when whole blood samples are run from Small Sample Containers (SSC). Siemens has confirmed that when whole blood samples are run from SSC, HbA1c values may exhibit a percent bias ranging from - 8.2% t | Device Design | An Urgent Medical Device Correction, dated August 2014, was provided to all Dimension VistaÂż HbA1c customers who received the affected lots to notify them of the issue, the potential risk to health, and actions to be taken by customers. |  | 20043 | 20043 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | August 10, 2015 | 361 | 20043 | Not_Computer
Z-0136-2015 | 69293 | [6 empty] | Animas Vibe Insulin Infusion Pump. |  |  | June 10, 2014 | 2014 | Animas Corporation | Animas determined that the Hebrew language translation in the Vibe Insulin Infusion Pump was incorrect and could lead to potential patient confusion. In addition, one line of Hebrew text was omitted from the display screen. | Employee error | A recall notification letter, dated June 2014, was sent to the recalling firm's sole customer, TriTech Biomed, in Israel, to notify them about the product, problem, and action to be taken. The distributor sent letters to their customers on July 16, 2014. |  | 1028 | 1028 | International Distribution in Israel only. | Terminated | September 02, 2015 | 449 | 1028 | Not_Computer
Z-0141-2015 | 69426 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | TITAN Operating Room Table |  |  | October 09, 2014 | 2014 | Trumpf Medical Systems, Inc. | Unintended movement of the operating table during surgical cases due to the use of damaged TITAN remotes. | Device Design | TRUMPF sent an Urgent Medical Device Correction letter dated October 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Actions to be taken:Complete and return the attached Confirmation of Receipt form with serial numbers of the TITAN table(s) and the serial number(s) for any remote controls you have. We will then arrange to send replacement remotes, which have been redesigned to eliminate this mode of failure. You will then be required to dispose of your recalled remotes.If you have any questions concerning this request or the procedure as outlined, please contact TRUMPF Medical Systems, Inc. at 888-474-9359. |  | 184 | 184 units | Worldwide Disttribution: US (nationwide) including the states of AL, AZ, CT, FL, GA, IA, IL, IN, KS, MD, MI, MN, MO, NC, ND, NJ, NM, NV, NY, OH, OR, PA, SC, TX, VA, WA and WY: and Internationally to Canada. | Terminated | March 23, 2016 | 531 | 184 | Not_Computer
Z-0139-2015 | 69532 | No | LPB |  |  | Cardiovascular | PMA | ThermoCool SmartTouch Navigation Catheters, Catalog No. D132701, D132702, D132703, D132704, D132705, D133601, D133602, D133603. |  |  | October 19, 2014 | 2014 | Biosense Webster, Inc. | The recall was initiated because Biosense Webster is providing additional labeling for the safe and effective use of the ThermoCool SmartTouch Catheter. | Use error | A customer notification letter dated October 19, 2014, will be sent to all customers who purchased the Biosense Webster's ThermoCool SmartTouch Catheter Family. The letter informs the customers of the additional information for the safe and effective use of the ThermoCool SmartTouch Catheter, which will be included in the updated labeling. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions related to the recall letter are instructed to contact their Biosense Webster sales representative or call (866) 473-7823, Monday through Friday from 7am to 8pm EST. |  | 173329 | 173,329 units total (21,812 units in US) | Nationwide Distribution | Terminated | October 06, 2015 | 352 | 173329 | Not_Computer
Z-0140-2015 | 69373 | No | NEW | Absorbable polydioxanone surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Monoswift PGCL [poly(glycolide-co-caprolactone)] Absorbable Surgical Suture; USP size 3/0. Product is labeled sterile. The product is sealed inside an inner foil pouch with an outside Tyvek pouch. It is an undyed, monofilament synthetic absorbable suture with needle size 3/0. |  |  | September 22, 2014 | 2014 | C P Medical, Inc. | CP Medical Inc. is recalling five lots of Monoswift PGCL [poly(glycolide-co-caprolactone)] Absorbable Surgical Suture for failing to meet the minimum tensile strength as dictated by USP was not met over the length of the product's shelf life. | Nonconforming Material/Component | CP Medical Inc.notified their customers between September 23 through October 3, 2014, of the recall via telephone or e-mail and followed up with a certified letter. Customers were advised to discontinue sale and/or use of the recalled product and return all existing stock to CP Medical. CP Medical also requested their customers to notify their downstream customers of the recall and request they return any unused recalled product, as well as provide information to them regarding distribution of the product. Customers are being asked to fill out and return a questionnaire regarding the recalled product and contact CP Medical Customer Service Representative at 1-800-950-2763 for return authorization and shipping account number. All shipping costs including those from downstream customers will be paid for by CP Medical. Replacement product or credit will be provided. |  | 888 | 74 boxes /12 sutures per box; plus an additional 13 sutures | US Distribution including the states of: AR, AZ, CA, FL, ME, MI, NE, OR, SC, and WA | Terminated | February 17, 2015 | 148 | 888 | Not_Computer
Z-0143-2015 | 69406 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Merge Healthcare, Merge Hemo Programmable diagnostic computer. |  |  | September 29, 2014 | 2014 | Merge Healthcare, Inc. | It has been reported that during use, the SpO2 value displayed on the Hemo Monitor may not update to reflect changes in the patient's oxygen value. It is also possible that if using the pulse rate of the SpO2 finger clip to calculate the patient's heart rate, it too may be subject to displaying a stale value. | Software design | Consignees were e mailed on 5/29/14 a Merge "Important Alert" letter. The letter described the problem and the product involved in the recall. Informed consignees that they had the opportunity to replace the Nellcor SpO2 module with a Masimo SpO2 module free of charge. The letter also provided Clinical Information and the Next Steps to follow. This included to contact Merge Technical Support if they wanted to replace their Nellcor module. For questions, they can contact Lee Harrop, Director Solutions Management, Cardiology, 386-239-2224, Lee.Harrop@Merge.com.A second letter was sent Sept. 11, 2014 via e-mail. Consignees were made aware of a firmware patch available. |  | 52 | 52 | Nationwide Distribution including the states of AZ, CA, FL, IL, IA, MA, MO, OH, SC, TN, and WI. | Open, Classified |  |  | 52 | Not_Computer
Z-0112-2015 | 69320 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Covidien Puritan Bennett 980 Ventilator, Rx ONLY. |  |  | October 01, 2014 | 2014 | Nellcor Puritan Bennett Inc. (dba Covidien LP) | A software issue may lead to ventilator inoperative situations. | Software design | The firm sent Urgent: Field Correction Action letters, dated October 3, 2014, to consignees. The letter identified the affected product and the reason for the recall. Customers were provided actions to be taken, as well as, important safety reminders. The letter informed customers that they may continue to use their ventilators pending the software correction as long as two gas sources are connected to the ventilator at all times. Customers are to complete the attached acknowledgement and receipt form and fax it to Covidien at the number provided. For further assistance, customers are to contact the Technical Support Department at 1-800-255-6774. |  | 324 | 324 units | Worldwide Distribution -- USA, including the states of CA, CO, FL, GA, KY, MA, MN, NC, NY, OH, OK, PA, SC, TN, TX, UT, WI; and, the countries of Canada, Mexico, Saudi Arabia, South Africa, and United Arab Emirates. | Terminated | May 02, 2016 | 579 | 324 | Computer
Z-0142-2015 | 64774 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | DMI Duro-Med Industries Rollator Transport Chair; Wheel Size: 7 1/4"; Overall Width: 23 1/2"; Seat Height: 21 1/4"; Seat Size: 14" x 12 3/4"; Handle Height: 31" - 36"; Width Between Handles: 18"; Weight: 19 Lbs; |  |  | June 05, 2014 | 2014 | Briggs Medical Service Company d.b.a. Mabis | There have been 5 customer reports of a 501-1018 series DMI Rollator Transport Chair collapsing during use. | Device Design | The firm, Briggs Healthcare, sent "URGENT PRODUCT RECALL" letters dated July 28, 2014 to all distributors, retailers and consumers. The letters included instructions for distributors and dealers to: 1) examine their inventory; 2) quarantine any affected products; 3) cease distribution of the affected products; and, 4) complete and return the attached Customer Recall Acknowledgement Form. Distributors and dealers who further distributed the products were also instructed to notify their customers of this recall by forwarding a copy of the enclosed consumer notification letter and its associated documents. Distributors and dealers with questions about this recall can contact Briggs at 800-342-4426 (Monday - Friday, 8:00 AM - 6:00 PM EDT) or by e-mail at RecallCoord@briggscorp.com. The consumer-directed letter dated July 28, 2014 included instructions for consumers to examine their rollator transport chair to determine whether or not it is subject to the recall. If the consumer has the affected product, they should discontinue using it immediately and contact Briggs at 800--342-4426 (Monday - Friday, 8:00 AM - 6:00 PM EDT) or by e-mail at RecallCoord@briggscorp.com to arrange for the return of the product. Consumers will be given the option of receiving a cash voucher or a replacement product in exchange for the recalled product. |  | 1869 | 1,869 units | Worldwide Distribution: US (nationwide) including states of:US: Nationwide (AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY) + District of Columbia + Puerto Rico; and country of : Canada.); FOREIGN: Canada | Open, Classified |  |  | 1869 | Not_Computer
Z-0144-2015 | 69265 | No | FQO | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | IMRIS Operating Room table (ORT200 and ORT300). |  |  | October 01, 2014 | 2014 | IMRIS Inc | One of the three pins that is used to hold the ORT200 table alignment block in place on the floor of the MR Suite went missing during testing at a health care facility. | Use error | Consignees were sent on 10/01/2014 an IMRIS "Urgent Field Safety Notice" letter dated September 25, 2014. The letter described the problem and the product involved in the recall. The letter informed consignees that the pins will be replaced with pins made of non-ferrous material. Customers are to ensure that users are aware of the changes and the existing procedure for table alignment and managing ferrous content in the magnet room. Contact customer support at 1-866-475-0525 for questions. |  | 14 | 14 | Worldwide Distribution -- USA, including the states of CA, FL, NH, NC, TN, and UT; and, the countries of AUSTRALIA, CANADA, and JAPAN. | Terminated | February 20, 2015 | 142 | 14 | Not_Computer
Z-0131-2015 | 67987 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Lipase Reagent, 7D80. |  |  | April 07, 2014 | 2014 | Abbott Laboratories, Inc | Package inserts have incorrect SmartWash parameters for Triglyceride. | Labeling Change Control | Firm issued a Product Correction Letter to all Lipase reagent customers who received lots 43827UQ12, 43819UQ03, 44449UQ05, 44127UQ06, 44334UQ07, 43994UQ07, 44596UQ08, 44775UQ09, and 44930UQ11. The letter identified the affected product and provided an explanation of the problem, as well as, the patient impact and actions to be taken. The Product Correction Letter provided guidance for the laboratory to correctly configure the Architect cSystem's Lipase assay parameters. Customers are to retain a copy of the letter for their lab records, and forward a copy to any laboratories that the product was further distributed to. Questions should be directed to Customer Service at 1-877-4ABBOTT or your local customer service for customers outside of the US. |  | 11916 | 11916 units | Worldwide Distribution: US (nationwide) and countries of: Algeria, Angola, Argentina, Armenia, Australia, Austria, Bangladesh, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Columbia, Croatia, Czech Republic, Dominican Republic, Egypt, Estonia, France, Germany, Greece, Guatemala, Hungary, India, Indonesia, Iraq, Ireland, Israel, Italy, Jamaica, Jordan, Kenya, Kuwait, Kyrgyzstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Moldova, Montenegro, Morocco, Netherlands, New Zealand, Norway, Pakistan, Panama, Paraguay, Philippines, Poland, Portugal, Qatar, Republic of Yemen, Romania, Saudi Arabia, Senegal, Serbia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Trinidad and Tobago, Tunisia, Turks and Caicos, Turkey, Uganda, United Arab Emirates, United Kingdom, and Uruguay. | Terminated | April 01, 2015 | 359 | 11916 | Not_Computer
Z-0147-2015 | 69446 | No | DRC | Trocar. | Cardiovascular | Cardiovascular | 510(K) Exempt | Blunt Hurd Cannula Blunt Tip 10 mm with 5 mm Reducer simplifies open laparoscopy. |  |  | September 13, 2013 | 2013 | CooperSurgical, Inc. D.B.A. Lone Star Medical Products | CooperSurgical is recalling two lots of 900-863 because the wrong package insert (IFU) was placed into the shipper. | Error in labeling | CooperSurgical sent an Urgent Field Safety Notice ( Correction ) dated September 27, 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The recalling firm provided correct package inserts (IFU) to affected consignees. The recalling firm intends to conduct effectiveness checks by reviewing return receipts from the initial recall communication.We apologize for any inconvenience this communication may cause. If you have any questions or concerns, please contact CooperSurgical at (203) 601-9818. |  | 90 | 90 units | Worldwide Distribution: US (nationwide) to North Carolina; and Internationally to Australia and Italy. | Terminated | September 23, 2015 | 740 | 90 | Not_Computer
Z-0020-2015 | 68846 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Proteus 235 |  |  | March 27, 2014 | 2014 | Ion Beam Applications S.A. | While the Proteus 235 is in the pencil beam scanning (PBS) treatment mode, on rare occasions, Low Level Radio Frequency (LLRF) pulsed dee voltage (i.e., voltage applied between two electrodes in the cyclotron to accelerate protons) may generate a hot spot on a pencil beam position within a tumor, generating a dose that might exceed the prescribed dose on a single voxel. | Device Design | A User Notice was sent to direct accounts in/about March/April 2014 to notify them of the issue and provide instructions to avoid the identified potential risk to patients. |  | 8 | 8 | Worldwide Distribution-USA (nationwide) including the states of PA, NJ, IL, WA, and TN, and the countries of Germany, Italy, and Czech Republic. | Terminated | December 10, 2014 | 258 | 8 | Not_Computer
Z-0148-2015 | 69372 | No | DIP | Diphenylhydantoin test system. | Toxicology | Toxicology | 510(k) | siemens DimensionÂż Phenytoin FlexÂż Reagent Cartridge |  |  | September 24, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Internal testing has confirmed a low absorbance range between the levels 1 and 5 calibrators with this lot resulting in the potential for a positive or negative bias and imprecision across the PTN assay range. | Under Investigation by firm | Siemens sent an Urgent Medical Device Recall Notification dated September 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed as follows: " Discontinue use and discard any remaining inventory of PTN lot FA5154, and complete the attached form indicating your replacement product needs in the attached Field Correction Effectiveness Form. " Siemens will replace any unused inventory of the affected lot at no charge.We apologize for the inconvenience this situation may have cause. If you have any questions, please contact Siemens Customer Care Center - Technical Solutions at (800) 441-9250, or your local Siemens technical support representative. |  | 1733 | 1733 | Worldwide Distribution - USA (nationwide) including Puerto Rico., and the countries of Australia, Brazil, Bahamas, Canada, Chile, India, Japan and Uruguay | Open, Classified |  |  | 1733 | Not_Computer
Z-0149-2015 | 69369 | No | JQW | Pipetting and diluting system for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Freedom EVO-2 100 Base Unit, with Air LiHa Arms |  |  | September 26, 2014 | 2014 | Tecan US, Inc. | Distortion (noise interference) of the liquid level capacitance can result in air aspiration instead of liquid aspiration. This might cause a false signal (air detection) with possible dispense errors resulting in erroneous results. | Nonconforming Material/Component | Tecan sent an Urgent Field Corrective Action letter dated September 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Tecan has instructed field engineers to replace the ZAAP Motion Board in the Freedom EVO instrument. Your instrument was identified as affected and a Tecan Service Representative will contact you to arrange replacement of the board.Please make sure that your internal controls would detect any possible dispense errors that might affect your application when using the instrument before replacement of the board. Alternatively, stop using the instrument until the board is replaced.We apologize for any inconvenience this issue may have caused,and are implementing corrective and preventive actions to prevent recurrence. If you have any further questions, please contact your local Tecan Helpdesk for assistance.For further questions please call (919) 361-5200 ext. 19519 |  | 2 | 2 units | US Distribution to IL. | Terminated | July 10, 2017 | 1018 | 2 | Not_Computer
Z-0146-2015 | 69323 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | ACCU-CHEK Spirit Combo insulin infusion pump sold as part of Accu-Chek Combo System |  |  | September 25, 2014 | 2014 | Roche Diagnostics Operations, Inc. | The insulin pump may lose time and date settings during a power interruption (such as a battery change) due to a faulty capacitor. If the capacitor fails, the time and date will return to default setting. Though the pump prompts customers to confirm the time and date, if the user does not recognize the time and date have returned to default, a shift of their basal rate time block could occur. | Equipment maintenance | Roche sent an "URGENT MEDICAL DEVICE CORRECTION" letter dated September 25, 2014 via UPS Ground (receipt required) to the consignees. The letter included a description of the reason for recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. No product returns were expected. For questions, contact the ACCU-CHEK Customer Care Service Center at 1-800-688-4578. |  | 8116 | 8116 | US Nationwide Distribution in the states of CA, CT, IA, NV, OH, IN, WI, GA, MO, NY, VA, LA, ME, TX, NC, ND, SD, NJ, CO, AL, SC, IL, KY, FL, OR, MN, MD, MS, AR, WV, WA, TN, CT, ID, KS, MI, PA, UT, MA, RI, NM, OK, DE, AZ, NE, MT, NV, WY, AK, and NH. | Open, Classified |  |  | 8116 | Not_Computer
Z-0145-2015 | 69224 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | Parker Eclipse Probe Cover, Latex Free, REF 38-01, 10 Eclipse Probe Covers, Size: 2.37/1.75"W x 9.5"L (60/44 mm x 241 mm), Lot #D424, Expiry date 04/2016, Pre-gelled inside with Aquasonic 100 Ultrasound Transmission Gel |  |  | August 06, 2014 | 2014 | Parker Laboratories, Inc. | Parker Laboratories, Inc. received a report of an incorrect expiration date on the inner boxes and on the plastic bags of the Eclipse Probe Covers. | Incorrect or no expiration date | Parker Laboratories Inc. sent an "Urgent Medical Device Product Recall" Letter/Response Form via Fedex to their customers. The letter indentified the affected product, problem and actions to be taken. For questions contact Parker Laboratories Quality Assurance at 973-276-9500, ext. 7142. |  | 367 | 367 boxes | Worldwide Distribution: US (Nationwide) and country of: Canada. | Terminated | April 29, 2015 | 266 | 367 | Not_Computer
Z-0150-2015 | 69433 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | APTUS Ulna Shortening 2.5 |  |  | October 01, 2014 | 2014 | Medartis, Inc. | A recall of the APTUS Ulna Shortening 2.5 instrument was initiated due to three (3) complaints of breakage of the temporary tension bolt. | Process design | The recall is being extended to end user level of distribution. The sales representatives, customers and a distributor were notified via phone or email on September 19th and notified of the recall via US mail on 10/1/14. US customers, distributors, and sales representatives were instructed to return the affected instrument kits to the importer Medartis, Inc. via transport agency. |  | 44 | 44 | US Nationwide Distribution in the states of WV, IL, CT, AZ, and WA. | Terminated | August 06, 2015 | 309 | 44 | Not_Computer
Z-0151-2015 | 69400 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | Joerns EasyCare Bed |  |  | September 24, 2014 | 2014 | Joerns Healthcare | The Easy-Care Care-Lock feature installed at the foot of the bed may potentially fail to stay engaged (locked) after the Care-Lock System has been engaged by stepping on the red (locked) pedal. | Device Design | Joerns sent an Urgent: Medical Device Correction letter dated September 24, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. The firm will provide the customers with a temporary solution kit to safegard the bed until a replacement Care-Lock assembly can be shipped and installed. Customers were instructed to complete the response form. For questions call 800-826-0270. |  | 3135 | 3135 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | February 01, 2016 | 495 | 3135 | Not_Computer
Z-0155-2015 | 69412 | bag, urine collection, leg, for external use, non-sterile | NNW | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Coloplast Conveen Urine Collection Leg Bag, Product Numbers 5170, 5171, and 5174. |  |  | September 29, 2014 | 2014 | Coloplast Manufacturing US, LLC | Coloplast is recalling the Conveen Contour Leg Bag due to potential leakage from the bag seam when filled. | Process control | Consignees were sent a Coloplast "Recall Voluntary Notification " letter on 9-30-2014. The letter described the problem and the product affected by the recall. Requested consignees to complete the attached Acknowledgement Form and return to Coloplast or contact Coloplast Customer Service at 800-533-0464 to communicate the status of any product in their possession. If the product was distributed, they advised consignees to contact customers immediately and advise them of the recall and have them return the product to them. For questions they can contact Coloplast at Customer Service at 800-533-0464. |  | 217540 | 217,540 | Nationwide Distribution including Puerto Rico and the states of AL, AZ, CA, CO, FL, GA, ID, IL, IN, IA, KS, KY, LA, MD, MA, MI, MN, MS,MO,NE,NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, and WV. | Terminated | March 10, 2015 | 162 | 217540 | Not_Computer
Z-0152-2015 | 69425 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Laksell Gamma Knife Perfexion |  |  | October 06, 2014 | 2014 | Elekta, Inc. | Elekta has become aware that the latches may be locked before they have been fully turned resulting in a poorly locked frame. | Device Design | Elekta sent an" Important Field Safety Notice" distributed on October 6, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. This Important Field Safety Notice describes the problem and possible work arounds for customers to follow until a permanent solution can be developed. They are:" Always assure that the plastic lever is only operated when at a right angle with the frame adapter." Do not force the plastic lever in place if it meets significant resistance when turned." Ensure that the plastic lever is completely flush with the frame adapter and that no angulations are present.Corrective Action #2 - Elekta will check and replace as necessary affected frame adapters in the field. This will be released at the end of October 2014, and Elekta will have six months to complete the checks and replacements for those products in the field. If you have any questions please contact your local Elekta office, or call (770) 300-9725. |  | 181 | 181 | Worldwide Distribution - USA (nationwide) including the states of AK, AZ, AR, CA, CO, CT, FL, GA, HI, IL, KY, LA, MA, MI, MN, MS, MO, MT, NJ, NM, NY, NC, OH, OR, PA, RI, SC, TX, TN, UT, VA, WA and WV, and the countries of Australia, Austria, Belgium, Brazil, Canada, China, Colombia, Czech Republic, Egypt, France, Germany, Greece, Hong Kong, India, Indonesia, Italy, Japan, Mexico, Morocco, Netherland, Norway, Poland, Portugal, Russia, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey and United Kingdom. | Open, Classified |  |  | 181 | Not_Computer
Z-0153-2015 | 69392 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Spline Twist Implant, HA 3.25mmD X 15mmL; Part No. 2135. |  |  | October 01, 2014 | 2014 | Zimmer Dental Inc | Zimmer Dental is recalling Tapered Screw-Vent Implants because the incorrect diameter size was printed on the inner vial. | Labeling Change Control | Customer Notification letters dated 10/6/14 were sent to all customers who purchased the Tapered Screw Vent Implant and the Spline Twist Implant because the incorrect diameter size was printed on the inner vial. The letters informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete the attached Business Reply Form and fax it to (574) 372-4265 or email to corporatequality.postmarket@zimmer.com. Customers with questions are instructed to contact Zimmer Dental Customer Service at (800) 854-7019 from 7am-5pm, Monday-Friday. |  | 1044 | 1,044 units | Worldwide Distribution-US (nationwide) and the countries of Poland, Canada, Belgium, Italy, France, Israel, Morocco, Spain, Costa Rica, and Japan. | Terminated | April 28, 2015 | 209 | 1044 | Not_Computer
Z-0161-2015 | 58185 | Yes | DTR | Cardiopulmonary bypass heat exchanger. | Cardiovascular | Cardiovascular | 510(k) | Terumo Custom Cardioplegia Delivery Sets |  |  | March 01, 2011 | 2011 | Terumo Cardiovascular Systems Corp | Defect in the luer connector can compromise the seal on the female side of the connector and the compromised seal can introduce air in the table line | Device Design | Terumo Cardiovascualr notified accounts by letter dated 3/1/2011 advising users users that the luer connector to the table line in certain lots of myocardial protection sets may leak. Users should not continue to use affected myocardial protection sets in a manner that exposes the connector to negative pressure without taking correction or replacement activities as described. Firm provided mitigating instructions:Clamp the table line distal to the luer connector when ending an infusion of cardioplegia thus isolating the leak from negative pressure. Contact Terumo CVS Customer Service to report the defect and arrange return of the unit: 800-521-2818. |  | 24 | 24 units | Worldwide Distribution: US (nationwide) and the countries of Canada, Brazil, Hong Kong, Korea, and Thailand. | Terminated | November 06, 2014 | 1346 | 46190 | Not_Computer
Z-0165-2015 | 69424 | No | HAB | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | UNIVERSAL Modular Electric/Battery Handpiece for surgery |  |  | October 06, 2014 | 2014 | Zimmer Surgical Inc | Non-injury complaints from Europe identified a malfunction that made it possible for a handpieceto start by itself when the power source is connected. | Device Design | Zimmer sent an Urgent Medical Device Correction letter via email to US distributors with traceable Fed Ex mailings to all consignees dated October, 2014. The letter identified the affected product, problem and actions to be taken. Customers were advised of the potential malfunction and the operating procedure for the handpiece per the Instructions for Use (IFU). Zimmer Surgical, as the importer, will provide notice to consignees in the US accordingly. In addition, units will receive an upgrade as part of the annual preventive maintenance per the Instructions for Use. |  | 828 | 828 | US Nationwide Distribution in the states of: AK, AL,AR, AZ,CA,CO,CT, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA,ME, MI, MN, MO, NC, NE, NJ, NY, OH, OR, PA, RI, SC, TN, TX, VA, WA, WI, WV, and WY. | Terminated | April 15, 2015 | 191 | 828 | Not_Computer
Z-0162-2015 | 69452 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Merit Maestro Microcatheter, 2.4F 130cm Straight, REF28MC24130ST, Rx Only, Non-pyrogenic Sterile if package is unopened or undamaged. |  |  | September 26, 2014 | 2014 | Merit Medical Systems, Inc. | These microcatheters were packaged with a 2.8 F distal tip, rather than the labeled 2.4 F. | Mixed-up of materials/components | Consignees were notified via letter on September 26, 2014. Written notification included product identification information, instruction to immediately quarantine any devices and discontinue use, ensure all personnel to whom devices were distributed are made aware of this field action, and instructions to contact their Merit representative. An email notification was sent to US Merit sales representatives with accounts that had received affected product. The affected sales representatives were provided a Customer Response Form applicable to their account, a representative copy of communications sent directly to their affected accounts, and a copy of the Frequently Asked Questions. Merit sales representatives were instructed to visit the account and assist with the return of the affected units back to the Merit Field Assurance Department. |  | 56 | 56 units | Nationwide and to India. No govt/military/VA consignees. | Terminated | November 19, 2014 | 54 | 56 | Not_Computer
Z-0163-2015 | 69462 | No | FFL | Ureteral stone dislodger. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Bard Dimension Stone Basket |  |  | October 07, 2014 | 2014 | C.R. Bard, Inc. | The product is non-sterile. | Release of Material/Component prior to receiving test results | Bard sent an Urgent: Medical Device Recall letter to each customer on 10/7/2014 with proof of delivery notice. Wholesalers were asked to conduct a sub-recall. Letter identify the affected product, problem and actions to be taken. Customers were instructed to examine their inventory and quarantine product subject to recall, do not use or further distribute any of the affected product, remove and return affected product to Bard Medical Division, and if the product have been further distributed identify your customers and notify them at once of the product recall. Customers were instructed to complete the enclosed Recall &amp; Effectiveness Check Form and fax to 1-770-784-6469. For questions call 1-770-784-6471. |  | 86 | 86 units | Worldwide Distribution in the states of AL, FL, NJ, NY, NC, PA, SC, TX, WI, and countries of Germany, Italy, Kuwait, and Luxembourg. | Terminated | June 12, 2015 | 248 | 86 | Not_Computer
Z-0085-2015 | 69181 | No | KPE | I.V. container. | General Hospital | General Hospital | 510(k) | Baxter 2B8013 Empty INTRA VIA Container with PVC Ports (Sterile, nonpyrogenic fluid path) ,500 ml. Admixture Products and Accessories |  |  | September 16, 2014 | 2014 | Baxter Healthcare Corp | Particulate matter found inside the fluid path. | Equipment maintenance | Baxter issued a letter to their consignees stating the following:Baxter is requesting that you take the following actions:Action to be taken if product was purchased directly from Baxter.1. Locate and remove all affected product from all locations in your facility(the product code and lot numbers can be found on the individualproduct package or shipping carton).2. Contact Baxter Healthcare Center for Service to arrange for return andcredit. The Center for Service can be reached at 888-229-0001between the hours of 7:00 am and 6:00 pm Central Time. Please haveyour Baxter eight-digit ship-to account number ready when calling.3. Complete the enclosed customer reply form , and return it to Baxter byeither tax or scanned e-mail.4. If you distributed any of the affected products to other facilities, pleaseconduct a recall with your end-user customers in accordance with yourcustomary procedures. Baxter distributed product 2B8011 lotU R 13D 15112 to customers between 4/26/2013 and 6/20/2013. Baxterdistributed product 2B8013 lot UR13K1 4095 to customers between 11 /27/2013 and 3/ 10/2014.Action to be taken if product was received from an alternate source.1. Locate and remove all of the affected products from all locations inyour facility (the product code and lot numbers can be found onthe individual product package or shipping carton).2. Contact Baxter Healthcare Center for Service to arrange for returnand credit. The Center for Service can be reached at 888-229-0001 between the hours of 7:00 am and 6:00 pm Central Time.Please have your ship-to account number ready when calling.3. Follow your supplier's reply and recall process. Please do not return the customer reply form to Baxter. |  | 20448 | 20448 | Nationwide Distribution including Puerto Rico, and the country of Canada. | Terminated | June 21, 2017 | 1009 | 114672 | Not_Computer
Z-0166-2015 | 69408 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | CADD-Solis Medication Safety Software, Administrator CD, Version 3.1, REF 21-2194-0301-01. |  |  | September 29, 2014 | 2014 | Smiths Medical ASD, Inc. | Smiths Medical has identified an issue with a single batch (Lot Number 2752712) of CADD&quot;-Solis Medication Safety Software Administrator 3.1 CDs. CADD&quot;-Solis Medication Safety Software Administrator 3.1 CDs were sent with the CADD&quot;-Solis Medication Safety Software Point of Care 3.1 software loaded on them. | Process control | Consignees were hand delivered by Smiths Medical Sales representatives on 9/29/2014 a Smiths Medical "Urgent Field Safety Notice" dated 25-September-14. The letter described the problem and the product being recalled. The letter also provided the Advice on Action to be Taken by the User, Transmission of this Urgent Medical Device Recall Notice and requested consignees to complete and return the Confirmation Form by fax to 1-800-237-8033 or via email to recall.response@smiths-medical.com. For questions customers can contact Smiths Medical's Customer Service Department at 1-800-258-5361. |  | 24 | 24 | US Nationwide Distribution in the state of AK, CA, CO, CT, GA, MI, MN, NH, NC, SC, TX, VA, WI. | Terminated | February 25, 2015 | 149 | 24 | Computer
Z-0037-2015 | 68777 | No | KNH | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | The Hulka Clip is a sterile packaged medical device that carries a two year expiration date. It is packaged two clips per procedural tray. Theses procedural trays are then sealed in a Tyvek/Mylar pouch and packaged in a cardboard shipping carton with ten pouches in each shipping carton. |  |  | July 09, 2014 | 2014 | Richard Wolf Medical Instruments Corp. | Richard Wolf Medical Instrument Corporation is recalling Hulka Clip Tubal Occluding Device lot 51000232. The devices are being recalled as a part of an investigation into a biocompatibility test discrepancy where the recalled lot (lot 51000232) passed final release testing, but the subsequent lot failed biocompatibility testing on a component that was also used to manufacture lot 51000232. | Nonconforming Material/Component | "Urgent: Medical Device Recall" letters dated July 9, 2014 were sent to all direct accounts (customers). The letters included instructions for customers to: 1) inspect their inventories for the recalled products; 2) contact Richard Wolf Medical Instruments at 800-323-9653 x 225 or 380 to notify Richard Wolf Medical Instruments if you have any devices in inventory; 3) review your records and report the number, if any, of Hulka Clips from lot 51000232 that were implanted in your patients on the Recall Response Form; 4) complete and return the Recall Response Form attached to the letter; and, 5) a RMA number will be provided to customers returning the recalled products and those customers will receive a credit. All returned product should be sent to the attention of Ms. Lisa Williams, Regulatory Specialist. Customers with questions can contact Richard Wolf Medical Instruments at 800-323-9653 x 225 or 380. |  | 73 | 73 cartons (1,460 clips) | Worldwide distribution: US (Nationwide) including states of: (AK, AL, CO, CT, FL, GA, IA, IN, KY, LA, MI, MO, MS, NC, NY, OH, OR, PA, SC, TN, TX, VA, WA, WV; and country of: Germany. | Terminated | April 24, 2017 | 1020 | 73 | Not_Computer
Z-0133-2015 | 69395 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | SAM Junctional Tourniquet (SJT) Auxiliary, Model SJT 102 and SJT 112. |  |  | August 14, 2014 | 2014 | Sam Medical Products | SAM Junctional Tourniquet (SJT) Auxiliary strap, model SJT 102 and SJT 112, is recalled because the snap hook on auxiliary strap broke during training demonstration. | Component design/selection | SAM Medical Products issued the press release on August 29, 2014 and the link is http://www.ireachcontent.com/news-releases/sam-medical-products-implements-voluntary-recall-of-its-accessory-axilla-strap-for-the-sam-junctional-tourniquet-273239381.html. SAM Medical Products sent the "URGENT: MEDICAL DEVICE RECALL" letter, dated September 10, 2014, to their consignees.The 5 additional lot numbers (7/1/14, 7/2/14, 8/4/14, 8/7/14, and 8/11/14) that were erroneously placed on the recall notification letter and they are NOT part of the recall.SAM Medical Distributors will need to notify their customers via this letter and the included FAX Back document of the actions being taken as well as ship the replacement straps to individual customers.Distributors will document and report to SAM Medical that all customers have been notified and replacement straps delivered to these customers. Distributors will notify SAM Medical that the replaced straps have been destroyed. This information will be documented on the included FAX Back document and then returned to SAM Medical (via FAX, e-mail, or mail, perinstructions on the FAX Back document).Customers will receive one complete Auxiliary Accessory Strap at no charge with a reinforced metal clip as a replacement. Once received, the existing strap should be removed from the SJT and replaced with the new strap. Customers/End Users need to complete and return the FAX Back document to the distributor from whom the FAX Back was sent.Customer can contact SAM Medical Customer Service at:-Email customerservice@sammedical.com-Phone (503) 639-5474, Monday through Friday, 7:30-4:30, Pacific Time-Toll free l-800-818-4726-Website: vww.sammedical.com |  | 3633 | total 3633 units (3342 units in US and 291 units Outside US) | Worldwide Distribution: US (nationwide) in states of Ohio, Oregon, New Mexico, North Carolina, Texas and to the following countries: Australia, Canada, Hong Kong, Iceland, Japan, Lebanon, Singapore, Thailand, and United Kingdom. | Terminated | July 21, 2015 | 341 | 3633 | Not_Computer
Z-0164-2015 | 69430 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Siemens Calibrator E; SMN 10309079 (2 pack), SMN 10321075 (6 pack), SMN 10335532 (6 pack Ref) |  |  | September 26, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Positive bias for Calibrator E lot kits ending in 38 and 39 with the ADVIA Centaur Systems Cortisol and Progesterone assays compared to the respective master curves. | Device Design | An Urgent Medical Device Correction (UMDC) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on September 26, 2014. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on September 26, 2014. These notices inform the customer of the positive bias with ADVIA Centaur Systems Cortisol and Progesterone when calibrated with Calibrator E kit lots ending in 38 and 39.Customers may continue use of the affected lots as the clinical utility of the assay is not impacted. The communications also provide information on the performance of the assays when they use Calibrator E kit lots ending in 40. The notices should be reviewed with the facility's laboratory or Medical director.All affected domestic customers were sent a hard copy of the Urgent Medical Device Correction via Federal Express. Fax-back forms will be used to confirm that the customers received the UMDC. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Medical Device Correction and completed the required actions.All affected customers outside the US will be provided a copy of the Urgent Field Safety Notice via hard copy, e-mail, and/or fax as determined by each country's local regulations and procedures. Each customer will be required to return a completed confirmation fax-back to Siemens. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Field Safety Notice and completed the required actions. |  | 6 | Total = 19,352 (13,881 - 2 pack, 4878 - 6 pack, 593 - 6 pack REF) | Nationwide including Puerto Rico. Worldwide: Afghanistan, Armenia, Austria, Bahrain, Belgium, Bulgaria, Burkina Faso, Canary Islands, Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Georgia, Germany, Great Britain, Greece, Hungary, Iran, Iraq, Ireland, Italy, Jordan, Kazakhstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Malta, Morocco, Netherlands, Norway, Oman, Pakistan, Poland, Portugal, Qatar, Romania, Russian Fed., Saudi Arabia, Serbia, Slovakia, Spain, Sweden, Switzerland, Turkey, Unit.Arab Emir., Vatican, Yemen, AR, AU, BD, BR, CA , CH, CL, CN, CO, DO, EC, EG, GT, HK, ID, IE, IL, IN, JP, KR, LY, MM, MX, MY, NI, NP, NZ, PE, PH, PY, SA, SG, SV, TH, TW, UY, VN, ZA | Terminated | January 29, 2016 | 490 | 6 | Not_Computer
Z-0125-2015 | 69326 | No | MIH |  |  | Cardiovascular | PMA | Ovation Prime Abdominal Stent Graft System - 29mm Aortic Body stent grafts. |  |  | September 20, 2014 | 2014 | Trivascular, Inc | TriVascular has identified select lots of Ovation Prime 29mm Aortic Body stent grafts that may be associated with rapid emptying of the fill polymer syringe when it is connected to the Autoinjector during the fill polymer injection step of the primary implant procedure. | Under Investigation by firm | Urgent Field Safety Notice letter was sent to all consignees on 9/19/14 via trackable mail system. Users are notified that TriVascular has identified select lots of Ovation Prime 29mm Aortic Body stent grafts that may be associated with rapid emptying of the fill polymer syringe when it is connected to the Autoinjector during the fill polymer injection step of the primary implant procedure. |  |  | 0 Devices | International Distribution Only -- Canada, Turkey, Austria, Ireland, Belgium, Italy, Switzerland, Spain, France, and Cyprus. | Terminated | December 04, 2014 | 75 | 0 | Not_Computer
Z-0169-2015 | 69463 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Elite XPL (Pulsed Light) Handpiece for the Cynosure Elite MPX Laser System |  |  | October 09, 2014 | 2014 | Cynosure, Inc. | If the laser system is turned on and the user adjusts, installs or removes the XPL pulsed light handpiece, there is a potential risk of an electrical shock | Device Design | Cynosure sent an Urgent Medical Device Recall Notification dated October 9, 2014, to all affected customers. They were also notified by phone. The letter identified the product the problem and the action needed to be taken by the customer.There are two methods for returning the Elite XPL handpieces. If the handpiece is NOT currently connected to the laser console, the customer will be asked to return it to the Cynosure office. If the handpiece is not connected there is no risk.If the handpiece is currently connected to the laser console, the customer is instructed to immediately discontinue use of the laser system and to turn off the system immediately. They are also being instructed to leave the handpiece attached to the console and to not touch it. A Cynosure Field service engineer will then come to the office and remove the handpiece and return it to the Cynosure factory. In addition, the Cynosure Field Service engineer will install an interlock into the connector receptacle on the laser console. This interlock plug is an existing component that is used on systems that do not have an IPL handpiece in place.If your XPL pulsed light handpiece is not attached to the laser system please call Cynosure customer service (1-877-222-6822) to arrange for the handpiece to be returned and repaired. |  | 322 | 322 units | Worldwide Distribution - Nationwide Distribution and the countries of : Australia, Canada, France, Germany, Italy, Netherlands, Ukraine and UK. | Terminated | March 26, 2015 | 168 | 322 | Not_Computer
Z-0168-2015 | 69583 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | COMPASS, Model No. SW Version 3.1, catalogue number CS10-100 (medical Linear accelerator) radiological |  |  | September 25, 2014 | 2014 | Iba Dosimetry Gmbh | Error in the software. A deviation between reconstructed and planned dose distribution may not be detected prior to treatment and this can result in an over or under-estimation of the pretreatment delivered dose. | Software design | The field safety notice including a confirmation of receipt was sent to the affected hospitals; by email and UPS letter on 09/25/2014. |  | 124 | 124 units | Worldwide Distribution-Distributed in the US including Guam and the countries of Germany, Austria, France, UK, Poland, Finland, The Netherlands, Spain, Turkey, Russia, and South Africa, Algeria, Mexico, Brazil, Venezuela, Columbia, Panama, China, Japan Philippines, Korea, India, Singapore, Thailand, Taiwan, Hong Kong, and Malaysia. | Terminated | January 26, 2015 | 123 | 124 | Computer
Z-0172-2015 | 69382 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Stryker SmartLife Large Aseptic Housing |  |  | October 08, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | The SmartLife Aseptic Housings are being recalled in that the top section of the housing may separate from the bottom section due to insufficient welding. | Process control | On October 8, 2014, Stryker Instruments issued "URGENT MEDICAL DEVICE RECALL NOTIFICATION" to their consignees. Notification included affected product description, reason for recall, risk to health, and actions to be taken by the customer/user. The customer/user were instructed to immediately review this recall notification, check all stock areas; discontinue use of any product which shows sign of cracking or damage at the weld; and complete and return the enclosed Business Reply Form (BRF) via fax to Stryker Instruments Regulatory Department, 866-521-2762 or scan and email a copy to kara.spath@stryker.com. Upon receipt of completed BRF, a replacement product will be shipped and the customer/users, upon receipt of replacement product, are to remove and return all recalled product from use and replace them with new product. For questions regarding this recall, please contact Stryker Instruments: at 269-389-4518 or email: kara.spath@stryker.com. |  | 3138 | 3,138 in total | Worldwide Distribution: US (nationwide) WI, MD, NY, NM, AZ, GA, CA, MN, DE, PA, TX, MT, MI, KS, KY, WY, IL; and countries of: Australia, Canada, Sweden, Netherlands, France Switzerland, India, Japan, Poland, South Africa, England, and Hong Kong. | Terminated | May 31, 2016 | 601 | 3138 | Not_Computer
Z-0178-2015 | 69535 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Cardiopulmonary bypass connector; P/N: 321334-000 DESC: CONNECTOR, 3/8 X 3/8 X 1/2 CLEAR PAY LOT: 295424 STERILE Expiration May 2019 Manufacturing Date May 2014 |  |  | September 18, 2014 | 2014 | Dokitz NovoSci | The single leg of the PAY connector has been observed to become detached. | Process control | The recalling firm notified the affected customers via letter and requested the affected lot be returned to NovoSci. |  | 240 | 240 units | MO, TX, PA, IL, MD, UT | Terminated | May 15, 2015 | 239 | 240 | Not_Computer
Z-0177-2015 | 69512 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | St. Jude Medical, Fixed Curve Response 6F Electrophysiology Catheter, CSL, Model 401353 |  |  | October 14, 2014 | 2014 | St Jude Medical | St. Jude Medical is performing a voluntary recall of batch 4671318, model number 401353 of the Response Fixed Curve DECA 6F 65cm CSL 2/8 Catheter. Specifically, the firm has identified that every device packaged within batch 4671318 contains the incorrect catheter (Supreme FC QUAD 5F catheter) within the packaging. To date, no serious injuries have occurred as a result of this issue | Labeling mix-ups | St. Jude Medical sent an "Urgent Medical Device Recall Notice" letter dated October 7, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Advised customers to discontinue the use of the recalled device and inform them that a St. Jude Medical sales representative will contact them in the near future to facilitate removal and reconciliation of the recalled device. For questions they can contact their St. Jude Medical Sales Representative. For further questions please call (651) 756-2295. |  | 26 | 26 | US Distribution including the states of CA, GA, NJ, NY, OR and PA. | Terminated | December 10, 2014 | 57 | 26 | Not_Computer
Z-0174-2015 | 69454 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Straumann Bone Level Implant 04.1 mm RC, SLActiveÂż 8mm, RoxolidÂż, Loxim |  |  | September 24, 2014 | 2014 | Straumann USA, LLC | Product transfer piece was fitted with a Narrow Connect (NC) instead of a Regular Connect (RC) | Packaging process control | Straumann USA sent a Field Corrective Action Notice letter dated September 24, 2014 and delivered via United Parcel Service. The letter identified the affected product, problem and actions to be taken. Straumann is requesting the return of unused implants. Returns and Customer Confirmation form is to be returned. Contact firm 781-747- 2514. |  | 111 | 111 units | US Nationwide Distribution | Terminated | June 26, 2017 | 1006 | 111 | Not_Computer
Z-0179-2015 | 68079 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Thyroglobulin (1D4) is a mouse anti-human monoclonal antibody produced as a tissue culture supernatant, and supplied in Tris buffered saline with carrier protein, containing 0.35% ProClin 950 as a preservative. Total volume = 7 mL. |  |  | April 18, 2014 | 2014 | Leica Microsystems, Inc. | Bond Ready to Use Primary Antibody Thyroglobulin does not function as intended up to the expiry date on the product labelling. This was detected through an on-going stability program. There is a link between the age of the product and staining intensity. | Unknown/Undetermined by firm | Leica Biosystems sent a Medical Device Recall Notification letter dated April 18, 2014, to all affected customers. The letter included instructions for customers to: 1) assess any results previously obtained by using the recalled products; 2) identify and destroy any unused or partly used recalled products or confirm that the recalled products are no longer in stock; and, 3) complete and return the attached Medical Device Recall Notification Acknowledgement Form. Customers with questions were instructed to contact Leica Microsystems by e-mail at LMGRA@leica-microsystems.com.For questions regarding this recall call 847-405-5413. |  | 180 | 180 units | Worldwide Distribution - USA including AR, AZ, CA, CO, CT, FL, GA, IL, KY, LA, MA, MD, MI, MN, MO, MT, MY, NC, NH, NJ, NY, OH, ON, OR, PA, TN, TX, UT, VA, WA, WI and the District of Columbia, Internationally to Belgium, Bulgaria, France, Italy, Pakistan, Philippines, Russia, Saudi Arabia, Spain, Turkey, and the United Kingdom. | Terminated | August 14, 2017 | 1214 | 180 | Not_Computer
Z-0175-2015 | 69437 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Access 2 Immunoassay Systems, Catalog No. 81600N, 386220, Instructions for Use Part No. B14253A. |  |  | October 06, 2014 | 2014 | Beckman Coulter Inc. | Beckman Coulter is recalling the Small Form Factor Console PC (SFF PC) (B23083) for use with the Access 2 Immunoassay systems because it may experience a "MFC Exception" error during normal operation of the Access 2 Immunoassay Systems. | Vendor change control | An Urgent Medical Device Recall Letter dated 10/6/14 was sent to all customers who purchased the Small Form Factor Console PC for use with the Access 2 Immunoassay Systems Family. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Customer Support Center at http://www.beckmancoulter.com/customersupport/support or call (800) 854-3633 in the US and Canada. Outside the US and Canada, customers are instructed to contact their local Beckman Coulter representative. |  | 1519 | 1,519 units total (403 units total) | Worldwide Distribution - US Nationwide in the countries of Algeria, Argentina, Australia, Austria, Bangladesh, Belarus, Bosnia-Herzegovina, Botswana, Brunei Darussalam, Bulgaria, Canada, China, Colombia, Costa Rica, Croatia, Czech Republic, Ecuador, Egypt, El Salvador, Ethiopia, France, Germany, Gibraltar, Greece, Guam, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Kenya, Korea, Lebanon, Libyan Arab Jamahiriya, Malaysia, Mexico, Morocco, Namibia, Pakistan, Palestinian Territory, Panama, Peru, Philippines, Poland, Portugal, Romania, Russian Federation, Saudi Arabia, Serbia, Slovakia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Kingdom, United States including Puerto Rico, Venezuela, and Vietnam. | Terminated | June 29, 2015 | 266 | 1519 | Not_Computer
Z-0183-2015 | 69480 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zeego systems, model number 10280959, with serial numbers 1000-1599 : Artis zeego is a dedicated angiography system developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed include cardiac angiography, neuro angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | September 23, 2014 | 2014 | Siemens Medical Solutions USA, Inc | A cable inside the C-arm systems may have been routed in an inappropriate manner. Improper routing of the cable may result in increased wear over time. Without additional measures, the potential exists for a cable inside the C-arm to break which could result in restricted functionality or system failure and an ongoing procedure could be terminated. Additionally, a subgroup of affected C-arm systems will receive a replacement of the cable harness. | Device Design | Siemens sent an " Important Safety Information " letter dated September 23, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Siemens will, without charge, remedy the defect or bring the product into compliance with each applicable Federal standard in accordance with a plan to be approved by the Secretary of Health and Human Services.Our service organization will contact with you to schedule an appointment to perform the corrective action. Please feel free to contact our service organization for an earlier appointment.Please observe this safety notice and comply with the corresponding measure until the update has been fully completed. For further questions, please call (610) 219-6300. |  | 164 | 164 | Nationwide Distribution including Puerto Rico. | Terminated | May 08, 2017 | 958 | 164 | Not_Computer
Z-0181-2015 | 68723 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Philips Medical Systems Allura Xper FD20C Radiological Imaging with software version R8.2.O; System Code: 722028. |  |  | July 08, 2014 | 2014 | Philips Medical Systems, Inc. | Faulty Automatic Motion Controller (AMC), a problem in the Power On Self Test (POST) error handling was detected, can result in a hazardous movement of the C-arc. system. | Software design | Philips Medical Systems issued certified letters on 7/8/14 and receipt returned for tracking purposes. A mandatory Field Change Order with reference FC072200253 is being released that requires Philips field service engineers to install Software release R8.2.0.3. If you need any further information or support concerning this issue, contact your local Philips representative: SUPPORT 1-800-722-9377. |  | 22 | 22 systems | Distributed in the states of IL, and TN, and the countries of Argentina , Australia, Belgium, Denmark, Egypt, France, Germany, Hong Kong, Italy, Japan, Korea, Republic of Netherlands, Norway, Saudi Arabia, Switzerland, United Kingdom, and Viet Nam. | Open, Classified |  |  | 24 | Computer
Z-0191-2015 | 69464 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | BardÂż PerFix Light Plug |  |  | October 08, 2014 | 2014 | Davol, Inc., Subs. C. R. Bard, Inc. | The patient record peel-off label product code and size did not match the other labeling and actual product in the packaging for the Bard Small PerFix Light Plug. | Error in labeling | Davol Inc. (Subsidiary of C.R. Bard, Inc)sent Customer Notification letters on Friday, October 10, 2014, via FedEx next day delivery with signature confirmation to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Action required:Examine your inventory and identify any product subject to this communication. A sample of the product labeling for the lot has been enclosed to assist in product code and lot number identification. The enclosure also depicts the supplied label as well as an image of the label with the correct information.If you have any remaining inventory, a copy of this communication can be kept with the units for future medical record documentation purposes.If you have already used some of the product lot, you may amend any patient record(s) based upon the content of this communication.Please complete and return the accompanying Effectiveness Check Form attached to this letter regardless of whether or not you have any remaining units of the affected product lot.If you may have further distributed or transferred this product, please identify the respective organizations and notify them of this product communication. Your notification to these organizations may be enhanced by including a copy of this product communication letter.This product communication is being made with the knowledge of the Food and Drug Administration. We apologize for any inconvenience this may have caused your organization. Please contact the Davol Customer Service Department at 1-800-556-6275 or C.R. Bard Medical Services &amp; Support Department at 1-800-562-0027 if you have any questions. |  | 78 | 78 units | US Distribution including the states of CA, NM, OH, DE, MN, UT, MD and MO. | Terminated | December 23, 2015 | 441 | 78 | Not_Computer
Z-0196-2015 | 69472 | No | GDC | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | The 5085 is a general surgical table with high patient weight capacity, extended width capability, and lower minimal table top elevation. |  |  | October 10, 2014 | 2014 | Steris Corporation | The 5085 and 5085SRT stainless steel column shrouds may become damaged as a result of misalignment during extreme table articulation, excessive side force on the table column, or pressure on the base of the table. Severe damage to the shrouds may prevent the user from raising or lowering the surgical table. | Device Design | STERIS will mail Consignee Notification Letters to the affected Consignees via FedEx with tracking numbers for delivery confirmation. The Consignee Notification mailing will initiate the week of October 13, 2014. |  | 999 | 999 | US distribution: AL, AZ, AR, CA, CO, CN, DC, FL, GA, HI, ID, IL, IN, IA, KN, KY, LA, ME, MD, MI, MN, MS, MO, NB, NV, NY, NC, ND, OH, OK, OR, PN, RI, SC, SD, TN, TX, VA, WV, WI. Foriegn: Austrialia, Brazil, Canada, China, Ecuador, India, Korea, Mexico, Pnama, and Thailand. | Terminated | June 26, 2015 | 259 | 999 | Not_Computer
Z-0192-2015 | 69474 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | OECÂż UroView 2800. |  |  | October 03, 2014 | 2014 | GE OEC Medical Systems, Inc | GE Healthcare initiated a field correction because prior revision of components were potentially installed on certain OEC 9800, FlexiView 8800, OEC MiniView 6800 and OEC UroView 2800 systems. | Component design/selection | On October 3. 2014. GE Healthcare Surgery began distributing an Urgent Medical Device Correction letter tocustomers to inform them of the issue and provide instructions for immediate mitigation of the issue. The letterwas sent to 3 titles at each facility affected: 1.1 Facility Administrator. 2.1 Director/Manager of Radiology andRadiologists, 3.) Radiology Department. Firm will track allservice requests for the field action to completion. These service requests will document the inspection and component replacement (if applicable) on the affected units. Exceptions to completion of the service request will be recorded. Records of field action completion will be retained per our processes and procedures. No product is being returned. |  | 2 | 2 | Worldwide Distribution. | Terminated | March 23, 2015 | 171 | 344 | Not_Computer
Z-0170-2015 | 69540 | No | LON | Fully automated short-term incubation cycle antimicrobial susceptibility system. | Microbiology | Microbiology | 510(k) | Rapid Neg BP Combo Panel Type 3; Catalog number B 1017-117. |  |  | October 17, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | An increase in false positive susceptible results on Rapid Neg BP Combo Panel Type 3 and Rapid Neg Urine Combo Panel Type 1 used with MicroScan Microbiology Systems. | Nonconforming Material/Component | Customer Recall notification letters were sent on October 17, 2014. Letters recommend that any remaining inventory be discarded and that consideration be given to reviewing previous test result, conduct patient follow up and/or repeat testing if isolates are still available.A second letter clarifying the recall was issued November 7, 2014 which provided additional information as requested by CDRH. |  | 1381 | 1381 | Distributed US (nationwide) including the states of IA, CA, NY, WI, MI, SD, NC, FL, HI, OH, WA, TX, WA, IN, NM, MS, KY, IL, UT, ND, and WY and the country of Canada. | Terminated | June 29, 2015 | 255 | 3620 | Not_Computer
Z-0198-2015 | 69397 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona Stemmed 5 Degree Cemented Tibia |  |  | October 09, 2014 | 2014 | Zimmer, Inc. | Cleaning process validation failure. | Nonconforming Material/Component | Zimmer sent an Urgent Medical Device Recall letter dated October 9, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to: 1. Review the notification and ensure affected personnel are aware of the contents. 2. Assist your Zimmer sales representative with the quarantine of any affected product. 3. Your Zimmer sales representative will remove the recalled product from your facility. 4. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-800-348-2759. Hours of operation are Monday through Friday, 8 a.m. through 8 p.m. EST. |  | 138 | 138 units | US Nationwide Distribution and countries of Austria ,Belgium, Switzerland, Germany, France, United Kingdom, Italy and South Africa. | Terminated | June 24, 2015 | 258 | 138 | Not_Computer
Z-0197-2015 | 69596 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Integra Titan Reverse Shoulder System right and left head cutting templates, 142.5 degrees. |  |  | October 24, 2014 | 2014 | Integra LifeSciences Corp. | A single lot of left and a single lot of right Reverse Shoulder System cutting templates were manufactured incorrectly. Specifically, the threaded handle has been welded backwards on the template resulting in the cutting angle being the reverse of what it should be. | Nonconforming Material/Component | The recalling firm will notify their consignees via letter delivered by traceable courier service. The consignee will be asked to complete and return a form to the recalling firm stating if they have affected product or not. |  | 12 | 12 units | Nationwide Distribution including the states of TX, CA, OH, MS, VA, IL, TN, GA, CA, FL, and AZ. | Terminated | July 21, 2015 | 270 | 12 | Not_Computer
Z-0200-2015 | 69521 | No | JSK | Supplement for culture media. | Microbiology | Microbiology | 510(K) Exempt | Oxoid Legionella BCYE Growth Supplement SR0110A, packaged in 100 ml vials, 10/pkg. |  |  | October 17, 2014 | 2014 | Remel Inc | Recalled product may not provide adequate recovery of Legionella pneumophila when used as directed and may result in false negative identification of Legionella pneumophila. | Under Investigation by firm | Thermo Fisher sent an Urgent Medical Device Recall letter dated October 17, 2014, to all affected customers. |  | 160 | 16/10/100 ml, vial, pkgs | Distributed in the states of ND, VT and TX. | Open, Classified |  |  | 160 | Not_Computer
Z-0199-2015 | 69271 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | ATTUNE Intuition Distal Femoral Jig; sold non-sterile; product is wrapped in bubble wrap and then sealed inside a polyethylene bag |  |  | October 15, 2014 | 2014 | DePuy Orthopaedics, Inc. | Issuing a device correction because if the pin bushing is over loaded or inverted, there is the potential for partial or full displacement and fracture of the plastic around the pin bushing. | Process control | DePuy sent an "URGENT DEVICE CORRECTION" notifications dated October 15, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, please call 574-371-4917. |  | 3082 | 3,082 | Worldwide Distribution - USA (nationwide) except AK, including Puerto Rico and the countries of Canada, Australia, Austria, Belgium, China, France, Germany, India, Ireland, Israel, Italy, Japan, Malaysia, Netherlands, New Zealand, Singapore, South Africa, South Korea, Spain, Switzerland, Thailand, UAE, United Kingdom and Chile. | Terminated | November 19, 2015 | 400 | 3082 | Not_Computer
Z-0201-2015 | 69251 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Licox Pt02 Monitor, Rx Only, Manufacturer: Integra LifeSciences (Ireland) Limited, IDA Business and Technology Park, Sragh, Tullamore, County Offaly, Ireland, Distributed by: Integra LifeSciences Corporation, 311 Enterprise Drive, Plainsboro, NJ 08536 |  |  | August 18, 2014 | 2014 | Integra LifeSciences Corp. | Integra LifeSciences Corporation received 2 complaints that the USB port on the Licox Pt02 monitor does not consistently provide the user the ability to extract the Pt02 trend data according to the User's Manual for the device. | Device Design | Integra LifeSciences sent an "Urgent Voluntary Medical Device Correction" letter/Return Response Form dated August 18, 2014, to their affected customers via traceable courier service. The letter identified the product the problem and the action needed to be taken by the customer.Complete the attached form. If you do not have an affected monitor, check the box: I do not have any monitors with an affected Serial Number listed below.If you do have affected monitors, check the box: I do have monitors with an affected Serial Number listed below. Circle all the Serial Numbersfor the affected monitors you have and record the total quantity.Complete the other information on the form and return the form by email or fax as indicated. Keep a copy of the form for your records.When your form is received, Integra will schedule this service with your facility and provide a temporary loaner monitor to minimize any disruption this action may cause.Should you have any questions, please contact Neuro Customer Service at 1-855-532-1723. We apologize for any inconvenience this may cause and thank you for your cooperation in this effort. |  | 43 | 43 (LCX02) and 17 (LCX02R) | US Distribution including the states of : CA, DC, FL, IL, IN, MI, MN, NC, NY, PA and TX | Open, Classified |  |  | 43 | Not_Computer
Z-0202-2015 | 69493 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | GMK Intramedullary Extension Rod 100 mm is a rod to obtain correct leg alignment during total knee replacement. |  |  | September 22, 2014 | 2014 | Medacta Usa | Medacta has registered in its database 3 cases of breakage of the GMK Intramedullary ExtensionRod 100 mm, code 02.07.10.0162. | Nonconforming Material/Component | Medacta sent an URGENT MEDICAL DEVICE letter dated September 22, 2014 to the sole US consignee. The letter included instructions for the consignee to: 1) check its inventory of the device; 2) quarantine and do not use the recalled device; and, 3) return the recalled device upon receipt of a RMA # from the Medacta QA Department. For questions regarding this recall call 805-910-6511. |  | 1 | 1 rod (US); 145 rods (global) | Worldwide Distribution - US including OH and Internationally to Australia, Austria, Belgium, France, Germany, Greece, Italy, Japan, Portugal, South Africa, Spain, and Switzerland. | Terminated | May 07, 2015 | 227 | 1 | Not_Computer
Z-0204-2015 | 69367 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | Transport Rollator with Padded Seat; Large 8-inch wheels with wheel locks (indoor and outdoor use), fold down footrests, ergonomic handles, curved padded backrest and handy wire basket for storage. Item# 30192 Burgundy each/Item# 30194 Blue each.. |  |  | October 01, 2014 | 2014 | Roscoe Medical Inc | The caster may separate completely from the device frame causing injury to user. | Device Design | The firm, Roscoe Medical, issued a "URGENT RECALL NOTICE-Timely Response Required" Quarantine Communication to customers between 09/08/2014 and 09/10/2014. The notice identified affected products as 30194, serial number range MH140407092MH140407291, the problem and actions to be taken. The customers were instructed NOT to use this product and to destroy any affected product that may be in their inventory/control; immediately complete and return the attached Recall Response Form/Certificate of Destruction to Roscoe Medical- "Transport Rollator Recall" within 15 days of receipt of the recall notification via fax to (440) 572-4261 or email: corporate@roscoemedical.com, even if you do not have any affected product; and notify your customers immediately of this product recall.Should you have any questions, contact Roscoe Medical Customer Support at (800) 376-7263 Monday-Friday 8:00 am EST - 5:00 pm EST. |  | 831 | 831 units | Nationwide Distribution: US including states of: AZ, CA, CO, FL, GA, IA, ID, IL, IN, KY, LA&lt; MA, MD, MI, MN, MO, NC, NJ, NM, NV, NY, OH, PA, SC, TN, TX, VA, WA, WI,and WY. | Open, Classified |  |  | 831 | Not_Computer
Z-0207-2015 | 69500 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical MILEX PESSARY KITS-Ring w/Support Folding Pessary- #1 |  |  | September 30, 2014 | 2014 | CooperSurgical, Inc. | Incorrect expiration date printed on the MILEX PESSARY KIT packaging | Incorrect or no expiration date | CooperSurgical notified customers a Recall Notification dated September 29, 2014, via UPS . The letter identified the product problem and the action needed to be taken by the customer. CooperSurgical's will replace any of the affected and provided further instruction on the disposition of the product at the time of replacement. If you have any further questions, contact 203.601.5200.. |  | 17 | 17 units | US Distibution in states of:OH, CA, WI, WA, FL, MN, NH, RI, MA, IL, MI, IA, and AL | Terminated | May 28, 2015 | 240 | 23 | Not_Computer
Z-0210-2015 | 69668 | No | FMT | Infant radiant warmer. | General Hospital | General Hospital | 510(k) | Fisher and Paykel Healthcare (FPH) IW900-Series Infant Warmer |  |  | November 06, 2014 | 2014 | Fisher &amp; Paykel Healthcare, Ltd. | Breakage of the nut that secures the heater head in place, which can cause the heater head to become partially detached and swing towards the bassinet. There is the potential that a heater head could contact a patient situated on the bassinet and cause a serious injury. | Manufacturing material removal | Fisher &amp; Paykel Healthcare sent an Urgent Medical Device Recall letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to:Step 1: Identify any affected Infant Warmers in your inventory by checking the REF and SERIAL number on the device label (see examples in Figures 1 and 2).Step 2: Identify any affected spare parts in your inventory by checking the REF and LOT number on the parts (see example in Figure 3).Step 3: Check your Infant Warmer service records to determine if affected spare parts were installed on any other Infant Warmers. If you are unable to identify specific Infant Warmers that may have had affected spare parts installed, then please contact FPH at the contact details below.Step 4: Replacement of Affected Nut. - Remove the affected Infant Warmer from use.- Switch the warmer off and disconnect it from the mains power supply.- Support the heater head during disassembly to prevent it from falling.- Follow Steps A to F in Figure 4 below to replace the nut.- Ensure the pivot washer remains in place when installing the new nut.- Ensure that the heater head is secure, can rotate smoothly, and the detent locates the head in the central position.- Reconnect the mains power supply, switch the warmer on, and ensure there are no error codes or alarms before returning to use.Step 5: Complete the Urgent Medical Device Recall Response Form and return it to your FPH Representative |  | 661 | Nationwide (U.S.): 661 Infant Warmers and 180 Spare Parts, Internationally: 2,451 Infant Warmers and 280 Spare Parts | Worldwide Distribution - USA (nationwide) and Internationally to Albania, Algeria, Australia, Austria, Bahrain, Bangladesh, Bekgium, Bermuda, Chile, China, Croatia, Dominican Republic, Estonia, Finland, France, Germany, Great Britain, Greece, Hong Kong, Iceland, India, Indonesia, Iran, Ireland, Israel, Jordan, Kenya, Kuwait, Latvia, Lebanon, Lithuania, Malaysia, Mauritius, Mexico. Morocco, New Zealand, Nigeria, Northern Ireland, Norway, Pakistan, Palestinian Territory, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, singapore, Slovenia, South Africa, Spain, Sri Lanka, Sultanate of Oman, Syria, Taiwan, Thailand, The Netherlands, Tunisia, Turkey, United Arab Emirates, Uruguay, Venezuela, Vietnam, and West Bank. | Terminated | April 27, 2015 | 172 | 661 | Not_Computer
Z-0213-2015 | 69504 | No | DQA | Oximeter. | Cardiovascular | Anesthesiology | 510(k) | Masimo SET uSpO2 Cable |  |  | September 30, 2014 | 2014 | Masimo Corporation | Firm has identified a small number of Oximetry Cables with crossed-internal wires. The effect on performance when these wires are crossed is that when a low SpO2 value is measured, a high SpO2 value would be displayed on the monitoring device. Additionally when a high SpO2 value is measured, a low SpO2 value would be displayed on the monitoring device. | Employee error | An "Urgent: Advisory Notice," dated 09/30/2014, was sent to customers. The letter informed customers of the problem identified and the actions to be taken. Customers were instructed to evaluate their inventory of Oximetry Cable(s) using the following method: Use the Oximetry Cable to take a SpO2 measurement on a healthy individual with an expected SpO2 value above 90% or review SpO2 measurements take with the Oximetry Cable and verify that measurement results were consistent with expected SpO2 values. If the SpO2 measurement(s) is consistent with expected SpO2 value, the Oximetry Cable is not affected and can continue to be used. Customers are instructed to complete and return Attachment 2, Tracking/Verification Form, to Masimo in order to document that the Oximetry Cable(s) is functioning as intented. If the assessment results in SpO2 measurement that is not consistent with expected SpO2 value, then the Oximetry Cable should be removed from service, segregated and returned to Masimo. Customers are instructed to call 1-800-326-4890 and select option 2 for Technical Services. Customers are to determine with Technical Services whether they will return the affected Oximetry Cable(s) to Masimo for replacement or refund. If returning for replacement customers are suppose to obtain a Return Material Authorization Number (RMA number). |  | 21754 | 21,754 units (21,222) in the U.S. and Puerto Rico) | Worldwide distribution including; nationwide (U.S. and Puerto Rico) and the countries of Argentina, Austria, Brazil, Canada, China, Denmark, France, Gernany, Hong Kong, India, Italy, Japan, Jordan, Kuwait, Mexico, Netherlands, Poland, Portugal, South Korea, Spain, Switzerland, Turkey, and United Kingdom. | Terminated | March 28, 2016 | 545 | 21754 | Not_Computer
Z-0211-2015 | 69557 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | Cardiohelp-I System |  |  | October 17, 2014 | 2014 | Maquet Cardiovascular Us Sales, Llc | Cardiohelp- i System may have a software issue that can potentially result in an erroneous display of a &quot;Battery Needs Service&quot; message after startup of the units when using either AC or DC power | Software Design Change | Maquet Inc. sent a Product Recall Urgent Medicall Device Field Correction letter to all of their customers via Federal Express. The letter identified the affected product, problem and actions to be taken. The letter stated that a software upgrade will be performed by an authorized Maquet Service Technician. A Maquet representative will contact customer to schedule an on-site service of the device. Any questions can be directed to the local Maquet representative or Maquet Customer Service at 1-888-627-8383 (option 2, option 2) Monday through Friday, between the hours of 8:00 am and 6:00 pm (EST). |  | 331 | 331 units | Worldwide Distribution - US Nationwide | Terminated | April 06, 2016 | 537 | 331 | Computer
Z-0209-2015 | 69536 | No | JCW | Penile inflatable implant. | Gastroenterology/Urology | Gastroenterology/Urology | PMA | AMS 700Âż Series Inflatable Penile Prosthesis Accessory Kit (AMS 700 Line Accessory Kit), 72401850. |  |  | October 10, 2014 | 2014 | American Medical Systems, Inc. | AMS' supplier has issued a recall of certain lots of Lightning Bolt Keith needles supplied in AMS 700 Accessory Kits. Based on this, AMS is recalling AMS 700 Accessory Kits which contain these lots of needles.The supplier made the decision to issue this recall after an internal investigation that was requested by AMS. The investigation identified a potential nonconformity that may lead to a bent | Nonconforming Material/Component | Consignees were sent via either Fed Ex mailing or hand delivery by AMS Field Personnel beginning on 10/10/2014 an AMS "Urgent Medical Device Recall" letter dated October 9, 2014. The letter described the problem and the product involved in the recall. Requested consignees to complete and return the Product Recall Acknowledgement Form. For questions consignees can contact AMS Customer Service at 1-800-328-3881 , FAX 1-800-366-9035, ore mail salesoperations@ammd.com. |  | 184 | 184 (110 US, 74 OUS) | Worldwide Distribution US (nationwide) including PR and the states of AL. AZ. AR, CA, CO, CT, IL, IN, KS, KY, MD, MA,MI, MN, MO, NY,NC, OH,OR, TN, TX, VA, and WA, and the countries of ARGENTINA, AUSTRALIA, BRAZIL, CANDA, GERMANY, FRANCE, GREAT BRITAIN, LEBANON, PANAMA, SOLOMON ISLANDS, SWEDEN, and COLOMBIA. | Terminated | September 28, 2015 | 353 | 184 | Not_Computer
Z-0205-2015 | 69488 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Launcher Coronary Guide Catheter; MODEL: LA6JR40K |  |  | October 13, 2014 | 2014 | Medtronic Vascular | Two reports have been received on the 6F launcher Guide Catheter lot (Model Number LA6JR40K; lot No 0007244836} for incorrect usable length of the Guide Catheter. It was reported that the catheters had a usable length of 90 cm instead of 55 cm, which was specified on the packaging label. Usable length is defined as the total length of catheter shaft from the distal end of the strain relief to the | Process control | Beginning October 13, 2014, Medtronic field representatives will being hand delivery of the customer communication to affected customers, notifying them of the issue, and requesting return of unaffected product. Customers will be asked to sign a certificate confirming they have received the notification and have taken the requested action.1. Identify and quarantine all listed, unused product in your inventory.2. Return all listed product in your inventory to Medtronic. Contact Customer Service at 888-283-7868 to initiate a product return and credit. Your Medtronic sales representative can assist you in the return of this product as necessary.3. Complete the attached Customer Confirmation Certificate and email it to RS.CFQFCA@Medtronic.com or fax it to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. |  | 11 | 11 devices | US Distribution including the states of NC, IA, IL and KS. | Terminated | February 18, 2015 | 128 | 11 | Not_Computer
Z-0212-2015 | 69588 | No | JJY | Quality control material (assayed and unassayed). | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE Calcitonin Control Module (IMMULITE, IMMULITE 1000, IMMULITE 2000, IMMULITE 2000 XPi) |  |  | October 20, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | A typographical error regarding the expiration date was identified in the control module instructions for use. | Error in labeling | Siemens Healthcare Diagnostics issued an Urgent Medical Device Correction letter October 17, 2014, to US customers, and an Urgent Field Safety Notice issued October 16, 2014, outside the US. The letter identified the product the problem and the action needed to be taken by the customer.Customers are advised to review the notice with their medical directors. Customers should discard the control module IFU listed in table 1 and refer to table 2 for the list of revised IFU versions as noted on the notice. Finally, the customers should check the system for the expiration date and if correct, no further action is required. Field Correction Effectiveness Check should be returned via fax (312) 275-7795 within 30 days. For further questions please call 914-631-8000. |  | 2212 | 2212 units (92 units domestically &amp; 2120 units internationally) | Worldwide Distribution - US Distribution including the states of FL, KY, MN, MA, PA, UT, NY, VA and CA., and the countries of Algeria, Australia, Argentina, Austria, Belgium, Brazil, Bulgaria, Cameron, Chile, China, Colombia, Costa Rica, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Great Britain, Greece, Hong Kong, Hungary, India, Iraq, Ireland, Israel, Italy, Jordan, Kazakhstan, Latvia, Macedonia, Mexico, Netherlands,Pakistan, Peru, Poland, Portugal, Rep. Korea, Romania, Russian Fed, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Tadjikistan, Taiwan,Turkey, Unit. Arab Emir and Vietnam. | Terminated | March 31, 2017 | 893 | 2212 | Not_Computer
Z-0219-2015 | 69685 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | 10 mm/135 degree Titanium Cannulated Trochanteric Fixation Nail 170 MM - Sterile: intended to treat stable and unstable fractures of the proximal femur |  |  | October 29, 2014 | 2014 | Synthes, Inc. | The 10 mm/135 degree Titanium Cannulated Trochanteric Fixation Nail 170 MM (Sterile) was assembled incorrectly. If the non conformance goes undetected, it may impact the locking function post-operatively and compromise the bone reduction and construct stability which may lead to non-union/malunion. | Employee error | An urgent medical device recall notice, dated October 29, 2014, was sent to end users that identified the product, problem and action to be taken. Customers were instructed to return the device to Synthes along with completing the verification section of the notice. |  | 6 | 6 | US Distribution in the states of: AK, FL, OH, OK, and TX. | Terminated | December 18, 2015 | 415 | 6 | Not_Computer
Z-0214-2015 | 69282 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA CentaurÂż and ADVIA CentaurÂż XP and ADVIA CentaurÂż CP Cleaning Solution (For in vitro diagnostic use with the ADVIAÂż Centaur systems) |  |  | September 12, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Stability of the ADVIA Centaur Cleaning Solution is not assured | Material/Component Contamination | Siemens issued an Urgent Medical Device Recall (UMDR) to customers in the US on September 12, 2014. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on September 12, 2014. These notices inform the customer of the degradation of hypochlorite inthese lots of cleaning solution. Customers were instructed to discontinue use of the affected lots and dispose of the product according to local regulations or procedures. The notices should be reviewed with the facility's laboratory or medical director. Fax-back forms will be used to confirm that the customers received the UMDR. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Urgent Medical Device RecallFor questions regarding this recall call 508-668-5000. |  | 4645 | 4645 (12 pk) | Worldwide Distribution - USA (nationwide) and Internationally to Canada. Grerat Britain, Saudi Arabia, Nepal, and Viet Nam. | Terminated | July 17, 2015 | 308 | 4645 | Not_Computer
Z-0216-2015 | 69487 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Transport Chair and Rollator in 1, Model HE21630011 Burgundy and HE21630011 Blue. |  |  | October 10, 2014 | 2014 | Silver Star Brands | Silver Star Brands is recalling the Transport Chair and Rollator in One because they have received several reports of the wheel assembly failing resulting in damaged wheels.. | Device Design | Consignees were sent on 10/10/2014 an Easy Comforts "Urgent: Rollator Recall" letter dated October 10, 2014. The letter described the problem and the product involved in the recall. Advised consignees to stop using the product immediately. The consignees can contact the Product specialist at Easy Comforts at 1-888-327-2390 Monday though Friday from 7:30 am to 3:30 pm central time. |  | 129 | 129 | Nationwide Distribution including the states of AL, AZ, AR, CA, CO, CT, FL, GA, HI, IL, IN, IA, KS, LA, MD, MI, MO, MT, NE, NV,NH, NM, NY, NC, OH,OK, and PA. | Open, Classified |  |  | 129 | Not_Computer
Z-0218-2015 | 69645 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | CyberKnife Robotic Radiosurgery System with the first generation IRIS Variable Aperture Collimator. |  |  | October 31, 2014 | 2014 | Accuray Incorporated | Software upgrade to correct potential safety issue related to CyberKnife System that occurs when upgrading the Treatment Delivery Software for the first generation Iris Variable Aperture Collimator. | Software design | Accuray sent an Urgent Safety Advisory Notification letter dated October 17, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check the radiaiton field size (beam width) for each Iris setting to ensure that it matches the beam commissioning data in the MultiPlan System to within the stated tolerance of 0.2 mm. Customers with questions were instructed to contact Accuray Customer Support at 1-877-668-8667 or visit the Accuray Technical Solutions Center at www.accuray.com or call North America: +1.877.668.8667, Europe: +800.4141.9595, Hong Kong +852.2247.8688, Japan +81.3.6269.9556.For questions regarding this recall call 408-716-4600. |  | 84 | 84 devices subject to correction. | Worldwide Distribution - USA including CA, MS, FL, AZ, DC, TX, CO, CT, IL, WI, RI, NV, OK, MA, AL, NY, MO, NC, TN, PA, DE, NJ, MI, IN, KS, OH. Foreign distribution to Turkey, Switzerland, Spain, Saudi Arabia, Russia, Poland, Mexico and Korea, Italy, India, Greece, Germany, France, England, Canada, and Belgium. | Terminated | February 20, 2015 | 112 | 84 | Computer
Z-0217-2015 | 69565 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Vent Catheters, . Packaged sterile and non-pyrogenic in a sealed, peel-type pouch. Model Numbers: E061, E060, PE062, EM012 |  |  | October 08, 2014 | 2014 | Edwards Lifesciences, LLC | Edwards Lifesciences is conducting a field action on the Vent Catheters due to the potential for disruption of the white plastic coating on the stylet. | Device Design | Edward Lifesciences sent an Urgent Medical Device Safety Notice date October 8, 2014, via Fed-Ex to all affected customers. On 3/4/15, letters were sent to surgeons. As part of the customer instructions in the safely notice, customers were requested to forward copies of the safely alert to any organization where the potentially affected devices have been transferred. No product is being returned. The distribution of the device has not been stopped as a result of this field correction. No additional production lots will be released for distribution until the IFU has been revised. Customers with questions were instructed to call (800) 424-3278.For questions regarding this recall call 801-553-7531. |  | 157305 | 157,305 | Worldwide Distribution - USA including Alabama, Arkansas, Arizona, California, Colorado, Connecticut, Florida, Georgia, Hawaii, Iowa, Idaho, Illinois, Indiana, Kansas, Kentucky, Louisiana, Massachusetts, Maryland, Maine, Michigan, Missouri, Mississippi, Montana, North Carolina, Nebraska, New Hampshire, New Jersey, Nevada, New York, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Vermont, Washington, Wisconsin, West Virginia. US Territory: Puerto Rico and Internationally to Australia, Bangladesh, Brazil, Canada, Chile, China, Colombia, Hong Kong, India, Japan, Malaysia, Mexico, Netherlands, Singapore, South Korea, Sri Lanka, Taiwan, and Venezuela. | Open, Classified |  |  | 157305 | Not_Computer
Z-0221-2015 | 69656 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MOSAIQ |  |  | November 03, 2014 | 2014 | Elekta, Inc. | A problem can exist in MOSAIQ resulting in the display of incorrect numeric data due to a dose rounding error on printed reports. | Software design | Elekta sent an Urgent Important Field Safety Notice letter to alert customers of the issue. This notice was distributed on November 3, 2014. A fix to the defect has been developed and has been released by Product Creation. For question call 408-830-8000. |  | 15 | 15 | US Nationwide Distribution in the states of AZ. CO, IL, MD, MO, NY, OK, OR, PA, RI, TX, VA | Terminated | July 27, 2015 | 266 | 15 | Not_Computer
Z-0222-2015 | 69531 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | MC-500 Multicolor Laser Photocoagulator using software version 2.20. |  |  | May 24, 2013 | 2013 | Nidek Inc | Retrospective review found that a Engineering Change Order initiated in May 2013 for a memory function problem was not reported as required. | Software design | Nidek Engineers visited all affected sites to install new software version.For questions regarding this recall call 510-353-7785. |  | 5 | 5 | US Distribution in the states of: CA, NE, and NY. | Terminated | November 17, 2014 | 542 | 5 | Computer
Z-0223-2015 | 69508 | No | LZS |  |  | Ophthalmic | PMA | Nidek Excimer Laser Corneal Surgery System, model EC-5000 CXIII. |  |  | August 06, 2012 | 2012 | Nidek Inc | Contact failure in F1 and F4 fuse holders was detected on the EC-5000 CXIII Excimer Laser Corneal Surgery System that could cause a serious adverse event for the patient. | Component design/selection | Nidek Engineers visited each site to perform correction - fuse holders have been inspected and replaced if necessary.An Engineering Change Order was issued on June 1, 2012, for inspection and/or replacement of affected fuse holders by Nidek Service Engineers.For further questions please call (510) 353-7785. |  | 14 | 14 units | US Distribution including the states of CO, NY, MO, LA, ME, CA, MI, AZ, CT and FL. | Terminated | November 17, 2014 | 833 | 14 | Not_Computer
Z-0224-2015 | 69537 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Nidek MC-500 Multicolor Laser Photocoagulator using software version 2.20. |  |  | October 24, 2012 | 2012 | Nidek Inc | When the MC-500 is turned on and coagulation is performed without changing the laser color, the delivery unit occasionally fails to emit the laser, or the power output is lower than the set program. | Device Design | Nidek Field Service engineers visited each affected location to upgrade software. |  | 21 | 21affected devices | US Distribution in the states of: NJ, SC, IL, MI, NY, CA, TX, OH, IN, UT, and MN. | Terminated | November 17, 2014 | 754 | 21 | Computer
Z-0225-2015 | 69693 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | The MectaLIF Oblique Handle, model number 03.22.10.0262, is used to insert the MectaLIF oblique spinal cage implants. |  |  | October 21, 2014 | 2014 | Medacta Usa | The MectaLIF Oblique Handle has the laser marking &quot;MEDIAL&quot; on the incorrect side and could result in incorrect insertion of the device during surgery. | Employee error | Medacta sent an URGENT MEDICAL DEVICE RECALL letter dated October 21, 2014 were via e-mail to all sales representatives and distributors. The letters included instructions for the sales represetantives and distributors to: 1) check their inventory for the recalled devices; 2) immedaitely sequester the recalled devices and do not use them during surgery; 3) if the devices have been further distributed, inform them of the recall; and, 4) replacement devices will be provided. Customers with questions can contact Medacta USA at 805-910-6511. |  | 32 | 32 oblique handles | Worldwide Distribution - US including CA, CO, ID and Internationally to Australia and Germany. | Terminated | March 14, 2016 | 510 | 32 | Not_Computer
Z-0232-2015 | 69579 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | EtestÂż Ceftaroline (CPT32) US B30. |  |  | October 16, 2014 | 2014 | Biomerieux | The products have a wrong expiration date on their labeling, 5 years instead of 2 years. | Labeling Change Control | Consignees of bioMÂżrieux, impacted by the identified issue, were notified of the FSCA and corrective action (FSCA notification) electronically on October 16th, 2014. Additionally, the bioMÂżrieux United States subsidiary notified US consignees of the FSCA notification via FedEx on October 23rd, 2014. |  | 652 | 652 devices | Worldwide Distribution-US (nationwide) including DC and the states of AL, AK, AR, AZ, CA, CO, DE, FL, GA, IA, IN, KY, MA, MD, MI, MO, NC, NE, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, and the countries of Australia, Austria, China, Czech Republic, France, Germany, India, Italy, Japan, Netherlands, Poland, Portugal, Switzerland, Sweden, Turkey, United Kingdom, Egypt, Pakistan, Slovenia and Uruguay. | Open, Classified |  |  | 652 | Not_Computer
Z-0227-2015 | 69514 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SideFire Directional Laser Probe, REF: NBP101-01, NBP201-01, NBP301-01, NBP401-01, NBP501-01. These laser probes are part of the NeuroBlate System, a unique collection of MRI-compatible laser devices and accessories that create an MRI guided delivery of precision thermal therapy in the brain. The NeuroBlate System components include Gas-cooled Laser Delivery Probes to deliver controlled energy to a target zone. |  |  | October 03, 2014 | 2014 | Monteris Medical Corp | Monteris Medical received a complaint of a separation of the distal sapphire tip from the laser delivery probe. | Process control | Monteris Medical sent an Urgent Advisory Notice of Product Recall letter dated October 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action to be taken by user:Our distribution records document that you were shipped products from the affected lots. Please check your inventory to confirm whether you have any stock affected by this notice. An acknowledgment form is included to assist you in the assessment of your inventory. Your Monteris Medical Sales Representative will be glad to assist you with this assessment. Once your inventory has been verified, please fax the form back to Monteris Medical at 763-253-4728. |  | 234 | 234 devices | US Distribution including the states of CA, CT, FL, GA, KS, MN, MO, NC, OH, PA, TX, and VA | Terminated | December 08, 2014 | 66 | 367 | Not_Computer
Z-0229-2015 | 69461 | No | NBC | B-type natriuretic peptide test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Centaur BNP Assay for in vitro diagnostic use in the quantitative determination of B-type Natriuretic Peptide (BNP) in human plasma using the ADVIA Centaur, Centaur XP, Centaur CP Systems; Catalog Numbers: |  |  | October 06, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc | Siemens Healthcare Diagnostics has confirmed that approximately 2000 ReadyPacks from the ADVIA CentaurÂż Systems BNP reagent contains an ADVIA Centaur Systems FT4 human-readable side label. | Labeling mix-ups | A Customer Notification (CN) was sent to all affected Siemens Healthcare Diagnostics customers in the United States on October 03, 2014 to be delivered to customers on October 06, 2014 via Federal Express. An Urgent Field Safety Notice (UFSN) was sent out for distribution to all customers outside the United States on October 06, 2014. These notices inform the customer of the potential to receive an ADVIA Centaur BNP ReadyPack with a FT4 human-readable side label. Siemens instructs customers to inspect their inventory and discard any BNP Readypacks with a FT4 side label. In addition, Siemens informs customers the end labels of the affected ReadyPacks correctly identify the pack as BNP and Siemens has confirmed the reagent within these packs is BNP Reagent. Therefore valid BNP results will be generated when calibrating and running quality control or patient samples. Fax-back forms will be used to confirm that the customers received the notification. For those customers that do not return a fax-back form, a Siemens representative will follow up with those customers to ensure they received the Customer Notification and completed the required actions. |  | 7078 | 100 test - 4,600 kits; 500 test - 1,878 kits | Worldwide Distribution - US Nationwide including Puerto Rico; and countries of: Canada, Austria, Belgium, Czech Republic, Denmark, Estonia, Finland, France, Fren.Polynesia, Germany, Great Britain, Greece, Guadeloupe, Italy, Latvia, Lithuania, Netherlands, Norway, Poland, Portugal, Reunion, Russian Fed., Slovakia, Slovenia, Spain, Switzerland, Turkey, Unit.Arab Emir., AR, AU, BR, CL, CN, CR, IN, JP, KR, LK, MY, SG, TW, UY, VN, and ZA. | Terminated | July 22, 2015 | 289 | 7078 | Not_Computer
Z-0416-2015 | 68780 | walker, mechanical | ITJ | Mechanical walker. | Physical Medicine | Physical Medicine | 510(K) Exempt | MediChoice Dual-Release Folding Walker with 5" wheel |  |  | June 25, 2014 | 2014 | Owens &amp; Minor Distribution, Inc. | Product wheels may be too loose, too tight, or fall off. | Employee error | Customers were notified via letter and email by June 27, 2014. to take the following actions: 1. Designate a quarantine location in the DC for the recalled lots and serial numbers. 2. Label this area very clearly as "RECALLED PRODUCT IN QUARANTINE". 3. Clearly segregate the area to prevent the product from being moved or handled and to further designate that the product is in quarantine. 4. Record each case lot# or walker serial number if the product is out of the case on the recall inventory report.5. Label each unit in the quarantine area with a label that clearly states RECALLED PRODUCT IN QUARANTINE. 6. Conduct a daily inventory verification check using to ensure the inventory still remains in quarantine. In addition, please forward a copy of this letter to hospital personnel in a position to notify patients who may have received and been discharged with a walker from the lots and serial numbers identified above. Hospital Customers: Contact your local DC Customer Service Representative who will assist in returning and replacing the recalled walker. Patients Who May Have Received a Recalled Walker from a Hospital: Contact Owens &amp; Minor Mainstreet Customer Service at 1-800-818-6664 who will assist you in returning and replacing a walker. |  | 4856 | 1214 cases; 4 units/case | US Distribution including the states of CA, CO, CT, GA, IL, MA, MD, ME, MI, NE, NH, NJ, NY, OH, OR, PA, SC, TX, UT, VA, VT, WA and WY. | Terminated | January 06, 2017 | 926 | 4856 | Not_Computer
Z-0235-2015 | 69601 | No | KPQ | Radiation therapy simulation system. | Radiology | Radiology | 510(k) | Siemens SimView NT and SimView NT with IDI simulation system components: |  |  | October 10, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Siemens became aware of a very rare occurrence where an intended movement of the SimView NT simulation system components could not be stopped with usual hand control commands or by use of the collision detection functionality.It is possible that simulation system components could collide with patient and/or system equipment. | Device Design | Customers were initially notified about the issue by letter on January 7, 2014 (this action was reported as Z-0786-2014 which has been terminated). Siemens promised to notify them when a fix became available. Customer letter, dated November 7, 2014, was sent to notify end users that a fix has become available and Siemens would contact them to schedule an appointment for the correction. |  | 6 | 6 | Worldwide Distribution - US Nationwide in the states of PA, KY, LA, MA, MI, and CA; and countries of: Argentina, Belarus, Bulgaria, Costa Rico, P.R. China, Curacao, Egypt, Repulic Korea, Germany, Greece, Hungary, Italy, Iraq, Jordan, Malaysia, Palestanian Ter, Poland, Thailand, Turkey, and Vietnam. | Terminated | June 29, 2017 | 993 | 6 | Not_Computer
Z-0420-2015 | 69434 | No | LCP | Glycosylated hemoglobin assay. | Hematology | Hematology | 510(k) | Siemens Dimension HB1C Flex reagent cartridge; Dimension Hemoglobin A1c (HB1C) Flex reagent cartridge on the Dimension clinical chemistry system is an in vitro diagnostic assay for the quantitative determination of hemoglobin A1c (HbA1c) in human anticoagulated whole blood. Measurements of hemoglobin A1c are effective in monitoring long-term glucose control in individuals with diabetes mellitus. |  |  | September 30, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Certain lots of the reagent cartridges show confirmed positive bias averaging 0.4% Hemoglobin A1c units and occasionally up to 1.0% HbA1c units for patient samples when compared to the National Glycohemoglobin Standardization Program (NGSP). QC samples may exhibit a similar bias. | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Correction letter dated September 2014 was provided to all Dimension VistaÂż HbA1c and DimensionÂż HB1C customers who received the affected lots to notify them of the issue, the potential risk to health, and actions to be taken by customers. Customers were instructed to discontinue use of and discard the affected devices.Customers need to review the letter with their Medical Director. Siemens will replace any unused inventoryof the affected lots at no charge. Customers were asked to complete the attached form and fax it to 302-631-7795. Customers should retain a copy of the letter with their laboratory records, and forward the letter to those who may have received this product. Customers with questions should contact their Siemens Customer Care Center - Technical Solutions or their local Siemens technical support representative.For questions regarding this recall call 800-441-9250. |  | 53046 | 53, 046 | Worldwide Distribution: US (nationwide) and internationally to: Canada. | Terminated | August 10, 2015 | 314 | 53046 | Not_Computer
Z-0418-2015 | 69632 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Patient Specific Instrument PERSONA CR Pin Guides (Femur and Tibia) is intended to be used as a surgical instrument to assist in the positioning of Total Knee Replacement components intraoperatively and in guiding the marking of bone before cutting provided that anatomic landmarks necessary for alignment and positioning of the implant are identifiable on patient imaging scans. |  |  | October 16, 2014 | 2014 | Materialise USA LLC | The box of guides for a specific case arrived intact but contained two (2) femoral guides instead of one (1) femoral guide and one (1) tibial guide. The device did not meet specification: One tibia guide missing and one femur guide contained an edge which was +0.04 mm over specification. | Nonconforming Material/Component | Materialise USA was notified on October 16, 2014 by the Field Representative of the distributor (Zimmer) that the box of guides for case JTOMA-JL-0636-L arrived intact but contained two (2) femoral guides instead of one (1) femoral guide and one (1) tibial guide. A new set of tibial and femoral guides were manufactured and sent to Zimmer (headquarters distribution center) on October 17, 2014 before the surgery date on October 20, 2014.For questions regarding this recall call 734-662-5057. |  | 1 | 1 | Nationwide Distribution to MN only | Terminated | February 17, 2015 | 124 | 1 | Not_Computer
Z-0424-2015 | 69366 | No | HWA | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Trabecular Metal Tibial Impactor Replacement Pads. |  |  | October 14, 2014 | 2014 | Zimmer, Inc. | To provide clarifying instructions relating to the assembly/disassembly of the Trabecular Metal Tibial Impactor with/from TM tibialImpactor Pads due to complaints regarding Tibal Impactor pad fracture. | Labeling design | On 10/14/2014, URGENT MEDICAL DEVICE CORRECTION notifications were sent to the affected distributors and hospital Risk Managers. All distributors were notified via electronic mail, and all hospital Risk Managers and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call the customer call center at 1-877-946-2761. |  | 4961 | 4961 units | Worldwide Distribution - US Nationwide in the state of AL, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, MA, MD, ME, MI, MN, MO, ND, NH, NJ, NV, NY, OH, OR, PA, SC, TX, VA, WI, NC, TN, OK, AZ, UT, AK, PR, DC. and the countries of: Germany, China, Canada, Australia, Japan, United Arab Emirates, Austria, Belgium, Switzerland, Germany, Denmark, Spain, Finland, France, United Kingdom, Greece, Iceland, Italy, Malta, Netherlands, Norway, Poland, Russian Fed., Saudi Arabia, Sweden, Slovakia, and South Africa. | Terminated | June 09, 2015 | 238 | 6126 | Not_Computer
Z-0421-2015 | 69505 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemStar infusion pump, list numbers 13000 and 13100, GemStar Infusion Pump (pain management) list number 13150. |  |  | March 19, 2013 | 2013 | Hospira Inc. | Some patients are obtaining access to the dosage reset codes required to change the settings on GemStar infusion pumps and are subsequently changing the infusion rate set by their doctors. Patients are obtaining access to these codes via website sponsors by downloading the GemStar User/Service Manual featured on the sponsors' websites. The GemStar User/Service Manual should be removed from these w | Other | Hospira sent a letters dated March 19, 2013 to Internet website sponsors to request removal of the GemStar User/Service Manual from their websites as soon as possible. The Internet website sponsors were requested to return a confirmation form after removing the GemStar User Service Manual from their site. For questions call 224-212-2891. |  |  | Not Applicable | Worldwide Distribution - US: Nationwide and the countries of Australia, Austria, Bahrain, Belgium, Croatia, Egypt, Finland, France, Germany, Gibraltar, Greece, Hungary, Icelanofd, Ireland, Israel, Italy, Jordan, Kuwait, Lebanon, Luxembourg, Malta, Netherlands, New Zealand, Oman, Portugal, Qatar, Romania, Saudi Arabia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, United Kingdom | Completed |  |  |  | Not_Computer
Z-0425-2015 | 69625 | No | LBH | Cavity varnish. | Dental | Dental | 510(k) | Relief ACP Oral Care Gel intended to relieve discomfort from dentin sensitivity. Sold separately in single, single syringe, 4 packs, 6 packs and in teeth whitening kits |  |  | October 24, 2014 | 2014 | Discus Dental LLC | Product sold after April 1, 2014 were inadvertently filled with hydrogen peroxide teeth whitening gel (6% concentration as dispensed) on one side of the dual barrel syringe instead of the intended component (phosphate gel). Consequently, use of this product by patients may result in a neutral effect or increased sensitivity. | Mixed-up of materials/components | Cease distribution of the product immediately and contact patients who may have received the product from April 1, 2014. Patients who are still in possession of Relief ACP Oral Care Gel should return affected product to dentinal office.Discus Dental is offering a non-cost replacement for all affected product, including free shipping. To remedy the situation take the following steps:1. Contact the patients who have received affected product after April 1, 2014 either by letter or telephone, and instruct them to cease use of the product immediately and return any remaining product.2. Please use Attachment A: Product Shipped, Acknowledgement and Response Form to document the quantity of affected lots of product being returned to Discus Dental LLC. Go to www.philips.com/discus-reliefgel and enter the requested information with contact details.3. Please contact Discus Dental customer service at CustomerServicePhone if you have any questions and to arrange for replacement products and prepaid return shipping label. |  | 34904 | 20,251 distributed in U.S. on nationwide level and 14,653 internationally | Worldwide distribution including Nationwide (U.S.) and the countries of New Zealand, Australia, Canada, Germany, Switzerland, Italy, Austria, Finland, South Africa, United Arab Emirates, Korea, Turkey, Egypt, Jordan, Thailand, Lebanon, Kuwait, Tunisia, India, and Dominican Republic. | Terminated | April 14, 2015 | 172 | 34904 | Not_Computer
Z-0433-2015 | 69638 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | 8MM,MARYLAND BIPOLAR FORCEPS; |  |  | November 03, 2014 | 2014 | Intuitive Surgical, Inc. | Deviations in reprocessing steps from those stated in the reprocessing instructions can cause surface degradation of the housing and/or accelerate mechanical wear of the instrument. | Labeling design | Urgent Medical Device Recall letters were sent to affected sites on November 3, 2014. Letters inform customers of the reason for recall, risk to health and the affected instrument. A table of all of the affected IS4000 instruments affected is attached to the letter indicating the part number and product name for each. Letters also inform customers that all instruments shipped since July 2014 are manufactured with a new material and not subject to recall.Customers are to take the following actions:1. Ensure that all affected personnel are aware of the information.2. Identify affected product in their possession3. Contact ISI rep for replacement.4. Complete and return the attached Acknowledgement From.Until replacement instruments are received, the old instruments may be used provided they do not show signs of degradation and that the reprocessing instructions are followed. Questions to 800-876-1310 Option 3 ( 6 am to 5 pm PST) or ics@intusurg.com. |  | 59 | 5-9 units - total all devices | US Distribution -- FL, TX, NY, NV, OR, IN, UT, and MA. | Terminated | February 03, 2015 | 92 | 59 | Not_Computer
Z-0463-2015 | 69542 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Abbott m2000sp, an automated system for performing sample preparation for nucleic acid testing. |  |  | October 23, 2014 | 2014 | Abbott Molecular | Some versions of Application Specifications (App Spec) are incompatible with m2000sp system software version 6.0 and 7.0. This issue will cause Error Code 9000 (An unexpected error has occurred: Software error) to be generated when the Available Reagent Packs screen is selected. | Software design (manufacturing process) | Field Correction Recall / Urgent Field Safety Notice letters dated September 23, 2014 were sent to all direct accounts (customers). The letters included instructions for customers to: 1) review the app spec files that are installed on the m2000sp system and delete any incompatible app specs; 2) if one or more app spec files were deleted, installation of new app spec file(s) is needed; and, 3) review this information with laboratory personnel and retain the letter for future reference. Customers with questions about this market action can contact Technical Support at 800-553-7042, Option 2, or the local Abbott Molecular representative. |  | 9 | 9 CD ROMs (US); 1,273 CD ROMs (Foreign) | Worldwide distribution. US nationwide (CA, CT, FL, NC, PA, SC, and DC), Algeria,Australia,Austria,Belgium,Bosnia and Herzegovina,Cameroon,Canada,Chad,Colombia,Costa Rica,Croatia,Denmark,Egypt,Ethiopia,Finland,France,Germany,Greece,Guinea,Honduras,Hong Kong,India,Indonesia,Ireland,Israel,Italy,Japan,Kenya,Korea,Lithuania,Mali,Mexico ,Netherlands,New Zealand,Nigeria,Norway,Pakistan,Paraguay,Poland,Portugal,Romania,Russia,Saudi Arabia,Senegal,Slovenia,South Africa,Spain,Sweden,Switzerland,Taiwan,Tanzania,Thailand,Turkey,Uganda,Ukraine,United Arab Emirates, United Kingdom,Vietnam | Terminated | March 08, 2016 | 502 | 9 | Computer
Z-0449-2015 | 69629 | stains, microbiologic | JTS | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Remel TB Potassium Permanganate (.5% Aqueous) 40192, packaged in 250 ml bottles, 5/pkg. |  |  | October 29, 2014 | 2014 | Remel Inc | The product may appear cloudy or discolored and may not perform correctly in presumptive diagnosis of mycobacterial disease. | Process control | Thermo Fisher Scientific sent a Medical Device Recall letter dated October 29, 2014, to all affected customers. The letter notified customers that Remel, Inc., a part of Thermo Fisher Scientific, was voluntarily recalling the affected product. Customers were instructed to notify any personnel who need to be notified of the potential for false negative results. Customers should review results reported using this lot and consider retesting and/or seek appropriate expert advice. Return the attached Medical Device Recall Return Response form acknowledging receipt of the notice and disposal of the affected product.Customers with questions were instructed to contact the Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International).For questions regarding this recall call 913-888-0939. |  | 395 | 79/5/250- packages/bottles/ml | Worldwide Distribution to USA including AL, CA, CO, IA, MO, NC, NY, NV, OH, PA, SC, SD, TN, TX, VA, WI and Internationally to Canada. | Terminated | June 01, 2015 | 215 | 395 | Not_Computer
Z-0471-2015 | 69748 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Cardiac Function Analysis Software |  |  | November 05, 2014 | 2014 | Toshiba American Medical Systems Inc | Potential problem with the cardiac function analysis software (CFA). Incorrect analysis results may be displayed in the Function Parameters for the Entire Heart displayed as analysis results of the CFA and in a Left-Ventricular Volume Curve generated based on some of those parameters. | Software design | Toshiba sent an Urgent Medical Device Correction letter dated November 5, 2014, to all affected customers. Customers were asked to refrain from using the function parameters for the entire heart and the left-ventricular volume curve until corrective actions have been taken.Customers were instructed to share this information with all users and reviewing radiologist as well as clinical engineering or Biomedical group at facility.Please complete and return the attached form and fax it to the toll free number at the top of the form 877-349-3054. This form can also be sent via email to raffairs@tams.com.Revised software to prevent occurrence of this problem will be installed on the system. When the new software becomes available the Toshiba service representative will contact for an appointment to install it on the system.Customers with questions were instructed to call 800-421-1968 or contact their local Toshiba Representative at 800-521-1968. |  | 1 | 1 unit within the U.S. only | US Nationwide Distribution to CA, MA, MD | Terminated | October 09, 2015 | 338 | 3 | Computer
Z-0226-2015 | 69603 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Synthes Titanium Polyaxial Reduction Head for Titanium Matrix Spine Screws: The Titanium Polyaxial Reduction Head is a part of the Matrix Spine System. |  |  | October 23, 2014 | 2014 | Synthes (USA) Products LLC | Two nonconforming parts of the Titanium Polyaxial Reduction Head for Titanium Matrix Spine Screws may have been inadvertently packaged and shipped. It is believed that the inner diameter of the Reduction Head may not have been machined to the correct specification. | Process control | DePuy Synthes sent an Urgent Notice: Medical Device Recall letter, dated October 23, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check inventory and remove affected stock, call Synthes to obtain a Return Authorization for the device, and to return the completed verification section of the response form by fax to (610) 430-7083 or Scan/email: Fieldaction@synthes.com. Customers with questions were instructed to call 610-719-5450. |  | 24 | 24 | Distributed to MI, MT, and CO. | Terminated | August 10, 2015 | 291 | 24 | Not_Computer
Z-0490-2015 | 69786 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Polyflux Revaclear Max Dialyzer, Model Number 110634. |  |  | November 05, 2014 | 2014 | Gambro Renal Products Inc | Gambro is initiating a field action for Polyflux Revaclear Dialyzer and Polyflux Revaclear Max Dialyzer due to potential tears in the individual packaging that may compromise the sterile barrier. | Process control | Recall letter was sent via UPS delivery next day to US customers on November 5th, 2014. Product that is returned will be quarantined at the distribution centers. |  | 20000000 | 20 million dialyzers | Worldwide Distribution: US (nationwide) including PR, and countries of: CA, BB, BM, BS, DO and LC. | Terminated | May 20, 2015 | 196 | 20000000 | Not_Computer
Z-0473-2015 | 69793 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | IntelliSpace Portal DX/HX/EX model number: 881001 |  |  | November 09, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Software defect. In certain circumstances, the application may display incorrect measurements of cardiac and aortic anatomy used to select and size the Transcatheter heart valve (THV) to be implanted . | Software design | Consignees were notified of the Field Safety Notice via certified mail. |  | 32 | 32 | Worldwide Distribution: US Distribution in states of: FL, NC, and VT; and the countries of Argentina, Austria, Belgium, Canada, Denmark, France, India, Italy, Japan, Netherlands, United Kingdom, and Vietnam. | Terminated | September 14, 2015 | 309 | 32 | Computer
Z-0468-2015 | 69631 | No | IYB | Diagnostic x-ray tube mount. | Radiology | Radiology | 510(K) Exempt | The J700 Floor Mounted Tube Stand (FMTS) is intended to support and position the diagnostic x-ray tube housing assembly for a medical radiographic procedure. The product consists of a vertical column (with counterbalanced vertical carriage, counterweight and pulley system), horizontal floor car and a floor mounted track to ride on. |  |  | October 28, 2014 | 2014 | Summit Industries Inc. | The welds on J700 tube stands may be insufficient resulting in tube stand breakage, damage to the X-Ray machine, and/or patient or healthcare provider injury. | Process control | Summit Industries sent an Urgent Medical Device Recall letter dated October 27, 2014, to all direct accounts (distributors). The letters included instructions for distributors to: 1) make every effort to locate the affected devices; 2) replace and return the defective parts; 3) contact the end user for any defective device(s) that were installed and determine if the affected device(s) is/are in service; 4) if the device(s) is/are still in service, deliver a copy of the "End User Letter", immobilize the tube arm, and place the REMOVE FROM SERVICE sticker on the device; 5) complete and return the attached questionnaire for a no-charge purchase order for the necessary replacement parts; and, 5) inform the end user that new replacement pats will be installed as soon as possible. Distributors with questions can contact Summit Industries at 800-729-9729 (Monday - Friday, 7:00 AM - 3:45 PM CST). The Urgent Medical Device Recall letter dated October 27, 2014, to end users advised them of the recall and the steps that Summit Industries and their distributor would be taking to correct the issue. |  | 32 | 32 | US Distribution including the states of: AZ, CO, IL, KY, LA, NE, NY, OH, PA, TN, TX, VA, WA and WV. | Terminated | May 21, 2015 | 205 | 32 | Not_Computer
Z-0483-2015 | 69627 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Cannulaide, Model 101, 102, 103. Respirator. |  |  | May 13, 2008 | 2008 | Beevers Manufacturing &amp; Supply, Inc. | Cannulaide recalled May 2008 due to improper seal of the pouch. The improper seal of the Cannulaide's package could potentially affect the device shelf life. | Equipment maintenance | Beevers Manufacturing &amp; Supply, Inc. (BMS) did contact the customers via letter and phone call in May 2008. Customers that didn't reply were further contacted via email on June 2008. Salter Labs provided BMS' recall report to FDA in October 2014.Salter Labs was unable to locate a recall letter issued by BMS to the affected customers. For further questions please call (503) 472-9055 |  | 16575 | 16,575 units total (15,625 units in US and 950 units outside US) | Worldwide Distribution - US Distribution including the states of LA, GA, OR, MA, NJ, FL and CA., and the countries of Canada, France, Japan, Netherlands and United Kingdom. | Terminated | November 28, 2014 | 2390 | 16575 | Not_Computer
Z-0485-2015 | 69728 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Esprit V1000 Ventilator and Esprit/V200 Conversion Ventilator |  |  | October 29, 2014 | 2014 | Respironics California Inc | If the power supply fan mounting screws are installed incorrectly, the ends of the screws may touch the power supply printed circuit board assembly (PCBA) and short to ground. This may prevent the ventilator to switch to battery power if AC power is disrupted. | Device Design | Philips Healthcare sent an Urgent - Field Safety Notice to all affected consignees to inform them that Respironics is recalling the Esprit V1000 and Esprit/V200 Conversion ventilators because if the power supply fan mounting screws are installed incorrectly, the ends of the screws may touch the power supply printed circuit board assembly (PCBA) and short to ground. This may prevent the ventilator to switch to battery power if AC power is disrupted. A high urgency audible alarm will sound alerting medical personnel of the loss of power. In the absence of AC power, failure of battery power will cause loss of ventilator support, which may result in hypercarbia or hypoxemia. The potential issue affects ventilators manufactured prior to 12/31/04 due to design of the fan/shroud assembly. Esprit V1000 and Esprit/V200 Conversion Ventilators manufactured after that affected date are not affected. Customers with concerns were instructed to contact a Philips Respironics representative at (800) 722-9377. For questions regarding this recall call 760-918-7300. |  | 3943 | 3,943 units total (3,071 units in US) | Worldwide Distribution - USA (nationwide) and Internationally to Argentina, Australia, Brazil, Bulgaria, Canada, Chile, China, Colombia, Ecuador, Egypt, Greece, Hong Kong, India, Italy, Korea, Lebanon, Malaysia, Mexico, Pakistan, Panama, Peru, Philippines, Saudi Arabia, Singapore, South Africa, Spain, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, Venezuela, Vietnam, and Yemen. | Open, Classified |  |  | 3943 | Not_Computer
Z-0486-2015 | 69602 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | The Artis zee / zeego Angiography System is designed as a set of components that |  |  | September 11, 2013 | 2013 | Siemens Medical Solutions USA, Inc | There is a potential issue on running Artis systems running software VC1x software if a network problem arises, the function cannot be deactivated again by pressing the &quot;Block Radiation&quot; key on the touch screen control. If this behavior occurs, the system image generation function is not available for patient examinations without switching the system off and back on again manually. This | Device Design | Siemens sent an Important Customer Safety Notice in September 2013, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. If the fault described above occurs, we urgently recommend switching the system OFF manually ("System shutdown") and then back ON again. The fault cannot generally be rectified sufficiently, either by executing the "System Restart" or "Application Restart" functions. A corresponding supplement to the Operator Manual is enclosed with this letter.The problem will be solved when the next Service Update is available (currently scheduled for 2014).We thank you for your cooperation in dealing with this safety notice and request that you promptly notify and instruct accordingly all the staff at your organization who have to be aware of this problem. Please also forward this safety information to other organizations that could be affected by this measure. Please observe this safety notice, and comply with the corresponding measures until the update has been fully completed.If this device has been sold and is therefore no longer in your possession, please forward this safety notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible.For further questions please call (610) 219-6300. |  | 1131 | 1131 | Worldwide Distribution - USA (nationwide) including Puerto Rico., and the countries of: United Arab Emirates, Armenia, Austria, Australia, Belgium, Bahrein ,Belarus, Switzerland, Chile, China, Colombia, Czech Republic,Germany, Denmark, Egypt, Spain, Finland, France, United Kingdom, Hong Kong, Hungary, Indonesia, Ireland, Israel, India, Italy, Jordan, Japan, Kenya, Korea, Kazakhstan, Malaysia, Netherlands, Norway, New Zealand, Philippines, Pakistan, Poland, Palestina, Portugal, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovakia, Thailand, Turkey, Taiwan, Ukraine, Vietnam and South Africa. | Terminated | March 30, 2015 | 565 | 1131 | Computer
Z-0450-2015 | 69783 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView product code: 882478 |  |  | November 04, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Software issues | Software design | Philips Healthcare sent an URGENT - Field Safety Notice Medical Device Correction letter dated November 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This Field Safety Notice 88200485 is intended to inform you about:" what the issues are and under what circumstances they can occur" the actions that should be taken by the customer / user in order to prevent risks for patients or users" the actions planned by Philips to correct these issues.If you need any further information or support concerning these issues, please contact your local Philips representative:For North America contact the Customer Care Solutions Center (1-800-722-9377: follow the prompts).Philips apologizes for any inconveniences caused by this problem. |  | 85 | 85 | Worldwide Distribution - USA including the states of AR, CA, FL, GA, ID, LA, ME, MI, MN, MO, MS, NJ, NY, OH, OK, PA, and TX. Government Facilities: AR, FL, MN, OH, and OK. and the countries of: Canada, China, Denmark, Germany, Italy, South Korea, Spain, Sweden, Switzerland, Taiwan, and United Kingdom. | Open, Classified |  |  | 85 | Computer
Z-0491-2015 | 69541 | No | LFL |  |  | General &amp; Plastic Surgery | 510(k) | Integra Selector Kit |  |  | September 29, 2014 | 2014 | Integra LifeSciences Corp. | Some specific lots of Selector Tubing may potentially leak irrigation fluid during use if not placed on the Selector handpiece in a specific manner. | Device Design | Integra sent an 'Urgent- Medical Device Correction' letters/acknowledgement forms dated 9/29/2014 by traceable courier service. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to review their inventory for the affected product, complete the attached form. Complete the other information on the form and return the form by email or fax.Customers with questions should contact Customer Service at 1-855-532-1723. |  | 6358 | 6,358 tubing kits | Worldwide Distribution - USA (nationwide) and Internationally to Algeria, Australia, Bulgaria, Canada, Germany, Great Britain, Greece, Hungary, Indonesia, Israel, Italy, Japan, Mexico, Portugal, Spain, Sweden, Thailand, Turkey, and Wales. | Terminated | May 27, 2015 | 240 | 6358 | Not_Computer
Z-0476-2015 | 69386 | No | LBH | Cavity varnish. | Dental | Dental | 510(k) | Dharma Cavity Varnish 15mL |  |  | September 29, 2014 | 2014 | Dharma Research, Inc. | During an FDA inspection it was found that the products are marketed without a cleared 510k. | No Marketing Application | On September 29, 2014 the firm sent a notification to its consignees informing them of the product being discountued and removed. The firm asked for the product to be returned directly to Dharma's office at 5220 NW 72 Avenue, Suite 15, Miami FL 33166. For any questions please do not hesitate to call (305) 482-9669. |  | 916 | 916 | Worldwide Distribution: US Distribution in states of: CA, PA, FL, IL, including Puerto Rico; and countries of: Honduras, Dominican Republic, Turkey, Nicaragua, Colombia, Pakistan, Trinidad &amp; Tobago, Guatemala, Phillipines, Saudi Arabia, and United Arab Emirates. | Open, Classified |  |  | 3758 | Not_Computer
Z-0470-2015 | 69839 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Non-sterile drape |  |  | November 04, 2014 | 2014 | Novadaq Technologies, Inc. | Non sterile-drape LUNA drape was inadvertently mislabeled with an additional contents sterile label. | Error in labeling | Novadaq initiated their recall by telephone and followed by an email and a Medical Device Recall Notice dated November 11, 2014, to all affected customers. Distributors were instructed to notify their customers. The firm is removing the recalled products and replacing them with correctly labeled product. Customers were instructed to confirm receipt of the Recall Notice and their understanding of the actions to be taken by completing the attached confirmation form and faxing it to Novadaq Quality Department at 1-800-886-2419. For questions regarding this recall call 604-432-9861. |  | 324 | 324 units | Nationwide Distribution to IA, LA, GA, MD, ID, FL, TN, OH, TX, NC, MA, KY, KS, WI, IN, SC, OR, CT, CA, WA, PA, OK, and PR. | Terminated | November 06, 2015 | 367 | 1152 | Not_Computer
Z-0493-2015 | 69683 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | GuardIVa Antimicrobial Hemostatic IV Dressing, REF FP-23-AD008, Manufacturer: Bard Access Systems, Inc. USA. Product Code FP23AD006. |  |  | November 03, 2014 | 2014 | Bard Access Systems | Bard Access Systems is recalling GuardIVa (Ref. No. FP23AD006) from lot ASXHT003 because they were non-sterile unit samples which potentially could be used clinically. | Mixed-up of materials/components | A customer letter was sent on 11/3/14 requesting immediate examination of inventory and to quarantine the products for return. |  | 2532 | 2532 units | US Distribution to state of: TX, OK, LA, and AR. | Terminated | May 12, 2015 | 190 | 2532 | Not_Computer
Z-0488-2015 | 69642 | No | HHK | Endometrial suction curette and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Cooper Surgical CANNULA-CURETTE 12MM; Model Number: MX562 |  |  | October 27, 2014 | 2014 | CooperSurgical, Inc. | The product has been identified to contain an incorrect curette type. | Packaging process control | CooperSurgical sent an Recall letter dated October 27, 2014, via UPS to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The subject device does not meet the required release specification as it contains an incorrect curette type. No adverse events or injuries have been reported to date in connection with the use of the CANNULA CURETTE. The event was discovered during internal investigation. Products that do not meet the above criteria are safe to use alongside the DFU guideline.The letter requests that customers return the recalled product for a full refund or exchange. If you have any further questions please feel free to call (203) 601.5200. |  | 30 | 30 | US Distribution including the states of NY, MO and NV. | Terminated | May 28, 2015 | 213 | 30 | Not_Computer
Z-0484-2015 | 69636 | No | HGI | Gynecologic electrocautery and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical LEEP RADIUS LOOP ELECTRODE (W 2.0 cm x D 1.0 cm D, 12 cm shaft ). Box of 5 |  |  | October 27, 2014 | 2014 | CooperSurgical, Inc. | Outer box is mislabeled as P/N R2010- Large Radius Loop Electrode instead of P/N R1010- Medium Radius Loop Electrode | Packaging process control | CooperSurgical issued recall letter dated October 27, 2014, to all affected customers via certified mail delivery (UPS). Acccounts were requested to return product for replacement and complete the return response form. Customers with questions were instructed to call 203-601-5200. |  | 595 | 595 units | Worldwide Distribution - USA (Nationwide) and Internationally to Canada, Portugal, and Hong Kong. | Terminated | May 29, 2015 | 214 | 595 | Not_Computer
Z-0487-2015 | 69648 | No | BZD | Noncontinuous ventilator (IPPB). | Anesthesiology | Anesthesiology | 510(k) | Cannulaide, Model 101. The Cannulaide product is composed of die-cut hydrocolloid coated polyurethane film with an integral hook strip (i.e., Velcro) to allow for the use of a loop material to help secure a nasal interface. |  |  | January 18, 2007 | 2007 | Beevers Manufacturing &amp; Supply, Inc. | Cannulaide model CA 101 was recalled because the firm discovered this model was packed and labeled as CA 102. | Labeling Change Control | Beevers Manufacturing and Supply, Inc. (BMS) of McMinnville, Oregon, began notifying consignees about this recall on 1/18/2007.Customers were advised to inspect blue box that say size 2 (CA 102) and check if size 1 (CA 101) was packed inside. Salter Labs provided BMS' recall report to FDA in October 2014. |  | 393 | 393 units | Worldwide Distribution: US (nationwide) and country of: ME-Musaffah. | Terminated | December 01, 2014 | 2874 | 393 | Not_Computer
Z-0474-2015 | 69289 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ARCHITECT Total T3 Reagent Kit (7K64) consists of: |  |  | September 11, 2014 | 2014 | Abbott Ireland Diagnostics Division | 17% of ARCHITECT Total T3 Reagent Kit, LN 7K64-20/-25 reagent lots 38901UI00 and 38901UI01 may exhibit lower Relative Light Units (RLUs) than expected, which in turn may result in controls out of range or patient results higher than expected. | Employee error | Abbott sent an Product Recall letter dated September 11, 2014, to all direct accounts (customers). The letters included instructions for customers to: 1) Discontinue use of and destroy any remaining inventory according to your laboratory procedures; 2) follow your laboratory protocol regarding the need for review of previously reported patient results; 3) replacement material is available; for ordering, please order through your normal ordering process; and, 4) if you have forwarded any of the affected lots to another laboratory, please provide them with a copy of this communication. Customers or healthcare providers with questions can contact Customer Service at 1-877-422-2688 (24 hours a day, 7 days a week). |  | 7415 | 7,415 kits | Worldwide Distribution - USA (nationwide) including Puerto Rico and the states of AL , AR , AZ , CA , CT , FL , GA , IL , IN , KS , LA , MA , MD , MI , MN , MO , MS , MT , NC , NJ , NV , NY , OH , OK , PA , SC , SD , TN , TX , UT , VA , WA , WI and WV., and the countries of : Andorra , Angola , Argentina , Armenia , Austria , Azerbaijan , Bangladesh , Barbados , Belarus , Belgium , Bosnia &amp; Herzegovina , Brazil , Burkina Faso , Canada , Cayman Islands , Chile , China , Colombia , Costa Rica , Croatia , Czech Republic , Ecuador , Egypt , El Salvador , Ethopia , Georgia , Germany , Greece , Greenland , Guatemala , Honduras , Hong Kong , India , Indonesia , Iran , Ireland , Israel , Italy , Jamaica , Japan , Kazakhstan , Kenya , Korea , Kuwait , Latvia , Lebanon , Libya , Malaysia , Mexico , Moldova , Montenegro , Oman , Pakistan , Gaza &amp; Jericho , Panama , Paraguay , Peru , Philippines , Poland , Portugal , Qatar , Romania , Russia , Saudi Arabia , Serbia , Singapore , Slovakia , Spain , Sri Lanka. , Switzerland , Taiwan , Tajikistan , Thailand , The Netherlands , Trinidad &amp; Tobago , Turkey , Ukraine ,United Arab Emirates , United Kingdom , Venezuela , Vietnam and Yemen. | Open, Classified |  |  | 7415 | Not_Computer
Z-0477-2015 | 69735 | No | PAV | Intravascular occluding catheter. | Neurology | Neurology | HDE - Humanitarian Device Exemption | 7F NeuroFlo Catheter, Model 1027 |  |  | April 01, 2014 | 2014 | ZOLL Circulation, Inc. | After aquisition of assets of CoAxia, Zoll discovered that one lot of NeuroFlo Catheters was manufactured by another firm and an HDE supplement listing as the owner of the product was not filed. | Other | Hospital was notified in April 2014 of the recall by UPS. |  | 1 | 1 device | One location in NC | Terminated | November 26, 2014 | 239 | 1 | Not_Computer
Z-0206-2015 | 69568 | protector, skin pressure | FMP | Skin pressure protectors. | General Hospital | General Hospital | 510(K) Exempt | Gel-E Donut : 92025-A (Extra Small), 92025-B (Small), |  |  | October 10, 2014 | 2014 | Philips Medical Systems, Inc. | Mold on gel filled Gel-E and Squishon products | Material/Component Contamination | Philips Healthcare/Children's Medical Ventures issued a Medical Device Recall Letter dated Oct 10, 2014. Customers were asked to discontinue use and dispose of all product in their facility, even if mold is not visible. Following these actions, customers will be asked to return a reply form indicating that these actions have been completed. Customers will be given credit for all disposed products. Further, the Gel-E Donut and Squishon 2 products will be withdrawn from the market. Customers with questions were instructed to contact their Philips/Children's Medical Ventures representative.For questions regarding this recall call 978-687-1501. |  | 336695 | 336,695 units | Worldwide Distribution - USA (nationwide) and Internationally to: Canada, Australia, Austria, Belgium, France, Germany, Iceland, Ireland, Italy, Japan, Kuwait, Netherlands, New Zealand, Norway, Portugal, Reunion, Romania, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Thailand, and the United Kingdom. | Terminated | November 29, 2016 | 781 | 336695 | Not_Computer
Z-0464-2015 | 69515 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | AlignRT- Intended for prescription use. The system is indicated for use during simulation, setup and stereotactic radiosurgery and precision radiotherapy for lesions, tumors and conditions anywhere in the body where radiation is indicated. |  |  | October 15, 2014 | 2014 | Vision Rt Inc | Potential failure of AlignRT to assert interlock. | Software in the Use Environment | The firm, visionrt. sent an "URGENT MEDICAL DEVICE CORRECTION NOTIFICATION-URGENT FIELD SAFETY NOTICE" dated October 15, 2014 to all affected customers on 21-Oct-2014. The notice describes the product, problem and actions to be taken. The customers were informed that they did not need to return their devices; they should pass the letter to all those who need to be aware of it within their organization; promptly inform Vision RT if they believe that patient harm occurred due this issue, and complete and return the acknowledgement form via email to service@visionrt.com.Vision RT is designing an upgrade that will resolve the issue, and will subsequently contact the customer in order to install the upgrade on their device.Should you have any queries on this letter, please do not hesitate to contact Vision RT by telephone on +44 20 83464300 (866 778-2379 from the US) or as per http://www.visionrt.com/contact/details. |  | 7 | 7 (Six Units distribited in the US and One in OUS.) | Worldwide Distribution: US distribution in states of: CA, MN, TX, and UT; and internationally to: Republic of Ireland. | Open, Classified |  |  | 7 | Not_Computer
Z-0467-2015 | 69621 | No | JKR | Calcitonin test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Calcitonin ELISA Kit, Code No. 7024, 7024BU. |  |  | September 18, 2014 | 2014 | Biomerica | A potential issue with false high values. | Material/Component Contamination | An Urgent Field Safety Notice dated 9/18/14 was sent to all their customers who purchased the Calcitonin ELISA Kit. The letter informed the customers of the problems identified and the actions to be taken. Customers were to examine their inventory immediately and quarantine any affected product. Also, customers were asked to notify any customers that the product was further distributed to. Customers with questions were instructed to contact their local Customer Service or Sales Representative regarding the recall notification. Customers were instructed to complete and return the enclosed response form by either fax or email as soon as possible and not later than 10/1/14. |  | 1062 | 1,062 units | Worldwide Distribution -- USA, including the states of NH, CA, NJ, FL; and, the countries of Russia, Ukraine, Switzerland, Spain, Italy, Serbia, China, France, Pakistan, Austria, United Kingdom, Egypt, and India. | Terminated | March 28, 2015 | 191 | 1062 | Not_Computer
Z-0215-2015 | 69619 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | AxoTrack I Sterile Procedure Kit 18G, K1001-18XTW, (24/case labeled for US, 5/case labeled for US, 24/case labeled for EU, 5/case labeled for EU). |  |  | March 24, 2014 | 2014 | Soma Access Systems LLC | There is a possibility the sterile packaging may fail before the expiration date specified on the package. | Packaging | Consignees were notified via USPS Certified Letter, email and FedEx on 3/24/2014. |  | 3530 | 3530 units | Worldwide Distribution-US (nationwide) including the states of AL, AZ, CA, CT, FL, GA, IN, KY, LA, MA, MI, MO, NC, NJ, NY OH, OR, PA, SC, TN, TX, VA, and WA and the countries of Belgium, Hong Kong, India, Japan, Panama and Singapore. | Terminated | August 06, 2015 | 500 | 3530 | Not_Computer
Z-0482-2015 | 69669 | No | NJE |  |  | Neurology | HDE - Humanitarian Device Exemption | Gateway OTW 2.75mm x 9mm; |  |  | October 22, 2014 | 2014 | Stryker Neurovascular | Units of GDC 360 coils, Target coil and Gateway PTA Balloon Catheter were shipped to US customers beyond the expiration date. | Process control | Stryker Neurovascular sent an Urgent Medical Device Recall Notification dated October 22, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Letters request that users take the following actions:1. Immediately locate the subject devices and quarantine them.2. Distribute this notice to all affected Departments in your facility.3. Inform Stryker Neurovascular if any of the subject devices have been distributed to other organizations.a. Please provide contact details so that Stryker Neurovascular can inform those recipients appropriately.4. Complete and sign the attached Customer Response Form and:a. Fax it to: 1-866-876-4355 orb. E-mail a copy to: NeuroComplaints@stryker.comc. Please complete this form even if you do not have any subject devices to return to prevent the need of follow up notices.5. Keep a copy of the completed, signed Customer Response Form for yourrecords.Stryker Customer Service reps will contact customers to arrange for replacement.Any returned product will be quarantined and destroyed. |  | 1 | 1 unit | US Distribution to the states of: VA, NH, OH and TX. | Terminated | February 12, 2015 | 113 | 1 | Not_Computer
Z-0237-2015 | 69448 | No | MCZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | Surgical packs (includes trays, bags and kits) |  |  | October 08, 2014 | 2014 | Customed, Inc | Customed products are being recalled because of compromised sterility, not limited to potentially damaged packaging. Serious deficiencies in the manufacturing processes (QSR) and uncontrolled and inadequate storage conditions increase the risk of contamination and may result in patient infection. | Packaging | Customed initiated a recall on October 8th, 2014 of sterile convenience surgical Packs-trays, bags (known as Convenience Packs) and new lots of products added to the initial June 3, 2014 recall of 233 sterile surgical packs. The products are being recalled because of significant compromised sterility, not limited to potentially damaged packaging. Customers were instructed to quarantine and cease use of the affected product identified. Customers are to return the response form included with the letter, to the contact information provided. Upon receipt, the firm will contact customers with further instructions to obtain credit, replacement product, and for return of the recalled product or any other disposition action. If customers have any questions, they should contact the Customed Recall Coordinator at 407-850-5558 extensions 8001 or 8005, or via e-mail: pablo.diaz@customedhealing.com. |  | 185000 | 185000 | US Distribution to Florida, New York and Puerto Rico.. | Terminated | August 14, 2017 | 1041 | 1079228 | Not_Computer
Z-0167-2015 | 69345 | No | HQG | Prescription spectacle lens. | Ophthalmic | Ophthalmic | 510(K) Exempt | Prescription eyeglass safety lenses. |  |  | September 12, 2014 | 2014 | Eyemart Express Ltd | Prescription eyeglass safety lenses did not meet specifications. | Employee error | Eyemart Express sent an Important Recall Notice dated September 12, 2014, to all affected customers. The notice identified the product, the problem, and the action to be taken by the customer. Customers were instructed to call Eyemart Express Customer Service at 1-888-DRBARNES to discuss their lens replacement options of if they have any questions. |  | 76 | 76 pairs of eyeglasses | Nationwide Distribution including AK, AL, AR, CA, LA, MD, MO, MS, MT, NC, ND, NM, OH, OK, TN, SC, TX, VA, WI, and WV. | Terminated | October 09, 2015 | 392 | 76 | Not_Computer
Z-0186-2015 | 69570 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Discovery NM/CT 670 dual detector free-geometry integrated nuclear imaging system. Models H3100NA, H3100PP, H3100RA, H2401LS, H2401LZ, H2402LS. |  |  | October 14, 2014 | 2014 | GE Healthcare, LLC | GE is issuing this recall due to a potential safety issue related to a portion of the system falling onto the patient during a scan due to fasteners being loose that secured the camera to the gantry. GE is updating the Preventative Maintenance procedure and schedule, and added an additional Preventative Maintenance check for loose fasteners on the Nuclear Medicaine systems. | Nonconforming Material/Component | The firm, GE Healthcare, sent an "URGENT MEDICAL DEVICE CORRECTION" letter, GEHC Ref# 40860, dated October 14, 2014 to its consignees. The letter described the Safety Issue, safety Instructions, Affected Product Details, Product Correction, actions to be taken and Contact Information. For questions, consignees in the United States, can contact 1-800-437-1171, for other countries they can contact their GE Healthcare Service Representative. |  | 255 | 255 | Worldwide Distribution: US (nationwide) including Guam, Puerto Rico and DC; and countries of: ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, BELGIUM, BELRARUS, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, ESTONIA, FINLAND, FRANCE, GEORGIA, GERMANY, HUNGARY, GREECE, GUADELOUPE, HONG KONG, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAPAN, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LITHUANIA, MALAYSIA, MALTA,MARTINIQUE, MEXICO, MOROCCO, NAMIBIA, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PANAMA, PERU, PHILIPPINES, POLAND, PORTUGAL QATAR, ROMANIA, RUSSIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SWEDEN,SWITZERLAND, TAIWAN, THAILAND, TUNISIA, TURKEY, UNITED KINGDOM, UKRAINE, URUGUAY, VENEZUELA, VIET NAAM. | Terminated | February 27, 2015 | 136 | 3410 | Not_Computer
Z-0466-2015 | 69566 | No | LIB | Microorganism differentiation and identification device. | Microbiology | Microbiology | 510(K) Exempt | Oxoid Legionella Latex Test, DR0800M, containing 50 tests per box. The box contains a variety of components, one of which is Oxoid L. pneumophila Latex Reagent, DR0802, Lot 1301912. The responsible firm on the label is Oxoid Ltd. |  |  | October 23, 2014 | 2014 | Remel Inc | A reagent within the test may return false negative results. | Other | ThermoFisher sent a Medical Device Recall letter dated October 23, 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify any personnel who need to be aware of the potential for false negative results for the affected product. Customers should review their inventory for the affected products and discard them. Return the attached Medical Device Recall Return Response acknowledging receipt of the notice and disposal of the affected product. If product was further distributed customers should contact those entities, advise them of the situation and provide them with a copy of this letter. Customers with questions were instructed to contact their Technical Services Department at 800-255-6730 (US) or 913-888-0939 (International). |  | 2950 | 59/50 test boxes | Nationwide Distribution including AZ, CA, DC, FL, IL, ME, NJ, NY, PA, VA, and VT. | Terminated | January 22, 2015 | 91 | 3050 | Not_Computer
Z-0190-2015 | 69473 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16-slice CT Scanner System and MX 16-slice CT Scanner System Phase 11 can be used as a whole body computed tomography X-ray system featuring a continuously rotating X-ray tube and detector array with multi-slice capability up to 16 slices simultaneously. The acquired X-RAY transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. The system is suitable for all patients. Access Dual CT system can be used as a Head and Whole Body Computed Tomography X-ray System featuring a continuously rotating X-ray tube and detector array with multi-slice capability. The acquired X-ray transmission data is reconstructed by computer into cross-sectional images of the body from the same axial plane taken at different angles. |  |  | October 09, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | &quot;The orientation of images is displayed incorrectly. When operator chooses gantry on the right side from operation console, and the patient position is HFDL, HFDR, FFDL or FFDR, the patient position picture and text is not consistent. &quot;L&quot; and &quot;R&quot; of scan images are reversed and this is not consistent with the patient actual position. &quot;When a surview which is longe | Software Design Change | Philips is informing the consignees of this issue through a Field Safety Notice (FSN) 72800615 on October 09, 2014. The US consignees will receive this notice via certified mail and Philips Key Market representatives will distribute the FSN internationally. Additionally, A Field Service Engineer will contact each site to schedule a time to install the software update. |  | 708 | 708 | Domestic Distribution: AR, CA, FL, GA, IL, IN, KY, MI, MS, NC, ND, NJ, NV, NY, OH, PA, PR, SC, TN, and VA. Foreign distribution: Albania, Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Chile, China, Colombia, Congo, Ecuador, Egypt, France, Germany, Ghana, Greece, Gutamala, Hungray, India, Indonesia, Ireland, Italy, Japan, Jordan, Latvia, Lybyan Arab JHamahiri, Lithuania, Malawi, Mexico, Myanmar, Namibia, Nepal, Netherlands, Oman, Palestine, Paraguay, Philippines, Poland, Portugal, Romania,Russia, Rwanda, Saudi Arabia, Senegal, South Agrica, South ZKorea, Spain, Sri Lanka, Switzerland, Tunisia, Turkey, Uganda, UAE, Venezuela, Vietnam, Zambia. | Terminated | July 02, 2015 | 266 | 708 | Computer
Z-0220-2015 | 69615 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | LANX, REF, PCR8XXX-XXX, LOT, RXONLY. Packaged within aluminum sterilization case. |  |  | March 09, 2011 | 2011 | Biomet Spine LLC. | Biomet Spine, LLC. announces voluntary recall of Durango Anchored ALIF Plate due to the set screw component potentially becoming dislodged from the plate assembly. | Device Design | Biomet notified customers by phone on March 9, 2011 to return the affected product. |  | 120 | 120 | US Nationwide Distribution in the states of California, Indiana, Louisiana, and Texas, | Terminated | January 07, 2015 | 1400 | 120 | Not_Computer
Z-0492-2015 | 69787 | No | GAP | Natural nonabsorbable silk surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon PERMA-HAND Silk siliconized black braided 10 strands per packet Ethicon LLC- Made in USA. |  |  | October 17, 2014 | 2014 | Ethicon, Inc. | Ethicon is unable to ensure that PERMA-HAND silk suture (siliconized black braid) meets the expected sterility assurance level for sterile products. | Other | Ethicon sent an Urgent Medical Device Recall Letter/Business Reply Form dated 10/17/2014 via UPS next day mail. |  | 2508 | 2,508 each | Nationwide Distribution | Terminated | August 14, 2015 | 301 | 2508 | Not_Computer
Z-0419-2015 | 69584 | No | KNH | Contraceptive tubal occlusion device (TOD) and introducer. | Obstetrics/Gynecology | Obstetrics/Gynecology | PMA | Gyrus ACMI Disposable Falope-Ring Band Applicator Kits. REF 006889-901 and REF 005280-901 (with 8 mm Disposable Trocan). Sterile EO. Single Use Only. The disposable Falope-Ring Band Applicator is a sterile, two-band, dual-incision instrument. The two-band applicator enables the physician to occlude both Fallopian tubes without removing the applicator from the abdominal cavity. The Falope-Ring Band and Applicator Kits are indicated for use for female sterilization (permanent contraception). |  |  | October 21, 2014 | 2014 | Gyrus Acmi, Incorporated | All packages of Falope Ring Band Applicator kits are being recalled due to compromises in sterile packaging that could be associated with bacterial contamination that might lead to patient infection. | Process control | Olympus sent an Urgent Medical Device Recall letter dated October 21, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Olympus requires you to take the following action: I. Immediately cease any further use of any affected product you have, remove it from your inventory and quarantine it until it is shipped back to us. 2. Call your Olympus customer service representative at 1-888-524-7266 to obtain a Returned Goods Authorization so that you may return the product with no charge to you. Olympus will issue a credit t or replacement to your facility y for any returned product. 3. Please note on the enclosed questionnaire that you have received this information. 4. Fax the completed questionnaire to 484-896-7128 regardless of whether you have any affected inventory at your facility. ln addition, if you may have further distributed this product, please identify your customers, notify them at once of this product recall, and appropriately document your notification process. Your notification to your customers may be enhanced by including a copy of this recall notification letter. |  | 74792 | 74,792 units | Worldwide distribution. US Nationwide, Canada, Vietnam, and Europe | Completed |  |  | 74792 | Not_Computer
t | o | o |  | b | i | g | [18 empty] | Not_Computer
Z-0902-2015 | 69773 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABX PENTRA Reagent Container, Model No. B1037307 (15 mL), B1034626 (10 mL), and B1034634 (4 mL) |  |  | November 13, 2014 | 2014 | Horiba Instruments, Inc dba Horiba Medical | HORIBA Medical is recalling the ABX PENTRA 400 Reagent Container because they do not always sit properly into the reagent rack. In low reagent level scenarios, this could lead to insufficient reagent volume being pipetted and could lead to incorrect results being reported without an alarm. The defect would result in a false high or false low results for HbA1c and false high results for magnes | Package design/selection | A recall letter was sent on 11/13/14 to all customers who purchased the ABX PENTRA 400 Clinical Chemistry Analyzer. The letter informs the customer of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed response form within 10 days. Customers with questions are instructed to contact their local HORIBA Medical representative. |  | 443 | 443 units | Nationwide in US | Terminated | March 19, 2015 | 126 | 443 | Not_Computer
Z-0880-2015 | 69964 | No | GJS | Prothrombin time test. | Hematology | Hematology | 510(k) | Alere INRatio PT/INR Test Strips, Alere INRatio |  |  | December 05, 2014 | 2014 | Alere San Diego, Inc. | In certain cases, an INRatio PT/INR Monitor system may provide an INR result that is significantly lower than a result obtained using a reference INR system (laboratory method). This issue can arise if the patient has certain medical conditions. It can also occur if the instructions in the labeling for performing the test are not followed. | Under Investigation by firm | An Urgent: Medical Device Correction, dated December 5, 2014, was sent to Healthcare Professionals. The letter identified the affected product and reason for recall. Customers were instructed to: -Ensure they have read and understand the precautions described in the current product labeling (a complete list of product labeling is provided in Appendix A) and the additional precautions in this notice describing medical conditions that may increase the risk of obtaining a (falsely or erroneously) lower than expected INR result. The INRatio PT/INR Monitor system should NOT be used if the patient has any of the medical conditions described. -Verify that the patient has hematocrit within the range of 30% to 55%. If not, immediately transition the patient to an alternate INR monitoring method.-Perform INR verification testing for the patients using a laboratory INR test method. Immediately transition any patient with a significantly discrepant low result on the INRatioÂżÂż PT/INR Monitor system to an alternative INR monitoring method.-If have forwarded product to another customer, please provide a copy of this letter to them.-Please complete and FAX or e-mail the enclosed Reply Form (Appendix B the last page of this notice) within 10 days to confirm your receipt of this notice. Customers can contact Alere INRatio Recall Hotline by phone at 1-877-929-2579, for questions regarding this notification. Additionally, a website has been established providing information and frequently asked questions. www.inr-care.comPlease FAX or e-mail the completed Reply Form to:Alere San Diego, Inc.Fax: 1-877-929-2580E-mail: Alere4319@stericycle.com. |  | 23894578874 | 77,342 units (12 test strips/box) and 308,947 units (48 test strips/box) | Worldwide Distribution -- US nationwide, Canada, South America, EU, Asia Pacific, Africa, and Middle East. | Open, Classified |  |  | 23895042808 | Not_Computer
Z-0910-2015 | 70144 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | McKesson Cardiology ECG Management |  |  | December 19, 2014 | 2014 | McKesson Israel Ltd. | An ECG procedure time stamp is incorrect, due to a software error in the McKesson Cardiology ECG Management with software versions 13.1 and 13.1.1. | Software design | McKesson sent an Urgent Field Safety Notice dated December 16, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed how to identify the potentially affected procedures and update the procedure date (if relevant) until a software update is made advailable for installation by McKesson. McKesson will proactively identify prior procedures affected by this issue and will notify customers of any issues identified. In addition to the Field Safety Notice, McKesson Israel will provide correction to the affected SW versions and will contact customers for deployment of a software patch that prevents this scenario from occurring. Customers with questions should contact McKesson Support at 1-866-777-0202.For questions regarding this recall call 404-338-3556. |  | 9 | 9 | Worldwide Distribution - USA including MA, MS, NH, NC, TX, WA and Internationally to the United Kingdom. | Terminated | February 25, 2016 | 433 | 9 | Computer
Z-0906-2015 | 69466 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Optetrak Asymmetric Hi-Flex Posterior Stabilized Cemented Femorals, size 3, Catalog #244-02-03, Catalog #244-03-03. |  |  | October 09, 2014 | 2014 | Exactech, Inc. | Mislabeled | Labeling mix-ups | The firm, Exactech, sent an "Urgent Field Safety Notice" letter dated September 30, 2014 to its Consignees/customers by electronic mail on October 9, 2014. The letter described the product, problem and actions to be taken. The customers were instructed to: immediately cease distribution or use of the products; extend information to your accounts that may have this product in their possession; verify whether you have any of the subject femoral components from the specified serial number ranges; complete and return the attached inventory RESPONSE FORM via fax to: 011 61 2 9638 0778 within 48 hrs of receipt of the notice, and contact your Exactech representative to confirm quantities at your location and arrangements for product return and inventory restocking.If you have any questions, contact Manager, Regulatory Affairs at 352-377-1140. |  | 96 | 96 | International Distribution only to countries of: Argentina, Austria, Australia, China, Cyprus, France, Greece, India, Japan, Lebanon, The Netherlands, Spain and Tunisia. | Open, Classified |  |  | 96 | Not_Computer
Z-0909-2015 | 69785 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Cypher MIS Screw System-The system includes screws, various types and sizes of rods, and set screws. Various instruments are also available for use by the surgeon to facilitate implantation of the device. |  |  | October 08, 2014 | 2014 | Ebi, Llc | Inoperative screw head seat splay and/or fracture as a result of positioning of the rod which is a part of the Cypher MIS Screw System. | Device Design | The firm, EBI LLC/Biomet, sent an "URGENT PRODUCT INFORMATION" letter dated 10/6/2014 via Fed Ex to its customers. The letter described the product, problem and actions to be taken. The customers were instructed to follow the steps outlined in the letter. If you have any questions, please feel free to contact Director of Regulatory Affairs at 303-501-8548. |  | 2816 | 2816 approx.units ( 2688 units US and 128 OUS) | Worldwide Distribution: US Distribution to states of: AL, CA, CT, IN, LA, NC, NJ TX &amp; WI; and the country of: The Netherlands. | Terminated | December 21, 2015 | 439 | 2816 | Not_Computer
Z-0908-2015 | 70046 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | I.M. Hole Locator Instrument. Orthopedic manual surgical instrument. |  |  | November 12, 2014 | 2014 | DePuy Orthopaedics, Inc. | Use of excessive force when impacting may lead to intra-operative femoral fracture. | Device Design | The firm, DePuy Synthes, sent an "URGENT INFORMATION - DEVICE CORRECTION NOTICE HP LCS IM Hole Locator" notification dated November 12, 2014 to the affected distributors. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. The distributors were instructed to: review the instructions and the information in this notice; provide this notice to your sales consultants, and complete and return the DISTRIBUTOR RESPONSE CARD: Verify Sales Consultant Notifications via fax to (574) 371-4939 or email: kearle@its.jnj.com. For questions, please contact Recall Coordinator, at 574-371-4917 (M-F; 8 a.m. - 5 p.m. EST). |  | 1130 | 1130 | Worldwide Distribution: US (nationwide) including states of: AZ, CA, CO, FL, GA, HI, IA, IL, IN, KY, LA, MA, MD, ME, MI, MN, MT, NC, NV, NY, OH, PA, TN, TX, VA, WA, and WI; and countries of: Canada, Australia, Austria, Belgium, Czech Republic, France, Germany, India, Israel, Italy, Japan, Malaysia, Russia, South Africa, Switzerland, and United Kingdom. | Terminated | February 24, 2015 | 104 | 1130 | Not_Computer
Z-0907-2015 | 70044 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Ti-Bond Marketing Literature/Labeling, Catalog No. LT10000, MP-14000 |  |  | December 18, 2014 | 2014 | Spinal Elements, Inc | Spinal Elements is recalling the Ti-Bond related brochures that includes LT-10000 and MP-14000 (all revisions) the content regarding the Ti-Bond coating within the literature is unsubstantiated in context and/or misleading for the coating's use in spine applications. | Labeling False and Misleading | Spinal Elements sent an Urgent Field Action letter dated December 22, 2014, to all affected customers. The letter informed the customers to cease all distribution of the literature material. The letter informed the customers that the content regarding the Ti-Bond coating in the literature is unsubstantiated in context and/or misleading for the coating's use in spine applications. In addition to the customer notiification letter, it will include a list of claims related to the field action. The letter also informed customers of the actions to be taken. Customers were instructed to immediately examine their literature inventory, discontinue the use and /or distribution of the affected brochure, destroy all copies or send them back to:Spinal Elements, Inc.3115 Melrose DriveCarlsbad, CA 92010If product was further distributed customers should identify those customers and notify them at once. Customers were also instructed to complete and return the enclosed "Customer Response Form."For questions regarding this recall call 760-607-0121, ext 206.Customers with questions were instructed to contact Customer Service at 1-877-SPINALS, ext 213.Customers with questions were instructed to contact Customer Service at (877) SPINAL5 Ext 213. |  | 1200 | 1200 units | Nationwide Distribution | Terminated | March 10, 2015 | 82 | 1200 | Not_Computer
Z-0905-2015 | 70043 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Femoral Impactor ATTUNE INTUITION Impactor. |  |  | November 10, 2014 | 2014 | DePuy Orthopaedics, Inc. | ATTUNE INTUITION Impactors (cat. nos. 254401003, 254401004, and 254401006) have broken and produced small pieces. Should a fracture occur and not be observed during surgery, there is the potential for these small fractured pieces of the instrument to be left in the patient. | Device Design | DePuy Synthes sent an Urgent Information letter dated November 10, 2014, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to complete the Hospital/User Facility Acknowledgement Car and return to their Deputy Orthopaedics Sales Consultant or fax to 574-371-4939. Customer swere also asked to retain a copy of the complete Acknnowlegement Card in their files along with the notice. Customers with questions were instructed to call 574-371-4917 (M-F; 8 a.m.  5 p.m. EDT).For questions regarding this recall cal 574-371-4917. |  | 4661 | 4661 | Worldwide Distribution: US (nationwide) including P.R. except AK and RI; and Internationally to: Chile, Canada, Australia, Austria, Belgium, China, France, Ireland, India, Israel, Italy, Japan, Malaysia, Netherlands, New Zealand, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Thailand, United Arab Emirates (UAE), and United Kingdom (UK). | Terminated | December 08, 2015 | 393 | 13077 | Not_Computer
Z-0922-2015 | 69927 | No | JEA | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | IR remote control JUPITER / TS7500, Material no.1601927, Serial number: to (incl.) 102222333. |  |  | December 03, 2014 | 2014 | Trumpf Medical Systems, Inc. | Complaints have been reported of unintended movement occurring during surgical cases due to the use of damaged JUPITER remotes | Component change control | Trumpf sent an Urgent Medical Device Correction letter on December 3, 2014, to all affected customers. Customers were instructed to complete and return the Confirmation of Receipt form with serial numbers of the affected table(s) and the serial number(s) for any remote controls they have. Replacement remotes will be sent. Customers were instructed to dispose of their recalled remotes. Customers with questions were instructed to call 888-474-9359.For questions regarding this recall call 843-534-0606. |  | 3303 | 3303 total all devices | Worldwide Distribution - USA including AL, AZ, CA, CT, FL, GA, IL, KS, LA, MD, MI, MK, MS, NC, NH, NJ, NY, OH, OK, PA, RI, SC, TX, VA, WA, WI and Internationally to Canada. | Open, Classified |  |  | 3440 | Not_Computer
Z-0913-2015 | 69114 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Symbia Intevo Excel is a non-diagnostic SPECT ICT system with CT support for only attenuation correction and anatomical localization. |  |  | August 12, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | Siemens Medical Solutions USA, Inc. has received reports of unintended radial motion of the detectors. This unintended motion, should it occur, may cause the patient compression related injuries. | Device Design | Siemens sent Customer Safety Advisory Notice letters to all customers of record. The letters instructed customers to: 1) continue to use their system(s); 2) place the safety advisory in the system's instructions for use; and, 3) to anticipate that their local service engineer will contact them to perform an inspection and schedule any necessary repairs by December 2014. Customers were also instructed to share the letter with their support staff and, if the equipment was sold or transferred, to forward the safety notice to the new owner of the equipment. Customers with questions can contact their local service representative or Siemens Medical Solutions at 800-888-7436. |  | 6 | 6 systems | Worldwide Distribution - USA (naitonwide) and Internationally to Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, Colombia, Costa Rica, Croatia, Czech Republic, Denmark, Ecuador, Egypt, Estonia, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Kuwait, Latvia, Lithuania, Malaysia, Mexico, Morocco, Myanmar, Netherlands, New Zealand, Norway, Oman, P.R. China, Pakistan, Peru, Philippines, Poland, Portugal, Qatar, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Venezuela, and Vietnam. | Open, Classified |  |  | 2213 | Not_Computer
Z-0914-2015 | 69992 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | e.cam -emission computed tomography system |  |  | April 08, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | There is a potential for patient finger injury on the e.cam systems. | Use error | The firm, Siemens, sent a "Customer Safety Advisory Notice" letter dated April 8, 2014 to all direct accounts (customers). The letter described the product, problem and actions to be taken. The letter instructed the customers to follow the instructions in the e.cam operator's manual and the e.cam Patient Bed Safety Addendum of the operator's manual; and to always monitor the patient at all times while they are on the system during initiation of any system movement. The letter further requested that customers forward the letter to the new owner of the equipment in the event that the system is sold or transferred. Customers with questions about this letter can contact Siemens Medical Solutions USA at 800-888-7463 (Monday - Friday, 7:00 AM - 8:00 PM EST). |  | 4462 | 4,462 systems | Worldwide Distribution: US (nationwide) including states of: AK, AR, AZ, CA, CO, CT, DE, FL, GA, HI, ID, IL, IN, KS, KY, MA, ME, MI, MO, MS, MT, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WY and Puerto Rico; and countries of: Algeria , Angola , Argentina , Australia , Austria , Azerbaijan , Bahrain , Bangladesh , Belarus , Belgium , Bolivia , Bosnia Herzeg. , Brazil , Bulgaria , Cameroon , Canada , Chile , Colombia , Costa Rica , Croatia , Cuba , Czech Republic , Dem. Rep. Congo , Denmark , Dominican Rep. , Ecuador , Egypt , El Salvador , Finland , France , Georgia , Germany , Ghana , Greece , Guam , Honduras , Hong Kong , India , Indonesia , Iran , Iraq , Ireland , Israel , Italy , Jamaica , Japan , Jordan , Kazakhstan , Kosovo , Kuwait , Latvia , Lebanon , Luxembourg , Macedonia , Malaysia , Mauritius , Mexico , Morocco , Myanmar , Nepal , Netherlands , New Zealand , Niger , Nigeria , Norway , Oman , P.R. China , Pakistan , Palestinian Ter , Peru , Philippines , Poland , Portugal , Qatar , Rep. of Yemen , Republic Korea , Romania , Russian Fed. , Saudi Arabia , Serbia , Singapore , Slovakia , Slovenia , South Africa , Spain , Sri Lanka , Sweden , Switzerland , Syria , Taiwan , Thailand , Trinidad,Tobago , Tunisia , Turkey , Uganda , Ukraine , United Arab Emirates , United Kingdom , Uruguay , Uzbekistan , and Vietnam. | Open, Classified |  |  | 4462 | Not_Computer
Z-0940-2015 | 69902 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | IS3000 Vision Side Cart (VSC) used in conjunction with the da Vinci Si Surgical System IS3000. |  |  | December 05, 2014 | 2014 | Intuitive Surgical, Inc. | Intuitive Surgical is initiating a voluntary correction related to da Vinci Si System (IS3000) Vision Side Carts, with threaded-in caster wheels, which can result in the caster wheel bolts gradually unthreading and the wheels loosening. | Device Design | The firm, Intuitive Surgical, sent an "Urgent Medical Device Correction" letter dated December 5, 2014 to all customers on December 5, 2014 by trackable method. The letter described the product, problem and actions to be taken. The customers were instructed to: ensure all affected personnel are fully informed of this notice; identify affected product, and retain a copy of this notice with your da Vinci Si user manual. Please note: the da Vinci Si Vision Side Cart can continue to be used safely. Ensure that all wheels are fully unlocked before attempting to move the Vision Side Cart. Avoid rolling the Vision Side Cart over uneven surfaces (i.e., over cables or a threshold).Intuitive Surgical is working on a newly designed wheel assembly. Uponavailability, Field Service will retrofit the wheel assembly of the affected VisionSide Carts with the new design.If you need further information or support concerning this Medical Device Correction, please contact your Clinical Sales Representative or Intuitive Surgical Customer Service at the numbers listed below:ÂżÂż North America and South America: 8008761310 Option 3 (6 a.m. to 5 p.m. PST)ÂżÂż Japan: 0120565635 or 0355751362 (9 a.m. to 6 p.m. JST)ÂżÂż South Korea: 0232713200 (9 a.m. to 6 p.m. KST)ÂżÂż Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020 (8 a.m. to 6 p.m.) or ics@intusurg.com |  | 202 | 202 | Worldwide Distribution: US (nationwide) and countries of: Belgium, France, Germany, Romania, India, Israel, Italy, South Korea, Norway, Slovak Republic, Spain, Sweden, Turkey, and United Kingdom. | Terminated | March 03, 2015 | 88 | 202 | Not_Computer
Z-0934-2015 | 70009 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | AlignRT is a video-based three-dimensional (3D) surface imaging system which is used to image the skin surface of a patient in 3D before and during radiotherapy treatment. |  |  | November 11, 2014 | 2014 | Vision Rt Inc | Potential use error when the external Gate Controller is turned ON after the Vision RT software has been opened. | Under Investigation by firm | VisionRT sent an Urgent Medical Device Correction Notification letter dated November 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers shall:A. Print the instructions in Appendix 3, place them in the front of the product guide and ensure these instructions are read and understood by all those who need to be aware of them within their organisation;B. Maintain awareness of this issue until all actions indicated in this letter have been successfully completed;C. Promptly inform Vision RT if they believe that patient harm occurred due to this issue; ANDD. Complete the acknowledgement in Appendix 1 and return it via email to corrections@visionrt.com.Contact Vision RTShould you have any queries on this letter, please do not hesitate to contact Vision RT by telephone on +44 20 83464300 (866 778-2379 from the US) or as per http://www.visionrt.com/contact/details.For questions regarding this recall call 866-594-5443. |  | 55 | 55 | Nationwide Distribution | Open, Classified |  |  | 55 | Computer
Z-0933-2015 | 69961 | No | GFD | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | CelluTome Epidermal Harvesting System Harvester (5.0 CM) STERILE EO Rx Only Manufactured for: KCI USA, Inc. Part Number CT-H50 or CT-H25 |  |  | November 19, 2014 | 2014 | KCI USA, INC. | KCI has received reports that, in a small number of cases, the CelluTome Harvester (a component of the CelluTome Epidermal Harvesting System), blades have the potential to drift during shipping and could prematurely protrude into openings of harvester top plate. | Device Design | KCI sent an Urgent - Voluntary Medical Device Field Notification letter to all affected consignees on December 8, 2014. The letter notified the customer of the issue, and provided instructions to inspect their harvester prior to use. If blade drift is identified, instructionswere provided to the customer to either return the device to KCI, or if they wish, how to reset the blade prior to use. The letter instructs the customer to respond to the recall notifcation by returning the product reconciliation form attached the letter. Follow-up notifcations will be send to non-responding consignees.Consignees with questions were instructed to contact their local KCI Representative or Customer Technical Service Center at 1-800-275-4524 and select option 3.For questions regarding this recall call 210-515-4108. |  | 5695 | 1,139/5 Packs | Worldwide Distribution - USA (Nationwide) and Internationally to Australia, Austria, Canada, France, Germany, Ireland, Italy, Isreal, Netherlands, South Africa, Spain, Sweden, Switerland, Turkey, UAE, and United Kingdom. | Terminated | September 18, 2015 | 303 | 5695 | Not_Computer
Z-0937-2015 | 69979 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Stryker Orthobiologics Imbibe Aliquot Needle Bone Cement Needle- Model Numbers: 2110-0505 Beveled Needle, 11 gauge x 4 inch |  |  | December 04, 2014 | 2014 | Orthovita, Inc., dBA Stryker Orthobiologics. | There is the potential for a breach in the inner or outer packaging pouches of all lots of the Stryker Orthobiologics Imbibe Aliquot Needle Bone Cement Needle. | Package design/selection | The firm, Stryker, sent an "URGENT PRODUCT RECALL" letter dated December 4, 2014, to end users/customers. The letter identified the product, problem, and actions to be taken. Customers were instructed to complete and return the URGENT MEDICAL DEVICE RECALL NOTIFICATION ACKNOWLEDGMENT FORM within 5 days via fax to 215-253-5020, contact other hospitals in your territory that have the affected products to arrange return of the product, and return the affected product to: Stryker Orthobiologics, 67 Great Valley Parkway, Malvern, PA 19355. If you have any questions, feel free to contact the Associate Manager, Quality Assurance and Regulatory Compliance at 610-640-1775 x 5299. |  | 2155 | 2155 | Worldwide Distribution: US (nationwide) and country of: Canada. | Terminated | July 28, 2015 | 236 | 91925 | Not_Computer
Z-0939-2015 | 69468 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HomeChoice Automated PD system and HomeChoice Pro Automated PD system Baxter's |  |  | October 06, 2014 | 2014 | Baxter Healthcare Corp. | There are additional and updated warning and cautions that are not in the Patient At-Home Guide for HomeChoice Automated PD Systems and HomeChoice PRO Automated PD Systems. | Labeling design | The firm, Baxter, sent an "IMPORTANT PRODUCT INFORMATION" letter dated October 6, 2014 to all peritoneal dialysis provider direct accounts. Baxter also sent an "IMPORTANT PRODUCT INFORMATION" letter dated October 9, 2014 to all peritoneal dialysis patient direct accounts The letters described the product, problem and actions to be taken. The letters included instructions for the recipients to complete the attached Customer Reply Form and to return it to Baxter via scanned e-mail to: fca@baxter.com or fax : 224-270-5457. In addition, the letter to patients instructed them to contact their healthcare provider to discuss this issue.Customers with clinical questions should contact Baxter's Renal Division Clinic Helpline at 888-736-2573, option 2 (8:00 AM - 5:00 PM CST). Adverse reaction or quality problems experienced with the use of these product can be reported by contacting Baxter Product Surveillance at 800-437-5176 (Monday - Friday, 8:00 AM - 5:30 PM CST) or e-mail to corporate_product_complaints_round_lake@baxter.com. Customers with general questions about the letter should contact The Center for One Baxter at 800-422-9837 (Monday - Friday, 8:00 AM - 5:00 PM CST). |  | 96458 | 96,458 systems | Worldwide Distribution: US (nationwide) Including states of: AK, AL, AR, AZ, CA, CO, CT, DE, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY and District of Columbia; and countries of: Algeria, Argentina, Aruba, Australia, Austria, Bahamas, Bahrain, Barbados, Belarus, Belgium, Bosnia, Brazil, Bulgaria, Cayman Islands, Chile, Colombia, Croatia, Cuba, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Estonia, Finland, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Iran, Iraq, Ireland, Israel, Italy, Jamaica, Japan, Jordan, Kazakhstan, Kuwait, Lebanon, Macedonia, Malaysia, Malta, Mexico, Morocco, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Poland, Portugal, Quatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Terminated | March 28, 2016 | 539 | 96458 | Not_Computer
Z-0936-2015 | 69422 | No | IKK | Isokinetic testing and evaluation system. | Physical Medicine | Physical Medicine | 510(K) Exempt | Cervical Extension Machine |  |  | December 03, 2013 | 2013 | MedX Holdings, Inc. | Review of DHRs revealed that HiPot, Continuity, load and friction re-testing was not completed on these devices during manufacturing. | Unknown/Undetermined by firm | Medx began contacting its consignees December 30, 2013, to schedule testing on the devices at their facility. The technician has been scheduled to complete required testing of these devices in the field.For questions regarding this recall call 322-351-2005. |  | 4 | 4 | Worldwide Distribution - USA including MO, IL, ID, and MN and Internationally to Germany and South Korea. | Open, Classified |  |  | 13 | Not_Computer
Z-0931-2015 | 69753 | No | HTC | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Synthes Small Electric Drive (SED), part number: 05.001.175: The Small Electric Drive is designed for use in general traumatology, especially hand and food applications involving surgical procedures such as drilling, burring, reaming, pin and wire placement, cutting of bone and hard tissue. |  |  | November 06, 2014 | 2014 | Synthes (USA) Products LLC | The device may operate solely in reverse mode, not operate in reverse mode when intended, not function, or unintentionally start. Unintended start of the hand piece or wrong mode/direction could cause bone/soft tissue, peripheral nerve, and /or vascular damage. Injuries may need surgical or medical intervention. Surgical delay and postoperative local infection are also potential risks. | Device Design | DePuy Synthes sent an Urgent Notice: Medical Device Removal letter dated November 6, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to notify their consignees using the hospital personnel letter, dated November 6, 2014, as their template for their consignee notification letter, review their inventory and remove the affected lots from stock.Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number. Complete the Verification Section (page 4 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided. Return the Verification Section (page 4 of this letter) with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132. Send a copy of the completed Verification Section to DePuy Synthes by: Fax: (610) 430-7083 or Scan/email: Fieldaction@synthes.comIf you DO NOT have the identified product, please take the following steps: Complete the attached Verification Section (page 4 of this letter) by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information. Return the documents to DePuy Synthes by: Fax: (610) 430-7083 or Scan/email: Fieldaction@synthes.com Note: If the Verification Section is answered on behalf of more than one facility and/or individual, please clearly indicate the name and address of the facility and/or individual on page 4 of the notification.Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 12 | 12 | US Nationwide Distribution to Florida only | Terminated | August 17, 2015 | 284 | 12 | Not_Computer
Z-0932-2015 | 69586 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter Amia Automated PD systems are used in the treatment of adult renal failure patients to perform automated peritoneal dialysis. |  |  | April 24, 2014 | 2014 | Baxter Healthcare Corp. | System error 01779 is produced when the battery cannot be charged due to specific voltage differences between the battery and the charger. | Under Investigation by firm | Baxter sent an IMPORTANT PRODUCT INFORMATION letter dated April 24, 2014 to all affected customers. The peritoneal dialysis provider consignees were instructed to: 1) complete the enclosed customer reply form and return it to Baxter; and, 2) contact Baxter Technical Services at 800-553-6898 should a System Event 01179 occur. Baxter Technical Services is available 24 hours a day, 7 days a week. Clinical questions can be directed to Baxter's Renal Clinical Helpline at 888-736-2543, option 2 (Monday - Friday, 8:00 AM - 5:00 PM CT). General questions regarding the Import Product Information letter can be direct to The Center for One Baxter at 800-422-9837 (Monday - Friday, 8:00 AM - 5:00 PM CT).For questions regarding this recall call 800-422-9837. |  | 55 | 55 units | Nationwide Distribution including GA, MA, NC, NY, PA, and WA. | Terminated | July 21, 2017 | 1184 | 55 | Not_Computer
Z-0941-2015 | 69806 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CSA Medical truFreeze System; Model: CC3-01, a cryogenic surgical device |  |  | November 10, 2014 | 2014 | CSA Medical | An increase in complaint trending for a loss of audible sound associated with the timer has been observed. | Software design | The firm, CSA Medical, sent a letter to its consignees 11/14/2014, via Fedex overnight . The letter explains the event and asks consignees to us a verbal call of time during the cryogen spray period until the root cause is determined and a permanent correction is implemented. Consignees are also asked to complete an acknowledgment form and return it to CSA Medical via, fax, email, telephone, etc. |  | 46 | 46 | US Distribution to states of: AR, CA, CO, CT, DC, FL, GA, IL, IN, LA, MA, MD, MI, NC, NJ, NY, OK, PA, RI, SC, TN, TX, UT, VA, and WV. | Terminated | March 25, 2015 | 135 | 46 | Not_Computer
Z-0930-2015 | 69987 | No | FST | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | HarmonyÂż vLED Surgical Lighting System, one unit packaged per box |  |  | November 26, 2014 | 2014 | Steris Corporation | STERIS has learned that groups of LED lights (modules) contained within the vLED light head may illuminate inconsistently. | Software Design Change | STERIS sent an Urgent Voluntary Field Correctionn Notice to all affected consignees via FedEx. The Consignee Notification mailing was initiated the week of December 15, 2014. Customers were instructed that their local STERIS representative would contact their facility to arrange for the installation of updated software. Customers with questions were instructed to call 1-800-548-4873.For questions regarding this recall call 440-392-7231. |  | 140 | 140 units | Worldwide Distribution - USA including Alabama, Alaska, Arkansas, California, District of Columbia, Florida, Illinois, Indiana, Kansas, Louisiana, Mississippi, Nevada, New Jersey, New York, North Carolina, Ohio, Oklahoma, Oregon, Pennsylvania, Puerto Rico, South Carolina, Tennessee, Texas, Utah, and Virginia and Internationally to Canada, China, Greece, Indonesia, Republic of Korea, Romania, Russian Federation, Taiwan, and Viet Nam. | Terminated | July 11, 2015 | 227 | 140 | Computer
Z-0946-2015 | 69937 | No | JIT | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension Vista Myoglobin Calibrator (MYO CAL), Lot 4FD085 |  |  | November 24, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Dimension Vista Myoglobin Calibrator (MYO CAL) may produce a positive shift in MYO QC and Patient test results that exceeds acceptance criteria for this product. A positive shift observed up to 12% at MYO concentrations within and above the reference range of the assay. Depending on quality control limits, this drift may not be detected. | Under Investigation by firm | The firm, Siemens, sent an "Urgent Medical Device Recall"letter dated November 2014, to all customers who received the affected lot. The letter described the product, problem and actions to be taken. The customers were instructed to: Review this letter with your Medical Director; discard your remaining inventory; recalibrate with an alternate in-date lot of MYO CAL, and complete and return the attached Field Correction Effectiveness Check Form and fax the form to (312) 275-7795. Siemens will replace any unused inventory of the affected lot at no charge.If you have any questions, please contact your Siemens Customer Care Center - Technical Solutions at 800-441-9250 or your local Siemens technical support representative. |  | 1230 | 1230 | Worldwide Distribution: US (nationwide) and country of: Canada. | Open, Classified |  |  | 1230 | Not_Computer
Z-0945-2015 | 70081 | No | OFA | Cardiovascular surgical instruments. | Cardiovascular | Cardiovascular | Enforcement Discretion | Pacemaker Tray (Cardiovascular Surgical Instruments convenience kit) |  |  | October 23, 2014 | 2014 | Centurion Medical Products Corporation | According to the recall notice received from Hospira, the recall was initiated due to a confirmed customer report of particulate in a single unit. Hospira identified the particulate as human hair, embedded in and attached to a pinched area of the stopper. | Nonconforming Material/Component | Centurion sent an Urgent Product Recall Notice dated October 23, 2014, to al affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify and remove all affected inventory that may be in their possession and to forward a copy of thie noctice to any cusotmers to whom the affected product may have been further distributed. Customers were asked to complete the enclosed accountability record whether not they have the product and fax to 517-546-3356. Customers with questions were instructed to call 517-546-5400 Ext. 1135.For questions regarding this recall call 517-546-5400, ext 1135. |  | 8 | 8 kits | Nationwide Distribution including MA, IN, NY, GA, AL, PA, and RI. | Terminated | June 08, 2015 | 228 | 1290 | Not_Computer
Z-0949-2015 | 70066 | Yes | NVZ |  |  | Cardiovascular | PMA | Merlin@home RF Remote Monitoring Transmitter, Model EX1150, with software versions 6.1M or 7.0. |  |  | December 18, 2014 | 2014 | St Jude Medical Cardiac Rhythm Management Division | Transmitter may initiate a software reset resulting in backup operation in some implanted St. Jude Medical Radio Frequency (RF) enabled Implantable Cardioverter Assura, Unify Assura, and Quadra Assura ICDs and Assurity and Allure Pacemakers. | Device Design | The firm, St. Jude Medical, sent an "Important Medical Device Correction" Dear Doctor Letter dated 12/18/14 to customers to inform them that the Merlin@home transmitters may initiate a software reset resulting in backup operation in some implanted St. Jude Medical Radio Frequency (RF) enabled Implantable Cardioverter Defibrillators (ICDs) and Pacemakers. The letter informs the customers of the problems identified and the actions to be taken. Customers with questions or concerns are instructed to contact their local St. Jude Medical representative or St. Jude Medical's Technical Services Department at 1-800-722-3774. |  | 257456 | 257,456 units | Worldwide Distribution: US (nationwide) and countries of: EU, Mideast, Canada, Australia/New Zealand, and Asia-Pacific (including Japan). | Terminated | March 30, 2016 | 468 | 257456 | Computer
Z-0948-2015 | 69300 | No | JJH | Clinical sample concentrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | The Abbott m2000sp is an automated system for performing sample preparation for nucleic acid testing. |  |  | November 23, 2013 | 2013 | Abbott Molecular | The x-drive assembly may fail prematurely. This could cause the manipulator or liquid handling arm to contact an object, shake and drip, or improperly pipet. | Nonconforming Material/Component | Abbott Molecular sent a different Field Correction Recall / Urgent Field Safety Notice letters dated October 24, 2014 to each of three sets of customers. Customers whose instruments' X-Drive Assembly belts were perviously inspected by a Field Service Representative and found to not require replacement were advised to contact their local Abbott Molecular Representative or Abbott Molecular Technical Support at 800-553-7042 x 2 should they hear any unusual noises associated with the LiHa or RoMa movement or misalignment of the LiHa or RoMa. Customers whose instruments' X-Drive Assembly belts were previously inspected and replaced by a Field Service Representative were notifieied that this maintenance activity took place. Customers with questions were instructed to contact Abbott Molelcular.For questions regarding this recall call 224-361-7000.Customers whose instruments' X-Drive Assembly belts have not been inspected were notified that a Field Service Representative is scheduled to visit their facility to inspect and if necessary, replace the X-Drive Assembly belt(s). Customers were advised to contact their local Abbott Molecular Representative should they hear any unusual noises associated with the LiHa or RoMa movement or misalignment of the LiHa or RoMa. |  | 176 | 176 devices | Worldwide Distribution - USA including CA, FL, GA, IL, MI, NC, NJ, NY, OH, OK, OR, and TX and Internationally to Australia, Austria, Belgium, Bosnia And Herzegovina, Brazil, Canada, China, Colombia, Costa Rica, Denmark, Dominican Republic, Egypt, Finland, France, Germany, Indonesia, Ireland, Israel, Italy, Kazakhstan, Kenya, Malawi, Mali, Mexico, Morocco, Mozambique, New Zealand, Nigeria, Norway, Pakistan, Poland, Russia, Saudi Arabia, Senegal, South Africa, South Korea, Taiwan, Thailand, Turkey, Ukraine, and United Kingdom. | Open, Classified |  |  | 176 | Not_Computer
Z-0492-2015 | 69787 | No | GAP | Natural nonabsorbable silk surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Ethicon PERMA-HAND Silk siliconized black braided 10 strands per packet Ethicon LLC- Made in USA. |  |  | October 17, 2014 | 2014 | Ethicon, Inc. | Ethicon is unable to ensure that PERMA-HAND silk suture (siliconized black braid) meets the expected sterility assurance level for sterile products. | Other | Ethicon sent an Urgent Medical Device Recall Letter/Business Reply Form dated 10/17/2014 via UPS next day mail. |  | 2508 | 2,508 each | Nationwide Distribution | Terminated | August 14, 2015 | 301 | 2508 | Not_Computer
Z-0954-2015 | 70063 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Proteus XR/a with wall stand model number 600-0301. |  |  | December 08, 2014 | 2014 | GE Healthcare | GE Healthcare has recently become aware of a potential safety issue involving the breakage of the two steel cables which support the bucky device in the wall stand of Proteus XR/a X-ray imaging systems.No injuries reported. | Nonconforming Material/Component | Consignees were sent on 12/8/2014 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 17107, dated December 8, 2014. The letter was addressed to Hospital Administrators /Risk Managers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product correction and Contact information (For questions in USA call 800-437-1171, For other countries contact your local GE Healthcare Service Representative. |  | 3096 | 3096 (2486 US, 610 OUS). | USA: All states in continental US including, Guam, DC.OUS; ARAB EMIRATES, AUSTRIA, BOSNIA AND HERZEGOVINA, BRAZIL, BULGARIA, CAMBODIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CZECH REPUBLIC, ECUADOR, EGYPT, FINLAND, FRANCE, GERMANY, GREECE, HONDURAS, HONG KONG, HUNGARY, INDONESIA, IRELAND, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, KUWAIT, LATVIA, LITHUANIA, MALAYSIA, MEXICO, NETHERLANDS, NEW ZEALAND, NORWAY, PAKISTAN, PHILIPPINES, POLAND, POLYNESIA, PORTUGAL, RUSSIAN FEDERATION, SAUDI ARABIA, SOUTH AFRICA, SOUTH KOREA, SPAIN, SWEDEN, SWITZERLAND, TAIWAN, THAILAND, UKRAINE, UNITED KINGDOM, URUGUAY, VENEZUELA, VIETNAM. | Terminated | February 03, 2016 | 422 | 3096 | Not_Computer
Z-0953-2015 | 70014 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | The ATTUNE INTUITION Impaction Handle is a re-useable instrument utilized in knee replacement procedures. The instrument interfaces with several related ATTUNE instruments: keel punch, tibial tower, evaluation bullet, fixed bearing tibial impactor, rotating platform tibial impactor, and femoral impactor, for the purpose of impaction, extraction, and insertion applications at steps within the surgical procedure. Due to the frequency of its use, two impaction handles are supplied in each kit. |  |  | November 10, 2014 | 2014 | DePuy Orthopaedics, Inc. | Complaints have indicated that the ATTUNE INTUITION Impaction Handle lever has fractured during impaction. If the lever fractures, there is the potential for fractured pieces of the device, or the inner spring, if released, to be left in the patient if not observed during surgery. | Device Design | On 11/10/2014, URGENT INFORMATION ATTUNE INTUITION IMPACTION HANDLE AND IMPACTORS DEVICE CORRECTION NOTICE notifications were sent to the affected distributors and medical professionals. All distributors were notified via electronic mail, and all medical professionals with affected inventory were also notified by distributor or Sales Consultant via in person delivery or mail or electronic mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, please contact Kim Earle, Recall Coordinator, at 574-371-4917 (M-F; 8 a.m. - 5 p.m. EDT). |  | 7944 | 7,944 | Worldwide Distribution-US (nationwide) except AK and RI, including PR and the countries of Chile, Canada, Australia, Austria, Belgium, France, Hong Kong, Ireland, India, Israel, Italy, Japan, Malaysia, Netherlands, New Zealand, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Switzerland, Thailand, United Arab Emirates (UAE), and United Kingdom (UK | Terminated | December 08, 2015 | 393 | 7944 | Not_Computer
Z-0950-2015 | 69419 | No | LPM | Soft (hydrophilic) contact lens. | Ophthalmic | Ophthalmic | PMA | Labeled in part: I Color Complete / Color Lens, blister packing, 14.5mm. |  |  | September 10, 2014 | 2014 | Intertrade Imports Inc. | A surveillance sample collected by FDA confirmed the presence of Bacillus Cereus and Bacillus Lentus in the contact lenses. I Color Color Complete brand color contact lenses are unapproved medical devices. | Material/Component Contamination | The firm verbally contacted each distributor on September 10, 2014 to notify them about the recall and actions that needs to be taken. |  | 750 | 750 units | Worlldwide Distribution - USA including Texas, California, Florida, and St. Lucia, Barbados and St. Marten. | Open, Classified |  |  | 3600 | Not_Computer
Z-0952-2015 | 70032 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | POWEREASE Instruments Set Screw Breakoff Instrument, REF 2346010, Medtronic Sofamor Danek, USA, Inc. Spinal surgical instrumrent. |  |  | December 03, 2014 | 2014 | Medtronic Sofamor Danek USA Inc | The retaining tabs component of the Set Screw Retaining Drive may be oversized, which will not allow proper mating between the Set Screw Retaining Drive and the set screw. | Process design | Medtronic sent an Urgent - Voluntary Market Removal letter dated December 3, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Instructions to Customers:1.) Please remove the product from normal storage locations. Do not use this product.2.) As applicable, contact your Medtronic Sales Representative to coordinate the return of affected product. Alternatively, you may call Medtronic's Global Quality Department at (800) 876 -3133, extension 3197 for assistanceIf you need assistance or have any related questions or concerns, please do not hesitate to contact your Medtronic Sales Representative or the Global Quality Department at (800) 876 - 3133, extension 3197. We sincerely apologize for this inconvenience, but this action is necessary to assure that our high standardof quality is maintained. |  | 24 | 24 units | US Distribution to the states of : CA, KS, KY, NC, TX, WA, MI, OH, IA, PA, OR, MS, FL, NY, GA and AR. | Terminated | September 24, 2015 | 295 | 24 | Not_Computer
Z-0947-2015 | 69872 | No | DHB | Radioallergosorbent (RAST) immunological test system. | Immunology | Immunology | 510(k) | Phadia 1000 Instrument (introduced as UniCAP 1000), Article number: 12-3800-01 (All instrument software versions since launch in 2003); Multiple ImmunoCAP assays for Allergen Testing, FDA Cleared for use on Phadia 1000 instrument with IU statement-see attachments. Software utilized in multiple ImmunoCAP assays for Allergen Testing, FDA Cleared for use on Phadia 1000 instrument with IU statement-see attachments. |  |  | November 20, 2014 | 2014 | Phadia US Inc | During an investigation of instrument logs it was determined that In specific circumstances involving multiple steps, a rack sequencing error may occur. This will result in a mismatch between the sample ID and the test result reported for all subsequent sample racks in that run. | Software design | ThermoFisher sent an Urgent Medical Device Product Correction letter to all affected customers. The firm will inform customers of the specific sequence of errors and operator actions that can trigger the event via letter and phone, and will request that they do not respond to the Barcode Read Error and let the sample with the error eject normally from the instrument and be rerun after the barcode read error has been resolved. The firm will issue a mandatory Phadia 1000 Instrument Software (ISW) update that will correct the issue and make efforts to install it within 45 days from date of availability. Customers were asked to return the written response form via fax to: 1-888-243-5214Attention: Customer Support, Phadia 1000 ISW Product Correction Response; EMAIL: Compliance-us.idd@thermofisher.comEmail Subject Line: Phadia 1000 ISW Product Correction Response.Customers with questions were instructed to contact Phadia US Technical Support at 800-346-4364, option 2.For questions regarding this recall call269-492-1957. |  | 146 | 146 | Nationwide Distribution including AZ, CA, CO, CT, FL, GA, HI IL, IN, KS, MD, MI, MN, NC, NJ, NV, NY, OH, OK, OR, PA, TN, TX, UT, VA, and WA. | Terminated | June 30, 2015 | 222 | 146 | Computer
Z-0958-2015 | 69587 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | Philips Expression MR200 MRI Patient Monitoring System. |  |  | October 22, 2014 | 2014 | Invivo Corporation | The MR 200 devices failed the initial power up test due to capacitor short and hi-pot testing; possible loss of Patient Monitoring or User/Service Electrical Shock. | Component design/selection | Philips Medical Systems sent an Urgent - Medical Device Correction Letter dated October 22, 2014, to all affected customers. The letter identified the product, the problem and the action needed to be taken by the customer. Our records indicate that you have an affected device. The following page provides additional instructions and actions that will be taken to address this problem. If you need any further information or support concerning this issue, please contact your local Philips representative: 877-468-4861 option 1, then option 2. |  | 20 | 20 | CA, FL, IN, KY, MN, NC, NJ, NY, OK, SD and TX Netherlands, Poland and United Kingdom. | Terminated | January 09, 2017 | 810 | 20 | Not_Computer
Z-0959-2015 | 69751 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | IC PROFILER-MR - Model 1123. |  |  | August 08, 2014 | 2014 | Sun Nuclear Corporation | Has a Potential Pinch Hazard to user performing Machine QA in an MRI environment. | Use error | Sun Nuclear Corp. sent a memo dated 08/08/2014 to their 3 customers of the possible manufacturing error. |  | 6 | 6 Units. | Distributed in the states of CA, CO, and OH. | Terminated | March 23, 2015 | 227 | 6 | Not_Computer
Z-0963-2015 | 69747 | No | MKJ | Automated external defibrillator system. | Cardiovascular | Cardiovascular | PMA | HeartStart MRx Monitor/Defibrillator; Model. Numbers: M3535A, M3536A, M3536J, M3536M, M3536MC, M3536M2, M3536M4, M3536M5, M3536M6 |  |  | November 19, 2014 | 2014 | Philips Medical Systems, Inc. | Device may malfunction, which could cause therapy to be delayed, disabled, or delivered inadvertently. | Software design | On November 19, 2014, Philips notified consignees of the recall via letter. The Urgent Medical Device Correction notification/Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.A software upgrade will be provided free of charge for all units affected by one or more of these issues. A Philips Healthcare representative will contact customers with affected devices to arrange for installation of the upgrade. Philips is asking customers to follow the Action to be Taken by Customer/User section of the Urgent Medical Device Correction notification/Field Safety Notice |  | 74499 | 74,499 devices | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 74499 | Computer
Z-0964-2015 | 70042 | No | MRD |  |  | General &amp; Plastic Surgery | 510(k) | MENTOR MemoryGel Resterilizable Gel Sizer Moderate Plus Profile STERILE. Breast prosthesis sizer. |  |  | December 11, 2014 | 2014 | Mentor Texas, LP. | The IFU provided in the affected units of the MENTOR MemoryGel Resterilizable Sizer is incorrect. Those units include the IFU for the MENTOR MemoryShape Resterilizable Sizer. | Employee error | Mentor sent an Important Labeling Correction letter dated December 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm plans to follow-up with sales representative visits. The recall letter notifies the customer of the issue, but does not instruct them to return the product. Should the customer chose to return the product, all returned product will be tagged to denote no further distribution. The recalled units under firm control have been placed on hold in the inventory system.If you have any questions related to this notice, please contact your local Mentor Sales Representative, or Mentor Customer Service at 1-800-235-5731. |  | 45 | 45 units | US Distribution to the states of : MA, WV, NY, FL, ME, NC, TX, IL, KS, TN, NJ, PA, CA, OR, WI, GA, and WA | Terminated | January 04, 2016 | 389 | 45 | Not_Computer
Z-0915-2015 | 69931 | Yes | NPT |  |  | Cardiovascular | PMA | Lotus TAVR 23mm, |  |  | November 19, 2014 | 2014 | Boston Scientific Corporation | Lotus valve became unlocked during release from the delivery system. This may lead to percutaneous insertion of another valve, resulting in a Valve in Valve (ViV) implantation. If this occurs, the first valve could embolize and it may be necessary to convert the patient to surgery. | Device Design | Boston Scientific notified consignees on November 19, 2014, of the problem via letter through overnight delivery directing consignees to segregate and return affected product to Boston Scientific utilizing the Field Safety Corrective Action (FSCA) Instructions provided. The FSCA details that consignees should: Immediately discontinue use of and segregate product; Complete and return the Account Reply Verification Tracking form and Package/Ship the affected product. |  | 278 | 278 units, all sizes | Distributed only in the countries of Finland, France, Germany, Great Britain, Italy, Norway, Spain, Sweden, and Switzerland. | Terminated | March 26, 2015 | 127 | 278 | Not_Computer
Z-0840-2015 | 69660 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE OEC Brivo 715 Prime Mobile C-Arm X-Ray Product |  |  | September 15, 2014 | 2014 | GE OEC Medical Systems, Inc | mA Accuracy may exceed the design and labeling specification of+ 10%. Per 1020.32(f): Deviation of x-ray tube potential and current from the indicated values shall not exceed the maximum deviation as stated by the manufacturer in accordance with 1020.30(h)(3). | Radiation Control for Health and Safety Act | Planned action(s) to repair defect or to bring product into compliance: GE Healthcare will, without charge, perform a mA calibration on all affected systems in the USA to bring the product back into compliance. The details of which will be included in a subsequent communication to customers or through a GE Healthcare field engineer site visit.Should customers have any questions or concerns regarding these corrections, please do not hesitate to contact the service team for further information at 800-874-7378 option 8. |  | 61 | 61 units installed in US | Worldwide Distribution | Terminated | May 20, 2015 | 247 | 61 | Not_Computer
Z-0965-2015 | 70010 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | MUM-1 (Multiple myeloma oncogene-1), catalog number PRM352 AA; |  |  | December 12, 2014 | 2014 | Biocare Medical Llc | Three lots of MUM-1 may not meet their labeled shelf life, as a drop in staining intensity over time has been observed. | Incorrect or no expiration date | Biocare sent a Medical Device Recall notifications letter to all customers by traceable method. The letters identified the affected product, problem and actions to be taken. Users are asked to immediately discontinue used of the affected lots and dispose of them according to applicable national, state and local laws. Users are requested to complete and return the response form so that proper credit can be issues 800-799-9499. |  | 59 | 59 | Worldwide Distrinution - US Nationwide in the states of TX, WI, MO, WV, SC, MA, WA, TN, AZ, CA, IL, NY, LA and countries of Kuwait, India, Venezuela, New Zealand, Turkey, Uruguay. | Terminated | March 30, 2015 | 108 | 59 | Not_Computer
Z-0960-2015 | 69995 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia T and Symbia T2 computer tomography systems. These systems are used to perform SPECT, CT, and SPECT + CT scans and imaging studies. |  |  | March 05, 2014 | 2014 | Siemens Medical Solutions USA, Inc. | There are confirmed reports of a component coming loose inside the CT enclosure and damaging the system.components. | Component change control | The firm, Siemens, sent an undated "Customer Safety Advisory Notice" letter to all customers. The described the product, problem and actions to be taken. The letters inlcuded instructions for customers to: 1) continue to use their systems while waiting for the field correction to be scheduled and performed; 2) ensure that the safety advisory is placed in the system's instruction for use; and, 3) closely observe the patient during studies. The letter also instructs customers of record to forward a copy of the letter to the new owner if the equipment was sold or transferred. Customers with any questions about this issue can contact Siemens Molecular Imaging at 800-888-7436 (US). |  | 295 | 295 systems | Worldwide Distribution: US Nationwide including Puerto Rico; Austria, Australia, Bangladesh, Belgium, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Czech Republic, Ecuador, Egypt, Finland, France, Germany, Great Britain, Hungary, Ireland, India, Iran, Italy, Japan, Mexico, Netherlands, New Zealand, Norway, Philippines, Poland, Portugal, Russia, Singapore, Slovakia, Slovenia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, and Tunisia. | Open, Classified |  |  | 295 | Not_Computer
Z-0971-2015 | 69957 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Force; the intended use of computed tomography is to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angels or spiral planes taken at different angles. |  |  | December 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Artifacts are found in acquired imaging on the SOMATOM Force. | Software Manufacturing/Software Deployment | A customer advisory notice, dated December 3, 2014, was sent to end users that identified the product, problem, and action to be taken. Customers were informed that new software would be installed via remote access. |  | 7 | 7 | PA, MA, MN, SC, MD, and TX. | Terminated | April 06, 2015 | 124 | 7 | Computer
Z-0978-2015 | 70060 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Timberline Anchored Lateral Variable Retractable Sleeve LITPCR8600-2101-07, Model No. 86300204 |  |  | August 22, 2013 | 2013 | Biomet Spine, LLC | The inner diameter of the Sleeve shaft is undersized, resulting in interference fit when inserting the Anchored Lateral Straight Retractable Awl or the Anchored Lateral Straight Retractable Drill into the Sleeve shaft. | Component design/selection | The firm, Lanx, sent their Consignees/customers a letter dated October 15, 2013. The letter described the product, problem and actions taken. Lanx and their distributors replaced the affected inventory with the conforming product.Please contact the Chief Compliance Officer at 303-501-8534 or by email: Randy.fesmire@lanx.com. |  | 12 | 12 units | US Distribution to states of: AZ, CA, NY, TN, and TX including PR. | Terminated | March 13, 2015 | 568 | 12 | Not_Computer
Z-0976-2015 | 69091 | No | JSG | Multipurpose culture medium. | Microbiology | Microbiology | 510(K) Exempt | Columbia Agar with 5% Sheep Blood |  |  | June 06, 2014 | 2014 | Becton Dickinson &amp; Co. | A low level of surface and subsurface contamination of Listeria monocytogenes on 9 lots of non-sterile prepared plated media containing sheep blood, typically visible immediately upon removal from the packaging. The contamination is typically observed as 1-3 CFUs per plate when present. | Material/Component Contamination | BD Diagnostic Systems sent a Field Correction Notice dated August 2014 to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to continue to follow all surveillance requirements of their laboratory upon confirmation of Listeria monocytogenes. Laboratory staff should notify their respective state health department. Customers with questions should contact BD Technical Services at 800-638-8663.For questions regarding this recall call 410-316-4000. |  | 775 | 775 cases | Worldwide Distribution - US nationwide including Puerto Rico, Belgium, Canada, Philippines, Singapore, and Taiwan. | Terminated | August 12, 2016 | 798 | 9922 | Not_Computer
Z-0970-2015 | 70082 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | EXp Tibial Insert ; Intended Use of Device(s): The Proven Gen-Flex Total Knee System is intended for: a. Total knee replacement due to osteoarthritis, osteonecrosis, rheumatoid arthritis, and/or post-traumatic degenerative problems; and b. Revision of failed previous reconstructions where sufficient bone stock and soft tissue integrity are present. |  |  | December 01, 2014 | 2014 | Stelkast Co | Difficulty locking EXp tibial inserts into tibial trays, caused by a manufacturing condition where a lead in radius on the anterior locking tab is not present on certain insert lots. The absence of this lead in may prevent or make assembly of the insert and tray difficult. | Process design | On December 1, 2014, phone calls were made to all direct accounts to notify them of the recall and to request immediate recovery and return of the SC3425 EXp PS Tibial Tray Inserts and SC3453 EXp CR-HF Inserts from their inventory. |  | 553 | 553 | TX, MD, IN, KS, CA, PA, GA, WA, NY, IL. | Terminated | May 06, 2015 | 156 | 553 | Not_Computer
Z-0973-2015 | 70018 | No | GAD | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Malleable C- Retractor, Part Number 397.232, used to provide retraction of the soft tissues during use of trocar instrumentation. Used in the following systems: the Basic Trocar System, Universal Trocar System, Craniofacial Modular Fixation System, Craniofacial (CMF) Distraction System, MatrixORTHOGNATHIC Plating System, and the Matrix MANDIBLE Plating System. |  |  | December 15, 2014 | 2014 | Synthes, Inc. | Potential failure and/or corrosion of the internal spring of the retractor, which may have been manufactured from an incorrect material. Surgical delay may occur. Spring particles could transfer from the instrument to the patient, and may lead to an adverse tissue reaction and/or infection. | Component design/selection | The firm, DePuy Synthes, sent an "URGENT NOTICE: MEDICAL DEVICE RECALL" letter dated December 15, 2014, to its end users and sales consultants. The letter identified the product, problem, and actions to be taken. The End users/sales staff were instructed to immediately remove affected stock from inventory; call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number; complete and return the attached Verification Section to DePuy Synthes via fax: 866-827-3421 or Scan/email: Synthes8645@stericycly.com, even if you do not have any of the affected product, and return any affected devices to Synthes.If you have any questions, please call 610-719-5450 or contact your DePuy Synthes sales consultant. |  | 1762 | 1762 | Worldwide Distribution: US (nationwide) and internationally to Canada. | Open, Classified |  |  | 1762 | Not_Computer
Z-0854-2015 | 69478 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Artis zee/ zeego systems, model numbers: 10094135, 10094137, 10094139, 10094141, 10280959: dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed include cardiac angiography, neuro angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | September 17, 2014 | 2014 | Siemens Medical Solutions USA, Inc | The Small Focus of a 3 focal spot x-ray tube may fail and further release of radiation will not be possible. The Large Focus and Micro Focus focal spots will still be operational without any restriction. Manual switchover is needed, resulting in a short interruption of the procedure. | Process control | Planned action (s) to repair defect or to bring product into compliance for Siemens Medical Solutions Inc. concerning the corrective action plan (CAP) for the Model(s) Artis Zee/Zeego Stationary Angiographic X-ray Systems.1.Notifying the customer of the defect.2.Implementing a corrective software patch to protect the system from damage inthe event of hardware failure3. Scheduling a service visit from a Siemens representative to perform a hardware replacement.4. Replacing the non-conforming X-ray tube with one manufactured after the nonconformity issue was resolved.For further questions please call (610) 219-6300. |  | 44 | 44 | USA Distribution to the states of MO, MI, NY, OK, CA, LA, CT, IN, VA, MD, WA, HI, ME, TX, NY, OH, PA and CO. | Terminated | August 06, 2015 | 323 | 44 | Computer
Z-0977-2015 | 70065 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | lnline Persuader- PedFuse Pedicle Screw System; Part Number: 11-80164 |  |  | December 19, 2014 | 2014 | SpineFrontier, Inc. | A crack may develop in the handle to the Inline Persuader, which may progress if the user continues to use the affected instrument. | Device Design | Spinefrontier sent an Urgent Voluntary Recall Notification letter dated December 16, 2015, to all affected customers. The notice informed customers of the recall and requests that they initiate the immediate return of the parts listed. In addition, SpineFrontier is requiring confirmation via recall acknowledgement form that consignees have received this notice and that they will initiated the immediate return of the parts. Customers can contact SpineFrontier at 978-232-3990 Monday through Friday 8:30AM to 5:30PM EST with any questions. |  | 4 | 4 devices | Nationwide Distribution including TX, OH, VA, and Washington, DC. | Terminated | January 11, 2016 | 388 | 4 | Not_Computer
Z-0967-2015 | 69830 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 3600 Immunodiagnostic System, Catalog Number 6802783, IVD --- Ortho Clinical Diagnostics. |  |  | November 19, 2014 | 2014 | Ortho-Clinical Diagnostics | Software Anomaly: the firm has identified an anomaly with VITROS System Software Version 3.1 and below, and determined that the software may not properly identify an expired calibration. | Software design | On 11/19/2014, IMPORTANT PRODUCT CORRECTION NOTIFICATION letter (Ref. CL2014-296, dated 11/19/2014) was sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to the US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and advise them of the required actions. Immediately install the most recent ADD using the All Assay Data option. If any of your calibrations are expired, the system will post PZB-005 condition codes. Perform calibrations prior to using those reagents. As expected, calibrations that will expire within one hour will also be flagged. Continue to install the most recent ADD at least once each week using the All Assay Data load option so that the system properly flags expired calibrations. --- Foreign affiliates were informed by e-mail on 11/19/2014 of the issue and instructed to notify their consignees of the issue and required actions. |  | 658 | 658 total: USA - 144, Foreign - 514 | Worldwide Distribution-US (nationwide) including Puerto Rico and the countries of Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Panama, Singapore, Spain and Venezuela. | Open, Classified |  |  | 2827 | Computer
Z-0972-2015 | 70016 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis Q, and Artis Q.zen ceiling systems; model numbers: 10848281; 10848354: The Artis family is dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed include cardiac angiography, neuro angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | December 03, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Suboptimal routing of the cable may result in increased wear over time. Without additional measures, damage of a cable may result in limited functionality and in rare cases, failure of a system. An ongoing procedure could be terminated in rare cases. | Device Design | Siemens sent an important customer safety information notice, dated December 3, 2014, to all affected customers. The letter identified the product, problem, and action to be taken. The following hardware modifications will be implemented with the corrective action AX042/14/S:1. A modified cable outlet to provide additional space for the cabling.2. A modified plastic cover at the cable outlet providing optimal cable guidance in the area of the arc. In general all systems will be checked for existing damage prior to the implementation of this action.For questions regarding this recall call 610-219-6300. |  | 50 | 50 | Nationwide Distribution | Terminated | April 06, 2015 | 124 | 50 | Not_Computer
Z-0966-2015 | 70073 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare, Discovery IGS 730 &amp; Discovery IGS 740. |  |  | December 10, 2014 | 2014 | GE Healthcare | Potential unintentional system motion or rotation due to a positioning user interface issue with Discovery IGS 730 and 740 products. | Software design | Consignees were sent a 12/10/2014 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 12223 dated December 10, 2014. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information ( US call 800-437-1171, other countries contact your local GE Healthcare Service Representative). |  | 24 | 24 ( 7 US, 17 OUS) | Worldwide Distribution-Distributed in the states of CO, CT, MS, NJ, NY, OH, and WA, and the countries of BELGIUM, CANADA, CHINA, FRANCE, GERMANY, JAPAN, MEXICO, PERU, UNITED ARAB EMIRATES, and UNITED KINGDOM. | Terminated | June 01, 2015 | 173 | 24 | Not_Computer
Z-0985-2015 | 70074 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Revolution CT., multi-slice (256 detector row) CT scanner. |  |  | December 19, 2014 | 2014 | GE Healthcare | Potential safety issue due to a required torque check on three bolted joints of the Revolution CT scanner.No injuries or illnesses reported. | Process control | Consignees were sent on 12/19/2014 a GE Healthcare "Urgent Medical Device Correction" letter GEHC # 25457 dated December 19, 2014. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology / Cardiology and Radiologists / Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information (US phone 800-437-1171, Japan 0120-055-919, other countries contact your local GE Healthcare Service Representative). |  | 13 | 13 (7 US, 6 OUS) | US: FL, WA, IL, UT, NY, CA.OUS: Canada, France, Hong Kong, Japan, Switzerland, United Arab Emirates. | Terminated | April 29, 2015 | 131 | 13 | Not_Computer
Z-0983-2015 | 70070 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Timberline Anchored Lateral Fixed Retractable Sleeve Model No. 8630-0205. Timberline Anchored Lateral Fixed Retractable Sleeves are used with the Timberline MPF System to help establish access into the vertebrae for the screws utilized to anchor the plate to the vertebrae. The Timberline Retractable Drill and the Timberline Retractable Awl fit within this Sleeve. |  |  | August 22, 2013 | 2013 | Biomet Spine, LLC | The inner diameter of the Sleeve shaft is undersized, resulting in interference fit when inserting the Anchored Lateral Straight Retractable Awl or the Anchored Lateral Straight Retractable Drill into the Sleeve. | Device Design | The firm, Lanx, sent a letter to their Consignees/customers dated October 15, 2013. The letter describes the product, problem and actions taken. Lanx and their distributors replaced the affected inventory with conforming product. Please contact Chief Compliance Officer at 303-501-8534 or by email: Randy.fesmire@lanx.com if you have any questions regarding this recall. |  | 14 | 14 units | US Distribution to states of: AZ, CA, NY, TN and TX including PR. | Terminated | March 13, 2015 | 568 | 14 | Not_Computer
Z-0990-2015 | 69997 | No | FMQ | Protective restraint. | General Hospital | General Hospital | 510(K) Exempt | Civco, MTSR03 Shoulder Retractor with Wrist Loops Kit with Base, plus two pairs of wrist loops (nylon and sherpa), Non-sterile re-usable wool wrist strap, REF: 20SR03, Rx |  |  | December 11, 2014 | 2014 | Med Tec Inc | The product (wool wrist loop) was manufactured 7cm shorter than product specifications. | Component design/selection | The firm sent e-mailed Field Safety Notice letters and contacted customers via phone on December 11, 2014, informing customers of the issue with the product. Customers were advised they may accept the product as is or request a replacement. A response form was enclosed for consignees to indicate the status of the product.Please contact a CIVCO insides sales representative for additional assistance. You may reach your representative by calling CIVCO Medical Solutions, Radiation Oncology, at 800-842-8688 or 712-737- 8688 between 8:00AM and 5:00PM CST. |  | 2 | 2 | Worldwide Distribution - US Distribution to the states of : AZ, MO, PA and WI., and the countries of : Austria, India and Japan. | Open, Classified |  |  | 13 | Not_Computer
Z-0957-2015 | 69960 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | iPlan RT is a radiation treatment planning system that is intended for use in stereotactic, conformal, computer planned, Linac based radiation treatment of cranial, head and neck, and extracranial lesions. |  |  | November 19, 2014 | 2014 | Brainlab AG | iPlan RT Radiation Treatment Planning Software: Potentially incorrect patient positioning when using multiple localized CT image data sets. | Software in the Use Environment | Brainlab sent a FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letter dated November 19, 2014, to all affected customers. The letter included instructions to: 1) If clinically not required, do not use multiple localized CT scans within one treatment plan (to avoid potentially incorrect Reference and Alignment Set assignments in the first place). 2) If you must use multiple localized CT scans, e.g. for recurrent treatment planning of the same patient, always guarantee that the latest (see note *) CT scan is defined as both Alignment Set and Reference Set during treatment planning. Brainlab will provide a software solution to prevent the described scenario from occurring. Brainlab will actively contact affected customers tentatively starting January 2016 to schedule the update. Brainlab will additionally refine the instructions for use of the iPlan RT treatment planning software in regard to implications of Alignment Set and Reference Set selections and will provide this user guide update to existing affected customers together with the software update.Customers with questions can contact Brainlab at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 1412 | 1412 systems total | Worldwide Distribution - USA including Alabama, Arizona, Arkansas, California, Colorado, Connecticut, Florida, Georgia, Idaho, Illinois, Indiana, Iowa, Kansas, Kentucky, Louisiana, Maine, Maryland, Massachusetts, Michigan, Minnesota, Mississippi, Missouri, Montana, Nebraska, Nevada, New Hampshire, New Jersey, New York, North Carolina, North Dakota, Ohio, Oklahoma, Oregon, Pennsylvania, Rhode Island, South Carolina, South Dakota, Tennessee, Texas, Utah, Vermont, Virginia, Washington, Wisconsin, Wyoming and Internationally to Algeria, Angola, Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Japan, Kazakhstan, Latvia, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, Norway, Oman, Panama, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Kingdom, Uruguay, and Venezuela. | Open, Classified |  |  | 1412 | Computer
Z-0982-2015 | 70056 | No | OVD | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Anchored Lateral Retractable Drill, Timberline IPF IFU IFU8621-0113, Model 8630-0102. |  |  | August 22, 2013 | 2013 | Biomet Spine, LLC | The outer diameter of the Drill shaft is oversized; thereby, resulting in interference fit when inserting the Drill into the Fixed or Variable Sleeve Assembly. | Component design/selection | Distributors and Direct Sales Reps notified via letter to return product. |  | 15 | 15 | Distributed in PR and the states of NY, AZ, TN, TX, and CA. | Terminated | March 13, 2015 | 568 | 15 | Not_Computer
Z-0986-2015 | 70093 | No | LWH | Surgical sponge scale. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | The SurgiCounter Charging Cradle |  |  | June 16, 2013 | 2013 | Stryker Instruments Div. of Stryker Corporation | Faulty SurgiCounter Charging Cradle can prevent the scanner from charging or downloading information after a procedure. | Device Design | On 6/16/2013, Improvements to the SurgiCounter" charging base notifications were emailed to the affected users with instructions for requesting a replacement product. The notification included a description of the reason for the recall, affected product, and instructions for responding to the notification. For questions, call the Support team any time at (877) 520-2300 option 1, or support@scmd.com. |  | 3625 | 3625 | US Nationwide Distribution in the states of AK, AL, AR, AZ, CA, CT, FL, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, ND, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, and WY. | Terminated | August 20, 2015 | 795 | 3625 | Not_Computer
Z-0979-2015 | 69969 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Various Brilliance and Ingenuity Computed Tomography X-ray systems; one unit per package |  |  | December 08, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Improper assembly of the vertical motor/brake system may lead to uncommanded vertical motion. | Process control | Philips sent an Urgent - Field Safety Notice letter dated December 4, 2014 to all affected customers. The US consignees will receive this notice via certified mail and Philips Key Market representatives will distribute the FSN internationally. Additionally, A Field Service Engineer will service the system to resolve the issue. For questions contact your local Philips representative or local Philips Healthcare office. |  | 500 | 500 units | US Nationwide Distribution - AR, AZ, CA, CO, FL, FL, GA, IL, IN, MD, MI, MN, NC, ND, NJ, NY, OH, OK, OR, PA, PA, SC, TN, TX, VA, and WA. The affected product was also distributed to one government consignee. The affected product was also distributed to the following foreign countries: Canada, Albania, Algeria, Australia, Austria, Bangladesh, Belgium, Brazil, Cayman Islands, China, Czech Republic, Denmark, Ecuador, Ethiopia, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Libyan Arab Jamahiriya, Malaysia, Mauritius, Mexico, Myanmar, Netherlands, New Zealand, Nicaragua, Norway, Oman, Panama, Poland, Russia, Saudi Arabia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Tunisia, United Kingdom, and Vietnam. | Terminated | October 01, 2015 | 297 | 500 | Not_Computer
Z-0987-2015 | 70078 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry XPT interfaced to the CentraLink Data Management System V14x in specific configurations: The CentraLink system software is a network solution provider and multi-system data manager for the instruments and lab automation systems (LAS) within the lab. The CentraLink software consolidates data from all connected instruments so that an operator can review and edit patient and quality control results from a single location. |  |  | December 16, 2014 | 2014 | Siemens Healthcare Diagnostics, Inc. | Enabling sending of preliminary/initial results on the ADVIA Chemistry XPT creates a risk that a critical result could be overwritten by the same result rather than showing the repeated result on CentraLink. A result could be erroneous and critical and appear to be verified upon repeat. | Under Investigation by firm | A Siemens service provider visited each site to verify and/or update customer configuration on the ADVIA Chemistry XPT and CentraLink Systems to ensure proper reporting of results. An Urgent Medical Device Correction letter was issued on December 16, 2014 to all affected customers. Customers were also provided a Customer Bulletin providing instructions on how configuration setting should remain. |  | 7 | 7 | Distributed in the state of WA. | Open, Classified |  |  | 7 | Not_Computer
Z-0981-2015 | 69767 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | Drill set long, drill stop compatible for 4.8mm diameter implant, Article Number 040.447S Endosseous dental implant drill set. |  |  | October 27, 2014 | 2014 | Straumann USA, LLC | Drill set may contain incorrect drill. | Process control | Straumann sent a Field Safety Corrective Action by letter dated October 29, 2014 which was delivered via United Parcel Service. The letter identified the product the problem and the action needed to be taken by the customer.Customers are asked to examine their inventory and return the product. In addition, customers are asked to complete a Customer Confirmation Form and return it to Straumann via fax to (978)747-0023. |  | 362 | 362 pieces | Nationwide in the following states: AZ, CA, CO,CT, FL, IA, IL, IN, LA, MA, MD, ME, MN, MO, NC, NJ, NY, OH, OK, OR, PA, RI, TX, VA and WV. Worldwide in the following countries by Institut Straumann AG distribution network: Austria, Australia, Canada, Germany, Spain, Finland, France, United Kingdom, Netherlands, Belgium, Switzerland, Ireland, Italy. | Terminated | April 10, 2015 | 165 | 362 | Not_Computer
Z-0992-2015 | 69739 | No | KHO | Fluorometer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Infinite 200PRO |  |  | November 10, 2014 | 2014 | Tecan US, Inc. | Incorrect lumi firmware version installed (E.027 instead of V2.00) | Employee error | Tecan sent an Urgent Field Corrective Action letter dated October2014, to all affected consignees by via FedEx on/or about November 10, 2014. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed not to use te kinetic luminescence measurement workflow requiring movement of th eplate transport out during kinectice measurements until the firmware has been upgraded to current version 2.0. A Tecan service representative will contact you to arrange a time to install the firmware. Additionally, they may want to review prior measurement results obtained with this workflow. Customers with questions were advised to contacte their local Tecan Helpdesk.For questions regarding this recall call 919-361-5200, ext 19524. |  |  |  | Worldwide Distribution - USA including MD and Internationally to Hong Kong, South Korea and the UK. | Terminated | December 21, 2015 | 406 |  | Computer
Z-0980-2015 | 69956 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens Linear Accelerator: ARTISTE, ONCOR Avant Garde, ONCOR Impression Plus, ONCOR Expression, ONCOR Impression, PRIUS HI, Mevatron M2/Primus Mid-Energy, MEVATRON PRIMART, MEVATRON KD-2, MEVATRON KDS-2, MEVATRON K2, MEVATRON 6323-2, MEVATRON M2 6740, MEVATRON M2 6700, MEVATRON 6300, MEVATRON MDX-2, MEVATRON MD-2, MEVATRON KD, MEVATRON KDS, MEVATRON 77, MEVATRON K7467, MEVATRON 7767, MEVATRON K8067, MEVATRON M 6300 (MX), MEVATRON M 6730, MEVATRON M 7140, MEVATRON M 7400, MEVATRON MDX, MEVATRON MD (MD, MDX), MEVATRON M 7445 (MXE), MEVATRON M 7400/7440 (MXE), MEVATRON M 6730/6740 (MXE), MEVATRON M 6700 (MX): the intended use of the MEVATRON , PRIMUS , ONCOR and ARTISTE MV System linear accelerator system is to deliver x-ray radiation for therapeutic treatment of cancer. |  |  | December 02, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Use of any unauthorized third-party components on Siemens LINACs may lead to patient or operator injury, damage to the LINAC, and can potentially compromise customer warranty. The dark current dose must be considered in the overall dose calculation, or an overdose and serious injury may occur. | Component design/selection | An urgent field safety notice, dated December 2, 2014, was sent to end users advising them to use only Siemens' authorized components for patient treatments as described in the operator's manual. |  | 320 | 320 | Worldwide Distribution - US Nationwide including Canada and Mexico. | Terminated | May 12, 2015 | 161 | 320 | Not_Computer
Z-0988-2015 | 70024 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Femto LDV Z-Generation Femtosecond Surgical Lasers are used to perform ophthalmic surgeries. |  |  | December 18, 2014 | 2014 | Ziemer Usa Inc | The devices were not shipped with the Starter Kits. | Labeling False and Misleading | Notification letters dated December 18, 2014 were sent to all direct account s(customers). The letters were included with the Starter Kit referred to in the Surgical Procedure Manual. |  | 26 | 26 lasers | US: Nationwide (AL, CA, FL, GA, IL, MN, MO, NC, ND, NH, NJ, NM, NY, PA, TX, VA, WA) | Open, Classified |  |  | 26 | Not_Computer
Z-0994-2015 | 70020 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Ysio Max, Luminos dRF Max and Agile Max systems with software version VE10 and SmartOrtho license. |  |  | December 15, 2014 | 2014 | Siemens Medical Solutions USA, Inc | Potential for composed images to be flipped before being sent to PACS on systems with software version VE10 and SmartOrtho license. Flipped images may be reversed so the associated annotations, e.g. labels (R/L), may be displayed incorrectly, potentially leading to misdiagnosis. | Software design | The firm, Siemens, sent a "Safety Advisory Notice - XP080/14/S" letter dated December 8, 2014, to end users that identified the product, problem, and actions to be taken. Siemens provided instructions to avoid the issue and indicated a software fix would be released as a permanent resolution. The customers were instructed observe this safety notice and comply with the corresponding measure until the update has been fully completed.If you have any questions, contact the Regulatory Technical Specialist at 610-219-2119. |  | 11 | 11 | US Distribution to states of: ID, IL, MI, MN, MO, ND and PA. | Open, Classified |  |  | 11 | Computer
Z-0996-2015 | 69671 | No | GBB | Operating tables and accessories and operating chairs and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Powered stretcher chair used during facial and eye surgeries. One chair per package. |  |  | November 14, 2014 | 2014 | TransMotion Medical Inc | The firm is finding that the stretcher chair controllers are moving the chairs involuntarily. | Component design/selection | TransMotion Medical Inc. sent an Urgent Medical Device Recall letter dated November 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were instructed:Discontinue use of the device and contact TransMotion Medical Inc. for a replacement. Suspected faulty pendants must be destroyed by severing the cord and discarding the control. Replacement of the pendant is a long-term solution to the issue due to replacement pendants being not of the same type or design as the failed pendant.Accompanying the recall letter will be an acknowledgment and receipt form that is required to be completed by the customer and returned to TMM via fax or email. This form will confirm that the product resides at the customer site where the unit was originally shipped. If the unit is no longer present at the customer site the customer can indicate this on the form and provide further information as to the whereabouts of the device if known. Ship replacement pendant control to the customer free of charge with installation instructions as to how to replace the pendant on the device. The firm will track non-responders and initiate follow-up contact.For further questions please call Phone: 330-239-4192 ext. 175 or Toll free: 1-866-860-8447 (7:30 AM to 4:00 PM Eastern Time Monday through Friday) |  | 1127 | 1,127 units | Worldwide Distribution - US Distribution and the countries including Canada, Australia, Chile, China, Costa Rica, Germany, Algeria, Israel, Italy, South Korea, Malaysia, Nigeria, Netherlands, Poland, Portugal, Saudi Arabia, Singapore, Thailand, and Venezuela. | Terminated | July 11, 2015 | 239 | 1127 | Not_Computer
Z-0993-2015 | 69560 | No | OUG | Medical device data system. | General Hospital | General Hospital | 510(K) Exempt | TEGRIS System manufactured by MAQUET GMBH in Germany |  |  | August 21, 2014 | 2014 | Maquet Medical Systems USA | The wrong movement results on the MAGNUS OR table. The button commands on the TEGRIS touchscreen for Lower Leg Up and Upper Leg Down are switched in the software for the integration with the MAGNUS operating table system. | Software design (manufacturing process) | Maquet sent an Urgent Medical Device Field Correction Letter dated August 21, 2014, to all affected customers via Federal Express. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following measures until the issue is corrected: The TEGRIS equipped with a software module version up to 2.1.X.X should not be used in conjunction wiht a MAGNUS operating table system. All other MAQUET operating tables are not affected.A MAQUET service representative would contact them and arrange to have their TEGRIS software module updated. Customers with questions were instructed to contact MAQUET Technical Support at 1-888-627-8383 (press option 3, followed by option 1, then option 1 again.For questions regarding this recall call 973-709-7000. |  | 6 | 6 units | US Nationwide Distribution including NE, NY and PA. | Terminated | November 17, 2015 | 453 | 6 | Computer
Z-0997-2015 | 69742 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | Philips IntelliVue TcG10; Monitor, carbon-dioxide, cutaneous. |  |  | November 19, 2014 | 2014 | Philips Medical Systems, Inc. | Philips has discovered that the Instructions for Use (IFU) for the IntelliVue TcG10 Transcutaneous Gas Measurement Module is missing the contraindication that the device is not to be used on patients under gas anesthesia. | Under Investigation by firm | Customers were notified of the recall by letter sent via UPS on November 19, 2014. The Field Safety Notice informs customers of the issue, identifies details of the units affected, gives instructions on actions to be taken by the customer and identifies what action Philips plans to take to remedy the issue.Philips asked customers to do the following: Upon receipt of this notification, ensure that the IntelliVue TcG10 addendum to the Instructions for Use is being reviewed and implemented by all members of your staff who operate the device. Review this information with all staff members who are involved in the operation of the IntelliVue TcG10 modules and need to be aware of the contents of this communication. The addendum should be stored with the IntelliVue TcG10 Instructions for Use. If you need any further information or support concerning this issue, please contact your local Philips representative or call us at 1-800-722-9377. |  | 86 | 86 | Distributed in the states of CA, HI, IL, MA, MI, MO, MT, NC, OH, and PA. | Terminated | September 29, 2015 | 314 | 86 | Not_Computer
Z-0995-2015 | 69978 | No | FPO | Wheeled stretcher. | General Hospital | General Hospital | 510(K) Exempt | Ferno Model 35X PROFlexx Stretchers, one unit per package |  |  | December 03, 2014 | 2014 | Ferno-Washington Inc | The wheel castor assemblies may be loosening on the stretchers. | Mixed-up of materials/components | Ferno initiated contact with customers via email and telephone on Wednesday, December 3, 2014. The firm will replace all wheel caster assemblies. |  | 46 | 46 units | The affected product was distributed to the following states: OH, MA, HI, NJ, TN, TX, MO, and GA. The affected product was distributed to the following foreign countries: Canada, Germany and Venezuela. | Terminated | December 01, 2015 | 363 | 46 | Not_Computer
Z-0998-2015 | 70080 | No | LIT | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Cordis POWERFLEX PRO PTA Dilatation Catheter |  |  | December 22, 2014 | 2014 | Cordis Corporation | One unit from a demo lot of POWERFLEX PRO PTA Dilatation Catheter was left at a hospital and was not sterilized. | Labeling design | The firm, Cordis, sent an "URGENT MEDICAL DEVICE RECALL (Removal)" letter dated December 29, 2014 to its customers. The letter describes the product, problem and actions to be taken. The customers were instructed to immediately check their inventory to confirm they do not have any units from the demo lot in possession; identify and set aside any unit from the identified lot; review, complete, sign and return the enclosed Acknowledgement Form directly to Cordis at the fax number on the form; return any affected product to Cordis; share this letter with others in your facility; contact any other facility that has been provided with units of POWERFLEX PRO product; maintain awareness of this notice until all affected product has been returned to Cordis and keep a copy of this notice witn any affected product.For any physician with medical concerns, please contact the Cordis Endovascular Medical Information Line 877-DEVICES Monday through Friday from 9am-5pm. For questions related to the recall that are not adequately addressed in the letter, please contact the Cordis Field Action Coordinator Hal Baden at (786) 313-2365. For questions related to the Acknowledgement Form and its return, please contact Cordis QA at (786) 313-8730. |  | 120 | 120 | Nationwide Distribution: US including states of: CT, NJ, NY, PA, MD, VA, WV, NC, SC, GA, FL, AL, TN, MS, KY, OH, IN, MI, IA, WI, SD, IL, MO, KS, NE, LA, OK, TX, CO, UT, AZ, CA, OR, WA, AR, IN, MN, WY, ND, NV, MA, RI, NH, CT, DE, District of Columbia and Hawaii. | Terminated | July 09, 2015 | 199 | 120 | Not_Computer
Z-1003-2015 | 70136 | apparatus, traction, non-powered | HST | Nonpowered orthopedic traction apparatus and accessories. | Orthopedic | Neurology | 510(K) Exempt | Fitness Gear Inversion table (Style STE00059FG or Style STE00118FG), Distributed by ASLI, Wilmington, DE (on carton), Paradigm Health &amp; Wellness, Inc. (in user manuals); The products are inversion tables, which allow the user to secure his or her ankles to the table and invert. The products were intended to be fitness equipment; i.e. consumer products. |  |  | December 29, 2014 | 2014 | Dick's Sporting Goods | Reports of customers falling from the Inversion Table and sustaining injuries. The Inversion Table is a fitness product, not a medical device. DSG removed medical claims from product labels and advertising. | Labeling False and Misleading | DSG sent emails and medical device correction notices, dated December 2014, to consumers who made purchases online and retail store consumers for whom DSG has contact information via its Scorecard customer loyalty program. DSG also posted the recall on their web site and notified their retail stores with instructions to post notices in DICKs stores. DSG plans to remind stores approximately monthly for three months to display the recall poster, beginning one month after the recall announcement. Consumers were asked to return the product to the stores where they were purchased for a full refund. A Press Release, dated January 8, 2015, was also issued about the recall. |  | 32581 | 32581 | Nationwide. | Terminated | September 02, 2015 | 247 | 32581 | Computer
Z-1001-2015 | 70012 | No | LCO | Platelet factor 4 radioimmunoassay. | Hematology | Hematology | 510(k) | PF4 Concentrated Wash (10x), REF 403624, packaged in the following assay kits: 1. ImmucorGamma LIFECODES PF4 IgG (catalog numbers HAT13G, HAT45G) and |  |  | December 15, 2014 | 2014 | Genetic Testing Institute,inc | Some Lots of PF4 Concentrated Wash (10X) are a slightly more dilute concentration. This may lead to slightly increased Optical Density (OD) values. Samples with a higher background reactivity may demonstrate false positive result (borderline reactivity) at or near the cutoff for the assay. | Employee error | Immucor sent an Urgent - Field Safety Notice letter dated December 15, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised to stop using the recalled product and to destroy it. Requested consignees to complete the Customer Response From and return to waukeshatechsupport@immucor.com. Please contact Technical Support should you need additional information at 262.754.1000 or waukeshatechsupport@immucor.com. |  | 6206 | 6206 kits | Worldwide Distribution - Nationwide Distribution to the states of : AL, AZ, AR, CA, CO, CT, DE, DC, FL, GA, IL, IN, IA, KS, KY, LA, ME, MD, MA, MI, MN, MS, MO, NE, NV, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN,TX, UT, VA, WA, WV and WI., and the countries of : Canada, Japan, Belgium, European Union , Australia, Taiwan, Thailand and Scotland. ( some countries not yet identified ) | Terminated | April 30, 2015 | 136 | 6206 | Not_Computer
Z-1002-2015 | 70106 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | Anti-Lambda APC-H7, Catalog Number 656157. |  |  | December 19, 2014 | 2014 | BD Biosciences, Systems &amp; Reagents | Two lots of the Anti-Lambda APC-H7 antibody are contaminated with CD38 antibody. | Vendor change control | BD sent a Recall notification letters dated December, 2014 to affected customers via Certified mail. The letter identified the affected product, problem and actions to be taken. Customers were instructed to inspect inventory for affected lots, immediately remove affected reagents from their laboratory, request replacement product and complete and return the enclosed Tracking/Verification form, per the instructions. For questions contact BD Customer Support at 855-236-2772. |  | 15 | 15 units | US Nationwide Distribution in the states of CA, NY, and MI | Terminated | March 11, 2016 | 448 | 15 | Not_Computer
Z-1004-2015 | 69727 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Software Version 3.1 utilized on the following systems: 1) 4600 Chemistry Systems (VITROS 5,1 FS System family member, 2) 5600 Integrated System |  |  | September 29, 2014 | 2014 | Ortho-Clinical Diagnostics | Calibration may not occur when using calibrator barcode labels supplied with VITROS Chemisty Products Calibrator Kit 2. | Software Manufacturing/Software Deployment | Ortho Clinical Diagnostics sent an Important Product Correction Notification letter dated September 29, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed Calibration may not occur when using calibrator barcode labels supplied with VITROS Chemisty Products Calibrator Kit 2.The resolution to this software anomaly will be in a future version of software. Until the new software is installed on their VITROSÂż System, customers were instructed not to use the calibrator barcode labels for VITROSÂż Calibrator Kit 2. Customers were advised that you manually program the calibration.Customers with questions were instructed to contact Customer Technical Services at 1-800-421-3311. |  | 12 | 12 units | Worldwide Distribution - USA including CO, NC, NJ, NY, PA, MI, WV and Internationally to Canada, Japan and Spain. | Terminated | April 03, 2017 | 917 | 12 | Computer
Z-1000-2015 | 70107 | No | GKL | Automated cell counter. | Hematology | Hematology | 510(k) | CD8 (Leu-2a) PE (In vitro diagnostic), Catalog number 340046. |  |  | December 19, 2014 | 2014 | BD Biosciences, Systems &amp; Reagents | One lot of CD8 PE is contaminated with CD4 FITC and may cause an unexpected staining pattern and incorrect patient data. | Process control | Recall notification letters were sent to all consignees on December 19, 2014 by certified mail. |  | 165 | 165 units | Worldwide distribution. US Nationwide including Puerto Rico, China, Brazil, Japan, Guatemala, India, New Zealand, Australia, Canada, Thailand, Chile and Malaysia. | Terminated | February 23, 2016 | 431 | 165 | Not_Computer
Z-1006-2015 | 70007 | No | BYI | Powered percussor. | Anesthesiology | Anesthesiology | 510(k) | FLUTTER Percussor respiratory device, packaged one unit per clear plastic package. |  |  | November 13, 2014 | 2014 | Aptalis Pharmatech Inc. | Customer notification that the device may be difficult to open or close. | Process control | Aptalis Pharmatech distributed a customer notification letter on November 13, 2014. The letter identified the affected product, problem and actions to be taken. Customers were instructed to forward complaints to Aptalis Pharmatech Customer Service department at 908-927-9600. |  | 49000 | Approximately 49,000 units | Worldwide Distribution - US Nationwide in the states of OH, NM, CO, NC, IL, MD, LA, MN, NJ, SC, AL, VA, MI, MA, and CA. in the countries of: Canada, Australia, Argentina, Guatemala, Serbia, Chile, and Finland. | Terminated | July 11, 2015 | 240 | 49000 | Not_Computer
Z-1007-2015 | 69938 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | The Arrow Edge Hemodialysis Catheterization Set Long-Term Access and NextStep Antegrade Hemodialysis Catheterization Set, Long-Term Access, part numbers: AC-15192-SFX, AC-15232-SFX, AC-15272-SFX, AC-15312-SFX, AC-15422-SFX, CS-15192-SFX, CS-15192-SFXM, CS-15232-SFX, CS-15242-I, CS-15272-SFX, CS-15282-I, CS-15312-SFX, CS-15322-I, CS-15362-I, CS-15422-SFX, CS-15502-SFX, CS-15552-I, CSD-15242-I, CSD-15282-I, and CSD-15322-I. |  |  | February 04, 2013 | 2013 | Arrow International Inc | The tunneling sheath packaged within the kits can crack during placement. If the tunneling sheath cracks, it can causea delay in treatment for the time it takes to retrieve a replacement kit. | Component design/selection | An Urgent Medical Device Recall Notification letter, dated February 4, 2014, was sent to direct consignees that identified the product, problem, and action to be taken. Customers were asked to discontinue use and quarantine affected inventory. The Recall Acknowledgement form should be completed and faxed back to 1-800-343-2935. A customer service representative will contact customers with a Return Goods Authorization (RGA) Number and provide instructions for the return of the product. Customers are asked to contact their local sales representative or Customer Service at 1-800-343-2935 with questions. |  | 2480 | 2,480 | Worldwide Distribution -- US and Canada. | Open, Classified |  |  | 2480 | Not_Computer
Z-1005-2015 | 70017 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | Multi-Lumen PI CVC Kit and PSI Kit: 3-L 7 FR X 8 product code: ASK-45703-POSU lot #23F13L0166, exp. 31-Jul-2015: |  |  | December 16, 2014 | 2014 | Arrow International Inc | the Instructions for Use (IFU) were inadvertently removed from the kit. | Labeling False and Misleading | An Urgent Medical Device Corrective Action Notification letter, dated December 16, 2014, was sent to Owens &amp; Minor, 1605 Worldwide Blvd., Hebron, KY 41048-8653, instructing them to immediately discontinue distribution and quarantine any products, use customer letter and Corrective Action Acknowledgement Form templates to communicate this corrective action to any of thier customers that have received product included within the scope of the corrective action, and to have their customers return a completed Corrective Action Acknowledgement Form outlining their remaining inventory of the affected product codes and lot numbers for which replacement IFUs is needed. |  | 420 | 420 | In the state of KY | Terminated | August 10, 2015 | 237 | 420 | Not_Computer
Z-1009-2015 | 69533 | No | MQN | Bone plate. | Dental | Dental | 510(k) | The Synthes Mandible External Fixator; all lots of part nos.: |  |  | October 17, 2014 | 2014 | Synthes, Inc. | Under new testing protocols for MR Environment safety and compatibility, metal devices are no longer designated MR Safe. Some components in the CMF Mandible External Fixator Systems were labeled and/or etched with information indicating they are MR Safe. They can no longer be labeled MR Safe. | Labeling design | Urgent notices of medical device labeling updates, dated October 17, 2014 and October 24, 2014 were sent to end users and sales consultants, which described the product, problem, and action to be taken. These consignees were instructed to inspect inventory and remove out-dated labeling inserts and replace with updated labeling information. |  | 43159 | 43159 | Worldwide Distribution. | Open, Classified |  |  | 43159 | Not_Computer
Z-1011-2015 | 69792 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | BrightView model number: 882478 |  |  | November 12, 2014 | 2014 | Philips Medical Systems (Cleveland) Inc | Unintended detector and gantry movement due to software issues. | Software design | Philips Healthcare sent an Urgent - Field Safety Notice dated November 4, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Philips is advising the customers/users to retain a copy of this Field SafetyBY CUSTOMER I USER Notice with the equipment Instruction for Use, until further notice.Problem #1:Philips is advising the customer/user to perform the following workaround prior tothe implementation of the field safety correction, to avoid contact between thedetector and the head holder as describe above in Problem #1:ONLY if the system is positioned in a Relative-90 configuration with one of thedetectors below the pallet level and the table in in the patient load position:Follow the sequence of operations prior to performing a study utilizing the headholder:- Deploy or store the x-ray panel if necessary;- Run the patient load PPM;- Install the head holder; and- Position the patient on the table and proceed with the studyIn addition, take the following preventative actions:" Monitor the patient during any and all system motions." Know locations of the multiple E-Stop button(s) to halt all system motion.(Emergency Stop: where the system immediately stops all detector andgantry motions by disconnecting power to the motors that control systemmotion) Customers with questions were instructed to contact the Customer Care Solutions Certer at 1-800-722-9377.For questions regarding this recall call 440-483-2119. |  | 1064 | 1064 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada, Algeria, Argentina, Australia, Austria, Bahrain, Belgium, Brazil, China, Czech Republic, Denmark, Dominican Republic, Egypt, Finland, France, Gabon, Germany, Greece, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Lebanon, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Palestine, Panama, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Syrian Arab Republic, Taiwan, Thailand, Turkey, Ukraine, United Arab Emirates, United Kingdom, United Kingdom, and Viet Nam. | Open, Classified |  |  | 1064 | Computer
Z-1010-2015 | 70102 | No | OCX | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Product Code 2C4013, Four Lead TUR Irrigation Set, for use with UROMATIC |  |  | January 14, 2015 | 2015 | Baxter Healthcare Corp. | Potential for failure of the pouch packaging seal at highaltitudes. | Package design/selection | 1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.Baxter Healthcare Center for Service can be reached at 888-229-0001 betweenthe hours of 7:00 AM and 6:00 PM Central Time, Monday through Friday. Pleasehave your Baxter 8-digit ship-to account number, product code, lot number(s) and quantity of product to be returned ready when calling.3. Complete the enclosed Customer Reply Form and return it to Baxter by faxing it to 224-270-5457 or scanning and e-mailing to fca@baxter.com . Returning the Customer Reply Form promptly will prevent you from receiving repeat notices.4. If you distribute this product to other facilities or departments within yourinstitution, please forward a copy of this communication to them.5. If you are a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, please conduct a consumer-level recall of the affected product that you distributed to customers.If product was purchased from a distributor: 1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit.The Baxter Healthcare Center for Service can be reached at 888-229-0001between the hours of 7:00 AM and 6:00 PM Central Time, Monday throughFriday.3. Follow your supplier's recall and response process. Please do not return thecustomer reply form to Baxter. |  | 26028 | 26,028 units | US Distribution. | Terminated | April 10, 2017 | 817 | 26028 | Not_Computer
Z-1013-2015 | 70200 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Syngo RT Therapist: The intended use of the SIEMENS branded ARTISTE, ONCOR and PRIMUS family of linear accelerator systems is to deliver X-Ray photon and electron radiation for the therapeutic treatment of cancer |  |  | January 02, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Combination of CTVision with syngo RT Therapist / syngo RT Oncologist 4.3.SP1 automatic registration in Adaptive Targeting might result in wrong offset calculations. Applying this offset can lead to patient mistreatment. Cone Beam imaging is not affected by this problem. | Software Design Change | A customer advisory notice, dated January 2, 2015, was sent to end users which identified the product, problem, and action to be taken. It was recommended that customers use only manual registration with the 6 degrees of freedom option disabled, which provides the correct offset values. A software update will be deployed once a solution is available. |  | 3 | 3 | UT, WI, NY | Terminated | September 10, 2015 | 251 | 3 | Computer
Z-1012-2015 | 70086 | No | JXG | Central nervous system fluid shunt and components. | Neurology | Neurology | 510(k) | Miethke Shunt System accessories |  |  | December 17, 2014 | 2014 | Aesculap, Inc. | Some Miethke Shunt System accessories may have an incorrect MR safety symbol on the shelf box label. The label should have NO symbol but instead is marked with the MR conditional symbol. The product is MR unsafe. This information is only required to be in the Instructions for Use (IFU). | Labeling False and Misleading | Aesculap sent an Important Correction and Removal and Field Safety Notification letter, dated December 17, 2014, to all affected customers. The letter identiifed the affected product, problem, and actions to be taken. Customers were requested to remove from inventory and contact Aesculap regarding return of the affected devices and acknowledgement form. |  | 772 | 772 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 772 | Not_Computer
Z-0955-2015 | 70133 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | The Revolution CT is a multi-slice (256 detector row) CT scanner consisting of a gantry, patient table, scanner desktop (operator console), system cabinet, power distribution unit (POU). and interconnecting cables.The system includes image acquisition hardware, image acquisition and reconstruction software, and associated accessories. |  |  | November 12, 2014 | 2014 | GE Healthcare | A required quality control test was not performed during installation associated with the software of the Revolution CT scanner. | Radiation Control for Health and Safety Act | GE Healthcare Planned Action (s) to repair defect or to bring product into Compliance letter dated January 16, 2015 to GE Healthcare.1. A GE Healthcare service representative will perform a noise specification and tolerance calibration to correct all affected systems via Field Action Instruction (FMI 25457).2. A customer notification letter which includes a statement that GE Healthcare will without charge, remedy the defect or bring the product into compliance.CDRH approves the CAP subject to the following conditions:1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. GE Healthcare Surgery will implement this CAP by May 1, 2015.You may proceed or continue with implementation of the CAP (subject to the conditions noted above).Call Center phone numbers: United States: 800 437-1171, Japan: 0120-055-919Information For other countries, please contact your local GE Healthcare Service Representative if you have any questions or concerns regarding this notification. |  | 13 | 13 units | US Distribution to the states of CA, FL, IL, NY, UT and WA. | Terminated | April 29, 2015 | 168 | 13 | Computer
Z-1019-2015 | 70141 | No | IOR | Mechanical wheelchair. | Physical Medicine | Physical Medicine | 510(k) | Zippie Zone Pediatric Rigid Wheelchair Model EIRZ1. |  |  | December 19, 2014 | 2014 | Sunrise Medical (US) LLC | There is the potential for the height adjustable handle to dislodge due to failing locking mechanisms on specific identified chairs. | Device Design | Sunrise Medical sent an Urgent Medical Device Field Correction letters dated December 19, 2014 to the two affected distributors. The letter identified the affected product, problem and actions to be taken. For question contact Sunrise Medical Regulatory Affairs at either (559)294-2840. |  | 3 | 3 | Nationwide Distribution in the states of CO and ME. | Terminated | March 10, 2015 | 81 | 9 | Not_Computer
Z-1014-2015 | 70137 | No | OQY | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Ebb Complete Tamponade System, BD-OTS Obstetric Catheter, REF CTS-1000, Rx Only, Sterile EO |  |  | December 30, 2014 | 2014 | Clinical Innovations, LLC | Potential for a balloon leak in all lots of (CTS-1000) manufactured since August 2014. | Process design | Clinical Innovations sent an "Urgent - Medical Device Recall" letter dated December 31, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They will be instructed to quarantine remaining product. A company representative will visit the hospital and retrieve the product in accordance to instructions detailed on the Representative Recall Reconciliation Instructions. The representative will reconcile the inventory on the Field Reconciliation Form and return the product to Clinical Innovations as indicated on the instructions. The address for the devices to be returned is included on the Representative Recall Reconciliation InstructClinical Innovations is working diligently to resolve this issue. In the meantime ,please quarantine any remaining product in your hospital. A company representative will contact you and make arrangements for the return of the product. Clinical Innovations will provide replacement product as soon as it becomes available. If you have any questions,call Clinical Innovations,at 1-(888)268-6222 or your service representative, as is the name of the Company Contact. |  | 827 | 827 units | Worldwide Distribution - USA (nationwide) and the countries of : Canada, Finland, Estonia, Belgium, Ireland, Sweden, France, Switzerland, Spain, Netherlands, Middle East, Scotland, Germany and Wales. | Terminated | December 28, 2015 | 363 | 827 | Not_Computer
Z-1015-2015 | 70027 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | ASSY, POWER SUPPLY 130V 50HZ. This component is used with the following machines: 2008T Hemodialysis Machine, 2008K Hemodialysis Delivery Equipment, 2008 K2 Hemodialysis Delivery Equipment |  |  | December 16, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Unqualified power assemblies shipped to customers. | Under Investigation by firm | Consignees were notified of the recall via letter sent on 12/18/2014 via certified mail. The letter informs consignees of the recall and requests that they take the following steps: Please check your stock for any 50 Hz Power Supplies (P/N 190092) with the name Electro Assemblies, LLC which is visible on the Transformer contained inside the Power Supply and discontinue use. Place the product in a secure, segregated area. Please contact your Fresenius Technical Services at 1-800-227-2572 option 4 then option 1, and reference the Field Action Number above. They will issue you a Return Good Authorization so you can return this product. that a reply form is completed and returned via fax or scanned email copy. In addition, consignees are asked to complete a reply form and return via fax/mail. |  | 10 | 10 | Distributed in OK and FL. | Open, Classified |  |  | 10 | Not_Computer
Z-1016-2015 | 69993 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Airo Mobile Intraoperative CT I Airo; Model #: MobiCT-32. Computed Tomography x-ray system. |  |  | December 08, 2014 | 2014 | Mobius Imaging, LLC | There is a risk that during the transfer of an image and navigation data to the Brainlab Curve Image Guided Surgery Navigation System after a CT scan, an error may occur, causing either no navigation data or incorrect navigation data being transferred to the curve. | Software design | Mobius sent an Urgent Field Safety Notice dated December 8, 2014, to all affected customers. The letter explains the recall and provides advise on action to be taken by the user. The letter also informs consignees that Mobius Imaging is working on a software fix for the software bug/communication error. The final fix is anticipated in early 2015.Advise on action to be taken by the user:The user should be prepared for the potential necessity to perform a manual navigation registration in order to use navigation.In general, do not use navigation data transmitted by the AIRO Mobile CT Scanner of any scan the Brainlab AIR App displayed the "Received Wrong Dataset" error for. To use such scans with the Curve, a manualnavigation registration is required.To avoid the possibility of incorrect data from being transferred to the Curve, Mobius Imaging is recommending that customers only select "NO" when they come to the "Continue with same patient" screen after each scan. They can select the same patient under the "Existing Patients" tab and restart the workflow. Additionally, always adequately verify the accuracy of the navigation registration data as required by the Brainlab Curve" Navigation system software and the AERO Operator's Manual. If you require further clarification, please feel free to contact your local Brain lab Customer SupportRepresentative.Customer Hotline: +49 89 99 15 68 44 or +1800 597 5911 (for US customers) or byE-mail: support@brainlab.com (for US customers: us.support@brainlab.com)Fax Brainlab AG: + 49 89 99 15 68 33Address: Brainlab AG (headquarters), Kapellenstrasse 12, 85622 Feldkirchen, Germany.For questions regarding this recall call 978-615-5025. |  | 315 | 21 devices distributed. 15 of these devices affected by the software v. 1.1.1+ patch | USA including AK, IL, WA, MA, NY, WV, NC, TX, CA, PA and Internationally to Germany, Netherlands, Belgium, Switzerland, Saudi Arabia, Egypt. | Terminated | July 06, 2015 | 210 | 315 | Computer
Z-1026-2015 | 70030 | No | KHP | Lactic acid test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Beckman Coulter, Lactate Assay, Catalog No. OSR6X93 |  |  | December 15, 2014 | 2014 | Beckman Coulter Inc. | N-Acetyl Cysteine (NAC), when administered in therapeutic concentrations for the treatment of acetaminophen overdose, interferes with the assays for these analytes: Cholesterol OSR6X16, Uric Acid OSR6X98, Lactate OSR6X93, and Lipase OSR6X30. NAC interference may lead to falsely low results. | Under Investigation by firm | A customer notification letter was sent to their customers on 12/15/14 to inform them that Beckman Coulter has recently confirmed that N-Acetyl Cysteine (NAC), when administered in therapeutic concentrations (for the treatment of acetaminophen overdose), has been determined to interfere with assays for the following analytes: Cholesterol OSR6X16, Uric Acid OSR6X98, Lactate OSR6X93, and Lipase OSR6X30. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to call (800) 223-0130 in the US and Canada. Customers outside the US and Canada are instructed to contact their local Beckman Coulter Representative. |  | 31799 | 31,799 units total (15,841 units in US) | Nationwide in US and Puerto Rico.World Wide: Albania, Algeria, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bolivia, Bosnia-Herzegovina, Botswana, Brazil, Bulgaria, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyrus, Czech Republic, Denmark, Ecuador, Egypt, El Salvador, Eritrea, Ethiopia, Finland, France, Germany, Ghana, Gibraltar, Greece, Guadeloupe, Guam, Honduras, Hong Kong, Hungary, India, Indonesia, Ireland, Israel, Italy, Jordan, Kazakhstan, Kenya, Korea, Kuwait, Lebanon, Lesotho, Liberia, Libyan Arab Jamahiriya, Lithuania, Malawi, Malaysia, Mexico, Morocco, Mozambique, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russian Federation, Saudi Arabia, Serbia, Sierra Leone, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, Sweden, Switzerland, Taiwan, Thailand, Tunisia, Turkey, Ukraine, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Viet Nam, Zambia, and Zimbabwe | Terminated | February 19, 2016 | 431 | 31799 | Not_Computer
Z-1027-2015 | 70318 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza: Syngo.plaza is a Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images. It supports the physician in diagnosis and treatment planning. |  |  | January 13, 2015 | 2015 | Siemens Medical Solutions USA, Inc | In case of a system crash, images may not be written to the hard disk and this may result in inconsistencies in the database. In case of a system crash (e.g. blue screen, power outage) images may not be written from cache to the hard disk and might get lost. | Software Manufacturing/Software Deployment | Siemens sent a Safety Advisory Notice dated January 13, 2015, to all affected customers. The Safety Notice identified the product, the problem, and the action to be taken. Siemens recommended checking patient data in the event of a system crash and advised customers that a software update is planned for mid-2015 that will resolve the issue.For questions regarding this recall call 610-219-6300. |  | 50 | 50 | Nationwide Distribution | Open, Classified |  |  | 50 | Computer
Z-1022-2015 | 70068 | No | MWI | Cardiac monitor (including cardiotachometer and rate alarm). | Cardiovascular | Cardiovascular | 510(k) | VC150 Vital Signs Monitor (VC150 monitor equipped with Masimo SpO2 technique only); Innokas Medical CARESCAPE VC150 Rx Only Innokas Medical Oy. Catalog numbers 2067980-002, 2067980-006, 2067980-010, and 2067980-014. |  |  | December 03, 2014 | 2014 | INNOKAS MEDICAL OY | If the SpO2 or RRa value is violating the respective alarm limit at the time of switching to monitoring mode, the monitor does not issue visual and audible alarms. | Software design | A notification letter was sent to the distributors on 12/3/2014. The letter requests the devices be returned to the manufacturer for repair. |  | 20 | 20 units | Distributed in the US and Europe. | Terminated | April 24, 2015 | 142 | 20 | Not_Computer
Z-1021-2015 | 70019 | No | HXW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 2.0 Mandible Locking Plate (MLP) Bender / Cutter (part number 329.143) |  |  | December 15, 2014 | 2014 | Synthes, Inc. | Certain lots have the potential for tip breakage while bending 2.0 mm thick plates. If device is used and breaks while in the Operating Room, surgical delay may occur while alternate bending or cutting instruments are located. | Finished device change control | An Urgent Notice: Medical Device Recall, dated December 19, 2014, was sent to end users and sales consultants that described the product, problem, and action to be taken. End users/sales staff were requested to immediately stop use and remove affected product from inventory. Customers were asked to call and obtain a Return Authorization Number, and complete and return the Verification Section of the letter along with the devices to Synthes per the instructions provided. Questions should be directed to 610-719-5450 or your local DePuy Synthes sales consultant. |  | 973 | 973 | Worldwide Distribution -- US and Canada. | Open, Classified |  |  | 973 | Not_Computer
Z-1032-2015 | 70050 | No | NZW | Intravascular catheter. | General Hospital | General Hospital | 510(k) | HEPFLUSH-10 (HEPARIN LOCK FLUSH SOLUTION, USP) 100 USP units per 10 mL (10 USP units per mL), Single Dose Vial, Rx only -- Fresenius Kabi USA, LLC -- NDC 63323-017-10, Product Code 1710 |  |  | December 19, 2014 | 2014 | Fresenius Kabi USA, LLC | The lot failed pH specification. | Component design/selection | Fresenius sent an URGENT DRUG RECALL Letter dated December 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.1. Examine their stock immediately to determine if they have any product from the affected lot. If the consignee is a distributor, immediately notify your customers that have been shipped or may have been shipped this product/lot of this recall and direct them to discontinue distributing or dispensing the affected lot. Please have them prepare to return the product to Fresenius Kabi as instructed in the Urgent Product Recall Response Form. Your customers may retrieve the recall letter and response form at http://www.fresenius-kabi.us /products/pharmaceutical-products/product-updates.html.2. If you have the affected lot available, immediately discontinue distributing or dispensing the lot, and return all units to Fresenius Kabi ,via FedEx Ground, using the enclosed return goods label and packing slip. A FedEx Ground labelcan be obtained by checking the box and noting your mailing address on the enclosed Urgent Product Recall Response Form. It will be mailed to the consignee upon receiving the request. A credit memo will be issued covering the quantity of the return to Fresenius Kabi.3. PLEASE COMPLETE THE ENCLOSED "URGENT PRODUCT RECALL RESPONSE FORM" AND SEND IT BACK TO US IMMEDIATELY VIA EMAIL OR FAX AT 1-708-649-8630. Use the following contact phone numbers as appropriate. Hours of operation: Monday through Friday 8:00 am to 5:00 pm CST: QA Department: (866) 716-2459 for information on how to return product; Vigilance &amp; Medical Affairs: 1(800) 551-7176 for clinical/technical information/Adverse Events (ADE's) reporting. |  | 69725 | Domestic: 69,725 | Nationwide Distribution | Terminated | January 26, 2016 | 403 | 69725 | Not_Computer
Z-1028-2015 | 70034 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Portex Anaerobic Pulsator Arterial Blood Sampling Kit, with Liquid Sodium Heparin for Gases and Electrolytes. |  |  | December 17, 2014 | 2014 | Smiths Medical ASD, Inc. | The outer box label is marked with the correct expiration date of 08/2016. The inner Tyvek Unit Label is marked with 08-2017 which exceeds the actual expiration date by one year. | Employee error | Smiths Medical notified American Medical Depot via telephone on 12 December 2014. Smiths Medical instructed them to put product at their location on hold. American Medical Depots customer will be instructed to return their unused product for replacement. |  | 900 | 900 units | US in the state of FL. | Terminated | June 25, 2015 | 190 | 900 | Not_Computer
Z-1031-2015 | 70232 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | HomeChoice/HomeChoice Pro Automated Personal Cycler, for Automated Peritoneal Dialysis. Product codes 5C8310, 5C8310R, R5C8320, and 5C4474. |  |  | October 21, 2014 | 2014 | Baxter Healthcare Corp. | The keypad buttons on HomeChoice devices may be activated without the operator pressing them. | Nonconforming Material/Component | An Urgent Device Correction notification was sent to consignees on October 31, 2014. Within the notification, Baxter Healthcare asked customers to do the following: 1) Contact Baxter Technical Services to arrange for the exchange (swap) of any affected HomeChoice PRO Automated PD System devices. 2) Complete the attached customer reply form and return it to Baxter by either fax or scanned e-mail. For general questions regarding this communication, contact The Center for One Baxter at 900-422-9837, Monday through Friday between the hours of 8:00 AM and 5:00 PM Central Time. |  | 200 | 200 | AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, LA, MA, MD, MI, MO, NC, NE, NJ, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, VT, WA, WI and District of Columbia | Terminated | September 01, 2015 | 315 | 200 | Not_Computer
Z-1029-2015 | 69963 | No | ONN | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Verasense Knee System, Model SYK-TRCR 02. Intelligent disposable tibial insert that measures dynamic loads during orthopedic knee surgery. |  |  | December 05, 2014 | 2014 | Orthosensor, Inc. | Presence of one inoperable (dead) sensor pad. | Under Investigation by firm | OrthoSensor sent an Recall letter dated December 4, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Required Action for Users" Remove the above-referenced serial numbers from your VERASENSE inventory and quarantine." Complete the enclosed Customer Response Form.Actions to be taken by OrthoSensor" You will be credited for these devices for your full purchase amount. The Returned Materials Authorization (RMA) number for each device in your possession is listed above." OrthoSensor will replenish these devices as soon as they are available.For further questions please call (954) 577-7770. |  | 22 | 22 devices | US Distribution to the states of : IL, MI, NY, CA, VA and NV. | Terminated | March 13, 2015 | 98 | 22 | Not_Computer
Z-1030-2015 | 70313 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | Arrow Select Kits (ASK), Maximal Barrier Precautions Kits (CDC) and Arterial Catheterization Kit (NA) with Biopatch Dressing |  |  | January 09, 2015 | 2015 | Arrow International Inc | Certain product codes and lot numbers of Kits with Biopatch Dressing were packaged with the incorrect size Biopatch Dressing. No risk to patients is anticipated. | Packaging process control | Arrow International, Inc. (Arrow) sent notification letters, dated January 9, 2015, to alert their direct account customers that it may be necessary to obtain an alternate Biopatch Dressing if a kit is used that contains an incorrectly sized Biopatch Dressing. Due to the inconvenience, Arrow will reimburse customers $7.50 per kit received. |  | 4484 | 4484 | US (nationwide) and PR. | Open, Classified |  |  | 4484 | Not_Computer
Z-1034-2015 | 69299 | No | OYC |  |  | Clinical Chemistry | PMA | Animas Vibe Insulin Infusion Pump. |  |  | September 06, 2011 | 2011 | Animas Corporation | Calibration factors in the pump overwritten during a programming step. The force sensor could send a lower signal value to the pump processor, with loss of prime warnings, occlusion alarms and the pump unable to detect a cartridge during the prime sequence. Field action initiated 8/29/2011. | Software design (manufacturing process) | The field action was initiated on 8/29/2011 and comprised communication to patients and distributors using email, letters and verbal communication. In addition, notifications were sent to those Health Care Professionals who had patients that were affected by this field action. Health Authorities were notified in those countries where the pump had been distributed (France, Germany, Sweden and United Kingdom) or was in the hands of patients. Replacement product was provided to distributors and end users who were identified as having affected product. |  | 1235 | 1235 | No US distribution, Distributors are located in France, Germany, Sweden and United Kingdom. | Terminated | August 10, 2015 | 1434 | 1235 | Not_Computer
Z-1033-2015 | 70008 | No | MGX |  |  | Toxicology | PMA | DrugCheck Ketamine Dip Test |  |  | November 21, 2014 | 2014 | Express Diagnostics Int'l., Inc. | This recall has been initiated due to potential false positives results when using the test to screen for Ketamine. Use of this product may give incorrect screening results. | Nonconforming Material/Component | Consignees were sent an Express Diagnostic Int. Inc., "Urgent" Letter dated 11//21/2014. The letter described the problem and the product involved in the recall. It advised consignees to quarantine the product. If the product has been further distributed, consignees were advised to notify their customers at once, Consignees were requested to complete and return the "Acknowledgement Form" either via fax to 1-507-526-2252 or e-mail Quality@drugcheck.com. For questions consignees can contact Jackie Gale, Director of Quality Systems and Regulatory Affairs at 1-507-526-3951. |  | 5050 | 5,050 | Worldwide Distribution -- US, including the states of CA, NY, MN, and SC; SWEDEN, CANADA, UNITED KINGDOM, and GUATEMALA. | Open, Classified |  |  | 5050 | Not_Computer
Z-0999-2015 | 70058 | No | OWD |  |  | Molecular Genetics | PMA | cobas 4800 KRAS AMP/DET 24T CE-IVD Mutation Test; CE-IVD 5852170190. |  |  | December 18, 2014 | 2014 | Roche Molecular Systems, Inc. | cobas KRAS Mutation Kit T10786 is generating invalid results for the Mutant Control and/or KRAS Calibrator. | Device Design | RMS issued an Urgent Field Safety Notice on 12/18/2014.A Quality Notification was provided to all users of the of cobasÂż KRAS Mutation Kit CE-IVD, lots T13091, T15435, T15475, T14777, and W00389, informing them of the issue and instructing them to discontinue use and discard any remaining inventory immediately. |  | 660 | 660 pieces | Foreign Distributed. | Open, Classified |  |  | 660 | Not_Computer
Z-1038-2015 | 70021 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftLab with SA HIS versions 4.0.7.0-4.0.7.1 |  |  | May 20, 2014 | 2014 | SCC Soft Computer | The interface fails to send abnormal flags for Reference Lab test results. | Device Design | The firm sent out 2 CORRECTION COMMUNICATIONS informing consignees of functionality in release of 4.0.7 HIS Add-on interfaces which SCC has designated to be a potential patient Risk-to-Health issue. Further the corrective action indicated HIS releases 4.0.7.2 and higher will be corrected. Hot Fixes have been developed for HIS Add-on patches 4.0.7.0 and 4.0.7.1. Lastly the communication requests that the user "please acknowledge receipt of this task and grant permission to load the mandatory hot fix(es). Thank you." |  | 48 | 48 | Distributed US (nationwide) in the states of OH, GA, NJ, WI, FL, NH, IL, NY, IN, CA, AL, VA, MO, MI, KY, MA, PA, CO, ME, TN, WA, MN, SD, LA, NC, and NV and the country of Canada. | Open, Classified |  |  | 48 | Not_Computer
Z-1039-2015 | 70062 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Baxter Integrated APD Set with Cassette 3-Prong, Sterile, nonpyrogenic fluid path. |  |  | December 08, 2014 | 2014 | Baxter Healthcare Corp. | Baxter is issuing a recall for one lot of Integrated Automated Peritoneal Dialysis set with Cassette 3 prong due to complaints received for leakage. | Device Design | Baxter sent an Urgent Product Recall notification letter dated January 8, 2015 . The letter identified the affected product, problem and action to be taken. The recalling firm has asked the consignees to 1) locate and remove all affected product from their facilities, 2) contact Baxter to arrange for return and credit, and 3) complete the attached customer reply form and return it to Baxter. For questions call 800-422-9637. |  | 8640 | 8,640 units | US Nationwide Distribution in the states of FL, WI, AR, TN, TX, NC, MO, NM, NJ, CT, PA, NV, NY, CA, MI, WI, KS, and OH. | Terminated | October 28, 2015 | 324 | 8640 | Not_Computer
Z-1036-2015 | 69450 | No | FNL | AC-powered adjustable hospital bed. | General Hospital | General Hospital | 510(K) Exempt | G- Series Beds; G29 full length bed rails assembly |  |  | December 05, 2014 | 2014 | Invacare Corporation | Invacare identified that the rail mounting bracket on the G29 full length bed rail was out of specification for entrapment safety. | Nonconforming Material/Component | Invacare sent an Urgent Medical Device Recall letter dated December 5, 2014, to all affected customers. The letter identified the product and the problem. Customers with questions should call 440-329-7996.For questions regarding this recall call 440-329-6875. |  | 2499 | 2499 units | US Nationwide Distribution including OH, CO, MO, and NY. | Terminated | June 18, 2015 | 195 | 2499 | Not_Computer
Z-1041-2015 | 70023 | No | CGA | Glucose test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Cobas b 221 (versions 5 and 6) Blood Gas and Electrolytes Analyzer |  |  | December 18, 2014 | 2014 | Roche Diagnostics Operations, Inc. | Chemical component of the F3 Fluid pack and environmental bacterial contamination causes a positive bias up to 39% at 4.1 mmol/l for the glucose measurement and 30% at 4.12 mmol/l for lactate. Urea readings are unaffected. | Material/Component Contamination | Roche Diagnostics has confirmed an issue with the calibration and QC stability of MSS parameters for cobas b 221 &lt;6&gt; systems. Internal investigations identified the formation of white deposits in the fluidic system, especially in the SD cartridge tubing, as the cause for the premature failure of the MSS parameters. These white deposits can potentially affect the fluidic pathways of both the Standby solution and the calibration solutions Cal1 to Cal4 (from the S3 Fluid Pack) and may reduce the lifetime of metabolite sensor cartridges.The cause of the white deposit has been identified as a chemical component of the S3 Fluid Pack solutions, as well as environmental bacterial contamination. Since this issue involves the S3 Fluid Pack, only the cobas b 221 &lt;6&gt; system is impacted.This issue may have an effect on the calibration of Glucose and Lactate parameters, which may lead to patient results being affected, as calibration drifts may not be detected by calibration and QC procedures. Glucose and Lactate are the only parameters affected by this issue.As a result, Roche is implementing an interim corrective action to be performed on your systems on a regular basis by Roche Field Engineering Specialist (FESs). Read the Urgent Medical Device Correction (UMDC) in its entirety." Continue to run and maintain your cobas b 221 &lt;6&gt; system as per the Instructions for Use." Complete the attached fax form and fax it to 1-888-656-6385." File this UMDC for future reference. Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-526-2272 if you have questions about the information contained in this UMDC. Public Contact is Todd Siesky, at todd.siesky@roche.com. |  | 103 | 103 | US Nationwide distribution. | Open, Classified |  |  | 103 | Not_Computer
Z-1040-2015 | 70237 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | MobileDiagnost wDR motorized portable diagnostic X-ray systems. |  |  | August 06, 2014 | 2014 | Sedecal S.A. | Reports of unexpected movement. | Other | URGENT - Field Safety Notice letters dated August 6, 2014 were sent to all consignees of record. The letters instructed consignees: 1) The unit is equipped with a safety switch in the drive handle. If the drive handle is released the unit will stop; 2) Do not try to correct unintended movement using the controls on the drive handle; 3) Instead, stop the system immediately by releasing the drive handle entirely; 4) Alternatively, the emergency button can be pressed; 5) Take special care when the system crosses metal strips or plates on the floor or enters an elevator, as these are situations where electrostatic discharge is likely to occur; and, 6) Should the customer feel uncertain regarding this action, he/she can contact Philips Healthcare. |  | 283 | 283 systems | US: Nationwide (AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NE, NM, NY, OH, OR, PA, SC, TN, TX, UT, VA, VT, WA, WI) | Open, Classified |  |  | 283 | Not_Computer
Z-1037-2015 | 70184 | No | HBC | Electric cranial drill motor. | Neurology | Neurology | 510(k) | 16cm Minimally Invasive Attachment. For use with Motor Systems for cutting and shaping bones, including the spine and cranium. Model MIA16-G1 |  |  | December 10, 2014 | 2014 | The Anspach Effort, Inc. | Customer complaint indicating the color bands on the MIA16-G1 were not correct. | Environmental control | Anspach sent an "URGENT MEDICAL DEVICE RECALL" (Removal) letter dated December 10, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm requested the customer to 1) Immediately identify and set aside all product listed below 2) Review, complete, sign and return the attached reply form 3) Share this letter with others in your facility that need to be made aware of this recall 4) If you distribute any of the affected products to other services or facilities, please forward this letter as appropriate 5) Maintain awareness of this notice 7) Keep a copy of this notice. If you have any questions regarding this removal, please contact the Complaint Handling Unit Manager at 561-494-3673 or contact your Anspach Sales Representative. |  | 8 | 8 | US Distribution to the states of: NC and CA., and Internationally to Switzerland. | Open, Classified |  |  | 8 | Not_Computer
Z-1047-2015 | 70288 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Fresenius Crit Line in a Clip (CLiC) with SW version 2.51 Model Number: CL10041001. A continuous real-time monitor for non-invasive hematocrit, oxygen saturation and percent change in blood volume calculation during hemodialysis treatment. |  |  | December 19, 2014 | 2014 | Fresenius Medical Care Holdings, Inc. | Potential for misinterpretation of the graphic display of the Blood Volume ( BV) slope | Device Design | FMCRTG, LLC representatives contacted Clinics on 12/19/14 by visit/telephone phone) and removed all CLiC units. |  | 35 | 35 units | CT, NY | Completed |  |  | 35 | Not_Computer
Z-1042-2015 | 70105 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer PERSONA Tibial Articular Surface Inserter; Item # 42-5299-001-00. Manual surgical instrument for knee arthroplasty (orthopedic). |  |  | November 25, 2014 | 2014 | Zimmer, Inc. | Potential for fracture of the tip of the product. The fractured tip is approximately 4mm x 7mm x 3.5mm in size. Retained tip may result in: pain, soft tissue irritation, soft tissue damage, or increased wear due to contact with the fractured tip; Revision surgery to remove the fractured tip | Process control | On 11/25/2014, URGENT MEDICAL DEVICE RECALL LOT SPECIFIC notifications were sent to the affected distributors, hospital Risk Managers and Surgeons with instructions for returning the affected product. All distributors were notified via electronic mail, and all hospital Risk Managers, Surgeons, and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call the customer call center at 1-800-348-2759. |  | 1091 | 1,091 | Worldwide distribution: US nationwide, including Puerto Rico, South Korea, Singapore, Dominican Republic, Canada, Japan, and Germany. | Terminated | December 01, 2015 | 371 | 1091 | Not_Computer
Z-1018-2015 | 70048 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Covidien Trellis-8 Peripheral Infusion System. Models CVT808015, CVT808025, CVT812015, and CVT812025. The following models are not offered for sale in the US: EVT808015, EVT808025, EVT812015, and EVT812025 Sterile EO. |  |  | December 10, 2014 | 2014 | Covidien | A manufacturing error resulted in the risk of incorrect proximal and distal balloon inflation port identification on the units. Units have been identified to have the distal balloon inflation port incorrectly labeled as proximal, and, the proximal balloon port incorrectly labeled as distal. | Mixed-up of materials/components | Consignees were sent an Urgent Product Recall letter dated 12/15/2014. The letter described the issue, identified the affected product, and the required actions. Affected product was to be returned directly to customers' Covidien Sales Rep; or customers were to contact Covidien Service at 1-800-716-6700 to arrange for product return. Questions can be directed to the Covidien Sales Rep or to Covidien Service at 1-800-716-6700. Customers were to complete and return the Verification Form along with the unused product as soon as possible. |  | 1032 | 1032 devices (919 US, 113 OUS) | Worldwide Distribution - US Nationwide and the countries Australia, Canada, Finland, France, Germany, Ireland, Italy, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | March 31, 2016 | 477 | 1248 | Not_Computer
Z-1043-2015 | 70100 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Trident PSL HA Solid Back 50 mm |  |  | December 15, 2014 | 2014 | Stryker Howmedica Osteonics Corp. | Five customer complaints were received for the same batch of Trident PSL Shells (part number 540-11-50E lot 49073501) alleging that when the packages were opened there were no dome holes plugs present. | Process control | Stryker sent an Urgent Medical Device letter dated January 7, 2015, with a Recall Notification/Acknowledgement Form via Fed Ex to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to complete and return the attached Product Recall Acknowledgement Form to so_m_product_field_action_response@stryker.com or fax to 1-855-251-3635.Customers should return the affected product to:Attention: Regulatory ComplianceStryker Orthopaedics325 Corporate DriveMahwah, New Jersey 07430Customers with questions should call 201-831-6365.For questions regarding this recall call 201-831-5000. |  | 7 | seven units | Nationwide Distribution including NY, OH, PA, and TX. | Open, Classified |  |  | 7 | Not_Computer
Z-1048-2015 | 70181 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Si Surgical System, IS3000; - Wall chart for EndoWrist Stapler 45 System. |  |  | January 15, 2015 | 2015 | Intuitive Surgical, Inc. | The Wall chart has been updated because it was noted that Wall Chart (PN 551524-03 Rev A) was inconsistent with the User Manual Adendum (PN 550986-07_B) with regards to neceassary sterilization dry time. | Error in labeling | Medical Device Recall Replacement letters were sent to all customers on January 15, 2015. The letter identified the reason for the recall, the risk to health, affected products, and actions to be taken by the customer or user and the firm. Customers were asked to discard the existing EndoWrist Stapler Documentation and replace it with the new Documentation provided by the firm. Intuitive Surgical will ship replacement instruments at no charge to customers. Then customers were asked to reunite the replacement instruments with the additional components. The firm will retrain customers on the updated instructions for use. A Clinical Sales Representative or Intuitive Surgical Customer Service at 1-800-876-1310, should be contacted if further information or support is needed concerning the recall notification. |  | 1034 | 1,034 | Nationwide Distribution | Terminated | March 03, 2015 | 47 | 1034 | Not_Computer
Z-1050-2015 | 68672 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Plum A+ Infusion System with Hospira MedNet List 20792-04-25, Pole Clamp Assembly Part Number 820-95443-008. A cassette based multi-function infusion system. Each system includes a pumping module and an assortment of disposable IV sets, optional accessories, and an operator's manual. The pole clamp assembly is used to provide a mechanism to tighten and loosen the grip of the pole clamp to remove the infuser from the IV pole. |  |  | June 11, 2014 | 2014 | Hospira Inc. | Plum A+ infusion system pole clamp assemblies were discovered to be cracked. If the knob on the pole clamp assembly is cracked, the infuser cannot be attached to the IV pole since the knob is used to tighten and loosen the grip of the pole clamp to the pole. | Vendor change control | Hospira contacted the the sole customer on June 11, 2014, to inform them of the issue of the potential impact to their Plum A+ infusers list 20792-04-25. Hospira requested that the customer to allow Hospira Field Service Representatives to inspect and replace the impact pole clamp assemblies.On June 12, 2013 Geneva Hospital approved Hospira's request to the inspection and replacement of the impacted pole clamp assemblies. On June 16, 2013 Hospira's Field Service personnel performed the on-site, visual inspection of the pole clamp assemblies and confirmed 15 units were cracked. Hospira determined to replace all 102 pole clamp assemblies that were received by Geneva Hospital. For further questions please call (800) 441- 4100. |  | 102 | 102 infusors | US Distribution to New York. | Completed |  |  | 102 | Not_Computer
Z-1054-2015 | 70185 | No | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | Lubricated Intubating Stylet 6FR that is inserted into an endotracheal tube to facilitate intubation. This product is used in surgical or medical emergency facilities. Primarily used on infants and pediatric patients. |  |  | December 23, 2014 | 2014 | Cardinal Health | A small piece of the stylet sheath has the potential toshear off upon removal of the stylet from the tracheal tube. | Process control | URGENT: PRODUCT RECALL letters dated December 23, 2014 were sent to all customers of record. The letters included instructions for customers to: 1) quarantine the affected product; 2) notify any customers whom you may have distributed the affected product to; 3) contact the appropriate Customer Service group to arrange for return and credit of any affected product in your inventory; and, 4) complete and return the enclosed Acknowledgement Form. Customers with questions about this recall can contact Cardinal Health Quality Systems at 800-292-9332. |  | 104432 | 104,432 stylets | US: Nationwide (AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NM, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, WA, WI, WV, WY) + District of Columbia + Guam + US Army Facility in South Korea | Open, Classified |  |  | 104432 | Not_Computer
Z-1049-2015 | 70365 | No | NFJ | Picture archiving and communications system. | Radiology | Ophthalmic | 510(k) | FORUM Archive and Viewer, version 3.1, v 3.1.1, (DVD Format) and v 3.2, v 3.2.1.(DVD and USB Flash Drive Format). |  |  | January 22, 2015 | 2015 | Carl Zeiss Meditec AG | Software defect in the FORUM Viewer versions 3.1 and 3.2 which may lead to misinterpretation of the optical coherence tomography (OCT) data. | Software in the Use Environment | Important Field Corrective Action for Forum Viewer letters were sent to all affected US customers on January 22, 2015 by certified mail. |  | 985 | 985 | Nationwide Distribution. | Terminated | March 04, 2016 | 407 | 985 | Computer
Z-1053-2015 | 70260 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | SmartSite Low Sorbing Infusion Set, Model No. 10015862 |  |  | January 13, 2015 | 2015 | CareFusion 303, Inc. | CareFusion is recalling the SmartSite Low Sorbing Infusion set due to disconnection and leakage issues. | Under Investigation by firm | CareFusion sent an Urgent Medical Device Recall Notification letter dated January 13, 2015, to all affected customers. The letter informs the customers of the problems identified and the actions to be taken. Customers were instructed to complete and return to CareFusion the recall response card. Customers with questions were instructed to CareFusion: Recall related questions, (888) 562-6018; Adverse Event Reports, (888) 812-3266; Technical Questions regarding the Alaris System, (888) 812-3229. |  | 7440 | 7,440 units | Nationwide Distribution | Terminated | March 15, 2016 | 427 | 7440 | Not_Computer
Z-1052-2015 | 70240 | No | OIY | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona (TASP) Tibial Articular Surface Provisional Shim, all sizes and Lots Catalog Number: 42-5279-002-00 thru 42-5279-009-04. |  |  | December 11, 2014 | 2014 | Zimmer, Inc. | Potential for delay in surgery and/or ball bearings to be left in the wound. Repeated cycles of sonic cleaning can break down the ball bearing retention feature (swage) on the shim. To prevent this, TASP shims must be manually cleaned and inspected per product insert, 87-6203-991-22. | Device Design | Zimmer sent an URGENT MEDICAL DEVICE CORRECTION notifications letter dated December 11, 2014 to affected distributors via electronic mail, and all hospital Risk Managers and distributors with affected inventory via courier. The letter identified the affected product, problem, instructions for responding to the formal recall notification.and actions to be taken. For questions call the customer call center at 1-800-348-2759. |  | 67963 | 67963 | Worldwide Distribution - USA Nationwide including DC, and PR and the countries of Australia, Austria, Belgium, Canada, Czech Republic, Dominican Republic, France, Germany, India, Iran, Israel, Italy, Japan, Luxembourg, Netherlands, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | October 23, 2015 | 316 | 67963 | Not_Computer
Z-1055-2015 | 70079 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | BD 10 mL Syringe BD Luer-Lok Tip with Blunt Plastic Cannula, Not for Skin Injection ; Catalog #303348, |  |  | December 19, 2014 | 2014 | Becton Dickinson &amp; Company | The stopper of the 10mL BD Luer-Lok syringe is not properly seated in the syringe. | Device Design | BD sent an Urgent Product Advisory Notice letter dated 12/19/2014. The letter identified the affected product, problem and actions to be taken. Customers are advised of the issue with the product and should notify their customers. The attached Product Advisory Response Form should be completed and faxed to 1-866-551-2707. Questions can be directed to 1-866-367-4703. |  | 564000 | 564,000 units | US Nationwide Distribution | Terminated | December 07, 2015 | 353 | 564000 | Not_Computer
Z-1057-2015 | 70414 | No | LYQ | Nonpowered goniometer. | Orthopedic | Orthopedic | 510(k) | KYPHON EXPRESS, Osteo Introducer System, Diamond and Bevel, Size 2, REF T34A, QTY: 1EA, Rx only, STERILE R. For use with spinal fixation devices. |  |  | January 13, 2015 | 2015 | Medtronic Sofamor Danek USA Inc | The product, labeled as containing, one bevel-tipped instrument and one diamond-tipped instrument, may contain either two bevel-tipped instruments or two diamond-tipped instruments. | Process control | On January 12, 2015, the Sales Managers and Sales Force were notified by Post-Market Quality regarding the imminent notification and follow-up activities. The recall notices were delivered to the Risk Managers at impacted consignee sites on January 13, 2015. These recall notices included a Risk Manager Letter informing them on the action, and a consignee-specific questionnaire to indicate their product on-hand. |  | 239 | 239 units | Nationwide and Canada | Terminated | July 13, 2015 | 181 | 239 | Not_Computer
Z-1056-2015 | 70362 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 Ultrasound System, Model 10433816, component 10040596 - UILK2. |  |  | January 12, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | The bolt holding the control panel of the ACUSON SC2000 in a fixed position may fail and the Control panel cannot be locked into position. | Component design/selection | Siemens Medical sent an Customer Safety Advisory Notification to all affected customers on January 12, 2015 , via Certified mail, return receipt requested.The letter identified the product the problem and the action needed to be taken by the customer. At this time there is not mitigation to avoid risk associated with this issue. Local Customer Service Engineers will be contacting each site to replace the console. Until the necessary repairs have been completed, please share the information with all uses of the device.If you have any questions, please contact your local service support person for information regarding timelines and status.Until the necessary repairs have been completed on your system, please share this information with all personnel within your organization who need to be aware of this issue.As always, patient safety issues are a very high priority. To date, no patient injury has been reported. This problem was discovered as part of our ongoing quality process.We sincerely regret any inconvenience this condition may cause in your daily operations. For further questions please call (650) 694-5398. |  | 230 | 230 | Worldwide Distribution - USA and the countries of Australia, Austria, Belarus, Canada, Ecuador, Finland, France, Georgia, Germany, India, Iraq, Ireland, Israel, Italy, Japan, Lithuania, Mexico, Nigeria, P.R. China, Philippines, Portugal, Republic Korea, Russian Fed., Saudi Arabia, Serbia, South Africa, Spain, Suriname, Sweden, Taiwan, Turkey, U.A.E. and United Kingdom. | Terminated | November 12, 2015 | 304 | 230 | Not_Computer
Z-1059-2015 | 70138 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | LINAC Accelerators: ARTISTE, ONCOR Avant Garde, ONCOR Impression Plus, ONCOR Expression, ONCOR Impression, PRIMUS HI, |  |  | December 19, 2014 | 2014 | Siemens Medical Solutions USA, Inc | A hardware update to replace the old PCB Assembly, Head Driver-G41 with PCB with overcurrent protection as a result of reports of the electric board inside the gantry showing signs of overheating, including thermal damage and smoke emission. | Component design/selection | Siemens sent a Customer Advisory Notice Letter TH016/14/S, dated December 19, 2014, to end users that identified the product, problem, and actions to be taken. End users were advised to read the Field Safety Notice and include it with the User's Manual until the hardware update is completed. |  | 310 | 310 | Worldwide Distribution - US Nationwide including Canada and Mexico. | Open, Classified |  |  | 310 | Not_Computer
Z-1058-2015 | 70239 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Puritan Bennett 980 Ventilator System, Model No. PB980 Ventilator (980xxxxxxxx), |  |  | January 12, 2015 | 2015 | Covidien LP (formerly Nellcor Puritan Bennett Inc.) | Covidien is issuing a voluntary field action for all Puritan Bennett 980 ventilators due to occasional GUI transient resets that last approximately 30 seconds. | Software design | Covidien sent an Urgent Field Safety Notice on January 12, 2015, to all affected customers. Customers were requested to forward the recall letter to anyone they may have transferred or assigned for use one of the potentially affected ventilators. Covidien is updating software that will modify ventilator alarms and status display messages during a GUI reset. The Puritan Bennett 980 Ventilator Operator Manual will be updated to delete,replace ventilator during a GUI transient reset, and to inform users to monitor the secondary status display until the primary display refreshes. Covidien will communicate with customers when the software and operator manual updates are available.Customers with questions were instructed to contact Technical Support Department at 800-255-6774, option 4, and then option 1.For questions regarding this recall call 303-876-8870. |  | 418 | 418 units | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | January 06, 2016 | 359 | 418 | Computer
Z-1060-2015 | 70149 | No | BTY | Predictive pulmonary-function value calculator. | Anesthesiology | Anesthesiology | 510(k) | Shape-HF Cardiopulmonary Testing System Disposable Patient Interface (DPI). Part Number 0004-4001. |  |  | December 23, 2014 | 2014 | Shape Medical Systems, Inc | This recall has been initiated because use of the impacted DPIs could result in erroneous testing results potentially leading to incorrect diagnosis and incorrect treatment. The magnitude of error for VO2 or VCO2 parameters could be approximately 20% with use of the impacted product. | Under Investigation by firm | An Urgent Medical Device Recall letter was mailed to consignees on 12/23/2014. The letter identified affected product, described the issue and asked for product to be returned. Shape Technical Service will contact customers to assist in the return of product. A response form was asked to be returned. Customers with questions can call Shape Technical Service at 1-888-906-6266. |  | 564 | 564 devices | Distributed in the states of AZ, FL, GA, MA, MN, NJ. NY, PA, TX, and WI. | Terminated | April 14, 2015 | 112 | 564 | Not_Computer
Z-1008-2015 | 70127 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | MAQUET Servo Humidifier 163; Model No.: 01-06-8125-8, catalog no. XKC01-06-8125-8. |  |  | December 04, 2014 | 2014 | Teleflex, Inc. | Connector mount cracks which may lead to a leak failure during use and necessitate immediate replacement in the breathing circuit. If leak exists and is left untreated, patient may be deprived of adequate anesthetic gases over a period of time and serious adverse health consequences may occur. | Process control | An Urgent - Field Safety Notice, dated December 4, 2014, was sent to end users that identified the product, problem, and action to be taken. Customers were requested to immediately check inventory, discontinue use and distribution, quarantine. Customers were to return the Acknowledgement form and request a return authorization number from their local Maquet representative for affected devices. A credit will be issued upon receipt of the returned affected product. For questions or for support concerning this issue, a local Maquet representative should be contacted at fieldactions@maquet.com. |  | 18 | 18 | Distributed in the states of FL, MI, WV, and KS. | Terminated | March 21, 2017 | 838 | 18 | Not_Computer
Z-1051-2015 | 70257 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | MAQUET, SERVO Humidifier 163, Rx Only, Inmed Mfg. Sdn Bhd. Catalog No: XKC01-06-8125-8, Model No. 01-06-8125-8. |  |  | January 12, 2015 | 2015 | Teleflex Medical | Defects: Cracks that may occur in the center of the corrugated flex tube of the 15 mm and 22 mm ISO female connector. | Process control | Consignees were notified by an Urgent - Field Safety Notice, dated 01/12/2015. The letter identified the affected devices, a description of the problem, and actions to be taken. Customers were asked to check their inventory for any affected product, to cease use and distribution, and immediately quarantine any affected product. Customers are to complete and return the Acknowledgement form and to their local MAQUET representative, who provide a return number. A credit will be issues upon receipt of the affected product. Customer should contact their local MAQUET representative at fieldactions@maquet.com with questions or for support this issue. |  | 100 | 100 units | Distributed to FL, KS, MI and WV. | Terminated | May 10, 2016 | 484 | 100 | Not_Computer
Z-1063-2015 | 70361 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ACUSON S2000 Automated Breast Volume Scanner (ABVS), model number 10434730. |  |  | January 12, 2015 | 2015 | Siemens Medical Solutions USA, Inc. | There is a low probability the scanner arm will become completely detached from the scanner column. | Vendor change control | Siemens sent a Customer Safety Advisory Notification to all affected sites on January 12, 2015, certified mail, with return receipt. The letter identified the affected product, the issue, and potential risk. Customers were asked to immediately discontinue use of the affected product. A Customer Service Engineer will contact customers to schedule an appointment to make the necessary repairs. Questions should be directed to a local service support person for information regarding timelines and status. Customers are to share the information with all personnel, who need to be aware of the issue. |  | 24 | 24 units | Worldwide Distribution -- US, including the states of TX and MN, and the territory of P.R.; and, the countries of China, Germany, Russian Fed, Saudi Arabia, Taiwan, France, Thailand, and Italy. | Terminated | June 10, 2015 | 149 | 24 | Not_Computer
Z-1062-2015 | 69863 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | CARDIOSAVE Intra-Aortic Balloon Pump and CS300 Intra-Aortic Balloon Pump; and, non-sterile, not for human use, SENSATION and SENSATION PLUS IABC demonstration units. Part Number 0998-00-0800-xx; 0998-00-3023-xx; 0998-UC-0800-xx; 0998-UC-3023-xx |  |  | December 16, 2014 | 2014 | Maquet Datascope Corp - Cardiac Assist Division | Ferrule Cap was not removed on some fiber-optic Sensation and Sensation Plus IABC demonstration units. If fiber-optic connector is inserted into the Maquet CS300 or CARDIOSAVE IABP, cap will prevent connection of unit and may become dislodged and remain in IABPs fiber-optic receptacle. | Device Design | Maquet Inc. sent a Product Recall Medical Device Field Correction letter/Field Correction response form, dated 12/16/2014 via Fed Ex. The letter identified the affected product and the reason for the recall. Users were asked to inspect their inventory for the affected units. If customers have any affected product, they are to follow the instructions provided. Additionally, users are to complete and return the attached Field Correction Response Form to acknowledge receipt of the letter and to indicate completion of the steps outlined. Completed forms should be returned by e-mailing a scanned copy to fieldactions@maquet.com or faxing a copy to 1-973-396-3607. Technical questions should be directed to the Technical Support Department at 1-800-777-4222. |  | 1202 | 838 units-US; 364 units - OUS | Worldwide Distribution | Terminated | June 09, 2017 | 906 | 1202 | Not_Computer
Z-1064-2015 | 69388 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | Digital Mobile X-ray Systems used to take diagnostic radiographic exposures or patients. |  |  | September 22, 2014 | 2014 | Virtual Imaging, Inc. | There is a potential for unintended and erratic motion. | Component design/selection | Virtual Imaging sent an FIELD CORRECTION NOTIFICATION letter dated September 22, 2014, to all affected customers. The letters advised direct accounts that a Virtual Imaging Field Engineer with contact them to schedule an appointment for the replacement of the Digital Motion Control Kit. Customers with any concerns about this situation can contact Virtual Imaging's Technical Support at 561-893-8500 or via e-mail at technicalsupport@vifla.com. |  | 316 | 316 (US) 6 (Foreign) | USA (nationwide) to the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MO, MS, MT, NC, NE, NH, NJ, NM, NY, OH, OK, OR, PA, SC, TN, TX, VA, WA, WI, WV) + US Army facility in South Korea; and the country of Canada. | Open, Classified |  |  | 316 | Not_Computer
Z-1067-2015 | 70157 | No | GWE | Evoked response photic stimulator. | Neurology | Neurology | 510(k) | EPWorks software used in the Protektor 32 |  |  | January 02, 2015 | 2015 | Natus Medical Incorporated | Software error occurs when using remote monitoring; if the remote user tries to stop the free run waveform group, the system will display a message informing the user that they do not have sufficient privilege. | Software design | Natus Medical Inc. DBA Excel-Tech Ltd (Xltek) issued medical device notification letters dated December 22, 2014 to their customers vial email and via courier service on January 2, 2015. Within the notice, customers are instructed are on to removal the previous installation and also on how to install the new program upgrade. Natus Medical Inc. has requested that the Previous Version Removed from Use Confirmation form be completed by each customer and returned. Customers who require assistance can contact Natus Technical service by calling 800-387-7516 or via email Oakville_Technical_Service@xltek.com or ots@natus.com. |  | 248 | 248 units | US Nationwide Distribution | Terminated | March 22, 2017 | 810 | 1176 | Computer
Z-1046-2015 | 70103 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ABL90 FLEX analyzer; Model number 393-090. A portable, automated analyzer that measures pH, blood gases, electrolytes, glucose, lactate, bilirubin and oximetry in whole blood. |  |  | December 08, 2014 | 2014 | Radiometer America Inc | The ABL90 analyzer does not always use the most recent calibration data to calculate patient results. This can in some cases lead to a biased patient result. | Software design | Radiometer sent a Field Safety Notice dated December 17, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. What you should do short term: " Please ensure that the operator always closes the inlet when prompted by the analyzer, as per the labeling. " In case the inlet, by mistake, is not closed as per the labeling, please ensure that the following is carried out before measuring a patient sample: " Close the inlet and wait for the analyzer to go to READY (approximately five minutes) " Open the inlet and then close again (to initiate a Rinse) " Wait for the analyzer to go to READY (approximately 1Âż minutes) " Please return the confirmation fax form no. 1 to your RADIOMETER representative when the above actions have been implemented Final Solution provided by RADIOMETER: A mechanism will be provided in an upgraded version of the analyzer software to eliminate the possibility of this error. The new software version will be installed by your local engineer when available.Customers were instructed to ensure that this letter is distributed to the final end-user. Customers with questions were instructed to contact Radiometer America Technical Support at 1-800-736-0600, opt 4. |  | 5002 | 5002 units | Worldwide Distribution -USA including CT, MN, FL, CA, PA, NY, KY, MA, WA, MI, TX, NC, IL, OH, SC, OK, NJ, MD, VA, UT, AL, NM, GA, LA, OR, IN, WI, AZ, ND, CO, MO, WV, ID, TN, VT, NE, IA, NH, MS, ME, AR and Internationally to Canada, Australia, Austria , China, Czech Republic, Denmark, Estonia, Lithuania, Romania, Slovenia, United Arab Emirates, Bolivia, Brazil, Indonesia, Israel, Kuwait, Malaysia, Mexico, Peru, Qatar, Saudi Arabia, Serbia and Montenegro, Syrian Arab Republic, Thailand, Vietnam, Finland, France, Germany, Hong Kong, Hungary, Ireland, Italy, Japan, Netherlands, New Zealand, Norway, Poland, Republic of Korea, Russia, Singapore, South Africa, Spain, Sweden, Switzerland, Taiwan, Turkey, and United Kingdom. | Terminated | September 23, 2015 | 289 | 5002 | Computer
Z-1070-2015 | 70426 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Aquarius iNtuition Client Viewer. Findings Workflow module, RECIST 1.1: |  |  | January 21, 2015 | 2015 | TeraRecon, Inc. | Software anomaly related to RECIST1.1 target lesion evaluation criteria in Findings Workflow Module within the Aquarius iNtuition Client Viewer. | Software design | Medical Device Correction letters were sent to all affected users on January 21, 2015 by both e-mail and physical mail - certified mail for us customers and FEDEX for international customers. |  | 91 | 91 | Worldwide Distribution-US (nationwide) and the countries of Canada, Switzerland, Germany, France and Italy. | Terminated | September 30, 2015 | 252 | 91 | Computer
Z-1069-2015 | 70381 | No | MAX | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Timberline Cranial Caudal and Posterior Blades, 40 - 180mm Model # 8734-2XXX, 8734-5XXX is a reusable device provided within an aluminum sterilization tray along with other instruments for use in a Timberline procedure |  |  | February 22, 2012 | 2012 | Biomet Spine, LLC | Biomet Spine, LLC is recalling the Timberline Cranial Caudal and Posterior Blades, 40-180mm due to oversize of the set screw hex feature. | Nonconforming Material/Component | Biomet initiated the recall on February 22, 2012, No letters were or will be sent to customers. 100% retrieval was achieved by working directly with the customers.For questions regarding this recall call 303-501-8548. |  | 213 | 213 | Distributed to TX, IN, MO, and NY. No foreign/govt/military/VA. | Terminated | March 23, 2015 | 1125 | 213 | Not_Computer
Z-1066-2015 | 70298 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | System 6 Aseptic Housing. |  |  | January 07, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Stryker is recalling System 6 Aseptic Housing, due to a failed Continuous Bond Line Test from a housing without a continual weld. Customers are instructed to return recalled product to Stryker. | Under Investigation by firm | Stryker sent notice to customers via Fed ex overnight on 1/16/2015, regarding its voluntary recall of Stryker System 6 Aseptic Housings because the laser welder may have failed to produce a continuous bond line for the four specific lots: 13205, 13209, 13210, 13212. Customers are instructed to discontinue use, complete the business reply form, and return all recalled product. Customers may contact Kara Spath at 269-389-4518, or kara.spath@stryker.com. |  | 399 | 399 | Worldwide Distribution-US (nationwide) including Puerto Rico and the states of AR, CA, CO, FL, GA, HI, MD, NC, NY, PA, SC, TX, and VA, and the countries of Australia, Canada, Switzerland, Hong Kong, South Africa, Chile, Korea, Poland, India, Colombia, The Netherlands, and France. | Terminated | March 09, 2016 | 427 | 399 | Not_Computer
Z-1065-2015 | 69837 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | SoftPath Laboratory Information System. Versions 4.3.0.8, 4.3.0.9, 4.3.0.10, 4.3.0.11, 4.3.0.12, 4.3.0.14, 4.3.0.15, and 4.4.0.0 |  |  | November 26, 2014 | 2014 | SCC Soft Computer | Modifications to diagnostic text may be: 1) Saved to the database but not appear on the report sent to the physician; or 2) Documented on the report, but not saved to the database. | Software design | On 11/26/2014 SCC Soft Computer sent a Correction Communication to their consignees: Corrective Action: Versions 4.3.0.8 through 4.4.0.2 will have a mandatory hotfix created. Due to the amount of versions affected, the range of availability dates will be 11/24/2014 through 12/30/2014. Utility: A utility will be made available to be run on your Live system that will identify cases in which there is a disparity between the diagnosis text in the document and the database. The utility output will be provided for client review and validation. |  | 108 | 108 units | Distribution US Nationwide and CANADA. | Open, Classified |  |  | 108 | Not_Computer
Z-1078-2015 | 70248 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | Hudson RCI, Sheridan, Preformed Tracheal Tubes, Cuffed and Uncuffed, Rx Only, Sterile, Teleflex Medical. |  |  | January 09, 2015 | 2015 | Teleflex Medical | Labeling; The affected product packaging is incorrectly labeled as being a cuffed oral ET tube; The correct product labeling is uncuffed nasal ET tube. | Labeling mix-ups | Teleflex Medical sent an " Urgent Medical Device Recall Notification" dated January 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They were instructed to immediately discontinue use and quarantine any of the affected products, return the product and to complete the enclosed Recall Acknowledgement Form even if they do not have the affected product in stock. Distributors were instructed to communicate the recall to any of their customers who may have received product included within the scope of the recall. Have the customers return any affected product to them, together with a completed Recall Acknowledgement Form.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 300 | 300 units | Worldwide Distribution - US Distribution to the states of FL, GA, IL, KY, MA, MI, MO, MS, OH, TX, WA and WV., and to the countries of Canada and the Netherlands. | Open, Classified |  |  | 300 | Not_Computer
Z-1087-2015 | 70259 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | The Siemens Symbia T series is intended to aid in detecting, localizing, diagnosing, staging and restaging of lesions, tumors, disease and organ function for the evaluation of diseases and disorders. The images produced by the system can also be used by the physician to aid in radiotherapy treatment planning and interventional radiology procedures. |  |  | October 23, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | There is a potential for the linear bearing car to fail allowing the ball bearings to fall out. | Component design/selection | Consignees were not notified of this issue. |  | 1075 | 1,075 systems (global) | Worldwide Distribution-US (nationwide) including DC and Puerto Rico and the states of AK, AL, AR, AZ, CA, CO, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, TN, TX, UT, VA, WA, WI, WV, and WY), and the countries of AUSTRALIA , AUSTRIA, BANGLADESH , BELGIUM , BRAZIL , BULGARIA , CANADA , CHILE , CHINA , COLOMBIA , CROATIA , CZECH REPUBLIC , DENMARK , ECUADOR , EGYPT , FINLAND , FRANCE , GERMANY , GREECE , HONG KONG , HUNGARY , INDIA , INDONESIA , IRAN , IRELAND , ISRAEL , ITALY , JAPAN , KUWAIT , LATVIA , LITHUANIA , MEXICO , MOROCCO , NETHERLANDS , NEW ZEALAND , NORWAY , PERU , PHILIPPINES , POLAND , PORTUGAL , QATAR , ROMANIA , RUSSIA , SAUDI ARABIA , SINGAPORE , SLOVAKIA , SOUTH AFRICA , SPAIN , SWEDEN , SWITZERLAND , TAIWAIN , THAILAND , TUNISIA , TURKEY , UNITED ARAB EMIRATES, UNITED KINGDOM , and VENEZUELA. | Completed |  |  | 1794 | Not_Computer
Z-1084-2015 | 70088 | No | FWZ | Manual operating table and accessories and manual operating chair and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | The MAYFIELDÂż Ultra 360 Base Unit for head positioner, one unit per package. |  |  | December 19, 2014 | 2014 | Integra LifeSciences Corporation | Through an investigation of complaints, Integra found that there is the possibility for the locking handle mechanism of the Base Units to fail during use. | Nonconforming Material/Component | Integra sent an Urgent - Voluntary Medical Device Recall letter dated December 19, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to take the following actions:1. Review your MAYFIELD equipment and determine if you have any MAYFIELD Ultra 360 Base Units2. Once identified we recommend that you discontinue use and remove them from service until they are repairedby an authorized Integra Repair Center3. Complete the attached Recall Acknowledgement and Repair Form4. Upon receipt of the Recall Acknowledgement and Repair Form, Customer Service will contact you and providean RMA number and instructions for returning the productor questions call 609-936-6822.For questions contact Customer Service:USA 800-654-2873Outside USA 609-936-5400 |  | 1053 | 1053 Units | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV, WY, and DC. (25 government consignees) and in the foreign countries: CANADA, ALGERIA, ARAB EMIRATES, AUSTRALIA, AUSTRIA, BELGIUM, BRAZIL, CHILE, CHINA, COLUMBIA, CYPRESS, DENMARK, DOMINICAN REPUBLIC, FRANCE, GERMANY, GREAT BRITAN, HONG KONG, HUNGARY, INDIA, INDONESIA, ISRAEL, ITALY, JAPAN , JORDAN, KOREA, LIBYA, MACEDONIA, MAYLASIA, MEXICO, NETHERLANDS, POLAND, PORTUGAL, PUERTO RICO, ROMANIA, SAUDI ARABIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, TAIWAN, THAILAND, and TURKEY. | Open, Classified |  |  | 1053 | Not_Computer
Z-1076-2015 | 70051 | No | FTL | Surgical mesh. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | The SurgiWrap Bioresorbable Adhesion Barrier Film. SurgiWrap 100x130x0.05mm |  |  | August 26, 2014 | 2014 | MAST Biosurgery USA, Inc | The expiration date of the bar code on the envelope label did not match the expiration date next to the hourglass. The printed expiration date was August of 2017, while the bar code date read as July 31, 2017. The correct expiration date was July 31, 2017. | Labeling Change Control | MAST Biosurgery labeled 49 new envelopes and sent them to the distributor, and subsequently had the distributor remove the contents of the outer envelope (instructions for use and the foil pouches containing the product) and place them into the newly supplied and properly labeled outer envelopes provided by them. The distributor requested to send back the original 49 enveloped.The second batch of new labels was inspected and attached to the new outer envelope to create a properly labeled outer envelope prior to shipping. The firm requested that once the distributor replaced the product and the IFU into the newly labeled envelopes that they provide documentation that the product had been repackaged into the new envelopes. The distributor was also requested to verify that they did not ship any product to customers with the incorrect label. The firm had the distributor return the 49 improperly with the incorrect labels and the first set of "corrected" labels to them for destruction.Further questions call (858) 550-8050. |  | 49 | 49 pieces | Worldwide Distribution to the country of China. | Terminated | February 24, 2015 | 182 | 49 | Not_Computer
Z-1088-2015 | 70279 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Precise Digital Accelerator |  |  | January 07, 2015 | 2015 | Elekta, Inc. | Elekta has identified a batch of securing bolts that fix the Gantry counter frame weight stack may fail. | Other | Elekta sent an Urgent Important Field Safety Notice on January 7, 2015, to all affected customers. Elekta recommended that customers replace the M20 x 50 MM bolts marked 'OF' on the gantry counterbalance frame immediately. The digital linear accelerator must not be used until these 'OF' bolts are replaced. Elekta recommended that customers put the machine in a parked position at 0, and attach the gantry locking straps. A Modification kit will be available as soon as possible which contains the correct bolts. Customers were instructed to complete the Important Field Safety Notice Acknowledgement form and return to their local Elekta Office or Representative. Customers with questions were instructed to contact their local Elekta representative.For questions regarding this recall call 404-993-5886. |  | 46 | 46 | Worldwide Distribution - USA including AZ, CA, FL, KS, LA, MN, MO, NY, NC and Internationally to Austria, Bulgaria, Canada, China, Germany, Hungary, Italy, Malaysia, Malta, Netherlands, Norway, Russia, Spain, Sweden, Taiwan and Ukraine. | Terminated | April 14, 2016 | 463 | 46 | Not_Computer
Z-1083-2015 | 70351 | No | LFJ | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Hemodialysis Catheters and PICCs: Product identification: MCDLTSL24P (SF X 24CM LT HEMO-CATH), MC061422 (14FX15CM SLX DBL LMN ST CATH), |  |  | January 21, 2015 | 2015 | Medical Components, Inc dba MedComp | Drug products contained within the kits may have been rendered ineffective by a duplicate sterilization process. The kits include one or more of the following drugs: 5ml ampule of lidocaine, 3ml Chloraprep one-step, 10ml pre-filled saline syringe, Hygea saline wipes, and 3ml tinted Chloraprep. No injuries or illnesses were reported. | Process control | Medcomp sent a Product Alert/Recall letter dated January 21, 2015 to customers via email. The letter identified the affected product, problem and actions to be taken. Medcomp requested the return of all affected product. Customers were instructed to contact their customer service representative for instructions for returning the product. |  | 243 | 243 | US Nationwide Distribution | Terminated | August 03, 2015 | 194 | 243 | Not_Computer
Z-1082-2015 | 70175 | No | FOZ | Intravascular catheter. | General Hospital | General Hospital | 510(k) | ARROWg+ard Blue PLUS Pressure Injectable Two-Lumen CVC Kit. The Multiple-Lumen Catheter permits venous access to the central circulation by way of the femoral jugular, or subclavian veins. |  |  | January 02, 2015 | 2015 | Arrow International Inc | Foreign material found within the kits. The foreign material, white or black particles and/or thread, is visible and can beidentified by the user. | Process control | An urgent Medical Device Recall Notification, dated January 2, 2015, was sent to end users requesting they check their inventory and return affected product per Arrow's instructions. |  | 1195 | 1195 | Nationwide | Open, Classified |  |  | 1195 | Not_Computer
Z-1077-2015 | 70332 | No | FYA | Surgical apparel. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Flyte Togas are components of the Stryker Flyte System and are intended to be worn over any Stryker Flyte Helmet. This toga provides Level 1 and 4 protection according to the Association for the Advancement of Medical Instrumentation (AAMI) Liquid Barrier Performance and Classification System (PB70:2012). |  |  | December 22, 2014 | 2014 | Stryker Instruments Div. of Stryker Corporation | A single lot (11121171) of Flyte Toga Peel-away X-Large (catalog number 0408-830-100) was labeled with an incorrect expiration date of 2412-12-01. The correct expiration date should have read 2014-12-01. | Error in labeling | Stryker sent an Urgent Medical Device Recall Notification letter dated December 22, 2014, to all affected customers. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and responding to the formal recall notification. Customers were instructed to do the following:1. Immediately review this Recall Notification2. Immediately check all stock areas and/or operating room storage and quarantine any affected product found.3. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify howmany, if any, affected items are currently in their inventory. Complete and return the BRF even if theydo not have any affected product on hand.Note: Your signature on the BRF indicates that you received and understand this Notification andhave followed the instructions in this Notification.4. If customers further distributed this product, they were instructed to forward this letter and the attached Business Reply Form (BRF) to all affected locations. Please indicate each location on the BRF. Completed Business Reply Forms should be faxed to 866-521-2762.Upon receipt of the Business Reply Form, Stryker will email a pre-paid shipper. This shipper can be used to return the recalled Flyte Toga(s) to Stryker.Upon receipt of the recalled Flyte Toga(s), credit will be applied to the customer's account.For questions regarding this recall customers were instructed to contact Stryker Instruments: Monday-Friday 8am-5pm (EST) at 269-389-2458.For questions regarding this recall call 269-389-2458. |  | 1530 | 1530 units (170 cases) | Nationwide Distribution including AL, AR, AZ, IA, IL, IN, LA, MA, NC, NE, NM, SC, TN, TX, WI, and WV. | Terminated | May 05, 2015 | 134 | 1530 | Not_Computer
Z-1081-2015 | 70036 | No | LZU |  |  | Ophthalmic | 510(k) | VisiPlugST for the Lacrimal Efficiency Test. It is also sold in the EU as CollaSyn Plugs. Model number 1813, 1814, 1815, 1903, 1904, and 1905. |  |  | November 26, 2014 | 2014 | Lacrimedics Inc | Moisture level may cause the VisiPlugST (or CollaSyn Plugs) to break apart, or crumble upon contact. The firm's data also suggests that a 4 year shelflife may not be sustained. | Nonconforming Material/Component | The firm notified their foreign consignees on 11/26/14. The firm made a decision to conduct a domestic recall on 12/4/14 and will notify US consignees Consignees will be provided with a return label and request any devices that they still have be returned to the firm.Any consignees that have gone out of business will be sent a certified letter to their last known address.CORRECTIVE ACTIONS:The firm has halted sales of the device until the drying method can be improved. Any remaining stock has been placed in quarantine and will be scrapped. |  | 781 | 781 units | Worldwide Distribution-US (nationwide) including the states of AR, CA, FL, ID, IL, KS, MA, MI, NC, NJ, NY, OR, PA, and TX and to the countries of Australia, Brazil, Czech Republic, France, Ireland, Mexico, Peru, Switzerland, and United Kingdom. | Terminated | March 23, 2015 | 117 | 781 | Not_Computer
Z-1080-2015 | 70177 | No | CAZ | Anesthesia conduction kit. | Anesthesiology | Anesthesiology | 510(k) | ARROW Epidural Needle Component (product number AN-05501). Epidural needles permit access to the epidural space and are used as a conduit for epidural catheter placement. |  |  | January 05, 2015 | 2015 | Arrow International Inc | The product labeling does not specify the 5 year shelf life of the product. | Labeling design | Urgent Medical Device Recall Notifications, dated January 2015, were sent to direct accounts that identified the product, problem, and actions to be taken. Customers were asked to discontinue use and return affected product. |  | 63022 | 63022 | Distribution US Nationwide, Canada and Mexico. | Open, Classified |  |  | 63022 | Not_Computer
Z-1085-2015 | 70258 | No | CAE | Oropharyngeal airway. | Anesthesiology | Anesthesiology | 510(K) Exempt | LMA (Laryngeal Mask Airway) Supreme (New Cuff) Size 5, Sterile, Single use, |  |  | January 13, 2015 | 2015 | Teleflex Medical | The labelling of units from the affected lot indicates that it contains a size 5 LMA Supreme (New Cuff), but may contain a size 4 LMA Supreme (New Cuff). | Other | Teleflex sent an Urgent Medical Device Recall Notification letter dated January 13, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediately discontinue use and quarantine any products and to return it to Teleflex Medical. Customers were asked to complete the enclosed Recall Acknowledgement Form via fax even if they do not have the affected product in stock. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 2640 | 2640 units | Worldwide Distribution - U.S Nationwide in the states of: AZ, CA, DE, FL, GA, IL, KY, MO, NY, OK, PA, TX, VA &amp; WA and the country of Canada. | Open, Classified |  |  | 2640 | Not_Computer
Z-1089-2015 | 70376 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Telluride Percutaneous Rod Inserter Long, Model #7706-1904. The Inserter is a reusable device provided within an aluminum sterilization case along with other orthopedic manual surgical instruments. The case consists of three trays of instruments, and the middle tray contains a quantity of two Inserters. |  |  | December 19, 2011 | 2011 | Biomet Spine, LLC | The Telluride Percutaneous Rod Inserter Long does not contain a feature that locks the Rod in place during use. There is a potential for the Rod to disengage from the Inserter during insertion, and could lead to harm to the patient. | Device Design | Consignees were informally contacted 12/16/2011 of units in the field requesting retrieval. All units were in Lanx's (Biomet's) possession at the time of the recall initiation on 12/19/2011. |  | 10 | 10 | Distributed to TX, OH, NE and NY. | Terminated | March 23, 2015 | 1190 | 10 | Not_Computer
Z-1090-2015 | 70166 | No | FWP | Chin prosthesis. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | MEDPOR Surgical Implant Contoured Two Piece Chin 5 mm; Part Number 86001; Lot Numbers A1405048, A1404026 |  |  | December 17, 2014 | 2014 | Stryker Craniomaxillofacial Division | Potential for implant damage, implant cracking off/breaking intra-operatively. | Process control | Stryker sent an Urgent Medical Device Recall Notice dated December 23, 2014, to all affected customers with a response form to be completed and returned upon receipt and review. The letter identified the affected product, potential hazards, risk mitigation and the actions to be taken. For questions contact Stryker CMF at 1-800-962-6558, or e-mail at CMF-custserv@stryker.com. |  | 1051 | 1051 | Worldwide Distribution - US Nationwide in the states of PA, KY, VA, CT, UT, MD, AR, FL, TX, AL, DE, IL, IA, MI, CO, OH, GA, CA, DC, MO, LA, WI, ID, SC, NC, NY, MA, NH, CA, OK, MN, KY, IN, OR, TN and the countries of: Argentina, Australia, Canada, Chile, China, Netherlands, France, Spain, South Korea, Colombia, Switzerland, Singapore, and South Africa. | Terminated | November 06, 2015 | 324 | 1051 | Not_Computer
Z-1094-2015 | 70055 | No | MTJ |  |  | General &amp; Plastic Surgery | 510(k) | Hemostatic Bone Putty, for use as a water-soluble implant material and for use in the control of bleeding from bone surfaces. |  |  | January 06, 2012 | 2012 | Synthes, Inc. | Revision A of the Hemostatic Bone Putty Brochure (J10847A) was sent to sales consultants instead of Revision B (J10847B), which corrected an error in the indication statement. | Labeling mix-ups | On/about 12/22/2011, emails containing Revision B brochure, as well as instructions to delete/destroy Revision A, were sent to 38 affected individuals (sales consultants). On/about 1/04/2012, follow up phone calls were made to each member of the sales force who had received Rev. A (110847 A) to confirm that they had replaced the Rev. A with the Rev. B (110847 B). In addition, the members of the sales force were asked whether or not the Rev. A had been disseminated to any HCP's. During the follow up conversation, one member of the sales force stated that he had provided the previous version to a surgeon. On 01/06/2012, the sales consultant spoke to the surgeon in person and the surgeon confirmed that he had disposed of the Rev. A. |  | 38 | 38 | Distributed in the state of WI | Terminated | March 30, 2015 | 1179 | 38 | Not_Computer
Z-1100-2015 | 70280 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | AQUAPAK 340 SW, 340 ML W/040 ADAPTOR, FREN. |  |  | January 14, 2015 | 2015 | Teleflex Medical | The packages may contain foreign material (black particles). | Material/Component Contamination | Teleflex sent and Urgent Medical Device Recall Notification letter dated January 14, 2015 to all affected customers. The letter identified the affect product, problem and actions to be taken. Customers were instructed to immediately discontinue use and quarantine any products with the affected catalog numbers and lot numbers. Customers were asked to complete and return the enclosed Recall Acknowledgement Form. For questions contact your local sales representative or Customer Service at 1-866-246-6990. |  | 15779607 | 15,779,607 units | Worldwide Distribution - U.S. Nationwide and the countries of: Australia, Bahrain, Belgium, Bermuda, Brazil, Canada, China, Ecuador, Germany, Great Britain, Hong Kong, Indonesia, Japan, Jordan, Malaysia, Mexico, New Zealand, Philippines, Singapore, South Korea, Saudi Arabia, Taiwan, Thailand, Uruguay and Zambia. | Open, Classified |  |  | 15779607 | Not_Computer
Z-1114-2015 | 69723 | No | HSD | Shoulder joint humeral (hemi-shoulder) metallic uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Arthrosurface HHXL (OVO) 15.6x32mm Taper Post Fixation Component; Catalog Number: 8156-0032, Part Number: 156-0032. HemiCAP Contoured Articular Prosthetic for Humeral head resurfacing. Orthopedic shoulder prosthesis component. |  |  | November 24, 2014 | 2014 | Arthrosurface, Inc. | Fully threaded Taper Post Fixation components not cleared for marketing in the US with current indications. | No Marketing Application | Arthosurface sent an Urgent Medical Device Recall letter on November 24, 2014 via UPS, to all affected customers. The letter informed consignees of the recall and asked them to return two completed forms, 1. Acknowledgement &amp; Receipt Form and 2. Inventory Record via fax, email, or mail. Customers with questions were instcted to contact their local sales representatiave or email RA@arthrosurface.com. For questions regarding this recall call 508-528-4604. |  | 187 | 187 | Worldwide Distribution - USA nationwide. Internationally distributed to Europe, Canada, and Australia.Only products in US affected by this recall. | Terminated | March 27, 2015 | 123 | 786 | Not_Computer
Z-1115-2015 | 70304 | No | NDP | Endosseous dental implant accessories. | Dental | Dental | 510(K) Exempt | GemLock Long Hex Driver, Catalog RHL2.5; Also distributed within the Tapered SwissPlus &amp; SwissPlus Implant Systems Kits, Surgical, Complete, Catalog OPCST; Tapered Screw-Vent Implant System Surgical Kit, Complete, Catalog TSVKIT. |  |  | January 22, 2015 | 2015 | Zimmer Dental Inc | The Gemlock Long Hex Driver may not fit into the Fixture Mount Transfer (FMT) or the internal hex of the implant. | Device Design | An Urgent Device Recall Notice was sent on 1/22/15 to their domestic and foreign customers who purchased the Gemlock Long Hex Driver. The letter informs the customers of the problems identified, risks, and the actions to be taken. Customers are instructed to review the notice and check their inventory for the affected units. All affected product is to be immediately quarantined and removed from service. Customers are to contact Simmer Dental Customer Service at 1-800-854-7019 to obtain a Return Authorization Number and set up return of the products. Customers should complete the attached Business Reply Form and fax it to (574) 372-4265 or e-mail to corporatequality.postmarket@zimmer.com. Customers with questions are instructed to contact Zimmer Dental Customer Service at (800) 854-7019, Mon-Fri, 7 am-5pm (Pacific).Zimmer Dental is expanding the recall to include 1 additional lot for catalog no.TSVKIT. Zimmer Dental plans to inform an amended customer notification letter to their US and foreign consignees. The letter will go out to customers on 4/20/15 to inform them of the expanded lot. |  | 627 | 627 units | Worldwide Distribution -- US, Argentina, Australia, Bulgaria, Canada, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, Egypt, France, Germany, Hong Kong, Hungary, Iran, Israel, Italy, Japan, Lebanon, Lithuania, Malaysia, Mexico, Morocco, Netherlands, Pakistan, Panama, Poland, Romania, Russia, Saudi Arabia, Serbia, Taiwan, Thailand, Tunisia, Turkey, and UAE.Expanded Distribution: Spain | Terminated | August 03, 2015 | 193 | 627 | Not_Computer
Z-1116-2015 | 70492 | Yes | DWC | Cardiopulmonary bypass temperature controller. | Cardiovascular | Cardiovascular | 510(k) | Terumo HX2 Temperature Management System provides temperature control of two independent water circuits that directly controls the temperature of patient blood and cardioplegia solution during cardiovascular surgery. |  |  | February 06, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Malfunctioning mixing valve and loss of cooling and/or heating capability. System goes to standby mode and shows error messages EOd (invalid mix valve measured range of motion) or EOE (valve unable to move or complete a motion). Mixing valves need to be replaced at regular intervals. | Device Design | On 2/6/2015, URGENT MEDICAL DEVICE RECALL notifications were mailed to the affected hospitals. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call customer service at 1-800-521-2818. |  | 103 | 103 units | US (nationwide) including the states of AR, CA, CO, CT, FL, IA, IL, KS, KY, MD, MI, MO, MS, NY, OH, PA, TX, UT, WA, and WI and the country of Japan. | Open, Classified |  |  | 103 | Not_Computer
Z-1117-2015 | 70275 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | HARMONIC ACE Curved Shears Instructions for Use, 5MM, IS1200; for use with the da Vinci Standard Surgical System. |  |  | January 21, 2015 | 2015 | Intuitive Surgical, Inc. | Internal labeling review found that the IFU for the Harmonic Ace incorrectly instructs the user on how to manually open the grips. | Error in labeling | A Field Safety Notice entitled, Urgent Medical Device Correction, and dated, January 21, 2015, were sent to customers. The letter identified the problem and the reason for the recall as well as the risk to health. Customers are asked to ensure that all affected personnel are informed of the notice. Additionally, customers are to discard their current Harmonic Ace IFUs. Customers are to refer to the specified manual, reference guide, and instruction card mentioned in the letter. A copy of the notice is to be retained with the da Vinci Standard, S or Si user manual. Intuitive Surgical will remove the Harmonic ACE IFU from all future shipments of the product. Representatives will be available by phone to answer questions related to the Correction. |  | 12467 | 12,467 - both sizes | Worldwide Distribution -- US, Australia, Austria, Belgium, Brazil, Canada, Chile, China, Colombia, Czech Republic, Denmark, Dominican Republic, Ecuador, Finland, France, Germany, India, Israel, Italy, Japan, Lebanon, Mexico, Monaco, Netherlands, Pakistan, Panama, Puerto Rico, Qatar, Romania, Russian Federation, Saudi Arabia, Singapore, Slovakia, South Korea, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, and the United Kingdom. | Terminated | June 01, 2015 | 131 | 12467 | Not_Computer
Z-1121-2015 | 69782 | No | KRB | Thermodilution probe. | Cardiovascular | Cardiovascular | 510(k) | Pulsiocath PiCCO ProAQT Monitoring Kit, single-use cardiac output (CO) sensor, one unit per plastic sterilizable package. |  |  | November 04, 2014 | 2014 | Pulsion Medical Inc | Internal testing and investigation of the packaging revealed cracks and holes in primary packaging | Packaging | Pulsion sent an Urgent Medical Device Recall notification letter on 11/04/14 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers where instructed to:" Piease examine your inventory immediately to determine if you have any of the products listed." lf so, please remove the affected products; quarantine them in a secure location." Piease complete and return the enclosed Field Recall Response Form, acknowleding your receipt and understanding of this communication " Send all affected products back to PULSION Medical Systems." Upon return of the affected product you will receive a credit note.For questions contact your PULSION Medical lnc Recall Coordinator. |  | 90 | 90 units | US Nationwide Distribution in the states of CT, NJ and NY. The affected product was distributed to one government consignee. | Terminated | October 02, 2015 | 332 | 90 | Not_Computer
Z-1124-2015 | 70087 | Yes | BTR | Tracheal tube. | Anesthesiology | Anesthesiology | 510(k) | RÂżSCH Curved Reinforced Endotracheal Tube: Rusch Reinforced Endotracheal Tube is designated for oral or nasal intubation and are indicated for airway management. The product may be used to reduce the potential for kinking whenever an unusual positioning of the head or neck is required following intubation. |  |  | December 12, 2014 | 2014 | Teleflex, Inc. | Teleflex Medical has issued an advisory notification for the RUSCH Curved Reinforced Endotracheal Tube because of a discrepancy between the labeled cuff diameter and the actual cuff diameter. The actual product cuff diameter meets required specification. | Labeling mix-ups | Teleflex Medical sent an Urgent Medical Device Notification dated December 12, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Please provide this Advisory Notice to all those who need to be aware of it within your organization. Please consider, clinicians, risk managers, supply chain/distribution centres, etc. in the circulation of this notice. There is no further action required.Clinical recommendation with regard to patient use is to continue the standard clinical practice of basing ET tube selection based on the tube size, carefully checking for the adequacy of ventilation and to change the ET tube size up or down as appropriate when intubating the patient.INSTRUCTION FOR DISTRIBUTORS OF AFFECTED PRODUCTIf you are a distributor, provide this advisory notice to all of your customers who have received product in scope of this action.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-246-6990. |  | 2292361 | 2,292,361 | Nationwide Distribution | Terminated | July 30, 2015 | 230 | 2292361 | Not_Computer
Z-1119-2015 | 70342 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis One; The Artis One is an angiography system developed for diagnostic interventional procedures including, but not limited to pediatric and obese patients. Procedures that can be performed with the Artis One include cardiac angiography, neuro-angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures as well as procedures next to the table for patient extremities. Additionally, angiographic procedures can be performed in the operating room, image guided surgery by x-ray, by image fusion, and by navigation systems. |  |  | January 16, 2015 | 2015 | Siemens Medical Solutions USA, Inc | The possibility exists that the monitor may fail and requires a power circle (shutdown and then power on) to resume operation. The problem is not systematic; but sporadic on single units. | Process control | Siemens notified their customer via Customer Safety Advisory Letter AX049/14/S, dated January 16, 2015, and has initiated a fix with Update Instruction AX048/14/S which will be completed by a Siemens service technician. Customers with questions should contact the Customer Care Center at 1-800-888-7436.For questions regarding this recall call 610-219-6300. |  | 1 | 1 | Nationwide Distribution to MI only | Terminated | April 21, 2015 | 95 | 1 | Not_Computer
Z-1122-2015 | 70438 | No | LPH | Hip joint metal/polymer/metal semi-constrained porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Mallory Head 4 Finned Acetabular Shell 50 mm, PN 13-104150, hip prosthesis component (Ringloc) |  |  | January 09, 2015 | 2015 | Biomet, Inc. | Biomet recalled the Mallory Head 4 Finned Acetabular Shell 50mm (Lot 364600) and Ringloc + Acetabular Shell Limited Hole Finned 52 mm (Lot 427840) following a mix up in packaging with the two. | Packaging process control | Biomet notified all domestic and foreign customers via FEDEx overnight on 1/9/2015. Response forms were provided, customers were instructed to notify hospital personnel if product was further distributed, and remove recalled products from circulation. Replacement parts were made available, customers may contact Audrey Daenzer, Field Action Specialist, Regulatory ComplianceBiomet, Inc., audrev.daenzer@biomet.com, or 574-372-1570 Monday- Friday 8 AM to 5 PM. |  | 17 | 17 | Distributed in the states of VA and TX and the country of Japan. | Terminated | October 22, 2015 | 286 | 17 | Not_Computer
Z-1125-2015 | 70363 | No | DXH | Telephone electrocardiograph transmitter and receiver. | Cardiovascular | Cardiovascular | 510(k) | Alive ECG App 2.1.2 ( a medical device application for the Apple iOS operating system, intended to be used with the AliveCor Heart Monitor. The AliveCor Heart Monitor is intended to record, store and transfer single-channel electrocardiogram (ECG) rhythms. The AliveCor Heart Monitor also displays ECG rhythms and detects the presence of atrial fibrillation and normal sinus rhythm (when prescribed or used under the care of a physician). |  |  | January 09, 2015 | 2015 | Alivecor SFO | Alive ECG App version 2.1.2 (intended to be used with the AliveCor Heart Monitor) crashed upon use of the application. | Software change control | AliveCor posted information on their website, Facebook Page and Twitter to alert users of the issue. |  | 5600 | 5600 active users with Alive ECG app for iOS . | Downloaded by Apple users - locations not shared by Apple. | Terminated | February 18, 2015 | 40 | 5600 | Computer
Z-1035-2015 | 70209 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | This correction applies to all GE Healthcare MRI systems with superconducting magnets: Discovery MR450, Discovery MR750, Signa HDx, Signa HDxt, Signa HDi, GE 1.5T Signa HDe, Optima MR360, Brivo MR355, Signa Excite-HD 1.5T &amp; 3T, Signa Excite 1.5T HD Twinspeed, Signa Excite 1.5T HD Echospeed, Signa Excite 1.5T HD Highspeed, 1.5T Signa Infinity TwinSpeed , 1.5T Signa Infinity EchoSpeed Plus, 1.5T Signa Infinity HiSpeed Plus, 1.0T Signa Infinity HiSpeed Plus. 1.0T Signa Infinity SmartSpeed, Signa EXCITE 3.0T, Signa EXCITE 3.0T HD, Signa Excite 1.5T TwinSpeed, Signa Excite 1.5T EchoSpeed, Signa Excite 1.5T HiSpeed, Signa Excite 1.5T SmartSpeed, Signa Excite 1.5T, Signa Excite 3.0T, Signa Contour/I, Signa OpenSpeed, Vectra, MR Max, Optima MR450w GEM, Discovery MR750W GEM, Optima MR450w. |  |  | January 07, 2015 | 2015 | GE Healthcare | At certain sites, the MRU may not be connected to the magnet. In emergency situations, a disconnected MRU could delay removal of a ferrous object from the magnet, potentially resulting in life-threatening injuries. The MRU must be connected to the magnet at all times. | Employee error | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref# 60876 dated January 6, 2015 to affected consignees. The letter was addressed to Hospital Administrators / Risk Managers, Radiology Department Managers, &amp; Radiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction &amp; Contact Information. Customers were instructed to do the following:As a preventative measure, confirm that MRU is connected to the magnet by performing the following four step test on the MRU.1. Verify the green CHARGER POWER LED is lit.2. Depress and hold the TEST BATTERY switch for 15 seconds. The green BATTERY TEST LED should light and remain lit while the TEST BATTERY switch is depressed.3. Place the TEST HEATER toggle switch in the A position. The green HEATER TEST LED should light. If it does not light, depress TEST HEATER LED switch to verify that the LED is functioning.4. Place the TEST HEATER toggle switch in the B position. Green HEATER TEST LED should light. If it does not light, depress TEST HEATER LED switch to verify that the LED is functioning.If the MRU test does not perform as described in each of the 4 steps above, GEHC strongly recommends that you stop using the system, and immediately call your GEHC representative.Customers with questions may contact their local service representative.For questions regarding this recall call 1-800-437-1171. |  | 12968 | 12,968 (5,708 US, 7,260 OUS). | Worldwide Distribution - US Nationwide in all states in continental USA including DC, PR, GU, and the countries of ALBANIA, ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BOLIVIA, BOSNIA &amp; HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR,K ESTONIA, FINLAND, FRANCE, FRENCH POLYNESIA, GEORGIA, GERMANY, GREECE, GUADELOUPE, GUATEMALA, GUYANA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG, MACEDONIA, MALAYSIA, MALTA, MARTINIQUE, MAURITIUS, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMAN, PAKISTAN, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, QATAR, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SLOVENIA, SOUTH AFRICA, SPAIN, SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TURKEY, UKRAINE , UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, UZBEKISTAN, VENEZUELA, VIETNAM, ZAMBIA. | Terminated | September 10, 2015 | 246 | 12968 | Not_Computer
Z-1061-2015 | 70238 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Echostar Spica 1.5T MRI system, for magnetic resonance imaging of the body. |  |  | January 09, 2015 | 2015 | Alltech Medical Systems America Inc | The MRI system safety feature, which lowers the magnetic field in emergency situations, may not work correctly. | Process control | Altech Medical Systems America Inc., notified customers on January 9, 2015, of the problem via telephone. Customers were instructed to immediately cease using the product until the correction can be made. A follow-up letter was sent through the mail, which stated that a service engineer would make a site visit to ramp down the magnet until the issue can be corrected. Customers were asked to complete the attached Acknowledgement and Receipt Form and fax it to (440) 424-2255. For questions contact AMSA at 1-855-200-2672. |  | 2 | 2 units | US Distribution in the state of Maryland. | Terminated | September 15, 2015 | 249 | 2 | Not_Computer
Z-1126-2015 | 70294 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | BodyGuard 323 pump, models |  |  | January 16, 2015 | 2015 | CME America, LLC | CME America is recalling the BodyGuard and BodyGuard 323 Infusion pumps due to the potential for an over delivery. | Software design | CME America sent an "Urgent Medical Device Recall" letter dated January 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Required Action for Users: To date there have been no reports of adverse events or serious injury. However, there is the potential for serious injury or death. CME America recommends that you DO NOT use the BodyGuard 323 pump for PCA infusion. If your devices are currently in use, be sure that the programmed parameters DO NOT involve a patient controlled bolus dose less than 0.1 miln volume (translated into meg or mg).Follow-up Actions by CME America: CME America has identified the root cause of this issue and is recommending installation of a new version of software, USA36B4, for affected BodyGuard 323 infusion pumps. This software is identical to the previously released version, USA36A4.2. However the designation of B4 in the software name will identify updated software in each device. CME America will contact all affected customers within 60 days to initiate the scheduling process for the remediation.Your assistance is appreciated and necessary to prevent use and facilitate remediation of the recalled product. CME America can address any questions regarding the recalled product by calling the CME America recall line at 720-839-2391. |  | 3186 | 3,186 | Worldwide Distribution - Nationwide Distribution and VA/military/govt consignees and the country of Canada. | Terminated | March 31, 2017 | 805 | 3186 | Computer
Z-1127-2015 | 70335 | No | HQF | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | MC-500 Multicolor Laser Photocoagulator; |  |  | March 27, 2012 | 2012 | Nidek Inc | The spot size control of the scan delivery unit of the MC-500 Multicolor Laser Photocoagulator may turn by itself. | Component design/selection | Nidek Japan issues a Engineering Change Order and Nidek reps visited affected sites to perform correction. |  | 16 | 16 units in commerce | Nationwide Distribution- including the states of CA, MI, NY, IL, NH, KS, NY, UT, IA, WI, TX, and FL. | Terminated | February 19, 2015 | 1059 | 16 | Not_Computer
Z-1128-2015 | 70374 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Restoris Partial Knee Application (PKA) RIO (TGS 2.0). |  |  | January 22, 2015 | 2015 | Mako Surgical Corporation | When using the MAKOplasty partial knee Arthroplasty application, the burr continues spinning outside of the stereotactic boundary and after the control switches (foot pedal and trigger) cease to be activated | Device Design | On January 22, 2015 the firm sent an IMPORTANT MEDICAL DEVICE CORRECTION asking consignees to complete the attached acknowledgement form and return them by email, fax or mail. The notice also noted the RIO system can continue to be used. The likelihood of occurrence is rare (.035% of cases) and the event is readily detectable and can be resolved. Should you have any questions regarding this Important Medical Device Recall Notice, to contact us at (954) 628-1721. |  | 228 | 228 units | Nationwide Distribution-including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, LA, MA, MD, MI, MO, MS, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, WA, WI, WV and Hawaii. | Open, Classified |  |  | 228 | Not_Computer
Z-1130-2015 | 70352 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | Haemonetics TEG Hemostasis System Level II Control, Catalog number: 8002. |  |  | January 26, 2015 | 2015 | Haemonetics Corporation | Specific lots of Level II Control Kits produce high Maximum Amplitude (MA) results, which are outside of the manufacturer's expected Quality Control range for MA parameter. | Process control | Haemonetics issued on 1/26/15, an Urgent Medical Device Recall to consignees The letter identified the affected product, the reason for the recall, as well as the risk to health. It instructs consignees to inspect their stock and determine if they have any of the affected lots. Customers are to immediately discontinue use and discard as instructed any affected product. The attached customer acknowledgment form should be completed and returned whether or not customers have any affected product. Customers are to contact their local service representative at 1-800-537-2802 if replacement product is needed. If any affected product has been forwarded to other locations, customers are to inform them of the recall. Please call the Technical Support team at 1-800-438-2834, with questions about this action. |  | 4143 | 4,143 kits | Worldwide Distribution -- US, Canada, Czech Republic, Denmark, Great Britain, and Sweden. | Terminated | June 05, 2015 | 130 | 4143 | Not_Computer
Z-1129-2015 | 70207 | No | ILO | Water circulating hot or cold pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | Cold Therapy Pad - cold water circulating pad, one unit per package. |  |  | January 07, 2015 | 2015 | Cincinnati Sub-Zero Products Inc | Seals in the water path of the CT-99 cold therapy pad may separate during use, which could lead to a water leak. | Process design | Cincinnati Sub-Zero sent an Urgent Medical Device Recall Field Action Notice dated January 7, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to immediately cease use and/or further distribution of the affected product. Customers were instructed to obtain a Return Material Authorization Number and return the pad(s) for exchange. Customers with questions were instructed to call 1-800-989-7373 or 513-772-8810. Customers were also asked to commplete and return the enclosed Response Form as soon as possible by fax to (513)772-9119 to acknowledge receipt of the notification and to inform CSZ that they have performed the requested actions.For questions regarding this recall call 513-772-8810, ext 6816 |  | 1390 | 1390 | Nationwide Distribution including FL, IL, TN, GA, VA, MO, NM, WI, MI, NC, IN, OH, MA, CA, WA, PA, KY, MD, HI, and AL. | Terminated | June 19, 2015 | 163 | 1390 | Not_Computer
Z-1137-2015 | 69378 | No | ITI | Powered wheelchair. | Physical Medicine | Physical Medicine | 510(k) | TDX Electric Wheelchair. Invacare FDX, TDX series and Storm Series (3G) power wheelchairs in 20 to 24 inch width configurations with power recline seating system and a conventional style back. |  |  | January 15, 2015 | 2015 | Invacare Corporation | If the slack in the wires is not routed and secured correctly, flexing of the back pan under the users weight may cause the headrest knob or clamp bracket to pinch, damage or cut the wire. Smoking, sparking, burning or fire may occur if a wire is repetitively pinched at the same place. | Device Design | The firm initiated the recall on 01/16/15 by sending a notification letter to customers through the mail. Communication was sent to providers to contact consignees via mail return receipt to inform them of the recall. Recall notification included a Provider letter, Consumer letter, Provider Response card, Unit Disposition form and tracking sheet and the Wiring guide. The notification instructed the provider to visually inspect all potentially impacted chairs. Not all potentially impacted chairs are expected to have improperly routed and secured wires. |  | 1942 | 1942 | Worldwide Distribution-US (nationwide) including DC and the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NH, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, and WV, and the countries of Canada and Australia. | Terminated | September 23, 2015 | 251 | 2527 | Not_Computer
Z-1135-2105 | 70401 | No | MEH | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | CoCr Femoral Head - indicated for hip arthroplasty for noncemented use in skeletally mature invididuals undergoing primary surgery for rehabilitating hips damaged as a result of noninflammatory degenerative joint disease |  |  | December 17, 2014 | 2014 | Zimmer Gmbh | Potential compromise of the inner sterile tray. If a compromised tray is detected, a delay in surgery may result. If undetected breach in both the inner and outer trays were to occur, and there is contamination present, a periprosthetic infection is possible, which could lead to revision. | Package design/selection | On 12/17/2014, URGENT MEDICAL DEVICE RECALL notifications were sent to the affected distributors and hospital Risk Managers with instructions for returning the affected product. All distributors were notified via electronic mail, and all hospital Risk Managers and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call the customer call center at 1-800-348-2759. |  | 6109 | 6109 | Domestic:AK, AL, AR, AZ, CA, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MS, MA, MD, ME, MI, MN, MO, NC, ND, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, WA, WI, WV, WYInternational: Canada, Japan | Terminated | September 23, 2015 | 280 | 18403 | Not_Computer
Z-1140-2015 | 70510 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Maquet CARDIOSAVE Hybrid Intra-Aortic Balloon Pump (IABP) |  |  | January 30, 2015 | 2015 | Maquet Datascope Corp - Cardiac Assist Division | Power supply malfunction complaints related to suboptimal thermal management. | Under Investigation by firm | Maquet Inc. sent a recall letter/return response form dated 1/30/2015. |  | 1300 | 1,300 units | US Nationwide distribution | Open, Classified |  |  | 1300 | Not_Computer
Z-1072-2015 | 69336 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | The Plum A+ Infusion System is designed to meet the fluid delivery requirements of today's evolving healthcare environments. It is a cassette based multi-function infusion system. The Plum A+ allows two lines in and one line out. The pump can be used for standard, piggyback, or concurrent delivery. Therapy modes include: Standard Infusions, Multistep Programming, Loading Dose, and Dose Calculation. The Plum A+ is designed to deliver parenteral, enteral, or epidural infusions over a broad range of infusion rates from multiple fluid container types. Each system includes a pumping module and an assortment of disposable IV sets, optional accessories, and an operator's manual. |  |  | July 02, 2014 | 2014 | Hospira Inc. | One lot of alarm assemblies used in Plum A+ and Plum A+3 infusion pumps may fail to sound at all volume levels. | Component change control | Hospira sent an Urgent Medical Device Correction letter dated July 2, 2014, to all affected customers. The letter notified customers that Hospira will be contacting them regarding the completion of an audible alarm test and will replace any alarm assemblies that fail to audible alarm test. Additionally, the letter included instructions for customers to: 1) inform potential users of this notification; 2) complete and return the attached reply form; and, 3) if the products were further distributed, notify those accounts of this recall and ask them to contact Stericycle at 888-912-7350 (Monday - Friday, 8:00 AM - 5:00 PM ET) to receive a reply form. A second URGENT MEDICAL DEVICE CORRECTION letter dated July 16, 2014 was sent to all customers of record because the initial letters may have contained errors in the attached data files listing the potentially affected devices at the customer's' facility. The instructions to the customer were unchanged from the original letter.For questions regarding this recall call 800-441-4100. |  | 3939 | 3,939 infusion pumps | Worldwide Distribution - USA including AK, AL, AR, CA, CO, CT, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico and Brazil, Canada, Colombia, Ecuador, France, Israel, Italy, Mexico, Oman, Peru, Saudi Arabia, Spain, Turkey, United Arab Emirates, United Kingdom, and Venezuela. | Open, Classified |  |  | 9109 | Not_Computer
Z-1141-2015 | 70467 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | EndoWrist One Vessel Sealer, IS3000, 6 pack used in conjunction with the da Vinci Surgical System Si IS3000 |  |  | February 11, 2015 | 2015 | Intuitive Surgical, Inc. | Intuitive Surgical has identified the potential for a certain, small portion of Vessel Sealers to exhibitinterference in the back of the instrument grips due to manufacturing variability in some jaws. | Process control | Intuitive Surgical sent an Urgent Medical Device Correction letter dated Feruary 11, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to take the following actions:1. Forward this letter to any members of your medical staff who perform da Vinci procedures in addition to your Risk Manager, OR Director, Purchasing and Biomedical Engineering staff.2. This field correction does not require you to quarantine or return the affected devices. Vessel Sealers can continue to be used safely by adhering to the following guidelines as provided in the Vessel Sealer User Manual:a. Only procedures that incorporate sealing and transection of thin tissue bundles and small diameter vessels are affected by this issue. Prior to the procedure, determine whether this issue may be a factor and choose whether or not to use the device per the following instructions. Alternative methods of sealing and transection are described below.b. During sealing, always inspect tissue for energy effect (i.e. blanching, steam, bubbling) along the entire jaw of the instrument. While audio tones designate seal cycle completion, tissue effect should always be seen prior to proceedingto cut.If no tissue effect is seen, then DO NOT CUT. Re-grasp tissue and re-attempt sealing as described above.If the issue persists, the instrument may be affected and an alternative means of sealing and transecting thin vascular tissue bundles should be used. These other methods include using another da Vinci Vessel Sealer, alternative da Vinci electrosurgical instruments, a da Vinci Harmonic instrument or hand-held laparoscopic sealing device.3. Complete the attached Acknowledgement Form and return it to Intuitive Surgical as instructed.4. Please retain a copy of this letter and the Acknowledgement Form for your files.5. If you choose to not continue to use affected Vessel Sealer instruments, you may retu |  | 3282 | 3282 total units, all part numbers | Worldwide Distribution - USA (nationwide) and Internationally to Australia, Belgium, China, Denmark, Ecuador, Finland, France, Germany, Greece, Italy, Netherlands, Norway, Puerto Rico, Saudi Arabia, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | June 01, 2015 | 110 | 3282 | Not_Computer
Z-1146-2015 | 70383 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Set Screw 5.5 mm Rod - Model #7703-1600. |  |  | November 30, 2012 | 2012 | Biomet Spine, LLC | Biomet Spine LLC is recalling the Set Screw 5.5 mm Rod due to the set screw being out of specification | Process control | Biomet Spine LLC (formerly Lanx) sent an Urgent Recall Notification letter to affected customers. The letter identified the affected product, problem, and actions to be taken. For questions contact Lanx Customer Service. |  | 1999 | 1999 | Worldwide Distribution - US Nationwide in the states of: AZ, CA, CO, CT, FL, GA, HI, IA, ID, IN, KY, LA, MA, MD, MI, MO, NC, NE, NY, OH, OR, PA, TN, TX, VA, WI. and the countries of: Italy, and Mexico. | Terminated | May 22, 2015 | 903 | 1999 | Not_Computer
Z-1143-2015 | 70388 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Starburst MRI Semi-Flex Electrode Device, 25cm/Attached Cable/14Ga/6.4F, Product No. H7877001039091, Catalog No. 700-103909; 25cm/14Ga/6.4F, Product No. H7877001026151, Catalog No. 700-102615; and 25cm/14Ga/6.4F, Product No. H7877001026151, Catalog No. 700-102615 |  |  | December 19, 2014 | 2014 | Angiodynamics | Printed mark on the trocar intended to indicate a 5cm distance from the distal tip of the needle is actually located 4cm from the distal tip of the needle. | Other | AngioDynamics sent an Urgent Medical Device Recall - Immediate Action Required letter dated December 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken:Immediately remove the recall product from your inventory (whether in labs, Central Supply, Shipping and Receiving or ANY other location).Segregate this product in a secure location for return to AngioDynamics, Inc.Immediately forward a copy of this recall notification to all sites to which you have distributed affected product.If affected product is located in your institution, please call AngioDynamics, Inc. Customer Service at 1-800-772-6446 between 8:00a.m. and 7:00p.m. (Monday- Friday: Eastern Standard Time) to obtain a replacement or credit for your returned product.Promptly complete, sign and return the enclosed Reply Verification Tracking Form (even if you do not have any product to return) following the directions on this page and the Reply Verification Form. |  | 43 | 43 | Worldwide Distribution - US Distribution to the states of AZ, AR, FL, MI, MN, MO, NY, TX, VA, WA and WI., and to the countries of Hungary and Spain. | Terminated | November 27, 2015 | 343 | 43 | Not_Computer
Z-1208-2015 | 70297 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Comprehensive Reverse Shoulder Glenosphere Mini Baseplate 25mm with Taper Adaptor, |  |  | January 07, 2015 | 2015 | Biomet, Inc. | The taper adaptor is missing from the packaging, which may result in a delay in surgery greater than 30 minutes. | Process control | Biomet issued written notices on 1/7/2015, Requested completion and submission of response/effectiveness forms for Product Part Number: 01 0000589 Product Description: Comprehensive Reverse Shoulder Glenosphere MiniBaseplate 25mm with Taper Adapter, Expiration October 2024. Questions related to this notice should be directed to Audrey Daenzer, Field Action Specialist (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 22 | 11- 6 consumed, 5 with distributors | Domestic: AL, FL, NY, TXForeign: None | Terminated | October 22, 2015 | 288 | 22 | Not_Computer
Z-1145-2015 | 70453 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs Healthcare qube Compact Monitor, Model 91390. |  |  | January 28, 2015 | 2015 | Spacelabs Healthcare Inc | Spacelabs Healthcare qube Bedside Monitors, Model 91390, are recalled because the firm has received multiple reports of qube monitors failing to boot up or returning to factory default configuration settings following power on or reset. | Device Design | Spacelabs began to send the Urgent Medical Device Correction XPREZZON Bedside Monitor, Model 91393 qube Compact Monitor, Model 91390 letter, dated 28, January 2015, to US consignees.Spacelabs began sending the Urgent Medical Device Correction XPREZZON Bedside Monitor, Model 91393 qube Compact Monitor, Model 91390 letter (translated as necessary) to international customers via email on 2/4/15.Actions to be Implemented by Users:-Do not rely solely on the monitor for critically ill patients.-If customer facility operates any XPREZZON or qube monitors on a standalone basis, rather than via central monitoring, please contact Spacelabs technical support as these monitors will be the first repaired.Spacelabs Healthcare will contact consignees to schedule a convenient time for Spacelabs to replace the PCBAs in all of your facility's affected XPREZZON and qube monitors at no cost.If you have any questions about this corrective action, please contact Spacelabs at the toll free number of 1-800-522-7025 and select 2 for Technical Support, or via direct dial at 1-425-363-5203. |  | 2955 | total 2955 units (2117 units in US and 838 units outside US) | Worldwide Distribution. | Terminated | December 30, 2015 | 336 | 2955 | Not_Computer
Z-1131-2015 | 70340 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HeartWare Controller) |  |  | January 16, 2015 | 2015 | HeartWare Inc | The affected clinical trial Controllers exhibit a higher susceptibility to ESD than newer commercial Controllers. | Device Design | HeartWare sent an URGENT MEDICAL DEVICE RECALL dated January 16, 2015 asking clinicians to quarantine the affected controllers and complete and return the acknowledgement form. HeartWare also asks clinicians to identify patients with affected controllers and review the applicable risks with the patient and as soon as reasonably possible if medically advisable, exchange the affected controller under medical supervision with a new controller. HeartWare also asks that for each patient using an affected controller as their backup to contact the patient and arrange to have the backup controller replaced with a new controller. Lastly, HeartWare asks the firm to complete the patient tracking form and return all quarantined controllers to HeartWare along with the completion form. Should you have any questions or concerns, please contact your local HeartWare representative. Our 24-Hour Clinical Support personnel are also available at 1-888-494-6365. |  | 4845 | 4,845 | Worldwide Distribution - US Nationwide in the states of: AL, AZ, CA, CO, FL, GA, IL, IN, KY, LA, MA, MD, MI, MN, MO, NC, NY, OH, OR, PA, TX, UT, VA, WA and District of Columbia and countries of: Canada, Poland, Australia, Austria, Italy, DE, New Zealand, South Africa, France, Switzerland, Germany, Brazil, UK, Belgium, Netherlands, Israel, Malaysia, Turkey, Greece, Japan, Singapore, Poland, Lithuania, Sweden, Norway, Finland, and India. | Open, Classified |  |  | 4845 | Not_Computer
Z-1212-2015 | 70545 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | COMPASS SW Version 3.1, Catalog Number CS10-100, medical linear accelerator, radiological |  |  | February 02, 2015 | 2015 | Iba Dosimetry Gmbh | Error in the software. During internal tests of the current development version of the Compass SW it was found that dose reconstruction for DMLC plans when the jaws move or mlc leaves reverse during beam on will not be scaled correctly for all control points. | Software design (manufacturing process) | IBA sent an Field Safety Notice dated January 28, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken:Please do not use the dose reconstruction function COMPASS 3.1 and 3.1a for DMLC plans from Monaco or Eclipse -11 when the jaws move or MLC leaves reverse during beam on. A correct Version ( COMPASS 3.1b) will be delivered as soon as it is available. The field safety notice including a confirmation of receipt was sent to the affected hospitals; by email and UPS letter on February 2, 2015.Should you have questions or require additional information, please contact the Customer Service Team at South and Middle America, USA and Canada - Phone. +1 901.386.2246Asia Pacific, Australia and New Zealand - Phone. +86 10 8080 9107All other countries - Phone. 49 9128 607 - 38 |  | 124 | 124 units | Worldwide Distribution - US Distribution to the states of AZ and OK., and the countries of Germany, Austria, France, UK, Poland, Finland, The Netherlands, Spain, Turkey, Russia, and South Africa, Algeria, North and South America, Mexico, Brazil, Venezuela, Colombia, Panama, Asia, China, Japan, Philippines, Korea, India, Singapore, Thailand, Taiwan, Hong Kong, Malaysia and Guam. | Terminated | August 18, 2015 | 197 | 124 | Computer
Z-1200-2015 | 70322 | No | OEZ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | Enforcement Discretion | NAMIC Convenience Kit, 3V WMMII KIT, UPN H965601852521, REF/Catalog No. 60185252, STERILE, Rx ONLY --- For single use only. --- Device Listing Number D136588. |  |  | January 09, 2015 | 2015 | Navilyst Medical, Inc | The labelled expiration dating of the affected product is not fully supported by shelf life verification testing, which may result in a sterility breach. Use of a Fluid Management Convenience Kit, where sterility has been breached, may result in an increased risk of infection. | Packaging | Navilyst Medical sent an Urgent Medical Device Recall Notification letter dated January 9, 2015 to US customers (end users) via Federal Express. On January 20, 2015, Recall Notification Packages were sent via Federal Express to 2 International consignees (distributors). the letter identified the affected product, problem and actions to be taken. Customers are instructed to: " Immediately remove the recall product from your inventory (whether in labs, Central Supply, Shipping and Receiving or ANY other location)." Segregate this product in a secure location for return to Navilyst Medical." Immediately forward a copy of this recall notification to all sites to which you have distributed affected product." Complete and return the Reply Verification Tracking Form. |  | 7 | 7 units | Worldwide Distribution - US Nationwide and the countries of: Canada, and Guam. | Open, Classified |  |  | 16508 | Not_Computer
Z-1215-2015 | 70291 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | The Brainlab Offset Cup Impactor Universal. Part of the Brainlab hip system. An intraoperative image-guided localization system to enable minimally invasive orthopedic surgery. (a stereotaxic instrument) |  |  | January 16, 2015 | 2015 | Brainlab AG | The recommended sterilization and drying parameters are not effective to achieve the minimum required sterilization level. | Reprocessing Controls | FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letters dated December 11, 2014 were sent to all direct accounts (customers). The letters included instructions for customers to utilize the sterilization cycles included in the letter. Customers were further advised to alert users of these devices to the sterilization process changes. Customers with questions can contact Brainlab by telephone at 800-597-5911 or via e-mail at us.support@brainlab.com. |  | 114 | 114 offset cup impactors | Worldwide Distribution - US: Nationwide (AZ, CA, ID, IL, MD, MI, MN, OH, OK, OR, TN, WA, WI, WV); Australia, Austria, Bahrain, Belgium, Germany, Hong Kong, India, Italy, Japan, Malaysia, Russian Federation, Saudi Arabia, South Korea, Spain, Switzerland, Thailand, United Arab Emirates, United Kingdom | Open, Classified |  |  | 114 | Not_Computer
Z-1214-2015 | 64784 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | The Centricity PACS Workstation is intended for use as a primary diagnostic and analysis tool for diagnostic images by trained healthcare professionals, including radiologists, physicians, technologists, clinicians and nurses. It is also intended for use as a clinical review workstation throughout the healthcare facility. The workstation interface provides the user with a means to display, manipulate, archive, print and export images when connected with the Centricity PACS infrastructure. |  |  | February 18, 2013 | 2013 | GE Healthcare | Using Merge Exam in single Study Mode may result in Missing Study Record (Cannot Display Exam, Send Exam, etc). Exam merge in Single Study Mode may fail on &amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp;amp | Software design | Consignees were notified of the issue via telephone and site visit by GE Healthcare IT Field Engineers. |  | 128 | 128 systems | Distributed US (nationwide) including the states of AK, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, ME, MN, MO, MS, NC, NE, NJ, NY, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA, WI, and WV, and the country of Bermuda. | Open, Classified |  |  | 128 | Not_Computer
Z-1213-2015 | 70456 | toothbrush, manual | EFW | Manual toothbrush. | Dental | Dental | 510(K) Exempt | Manual Toothbrush; SilverCare Junior 2/6 years Antibacterial Toothbrush Made in Italy |  |  | February 03, 2015 | 2015 | Norwex Usa Inc | The toothbrush head can potentially disengage from the handle. | Under Investigation by firm | Norwex sent an Urgent Product Recall letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with questions were instructed to call 1-866-450-7499. For questions regarding this recall call 204-647-0798. |  | 35214 | 35,214 | Nationwide Distribution | Open, Classified |  |  | 35214 | Not_Computer
Z-1044-2015 | 69870 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Multi Diagnost Eleva II |  |  | October 14, 2013 | 2013 | Philips Medical Systems, Inc. | The possibility of the Reference Air Kerma Rate for low dose mode will not be reduced to 50% of the value for the normal mode, but will be approximately 70-80%, leading to a higher patient dose than expected. | Radiation Control for Health and Safety Act | Philips Healthcare Planned Action: 1. Software updates to SW R6.1.3 through Field Change Orders (FCOs) FCO70800140 and FCO70800142.2. A customer notification letter which includes a statement that Philips Healthcare will without charge, remedy the defect or bring the product into compliance.1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. Philips Healthcare will implement this CAP by September 1, 2015.For further questions please call (978) 687-1501. |  | 238 | 238 | Worldwide Distribution - USA ( nationwide ) to the states of : AR, AZ, CA, CO, DE, FL, GA, HI, IL, IA, KY, KS, LA, MI, MO, MS, NC, NY, OH, PA, TN, TX, VA, VT, WA, WY and WV., and the countries of Austria, Australia, Colombia, Germany, Switzerland, Turkey, Ireland, Dubia and United Kingdom. | Open, Classified |  |  | 238 | Computer
Z-1224-2015 | 70224 | No | CBN | Gas-scavenging apparatus. | Anesthesiology | Anesthesiology | 510(k) | Servo-I ventilator system CO2 analyzer. Part number 65 23 588. |  |  | January 12, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | A calibration problem was detected in a specific number of the Servo-I ventilator system CO2 analyzer models that may have been caused by the module's internal circuit board function. | Device Design | Maquet sent an Urgent Product Recall Medical Device Action Field Removal Letter dated 1/12/2015 via FedEx. |  | 486 | 486 units | Worldwide distribution - USA nationwide, Austria, Australia, Belgium, Chile, Canada, Colombia, Ecuador, Egypt, Finland, France, Germany, Hungary, Iceland, India, Indonesia, Iraq, Italy, Japan, Libya, Netherlands, Norway, Panama, Peru, Portugal, Romania, Russia, Saudi Arabia, South Africa, Spain, Sweden, Thailand, Tunisia, Turkey, Ukraine, and United Kingdom. | Open, Classified |  |  | 486 | Not_Computer
Z-1219-2015 | 70500 | No | LZO | Hip joint metal/ceramic/polymer semi-constrained cemented or nonporous uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Zimmer Segmental System (ZSS) Cemented Stem / ZSS Cemented Stem, Smooth Stem |  |  | February 05, 2015 | 2015 | Zimmer, Inc. | PMMA coating does not meet specifications. Area of the stem may be missing PMMA coating. | Process control | Zimmer sent an URGENT MEDICAL DEVICE RECALL - Lot Specific letter dated February 5, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Your Responsibilities1. Review the notification and ensure affected personnel are aware of the contents.2. Locate and remove from inventory the affected products identified above.3. Return any affected product with missing PMMA coating per the PER process4. Return the Notification Acknowledgment Form (Attachment 1) to corporatequality.postmarket@zimmer.com.5. Please notify Zimmer if the hospital that you have distributed the affected product to has implanted the product. Inaddition, identify the surgeons that have used this product.6. If after reviewing this notification you have further questions or concerns please call the customer call center at1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 1 | 1 unit | US Distribution to the state of TX | Terminated | September 23, 2015 | 230 | 1 | Not_Computer
Z-1216-2015 | 69970 | No | IKZ | Powered patient rotation bed. | Physical Medicine | Physical Medicine | 510(K) Exempt | RotoRest Delta Advanced Therapy System. |  |  | December 03, 2014 | 2014 | Arjo Hospital Equipment AB | The recalled devices labeling and instructions for use contain unapproved medical claims. | Labeling design | Devices in the rental fleet will be delivered to the customer with revised documents including the updated information. The customers who purchased the devices will be notified of the recall via letter sent by registered mail. The field correction notice will include a recall letter, response form, and updated labeling. |  | 271 | 271 units | Worldwide Distribution. US nationwide, Germany, Austria, France, Italy, India, EU, Mexico, Switzerland, Canada, Saudi Arabia, Kuwait, Middle East, China, Japan, and Qatar. | Terminated | October 30, 2015 | 331 | 1064 | Not_Computer
Z-1223-2015 | 70395 | Yes | MRN | Nitric oxide administration apparatus. | Anesthesiology | Anesthesiology | 510(k) | Ikaria, INOmax DSIR (Delivery System), Model 10007. Nitric oxide delivery system for use with ventilators. |  |  | January 14, 2015 | 2015 | INO Therapeutics (dba Ikaria) | Potential delivery failure alarm condition. INOmax DSIR with software version 2.0.4 and a certain variant of the monitor display may trigger a delivery failure alarm when the display brightness is set to its lowest level. If this condition occurs, therapy will be interrupted. | Process change control | Consignees were sent on 1/14/2015 an Ikaria "Urgent Field Correction Action" letter dated January 14, 2015. The letter described the problem and the product involved in the recall. The letter provided a number of recommendations to the consginees in regard to use and return of the product. For questions or adverse events they can contact Ikaria Customer Care at 1-877-KNOW-INO (1-877-566-9466) 24 hours a day, 7 days a week. |  | 45 | 45 units | Distribution US nationwide, including AZ, CA, CT, FL, ID, LA, MD, MI, MS, NH, NJ, NV, NY, OK, PA, TX, UT, VA, WV. | Terminated | April 15, 2015 | 91 | 45 | Computer
Z-1226-2015 | 70441 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Vacuum Source Tubing Set for the EndoWrist Stabilizer for da Vinci S (IS2000) and Si (IS3000); |  |  | February 16, 2015 | 2015 | Intuitive Surgical, Inc. | There is the possibility that the Vacuum Source Tubing set for the Endo Wrist Stabilizer may have the outer pouch compromised during shipping, resulting in small pinholes which compromise the sterility of the pouches outer surface. | Under Investigation by firm | Urgent Medical Device Recall letters were sent on February 16, 2015 |  | 706 | 706 units | US and Belgium, Brazil, Canada, China, Finland, France, Greece, India, Italy, Japan, Qatar, Russia, Saudi Arabia,South Korea, Spain, Switzerland, Taiwan, Turkey, United Kingdom | Terminated | July 01, 2015 | 135 | 706 | Not_Computer
Z-1227-2015 | 70128 | dispenser, liquid medication | KYX | Liquid medication dispenser. | General Hospital | General Hospital | 510(K) Exempt | Baxter, Self-Righting Syringe Tip Caps, Self-Righting DISCPAC Syringe Tip Caps. Product codes H93866100, H93869100, H93866025, H93867025, H93867100, H938671025, H938671100, H938673025, H938673100, H938674025, H938674100, H938676025, H938676100, H938677025, H938677100, H938679025, H938679100, H938690025, H93869025, H938693025, H938696025, H93868025, H93868100, H938672025, H938672100, H938675025, H938675100, H938678025, and H938678100. |  |  | December 29, 2014 | 2014 | Baxter Corporation Englewood | Self-Righting Luer Tip Caps were manufactured and shipped with a missing latex symbol on both single and case labels. | Labeling Change Control | Baxter sent an Safety Alert letter dated January 15, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Baxter is requesting that you take the following actions:1.Complete the enclosed customer reply form and return it to Baxter by faxing it to 224-270-5457 or scanning and e-mailing it to fca@baxter.com. Returning the customer reply form promptly will prevent you from receiving repeat notices.2. Please forward a copy of this Safety Alert letter to other facilities or departments within your institutions to ensure that those locations are aware of this action.Action to be taken if you are an end-user and purchased product from a distributor or reseller:1. Please forward a copy of this Safety Alert letter to other facilities or departments within your institutions to ensure that those locations are aware of this action.2. Follow your supplier's reply process. Please do not return the customer reply form to Baxter.Action to be taken if you are a distributor:If you are a dealer, wholesaler, or distributor/reseller that distributed any product to other facilities, please notify your customers of this Safety Alert in accordance with your customary procedures. Further information and supportFor general questions regarding this communication, contact The Center for One Baxter at 800-422-9837, Monday through Friday, between the hours of 8:00 AM and 5:00 PM Central Time. |  | 1404230 | 1,404,230 | Nationwide Distribution | Open, Classified |  |  | 1404230 | Not_Computer
Z-1222-2015 | 70397 | No | BSY | Tracheobronchial suction catheter. | Anesthesiology | Anesthesiology | 510(K) Exempt | KimVent Wet Pak, Closed Suction System for Adults, Double Swivel Elbow, 14 Fr. Sterile, Distributed by Kimberly Clark Global Systems, LLC. |  |  | January 16, 2015 | 2015 | Halyard Health | Defect: A hole or crack was found which may cause air to leak. | Packaging process control | Consignees were notified by letter via FedEx on 01/16/2015. |  | 83 | 83 cases (20 tubes per case) | CA, FL, IL, LA, MA, MK, MN, MS, NM, NY, OK and TX and Japan. | Terminated | July 07, 2015 | 172 | 93 | Not_Computer
Z-1147-2015 | 70271 | No | LGW |  |  | Neurology | PMA | The Medtronic, Quadripolar ITREL 3 IPG, Model 7425, is a multiprogrammable device designed for use in the Stimulation of the Central Nervous System, Peripheral Nervous System, or muscles when connected to an extension and/or lead. |  |  | January 09, 2015 | 2015 | Medtronic Neuromodulation | Current labeling indicates that cycling improves device longevity in every programming scenario. However, enabling cycling at certain parameter settings may decrease the device battery longevity. With cycling enabled, longevity estimate displayed on the Model 8840Programmer may be inaccurate. | Other | Consignees were sent a Medtronic {Medical Device Correction" letter dated January 9, 2015. The letter was addressed to Healthcare Professional. The letter described the issue and the product involved in the the recall. The letter also listed the Occurrence of the problem and labeling update. A Physician Reply Form was included which requested consignees to return the form to Medtronics. For questions they can contact Medtronic Neuromodulation Technical Services at 1-8007070944 weekdays 7 am 6 pm Central time. A second "Physician Notification Confirmation Form" was included to be completed by Medtronic Representatives. |  | 396975 | 396,975 (258,805 USA, 138,170 OUS) | Worldwide Distributon- USA( nationwide including DC and PR) and Internationally to Mexico and Canada | Open, Classified |  |  | 396975 | Not_Computer
Z-1232-2015 | 70471 | No | JSK | Supplement for culture media. | Microbiology | Microbiology | 510(K) Exempt | Oxoid Legionella BCYE Growth Supplement SR0110A packaged in packs of 10- 100 ml vials. The responsible firm name on the label is Oxoid Ltd., Basingstoke, Hampshire, England and Oxoid Ltd., Basingstoke, Hants, England. |  |  | February 05, 2015 | 2015 | Remel Inc | Use of the product may not provide adequate recovery and could result in false negative identification of Legionella pneumophila. | Under Investigation by firm | ThermoFisher Scientific sent an Urgent Medical Device Recall letter dated February 5, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The letter informed the consignees of the recall and the associated risk of false negative identification of Legionella pneumophila. The consignees were instructed to notify any personnel who need to be aware of the potential for false negative id. of Legionella spp. Further, the consignees were instructed to discard any affected product remaining in their inventory and review all results reported using the recalled lot. The recalling firm requested the consignees complete and return the Medical Device Recall Return Response form acknowledging receipt of the notice and disposal of the affected product. If the affected lot was further distributed, the consignees were instructed to notify those entities of the recall and provide a copy of the recall notification letter.If you have any questions, please contact our Technical Services Department at 800.255.6730 (US) or 913.888.0939 ( International) |  | 24 | 24 packs (10 x 100 ml vials) | US Distribution to the states of : MI, ND, SC, TN, and TX. | Open, Classified |  |  | 24 | Not_Computer
Z-1120-2015 | 70011 | No | IZI | Angiographic x-ray system. | Radiology | Radiology | 510(k) | Philips Medical System Allura Xper X-Ray Angiographic |  |  | June 06, 2014 | 2014 | Philips Medical Systems, Inc. | In certain circumstances, a software error can lead to a situation where the five minute fluoroscopy audible signal does not sound. | Radiation Control for Health and Safety Act | Philips Medical Planned Action: 1. The affected customers will be sent a customer notification letter that identifies the software corrective actions and indicates that these corrective actions will be provided to the end-user free of charge2. A mandatory Field Change Order with reference FCO72200270 will be released that requires Philips field service engineers to install Software release R7.2.8 which addresses the buzzer issue. The expected date of this FCO will be February 2015. The software release will contain updated timer function code to correct the intermittent fluoroscopy timer malfunction.3. Philips Healthcare Field Service Engineers will schedule an appointment with customers to install the software update.4. The software solutions will be rolled out consecutively, each within a maximum completion time of 6 months. The software releases will be corrected in descending order of the number of affected systems in the field, with approximately 80% of the total installed base corrected by 4 2015. |  | 7439 | 7439 | Worldwide Distribution - USA nationwide including Puerto Rico, Austria, Azerbaijan, Australia, Algeria, Argentina, Albania, Brazil, Belgium, Bangladesh, Bahamas, Bulgaria, Cambodia, Chile, Canada, China, Croatia, Czech Republic, Netherlands, New Zealand, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Reunion, Singapore, Slovenia, South Africa, Spain, Sudan, Sweden, Saudi Arabia, Switzerland, Syria, Taiwan, Thailand, France, Dominican Republic, Colombia, Turkey, Germany, Greece, Hong Kong, Denmark, Ecuador, Egypt, Estonia, Finland, Vietnam, Russian Federation, Iran, Iraq, Ireland, Israel, Italy, India, Indonesia, Japan, Jordan, Kenya, South Korea, Kuwait, Latvia, Lebanon, Lithuania, Macedonia, Malaysia, Mauritius, Mexico, Moldavia, Mongolia, Morocco, Nepal, Ukraine, United Arab Emirates, and United Kingdom | Open, Classified |  |  | 7439 | Computer
Z-1237-2015 | 70434 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Monaco radiation therapy treatment planning system. |  |  | December 19, 2014 | 2014 | Elekta, Inc. | Incorrect CT to ED when using the Monaco Image Statistics Tool. | Other | Elekta sent an Important User Notice on December 19, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.This notice provided a workaround until new software releases. The notice included a acknowledgement form to be returned.For further questions please call (770) 300-9725 |  | 1354 | 1354 | Worldwide Distribution - USA (nationwide) Distribution including Puerto Rico and to the states of : AL, AK, AZ, CA, CO, CT, FL, ID, IL, IN, KS, KY, LA, MD, MI, MN, MO, MT, NH, NJ, NM, NY, NC, ND, OH, OK, OR, PA, SC, SD, TN, TX, VA, WA and WI., and to the countries of Angola, Argentina, Austria, Australia, Bangladesh, Belgium, Bulgaria, Bahrain, Brazil, Canada, Switzerland, Chile, China, Colombia, Cyprus, Czech Republic, Germany, Algeria, Ecuador, Spain, Finland, France, Great Britain, Greece, Hong Kong, Ireland, Israel, India, Iraq, Italy, Jordan, Japan, South Korea, Kazakhstan, Sri Lanka, Lithuania, Libya, Morocco, Myanmar, Malta, Mexico, Malaysia, Namibia, Nicaragua, Netherlands, New Zealand, Peru, Philippines, Poland, Portugal, Romania, Russia, Saudi Arabia, Sweden, Singapore, Slovenia, Thailand, Turkey, Taiwan, Venezuela and South Africa. | Open, Classified |  |  | 1354 | Computer
Z-1209-2015 | 70506 | Yes | OAD |  |  | Cardiovascular | PMA | IntellaTip MiFi Open-Irrigated radiofrequency Ablation Catheters: |  |  | February 03, 2015 | 2015 | Boston Scientific Corp | Reports of formation of char adherent to the proximal part of the distal tip electrodes | Device Design | Customer notification letters were sent by overnight mail on February 10, 2015. |  | 14 | 14 units | Distributed in Germany, Great Britain, and the Netherlands. | Terminated | March 21, 2017 | 777 | 14 | Not_Computer
Z-1210-2015 | 69939 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Xoran MiniCAT X-ray imaging device |  |  | September 08, 2014 | 2014 | Xoran Technologies, LLC | Out of tolerance kVp, Unexposed (black) frames, and Out of tolerance Dose. | Radiation Control for Health and Safety Act | Xoran Planned Action :A Xoran Technologies service representative will re-measure and perform calibration to correct all affected systems for kVp and dose issues. For the image quality issue, the service representative will mark the black frame and then attempt to reconstruct an adequate image. CDRH also understands that you are continuing your investigation to determine the root cause and a more permanent corrective action for the 3 identified defects.1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. A customer notification letter which includes a statement that Xoran Technologies, LLC will without charge, remedy the defect or bring the product into compliance.3. Xoran Technologies, LLC will continue to evaluate why the kVp is not always within Xoran specifications, why the unexposed frames continues to occur, and why the dose is not always within Xoran specifications and identify the actions needed to resolve these issues by July 1, 2015. Xoran will communicate the results of this evaluation and any necessary actions to both customers and CDRH.Further questions please call (734) 663-7194. |  | 432 | 432 | Nationwide Distribution | Terminated | February 02, 2016 | 512 | 432 | Not_Computer
Z-1233-2015 | 70436 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | ARROW FiberOptix Intra-Aortic Balloon Catheter Kit, 40 cc 7.5 Fr |  |  | February 05, 2015 | 2015 | Arrow International Inc | Outer package product label incorrectly identifies the product as FiberOptix Flex IAB 7.5Fr rather than the correct FiberOptix IAB 8Fr | Labeling mix-ups | Arrow International notified Users, Distributors, Risk Managers by letter dated February 5, 2015. The letter explains the labeling error and requests immediately discontinue use and quarantine product. Customers were instructed to complete the Recall Acknowledgement Form and return by fax to 1-855-419-8507 whether they have or do not have the affected product. Customers with questions were instructed to contact their local sales representataive or Customers Service at 1-866-246-6990.For questions regarding this recall call 610-378-0131. |  | 165 | 165 units | Worldwide Distribution - USA (nationwide) and Internationally to Belgium | Open, Classified |  |  | 165 | Not_Computer
Z-1228-2015 | 70424 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | CareLink Pro |  |  | November 10, 2014 | 2014 | Medtronic MiniMed Inc. | The firm is informing customers of a software error that might result in an inaccurate display of the Temp Basal time and duration in CareLink Clinical reports with data uploaded from the MiniMed 620G or 640G insulin pump. | Software design | Medtronic MiniMed Inc. sent an Urgent Medical Device Correction letter in December 2014 to all affected customers. The firm states in their letter that patients can continue to use their MiniMed 620G and 640G insulin pump systems, and that there is no risk to patients caused by this issue. The firm instructs that when reviewing the CareLink Pro 4.0 or CareLink Clinicial reports from the MiniMed 620G and 640G insulin pumps not to use the Temp Basal graphic for clinical consideration.For questions regarding this recall call 818-567-4700. |  | 9 | 5 CareLink Pro 4.0 CD ROMs | Worldwide Distribution to Japan only.CareLink Clinical is for clinical trials only (US and International) | Terminated | March 10, 2015 | 120 | 9 | Computer
Z-1229-2015 | 70470 | No | JLS | Progesterone test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Centaur Systems Progesterone Kit, (1 -pack) |  |  | January 27, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc | ADVIA Centaur Systems Progesterone Kit Lots Ending in 268 have a high bias on results greater than 30 ng/mL(95.4 nmol/L) | Process control | Siemens issued an Urgent Medical Device Correction (UMDC) was sent to affected Siemens Healthcare Diagnostics customers in the United States on January 27, 2015 to be delivered to customers on January 28, 2015. An Urgent Field Safety Notice (U FSN) was sent out for distribution to all customers outside the United States on January 27, 2015. These notices inform the customer of the potential to observe a high bias on patient sample with results greater than 30 ng/ml when using ADVIA Centaur Systems Progesterone kit lots ending in 268. To mitigate the bias, Siemens instructs customers to dilute samples with values greater than 30 ng/ml when using ADVIA Centaur Systems Progesterone kit lots ending in 268. In addition, Siemens informs customers to continue to follow Urgent Medical Device Correction 10819675, Rev. A and Urgent Field Safety notice 10819674, Rev A (Information Regarding the ADVIA Centaur Systems Calibrator E), dated September 2014.Fax-back forms to be returned to confirm that the customers received theUrgent Medical Device Correction. |  | 9778 | 9778 kits | Worldwide Distribution-US (nationwide), Austria, Bahrain, Belgium, Bulgaria, Brazil, Canada, Chile, Colombia, Croatia, Cyprus, Czech, Republic, Denmark, Egypt, Estonia, France, Georgia, Germany, Great Britain, Greece, Hong Kong, Hungary, India, Iran, Iraq, Ireland, Italy, Japan, Jordan, Kazakhstan, Latvia, Lebanon, Libya, Lithuania, Luxembourg, Macedonia, Mexico, Morocco, Netherlands, Norway, Oman, Pakistan, Peru, Poland, Portugal, Qatar, Romania, Russian Federation, Serbia, Slovakia, Spain, South America, Sweden, Switzerland, Taiwan, Turkey, United Arab Emirates, and Vatican. | Terminated | March 31, 2016 | 429 | 23477 | Not_Computer
Z-1234-2015 | 70466 | No | KGX | Medical adhesive tape and adhesive bandage. | General Hospital | General &amp; Plastic Surgery | 510(K) Exempt | Zip 8i Surgical Skin Closure Device; |  |  | February 05, 2015 | 2015 | Zipline Medical | The manufacturers seal may not be present on the Nylon pouch resulting in lack of assurance for sterility. | Packaging process control | Zipline sent an Urgent Voluntary product Recall letters to customers on February 5, 2015 by e-mail with possible follow up on site by firm representatives. The letter identified the affected product, problem and actions to be taken. Customers are asked to quarantine inventory immediately and examine the affected lots for complete pouch seals. Any product found with an incomplete seal should be destroyed or returned to ZipLine Medical. Please complete the enclosed response form as soon as possible, even if you no longer have stock of the lot numbers subject to this correction. This information is required to reconcile the inventory. Please contact Customer Service at 1-888-326-8999 to coordinate any product returns, replacement stock, or credit notes. |  | 3109 | 3109 - total - all model numbers | Worldwide Distribution - US Nationwide and the countries of Panama, Tehran, Ireland, Sweden, Japan, Hong Kong, Finland, Pakistan, Taiwan, Malaysia. | Terminated | April 10, 2015 | 64 | 3109 | Not_Computer
Z-1240-2015 | 70220 | No | KMI | Bed-patient monitor. | General Hospital | General Hospital | 510(K) Exempt | The I-Cart is used to store, transport, and power video equipment used in medical procedures. |  |  | December 10, 2014 | 2014 | Richard Wolf Medical Instruments Corp. | The transformers installed on carts used for Richard Wolf medical video systems may overheat causing smoke and black plastic material to leak from the cart chassis. | Device Design | Field Action Notice letters dated December 10, 2014 were sent to both direct accounts (customers). The letters instructed customers to: 1) Instruct the biomed to bypass the integrated power grid on the cart and not to use the electrical system on the affected carts until a qualified Omni Corporation service technician is able to inspect and repair the carts; and 2) Omni Corporation will schedule a qualified Omni field service technician, in conjunction with Richard Wolf and the customer's biomed unit, to visit and replace the transformer on the cart. |  | 83 | 83 carts | Nationwide | Terminated | July 31, 2017 | 964 | 83 | Not_Computer
Z-1241-2015 | 70338 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | LFT v40 Femoral Head; Catalog Number 6260-9-032; |  |  | February 03, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Three customer complaints were received for the same batch of V40 LFIT Vitallium femoral heads (part number 6260-9-032, lot 48681201) alleging that the femoral head couldn't be assembled with its corresponding V40 stem trunnion. | Under Investigation by firm | Branches/Agencies were notified of this action by e-mail on December 19, 2014 and were asked to quarantine affected devices. Notification Letters and Product Accountability Forms were sent to Hospital Risk Management and Chief of Orthopaedics via FedEx (with return receipt) by 2/3/2015. |  | 9 | 9 units | Distributed in the states of GA, MA, NC, NY &amp; PA. | Open, Classified |  |  | 9 | Not_Computer
Z-1242-2015 | 70590 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | LASSO NAV Duo Loop eco Catheters, Catalog No. D-1354-01, D-1354-02, D-1354-03, D-1354-04, D-1354-05, D-1354-06. |  |  | February 25, 2015 | 2015 | Biosense Webster, Inc. | Biosense Webster is recalling the LASSO NAV Duo Loop eco Catheters because of reports of difficulty retracting and damage to the spine cover. | Device Design | Biosense Webster sent an Urgent Voluntary Medical Device Recall letter dated February 24, 2015 , to all affected customers. The letter informed customers that Biosense Webster is recalling the LASSO NAV Duo Loop eco Catheters because of reports of difficulty retracting and damage to the spine cover. The letter informed the customers of the problems identified and the actions to be taken. Customers with questions were instructed to contact Biosense Webster sales representatives or call (866) 473-7823, Monday through Friday from 7am to 8pm EST. Customers with questions related to the Voluntary Field Removal Certification Form and its return, are instructed to call (909) 839-8468.For questions regarding this recall call 909-839-8500. |  | 464 | 464 units | Nationwide Distribution | Terminated | October 16, 2015 | 233 | 464 | Not_Computer
Z-1239-2015 | 70502 | No | HHW | Vaginal pessary. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | CooperSurgical PESSARY KIT-CUBE#2, 1-3/8 in , 35 mm |  |  | February 09, 2015 | 2015 | CooperSurgical, Inc. | Incorrect size printed on the Milex Pessary Kit packaging label | Error in labeling | CooperSurgical sent an Recall Letter dated February 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter asked customers to contact CooperSurgical Product Surveillance Department if they had any questions.At CooperSurgical's expense, arrangements were made to replace any of the affected products that customers had on hand. CooperSurgical provided further instruction on the disposition of the product at the time of replacement.Questions contact CooperSurgical 203.601.5200. |  | 95 | 95 units | Nationwide Distribution | Terminated | January 15, 2016 | 340 | 95 | Not_Computer
Z-1243-2015 | 70464 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | DynaLOC Breast Interventional Planning Software Versions 3.1 and 3.2 |  |  | January 21, 2015 | 2015 | iCAD, Inc. | A device malfunction may cause the biopsy needle to fail to reach the intended target. If location confirming radiographs are not taken and properly evaluated, biopsy site mislocation can result in non-target tissue being sampled rather than target tissue. May result in an incorrect diagnosis. | Device Design | iCAD notified Invivo of the recall via letter on January 21, 2015. Invivo sent a Customer Information letter dated January 23, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letters from both Invivo and iCAD (device manufacturer) recommend that DynaCAD Breast Interventional Software not be used with Invivo ds Breast 16 channel lateral grid plate part number 4535-303-31491. Invivo and iCAD are developing DynaCAD Breast Interventional 3.3 that corrects this error and will be available for installation beginning in March 2015.If you need any further information or support concerning this issue, please contact your local Invivo representative:1-877-Invivo1 (1-877-468-4861, option #3)Invivo apologizes for any inconveniences caused by this problem. |  | 153 | 153 devices | Worldwide Distribution - Nationwide Distribution to the states of : AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, NC, NE, NH, NJ, NY, OR, PA, SC, TN, TX, VA, VT, WA, WI, WV and W., and to the countries of : BELGIUM, GERMANY, GREAT BRITAIN, INDIA, ISRAEL, NETHERLANDS and SWITZERLAND. | Open, Classified |  |  | 153 | Computer
Z-1244-2015 | 70433 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Focal Sim radiation therapy treatment planning system |  |  | December 19, 2014 | 2014 | Elekta, Inc. | Incorrect CT to ED when using the Monaco Image Statistics Tool. | Software design | An Important User Notice (IUN 382-03-FCL-001) was issued to customers on 12/19/2014. The notice included a workaround to be used until a new software release. The notice also included an acknowledgement form to be returned. |  | 907 | 907 | Worldwide distribution. US nationwide (AL, AK, CA, CO, CT, FL, IL, IN, LA, MD, MA, MI, MO, MT, NJ, NY, NC, OH, OK, OR, PA, SC, TN, TX, VA, WA, WY, DC), Puerto Rico, Argentina, Austria, Australia, Bulgaria, Bahrain, Brazil, Belarus, Canada, Chile, China, Colombia, Cuba, Cyprus, Czech Republic, Germany, Algeria, Egypt, Spain, France, Great Britain, Greece, Croatia, Indonesia, Israel, India, Iraq, Italy, Jordan, Japan, South Korea, Sri Lanka, Lithuania, Libya, Morocco, Macedonia, Mexico, Malaysia, New Zealand, Philippines, Portugal, Romania, Serbia, Russia, Singapore, Slovenia, Suriname, Thailand, Turkey, Ukraine, Venezuela, South Africa, Kosovo. | Terminated | August 07, 2017 | 962 | 907 | Computer
Z-1245-2015 | 70194 | Yes | FIK | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | AMIA v1.0 APD System. For Automated Peritoneal Dialysis (APD) therapy for pediatric and adult renal patients. |  |  | April 03, 2014 | 2014 | Baxter Healthcare Corp. | Additional warning in the AMIA APD System labeling for patients vulnerable to unintended Increased Intraperitoneal Volume (IIPV), also known as overfill. The patients include neonates, infants, small children, and any patient with significant heart and/or lung disease. | Error in labeling | Baxter sent IMPORTANT PRODUCT INFORMATION letters dated April 3, 2014 to all peritoneal dialysis healthcare providers of record. The letters advised healthcare providers of the potential to develop Increased Intraperitoneal Volume (IIPV) and instructed the healthcare providers to complete and return the attached Customer Reply Form. Healthcare providers with clinical questions can contact Baxter's Renal Division Clinical Helpline at 888-736-2543, option 2, 8:00 AM - 5:00 PM CT. Healthcare providers with general questions can contact The Center for One Baxter by telephone at 800-422-9837, Monday - Friday, 8:00 AM - 5:00 PM CT. Additionally, an IMPORTANT PRODUCT INFORMATION letters dated April 8, 2014 were sent to all peritoneal dialysis patients of record. The letters advised patients of the potential to develop Increased Intraperitoneal Volume (IIPV) and instructed the patients to complete and return the attached Home Patient Reply Form. Patients with questions about the letter can contact The Center for One Baxter by telephone at 800-422-9837, Monday - Friday, 8:00 AM - 5:00 PM CT. |  | 55 | 55 systems | US: Nationwide in the states of (GA, MA, NC, NY, PA, WA) | Terminated | July 21, 2017 | 1205 | 55 | Not_Computer
Z-1246-2015 | 70507 | No | MBB | Polymethylmethacrylate (PMMA) bone cement. | Orthopedic | Orthopedic | 510(k) | SMARTSET GHV Gentamicin Bone Cement. A self-curing, radiopaque, polymethyl methacrylate based cement containing antibiotic. It is used for securing a metal or polymeric prosthesis to living bone in arthroplasty procedures. The bone cement has no intrinsic adhesive properties but relies instead on close mechanical interlock between the irregular bone surface and the prosthesis. |  |  | February 11, 2015 | 2015 | DePuy Orthopaedics, Inc. | One lot of SMARTSET GHV Gentamicin Bone Cement is partially agglomerated . This may result in surgical delays while a replacement package of cement is retrieved and mixed. | Under Investigation by firm | DePuy Orthopaedics, Inc. sent an Urgent Information - Recall Notice dated February 6, 2015, to all affected consignees for a voluntary recall of one lot of SMARTSETÂż GHV Gentamicin Bone Cement because the affected cement is partially agglomerated (clumped together). Consignees were instructed to discontinue use of the affected lot of bone cement immediately and destroy according to the consignees hazardous materials protocol or return if Certified Dangerous Goods Shipper, holding an International Air Transport Association (IATA) certification. Customers were instructed to call 574-371-4917 (M-F; 8 am  5 pm EST).For questions regarding this recall call 574-371-4577. |  | 1267 | 1267 | Nationwide Distribution including CA, CO, CT, FL, GA, IL, IN, KS, LA, MA, MD, MI, MS, MT, NC, NV, NY, OH, OR, PA, TN, TX, UT, VA, WA, and WV. | Terminated | May 19, 2015 | 97 | 1267 | Not_Computer
Z-1249-2015 | 70477 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | EVOLUTION MP TIBIAL BASE, REF ETPKN8SR, SIZE 8 RIGHT, TYPE PRIMARY, FIT CEMENTED, STYLE KEELED, STERILE R, Rx ONLY, MicroPort Orthopedics, Inc. Knee prosthesis component. |  |  | February 02, 2015 | 2015 | Microport Orthopedics INC. | Some units of the EVOLUTION Tibial Base was affected by a casting tool issue which caused varying retaining wall thickness and keel offset in the finished product. | Device Design | MicroPort Orthopedics sent an Urgent Action Required letter dated January 29, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to check their stock including all consignment stock to determine if they have any affected product. If they have affected product they should return immediately to the MicroPort Orthopedics Inc. Distribution Center at 11481 Gulf Stream, Arlington, TN 38002. Mark all return shipping boxes with "RECALL" for identification and processing. Customers with questions were instructed to contact MPO's Customer Experience Department at 1-866-872-0211. For questions regarding this recall call 901-867-4324. |  | 17 | 17 units | Worldwide Distribution - USA including AL, AZ, ID, KS, KS, MO, TX, IA, WV and Internationally to Canada and The Netherlands. | Terminated | March 19, 2015 | 45 | 17 | Not_Computer
Z-1252-2015 | 70203 | No | MQM | Bilirubin (total and unbound) in the neonate test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | RAPIDPoint 405 Blood Gas Analyzer; Siemens Material Number: 10310464, 10322347, 10314817, 10317193, 10318999, 10320055, 10322347, 10328278, 10328302, 10336784. |  |  | December 30, 2014 | 2014 | Siemens Healthcare Diagnostics Inc | Neonatal bilirubin (nBili) parameter may have increased variability when the nBili concentration is greater than 12 mg/dl and the tHb concentration exceeds the upper reportable range of greater than 25 g/dl. | Software design | An Urgent Field Safety Notice was sent by e-mail to all affected Siemens Healthcare Diagnostics Regional offices on December 30, 2014 both in the United States and elsewhere for communication with affected customers. In addition, all affected domestic customers were sent a hard copy of the Urgent Field Safety Notice via Federal Express on December 30, 2014. All affected customers outside the US will be provided a copy of an Urgent FieldSafety Notice via hard copy, e-mail, and/or fax as determined by each countrieslocal regulations and procedures. This notice informs consignees of the issue. Consignees are also asked to complete a field effectiveness check form and return it Siemens via fax (312) 275-7795. |  | 3420 | 3420 devices | Worldwide Distribution - US nationwide including Puerto Rico, Canada, Mexico, Albania, Algeria, Argentina, Austria, Bahrain, Belgium, Bosnia Herzegovina, Botswana, Brazil, Bulgaria, Chile, Colombia, Croatia, Czech Republic, Denmark, Egypt, Estonia, Finland, France, Georgia, Germany, Greece, Hungary, Iceland, Iraq, Ireland, Israel, Italy, Ivory Coast, Kazakhstan, Kuwait, Latvia, Lebanon, Lesotho, Lithuania, Luxembourg, Macedonia, Netherlands, New Caledonia, Norway, Peru, Poland, Portugal, Qatar, Romania, Russian Fed., Saudi Arabia, Serbia, Slovakia, Slovenia, South Africa, Spain, Sweden, Switzerland, Turkey, U.A.E., United Kingdom, Vatican City, Venezuela, White Russia/Belarus. | Terminated | August 02, 2017 | 946 | 10900 | Not_Computer
Z-1248-2015 | 70514 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Proteus XR/a with wall stand model number 2260354. |  |  | January 30, 2015 | 2015 | GE Healthcare | Potential breakage of the steel cable which supports the bucky device in certain wall stands of Proteus XR/a X-ray imaging systems. A fall of a bucky while the system is in use could result in an injury to a patient or operator. | Device Design | Consignees were sent on 1/30/2015 a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref#17108 dated January 30, 2015. The letter was addressed to Hospital Administrators/Risk Managers, Managers of Radiology/Cardiology, and Radiologists/Cardiologists.The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact 800-437-1171. Outside the US contact your local GE Healthcare Service Representative. |  | 192 | 192 (131 US, 61 OUS). | Worldwide Distribution-US (nationwide) including the states of AK, AZ, CA, CO, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MD, MI, MO, NC, NE, NH, NJ, NM, NV, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI,and WV, and the countries of ARGENTINA, CANADA, CHILE, CZECH REPUBLIC, FRANCE, GERMANY, GREECE, HONG KONG, HUNGARY, IRELAND, ITALY, JAPAN, NEW ZEALAND, PORTUGAL, RUSSIA, SAUDI ARABIA, SLOVAKIA, SPAIN, and UNITED KINGDOM. | Terminated | September 15, 2015 | 228 | 192 | Not_Computer
Z-1079-2015 | 70241 | No | KXJ | Radiologic table. | Radiology | Radiology | 510(K) Exempt | Juno DFR X-ray system |  |  | December 14, 2014 | 2014 | Villa Radiology Systems LLC | The system does not provide the appropriate audible signal, permanent activation, and manual override, although the system is in high-level control functionality | Radiation Control for Health and Safety Act | Planned Action by the firm: 1. You will send out a customer notification letter that identified the software corrective actions and that indicates that these corrective actions will be provided to the end-user free of charge.2. You will implement a corrective software patch so that kV and mA manual adjustments from the generator console are inhibited, and consequently no high-level control modality is allowed. In AERC, high-level control is already disabled. A Field Service Engineer is to implement the software patch according to the corrective action plan1. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFRThe system shall be modified upgrading the generator power firmware which will inhibit the high-level control bringing the system into compliance.This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.2. You will implement the corrective software patch to inhibit high-level control according to the corrective action plan by June 22, 2015.For further questions please 203.262-8836 |  | 30 | 30 units distributed in the US | US Distribution to the states of: AZ, AL, CA, CO, CT, DC, FL, IA, IL, IN, MI, MD, MS, MA, MO, NY, OH, PA, UT, TN and TX. | Terminated | March 16, 2017 | 823 | 30 | Computer
Z-1251-2015 | 70432 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips DuraDiagnost X-Ray system, Stationary X-Ray System Radiographic |  |  | January 23, 2015 | 2015 | Philips Medical Systems, Inc. | Improper installation of Tube arm, and Wall stand may cause it to fall down, Improper installation of table may cause it to float. | Process control | Philips Healthcare issued notice to foreign Distributors on 1/23/15 to issue and advise customers to stop using the device until a Philips Engineer checks the device. After the engineer verifies that the tube arm, Wall stand and table were properly assembled, the device may be used again.Philips will subsequently install a new Tube Arm and Wall stand fixing structure that is designed to prevent this hazard if the tube arm or Wall stand is incorrectly reinstalled in the future.Philips will subsequently provide the method to adjust the T-nuts of table to make sure it was properly assembled. A Philips Service Engineer will contact affected customers to install the solution. |  | 71 | 71 units | Worldwide Distribution. Argentina, Australia, Burkina-Faso, Chile, China, Colombia, Ecuador, France, Germany, Hungary, Indonesia, Iraq (NL), Kuwait, Libya, Macedonia, Malaysia, Mayotte, Poland, Saudi Arabia, Slovenia, Spain, Switzerland, Turkey, United Kingdom, and Uzbekistan. | Open, Classified |  |  | 71 | Not_Computer
Z-1247-2015 | 68235 | mattress, flotation therapy, non-powered | IKY | Nonpowered flotation therapy mattress. | General Hospital | General Hospital | 510(K) Exempt | Pressure IQ Evolve Mattress; Non-powered mattress replacement system. |  |  | September 01, 2014 | 2014 | Arjo Hospital Equipment AB | The firm received complaints regarding multiple mattresses were showing failures resulting in pressure ulcers. | Device Design | On 9/1/14 the recalling firm sent out a letter to their customers notifying them of an updated IFU for the recalled device. The letter also requested the customers to verify the quantity of the product they currently have on hand. |  | 1034 | 1,034 units | Nationwide Distribution. | Terminated | January 04, 2016 | 490 | 1034 | Not_Computer
Z-1258-2015 | 70287 | No | JKA | Blood specimen collection device. | Clinical Chemistry | Clinical Chemistry | 510(k) | Siemens Healthcare Caps for capillary 50/60 uL, Package count 200 |  |  | January 22, 2015 | 2015 | Siemens Healthcare Diagnostics Inc | Capillary Caps for RAPIDLyte Multicap and Multicap-S Blood Collection Capillaries may not attach and result in breakage of the glass capillary | Nonconforming Material/Component | Siemens Healthcare issued an Urgent Field Safety Notice via e-mail to affected Siemens Healthcare Diagnostics Regional offices on January 22, 2015 both in the United States and elsewhere for communication with affected customers. Siemens advises customers to use caution whenapplying the Capillary Caps with Siemens Multicap glass and Multicap-S plastic capillaries to minimize the potential for breakage or bending. |  | 6184 | 6184 | Worldwide Distribution. US Nationwide, Australia, Austria, Belgium, Bosnia, Herzegovina, Canada, Colombia, Croatia, Curacao, Czech Republic, Estonia, Finland, France, Georgia, Germany, Great Britain, Hong Kong, Ireland, Italy, Japan, Jordan, Kazakhstan, Kyrgyzstan, Latvia, Lithuania, Mexico, Netherlands, New Zealand, Norway, Poland, Portugal, Russian Federation, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand, United Arab Emirates. | Terminated | April 13, 2016 | 447 | 49309 | Not_Computer
Z-1250-2015 | 70473 | No | ILY | Infrared lamp. | Physical Medicine | Physical Medicine | 510(K) Exempt | Apollo (cold) Laser Desktop Control Units, Model AP2-DT. |  |  | December 17, 2014 | 2014 | Pivotal Health Solutions | Control units were equipped with an internal mounting kit that does not meet medical safety standards, and are conductive, increasing the risk of electric shock to the user and patient. These units were manufactured prior to Pivotal Health Solutions acquisition of the Apollo product line. | Device Design | The firm, Pivotal Health Solutions, sent a letter dated December 17, 2014 on 12/17/2014 and an amended Pivotal "Urgent Medical Device Recall" letter dated February 9, 2015 to its consignees/customers. The letters described the product, problem and actions to be taken. The consignees/customers were instructed to stop using the unit and contact Pivotal Health Solution's at 1-800-743-7738 to arrange for return and repair asap; immediately examine your device inventory and quarantine any product subject to recall; if you have further distributed the product, identify your customers and notify them at once of this product recall, and complete and return the enclosed RETURN AUTHORIZATION FORM with the units to Pivotal Service Center, 1654 Mardon Drive, Dayton, OH 45432 and the DECLARATION OF CONTAMINATION STATUS form via Fax to: 605-882-8398.If you have any questions, contact Pivotal Health Solutions Service Repair Coordinator at 800-743-7738 or email Robin@PivotalHealthSolusitons.com. |  | 35 | 35 | US Distribution to states of: AZ, AR, CA, GA, IL, IN, IA, KS, MI, MN, MO, NY, OH, OR, PA, UT and WA. | Terminated | March 09, 2016 | 448 | 35 | Not_Computer
Z-1259-2015 | 70603 | No | DSP | Intra-aortic balloon and control system. | Cardiovascular | Cardiovascular | 510(k) | Arrow Intra-Aortic Balloon Catheter Kit Fiberoptix (IAB), |  |  | November 08, 2013 | 2013 | Arrow International Inc | Potential for a leak between the catheter and the bifurcate which would allow helium to escape to the atmosphere. | Process control | Arrow International Inc notified accounts by letter dated 11/8/13 describing the problem and requested users to immediately discontinue use and return product. A Recall Acknowledgement form was to be completed and returned. |  | 166 | 166 units | Distribution US nationwide (including AZ, NJ, ME, TX, WA), and Germany. | Terminated | March 16, 2015 | 493 | 166 | Not_Computer
Z-1260-2015 | 70580 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Synthes Inserter for Titanium Elastic Nails; The Inserter for Titanium Elastic Nails (TEN) is intended for fixation of diaphyseal fractures where the canal is narrow or flexibility of the implant is paramount. This includes: lower extremity fractures in pediatric and small-statured patients and upper extremity fractures in patients. |  |  | February 03, 2015 | 2015 | Synthes, Inc. | The Inserter has the potential for mechanical failures such as breakage of the main shaft of the Inserter for the TEN, malfunction of the chuck of the Inserter for the TEN in the form of the chuck getting stuck/jammed, breakage of the cross bar of the Inserter for the TEN. | Device Design | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated February 3, 2015, to all affected customers to alert them about the recall and explain the benefits and risks associated with removing this device from the market. It stated that the immediate removal of the product would prevent the performance of emergency surgery, thus DePuy Synthes is not requiring an immediate removal of affected product(s). DePuy Synthes is in the process of developing the replacement &amp; recovery plan for indicated devices. Due to their current but temporary lack of replacement inventory, Synthes proposed to remove and replace the device in increments to prevent any adverse impact on the availability of the device for use on patients in trauma-related surgery events.DePuy Synthes will contact customers as replacements become available.Customers were advised to complete the Verification Section by checking the appropriate box indicating affected product has been located. Also,indicate the number of devices found. If your facility chooses to return the product(s) at this time, please take the following additional steps:Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number, include a copy of Verification Section with the product to: Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132.Send the completed Verification Section to the following fax or email address: Fax: (866) 782-7287 or Scan/email: Synthes5095@stericycle.com Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales Consultant.For questions regarding this recall call 810-719-5000. |  | 1550 | 1550 | Nationwide Distribution | Open, Classified |  |  | 1550 | Not_Computer
Z-1264-2015 | 70562 | stain, papanicolau | HZJ | Dye and chemical solution stains. | Hematology | Pathology | 510(K) Exempt | Carton: Richard-Allan Scientific CYTO-STAIN; REF 7501R; Quantity: 1x1 Gallon; Size: 1 Gallon/3.8L; LOT 313833; 11/2016; Bottle: Richard-Allan Scientific CYTO-STAIN G; REF 7501G; 1 Gallon/3.8L; Thermo Scientific. |  |  | February 19, 2015 | 2015 | Richard-Allan Scientific Company | The case was correctly labeled as REF 7501R, however the gallon bottle was incorrectly labeled as REF 7501G Cyto-Stain G. The bottle label has the correct lot number of 313833 and EXP date 11/2016. The gallon bottle contains the correct intended contents of 7501R: Cyto-Stain R Stain. | Error in labeling | Urgent Medical Device Recall letters sent to Richard Allan distributors and customers. The recall notification includes the reason for the voluntary recall, risk to health, product information and actions to be taken by the distributor/customer. |  | 132 | 132 | US, Canada and Europe | Terminated | January 12, 2016 | 327 | 132 | Not_Computer
Z-1263-2015 | 70264 | [6 empty] | VITROS Hand-held Barcode Scanner, Catalog Number 6844210, for use with the following VITROS Systems: VITROS 5,1 FS Chemistry System, VITROS 3600 Immunodiagnostics System, VITROS 4600 Chemistry System, VITROS 5600 Integrated System. This product is an accessory for use with VITROS Analyzer Systems. |  |  | December 29, 2014 | 2014 | Ortho-Clinical Diagnostics | In combination with the Batch Programming option (only) for programming samples, the hand-held barcode scanner unexpectedly skips sample cup positions within a sample tray. This issue only occurs when using the Batch Programming option; all other sample programming options function as intended. | Software design | On 12/29/2014, Customer letter (Ref. CL2014-315) was sent via FedEx overnight courier or ORTHO PLUS e-Communications and/or US Postal Service Priority Mail (for PO Boxes only) to US consignees (OCD Direct, Drop-Ship and US Federal Government) to inform them of the issue and provided with instructions for reconfiguring their Barcode Scanners. |  | 72 | Domestic: 62 units; Foreign: 10 units | Worldwide Distribution-US (nationwide) and the countries of Canada, Australia and England. | Terminated | March 27, 2017 | 819 | 72 | Not_Computer
Z-1262-2015 | 70544 | No | FMF | Piston syringe. | General Hospital | General Hospital | 510(k) | - NTL-30 30ml Nutrio Enteral Feeding Syringe (sterile) |  |  | February 13, 2015 | 2015 | Medela Inc. EF Division | The firm is recalling sterile enteral feeding syringes due to determining that the sterility of the affected products may be compromised as a result of a potential breach in packaging. | Packaging | Medela sent a Device Recall letter dated February 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Requested Action of Distributors:1. Stop sale and distribution of product with affected lot numbers.2. Conduct a physical count and record the data on the Product Inventory Form included with the letter.3. Return the Product Inventory Form via fax to 800-995-7867 or e-mail it to syringe@medela.com. The firm should contain a contact name and signature.4. If have the affected product in stock, Medela Customer Service will issue a RMA and schedule a pick-up of the affected product. Contact will be made to ensure product is ready. For further questions please call (888) 399-2455. |  | 1959187 | 1,959,187 | Nationwide Distribution. | Terminated | November 09, 2015 | 269 | 1959187 | Not_Computer
Z-1266-2015 | 70523 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Persona Trabecular Metal Tibial Plate / Persona TM Tibia-Prosthesis, Knee, Patello/Femorotibial, Semi-Constrained, Uncemented, Porous, Coated, Polymer/Metal/Polymer |  |  | January 28, 2015 | 2015 | Zimmer, Inc. | An increase in complaints of loosening and radiolucent lines. | Under Investigation by firm | Zimmer Inc. is initiating a voluntary recall of Persona Trabecular Metal Tibial following an increase in complaints of radiolucent lines and loosening. All sizes and lots of the affected devices are being removed from distribution. Urgent Medical Device Recall notices were issued to affected distributors, hospitals, and surgeons on 2/16/2015 via mail. Customers are asked to review the notification and ensure affected personnel are aware of the contents. All affected product are to be located and quarantined immediately. The Inventory Return Certification Form should be completed and returned along with the recalled product. Customers may contact the following with any questions: 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 11658 | 11,658 | Worldwide Distribution -- US, including the states of AK, AL, AZ, CA, CO, FL, IL, IN, KS, MA, MI, MN, MO, NC, NJ, NY, OH, OK, PA, TN, TX, UT, VA, WA, and WI; Foreign:Canada, Australia, New Zealand, Korea, Austria, Belgium, Germany, Italy, Luxembourg, Netherlands, South Africa, Sweden, Switzerland, and United Arab Emirates. DOD/VA: HOSPITAL- VA MAINE HLTCR SYTM (402)VA MROC (402) ATTN: PROSTHETICS DEPT TOGUS ME 04330 207-672-9425HOSPITALVA MED CTR (438) 2501 W 22 SIOUX FALLS, SD 57105 605-336-8922HOSPITALVA MED CTR (509)1 FREEDOM WAY AUGUSTA , GA 30904 706-721-6892HOSPITALVA MED CTR (521) 700 S 19TH ST, BIRMINGHAM, AL 35233 378-962-1599HOSPITALVA MED CTR (523) 1400 VFW PARKWAY, WEST ROXBURY, MA 02132 617-862-5977HOSPITALVA MED CTR (528) 3495 BAILEY AVENUE BUFFALO, NY 14215 716-898-6922HOSPITALVA MED CTR (552) 4100 WEST 3RD STREET BLDG # 330 RECEIVING, DAYTON, OH 45428 937-226-5987HOSPITALVA MED CTR (583), 1481 WEST 10TH STREET INDIANAPOLIS, IN46202 317-267-8765HOSPITALVA MED CTR (595) 1700 SOUTH LINCOLN AVENUE, LEBANON, PA17042 717-272-6621HOSPITALVA MED CTR (657) 915 N GRAND BLVD, ST LOUIS, MO63106 314-652-4100HOSPITALVA MED CTR (664) 350 LA JOLLA VILLAGE DR SAN DIEGO, CA92161 858-552-8585HOSPITALVA MED CTR (693) 1111 EAST END BOULEVARD, WILKES BARRE, PA 18711 717-824-3521HOSPITALVA SNHCS (654/118) 975 KIRMAN AVENUE RENO NV89502 702-597-1699 | Open, Classified |  |  | 11658 | Not_Computer
Z-1267-2015 | 70556 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SIEMENS SOMATOM Force with software version VA50A and /or VA50A_FP1 and/or VA50A_SP0; Computed tomography system. |  |  | February 06, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Possibility of image artifacts during data acquisition when using Adaptive Cardio Sequence, Turbo Flash, and Head modes. The update improves system start-up behavior and resuming, ECG handling with visual feedback of correct placement of ECG leads and contact quality, and other safety related issues. | Software design | Siemens sent a Customer Advisory Notice dated February 6, 2015, to all customers via USPS Certified Mail. The notice identified the affected product, problem and actions to be taken. |  | 12 | 12 | US Nationwide Distribution in the states of PA, MA, MN, NY, SC, MD, NY, IL, OH, TX and KY. | Terminated | August 31, 2015 | 206 | 12 | Not_Computer
Z-1265-2015 | 70524 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | MAX-TL 9.0 x 7.0 mm Implant, Catalog No. 15711K; Z-MAX Implant 9.0 x 7 mm Length, Catalog No. Z-MAX9-7. Dental implant. |  |  | January 23, 2015 | 2015 | Southern Implants, Inc | Z-MAX Implant, 9mm diameter, 7mm length labeled package, catalog number Z-MAX9-7, contained a MAX-TL Implant, 9mm diameter, 7mm length. And vice versa. | Mixed-up of materials/components | A customer notification letter dated 1/23/15 was sent to customers to inform them that Southern Implant is recalling the MAX-TL 9.0x7.0mm Implant because it may contain a Z-MAX9-7 Implant, 9mm diameter, 7mm length. Customers are instructed to return the fax back form and to contact customer relations depart at (866) 902-9272. |  | 130 | 130 units | US Nationwide Distribution. | Terminated | May 20, 2015 | 117 | 130 | Not_Computer
Z-1278-2015 | 70270 | applicator, absorbent tipped, sterile | KXG | Absorbent tipped applicator. | General Hospital | General Hospital | 510(K) Exempt | Puritan absorbent foam tipped applicator; Catalog Number 25-1406 1 PF BT. |  |  | December 22, 2014 | 2014 | Puritan Medical Products Co., LLC | Shaft breaks so that adequate patient specimen sample cannot be obtained. | Process control | Consignees were initially notified of the recall via phone on 12/23/2014 and 12/29/2014. They were informed that the product may have been over exposed to gamma radiation, resulting in brittle sticks. They were instructed to not use the product for patient sampling and to return any product remaining in inventory to Puritan. |  | 4500 | 9 cases of 500 applicators each | Distributed in the states of MS, ME, GA, and NC. | Terminated | March 18, 2015 | 86 | 4500 | Not_Computer
Z-1269-2015 | 70512 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | MiniCap with Povidone-Iodine Solution, item number 5C4466P, for hemodialysis. |  |  | January 22, 2015 | 2015 | Baxter Healthcare Corp. | Product may have separating or protruding sponges | Under Investigation by firm | Baxter sent an Urgent Product Recall letter dated January 22, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Baxter is requesting that you take the following actions:1. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product package or shipping carton.2. Contact Baxter Healthcare Center for Service to arrange for return and credit. Center for Service can be reached at 888-229-0001, Monday through Friday, between the hours of 7:00 AM and 6:00 PM Central Time. Please have your Baxter eight-digit ship-to account number ready when calling.3. If you received your product directly from Baxter, please complete the enclosed customer reply form, and return it to Baxter by either fax or scanned e-mail.1. If you ordered product from a dealer, wholesaler, or distributor/reseller, please follow your supplier's reply and recall process.2. Locate and remove all affected product from your facility. The product code and lot number can be found on the individual product package or shipping carton.3. Contact Baxter Healthcare Center for Service to arrange for return and credit.Center for Service can be reached at 888-229-0001, Monday through Friday, between the hours of 7:00 AM and 6:00 PM Central Time. Please have your ship-to account number ready when calling.4. Please do not return the Baxter customer reply form to Baxter. Reply forms should be returned to your supplier.For clinical questions, contact Baxter's Renal Clinical Helpline at888-736-2543, option 2, Monday through Friday, between the hours of8:00 AM and 5:00 PM Central Time.For general questions regarding this communication, contact The Center for One Baxter at 800-422-9837, Monday through Friday, between the hours of 8:00AM and 5:00 PM Central Time. |  | 4391520 | 4,391,520 units | Worldwide Distribution - Nationwide Distribution and to the countries of : Canada, Austria, Norway, Portugal, Saudi Arabia and Spain. | Open, Classified |  |  | 4391520 | Not_Computer
Z-1275-2015 | 70430 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | BD Posiflush Saline Flush Syringe 10 ml inlcuded in the following device kits: Groshong Midline Bard MYPICC Kit 4F (Single-Lumen) |  |  | August 08, 2014 | 2014 | Bard Access Systems | Some unit packages of BD PosiFlush SF Saline Flush Syringe 10 mL may exhibit open seals, which could impact product sterility. Bard includes these pouched syringes, attached to some configurations of PICC kits, Port Access kits (PAK), PowerGlide kits, Midline kits, and Convenience kits. | Nonconforming Material/Component | The firm, BD Medical, sent an "URGENT VOLUNTARY PRODUCT RECALL" letter dated July 22, 2014 and a follow-up "URGENT: Voluntary Product Recall-BD PosiFlush SF Saline Flush Syringe 10mL with rework instructions for applicable Bard Kits" letter dated August 8, 2014 to its customers and distributors. The letters described the product, problem and actions to be taken. The customers were instructed to immediately review your inventory and quarantine product subject to the recall; complete and return the Recall Response Card form via fax to BD at 1-201-847-6990 or email to Becky_Saggau@bd.com and return all affected products. Note: If you do not have any of the affected lots in your inventory, please complete the Recall Response Card form indicating you have (0) quantity and fax or email.The distributors were instructed to immediately examine your inventory and identify any product subject to this notification. In addition, if you have further distributed this product, please identify your customers and notify them at once of this notification. Complete the attached Recall Reply Form and FAX to Bard Access Systems, Inc. at 1-801-522-5674, or EMAIL to BASD.fieldaction@crbard.com. Even if you do not have any inventory of the affected syringe lots.Please direct any return requests and questions to our coordinator at 800-290-1689. If you have any questions or require assistance with the return of the recalled product and/or availability of replacement product, please contact BD Customer Service at 1-888-237-2762 option 3. |  | 48 | 48 units | US Nationwide Distribution. | Terminated | August 05, 2015 | 362 | 290671 | Not_Computer
Z-1268-2015 | 70348 | No | HNO | Keratome. | Ophthalmic | Ophthalmic | 510(k) | ML 7090 Calibrated LASIK Blade (CLB); CLB CALIBRATED LASIK BLADE 10. |  |  | January 13, 2015 | 2015 | Med-logics Inc | The ML 7090 Calibrated LASIK Blade (CLB) will not fit into the AMO Amadeus Metal Blade Holder. | Process design | The firm notified their consignees by letter on 1/13/15, sent via FedEx. The US customers were instructed to return any remaining product to MED-LOGICS and complete the response form. |  | 360 | 360 units | Distributed in the states of CA, AL, GA, and TN, and the countries of Chile and Germany. | Terminated | October 28, 2015 | 288 | 360 | Not_Computer
Z-1279-2015 | 70673 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | GE Healthcare Millennium VG, Discovery VH and VariCam Nuclear Medicine systems, all configurations. |  |  | March 05, 2015 | 2015 | GE Healthcare | GE Healthcare has recently become aware of a potential safety issue due to a collimator falling off a VG system during a Quality Control check, while the detectors were rotating. No injuries have been reported to date | Use error | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref# 40865 dated March 5, 2015. The letter was addressed to Managers of Nuclear Medecine, Hospital Administrators / Risk Managers and Managers of Radiology /Cardiology. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact in the United States: 1-800-437-1171. For other countries, contact your local GE Healthcare Service Representative. |  | 472 | 472 (127 units US, 345 units OUS) | Worldwide Distribution - US Nationwide in the states of: AL, AZ, CA, CO, CT, FL, GA, ID, IL, IN, IA, LA, ME, MD, MA, MI, MN, MS, NE, NY, NC, OK, PA, PR, SD, TN, TX, VA, WA, WV, WI. and the countries of URUGUAY, UNITED KINGDOM,UKRAINE, TURKEY, THAILAND, TAIWAN, SWEDEN, SPAIN, SLOVAKIA, RUSSIA,PORTUGAL, POLAND, PERU, NORWAY, NETHERLANDS, MEXICO, REPUBLIC OF KOREA, JAPAN, ITALY, ISRAEL, INDIA, GREECE, GERMANY, FRANCE, ESTONIA, DENMARK, CZECH REPUBLIC, CYPRUS, COLOMBIA, CHINA, CANDA, BRAZIL, BELGIUM, AUSTRIA, AUSTRALIA, ARGENTINA. | Terminated | August 17, 2015 | 165 | 472 | Not_Computer
Z-1277-2015 | 70143 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Medstream 81" (206 cm) 20 drop Universal Administration Sets. |  |  | October 10, 2014 | 2014 | US Infusion Inc dba Trucare Biomedix-USA | Several complaints regarding the IV sets leaking at the Y site. | Unknown/Undetermined by firm | On December 10, 2014 the firm sent an URGENT DRUG RECALL to its consignees. Instructions were to immediately examine inventory and quarantine all product subject to the recall. Customers were asked to complete and return the response form. For returned goods instructions contact Roberta Parness at 866-593-8444 or via email at rparness@tcbiomedix.com. |  | 62700 | 62,700 IV sets | Distributed in the states of TX, IL, MA, OH, CA, and GA. | Terminated | June 23, 2015 | 256 | 62700 | Not_Computer
Z-1261-2015 | 70558 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Autostainer Link 48 (AS480), Autostainer Plus Link Instrument (AS100), Autostainer (S3400), and Autostainer Plus (S3800). Automated slide stainer for in vitro diagnostic use. |  |  | February 12, 2015 | 2015 | Dako North America Inc. | A false negative result affecting the diagnosis may occur. A defect in the syringe assembly stopcock component may, in some circumstances, cause leakage of excess buffer onto slide location 34 on the AS480, or slide locations 35 and 36 on the AS100, S3400, and S3800. Affected dates are from 12/2013-08/2014. | Nonconforming Material/Component | The firm, Dako, sent an "Urgent Medical Device Recall Notification" letter dated 2/12/15 to customer on 02/12/15.The firm's letter advised customers to do the following:The firm stated that customers can continue to use the instrument however should take the following actions:- Discontinue to use the HercepTest on slide location 34- As set forth in Dako's staining protocol, use on-slide positive controls and two drop zones- Watch for any indication of buffer leakage from they syringe assembly. This can be done easily when the tubes are checked for air bubbles in connection with exchange of the buffer bottle- Check for weak staining in regards to HercepTest in slide position 34.Dako's Actions:Dako states that a representative will contact the customers within the next few weeks in order to determine whether the HercepTest has been used for staining slides in position 34.The firm has identified a longer term corrective action plan of installing a drip tray under the stopcock and states that a representative will contact the customer when it is available.Transmission of Notice:The firm requests that customers inform those that need to be aware of the notification and to where the instruments might have been transferred to. The firm states that customers should ensure the organizations maintains awareness of the noticed and recommended steps until the corrective actions have been completed.The firm also asks customers to complete, sign, and return the attached recall form as soon as possible. On 02/23/15 the firm sent out a second notification letter to customers with the following additional recommendations:- Discontinue the use of HercepTest in slide location 34, 35, and 36 depending on the make and model of the instrumentIn addition the firm recommended that customers pay special attention to any indication that leakage has compromised staining.-When staining using the HercepTest protocols, excess leakage from the sy |  | 409 | 409 | Worldwide Distribution: US (nationwide) and countries of: AT, AU, BE, CA, CH, DE, DK, ES, FR, GB, IE, IT, JP, NL, NO, PL, and SE. | Terminated | February 11, 2016 | 364 | 409 | Not_Computer
Z-1282-2015 | 70384 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Siemens LANTIS Oncology Information System Servers; allows the radiation therapist to deliver treatment to patient using the MEVATRON and all available accessories |  |  | January 21, 2015 | 2015 | Siemens Medical Solutions USA, Inc | There is a potential safety risk when using LANTIS server software with operating systems with which it has not been validated or released which can lead to an incorrect treatment to the patient. | Use error | Siemens sent an Customer Advisory Notice, dated January 21, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were informed the risk associated with use of more recent operating systems with LANTIS Server Software. Customers were warned against the use of more recent operating systems to ensure the LANTIS system remains in the validated configuration.In the interest of safety, we ask you to perform the above preventive measures and inform all affected personnel immediately.Please include this Field Safety Notice in your System Owner Manual chapter "Safety Advisory Letters" where it should remain.The relevant National Competent Authority will be informed of this Field Safety Notice.Please forward this safety information to other organizations that could be affected by this measure.If the device has been sold and therefore no longer in your possession, please forward this customer notice to the new owner. We would also request that you inform us of the identity of the device's new owner where possible. Please observe this notice until the update has been fully completed.We apologize for any inconvenience and thank you for your cooperation.For further questions please call (610) 219-6300. |  | 87 | 87 | Nationwide Distribution | Terminated | July 30, 2015 | 190 | 87 | Computer
Z-1280-2015 | 70546 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Passport V Monitor, Multiparameter Patient Monitor(with Arrhythmia Detection and Alarms) |  |  | January 23, 2015 | 2015 | Mindray DS USA, Inc. dba Mindray North America | An issue has been identified with Passport V Monitors invasive blood pressure function (IBP) which may provide an incorrect IBP measurement | Process control | Mindray DS USA Inc sent Urgent: Passport V Monitor Field Corrective Action letters dated 1/23/2015 to hospital administrators. The letters informed customers that the IBP module may provide an incorrect IBP measurement. The letter also explained the adverse effects on patients. Customers are asked to either calibrate the IBP following the instructions provided; or, contact a Mindray Service Representative, who can perform the calibration, at 1-800-288-2121. In the case of the second option, it is recommended that the monitor be removed from service until the calibration is performed. If the first option is chosen, customers are to notify a Service Representative upon completion of the action. Customers can direct their questions to Diane Arpino at 201-995-8407. |  | 519 | 519 units | Nationwide Distribution | Open, Classified |  |  | 519 | Not_Computer
Z-1283-2015 | 70341 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Triathlon Distal Capture Assembly |  |  | February 12, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Potential for weld disassociation of the cross pin, which could lead to loose or disassociated trigger mechanism on the Distal Capture Assembly. | Device Design | On March 20, 2015, Stryker sent Urgent Medical Device Recall Notifications/Urgent Medical Device Recall Notification Acknowledgement Forms dated March 19, 2015 to all affected customers. The letter identified the product, the related issue and the action to be taken by the customers. Customers were asked to return the attached acknowledgement form confirmed they received and reviewed the notification. Customers with questions were instructed to call 201-831-6365. Stryker sent an Urgent Medical Device Correction letter dated February 11, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were asked to return the attached acknowledgement form confirming they have received and reviewed the letter. Customers with questions were instructed to call 201-831-6365. |  | 2030 | 2030 units | Worldwide Distribution - USA (nationwide) Puerto Rico, and Internationally to Canada, Hong Kong, Netherlands, Sweden, Switzerland, Spain, France, Italy, UK, India, Australia, Korea, Panama, Polan, Ireland, Kuala Lumpar, New Zealand and South Africa. | Open, Classified |  |  | 2030 | Not_Computer
Z-1281-2015 | 70455 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Thumb Screw for the Synthes Recon Locking Aiming Arm for Lateral Entry Femoral Nails-EX (part number 03.010.048) and the Thumb Screw for the Aiming Arm for Titanium Cannulated Tibial Nails-EX (part number 03.010.052). |  |  | February 17, 2015 | 2015 | Synthes, Inc. | For certain lots, the incorrect raw material of annealed 17-4PH Stainless Steel was used instead of the correct material of heat-treated 304 Stainless Steel. This may lead to stress-corrosion cracking, which can lead to surgical delay or unintended debris entering the surgical site. | Process control | An Urgent Notice: Medical Device Recall, dated February 17, 2015, was sent to all direct consignees and sales consultants instructing them to take the following actions:Take note of the quantity of affected parts in your possession, and then follow ONE of the options for replacement of the affected product:Option 1: Call DePuy Synthes Repair Department at 1-800-288-6698 to schedule a repair for the affected products.Complete the Verification Section following the instructions provided and send a copy of the completed Verification Section to DePuy Synthes by: Fax: 866-723-2823 or Scan/e-mail: Synthes6734@stericycle.comOption 2:Call DePuy Synthes Repair Department at 1-800-288-6698 to order repair part. The replacement part will be sent directly to your facility.Questions should be directed to 610-719-5450 or a local DePuy Synthes sales consultant. |  | 183 | 183 | Nationwide Distribution | Open, Classified |  |  | 183 | Not_Computer
Z-1289-2015 | 70385 | No | CFN | Immunoglobulins A, G, M, D, and E immunological test system. | Immunology | Immunology | 510(k) | N Latex IgM; In-vitro diagnostic reagent for the quantitative determination of IgM in human cerebrospinal fluid (CSF) and in paired CSF/serum samples by means of particle-enhanced immunonephelometry using the BN II and BN ProSpec Systems. The determination of IgM aids in the evaluation of the patient's immune system. |  |  | January 15, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Complaints have been confirmed on failed proficiency testing due to serum samples recovering too low when using some N Latex IgM reagent kit lots 44018, 44019 and 44042. | Nonconforming Material/Component | The firm, Siemens, sent an "Urgent Medical Device Recall" letter dated January 2015 to its customers on January 13, 2015. The letter described the product, problem and actions to be taken. The customers were instructed to do the following: Discontinue use of the affected lots and discard any remaining material; Identify if IgM quotient was overestimated due to falsely low IgM results with serum sample; Reconsider patient diagnosis taking other patient data into consideration.; When other patient data does not support the result for an intrathecal IgM synthesis or inflammation, re-measurement of samples may be considered using and alternative lot. In addition, Siemens requested customers to review this letter with their Medical Director; to complete and return the attached Effectiveness Check/Product Replacement form via fax to Customer Care Center at (312) 275-7795; retain this letter with your laboratory records, and forward this letter to those who may have received this product. If you have any questions, please contact your Siemens Customer Care Center 800-441-9250 or email: aydee.I.crawford@siemens.com or your local Siemens technical support representative. |  | 26 | 26 | Worldwide Distribution: US distribution in states of: UT , CO, CA , FL , NY , OH , NC and country of : Canada. | Open, Classified |  |  | 26 | Not_Computer
Z-1285-2015 | 70535 | No | IOL | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | GE Healthcare, Treadmill T2100 and Treadmill T2000 with power cord part numbers 408930-001 Rev E, 408930-002 Rev C, or 2028038-008 Rev B. Also Field Replaceable Unit P/Ns 408930-001 Rev E, 408930-002 Rev C. |  |  | February 16, 2015 | 2015 | GE Healthcare | Potential safety issue due to improper assembly associated with the power cords of T2100 and T2000 Treadmills. | Process control | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref#30073 dated February 16, 2015. The letter was addressed to Director of Biomedical Engineering, Chief of Nursing, &amp; Healthcare Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative.UPDATE 3-30-2015:A second revised letter was sent on March 31, 2015. NOTE: this revised notification includes an update to the Affected Product Details section to include all revisions of power cords manufactured within the date range indicated. (The previous notification referenced only revisions B , C , &amp; E ). |  | 13345 | 13,345 (5,495 units US, 7,850 units OUS) | CONTINENTAL USA DISTRIBUTION INCLUDING DC AND PR.OUS: ARGENTINA AUSTRALIA, AUSTRIA, BANGLADESH, BARBADOS, BHUTAN, BOLIVIA, BRAZIL, BRUNEI, CAMBODIA, CAYMAN ISLANDS, CHILE, CHINA, COLOMBIA, COSTA RICA, DOMINICAN REPUBLIC, ECUADOR, EL SALVADOR, SPAIN, FINLAND, FRANCE, UNITED KIGDOM, GERMANY, GUAM, GUYANA, HONG KONG, INDIA, INDONESIA, IRELAND, ISRAEL, ITLAY, JAMAICA, JAPAN, REPUBLIC OF KOREA, MACAU, MALYAYSIA, MARTINIQUE, MEXICO, MONACO, MYANMAR, NEPAL, NEW ZEALANAD, NICARAGUA, PANAMA, PARAGUAY, PERU, PHILIPPINES, PORTUGAL, RUSSIA. SINGAPORE, SOUTH KOREA, SRILANKA, TAIWAN, THAILAND, TRINIDAD AND TOBAGO, URUGUAY, UZBEKISTAN, VENEZUELA, VIET NAM. | Terminated | June 23, 2017 | 858 | 13345 | Not_Computer
Z-1276-2015 | 70644 | No | DYB | Catheter introducer. | Cardiovascular | Cardiovascular | 510(k) | Medtronic, 10 Fr FlexCath Select Steerable Sheath 990065. |  |  | February 23, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | Clinicians observed debris, appearing to originate from the hemostasis valve on the proximal end of the Steerable Sheath, outside of the patients bodies. | Nonconforming Material/Component | The firm, Medtronic, Inc., sent an "URGENT MEDICAL DEVICE RECALL" letter dated February 2015 to its consignees/customers. Medtronic representatives, beginning Monday, Feb 23, 2015 hand delivered the letter to US consignees/customers.. The letter was addressed to Physician, Risk Manager and Health Care Professional. The letter described the product, problem and actions to be taken. The consignees/customers were instructed to immediately quarantine all unused product; return all affected product to Medtronic; contact Customer Service at 1-800-848-9300 to initiate product return and to complete and return the Customer Confirmation Certificate to via email to: RS.CFQFCA@Medtronic.com or fax to Medtronic at 651-367-0612 to the attention of Customer Focused Quality. In additions, consignees/customers should share this notification with others in your organization.Your Medtronic field representative may assist the customers with identifying suitable replacement product.If you have any questions, contact Customer Service at 1-800-848-9300. |  | 151 | 151 (4 US, 147 OUS) | Worldwide Distribution: US (nationwide) in states of: NC and IA; and countries of: AUSTRIA, BELGIUM, DENMARK, FRANCE, GERMANY, NETHERLANDS, POLAND, SPAIN, and UNITED KINGDOM. | Terminated | October 07, 2015 | 226 | 151 | Not_Computer
Z-1288-2015 | 70648 | component, external, limb, ankle/foot | ISH | External limb prosthetic component. | Physical Medicine | Physical Medicine | 510(K) Exempt | LP Rotate Foot System prosthetic device for fitting lower extremity amputations. |  |  | March 04, 2015 | 2015 | Ossur H / F | The threads of the outer tube of the shock module of LP Rotate are too short, resulting in insecure assembly. Use of the product may cause the patient to fall with continued use, as the shock module can become loose from the foot blade. | Nonconforming Material/Component | The firm's recall letter states the following instructions:Immediately examine inventory and quarantine product subject to the recall. In addition the firm asks if further distributed to identify the customers and notify them of the recall. The firm states that notification may be enhanced by providing a copy of the letter.The firm states that the recall notice needs to be passed along to all those who need to be aware.The firm also states that if customers are wearing a faulty product that they recommend them scheduling a visit to have the product replaced with a new LP Rotate that is currently available. The firm states that they can also arrange pick-up of defected products.For any questions the firm provides a customer service contact: Jamie Dutter, Senior Customer Service Representative. |  | 2194 | 2,194 (1,464 in U.S. and 730 internationally) | Nationwide within the U.S. and internationally to the following foreign countries; Belgium, United Kingdom, Germany, Czech Republic, France, Israel, Poland, Spain, Australia, Japan, Mexico, Bulgaria, Canada, South Korea, Sweden, Finland, | Terminated | October 28, 2015 | 238 | 2194 | Not_Computer
Z-1287-2015 | 70664 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Signature TKA GDE/MDL Set 04-05 (Femur &amp; Tibia) |  |  | February 04, 2015 | 2015 | Materialise N.V. | Images belonging to a different patient were used for the production of the patient specific surgical guide. The incorrect guide may cause a delay in surgery to accommodate the preparation and use of traditional instrumentation. | Employee error | Materialise N.V. sent an email on February 4, 2015, to the the affected distributor regarding the reason for the recall, affected product, and instructions for removing the affected product. For questions regarding this recall call 321-639-6611. |  | 1 | 1 unit | Worldwide Distribution to the Netherlands only. | Terminated | October 30, 2015 | 268 | 1 | Not_Computer
Z-1284-2015 | 70490 | Yes | OAE |  |  | Cardiovascular | PMA | TactiSys Quartz Pack, PN-004 400. |  |  | November 25, 2014 | 2014 | St Jude Medical | A configuration update needs to be done on TactiSys to appropriately recognize all TactiCath catheters. | Software design | St. Jude Medical sent a Service Bulletin to inform of a TactiSys Configuration File update that needs to be performed. A technical service personnel will visit each consignee and performed the software change on-site. For questions about this Service Bulletin, contact Technical Support(ETechsupport@sjm.com). |  | 89 | 89 | US Nationwide Distribution in the states of: AL, AZ, CA, CO, CT, FL, GA, ID, IL IN, KS, LA, MD, MA, MI, MN,MS, NE, NV,NH, NM, NY, NC, OH, OK OR, PA, TN, TX, UT, VA,WA. | Terminated | May 06, 2015 | 162 | 89 | Computer
Z-1286-2015 | 70402 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | CADD" Medication Cassette Reservoir, 50mL, REF 21-7001-24 &amp; 21-7301-24, Sterile EO, Rx only. |  |  | February 02, 2015 | 2015 | Smiths Medical ASD, Inc. | Smiths Medical has become aware of an issue with specific lots of 50mL CADD" Medication Cassette Reservoirs ( Cassette ). Some Cassettes may leak at the sealing area of the pump tube and medication bag. Smiths Medical has received no reports ofserious injury or death related to this issue. | Under Investigation by firm | Smiths Medical sent an "Urgent Medical Device Recall" letter dated February 2, 2015, to all affected customers. The letter was addressed to Risk Safety Manger, Biomedical Professionals, Clinicians who oversee the use of CADD pumps, Distributors, and other users of these devices. The letter described the problem and the product involved in the recall. The letter also addressed the "Advice on Action to be Taken by the User", and "Transmission of this Urgent Recall Notice". Customers were requested to complete and return the Confirmation Form by Fax to 1-800-237-8033 or be email to recall.response@smiths-medical. For issues with these products they can contact Smiths Medical's Global Complaint Department at 1-866-216-8806 or globalcompaints@smiths-medical.com. For questions on the information provided they can contact Customer Service Department at 1-800-258-5361. |  | 48208 | 48,208 (17,040 US, 31,168 OUS). | Worldwide Distribution - US (nationwide) Distribution to the states of : AL, AZ, AR, CA, CO, CT, FL, GA, HI, IL IN, IA, MD, MA, MN, MS, MO, NE, NV, NH ,NJ, NM, NY, NC, OH, OR, PA, SD, TN, TX, UT, VA and WA., and to the countries of : Arab Emirates, Austria, Australia, Belgium, Canada, Switzerland, Colombia, Czech Republic, Germany, Denmark, Spain. Finland, France, United Kingdom (GB) Hungary, Indonesia, Ireland, Italy, Netherlands, Portugal, Sweden, Singapore and South Africa. | Terminated | March 18, 2016 | 410 | 48208 | Not_Computer
Z-1291-2015 | 70551 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | Outer cardboard box labeling: |  |  | February 11, 2015 | 2015 | Sybron Dental Specialties | Cases of Cure Sleeve labeled with part number 4513, which fits 13mm tips, have inner boxes labeled with part number 4500 which fit 7mm and 8mm tips. There is a low risk associated with the use of the product as the sleeve may tear causing a potential for cross contamination. | Packaging | The firm sent out letters to customers on 02/11/15.The firm's letter states that they are willing to replace or credit back customers if they have any of the affected product, and instruct them to contact Metrex Research Customer Care at 1-800-841-1428 in order to receive an RMA number. The firm also requests that customer complete and fax back the enclosed reply form in order to confirm receipt of this advisory notification. The firm also states that if the customer is a distributor that they identify the customers that might have received the affected product, and contact them within 48 hours. |  | 172 | 172 | Worldwide distribution. US Nationwide, Canada, Australia, Netherlands, Switzerland, Japan, | Terminated | May 21, 2015 | 99 | 172 | Not_Computer
Z-1292-2015 | 66737 | Yes | FKX | Peritoneal dialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | The HomeChoice Pro APD System is a peritoneal dialysis system. |  |  | October 16, 2013 | 2013 | Baxter Healthcare Corp. | The occluder in affected devices 1) may have sharp edges that could cut the pneumatic tubing that connects the occluder blade to the manifold causing an unrecoverable system error, and/or 2) blade may separate from the back plate and cause Slow / No Flow alarm requiring therapy termination. | Process design | An Urgent Device Correction letter dated 10/16/13 was mailed to dialysis centers and medical facilities. A similar Urgent Device Correction letter dated 10/31/13 was mailed to dialysis home patients. The letters requested for customers to examine their devices to determine if the devices are being recalled. If the devices are on the attached list of recalled devices, the customer is to contact Baxter Technical Services at 1-800-553-6898 (7 days a week, 24 hours a day) to arrange for a swap cycler. The letter further requests that customers complete the attached Customer Reply Form and return it to Baxter by fax or e-mail. |  | 75 | 75 units (US) | Worldwide Distribution-US (nationwide) including the states of AK, AR, AZ, CA, FL, GA, HI, IA, ID, IL, IN, KS, LA, MA, MD, MI, MO, NC, NJ, NV, NY, OH, PA, SC, TN, TX, UT, VA, VT, WA, and WI, and the countries of Australia, Canada, and Mexico. | Terminated | July 21, 2017 | 1374 | 75 | Not_Computer
Z-1293-2015 | 69813 | No | CGL | Creatinine test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Crea A and B membranes, packaged 4 per box |  |  | October 29, 2014 | 2014 | Radiometer America Inc | Negative drift on QC during in-use lifetime of CREA membranes can be observed. | Under Investigation by firm | Radiometer sent a Product Notification letter dated November 10, 2014, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to perform the following steps immediately:" Perform one quality control measurement using S7835 AutoCheck 6+, level 1 on already installed Crea membrane units to verify the Crea performance. If the QC result is out of range, new Crea A and Crea B membranes must be installed.Until further notice the following must be carried out every 24 hours:" Perform one quality control measurement using S7835 AutoCheck 6+, level 1 to verify the Crea performance. If the QC result is out of range, new Crea A and Crea B membranes must be installed.Please do not report Crea results unless the membrane unit has been verified as described above.If you already run S7835 AutoCheck 6+ level 1, every 24 hours, as part of your regular QC schedule, you do not need to do anything additional. Please review your QC results to ensure the results are within the QC range.Please complete and return the attached fax form, with your signature as soon as possible or within 7 days.Customers with questions were instructed to contact Radiometer America Technical Support at 1-800-736-0600, option 4.For questions regarding this recall call 440-871-8900, ext 209. |  | 10824 | 10,824 boxes | Worldwide Distribution - USA including NM, NC, TN, OH, NE, CA, MO, GA, MA, TX, WV, PA, NY, VA, MN, IN, ME, MD, IL, MI, KS, WI, FL, WA, CT, OK, CO, and Internationally to Canada, Denmark, Spain, Italy, India, Sweden, Germany, Australia, Hungary, Great Britain, Norway, Uruguay, France, Poland, Switerland, Russian Federation, Netherlands, Czech Republic, Mexico, Brazil, Czechoslovakia, Austria, Singapore, Syria, Phillipines, Egypt, Bolivia, Japan, New Zealand, Turkey, Saudi Arabia, Algeria, Qatar, and Finland. | Terminated | July 02, 2015 | 246 | 10824 | Not_Computer
Z-1295-2015 | 70647 | Yes | MQS | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Transonic Hemodialysis Monitor, Model HD03 -- |  |  | February 17, 2015 | 2015 | Transonic Systems Inc | Customer unable to obtain measurements due to the device seeing no flow or irrelevant flow condition during dialysis procedures. The user saw a zero or a very low flow value displayed on HD03 and/or no measurements produced in other modes (recirculation; access flow; or cardiac output). | Process control | Transonic sent a Medical Device Recall Letters and Response Forms dated February 17, 2015 to the users via e-mail and fax. The recall is being conducted to the user level. The letter identified the affected product, problem and actions to be taken. For questions call 607-257-5300 ext. 326. |  | 31 | Domestic: 21 units, Foreign: 10 units | Worldwide Distribution - USA Nationwide and the countries of: Australia, Canada, and South Korea | Terminated | May 31, 2016 | 469 | 31 | Not_Computer
Z-1294-2015 | 70698 | No | DWQ | External vein stripper. | Cardiovascular | Cardiovascular | 510(k) | LeMaitre Vascular Tumescent Catheter Inversion (TCI) System. The TCI System is intended to provide a system to the removal of an incompetent greater saphenous vein. |  |  | February 24, 2015 | 2015 | LeMaitre Vascular, Inc. | Mislabeled with an incorrect expiration date | Employee error | LeMaitre Vascular issued a Field Safety Notice/Dear Doctor dated 2/24/15 to customers The letter provides information on the recalled product and instructions on how the customer can inspect for the defective devices and return the affected products for replacement/credit. The letter contains a form that is requested to be returned to LeMaitre Vascular, Inc. as a record of notification and reconciliation.Questions, contact 781-221-2266 ext. 183 |  | 54 | 54 units | CA, CO, KS, MN, UT, WA, WI | Terminated | May 21, 2015 | 86 | 54 | Not_Computer
Z-0984-2015 | 69880 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ARTISTE MV System, Linear Accelerator (LINAC) with SysVC10A software. |  |  | November 19, 2014 | 2014 | Siemens Medical Solutions USA, Inc | There may be an existing dark current phenomenon on ARTISTE LINAC in combination with IMRT or mARC treatments using unflat beams. Software issue. | Software design | Siemens Planned Action: 1. Providing recommendations for user actions in terms of routine dosimetric quality control checks for any IMRT treatment at very high dose rates (Multiple X) as well as for mARC treatments.2. Software update for the Function Controller #3, which controls parts of the beam generation. This software update enables the correct switching between Pulse Forming Network (PFN) and reduced Pulse Forming Network (iPFN) in pause state, to reduce the dark current during the movement as much as possible. This software will be installed by a Siemens service representative.3. Siemens will contact the user to make an appointment to implement the modification system at a date and time when the system will be available. These corrections shall be made free of charge.The deployment of the permanent correction shall be completed by July 1, 2015.For further questions please call (610) 219-6300. |  | 3 | 3 | US Distribution to the states of UT, WI, and NY. | Terminated | April 28, 2015 | 160 | 3 | Computer
Z-1301-2015 | 70286 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | NovaPlus Brand Warm Gel Infant Heel Gel Warmer, Item # V6391 |  |  | January 26, 2015 | 2015 | CooperSurgical, Inc. | Products without expiration date have the remote potential for minor burn | Package design/selection | CooperSurgical issued Recall Notification letter on January 26, 2015, via Fedex confirmed receipt delivery, to all affected customers. Customers were requested to review inventory and return/destroy product that does not have a expiration and lot number. A fax back form to be returned. Customers with questions should call 203.601.5200.CooperSurgical's expense, arrangements were made to replace any of the affected products that customers had on hand. CooperSurgical provided further instruction on the disposition of the product at the time of replacement. |  | 4938266 | 4938266 units | Worldwide Distribution - USA (nationwide) and Internationally to AUSTRALIA, BAHRAIN, BERMUDA, CANADA, FRANCE, LATVIA, LIBYA, MEXICO, NEW ZEALAND, QATAR, SAUDI ARABIA, SINGAPORE, and UNITEDKINGDOM. | Open, Classified |  |  | 52899493 | Not_Computer
Z-1238-2015 | 70458 | No | JAI | Powered radiation therapy patient support assembly. | Radiology | Radiology | 510(k) | Elekta Precise Treatment Table |  |  | January 29, 2015 | 2015 | Elekta, Inc. | It is possible to position the Treatment Table with errors greater than 5mm and no inhibits are displayed, if the positional sensors are not correctly installed during corrective maintenance. Clinical mistreatment may occur if a position error is not detected. | Radiation Control for Health and Safety Act | Planned Action by Elekta:1. Notifications have been sent out. The notification informs users of the possibility treatment table position errors might occur if there is a maintenance error when replacing a potentiometer. The notification recommends establishing a procedure to verify the accuracy of the treatment table position as shown by the treatment room monitor which is performed as part of the daily machine checks. An example procedure is provided.2. Elekta will establish a procedure service engineers will follow after maintenance involving the relevant potentiometers to ensure proper installation. The procedure will be taught using a video which includes a step by step guide to the correct methodology for installation of the potentiometer assembly, a description of the impact for not installing correctly, and a guide to the new verification of correct installation procedure. The video will be distributed through Elektas learning management system (LMS). The LMS maintains a record of service engineers who successfully complete the training.3. Elekta will implement an engineering change to the potentiometer assembly which applies to the manufacture of all new Precise Treatment Tables. This change will eliminate the potential for incorrect installation and will prevent a loose grub screw from causing backlash. This change is expected to consist of a 'D' shaped bore that cannot spin on the shaft of the potentiometer.All actions associated with this correction, including service engineer training, will be completed by June 2015. For further questions please call (770) 300-9725. |  | 3480 | 529 US 2951 Worldwide | Worldwide Distribution - USA (nationwide) Distribution and to the countries of Albania, Algeria, Argentina, Australia, Austria, Bahrain, Bangladesh, Belarus, Belgium, Bosnia and Herzegovina, Brazil, Bulgaria, Canada, Chile, China, Colombia, Croatia, Cuba, Cyprus, Czech Republic, Denmark, Ecuador, Egypt, Finland, France, Germany, Greece, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Lebanon, Libya, Lithuania, Malaysia, Malta, Mexico, Morocco, Myanmar, Namibia, Netherlands, New Zealand, Nigeria, Norway, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Senegal, Serbia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland and Syria. | Open, Classified |  |  | 3480 | Not_Computer
Z-1296-2015 | 70246 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Symbia E is a multi-purpose SPECT system used to detect or image the distribution of radionuclides in the body or organ, using the following techniques: planar imaging, whole body imaging, tomographic imaging for isotopes with energies up to 588 keV. |  |  | April 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | Radial motor drives replaced during a customer service action were incorrectly assembled backwards. | Employee error | Urgent Field Correction- Recall letters dated April 11, 2012 were sent to all potentially affected customers. The letters included instructions for customers to immediately remove the patient from the bed and contact the customer's Service Engineer in the event that one of the following errors is displayed "A motion control error has occurred (E112)" or "A motion control error has occurred (E113)." Customers were advised that a Service Engineer will be inspecting the radial drive gear on the customer's camera system to ensure it is functioning correctly and that the machines can continue to be used as usual until the Service Engineer can perform the inspection of the machine(s).Customers with questions were advised to contact Siemens as 800-888-7436. |  | 419 | 419 systems (worldwide) | Worldwide Distribution-US (nationwide) including the District of Columbia and the states of AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, VA, VT, WA, WI, and WV and the countries of AE, AL, AO, AR, AT, AU, BA, BD, BE, BG, BH, BO, BR, BW, BY, CA, CH, CL, CM, CN, CO, CR, CY, CZ, DE, DK, DO, DZ, EC, EG, ES, FA, FI, FR, GB, GE, GR, HK, HR, ID, IE, IL, IN, IQ, IR, IT, JO, JP, KR, KW, KZ, LB, LK, LU, LV, MA, MK, MM, MU, MX, MY, NE, NG, NL, NO, NP, NZ, OM, PE, PH, PK, PL, PS, PT, QA, RO, RS, RU, SA, SE, SG, SI, SK, SV, SY, TH, TN, TR, TW, UA, UG, UY, UZ, VN, YE, and ZA. | Open, Classified |  |  | 5275 | Not_Computer
Z-1306-2015 | 70151 | No | MMP | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General Hospital | 510(k) | Penner Patient Stretcher/Transfer Lift System Models 391000-1 and 392000-1 |  |  | January 30, 2015 | 2015 | Penner Mfg Inc | If the patient sits on either wing, not squarely on the center section of the Stretcher Lift, tipping and injury to the operator or patient could result. The device manual was updated with an additional warning statement. | Use error | Beginning on 1/30/2015, the recalling firm sent letters via certified mail to their customers informing them of the change to the stretcher manual along with a revised manual which incorporated those changes. For questions regarding this recall call 402-694-5003. |  | 158 | 158 | Worldwide Distribution - USA (nationwide) and Internationally to Canada. | Terminated | March 10, 2016 | 405 | 158 | Not_Computer
Z-1307-2015 | 69812 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | Lithium Polymer Fiber Optic Module Batteries used in the Intera MR System upgraded to CDAS, Part Number: 9065 and 9067. These rechargeable batteries are used to supply power to the following products: Wireless ECG Module used with the Precess MRI compatible Patient Monitoring System, Wireless SpO2 Module used with the Precess MRI Compatible Patient Monitoring System, Philips Basic Triggering Set, Philips Monitoring Sensor Set, Philips Battery Set |  |  | November 20, 2014 | 2014 | Philips Medical Systems, Inc. | VCG battery ignited in a VCG unit when recharging. | Component design/selection | Consignees were informed of the recall via letter sent on November 20, 2014 via certified mailing. Field Service Engineers (FSE) will replace all affected batteries. |  | 79 | 79 | Distributed in the following states: AL, CA, FL, GA, IL, MA, MN, NC, ND, NE, NJ, NY, OH, OR, PA, TN, TX, UT, WA, WI; OUS: Australia, Austria, Belgium, China, Ecuador, Finland, France, Germany, Italy, Japan, Mexico, Netherlands, slovenia, South Africa, Sweden, Switzerland, United Kingdom | Open, Classified |  |  | 79 | Not_Computer
Z-1211-2015 | 70478 | No | RCE |  |  |  |  | GE Inspection Technologies x¦tome¦x L240 an Cabinet X-ray system |  |  | January 13, 2015 | 2015 | GE Inspection Technologies, LP | It was discovered that the upper left corner-shaped lead shield in the window in one of the doors of the v¦tome¦x L240 installed had lost adhesion and fall out of position, which could lead to gaps in shielding. | Device Design | GE Inspection Technologies, Planned Action: 1. GE will notify customers of the affected cabinet x-ray systems through a Technical Information Letter. GE notified its affected customers via email on January 13, 2015, via letter dated January 14, 2015, and via telephone on January 19, 2015. 2.GE will install an interim insert within 30 days of receipt of this letter. The insert will cover any existing or interim gaps in the original lead shielding on the door.3. GE will replace the doors of affected cabinet x-ray systems no later than 180 day of receipt of this letter.4. These corrections will be made free of charge.The Technical Information Letter used for notification of all dealers and purchasers is to be made within 15 working days of receipt of this letterFor further questions, please call (717) 447-1278 |  | 2 | 2 | US Distribution to the states of OH and PA. | Terminated | August 17, 2015 | 216 | 2 | Not_Computer
Z-1310-2015 | 70686 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RayStation Radiation Therapy Treatment Planning System; -- RayStation 3.5, RayStation 4.0, RayStation 4.5 and RayStation 4.7. For RayStation 4.7, the issue applies also to machines with fixed jaws, regardless of MLC/jaw position. |  |  | March 04, 2015 | 2015 | RAYSEARCH LABORATORIES AB | An issue with photon dose calculation for DMLC (Dynamic MLC) plans for machines where the MLC is positioned above the jaws, e.g. some Elekta linacs. The magnitude of the error depends on the beam model output factor corrections and on the individual DMLC plan characteristics. | Under Investigation by firm | The firm, Raysearch Laboratories, sent a "Medical Device Correction/Field Safety Notice" and Reply Form(dated 3/3/2015) on 3/4/2015 via e-mail to the affected customers. The notice describes the product, problem and actions to be taken. The customers were instructed to deprecate any beam models that have been commissioned as DMLC capable for the affected linac types; do not commission beam models as DMLC capable for the affected linac types; educate physics staff and all users; inspect your system and identify all affected units of RayStation, acknowledge this notice by email or complete and return the Reply Form via fax to Americas market, RaySearch Americas Inc., at 888 501 7195, email: freddie.cardel@raysearchlabs.com or Rest of the world, RaySearch Laboratories AB, to email support@raysearchlabs.com.If you have any questions, contact Quality and Regulatory Affairs Manager, at +46 722 366 110 or email david.hedfors@raysearchables.com. |  | 245 | Domestic: 245 units | Nationwide Distribution. | Open, Classified |  |  | 245 | Not_Computer
Z-1290-2015 | 70727 | No | MCX |  |  | Cardiovascular | PMA | Diamondback 360 Coronary Orbital Atherectomy System, Model Number: DBEC-125. |  |  | March 04, 2015 | 2015 | Cardiovascular Systems, Inc. | The OAD saline sheath outer layer may flake off from the catheter due to friction between the outer material layer and corresponding materials it may come in contact with during clinical use. If the material is large enough, it has the potential to cause a flow limiting embolism. | Nonconforming Material/Component | Consignees were sent on 3/4/2015 a CSI "Urgent Medical Device Recall" letter dated March 04, 2015. The letter described the problem and the affected product. Consignees were advised to remove and return the affected product to CSI. The letter also requested consignees to complete and return the Customer Acknowledgement Form by Fax or with the RGA (Returned Goods Authorization). For further information they can contact Customer Service, Cardiovascular Systems, Inc., 1225 Old Hwy 8, New Brighton, MN 55112, 877-274-0901 -Tel, 612-677-3355- Fax. |  | 30 | 30 | Nationwide Distribution-including the states of AZ, FL, GA, LA, MA, MS, MO, NJ, NY, OH, and RI. | Terminated | April 26, 2016 | 419 | 30 | Not_Computer
Z-1304-2015 | 70705 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | da Vinci Xi Surgical System, IS4000; |  |  | March 03, 2015 | 2015 | Intuitive Surgical, Inc. | Correction due to a fault that causes distal SUJ (DSUJ) failures that indicates a mismatch between the positions reported by the redundant position sensors on one of the patient cart arm joints. | Under Investigation by firm | The firm, Intuitive Surgical, sent an "Urgent Medical Device Correction" letter dated March 11, 2015 to its consignees/customers on March 11, 2015 via Fed Ex. The letter described the product, problem and actions to be taken. The customers were instructed to ensure that all affected personnel are fully informed of this notice; forward this notice to your Risk Manager, OR Director, Purchasing Manager, biomedical Engineering staff and members of your medical staff who perform da Vinci Surgery procedures; Identify systems containing affected product at your site; if you encounter this error during a procedure and the fault cannot be recovered, disable the affected arm and contact dVStat; Contact Intuitive Surgical Customer Service; Complete and return the attached Acknowledgment Form to Intuitive Surgical via U.S. Fax +1 (408) 716-3040, or Scan and Email: isi.compliance@intusurg.com, ATTN: REGULATORY COMPLIANCE-Subject line for email: Pot to Encoder;and retain a copy of this notice for your records.An Intuitive Surgical representative will contact you to schedule a remote assessment to determine the health of your system and to schedule an appointment to replace the affected arm.If you need further information or support concerning this Medical Device Correction, please contact your Clinical Sales Representative or Intuitive Surgical Customer Service at the numbers listed below:-North America and South America: 800-876-1310, Option 3 (6 AM to 5 PM PST)-Japan: 0120-56-5635 or 03-5575-1362 (9 AM to 6 PM JST)-South Korea: 02-3271-3200 (9 AM to 6 PM KST)-Europe, Middle East, Asia and Africa: +800 0821 2020 or +41 21 821 2020(8 AM to 6 PM CET) or ics@intusurg.com. |  | 167 | 167 total affected da Vinci Xi systems | Worldwide Distribution: US (nationwide) and countries of: Australia, Belgium, China, France, Germany, India, Israel, Italy, Norway, South Korea, Spain, Sweden, Switzerland, Turkey, and United Kingdom. | Terminated | September 08, 2015 | 189 | 167 | Not_Computer
Z-1308-2015 | 70712 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | MAGNETOM systems Aera/Skyra/Avanto/Verio with software syngo MR D13A ; indicated for use as magnetic resonance diagnostic devices (MRDD) that produce transverse, sagittal, coronal and oblique cross-sectional images, spectroscopic images and/or spectra, and that display the internal structure and/or function of the head, body, or extremities. Other physical parameters derived from the images and/or spectra may also be produced. Depending on the region of interest, contrast agents may be used. These images and/or spectra and the physical parameters derived from the images and/or spectra, when interpreted by a trained physician, yield information that may assist in diagnosis. The MAGNETOM systems may also be used for imaging during interventional procedures when performed with MR compatible devices such as in-room display and MR-safe biopsy needles. |  |  | March 09, 2015 | 2015 | Siemens Medical Solutions USA, Inc | The gradient output supervision was permanently turned off on the MAGNETOM system, meaning that gradient outputs could exceed IEC60601-2-33 limits and peripheral nerve stimulation could occur. | Nonconforming Material/Component | A Customer Safety Advisory Notice, dated March 6, 2015, was sent to end users to inform them of the potential issue and what measures were being taken to mitigate possible risks. |  | 132 | 132 | Nationwide Distribution. | Open, Classified |  |  | 132 | Not_Computer
Z-1305-2015 | 70687 | No | LNH | Magnetic resonance diagnostic device. | Radiology | Radiology | 510(k) | GE Healthcare, SIGNA 1.5T TWINSPEED, SIGNA INFINITY MRISYSTEM, SIGNA 3.0T INFINITY WITH EXCITE, SIGNA 3.0T WITH EXCITE, SIGNA EXCITE 1.5T, SIGNA EXCITE 3.0T, GE 1.5T AND 3.0T SIGNA HDX MR SYSTEM, Discovery MR750 3.0T, Discovery MR450 1.5T, Discovery MR750w 3.0T, Optima MR450w 1.5T, SIGNA Contour, Brivo MR355, Optima MR360, GE 1.5T SIGNA HDE MR SYSTEM, Signa Openspeed 0.7T MR SYSTEM, Vectra, SIGNA 0.35T OVATION WITH EXCITE, SIGNA 0.2T PROFILE With EXCITE |  |  | March 09, 2015 | 2015 | GE Healthcare | GE Healthcare has become aware of a potential safety issue involving MRI systems due to software versions not being maintained properly at some sites. | Software Manufacturing/Software Deployment | GE Healthcare sent an "Urgent Medical Device Correction" letter dated March 9, 2015, to all affected customers. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact your local service representative or call 800-437-1171 (available 24 hours a day). |  | 9369 | 9,369 (2,937 US, 6,432 OUS). | Worldwide Distribution - USA (nationwide) Distribution including PR, GU and DC., and to the countries of: ALBANIA. ALGERIA, ARGENTINA, AUSTRALIA, AUSTRIA, AZERBAIJAN, BAHRAIN, BELARUS, BELGIUM, BOLIVIA, BOSNIA AND HERZEGOVINA, BRAZIL, BULGARIA, CANADA, CHILE, CHINA, COLOMBIA, COSTA RICA, COTE DLVOIRE, CROATIA, CYPRUS, CZECH REPUBLIC, DENMARK, DOMINICAN REPUBLIC, ECUADOR, EGYPT, EL SALVADOR, ESTONIA, FINLAND FRANCE, FRENCH POLYNESIA, GEORGIA, GERMANY. GHANA, GREECE, GUADELOUPE, GUATEMALA, HONDURAS, HONG KONG, HUNGARY, ICELAND, INDIA, INDONESIA, IRAN IRAQ, IRELAND, ISRAEL, ITALY, JAMAICA, JAPAN, JORDAN, KAZAHHSTAN,KENYA,REPUBLIC OF KOREA, KUWAIT, LATVIA, LEBANON, LIBYAN ARAB JAMAHIRIYA, LITHUANIA, LUXEMBOURG MACEDONIA, MALAYSIA, MEXICO, MOROCCO, NETHERLANDS, NEW CALEDONIA, NEW ZEALAND, NICARAGUA, NIGERIA, NORWAY, OMA N,PAKISTAN, PALESTINE, PANAMA, PARAGUAY, PERU, PHILIPPINES, POLAND, PORTUGAL, REUNION, ROMANIA, RUSSIA, SAUDI ARABIA, SERBIA, SIERRA LEONE, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN SUDAN, SWEDEN, SWITZERLAND, SYRIA, TAIWAN, THAILAND, TUNISIA, TURKEY, UGANDA , UKRAINE, UNITED ARAB EMIRATES, UNITED KINGDOM, URUGUAY, UZBEKISTAN, VENEZUELA, VIETNAM and ZAMBIA. | Terminated | September 08, 2015 | 183 | 9369 | Computer
Z-1309-2015 | 70557 | No | HQL | Intraocular lens. | Ophthalmic | Ophthalmic | PMA | Intraocular Lens (IOL), Part No. EC-3 PAL, Diopter: 17.5, 18.0, 18.5, 19.0, 19.5, 21.5, 23.5, 9.0, 22.5, 23.0, 21.0, 24.0. |  |  | February 13, 2015 | 2015 | Aaren Scientific, Incorporated | Aaren Scientific is recalling intraocular lenses because the instructions for use (IFU) are not fully compliant with US requirements. | Labeling design | A customer notification letter dated 2/13/15 was sent to all customers who received the Intraocular Lens model EC-3 PAL. The letter provides the customers with the problems identified and the actions to be taken. Customers are instructed to confirm receipt of the recall letter and to contact Mark Scheib at (909) 906-5119 or via email at mark.scheib@zeiss.com with any questions or comments. |  | 28 | 28 units | Distributed in the states of OK, KS, FL, and AL. | Terminated | June 29, 2015 | 136 | 28 | Not_Computer
Z-1312-2015 | 70216 | No | FRG | Sterilization wrap. | General Hospital | General Hospital | 510(k) | Tyvek Pouch/Roll with STERRAD Chemical Indicator, P/N 12425 |  |  | December 24, 2014 | 2014 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the Tyvek Rolls with STERRAD Chemical Indicator because they were shipped to customers after its expiration date of May 18, 2013. | Under Investigation by firm | A customer notification letter was sent on 12/24/14 to all customers who purchased the expired Tyvek Rolls with STERRAD Chemical Indicator. The letter informs the customers of the problems identified and the actions to be taken. |  | 2 | 2 units | Saudi Arabia and Portugal | Terminated | September 04, 2015 | 254 | 2 | Not_Computer
Z-1313-2015 | 70247 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Electric Pen Drive (EPD) System, The Synthes Electric Pen Drive is indicated for screw insertion, pin and wire replacement, cutting of bone and metal, drilling, decorticating, shaping, and smoothing of bones and teeth in a wide variety of surgical procedures, including, but not limited to general orthopaedic trauma, foot, hand, maxillofacial, neurosurgical, oral, otolaryngolical, reconstructive, and spine surgery. |  |  | December 19, 2014 | 2014 | Synthes, Inc. | In May 2013, the original recall was initiated due to complaints the Hand Switch for the Electric Pen Drive was not labeled correctly, and the on and lock labels were interchanged. | Process design | Synthes sent an Urgent Medical Device Recall Notice, dated December 19, 2014, to all affected customers to provide updated information on additional lots identified. Customers were instructed to examine their inventory for the affected product, remove from inventory, and return them to Synthes immediately for replacement. Customers were instructed to call Synthes to obtain a Return Authorization (RA) Number, complete the Verification Section at the end of the letter, and return the Verification Form with the product to: Credit/Returns, Synthes 1101 Synthes Avenue Monument, CO 80132. Customers with questions were instructed to call 610-719-5450 or email FieldAction@synthes.com. For questions regarding this recall call 610-719-5000. |  | 249 | 249 | Nationwide Distribution | Open, Classified |  |  | 249 | Not_Computer
Z-1311-2015 | 70608 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC units, Model No. 8015. Infusion pump. |  |  | March 12, 2015 | 2015 | CareFusion 303, Inc. | CareFusion is recalling the Alaris PC unit because of an error code. The error code may occur upon power on during the "Power-On Self Test" due to a keypad issue. | Process control | CareFusion sent an Urgent: Medical Device Recall Notification letter dated March 12, 2015, to all affected customers. The letter informs the customers that CareFusion has received service reports of a system error on the Alaris PC unit 8015. The letter informs customers of the problems identified and the actions to be taken. Customers with recall related questions are instructed to contact CareFusion Support Center at (888) 562-6018, supportcenter@carefusion.com. Customers with adverse event reports are instructed to contact customer advocacy at (888) 812-3266, customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris system are instructed to contact technical support at (888) 812-3229, DL-US-INF-TechSupport@carefusion.com. |  | 56015 | 56,015 units | Worldwide Distribution - USA (nationwide), Canada, European Union, Singapore and Australia. | Open, Classified |  |  | 56015 | Not_Computer
Z-1314-2015 | 70578 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Variable Angle Locking Peg, D:1.8mm x L:14.0mm; Variable Angle Locking Peg, D:1.8mm x L:16.0mm, Variable Angle Locking Peg, D:1.8mm x L:18.0mm; Variable Angle Locking Peg, D:1.8mm x L:20.0mm; Variable Angle Locking Peg, D:1.8mm x L:22.0mm, Variable Angle Locking Peg, D:1.8mm x L:24.0mm; Variable Angle Locking Peg, D:1.8mm x L:26.0mm. |  |  | February 16, 2015 | 2015 | Flower Orthopedics Corporation | Locking pegs were not locking during distal radius procedures. There has been one (1) report of revision surgery scheduled to remove a loosened peg. | Device Design | On February 16, 2015, Flower Orthopedics generated and distributed a field memo/advisory notice (FOCMB-15-02) to all distributors with locking peg inventory on February 6, 2015. The field memo provided technique guidance to end users. All 12 consignees with remaining inventory (ie.not reported as used) were contacted directly via phone and all have agreed to ship the locking pegs back to Flower Orthopedics. Flower Orthopedics determined that a generic statement to all customers was not necessary for the following reasons detailed above and summarized again here: product was distributed only through direct channels, the health hazard evaluation has determined there is a very low risk to health, and all consignees have been directly contacted already and product inventory is being returned. |  | 901 | 901 | Distributed in the states of VA, TX, IL, OK, PA, CA, NY, KY, and FL. | Terminated | August 06, 2015 | 171 | 901 | Not_Computer
Z-1315-2015 | 70090 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | Liquid Mouse Monoclonal Antibody CD15 used is laboratories to perform immunohistochemical (IHC) staining. |  |  | December 16, 2014 | 2014 | Leica Microsystems, Inc. | The staining intensity decreases over the shelf life. | Labeling design | Leica sent an Urgent Medical Device Recall Notification letters dated December 16, 2014 to all customers of record. The letters included instructions for customers to: 1) destory any usused and partially used recalled materials or confirm that the materials are no longer in stock; 2) complete and return the attached Field Safety Correction Notice Acknowledgement Form; 3) contact your local Leica representative immediately if a replacement is required; and, 4) pass this notice primarily to the end users where the product has been sold and to all those within your organisation who need to be aware of this manufacturing issue. Customers with questions can contact Leica Microsystems via e-mail at LMGRA@leica-microsystems.com. |  | 309 | 309 units | Worldwide Distribution - US Nationwide in the states of CA, CT, FL, IA, ID, IL, MT, NC, NY, TX, WA and in the countries of: Argentina, Australia, Bosnia and Herzegovina, Brazil, Canada, Chile, Denmark, El Salvador, France, Germany, Hungary, India, Italy, Japan, Netherlands, Nigeria, Portugal, Romania, Russia, Saudi Arabia, Serbia, South Africa, Spain, Sweden, Switzerland, Taiwan, Tunisia, Turkey, United Arab Emirates, United Kingdom. | Open, Classified |  |  | 309 | Not_Computer
Z-1270-2015 | 70443 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare Engstrom Carestation, Engstrom Pro, Cart Only S/5 EC. GE model numbers 1505-8538-000, 1505-9000-000, and 1505-9003-000. Ventilators. |  |  | January 27, 2015 | 2015 | GE Healthcare | Potential failure of the caster mounting hardware, which may result in the ventilator falling/tipping (overbalancing). | Device Design | GE Healthcare sent an Urgent Medical Device Correction letter dated January 26, 2015, Ref#34063. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171, or your local Service Representative. |  | 3213 | 3,213 (437 US, 2,776 OUS) | Worldwide Distribution. US nationwide, ALBANIA, ARGENTINA, AUSTRALIA, AUSTRIA, BAHRAIN, BANGLADESH, BELGIUM, BOSNIA &amp; HERZEGOVINA, BRAZIL, CANADA, CAMEROON, CANDA, CHILE, COLOMBIA, COSTA RICA, CROATIA, CZECH REPUBLIC, DENMARK, ECUADOR, EGYPT, EL SALVADOR, FINLAND, FRANCE, GERMANY, CHANA, GREECE, GUATEMALA, HONDURAS, INDIA, INDONESIA, IRAQ, ISRAEL, ITALY, JAPAN, KAZAKHSTAN, KENYA, REPUBLIC OF KOREA, KUWAIT, LEBANON, LIBYAN ARAB JAMAHIRIYA, MACEDONIA, MALAYSIA, MALTA , MEXICO, MOROCCO, NETHERLANDS, OMAN, PANAMA, PAPUA NEW GUINEA, PERU, POLAND, PORTUGAL, QATAR, RUSSIAN FEDERATION, SAUDI ARABIA, SERBIA, SINGAPORE, SLOVAKIA, SOUTH AFRICA, SPAIN, SRI LANKA, SWITZERLAND, SYRIAN ARAB REPUBLIC, THAILAND, TUNISIA, TURKEY, UNITED ARAB EMIRATES, UNITED KINGDOM, VIETNAM, VENEZUELA, and URUGUAY. | Terminated | February 03, 2016 | 372 | 3213 | Not_Computer
Z-1322-2015 | 70693 | No | ISW | External assembled lower limb prosthesis. | Physical Medicine | Physical Medicine | 510(K) Exempt | The Rheo Knee bionic prosthetic is to be used exclusively for fittings of lower extremity amputations. RHEO KNEE uses Artificial Intelligence to continuously adapt to the users walking style and environment. The RHEO KNEE recognizes and responds immediately to changes in speed, load and terrain, restoring the users ability to walk naturally, comfortably and confidently at any speed. |  |  | February 26, 2015 | 2015 | Ossur H / F | The firm is recalling Rheo Knee bionic prosthetic due to it being discovered during an internal audit of the service line that devices were released for distribution without fully going through the assembly process. | Employee error | The firm, Ossur, notified their customers of the Recall Notice via phone on February 26, 2015. The firm informed the customers about the recalled product, the problem and actions to be taken. The customers were instructed to send the unit in for service so the Acceptance testing can be performed for the unit in accordance to their process. If you have any questions, please contact Regulatory Affairs &amp; Quality Assurance specialist at 949-382-3741 or email: kmontes@ossur.com. |  | 4715 | 4715 units (2 units affected) | Worldwide distribution: US distribution in TN and country of: Austria. | Terminated | November 20, 2015 | 267 | 4715 | Not_Computer
Z-1316-2015 | 70657 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac versions 6.x patient positioning systems are used to position patients during radiosurgery or radiotherapy procedures.. |  |  | February 16, 2015 | 2015 | Brainlab AG | The ExacTrac 6.x Patient Positioning System may incorrectly position the patient when using the ExacTrac Cone Beam CT (CBCT) with a TrueBeam-specific optional subvolume-CBCT. | Software in the Use Environment | BrainLab sent an FIELD SAFETY NOTICE / PRODUCT NOTIFICATION letters dated February 16, 2015, to all affected customers. The letters included instructions for customers to: 1) Do not use any actively re-reconstructed CBCT subvolumes from the Varian TrueBeam system with the Brainlab ExacTrac 6.x CBCT Import &amp; Alignment Software module; 2) Import exclusively original, not modified, CBCT volumes into ExacTrac to use for patient positioning; and, 3) Continue to verify any ExacTrac CBCT-based correction using the ExacTrac X-ray verification and/or an external IGRT procedure as mandatory. Brainlab provided existing potentially affected ExacTrac v.6.x CBCT Import &amp; Alignment Software customers (with a Varian TrueBeam system) with the product notification information. Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact the customer starting August 2015 to schedule the update installation.For further questions please call Customer Hotline + 49 89 99 15 68 44 or 1-800 597-5911 ( For US Customers only) |  | 24 | 24 systems (US); 36 systems (Foreign) | Worldwide Distribution - USA and to the countries of : Australia, Austria, Canada Denmark, France, Germany, Hong Kong, Japan, Russia, Singapore, South Korea and United Kingdom. | Open, Classified |  |  | 24 | Computer
Z-1317-2015 | 70593 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | BD Q-Syte Luer Access Split Septum, 0.16 ml, Sterile, Rx only, Catalog REF #385100, Manufacturer: Becton Dickinson Infusion Therapy Systems Inc. |  |  | February 19, 2015 | 2015 | Becton Dickinson &amp; Company | Becton Dickinson received an elevated number of complaints for blood or infusate/medication leakage due to tearing of the silicone septum. | Process control | Becton Dickinson sent an Urgent Voluntary Product Recall Letter/Recall Response Card/Packing Instructions, dated 2/19/15, to their affected customers via UPS next day mail. The letters identified the affected device and explained the reason for the recall. Customers were to review their inventory for the specific catalog and lot numbers provided and quarantine the affected product. The Recall Response Card form should be completed and faxed back to BD at 1-866-514-2113 or e-mail bd7043@stericycle.com. All affected product should be returned with the completed Recall Response Card following the steps on the enclosed packing instruction. Upon receipt of the returned product, BD will send customers replacement product. If customers have questions or require assistance, they should contact 1-866-800-2917. |  | 315800 | 315,800 units | Distributed to the states of CA, CT, FL, IL, MA, ND, NY, OH and PA. | Terminated | May 25, 2017 | 826 | 315800 | Not_Computer
Z-1320-2015 | 70355 | No | OWD |  |  | Molecular Genetics | PMA | cobasÂż EGFR Mutation Test |  |  | January 13, 2015 | 2015 | Roche Molecular Systems, Inc. | False positive results for Exon 20 insertion mutations are being detected with the cobasÂż EGFR MutationTest, kit batch T08661. | Under Investigation by firm | RMS notified their customers on 1/13/2015. |  | 190 | 190 kits | Foreign | Open, Classified |  |  | 190 | Not_Computer
Z-1318-2015 | 70736 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | DUAFIT(R) PIP 0o CA SIZE 2, REF A60 SP002, STERILE R, Rx only. |  |  | March 09, 2015 | 2015 | In2Bones, SAS | Dimensional non-conformity leading to the impossibility to introduce the implant in the dedicated holder. | Nonconforming Material/Component | The firm notified their sole consignee in the US by email on 03/09/2015 requesting the return of the product, and to notify his consignee(s) to quarantine the unused product. |  | 35 | 35 units | Distributed in the state of TN. | Open, Classified |  |  | 50 | Not_Computer
Z-1321-2015 | 70410 | No | JIX | Calibrator. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry Systems Drug Calibrator I; For in vitro diagnostic use in the calibration of Phenobarbital_2 (PHNB_2), Phenytoin_2 (PHNY_2), and Theophylline_2 (THEO_2) methods on the ADVIA Chemistry systems. |  |  | January 19, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Incorrect calibrator lot values were assigned for ADVIA Chemistry Systems Drug Calibrator I, lot 3JD018, for the ADVIA Chemistry Systems Phenytoin 2 (PHNY_2) and Phenobarbital 2 (PHNB_2) assays which was the result of calibrator value miscalculation. | Device Design | Siemens issued an Urgent Medical Device Correction, on/about January 19, 2015, notifying customers about the product, problem, and action to be taken. Customers were instructed to update the calibrator values for Drug Calibrator 1, Lot 3JD018. All customerswere requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within (30) days. Field service personnel have been sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. Customers with questions were instructed to contact their local Siemens technical support representative. For questions regarding this recall call 914-631-8000. |  | 657 | 657 | Worldwide Distribution - US (nationwide) and Internationally to Canada and Mexico. | Terminated | September 21, 2015 | 245 | 657 | Not_Computer
Z-1324-2015 | 70576 | No | OVE | Intervertebral body fusion device. | Orthopedic | Orthopedic | 510(k) | Endo TCS Driver Straight Hexalobe #6 P/N 5210-1004, Non-Sterile, Manufactured by Titan Spine, LLC. |  |  | February 20, 2015 | 2015 | TITAN SPINE, LLC | The TCS Straight Hexlobular #6 Drivers (5210-1004) tip sheared off during use in procedures that involved patients with dense bones. As a result of this malfunction, Titan Spine, LLC is removing all of the TCS Straight Hex lobular #6 Drivers (5210-1004) Lot# G150101 from the field. | Nonconforming Material/Component | Titan Spine sent an "Urgent Medical Device Recall" letter via Fed Ex dated February 20, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.A replacement TCS Straight Hex lobular# 6 Drivers (5210-1004) will be shipped out to each Titan Spine Area Vice President and Sales Rep. within 4-6 weeks.For further questions please call (262) 242-7801. |  | 26 | 26 | US Distribution to the states of : AZ, CA, CT, FL, GA, ID, IL, KY, MD, MO, NV, NY, OH, OK, OR, TX, UT and WI. | Terminated | April 17, 2015 | 56 | 26 | Not_Computer
Z-1327-2015 | 70599 | No | HGY | Nonpowered breast pump. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | Ameda Dual Short Term HygieniKit Milk Collections System (SKU 17115) - Primary packaging is 12 pouches with Tyvek header (10.75 x 16.375 x 3") within a secondary corrugated shipper (20.75 x 15.50 x 8.5" inner dimensions). |  |  | February 12, 2015 | 2015 | Ameda, Inc. | Devices were not sterilized | Under Investigation by firm | Ameda sent an URGENT: MEDICAL DEVICE RECALL letter dated February 12, 2015 to all affected customers. The letter included instructions for distributors to: 1) quarantine the recalled products; 2) complete and return the enclosed Recall Response Forms within 3 days of receiving the letter; 3) upon receipt of the completed Recall Response Form, an Ameda representative would contact the distributor to make arrangements for the return of the recalled products; 4) Ameda will use the information provided in the returned Recall Response Form to contact the distributor's customers of record to arrange for the return of recalled product that was already shipped to the customers of record; and, 5) replacement product will be shipped to both distributors and their customers of record upon receipt of the completed Recall Response Form. Questions about this recall will be answered by Ameda personnel who can be reached at 847-964-2645. |  | 12 | 12 collection systems | Nationwide Distribution including CA, FL, GA, IA, IL, IN, LA, MA, MD, MO, MS, NC, NJ, NY, OK, PA, RI, SC, TN, TX, UT, VA, and WA. | Terminated | July 21, 2017 | 890 | 1775 | Not_Computer
Z-1323-2015 | 70594 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Coblator II System, 120V. |  |  | February 20, 2015 | 2015 | ArthroCare Medical Corporation | The recalling firm was notified by the contract manufacturer that they had reversed the L2, L3, L4, or L5 components on the P/N09010 RoHS boards of the Coblator II Controller (Model RF8000E). | Packaging process control | ArthroCare sent an Urgent Medical Device Recall letter dated February 2015 to customers. A response form will be included in the letter and a follow up notifcation letter will be sent to non-responding consignees. The letter identified the affected product, problem and actions to be taken. For questions contact your local Returns Department Representative at 800-343-5717, option 3. |  | 15 | 15 units | US Nationwide Distribution in the states of TN, IA, AZ, NE, SD, NY, and OK. | Terminated | February 26, 2016 | 371 | 15 | Not_Computer
Z-1325-2015 | 70605 | No | MBH | Knee joint patellofemorotibial metal/polymer porous-coated uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen TRABECULAR METAL FEMORAL AUGMENT BLOCK DISTAL ONLY, 20 MM Augment, Size D, Sterile, Rx Only. |  |  | February 04, 2015 | 2015 | Zimmer Trabecular Metal Technology, Inc. | One unit of the 20 mm augments was mislabeled as 5 mm and it was subsequently distributed. | Error in labeling | Zimmer Trabecular Metal Technology sent an email dated 12/10/2014 to their sole customer. The customer was requested to return the affected product. |  | 1 | 1 unit | US Distribution in the state of CA | Terminated | July 16, 2015 | 162 | 1 | Not_Computer
Z-1336-2015 | 70632 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Boston Scientific: Expel" Drainage Catheter with Twist-Loc" Hub APDL Drainage Catheter and APDL Drainage Catheter Kit, Expel" Drainage Catheter APD Drainage Catheter, Expel" Large Capacity Drainage Catheter APD Large Capacity Drainage Catheter. |  |  | February 25, 2015 | 2015 | Boston Scientific Corporation | Boston Scientific is initiating this Medical Device Field Correction of the Expel" APD and APDLDrainage Catheters. To date, Boston Scientific has received one complaint for device fragmentationafter an Expel" APD catheter was implanted in the biliary system. The most serious foreseeable patientrisk for this issue is additional intervention for fragment retrieval using minimally invasive meth | Nonconforming Material/Component | Boston Scientific sent an "Urgent Medical Device Correction" letter dated February 25, 2015, to all affected customers. The letter described the problem and the product involved in the recall. The letter informed customers not to use the Expel APD and APDL Drainage catheters for bile drainage and to complete and return the Customer Acknowledgement Form. The letter also noted that the devices are not being retrieved and they are not required to return them to Boston Scientific. For further questions please call (763) 494-7971. |  | 918 | 918 | Worldwide Distribution - US Distribution to the states of : CA, LA, FL, WI, NY, IL, MD and OH., and to the countries of: Belgium, Czech Republic, Finland, France, Germany, Great Britain, Italy, Spain, Switzerland, Austria, Hungary, Netherlands and Portugal. | Open, Classified |  |  | 918 | Not_Computer
Z-1348-2015 | 70721 | No | FYD | Air-handling apparatus for a surgical operating room. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | LASER RESISTANT WAND - STERILE, 7/8" X 8" (2.22cm x 20.32cm), REF 24013, QTY 10 EA, STERILE EO, NO LATEX |  |  | March 10, 2015 | 2015 | Microtek Medical Inc | This recall is being conducted because a cohesive band used to package the devices contains latex and the product is not labeled as required by 21 CFR 801.437; it does not contain Caution: The Packaging of This Product Contains Natural Rubber Latex Which May Cause Allergic Reactions. | Error in labeling | The firm, Ecolab - Microtek Medical, sent an "Urgent: Recall - Mislabeled Product Immediate Action Required" letter dated March 2015 to its US customers on 03/10/2015 via letter and email. The letter describes product, problem and actions to be taken. The customers were instructed to immediately examine your inventory and promptly quarantine the lot(s) listed; complete and return the enclosed Recall Response Form via fax to Quality Assurance Manager, Ecolab-Microtek Medical at (662) 244-3011 or email to: recall@ecolab.com RE: Recall - Mislabeled Product ASAP, and if you have further distributed the products subject to this recall, please notify your customers at once. Upon receipt of the completed response forms, Ecolab-Microtek Medical will issue a return material authorization (RMA) and ask that you return any unopened product listed above to Ecolab-Microtek Medical at the address listed in the attached form.If you have any questions in regard to this recall, please contact Customer Service by phone at 1-800-824-3027 or email; customerservice@microtekmed.com. |  | 6540 | 6540 units | US (nationwide) Distribution. | Open, Classified |  |  | 145758 | Not_Computer
Z-1350-2015 | 70704 | No | DRX | Electrocardiograph electrode. | Cardiovascular | Cardiovascular | 510(K) Exempt | Skintact ECG Electrode FS-50, Catalog no. 58028 , &amp; Stable Base ECG Electrode SBT-60, Catalog no. 50547. |  |  | February 24, 2015 | 2015 | Leonhard Lang Medizintechnik GmbH | The electrodes may fail to transmit signals. | Nonconforming Material/Component | Consignees were sent via e-mail a Leonhard Lang "Important Safety Notice" letter, dated 24 February 2015. The letter provided the Product trade name, Affected lot numbers, Required Action, Target audience, Details of the affected product which included Description of the problem and Measures to be taken by the user. The letter advised consignees to quarantine the affected product, and to complete and return the acknowledgment form to Leonhard Lang. For questions contact either Sophie Hurmann, Sales Department (tel: +43 512 33425 - 45, fax: +43 512 39 22 10, e-mail: sophia.hurmann@leonhardlang.at) , Stefan Rudig, Quality Department ,(tel: +43 512 33425 -15, fax: +43 512 39 22 10 e-mail: stefan.rudig@leonhardlang.at), or Bernhard Ladner,Qualitats Manager (tel: +43 512 33425- 39, fax: +43 512 39 22 10, e-mail: bernhard.ladner@leonhardlang.at) . |  | 364800 | 364,800 | Distributed to FL. | Terminated | July 26, 2017 | 883 | 364800 | Not_Computer
Z-1353-2015 | 70609 | No | LMH |  |  | General &amp; Plastic Surgery | PMA | Vitagel; indicated in surgical procedures (other than in neurosurgical and ophthalmic) as an adjunct to hemostasis when control of bleeding by ligature or conventional procedures is ineffective or impractical. |  |  | February 26, 2015 | 2015 | Orthovita, Inc., dBA Stryker Orthobiologics. | Several shipments of Vitagel product (a surgical hemostat) required to be stored at 2 -8 degrees Celsius (C) were not delivered to customers within the validated delivery time. | Storage | Orthovita, Inc., dBA Stryker Orthobiiologics made phone notifications in February 2015 to all affected customers. Customers were told that the Vitagel they ordered on x date has an internal procedural mandate that says they must deliver in 35 hours and it has now passed that window. As a precaution until we can confirm that Vitagel delivered after shipping for longer is safe for use we are asking that they quarantine the order in question in the event the delivery occurs.For questions regarding this recall call 610-640-1775. |  | 274 | 274 | Nationwide Distribution including NY PA WV MI IL MA TX AR KS FL WI CA MO LA VA IA OH and NC. | Terminated | July 28, 2015 | 152 | 274 | Not_Computer
Z-1354-2015 | 70596 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.plaza. A Picture Archiving and Communication System (PACS) intended to display, process, read, report, communicate, distribute, store and archive digital medical images. It supports the physician in diagnosis and treatment planning. |  |  | February 18, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Possibly incomplete archived studies during pre-fetch. In a server farm setup, when pre-fetch/retrieve operation is performed for partially archived studies, the series that have not yet been archived, will remain unarchived. | Software design | Safety Advisory Notice, dated February 18, 2015, was sent to end users that identified the product, problem, and action to be taken. Users were advised not to manually remove the delete protection in the STS. |  | 4 | 4 | US Distribution: MA, WI, FL | Open, Classified |  |  | 4 | Not_Computer
Z-1351-2015 | 70472 | Yes | MSD | Blood access device and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Medcomp Duo-Flow 400XL Catheter, Hemo-Flow Catheter, SLX Catheter, Triple Lumen Infusion Catheter, CT rated PICC, Split Cath III, PC Split Cath III; |  |  | February 05, 2015 | 2015 | Medical Components, Inc dba MedComp | Medcomp has initiated the recall of Duo-Flow 400XL Catheter, Hemo-Flow Catheter, SLX Catheter, Triple Lumen Infusion Catheter, CT rated PICC, Split Cath III, PC Split Cath III because the product shipped was out of specification. One BI (biological indicator) out of 21 on the sterilization load was out of specification. | Process control | On or about 02/03/2015, Medcomp sales representatives notified the affected medical facilities and providing Medcomp's recall letter. This was accomplished by e-mail, telephone, or visit. Distributors for the international accounts were notified by Medcomp via e-mail. A copy of Medcomp's recall letter was attached. For AngioDynamics product- Medcomp notified AngioDynamics by e-mail of the product recall. Every consignee was requested to complete and return Page 2 of the recall letter verifying receipt of the recall and disposition of the product they received. AngioDynamics was requested to provide an account of the disposition for the total number of product they received. |  | 3152 | 3152 | Worldwide Distribution. | Terminated | August 06, 2015 | 182 | 3152 | Not_Computer
Z-1352-2015 | 70579 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Synthes 130 Degree Aiming Arm and 125 Degree Aiming Arm; The 130 Degree Aiming Arm and 125 Degree Aiming Arm are a part of the TFN-ADVANCED (TFNA) System. The TFN-ADVANCED Proximal Femoral Nailing System is intended for treatment of fractures in adults and adolescents (12-21) in which the growth plates have fused. |  |  | February 24, 2015 | 2015 | Synthes, Inc. | The guide sleeve has the potential to jam in the Aiming Arms, which may delay disassembly and removal of instruments. | Component design/selection | DePuy Synthes sent an Urgent Notice: Medical Device Recall letter, dated March 2, 2015, to affected customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to review inventory for the affected parts and following the actions provided in the letter for removal and replacement of these devices. For questions call 610-719-5450. |  | 131 | 131 | US Nationwide Distribution - NY, TX, MD, OR, MS, ND, CA, AZ, MI, IL, PA, TN, AR, SD, SC, FL, VA, OK, and GA | Terminated | September 21, 2015 | 209 | 131 | Not_Computer
Z-1356-2015 | 70763 | No | DTE | External pacemaker pulse generator. | Cardiovascular | Cardiovascular | 510(k) | Medtronic Dual Chamber Temporary External Pacemaker, Model 5392. |  |  | March 17, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | Possible performance issue when used with specific AA-sized (LR6) batteries. The negative terminal of a commercially available AA (LR6) battery did not maintain a sufficient connection with the battery drawer electrical contact. This issue could prevent the EPG from powering on or cause the EPG to abruptly lose primary battery power, potentially stopping delivery of pacing therapy. | Device Design | Medtronic sent an "Urgent Medical Device Correction" letter dated March 2015 to affected customers. The letter was addressed to Risk Manager or Health Care Professional. The letter described the Issue and Customer Actions to be taken. The letter requested customers to complete and return the Customer Confirmation Certificate. For questions contact your Medtronic representative, for technical questions call Instrument Technical Service at 800-638-1991. |  | 6040 | 6040 devices | Worldwide Distribution - US Nationwide in the states of: AL, AK, AZ, AR, CA, CO, CT, DC, FL, GA, ID, IL, IN, KS, KY, LA, MA, MI, MM, MS, MO, MT, NE NV,N, NM, NY, NC, ND, OH,OK, PA, RI, SC, TN, TX, UT, VT, VA, WA, WV, WI, WY, PR, VI and the countries of:Australia, Austria, Bahamas, Belgium, Bolivia ,Botswana, Brazil ,Brunei, Darussalam, Bulgaria, Burkina Faso, Canada, Colombia, Croatia, Cyprus, Czech Republic, Denmark,, Dominican Republic, Egypt, France, Georgia, Germany, Greece, Guyana, Honduras, Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Jordan, Kazakhstan, Korea, Kuwait, Lebanon, Luxembourg, Macedonia, Malaysia, Mauritius, Morocco, Mozambique, Namibia, Nepal, Netherlands, New Zealand, Norway, Oman, Pakistan, Papua, New Guinea, Peru, Philippines, Poland, Portugal, Qatar, Saudi Arabia, Serbia, Singapore, Slovakia, Slovenia, South Africa, Spain, Sri Lanka, St. Pierre and Miquelon,Sweden, Switzerland, Tanzania, Thailand, Trinidad and Tobago, Tunisia, United Arab Emirates, United Kingdom, Uruguay, Uzbekistan, Venezuela, Viet Nam, and Yemen.. | Terminated | October 08, 2015 | 205 | 6040 | Not_Computer
Z-1357-2015 | 70692 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | RECLAIM Assembled Implant Inserter Adaptor orthopedic manual surgical instrument for hip implantation. |  |  | February 23, 2015 | 2015 | DePuy Orthopaedics, Inc. | This lot of the RECLAIM Assembled Implant Inserter Adaptor Instrument may fracture during surgery and there is the potential for plastic pieces to be left in the patient. | Under Investigation by firm | The firm issued a voluntary device removal on 02/23/2015 for lot no. SO2018074 of the RECLAIMÂż Assembled Implant Inserter Adaptor. The recall notice informs US Distributors and Medical Professionals about the issue and any actions to take. US Distributors will return affected instruments and notify one Hospital/User Facility with the Hospital Notice and Reconciliation Forms. Reconciliation forms are required for the end-user Hospital/Medical Facility Consignee. Customers may contact Mindy K. TinsleyDirector Wordwide CommunicationsTel: (574) 372-7136Fax: (574) 372-7717E-mail: mtinsley@its.jnj.com |  | 54 | 54 of Lot No: SO2018074 | Worldwide Distribution: US (nationwide) in CA; Ireland and Israel. | Terminated | September 28, 2015 | 217 | 54 | Not_Computer
Z-1355-2015 | 70604 | Yes | FGE | Biliary catheter and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Various Lots of Expel Drainage Catheter with TwistLoc Hub Biliary Drainage Catheter and Biliary Drainage Catheter Kits |  |  | February 25, 2015 | 2015 | Boston Scientific Corporation | Complaints of device fragmentation after the catheter was implanted in the biliary system. The most serious patient risk for this issue is additional intervention for fragment retrieval using minimally invasive methods and, in some of the reported cases, this was performed to retrieve the device fragments. | Device Design | Boston Scientific sent an Urgent Medical Device Correction - Immediate Action Required letter dated February 5, 2015. to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter described the problem and the product involved in the recall. Advised consignees to complete and return the "Recall Removal Reply Verification Tracking Form".Your local Sales Representative can answer any questions you may have regarding this Field Correction notice.For further questions please call (763) 494- 7971. |  | 289 | 289 | Worldwide Distribution - US Distribution to the states of : CA, WI, FL and OH., and to the countries of: Belgium, Czech Republic, Finland, France, Germany, Great Britain, Italy, Spain, Switzerland, Austria, Hungary, Netherlands and Portugal. | Terminated | June 30, 2017 | 856 | 289 | Not_Computer
Z-1359-2015 | 70683 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Computed Tomography X-ray Systems (Brilliance CT 64-channel w/Essence technology, Brilliance iCT, Brilliance iCT SP, Ingenuity Core, Ingenuity Core128 &amp; Ingenuity CT), Philips Medical Systems, Cleveland, OH |  |  | February 05, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | Philips discovered that a software defect exists in marketed product wherein the sign indication of the longitudinal position of some types of scan is inverted. | Software design | Philips sent an Customer Information letter dated February 5, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. They informed their customers that they should refrain from using Multi Planar Reconstruction (MPR) derived images for identifying the absolute position of the intervention for feet-first patient orientation. A Philips Field Service Engineer will contact you to schedule the software installation at your site.If you need any further information or support concerning this issue, please contact your local Philips representative.Philips apologizes for any inconveniences caused by this problem. For further questions please call (440) 483-7600. |  | 424 | 424 units | Worldwide Distribution - Nationwide Distribution to the states of : AR, AZ, CA, CT, DE, FL, GA, IA, IN, KS, LA, MA, MD, MI, MN, MO, MS, MT, NH, NJ, NM, NV, NY, OH, OR, PA, SC, TN, TX, UT, VT, WV and WY., and to the countries of : Argentina, Australia, Austria, Belgium, Bolivia, Brazil, Canada, Cambodia, China, Colombia, Denmark, Finland, France, Germany, Hong Kong, India, Indonesia, Iran, Iraq, Ireland, Italy, Japan, Latvia, Libyan Arab Jumeirah, Lithuania, Malaysia, Mexico, Myanmar, Netherlands, New Zealand, Norway, Panama, Philippines, Russia, Singapore, Slovakia, South Africa, South Korea, Spain, Sweden, Switzerland, Taiwan, Tanzania, Thailand, Training, Tunisia, Turkey, United Arab Emirates,United Kingdom, Uzbekistan and Viet Nam. | Open, Classified |  |  | 424 | Computer
Z-1361-2015 | 70610 | No | BZA | Airway connector. | Anesthesiology | Anesthesiology | 510(K) Exempt | Double Swivel Connector with Suction Port |  |  | February 17, 2015 | 2015 | King Systems Corp. | Product is mislabeled as Double Swivel Connector with Suction Port and is actually Double Swivel Connector without Suction port. If the incorrect component is not identified prior to clinical use, the incorrect connector would be replaced during the case without risk to the patient and could result in a short delay of the procedure. | Mixed-up of materials/components | Kingsystems sent an Urgent Medical Device Recall letter dated February 17, 2015 to affected customers. The letter identified the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, call your King Systems customer service representative at 800.642.5464. |  | 900 | 900 units | US Nationwide in the states of IL, IN, KY, MI, PA, and TN | Terminated | October 27, 2015 | 252 | 900 | Not_Computer
Z-1358-2015 | 70666 | No | KNX | Urine collector and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(K) Exempt | Bard 100% Latex-Free Urinary Drainage Bag with Anti-Reflux Chamber and Bard EX-Lok Sampling; Product Code 154006: Lot # NGYJ0603 and NGYJ1671 |  |  | January 30, 2015 | 2015 | C.R. Bard, Inc. | Potential breach of the sterile barrier packaging. | Unknown/Undetermined by firm | Bard sent an Urgent Medical Device Recall letter on January 30, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter also requested that a sub-recall should be conducted if the product was further distributed. The letter also included a response form which is to be returned to Bard.This voluntary recall is only for the lots noted above with the affected product catalog number. Notifications of this recall are being sent to all affected accounts of Bard Medical Division, Inc. If you have further distributed this product, you are required to notify your accounts of the voluntary recall. If you have any questions or need assistance in notifying your accounts about the correction, please call 1-770-784-6471. We appreciate your cooperation and assistance in dealing with this matter and sincerely apologize for any inconvenience that may result from this action. |  | 9620 | 9,620 units | Nationwide Distribution to the states of : AL, AZ, CA, FL, GA, IL, IN, KY, LA, MD, MA, MI, MN, MS, MO, NJ, NY, NC, OH, OK, OR, PA, RI, SC, TN, TX, VA, WV and WA. | Open, Classified |  |  | 9620 | Not_Computer
Z-1360-2015 | 70676 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Brand Name: PersonaÂż Anterior Referencing Sizer with Locking Boom |  |  | March 10, 2015 | 2015 | Zimmer, Inc. | Zimmer, Inc is voluntarily recalling 22 lots of PersonaÂż Anterior Referencing Sizer with Locking Boom, part number 42-5099-088-10. Devices are being removed from distribution due to nonconforming to print specifications. | Nonconforming Material/Component | Zimmer sent an Urgent Medical device recall letter to all customers on March 10, 2015, via letter and email. The letter identified the product, the problem, and the action to be taken by the customer. Affected units will be inspected. Units that fail inspection will be physically destroyed/scrapped. Units that pass inspection will be re-etched and redistributed. If after reviewing this notification customers have further questions or concerns they were instructed to call the customer call center at 1-877-946-2761. Hours of operation are Monday through Friday, 8:00 a.m. through 5:00 p.m. EST.For questions regarding this recall call 574-372-4487. |  | 688 | 688 | Worldwide Distribution - US including Domestic: Puerto Rico, AZ, NY. OR, ME, NC, OH, TX, CO, IL, TN, PA, HI, CA, KS, WI, VA, MN, NJ, UT, MA, GA, MI, OK, MS, FL, AL, IN, and Internationally to Germany, India, Dominican Republic, Singapore, Korea, Japan, Austrailia, Canada. | Terminated | November 06, 2015 | 241 | 688 | Not_Computer
Z-1364-2015 | 70667 | No | JDR | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Ziploop Button |  |  | February 26, 2015 | 2015 | Biomet, Inc. | The firm is recalling ToggleLocs and Ziploops manufactured in May and July of 2014 as they were not processed following a standard work process, which resulted in complaints of sterile packaging not being sealed. | Employee error | Biomet sent an Urgent Medical Device Recall Notice letter dated March 4, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to immediate locate and discontinued use of the product and its return to Biomet. Customers may contact Field Action Specialist, Regulatory Compliance,at (574) 372-1570. |  | 50 | 50 Units | Worldwide Distribution: US (nationwide) Distribution in the states of:TX, GA, NM, NC and OK and the countries of: Japan, and the Netherlands. | Terminated | January 12, 2016 | 320 | 50 | Not_Computer
Z-1362-2015 | 70630 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | GAP Nail 4.8mm (diameter size) of the GAP ENDO-EXO MEDULLARY SYSTEM. Catalog number GAP-N48-16; GAP-N48-18; GAP-N48-26; GAP-N48-28. |  |  | February 20, 2015 | 2015 | Pega Medical Inc. | Failures of the 4.8mm of the GAP Endo-Exo Medullary System have been identified at the level of the distal cortical screw holes. Failures mainly occurred in patients with tibial pseudoarthrosis (CPT) or Osteogenesis imperfecta (OI). Higher failure rates have been detected in tibias with distal fractures that are close to the distal locking holes. | Other | Pega Medical sent the Urgent Field Safety Notice-Recall/Advisory Notice letter, dated 2015/02/20, to US consignees (surgeons) via email.The following actions were advised:1. It is recommended to monitor patients as considered necessary in order to ensure that weight bearing does not take place until full consolidation of the fracture/osteotomy is completed and that weight bearing is applied progressively.2. Report to the distributor or Pega Medical any adverse event observed. In cases where the consolidation has occurred, normal patient follow up can be continued with no serious consequences to the patient. In cases where consolidation has not been achieved after six weeks, replacement of the nail might be necessary.3. All surgeons receiving this notice must acknowledge reception by a written confirmation sent to the distributor or to the contact reference person indicated below.US consignees with questions can call Pega Medical at 450-688-5144 ext. 242, or send fax to 450-688-1977, or send email to egarcia@pegamedical.com. |  | 7 | 7 units distributed in the US; | US Nationwide Distribution in the states of: California, Florida, New York, and Puerto Rico (US). | Terminated | April 30, 2015 | 69 | 7 | Not_Computer
Z-1365-2015 | 70846 | No | KRA | Continuous flush catheter. | Cardiovascular | Cardiovascular | 510(k) | Covidien Trellis 6 and Trellis 8 Peripheral Infusion Systems, Sterile EO, Rx only. Models BVT608010V01, BVT608030V01, BVT612010V01, BVT612030V01, CVT808015, CVT808025, CVT812015, CVT812025, EVT808015, EVT808025, EVT812015, EVT812025, BVT808015, BVT808030, BVT812015, BVT812030, EUT808015, EUT808030, EUT812015, EUT812030. |  |  | March 23, 2015 | 2015 | Medtronic Inc. Cardiac Rhythm Disease Management | Medtronic is conducting a voluntary recall of all former Covidien Trellis 6 and Trellis 8 Peripheral Infusion Systems due to potential for a sterility breach of the outer packaging or pouch material. A breach of the outer pouch may compromise the outside surface sterility of the inner pouch and does not directly affect the sterility of the device components within the inner pouch. | Package design/selection | Consignees were sent on 3/23/2015 a Medtronic "Urgent Medical Device Recall" letter dated March, 2015. The letter described the problem and the product involved in the recall. Consignees were advised to quarantine and return the product and to complete and return the Customer Confirmation Certificate to Medtronic. For questions contact your Medtronic Sales Representative or Customer Service a 1-800-716-6700. |  | 8269 | 8,269 ( 6,786 US, 1,483 OUS) | Worldwide Distribution: US (nationwide) including all states in continental USA, except DE and including DC and PR; and countries: AUSTRIA, AUSTRALIA, NEW ZEALAND, BELGIUM, CZECH REPUBLIC, DENMARK, EGYPT, FINLAND, FRANCE, GEORGIA, GERMANY, IRELAND, ITALY, NETHERLANDS, POLAND, PORTUGAL, ROMANIA, RUSSIA, SPAIN, SWEDEN SWITZERLAND, TURKEY, UNITED KINGDOM, CHILE, EL SALVADOR, PANAMA, and CANADA. | Open, Classified |  |  | 8269 | Not_Computer
Z-1376-2015 | 70677 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | GE Healthcare, Aespire 7900, Aespire View Anesthesia Machines. Model numbers 1009-9012-000. |  |  | February 27, 2015 | 2015 | GE Healthcare | Specific Aespire 7900 Anesthesia Machines have an increased likelihood of failure of the CPU circuit board. This could result in a blank display or loss of mechanical ventilation. | Process control | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 34064, dated February 27, 2015. The letter was addressed to Chief of Anesthesia, Health Care Administrator / Risk Manager and Director of Biomedical / Clinical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The letter stated that customers can continue to use the affected device in Manual Ventilation mode, and contact their local GE Healthcare Service Representative. GE Healthcare will correct all affected devices at no cost to the customers. A Service Representative will contact customers to arrange for the correction. For questions or concerns contact GE Healthcare Service at 1-800-437-1171, or your local Service Representative. |  | 46 | 46 (10 units US, 36 units OUS) | Worldwide Distribution -- US, including the states of NC and OK; and the countries of CHINA, INDIA, JAPAN, MEXICO, and VENEZUELA. | Terminated | April 16, 2015 | 48 | 46 | Not_Computer
Z-1375-2015 | 70811 | No | MVU | Analyte specific reagents. | Pathology | Hematology | 510(K) Exempt | CD123 (9F5) PE |  |  | March 23, 2015 | 2015 | BD Biosciences, Systems &amp; Reagents | One lot of BD CD123 PE (ASR) has been determined to contain a low amount of CD4 antibody and may result in an unexpected staining pattern. | Under Investigation by firm | Recall letters were sent to affected sites on March 23, 2015 by certified mail. All product in house was placed on ship hold and remaining inventory is in process of being scrapped. |  | 9 | 9 | US Distribution in states of: MO, GA, AZ, NE, TX, VA | Terminated | February 26, 2016 | 340 | 9 | Not_Computer
Z-1371-2015 | 70820 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences Fem-Flex II Femoral Arterial Cannula 8, 10, 12 French,Sterile EO, Rx Only, Manufacturer Edwards Lifesciences LLC, Irvine, CA.-Model Numbers: FEMII008A, FEMII008AT, FEMII008V, FEMII010A, FEMII010AT, FEMII010V, FEMII012A, FEMII012AT, and FEMII012V. |  |  | March 23, 2015 | 2015 | Edwards Lifesciences, LLC | Edwards Lifesciences is recalling Fem-Flex II Pediatric Femoral Arterial Cannula sizes 8, 10, 12 French because of the potential of tissue damage caused by a protruding wire located at the tip of the cannula. | Process change control | The firm, Edwards, sent an "URGENT - PRODUCT RECALL - ACTION REQUIRED" letter dated March 23, 2015 via FED EX on March 27, 2015 to their customers. The letter described the product, problem, and actions to be taken. The customers were instructed to review entire inventory for the lots listed; complete and return attached acknowledgment form via fax to Edwards Customer Service at 800.422.9329. within three days of receipt of this Field Safety Notice; contact Customer Service at 800.422.3278 to obtain an RGA number and replacement product; return affected product to Edwards Lifesciences, Attn: Cirilo Chaparro, 12050 Lone Peak Drive, Draper, UT 84020, Attention: RECALL, RGA#XXX; and transfer this notice to other organizations if the affected devices have been transferred to any another facilities.If you have any questions that have not been answered by this letter, please call Edwards Customer Services at 800.424.3278 from the hours of 6:00AM - 4:30PM PST; Edwards Customer Service at (800) 268-3993 from 8:00AM  4:30PM Eastern Time or contact your Edwards sales representative concerning the recall. |  | 5125 | 5125 units | Nationwide Distribution. | Terminated | April 17, 2017 | 756 | 5125 | Not_Computer
Z-1374-2015 | 70445 | No | FGY | Introduction/drainage catheter and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Bio-CORE Cannula (Needle) Set |  |  | January 23, 2015 | 2015 | Biologic Therapies, Inc | Devices are misbranded - failure to properly classify and obtain FDA clearance prior distribution | Use error | On January 23, 2015 the firm sent an URGENT: Class II MEDICAL DEVICE RECALL notification asking its consignees to: 1) Quarantine all remaining devices in their inventory and return the product to Biologic Therapies Inc. via UPS Ground. Product returns should be shipped under UPS Account #A197E9 and must reference RMA#15-0115-01. Please ship to the Biologic Therapies Inc. address listed at the bottom of Page 5. To coordinate return shipping, please contact Biologic Therapies, Inc. via the contact information provided.Contact hours: Monday - Friday 8am to 5pm Eastern Standard Time. (352) 304-5149 2) Adverse reactions or quality problems experienced with the use of this product may be reported to the FDAs MedWatch Adverse Event Reporting program either online, by regular mail or by fax. |  |  | 0 (distributed or released) | US Distribution in states of: MA, CA, WA, NJ, TX, IL, and FL. | Terminated | August 06, 2015 | 195 | 2855 | Not_Computer
Z-1369-2015 | 70775 | No | JQP | Calculator/data processing module for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ACCU-CHEK Connect Diabetes Management App; Instruction Manual |  |  | October 30, 2014 | 2014 | Roche Diabetes Care, Inc. | Roche Diabetes Care has become aware of an issue with the Accu-Chek Connect diabetes management app that could potentially lead to inaccurate bolus advice being provided to the user. A thorough investigation of the situation revealed that this issue may occur if the user changes the screen orientation of the phone from portrait to landscape or vise versa while looking at the Bolus Advisor or Carb | Software design | Urgent Field Safety Notices were sent to the customer and healthcare professional starting on 10/30/2014. The letter provided a description of the reason for the notification; and recommended that all users should update their app immediately. |  | 72772 | The application has been downloaded total of 644 times (Italy-219, South Africa-24, Germany-401) and there are 113 bolus advice activations. | There are no US distributors as the affected device is not marketed in the US. The affected device was distributed in Germany, Italy and South Africa. The application has been downloaded total of 644 times (Italy-219, South Africa-24, Germany-401) and there are 113 bolus advice activations. | Terminated | June 08, 2015 | 221 | 72772 | Not_Computer
Z-1367-2015 | 70571 | No | LXS | Nd:YAG laser for posterior capsulotomy and peripheral iridotomy. | Ophthalmic | Ophthalmic | 510(k) | YC-1800 OPHTHALMIC YAG LASER SYSTEM; |  |  | March 06, 2015 | 2015 | Nidek Inc | Dear Doctor letter and Quick Reference Guide is being sent to users of the YC-1800 to reiterate important safety and cleaning information after complaints of pitting lens that may cause unclear images. | Process control | Dear Doctor letters were sent along with a Quick Reference guide to all affected sites in February. |  | 799 | 799 units | Nationwide Distribution. | Terminated | September 03, 2015 | 181 | 799 | Not_Computer
Z-1368-2015 | 70633 | No | HNL | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | LIS27T - Lacrimal Intubation Set RX Only Sterile |  |  | February 26, 2015 | 2015 | Quest Medical, Inc. | Expiration date incorrectly printed in Manufacturing Date field on both pouch and carton labels | Labeling Change Control | The consignees were notified by letter on 3/2/2015. The recalling firm requested the consignees to return the recalled product for a replacement device. |  | 106 | 106 units | Worldwide Distribution: US (nationwide) including DC and the states of IL, KY, TX, CA, FL, MI, GA, MD, NH, OH, NY, UT, NJ, MO, TN, VA, NC, CO, MI, and PA, and the country of Ireland. | Terminated | February 26, 2016 | 365 | 106 | Not_Computer
Z-1366-2015 | 70665 | No | KZH | Syringe needle introducer. | General Hospital | General Hospital | 510(k) | Autoject E1 Fixed needle device, Model Number AJ1310, UPC Code 384701310010 |  |  | November 24, 2014 | 2014 | Owen Mumford USA, Inc. | The syringe carrier is missing components: a damper and spring. | Process control | Recall letters were issued to customers via FedEx. The letter instructed the customer to check their shelves for the product and specific lot number and to visually inspect (pictures included in letter of a correctly assembled device and one with the missing spring &amp; damper) the device if they have the affected lot. A form was included which was to be filed out and returned. |  | 3100 | 3,100 | US nationwide distribution in the states of AL, AR, CA, FL, GA, IL, MO, NC, OK, TN, and TX. | Terminated | July 31, 2017 | 980 | 3100 | Not_Computer
Z-1377-2015 | 70649 | No | MUH | Extraoral source x-ray system. | Dental | Radiology | 510(k) | CS 8100, CATALOG # (s): 5303045, 5311129, 5303060, 5303110 |  |  | January 08, 2015 | 2015 | Carestream Health, Inc. | Units device head descended unexpectedly | Component design/selection | The firm sent customers "Urgent: Medical Device Recall" letters, dated December 29, 2014. The letter described the problem, as well as, the actions to be taken. A Carestream Health service representative will contact customers to schedule a visit to inspect the affected devices. Questions or concerns should be directed towards the Carestream Customer Care Center in the U.S. at 1-800-328-2910. |  | 3110 | 3,110 units (863 domestically &amp; 2,247 internationally) | Worldwide Distribution -- US, including the states of AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NV, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, TN, TX, UT, VA, VT, WA, WI, and WV; and the countries of India and Switzerland. | Open, Classified |  |  | 3110 | Not_Computer
Z-1379-2015 | 70690 | No | JPA | Multipurpose system for in vitro coagulation studies. | Hematology | Hematology | 510(k) | ROTROL P Control for ROTEM delta Thromboelastometry System, Manufactured for: Tem Innovations GmbH. |  |  | October 23, 2014 | 2014 | TEM Systems Inc | Some of vials were found to be partially filled. | Nonconforming Material/Component | Tem sent an Dear Customer letter dated October 23, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. You are entitled to return products delivered to you in the event that they are defective or otherwise not in conformity with your order when you received them. If you observe a low level of diluent in a box of Rotrol P lot number 41819801 at your location, please contact us for a free of charge replacement box from a new lot of Rotrol P.You can reach us at:Email: support@roteminc.comPhone: 919-941-7777, Option 2Toll Free: 1-866-597-1652, Option 2We apologize for any inconvenience you may experience. |  | 7870 | Total 1574 x 5 vials per box | Nationwide Distribution | Terminated | August 13, 2015 | 294 | 7870 | Not_Computer
Z-1380-2015 | 70777 | No | KKX | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | da Vinci S Surgical System IS2000, Arm Drape, 20 Pack. |  |  | March 16, 2015 | 2015 | Intuitive Surgical, Inc. | Cloudy/waxy appearance with potential to transfer waxy substance to the patient, and potential for drape to tear. | Packaging process control | Firm issued field safety notice to all customers on March 16, 2015. Field safety notice includes acknowledgement form to be returned to recalling firm. A second letter was sent on May 12, 2015 to inform users that the date that the drape product unaffected by the issue will not be availalble untisl October/November 2015, not May 2015 as stated in the original communication. |  | 442475 | 442,475 total all drape models | Worldwide Distribution-Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, Chile, China (including Hong Kong), Colombia, Cyprus, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Finland, France, France, Germany, Greece, Iceland, India, Indonesia, Ireland, Israel, Italy, Japan, Kuwait, Lebanon, Luxembourg, Malaysia, Mauritius, Mexico, Monaco, Netherlands, New Zealand, Norway, Pakistan, Panama, Philippines, Poland, Portugal, Puerto Rico, Qatar, Romania, Russia, Saudi Arabia, Singapore, Slovakia, Slovenia, South Africa, South Korea, Spain, Sweden, Switzerland, Switzerland, Taiwan, Thailand, Turkey, United Arab Emirates, United Kingdom, United States, Uruguay, Venezuela, Vietnam. | Terminated | December 03, 2015 | 262 | 442475 | Not_Computer
Z-1394-2015 | 70748 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | INFX-8000V Interventional Angiography System. CAT-880B catheterization table. |  |  | March 11, 2015 | 2015 | Toshiba American Medical Systems Inc | The table base rotational brake that holds the table top may not always be engaged when any force or side impact is applied to the table top. | Device Design | Toshiba sent a Urgent Medical Device Correction letter dated March 11, 2015, to all customers who purchased the CAT-880B catherization table for the Tosiba Infinix X-Ray Interventional System. The letter informed the customers of the problems identified and the actions to be taken. Customers were instructed to complete and return the attached form and fax it to (877) 349-3054 or email to raffairs@tams.com. Customers with questions were instructed to call (800) 421-1968 or contact their local Toshiba Representative at (800) 521-1968.For questions regarding this recall call 714-730-5000. |  | 28 | 28 units | Nationwide Distribution | Terminated | September 18, 2015 | 191 | 28 | Not_Computer
Z-1399-2015 | 70591 | No | JGS | Sodium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | ADVIA Chemistry XPT System, Siemens Material Number 10723034, Software Version 1.0.2, Siemens Material Number 11219493, IVD. |  |  | February 11, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens Healthcare Diagnostics has confirmed an issue with ADVIA Chemistry XPT software version 1.0.2. The system may cause samples to remain in an Inprocess state. Test resultson a sample that is held Inprocess will not transmit to the LIS. Manual intervention is necessary to complete the processing of the samples that are held Inprocess. | Software design | Siemens Healthcare Diagnostics sent Urgent Medical Device Correction Letters (dated 2/11/2015) via FedEx to the domestic consignees informing users of the software issue with ADVIA Chemistry XPT Software Version 1.0.2. The letters included work instructions to handle the bug. Customers are requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within 30 days. For questions about in process samples, the event message 02 981 8728, or need assistance with mitigation solutions, please contact your Siemens Customer Care Center or your local Siemens Technical Support Representative. Urgent Field Safety Notice Letters were sent to the foreign consignees (distribution is determined at the country level). Field service personnel were notified of the Urgent Medical Device Correction Letter. |  | 32 | 16 units: 3 Domestic, 13 Foreign | Distributed in the states of AZ, CA, and WA, and the countries of Germany, Italy, Spain, and UK. | Open, Classified |  |  | 32 | Computer
Z-1396-2015 | 70744 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Instrument Kit for use with OsteoSponge SC grafts. The Instrument Kit is used to insert OsteoSponge SC grafts to fill bone defects in the subchondral region of articulating joints. |  |  | March 09, 2015 | 2015 | Bacterin International, Inc. | Instrument Kits for OsteoSponge SC grafts, OsteoLock Facet Stabilization Dowel, and BacFast HD Facet Dowel have incorrect sterilization parameters in the IFU and on the lid of the kits. Sterilization parameters for the Steam Gravity 121 degree C (250 F) cycle were not validated and the dry time for the Steam Pre-Vacuum 132 degree C (270 F) cycle was omitted. | Process control | The firm, Bacterin, sent a "VOLUNTARY PRODUCT REMOVAL NOTIFICATION" letter and Steps for Voluntary Removal instructions, dated 03/03/2015, to their consignees on 3/9/2015. The letter described the product, problem and actions to be taken.The Consignees were instructed to check with your Inventory and Central Supply Departments to determine if you have the Instrument Kit for use with OsteoLock Facet Stabilization Dowel and BacFast HD Dowel listed; immediately remove all affected product and segregate the product in a secure location for return to Bacterin International, Inc.; complete and return the Tracking Verification Form to Bacterin (even if they do not have affected product to return) to listed address, by email tkarlberg@bacterin.com or fax to 1-406-388-3380; indicate returned products on tracking/verification form; contact the firm via email at tkarlberg@bacterin.com or call 1-406-388-0480 ext. 131 from 8am to 5pm Mountain Time to obtain instructions for returning the Instrument Kits to Bacterin; package the returned goods and send to Bacterin International, Inc. 664 Cruiser Lane, Belgrade, MT 59714.Consignees with further questions about recall, please call Regulator Affairs Manager at 406-388-9480 ext. 128 or send email to cming@bacterin.com |  | 63 | 63 units | US Distribution in states of: Alabama, Arizona, Arkansas, Colorado, Florida, Georgia, Illinois, Indiana, Louisiana, Maryland, Michigan, Minnesota, Mississippi, Missouri, Nebraska, New Jersey, New York, North Carolina, Ohio, Pennsylvania, Puerto Rico, Texas, and Virginia. | Terminated | July 01, 2015 | 114 | 63 | Not_Computer
Z-1398-2015 | 70627 | No | CEM | Potassium test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | QuikLYTE Dilution Check: used to monitor and adjust for the dilution ratio on Dimension clinical chemistry systems with the QuikLYTE module. |  |  | February 24, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | A positive bias beyond 5% resulting in a dilution check failure or an unnecessary dilution check correction if the positive bias is greater than 1% and less than 5%. The dilution check correction would be detected by QC; however, if not detected by QC, the calibration curve would be decreased and cause elevated sodium, potassium and chloride patient results. | Under Investigation by firm | An Urgent Medical Device Recall letter, dated February 2015, was provided to all DimensionÂż system customers who received the affected lot to notify them that Siemens confirmed customer complaints of failing dilution checks when using QuikLYTE Dilution Check lot 4MD707. The letter also instructed them to immediately discontinue use and discard all inventory of Dimension QuikLYTEÂż Dilution Check lot 4MD707. Siemens offered a no charge replacement with a non-impacted lot. |  | 5981 | 5981 | Worldwide Distribution. | Open, Classified |  |  | 5981 | Not_Computer
Z-1397-2015 | 70368 | No | IMD | Hot or cold disposable pack. | Physical Medicine | Physical Medicine | 510(K) Exempt | CooperSurgical TransWarmer Infant Transport Mattress with WarmGel |  |  | May 01, 2013 | 2013 | CooperSurgical, Inc. | Trans Warmer Infant and Transport Mattress not labeled with exp.date may not meet specification | Package design/selection | CooperSurgical sent a formal recall notification letter dated April 4, 2015 to all affected customers. The letter identified the product problem and the action needed to be taken by the customer.CooperSurgical notified consignees affected by this recall via Fedex with confirmed delivery receipt. The letter identified the affected product and detailed the necessary actions to be taken by customers who carried the affected lots in their inventory. CooperSurigcal will replace any of the affected products thatcustomers had on hand. CooperSurgical provided further instruction on the disposition of the product at the time of replacement.Contact CooperSurgical at 203 601 5200 Ext. 3100 if you have questions.Previously CooperSurgical notified consignees affected by this recall on 5/1/13. |  | 4807566 | 4,807,566 | Worldwide Distribution - US Nationwide - in the countries of:ARGENTINA, AUSTRALIA,BAHRAIN, ,BERMUDA,CANADA,COLOMBIA ,ENGLAND,GERMANY,HUNGARY,INDONESIA, KUWAIT,LATVIA,LIBYA,MALAYSIAMEXICO,NEW ZEALAND,SAUDI ARABIA. TURKEY, UNITED ARAB EMIRATES | Open, Classified |  |  | 4807566 | Not_Computer
Z-1400-2015 | 70719 | No | ITX | Diagnostic ultrasonic transducer. | Radiology | Radiology | 510(k) | CIVCO Surgi-Tip Transducer Cover Sterile (15.2 x 244 cm) polyethylene telescopically-folded cover with attached (3 cm) NeoGuard tip, REF 610-833. Each cover is packaged in a Tyvek pouch with accessories of gel and fasteners. Individual Tyvek pouches are packaged as a box of 12. Rx only. |  |  | February 26, 2015 | 2015 | Civco Medical Instruments Inc | Leaks were detected in the recalled lot which may compromise the sterility of the product. | Process design | On 3/12/2015, the Urgent Medical Device Recall notice was emailed to all of the affected customers. The notice identified the affected product and stated that the sterile covers from the affected manufacturing lot were not meeting quality levels and may leak. Customers were requested to return all of surgi tip covers from the production lot identified for replacement. To schedule the return and arrange for replacement product to be sent, customers were asked to contact Ms. Amy King, Customer Service Team Leader at 319-248-6520 or info@civco.com. |  | 120 | 120 pouches | Distributed to the states of CA, IL, MA, MI, MO, and WI. | Open, Classified |  |  | 120 | Not_Computer
Z-1402-2015 | 70645 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Vitek 2 AST-P640 REF 418 579 Gram Positive Susceptibility Card, 20 cards per container, IVD. |  |  | February 25, 2015 | 2015 | Biomerieux Inc | The outer carton label and package insert for the recalled product incorrectly lists Streptococcus pneumoniae as an organism for intended use. | Employee error | Biomerieux, Incs sent an Urgent Product Correction Notice on February 25, 2015, to all affected customers. The letter identified the product, the problem, and the actions to be taken by the customer. Customers were instructed to do the following: " Please confirm this letter has been distributed and reviewed by all appropriate personnel within yourorganization." For the referenced test kits (VITEKÂż 2 ASTP635, ASTP640 or ASTP641), review the claimed organisms inthe Package Insert antibiotic table to confirm which organisms can be tested." Refrain from using the referenced test kits to determine the susceptibility of Streptococcus pneumoniae." Reference your Laboratory Standard Operating Procedures to identify card type used for specific organisms." Please store this letter with your bioMÂżrieux instrument documentation.Please complete the attached Acknowledgement Form and return it to bioMÂżrieux, Inc.Customers with questions were instructed to contact their local BioMerieux representative.For questions regarding this recall call 314-731-8526. |  | 2053 | 2053 cartons | Worldwide Distribution to United Kingdom and Turkey. NO U.S. distribution. | Terminated | March 22, 2017 | 756 | 13967 | Not_Computer
Z-1404-2015 | 70723 | No | KDQ | Vacuum-powered body fluid suction apparatus. | General Hospital | General Hospital | 510(K) Exempt | Pluer-evac Sahara Chest Drainage System, S-1100-08LF, Dry Suction/Dry Seal Control, Rx Only, Teleflex Medical. |  |  | March 13, 2015 | 2015 | Teleflex Medical | Package contains incorrect sized unit: The adult sized S-1100-08LF product contains an infant sized chest drainage system. | Other | Teleflex Medical sent an Urgent Medical Device Recall Notification dated March 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to immediately discontinue use and quarantine any of the affected listed catalog numbers and to return the product. Complete the enclosed acknowledgement form to Teleflex Medical. The form is to be returned even if they have not products.For further questions contact your local sales representative or Customer Service at 1-866 246-6990. |  | 2148 | 2,148 units | US Distribution to the states of : AK, AL, AR, CA, CO, FL, GA, IL, IN, KY, KS, MA, MD, MI, MN, NY, OH, PA, RI, SC, TX, VA WI and WV. | Terminated | August 14, 2017 | 885 | 2148 | Not_Computer
Z-1408-2015 | 70927 | No | OBJ | Diagnostic intravascular catheter. | Cardiovascular | Cardiovascular | 510(k) | VasoNova Vascular Positioning System (VPS ) G4 Console |  |  | March 30, 2015 | 2015 | VasoNova, Inc. | A limited number of VPS G4 Consoles are being recalled due to noncompliance with International Electrotechnical Commission Standard 60101-1-1. | Device Design | Teleflex sent an Urgent Medical Device Recall Notification letters dated on March 30, 2015 to all affected customers. The letters identified the affected product, problem and actions to be taken. The risk to users is identified and users are asked to discontinue use and quarantine affected serial numbers. Teleflex sales reps will help with return of product. The Acknowledgeable form should be completed and returned as per instructions. Questions should be directed to 1.877.236.6869. |  | 17 | 17 | US Nationwide Distribution in the states of AZ, DC, PA, IN, PA. GA, NY, KS, CA, VA, NC. | Terminated | April 11, 2016 | 378 | 17 | Not_Computer
Z-1329-2015 | 70668 | [6 empty] | Ebola GP IgX Blood, Serum, Plasma, Cassette, Catalog No. I-101 |  |  | March 13, 2015 | 2015 | Lusys Laboratories, Inc. | LuSys Laboratories is recalling Ebola Virus One Step Test Kits to stop and prevent further use of these devices. The Ebola Virus One-Step Test Kits have not yet been cleared, approved or authorized by the FDA for diagnostics purposes. | No Marketing Application | LuSys Laboratories sent an Urgent Medical Device Recall letter dated March 13, 2015, to all affected customers informing them that the Ebola Virus One Step Test Kits are not approved by FDA for diagnostics purposes. The letter informs the customers of the problems identified and the actions to be taken. Customers were instructed to discontinue use and return the products to LuSys Laboratories. Customers are instructed to fax back the acknowledgement and Receipt Form to 1-858-866-1688.For questions regarding this recall call 858-546-0902. |  | 1015 | 1,015 units | Worldwide Distribution - US including CA and Internationally to Canada, Sierra Leone, and Denmark. | Terminated | May 31, 2017 | 810 | 2275 | Not_Computer
Z-1413-2015 | 70839 | No | DQY | Percutaneous catheter. | Cardiovascular | Cardiovascular | 510(k) | Bard ConQuest PTA Balloon Dilatation Catheter, 12mm Balloon Size, Product Code: CQ-75124. |  |  | March 23, 2015 | 2015 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular is recalling the Bard ConQuest PTA Balloon Dilatation Catheter because it may be at risk of having deflation related issues. | Process control | Bard sent an Urgent Medical Device Recall Notification letter dated March 23, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter informs the customers that the product code and lot combination listed on the attachment 1 of the recall letter may be at risk of having deflation related issues. Customers are instructed to complete and return the Effectiveness Check Form. Customers with questions are instructed to call Bard Peripheral Vascular's (BPV) at (800) 321-4254 option #2 Ext 2501 (M-F 7am-4pm MST) Customers are instructed to fax the recall and effectiveness check form to BPV at (800) 994-6772. |  | 2620 | 2620 units | Worldwide Distribution - US Distribution including Puerto Rico and the countries of : Austria, Belgium, Germany, France, Greece, Iran, Ireland, Israel, Martinique, Netherlands, Portugal, Russia, Slovenia, South Africa, Spain, United Arab Emirates (UAE), United Kingdom, Canada, Korea, Taiwan, Singapore, Thailand, Chile, Guam, Mexico, Malaysia and Australia. | Terminated | August 28, 2015 | 158 | 2620 | Not_Computer
Z-1412-2015 | 70810 | No | LBI | Neonatal phototherapy unit. | General Hospital | General Hospital | 510(k) | Natus neoBLUE blanket LED Phototherapy system which includes the following items: |  |  | March 16, 2015 | 2015 | Natus Medical Incorporated | neoBLUE blanket LED Phototherapy System is recalled because Natus has received feedback from some customers about early failure of the neoBLUE blanket system fiberoptic pads. These failures involve discoloration and melting of the fiberoptic bundle at the connector that is inserted in the neoBLUE blanket light box. | Other | Natus sent the PRODUCT RELIABILITY ALERT neoBLUE blanket system performance along with an attached Technical Bulletin, Instructions for Use, and Service Manual to consignees via email during the week of March 16, 2015.Natus emailed the Distribution Partner (DP) letter to all Natus InternationalDistribution Partners, requiring that they notify their customers (at Biomedical Department and Neonatal Intensive Care Unit) with copies of the Field Safety Notice, customer letter, the Technical Bulletin, Instructions for Use, and Service Manual.The technical bulletin asks the customer to visually and tactilely inspect the fiber optic cable connector before use, and to call Natus if they notice a failingpad. This procedure will prevent the complete degradation and therapy will not be compromised.U.S. customers should contact Natus Technical Service at 888-496-2887 with any questions regarding this notification. International customers should contact their local distributors or contact Natus Technical Service at (888) 496-2887 or Technical_Service@natus.com. |  | 2611 | 2611 units (1911 units in the US and 700 units outside the US) | distributed worldwide | Open, Classified |  |  | 2611 | Not_Computer
Z-1410-2015 | 70628 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | OSSEOTITE Tapered Certain Implant Rx Only; 3.25 x 18.5 - 15.0 mm. |  |  | November 07, 2014 | 2014 | Biomet 3i, LLC | Endosseous Dental Implants in contact with a residual machine fluid caused the implants to be discolored. | Manufacturing material removal | Biomet 3i sent an URGENT MEDICAL DEVICE RECALL NOTICE dated November 7, 3014 to all affected customers. The letter identified the product, problem, and actions to be taken by the customer. |  | 620 | 620 implants in total | Worldwide Distribution-US (nationwide) including the states of AL, AZ, CA, CO, CT, DE, FL, GA, IL, KY, LA, MA, MD, MI, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, VA, WA and HI and the countries of Austria, Australia, Canada, France, Germany, Italy, Japan, Netherlands, New Zealand, Portugal, Spain, Turkey, Taiwan, and United Kingdom. | Terminated | May 22, 2015 | 196 | 620 | Not_Computer
Z-1414-2015 | 70821 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Replace Select Tapered TiU NP 3.5x10mm, Article No. 29401. |  |  | March 09, 2015 | 2015 | Nobel Biocare Usa Llc | A material used in the packaging of the dental implants may impart a sticky substance on the top surfaces of the implants and in some cases on the first threads. | Nonconforming Material/Component | Nobel Biocare sent an Important - Field Correction letter dated March 6, 2015, to all affected customers and another recall letter dated dated March 10, 2015, to all affected customers to inform them that Nobel Biocare Quality Management Systems observed an issue with a material used in the packaging of their implants. The firm's internal analysis has confirmed that this material is not according to their standards. The letters provide the problems identified and the actions to be taken. Customers were instructed to return the affected product(s) not yet used to Nobel Biocare to receive a replacement product. Customers were also instructed to complete the attached Acknowledgement form and fax it to 714-282-5073 or email it to aeinfo@nobelbiocare.com. Customers with questions were instructed to contact their customer service department.For questions regarding this recall call 714-282-4800. |  | 131 | 131 units | Worldwide Distribution - US (nationwide) and Internationally to Austria, Australia, Belgium, Bulgaria, Canada, Switzerland, Czech Republic, Germany, Denmark, Estonia, Spain, Finland, France, United Kingdom, Greece, Croatia, Hungary, Ireland, Italy, Lithuania, Luxembourg, Mexico, Netherlands, Norway, Poland, Portugal, Romania, Russian Fed., Saudi Arabia, Sweden, Turkey, and United Arab Emirates. | Terminated | December 30, 2015 | 296 | 7594 | Not_Computer
Z-1442-2015 | 70896 | No | DQX | Catheter guide wire. | Cardiovascular | Cardiovascular | 510(k) | ProTrack Microcatheter; |  |  | March 30, 2015 | 2015 | Baylis Medical Corp * | Baylis Medical Company is recalling two lots of ProTrack Microcatheter due to the possibility that the catheter may have circumferential defects along its shaft. | Process control | A recall notification, marked "URGENT" will be sent by Baylis Medical via overnight courier to all customers affected by the recall with a fax-back form. The form is to be faxed back to Baylis Medical confirming receipt of the recall notice. As per instructions on the fax-back form, affected devices are requested to be returned. Baylis Medical will replace the returned affacted catheters with a new catheter. |  | 58 | 58 units | US Nationwide Distribution - US Nationwide in the states of MO, MA, OH, CA, AR, MA, FL, GA, CO, TX, KY, NY, MI, WI, OR | Open, Classified |  |  | 58 | Not_Computer
Z-1441-2015 | 70749 | No | KPY | Personnel protective shield. | Radiology | Radiology | 510(K) Exempt | Kenex Radiation Shield installed with GE Healthcare Interventional, Surgery, CT and X-Ray Imaging systems. |  |  | March 09, 2015 | 2015 | GE Healthcare | GE Healthcare has recently become aware of an improper installation of a ceiling suspended pendent supporting a radiation shield and/or surgical lamp. The pendent system was not installed per specified installation requirements. There has been a reported incident of a radiation shield falling from the ceiling causing a minor injury. | Equipment maintenance | GE Healthcare sent an Urgent Medical Device Correction letter dated March 9, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions in USA contact 1-800-437-1171. For other countries contact your local GE Healthcare Service Representative. |  | 797 | 797 (3 US, 794 OUS). | Worldwide Distribution - US Distribution to the states of : CA and MA., and to the countries of : SPAIN, INDIA, JAPAN, AZERBAIJAN, BULGARIA, BELARUS, SWITZERLAND, CZECH REPUBLIC, GERMANY, EGYPT, FINLAND, FRANCE, UNITED KINGDOM, GUYANA, IRELAND, ISRAEL, ITALY, NIGERIA and NORWAY. | Terminated | December 21, 2015 | 287 | 797 | Not_Computer
Z-1443-2015 | 70822 | No | GDZ | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Bard Peripherally Inserted Central Catheter (PICC) and Dialysis catheter kits containing Safe-Cut Safety Scalpels. |  |  | June 24, 2014 | 2014 | Bard Access Systems | Potential for inadvertent scalpel stick injury involving the safety scalpel. | Component change control | Bard sent an Urgent Notification letter date June 26, 2014 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed on the proper device use and safety mechanism activation of the Safe-Cut Safety Scalpel. For questions call coordinator, Shelly Gilbert, at 1-800-290-1689. |  | 51025 | 51,025 units | Worldwide Distribution - US Nationwide | Terminated | November 05, 2015 | 499 | 51025 | Not_Computer
Z-1444-2015 | 70767 | No | JSI | Selective culture medium. | Microbiology | Microbiology | 510(K) Exempt | Hardy Diagnostics cat no. G327 HardyCHROM SS Agar |  |  | March 04, 2015 | 2015 | Hardy Diagnostics | The firm is recalling HardyCHROM SS, Cat. no. G327 Lot no. 15037 due to a report of possible lack of color development in Shigella organisms in the referenced lot. | Employee error | On 03/12/2015 the firm sent recall notification letters to customers. The firm's letter informs customers of the reason for recall, and requests that they discard any remaining stock of the recalled lot.The firm continues to request that even though the customer has been informed of the necessary actions that they still take a moment to complete, sign, and fax back the enclosed form stating their compliance with the above action to the attention of their Quality Assurance Department at (805) 614-9274. The firm stated that the form is required for their records and need to document the full quantity received, used and/or discarded. For replacements or credit to please contact the Customer Service Department at (800) 266-2222 and select option 1. The firm states that replacements will be sent out at no charge to customers, or a credit will be applied to the invoice.For any additional questions the firm states to contact Technical Services at (800) 266-2222, option 2 or e-mail at techservice@hardydiagnostics.com. |  | 840 | 840 (84 pk/10) | Worldwide Distribution -- U.S., including the states of TX, CA, AZ, WA, PR, and IA; and the countries of Canada and NSW. | Terminated | July 13, 2015 | 131 | 840 | Not_Computer
Z-1446-2015 | 70837 | No | LXT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | LOW PROFILE NON-LOCK SCREW; 2.7mm x 22 mm; REF 131227222; |  |  | March 18, 2015 | 2015 | Biomet, Inc. | A Biomet investigation found that the Low Profile Non-Locking Screw 2.7mmx22mm screw (Part Number: 131227222 Lot: RM105E) was anodized green and should have been anodized gold per the print. | Nonconforming Material/Component | On 3/18/2015, Biomet issued an URGENT MEDICAL DEVICE RECALL NOTICE to affected distributors and one medical facility. Notification included the reason for recall, description of affected product, statement of no expected patient health risk, and actions required of the addressee. Questions related to the recall notification should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 45 | 45 | Nationwide Distribution-including the states of CA, IA, NY, FL, PA, UT, NC, SC, KY, CO, MA, GA, WA, TX, OK, and IL. | Terminated | December 11, 2015 | 268 | 45 | Not_Computer
Z-1445-2015 | 70830 | No | OWB | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | Artis Q/ Q zen systems; Artis Q/ Q zen systems are dedicated angiography systems developed for single and biplane diagnostic imaging and interventional procedures including, but not limited to, pediatric and obese patients. Procedures that can be performed include cardiac angiography, neuro angiography, general angiography, rotational angiography, multipurpose angiography and whole body radiographic/fluoroscopic procedures. |  |  | March 19, 2015 | 2015 | Siemens Medical Solutions USA, Inc | potential issue with the Artis Q/ Q zen systems where it is possible that an electrical connection in the equipment cabinet has not been installed correctly. The potentially defective connection in the equipment cabinet will be inspected and, if necessary, corrected. | Employee error | A customer notification, dated March 19, 2015, was distributed to all affected customer via AX018/15/S. Additionally, Siemens has initiated an inspection of Artis Q/ Q zen systems via Update Instruction AX017/15/S. |  | 139 | 139 | Nationwide | Terminated | August 06, 2015 | 140 | 139 | Not_Computer
Z-1447-2015 | 70733 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Bard Monopty Disposable Core Biopsy Instruments and Kits, Monopty 14g x 16cm (22mm), Product Code: 121416; Monopty Biopsy, Product Code: 000441; Monopty 14g x 10cm (22mm), Product Code: 121410; Monopty 14g x 10cm (11mm), Product Code: 211410. |  |  | March 17, 2015 | 2015 | Bard Peripheral Vascular Inc | Bard Peripheral Vascular is recalling the Bard Monopty Disposable Core Biopsy Instruments and Kits because they may be at risk of having activation-related issues. | Nonconforming Material/Component | The firm, Bard Peripheral Vascular, sent an "URGENT MEDICAL DEVICE RECALL NOTIFICATION" letter dated 3/17/15 to customers. The letter informs customers of the problems identified and the actions to be taken. The customers were instructed to do the following: Do not use or further distribute any affected product; check all inventory locations within your institution; if you have further distributed any of the product, advise them of the recall and have them return the affected product to BPV; remove any idenfied product from your shelves; and complete and return the attached Recall and Effectiveness Check Form to BPV by fax at (800) 994-6772, if you cannot FAX the form, call the BPV telephone number provided on the form and report the required information verbally. BPV will provide replacement product for your returned product.Customers are instructed to call BPV's Recall Coordinator at 1-800-321-4254 Option #2 Ext 2988 (M-F 7am to 4pm MST) or email at HUrandy.hopkins@crbard.comUH. |  | 29342 | 29,342 units | Worldwide Distribution: US (nationwide) and countries of: Denmark, Spain, Finland, France, Ireland, Israel, Italy, Netherlands, Poland, Japan and Pakistan. | Terminated | August 03, 2015 | 139 | 29342 | Not_Computer
Z-1448-2015 | 70131 | No | KTO | Immunoglobulin G (Fc fragment specific) immunological test system. | Immunology | Immunology | 510(K) Exempt | SpermMar Test |  |  | December 22, 2014 | 2014 | Vitrolife Inc | The firm is recalling three batches (FP14A09, FP14A10, FP14A11) of SpermMAR IgA test due to new formulation causing the motility of sperm cells to reduce in certain samples when the new bead solution was added. | Component design/selection | FertiPro sent an Urgent: Medical Device Recall letter dated December 22, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The customers were instructed to check if they have stock of the concerned lots and provide a confirmation that they have destroyed the affected products.Distributors were asked to inform their customers and instruct them to do the same.For further questions you may call (858) 824-0888. |  | 45 | 45 kits | US Distribution to the state of : OR., and Internationally to Argentina. | Terminated | May 05, 2015 | 134 | 45 | Not_Computer
Z-1449-2015 | 70847 | No | DXX | Steerable catheter control system. | Cardiovascular | Cardiovascular | 510(k) | Magellan Robotic System; Model No(s). 11132 and 11139. A steerable catheter control system used at clinical sites. |  |  | March 18, 2015 | 2015 | Hansen Medical Inc | Based on investigation of two complaint incidents, all Hansen Medical Magellan Robotic System (Model No. 11132 and 11139) may experience an unexpected event during a retraction operation of catheter. Retraction of catheter may not stop when equipment button is released. | Software design | Hansen Medical sent an Urgent Notice - Field Safety Notice dated March 24, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. We request that you review the information in pages 1 and 2 with Hansen Medical personnel and acknowledge your understanding by signing the Field Safety Notice Training Form on page 3.If you need any further information regarding this issue, please do not hesitate to contact me via telephone at 650-404-2784. |  | 19 | 19 total installed systems | Worldwide Distribution - US Distribution and the countries of : Saudi Arabia, Australia, France, Italy, Germany, and United Kingdom. | Terminated | September 23, 2015 | 189 | 19 | Not_Computer
Z-1451-2015 | 70658 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Master Drug Library Software version 8.0, Product Code 35723V080, to be used with SIGMA Spectrum Infusion System (Pump) version 8.0, Product Code 35700BAX2 |  |  | March 02, 2015 | 2015 | Baxter Healthcare Corp | Loading/Bolus default dose settings in the Master Drug Library and the values shown on the pump during programming may differ. MDL drug dose time in seconds will round to the nearest integer in minutes on the pump dose setup screen (20 sec may show as 1 min on the pump display). The pump will administer drugs as configured. The discrepancy may cause therapy delay or unintended rate of delivery. | Under Investigation by firm | The firm, Baxter, sent an "Urgent Device Correction" letter dated 3/05/2015 with MDL Instructions were sent to all affected users via USPS First Class Mail on 3/05/2015. All four (4) limited launch locations with Master Drug Library (MDL) version 8.0 with SIGMA Spectrum Infusion System version 8.0 were notified via phone call on 3/02/2015. The letter describes the product, problem and actions to be taken. Baxter is requesting customers take the following actions: 1. Follow the steps described in Attachment 1 to confirm that this discrepancy is properly corrected. 2. Complete the enclosed customer reply form and return it to Baxter by faxing it to 224-270-5457 or scanning and e-mailing it to fca@baxter.com. --- Baxter sent a follow-up Communication Letter (dated 3/17/2015) to the initial Urgent Device Correction Letter (dated 3/5/2015) to provide additional clarification and instructions to ensure continued mitigation of this discrepancy as users implement the Spectrum infusion System in their facility. The four launch sites received the follow-up letters via hand delivery; remaining customers received the follow-up letter via hand delivery and/or e-mail.If you have additional questions, please contact your Baxter sales representative, Baxter Healthcare Medina at 800-356-3454 (choose option 1) Monday through Friday, during the hours of 7:00 am to 7:00 pm Eastern Time, or The Center for One Baxter at 800-422-9837, Monday through Friday during the hours of 8:00 am to 5:00 pm Central Time. |  | 58 | 58 MDLs | Worldwide Distribution: US Distribution to states of: CA, CT, FL, IA, IL, IN, LA, MA, MD, MI, MN, MO, NC, NE, NH, NY, SD, TX, UT, WV, including Puerto Rico and country of: Canada. | Open, Classified |  |  | 58 | Not_Computer
Z-1452-2015 | 70758 | No | KGN |  |  | General &amp; Plastic Surgery | 510(k) | IntegraÂż Flowable Wound Matrix |  |  | March 13, 2015 | 2015 | Integra LifeSciences Corp. | A lot of Flowable Wound Matrix that has been distributed to customers in late 2013 / early 2014 was verified as not meeting stability acceptance criteria for the attributes of visual appearance during dispensing and force vs. displacement (force needed for product to be extruded from the syringe). | Unknown/Undetermined by firm | Integra Life Sciences Inc. sent a Recall letter/Recall acknowledgement and return form dated March 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Integra asks customers to review inventory and determine if any of the affected inventory is in possession. The attached Recall Acknowledgement and Return Form should be completed and returned. Customer service will contact the customer upon receipt of the form and provide an RMA number. Questions can be directed to Customer Service at 1-888-301-0203. |  | 287 | 287 packages ( 179 US 108 OUS ) | Worldwide Distribution - US Distribution and to the countries of and to the countries of Greece, France, Italy, Spain, Switzerland, South America, Turkey and United Kingdom. | Terminated | April 13, 2016 | 397 | 287 | Not_Computer
Z-1453-2015 | 70872 | No | OCZ | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Talon Grasping Device, 160 cm, US endoscopy. Used for retrieval of foreign bodies, tissue specimens, stones or calculi in endoscopic procedures of the gastrointestinal tract. |  |  | March 10, 2015 | 2015 | US Endoscopy Group Inc | A wire component on the distal grasping assembly of the device became detached.. | Under Investigation by firm | On 03/25/2015 the firm sent recall notification letters to their customers via email. |  | 1695 | 1695 units | Product was shipped to the following states: AL, AR, AZ, CA, CO, CT, DC, FL, GA, IA, ID, IL, IN, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI &amp; WV. Product was also sent to the following countries: ARGENTINA, AUSTRALIA, BELGIUM, CANADA, ESTONIA, FRANCE, MEXICO, NEW ZEALAND, SAUDI ARABIA &amp; SWEDEN. | Open, Classified |  |  | 1695 | Not_Computer
Z-1450-2015 | 70829 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | WALLACH LL 100 cryosurgical units: LL100 Cryosurgical - Part Number 900001; LL110 with Pinned Yoke - Part Number 900019-70; LL100 Multi Tip W/TC - Part Number: 900629; LL100 - CO2 - Part Number 900162 |  |  | March 09, 2015 | 2015 | CooperSurgical, Inc. | Internal gas line may malfunction causing the line to burst under pressure. | Under Investigation by firm | CooperSurgical sent an Recall letter dated March 13, 2015, shipped via Fedex with confirmed delivery receipt. The letter identified the product the problem and the action needed to be taken by the customer. No adverse events or injuries have been reported to date in connection to this recall. As stated above, the risks associated with the use of this product involve the potential for a hose to burst while under pressure. The potential hazard associated with the failure mode is negligible; there is no foreseen adverse event due to this failure mode. The event was discovered internally during an investigation. Please complete the attached Acknowledgement and Receipt Form to make repair arrangements. If you have any further questions please feel free to contact me at 203.601.5200 Monday through Friday 8:00am to 5:00pm EST. |  | 255 | 255 units | Worldwide Distribution - Nationwide Distribution to the states of : AK, AL, AR, AZ, CA, CO, DC, FL, GA, HI, IL, IN, KY, LA, MA, MD, MI, MO, MS, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA and WV,. and to the countries of : CANADA, COLOMBIA, EL SALVADOR, ENGLAND, GUATEMALA, HONG KONG, INDONESIA, SINGAPORE and ZIMBABWE. | Terminated | February 01, 2016 | 329 | 255 | Not_Computer
Z-1454-2015 | 70856 | No | OYP | John Cunningham Virus serological reagents. | Microbiology | Microbiology | 510(k) | STRATIFY JCV DxSelect, Model No. EL1950. Assay for detection of antibodies to JC Virus in human serum or plasma. In Vitro diagnostic. |  |  | March 15, 2015 | 2015 | Focus Diagnostics Inc | Focus Diagnostics is recalling the Stratify JCV Dx Select due to the use of a non-conforming batch in the manufacture of JCV DxSelect EL1950 kit lot #27333. This could lead to a potential for false positive JCV results. | Nonconforming Material/Component | Focus Diagnostics sent an Urgent Device Customer Notification letter dated March 15, 2015, was sent to the customer to inform them that Focus Diagnostics is recalling the Stratify JCV DxSelect (EL1950) lot#27333 due to the potential for false positive JCV results. The letter informs the customer of the issues and the actions to be taken. The customer is instructed to complete the enclosed acknowledgement form and email the form to DxTS@focusdx.com or fax back to Focus Diagnostics Technical Services at (562) 240-6526 within 10 business days. If the customer has any questions or require additional information, then they are instructed to contact Focus Diagnostic's Technical Services department at (800) 838-4548, select option 3, between the hours of 7am to 5pm (PST) or send an email to DxTS@focusdx.com. |  | 300 | 300 units | Internationally: Denmark only. | Terminated | July 23, 2015 | 130 | 300 | Not_Computer
Z-1455-2015 | 70739 | No | MQV | Resorbable calcium salt bone void filler device. | Orthopedic | Orthopedic | 510(k) | DBX Mix, 20 cc, Product Code: 058200. Demineralized Bone Matrix Mix, bone void filler. |  |  | December 04, 2014 | 2014 | Musculoskeletal Transplant Foundation, Inc. | The seal on the DBX Mix outer PETG/Aclar tray with foil lid packaging for lot identified above may not be fully intact. | Packaging process control | MTF notified their customers on 12/4/2014 and 12/8/2014 via phone and letter. |  | 25 | 25 units | US Nationwide distribution to CA, KY, IL, LA, PA, NC, TN, WV, and Puerto Rico. | Open, Classified |  |  | 25 | Not_Computer
Z-1459-2015 | 70722 | No | KWS | Shoulder joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Equinoxe 44mm Short Humeral Head |  |  | February 13, 2015 | 2015 | Exactech, Inc. | Incorrectly packaged. Outer and inner labeling may incorrectly identify the size of the enclosed device. | Labeling mix-ups | Exactech Agents sent a Product Recall Advisory Notice dated February 12, 2015 by mailed to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete the Recall Inventory Respponse form and return it to Exactech, per the instructions provided on the notice. |  | 112 | 112 devices | Worldwide Distribution - US Nationwide in the states of AK, AZ, CA, CO, FL, ID, ME, MN, NC, OH, OK, PA, RI, TN, VA and WA. and countries of Brazil and Canada. | Open, Classified |  |  | 112 | Not_Computer
Z-1458-2015 | 70754 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Spacelabs Pediatric Flow Sensor Kit, PN: 376-0561-00. This kit is for Spacelabs Healthcare Blease 700/900 Series Ventilators. |  |  | March 05, 2015 | 2015 | Del Mar Reynolds Medical, Ltd. | Reports of inaccurate low flow readings. Monitored inspiratory tidal volume (VTi) and expiratory tidal volume (Vte) measurements from the pediatric flow sensor are reporting out of specification low compared to the actual delivered volumes being administered to the patient. | Vendor change control | Spacelabs started mailing the Urgent Medical Device Correction- Pediatric Flow Sensor kit, PN: 376-0561-00 used with Spacelabs 700 and 900 series ventilators letter, dated 17 March 2015, via priority service to all US customers.Spacelabs will be emailing the customer letter (translated as necessary) to all international subsidiaries and distributors of record on March 20, 2015.Customers are asked to dispose of all pediatric flow sensors purchased from Spacelabs (even if they shipped outside of the time frame in which affected product shipped). Customers are requested to complete and return the Business Reply Form to Spacelabs Healthcare either by fax to Global Tech Support at 425-363-5758or scan the form and email to techsupport@spacelabs.com Spacelabs will send you at no charge replacement pediatric flow sensors (376-0561-01) for all Spacelabs pediatric flow sensors you purchased. These NEW sensors will be marked with a "-01" molded into the body so you can tell the difference between the old sensors and the new ones.For additional information or technical assistance, please contact:Technical SupportSpacelabs Healthcare, Inc.35301 SE Center StreetSnoqualmie, WA 980651-800-522-7025 and select 2 for Technical Support |  | 1040 | 1040 units total (398 in the US and 642 international) | Worldwide Distribution-US including the states of Louisiana, North Carolina, and Washington and the countries of Argentina, Brazil, China, Colombia, Ecuador, Finland, Great Britain, India, Libya, Mexico, Nicaragua, Oman, Philippines, Saudi Arabia, Slovakia, Sri Lanka, United Arab Emirates, and Vietnam. | Terminated | November 16, 2015 | 256 | 1040 | Not_Computer
Z-1464-2015 | 70773 | No | MLM | Tacrolimus test system. | Clinical Chemistry | Toxicology | 510(k) | DiaSorin, PRO-Trac II Tacrolimus ELISA kit, REF 32400 and 32400-CN. |  |  | March 16, 2015 | 2015 | Diasorin Inc. | Instability in the PRO-Trac II Tacrolimus ELISA kit which may cause falsely elevated patient results. | Under Investigation by firm | DiaSorin sent a Customer Notification Letter" to affected customers. The letter identifed the affected product, problem and actions to be taken. Customers were asked to complete and return the confirmation form to the manufacturer. For questions contact DiaSorin Inc. Product Support at 1-800-328-1482. |  | 451 | 451 kits | Worldwide Distribution - US in the state of CA and the countries of: CHINA, BELGIUM, ARGENTINA, and INDIA. | Terminated | September 15, 2015 | 183 | 451 | Not_Computer
Z-1457-2015 | 70835 | No | CHI | Lipase test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | ADVIA Chemistry Lipase Reagent , Catalog Number B01-4840-01 REF: 01984894. |  |  | March 23, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Siemens internal investigation confirmed that current contamination avoidance settings do not prevent carryover resulting in falsely elevated ADVIA Chemistry Lipase outlier results when running with Triglycerides, Triglycerides_2, Triglycerides_c, Cholesterol concentrated, and DLDL reagents. | Device Design | Siemens issued an Urgent Medical Device Correction on March 23, 2015, notifying direct consignees about the product, problem, and action to be taken. Customers were instructed to follow the directions provided in the customer letter in the event that a Lipase outlier is observed. All customers were requested to complete and return the Effectiveness Check questionnaire attached to the Urgent Medical Device Correction letter within (30) days. Field service personnel have been sent an SKB (Service Knowledge Base) describing the issue and instructing them how to deal with customer questions. For further questions please call (312) 275-7795. |  | 7814 | 7814 | Worldwide Distribution - US (nationwide) Distribution and to the countries of : Thailand, Pakistan, Vatican, Guadeloupe, Reunion, French Polynesia , Serbia, Martinique, India, China, Singapore, Brazil, South Africa, Mexico, Canada, Israel, Australia, Egypt, Russian Federation, Indonesia, Rep. Korea (S) ,Unit.Arab Emir., Taiwan ,Argentina ,Chile, Malaysia ,Colombia, Peru, Kazakhstan, Paraguay, Vietnam, Algeria, Austria, Belgium, Switzerland, Cyprus, Czech Republic, Germany, Estonia, Spain, Finland, France, United Kingdom, Greece, Hungary, Italy, Lithuania, Luxembourg, Latvia, Netherlands, Norway, Poland, Portugal, Romania, Slovenia, Slovakia and Turkey. | Terminated | December 18, 2015 | 270 | 7814 | Not_Computer
Z-1465-2015 | 70832 | No | HXC | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Lineum HF Torque Handle Limit (Lineum) Model Number 14-525035 |  |  | March 12, 2015 | 2015 | Biomet Spine, LLC | Biomet is recalling the Lineum HF Torque Handle Limit (Lineum) and Torque Limiting Handle (Aspen/ Alpine) due to the supplier withdrawing it from the market for component updates or reconfiguration. | Component design/selection | Biomet sent an Urgent Medical Device Recall Notice dated March 12, 2015 via phone and customer letters. The letter identified the affected product, problem and actions to be taken. Customers were asked to follow the instructions on the enclosed "Response Form". For questions call 858-232-2672. |  | 148 | 148 | Worldwide Distribution - US Nationwide | Terminated | September 16, 2015 | 188 | 1022 | Not_Computer
Z-1463-2015 | 69996 | No | HIG | Hysteroscopic insufflator. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Fluid Management System P4000; for irrigation and fluid warming in laparoscopic procedures, and distention, fluid warming, and volume/deficit measurements in endoscopic procedures within gynecology, urology, and orthopedic procedures. |  |  | September 01, 2014 | 2014 | Thermedx LLC | To correct software bugs that could affect the ability to accurately measure fluid deficit. | Software design | Thermedx sent an Urgent Product Correction letter dated March 13. 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm informed their customers that the software update will be performed at customer locations by Thermedx trained and authorized personnel. A Thermedx and/or your Sales Representative will contact customers to schedule a time to perform the field service update.Should you have any questions regarding this notification, please call a Thermedx technical service representative. Support is available toll free at 1-888-542-9276 from 8AM to 5PM Monday to Friday Eastern Standard Time or by e-mailing techservice@thermedx.com. Our website may be found at www.thermedx.com. We sincerely apologize for the inconvenience and thank you in advance for your cooperation. |  | 127 | 127 units | US Distribution to the states of : OH, MS, MA, WA, MI, NC, NY, WV, LA, TX and IL. | Terminated | September 14, 2015 | 378 | 127 | Computer
Z-1467-2015 | 70369 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 6" Trifurcated Extension Set, Product Code: AMS-375-1 |  |  | January 15, 2015 | 2015 | Churchill Medical Systems, Inc. | Potential for leaking caused by insufficient bond between needleless device and female luer. | Process design | Vygon sent an Urgent Medical Device Recall letter dated January 15, 2015, to all affected customers via email or courier service. The letter instructs consignees to check stock for products included in the scope of the recall, count and document the affected inventory on the "Recall Acknowledgement and Inventory Return Form" and fax or email the form to Vygon. Vygon customer service will then contact the customer to issue a Return Merchandise Authorization Number so the inventory can be replaced at no cost. Distributors were provided with instructions on how to handle product that was further distributed by them. If after 30-days, product is not returned from each consignee, Vygon will follow-up with the consignee to obtain a status of the product return.Customers with questions were instructed to contact Vygon Customer Service at 1-800-473-5414, or by e-mail at customerservice@vygonus.com. |  | 5200 | 5,200 | Nationwide Distribution including TX, OH, NY, CA, UT, PA, MA, NY, VA, and IL. | Terminated | February 25, 2016 | 406 | 158850 | Not_Computer
Z-1480-2015 | 70710 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Captiva's lumbar intervertebral fusion system Reamers (color code rings on the Reamers). |  |  | February 25, 2015 | 2015 | Captiva Spine, Inc | It is possible for the titanium color-coded ring to detach from the Reamers. | Device Design | On 02/26/2015 the firm sent a Voluntary Correction Notification to its consignees via FedEx. Within that notification the firm asks to please check with your product representative and Central Supply Departments to determine if you have any of the products listed on the notificaiton. If any product is located please call 877-772-5571 from 8:30am (Eastern Time) to receive instructions as to how to remove the color-coded ring OR please follow the enclosed steps for the correction in their entirety. They contain information about returning the removed rings. Enclosed is Form 30.4 Product Tracking and Verification Form. You must complete and return this form even if you do not have any rings to remove and return. Please contact your Captiva Spine customer service representative if you have any questions regarding this correction, any of our products, or would like assistance with the correction. |  | 233 | 233 devices | GA, IN, FL, UT, AZ, CA, and NY | Terminated | June 01, 2015 | 96 | 233 | Not_Computer
Z-1479-2015 | 70488 | No | FTY | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | SpineFrontier MISquito Percutaneous Pedicle Screw System Surgical Technique; Part Number: DOC80023. The Rod Caliper is a manual surgical instrument that may be used as part of the PedFuse Pedicle Screw System and its MISquito components during percutaneous, minimally invasive surgical (MIS) technique of posterior pedicle screw fixation. |  |  | March 05, 2015 | 2015 | SpineFrontier, Inc. | The PedFuse Pedicle Screw Rod Caliper included with the MISquito Pedicle Screw System may be used incorrectly as the Surgical Technique does not adequately convey the intended use of the design. | Labeling design | SpineFrontier sent an Urgent Advisory Notice letter to customers via email. The letter identified the affected product, problem and actions to be taken. Customers were asked complete and return the Acknowledgement Form to confirm receipt of the recall notification. For questions call 978-232-3990. |  | 5 | 5 devices | US Nationwide Distribution in the states of TX, KS, CO, and MA | Terminated | June 15, 2015 | 102 | 5 | Not_Computer
Z-1481-2015 | 70948 | No | HGS | Intrauterine pressure monitor and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | HSG Trays/HS 5F Catheter Combo; HSG Procedure Tray+ Shapeable 5F HSG catheter. For hysterosonography and hysterosalpingography, gynecological use. |  |  | April 07, 2015 | 2015 | Catheter Research, Inc | HSG Trays and Catheters were mislabeled. Trays with the flexible HS catheter were labeled as containing the shapeable, stiffer HS catheter, and vice versa. | Packaging process control | Catheter Research, Inc. (CRI) has initiated a voluntary medical device removal of HSG Tray/ 5F HS Catheter Combo, and issued letters to customer son 04/07/2015 by e-mail and certified mail. Customers are asked to inspect their inventory and return recalled devices to CRI for replacement. If the devices were further distributed CRI requested customers provide notice of the recall. Customers may contact Babacar Diouf, VP of RA and QS, at (317) 872-0074 ext. 3512. |  | 880 | 880 units | Domestic: MN, CA, OH, IL, FL, MA, WI, VAForeign: NoneDOD/VA: None | Terminated | July 17, 2015 | 101 | 880 | Not_Computer
Z-1512-2015 | 70774 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Xpress System Pedicle Screw Assembly, 5.5 x 55mm. An internal fixation device for spinal surgery consisting of rods, pedicle screw assemblies (screw, cap, yoke) and cross bar connector assemblies. |  |  | March 11, 2015 | 2015 | X Spine Systems Inc | One lot of instruments have etched on the Yoke of the Pedicle Screw a laser marking size of 7.5 x 55mm instead of 5.5 x 55mm as described on the product label. | Nonconforming Material/Component | The firm, X-spine Systems, Inc., sent an "Urgent Medical Device Recall" letter dated March 19, 2015 to their customers. The letter describes the product, problem and actions to be taken. The customers were instructed to notify affected hospital and surgeon of the recall immediately; complete the attached Medical Device Recall Return Response Form and return it with any affected product in inventory to X-Spine Systems immediately via mail: 725 Alexandersville Road, Miamisburg OH 45342-RMA #5288; and No action should be taken with product that has been implanted.Please contact customer service at 800.903.0640 or email orders@x-spine.com with any questions. |  | 14 | 14 units | US Distribution to states of: KS, NV &amp; TX. | Terminated | October 21, 2015 | 224 | 14 | Not_Computer
Z-1501-2015 | 70878 | No | LKD | Cutaneous carbon dioxide (PcCO2) monitor. | Anesthesiology | Anesthesiology | 510(k) | Sentec, V-Sign Sensor 2 (REF VS-A/P/N), Reusable / Nonsterile / Waterproof, SenTec AG, CH-4106, Therwil, Switzerland. The SenTec Digital Monitoring System (SDMS) is used for continuous and non-invasive monitoring of transcutaneous carbon dioxide tension (tcPCO2), oxygen saturation (SpO2), |  |  | March 25, 2015 | 2015 | SenTec AG | SenTec is voluntarily replacing V-Sign Sensors 2 due to pre-mature deterioration of a part of the PCO2 measuring unit and thus disabling PCO2-monitoring.This means that the affected sensors reach their end-of-life earlier than normally expected requiring the sensor`s replacement.There is no increased risk for the patient or operator, because the aforementioned deterioration results in the sens | Under Investigation by firm | Sentec sent an Important Information letter via email dated March 25, 2015. The letter described the problem and the product involved in the recall. Customers were instructed to locate the affected sensors at their end-customers and replace them with the new sensors. Replaced sensors are to be returned to SenTec. For questions contact SenTec at + 41 61 726 97 60.Consignee was sent an emai in regard to the recall being expanded. The email was sent on May 08, 2014. The email informed the consignee of the additional lot being affected and provided a list of the serial number of V-Sign Sensors to be replaced. |  | 188 | 129 US, 59 OUS | US Distribution to the state of : MO | Terminated | February 05, 2016 | 317 | 188 | Not_Computer
Z-1486-2015 | 70859 | No | HRS | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | CROSSCHECK(R) Plating system, LAPIDUS CX Plate, 1mm Step, Left, REF CCP-LPX1L. |  |  | March 06, 2015 | 2015 | Wright Medical Technology, Inc. | The minor diameter, major diameter, thread pitch, and pitch on the Crosscheck are out of specification, which may result in intraoperative malfunction | Nonconforming Material/Component | Firm has initiated the recall on 03/06/2015 through written communication via Fed Ex P1 delivery to all Wright Medical Technologies distributors (hubs). Follow-up to non-responding consignees will commence 30 days after initiation. Effectiveness checks will be performed through the confirmation of returned product via a validated ERP system, a returned/signed verification form, and in the case of non-responders, a phone call, e-mail or follow up letter, depending on the region and availability of consignee. |  | 102 | 102 units | Nationwide Distribution. | Open, Classified |  |  | 3271 | Not_Computer
Z-1502-2015 | 70787 | No | ECN | Orthodontic appliance and accessories. | Dental | Dental | 510(K) Exempt | Herbst Universal Nut which is used in a Herbst Appliance |  |  | March 09, 2015 | 2015 | Specialty Appliance Works, Inc. | The Universal Nut, a component of the orthodontic Herbst Appliance was made with 416 SS and not the specified 303 SS. | Under Investigation by firm | Specialty Applicances sent an Urgent Voluntary Productionn Field Correction Notification dated March 9, 2015, to all affected customers who received the Herbst Appliances and for sales supply orders. The letter advised doctors to monitor the appliance during routine follow-up appointments for signs of corrosion, and for sales supply orders to return affected components to the lab for replacement. Customers were asked to complete and return the enclosed Field Correctivve Action Response Form and return via fax to 678-513-7345. To return affected devices call 800-522-4636 to receive instruction as to how to return it to Speciality Applicances. In the event that customers observe signs of corrosion or a fracture, they should contact customer service at 800-522-4636 or their sales representative. For questions customers should call 800-522-4636 For questions regarding this recall call 678-513-4408, ext 226. |  | 14428 | 14,428 Universal Nuts | Worldwide Distribution - US nationwide including Puerto Rico, Brazil, China, Colombia, Germany, Ireland, Mexico, Switzerland, and United Kingdom. | Terminated | January 12, 2017 | 675 | 14428 | Not_Computer
Z-1515-2015 | 70679 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | OptiMedica Catalys Precision Laser System; Model No(s). Catalys-U (US), and Catalys-I (Outside US). |  |  | February 20, 2015 | 2015 | Abbott Medical Optics, Inc. | Device may experience loss of suction during treatment resulting in scoring the cornea during lens fragmentation. | Device Design | Abbott Medical Optics (AMO) sent an Advisory Notice letter dated March 5, 2015 via FedEx overnight delivery to each Catalys Precision Laser System owner. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete and return the attached Acknowledgement Form. For questions contact AMO Technical Support at 1-855-662-0911. |  | 202 | 202 | Worldwide Distribution - US Nationwide and the countries of, Canada, Germany, Switzerland, Australia, Spain, Japan, Great Britain, France, Netherlands, Austria, Singapore, South Korea, India, Thailand, Colombia, Ecuador, Italy, Brazil, Egypt, Tunsia, Belgium, Mexico, Greece, Guatemala, Chile, Turkey, Peru, Hong Kong, Saudi Arabia, Bangladesh, and Dominican Rep. | Open, Classified |  |  | 202 | Not_Computer
Z-1513-2015 | 70851 | No | GEH | Cryosurgical unit and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | CSA Medical truFreeze Console- Cryosurgical Unit Cryogenic Surgical Device |  |  | March 25, 2015 | 2015 | CSA Medical | TruFreeze Console caused a higher rate of liquid nitrogen (cryogen) to be delivered and may cause: stricture,scarring, bradycardia, or pneumothorax | Software design | CSA Medical issued letter dated 3/25/15 advising users of the problem. Users provided with:The mitigations available to the active venting procedures coupled with the extremely unlikely probability of injury, rare risk, thus allowing the physician to continue with active venting procedures.In regards a passive venting procedure may not have sufficient mitigation to allow the physician to identify the potential hazard with sufficient time to preclude a potential for injury. Therefore, passive users will be instructed not to use the system until a software improvement is put into place. Additionally, no catheters or consoles will be shipped to passive venting users of truFreeze.A response form to be signed and returned confirming receipt of the notification.Questions contact: Stephen Mascioli, MD 781-538-4755 smascioli@csamedical.com. |  | 82 | 82 units | Nationwide | Terminated | August 21, 2015 | 149 | 82 | Computer
Z-1456-2015 | 70988 | No | DWF | Cardiopulmonary bypass vascular catheter, cannula, or tubing. | Cardiovascular | Cardiovascular | 510(k) | OriGen BIOMEDICAL VV13F Reinforced Dual Lumen Catheter STERILE DISPOSABLE: DO NOT RE-USE Manufactured by: OriGen Biomedical. |  |  | March 30, 2015 | 2015 | OriGen Biomedical, Inc. | Potential for clear extension tube to separate from the hub, which could result in blood loss and may require intervention to prevent permanent impairment/damage. | Mixed-up of materials/components | The firm, OriGen, sent an "URGENT; MEDICAL DEVICE RECALL" letter dated 3/30/2015 and response form to all consignees via e-mail. The conisgnees were notified on 3/31/15 and 4/1/15. The recalling firm also posted a recall notification on the ELSO website, which is a public site mainly used by professionals within the ECMO community and press release dated 4/15/2015. The letter described the product, problem and actions to be taken. The consignees were instructed to discontinue use of all OriGen W13F Reinforced Dual Lumen ECMO Catheters from the lot; immediately return any unused product to OriGen Biomedical; and complete and return the attached acknowledgement and receipt form even if you do not currently have product in your inventory. The firm recommends that the consignee use an alternate product in the interim. If you have any questions regarding the return of product that you have on hand or the status of the investigation, please do not hesitate to conact the company at: OriGen Biomedical, Attn: Director, Quality and Regulatory Affairs, 7000 Burleson Rd, Bldg D, Austin, TX 78744 or call +1 512 474 7278 CST Monday-Friday 8:00-5:00. |  | 51 | 51 units | Worldwide Distribution: US (nationwide) and countries of: Spain, UK, Germany, Sweden, and Netherlands. | Terminated | May 03, 2016 | 400 | 51 | Not_Computer
Z-1514-2015 | 70702 | No | KQM | Surgical camera and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Sony Medical Monitor LMD-1951MD 19 inch |  |  | March 05, 2015 | 2015 | Sony Latin America | Monitor has either not turned on or has lost image during clinical use. | Nonconforming Material/Component | SOLA sent an URGENT MEDICAL DEVICE FIELD ACTION dated March 2, 2015, to all affected customers. Within that notification SOLA requests that you send the customer letter included in Attachment 1 of this notification to all customers who have purchased the affected monitors from you. The letter instructs customers who already possess an AC adapter for each of their affected monitors to only use the AC adaptor when the monitors are in use. If you have units in inventory that can be provided with an AC adapter, we will provide under separate cover kits that contain: 1) a sticker warning and 2) an adhesive seal to go over the AC outlet and 3) correction instructions document. If you have units in inventory that cannot be accompanied by an AC adapter, for each of these affected monitors, do not ship the unit until the permanent corrective action described below has been applied. Please complete the MEDICAL DEVICE RECALL RETURN RESPONSE included in Attachment 2 to the contact person identified on the form. Please contact me below if you have any questions. 305-260-4000 Hours of operation: 8am to 8pm (Monday through Friday, excluding holidays). |  | 4764 | 4,764 units | Worldwide Distribution - US Distribution to the states of : FL and NC., and to the countries of : Ecuador, Costa Rica, Brazil, Nicaragua, Venezuela, Peru, Colombia, Dominican Republic, and South Korea. | Open, Classified |  |  | 4764 | Not_Computer
Z-1516-2015 | 70838 | No | MGQ |  |  | General &amp; Plastic Surgery | 510(k) | Iamin Moist Dressing and the Iamin Hydrating Gel: intended to cover a wound or burn on the skin, provide a moist wound environment and protect against abrasion, friction, desiccation and contamination. |  |  | March 26, 2015 | 2015 | PhotoMedex, Inc. | The labeling for the Iamin family of medical devices (e.g. printed and internet text) is not consistent with the intended uses and limitations identified at the time of the clearance to market these devices. | Labeling design | PhotoMedex sent an Urgent product Recall notification, dated March 2015, to direct consignees requesting they discontinue use, complete the response form, and return affected products. |  | 200188 | 200188 | Nationwide and Internationally, including Canada. | Terminated | August 18, 2015 | 145 | 200188 | Not_Computer
Z-1461-2015 | 70853 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | TIGERPAW System II, |  |  | March 25, 2015 | 2015 | Laax, Inc. | Increase of field reports involving issues with the TIGERPAW System II resulting in possible tissue tearing on left atrial wall and bleeding during use of the device, which may or may not be caused by incomplete closure of the TigerPaw System II fastener. | Unknown/Undetermined by firm | The firm, LAAx, Inc.on behalf of Maquet Medical Systems, sent an "URGENT MEDICAL DEVICE RECALL (Removal) IMMEDIATE ACTION REQUIRED " letter dated March 30, 2015 to all affected sites by express mail. The letter describers the product, problem and actions to be taken. The customers were instructed to examine your inventory immediately to determine if you have any TIGERPAW System II products; remove the products, quarantine them and place in a secure location; complete and return the MEDICAL DEVICE RECALL (Removal) RESPONSE FORM by e-mailing a scanned copy to:Tigerpaw2015@maquet.com or FAX to: 1-(973)-396-3607. If you have acted as a distributor for the TIGERPAW System product II, immediately contact those accounts, advise them of the recall situationa and have them return all stock to you.Your Marquet Medical System USA Sales Representative will contact you to assist with the return of any product.If you have any additional questions, contact your local Maquet Medical Systems USA Sales Representative or our Customer Service at 1-888-880-2874 (6:00 AM - 5:00 PM PST). |  | 4154 | ~ 4,154 total units in US - both model numbers | Worldwide Distribution: US (nationwide) and countries of:Australia, Bahrain, Belgium, Czech Republic, Denmark, Finland, France, Germany, Hong Kong, Ireland, Israel, Italy Netherland, Norway, Poland, Qatar, Slovenia, Spain, Saudi Arabia, Sweden, Switzerland, Turkey and United Kingdom. | Terminated | March 17, 2017 | 723 | 4154 | Not_Computer
Z-1517-2015 | 64225 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | The PCA 3 infusion pump allows clinicians, or patients, to (self) Administer analgesia safely and effectively within clinician programmed limits. |  |  | February 27, 2013 | 2013 | Hospira Inc. | The infusion pump has a clear plastic door that, due to its design, allows for excessive flexing. This in turn permits unauthorized access to the vial of medication. | Device Design | Hospira will contract with Stericycle Inc., in Indianapolis, IN to send letters and a reply form to Hospira customers. Receipt of the letters will be confirmed by the UPS tracking. Customers will be requested to return a reply form to Stericycle indicating the number of pumps on hand and that they have received the letter and notified the users.For questions regarding this recall call 800-441-4100. |  | 8214 | 8,214 units | Worldwide Distribution - US including AK, AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, WY, District of Columbia, Puerto Rico, American Virgin Islands; and Internationally to Canada. | Open, Classified |  |  | 47096 | Not_Computer
Z-1519-2015 | 70863 | No | GEX | Laser surgical instrument for use in general and plastic surgery and in dermatology. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Light Sheer Desire Diode Laser System with XC Handpiece Accessory options. |  |  | March 25, 2015 | 2015 | Lumenis Limited | Device software treatment preset parameters for the XC treatment handpiece do not match the Operator Manual, and exceed recommended settings. Operator Manual parameters are lower than indicated for specific hair color and Fitzpatrick skin type resulting in insufficient treatment effect. May result in patient burns and hypopigmentation. | Software design | Lumenis Limited sent an Urgent Safety Advisory Notice dated March 25, 2015, to all affected customers. Customers were instructed that Lumenis is eliminating the potential risk of patient harm through a software upgrade and Operator Manual replacement. The software upgrade must be completed by a Lumenis-certified engineer as soon as possible. In the Interim, customers can continue to use their LightSheer DESIRE system by following the recommendations in the operator manual to always perform a test patch prior to a full treatment. Customers with questions were instructed to contact the Aesthetic Business Unit Product Manager.For questions regarding this recall call 801-656-2690. |  | 117 | 117 units | Worldwide distribution. US nationwide (to OR, NM, NY, IL, and CO), Switzerland, Canada, Italy, Great Britain, United Arab Emirates, France, Belgium, India, Israel, Japan, Portugal, Czech Republic, Kuwait, Germany, Spain, Georgia, Armenia, Slovenia, New Zealand, Saudi Arabia, and Russian Federation. | Terminated | June 15, 2016 | 448 | 117 | Computer
Z-1520-2015 | 70888 | exerciser, non-measuring | ION | Nonmeasuring exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | AmTryke Model 1424 Community Cruiser Hand Cycle. Therapeutic tricycle. |  |  | March 27, 2015 | 2015 | Amtryke LLC | Product is being recalled due to complaints of cracked frames. | Under Investigation by firm | Consignees were notified via letter on 03/27/2015. |  | 53 | 53 units | US nationwide distribution to AL, CA, FL, HI, IL, KY, KS, LA, MI, NY, OK, OR, PA, SC, TX, VA, VT and WV. | Terminated | July 10, 2017 | 836 | 53 | Not_Computer
Z-1521-2015 | 70685 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Software Version 2.8 &amp; Below on VITROS 5,1 FS Chemistry Systems, Catalog Number 6801375, Global Trade Item Number 10758750001132, and VITROS 5,1 FS Chemistry System Refurbished, Catalog Number 6801890, Global Trade Item Number 10758750001644; IVD. |  |  | March 10, 2015 | 2015 | Ortho-Clinical Diagnostics | Software Anomaly; It is possible for the device to process samples with a cartridge other than the intended cartridge, potentially leading to erroneous patient results. If this anomaly occurred previously, an indication would be a series of results that were similar across multiple samples (i.e., results could be believable or outside of the Reportable Range for the intended assay). | Software in the Use Environment | On 3/10/2015, URGENT PRODUCT CORRECTION NOTIFICATION letters (Ref. CL2015-047, dated 3/10/2015) were sent to users notifying them of the issue and advising them of the required actions. Customers are to complete and return the Confirmation of Receipt no later than March 23, 2015. The firm is currently working on a resolution and will issue a follow up notification upon the software's availability. For questions, please contact Customer Technical Services at 1-800-421-3311. |  | 274.1 | VITROS 5.1 System: Domestic - 909; Foreign - 1250; 5,1 Refurbished: Domestic - 66 units, Foreign - 152 | Worldwide Distribution -- USA, Puerto Rico, Australia, Brazil, Canada, Chile, China, Colombia, England, France, Germany, India, Italy, Japan, Mexico, Republic of Panama, Singapore, Spain, and Venezuela. | Open, Classified |  |  | 274.1 | Computer
Z-1522-2015 | 70874 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan -Neg Breakpoint Combo Panel Type 34; |  |  | May 14, 2014 | 2014 | Beckman Coulter, Inc. | Siemens Healthcare Diagnostics has received complaints regarding panel degradation with a portion of the MicroScan Neg Breakpoint Combo 34 (NBP34) panel lot 2014-05-28. | Equipment maintenance | Siemens sent an Urgent Medical Device Recall letter dated May 14, 2014, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The risk to health is explained and users are asked to discontinue use and discard any of the recalled lot still in inventory. No other lot numbers of the NBPC 34 panels are affected and may still be used. The Field Correction Effectiveness Check form should have been completed and returned to Siemens.PLEASE FAX THIS COMPLETED FORM TO 302-631-8467. Contact Siemens Customer Care Center or local Siemens Tech support. For further questions please call (916) 374-3031 |  | 681 | 681 boxes | Worldwide Distribution - Nationwide Distribution including Puerto Rico and to the states of: KS, LA, MD, NY, WV,NM, NH, MO, CT, MA, OK, NC. IA, OH. WI, TX, PA, MS and MN and to the countries of : Canada, Kuwait, South Africa and Turkey. | Terminated | April 27, 2015 | 348 | 681 | Not_Computer
Z-1532-2015 | 70893 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Microbiology Systems: Dried Gram Negative MIC/Combo Panels |  |  | August 08, 2014 | 2014 | Beckman Coulter, Inc. | Technical Support Bulletin issued in February 2013 did not mention performance limitation for Proteus mirabilis and Imp when using the CLSIM100-S20 revised interpretive criteria. | Labeling False and Misleading | The firm, Siemens, sent an "Urgent Medical Device Correction" letter dated August 7, 2014 to affected customers by trackable method. The letter describe product, problem and actions to be taken. The customers were instructed to confirm receipt of the letter; complete the attached "Field Correction Effectiveness Check" form via fax to Siemens Healthcare Diagnostics at 302-631-8467;TSB 190 has been removed from use and archived from the Siemens Document Library, if you have a hardcopy, discard it appropriately;and follow the listed in instructions and retain and forward this letter to those who may have received this product.If you have any questions, please contact your Siemens Customer Care Center; Quality Systems &amp; Compliance Director at 916-374-3031 or email: juntalan@beckman.com or your local Siemens technical support representative. |  | 84249 | Approx ~(216240 boxes U.S. and 84249 boxes OUS) total | Worldwide Distribution: US (Nationwide) and to countries of: Argentina, Bahamas, Brazil, Canada, China, Ecuador, Egypt, Germany, Lithuania, Mexico, Paraguay, Peru Poland, Portugal, Romania, Russian Federation, Singapore, South Africa, Spain, Taiwan, Thailand, Trinidad and Tobago, United Arab Emirates, Uruguay, and Venezuela. | Terminated | April 27, 2015 | 262 | 84249 | Not_Computer
Z-1526-2015 | 70804 | No | LYJ |  |  | Neurology | PMA | VNS Therapy DemiPulse Generator Model 103 and VNS Therapy DemiPulse Duo Generator Model 104; CYBERONICS, INC. |  |  | March 06, 2015 | 2015 | Cyberonics, Inc | The pulse generators have a lower battery longevity than specified in their design requirement as a result of the devices being inadvertently left in a programmed ON state during manufacture. | Nonconforming Material/Component | On 03/6/2015, the recalling firm sent a recall notification letter to their consignees. The letter included an effectiveness card which is to be signed by the consignee and returned to the recalling firm. |  | 13 | 13 units | AR, TX, GA, CA, MI, IN, FL, AZ, NV | Terminated | October 09, 2015 | 217 | 13 | Not_Computer
Z-1525-2015 | 71021 | No | DRF | Electrode recording catheter or electrode recording probe. | Cardiovascular | Cardiovascular | 510(k) | WEBSTER HIS Catheter, 4 Pole Fixed Curve with Auto ID Technology, for mapping of cardiac structure; i.e., stimulation and recording only. |  |  | April 17, 2015 | 2015 | Biosense Webster, Inc. | The Webster HIS catheter, 4Pole Fixed Curve with Auto ID, is not being recognized by the CARTO 3 navigation System. | Software Manufacturing/Software Deployment | A recall letter dated 4/17/15 will be sent to customers by FedEx to inform them that Biosense Webster is recalling Webster HIS catheter, 4Pole Fixed Curve withAuto ID (Catalog No. D1085413). The letter informs the customers of the problems identified and the actions to be taken. Customers with questions related to the recall letter are instructed to contact their Biosense Webster salesrepresentative or call (866) 473-7823, Monday through Friday from 7:00AM to 8:00PM EST. Customers with questions related to the Voluntary Recall Certification Form and the process for returning the form, are instructed contact Elie Ghattas, Field Action Coordinator, at (909) 839-7281. |  | 503 | 503 units | US nationwide distribution. | Terminated | December 10, 2015 | 237 | 503 | Not_Computer
Z-1527-2015 | 70457 | No | LJT | Subcutaneous, implanted, intravascular infusion port and catheter. | General Hospital | General Hospital | 510(k) | 9F Plastic Dual Port; The port systems are totally implantable vascular access devices. They consist of an injection port, a self-sealing membrane and a detachable catheter. The port systems provide repeated vascular access and are applicable for delivery of: medications, intravenous fluids, blood products, parenteral nutrition, and for venous blood sampling. |  |  | January 28, 2015 | 2015 | Medical Components, Inc dba MedComp | The dual port with catheters was not covered under FDA clearance. | PMA | Sales reps were notified by email on/about January 28, 2015, and instructed to notify their medical facilities and provide the recall letter. This will be accomplished by either email, telephone, or visit. Every consignee will be requested to complete and return Page 2 of the recall letter verifying receipt of the recall and disposition of the product they received. |  | 325 | 325 | Worldwide distribution. US nationwide, Canada, Brazil, and United Arab Emirates. | Terminated | January 06, 2017 | 709 | 325 | Not_Computer
Z-1523-2015 | 70993 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | Anatomical Shoulder Handle for Rasp |  |  | March 18, 2015 | 2015 | Zimmer Gmbh | Complaints of difficulties to attach the Anatomical Shoulder Rasps (high resistance) or, once attached, due to difficulties to remove the handle (seizing up of the two components). This has the potential for delay in surgery. | Nonconforming Material/Component | On 3/18/2015, URGENT MEDICAL DEVICE RECALL Notification - LOT SPECIFIC letters were sent to the affected distributors with instructions for locating, quarantining, and returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 30 | 30 units | Distributed to the states of IN, GA, WI, MO, PA, KS, TN, and NY. | Terminated | September 23, 2015 | 189 | 30 | Not_Computer
Z-1536-2015 | 70641 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | The Patient Control Module (PCM) is used in conjunction with a Baxter infusor as a single use device for the control of intermittent bolus doses of medication based on patient demand. |  |  | April 17, 2014 | 2014 | Baxter Healthcare Corp. | Potential for device malfunction resulting in flow when the device should not be flowing | Process control | Baxter Healthcare sent an Urgent Product Recall letter dated April 17, 2014, to the sole US consignee. The letter included instructions for the consignee to: 1) locate and remove all affected product; 2) contact Baxter's Center for Service to arrange for the return and credit; and 3) to complete and return the Customer Reply Form included with the recall notification. Questions about the recall can be answered by calling The Center for One Baxter at 800-422-9837. |  | 7596 | 7,596 patient control modules | Worldwide Distribution - US Distribution to the state of FL., and to the countries of : Belgium, Denmark, Germany, Italy, Japan, Norway, Poland, South Korea, Spain, Sweden and United Arab Emirates. | Terminated | July 21, 2017 | 1191 | 7596 | Not_Computer
Z-1538-2015 | 70960 | No | FAE | Penile rigidity implant. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Genesis Malleable Penile Prosthesis, 9.5 mm, Catalog /Item Number 91-9509SC/519250. |  |  | April 09, 2015 | 2015 | Coloplast Manufacturing US, LLC | Several lots of the 9.5 mm Genesis Malleable Penile Prosthesis were mislabeled. The Retail Box and Patient Labels indicate a size 9 mm, instead of 9.5 mm. The device size is printed directly on the individual prosthesis tray during the packaging process and states the correct size, 9.5 mm. | Error in labeling | Consignees were notified via Field Representatives starting on 4-9-2015 with a Coloplast "Urgent Correction Notification" letter dated April 7, 2015. The letter described the problem and the product involved in the recall. Requested consignees to complete the Return Form if returning inventory. For questions contact Customer Service at 800-258-3476. |  | 11 | 11 (6 USA, 5 OUS) | Worldwide - US Nationwide Distribution in the states of FL, NY, TX and the country of FRANCE | Terminated | September 11, 2015 | 155 | 11 | Not_Computer
Z-1405-2015 | 70924 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | GE Healthcare Automatic Mobile X-Ray (AMX) Series: Brivo XR285amx |  |  | December 19, 2014 | 2014 | GE Healthcare | An issue was identified with the mAs accuracy check performed on the system that could result in inaccuracies in the 10% to 20% range for exposure times longer than 15 msec. | Radiation Control for Health and Safety Act | GE Healthcare planned action:1. The customer notification letter will be distributed to the facilities that have installed the following AMX series X-ray systems: Models Brivo XR285 AMX, Optima XR200 AMX, and Optima XR220 AMX.2. As stated in the customer notification letter, the correction will be provided free of charge and subsequent communication between the firm and the facility will result in a site visit by a GE Healthcare field engineer to bring the X-ray system into compliance.3. The GEHC field engineer will follow Field Modification Instruction 10899 which provides instructions on how to perform a mAs accuracy check and how to determine if the test results pass predetermined acceptance criteria.If you have any questions or concerns regarding this notification, please call theInformation following phone number: 1-800-437-1171. |  | 6 | 6 units US | US Distribution to the states of: AZ, FL, GA, MO and WI. | Terminated | April 29, 2015 | 131 | 6 | Not_Computer
Z-1535-2015 | 70823 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | CD99 |  |  | March 17, 2015 | 2015 | Biocare Medical, LLC | A drop in staining intensity over time has been observed. These lots may not meet their labeled shelf life. | Unknown/Undetermined by firm | Biocare sent a Medical Device Notification letters dated March 17, 2015 to all affected sites. The letters identify the affected product, problem, actions to be taken and the fact that it affects all identified lots shipped between 02/13/13 and / 01/28/15. The letters instructed customers to immediately discontinue further use of the recalled lots, dispose of them according to state and local regulations and complete the response form and returned by April 7, 2015. Credit will be issued upon receipt of the response form. |  | 887296 | 44: 8 of lot 110912, 15 of lot 041913; 17 of lot 120310 and 4 of lot 110314 | Worldwide Distribution - US nationwide in the states of TX, KS, OK, NJ, LA, GA, NC, SC, VA, and the countries of Spain, Austria, Ecuador, Dominican Republic, Vietnam, Kuwait, Hungary and Israel. | Terminated | June 15, 2015 | 90 | 887296 | Not_Computer
Z-1524-2015 | 70904 | No | JWY | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | MicroScan Pos Combo Panel Type 33 (PC33): |  |  | April 08, 2014 | 2014 | Beckman Coulter, Inc. | Siemens investigation confirmed that a portion one lot of panel Pos Combo 33 lot was packaged without desiccants.The substrates on a panel packaged without a desiccant may deteriorate over time due to exposure to additional moisture in the panel pouch. | Process control | Siemens sent an Urgent Medical Device Recall letter dated April 8, 2014, to all affected customers. Letters advised users of the recalled product with catalog number, Siemens Material Number and Lot number as well as the reason for recall. The risk to health was explained as well as the recommendation for discussing the content of the letter with the medical director regarding the need to review previous test results.Customers are asked to discontinue use and discard any remaining inventory of the PC33 panel 2014-09-12. All other lot numbers are not impacted.The attached confirmation receipt is to be completed and returned and the letter should remain with laboratory records. Questions should be directed to Siemens Customer Care Center. For further questions call (916) 374-3031. |  | 664 | 664 panel boxes | Nationwide Distribution including Puerto Rico and to the states of : CT, GA, NC, OK, NY, SC, KY, TN, CA, WA, OR, NJ, OH, PA, FL, MI, TX, WI, MO and NH. | Terminated | April 27, 2015 | 384 | 664 | Not_Computer
Z-1537-2015 | 70836 | No | MYN |  |  | Radiology | PMA | Logicon Caries Detector |  |  | April 06, 2015 | 2015 | GA Industries | GA Industries is recalling the Logicon Caries Detector, version 5.0, for Carestream Dental because it was not submitted for approval to the FDA prior to product release. | Unknown/Undetermined by firm | A Medical Device Correction and Removal Notice dated 3/25/15 will be sent to customers on 4/7/15 to inform them that GA Industries is providing them with a new approved version of Logicon Caries Detector, Version 5.1. The letter informs customers that the original version 5.0 should be uninstalled from their computers and the installation CD be destroyed. The letter informs the customers of the changes and improvements to the product. Customers are instructed to join training sessions conducted by Carestream that demonstrates the new features. Customers are instructed to contact LogiconSupport@Carestream.com for webinar schedule. Customers with questions are instructed to contact GA Industries at logiconsupport@carestream or at (800) 944-6365. |  | 825 | 825 units | US Nationwide Distribution | Terminated | November 23, 2015 | 231 | 825 | Computer
Z-1542-2015 | 71074 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Command Modules, Model 91496-M, upgraded with the Masimo SpO2 Option. |  |  | April 17, 2015 | 2015 | Spacelabs Healthcare Inc | Ultraview SL Command Modules which were upgraded with the Masimo SpO2 Option, Model 91496-M was affected. Customer reported that during the time that the module is resetting (~10 seconds), monitoring of all parameters will be suspended. Upon completion of this sequence, alarms will be reset to their default values. | Software design | Spacelabs Healthcare sent an Urgent - Medical Device Correction letter, dated 17 April 2015, to the US consignees via priority mail. On 4/24/2015, the firm will send the Urgent - Medical Device Correction letter (translated as necessary) via email to international subsidiaries and distributor of record. The Identified the affected product, problem and actions to be taken. Ultraview SL Command Modules which were upgraded with the Masimo SpO2 Option, Model 91496-M are affected. Ultraview SL Command Modules originally shipped with Masimo SpO2 are NOT affected. Spacelabs Healthcare will contact you to schedule a convenient time for Spacelabs to correct all of your facility's affected Command Modules at no cost. For additional information or technical assistance, please contact:Technical SupportSpacelabs Healthcare, Inc.35301 SE Center St.Snoqualmie, WA 980651-800-522-7025 and select 2 for Technical Support |  | 1999 | 1999 units total (1932 in the US and 67 international) | Worldwide Distribution- US Nationwide in the states of AL, DE, GA, ID, IL, KY, LA, MA, ME, MI, MN, MO, NC, NJ, NY, OH, OR, PA, SC, TN, VA, WA and the country of France. | Terminated | November 13, 2015 | 210 | 1999 | Not_Computer
Z-1541-2015 | 70737 | [6 empty] | Portable FAR Infrared Sauna |  |  | December 23, 2014 | 2014 | Therasage, LLC. | Product distributed without 510(k) approval | Labeling False and Misleading | On January 20, 2015 Therasage sent a notification to its consignees asking them to discard any of the labeling materials that state the following: 1) Use of the sauna allows one to "naturally detoxify." 2) Use of the sauna expels toxins from the body 3) Any reference to an exact amount of calories one uses in an hour when enjoying the benefits of this product 4) Although this product is enhanced with a negative ion environment, we cannot say that the product is enhanced with negative ion therapy 5) Any reference to the maximum infrared absorption available cannot be explicitly stated, as it may vary for each individual person. Customers with questions were instructed to call 888-416-4441 or email info@therasage.com.For questions regarding this recall call 561-416-9300. |  | 8 | 8 devices | Nationwide Distribution including SD, IL, TX, OR, NY, WI, GA, and CA. | Terminated | August 06, 2015 | 226 | 8 | Not_Computer
Z-1548-2015 | 70898 | No | KPA | Automated slide stainer. | Pathology | Pathology | 510(K) Exempt | Test Request Distributor (TRD 1.3 and TRD 1.4), a software module used on the Dako Omnis System, an automated slide stainer for in vitro diagnostic use. |  |  | April 01, 2015 | 2015 | Dako North America Inc. | If a user requests slides from the LIS or TPID, then updates a request by changing the test, the TRD software will reject the update. The LIS and printed slide label will display the updated test with the requested change, but the Omnis system will execute the initial test, not reflecting the change. The system will not warn the user that the test request was rejected. | Other | Dako sent a Voluntary Product Recall Notification letter starting the week of April 1, 2015 along with a acknowledgement attachment. The letter identified the affected product, problem and actions to be taken. The firm requests that customers inform those in their organization of the notification including organizations where the potentially affected instruments have been transferred. The firm requests that customers complete, sign, and return the acknowledgement form within one week after receiving the notification. For questions contact your Dako sales representative.On 04/29/15 the firm sent out an updated customer notification letter in order to provide additional clarification and instructions to customers in regards to the recall situation. |  | 33 | 33 (7 in U.S. and 26 Internationally) | Worldwide Distribution - U.S Nationwide in the states of IN and PA and the countries of: Spain, Belgium, France, Netherlands, Denmark, Finland, Sweden, Germany, UK, and Korea. | Terminated | August 05, 2015 | 126 | 33 | Computer
Z-1544-2015 | 70873 | No | BSP | Anesthesia conduction needle. | Anesthesiology | Anesthesiology | 510(k) | Integra Long Epidural Steroid Injection Needle (Tuohy Needle), 18 Gauge x 5 inch Removable Wing, Metal Stylet, Calibrations (10BX), Catalogue No. PITRW1850. |  |  | March 30, 2015 | 2015 | Integra LifeSciences Corp. d.b.a. Integra Pain Management | A single lot of 18 Gauge, 3.5 inch Tuohy Needles were incorrectly packaged and labeled as 18 Gauge, 5 inch needles. | Employee error | US Consignees were notified by an Urgent: Voluntary Medical Device Recall letter on 3/30/15 delivered by traceable courier service. The letter identified the affected product and explained the reason for the recall. Customers were asked to review their inventory, identify affected products, and immediately stop use and remove them from service. Customers are to complete the attached form and return via mail or fax. Upon receipt of the form, Customer Service will provide an RMA number and directions for returning the affected products to the firm. Customers should contact Customer Service at 1-800-241-2210 with any questions regarding the instructions. |  | 900 | 90 Tuohy needles (9 cases of 10 each) | Nationwide Distribution. | Terminated | June 24, 2015 | 86 | 900 | Not_Computer
Z-1543-2015 | 71015 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Remote Network Station, Catalog RNS-9703, Model No. RNS-9703-19, RNS-9703-24, Software version 02.40. |  |  | April 14, 2015 | 2015 | Nihon Kohden America Inc | Nihon Kohden America (NKA) is recalling the Remote Network Station (RNS) 9703 because it may fail to sound. | Under Investigation by firm | The firm, Nihon Kohden, sent an Urgent Medical Device Alert" letter dated 4/14 /15 to customers. The letter describes the product, problem and actions to be taken. The customers were instructed to stop using the RNS for primary alarm notification and instead utilize the CNS and or the Bedside monitor for primary alarm notification, until such time NKA can provide corrective measures to your facility; Inform potential users in your organization of this notification, and complete the attached reply form and return it by fax (949) 580-1550 or email it to RNS_Replay@nkusa.com. Customers with questions are instructed to contact NKA Technical Support Center at (800) 325-0283 Ext 6 (available 24 hours a day/7 days/week). |  | 208 | 208 units | US Distribution to states of: AL, CO, DE, FL, GA, HI, KY, IL, IN, MA, MD, MI, MN, MO, NC,, ND, NJ, NY.OH, PA, and WV. | Terminated | October 01, 2015 | 170 | 208 | Not_Computer
Z-1545-2015 | 70906 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Insorb Subcuticular Skin Stapler. INSORB staples are made from an absorbable copolymer which is a synthetic polyester derived from lactic and glycolic acids. It is chemically similar to other surgical glycolide/lactide-based copolymers. Polyglycolic/polylactic acid copolymers degrade in vivo by hydrolysis to glycolic acid and lactic acid which are then absorbed and metabolized by the body. |  |  | March 19, 2015 | 2015 | Incisive Surgical Inc | A previous Turkish distributor over-labeled a portion of the lot with an expiration date of 2016-05. The expiration date at the time of manufacturing was 2014-05. | Labeling False and Misleading | Firm is working with EU Rep on English and Turkish notification letters. |  | 1260 | 210 cartons (containing 6 INSORB staplers per carton) | International Distribution Only - Turkey. | Terminated | July 08, 2015 | 111 | 1260 | Not_Computer
Z-1550-2015 | 70574 | Yes | KDI | High permeability hemodialysis system. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | 2008 Series Hemodialysis Machines: 2008T with the following product code and description: 190395 2008K at Home System - 1st Gen. -Âż Canada; |  |  | February 20, 2015 | 2015 | Fresenius Medical Care Holdings, Inc. | The Acetate value for GranuFlo on the Select Concentrate screen should be 8mEq/l but it is displayed as 4mEq/l. The correct value is shown on the dialysate screen. | Software design | The firm, Fresenius Medical Care, sent an "URGENT - MEDICAL DEVICE CORRECTION" letter dated February 19, 2015 to its Consignees via certified mail on February 26, 2015. The letter describes the product, problem and actions to be taken. The consignees were instructed to sign and return the attached "Reply Form" via fax to: 781-699-9769 Attn: Quality Department or Scan and E-mail to: NOTIFYRA@fmc-na.com to confirm that they received and understand the customer notification. The firm will be providing a software update and stated that in the near future, your local clinic will be contacting you about installing the software update.If you have any questions pertaining to this Customer Notification, please contact your local clinic, 800-227-2572 or the Corrections and Removals Manager at 800-662-1237. |  | 133751 | 133,751 in total (US: 130,337 and OUS: Canada 1,201 and Mexico 2,213) | Worldwide Distribution: US (Nationwide) and countries of: Canada and Mexico. | Open, Classified |  |  | 133751 | Computer
Z-1547-2015 | 70958 | No | KWQ | Spinal intervertebral body fixation orthosis. | Orthopedic | Orthopedic | 510(k) | 4.0mm Variable Angle, Self-Drilling Hexalobe Screw, 14mm |  |  | April 10, 2015 | 2015 | Alphatec Spine, Inc. | During a field audit it was discovered by an employee that a 16mm Trestle Luxe Screw has been both laser marked and labeled as a 14mm Trestle Luxe Screw (4.0mm Variable Angle, Self-Drilling Hexalobe Screw, 14mm). As a result the firm recalling the affected lots codes for this mislabeling. | Employee error | Alphatec Spine sent an Urgent-Medical Device Recall Notification dated April 7, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The firm first requested that customers notify their consignees of the recall and the risk, and risk mitigation factors.The firm requests that customers take the following field actions:-Upon receipt of the Field Safety Notice to review inventory to determine if affected screw --part number 71240-14, lot 685904 is in their possession.-If the product is in their possession to please abstain from use and contact Alphatec Spine Customer Service (customerservice@alphatecspine.com) immediately for instructions on how to return the affected product. All shipping instructions will be provided, including replace product and/or credit.-Fill out page 3 of the letter to confirm that they have read the notification and have take all necessary actions stated in the notification.-Return a signed copy using one of the methods below:Mail to: Alphatec Spine, Inc. 5818 El Camino Real Carlsbad, CA 92008 Fax to: 1-800 431-7729 For further questions call ( 760) 431-9286. |  | 61 | (244 units in US and 61 units OUS) | Worldwide Distribution - US Distribution to the states of: KS, LA, TN, NJ, KY, CA, FL, OR, OK, AZ, WI, CO, IN, NH, AR, UT, and AL., and to the countries of : Spain, Japan, Belgium, and Saudi Arabia. | Terminated | July 10, 2015 | 91 | 61 | Not_Computer
Z-1546-2015 | 71061 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Female Driver, Part number FDr101, instrument used for the insertion of the Fassier-Duval IM Telescopic System. For orthopedic procedures. |  |  | April 10, 2015 | 2015 | Pega Medical Inc. | The hexagonal tip of the Female Driver instrument for the Fassier-Duval IM Telescopic System is 0.06 mm under the specifications. This may result in a reduction of the total strength of the instrument, possibly leading to deformation, wear, and tip breakage. | Nonconforming Material/Component | The firm sent the Urgent Field Safety Notice-Recall letter, dated April 10, 2015 to consignees via email. The letter listed the affected product as FEMALE DRIVER FASSIER-DUVAL IM TELESCOPIC SYSTEM.Consignees are requested to return devices back to the firm for disposal. Consignees with questions should contact the firm at 450-688-5144 ext. 242 or send email to egarcia@pegamedical.com |  | 9 | 9 units (5 units in the US and 4 units international) | Worldwide distribution. US in Delaware, Florida, Nebraska, and New York; Australia, Brazil, and United Kingdom. | Terminated | June 29, 2015 | 80 | 9 | Not_Computer
Z-1557-2015 | 70753 | No | BSR | Tracheal tube stylet. | Anesthesiology | Anesthesiology | 510(K) Exempt | Introduction Pack. A device inserted into an endotracheal tube to aid in the intubation procedure. |  |  | March 13, 2015 | 2015 | Vital Signs Colorado Inc. | Vital Signs Colorado (dba CareFusion) is recalling Vital Signs LightWand Stylet catalog numbers 3910 and 3960 due to a potential patient safety risk involving a breach in the sterile packaging that may not be recognized prior to use. | Packaging | Customers were notified on 3/13/15 via letter to perform an inventory, complete and return the Recall Response Form within 15 days, and to destroy the product on site (do not return to CareFusion). |  | 399 | 399 | Worldwide Distribution. US Nationwide, Australia, Canada, Chile, Colombia, Dominican Republic, Ecuador, Egypt, Germany, Guatemala, Hong Kong, India, Israel, Italy, Japan, South Korea, Lebanon, Mexico, Netherlands, New Zealand, Panama, Peru, Poland, Russia, Slovenia, Turkey, United Kingdom, and Qatar. | Terminated | October 07, 2015 | 208 | 83219 | Not_Computer
Z-1560-2015 | 70834 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | PowerLoc MAX Safety Winged Infusion Set, Product Code 0132215 |  |  | March 13, 2015 | 2015 | Bard Access Systems | Bard Access Systems is conducting a field action due to the potential that an incorrect Maximum Flow rate label associated to the power injection IFU for the PowerLoc MAX Safety Infusion Set (SWIS) was attached to the outside of the unit pouch. | Labeling False and Misleading | The firm, Bard, sent an "URGENT: Notification-Potential Incorrect Infusion IFU Label for PowerLocÂż MAX Safety Winged Infusion Set (SWIS)" letter dated March 25, 2015 to its customers. The letter described the product, problem and actions taken. The firm has corrected the cause of the issue and all future lots will be manufactured with the correct maximum flow rate of 2mL/second. Customers may return for a replacement or keep items.Please direct any return requests and questions to our coordinator at 801-522-5640. |  | 700 | 700 units | US Distribution to states of: GA and MN. | Terminated | June 24, 2015 | 103 | 700 | Not_Computer
Z-1561-2015 | 70938 | No | FLE | Steam sterilizer. | General Hospital | General Hospital | 510(k) | AMSCO C and AMSCO 400 Steam Sterilizers, STERIS Corporation. |  |  | February 19, 2015 | 2015 | Steris Corporation | STERIS has identified that the control board software in select AMSCO 400 and AMSCO C units will interrupt and cancel a processing cycle should the selected sterilizer shutdown time coincide with a processing cycle. This may result in a procedure delay. | Software design (manufacturing process) | On 3/9/2015, the firm sent Urgent Field Correction Notices, to their customers. The letter identified the affected product and gave a description of the problem. The letter states that the control software is being revised to ensure the utility shutdown feature will not interrupt an ongoing processing cycle. STERIS Field Representatives will contact customers to schedule installation of the new control board on affected units. Customers may direct questions regarding the issue to STERIS Customer Service at 1-800-548-4873; Marie LaFrance, Product Manager, at 440-392-7648, or their local sales representative. |  | 701 | 701 units | Worldwide Distribution -- US, including the states of AK, AL, AR, AZ, CA, CO, CT, DC, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, I, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, &amp; WY; and, the countries of Afghanistan, Canada, Japan, Kuwait, Mexico, Republic of Korea, &amp; United Arab Emirates. | Terminated | January 20, 2016 | 335 | 701 | Computer
Z-1558-2015 | 70895 | No | BTT | Respiratory gas humidifier. | Anesthesiology | Anesthesiology | 510(k) | Teleflex Medical Hudson RCI ConchaTherm Neptune Heated Humidifier, Rx Only. |  |  | April 01, 2015 | 2015 | Teleflex Medical | Customer complaints were received of power to the unit ceasing and the display going blank. | Nonconforming Material/Component | Teleflex sent an Urgent - Medical Device Notification letter dated April 1, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.They were informed that Teleflex Medical would exchange all of the potentially impacted units in the field with the NeptuneÂż Heated Humidifier with ConchaSmart" Technology. In the meantime, they were instructed to continue to follow hospital and departmental protocol regarding the regular monitoring of the patient and the product during treatment. If the unit display is blank, replace it with another unit and call Teleflex for an immediate replacement. A Recall Acknowledgement Form was attached to be completed, scanned and returned to Teleflex Medical to document receipt of this Notification.Teleflex Medical is committed to providing high quality, safe and effective products. We sincerely apologize for any inconvenience this action may cause your operations. If you have any other questions, feel free to contact your local sales representative or Customer Service at 1-866-396-2111. |  | 7228 | 7,228 units | US Nationwide and Internationally to the countries of: Belgium and Japan | Open, Classified |  |  | 7228 | Not_Computer
Z-1553-2015 | 70919 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Synthes Small Battery Drive (SBD) Oscillating Saw Attachment Part Number: 532.021 for orthopedic surgery. |  |  | March 24, 2015 | 2015 | The Anspach Effort, Inc. | The potential exists for the Oscillating Saw Attachment to disengage from the Small Battery Drive handpiece. If the attachment disengages from the handpiece intra-operatively, there is a potential for injury to the patient and/or user. | Device Design | DePuySynthes sent a "URGENT NOTICE: MEDICAL DEVICE RECALL- CORRECTION" letter dated March 24, 2015 to all affected customers. The letter describes the product, problem, and the actions to be taken by the customers. The firm requests that customers review the inventory and perform the in-process inspection outlined in the letter attachment, a review and repair will be implemented if necessary during the next visit schedule maintenance interval at the customer site.Customers having questions regarding this notice should contact the firm Complaint Handling Unit Manager at 561-494-3673 or to contact their Depuy Synthes Sales Consultant.For questions regarding this recall call 561-627-1080. |  | 12085 | 12,085 pieces | Worldwide Distribution - US (nationwide including DC, Hawaii, and Puerto Rico) and Internationally to Argentina, Austria, Australia, Belgium, Bahrain, Bolivia, Brazil, Canada, China, Chile, Costa Rica, Czech Republic, Germany, Denmark, Ecuador, Egypt, Spain, Finland, France, Great Britain, Greece, Guatemala, Hong Kong, Croatia, Hungary, Indonesia, Ireland, Israel, India, Iraq, Iran, Italy, Jordan, Japan, Kenya, Korea, Kuwait, Lebanon, Sri Lanka, Libyan Arab Jamahiriya, Morocco, Macedonia, Mexico, Malaysia, Netherlands, Norway, Nepal, New Zealand, Oman, Ontario, Peru, Philippines, Palestine, Portugal, Romania, Serbia, Russian Federation, Saudi Arabia, Sweden, Singapore, Slovenia, Slovakia, South Africa, El Salvador, Thailand, Turkey, Taiwan, Ukraine, Uruguay, Venezuela, Vietnam and United Arab Emirates. | Open, Classified |  |  | 12085 | Not_Computer
Z-1559-2015 | 70899 | No | MDZ | Medical washer-disinfector. | General Hospital | General Hospital | 510(k) | Caviwave Pro Ultrasonic Cleaning System, |  |  | February 19, 2015 | 2015 | Steris Corporation | Should the Caviwave Pro Ultrasonic Cleaning System experience a temporary interruption in electrical power during a processing cycle, the circulation pump may stop functioning. In addition, it is possible that the ultrasonic generator may stop working. The machine does not provide an alert notifying the user of this device malfunction. | Software design | STERIS sent an Urgent Field Correction Notice letter dated March 9, 2015 customers. The letter identified the description of the product, description of the problem and actions to be taken. For questions contact STERIS Customer Service at 1-800-548-4873 or local sales representative. |  | 491 | 491 units | Worldwide Distribution - US Nationwide in the states of: AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, ND, NH,NJ, NM, NV, NY, OH, OK, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI &amp; WV and the countries of: Australia, Brazil, Canada, France, Mexico, Saudi Arabia, Spain, Taiwan &amp; United Arab Emirates. | Terminated | October 21, 2015 | 244 | 491 | Not_Computer
Z-1574-2015 | 70638 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | MHI-TM2000 Linear Accelerator System (Software Version 3.5.0 and 3.5.1) |  |  | April 17, 2015 | 2015 | MITSUBISHI HEAVY INDUSTRIES, LTD., HIROSHIMA MACHINERY WORK | The operator console allows users to deliver therapeutic radiation to patients even though a specific communication error relevant to imaging conditions has occurred. In addition, the treatment record (the delivered radiation record) cannot be saved. | Software design (manufacturing process) | MITSUBISHI HEAVY INDUSTRIES, LTD. has issued a Safety Alert for the MHI-TM2000 Linear Accelerator System dated April 10, 2015, due to a software issue.Customers were advised not to deliver therapeutic X-ray if popup warning message show is displayed during treatment. Instructions were provided if therapeutic radiation is delivered during popup warning. Customers with questions should contact the Customer Service Representative by calling 1-800-597-5911, Monday through Friday or email: us.support@brainlab.com. **CAPA** Correction software for this issue has a tentative availability date of June 2015. Estimated software update completion date would be end of October 2015. |  | 11 | 11 units total | Worldwide Distribution - US including FL, NY, OH, &amp; TX, **Internationally - Belgium, Italy, Korea, France &amp; Germany. | Terminated | April 11, 2017 | 725 | 11 | Computer
Z-1562-2015 | 70841 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | 3/16 in. Trocar with Single Drain |  |  | March 26, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Stryker Instruments is voluntarily recalling the CBCII Wound Drains and Ambulating Bulb for a potential sterility breach in the packaging | Package design/selection | An URGENT MEDICAL DEVICE NOTIFICATION letter was sent to all customers on 03/31/2015. The letter identified the affected product, the reason for the recall, the risk to health, and actions to be taken by the customer or user. All are asked to sign the Business Reply Form stating they have read the Notification and followed all instructions within. Customers were asked to immediately check all inventory and quarantine any affected product, return affected product to Stericycle for destruction using a return label provided by Stericycle. Upon receipt of affected product a credit will be issued to customer account. |  | 2711 | 2,711 boxes (27,110 each) | Worldwide Distribution. US nationwide, Australia, Great Britain, The Netherlands, Spain, Romania, Greece, Switzerland, South Africa, and Japan. | Terminated | December 31, 2015 | 280 | 42806 | Not_Computer
Z-1395-2015 | 70701 | No | IWE | Light beam patient position indicator. | Radiology | Radiology | 510(K) Exempt | GE Discovery MR950 MRI system |  |  | October 24, 2014 | 2014 | GE Healthcare | The alignment lasers are missing the labels required by radiation safety regulations. | Error in labeling | GE Planned action(s) to repair defect or to bring product into compliance:The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 5 units consists of the following:1. New labels will be mailed to all purchasers, with instructions to apply the label in the appropriate location on the unit, or a GE field engineer will be sent to each user site to complete the correction.2. All mailings and service calls will be made at no cost to the purchaser.3. The corrections will be completed by August 1, 2015.If you have any questions, please contact your local Service Representative or call 1 (800) 437-1171. |  | 5 | 5 units (2 in US) | US Distribution to the state of CA., and Internationally to Italy and Japan. | Terminated | May 06, 2015 | 194 | 5 | Not_Computer
Z-1575-2015 | 71009 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | NexGen Augment; 5mm Distal Femoral Augment;Part number 00-5990-036-10. |  |  | April 02, 2015 | 2015 | Zimmer, Inc. | The affected lot is being recalled because it was packaged without the locking screw. The locking screw packaged with the augment is used to fix the augment to the femoral component. This has the potential to delay surgery | Mixed-up of materials/components | On 4/2/2015, URGENT MEDICAL DEVICE RECALL LOT SPECIFIC notifications were sent to the affected distributors with instructions for locating, quarantining, and returning the affected product. All distributors were notified via electronic mail, and distributors with affected inventory were also notified via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-877-946-2761 between 8:00 am and 5:00pm EST. |  | 20 | 20 units | Distributed in the states of IN, MN, FL, and IL, and the countries of Argentina and Germany. | Terminated | September 23, 2015 | 174 | 20 | Not_Computer
Z-1370-2015 | 70678 | No | OYE | Automated differential cell counter. | Hematology | Hematology | 510(k) | BD FACSCalibur; |  |  | March 06, 2015 | 2015 | BD Biosciences, Systems &amp; Reagents | The circuit breakers on some BD FACSCalibur instruments may have the incorrect rated amperage (20 amps instead of 2 amps) causing the circuit breaker to fail to open in an over-current situation. | Mixed-up of materials/components | BD Biosciences sent an Important Product Recall letter on March 6, 2015 by certified mail and Field Change Notification (FCN FCB15-01-660312 directing Field Service Engineers to correct the circuit breaker is in process. Customers with questions were instruct to contact BD Customer Support at 1-877-232-8995 (prompt 2) in the United States. For customers outside the US were instructed to contact their local BD Biosciences representative or distributor. For questions regarding this recall call 408-954-6080. |  | 72 | 72 devices | Worldwide Distribution -US including TX, NJ, and OK and Internationally to Hong Kong, Japan, Singopore, and Belguim | Terminated | April 04, 2016 | 395 | 72 | Not_Computer
Z-1576-2015 | 70870 | No | BRY | Anesthetic cabinet, table, or tray. | Anesthesiology | Anesthesiology | 510(K) Exempt | CONTIPLEX Continuous Peripheral Nerve Block Tray, CONTIPLEX Stirn |  |  | March 27, 2015 | 2015 | B. Braun Medical, Inc. | There is a potential for a hole to be present in the outer tray of a limited number of pain control trays. This issue may cause the integrity of the sterile barrier to be compromised. A compromised sterile barrier poses a risk of inflammatory response and/or possible contamination of the spinal fluid or epidural space. | Nonconforming Material/Component | BBMI provided a recall notification letter, dated March 27, 2015, to distributors and end users in receipt of suspect product, along with acknowledgement forms detailing the suspect product that was distributed to each account. All customers were provided with instructions for disposition of suspect product in their inventory. Distributors were reminded of their responsibility in executing the recall. Returned product will be sent to the BBMI Distribution Center in Breiningsville, PA. |  | 5000 | 5000 | Nationwide Distribution. | Open, Classified |  |  | 5000 | Not_Computer
Z-1577-2015 | 71118 | No | DSI | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Ultraview SL Command Modules, Model 91496, with the Masimo SpO2 PCBA, PN: 010-1636-02. |  |  | April 21, 2015 | 2015 | Spacelabs Healthcare Inc | Modules resetting. This failure mode happens when the Spacelabs Command Module is configured with the Masimo SpO2 option (-M), the Spacelabs Respiration option (-R) and the Masimo SpO2 PCBA, PN: 010-1136-02. With these two options and the Masimo SpO2 PCBA present, the module may experience random resets. | Software design | Spacelabs expanded the recall to include additional 29 units and notified one additional new consignee on June 4, 2015.*******************************************************************************************The firm, Spacelabs Healthcare, sent an "URGENT-MEDICAL DEVICE CORRECTION" letter dated 27 April 2015, via priority mail to US consignees and email to foreign consignees. The letter described the product, problem and actions to be taken. The consignees were instructed to circulate this notification to all persons concerned by the information it contains; add a copy to the appropriated system user manuals until such time as corrections have been implemented by Spacelabs, and weigh the benefits versus the risks when deciding whether or not to continue to use your modules until they can be upgraded.Only modules shipped or repaired after March 8, 2014 are affected by this recall.Affected command modules will be upgraded with a new version of software by Spacelabs at no cost.Customer with questions about this corrective action, please contact Spacelabs Healthcare at 1-800-522-7025 and select 2 for Technical Support. |  | 1973 | 1973 units total (1345 modules in the US and 604modules international and 3 PCBAs in US and 21 PCBAs international) | Worldwide Distribution:US (nationwide) including states of: AL, AR, AZ, CA, FL, GA, IA, ID, IL, LA, ME, MI, MN, MO, NE, NJ, NY, OH, OR, PA, SC, TN, TX, VA, WA, WY, and Puerto Rico and countries of: Argentina, Australia, Bangladesh, Canada, Chile, France, Great Britain, India, Kuwait, Panama, Poland, Qatar, Saudi Arabia, and Turkey.. | Terminated | November 16, 2015 | 209 | 1973 | Computer
Z-1583-2015 | 71029 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | AXIOM ARISTOS FX; Multipurpose Radiography System (MPRS) is a dedicated x-ray system with a flat panel detector which allows the acquisition of x-ray exposures without the use of conventional film/screen systems. The MPRS allows radiographic exposures of the whole body including skull, spinal column, chest, abdomen, extremities, and excluding mammography. Radiographic exposures may be taken with the patient in the sitting, standing, or supine positions. The intended use and indications for use of this modified device as described in its labeling have not changed from its predicate device. |  |  | April 14, 2015 | 2015 | Siemens Medical Solutions USA, Inc | If the cable cart and the corrugated hose are unable to move freely, the cable holder could drop down and hit patients or system operators. | Device Design | The firm, Siemens Healthcare, sent a "Safety Advisory Notice" dated April 14, 2015, to its customers. The notice described the product, problem and actions to be taken. The customers were instructed to follow the instructions in the notice; promptly notify and instruct all the staff at your organization making them aware of the problem; and forward this safety information to other organizations that could be affected. If the device has been sold and therefore no longer in your possession, forward the safety notice to the new owner. Additionally, customers were instructed observe this safety notice and comply with the corresponding measures until the update has been fully completed. If you have any questions, contact customer support at 610-219-6300. |  | 43 | 43 | Nationwide Distribution. | Open, Classified |  |  | 43 | Not_Computer
Z-1460-2015 | 70732 | No | MQB | Stationary x-ray system. | Radiology | Radiology | 510(k) | EOS, Digital radiography system used in general radiographic examinations. |  |  | February 17, 2015 | 2015 | Eos Imaging Inc | When performing calibration, an alert message on the spectral filtration of the X-ray beam may be suppressed. Improper filtration of the X-ray Beam can then occur in exams set up with copper filtration. | Radiation Control for Health and Safety Act | Planned Action of EOS Imaging Inc.:1. Install a new acquisition software version, which will prevent use of the system when improper filter positioning is detected.2. A customer notification letter which includes a statement that EOS Imaging will without charge, remedy the defect or bring the product into compliance.CDRH approves the CAP subject to the followingCDRH approves the CAP subject to the following conditions:1. Provide documentation to show that all filter positioning subassemblies have been replaced, which you state was implemented to prevent improper filter positioning.2. Notification of all dealers and purchasers is to be made within 15 working days of receipt of this letter in the manner specified in 21 CFR 1003.21 and 1003.22. This office and the Food and Drug Administration (FDA) district office coordinator noted below are to be included in the notification.3. EOS Imaging will implement this CAP by September 1, 2015.For further questions call (678) 564-5400. |  | 13 | 13 units installed in US | US Distribution to the states of: CA, DE, PA, MN, FL, MO, OH, IN and IL. | Terminated | May 31, 2017 | 834 | 13 | Computer
Z-1581-2015 | 71116 | No | IYO | Ultrasonic pulsed echo imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 volume imaging ultrasound systems at software versions VA16C, VA16D and VA16E. Model 10433816. |  |  | May 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | In some cases, the system is unable to capture a clip or image during a routine scan. | Other | Siemens sent a Customer Advisory Notice letter to affected customers. The letter identified the affected product. problem and how this issue is being resolved. To resolve this issue on all affected systems, a mandatory software update has been developed and is being installed free of charge. |  | 186 | 186 devices | Worldwide Distribution - US Nationwide and the countries of:Angola, Armenia, Australia, Austria, Belgium, Bosnia Herzeg., Brazil, Brunei, Bulgaria, Colombia, Czech Republic, Finland, France, Germany, India, Japan, Kazakhstan, Kenya, Kuwait, Mauritius, Mexico, Norway, Oman, P.R. China, Poland, Rep. of Yemen, Republic Korea, Romania, Russian Fed., Saudi Arabia, Singapore, Slovakia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E., United Kingdom. | Terminated | May 06, 2015 | 1090 | 186 | Computer
Z-1578-2015 | 71084 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ 5 Ultrasound System with Pediatric Cardiology option, Model: EPIQ 5G, EPIC 5C, EPIQ 5W. Catalog Number: 795204 / 795205. Part Number: 989605408541 With |  |  | April 10, 2015 | 2015 | Philips Ultrasound, Inc. | When Epiq 5 Ultrasound System, WITH the Pediatric Cardiology option, is set up to Metric and weight and/or height is entered, a unit conversion error may result in the incorrect calculation of the patients Body Surface Area (BSA). | Software design | Philips sent the Urgent-Medical Device Correction-Philips Model EPIQ5 and Ultrasound system letter, dated 2015 APRIL 10, and the Urgent-Medical Device Correction-Philips Model EPIQ7 and Ultrasound system letter, dated 2015 APRIL 10, letters to their consignees.These systems WITH Pediatric Cardiology option are affected.Philips issued two Field Change Orders (FCO) to their Field Service Engineer to install upgraded software: -FCO 79500336 for EPIQ5-FCO 79500337 for EPIQ7Customers with questions can call Philips Customer Service at 1-800-722-9377. |  | 751 | 751 units total (246 units in the US and 505 units outside the US) | Worldwide Distribution - US (nationwide and DC) and Internationally to Norway, India, New Zealand, Bahrain, Sweden, Italy, Australia, France, Netherlands, Switzerland, Germany, Austria, and Turkey. | Terminated | February 29, 2016 | 325 | 751 | Computer
Z-1580-2015 | 71050 | No | IYE | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | ViewRay System, Radiation Therapy System |  |  | January 15, 2015 | 2015 | Viewray Incorporated | The software was not correctly using the RT (Radiation Therapy) to MR (Magnetic Resonance image) coordinate correction for non HFS (Head First Supine) patient orientations, resulting in slice mismatch error. | Software Design Change | On January 15, 2014, ViewRay personnel discussed the investigation of the issue (slice mismatch error) with the customer. The firm advised the customer that until the issue was corrected, to not treat patients in non HFS orientations, or to use the MR acquired from the ViewRay system as the primary plan series for non HFS scans. ViewRay installed software version 3.4.5.28 on January 20 and 21, 2014, in order to prevent this slice mismatch error from occurring in any patient position. |  | 1 | 1 | US distribution to MO. | Terminated | September 23, 2015 | 251 | 1 | Computer
Z-1585-2015 | 70854 | No | MUE | Full-field digital mammography system. | Radiology | Radiology | 510(k) | Aspire HD, Full Field Digital, System, x-Ray, Mammographic; |  |  | February 03, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | User Manual lacks a description for handlingsome errors even though the system can be recovered from such errors | Labeling Change Control | Fuji Film on 2/3/15, issued a User Notification letter and the recovery procedure were hand carried to the user. A copy of Notification Letter is user's acknowledgement of receipt of the manual . |  | 1 | 1 unit | Distributed in HI. | Terminated | May 20, 2015 | 106 | 1 | Not_Computer
Z-1584-2015 | 71122 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | ACUSON SC2000 ultrasound systems at software versions VA20A or VA20B. Model number 10433816. |  |  | July 11, 2012 | 2012 | Siemens Medical Solutions USA, Inc. | In certain situations, loss of post exercise images can occur in Stress Echo. | Other | On 07/11/2012 a field corrective action was initiated via update programs US016/12/P to update all affected SC2000 customers to software version VA20C, which contains the fix to the described issue. Each customer received a site visit to install the software update. |  | 476 | 476 | Worldwide Distribution -- US and Albania, Australia, Austria, Bangladesh, Bulgaria, Canada, Colombia, Denmark, Ecuador, Egypt, France, Germany, Hungary, Iceland, India, Israel, Italy, Japan, Kenya, Kuwait, Malaysia, Mauritius, Mexico, Netherlands, P.R. China, Philippines, Portugal, Rep. of Yemen, Republic Korea, Russian Fed., San Marino, Saudi Arabia, Singapore, Spain, Suriname, Sweden, Switzerland, Taiwan, Thailand, Turkey, U.A.E., United Kingdom, and Venezuela. | Terminated | May 06, 2015 | 1029 | 476 | Computer
Z-1582-2015 | 70951 | No | IYH | X-ray radiation therapy system. | Radiology | Radiology | 510(k) | ExacTrac 6.x. is software used to place patients at an accurately defined point within the treatment beam of a medical accelerator for stereotactic radiosurgery or radiotherapy procedures. |  |  | March 06, 2015 | 2015 | Brainlab AG | ExacTrac 6.x Patient Positioning System: Potentially incorrect patient positioning when using the ExacTrac Cone Beam CT (CBCT) with a CBCT acquired at a couch angle other than 0.0 degrees. | Software in the Use Environment | FIELD SAFETY / PRODUCT NOTIFICATION letters dated March 6, 2015 were sent to all customers of record. The letters included instructions for customers to: 1) With immediate effect, always make sure the couch is at exactly 0.0 degrees (displayed on the Varian monitor) when acquiring a CBCT set to import into the ExacTrac software; 2) For Varian systems providing an option to prevent a CBCT acquisition for couch angles other than 0.0 degrees, Brainlab recommends selecting that option for use in combination with ExacTrac; and, 3) Continue to verify any ExacTrac CBCT based correction using the ExacTrac X-ray verification and/or an external IGRT procedure as mandatory. The letters further advised that Brainlab will be taking the following actions: 1. Brainlab provides existing potentially affected ExacTrac v.6.x CBCT Import &amp; Alignment Software customers (with a Varian C-Series linac) with this product notification information; and, 2. Brainlab will provide a software update with this issue solved to affected customers. Brainlab will actively contact the customers starting August 2015 to schedule the update installation. Customers with questions about this issue were advised to contact Brainlab at 800-597-5911 or via e-mail by writing at us.support@brainlab.com. |  | 25 | 25 systems (US only); 53 systems (Foreign) | *** US: Nationwide; *** FOREIGN: Austria, China, Denmark, France, Germany, Hong Kong, Japan, Netherlands, Qatar, Russian Federation, South Korea, Switzerland | Open, Classified |  |  | 25 | Computer
Z-1586-2015 | 70952 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | BARD MAX-CORE Disposable Core Biopsy Instrument |  |  | April 09, 2015 | 2015 | Bard Peripheral Vascular Inc | The firm is recalling Bard MAX-CORE Disposable Core Biopsy Instrument due to the risk of having self-activation-related issues. | Process design | Bard sent an Urgent Medical Device Recall Notification letter dated April 9, 2015 to their customers. The firm's letter provided customers with the following instructions:-Do not use or further distribute any affected product.-Check all inventory locations within the institution for BardÂż Max-Core DisposableCore Biopsy Instruments for the recalled product code/lot number combination.-If further distributed any of the product code/lot numbers listed, please immediately contact that location, advise them of the recall and have them return the affected product to BVD address listed.-Remove any identified product from the shelves.-If used the affected product, complete, and return the attached Recall and Effectiveness Check Form indicating no product will be returned.Report any new and/or previously unreported adverse events associated with this recall to the US Food and Drug Administration's ("FDA") MedWatch Program by phone at 1-800-FDA-1088; by fax at 1-800- FDA-0178; by mail to MedWatch. Food and Drug Administration. 5600 Fishers Lane. Rockville. MD 20857-9787: or on line at http://wwvdda.gov/medwatch/report.htm.Once the product affected by this recall has been removed from inventory:-Fill out the Recall and Effectiveness Check Form. Be sure to state the quantities and lot numbers of each recalled product that are in stock. -Call the firm's Recall Coordinator Regan Gealy at 1-770-784-6471 or email at Regan.Gealy@crbard.com. The Recall Coordinator will issue either a Return Authorization (RCL) Number or Consignment Recall Number (CRC) to facilitate the expedient return of the product.-Fax the Recall and Effectiveness Check Form back to at 1-770-784-6469.-BPV will provide replacement product for the returned product.A mailing label is going to be enclosed to return the affected product. Mark the outside package as "RECALLED PRODUCT" and include the RCL or CRC number. All products should be returned to the follo |  | 7995 | 7,995 units | Worldwide Distribution - US Nationwide in the states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IL, IN, KS, KY, LA, MA, MD, MI, MN, MO, MS, MT, NC, NH, NJ, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WI, WV and the countries of: France, Mexico, Netherlands, Panama, United Kingdom, Switzerland, Spain, Poland, Norway, Iceland, Italy, Greece, Finland, Belgium, Germany, Austria, China, Taiwan, Singapore, Korea, Australia, and Japan. | Terminated | September 25, 2015 | 169 | 7995 | Not_Computer
Z-1591-2015 | 71028 | No | FMG | Intravascular administration set. | General Hospital | General Hospital | 510(k) | 113" (287cm) Transfer Set w/MicroClave T-Connector, Check Valve, Anti-Siphon Valve, Smallbore Quadfuse Ext Set w/MicroClave, 4 Clamps (Blue, 2 Red, White), Rotating Luer, Non-DEHP Tubing, Item No. B33303 |  |  | April 20, 2015 | 2015 | ICU Medical, Inc. | ICU Medical is recalling MicroClave T-Connectors because they have potential connection separation issues that could lead to leakage. | Under Investigation by firm | An Urgent Medical Device Recall letter dated 4/17/15 will be sent to customers on 4/20/15 to inform them that ICU Medical Inc. has identified a potential connection separation issue in a small number of isolated Intravenous Administration Transfer Sets. While the separation may be visually detected, if used, this failure mode could lead to leakage. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to return completed recall response form to ICU Medical by Fax at (801) 264-1755 or email to recall@icumed.com. Customers with questions or require assistance relating to the recall are instructed to contact ICU Customer Service Monday through Friday between the hours of 8:30 AM and 4:00 PM Pacific time: 866-829-9025 and select option 8, or e-mail the following address: customerservice@icumed.com. |  | 900 | 900 units | Nationwide in US: SC, NC, MN, CT, KY, MD, FL, VA, TN | Terminated | November 24, 2015 | 218 | 23100 | Not_Computer
Z-1587-2015 | 70933 | No | HWE | Surgical instrument motors and accessories/attachments. | General &amp; Plastic Surgery | Orthopedic | 510(K) Exempt | Universal Charger |  |  | March 25, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | The Stryker Universal Battery Charger is not transmitting usage data to the Stryker Cloud as designed. The firm is initiating a software correction to address the issue. | Software design | Stryker issued e-mail notices to distributors and reps on March 26, 2015 advising that all battery charges would receive a software update to provide a field correction for the upgrade to units to be able to track device usage data to Stryker as designed. There is no customer action required, the chargers will transmit a message to Stryker once the upgrade is complete. Please contact Kelly Jo Davis at 269-389-2921 with any questions. |  | 223 | 223 chargers | US Nationwide in the states of AL, AR, AZ, CA, CO, FL, GA, IA, IL, IN, KS, KY, MD, MI, MO, MS, MT, NC, NH, NV, NY, OH, OK, PA, SC, TN, TX, UT, VA, VT, WA, WI and the foreign country of: Canada | Terminated | July 08, 2015 | 105 | 223 | Computer
Z-1596-2015 | 70662 | No | BYD | Heat and moisture condenser (artificial nose). | Anesthesiology | Anesthesiology | 510(K) Exempt | ThermoFlo 1 HME, Disposable, Non-sterile, Expiry date 04/17. |  |  | February 19, 2015 | 2015 | Arc Medical Inc | Outer shipping carton was marked with an incorrect expiry date, however inner product carton was labeled correctly. | Process control | Consignees were contacted by telephone on 02/19/2015. |  | 1810 | 1810 units | Distributed in CA and NC. | Completed |  |  | 1810 | Not_Computer
Z-1597-2015 | 70950 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Spinal Elements 5.5 x 45MM Mercury lordosed rod. A component of the Mercury Spinal System, comprised of a variety of screws, hooks, rods, connectors, and staples that are used for attachment to the non-cervical spine. Rods span the distance between screws and hooks and achieve fixation by the mechanical joining of the rods with the screws or hooks. Lordosed rods are pre-curved to better fit the patient anatomy. |  |  | April 09, 2015 | 2015 | Spinal Elements, Inc | Some rods may not meet the appropriate traceability requirements including the lot number. Some rods were found to not include the product identification number (product number and lot number) marked directly on the rod. The lack of marking interrupts the ability to maintain traceability of the device. | Under Investigation by firm | The firm sent customer notification letters on 04/09/15. In this letter the firm requested that all affected lots be returned, and replacement inventory will be made available.The firm requested that customers take the following actions:1. Immediately examine inventory for part number 60100-045.2. Immediately discontinue the use and/or distribution of the affected lots3. Send all affected product back to:Spinal Elements, Inc.ATTN: Julie Lamothe3115 Melrose DriveCarlsbad, CA 920104. If further distributed the product to please identify the customers and notify them as once.5. Complete and return the enclosed "Customer Response Form." Any questions the firm states to contact the Customer Service department at 1.877.SPINAL5 Ext. 213. |  | 588 | 588 | Nationwide Distribution including the states of GA, CA, TX, CO, IN, NC, FL, UT, TN, MO, MI, PA, OH, and OK. | Terminated | June 25, 2015 | 77 | 588 | Not_Computer
Z-1598-2015 | 71190 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Nuclear Gamma Cameras (ADAC Vertex Classic, ADAC VERTEX PLUS/SOLUS/CARDIO, ADAC VERTEX V60) Philips Medical Systems, Cleveland, OH |  |  | March 13, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | During clinical use of a Vertex Plus Gamma Camera, as the detector heads were in the relative 180 degree position, Detector #1 drifted down and made contact with the patients chin during a whole body study acquisition. | Employee error | Philips sent an Customer Information letter dated March 16, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were instructed to :1. Do not lubricate the clutch(es) with oil, grease or any other type of lubricant..2. Follow product Instructions For Use (IFU), CARDIO. SOLUS, ADAC VERTEX, AND ADAC VERTEX PLUS WITH ATLAS USERS MANUAL, 9201-0227E-ENG, REV A- General Precautions that instructs user to monitor the patient during scanning.3. Follow product Instructions For Use (IFU), CARDIC, SOLUS, ADAC VERTEX, AND ADAC VERTEX PLUS WITH ATLAS USER'SMANUAL, 9201-0227E-ENG, REV A- Emergency Patient Removal , Section Safety and Regulatory Agency Compliance, page 2-11, instruction(s) on emergencypatient removal.4. Ensure completion of product Preventive Maintenance (PM) procedure requiring inspection of the radius drive clutches every 3 months for contamination and replacement if contaminated and, testing of the gantry system.If you need any further information or support concerning this issue, please contact your local Philips representative: Technical Support Line: 1-800-722-9377. |  | 918 | 918 units | Worldwide Distribution - US (nationwide) including Puerto Rico and to the states of : AL, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI and WV., and to the countries of : Argentina, Australia, Austria, Belgium, Brazil, Canada, China, Costa Rica, Denmark, Egypt, France, Germany, Ireland, Italy, Japan, Lebanon, Malaysia, Netherlands, Norway, Pakistan, Portugal, Republic of Korea, Russia, South Korea, Spain, Sweden, Syrian Arab Republic, Thailand and United Kingdom. | Open, Classified |  |  | 918 | Not_Computer
Z-1600-2015 | 70975 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | TFN-ADVANCED Proximal Femoral Nailing System (TFNA); for treatment of fractures in orthopedics and trauma. |  |  | April 09, 2015 | 2015 | Synthes, Inc. | TFNA nails from certain lots were assembled with a locking mechanism too close to the top of the nail. Potential patient impact: surgical delay. | Process control | An urgent Field Safety Notification, dated April 9, 2015, was sent to end users and sales consultants to alert them about the issue and request that they take the following actions: 1) review the instructions listed in this notification; 2) forward this notice to anyone in your facility that needs to be informed; and 3) complete the attached Verification Section (page 4 of this notification) and send a copy of the completed Verification Section by: Fax: (888) 446-5160 orScan/email: Synthes3152@stericycle.com. |  | 8175 | 8175 | Distribution US Nationwide and Switzerland. | Open, Classified |  |  | 8175 | Not_Computer
Z-1599-2015 | 71138 | No | PJQ |  |  | Cardiovascular | PMA | VenaSeal, Sapheon Closure System, Model # SP-101, VS-301 and VS-402. For vascular use. |  |  | April 28, 2015 | 2015 | Covidien LLC | Possible sterility breach of the outer packaging (pouch) material which may compromise the outside surface sterility of the sealed inner tray. | Package design/selection | Consignees were contacted by letter via email on April 28, 2015. |  | 1661 | 1661 units | Worldwide distribution. No U.S. accounts; Australia, Canada, France, Germany, Hong Kong, Ireland, Italy, Latvia, Netherlands, Norway, Portugal, Switzerland, UAE and the United Kingdom. | Terminated | March 29, 2016 | 336 | 1661 | Not_Computer
Z-1605-2015 | 71125 | No | DZE | Endosseous dental implant. | Dental | Dental | 510(k) | Tapered HD Upgrade Set, REF TSKHDUS, Non-Sterile, Rx Only, |  |  | April 16, 2015 | 2015 | BioHorizons Implant Systems Inc | BioHorizons Tapered HD Upgrade set was packaged with an incorrect drill component manufactured of stainless steel in place of the correct component manufactured from titanium. | Process control | BioHorizons sent an Urgent Medical Device Field Correction letter on 04/16/2015 to affected customers. Non-responding customers will be notified after 30 days. The letter identified the affected product, problem and actions to be taken. For questions contact Customer Care Department at 888-246-8338. |  | 48 | 48 units | US Nationwide in the states of: AL, FL, MS, VA, LA, CA, NC, TX, TN, OH, OK, VT, WA, MD, NY, SD, and SC. | Terminated | August 07, 2015 | 113 | 48 | Not_Computer
Z-1601-2015 | 70868 | No | LTW | Antimicrobial susceptibility test powder. | Microbiology | Microbiology | 510(k) | Vitek 2 AST-YS06 REF 412 610, Fungal Susceptibility Card, 20 cards per carton, IVD. |  |  | March 27, 2015 | 2015 | Biomerieux Inc | The test kits have Out-of-Range-Low (OORL) Flucytosine QC results for Candida krusei ATCC 6258 | Other | The notification, consisting of the Field Safety Corrective Action (FSCA #2488) packet, was issued electronically to the subsidiaries and distributors on 3/27/15. FSCA #2509 was issued electronically to the subsidiaries and distributor on 4/13/15, which superseded FSCA #2488. |  | 1025 | 1,025/20-card cartons | Distributed only to foreign subsidiaries and distributors, including Canada and Mexico. | Terminated | March 24, 2017 | 728 | 46164 | Not_Computer
Z-1606-2015 | 71211 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | GE Healthcare, CARESCAPE Monitor B850, B650 and B450 (Bx50). |  |  | May 04, 2015 | 2015 | GE Healthcare | The heart rate could be calculated from pacer pulses without indication that pacemaker detection is OFF in some combinations of Bx50 monitors, a PDM, and CIC/CSCS. Then the monitor does not show the Pacer Off message and there may not be an alarm for asystole. Undetected asystole could result in irreversible changes in the patientÂżs condition and delayed or missed life sustaining patient treatment | Software design | GE Healthcare sent an Urgent Medical Device Correction letter dated May 1, 2015, to all affected consignees. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, and Director of Biomedical Engineering. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 110 | 110 (83 units US, 27 units OUS) | Worldwide Distribution - US including CO, MI, WI and Internationally to NETHERLANDS, FINLAND, and BELGIUM. | Terminated | September 15, 2015 | 134 | 110 | Not_Computer
Z-1607-2015 | 71135 | No | LMC | Medical image hardcopy device. | Radiology | Radiology | 510(K) Exempt | Nuclear gamma cameras (ADAC VERTEX PLUS/SOLUS/CARDIO, ADAC VERTEX V60 &amp; ADAC CARDIO C60) |  |  | March 11, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm was notified by a customer that the collimator exchange carriage and the collimator storage cabinet were misaligned. | Device Design | Philips sent an Urgent - Medical Device Correction letter dated March 12, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The firm requested their customers to do the following: " If a misalignment on the top and/or bottom of the guide rail between the cabinet and carriage can be visually detected do not continue with the collimator exchange procedure," If the misalignment is visually detected, be aware that the customer/user cannot remedy this issue and will require service interaction prior to the continued use of the system." This letter should be placed in your User Documentation until otherwise notified.If you need any further information or support concerning this issue, please contact your local Philips Healthcare representative: For North America and Canada contact the Customer Care Solutions Center (1-800-722-9377, Option 5: Enter site ID or follow the prompts). |  | 975 | 975 units | Worldwide Distribution - US (nationwide) to the states of: AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NH, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV and WY., and to the countries of : Australia, Austria, Belgium, Brazil, Canada, China, Costa Rica, Egypt, France, Germany, Hong Kong, Ireland, Italy, Japan, Korea, Lebanon, Netherlands, Norway, Pakistan, Portugal, Russia, Singapore, South Korea, Spain,Thailand and United Kingdom. | Open, Classified |  |  | 975 | Not_Computer
Z-1610-2015 | 71057 | No | GCY | Nonpowered, single patient, portable suction apparatus. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Zimmer 100 mL Hemovac Device Hemovac Mini Evacuator. 00-2568-000-10. Zimmer Surgical, Inc. |  |  | April 21, 2015 | 2015 | Zimmer Surgical Inc | Devices were placed in a shipper carton without the appropriate number of the Instructions for Use (IFU). Packaging specifications require that each sales unit is accompanied by an IFU. Only one IFU was present in the shipper carton, instead of one IFU per each of the two sales units. | Other | On 4/28/2015 the firm sent Urgent: Medical Device Removal letters to their customers. The letters identified the affected product in addition to stating that the reason for the recall is due to missing Instructions for Use from sales units. The letters discussed the clinical implications or risk to health, and actions required. Customers are to review the notification, identify and quarantine the affected product to prevent further distribution or use, and return the completed response form to CorporateQuality.PostMarket@zimmer.com and the affected product to the address provided. A credit will be issued for unused devices upon receipt. Questions regarding this information should be directed to 330-364-0989. |  | 36 | 36 units | Worldwide Distribution -- US, including the states of AZ, CA, ID, IL, MD, NC, RI, and SC; and, Canada. | Terminated | November 06, 2015 | 199 | 36 | Not_Computer
Z-1608-2015 | 71233 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Spacelabs qube Compact Monitor, Model 91390. The monitor is functioning as either bedside or central monitors; passively displays data generated by Spacelabs Healthcare parameter modules, Flexports interfaces, and other SDLC based products in the form of waveform and numeric displays, trends and alarms. |  |  | May 08, 2015 | 2015 | Spacelabs Healthcare Inc | The input power can fail even though the power source at the wall has not been interrupted. In addition, there is the potential for the qube Compact Monitor to fail to turn on. This is a retrospective recall related to a recall initiated in 2012 (Z-0403-2013). | Software design | The firm sent the Urgent Medical Device Correction- Spacelabs Healthcare Inc.-qube Compact Monitor, Model 91390, Power Input PCBA Failure letter, dated 8 May 2015, via priority service mail to US consignees.This letter described the affected system, problem, actions to be taken by users and corrective actions conducted by Spacelabs.For additional information or technical assistance, please contact:Technical SupportSpacelabs Healthcare, Inc.35301 SE Center St.Snoqualmie, WA 980651-800-522-7025 and select 2 for Technical Support |  | 5 | 5 units distributed in the US | distributed in North Carolina, Texas, and Wyoming | Terminated | November 05, 2015 | 181 | 5 | Not_Computer
Z-1604-2015 | 71099 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD) Pump Driveline Splice Kit |  |  | April 29, 2015 | 2015 | HeartWare Inc | Failures of the splice repair kit if exposed to excessive force. | Device Design | HeartWare, Inc. sent an Urgent Medical Device Correction letter to all affected consignees on May 11, 2015. The letter identified the affected product, problem, and actions to be taken by the consignees. The firm requested that consignees: 1) review the notice and the Patient Communication contents, 2) forward the notice to individuals within their organization who need to be aware of the notice, 3) identify patients that are currently on support and have undergone a driveline splice repair, 4) distribute the Patient Communication to the affected patients in person, 5) continue to reinforce the messages described in the notice with patients who have experienced a splice repair during their regularly scheduled appointments, 6) complete, sign, and return the "Acknowledge and Completion Form to HeartWare within 30 days of receipt of the letter. Should any questions or concerns arise, please contact there local HeartWare representative. Their 24-hour Clinical Support personnel are also available at 1 (888) 494-6365. |  | 35 | 35 | Worldwide Distribution - US Nationwide in the states of FL, KY, MO, NY, TN,TX, and the countries of: Canada, Germany, Italy, Netherlands, Turkey and United Kingdom. | Open, Classified |  |  | 35 | Not_Computer
Z-1614-2015 | 70784 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Kinemax Stem Extenders, 80 mm, |  |  | February 27, 2015 | 2015 | Stryker Howmedica Osteonics Corp. | Stryker received four customer complaints reporting that inside the box labeled part #6476-8-260 (Kinemax Stem Extenders, 80 mm) there was a different part #6476-8-250 (Kinemax Stem Extenders, 40 mm. | Mixed-up of materials/components | Stryker Orthopaedics an Urgent Medical Device Recall Letter/Notification Acknowledgement Form dated April 9, 2015 to their affected customers. The recall letter identified the affected product, problem and actions to be taken. Customers are asked to contact their Stryker Orthopaedics Sales Representative to arrange for return of the product if they have any of the affected items. Customers are asked to complete and return the acknowledgement form within 5 days by emailing to: so_m_product_field_action_response@stryker.com or faxed to: 1-855-251-3635. For question call 201-831-6365. |  | 4 | 4 units | US Nationwide distribution in the states of MA and RI. | Open, Classified |  |  | 4 | Not_Computer
Z-1615-2015 | 70913 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Dimension EXL with LM, Dimension EXL 200, Dimension RxL, Dimension RxL HM, Dimension RxL Max, Dimension RxL Max HM, Dimension Xpand, Dimension Xpand HM, Dimension Xpand Plus, Dimension Xpand Plus HM. |  |  | April 02, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Certain Dimension Integrated Chemistry Systems Instrument Power Configurations may have been installed with a non-standard instrument power plug.The use of a non-standard plug configuration as described may lead to a potential electrical hazard. The likelihood of potential clinical impact or injury due to this issue is extremely low. | Under Investigation by firm | Siemens sent an Urgent Field Safety Notice, dated April 2015, was issued for distribution to affected customers in Japan to notify them of the issue. A Siemens Service Representative will visit customer sites to inspect the Dimension instrument power plug and wall outlet configuration. If the instrument is configured with a non-standard 15 Amp plug/outlet or a 20 Amp twist lock plug, Siemens will correct the issue identified, ensuring that the system is installed according to the instrument requirements. |  | 17801 | 17801 | Distributed in the country of Japan. | Open, Classified |  |  | 17801 | Not_Computer
Z-1612-2015 | 70890 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | Fujifilm Synapse Cardiovascular I ProSolv CardioVascular |  |  | February 04, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | Under certain circumstances the patientDemographics in a report exported into the EMR may not match the demographics shown in the corresponding Synapse CV clinical report. | Software in the Use Environment | Fujifilm Medical on 2/4/14 , contacted users and advised to discontinue using HL7 reports pending FMSU determination of root cause. Firm issued follow-up letters on 2/24/14 and 7/24/14 advising users of a mandatory upgrade to version 5.1.1 I base, in-place upgrade, to correct the problems. |  | 16 | 16 units | Nationwide Distribution including PR and the states of IL, IN, FL, GA, MD, MI, MO, MS, NC, NY, OR, and PA, PR and the country of Canada. | Open, Classified |  |  | 16 | Not_Computer
Z-1611-2015 | 70891 | No | DQK | Programmable diagnostic computer. | Cardiovascular | Cardiovascular | 510(k) | FujiMedical Synapse Cardiovascular I ProSolv CardioVascular. |  |  | February 04, 2015 | 2015 | Fujifilm Medical Systems U.S.A., Inc. | Measurement values in the EMR may not accurately reflect the actual measurement value on the Synapse Cardiovascular clinical report. | Software in the Use Environment | FujiMedical Systems North America issued intial recall letter dated February 4, 2014, advising users of the HL7 report exported to the EMR wherein measurement values presented in the Hl7 reports in your EMR may not accurately reflect the actual measurement value on the Synapse Cardiovascular clinical report.Follow-up letters were issued February 12, 2014, and July 24, 2014. A mandatory upgrade to software version 5.1.1/ base, in-place upgrade required. A customer response form attached to acknowledge receipt and arrange for the on site upgrade. For questions regarding this recall call 203-324-2000. |  | 30 | 30 units | Worldwide Distribution - US including CA, FL, IL, IN, MD, MI, MN, MO, NC, NY, OH, OR, PA, PR, TX, VA, WIForeign: Canada | Open, Classified |  |  | 30 | Computer
Z-1613-2015 | 70995 | aid, transfer | IKX | Daily activity assist device. | Physical Medicine | Physical Medicine | 510(K) Exempt | ReTurn belt, wipeable, sizes Small through Extra Large, Plastic pouch, 1 item per pouch. Patient transfer aid. |  |  | March 26, 2015 | 2015 | Handicare USA, Inc. | Complaints were received that the heat-welded seam on the wipeable belt's strap tore when force was applied. | Device Design | Handicare sent an Voluntary Product Recall Notification letter dated April 2, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Customers were requested to quarantine any remaining stock and to contact customers if product was further distributed and to return the affected product for a replacement and/or full refund.Also, enclosed is a response form. Please complete and return the form even if you do not have any product to return. Your Sales representative can assist you in completing this form. This information is essential in order to maintain recall effectiveness information (required by FDA).Thank you for your support. We are committed to re-releasing the wipeable ReTurn belt as soon as we believe that it meets the quality standards that you expect from Handicare. Please contact our Customer Experience team at 1-866-276-5438 if you have any questions regarding the recall or would like assistance with the recall. We regret any inconvenience that this action may cause, but we appreciate your understanding as we take action to ensure patient and customer satisfaction. |  | 64 | 64 | US Distribution to the states of : FL, CA, NJ, NC and MD. | Terminated | August 17, 2015 | 144 | 64 | Not_Computer
Z-1620-2015 | 71027 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | The Integra Licox Pt02 Monitors are compact, portable device for use with Pressure/Temperature, Pressure and Oxygen Partial Pressure Transducer-Tipped Catheters. |  |  | April 14, 2015 | 2015 | Integra LifeSciences Corp. | Integra LifeSciences Corporation is recalling Integra's Camino ICP Monitor (CAM 02) &amp; Licox PtO2 Monitor ( LCX02 &amp; LCX02R) due to complaints that the audible speaker malfunctions. | Device Design | Integra initiated this recall by sending an Urgent Voluntary Medical Device Recall Letter/Recall Acknowledgement Form dated 4/14/15 via traceable courier service to their affected customers. |  | 59 | 59 monitors | Worldwide Distribution. | Open, Classified |  |  | 1814 | Not_Computer
Z-1616-2015 | 71045 | No | GIF | Blood cell diluent. | Hematology | Hematology | 510(K) Exempt | COULTER DxH Diluent, Catalog No. 628017, Part No. A59956. |  |  | April 13, 2015 | 2015 | Beckman Coulter Inc. | Beckman Coulter has received an increased number of complaints related to Hemoglobin (HGB) on UniCel DxH 800 and DxH 600 Coulter Cellular Analysis Systems. The use of the DxH Diluent lots listed in the recall notice may eventually result in a compromised HGB chamber. | Nonconforming Material/Component | An Urgent Medical Device Recall letter dated 4/13/15 was sent to customers who purchased the Coulter DxH Diluent. The letter informs customers that Beckman Coulter has received an increased number of complaints related to Hemoglobin (HGB) on UniCel DxH 800 and DxH 600 Coulter Cellular Analysis Systems. The letter informs customers of the problems identified and the actions to be taken.Customers are instructed to complete and return the enclosed Response Form within 10 days. Customers with questions are instructed to contact Customer Support Center via http://www.beckmancoulter.com/customersupport/support or call (800) 526-7694 in the US and Canada. Customers outside the US and Canada are instructed to contact their local Beckman Coulter Representative. |  | 829888 | 829,888 units total (659,584 units in US) | Worldwide Distribution -- Argentina; Brazil; Brunei Darussalam; Canada; Chile; Colombia; Hong Kong; India; Lebanon; Macao; Mexico; Panama; Philippines; Qatar; Singapore; Taiwan; Thailand; Trinidad and Tobago; United States, including Puerto Rico; and Uruguay. | Terminated | October 01, 2015 | 171 | 829888 | Not_Computer
Z-1622-2015 | 71034 | No | LPS | Urinary homocystine (nonquantitative) test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Homocysteine test system - Homocysteine Reagent. |  |  | March 23, 2015 | 2015 | Roche Diagnostics Operations, Inc. | Customers complained about under-recovery of non-Roche controls and discrepant low patient results with certain Homocysteine reagent lots. This negative bias could, in the worst case, lead to inaccurately low Homocysteine results. An elevated level of Homocysteine is considered an important risk factor in the assessment of peripheral vascular disease. Falsely low values could lead to a delay of di | Nonconforming Material/Component | Roche Diagnostics sent URGENT MEDICAL DEVICE CORRECTION notifications dated 3/23/2015 to affected direct accounts via UPS. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Please contact the Roche Support Network Customer Support Center, 24 hours a day, seven days a week at 1-800-428-2336 if you have questions. |  | 2558 | 2558 units | Nationwide Distribution-including PR and the states of AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KY, MA, MD, ME, MI, MO, MT, NC, ND, NE, NH, NJ, NY, OH, OK, OR, PA, , SC, TN, TX, VA, WA, and WV. | Terminated | January 19, 2016 | 302 | 2558 | Not_Computer
Z-1623-2015 | 71167 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | CT LightSpeed 16 and Discovery ST Pet/CT DST. Catalogue # 2339985. |  |  | April 22, 2015 | 2015 | GE Healthcare | GE Healthcare has become aware that 3 bolts that secure a power supply to the rotating side of the gantry were over-torqued on certain Lightspeed 16 CT scanner or PET/CT DST Scanner.. In the unlikely event that the component becomes loose on the CT gantry during operation, it could result in serious bodily injury if the component were expelled. There were no incidents or injuries reported as a res | Process design | Consignees were sent a GE Healthcare "Urgent Medical Device Correction" letter GEHC Ref# 25461 dated April 22, 2015. The letter was addressed to Hospital Administrators / Risk Managers, Managers of Radiology/Cardiology, and Radiologists/Cardiologists. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. The letter stated that customers can continue to use their device. If the error occurs, then stop using the system and call a service representative to bring the system back into proper operation. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative.UPDATED:Consignees were sent on 5/14/2015 a second GE Healthcare "Urgent Medical Device Correction" letter dated May 14th, 2015. The letter was to incorporate and additional unit being recalled. |  | 4 | 4 USA | Distributed to the states of MN, FL &amp; TX. | Terminated | July 15, 2015 | 84 | 4 | Not_Computer
Z-1621-2015 | 71049 | [6 empty] | MYLA server HP Proliant: DL380-G8 and ML350-G6 version 3.xx computer application software. |  |  | April 14, 2015 | 2015 | bioMerieux, Inc. | The MYLAÂż server could slow down due to the volume ( weight ) of the data to manage and it could also stop processing, impacting the server start-up phase. The problem occurs in high volume settings. The MYLA server slows down and will stop processing completely if any task requires more than 5 minutes to complete. The server then fails to reboot and requires service, provided remotely or on-site | Software design | Consignees were notified electronically on April 14, 2015 and letter via FedEx on 04/20/2015. |  | 730 | 730 units | Worldwide Distribution-Algeria, Argentina, Australia, Austria, Brazil Canada, Chile China, Colombia, Czech Republic, France Germany , Greece, Hungary, India, Italy Japan, Korea, Mexico, Netherlands, Poland, Portugal, Russia, Serbia, Singapore, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, Angola, Armenia, Kocaeli, Bahrain, Bangladesh, Republic of Belarus, Benin, Bolivia, Bosnia, Brunei, Bulgaria, Burkina Faso, Cambodia, Cameroon, Chad, Democratic Republic of Congo, Costa Rica, Germany Dominican Republic, Dutch Antilles, Ecuador, Egypt, El Salvador, Estonia, Ethiopia, Gabon, Georgia, Guam, Guatemala, Guinea, Honduras, Indonesia, Iran, Iraq, Israel, Ivory Coast, Kazakhstan, Kenya, Kuwait, Laos, Latvia, Lebanon, Libya, Lithuania, Macedonia, Malaysia, Mali, Malta, Morocco, Portugal, Myanmar, Managua, Niger, Nigeria, Oman, Palestine, Panama, Paraguay, Peru, Philippines, Qatar, Romania, Kingdom of Saudi Arabia, Senegal, Slovakia, Sri Lanka, Tanzania, Togo, Tunisia, Ukraine, UAE, Uraguay, Venezuela, Vietnam, Yemen and Zimbabwe. | Open, Classified |  |  | 730 | Not_Computer
Z-1617-2015 | 71016 | No | MHY |  |  | Neurology | PMA | Medtronic DBS Extension Kit for Deep Brain Stimulation, Models 37085 and 37086. Sterile |  |  | April 16, 2015 | 2015 | Medtronic Neuromodulation | Medtronic is providing Important information regarding extension handling during implant procedure after receiving reports of high impedances. This applies to extension Models 37085 and 37086 which can be used with the followingimplantable Neurostimulators: ActivaÂż PC (Model 37601), ActivaÂż RC (Model 37612), and the ActivaÂżSC (Model 37603). | Under Investigation by firm | Medtronic began notifying consignees with an Urgent Medical Device Safety Notification letter beginning April 16, 2015. Consignees will be notified by mail and/or in person, by visit of a Medtronic field representative. The letter described the issue and provided recommendations of following the instructions defined within the Neurostimulator implant manual. A Physician Reply Form was asked to be returned via fax at 800-897-3899 or email to neuro.quality@medtronic.com. Consignees with questions can contact Medtronic Neuromodulation Technical Services at 1-800-707-0933 weekdays 7am - 6pm CST. |  | 3294518983 | 86,851 (US 37,933 / OUS 48,918) | worldwide | Open, Classified |  |  | 3294518983 | Not_Computer
Z-1628-2015 | 70946 | No | NEP | Intravascular administration set. | General Hospital | General Hospital | 510(K) Exempt | ABACUS SE (Single-Workstation Edition) and ABACUS ME (Multi-Workstation Edition) |  |  | April 08, 2015 | 2015 | Baxter Corporation Englewood | Baxter Corporation is conducting a field action for the ABACUS SE and ME models due to the possibility that the compounder will load the incorrect formula upon scanning the bag label when two or more different formula files have been created with the same order number. | Software design | Baxter sent a Safety Alert letter dated April 8, 2015 to affected customers via USPS first class mail. The letter identified the affected product, problem and actions to be taken. Customers were asked to 1) contact Baxter Technical Services if you need to restore the ABACUS databse, 2) foword a copy of this Safety Alert letter to other facilities or departments within your institution to ensure that those locations are awae of this action, and 3) complete the customer reply form and return it to Baxter by faxing to 224-270-5457 or scanning and e-mailing it to fca@baxter.com. For questions contact the Center for One Baxter at 800-422-9837. |  | 882 | 882 | Worldwide Distribution - US Nationwide | Open, Classified |  |  | 882 | Not_Computer
Z-1627-2015 | 71143 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | PHALINX (TM) HAMMERTOE HANDLE, SIZE: LARGE, STYLE: ANGLED, REF 45304014, CONTENTS 1 EACH, RX ONLY |  |  | April 16, 2015 | 2015 | Wright Medical Technology, Inc. | The term DORSAL was laser etched on the incorrect side on some of the angled handles. | Process control | Wright sent an Urgent Medical Device Recall letter dated April 13, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to examine their inventory for the affected product, send it back to Wright Medical as soon as possible. Contact Customers Service at 901-867-4575 for return instructions and replacement inventory. Customers were instructed to complete the attached response immediately, confirming their receipt of this notice and the status of any product(s) on had, and return a copy by email to complaints@wmt.com or fax to 901-867-7401. Customers with questions should call 901-451-6144. |  | 126 | 126 units | Nationwide Distribution | Open, Classified |  |  | 497 | Not_Computer
Z-1555-2015 | 70916 | No | KPR | Stationary x-ray system. | Radiology | Radiology | 510(k) | Philips Healthcare DuraDiagnost X- Ray |  |  | February 09, 2015 | 2015 | Philips Medical Systems, Inc. | The system is designed to emit a beep upon termination of an exposure. However, if the system has been powered on for over 12 hours, the system will no longer emit this signal. This is a failure to comply with 21CFR 1020.31(j). | Software design (manufacturing process) | Philips Healthcare Planned action: 1. Philips healthcare will contact customers and initiate a software update to correct the software deficiency causing the problem.2. Philips engineers will install the software.3. Philips engineers will perform testing to ensure the software update was effective.4. Corrections are to be made free of charge5. Corrective actions should be completed by August 15, 2015For further questions call (978) 659-7892 |  | 3 | 3 DuraDiagnost | Worldwide Distribution - Nationwide Distribution and the countries of Australia, Austria, Argentina, Canada, China, Chile, Colombia Czech Republic, Great Britain, Germany, Egypt, Finland, France, India, Ireland, Japan, Malaysia, Macedonia, Mexico, Netherlands, New Zealand, Nepal, Peru, Poland, Russia, Saudi Arabia Singapore, Switzerland, Sweden, Spain, United Arab Emirates and United Kingdom. | Open, Classified |  |  | 57 | Computer
Z-1629-2015 | 71006 | No | CEE | Thyroxine-binding globulin test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | IMMULITE/IMMULITE 1000 Systems TBG, catalog LKTB1 |  |  | April 09, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | IMMULITE/IMMULITE 1000 TBG catalog LKTB lot 257 is failing the adjustment slope instrument parameter of &lt;1.80. The resulting slope of the kit is &gt;1.80 resulting in the potential to produce depressed TBG patient values. | Unknown/Undetermined by firm | Siemens sent an Urgent Field Safety Notice dated April 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be Taken by the CustomerDiscontinue use and discard kit lot listed in Table 1.Review your inventory of these products to determine your laboratory's replacement needs.Please review this letter with your Medical Director.Complete and return the Field Correction Effectiveness Check Form attached to this letter within thirty (30) days.Please retain this letter with your laboratory records, and forward this letter to those who may have received this product.We apologize for the inconvenience this situation may cause. If you have any questions, please contact your Siemens Customer Care Center or your local Siemens technical support representative. For further questions call (914) 631-8000. |  | 155 | 155 kits | Worldwide Distribution to the countries of : Denmark, Spain, Great Britain, Germany, Bosnia Herzeg, China and Russian Federation. | Terminated | March 22, 2017 | 713 | 155 | Not_Computer
Z-1635-2015 | 70994 | No | JWH | Knee joint patellofemorotibial polymer/metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Stryker Orthopaedics Triathlon femoral component. |  |  | August 15, 2014 | 2014 | Stryker Howmedica Osteonics Corp. | The peel strength of the inner blister may have been below internal validated requirements. The strength of the seal is directly related to long term shelf life of the sterile barrier. | Under Investigation by firm | The firm, Stryker Orthopaedics, notified their Branches/Agencies via email on 8/15/2014 and sent an "Urgent Product Recall" letter dated 8/15/2014 with a Product Recall Acknowledgement Forms to their Branches/Agencies and consignees. The letter describes the product, issue, potential hazards and actions to be taken. The consignees were instructed to complete and return the attached Product Recall Acknowledgment Form within 5 days via fax to 855-251-3635; and return the affected product to the attention of Regulatory Compliance, Stryker Orthopaedics, 325 Corporate Drive, Mahwah NJ 07430.If you have any questions, feel free to contact Sr. Regulatory Compliance Specialist to 201-831-5826. |  | 1147 | 1,147 in total | Worldwide Distribution: US (nationwide) including PR and countries of: Australia, Canada, China, Brazil, France, Hong Kong, India, Italy, Mexico, Netherlands, Singapore, South Africa, Sweden, and Spain. | Open, Classified |  |  | 1147 | Not_Computer
Z-1632-2015 | 71105 | No | IYN | Ultrasonic pulsed doppler imaging system. | Radiology | Radiology | 510(k) | EPIQ 7 Ultrasound System versions 1.3.2 or lower, WITHOUT the Pediatric Cardiology option, Model: EPIQ 7G, EPIC 7C, EPIQ 7W. Catalog Number: 795200 / 795201 |  |  | April 14, 2015 | 2015 | Philips Ultrasound, Inc. | When EPIQ 7 Ultrasound System, WITHOUT the Pediatric Cardiology option, is set up to "Metric" and weight and/or height is entered, a unit conversion error may result in the incorrect display of their values and incorrect calculation of the patients Body Surface Area (BSA). | Software design | Philips sent the Urgent-Medical Device Correction-Philips Model EPIQ5 Ultrasound system and the Urgent-Medical Device Correction-Philips Model EPIQ7 Ultrasound system letters, both dated 2015 APR 14, to their US consignees. Philips sent the Urgent-Field Safety Notice-Philips Model EPIQ5 Ultrasound system letter and the Urgent-Field Safety Notice-Philips Model EPIQ7 Ultrasound system letter (both letters dated 2015 APRIL 14) to their foreign consignees. The EPIQ5 or EPIQ7 Ultrasound systems WITHOUT the Pediatric Cardiology option are affected. Philips issued two Field Change Orders (FCO) to their Field Service Engineer to install upgraded software: -FCO 79500340 for EPIQ5 -FCO 79500341 for EPIQ7 Customers with questions can call Philips Customer Service at 1-800-722-9377.Field Service Engineers (FSE) to install software version 1.3.3 per FCO 79500340 instructions for EPIQ5 5 and per FCO 79500341 instructions for EPIQ7. |  | 1692 | 1692 units total (902 units in the US and 790 units outside the US) | Worldwide Distribution | Open, Classified |  |  | 2800 | Computer
Z-1630-2015 | 70828 | No | NLM | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | Carter-Thomason II Port Closure System; Model Number(s) CTI-1015P, for use in endoscopic and laparoscopic surgery. |  |  | March 20, 2015 | 2015 | CooperSurgical, Inc. | The pad printing around the suture entry holes on the 15mm Suture Guide included in the kit may not have been appropriately manufactured resulting in the padprinting flaking off around the entry holes. | Component design/selection | Consignees were notified of the recall by letter sent via FedEx dated March 20, 2015. The letter instructs consignees to do the following: discontinue use of the affected product, complete the attached acknowledgement and receipt form to return the product and schedule a replacement. |  | 430 | 430 units | US Nationwide distribution. | Terminated | July 29, 2015 | 131 | 430 | Not_Computer
Z-1636-2015 | 71127 | No | OYD |  |  | General &amp; Plastic Surgery | PMA | MelaFind Device, Model MF 100 US, optical diagnostic device for melanoma detection. |  |  | April 02, 2015 | 2015 | MELA Sciences, Inc. | The probability and histogram data within the Melafind's device displayed user interface lacks approval (not included in the PMA supplement). | Software design | The probability and histogram data within the Melafind's device displayed user interface is not included in the PMA supplement |  | 65 | 65 units | AL, CA, CT, CO, FL, HI, IL, MA, MD, MI, NH, NJ, NY, OH, PA, RI, SC, TN, TX, VA and WA | Open, Classified |  |  | 65 | Not_Computer
Z-1642-2015 | 71193 | No | GCJ | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | DeRoyal(R) Medium Specimen Retrieval Bag, REF 5500-000-003, Rx Only |  |  | April 17, 2015 | 2015 | DeRoyal Industries Inc | The firm received reports of specimen retrieval bags tearing, ripping, and coming apart during use. | Device Design | DeRoyal sent an Urgent: Voluntary Recall letter dated April 17, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to 1. Check your inventory for the product and specific lot number from the attached Recall Response Form.2. If you have affected inventory, please destroy it and dispose of it according to your facilitys procedures.3. Record all information on the included Recall Response and Notice of Destruction forms.4. If you purchase through a distributor, please indicate their name/location on the Notice of Destruction Form so that we may issue credit appropriately. 5. Return both forms to recalls@deroyal.com or 865-362-3716 (fax) even if you no longer have affected inventory.6. Once we have received your Recall Response and Notice of Destruction forms, you will receive credit for destroyed product as indicated on the form.7. If you are a distributor of these products, please forward this recall notice on to your end-users, or provide us with your customer listing and we will gladly contact them for you.For questions regarding this recall, please contact your DeRoyal Sales Representative or DeRoyal Customer Service at 1-800-251-9864. |  | 5359 | 5359 units | US Nationwide Distribution | Terminated | December 29, 2015 | 256 | 5359 | Not_Computer
Z-1643-2015 | 71117 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Pinnacle3 Software Version 10.0, Model 459800200841. |  |  | April 23, 2015 | 2015 | Philips Medical Systems, Inc. | A problem (the dose computed in planning mode is calculated incorrectly ) has been detected in the Philips Pinnacle3 software version 10.0 that, if it were to re-occur, could pose a risk for patients or users. Specifically, the dose engine is being passed the wrong snout position. The snout position is used to determine the penumbra of the beam which includes the calculation of the source si | Software design | Philips Healthcare sent a "Field Safety Notification PROST-T-300-01-06" letter dated 20 April 2015 to their customers. The letter described the Problem Description, Hazard Involved, How to Identify Affected Products, Action to be Taken by Customer / User, Actions Planned by Philips and Further Information and Support (Contact Philips Customer Care Solutions Center at 1-800-722-9377). |  | 4 | 4 | US Nationwide Distribution in the states of MO, FL, OH | Open, Classified |  |  | 4 | Computer
Z-1639-2015 | 71178 | preservative, cytological | LEA | General purpose reagent. | Pathology | Pathology | 510(K) Exempt | Mucolytic Agent 1 pint/473 ml; 1 gallon/3.8 l |  |  | April 29, 2015 | 2015 | Richard-Allan Scientific Company | Richard Allan Scientific (RAS) is initiating a recall of Mucolytic agent and Mucolexx following a customer complaint reporting that a lab had found fungal contamination in 2 lots of Mucolexx. | Material/Component Contamination | Richard Allan Scientific (RAS) initiated a voluntary recall of Mucolexx and Mucolytic Agent due to fungal contamination on 5/1/2015 via certified mail to all customers who received the affected. Customers were asked to return a repsonse form, and dispose or return all recalled products. Please contact Sarah Rickert at 269-544-5628 or sarah.rickert@thermofisher |  | 152478 | Product 76166- 116; Product Code 76167-29 | Worldwide Distribution - US Nationwide in the states of: AR; AZ; CA; CO; FL; GA; I A; IL; IN; KY; MA; MN; MO; MS; MT; NC; ND; NJ; NV; NY; OH; OK; PA; PR; SD; TN; TX; VA; WA; WI and the countries of: Bangladesh, Canada, China, Costa Rica, Dubai, France, Germany, Greece, India, Ireland, Israel, Italy, Japan, Korea, Netherlands, Panama, Poland, Portugal, Saudi Arabia, South Africa, Spain, Sweden, Switzerland, Thailand, Turkey, and UK. | Terminated | March 09, 2016 | 315 | 10142853 | Not_Computer
Z-1641-2015 | 71215 | No | GDW | Implantable staple. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ENDOPATH ECHELON" Vascular White Reload for Advanced Placement Tip, |  |  | February 26, 2015 | 2015 | Ethicon Endo-Surgery Inc | During an internal inspection of the production process the firm discovered an issue which may cause the cartridge to deploy an incomplete staple line. | Software in the Use Environment | Ethicon sent an Urgent Medical Device Recall letter dated February 27, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to examine your inventory immediately to determine if you have recalled product on hand, remove and quarantine the recalled product, communicate the issue to relevant operating room or materials management personnel or anyone else in your facility whoneeds to be informed, if any product included in this recall has been forwarded to another facility, contact that facility to arrange return, and complete the Business Reply Form (BRF) (Attachment B) confirming receipt of thisnotice within three (3) business days. The BRF may be sent to Stericycle by fax at 1-800-807-5967or by email at ees8760@stericycle.com. Please return the BRF even if you do not have affected product.For clinical or product support, please contact your local sales representative orcall our Customer Support Center at 1-877-ETHICON (1-877-384-4266) (7:30 a.m.  6:30 p.m. EST). If you need additional shipping labels or a communications package, contact Stericycle at 1-877-643-8419 (8:00 a.m.  5:00 p.m. EST) and reference Event8760. |  | 6744 | 6,744 devices | Worldwide Distribution - US Nationwide in the states of: AR, AZ, CA, CO, DE, FL, GA, IA, IL, KS, KY, LA, MA, ME, MI, MN, MO, NC, NE, NJ, NY, OH, OK, OR, PA, SC, TN, TX, UT, VA, WA &amp; WI and the country of Hong Kong. | Terminated | November 17, 2015 | 264 | 6744 | Not_Computer
Z-1645-2015 | 71101 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | SOMATOM Emotion 16, the intended use of computed tomography is to produce cross-sectional images of the body by computer reconstruction of x-ray transmission data from either the same axial plane taken at different angels or spiral planes taken at different angles. |  |  | April 20, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Software bugs in VC20B SP0a or SP1 software versions may cause issues that could make it necessary to rescan patients. Syngo Main UI may crash if the Patient Browser is scrolled with arrow keys of the keyboard. Tomo images may freeze on the screen when zoomed in or out under the mode of CAREVision. Sporadic displaying error in WorkStream4D application. DB may lock under heavy and multitasks. | Software design | Siemens sent an Customer Safety Advisory Notice dated April 20, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. A remote software update is scheduled for release that will resolve the issue. For further questions, please call (610) 219-6300. |  | 75 | 75 | Nationwide Distribution including Puerto Rico. | Open, Classified |  |  | 75 | Computer
Z-1656-2015 | 71177 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT. |  |  | March 16, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm was informed that while raising the patient couch on the system to perform an exam, the couch unexpectedly descended to the lowest point without being commanded to do so. | Employee error | On 3-18-2015 the firm sent Urgent Field Safety Notices to their customers. |  | 153 | 153 units total | Distributed in China. | Open, Classified |  |  | 153 | Not_Computer
Z-1663-2015 | 71022 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Liquid Waste Bottle Kit, Catalog#10386287. |  |  | March 05, 2015 | 2015 | Siemens Medical Solutions Diagnostics | The bottle assemblies were manufactured with a quality issue where in some cases the smaller diameter opening on both IMMULITE 2000 and IMMULITE 2000XPi Water Bottles and Liquid Waste Bottles is undersized and/or deformed, preventing the cap from closing or fastening securely to the bottle. | Nonconforming Material/Component | Siemens Healthcare Diagnostics sent an Urgent Device Correction Letter/Field Correction Effective Check Form dated 3/5/15 via FedEx to their affected customers. |  | 15 | 15 bottles | Worldwide Distribution. | Open, Classified |  |  | 1417 | Not_Computer
Z-1669-2015 | 71115 | No | JAA | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | SIEMENS Luminos Agile Max; a universal imaging system for radiographic and fluoroscopic studies. Using either film cassettes or a digital mobile flat detector, it can perform a range of applications including general R/F, angiography and pediatric examinations. Luminos Agile is applicable to emergency treatment on an outpatient basis, as well as for bedside examinations. |  |  | April 13, 2015 | 2015 | Siemens Medical Solutions USA, Inc | Study and all acquired images deleted when using systems with software version VE10E. Acquisitions have to be repeated. Connection with portable detectors (MAX wi-D, MAX mini) is sporadic in image recovery. Pressing undo button on the RAD subtask card or previous series button on the Image subtask card during image readout. Sporadically, during an automatic or a manual RIS update. | Software design | A safety advisory notice, dated April 13, 2015, was sent to end users with affected devices to alert them about the issue and associated risks,and to provide instructions for avoiding potential risk. A new software version will be released to permanently resolve the issue. |  | 133 | 133 total | Nationwide Distribution. | Terminated | March 20, 2017 | 707 | 133 | Computer
Z-1664-2015 | 71256 | joint, ankle, external brace | ITW | Limb orthosis. | Physical Medicine | Physical Medicine | 510(K) Exempt | Elastic Ankle Support. |  |  | April 13, 2015 | 2015 | United Surgical Associates | The product does not bear the caution label regarding latex content. Potential for Allergic skin reaction | Error in labeling | On 4/13/2015 United Ortho emailed the affected distributor regarding the need for correction due to latex in the product and that updated labeling and instructions would be provided. United-Ortho personnel will travel to Hawaii and accompany the distributor to label the product at the sites to which it was distributed. Records will be maintained at each location. |  | 1278 | 1278 | Domestic: HIInternational; None | Terminated | July 14, 2015 | 92 | 3294 | Not_Computer
Z-1677-2015 | 71059 | No | OCC | Respiratory viral panel multiplex nucleic acid assay. | Microbiology | Microbiology | 510(k) | Alere i Influenza A &amp; B |  |  | April 22, 2015 | 2015 | Alere Scarborough, Inc. dba Binax, Inc. | High invalid rate and an increased risk of false negative results with certain lots of the Alere i Influenza A &amp; B kit. | Under Investigation by firm | Alere sent an Urgent Medical Device Recall letter dated April 23, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to stop using the product and discard any unused product. Customers asked to complete and return the Verification form within 10 days to confirm receipt of the recall notification and to request credit.For questions about the information contained in the notification, contact Alere Technical Services, Monday through Friday, 8am to 8pm EST. Telephone: (855) 731-2288 or E-mail: t s.scr@alere.com.For questions regarding the status of your credit, please contact your AI ere i distributor or E-mail: ssc.alerefieldactions@alere.com. |  | 216 | 213 kits plus 3 - evaluation use only kits | Worldwide Distribution - US Nationwide in the states of AL CA CT FL IA ID IL IN KY LA MD MO MS NC NE NJ OH PA SC TN TX and the countries of: Hong Kong and Singapore. | Terminated | January 14, 2016 | 267 | 216 | Not_Computer
Z-1674-2015 | 71067 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | EXPEDIUM Spine System/Synapse System; MODEL: EXPEDIUM Offset Conical Lock Nut 8mm, CATALOG NO. : 1754-91-150, UDI NUMBER (if applicable): GTIN: 10705034032467 |  |  | April 21, 2015 | 2015 | DePuy Spine, Inc. | Device does not mate as intended. | Process design | Consignees were notified of the recall via letter sent by overnight shipment for delivery on April 23, 2015. The letter instructs distributors and customers to review their inventory and return any products with the listed lot number. |  | 183 | 183 devices | US: PA and OUS: SE, BE, IT, PT | Completed |  |  | 183 | Not_Computer
Z-1678-2015 | 71110 | No | BSY | Tracheobronchial suction catheter. | Anesthesiology | Anesthesiology | 510(K) Exempt | KimVent Turbo-Cleaning Closed Suction System for Adults, 10 F T-Piece (Product Code 227105); KimVent Turbo-Cleaning Closed Suction System for Adults, 12 F, DSE (Product Code 2271603); KimVent Turbo-Cleaning Closed Suction System for Adults, 12 F, T-Piece (Product Code 227135). |  |  | April 14, 2015 | 2015 | Halyard Health, Inc | The internal blue flapper valve of certain KimVent Turbo-Cleaning Closed Suction System for Adults may have a potential defect. If present, the defect potentially allows air to leak out of the closed ventilator circuit. | Device Design | Halyard Health Regulatory Affairs sent an Urgent Voluntary Product Recall letter dated April 16, 2015, to all affected consignees. Consignees were instructed to quarantine and discontinue use of the recalled product. Specific return instructions were provided in the Notice. The consignees were also requested to provide a response to to the recall notification. Wholesalers were instructed to follow their internal recall procedures to identify and address affected products. Consignees with questions were instructed to call Halyard Health Regulatory Affairs at 470-448-5700. |  | 220 | 11 cases each containing 20 individual endotracheal tubes | Nationwide Distribution including CA, MA, MI, NC, TX, VA , and WA. | Terminated | July 07, 2015 | 84 | 220 | Not_Computer
Z-1672-2015 | 71085 | No | NJL |  |  | Orthopedic | PMA | LCS COMPLETE RPS Femoral Implants, various sizes and orientation. Knee prosthesis component for orthopedic surgery. |  |  | March 04, 2015 | 2015 | DePuy Orthopaedics, Inc. | The LCS COMPLETE RPS Knee System has higher rates of revision in Australia when the native patella is not resurfaced when compared to other unresurfaced PS implants. If the LCS COMPLETE RPS Knee System is implanted and the native patella is not resurfaced, the patient may experience patello-femoral pain, potentially requiring a secondary procedure. | Unknown/Undetermined by firm | On 3/4/2015, URGENT INFORMATION - DEVICE CORRECTION NOTICE for the LCS COMPLETE RPS Knee System and on 3/19/2015 REVISED URGENT INFORMATION - DEVICE CORRECTION NOTICE for the LCS COMPLETE RPS Knee System (Revised) notifications were sent to the affected distributors and medical professionals. All distributors were notified via electronic mail, and all medical professionals with affected inventory were also notified by distributor or Sales Consultant via in person delivery or mail or electronic mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. For questions, please contact Kim Earle, Recall Coordinator, at 574-371-4917 (M-F; 8 a.m. to 5 p.m. EDT). |  | 19059 | 19,059 | Worldwide Distribution-US (nationwide) and the countries of CANADA, AUSTRALIA, AUSTRIA, BELGIUM, UK, FRANCE, GREECE, IRELAND, ISRAEL, ITALY, KOREA, LUXEMBOURG, MALAYSIA, NETHERLANDS, NEW ZEALAND, NORWAY, SOUTH AFRICA, SWITZERLAND, THAILAND. | Terminated | October 30, 2015 | 240 | 47791 | Not_Computer
Z-1676-2015 | 71229 | No | NJT | Immunohistochemistry reagents and kits. | Hematology | Pathology | 510(K) Exempt | D2-40 (Lymphatic Marker); For In Vitro Diagnostic; |  |  | May 06, 2015 | 2015 | Biocare Medical, LLC | Firm discovered that incorrect Instructions for Use were shipped with one lot of D2-40. | Labeling mix-ups | Biocare Medical sent a Medical Device Notification letter dated May 6, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed to complete all actions listed on the enclosed Return Response Form. |  | 19 | 19 units | Worldwide Distribution - US Nationwide in the states of: MI, SC, NC, AZ, CA, MO, WI and Hungary, Spain and Austria. | Terminated | June 04, 2015 | 29 | 19 | Not_Computer
Z-1675-2015 | 70959 | No | MHX | Arrhythmia detector and alarm (including ST-segment measurement and alarm). | Cardiovascular | Cardiovascular | 510(k) | Panorama Patient Monitoring Network, Multi- Parameter Patient Monitor (with Arrhythmia Detection and Alarm. Part number 0998-00-0708-01. |  |  | March 12, 2015 | 2015 | Mindray DS USA, Inc. dba Mindray North America | Software anomaly in the Panorama System software version 8.9 that manifests when the Panorama is in use with Mindray Passport V Monitor. Arrhythmia alarms which were previously displayed on both the Passport V and Panorama will not be displayed on the Panorama following a restart of either system (due to communication loss, restart, etc.). | Software design | Mindray sent a Corrective Action letter dated March 12, 2015 to their affected customer via FedEx. The letter identified the affected product, problem and actions to be taken. For questions call 201-995-8407. |  | 4 | 4 units | Worldwide Distribution - US Nationwide in the state of TX. | Open, Classified |  |  | 4 | Computer
Z-1609-2015 | 70920 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | CareFusion AVEA ventilator, all models. |  |  | April 03, 2015 | 2015 | Carefusion 211 Inc dba Carefusion | A pressure transducer failure can develop, activating a false Extended High Ppeak or Circuit Occlusion Alarm. The Safety Valve will open to ambient air and the unit will stop ventilating, allowing spontaneous breathing patients to breathe. The malfunction may delay initiation or cease ventilation. Nonbreathing patients will need manual ventilation or to be connected to another ventilator. | Process control | A customer letter dated 4/21/15 will be sent to all customers who received the AVEA Ventilator, all models, to inform them that CareFusion has identified a potential risk associated with AVEA Ventilator caused by a malfunction of a pressure transducer. The malfunctioning pressure transducer is detected and by design ceases ventilation and opens the safety valve to atmosphere allowing spontaneous breathing of patients capable of doing so. The non-breathing patient would need to be manually ventilated.The letter informs the customers of the issues, the potential risk, and the actions to be taken. Customers are instructed to contact CareFusion Recall Support Center at (888) 562-6018, (858) 617-5300, or email supportcenter@carefusion.com for recall related questions.Customers are instructed to contact CareFusion Technical Support/Customer Advocacy at (800) 231-2466, (714) 283-2228, or email support.vent.us@carefusion.com for Product Technical Support Adverse Event Reporting.CareFusion posted a press release on their website on 5/27/15 and it was sent to PRNewswire. The press release provided an update on a global voluntary recall that was initiated on April 21, 2015 to address an issue with certain units of AVEA ventilators. The press release provide the problems and the actions to be taken. The press also list the countries that are affected by the recall. |  | 15905 | 15,905 units total (5,670 units in US) | Worldwide Distribution. US Nationwide including Guam and Puerto Rico, Canada, Albania, American Samoa, Argentina, Australia, Austria, Bahamas, Bahrain, Bangladesh, Barbados, Belgium, Belize, Bermuda, Bolivia, Bosnia Herzegovina, Brazil, Bulgaria, Chile, China (Peoples Rep), Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Dominican Republic, Ecuador, Egypt, El Salvadore, Estonia, France, Georgia, Germany, Greece, Guatemala, Honduras, Hong Kong, Hungary, India, Indonesia, Iran, Ireland, Israel, Italy, Japan, Jordan, Korea (Rep), Kuwait, Latvia, Lebanon, Libyan Jamahiriya, Lithuania, Malaysia, Mexico, Moldova, Morocco, Nepal, Netherlands, Nicaragua, Norway, Occupied Palestinian Territory, Oman, Pakistan, Panama, Paraguay, Peru, Philippines, Poland, Portugal, Qatar, Romania, Russia, Rwanda, Saudi Arabia, Singapore, Slovakia, Slovenia, Spain, South Africa, South America, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Uruguay, Venezuela, Vietnam, and Yemen. | Open, Classified |  |  | 15905 | Not_Computer
Z-1671-2015 | 71185 | No | KNW | Gastroenterology-urology biopsy instrument. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EliteCore Full Core Biopsy Device. EliteCore 18G, 20cm with and without HiLiter tip, 15 cm with and without HiLiter tip, and 10 cm with and without HiLiter tip. |  |  | April 24, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Potential for the device cannula to overthrow past the intended length. | Device Design | Stryker sent an Urgent Medical Device Recall Notification letter dated May 5, 2015, to all affected customers via Fed Ex overnight. Customers were instructed to qaurantine any product found, return the repsonse form even if there is no product on hand, forward the letter if further distribution of recalled product occurred, return recalled product to Stryker. Customers were instructed to 269-389-2458 with any questions. |  | 1420 | 1,420 Units | US Nationwide Distribution including AL, AR, AZ, CA, CO, FL, GA, IL, IN, KS, KY, LA, MI, MN, MO, MS, NJ, NM, NV, NY, OK, OR, PA, TN, TX, UT, VA, and WI. | Terminated | October 15, 2015 | 174 | 1420 | Not_Computer
Z-1679-2015 | 71289 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | RaySearch RayStation 2.5, 3.0, 3.5 and 4.0; Radiation Therapy Treatment |  |  | May 08, 2015 | 2015 | RAYSEARCH LABORATORIES AB | RaySearch Laboratories AB became aware of an issue with the dose calculation for some imported VMAT plans where controls points are defined with wide gantry angle spacing. | Software Design Change | RaySearch sent a Device Correction letter dated May 8, 2015 via e-mailed to their customers. The letter identified the affected product, problem, actions to be taken by users and contact information. |  | 539 | Domestic: 539 units | US Nationwide Distribution | Terminated | April 11, 2017 | 704 | 539 | Not_Computer
Z-1683-2015 | 70672 | No | OOE | Ophthalmic laser. | Ophthalmic | Ophthalmic | 510(k) | Catalys Precision Laser System-Catalys - U - US product; Catalys-1 - International product; software version 3.00.05 |  |  | February 20, 2015 | 2015 | Optimedica Corporation | Software anomaly on OptiMedica CATALYS System version 3.00.05 that may result in incorrect parameters provided for cataract incision surgeon templates and both eyes have the same parameters. | Software design | Abbott sent an Advisory Notice dated March 5, 2015, vis Fed Ex overnight to all affected customers. Out of State accounts will be notified through their distributor after letters are translated into local language.Recall notification letters advise customers of issue with CATALYS System using software version 3.00.05. The letter actually covers two issues, the cataract templates being the same for both eyes, and a loss of suction during treatment that may result in scoring the cornea.In both cases the issue is explained and users are advised to use mitigation listed to prevent injury. AMO is developing an update to software that will correct these issues and users will be advised when it is available for installation.All users of the device should be made aware of the issues and each surgeon is requested to sign the acknowledgement form to verify their receipt and understanding. The acknowledgement for should be returned to Abbott Medical Optics per instructions. Questions in the US may be directed to AMO Technical support at 1-855-662-0911. |  | 171 | 171 distributed Worldwide | Worldwide Distribution - US Nationwide, Germany, Switzerland, Australia, Spain, Japan, Great Britain, France, Netherlands, Austria, Singapore, South Korea, India, Brazil, Italy, Belgium, Egypt, Colombia, Hong Kong, and Canada. | Terminated | December 08, 2015 | 291 | 171 | Computer
Z-1689-2015 | 70681 | Yes | DTZ | Cardiopulmonary bypass oxygenator. | Cardiovascular | Cardiovascular | 510(k) | Quadrox-iD Adult Diffusion Membrane Oxygenator |  |  | March 05, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | A potential issue was detected that could result in blood leakage from the gas outlet port in all lots of Quadrox-iD Adult Diffusion Membrane Oxygenator with Bioline and Softline Coatings, Custom Tubing Packs and Diffusion Membrane HLS Set Advanced and HIT Set Advanced Models. | Nonconforming Material/Component | Maquet sent a "Field Safety Notice" accompanied by a Response Form to their affected customers on 3/5/2015 via FedEx. The Field Safety Notice identified the affected products/Part numbers and actions to be taken by the user. The customers were asked to complete the Field Safety Notice Response Form to acknowledge that they have received the notice. The completed notice can be faxed to 1-973-396-3607 or sent via email to fieldactions@maquet.com. For any questions, customers can contact their Maquet Sales Representative or Maquet Customer Service at 1-888-627-8383 (press option 2, followed by option 2), Monday through Friday, between the hours of 8:00 a.m. and 6:00 p.m. EST. |  | 10199 | 10,199 units | Worldwide Distribution - US Nationwide and the countries of: Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Canada, China, Colombia, Croatia, Cyprus, Czech Republic, Denmark, Estonia, Finland, France, Germany, Great Britain, Greece, Hungary, Iceland, Ireland, Israel, Italy, Japan, South Korea, Latvia, Lithuania, Luxembourg, Malta, Mexico, Netherlands, Norway, Poland, Portugal, Romania, Saudi Arabia, Serbia Montenegro, Singapore, Slovakia, Slovenia, Spain, Sweden, Switzerland, Taiwan, Thailand, Turkey and Uruguay | Open, Classified |  |  | 10199 | Not_Computer
Z-1691-2015 | 70772 | Yes | OAE |  |  | Cardiovascular | PMA | SMARTABLATE RF Generator System, Catalog No. M490006, M490007, M4900106, and M4900107. |  |  | March 19, 2015 | 2015 | Biosense Webster, Inc. | Biosense Webster is recalling the SMARTABLATE Generator System because Biosense Webster has received two complaints reporting that during an RMT procedure the SMARTABLATE RF Generator inadvertently delivered RF energy. | Device Design | Biosense Webster sent an Urgent Field Safety Notification Letter dated March 18, 2015 via FedEx to all their customers that received the affected SMARTABLATE Generator System. The letter informs the customers of the problems identified and the actions to be taken. Biosense Webster is taking precautionary safety measures of reinforcing the Warnings in the Instructions for Use (IFU) for the SMARTABLATE System Foot Pedal. Customers with questions related to the are instructed to contact their local Biosense Webster sales representative or call (866) 473-7823, Monday through Friday from 7:00AM to 8:00PM EST. Customers with questions related to the voluntary field removal certification form and its return, are instructed to contact Elie Ghattas Field Action Coordinator at (909) 839-7281. |  | 115 | 115 units | Worldwide Distribution - US Nationwide in the state of: CA, IL, KY, NJ, PA and the countries of: Czech Republic, Denmark, Finland, France, Germany, and Sweden. | Terminated | June 17, 2015 | 90 | 115 | Not_Computer
Z-1681-2015 | 71047 | No | LKK |  |  | General Hospital | PMA | Medtronic SynchroMed II Implantable Infusion Pumps, models 8637-20 and 8637-40. Sterile. |  |  | April 10, 2015 | 2015 | Medtronic Neuromodulation | Medtronic is conducting a recall of a specific subset of Model 8637-20 and 8637-40 SynchroMed II implantable drug pumps because the audible alarm could cease to function. | Process control | Customers were sent a Medtronic "Urgent Medical Device Removal" letter dated April 2015. The letter described the problem and the product involved in the recall. The letter described the "Required Actions" which included to quarantine and return the product, complete and return the Customer Confirmation Form, and to share the notification with other customers. For questions, they can contact Medtronic Neuromodulation Technical Services at 1-800-707-0933.A second letter was sent to Healthcare Professional. The "Urgent Medical Device Correction" letter dated April 2015 described the Nature of the Device Issue, Scope and Likelihood of the Issue, Recommendations, Patient Management Recommendations and Additional Information.Advise consignees to report any malfunction or adverse event related to a device to Medtronic Neuromodulation Technical Services, at 1-800-707-0933 available weekdays from 7 am - 6pm Central Time, and to FDA'ss MedWatch Program (www.fda.gov/medwatch). |  | 1901 | 1901 (1576 US, 325 OUS) | Worldwide Distribution - US Nationwide including DC AND PR and in the countries of : Argentina, Austria, Belgium, France, Germany, Ireland, Italy, Japan, Luxembourg, Malta, Netherlands, Norway, Panama, Poland, Portugal, Russian Federation, San Marino, Saudi Arabia, South Korea, Spain, Switzerland, and United Kingdom. | Terminated | November 23, 2015 | 227 | 1901 | Not_Computer
Z-1680-2015 | 71044 | No | CCL | Oxygen gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | D-fend, (p/n 876446-HEL), and D-fend+ (p/n 881319-HEL) water trap. Used with Compact Airway Modules E/M-C(Ai)O(V)(X) and Cardiocap/5 monitors. |  |  | April 24, 2015 | 2015 | GE Healthcare | A potential safety issue due to loss of gas monitoring associated with occlusion of specific D-Fend / D-Fend+ water traps used with Compact Airway Modules E/M-C(Ai)O(V)(X) and Cardiocap/5 monitors. This water trap occlusion could cause a loss of respiratory airway gas monitoring. | Nonconforming Material/Component | GE Healthcare sent an "Urgent Medical Device Correction" letter GEHC Ref#36111 dated April 24, 2015. The letter was addressed to Healthcare Administrator / Risk Manager, Chief of Nursing, &amp; Director of Biomedical Engineering, The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction and Contact Information. A confirmation form is asked to be returned and GE will replace all affected product. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 2948 | 5,018 (US = 2,948 / OUS = 2,070) | Worldwide Distribution - US Nationwide including Washington DC, Puerto Rico. and the countries of: Algeria, Argentina, Australia, Austria, Bahrain, Belgium, Bolivia, Brazil, Canada, Chile, China, Colombia, Costa Rica, Croatia, Cyprus, Czech Republic, Denmark, Dominican Republic , Egypt, Estonia, Finland, France, French Polynesia, Germany, Greece, Guadeloupe, Guernsey , Guernsey , Hong Kong, Hungary, Iceland, India, Indonesia, Ireland, Israel, Italy, Jamaica, Japan, Jersey, Jordan, Korea (Republic of), Kuwait, Lithuania, Malaysia, Mexico, Netherlands, New Zealand, Norway, Oman, Pakistan, Panama, Peru, Philippines, Poland, Portugal, Russia, Saudi Arabia, Singapore, Slovenia, South Africa, Spain, Sweden, Switzerland, Taiwan, Thailand , Turkey, Ukraine, United Arab Emirates, United Kingdom, Vietnam, | Open, Classified |  |  | 2948 | Not_Computer
Z-1684-2015 | 70659 | No | LKK |  |  | General Hospital | PMA | CODMAN 3000 Series Refill Set- Huber needles. |  |  | March 05, 2015 | 2015 | Codman &amp; Shurtleff, Inc. | Needle is leaking from hub/needle cannula connection | Process control | Codman Neuro issued recall letter on 3/05/15 requested to immediately check all inventory to determine if you have affected product. Complete the Acknowledgement Form and follow the instructions to report your inventory status and return affected product for credit and replacement (if requested). General inquiries, customer service or field action logistics - please contact Codman Neuro Quality at: CodmanProductComplaints@DPYUS.JNJ.com or 866-491-0974. |  | 5463 | 5463 units | Worldwide Distribution-US (nationwide) and the countries of Canada, France, Germany, and Switzerland. | Open, Classified |  |  | 5463 | Not_Computer
Z-1687-2015 | 67989 | Yes | DTQ | Cardiopulmonary bypass heart-lung machine console. | Cardiovascular | Cardiovascular | 510(k) | 100/120V AC, Advanced Perfusion System Platform (APS) The TerumoÂż Advanced Perfusion System 1 is indicated for use for up to 6 hours on the extracorporeal circulation of blood for arterial perfusion, regional perfusion, and cardiopulmonary bypass procedures, when used by a qualified medical professional who is experienced in the operation of this or similar equipment. |  |  | May 01, 2015 | 2015 | Terumo Cardiovascular Systems Corporation | Terumo CVS is implementing field correction activities to address the identified causes of the Terumo System 1 experiencing loss of system power. These activities will be implemented in a phased approach beginning late 2015:- Implementing design changes and upgrades to improve the reliability of power switching and battery backup.- Adding a switch protector to the On/Off power switch to preven | Device Design | On 5/1/2015 Terumo issued an URGENT MEDICAL DEVICE CORRECTION letter to users explaining that Terumo has identified causes of the Terumo System 1 experiencing loss of system power. These activities will be implemented in a phased approach beginning late 2015: - Implementing design changes and upgrades to improve the reliability of power switching and battery backup.- Adding a switch protector to the On/Off power switch to prevent inadvertent shut off.- Replacing the On/Off power switch with a new design to prevent switch failure. A Terumo Field Service Representative will contact users to schedule the field correction activities once the activities are ready to be implemented.CUSTOMER INSTRUCTIONS: Review this Medical Device Correction and assure that all users have received notice of this issue.Âż- Complete and return the enclosed Customer Response Form.Âż- Terumo CVS will contact users to schedule the field correction activities.We encourage you to contact us with any questions or concerns:Terumo CVS Customer Service1-800-521-2818Fax 1-734-741-6149Customer Service Hours:Monday  Friday, 8 a.m.  6 p.m. ET |  | 1713 | 1713 units total | Worldwide Distribution-US (nationwide) including DC and the states of AL , AR , AZ , CA , CO , CT , DE , FL , GA , HI , IA , ID , IL , IN , KY , LA , MA , MD , ME , MI , MN , MO , MS , MT , NC , ND , NE , NJ , NM , NV , NY , OH , OK , OR , PA , SC , TN , TX , UT , VA , VT , WA , WI , and WV, and the countries of Mexico , AUSTRALIA , UNITED ARAB EMIRATES (UAE) , Hong Kong , Indonesia , Singapore , Taiwan , Thailand , Philippines , South Korea , CHILE , Vietnam , Malaysia , BELGIUM , Japan , and CANADA. | Open, Classified |  |  | 1713 | Not_Computer
Z-1685-2015 | 71014 | No | HJK | Polymethylmethacrylate (PMMA) diagnostic contact lens. | Ophthalmic | Ophthalmic | 510(k) | Volk 25D Large Clear BIO Lens with case UPC: 8 11608 01010 6, VOLK OPTICAL INC. The Volk Optical BIO Lenses are intended for visualization during diagnosis and laser therapy of the human retina (fundus) using a binocular indirect ophthalmoscope. |  |  | February 27, 2015 | 2015 | Volk Optical Inc | Mislabeling. The incorrect lot number was engraved on the anodized aluminum ring. | Process control | Volk Optical, Inc, sent an Urgent Medical Device Recall letter dated March 30, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.The Firm requests that you return the device(s) for replacement, at our cost. The recalled device(s) should be returned to Volk with an assigned Return Merchandise Authorization (RMA) number. Please contact Volk as soon as you receive this notification to setup a RMA for your device(s). Please contact Volk at (440) 942-6161 or toll-free at (800) 345-8655 to speak with a Volk RMA Coordinator. To return a recalled device(s) to Volk, please ask Volks RMA Coordinator for the appropriate Volk shipping account number (e.g., FedEx, UPS, etc.). |  | 6 | 6 units | Worldwide Distribution - US Distribution to the states of : MA and VA., and to the countries of Mexico and Sweden. | Terminated | February 19, 2016 | 357 | 6 | Not_Computer
Z-1682-2015 | 71228 | No | KPS | Emission computed tomography system. | Radiology | Radiology | 510(k) | Ingenuity TF Positron Emission Tomography (PET) and Magnetic Resonance Imaging (MRI), 882380, Philips Medical Systems. |  |  | March 11, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | It was noticed that the spectroscopy voxel map did not align with the anatomical images of the phantom. This resulted in incorrect orientation of the Chemical Shift Imaging (CSI) data. | Software in the Use Environment | On 3/11/2015 the firm sent Urgent - Medical Device Correction letters to their customers. The letters identified the affected device, explained the reason for the recall, how to identify affected product, advice on actions to be taken by customer/user, and any actions planned by the firm. Customers were advised to use software provided by Philips to view spectroscopy images; and, not to export and view images using third party software. The firm stated that it will be issuing a software update. If further information or support concerning the issue is needed, customers are asked to contact their local Philips representative or local Philips Healthcare office. |  | 10 | 10 units | Worldwide Distribution -- US, including the states of NY, OH, &amp; TX; and, the countries of Finland, Germany, Japan, Netherlands, South Korea, Spain, &amp; Switzerland. | Open, Classified |  |  | 10 | Computer
Z-1686-2015 | 71296 | No | DXT | Angiographic injector and syringe. | Cardiovascular | Cardiovascular | 510(k) | MEDRADÂż Stellant CT Injector System with Certegra |  |  | September 20, 2013 | 2013 | Medrad Inc dba Bayer R&I; | The root cause investigation showed that the under-volume hazard can occur when the unit has been configured to prime with contrast, but instead primes with saline due to the configuration data not being sent to the injector head. This may create a non-optimal priming condition that could lead to under-infusion. | Software design | Bayer Healthcare sent a Medical Device Field Corrective Action letter to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to identify any affected device. Bayer Service will schedule time to update their software free of charge. Customers with questions were instructed to contact the customer support team at 1-800-633-7237. For questions regarding this recall call 412-767-2400. |  | 133 | 133 | Nationwide Distribution | Terminated | June 16, 2015 | 634 | 133 | Computer
Z-1692-2015 | 70793 | No | CCK | Carbon dioxide gas analyzer. | Anesthesiology | Anesthesiology | 510(k) | Smiths Medical, BCI Advisor Vital Signs Monitor, Product Reorder No 92M774325A. |  |  | March 21, 2015 | 2015 | Smiths Medical ASD, Inc. | The Auxiliary Serial Input/ Output Modes on the Advisor do not function as expected, resulting in data formatting issues orincorrect data being transmitted in or out of the serial port into a data collection system. | Nonconforming Material/Component | Consignee was sent on 3/20/2015 a Smiths Medical "Urgent Medical Device Recall" letter dated 13 March 2015. The letter was addressed to Supplier Quality. The letter described the issue and the product affected by the recall. Requested consignees to quarantine the affected product and to complete and return the Confirmation Form by Fax to 1-800-237-8033 or by email to recall.response@smiths-medical.com within 10 days of receipt of this letter. Customers shall report any issues with these products to Smiths Medical's Global Complaint Department at 1-866-216-8806 or globalcompaints@smiths-medical.com. For questions they can contact Smiths Medical's Customer Service Department at 1-800-258-5361. UPDATE 4-20-2015 Smiths Medical sent an Urgent Medical Device Recall letter via email to the one consignee on April 10, 2015. The letter identified the additional monitors and requested that those devices be quarantined, return the attached confirmation form, and a customer service representative will contact the consignee to arrange for exchange of the affected Advisor monitors for credit or replacement. |  | 4202030 | 5 UPDATE 4-20-2015 to include 10 additional devices | MN | Terminated | November 20, 2015 | 244 | 4202030 | Not_Computer
Z-1637-2015 | 71191 | No | DXZ | Intracardiac patch or pledget made of polypropylene, polyethylene terephthalate, or polytetrafluoroethylene. | Cardiovascular | Cardiovascular | 510(k) | Synovis VASCU-GUARD Peripheral Vascular Patch in the following sizes and product codes: 1x6cm, product code 1504026, 0.8x8cm, product code 1504028, 1x10cm, product code 1504030, and 2x9cm, product code 1504032. VASCU-GUARD is packaged between two pieces of foam within a double sterile barrier pouch system. The contents of the unopened, undamaged package are sterile. VASCU-GUARD is intended for use in peripheral vascular reconstruction including the carotid, renal, iliac, femoral, profunda, and tibial blood vessels and arteriovenous access revisions. |  |  | May 02, 2015 | 2015 | Synovis Surgical Innovations, Inc. | Baxter healthcare is recalling specific product codes of Vascu-Guard Peripheral Vascular Patch due to complaints received for difficulty in distinguishing the smooth from rough surface. | Under Investigation by firm | The firm, Baxter, sent an "Urgent Product Recall" letter dated 5/2/2015 via USPS overnight delivery to its customers. The letter identified the affected product, problem and actions to be taken. The customers were instructed to immediately discontinue use and segregate the affected product, locate and remove all affected product from your facility and to return it to Baxter. Customers are to contact Baxter Healthcare Center, 888-229-0001, for Service to arrange for return and credit. In addition, the customers were to complete and return the Customer Reply Form by faxing to 224-270-5457 or emailing to fca@baxter.com.If you have any questions, contact Director, Quality at 651-796-7543 or email: heidi_drafall@baxter.com. |  | 3974 | 3,974 | Nationwide Distribution. | Terminated | February 18, 2016 | 292 | 3974 | Not_Computer
Z-1695-2015 | 71231 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Brilliance iCT Computed Tomography X-ray system |  |  | April 13, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | A calibrated torque wrench malfunctioned at the supplier leading to possible incorrect torque of x-ray tube mounting fasteners. | Other | Philips sent an Urgent - Field Safety Notice dated April 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Philips Healthcare is recommending to stop the use of your system until the X-ray tube mounting fasteners are replaced. In addition, a Philips Field Service Engineer (FSE) will service your system via aMandatory Field Change Order (FCO 72800638). A Philips FSE will contact youto schedule this appointment.If you need any further information or support concerning this issue, please contact your local Philips representative or your local Philips Healthcare office.This notice has been submitted to the appropriate Regulatory Agency.Philips apologizes for any inconveniences this problem has caused your organization. For further questions please call (440) 483-7600. |  | 2 | 2 | Internationally to the countries of China and Norway. | Terminated | October 01, 2015 | 171 | 2 | Not_Computer
Z-1694-2015 | 71243 | No | NDJ | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | Cortical Screw for M/DN Intramedullary Fixation. Metaphyseal/Diaphyseal Nail System, Orthopedic surgical device. |  |  | April 29, 2015 | 2015 | Zimmer, Inc. | The entire scope of 4.5mm Cortical Screws listed as compatible with the M/DN system is not compatible with the M/DN system. This has the potential to lead to a situation in which the screw would bind upon entering the nail and have to be removed during the surgery or removal of the screw after fracture healing. | Component design/selection | Zimmer sent an Urgent Medical Device Recall letter dated April, 29, 2015, to all affected customers.On 4/29/2015. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to review the notification and ensure affected personnel are aware of the contents. Locate all affected product and quarantine them immediately. Return the recalled product along with completed Inventory Return Certification Form. Notify Zimmer of any further distribution of the affected product. In addition, identify the surgeons and hospitals that you have identified using the provided spreadsheet. For further questions call the customer call center at 1-877-946-2761. |  | 289616 | 289,616 | Worldwide Distribution - US (nationwide) and Internationally to Canada, Brazil, Mexico, Bahamas, Nicaragua, Singapore, China, Japan, Malaysia, Taiwan, Thailand, Korea, Spain, Belgium, Switzerland, India, Germany, England, Italy, Saudi Arabia, Lebanon, Libya, United Arab Emirates, Algeria, France, United Kingdom, Canary Islands, Iraq, Jordan, Lithuania, Morocco, Macedonia, Malta, Mauritius, Serbia, Russian Federation, and Yemen. | Terminated | December 08, 2015 | 223 | 289616 | Not_Computer
Z-1696-2015 | 71250 | No | LXH | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | PRI Femoral Impactor Block for knee prosthesis surgery. |  |  | April 30, 2015 | 2015 | Zimmer, Inc. | Potential for uncured adhesive between the metal threaded insert and the impactor block to leak out of the assembled part following the autoclave sterilization process. The Master Bond epoxy serves as a secondary locking mechanism to fix the metal insert to the impactor head. Foreign material has been observed on the impactor block. This has the potential for a biological response. | Process control | On 4/30/2015, URGENT MEDICAL DEVICE RECALL LOT SPECIFIC notifications dated April 30, 2015 were sent to the affected distributors, hospital Risk Managers, and surgeons via courier; all distributors were also notified via electronic mail. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. If after reviewing this notification you have further questions or concerns please call the customer call center at 1-800-348-2759 between 8:00 am and 8:00 pm EST. |  | 1689 | 1,689 | Worldwide Distribution-US (nationwide) including PR and the states of AL, AZ, CA, CO, FL, IL, IN, KS, MA, MI, MN, MO, NC, NJ, NY, OH, OK, PA, TN, TX, UT, VA, WA, WI, and the countries of Canada, Korea, Singapore, Australia, Japan, Malaysia, India, Austria, Belgium, Canary Islands, Czech Republic, France, Germany, Iceland, Iran, Israel, Italy, Norway, Russian Federation, Saudi Arabia, South America, Spain, Sweden, Switzerland, United Kingdom, and United Arab Emirates. | Terminated | March 15, 2016 | 320 | 1689 | Not_Computer
Z-1697-2015 | 71279 | No | FRO |  |  | General &amp; Plastic Surgery | 510(k) | ActCel |  |  | May 18, 2015 | 2015 | Coreva Health Science LLC | Coreva Health Science, LLC. is recalling ActCel Hemostatic Dressing due to customer literature and web sites promoting that it can help control bleeding from open wounds and in body cavities. ActCel is indicated as a prescription device for management of topical wounds and to temporarily control external surface bleeding. | Labeling False and Misleading | On 05/18/15 the firm sent out customer notification letters to those distributors that were identified as marketing ActCel Hemostatic Dressing for use to help control bleeding from open wounds in body cavities.The firm's notification letter provided the following actions to be taken by the customer:- Review their literature and labeling (including internet websites) to determine whether there are any marketing statements for usage in body cavities; -Destroy all hardcopy literature and labeling with the statement for usage in body cavities - Correct all Internet websites eliminating his statement; and- Complete the attached Medical Device Labeling Clarification Return Response form and return it to Coreva Health Science.The firm states to direct any questions to: Ms. Casssie Inglis, Vice President of Marketing.email: cassie@coreva.netphone: 818-685-0951 ext. 119toll free: 800-808-9094 ext. 119Monday through Friday, 9am-4pm PST |  | 180900 | 180,900 | U.S. distribution to the following; TX, NY, FL, VA, WA, and AZ. No foreign distribution.NOTE: The distribution list the firm provided does not include all consignees that purchased the product; rather target the consignees (distributors) who were observed to be using an unapproved marketing use. | Terminated | August 26, 2015 | 100 | 180900 | Not_Computer
Z-1704-2015 | 71206 | No | LMD | Medical image communications device. | Radiology | Radiology | 510(K) Exempt | syngo Workflow SLR; The information system syngo Workflow SLR is a digital radiology information system (RIS) with integrated modules for patient administration, examination, reporting, statistics and system administration. The system electronically displays stores, retrieves, transfers, exchanges, and prints. |  |  | May 06, 2015 | 2015 | Siemens Medical Solutions USA, Inc | A potential exists for order transactions from interfaced HIS (Hospital Information System) systems to be lost during a restart of interface processes when using the Order Batching feature, leading to diagnosis delay. | Software design | Siemens sent an Safety Advisory Notice, dated May 6, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were given work around instructions to follow and informed that a software fix would be implemented to permanently resolve the issue.For further questions please call (610) 219-6300. |  | 131 | 131 | Nationwide Distribution | Open, Classified |  |  | 131 | Computer
Z-1701-2015 | 71202 | No | OLO | Stereotaxic instrument. | Neurology | Orthopedic | 510(k) | Tibial Alignment Guide; Orthopedic Stereotaxic Instrument; |  |  | April 19, 2015 | 2015 | Zimmer CAS | Zimmer CAS has determined that the potential exists for the spikes on the iASSIST Tibial Alignment Guide to bend or break during insertion or extraction from the tibia during use. | Device Design | Zimmer distributed notices via certified mail on May 14, 2015. Zimmer is removing affected Tibial Alignment Guide iAssist Knee System product from distribution in a two-phased process. Based upon complaint investigation, Zimmer CAS has determined that the potential exists for the spikes on the iASSIST Tibial Alignment Guide to bend or break during insertion or extraction from the tibia during use. Customerswere asked to review the notifictaion, follow the iASSIST Surgical Knee Technique provided, further distribute the notice for any distributed product, inspect devices in use before and after procedures, report any adverse events, and complete the certificate of acknowledgment and return it to CorporateQuality.PostMarket@zimmer.com.Customers with questions concerning this notice, please contact Customer Service at the following address/phone number.Customer Service / Zimmer CAS75 Queen Street, Suite 3300Montreal, Quebec, Canada H3C 2N6Email : cas-support@zimmercas.com, Telephone : 1-514-395-8883, toll free for North America 1-866-336-7846, Fax : 1-866-978-3801.For questions regarding this recall 574-372-4487. |  | 78 | 78 devices | Nationwide Distribution including AL, AZ, CA, CA, CO, FL, IN, MI, NJ, NY, OH, OR, PA, TX, VA, and WI. | Open, Classified |  |  | 78 | Not_Computer
Z-1700-2015 | 71169 | No | DYG | Flow-directed catheter. | Cardiovascular | Cardiovascular | 510(k) | Swan-Ganz CCO/Sv02/CEDV/VIP Thermodilution Catheters model number 744HF75. Flow-directed pulmonary artery catheters for monitoring cardiac output, mixed venous oxygen saturation, and end diastolic volume when used with the Vigilance Monitor. |  |  | May 05, 2015 | 2015 | Edwards Lifesciences, LLC | The product failed to meet the internal established statistical test requirement for heparin activity. | Process control | On 05/05/2015 Edwards sent their consignees a letter including the following instructions: Once you have verified your inventory, please complete the attached acknowledgement form (evenif you have no inventory) and fax it to Edwards Customer Service at 1-800-422-9329 within three days of receipt of this notification. This is to confirm that you have reviewed this notice and have taken appropriate action. Please call Edwards Customer Service at 1-800-424-3278 to arrange for return of any unused product from the identified lot and obtain information about replacement product. Please call for an RGA prior to returning any unused product and return unused product to the following address:Edwards LifesciencesReturn Goods/RGA # XXXXXX1212 Alton Pkwy.,Irvine, CA 92606-4838 |  | 978 | 978 | Worldwide distribution US nationwide, Canada and Europe (IC, BE, IT, GB, FR, DK, CZ, DE, AT, ES, IE, NO, PL) | Open, Classified |  |  | 978 | Not_Computer
Z-1703-2015 | 71273 | No | GFO | Partial thromboplastin time tests. | Hematology | Hematology | 510(k) | Dade Actin Activated Cephaloplastin Reagent; Liquid rabbit brain cephalin with plasma activator for use in the determination of the activated partial thromboplastin time (APTT) and other coagulation procedures requiring an activated partial thromboplastin reagent. |  |  | March 25, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Recovery for assayed controls are out of assigned ranges. A drift &gt;3 seconds in the normal APTT range and/or &gt;15% in the pathological range has been observed. Patient samples with recoveries close to the medical decision points could exhibit a deviation up to 4 seconds for the normal range and up to 33% for the pathological range. | Under Investigation by firm | On March 24, 2015 an Urgent Medical Device Recall letter was mailed to all affected customers to notify them about the issue, associated risks, and actions to be taken. |  | 27501 | 27501 | Worldwide Distribution. US naionwide including Puerto Rico, Barbados and Mexico. | Open, Classified |  |  | 27501 | Not_Computer
Z-1705-2015 | 71314 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Syngo.via and Syngo.x ; a software solution intended to be used for viewing, manipulation, communication, and storage of medical images. |  |  | May 15, 2015 | 2015 | Siemens Medical Solutions USA, Inc | measurements drawn on the 2nd and subsequent images of the series are not visible on printouts when the series is sent to print. | Software design | Siemens sent a Customer Advisory Notice, dated May 15, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were advised to check the software version of their syngo.via system and consider whether the issue described in the advisory notice impact their clincal coutine and whether to arrange the installation of SY027/15/P (VA30A_HF07). Customers were advised to include the advisory notice in their operator's manual. Customers were also advised to notify anyone at their organization who should be aware of this information.For questions regarding this recall call 610-219-6300. |  | 167 | 167 | Nationwide Distribution | Terminated | March 24, 2017 | 679 | 167 | Computer
Z-1708-2015 | 71096 | No | JLW | Thyroid stimulating hormone test system. | Clinical Chemistry | Clinical Chemistry | 510(k) | Dimension Vista Systems LOCI Thyroid Stimulating Hormone; the TSH method is an in vitro diagnostic test for the quantitative measurement of Thyroid Stimulating Hormone in human serum and plasma on the Dimension Vista System. |  |  | April 20, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Dimension Vista LOCI Thyroid Stimulating Hormone (TSH) Lot 14237AA may produce falsely depressed results ranging from -23% to -96% near the end of the well set due to inadequate reagent volume in a Flex well. | Nonconforming Material/Component | Siemens sent an Urgent Medical Device Recall notice, dated April 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. - Review this letter with your Medical Doctor.- Discontinue use and discard your remaining inventory of the reagent lot listed in Table 1.- Indicate your replacement product needs on the attached Field Correction Effectiveness Check form and fax to (312) 275-7795 within 30 days. Siemens will replace any unused inventory of the affected lot at no charge.We apologize for the inconvenience this situation may cause.. If you have any questions, please contact your Siemens Customer Care Center Technical Solutions at 800-441-9258, or your local Siemens technical support representative. |  | 2256 | 2256 | Worldwide Distribution - Nationwide and to the countries of : Australia, Japan, New Zealand, Canada, Austria, Belgium, Switzerland, Germany, Denmark, Spain, Finland, France, Great Britain, Italy, Netherlands, Portugal, and Slovakia. | Open, Classified |  |  | 2256 | Not_Computer
Z-1709-2015 | 71175 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | 0.9% Sodium Chloride Injection, USP, Flush Syringe, 10 mL in 12 mL, Sterile, Rx only, Code No. 2T0806 |  |  | April 27, 2015 | 2015 | MRP, LLC dba AMUSA | Incorrect expiration date on label | Process control | The firm, AMUSA, sent a "Voluntary Recall Notification" dated 4/24/2015 to their consignees via Fed Ex on 04/27/2015. The letter described the product, problem and actions to be taken. The consignees were instructed to return any product remaining in the distribution channel to AMUSA; check with your product representative to determine if you have any of the listed products; if any product is located, call (615) 833-2633 from 8:30 am to 5:00 pm (Eastern Time) to receive instructions as to how to return the product; follow the enclosed Steps for Voluntary Recall in their entirety, and complete and return the enclosed Product Tracking and Verification Form.Please contact your AMUSA customer service representative if you have any questions regarding this correction, any of our products, or would like assistance with the corrections. Customers can also contact the Director, Quality Assurance at 513-325-1940 or email qadirector@amusa.us. |  | 200000 | 200,000 units | US Distribution in states of:OH, PA, AL, IL, TN, and CT. | Terminated | May 02, 2016 | 371 | 200000 | Not_Computer
Z-1710-2015 | 71098 | No | LZS |  |  | Ophthalmic | PMA | EC-5000 Excimer Laser Corneal Surgery System (EC-5000, EC 5000 CXIII, EC-5000 Quest. |  |  | March 24, 2015 | 2015 | Nidek Inc | Multi-stage treatment option for the EC-5000 Operator's Manual was not approved by the FDA. | No Marketing Application | Nidek sent an Device Correction letter dated April 30th, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Nidek will contact customers to schedule an appointment to replace the operators manual and implementing a software lockout device on the device so that the multi stage feature is no longer available.If a Nidek rep does not contact you before April 30, 2015, please call 800-223-9044 to schedule an appointment, All non responding customers.will be contacted three time via Fed Ex letter |  | 114 | 114 units | Us distribution only. | Terminated | April 05, 2016 | 378 | 114 | Computer
Z-1707-2015 | 71176 | No | CDP | Total triiodothyronine test system. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS Immunodiagnostic Products Signal Reagent, Product Code 1072693, Unique Device Identifier No. 10758750008360 VITROS Immunodiagnostics Products Signal Reagent (SR) is supplied to users as a sales unit box containing two Signal Reagent Packs. Each Signal Reagent Pack is labeled with a barcode label. The label should include the name of the reagent, a barcode, lot information and expiry dating. |  |  | May 06, 2015 | 2015 | ORTHO-CLINICAL DIAGNOSTICS | The Immunodiagnostic Products Signal Reagent, Lot 3082 as the reagent packs may contain blank labels. | Error in labeling | On 5/06/2015, Customer Letter and Verification Form (Ref. CL15-095) were sent notifying the users of the issue and advised them of the required actions. The notification was sent via FedEx overnight courier and/ or ORTHO PLUS e-Communications and/ or US Postal Service Priority Mail (only to those consignees who provided a PO Box as an address) to all customers who received the affected product and advised them of the issue and the required actions. On 5/06/2015, Distributor Letter (Ref. DL15-095) was sent via US Postal Service Priority Mail and/or FedEx overnight courier to US Distributor consignees and instructed to notify their consignees of the issue and actions. For questions,questions, please contact Customer Technical Services at 1-800-421-3311 at any time. |  | 898 | Domestic: 898 units | Nationwide Distribution. | Open, Classified |  |  | 898 | Not_Computer
Z-1712-2015 | 71318 | No | MUJ | Medical charged-particle radiation therapy system. | Radiology | Radiology | 510(k) | Oncentra Brachy radiation therapy planning software designed to analyze and plan radiation treatments in three dimensions for the purpose of treating patients with cancer. |  |  | May 21, 2015 | 2015 | Elekta, Inc. | Incorrect dose calculation for Regions of Interest (ROIs) defined on a secondary image series. | Other | Important User Notice, 806-03-BTP-001, was distributed to customers on 5/21/2015, A software patch is in development and is expected to be released end of November, 2015. Elekta service will work with customers to implement the patch on affected devices within 6 months of the patch release. |  | 351 | 351 | AL, AZ, AR, CA, CO, CT, DC, FL, GA, HI, IL, IN, KS, KY, LA, MD, MA, MI, MN, MS, MO, MT, NE, NC, NH, NJ, NY, NC, ND, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, WV, WI, Argentina, Australia, Bangladesh, Belarus, Botswana, Brazil, Canada, Chile, China, Colombia, Ecuador, Egypt, Guatemala, Hong Kong, India, Indonesia, Israel, Japan, Jordan, Kosovo, Libya, Malaysia, Mexico, Morocco, Namibia, Nepal, New Zealand, Oman, Pakistan, Peru, Philippines, Qatar, Russia, Saudi Arabia, Serbia, Singapore, South Africa, South Korea, Sri Lanka, Suriname, Taiwan, Thailand, Tobago, Ukraine, Venezuela, Vietnam, Zambia | Open, Classified |  |  | 351 | Computer
Z-1711-2015 | 71277 | No | HTW | Orthopedic manual surgical instrument. | Orthopedic | Orthopedic | 510(K) Exempt | 3.2 mm Cannulated Drill Bit/ QC 170 mm |  |  | May 13, 2015 | 2015 | Synthes, Inc. | affected parts and lots of the Cannulated Drill Bits have the potential to break during use | Component design/selection | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated May 13, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed if they have any of the identified devices, please take the following steps:" Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number." Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affected product has been located. Also, please indicate the number of devices found and note the Return Authorization Number. Please include your name, title, address, telephone number and signature in the spaces provided." Return the Verification Section (page 3 of this letter) with the product to:" Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132." Send a copy of the completed Verification Section to DePuy Synthes by:" Fax: (888) 943-4893 or" Scan/email: Synthes6230@stericycle.comIf you DO NOT have the identified product, please take the following steps:" Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicating that no affected product has been located. Please include your name, title, address, telephone number and signature in the spaces provided. This return documentation acknowledges your receipt of medical device removal information." Return the documents to DePuy Synthes by:" Fax: (888) 943-4893 or" Scan/email: Synthes6230@stericycle.com" Note: If the Verification Section is answered on behalf of more th An urgent medical device recall notice, dated May 13, 2015, was sent to consignees and sales consultants to in form them of the issue and associated risks. All in receipt of the device were asked to discontinue use and return affected devices along with the response form.Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes Sales |  | 1495 | 1495 US (37 OUS) | Worldwide Distribution - US (nationwide) and Canada. | Open, Classified |  |  | 1495 | Not_Computer
Z-1715-2015 | 71209 | No | FEB | Endoscope and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | EVOTECH Endoscope Cleaner and Reprocessing System 220-230V, P/N 50004-002. |  |  | May 11, 2015 | 2015 | Advanced Sterilization Products | Advanced Sterilization Products (ASP) is recalling the EVOTECH Endoscope Cleaner and Reprocessor (ECR) because it may not detect an improper endoscope connection on two channels, and subsequently pass the cycle without cancellation if there is a disconnected scope. | Nonconforming Material/Component | A customer notification letter dated 5/7/15 will be sent to all customers on 5/13/15 to inform them that ASP is recalling the EVOTECH Endoscope Cleaner and Reprocessor (ECR) because it may not detect an improper endoscope connection on two channels and subsequently pass the cycle without cancellation if there is a disconnected scope. The letter informs the customers of the problems identified and the actions to be taken. Customers are instructed to complete and return the attached Business Reply Card (double-sided) via mail or by fax to Stericycle: 888-879-8144. Customers with questions regarding the business reply card are instructed to contact Stericycle at 888-345-4671 and reference event #7781. |  | 44 | 44 units | Worldwide Distribution-US (nationwide) and the countries of Australia, Singapore, Canada, and Thailand. | Open, Classified |  |  | 869 | Not_Computer
Z-1718-2015 | 71058 | No | ITH | Power traction equipment. | Physical Medicine | Physical Medicine | 510(k) | Extentrac Elite User Manual DCN: 14-001 (Powered decompression table, for traction) |  |  | May 13, 2015 | 2015 | Advanced Back Technologies, Inc. | Item is marketed for use in patients weighing up to 350 lbs, but the device was only approved for use in patients up to 300 lbs | Error in labeling | Advanced Back Technologies, Inc. sent an Urgent Medical Device Brocure &amp; User Manual Recall letter dated May 13, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed the discard the DCN's identified as DCN 13-001, DCN 09-001, DCN 14-001 and replace with the newly issued DCN 15-001, DCN 15-002. The new DCN's would be provided to customers in pdf format via email for issuance. Customers were asked to contact their customers who purchased the affected product and inform them that they should not treat patients with a maximum weight exceeding 300 lbs until the firm is able to make an amendment to their 510(k). Customers with questions should call 631-231-0076. |  | 18 | 18 units | Worldwide Distribution - US including CA, FL, NJ, TX, WA, Internationally: Korea, Russia, Kuwait, Canada, Greece, Lebanon, Turkey, Egypt, Saudi Arabia, China, Armenia, Azerbaijan, Australia. | Open, Classified |  |  | 54 | Not_Computer
Z-1713-2015 | 71267 | No | GJT | Factor deficiency test. | Hematology | Hematology | 510(k) | Factor deficient plasma for the Siemens BCSÂż / BCSÂż XP System. |  |  | May 11, 2015 | 2015 | Siemens Healthcare Diagnostics, Inc. | Discrepancies for Siemens BCS / BCS XP with dilution studies; erroneously high Factor IX results at the lower end of the measuring range. | Under Investigation by firm | An Urgent Medical Device Recall notice, dated May 2015, was sent to affected customers to inform them about the issue and associated risks. Customers were asked to discontinue use, distribute this information to all appropriate personnel, and to complete and fax back to Siemens the Effectiveness Check form. |  | 3575 | 3,575 | Worldwide Distribution -- Austria, Bulgaria, Belgium, Switzerland, Cyprus, Czech Republic, Germany, Denmark, Italy, Latvia, Netherlands, Poland, Portugal, Romania, Sweden, Slovenia, Slovakia, Turkey, Croatia, Bahrain, Belarus, Bosnia-Herzegovina, Guadeloupe, Kuwait, Lebanon, Libya, Macedonia, Montenegro, Qatar, Russia, San Marino, Saudi Arabia, Serbia, South Africa, United Arab Emirates, Argentina, Brazil, Chile, El Salvador, Peru, China, Mexico, Canada and United States. | Open, Classified |  |  | 3575 | Not_Computer
Z-1719-2015 | 71295 | No | IWB | Radionuclide radiation therapy system. | Radiology | Radiology | 510(k) | Laksell GammaPlan, a computer based dose planning system specifically designed for use with Leksell Gamma Knife, radiation therapy treatment. |  |  | June 01, 2015 | 2015 | Elekta, Inc. | Memory can become corrupted when creating a fused study via drag and drop in Leksell GammaPlan 10.2. | Software in the Use Environment | The firm, Elekta, sent an "URGENT IMPORTANT FIELD SAFETY NOTICE 100-01-102-015" to all customers on 6/1/2015. The notice described the product, problem and actions to be taken. The customers were instructed to not use drag-and-drop to fuse images as a solution until a permanent fix can be developed and implemented on all affected devices. A solution is in development that will resolve the issue and is expected to be released September 2015. Elekta Service will then work with customers to implement on affected devices - estimated 6 month completion. Customers were also instructed to complete and return the IMPORTANT FIELD SAFETY NOTICE ACKNOWLEDGEMENT form to your local Elekta Office or Representative as soon as possible and within 30 days at the latest. If you have any queries about this Notice, please contact your local Elekta representative or the Global Post market Surveillance Manager at 770-300-9725 or email: Elekta_Global_PMS@elekta.com. |  | 10 | 10 | Worldwide Distribution: US (nationwide) including states: IL, OH, PA, TX, and WI; and countries of: Germany, Japan and United Kingdom. | Terminated | January 12, 2016 | 225 | 10 | Not_Computer
Z-1716-2015 | 71093 | Yes | KOC | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | CRIT-LINE Blood Chamber- an accessory to the Crit line III- Monitors in real time blood parameters |  |  | April 27, 2015 | 2015 | Fresenius Medical Care Holdings, Inc. | Blood chamber connection leaks during the use of the CRIT-LINEÂż Blood Chamber | Under Investigation by firm | Fresenius LLC North America notified accounts by letter sent certified mail on May 22, 2015, expanding the recall to All lots. Previously recall letters were on April 30, 2015, 2nd notice on May 14, 2015. Users are requested to examine inventory, discontinue use and return product. Contact FMCRTG, LLC Customer Service Team 1-800-323-5188 for instructions. A reply form is included as part of the customer communication with instructions to fill out and return via fax or scanned/emailed copy.Confirmation of delivery will be obtained through the use of registered mail. |  | 22640288 | 22,640,288 | Worldwide Distribution - Nationwide Distribution including Puerto Rico., and the countries of : Ireland, Spain, Slovenia, Great Britain, Netherlands, Norway, Mexico, Egypt, Czech Republic and Spain. | Open, Classified |  |  | 22640288 | Not_Computer
Z-1720-2015 | 71112 | No | CAF | Nebulizer. | Anesthesiology | Anesthesiology | 510(k) | Roscoe Mini Neb Compressor User Manual for the NEB-ROS |  |  | April 20, 2015 | 2015 | Compass Health Brands | Roscoe Medical has recently identified the need to update its user manual with additional warning statements regarding proper handling and storage of the nebulizer. These warnings are to better inform end users of proper care when using and storing a Roscoe Medical Nebulizer to reduce user damage (wear &amp; tear) to the power cord. | Use error | Compass Health sent an Urgent Recall Field Correction letters dated April 20, 2015 to their customers. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to: Âż Ensure all affected personnel are fully informed of this notice.Âż Forward this notice to your Regulatory Compliance Manager, Purchasing Manager, Customer Service Manager and Field Technicians.Âż Advise your customers regarding the proper care and storage of the nebulizers.Âż Your customers should discard their NEB-ROS User Manual and immediately download the latest versions HERE.Âż If they are unable to download the most recent version, hard copies can be provided. Please contact Roscoe Medical Customer Support at (800) 376-7263 Monday - Friday 8:00 am EST - 5:00 pm EST.Âż Complete and return the attached Recall Field Correction Response Form to Roscoe Medical within fifteen (15) days of receipt of this field correction notification confirming your acknowledgement. Send completed form to Roscoe Medical via fax number (440) 572-4261 or via email at corporate@roscoemedical.com. |  | 219353 | 219,353 Manuals | Worldwide Distribution US nationwide in the states of: AL, AR, AZ, CA, CO, CT, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, PR, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV &amp; WY and countries of:: Australia, Canada, Chile, Colombia, Costa Rica, Dominican Republic, Ecuador, France, Hong Kong, India, Norway, Spain, Trinidad and Tobago &amp; United Kingdom. | Completed |  |  | 219353 | Not_Computer
Z-1730-2015 | 71142 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Low Volume Extension Set, 72' w/Clamp FLL/MLL APV=1.3ml, STERILE, AND FLUID PATH NONPYROGENIC UNLESS PACKAGE IS DAMAGED OR OPENED. LATEX FREE, SINGLE USE Tyvek pouches, 10 per box, Smiths Medical ASD, Inc. |  |  | May 04, 2015 | 2015 | Smiths Medical Asd Inc | Smiths Medical has become aware of an issue with the Low Volume Extension Set, (72'' Extension Set) where under-sized tubing was used in the production of 1 lot of product. | Process control | Smiths Medical sent Urgent Medical Device Recall Notice letters, dated 4 May 2015, requiring its customers to return all affected unused Extension Sets. Please take the following actions: 1. Customers are to inspect their inventory and quarantine any affected Extension Sets.2. Complete and return the attached Confirmation Form by Fax or by e-mail3. Upon receipt of the completed form, a customer service representative will contact customers to arrange for exchange of their unused affected Extension Sets for credit or replacement.The notice shall be passed on to all personnel who need to be aware, including points of use. If these affected Extension Sets have been further distributed to other persons or facilities, please promptly forward the recipients a copy of this Urgent Medical Device Recall Notice. |  | 6000 | 6,000 units | Distributed to the states of KY, MN, NJ, UT, &amp; WI. | Terminated | January 20, 2016 | 261 | 6000 | Not_Computer
Z-1724-2015 | 71174 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | enGen Laboratory Automation System (enGen System), Product Code ENGEN, configured with all of the following system components: Recapper Caps (Product Codes 952238 or 6844286 only), Recapper Module (230V) or Recapper Module (110V) (Product Codes 952036-EG or 952136-EG), Rack Entry and Exit Module and/or Rack Exit Module (Product Codes 952020-EG/6844086 or 952019-EG/6844103), Storage Rack Slim 100 positions or Storage Rack 180 positions (Product Codes 952075 or 952073). |  |  | May 06, 2015 | 2015 | Ortho-Clinical Diagnostics | Recapper Caps may fall off of, or be knocked from sample tubes placed in the storage racks of entry/exit modules. The gripper may grab a tube at the wrong height and drop the tube in the Rack Entry/Exit or Rack Exit module. These situations may result in splashing of bio-hazardous fluid. Sample cross-contamination may occur with an uncapped tube in the rack. | Component change control | On 5/06/2015, a customer letter (Ref. CL2015-092) with recommendation to prevent this issue was sent to all customers in the US with enGen Systems via FedEx or US Priority Mail (for Post Office Box addresses) or by email to subscribers of ORTHO PLUS e-Communications. Foreign affiliate consignees were notified by e-mail on 5/06/2015 informing them of the issue. For questions, please call Customer Technical Services at 1-800-421-3311 at any time. |  | 45 | 45 units worldwide (Domestic: 19 units, Foreign: 26) | Worldwide distribution. US Nationwide, Australia, Bermuda, Brazil, Canada, Chile, China, France, India, Italy, Norway, Portugal, Spain, Sweden, and the United Kingdom. | Open, Classified |  |  | 45 | Not_Computer
Z-1706-2015 | 71054 | Yes | BSZ | Gas machine for anesthesia or analgesia. | Anesthesiology | Anesthesiology | 510(k) | FLOW-i Anesthesia System |  |  | April 02, 2015 | 2015 | Maquet Cardiovascular Us Sales, Llc | In some Flow-i Anesthesia Systems, patient cassette can get dislodged which could result in gas leakage within the system internal circuit potentially leading to a stop in ventilation. | Under Investigation by firm | The firm, Maquet Inc., sent an "Urgent Medical Device Field Action Field Correction for FLOW-i Anesthesia Systems" letter dated 4/2/2015 via FedEx. The letter describes the product, problem and actions to be taken. The customers were instructed to read the letter and inform staff, who may use the FLOW-i Anesthesia System of the patient cassette locking device issue. Customers are to obtain the serial number of their system; verify the serial number obtained against the serial number listed in the box in the Appendix. A Maquet Service Representative will be contacting you to arrange a convenient time to replace your FLOW-i Anesthesia System patient cassette locking device.If you have any questions regarding this field action, please contact your Maquet Service Representative or Maquet Technical Support at 1-888-627-8383 (Press option 3, followed by option 1 and then option 1 again), Monday through Friday, between the hours of 8:00 a.m. and 5:00 p.m. EST. |  | 1641 | 1641 [80 units (US) 1561 units (OUS)] | Class I Recall - Worldwide distribution -- US, including the states of AL, AZ, DC, FL, MI, MO, NJ, NY, OR, PA, TX, and WV; and, countries of Argentina, Australia, Austria, Belgium, Brazil, Bulgaria, Chile, China, Colombia, Croatia, Czech Republic, Ecuador, Finland, France, Germany, Hong Kong, Hungary, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kuwait, Latvia, Lebanon, Libya, Mexico, Mongolia, Mozambique, Netherlands, Norway Oman, Panama, Peru, Poland, Portugal, Qatar, Romania, Russia, Saudi Arabia, Serbia, Singapore, Slovakia, South Africa, South Korea, Spain Sweden, Switzerland, Syria, Taiwan, Thailand, Tunisia, Turkey, United Arab Emirates, United Kingdom, Unites States, and Venezuela. | Terminated | July 06, 2016 | 461 | 1641 | Not_Computer
Z-1698-2015 | 71164 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD). |  |  | April 29, 2015 | 2015 | HeartWare Inc | HeartWare has received complaints relating to damage or bent connection pins within the power supply ports of the VAD controller. | Device Design | HeartWare sent an Urgent Medical Device Correction letter to customers on May 25, 2015. The letter identified the affected product, problem and actions to be taken. The letter instructed customers to:1. Review the enclosed notice and Patient Communication .2. Forward the notice to those individuals within their organization who need to be aware of its contents.3. Identify the patients currently supported by the HVAD System.4. Distribute the "Patient Communication" to their patients directly via FedEx or some other reliable means of communication.5. Continue to reinforce the messages set forth in the notice with their patients during their regularly scheduled appointments.6. Complete, sign and return the Acknowledgement and Completion Form to HeartWare within 30 days of receipt of this letter.For questions contact your local HeartWare representative or the 24-Hour Clinical Support personnel at 1-888-494-6365. |  | 3747 | 3,747 currently implanted worldwide | Worldwide Distribution - US Nationwide in the states of AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, VA and countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Not_Computer
Z-1728-2015 | 70918 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 5,1 FS Chemistry System, Software Version 2.8 &amp; Below |  |  | April 06, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly regarding urine samples that require acidified pretreatment. The software has allowed urine samples for assays that require acidified pretreatment to be metered from the same sample container as assays that required no pretreatment when the sample was programmed. | Software design | Ortho Clinical Diagnostics distributed correction notification letters dated April 6, 2015, and response forms to their customers via courier service. All urine patient results that require pretreatment was advised to refer to the notice of correction for instructions. Correction notices should also be posted beside the device user documentation for easy reference. Customers were advised to complete and return the Confirmation of Receipt before April 20, 2015. Customers with questions may contact the Customer Technical Service at 1- 800-421-3311 anytime. |  | 2387 | 2,387 units Total (987 domestically &amp; 1400 internationally) | Worldwide Distribution - US Nationwide, Brazil, Chile, Mexico, Australia, Hong Kong, India, Taiwan, Belgium, Denmark, France, Ireland, Italy, Portugal, Romania, Spain, Sweden and the UK. | Open, Classified |  |  | 2387 | Computer
Z-1699-2015 | 71272 | No | KWA | Hip joint metal/metal semi-constrained, with an uncemented acetabular component, prosthesis. | Orthopedic | Orthopedic | PMA | M/L Taper with KinectivÂż Technology. |  |  | May 18, 2015 | 2015 | Zimmer, Inc. | Zimmer is initiating a voluntary recall of 64 lots (752 implants total) of M/L Taper with KinectivÂż femoral stems and modular necks due higher than allowed cytotoxicity levels found with the product. Reasonable probability of adverse biological response and subsequent revision | Environmental control | Zimmer, Inc. sent an URGENT MEDICAL DEVICE RECALL letter dated May 18, 2015, to all affected consignees. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to review the notification and ensure affected personnel are aware of the contents. Locate all affected product identified and quarantine them immediately. Carry out a physical count of all affected product in their inventory and complete the Inventory Return Certification Form. Email a completed copy to corporatequality.postmarket@zimmer.com. Return the recalled product along with the completed Inventory return Certification Form. Notify Zimmer of any hospitals that they have further distributed the affected product to. In addition, identify the surgeons that have implanted the product. Supply the information for any hospitals that they have identified, as well as the affected surgeons using the provided spreadsheet template. Customers with questions or concerns should call the customer call center at 1-877-946-2761. |  | 752 | 752 | Worldwide Distribution - US Nationwide in the states of AK, AL, AZ, CA, FL, GA, IL, IN, KS, MA, MI, MN, MO, NC, NY, OH, OK, PA, TX, UT, VA, WA, WI and countries of Canada, Australia, Japan, Taiwan, France, Germany, Spain, and Italy. | Terminated | August 02, 2016 | 442 | 752 | Not_Computer
Z-1638-2015 | 71298 | No | OXO | Image-intensified fluoroscopic x-ray system. | Radiology | Radiology | 510(k) | GE OEC Brivo 865 plus, Mobile X-ray system laser aimer assemblies (part number 520002DW) |  |  | May 22, 2014 | 2014 | GE OEC Medical Systems, Inc | Laser aimer assemblies were shipped without certification of the overall assembly or submitting a laser product report and therefore do not comply with the US regulations for laser products. | Radiation Control for Health and Safety Act | GE Healthcare Surgery Planned Action(s) to repair the defect or to bring the product into compliance: The Center for Devices and Radiological Health (CDRH) understands that the CAP involving 9 units consisted of the following:1. Replacing the tube side cover with the laser aimer assembly with a tubeside cover without a laser aimer assembly.This correction brought the systems into compliance with applicable US regulations. For further questions please call (801) 536-4615. |  | 9 | 9 | US Distribution to the states of: MD, NC, MT, NJ, UT and NY | Terminated | June 24, 2015 | 398 | 9 | Not_Computer
Z-1725-2015 | 71244 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | Alkaline Wash, List number 9D31-20, a cleaning solution used by the Abbott Architect analyzers. It is supplied in 500 mL bottles and contains NaOCl, NaOH, and KOH. It is shipped and stored at room temperature and is shipped ready to use. |  |  | May 11, 2015 | 2015 | Abbott Laboratories | The product is leaking from the cap and crystallization was noted as present around the cap. | Under Investigation by firm | Abbott Laboratories sent a Product Recall letter dated May 11, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to wear the appropriate Personal Protective Equipment (PPE) to determine if they have the affected product in their inventory. If customers observe any signs of leakage or loose caps discontinue use and disppose of any remaining inventory of those kits in accorance with their institutions's policies and procedures. If customers do not observe any signs of leakage or loose caps they may continue to use the product following the precautions per the ARCHITECT System Operations Manual and the Safety Data Sheet. Customers were also asked to complete and return the enclosed Customers Reply form. If affected product was forwarded to other laboratories, customers need to inform them of the Product Recalls and provide them a copy of tthe product recall letter. Customers should also maintain a copy of the recall letter. Customers with questions were instructed to contact Customer Service at 1-877-4ABBOTT. For questions regarding this recall call 224-668-1926. |  | 8868 | 8868 units | Worldwide Distribution - US (nationwide), Mexico, Colombia, Brazil, Germany, Hong Kong, Belgium, Czech Republic, Finland, Hungary, Ireland, China, New Zealand, Uruguay, Puerto Rico, St. Lucia, Austria, Croatia, Denmark, France, India, Isreal, Bermuda, Curacao, Cayman Islands, Jamaica, Canada, Bangladesh, Cyprus, Estonia, Greece, Indonesia, Italy, Kazakhstan, Latvia, Macedonia, Norway, Portugal, Russia, Slovakia, Spain, Taiwan, UAE, Kenya, Lithuania, Malaysia, Philippines, Qatar, Saudi Arabia, Slovenia, Sweden, Tunisia, United Kingdom, Kuwait, Luxembourg, Netherlands, Poland, Romania, Singapore, South Africa, Switzerland, Turkey, and Vietnam. | Completed |  |  | 8868 | Not_Computer
Z-1723-2015 | 71344 | No | HSB | Intramedullary fixation rod. | Orthopedic | Orthopedic | 510(k) | Vilex, Inc., FUZE 10.0 mm x 250 mm, Product # FZ100-250T-25. |  |  | April 22, 2015 | 2015 | Vilex In Tennessee Inc | Incidence of breakage is higher than expected. | Device Design | The initial telephone call (4/9/2015) informed the distributor that the product was on hold (not to be used or transferred). The second telephone call (4/29/2015) and letter informed the customer that the product was being removed from market and that the customer was to return the product to the corporate office.The letter was disseminated on 05/05/2015 to distributors via certified mail. The letter contains the instructions for them to follow, which as summarized are: inventory all products on hand, sign off on inventory sheet, send products back to the corporate office with the identified return number on the shipping container, inform the corporate office that the process has been completed and provide the corporate office with the list of the products returned.A second letter was disseminated to sales representatives (distributors), surgeons, and hospitals on 05/28/2015. |  | 4 | 4 units | Distributed in the states of MS, MD, PA, and TX. | Open, Classified |  |  | 24 | Not_Computer
Z-1727-2015 | 71308 | No | LKF | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | Ally Uterine Positioning System (Ally UPS) used in the mounting and positioning and holding of uterine manipulators |  |  | May 18, 2015 | 2015 | CooperSurgical, Inc. | Design of device may expose user to injury to fingers or body parts | Device Design | CooperSurgical initiated the recall on May 18, 2015 via Fedex with confirmed delivery receipt. The letter identified the affected product and asked users to discontinue use and complete the Acknowledgement and Receipt Form to schedule a replacement.Questions contact 203.601.5200. |  | 5 | 5 units | Distributed in the states of CA, IL, NC, and NY. | Terminated | August 07, 2015 | 81 | 5 | Not_Computer
Z-1702-2015 | 71156 | No | LRO | Surgical drape and drape accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | Enforcement Discretion | The Alcon Custom Pak. The Alcon Customer Pak is an assemblage of single-use medical devices and accessories designed by medical professionals for specific ophthalmic surgical procedures. Some Paks contain the Devon Light Glove. |  |  | April 28, 2015 | 2015 | Alcon Research, Ltd. | The supplier of the Devon Light Glove, a component of the Alcon Custom PAK, recalled the gloves due to the potential for splits and holes on the glove. Should the user be unaware that the Light Glove is torn/split, surgical site infections are possible. | Nonconforming Material/Component | Alcon sent a Recall letter dated April 28, 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. The letter included a return form and instructions for credit. Customers were instructed to review inventory for any affected units, segregat the product to ensure it is not used, and return the attached Response Form via fax or email to Alcon.For questions or need assistance with obtaining replacement products contact Alson Customer service at 800-862-5266. |  | 219680 | 219,680 units | Worldwide Distribution - US Nationwide and the countries of Australia and Canada. | Terminated | March 24, 2016 | 331 | 219680 | Not_Computer
Z-1742-2015 | 70917 | No | JJE | Discrete photometric chemistry analyzer for clinical use. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | VITROS 4600 Chemistry System, Software Version 3.2 &amp; Below. |  |  | April 06, 2015 | 2015 | Ortho-Clinical Diagnostics | Software anomaly allows testing of multiple assays using a single urine specimen, regardless of the pretreatment requirements for each selected assay, potentially leading to biased patient results. | Software design | On April 6, 2015, Ortho Clinical Diagnostics distributed correction notification letters and response forms to their customers via courier service. All urine patient results that require pretreatment was advised to refer to the notice of correction for instructions. Correction notices should also be posted beside the device user documentation for easy reference. Customers were advised to complete and return the Confirmation of Receipt before April 20, 2015. Customers with questions may contact the Customer Technical Service at 1-800-421-3311 anytime..***CAPA***Since the Assay Data Disk (ADD) provides data tables that determine which assays can be tested from the same sample container, an updated ADD is being created to address the anomaly on the VITROS 4600 and 5600. The anomaly will be addressed with an updated ADD is uploaded and the revised ADD mitigates the issue totally. Anticipated date of availability TBD. |  | 255 | 255 units total (131 domestically &amp; 255 internationally) | Worldwide Distribution-US (nationwide), Brazil, Chile, Mexico, Australia, Hong Kong, India, Taiwan, Belgium, Denmark, France, Ireland, Italy, Portugal, Romania, Spain, Sweden and the UK. | Open, Classified |  |  | 255 | Computer
Z-1732-2015 | 71246 | No | DXN | Noninvasive blood pressure measurement system. | Cardiovascular | Cardiovascular | 510(k) | ClearSight, Model No. EVHRS Heart Reference Sensor. The EV1000 Clinical Platform Nl and ClearSight finger cuffs noninvasively measures blood pressure and associated hemodynamic parameters. |  |  | May 18, 2015 | 2015 | Edwards Lifesciences, LLC | Edward Lifesciences is recalling the ClearSight System because the finger component of the Heart Reference Sensor may be damaged. | Device Design | Edwards Lifesciences sent an URGENT Product Notification letter dated May 12, 2015, to all affected customers. The letter informed customers that Edwards Lifesciences has received reports that a component of the ClearSight System can be damaged if not handled properly. The letter informs customers of the problems identified and the actions to be taken. Customers with questions are instructed to call Edwards Technical Support. For questions regarding this recall call 949-250-2500. |  | 1309 | 1,309 units total (696 units in US) | Worldwide Distribution - US (nationwide) and Internationally to Australia, Estonia, Italy, Romania, Sweden, Austria, Finland, Kuwait, Saudi Arabia, Switzerland, Bahrain, France, Netherlands, Slovenia, Turkey, Belgium, Great Britain, Oman, Slovakia, South Africa, Bulgaria, Greece, Poland, South Africa, United Arab Emirates, Czech Republic, Hungary, Portugal, Spain, Denmark, Ireland, and Qatar. | Terminated | November 06, 2015 | 172 | 1309 | Not_Computer
Z-1734-2015 | 71235 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | AccuChek Spirit insulin infusion pump. |  |  | April 22, 2015 | 2015 | Roche Diabetes Care, Inc. | Roche has become aware of an increase of mechanical errors experienced by customers using the Accu-Chek Spirit Combo insulin infusion pump and Accu-Chek Spirit insulin infusion pump when the insulin pumps display E6 and El0 error messages.If the user does not act upon the E6 and E10 error messages appropriately, insulin delivery will be stopped and if unnoticed may lead to severe hyperglycemia u | Device Design | Roche sent an URGENT MEDICAL DEVICE CORRECTION notification dated April 22, 2015, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer. The recall notification included a description of the reason for the recall, customer responsibilities, training chart, and instructions for responding to the formal recall notification. Please contact the ACCU-CHEK Customer Care Service Center, 24 hours a day, seven days a week at 1-800-688-4578, if you have questions about the information contained in this notification. |  | 35795 | 35,795 | Nationwide Distribution | Open, Classified |  |  | 35795 | Not_Computer
Z-1739-2015 | 71330 | No | KWZ | Hip joint metal/polymer constrained cemented or uncemented prosthesis. | Orthopedic | Orthopedic | 510(k) | R3(TM) 28 MM I.D., 62 MM O.D. CONSTRAINED LINER, REF 71339162, QTY: (1), STERILE EO. |  |  | May 06, 2015 | 2015 | Smith &amp; Nephew, Inc. | The R3 Constrained outer liner locking groove diameter is 0.030 inches oversized and may not allow the locking ring to engage into the R3 shell. | Process change control | Smith &amp; Nephew sent an Urgent-Product Recall Notification letter dated May 6, 2015. to Sales offices by E-mail and Hospitals via Fed- Ex. The letter identified the affected product, problem and actions to be taken. Customers were asked to quarantine the product and return it. Distributors were asked to notified their customers of the recall and ask them to return the product. |  | 32 | 32 units | Worldwide Distribution - US Nationwide in the states of CA, FL, IL, IN, MI, NE, NY, TN, TX, VA, WI and the countries of Australia, Switzerland, Chile, Canada, and Spain | Open, Classified |  |  | 101 | Not_Computer
Z-1741-2015 | 71255 | No | KTT | Single/multiple component metallic bone fixation appliances and accessories. | Orthopedic | Orthopedic | 510(k) | Unilateral external fixation ankle clamp |  |  | May 01, 2015 | 2015 | Biomet, Inc. | Ankle clamp was assembled incorrectly; the offset of the pin to the pivot point is reversed. This prevents complete range of motion and if pushed too far may loosen the pin from the bone. A revision surgery may be necessary to adequately reconnect the fixator to the bone. | Process control | Biomet sent an URGENT MEDICAL DEVICE RECALL NOTICE letter dated May 1, 2015 to the affected distributors and implanting surgeons. The recall letter included a description of the reason for the recall, affected product, consignee responsibilities, and distributor instructions for responding to the formal recall notification. Consignees were instructed to return the affected product and continued monitoring of patients. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 19 | 19 | US Nationwide Distribution in the states of CA, TX, FL, NC, MI, GA. | Terminated | December 08, 2015 | 221 | 19 | Not_Computer
Z-1740-2015 | 71232 | No | KAM | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | RELIEVA FLEX Sinus Guide Catheter F-70. |  |  | April 28, 2015 | 2015 | Acclarent, Inc. | Acclarent determined that one single RELIEVA FLEX Sinus Guide Catheter was distributed past it's expiry date. | Nonconforming Material/Component | Acclarent sent an Urgent Medical Device Recall letter dated May 2015 to affected customers. The letter identified the affected product, problem and actions to be taken. Customers were instructed that if the affected catheter is still in inventory, it should be isolated and returned to Acclarent per instructions.Questions should be addressed to Gabriel Alfageme, Field Action Coordinator at 1-650-687-5398. |  | 986 | 986 | US in the state of Texas | Terminated | June 10, 2015 | 43 | 986 | Not_Computer
Z-1746-2015 | 69954 | [6 empty] | BBL Port A Cul tubes, catalog number 221606. Patient sample collection tubes. BBL Port A Cul tubes, jars and vials contain a reduced transport medium and are intended to maintain the viability of anaerobic, facultative and aerobic microorganisms during transport from the patient to the laboratory. Sterile packages are for collection of specimens in clean areas; e.g., surgical suites. |  |  | November 21, 2014 | 2014 | Becton Dickinson &amp; Co. | The products may exhibit indications of excessive oxidation. Excessive oxidation may affect the viability of anaerobic microorganisms. | Process control | BD Diagnostic Systems sent an Urgent Product Recall letter dated December 2014 to all affected customers. Recall notification was initiated on November 21, 2014, the notification was sent by fax and letters by UPS. Customer notification for the expansion of the recall was initiated on December 18, 2014, the notification was sent by fax and letters by UPS. For questions regarding this recall call 800-638-8663. |  | 145860 | 145,860 tubes | US Nationwide distribution in AK, AL, AR, AZ, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MN, MO, MS, MT, NC, ND, NE, NJ, NV, NY, OH, OR, PA, SC, SD, TN, TX, UT, VA, VT, WA, WI, WV, and WY. | Terminated | January 07, 2016 | 412 | 238570 | Not_Computer
Z-1748-2015 | 71198 | No | JSG | Multipurpose culture medium. | Microbiology | Microbiology | 510(K) Exempt | Remel Blood Agar (TSA w/sheep blood) plate REF R01202 packaged 10 plates per pack, 10/10-packs per case (100 plates). |  |  | May 04, 2015 | 2015 | Remel Inc | Product may be contaminated with Listeria monocytogenes. | Material/Component Contamination | The recalling firm sent the recall notification letter (dated May 1, 2015) to all customers on 5/4/2015 via standard mail. |  | 8450 | 845/10/10-plate pack cases | Nationwide Distribution. | Open, Classified |  |  | 8450 | Not_Computer
Z-1750-2015 | 71100 | Yes | CBK | Continuous ventilator. | Anesthesiology | Anesthesiology | 510(k) | Astral 100 ( Model No. 27001) &amp; Astral 150 (Model No. 27003) ventilator. |  |  | May 05, 2015 | 2015 | ResMed Ltd. | The Astral device allows clinicians to disable all alarms including those that detect circuit disconnection. Inappropriate disabling of alarms in ventilator dependent patients can lead to hazards of insufficient ventilation and insufficient expiratory pressure which can result in major severity harm. | Use error | A customer notification letter dated 5/5/15 will be sent to customers that the Astral device allows clinicians to disable all alarms including those that detect circuit disconnection. The letter provides the customers with the hazards involved and the actions to be taken. Customers are instructed to ensure alarms are appropriately set up for ventilator dependent patients. Customers with questions are instructed to contact Larissa D'Andrea at (858) 836-6837 or email larissa.d'andrea@resmed.com. |  | 1519 | 1519 units | Worldwide distribution. US nationwide, Uruguay and Canada. | Terminated | April 07, 2016 | 338 | 1519 | Not_Computer
Z-1745-2015 | 71393 | No | IZL | Mobile x-ray system. | Radiology | Radiology | 510(k) | AGFA DX-D 100 is a mobile digital radiography x-ray system |  |  | May 29, 2015 | 2015 | AGFA Healthcare Corp. | When using the DX-D 100 unit, sporadic unintended movements caused by an electrostatic discharge of the unit to the ground may occur. | Other | AGFA Healthcare sent an "URGENT FIELD SAFETY NOTICE" letter dated May 29, 2015, to all affected customers. The letter was either emailed or sent via FedEx to the US customers. When using the DX-D 100 unit, sporadic unintended movements caused by an electrostatic discharge of the unit to the ground may occur.Please distribute this information within your facility to all those who need to be aware of the notice. Please complete the feedback form as soon as possible and return it to us.We thank you for your careful attention to this issue and your continued support. If you have any questions about this matter, please contact your local Agfa HealthCare organization: 1-877-777-2432 and please reference PR1411200006 and PR1501190003. |  | 233 | 233 units | US Distribution to the states of : AL, CA, DC, FL, GA, IL, IA, KY, LA, MD, MA, MI, MN, MS, MO, NE, NJ, NY, NC, OH, OK, OR, SC, TN, TX, VA, WV, WI and WY. | Open, Classified |  |  | 233 | Not_Computer
Z-1751-2015 | 71241 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD) For use as a bridge-to-cardiac transplantation in patient who are at risk of death from refractory end-stage left ventricular heart failure. |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Discolored and Cracked Driveline Outer Sheath. | Device Design |  |  | 3747 | 3,747 currently in use | AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, and VA, Foreign: Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxemburg,Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK, | Open, Classified |  |  | 3747 | Not_Computer
Z-1744-2015 | 71365 | No | KIH | Cement dispenser. | Orthopedic | Orthopedic | 510(K) Exempt | The Optigun Ratchet is a hand-held cement gun for use with Optivac. The Optigun Ratchet is used to deliver bone cement from the Optivac. |  |  | May 14, 2015 | 2015 | BIOMET FRANCE S.A.R.L. | The pin which maintains the knob button, may disconnect and become lost. This may result in a delay of surgery to obtain a new Optigun and detect/retrieve the pin, and may necessitate revision surgery if the pin falls into the surgical site and is not detected. | Nonconforming Material/Component | BIOMET sent an URGENT ADVISORY NOTICE dated May 14, 2015, to all affected customers via courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers were instructed to check the product to identify to determine if the pin is defective, if defective they should return it to Biomet Warsaw, following the directions on the Response Form and confirm receipt of the notice by sending back the response form. Questions related to this notice should be directed to (574) 372- 1570 Monday through Friday 8 a.m. to 5 p.m. |  | 3 | 3 | Nationwide distribution to MD only | Terminated | October 27, 2015 | 166 | 3 | Not_Computer
Z-1749-2015 | 71297 | No | NKB | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | ORTHOFIX Firebird Spinal Fixation System 6.5mm x 55mm Multi-Axial Screw, Standard Implant Grade Titanium Alloy Rx Only Provided Non-Sterile Do Not Reuse |  |  | May 18, 2015 | 2015 | Orthofix, Inc | The Firebird Spinal Fixation System 6.5 mm x 55 mm Bone Screw, Self-Tapping is currently mislabeled at the distal tip of the bone screw showing a length measurement of 50 mm instead of 55 mm. | Unknown/Undetermined by firm | Orthofix sent an Urgent Medical Device Recall Notification letter dated May 15, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Affected customers were instructed to inspect their inventory, and identify and remove the recalled product to send back to Orthofix. The distributors are also instructed to provide recall notification to hospitals, surgeons, or other distributors that they have further distributed affected product to.Contact your local Orthofix representative or Orthofix customer service representative at 888-298-5700 for return instructions and shipment information. |  | 45 | 45 units | US Distribution to the states of : OK, MD, CA, FL, MI, SC, CO, NY and PA., and Internationally to Asturias Spain. | Terminated | October 30, 2015 | 165 | 45 | Not_Computer
Z-1759-2015 | 71134 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | MX 16 Slice (Uninterruptible Power Supply (UPS) accessory devices) Computed Tomography X-ray systems and Diagnostic Imaging systems. |  |  | March 06, 2015 | 2015 | Philips Medical Systems (Cleveland) Inc | The firm was notified that the failure of the UPS accessory devices used in conjunction with Computed Tomography X-ray systems and Diagnostic Imaging systems. may result in battery acid leakage, overheating and/or emission of fumes. | Nonconforming Material/Component | On 3/9/2015 the firm sent Urgent Field Safety Notices Medical Device Correction letters to their customers. |  | 125 | 125 units | Worldwide Distribution-US (nationwide) including DC &amp; Puerto Rico and the following countries: Albania, Algeria, Argentina, Australia, Austria, Bangladesh, Belgium, Bolivia, Brazil, Bulgaria, Cambodia, Cameroon, Cayman Islands, China, Colombia, Democratic Congo, Croatia, Cuba, Czech Republic, Denmark, Dominican Republic, Ecuador, Egypt, Ethiopia, Finland, France, Gabon, Georgia, Germany, Ghana, Greece, Guatemala, Hong Kong, Iceland, India, Indonesia, Iran, Iraq, Ireland, Israel, Italy, Japan, Jordan, Kazakhstan, Korea, Kuwait, Kyrgyzstan, LAO People's Dem. Rep., Latvia, Lebanon, Libya, Lithuania, Luxembourg, Malawi, Malaysia, Malta, Mauritius, Monaco, Mongolia, Morocco, Myanmar, Namibia, Nepal, Netherlands, New Zealand, Nicaragua, Nigeria, Norway, Oman, Pakistan, Palestine, Panama, Paraguay, Philippines, Poland, Portugal, RÂżunion, Romania, Russia, Rwanda, Saudi Arabia, Senegal, Serbia &amp; Montenegro, Singapore, Slovakia, South Africa, South Korea, Spain, Sri Lanka, Sudan, Sweden, Switzerland, Syria, Syrian Arab Republic, Taiwan, Tanzania, Thailand, Tunisia, Turkey, Uganda, Ukraine, United Arab Emirates, United Kingdom, Uzbekistan, Venezuela, Vietnam, Yemen, and Zambia. | Terminated | March 13, 2017 | 738 | 4590 | Not_Computer
Z-1757-2015 | 71171 | No | JDI | Hip joint metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Biomet Cobalt-Chrome Femoral Components, 32MM MOD HEAD COCR -6MM NECK; 32MM MOD HEAD COCR -3 MM NECK; 32MM MOD HEAD COCR STD NECK ; 32MM MOD HEAD COCR +3MM NECK; 32MM COCR MOD HD +6MM NO SKIRT |  |  | April 16, 2015 | 2015 | Biomet, Inc. | Biomet hip joint, knee joint and shoulder joint replacement components device labeling contains incorrect GTIN (Global Trade Identification Number) codes. | Labeling mix-ups | Biomet sent an Urgent Medical Device Recall Notice letter dated April 29, 2015 to affected customers via Fed-Ex. The letter identifed the affected product, problem and actions to be taken. Customers were asked to complete the attached response form and return it along with recalled product to Biomet. For questions callt (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 1315 | 1315 | Worldwide Distribution - US Nationwide in the states of AR, CA, IL, IN, PA, FL, KS, LA, WI, VA, MA, MD, NJ, UT, NC, TX, AL, MI, KY, OH, SD, MO, GA, NV, WA and the countries of: Europe, Canada, Japan, Chile, Korea, Mexico, New Zealand, Trinidad/Tobago, Thailand, Argentina, Australia, and China. | Open, Classified |  |  | 2632 | Not_Computer
Z-1758-2015 | 71269 | No | DXC | Vascular clamp. | Cardiovascular | Cardiovascular | 510(k) | Edwards Lifesciences IntraClude Intra-Aortic Occlusion Device, model number ICF100, 10.5 Fr (3.5 mm) x 39.4" (100 cm) ThruPort System. Packaged sterile and non-pyrogenic in a sealed, peel-type pouch |  |  | May 05, 2015 | 2015 | Edwards Lifesciences, LLC | Potential device damage on the strain relief near the hub area, which may result in blood leakage. | Device Design | Urgent Medical Device Safety Notice letters were sent via Fed-Ex by May 05, 2015. As part of the customer instructions in the safety notice, customers were requested to review the notice to understand the potential hazard, complete and return the acknowledgement form, and forward copies of the Safety Notice to any organization where the potentially affected devices have been transferred. The device will only be distributed to customers who have returned the acknowledgement form. Questions should be directed to Edwards Customer Service at 800-424-3278. |  | 3360 | 3,360 units. | Nationwide Distribution -- US, including the states of Alabama, Arizona, California, Colorado, Connecticut, Florida, Georgia, Idaho, Illinois, Kentucky, Louisiana, Massachusetts, Maine, Michigan, Missouri, Mississippi, North Carolina, Nebraska, New Jersey, New York, Ohio, Oklahoma, Oregon, Pennsylvania, South Carolina, South Dakota, Tennessee, Texas, Utah, Virginia, Washington, Wisconsin, and West Virginia. | Open, Classified |  |  | 3360 | Not_Computer
Z-1726-2015 | 71242 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD). |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Retraction of Pins within the driveline connector. | Device Design | Urgent Medical Device Correction letters, dated APR2015, were sent to users. The letters identified the affected device, as well as the reason for the recall and risks involved. Users were asked to review and familiarize themselves with the notice and "Patient Communication." The notice should be forwarded to anyone within customers' organization that need to be aware. Customers are to identify their patients who are currently supported by the HVAD System and distribute the "Patient Communication" to them. A HeartWare representative can assist with that process is needed. Customers should continue to reinforce the messages in the notice with their patients. The "Acknowledgement and Completion Form" should be completed, signed, and returned within 30 days of receipt of the notice. Questions or concerns should be directed to a local HeartWare representative. Clinical Support is also available at 1-888-494-6365. |  | 3747 | 3,747 currently in use | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, and VA; and, the countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Not_Computer
Z-1753-2015 | 71238 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | Gyrus ACMI Diego Elite Tubeset, intended for cutting, coagulation, drilling, debriding, and removal of bone, and soft and hard tissue in general ENT, Sinus/Rhinology, Nasopharyngeal/Laryngology, and Head and Neck procedures. Catalog No. TS101DC, Lot numbers JC924165 and below; Rx Only, STERILE EO |  |  | May 08, 2015 | 2015 | Gyrus Acmi, Incorporated | Potential lack of sterility assurance. | Under Investigation by firm | The firm notified their direct accounts of the recall by Fed Ex Second Day on 05/08/2015. The letter requesting disposal of the product was extended to hospitals or medical practices that purchased the product. Non-responding consignees will be contacted again if necessary. |  | 22250 | 22,250 boxes | Nationwide and Canada, Australia, Germany, New Zealand, Japan, India, Great Britain | Open, Classified |  |  | 34270 | Not_Computer
Z-1797-2015 | 71161 | No | MGQ |  |  | General &amp; Plastic Surgery | 510(k) | Amerigel Hydrogel Gauze Dressing 2 x 2, A5002. |  |  | February 25, 2015 | 2015 | Amerx Health Care Corp. | Nonapproved labeling claims (Sterile) | Labeling False and Misleading | On March 9, 2015 the firm contacted its customers via telephone. Within that conversation the firm asked its customers to examine their inventory to determine if they have any of the affected lot numbers on hand and notify the firm for an immediate product substitution, credit or refund. Customers were asked to discontinue dispensing any product on hand, and notify any customers that the product may have been further distributed to. The firm also asked that the response form be completed and returned. If no contact was made via telephone, the firm sent a certified letter asking the firm(s) to do the same actions as the phone request. If you have any questions please call 1-800-448-9599 or e-mail marketcorrection@amerxhc.com. |  | 6491 | 6,491 units | Worldwide Distribution -- US, including the states of AK, AL, AR, AZ, CA, CO, CT, FL, GA, IA, ID, IL, IN, KS, KY, LA, MA, MD, ME, MI, MO, MS, MT, NC, NE, NJ, NM, NV, NY, OH, OK, OR, PA, RI, SC, SD, TN, TX, UT, VA, VT, WA, WV, WY and Hawaii; and, the countries of Canada and Bahrain. | Open, Classified |  |  | 6491 | Not_Computer
Z-1794-2015 | 71278 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Medtronic Passive Planar Blunt Probe Part Number 960-556, also included in Part Number Kits: 9732540, 9732622, 9733174, 9733935. As part of kits, the product is either packaged with other instruments or into a holding tray. The Passive Planar Blunt Probe is an optically tracked, reusable instrument used in navigated surgery and is included with Medtronic cranial instrument kits. It is designed for use in patient registration and palpating brain tissue during surgical procedures. |  |  | May 05, 2015 | 2015 | Medtronic Navigation, Inc. | Passive Planar Blunt Probe (Lot# 150223) was manufactured with a sharp tip rather than a blunt tip. | Mixed-up of materials/components | Medtronic sent an Medical Device Safety Alert letter dated May 5, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer.Action to be taken by the user:"Examine your inventory and if any of the affected products are found, please quarantine them for return to Medtronic."A Medtronic representative will follow-up to arrange for removal and replacement.If you have any questions regarding this action, please contact Medtronic Navigation Technical Services at 1-800-595-9709. |  | 18 | 18 | US Distribution to the states of : CA, MO, LA and PA., and Internationally to Ireland. | Terminated | July 02, 2015 | 58 | 18 | Not_Computer
Z-1796-2015 | 71405 | No | DPS | Electrocardiograph. | Cardiovascular | Cardiovascular | 510(k) | ELI 380 Electrocardiograph. |  |  | May 29, 2015 | 2015 | Mortara Instrument, Inc | Mortara Instrument, Inc. has recently becomeaware of a potential safety hazard involving our ELI380 electrocardiograph. When used in a particularworkflow, acquired ECG waveforms for one patientmay become associated with the patientdemographics for a different patient when therecord is transmitted to a records managementsystem. | Software design | Mortara Instrument, Inc. sent an "Urgent Medical Device Correction" letter dated May 29, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Customers were advised that users in their facilities be made aware of the safety notification and that required actions listed on the letter be taken immediately. A Field service Engineer will be dispatched to their facility to update the affected unit(s) on-site. For questions they can contact Technical support at 1-888-667-8272 or Techsupport@mortara.com or contact your local service Representative. |  | 176 | 176 | Worldwide Distribution - US Distribution to the states of : AZ, CA, FL, GA, IL, MI, NE, SC, UT, WA and WI., and to the countries of : CHINA, PORTUGAL, CANADA, OMAN, MALAYSIA, NEW ZEALAND, INDIA, POLAND, LEBANON, SWITZERLAND, FINLAND, ROMANIA,UNITED ARAB EMIRATES, UNITED KINGDOM, AUSTRALIA, THE NETHERLANDS, ITALY, FRANCE, GERMANY, THAILAND,SAUDI, ARABIA, NORWAY and IRELAND. | Terminated | August 05, 2015 | 68 | 176 | Not_Computer
Z-1795-2015 | 71264 | No | GWN | Nystagmograph. | Neurology | Neurology | 510(k) | I-PortalÂż NOTC and VNG; a device used to measure, record, or visually display the involuntary movements (nystagmus) of the eyeball, indicated for use as a diagnostic tool to assist trained clinicians in their analysis of vestibular disorders, which requires the separation of central vs. peripheral nervous system deficits. |  |  | April 30, 2015 | 2015 | Neuro Kinetics, Inc. | complaints of system malfunction and unintended, sudden movement at start up. No injuries reported. | Software design | Neuro Kinetics sent a Field Correction Advisory letter dated April 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers with the affected product were instructed to contact Customer Service if a malfunction is encountered and not to use the device until it has been assessed and corrected by Nuerokinetics. Customers with questions should call 412-963-6649. |  | 140 | 140 | Worldwide distribution - US (nationwide) and Internationally to Canada and Mexico. | Terminated | April 17, 2017 | 718 | 140 | Not_Computer
Z-1731-2015 | 71239 | No | DSQ |  |  | Cardiovascular | PMA | HeartWare Ventricular Assist System (HVAD). |  |  | April 29, 2015 | 2015 | HeartWare Inc | Complaints with the HVAD Internal Controller Alarm Battery failures. | Device Design | Urgent Medical Device Correction letters were sent to customers. The letter identified the affected product, as well as the reason for the recall and risks involved. HeartWare requests that customers review and familiarize themselves with the letter and "Patient Communication," and ensure that all necessary individuals are aware. Patients currently being supported by the device should be identified and given the "Patient Communication." A HeartWare representative can provide assistance with that process if needed. The "Acknowledgement and Completion Form" should be completed, signed, and returned within 30 days of the receipt of the letter. Customers can contact their local HeartWare representative with questions or concerns. Clinical Support personnel are also available at 1-888-494-6365. |  | 3747 | 3,747 curently implanted | Worldwide Distribution -- US, including the states of AL, AZ, CA, CO, CT, DC, DE, FL, GA, IA, IL, IN, KY, LA, MA, MD, MI, MN, MO, MS, NC, NH, NJ, NY, OH, OR, PA, SC, TN, TX, UT, and VA; and, the countries of Argentina, Australia, Austria, Belgium, Brazil, Canada, Chile, Croatia, Czech Republic, Denmark, Egypt, Finland, France, Germany, Greece, Hong Kong, Hungary, India, Israel, Italy, Japan, Korea, Lebanon, Lithuania, Luxembourg, Malaysia, Netherlands, New Zealand, Norway, Poland, Romania, Singapore, Slovakia, South Africa, Spain, Sweden, Switzerland, Turkey, and UK. | Open, Classified |  |  | 3747 | Not_Computer
Z-1808-2015 | 70974 | No | NGT | Intravascular catheter. | General Hospital | General Hospital | 510(k) | T/S2 0.9% Sodium Chloride Injection, USP, AMUSA Sterile Field Flush Syringe, 10 mL in 12 mL Single Use Syringe, Rx only, Part no. 2S0706 |  |  | February 09, 2015 | 2015 | MRP, LLC dba AMUSA | Incorrect labeling of the expiration date | Process control | The firm, AMUSA, Inc., initiated the recall by telephone to the sole consignee on 02/09/2015. The consignee were informed of product, the problem and actions to be taken. The consignees verified that they were in possession of the product and that all the product has been destroyed.If you have any questions, call 513-235-1940 or email: qadirector@amusa.us. |  | 1080 | 1,080 units (3 cases) | US distribution to CA only. | Terminated | June 22, 2015 | 133 | 1080 | Not_Computer
Z-1801-2015 | 71247 | No | FMI | Hypodermic single lumen needle. | General Hospital | General Hospital | 510(k) | T.A.L.O.N (Tactically Advanced Lifesaving Intraosseous Needle) 15G x 38.5 mm 10 NEEDLE SET STERILE EO Single Use Rx only Vidacare Corporation |  |  | May 12, 2015 | 2015 | Vidacare Corporation | Potential breach in sterility | Device Design | Teleflex sent an Urgent Medical Recall Notification dated May 12, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. The letter directed the customers to quarantine the affected product and return it to the manufacturer. If you have any further questions please call 1-866-246-6990. |  | 2104 | 2,104 boxes | US Distribution to the states of : NC, MD, MS, TX and MO., and Internationally to Canada and Denmark. | Terminated | April 11, 2016 | 335 | 2104 | Not_Computer
Z-1800-2015 | 71276 | No | HWC | Smooth or threaded metallic bone fixation fastener. | Orthopedic | Orthopedic | 510(k) | 3.7mm Cannulated Locking Screws and 3.7mm Cannulated Conical Screws; . the Cannulated Screws are indicated for fixation of fractures, osteotomies, and non-unions of the clavicle, scapula, olecranon, humerus, radius, ulna, pelvis, tibia, and fibula, particularly in osteopenic bone. |  |  | May 13, 2015 | 2015 | Synthes, Inc. | Certain lots of the 3.7mm Cannulated Locking Screws and 3.7mm Cannulated Conical Screws were found to have been labeled incorrectly. | Employee error | DePuy Synthes sent an Urgent Notice Medical Device Recall letter dated May 13, 2015, to all affected customers to advise them about the issue and associated risks. These accounts were asked to review inventory, immediately remove the affected lots that are still in the original packaging from stock, and to complete and return the response and return authorization forms. Customers with questions were instructed to call 610-719-5450 or contact their DePuy Synthes consultant. |  | 54 | 54 | Nationwide Distribution | Open, Classified |  |  | 54 | Not_Computer
Z-1798-2015 | 71253 | No | HRY | Knee joint femorotibial metal/polymer semi-constrained cemented prosthesis. | Orthopedic | Orthopedic | 510(k) | Vanguard M Unicondylar Tibial Trial |  |  | April 22, 2015 | 2015 | Biomet, Inc. | The Left and Right orientation of the trial was incorrectly etched on one side. When the instrument is rotated about the axis of the handle 180 degrees to view the other side, the etch on the handle now shows R on the same side as the keel marked LM and the L on the same side as the keel marked RM. This is incorrect. Please note that the keel is not etched on this side by design. | Process control | On 4/22/2015, URGENT MEDICAL DEVICE RECALL NOTICE notifications, dated April 22, 2015 were sent to the affected distributors with instructions for returning the affected product. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers are required to immediately locate and remove the affected product from circulation, follow the instructions on the "Response Form," and e-mail a copy of the completed form to the designated e-mail address prior to returning the product. Additionally, if the affected product has been further distributed, the necessary personnel must notified via the "Dear Risk/Recall Manager" notice. Questions related to this notice should be directed to (574) 372-1570, Monday through Friday, 8 a.m. to 5 p.m. |  | 27 | 27 | Worldwide Distribution -- US, including the state of IL, NJ, Canada, Netherlands, and Chile. | Terminated | March 22, 2016 | 335 | 27 | Not_Computer
Z-1806-2015 | 71293 | No | FSY | Surgical lamp. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | BVI Custom EyesÂż Ophthalmic Procedure Kits with the component -Devon Light Glove |  |  | May 01, 2015 | 2015 | Beaver-Visitec International Inc. | Devon Light Glove in procedure kits may have holes/slits compromising sterility of the product | Nonconforming Material/Component | Beaver-Visitec International Sales Ltd. (BVI) sent an Urgent Field Safety Notice dated May 2, 2015, to all affected customers. Customers were instructed to quarantine the affected kits and return to BVI for subsequent replacement of product. A Verification Form was provided to confirm receipt of the recall notice. For questions regarding this recall call 791-906-7950. |  | 268 | 268 kits | Nationwide Distribution including KY, GA, MI, and NJ. | Open, Classified |  |  | 2072 | Not_Computer
Z-1799-2015 | 71364 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | Pull Reduction Device For 4.3mm Percutaneous Drill Guide; The pull reduction device with quick coupling is placed through the guide sleeve and plate holes to pull or push bone fragments relative to the plate. This instrument can be used for: Minor Varus-valgus adjustments (approximately 2Âż - 4Âż), Translational adjustments, Stabilization of plate bone orientation during insertion of the first screws, alignment of segmental fragments and pre-drilling dense or thick cortical bone before placing a 5.0 mm locking screw. |  |  | May 28, 2015 | 2015 | Synthes, Inc. | Certain affected parts and lots of the Pull Reduction Devices may have been manufactured to an incorrect hardness specification which could result in intraoperative breakage. | Component design/selection | DePuy Synthes sent an Urgent Notice Urgent notice dated May 28, 2015, to all affected customers to advise them of the defect and patient risks associated with its use. Customers were instructed if they have any of the identified devices, please take the following steps:Âż Call DePuy Synthes at 1-800-479-6329 to obtain a Return Authorization (RA) Number.Âż Complete the Verification Section (page 3 of this letter) by checking the appropriate box indicating affectedproduct has been located. Also, please indicate the number of devices found and note the ReturnAuthorization Number. Please include your name, title, address, telephone number and signature in thespaces provided.Âż Return the Verification Section (page 3 of this letter) with the product to:Âż Credit/Returns, DePuy Synthes, 1101 Synthes Avenue, Monument, CO 80132.Âż Send a copy of the completed Verification Section by:Âż Fax: (888) 943-5206 orÂż Scan/email: Synthes2912@stericycle.comIf you DO NOT have the identified product, please take the following steps:Âż Complete the attached Verification Section (page 3 of this letter) by checking the appropriate box indicatingthat no affected product has been located. Please include your name, title, address, telephone number andsignature in the spaces provided. This return documentation acknowledges your receipt of medical deviceremoval information.Âż Return the documents by:Âż Fax: (888) 943-5206 orÂż Scan/email: Synthes2912@stericycle.comÂż Note: If the Verification Section is answered on behalf of more than one facility and/or individual, pleaseclearly indicate the name and address of the facility and/or individual on page 3 of the notification.Customers with questions were instructed to call 610-719-5450.For questions regarding this recall call 610-719-5000. |  | 3766 | 3747 US, 19 Foreign | Worldwide Distribution-US (nationwide) and internationlly to Canada | Open, Classified |  |  | 3766 | Not_Computer
Z-1807-2015 | 71432 | No | LLZ | Picture archiving and communications system. | Radiology | Radiology | 510(k) | Philips Ultrasound, Model Q-Station, with software version 3 or higher, Catalog number: 795088; Part Number: 989605382391 |  |  | May 15, 2015 | 2015 | Philips Ultrasound, Inc. | Philips Healthcare has discovered a problem in the Philips Ultrasound Q-Station version 3.0 or higher that couldresult in measurements from a study (Structured Report) to be appended to subsequent studies for other patients. | Software design | PhilipsHealthcare sent an URGENT - Medical Device Correction Philips Ultrasound Q-Station, Q-STATION STRUCTURED REPORTING AND MEASUREMENTS letter, FCO 79500335, dated 2015 May 15, to all affected consignees.Philips is sending a CD with Q-Station software version 3.3.2 to consignees for a free of charge. Consignees must uninstall previous versions of Q-Station and install version 3.3.2 on your system carefully following the installation instructions that accompany the CD. It is important that consignees register their Q Station 3.3.2 installation as described in those instructions, so that Philips can confirm that they have completed this correction.Consignees were instructed to complete the enclosed Customer Reply Form to confirm that they have (1) read and understood this important Medical Device Correction and (2) have either installed Q-Station version 3.3.2 or declare that this FCO 79500335 does not apply to their institution. Consignees were asked to return the enclosed Customer Reply form to Philips Healthcare within 5 days of receipt of the replacement software. Consignees with questions can call local Philips representative or Philips Customer Service at 1-800-722-9377. |  | 125 | 125 units (30 in US and 95 international) | Worldwide Distribution - US (nationwide) and Puerto Rico; internationally to Australia, Canada, Colombia, France, Germany, Great Britain, India, Italy, Mexico, Netherlands, Norway, Pakistan, Poland, Russia, Singapore, Spain, Switzerland, Thailand, and Zimbabwe. | Open, Classified |  |  | 125 | Computer
Z-1811-2015 | 71081 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | Alaris PC unit model 8015, Infusion pump. software version 9.17 |  |  | May 13, 2015 | 2015 | CareFusion 303, Inc. | An issue with the cancel functionality used during atypical infusion programming to cancel user inputted values. An infusion may start that is greater than or less than the hospital established limits for the specific medication. | Software design | CareFusion sent an Urgent Medical Device Recall Notification letter dated May 13, 2015 to customers. The letter identified the affected product, problem and the recommended actions to be taken. Customers are informed that CareFusion has released a software upgrade to address this issue. CareFusion will contact all affected customers within 60 days to initiate the scheduling process for the software upgrade installation. Customers with recall related questions are instructed to contact CareFusion Support Center at (888) 562-6018, 7am-4pm PT, supportcenter@carefusion.com. Customers with adverse event report questions are instructed to contact customer advocacy at (888) 812-3266, 24 hrs/day, 7 days a week, customerfeedback@carefusion.com. Customers with technical questions regarding the Alaris System are instructed to contact Technical Support at (888) 812-3229, 6am-5pm PT, Mon-Fri, dl-us-inf-techsupport@carefusion.com. |  | 63 | 63 units | US Nationwide Distribution | Open, Classified |  |  | 63 | Computer
Z-1693-2015 | 71306 | No | JAK | Computed tomography x-ray system. | Radiology | Radiology | 510(k) | Neusoft NeuViz 64 Multi-slice CT Scanner System |  |  | March 08, 2014 | 2014 | Neusoft Medical System Co | The A-plane compensator is made of Teflon which will become fragile due to X-ray exposure passing through it. The thinnest part of the compensator is subject to radiation degradation and as a result has been found to fail prematurely. Failure is indicated by pin or ring artifact in the image | Radiation Control for Health and Safety Act | Neusoft planned action:1. Notification of all dealers and purchasers is to be made within 15working days of receipt of this letter in the manner specified in 21 CFR1003.21 and 1003.22. This office and the Food and Drug Administration(FDA) district office coordinator noted below are to be included in thenotification.2. Neusoft Medical Systems Co., Ltd will implement actions to monitor theexposures of the tube and when the exposure number is close to300,000 seconds, the A-plane compensator will be replaced. The Foodand Drug Administration (FDA) district office coordinator should benotified when the last A-plane compensator is replaced.3. The Customer Information Letter has been sent to customers as acorrective action. The maximum exposure number of 5 US systems is191830 seconds as of Apr.20, 2015, so the firm will monitor the datacontinuously and replace the A-plane compensators before the definedthreshold.4. Neusoft will replace all compensators that have exceeded the 300,000seconds free of charge and will replace all compensators that fail duringwarranty and the compensators failed less than 300,000 secondsthough out of warranty free of charge.You may proceed or continue with implementation of the CAP (subject to the conditions noted above) For further questions please call (281) 453-1205 |  | 5 | 5 installed in US | Worldwide Distribution - US Distribution including Puerto Rico and to the states of : IL, MA, TX and SC., and to the countries of : China, Malaysia, Lebanon, Ecuador, Egypt, United Arab Emirates,Kazashstan, Tunisia, Angola, Brazil, Uzbekistan, Lebanon, Cote D' ivoire, South Africa, Montenegro and Tajikistan. | Open, Classified |  |  | 5 | Not_Computer
Z-1810-2015 | 71268 | No | HAW | Stereotaxic instrument. | Neurology | Neurology | 510(k) | Navigation System II-Cart, 7700-100-000 with Articulated Arm Camera, 7700-103-001 and Articulated Arm Camera, 7700-103-001U |  |  | May 04, 2015 | 2015 | Stryker Instruments Div. of Stryker Corporation | Stryker initiated a device recall of the Navigation System II-Cart, Camera Articulated Arm because there is a potential that the weld seam between the main stud and the two flaps on the Articulated Arm was not welded in the correct location, which may result in a potential failure of the arm joint. | Process control | Stryker sent an Urgent Medical Device Recall Notification letter dated May 18, 2013 to Domestic customers notice via FedEx overnight and Sales Representatives via email. The International Subsidiaries will also be notified via email. Customers are instructed to immediately review the Recall Notification adn complete the following: Immediately check all stock areas and/or operating room storage for affected equipment found. Complete the enclosed Business Reply Form (BRF) to confirm receipt of this notification and identify how many affected items are currently in inventory. Please complete and return the BRF even if they dont have any affected product on hand. If customers have further distributed this product, they are asked to forward this letter and the attached BRF to all affected locations. Please indicate each location and serial number(s) on the BRF.Fax (866-521-2762) or email (kellyjo.davis@stryker.com) the completed BRF to Stryker Instruments Regulatory Department, Attn: Kelly Jo Davis. A Stryker Representative will contact the facility to set up a time to perform the additional inspection of the weld seam on the Navigation System II-Cart, Articulated Arm Camera. |  | 72 | 72 | Worldwide Distribution - US Nationwide in the states of AR, CO, FL, GA, IN, KS, KY, MA, MI, NC, NE, NJ, NY, OH, OK, PA, TX, WA, WY and the countries of Canada, Europe, and Asia. | Terminated | May 16, 2016 | 378 | 72 | Not_Computer
Z-1812-2015 | 71285 | No | NAY | Endoscope and accessories. | Gastroenterology/Urology | General &amp; Plastic Surgery | 510(k) | Thoracic Grasper, 5 MM, for use with da Vinci S (IS2000) and SI (IS3000) Surgical Systems. |  |  | May 21, 2015 | 2015 | Intuitive Surgical, Inc. | The coating of the 5 mm Thoracic Grasper main tube may experience damage with the possibility of fragments of coating material dislodging and potentially falling into patient. | Nonconforming Material/Component | Urgent Medical Device Recall letters were sent to all consignees on May 21, 2015.The letter describes the affected product and issue involved. Additionally, the risk to health, as well as the actions to be taken were also discussed in the letter. Customers are to ensure that all affected personnel were informed of the notice. All affected product are to be identified and returned by contacting Customer Service for a RMA. The Acknowledgment Form is to be completed and returned using the instructions provided. If further information or support concerning this recall is needed, customers are to contact their local Clinical Sales Representative or Intuitive Surgical Customer Service. |  | 2436 | 2,436 units | Worldwide Distribution -- US, Belgium, Brazil, Canada, Chile, Colombia,, Denmark, France, Germany, India, Netherlands, Panama, Singapore, South Korea, Switzerland, Taiwan, Turkey, and United Kingdom. | Open, Classified |  |  | 2436 | Not_Computer
Z-1816-2015 | 70164 | No | OOI | Instrumentation for clinical multiplex test systems. | Clinical Chemistry | Clinical Chemistry | 510(K) Exempt | BD Viper LT System, catalog number 442839 when used in conjunction with the BD ProbeTec Chlamydia trachomatis (CT) Qx Assay (cat. #442959) and the BD ProbeTec Neisseria gonorrhoeae (GC) Qx Assay (cat. #442842). In vitro diagnostic. |  |  | November 07, 2014 | 2014 | Becton Dickinson &amp; Co. | Potential for false positive Chlamydia trachomatis (CT) results and Neisseria gonorrhoeae (GC) results when processing the CT Qx and GC Qx Assays on the BD Viper LT Instrument. | Device Design | BD sent an Urgent Field Corrective Action letter dated November 2014, to all affected customers. The letter identified the problem the product and the action needed to be taken by the customer. Customers were advised: a. Include the date recall was initiated, if it is already underway or the date your firm plans to start the recalling process. b. How do you plan to notify all the consignees affected by this recall? (press release, letter, telefax, telephone, e-mail, visit, etc.) The customer was initially contacted via a phone call and subsequently provided with a written communication. c. How do you plan to monitor the number of consignees non-responding to the recall communication? The single US customer was contacted by phone and provided the written communication.d. How do you plan to do effectiveness checks of this recall? The affected instrument will be corrected.e. Date your firm ceased further distribution of the product(s). Product was not placed on hold. This issue does not affect new customers. f. How do you plan to store the recall product? BD has not requested that instruments or reagents be returned.g. How do you plan to dispose of the recall products? BD is not requesting that customers return the instruments or reagents.If you require further assistance or clarification, please contact the BD Technical Services Department at 1-800 638-8663. |  | 7 | 7 Viper LT Systems | Worldwide distribution to US (state of LA), Hong Kong, Austria, China, France, and the UK. | Terminated | May 09, 2016 | 549 | 7 | Not_Computer
Z-1814-2015 | 71329 | No | HDW | Contraceptive diaphragm and accessories. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(k) | MILEX OMNIFLEX DIAPHRAGM SIZE 80 |  |  | May 20, 2015 | 2015 | CooperSurgical, Inc. | A diaphragm size 80 labeled box was incorrectly packaged with a size 85. | Labeling mix-ups | CooperSurgical notified consignees by letter on 5/18/15 via Fedex with confirmed delivery receipt.. Accounts requested to discontinue use and complete the attached Acknowledgement and Receipt Form for replacement. CooperSurgical provided further instruction on the disposition of the product at the time of replacementIf you have any further questions please feel free to contact 203.601.5200.. |  | 200 | 200 units | AK, CA, DC, FL, IL, ME, MT, NE, NJ, NM, NY, OK, OR, P A, VA, W A, and WIForeign: CANADA, ENGLAND, and POLAND | Terminated | January 20, 2016 | 245 | 200 | Not_Computer
Z-1817-2015 | 71398 | No | DRQ | Transducer signal amplifier and conditioner. | Cardiovascular | Cardiovascular | 510(k) | CLEARSIGN II Amplifier for LABSYSTEM PRO EP Recording System, 120 channels, using firmware version 2.08. The channel numbers in the amplifier are as follows: |  |  | June 02, 2015 | 2015 | Boston Scientific Corporation | Boston Scientific has received complaints indicating that, when using the blood pressure (BP) channels on the CLEARSIGN II Amplifier, the surface Electrocardiogram (ECG) channels become over-written to a variable degree, with the result that it appears shifted from baseline on the system's output screen. This may, in turn, manifest as an uninterpretable ECG signal in the affected channel. | Software design | Boston Scientific sent an Important Field Safety Notification letter to all affected customers beginning June 3, 2015, by 24 hour overnight delivery. The letter identified the product the problem and the action needed to be taken by the customer. INSTRUCTIONS:1. Please read carefully the Field Safety Notice letter and immediately post this information in a visible location near the product to ensure it is easily accessible to all users of the device.2. Please complete the attached Verification Form even if you do not have any product to return.3. When completed, please return the Verification Form to your local Boston Scientific Office to the attention of ÂżCustomer_Service_Fax_NumberÂż on or before DATE 2015.4. If you have products to return, and once a loaner unit is available, a Boston Scientific representative will coordinate with you to package the product being returned using a special shipping container.Your Competent Authority is being notified of this Field Safety Notice.We regret any inconvenience that this action may cause and we appreciate your understanding as we take action to ensure patient safety and customer satisfaction. If you have any questions or would like assistance with this Field Safety Notice, please contact your local Sales Representative. For further questions please call (978)-805-3200 |  | 43 | 43 units | No US distribution. Worldwide Distribution to the countries of : Austria, Belgium, France, Germany, Great Britain, Greece, Italy, Netherlands, Poland, Romania, South Africa, Spain, Switzerland and Turkey. | Open, Classified |  |  | 43 | Not_Computer
Z-1815-2015 | 71343 | No | GWM | Intracranial pressure monitoring device. | Neurology | Neurology | 510(k) | Integra Camino Intracranial Pressure Monitoring Catheter with Licox IMC Bolt Fitting, Catalog No. 110-4L |  |  | May 27, 2015 | 2015 | Integra LifeSciences Corp. | Integra LifeSciences is recalling the Camino Intracranial Pressure Monitoring Catheters because they may have been incorrectly assembled. Specifically, the o-ring in the bolt adapter may be in the wrong location and as a result, there is a potential CSF/blood could leak from the bolt adapter. | Under Investigation by firm | Integra sent an Urgent Medical Device Recall letter dated May 27, 2015, to all customers who received the Camino Intracranial Pressure Monitoring Catheters because they may have been incorrectly assembled. Specifically, the o-ring in the bolt adapter may be in the wrong location and as a result, there is a potential CSF/blood could leak from the bolt adapter. The letter informs the customers of the problems identified and the actions to be taken. Customers were instructed to complete the customer response form and return it by email or fax to: FCA3@integralife.com or FAX to 1-609-275-9445. Customers with any questions regarding the instructions are instructed to contact Customer Service at 1-888-601-0203. For questions regarding this recall call 609-936-6822. |  | 1218 | 1218 units | Worldwide Distribution - US (Nationwide) and Internationally to ARGENTINA, BRAZIL, BULGARIA, CANADA, COLUMBIA, COSTA RICA, FRANCE, GERMANY, ITALY, MEXICO, GERMANY, ITALY, MEXICO, PANAMA, SOUTH AFRICA, SPAIN, SWITZERLAND, UNITED ARAB EMIRATES, AND UNITED KINGDOM | Terminated | December 09, 2015 | 196 | 1218 | Not_Computer
Z-1813-2015 | 71304 | No | MEA | Infusion pump. | General Hospital | General Hospital | 510(k) | CADDÂż- Solis Ambulatory Infusion Pump, Model 2111 &amp; 2112, Version 1.0, Version 2.0 and Version 3.0 &amp; Model 2120, Version 1.0. |  |  | May 26, 2015 | 2015 | Smiths Medical ASD, Inc. | Smiths Medical has become aware of an issue with an intermittent occurrence of binding of the locking assembly on some CADDÂż-Solis pumps. Binding of the Cassette/ Keypad Lock can occur after latching the CADDÂż-Solis Medication Cassette Reservoir or Administration Set disposable to the pump. When binding occurs, it can prevent the key from fully rotating the Cassette/ Keypad Lock to the locked (e | Nonconforming Material/Component | Consignees were sent on 5/26/2015 a Smiths Medical "Urgent Medical Device Recall Notice" letter dated 26 May 2015. The letter described the problem and the device involved in the recall. The letter provided "Advise on Action to be taken by the User" and " Transmission of this Urgent Medical Device Recall Notice". Consignees were requested to complete and return the Confirmation Form. For questions they can contact Smiths Medical's Customer Service Department at 1-800-258-5361, press 2, then press 1. |  | 4426 | 4426 (2427 US, 1999 OUS) | Worldwide Distribution-US (nationwide) including the states of AL, AZ, AR, CA, CO, CT, FL, GA, ID, IL, IN, IA, KS, LA, MA, MI, MN, MS, MO, NE, NM, NY, NC, ND, OH, OK, PA, SC, SD, TN, TX, UT, VA, WA, and WI, and the countries of AT (Austria), AU (Australia), BE (Belgium), CA (Canada), CH (Switzerland), DE (Germany), DK (Denmark), FR (France), GB (Great Britain), HK (Hong Kong), LB (Lebanon), NL (Netherlands), NZ (New Zealand), QA (Qatar), SA (Saudi Arabia), and SE (Sweden). | Open, Classified |  |  | 4426 | Not_Computer
Z-1820-2015 | 70675 | No | HMX | Manual ophthalmic surgical instrument. | Ophthalmic | Ophthalmic | 510(K) Exempt | Steriseal-Cannula Ophthalmic; |  |  | March 02, 2015 | 2015 | Hill-Rom, Inc. | The metal cannula attached to the handpiece can rotate and become dislodged from the plastic handpiece. | Use error | Customers were notified via telephone and e-mail by 03/02/2015. The firm is requiring customers to complete a submitted response form and certificate of destruction for recalled devices by 03/31/2015. If recalled devices were further distributed the firm is asking customers to conduct a sub recall and forward notice of the recall. Public contact is Aspen Medical Europe Ltd, QA Department, Lynda McConaghy-QARA Manager, Lynda.McConaghy@aspenmedicaleurope.com , or +44 (0)1527 587716. |  | 66691 | 66691 units | Worldwide Distribution-US (nationwide) including the states of MI &amp; PA, and the countries of Belgium, England, Italy, Czech Republic, Turkey, Greece, Ireland, Scotland, Spain, and Lebanon. | Open, Classified |  |  | 66691 | Not_Computer
Z-1818-2015 | 71132 | No | FSA | Non-AC-powered patient lift. | General Hospital | General Hospital | 510(K) Exempt | Viking M, L, and XL Mobile Lifts. |  |  | March 20, 2015 | 2015 | Hill-Rom, Inc. | Complaints (including one reported death) allegedly of the lift arm drifting down suddenly. If the lift arm assembly is manually lifted, the actuator can become damaged and get stuck in the highest position. If a patient is lifted into the sling while the actuator is stuck, there is a potential risk of a free fall of the patient, resulting in minor or potentially catastrophic injuries. | Labeling design | On 3/20/2015, URGENT Field Safety Notice notifications, dated 3/19/2015 were sent to US customers by e-mail and or courier. The recall notification included a description of the reason for the recall, affected product, consignee responsibilities, and instructions for responding to the formal recall notification. Customers are forward a copy of the letter to any other appropriate facility personnel. Additionally, the attached form should be completed and returned to Hill-Rom. Upon receipt, the firm will contact customers to schedule an inspection of the potentially affected devices. For questions, call 800-445-3720 (Mon - Thurs, 8 am - 6:30 pm; Friday, 8 am - 6 pm). |  | 24798 | 24, 798 | Worldwide Distribution -- AE, AO, AR, AT, AU, AZ, BD, BE, BG, BH, BR, CA, CH, CL, CN, CO, CS, CZ, DE, DK, EG, ES, FI, FR, GR, HK, HR, HU, ID, IE, IL, IN, IQ, IR, IS, IT, JO, JP, KR, KW, LB, LI, LT, LU, LV, MA, MT, MX, MY, NL, NO, NZ, OM, PH, PL, PT, QA, RO, SA, SE, SG, SI, SK, SY, TH, TN, TR, TW, UK, US, VE, and ZA. | Terminated | August 04, 2017 | 868 | 24798 | Not_Computer
Z-1821-2015 | 71332 | No | FPA | Intravascular administration set. | General Hospital | General Hospital | 510(k) | Sure T, contact, contact detach, neria, neria detach, neria multi and thalaset |  |  | May 21, 2015 | 2015 | Unomedical As | Unomedical has found that in rare cases the steel needle can break during use, interrupting the delivery of medication including insulin. | Nonconforming Material/Component | Each consignee was sent an "Urgent Field Safety Notice Regarding Steel Cannula Infusion Sets" on 5/21/2015 via Express Mail. This notification identified the affected product as well as the reason for the recall. It also contained instructions on how to use the product. Customers were to contact their health care provider if they suspected that a needle had broken off and remained under the skin. Diabetic patients were to monitor their blood sugar and non diabetic patients were to monitor symptoms which would indicate drug delivery had failed. Questions should be directed to the persons identified in the letter. |  | 22976870 | 22,976,870 units | Worldwide Distribution -- US, including the states of CA, IL, MN, and PA; and, the countries of Australia, Bahrain, Belgium, Canada, China, Cypress, Denmark, Estonia, France, Germany, Greece, Hong Kong, Hungary, Israel, Italy, Kuwait, Lebanon, Malaysia, Maldives, Oman, Schweiz, South Africa, Spain, Sweden, Switzerland, Netherlands, Turkey, United Arab Emirates, and United Kingdom. | Open, Classified |  |  | 22976870 | Not_Computer
Z-1819-2015 | 71262 | No | FZX | Manual surgical instrument for general use. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | ORTHOFIX Centurion POCT System Drill Guide PN 69-1010 and PN 69-1012. Drill Guides are used to interface with various drill tips during the drilling procedure of bone preparation. |  |  | May 11, 2015 | 2015 | Orthofix, Inc | There is a possibility that the thumb button remains in the depressed (unlocked) position allowing the threaded sleeve (drill stop) to unintentionally rotate and change the depth setting. | Device Design | Orthofix sent an Urgent Medical Device Product Advisory Notification letter dated May 11, 2015, to the affected distributors and hospitals. The consignees were directed to inspect thier inventory and perform checks before surgical use of the drill guides. Customers were instructed to provide a copy of the Product Advisory Notification letter if product was further distributed. Customers with questions were instructed to contact their local Orthofix representative and to for returns they should contact Orthofix customer service representative at 888-298-5700. |  | 36 | 36 units | Nationwide Distribution including CA, VA, FL, ID, CO, OR, PA, OK, SC, and TX. | Terminated | February 17, 2016 | 282 | 36 | Not_Computer
Z-1827-2015 | 71160 | Yes | KPO | Hemodialysis system and accessories. | Gastroenterology/Urology | Gastroenterology/Urology | 510(k) | NaturaLyte Liquid Bicarbonate Concentrate (Dialysate Concentrate for Hemodialysis (liquid), 6.4 Liter Bottle |  |  | May 20, 2015 | 2015 | Fresenius Medical Care Holdings, Inc. | Bacterial contamination. | Process control | Fresenius Medical Care sent an Urgent Medical Device Recall letter dated May 15, 2015, to all affected customers. Users were requested to Immediately examine stock to determine whether they have any NaturalyteÂż Liquid Bicarbonate Concentrate of the lots." If any product of these lots is found, discontinue use immediately." Place all units in a secure, segregated area." If affected product was on the machine prior to patient treatment, perform a[Heat Disinfect] program." Your dialysis schedule should not be interrupted. If interruption of yourdialysis schedule is expected, please discuss your options with your healthcare provider." Contact FMCNA Customer Service Team at 1-800-323-5188 forinstructions on how to return the recalled product." Promptly fill out and return the attached reply formAdditional medical concerns or questions, please contact Medical Informationand Communication: 855-616-2309 or Website: www.fresenius-medinfo.com.For questions regarding this recall call 800-662-1237. |  | 1856619 | 1,856,619 Bottles | Nationwide Distribution | Open, Classified |  |  | 1856619 | Not_Computer
Z-1826-2015 | 71519 | No | DRE | Vessel dilator for percutaneous catheterization. | Cardiovascular | Cardiovascular | 510(k) | Merit Medical Prelude Pro Sheath Introducter, 6F (2.0mm) 11cm, REF PRO-6F-11-038, Sterile EO |  |  | June 04, 2015 | 2015 | Merit Medical Systems, Inc. | Some of these introducer sheaths were packaged with a 0.035 inch dilator, rather than the labeled 0.038 inch. This affects one lot, H767887, of catalog number PR0-6F-11-038. | Packaging process control | Merit Medical sent an Urgent Product Recall letter on June 4, 2015, to all affected consignees. Consignees were instructed to immediately quarantine any affected devices and discontinue use, ensure all personnel to whom devices were distributed are made aware of this field action, and instructions to contact their Merit representative. An email notification was sent to Merit sales representatives with accounts that had received affected product. The affected sales representatives were provided a Customer Response Form applicable to their account and a representative copy of communications sent directly to their affected accounts. Merit sales representatives were instructed to visit the account and assist with the return of the affected units back to the Merit Field Assurance Department. Customers with questions were instructed to contact their Merit Sales Representative or call (801) 208-4365. |  | 1445 | 1,445 | Worldwide Distribution - US including TN, NE, MA, SC, WI, UT and Internationally to Belgium, Switzerland, United Kingdom, and Germany. | Terminated | October 26, 2015 | 144 | 1445 | Not_Computer
Z-1823-2015 | 71357 | No | MNH | Thoracolumbosacral pedicle screw system. | Orthopedic | Orthopedic | 510(k) | Fortex" Pedicle Screw System Continuous Radius Rod 5.5 - 50mm, X-Spine Systems, Inc., 452 Alexandersville Road Miamisburg, OH 45342 Phone: 937-847-8400 |  |  | April 07, 2015 | 2015 | X Spine Systems Inc | The Continuous Radius Rod should measure 50mm in length but actually measures 45mm in length. | Other | X-Spine sent an Urgent Medical Device Recall Letter dated April 13, 2015, to all affected customers. The letter identified the product the problem and the action needed to be taken by the customer. Actions to be taken:1. The affected hospital and surgeon should be notified of this recall immediately.2. Complete the attached Medical Device Recall Return Response Form and return it with any affectedproduct in inventory to X-Spine Systems immediately.a. Reference RMA # 5410b. Use enclosed shipping label3. No action should be taken with product that has been implanted.Please feel free to contact me at 800.903.0640 ext. 137. |  | 75 | 75 units | Worldwide Distribution - US Distribution to the states of : FL, ID, IL, KS, LA, MI, MO, NJ, OH, PA and TX., and to the countries of Austria, Germany and Italy. | Terminated | October 01, 2015 | 177 | 75 | Not_Computer
Z-1825-2015 | 71449 | No | IOL | Powered exercise equipment. | Physical Medicine | Physical Medicine | 510(K) Exempt | GE Healthcare, Treadmill T2100. |  |  | March 12, 2015 | 2015 | GE Healthcare | GE Healthcare has recently become aware of a potential safety issue due to uncontrolled motion of the walking belt on your T2100 Treadmill. | Device Design | General Healthcare sent an Urgent Medical Device Correction letter (GEHC Ref# 30074) dated March 12, 2015, to all affected consignees. The letter was addressed to Director of Biomedical Engineering, Chief of Nursing and Healthcare Administrator / Risk Manager. The letter described the Safety Issue, Safety Instructions, Affected Product Details, Product Correction, and Contact Information. For questions contact GE Healthcare Service at 1-800-437-1171 or your local Service Representative. |  | 11423 | 11,423 (5,540 units US, 5,883 units OUS | Nationwide Distribution including Guam, DC, and Puerto Rico.OUS: VIET NAM, VENEZUELA, UZBEKISTAN, URUGUAY, UNITED KINGDOM, UNITED ARAB EMIRATES, TURKEY, TRINIDAD &amp; TOBAGO, THAILAND, TAIWAN EL SALVADOR, SRILANKA, SPAIN, SOUTH AFRICA, SINGAPORE, SAUDI ARABIA, RUSSIA, PORTUGAL PHILIPPINES, PERU, PARAGUAY, PANAMA, PAKISTAN, NEW ZEALAND, NICARAGUA, NEPAL, MYANMAR, MOROCCO, MEXICO, MARTINIQUE, MALAYSIA, LEBANON, KUWAIT, REPUBLIC OF KOREA, KAZAKHSTAN, JAPAN, JAMAICA, ITALY, ISRAEL, IRELAND, IRAQ, INDONESIA, INDIA, HON GONG, GERMANY, FRANCE, FINLAND, ETHIOPIA, EGYPT, ECUADOR, DOMINICAN REPUBLIC, COSTA RICA, COLOMBIA, CHINA, CHILE , CAYMAN ISLAND , CANADA, CAMBODIA, BRUNEI DARUSSALAM, BRAZIL, BOLIVIA, BHUTAN, BANGLADESH, BAHRAIN, AUSTRIA, AUSTRALIA, ARGENTINA, ALGERIA. | Open, Classified |  |  | 11423 | Not_Computer
Z-1822-2015 | 71257 | No | GAT | Nonabsorbable poly(ethylene terephthalate) surgical suture. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | ULTRA-FAST FIX Knot Pusher Suture Cutters |  |  | April 30, 2015 | 2015 | Smith &amp; Nephew, Inc. | Sterility of device maybe compromised | Package design/selection | Smith &amp; Nephew sent an Urgent - Product Recall 1st Notification letter dated April 27, 2015, via Federal Express on April 30, 2015. Users are requested to inspect inventory and locate any unused devices from the above listed product and batch numbers, and quarantine them immediately. Complete the the Inventory Return Certification and obtain a RGA and obtain replacement product.Contact Smith &amp; Nephews Field Action Department via email at FieldActions@smith-nephew.com or fax +1-901-566-7975 to obtain a return authorization (RA) number |  | 202279 | 202,279 units | Worldwide Distribution - US (nationwide) Distribution including Puerto Rico and to the countries of : Argentina, Austria, Australia, Belgium, Brazil, Canada, Switzerland, Chile, China, Colombia, Costa Rica, Dubia UAE, Germany, Denmark, Spain, Finland, France, United Kingdom, Greece, Guam, Hong Kong, Israel, India, Italy, Japan, Korea, Mexico, Malaysia, Netherlands, Norway, Portugal, Qatar, Sweden, Singapore,Thailand, Turkey, Taiwan, Venezuela and South Africa. | Open, Classified |  |  | 202279 | Not_Computer
Z-1824-2015 | 71313 | No | HIB | Obstetric-gynecologic specialized manual instrument. | Obstetrics/Gynecology | Obstetrics/Gynecology | 510(K) Exempt | McKesson Medi-Pak Vaginal Speculum, 10/bag |  |  | April 13, 2015 | 2015 | CooperSurgical, Inc. | Small sized specula were incorrectly packaged in printed bags for the medium sized speculum | Packaging change control | CooperSurgical sent an Urget Medical Device Recall letter dated April 13, 2015, to all affected customers. The letter identified the product, the problem, and the actions to be taken by the customer. Customers were instructed to discontinue use, discard, or return the affected to product to Cooper Surgical. Customers were asked to complete the attached Acknowledgment and Receipt Form for replacement. Customers with questions were instructed to call 203-601-5200. |  | 9400 | 9400 units | Nationwide Distribution including IL, FL, GA, MA, MO, TX, and WA. | Terminated | February 05, 2016 | 298 | 9400 | Not_Computer
Z-1828-2015 | 71196 | No | DTL | Cardiopulmonary bypass adaptor, stopcock, manifold, or fitting. | Cardiovascular | Cardiovascular | 510(K) Exempt | Custom Cardiovascular Procedure Kits; Over-Pressure Safety Valves. |  |  | May 06, 2015 | 2015 | Terumo Cardiovascular Systems Corp | The Overpressure Safety Valve is not cleared with an indication for use on the arterial side of the extracoporeal bypass circuit. It is intended for use only on the venous side. There is a remote possibility of minimal blood loss, fluid weeping, or dripping on the user. | Component design/selection | Consignees were notified of the recall by phone beginning on 5/7/2015 and a Urgent Medical Device Recall letter, dated 5/7/2015. The letter identified the affected device, reason for recall, potential hazard, and correction. Customers were asked to review the notice and ensure that all users are aware of the issue. In addition, they are to confirm the quantity of procedure kits for return; and, complete and return the Customer Response Form. Customers can call Terumo CVS Customer Service at 1-800-521-2818 with questions. |  | 163 | 163 devices | Distributed to the states of MO and FL. | Terminated | July 27, 2015 | 82 | 163 | Not_Computer
Z-1829-2015 | 71450 | No | JAQ | Remote controlled radionuclide applicator system. | Radiology | Radiology | 510(k) | Surface Applicator Set with Leipzig-style Cone. |  |  | May 18, 2015 | 2015 | Varian Medical Systems Inc | Varian has discovered that there is a discrepancy in the absolute dose rate given with the Leipzig-style Surface Applicator's Instructions for Use (IFU): Dose CharacterizationGM11010080 2012-09-06. The actual dose rate of the applicator is approximately 14% higher than the rate published within the IFU. Use of the dose rate as provided by the IFU without modification or independent confirmat | Process design | All the consignees affected by this recall will be notified by:The Urgent Medical Device Correction Letter/ Field Safety Notification will bedistributed to all affected users, with a description of the problem.The Letter will also be distributed to the Varian Sales, Marketing, and Serviceorganizations, informing them of the issue.Additionally, the letter will be posted to the customer support site:http:/ /www.MyVarian.com. |  | 46 | 46 | Worldwide Distribution. | Terminated | January 26, 2016 | 253 | 46 | Not_Computer
Z-1830-2015 | 71287 | No | KGY | Organ bag. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(K) Exempt | LapSac Surgical Tissue Pouch. |  |  | May 14, 2015 | 2015 | Cook Inc. | Includes Information not stated in the intended use. | Labeling Change Control | Cook Medical sent an "Recall Communication Updated Instructions for Use" letter dated May 15, 2015, to all affected customers . The letter identified the product the problem and the action needed to be taken by the customer. Cook has requested that customers do the following: The intended use for the device is being modified from "Used to isolate tissue" during, or prior to, surgical removal and/or morcellation to "Used to isolate tissue" during, or prior to, surgical removal .A contraindication being added is as follows:" Not for use with power morcellators.The warnings are being modified to read as follows:" The LapSac Surgical Tissue Pouch is not intended for use with sharp surgical instruments." Care should be taken to avoid damaging, puncturing or tearing the LapSac Surgical Tissue Pouch." Use caution when removing potentially malignant or harmful tissues, to isolate contamination from surrounding tissue.The instructions for use are being modified to read as follows:1. Following aseptic technique, introduce the LapSac Surgical Tissue Pouch into the abdominal cavity.2. Place tissue to be removed from the body into the pouch. 3. Using the drawstring, carefully close the LapSac Surgical Tissue Pouch. Please note that if excess force is used, the drawstring may break or tear the LapSac.4. Remove the LapSac Surgical Tissue Pouch from the body, taking care not to spill contents of the bag.Please report any Adverse Event to Cook Medical Customer Relations (800) 457-4500 or 1-812-339-2235. Monday through Friday between 7:30 a.m. and 5:00 p.m. Eastern Daylight Time or email at CustomerrelationsNA@cookmedical.com; and to FDA. The customer notice must be passed on to appropriate personnel within each organization or to any organization where the potentially affected devices have been transferred.If customers have further questions regarding labelling changes for this device, please contact Cook Sales Representative or Cook M |  | 123184 | 123,184 devices were distributed in commerce | Worldwide Distribution - US (nationwide) including PR and to the states of: AZ, AK, AL, AR, CA, CO, CT, DC, DE, FL, GA, HI, IA, ID, IL, IN, KS, KY, LA, MA, MD, MI, MN, MS, MO, ME, MT, NC, ND, NE, NH, NJ, NY, NM, OH,OK,OR, PA, RI, SC, SD, TN, TX, UT, VA, WA, WI, WV and WY., and to the countries of : China, Korea, Denmark, Mexico, New Zealand, India and Colombia. | Open, Classified |  |  | 123184 | Not_Computer
Z-1831-2015 | 71385 | No | GEI | Electrosurgical cutting and coagulation device and accessories. | General &amp; Plastic Surgery | General &amp; Plastic Surgery | 510(k) | SERFAS 90 degree Energy Probe, Part Number 279-350-101; |  |  | June 03, 2015 | 2015 | Stryker Endoscopy | Stryker Endoscopy is recalling all non expired SERFAS 90 degree Energy Probes due to reports of fragments of the probe breaking off into the patient. | Device Design | Stryker sent an Urgent Medical Device Recall letter dated June 5, 2015, to all affected customers. The letter identified the product, the problem, and the action to be taken by the customer. Customers were instructed to do the following:1. Inform individuals within their organization who need to be aware of this device removal.2. Review inventory of lots of part number 279-350-101 and determine if you have the affected product (all non-expired devices) in stock. Response is required.3. If no product is found, complete acknowledgement form located on the Stryker Endoscopy recall website endorecall.stryker.com by logging in using the account number and zip code on this letter.4. If you do have product, segregate the product and call Stryker customer service at 1-800-624-4422 Option 3 to arrange for product return and issuance of credit.For questions regarding this recall call 408-754-2664. |  | 22063 | 22,063 devices | Worldwide Distribution - US (nationwide) and Internationally to US Argentina, Australia, Bolivia, Brazil, Chile, China, Colombia, Hong Kong, India, Italy, Japan, Republic of Korea, Malaysia, Netherlands, Poland, Romania, Singapore, South Africa, Spain, Sweden, Switzerland, United Arab Emirates, and United Kingdom. | Terminated | February 10, 2016 | 252 | 22063 | Not_Computer
Z-1809-2015 | 71254 | Yes | BTM | Manual emergency ventilator. | Anesthesiology | Anesthesiology | 510(k) | Lifesaver Single Patient Use Manual Resuscitator |  |  | May 14, 2015 | 2015 | Teleflex Medical | The intake port may be blocked which can cause the bag to fail to fill. | Nonconforming Material/Component | Teleflex Medical sent an Urgent Medical Device Recall Notification letter dated May 14, 2015, to all consignees asking them to immediately discontinue distribution and quarantine the recalled product. The letter also requested a sub-recall of the product. The recall letter also included a response form which is to be returned to Teleflex by emailing it to recall@teleflex.com or fax it to 1-855-419-8507, Attn: Customer Service. Customers with questions were instructed to contact their local sales representative or Customer Service at 1-866-246-6990.For questions regarding this recall call 610-378-0131. |  | 9333 | 9,333 units | Worldwide Distribution - US Nationwide in the states of AL, AK, AZ, AR, CA, CO, CT, FL, GA, IL, IA, KS, LA, MA, MI, MN, MS, MO, NE, NH, NY, NC, OH, OK, PA, RI, SD, TN, TX, VA, WA, including Puerto Rico and the countries of Australia, Bahamas, Canada, Guatemala, and Mexico. | Open, Classified |  |  | 9333 | Not_Computer
Z-0351-03 | 25150 | [6 empty] | Medtronic Cardioblate Radio Frequency Ablation Pen, model number 60811C47 |  |  | November 11, 2002 | 2002 | Medtronic Cardiac Surgery Technologies | The radio frequency ablation pens contained unreacted epoxy resin which was observed to leak out of the back-end of the device. | Other | A letter dated November 11, 2002 was sent to the affected consignees who were requested to quarantine any unused pens for retrieval by Medtronic representatives. |  | 148 | 148 pens | All of the consignees were in foreign countries in Europe, Asia and Latin America. | Terminated | July 02, 2004 | 599 | 148 | Not_Computer
Z-0494-05 | 31217 | No | RAD |  |  |  |  | X-Ray High Voltage Generator, Model KXO-80G used on Image Intensified Fluoroscopoc X-Ray System DUA-450D |  |  | September 07, 2004 | 2004 | Toshiba American Med Sys Inc | Non-compliance with Radiation Standards. Incorrect display of the radiographic time, due to software anomaly. | Other | Firm sent recall letters to customers and will be following up with field corrections which will be made on-site by the firm's own employees. |  | 494 | 494 | Nationwide. | Terminated | February 10, 2012 | 2712 | 494 | Computer
Z-0352-03 | 25226 | [6 empty] | Surgical Mesh |  |  | December 11, 2002 | 2002 | Synovis Surgical Inovation Div. of | A gap in the exterior seal may lead to sterility concerns with the interior pouch. | Other | The sites and/or distributors that received suspect devices were contacted by telephone on 12/11/02 and given instructions on how to return/replace the devices. In addition, a letter dated 12/11/02 was sent to each consignee. |  | 28 | 28 devices | Product shipped nationwide, no foreign consignees | Terminated | May 28, 2003 | 168 | 28 | Not_Computer
Z-0792-03 | 26046 | No | FRN | Infusion pump. | General Hospital | General Hospital | 510(k) | GemstarÂż 7 Therapy I.V. Infusion Pump, List#: 13000-04; 2.9 Software Version |  |  | April 01, 2003 | 2003 | Abbott Laboratories | Products contain software deficiencies which could allow over infusion thereby administering excess medication. | Other | On 4/1/03, the firm issued letters via Federal Express Overnight Delivery to all its direct consignees, informing them of the affected devices and providing instructions on the recall. |  | 1310 | 1310 units | Product received nationwide distribution to approx. 120 direct consignees. The firm identified 3 Govt. accounts: (1) Nthrn VA Comm. Hospital, 601 S. Carlin Spg. Rd., Arlington VA 22204, (2) Veterans Admin Med Ctr., 4100 W. 3rd St., Dayton OH 45428, (3) GW University Hospital, 901 23rd St. NW # 5306N, Washington DC 20037. The recall was appropriately extended to the user level; i.e., the hospital/medical centers, physicians and nurses who received the recalled product.There is no known Canadian distribution. | Terminated | October 09, 2003 | 191 | 1310 | Computer
Z-0412-03 | 4397 | [6 empty] | Cryo-Tip T-1910 Exo-Endocervical, Part Number 90501 (as labeled) |  |  | June 19, 2002 | 2002 | ACMI Corporation | Derm Bevel Cryo-Tips, Part Number 90512, were mislabeled as Exo- Endocervical Cryo-Tips, Part Number 90501. | Other | The recalling firm telephoned the 2 consignees on June 19, 2002, informed them of the recall and asked them to return the product they had. |  | 7 | 7 Cryo-tips | Seven units had been distributed to five consignees, but three consignees had returned four units on 5/22, 5/31 and 6/13/02 prior to the recall because they were mislabeled and not the devices needed by the consignees. This left 3 units at 2 consignees in the States of California and Texas. | Terminated | January 01, 1990 | -4552 | 7 | Not_Computer
Z-0382-03 | 3265 | [6 empty] | Medtronic Model 550 Bio-Console Extracorporeal Blood Pumping |  |  | September 25, 2001 | 2001 | MEDTRONIC PERFUSION SYSTEMS | Motor controller printed circuit boards with a capacitor installed in reverse can result in a shut down of the product pump. | Other | September 25th, 2001 letter about the problem followed by field correction of the units. |  | 126 | 126 units | The devices were distributed to 30 domestic consignees (mostly hospitals) located nationwide in the United States. The product was also distributed to 4 Medtronic foreign subsidiaries and from there to foreign consignees located worldwide. | Terminated | January 01, 1990 | -4285 | 126 | Not_Computer
Z-0353-03 | 4411 | [6 empty] | AXIS/IRIX Gamma Camera System, controlled with Odyssey computers, |  |  | October 08, 2001 | 2001 | PHILIPS MEDICAL SYSTEMS | A broken gantry wire may prevent the contact sensors from stopping motion of the unit. | Other | The firm sent a "Manditory Service Letter" (ML#437), dated 10/8/2001, to their Field Service Engineers, instructing them to immediately contact their accounts and schedule the units for repair. A software upgrade, version 8.9B, will be installed on all affected units in the field. The 8.9B retrofit kit (part #53000011) will be sent out after a service order has been generated. |  | 653 | 653 devices installed | The units were shipped to consignees located nationwide and worldwide. The following U.S. VA Medical Centers are involved: BUFFALO, NY; HINES, IL; CHICAGO, IL; MONTGOMERY, AL; TUCSON, AZ; SAN DIEGO, CA; SAN FRANCISCO, CA; LONG BEACH, CA; WILMINGTON, DE; DECATUR, GA; IOWA CITY, IA; N. CHICAGO, IL; BILOXI, MS; CLEVELAND, OH; CINCINNAYI, OH; DAYTON, OH; BIG SPRING, TX; HUNTINGTON, WV; ALBUQUERQUE, NM; COLUMBIA, SC; DENVER, CO; DES MOINES, IA; PHOENIX, AZ; PRESCOTT, AZ; FRESNO, CA; LEXINGTON, KY; SAN JUAN, PR; OKLAHOMA CITY, OK; and ENRICO, CA. | Terminated | April 15, 2004 | 920 | 653 | Computer
Z-0407-03 | 4607 | [6 empty] | EXPO 32 Software versions 1.0, 1.1B, and 1.2 |  |  | May 20, 2002 | 2002 | BECKMAN COULTER, INC. | Software contains errors that might affect Cytometry instruments | Other | Recall letters were sent May 20, 2002 which indicate potential problems that would be noted depending on the equipment being used and communicates instructions to prevent the problems until the next version of the software is issued. |  | 33 | 33 | Only domestically to the states of CA, FL, HI, KS, KY, MA, MD, MI, NC, NJ, OH, OR, SC, TX, UT, WA. | Terminated | November 22, 2004 | 917 | 33 | Computer
Z-0471-03 | 4795 | [6 empty] | Kendall Argyle 3.7 Fr Polyurethane Umbilical Vessel Catheter Kendall Argyle 3.7 Fr Polyurethane Umbilical Vessel Catheter |  |  | October 02, 2002 | 2002 | TYCO HEALTHCARE GROUP LP/KENDALL | Lumen of Catheter may not accomodate the designated blood gas sensor | Other | Tyco/Kendall notified Direct Accounts (Hospitals and Distributors) on 10/2/02 by letter advising accounts to examine inventory, quarantine recalled and return. Distributors are requested to subrecall.On 1/22/03, the firm extended the recall to ALL Lots which included the addition of one lot number: 426038. |  | 734 | 734 cases (10/case) | Nationwide. AR, GA, IL, IN, MO, MN, NC, NJ, SC, TX, WI, VA.International: UK | Terminated | May 02, 2012 | 3500 | 734 | Not_Computer
Z-0445-03 | 4403 | [6 empty] | IMx Glycated Hemoglobin Reagent Sets, list 1A86-88; |  |  | July 05, 2002 | 2002 | ABBOTT LABORATORIES, INC. | Low Control reads high, out of insert range | Other | Recall letters dated 7/2/02 were sent to all customers receiving IMx Glycated Hemoglobin Reagent Sets on 7/5/02. Customers were requested to cease use of the affected lots and destroy the remaining product. |  | 38877 | 38,877 kits | The sets were distributed nationwide to 1101 hospital/laboratory accounts, including 22 direct sale government accounts. The sets were also shipped through Abbott subsidiaries internationally to Canada, Bermuda, Korea, India, Hong Kong, Honduras, Taiwan, Bahamas, Japan, Argentina, Costa Rica, Nicaragua, Mexico, Panama, Brazil, Jamaica, Guatemala, Colombia, Chile, Uruguay, Dominican Republic, El Salvador, Great Britain, Pakistan, Guyana, Arabia, Haiti, Trinidad, Grenada, St. Lucia, Antigua, Curacao, Aruba, Tortola, Anguilla, Cayman, Virgin Islands and Australia. | Terminated | January 01, 1990 | -4568 | 38877 | Not_Computer
Z-0395-03 | 4719 | [6 empty] | Digene's Rapid Capture System (RCS) robotic microplate processor, |  |  | September 11, 2002 | 2002 | DIGENE CORP. | IVD instrument may measure incorrect sample volume for clinical testing | Other | Recalling firm notified consignees by letter dated 9/11/2002. Firm advises in the letter that users should not use RCS to process cervical brush specimens until firm has adjusted the instrument. |  | 5 | 5 units | Product distributed to end users in MD, IN and OR and to foreign consignee in Japan.US Gov't consignee:Oregon Public Health Lab1717 Southwest 10th St.Portland, OR 97201Foreign consignee:Mitsubishi8-5-1 Chuo, Ami-machiInashiki-gun, Ibaraki300-0332 Japan | Terminated | January 01, 1990 | -4636 | 5 | Not_Computer
Z-0334-03 | 4713 | [6 empty] | HomeChoice and HomeChoice PRO Automated Peritoneal Dialysis Systems; Manufactured for Baxter Healthcare Corporation, Deerfield, IL 60015 U.S.A., Made in U.S.A.; the following models are affected: |  |  | September 19, 2002 | 2002 | BAXTER HEALTHCARE CORP. | possible overfill | Other | Safety Alert letters dated 9/19/02 were mailed to all Home Care Coordinators on 9/19/02 and to all Home Dialysis Patients on 9/20/02, informing them of the possible overfill situation and advising them that if a manual drain is required during fill, to make sure that the manual drain is completed without interruption in order to prevent an overfill situation. Addendum pages were provided for the patient Help Cards, Patient At-Home Guide and Trainer's Guide to include the warning not to stop or bypass a manual drain during fill. |  | 12667 | 12,667 units | Nationwide to 892 hospitals, 88 VA/government accounts, 1 dealer, 9 physicians, 5 home care suppliers, 2,041 dialysis centers and 14,707 home patients. The device was also distributed internationally through Baxter subsidiaries in Canada, Japan, Argentina, Australia, Belgium, Brazil, Chile, China, Colombia, Czech Republic, Guatemala, Honduras, Hong Kong, India, Israel, Italy, Jordan, Lebanon, Malaysia, Malta, Mexico, New Zealand, Oman, Panama, Peru, Philippines, Qatar, Saudi Arabia, Singapore, South Africa, South Korea, Spain, Sweden, Taiwan, Thailand, Turkey, United Arab Emirates and the United Kingdom. | Terminated | July 16, 2015 | 4683 | 12667 | Not_Computer
Z-0406-03 | 4712 | [6 empty] | Myocardial Temperature Probe, YSI 400 Series, 15mm, packaged in |  |  | September 17, 2002 | 2002 | DEROYAL SURGICAL | Surgical device is mislabeled with incorrect size. | Other | Firm notified consignees 9/17/2002 by certified letter. The letter instructed to return product to firm's distribution center for replacement. Firm included a response form for consignees to fax numbers of product to be replaced. The letter additionally instructed distributors to further notify their customers or provide customer contact information. |  | 3000 | 3000 units | Product was distributed to health care facilities and distributors nationwide | Terminated | January 01, 1990 | -4642 | 3000 | Not_Computer
Z-0383-03 | 4419 | [6 empty] | BD DTX Plus Pressure Monitoring Kits and Extension Sets. Packaged |  |  | July 08, 2002 | 2002 | BECTON DICKINSON INFUSION THERAPY SYSTEMS, INC. | Bond between the tubing and female or male luer fitting may fail. | Other | On 7/8/2002, all consignees were notified by telephone and FAX. |  | 8572 | 7122 pressure monitoring kits, 1450 extension sets | Nationwide.Domestic distribution to AR,CA,CO,LA,MO,NJ,NY,OH,OR,PA, RI,TN,VA, WA,WV. Includes VA facility in OH. | Terminated | January 01, 1990 | -4571 | 8572 | Not_Computer
Z-0329-03 | 4792 | [6 empty] | Trade Name: ABC Fixation Pin. System Name: ABC Cervical Plating |  |  | September 06, 2002 | 2002 | AESCULAP, INC. | pin manufactured with sharp tip and not hardened, pin may be susceptible to bending | Other | The recalling firm issued recall leters dated 9/11/02 to the direct accounts via certified mail. Recall notifications dated 9/11/02 were also sent the their sales representatives. |  | 216 | 216 kits | The kits were shipped to hospitals in CA, GA, IN, MN, MS, NJ, PA, SC, TN, TX, UT, and WA. The kits were also shipped to 16 sales representatives. There were no foreign and/or government/military accounts. | Terminated | January 01, 1990 | -4631 | 216 | Not_Computer
Z-0331-03 | 4853 | [6 empty] | Stryker Wedge turning frame stretcher, model 965. 965 |  |  | October 23, 2002 | 2002 | STRYKER CORP., | Weld failure-anterior and posterior frame may dislodge from stretcher litter | Other | The firm sent each customer an undated urgent device correction letter on October 23, 2002 asking them to remove the stretchers from service until corrected. Stryker service technicians will visit each account with upgrade kits around November 2002. |  | 29 | 29 | Arkansas, District of Columbia, Georgia, Illinois, Louisiana, Maine, Missouri, Ohio, Utah, West Virginia, Hong Kong, Japan, Latin America, Sweden and Switzerland.U.S. GOVERNMENT ACCOUNTSWalter Reed Army Medical Center, Bldg 2 Room 2E07A, Washington, DCFort Gordon, Fort Gordon, GAHill AFB, Hill AFB, UT | Terminated | January 01, 1990 | -4678 | 29 | Not_Computer
Z-0333-03 | 4839 | [6 empty] | Cryolife Valve Allograft. |  |  | October 04, 2002 | 2002 | CRYOLIFE, INC. | Donor tissue had possible sepsis secondary to pyogenic brain abscess. |  | The implanting physician was notified of the circumstances via letter on/about October 4, 2002 |  | 1 | One individual heart valve allograft. | One consignee located in INJohn Fehrenbacher, M.D., Methodist Professional Center, 1801 North Senate Blvd., #755, Indianapois, IN 46202-5300. | Terminated | January 01, 1990 | -4659 | 1 | Not_Computer
Z-0328-03 | 4828 | [6 empty] | Epidural Catheterization Kit, model number AK-05503L. The kit is |  |  | October 15, 2002 | 2002 | ARROW INTL., INC. | mislabeled - package may contain incorrect gauge catheter | Other | The recalling firm issued recall letters dated 10/14/02 informing their customers of the problem and the need to return it. The product label states that a 19 gauge needle is part of the kit when a 20 gauge needle is actually in the kits for this recalled lot. |  | 720 | 720 kits | The kits were sold to two distributors in IL, eight hospitals in CA, IN, KY, and OH; and one government account in NE. There are no foreign accounts.The government/military account is VAMC, Omaha, NE. | Terminated | January 01, 1990 | -4670 | 720 | Not_Computer
Z-0365-03 | 4825 | [6 empty] | LC+, LCV+, and LCA Angiographic X-ray Systems, Models 2168678, 2168678-2, 2219151, 2219151-2, 2219151-3, 2124824, 2168678-3, 2219150, 2219150-2, 2219150-3, 2107015, 2115906, 2168677, 2219152, |  |  | October 28, 2002 | 2002 | GENERAL ELECTRIC MEDICAL SYSTEMS | The motor gearshaft for the image intensifier may break allowing the image intensifier to slide down to its end stop. | Other | Field modification instructions were issued on 10/28/02 to GE field engineers who then corrected the devices at hospitals. |  | 1361 | 1361 systems | The affected products had been shipped to 610 hospitals located nationwide in the U.S. and to 675 foreign hospitals located worldwide. | Terminated | July 02, 2004 | 613 | 1361 | Not_Computer
Z-0397-03 | 4820 | [6 empty] | Medex Vented Flow Clip Nitroglycerin Set, Pv22ml, sterile, packaged |  |  | October 16, 2002 | 2002 | MEDEX, INC. | Product contains latex, however the label incorrectly states 'Latex Free'. | Other | On 10/17/2002, the firm notified their direct accounts and sub-accounts by telephone and facsimile. The customers were instructed to remove the product from service and return to Medex. |  | 304 | 304 cases (6,080 units) | The product was shipped to distributors and hospitals located in IL, AL, FL, UT, NY, and CA. | Terminated | December 30, 2003 | 440 | 304 | Not_Computer
Z-0411-03 | 4819 | [6 empty] | Side Fire Laser Fibers, Models DBLF-60SF and HBLF-60SF DBLF-60SF and HBLF-60SF |  |  | October 15, 2002 | 2002 | Laser Peripherals, LLC | The laser fibers were manufactured with an expired epoxy sealant which could result in degradation and misdirection of the laser beam. | Other | The recalling firm telephoned all of the consignees on October 15, 2002 and then sent recall letters, dated October 15, 2002, to the consignees requesting that the consignees return unused product. |  | 159 | 159 laser fibers | The products were shipped to 17 consignees (mostly hospitals and some distributors) located nationwide in the United States. There are no U.S. Government accounts. The products were also shipped to 2 distributors in Canada and Japan. The foreign consignees are:InstruMed Surgical2255 Dunwin Drive, Unit # 1Mississauga, Ontario L5L 1A3Canada905-820-0902M &amp; M Co. LtdMutoh Asakusabashi Bldg.1-16-6 Yanagigashi, Taito-kuTokyo, 111-0052 Japan81-3-3865-6575 | Terminated | January 01, 1990 | -4670 | 159 | Not_Computer
Z-0326-03 | 4814 | [6 empty] | Qwikwash Tubing (Air, Waste, Water), list 6208-23; an accessory for |  |  | October 10, 2002 | 2002 | ABBOTT LABORATORIES, INC. | visual discoloration of water tubing assembly | Other | Recalled by telephone calls on 10/10/02 and follow-up letters sent 10/11/02. Accounts were requested to examine their Abbott Commander PPC and Qwikwash systems, as well as their stocks of the tubing, for tubing that does not have the TYGON logo imprinted on the tubing. They were instructed to destroy any found and report the number destroyed back to Abbott. |  | 1343 | 1343 | The tubing was shipped intermingled with tubing from another supplier nationwide to 248 laboratory accounts, including 20 government direct sale accounts. The tubing was also shipped to Abbott subsidiaries in Canada, Germany and Singapore. | Terminated | January 01, 1990 | -4665 | 1343 | Not_Computer
Z-0454-03 | 4812 | [6 empty] | A) Roche/Hitachi MODULAR ANALYTICS E 170 Module |  |  | October 18, 2002 | 2002 | ROCHE DIAGNOSTICS CORP., | Fluid leak may cause smoke/damage analyzer | Other | The firm sent an urgent product correction notice dated 10/18/02 to all customers alerting them of the problem and advising them to be alert for the smell or sight of smoke. The firm is replacing the fittings as part of the correction. |  | 116 | 116 | United States and CanadaU.S. Government accountsBrooke Army Medical Center, 3851 Roger Brooke Dr., Rm 430-1, Ft. Sam Houston, TXPortland WV Med Ctr, 3710 SW US Veterans Hosp. Rd., Portland, ORVA Med Ctr, 700 S 19th St., Birmingham, ALVA Med Ctr, 1601 SW Archer Rd, Gainesville, FLVA Med Ctr, 10 Calle Casia, San Juan, PR Canadian consignee: Roche Diagnostics Canada, Laval, QC | Terminated | January 01, 1990 | -4673 | 116 | Not_Computer
Z-0396-03 | 4803 | [6 empty] | Holdex tube holder Item#450261 |  |  | September 20, 2002 | 2002 | Greiner Vacuette North America, Inc. | Holdex tube holder malfunction | Other | Letters were sent to consignees on/about 09/20/2002. |  | 103921 | Lot# A020208-9800 pieces, Lot#A050113-23800 pieces | Nationwide.AR, CA, CT, FL, GA, KS, LA, MA, MD, ,MI, MO, NH, NJ, NM, NY, OK, OH,OR, PA, SC, TX, VA.Veterans Administration, Medical Center, Calm VA-1 689, West Haven, CT. State of Georgia, Central State Hosptial, Lawrence Building, Milledgeville, GA. Sate of California, California State University, Dr. Larry Mai/Anthropology, Long Beach, CA.It appears no foreign accounts are involved. | Terminated | January 01, 1990 | -4645 | 103921 | Not_Computer
Z-0327-03 | 4680 | [6 empty] | Alumina V40 Head |  |  | June 26, 2002 | 2002 | HOWMEDICA OSTEONICS CORP. | V40 Alumina Heads may not assemble properly onto V40 hip stems | Device Design | On June 26, 2002, Howmedica sent out recall letters to their customers instructing them to remove and quarantine affected product from their inventory. |  | 190 | 190 (+676 28mm heads) | Howmedica distributed this product throughout the nation to various hospitals. There were 4 international custmers in United Kingdom, Australia, France, and Italy. | Terminated | June 29, 2010 | 2925 | 190 | Not_Computer
Z-0332-3 | 4498 | [6 empty] | Little Soothers Cold pack. Each boxed set (unit) contains 3 patches. |  |  | July 25, 2002 | 2002 | Avon Products, Inc. | Microbial contamination. | Other | Voice message on 7/25/02, letter on 7/29/02 to Avon Sales Representatives advising of the problem and requesting return of product. |  | 1609 | 1609 units, each containing three Little Soothers. | Sales representatives ("Avon Ladies") nationwide, for further distribution/sales to their customers (consumers). | Terminated | August 24, 2012 | 3683 | 1609 | Not_Computer
Z-0811-03 | 4552 | [6 empty] | Phillips ADAC brand C-PET and Allegro Imaging Systems DICOM 3.0, C-PET rev. 7.3.5. |  |  | July 24, 2002 | 2002 | PHILLIPS ADAC | Resulting images may be labeled incorrectly on PETView and receiving system. | Other | On 7/24/02, the firm issued letters via Federal Express to all its direct consignees, informing them of the affected devices and providing instructions on the recall. |  | 104 | 104 | Distribution of the product is worldwide and nationwide. Distribution of CPET was made between 6/7/02 to present. Distribution of Allegro was made between 4/15/02 to present.Product was distributed to the following foreign countries: China, Canada, Australia, Egypt, Spain, Belgium, Italy, Germany, Austria, France, Japan, and Korea. Product was also distributed to: National Naval Medical Center1602 Skipwith Road Richmond, VA301-295-2357and Wilford Hall Air Force2200 Bergguist DriveLackLand AFBSan Antonio, TX210-292-2165 | Terminated | January 01, 1990 | -4587 | 104 | Not_Computer
Z-0011-05 | 30059 | No | LZG | Infusion pump. | General Hospital | General Hospital | 510(k) | Medtronic MiniMed Paradigm Model 512 |  |  | September 14, 2004 | 2004 | Medtronic MiniMed | Several complaints of pumps exhibiting A47 alarms were received. Investigation revealed the alarms occurred when the pump attempted to display text that exceeded the maximum display length of the pump screen. The investigation revealed this only occurs when the pump is programmed to display information in Spanish. | Other |  |  |  |  | Nationwide &amp; Spain | Terminated | March 03, 2005 | 170 |  | Not_Computer